## Abbreviations
Table 1. Abbreviations used for OREAS 134a
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*AAS | 3-acid (HNO3-HCIO4-HCI) digest with AAS finish
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with AAS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with either AAS or ICP-OES finish as appropriate
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
ABL*MS | Laser Ablation with Mass Spectrometry: ICP-MS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/AAS | ore grade aqua regia digest with either AAS or ICP-OES finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
IRC | infrared combustion furnace
LOI*TGA | loss on ignition with Thermal Gravimetric Analyser finish
MAR*OES | modified aqua regia digest with ICP-OES finish
OxBF*XRF | oxidising lithium borate fusion with X-Ray Fluorescence finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PPP*XRF | pressed powder pellet with XRF finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 134a
Lab Slot | Laboratory Name and Location
Acme (BV), Vancouver, BC, Canada
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Johannesburg, South Africa
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Amdel (BV), Adelaide, SA, Australia
Bureau Veritas Amdel Laboratories, Perth (Wangara), WA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Jakarta, Indonesia
McPhar Geoservices Inc., Makati City, Manila, Philippines
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
Ultra Trace Pty Ltd (BV), Perth, WA, Australia

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 134a
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate / Peroxide Fusion ICP
Ag, Silver (ppm) | 197.42 | 9.9841 | 186.2456 | 208.5944 | 190.1211 | 204.7189
Al2O3, Aluminium(III) oxide (wt.%) | 2.6025 | 0.13113 | 2.4942 | 2.7109 | 2.5397 | 2.6653
As, Arsenic (ppm) | 217.7 | 21.9748 | 196.3606 | 239.0394 | 208.1937 | 227.2063
Ba, Barium (ppm) | 1500.2639 | 155.6018 | 1333.5014 | 1667.0264 | 1479.9447 | 1520.5831
CaO, Calcium oxide (wt.%) | 6.3848 | 0.19651 | 6.2066 | 6.5631 | 6.3001 | 6.4696
Cd, Cadmium (ppm) | 545.8 | 35.8062 | 510.4076 | 581.1924 | 528.6426 | 562.9574
Co, Cobalt (ppm) | 100.1622 | 13.1722 | 89.9109 | 110.4136 | 97.6388 | 102.6856
Fe, Iron (wt.%) | 12.3213 | 0.79513 | 11.7061 | 12.9366 | 12.0125 | 12.6302
MgO, Magnesium oxide (wt.%) | 3.5548 | 0.11891 | 3.4534 | 3.6562 | 3.4982 | 3.6114
Pb, Lead (wt.%) | 12.9451 | 0.62283 | 12.4877 | 13.4024 | 12.6693 | 13.2208
SiO2, Silicon dioxide (wt.%) | 14.8131 | 0.8451 | 14.0971 | 15.529 | 14.5137 | 15.1124
Zn, Zinc (wt.%) | 17.4916 | 0.50511 | 17.1018 | 17.8815 | 17.1995 | 17.7838
4-Acid Digestion
Ag, Silver (ppm) | 200.8218 | 6.7526 | 196.2408 | 205.4028 | 196.7479 | 204.8957
Al2O3, Aluminium(III) oxide (wt.%) | 2.5326 | 0.11686 | 2.4573 | 2.6079 | 2.4866 | 2.5785
As, Arsenic (ppm) | 227.7698 | 22.911 | 216.506 | 239.0335 | 220.119 | 235.4205
CaO, Calcium oxide (wt.%) | 6.0361 | 0.3847 | 5.786 | 6.2862 | 5.9425 | 6.1297
Cd, Cadmium (ppm) | 535.6146 | 36.0705 | 514.1836 | 557.0457 | 523.0868 | 548.1424
Co, Cobalt (ppm) | 104.9669 | 8.2381 | 100.2586 | 109.6753 | 101.1469 | 108.787
Cu, Copper (ppm) | 1290.6657 | 45.0472 | 1269.9223 | 1311.4091 | 1266.8329 | 1314.4985
Fe, Iron (wt.%) | 12.0394 | 0.41049 | 11.8297 | 12.2492 | 11.8662 | 12.2126
MgO, Magnesium oxide (wt.%) | 3.4938 | 0.18412 | 3.3906 | 3.597 | 3.4256 | 3.5621
Pb, Lead (wt.%) | 12.7865 | 0.76631 | 12.3405 | 13.2324 | 12.4908 | 13.0821
S, Sulphur (wt.%) | 19.2749 | 0.50272 | 18.8476 | 19.7021 | 18.7185 | 19.8312
Sb, Antimony (ppm) | 115.4795 | 14.8451 | 106.9077 | 124.0514 | 111.607 | 119.3521
Zn, Zinc (wt.%) | 17.2671 | 0.5526 | 16.9491 | 17.5852 | 16.9654 | 17.5688
Aqua Regia Digestion (sample weights 0.15-1.0g)
Ag, Silver (ppm) | 193.5817 | 13.1673 | 185.5139 | 201.6494 | 188.4216 | 198.7417
Al2O3, Aluminium(III) oxide (wt.%) | 0.77352 | 0.085461 | 0.71446 | 0.83257 | 0.75478 | 0.79225
As, Arsenic (ppm) | 225.505 | 24.0738 | 212.394 | 238.616 | 219.2755 | 231.7345
Cd, Cadmium (ppm) | 546.7723 | 44.4356 | 520.4541 | 573.0905 | 537.1548 | 556.3898
Co, Cobalt (ppm) | 101.1341 | 8.1863 | 96.4734 | 105.7949 | 98.6548 | 103.6135
Cu, Copper (ppm) | 1279.4917 | 78.3751 | 1230.973 | 1328.0103 | 1250.3372 | 1308.6461
Fe, Iron (wt.%) | 11.8925 | 0.97309 | 11.3036 | 12.4814 | 11.6482 | 12.1367
MgO, Magnesium oxide (wt.%) | 3.425 | 0.13741 | 3.3322 | 3.5179 | 3.3698 | 3.4803
Pb, Lead (wt.%) | 12.6375 | 0.759 | 12.1095 | 13.1655 | 12.3449 | 12.9301
S, Sulphur (wt.%) | 18.4925 | 1.3285 | 17.3501 | 19.6349 | 18.0834 | 18.9016
Zn, Zinc (wt.%) | 16.9999 | 0.86337 | 16.4362 | 17.5636 | 16.5681 | 17.4317
Infrared Combustion
S, Sulphur (wt.%) | 19.1121 | 0.41288 | 18.8063 | 19.4179 | 18.8117 | 19.4125

## Indicative Values
Table 3. Indicative Values for OREAS 134a
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Oxidising Fusion XRF
Al2O3 | wt.% | 2.57 | Fe2O3 | wt.% | 17.6854 | SnO2 | ppm | < 13
As | ppm | 240 | K2O | wt.% | 0.9615 | SO3 | wt.% | 49.2947
BaO | ppm | 1501.6925 | MgO | wt.% | 3.5515 | SrO | ppm | 35.478
CaO | wt.% | 5.941 | MnO | wt.% | 0.4585 | TiO2 | wt.% | 0.093
Cl | ppm | 190 | NiO | ppm | 31.8125 | V2O5 | ppm | 71.408
CoO | ppm | 139.865 | P2O5 | wt.% | 0.07906 | ZnO | ppm | 215350.4
Cr2O3 | ppm | 51.1525 | PbO | ppm | 137773.88 | ZrO2 | ppm | 40.524
CuO | ppm | 1646.117 | SiO2 | wt.% | 14.82
Thermogravimetry
LOI1000 | wt.% | 14.59
Laser Ablation ICP-MS
Ag | ppm | 193 | Ge | ppb | 18150 | Se | ppm | < 5
As | ppm | 235.5 | Hf | ppb | 865 | Sm | ppm | 1.34
Ba | ppm | 1345 | Ho | ppb | 180 | Sn | ppm | 1
Be | ppm | 1.8 | In | ppm | 0.975 | Sr | ppm | 25.45
Bi | ppm | 0.19 | La | ppm | 8.64 | Ta | ppb | 205
Cd | ppm | 531.5 | Lu | ppb | 75 | Tb | ppb | 180
Ce | ppm | 16.25 | Mo | ppm | 6.6 | Te | ppb | < 200
Co | ppm | 103.35 | Nb | ppm | 2.05 | Th | ppm | 2.555
Cr | ppm | 10 | Nd | ppm | 6.875 | Tl | ppm | 80.5
Cs | ppm | 1.08 | Ni | ppm | 22 | Tm | ppb | 95
Cu | ppm | 1340 | Pb | wt.% | 12.6 | U | ppm | 1.365
Dy | ppm | 0.81 | Pr | ppm | 1.86 | V | ppm | 13.05
Er | ppm | 0.48 | Rb | ppm | 24.85 | W | ppm | 0.5
Eu | ppb | 320 | Re | ppb | 20 | Y | ppm | 5.73
Ga | ppm | 4.55 | Sb | ppm | 118 | Yb | ppb | 515
Gd | ppm | 1.2 | Sc | ppm | 1.8 | Zr | ppm | 27.75
Borate / Peroxide Fusion ICP
Be | ppm | 1 | Na | ppm | 652.819 | Ti | ppm | 612.6731
Cr | ppm | 80 | P | ppm | 296.5491 | V | ppm | 16.4
Cu | ppm | 1277.7 | S | wt.% | 19.3944 | Y | ppm | 5.4
K | wt.% | 0.92938 | Sb | ppm | 118.1667 | Zr | ppm | 30.2
LOI1000 | wt.% | 15.12 | Sc | ppm | 2
Mn | ppm | 3371.0409 | Sr | ppm | 28.8
4-Acid Digestion
B | ppm | 21 | K | wt.% | 0.795 | Sm | ppm | 1.38
Ba | ppm | < 1400 | La | ppm | 6.1 | Sn | ppm | 1.13
Be | ppm | 2 | Li | ppm | 17.84 | Sr | ppm | 18.3
Ce | ppm | 13.15 | Lu | ppb | 100 | Ta | ppb | < 500
Cr | ppm | 16.57 | Mn | ppm | 3424.1 | Tb | ppb | 180
Cs | ppm | 1.14 | Mo | ppm | 6.25 | Te | ppb | 90
Dy | ppm | 0.8 | Na | ppm | 303 | Th | ppm | 1.4425
Er | ppm | 0.56 | Nb | ppm | 1.86 | Ti | ppm | 479
Eu | ppb | 300 | Nd | ppm | 7.42 | Tl | ppm | 93
Ga | ppm | 4.88 | Ni | ppm | 21.29 | Tm | ppb | 100
Gd | ppm | 1.24 | P | ppm | 294 | U | ppm | 1.47
Ge | ppb | 2080 | Pr | ppm | 1.68 | V | ppm | 12.2
Hf | ppb | 880 | Rb | ppm | 27.38 | W | ppm | 0.52
Hg | ppb | 2000 | Re | ppb | < 1 | Y | ppm | 5.18
Ho | ppb | 200 | Sc | ppm | 1.4 | Yb | ppb | 520
In | ppm | 1.16 | Se | ppm | 7.44 | Zr | ppm | 27.75
Aqua Regia Digestion (sample weights 0.15-1.0g)
Au | ppb | < 0.5 | In | ppm | 1.202 | Se | ppm | 6.92
B | ppm | < 10 | K | wt.% | 0.15 | Sm | ppm | 1.44
Ba | ppm | < 600 | La | ppm | 6.54 | Sn | ppm | 0.932
Be | ppm | 0.7 | Li | ppm | 15.28 | Sr | ppm | 23.835
CaO | wt.% | 6.0901 | Lu | ppb | < 100 | Ta | ppb | < 50
Ce | ppm | 13.048 | Mn | ppm | 3854.1 | Tb | ppb | 200
Cr | ppm | 12.88 | Mo | ppm | 6.8 | Te | ppb | 100
Cs | ppm | 0.684 | Na | ppm | 161 | Th | ppm | 2.18
Dy | ppm | 1.11 | Nb | ppm | < 0.1 | Ti | ppm | 60
Er | ppm | 0.56 | Nd | ppm | 6.366 | Tl | ppm | 66.81
Eu | ppb | 360 | Ni | ppm | 21.64 | Tm | ppb | < 100
Ga | ppm | 3.13 | P | ppm | 281 | U | ppm | 0.96
Gd | ppm | 1.34 | Pr | ppm | 1.54 | V | ppm | 7.4
Ge | ppb | 475 | Rb | ppm | 11.7 | W | ppm | 0.24
Hf | ppb | 280 | Re | ppb | 2.6 | Y | ppm | 3
Hg | ppb | 2260 | Sb | ppm | 101.3958 | Yb | ppb | 480
Ho | ppb | 200 | Sc | ppm | < 0.5 | Zr | ppm | 10.4

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 134a
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate / Peroxide Fusion ICP
Ag, ppm | 197.42 | 9.9841 | 177.4518 | 217.3882 | 167.4677 | 227.3723 | 5.057% | 10.11% | 15.17% | 187.549 | 207.291
Al2O3, wt.% | 2.6025 | 0.13113 | 2.3402 | 2.8648 | 2.2091 | 2.9959 | 5.039% | 10.08% | 15.12% | 2.4724 | 2.7326
As, ppm | 217.7 | 21.9748 | 173.7504 | 261.6496 | 151.7756 | 283.6244 | 10.09% | 20.19% | 30.28% | 206.815 | 228.585
Ba, ppm | 1500.2639 | 155.6018 | 1189.0602 | 1811.4676 | 1033.4584 | 1967.0694 | 10.37% | 20.74% | 31.11% | 1425.2507 | 1575.2771
CaO, wt.% | 6.3848 | 0.19651 | 5.9918 | 6.7779 | 5.7953 | 6.9744 | 3.078% | 6.155% | 9.233% | 6.0656 | 6.7041
Cd, ppm | 545.8 | 35.8062 | 474.1876 | 617.4124 | 438.3814 | 653.2186 | 6.56% | 13.12% | 19.68% | 518.51 | 573.09
Co, ppm | 100.1622 | 13.1722 | 73.8178 | 126.5066 | 60.6456 | 139.6788 | 13.15% | 26.3% | 39.45% | 95.1541 | 105.1703
Fe, wt.% | 12.3213 | 0.79513 | 10.7311 | 13.9116 | 9.9359 | 14.7067 | 6.453% | 12.91% | 19.36% | 11.7053 | 12.9374
MgO, wt.% | 3.5548 | 0.11891 | 3.3169 | 3.7926 | 3.198 | 3.9115 | 3.345% | 6.69% | 10.04% | 3.377 | 3.7325
Pb, wt.% | 12.9451 | 0.62283 | 11.6994 | 14.1907 | 11.0766 | 14.8135 | 4.811% | 9.623% | 14.43% | 12.2978 | 13.5923
SiO2, wt.% | 14.8131 | 0.8451 | 13.1229 | 16.5033 | 12.2778 | 17.3484 | 5.705% | 11.41% | 17.12% | 14.0724 | 15.5537
Zn, wt.% | 17.4916 | 0.50511 | 16.4814 | 18.5019 | 15.9763 | 19.007 | 2.888% | 5.775% | 8.663% | 16.6171 | 18.3662
4-Acid Digestion
Ag, ppm | 200.8218 | 6.7526 | 187.3166 | 214.327 | 180.564 | 221.0796 | 3.362% | 6.725% | 10.09% | 190.7807 | 210.8629
Al2O3, wt.% | 2.5326 | 0.11686 | 2.2989 | 2.7663 | 2.182 | 2.8831 | 4.614% | 9.228% | 13.84% | 2.4059 | 2.6592
As, ppm | 227.7698 | 22.911 | 181.9478 | 273.5917 | 159.0368 | 296.5027 | 10.06% | 20.12% | 30.18% | 216.3813 | 239.1583
CaO, wt.% | 6.0361 | 0.3847 | 5.2667 | 6.8055 | 4.882 | 7.1902 | 6.373% | 12.75% | 19.12% | 5.7343 | 6.3379
Cd, ppm | 535.6146 | 36.0705 | 463.4737 | 607.7556 | 427.4032 | 643.826 | 6.734% | 13.47% | 20.2% | 508.8339 | 562.3953
Co, ppm | 104.9669 | 8.2381 | 88.4908 | 121.4431 | 80.2527 | 129.6812 | 7.848% | 15.7% | 23.54% | 99.7186 | 110.2153
Cu, ppm | 1290.6657 | 45.0472 | 1200.5714 | 1380.76 | 1155.5242 | 1425.8072 | 3.49% | 6.98% | 10.47% | 1226.1324 | 1355.199
Fe, wt.% | 12.0394 | 0.41049 | 11.2184 | 12.8604 | 10.8079 | 13.2709 | 3.41% | 6.819% | 10.23% | 11.4374 | 12.6414
MgO, wt.% | 3.4938 | 0.18412 | 3.1256 | 3.8621 | 2.9415 | 4.0462 | 5.27% | 10.54% | 15.81% | 3.3191 | 3.6685
Pb, wt.% | 12.7865 | 0.76631 | 11.2538 | 14.3191 | 10.4875 | 15.0854 | 5.993% | 11.99% | 17.98% | 12.1471 | 13.4258
S, wt.% | 19.2749 | 0.50272 | 18.2694 | 20.2803 | 17.7667 | 20.783 | 2.608% | 5.216% | 7.824% | 18.3111 | 20.2386
Sb, ppm | 115.4795 | 14.8451 | 85.7892 | 145.1698 | 70.9441 | 160.015 | 12.86% | 25.71% | 38.57% | 109.7056 | 121.2535
Zn, wt.% | 17.2671 | 0.5526 | 16.1619 | 18.3723 | 15.6093 | 18.9249 | 3.2% | 6.401% | 9.601% | 16.4037 | 18.1305
Aqua Regia Digestion (sample weights 0.15-1.0g)
Ag, ppm | 193.5817 | 13.1673 | 167.2471 | 219.9163 | 154.0798 | 233.0836 | 6.802% | 13.6% | 20.41% | 183.9026 | 203.2608
Al2O3, wt.% | 0.77352 | 0.085461 | 0.60259 | 0.94444 | 0.51713 | 1.0299 | 11.05% | 22.1% | 33.15% | 0.73484 | 0.81219
As, ppm | 225.505 | 24.0738 | 177.3573 | 273.6527 | 153.2835 | 297.7265 | 10.68% | 21.35% | 32.03% | 214.2297 | 236.7802
Cd, ppm | 546.7723 | 44.4356 | 457.9011 | 635.6435 | 413.4655 | 680.0791 | 8.127% | 16.25% | 24.38% | 519.4337 | 574.1109
Co, ppm | 101.1341 | 8.1863 | 84.7615 | 117.5067 | 76.5753 | 125.693 | 8.094% | 16.19% | 24.28% | 96.0774 | 106.1908
Cu, ppm | 1279.4917 | 78.3751 | 1122.7415 | 1436.2418 | 1044.3664 | 1514.6169 | 6.125% | 12.25% | 18.38% | 1215.5171 | 1343.4662
Fe, wt.% | 11.8925 | 0.97309 | 9.9463 | 13.8386 | 8.9732 | 14.8117 | 8.182% | 16.36% | 24.55% | 11.2978 | 12.4871
MgO, wt.% | 3.425 | 0.13741 | 3.1502 | 3.6999 | 3.0128 | 3.8373 | 4.012% | 8.024% | 12.04% | 3.2538 | 3.5963
Pb, wt.% | 12.6375 | 0.759 | 11.1195 | 14.1555 | 10.3605 | 14.9145 | 6.006% | 12.01% | 18.02% | 12.0056 | 13.2694
S, wt.% | 18.4925 | 1.3285 | 15.8355 | 21.1496 | 14.5069 | 22.4781 | 7.184% | 14.37% | 21.55% | 17.5679 | 19.4171
Zn, wt.% | 16.9999 | 0.86337 | 15.2732 | 18.7266 | 14.4098 | 19.59 | 5.079% | 10.16% | 15.24% | 16.1499 | 17.8499
Infrared Combustion
S, wt.% | 19.1121 | 0.41288 | 18.2864 | 19.9379 | 17.8735 | 20.3507 | 2.16% | 4.321% | 6.481% | 18.1565 | 20.0677

## OxFusion XRF
| Analytical results for Al2O3 in OREAS 134a (Indicative Value 2.57 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 2
2
| 1 | 1 | 2.563 | 1
| 1 | 2 | 2.577 | 1
| Mean |  | 2.57 | 16
| Median |  | 2.57 | 2.57
| Std Dev. |  | 0.0098995 | 7
| Rel.Std.Dev. |  | 0.3852%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 134a (Indicative Value 240 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 240 | 1
| 1 | 2 | 240 | 2
| Mean |  | 240 | 16
| Median |  | 240 | 240
| Std Dev. |  | 0 | 8
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 134a (Indicative Value 1502 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 1496 | 1
| 1 | 2 | 1507 | 3
| Mean |  | 1501.5 | 16
| Median |  | 1501.5 | 1501.6925
| Std Dev. |  | 7.7782 | 9
| Rel.Std.Dev. |  | 0.518%
| PDM3 |  | -0.01282%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 134a (Indicative Value 5.94 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 2
2
| 1 | 1 | 5.923 | 1
| 1 | 2 | 5.959 | 4
| Mean |  | 5.941 | 16
| Median |  | 5.941 | 5.941
| Std Dev. |  | 0.025456 | 10
| Rel.Std.Dev. |  | 0.4285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 134a (Indicative Value 190 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 200 | 5
| Mean |  | 190 | 16
| Median |  | 190 | 190
| Std Dev. |  | 14.1421 | 11
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CoO in OREAS 134a (Indicative Value 140 ppm) |  |  | Indicative
CoO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 140 | 1
| 1 | 2 | 140 | 6
| Mean |  | 140 | 16
| Median |  | 140 | 139.865
| Std Dev. |  | 0 | 12
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.09652%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 134a (Indicative Value 51 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 58 | 1
| 1 | 2 | 44 | 7
| Mean |  | 51 | 16
| Median |  | 51 | 51.1525
| Std Dev. |  | 9.8995 | 13
| Rel.Std.Dev. |  | 19.41%
| PDM3 |  | -0.2981%
| Z-Score (Absolute) |  | NA
| Analytical results for CuO in OREAS 134a (Indicative Value 1646 ppm) |  |  | Indicative
CuO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 1652 | 1
| 1 | 2 | 1640 | 8
| Mean |  | 1646 | 16
| Median |  | 1646 | 1646.117
| Std Dev. |  | 8.4853 | 14
| Rel.Std.Dev. |  | 0.5155%
| PDM3 |  | -0.007108%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 134a (Indicative Value 17.69 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 2
2
| 1 | 1 | 17.6711 | 1
| 1 | 2 | 17.6997 | 1
| Mean |  | 17.6854 | 16
| Median |  | 17.6854 | 17.6854
| Std Dev. |  | 0.020223 | 7
| Rel.Std.Dev. |  | 0.1143%
| PDM3 |  | 6.22e-05%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 134a (Indicative Value 0.962 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 3
3
| 1 | 1 | 0.96 | 1
| 1 | 2 | 0.963 | 2
| Mean |  | 0.9615 | 16
| Median |  | 0.9615 | 0.9615
| Std Dev. |  | 0.0021213 | 8
| Rel.Std.Dev. |  | 0.2206%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 134a (Indicative Value 3.55 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 2
2
| 1 | 1 | 3.55 | 1
| 1 | 2 | 3.553 | 3
| Mean |  | 3.5515 | 16
| Median |  | 3.5515 | 3.5515
| Std Dev. |  | 0.0021213 | 9
| Rel.Std.Dev. |  | 0.05973%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 134a (Indicative Value 0.459 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 3
3
| 1 | 1 | 0.459 | 1
| 1 | 2 | 0.458 | 4
| Mean |  | 0.4585 | 16
| Median |  | 0.4585 | 0.4585
| Std Dev. |  | 0.00070711 | 10
| Rel.Std.Dev. |  | 0.1542%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for NiO in OREAS 134a (Indicative Value 31.8 ppm) |  |  | Indicative
NiO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 1
1
| 1 | 1 | 25 | 1
| 1 | 2 | 38 | 5
| Mean |  | 31.5 | 16
| Median |  | 31.5 | 31.8125
| Std Dev. |  | 9.1924 | 11
| Rel.Std.Dev. |  | 29.18%
| PDM3 |  | -0.9823%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 134a (Indicative Value 0.079 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 3
3
| 1 | 1 | 0.0802 | 1
| 1 | 2 | 0.0779 | 6
| Mean |  | 0.07905 | 16
| Median |  | 0.07905 | 0.07906
| Std Dev. |  | 0.0016263 | 12
| Rel.Std.Dev. |  | 2.057%
| PDM3 |  | -0.0129%
| Z-Score (Absolute) |  | NA
| Analytical results for PbO in OREAS 134a (Indicative Value 137774 ppm) |  |  | Indicative
PbO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 137666 | 1
| 1 | 2 | 137882 | 7
| Mean |  | 137774 | 16
| Median |  | 137774 | 137773.88
| Std Dev. |  | 152.7351 | 13
| Rel.Std.Dev. |  | 0.1109%
| PDM3 |  | 8.71e-05%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 134a (Indicative Value 14.82 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 2
2
| 1 | 1 | 14.8 | 1
| 1 | 2 | 14.84 | 8
| Mean |  | 14.82 | 16
| Median |  | 14.82 | 14.82
| Std Dev. |  | 0.028284 | 14
| Rel.Std.Dev. |  | 0.1909%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SnO2 in OREAS 134a (Indicative Value < 13 ppm) |  |  | Indicative
SnO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 1
1
| 1 | 1 | < 13 | 1
| 1 | 2 | < 13 | 1
| Mean |  |  | 16
| Median |  |  | < 13
| Std Dev. |  |  | 7
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 134a (Indicative Value 49.29 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 2
2
| 1 | 1 | 49.145 | 1
| 1 | 2 | 49.445 | 2
| Mean |  | 49.295 | 16
| Median |  | 49.295 | 49.2947
| Std Dev. |  | 0.21213 | 8
| Rel.Std.Dev. |  | 0.4303%
| PDM3 |  | 0.0005518%
| Z-Score (Absolute) |  | NA
| Analytical results for SrO in OREAS 134a (Indicative Value 35.5 ppm) |  |  | Indicative
SrO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 1
1
| 1 | 1 | 35 | 1
| 1 | 2 | 35 | 3
| Mean |  | 35 | 16
| Median |  | 35 | 35.478
| Std Dev. |  | 0 | 9
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.347%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 134a (Indicative Value 0.093 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | OxBF*XRF | 3
3
| 1 | 1 | 0.093 | 1
| 1 | 2 | 0.093 | 4
| Mean |  | 0.093 | 16
| Median |  | 0.093 | 0.093
| Std Dev. |  | 0 | 10
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 134a (Indicative Value 71 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 71 | 1
| 1 | 2 | 71 | 5
| Mean |  | 71 | 16
| Median |  | 71 | 71.408
| Std Dev. |  | 0 | 11
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.5714%
| Z-Score (Absolute) |  | NA
| Analytical results for ZnO in OREAS 134a (Indicative Value 215350 ppm) |  |  | Indicative
ZnO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 0
0
| 1 | 1 | 215101 | 1
| 1 | 2 | 215599 | 6
| Mean |  | 215350 | 16
| Median |  | 215350 | 215350.4
| Std Dev. |  | 352.1392 | 12
| Rel.Std.Dev. |  | 0.1635%
| PDM3 |  | -0.0001857%
| Z-Score (Absolute) |  | NA
| Analytical results for ZrO2 in OREAS 134a (Indicative Value 40.5 ppm) |  |  | Indicative
ZrO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | OxBF*XRF | 1
1
| 1 | 1 | 41 | 1
| 1 | 2 | 41 | 7
| Mean |  | 41 | 16
| Median |  | 41 | 40.524
| Std Dev. |  | 0 | 13
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 1.175%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 134a (Indicative Value 14.59 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 14.61 | 1
| 1 | 2 | 14.57 | 10
| Mean |  | 14.59 | 16
| Median |  | 14.59 | 14.59
| Std Dev. |  | 0.028284 | 16
| Rel.Std.Dev. |  | 0.1939%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 134a (Indicative Value 193 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 194 | 1
| 1 | 2 | 192 | 12
| Mean |  | 193 | 16
| Median |  | 193 | 193
| Std Dev. |  | 1.4142 | 18
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 134a (Indicative Value 236 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 237 | 1
| 1 | 2 | 234 | 2
| Mean |  | 235.5 | 16
| Median |  | 235.5 | 235.5
| Std Dev. |  | 2.1213 | 19
| Rel.Std.Dev. |  | 0.9008%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 134a (Indicative Value 1345 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1340 | 1
| 1 | 2 | 1350 | 14
| Mean |  | 1345 | 16
| Median |  | 1345 | 1345
| Std Dev. |  | 7.0711 | 20
| Rel.Std.Dev. |  | 0.5257%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 134a (Indicative Value 1.8 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.8 | 15
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 134a (Indicative Value 0.19 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.18 | 1
| 1 | 2 | 0.2 | 16
| Mean |  | 0.19 | 16
| Median |  | 0.19 | 0.19
| Std Dev. |  | 0.014142 | 22
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 134a (Indicative Value 532 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 521 | 1
| 1 | 2 | 542 | 17
| Mean |  | 531.5 | 16
| Median |  | 531.5 | 531.5
| Std Dev. |  | 14.8492 | 23
| Rel.Std.Dev. |  | 2.794%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 134a (Indicative Value 16.3 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.1 | 1
| 1 | 2 | 16.4 | 18
| Mean |  | 16.25 | 16
| Median |  | 16.25 | 16.25
| Std Dev. |  | 0.21213 | 24
| Rel.Std.Dev. |  | 1.305%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 134a (Indicative Value 103 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 98.7 | 1
| 1 | 2 | 108 | 19
| Mean |  | 103.35 | 16
| Median |  | 103.35 | 103.35
| Std Dev. |  | 6.5761 | 25
| Rel.Std.Dev. |  | 6.363%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 134a (Indicative Value 10 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12 | 1
| 1 | 2 | 8 | 20
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 2.8284 | 26
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 134a (Indicative Value 1.08 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.05 | 1
| 1 | 2 | 1.11 | 21
| Mean |  | 1.08 | 16
| Median |  | 1.08 | 1.08
| Std Dev. |  | 0.042426 | 27
| Rel.Std.Dev. |  | 3.928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 134a (Indicative Value 1340 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1350 | 1
| 1 | 2 | 1330 | 22
| Mean |  | 1340 | 16
| Median |  | 1340 | 1340
| Std Dev. |  | 14.1421 | 28
| Rel.Std.Dev. |  | 1.055%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 134a (Indicative Value 0.81 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.78 | 23
| Mean |  | 0.81 | 16
| Median |  | 0.81 | 0.81
| Std Dev. |  | 0.042426 | 29
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 134a (Indicative Value 0.48 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.51 | 24
| Mean |  | 0.48 | 16
| Median |  | 0.48 | 0.48
| Std Dev. |  | 0.042426 | 30
| Rel.Std.Dev. |  | 8.839%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 134a (Indicative Value 320 ppb) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 320 | 1
| 1 | 2 | 320 | 25
| Mean |  | 320 | 16
| Median |  | 320 | 320
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 134a (Indicative Value 4.55 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 4.5 | 26
| Mean |  | 4.55 | 16
| Median |  | 4.55 | 4.55
| Std Dev. |  | 0.070711 | 32
| Rel.Std.Dev. |  | 1.554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 134a (Indicative Value 1.2 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.24 | 1
| 1 | 2 | 1.16 | 27
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0.056569 | 33
| Rel.Std.Dev. |  | 4.714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 134a (Indicative Value 18150 ppb) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 18100 | 1
| 1 | 2 | 18200 | 47
| Mean |  | 18150 | 16
| Median |  | 18150 | 18150
| Std Dev. |  | 70.7107 | 18
| Rel.Std.Dev. |  | 0.3896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 134a (Indicative Value 865 ppb) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 780 | 1
| 1 | 2 | 950 | 48
| Mean |  | 865 | 16
| Median |  | 865 | 865
| Std Dev. |  | 120.2082 | 19
| Rel.Std.Dev. |  | 13.9%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 134a (Indicative Value 180 ppb) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 180 | 50
| Mean |  | 180 | 16
| Median |  | 180 | 180
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 134a (Indicative Value 0.98 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.95 | 1
| 1 | 2 | 1 | 51
| Mean |  | 0.975 | 16
| Median |  | 0.975 | 0.975
| Std Dev. |  | 0.035355 | 21
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 134a (Indicative Value 8.64 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.57 | 1
| 1 | 2 | 8.71 | 16
| Mean |  | 8.64 | 16
| Median |  | 8.64 | 8.64
| Std Dev. |  | 0.098995 | 22
| Rel.Std.Dev. |  | 1.146%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 134a (Indicative Value 75 ppb) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 70 | 1
| 1 | 2 | 80 | 17
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 134a (Indicative Value 6.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 6.6 | 18
| Mean |  | 6.6 | 16
| Median |  | 6.6 | 6.6
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 134a (Indicative Value 2.05 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 2.1 | 19
| Mean |  | 2.05 | 16
| Median |  | 2.05 | 2.05
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 134a (Indicative Value 6.88 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.01 | 1
| 1 | 2 | 6.74 | 20
| Mean |  | 6.875 | 16
| Median |  | 6.875 | 6.875
| Std Dev. |  | 0.19092 | 26
| Rel.Std.Dev. |  | 2.777%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 134a (Indicative Value 22 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 24 | 21
| Mean |  | 22 | 16
| Median |  | 22 | 22
| Std Dev. |  | 2.8284 | 27
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 134a (Indicative Value 12.6 wt.%) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 2
2
| 1 | 1 | 12.5 | 1
| 1 | 2 | 12.7 | 22
| Mean |  | 12.6 | 16
| Median |  | 12.6 | 12.6
| Std Dev. |  | 0.14142 | 28
| Rel.Std.Dev. |  | 1.122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 134a (Indicative Value 1.86 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.84 | 1
| 1 | 2 | 1.88 | 23
| Mean |  | 1.86 | 16
| Median |  | 1.86 | 1.86
| Std Dev. |  | 0.028284 | 29
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 134a (Indicative Value 24.9 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25.5 | 1
| 1 | 2 | 24.2 | 24
| Mean |  | 24.85 | 16
| Median |  | 24.85 | 24.85
| Std Dev. |  | 0.91924 | 30
| Rel.Std.Dev. |  | 3.699%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 134a (Indicative Value 20 ppb) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | < 10 | 25
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  |  | 31
Rel.Std.Dev.
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 134a (Indicative Value 118 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 117 | 1
| 1 | 2 | 119 | 26
| Mean |  | 118 | 16
| Median |  | 118 | 118
| Std Dev. |  | 1.4142 | 32
| Rel.Std.Dev. |  | 1.198%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 134a (Indicative Value 1.8 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.8 | 27
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 134a (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 51
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 18
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 134a (Indicative Value 1.34 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.38 | 1
| 1 | 2 | 1.3 | 13
| Mean |  | 1.34 | 16
| Median |  | 1.34 | 1.34
| Std Dev. |  | 0.056569 | 19
| Rel.Std.Dev. |  | 4.222%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 134a (Indicative Value 1 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 1.2 | 14
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0.28284 | 20
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 134a (Indicative Value 25.5 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25.3 | 1
| 1 | 2 | 25.6 | 34
| Mean |  | 25.45 | 16
| Median |  | 25.45 | 25.45
| Std Dev. |  | 0.21213 | 21
| Rel.Std.Dev. |  | 0.8335%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 134a (Indicative Value 205 ppb) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 230 | 16
| Mean |  | 205 | 16
| Median |  | 205 | 205
| Std Dev. |  | 35.3553 | 22
| Rel.Std.Dev. |  | 17.25%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 134a (Indicative Value 180 ppb) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 180 | 17
| Mean |  | 180 | 16
| Median |  | 180 | 180
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 134a (Indicative Value < 200 ppb) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | < 200 | 1
| 1 | 2 | < 200 | 18
| Mean |  |  | 16
| Median |  |  | < 200
| Std Dev. |  |  | 24
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 134a (Indicative Value 2.56 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.71 | 19
| Mean |  | 2.555 | 16
| Median |  | 2.555 | 2.555
| Std Dev. |  | 0.2192 | 25
| Rel.Std.Dev. |  | 8.579%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 134a (Indicative Value 81 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 80.4 | 1
| 1 | 2 | 80.6 | 20
| Mean |  | 80.5 | 16
| Median |  | 80.5 | 80.5
| Std Dev. |  | 0.14142 | 26
| Rel.Std.Dev. |  | 0.1757%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 134a (Indicative Value 95 ppb) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 100 | 1
| 1 | 2 | 90 | 21
| Mean |  | 95 | 16
| Median |  | 95 | 95
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 134a (Indicative Value 1.37 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.33 | 22
| Mean |  | 1.365 | 16
| Median |  | 1.365 | 1.365
| Std Dev. |  | 0.049497 | 28
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 134a (Indicative Value 13.1 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.3 | 1
| 1 | 2 | 13.8 | 23
| Mean |  | 13.05 | 16
| Median |  | 13.05 | 13.05
| Std Dev. |  | 1.0607 | 29
| Rel.Std.Dev. |  | 8.128%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 134a (Indicative Value 0.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.5 | 24
| Mean |  | 0.5 | 16
| Median |  | 0.5 | 0.5
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 134a (Indicative Value 5.73 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.88 | 1
| 1 | 2 | 5.58 | 25
| Mean |  | 5.73 | 16
| Median |  | 5.73 | 5.73
| Std Dev. |  | 0.21213 | 31
| Rel.Std.Dev. |  | 3.702%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 134a (Indicative Value 515 ppb) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 510 | 1
| 1 | 2 | 520 | 26
| Mean |  | 515 | 16
| Median |  | 515 | 515
| Std Dev. |  | 7.0711 | 32
| Rel.Std.Dev. |  | 1.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 134a (Indicative Value 27.8 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27.5 | 1
| 1 | 2 | 28 | 27
| Mean |  | 27.75 | 16
| Median |  | 27.75 | 27.75
| Std Dev. |  | 0.35355 | 33
| Rel.Std.Dev. |  | 1.274%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 134a (Certified Value 197 ppm) | [11 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.10 | 1.12 | 1.16 | [5 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES | [5 empty] | 0
0
| 1 | 1 | 198.5 | 202 | 185 | 200 | 196 | [5 empty] | 1
| 1 | 2 | 199.2 | 198 | 190 | 210 | 190 | [5 empty] | 12
| 1 | 3 | 197.6 | 199 | 195 | 210 | 190 | [5 empty] | 16
| 1 | 4 | 201.7 | 198 | 200 | 220 | 177 | [5 empty] | 197.42
| 1 | 5 | 197.5 | 198 | 200 | 210 | 173 | [5 empty] | 7
| Mean |  | 198.9 | 199 | 194 | 210 | 185.2
| Median |  | 198.5 | 198 | 195 | 210 | 190
| Std Dev. |  | 1.7132 | 1.7321 | 6.5192 | 7.0711 | 9.7314
| Rel.Std.Dev. |  | 0.8613% | 0.8704% | 3.36% | 3.367% | 5.255%
| PDM3 |  | 0.7497% | 0.8003% | -1.732% | 6.372% | -6.19%
| Z-Score (Absolute) |  | 0 | 0.01 | 0.67 | 1.53 | 1.89
| Analytical results for Al2O3 in OREAS 134a (Certified Value 2.6 wt.%) | [11 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.16 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  | 2
3
| 1 | 1 | 2.618 | 2.74 | NVL | 2.456 | 2.57 | 2.51 | 2.4696 | 2.73 | 2.683 |  | 1
| 1 | 2 | 2.539 | 2.626 | NVL | 2.437 | 2.551 | 2.53 | 2.4923 | 2.68 | 2.853 |  | 1
| 1 | 3 | 2.632 | 2.683 | NVL | 2.419 | 2.702 | 2.59 | 2.4923 | 2.73 | 2.853 |  | 16
| 1 | 4 | 2.662 | 2.759 | NVL | 2.381 | 2.608 | 2.47 | 2.4923 | 2.7 | 2.74 |  | 2.6025
| 1 | 5 | 2.682 | 2.664 | NVL | 2.324 | 2.608 | 2.51 | 2.4847 | 2.7 | 2.815 |  | 8
| Mean |  | 2.6266 | 2.6944 |  | 2.4034 | 2.6078 | 2.522 | 2.4862 | 2.708 | 2.7888
| Median |  | 2.632 | 2.683 |  | 2.419 | 2.608 | 2.51 | 2.4923 | 2.7 | 2.815
| Std Dev. |  | 0.054998 | 0.054738 |  | 0.05231 | 0.058157 | 0.043818 | 0.009867 | 0.021679 | 0.075008
| Rel.Std.Dev. |  | 2.094% | 2.032% |  | 2.176% | 2.23% | 1.737% | 0.3969% | 0.8006% | 2.69%
| PDM3 |  | 0.9256% | 3.531% |  | -7.651% | 0.2032% | -3.094% | -4.468% | 4.053% | 7.158%
| Z-Score (Absolute) |  | 0.07 | 0.56 | NA | 1.55 | 0.07 | 0.69 | 0.95 | 0.66 | 1.25
| Analytical results for As in OREAS 134a (Certified Value 218 ppm) | [11 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 |  |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | 185 | 237 | 230 | 200 | 220 | 221 | 200 | < 1200 |  |  | 1
| 1 | 2 | 186 | 251 | 250 | 200 | 220 | 224 | 200 | < 1200 |  |  | 2
| 1 | 3 | 193 | 238 | 240 | 200 | 235 | 220 | 200 | < 1200 |  |  | 16
| 1 | 4 | 183 | 243 | 240 | 200 | 235 | 219 | 200 | < 1200 |  |  | 217.7
| 1 | 5 | 184 | 247 | 250 | 200 | 250 | 227 | 200 | < 1200 |  |  | 9
| Mean |  | 186.2 | 243.2 | 242 | 200 | 232 | 222.2 | 200
| Median |  | 185 | 243 | 240 | 200 | 235 | 221 | 200
| Std Dev. |  | 3.9623 | 5.933 | 8.3666 | 0 | 12.5499 | 3.2711 | 0
| Rel.Std.Dev. |  | 2.128% | 2.44% | 3.457% | 0% | 5.409% | 1.472% | 0%
| PDM3 |  | -14.47% | 11.71% | 11.16% | -8.13% | 6.569% | 2.067% | -8.13%
| Z-Score (Absolute) |  | 1.28 | 0.5 | 0.47 | 0.85 | 0.15 | 0.15 | 0.85 | 11.64
| Analytical results for Ba in OREAS 134a (Certified Value 1500 ppm) | [11 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.16 |  | ppm
|  |  | PF*OES | PF*MS | PF*OES | PPP*XRF | PPP*XRF | PF*MS | BF*OES | BF*OES | PF*OES |  | 0
0
| 1 | 1 | 1022 | 1516 | NVL | 1270 | 900 | 1590 | 1454 | 1396 | 1680 |  | 1
| 1 | 2 | 1003 | 1532 | NVL | 1320 | 900 | 1600 | 1452 | 1385 | 1750 |  | 14
| 1 | 3 | 1072 | 1529 | NVL | 1300 | 900 | 1670 | 1455 | 1301 | 1800 |  | 16
| 1 | 4 | 1090 | 1532 | NVL | 1310 | 910 | 1615 | 1451 | 1491 | 1740 |  | 1500.2639
| 1 | 5 | 1044 | 1481 | NVL | 1230 | 900 | 1650 | 1458 | 1381 | 1640 |  | 10
| Mean |  | 1046.2 | 1518 |  | 1286 | 902 | 1625 | 1454 | 1390.8 | 1722
| Median |  | 1044 | 1529 |  | 1300 | 900 | 1615 | 1454 | 1385 | 1740
| Std Dev. |  | 35.4993 | 21.7141 |  | 36.4692 | 4.4721 | 33.9116 | 2.7386 | 67.5663 | 62.6099
| Rel.Std.Dev. |  | 3.393% | 1.43% |  | 2.836% | 0.4958% | 2.087% | 0.1884% | 4.858% | 3.636%
| PDM3 |  | -30.27% | 1.182% |  | -14.28% | -39.88% | 8.314% | -3.084% | -7.296% | 14.78%
| Z-Score (Absolute) |  | 1.5 | 0.38 | NA | 0.54 | 2.07 | 0.81 | 0.13 | 0.13 | 1.19
| Analytical results for Be in OREAS 134a (Indicative Value 1 ppm) | [11 empty] | Indicative
Be | Round | Replicate | Lab | [9 empty] | 1
| No | No | 1.12 | [9 empty] | ppm
|  |  | BF*OES | [9 empty] | 2
2
| 1 | 1 | 1 | [9 empty] | 1
| 1 | 2 | 1 | [9 empty] | 15
| 1 | 3 | 1 | [9 empty] | 16
| 1 | 4 | 1 | [9 empty] | 1
| 1 | 5 | 1 | [9 empty] | 35
| Mean |  | 1
| Median |  | 1
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 134a (Certified Value 6.38 wt.%) | [11 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES | 2
3
| 1 | 1 | 6.244 | 6.44 | > 55.96 | 6.688 | 5.765 | 6.534 | 6.3 | 6.0194 | 6.63 | 6.842 | 1
| 1 | 2 | 6.383 | 6.16 | > 55.96 | 6.66 | 5.401 | 6.478 | 6.3 | 6.0501 | 6.5 | 7.024 | 4
| 1 | 3 | 6.392 | 6.3 | > 55.96 | 6.548 | 5.849 | 6.674 | 6.4 | 6.0571 | 6.64 | 7.178 | 16
| 1 | 4 | 6.38 | 6.3 | > 55.96 | 6.436 | 5.723 | 6.548 | 6.3 | 6.0361 | 6.63 | 6.968 | 6.3848
| 1 | 5 | 6.322 | 6.3 | > 55.96 | 6.296 | 5.597 | 6.506 | 6.3 | 6.0529 | 6.59 | 6.604 | 11
| Mean |  | 6.3442 | 6.3 |  | 6.5256 | 5.667 | 6.548 | 6.32 | 6.0431 | 6.598 | 6.9232
| Median |  | 6.38 | 6.3 |  | 6.548 | 5.723 | 6.534 | 6.3 | 6.0501 | 6.63 | 6.968
| Std Dev. |  | 0.06246 | 0.098995 |  | 0.16254 | 0.1743 | 0.075392 | 0.044721 | 0.015425 | 0.058052 | 0.21539
| Rel.Std.Dev. |  | 0.9845% | 1.571% |  | 2.491% | 3.076% | 1.151% | 0.7076% | 0.2552% | 0.8798% | 3.111%
| PDM3 |  | -0.6366% | -1.329% |  | 2.204% | -11.24% | 2.555% | -1.016% | -5.352% | 3.338% | 8.432%
| Z-Score (Absolute) |  | 0 | 0.16 | NA | 0.6 | 2.24 | 0.67 | 0.08 | 0.99 | 0.84 | 1.91
| Analytical results for Cd in OREAS 134a (Certified Value 546 ppm) | [11 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.09 | 1.10 | 1.16 |  |  |  |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES |  |  |  |  | 0
0
| 1 | 1 | 533 | 628 | 528 | 600 | 550 | 582 |  |  |  |  | 1
| 1 | 2 | 533 | 624 | 512 | 600 | 550 | 611 |  |  |  |  | 17
| 1 | 3 | 528 | 629 | 518 | 500 | 560 | 620 |  |  |  |  | 16
| 1 | 4 | 525 | 639 | 513 | 500 | 550 | 583 |  |  |  |  | 545.8
| 1 | 5 | 528 | 639 | 517 | 500 | 540 | 564 |  |  |  |  | 12
| Mean |  | 529.4 | 631.8 | 517.6 | 540 | 550 | 592
| Median |  | 528 | 629 | 517 | 500 | 550 | 583
| Std Dev. |  | 3.5071 | 6.8337 | 6.3482 | 54.7723 | 7.0711 | 22.9674
| Rel.Std.Dev. |  | 0.6625% | 1.082% | 1.226% | 10.14% | 1.286% | 3.88%
| PDM3 |  | -3.005% | 15.76% | -5.167% | -1.063% | 0.7695% | 8.465%
| Z-Score (Absolute) |  | 0.49 | 2.72 | 0.86 | 0.16 | 0.16 | 1.47
| Analytical results for Co in OREAS 134a (Certified Value 100 ppm) | [11 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*MS | PF*OES | PF*OES | 0
0
| 1 | 1 | 102 | 107 | 107 | 90 | 100 | 100 | 120 | 87.9 | 80 | < 200 | 1
| 1 | 2 | 102 | 110 | 105 | 90 | 110 | 100 | 120 | 88.4 | 80 | < 200 | 19
| 1 | 3 | 101 | 109 | 104 | 90 | 120 | 100 | 120 | 87.6 | 80 | < 200 | 16
| 1 | 4 | 106 | 112 | 104 | 80 | 100 | 100 | 130 | 88 | 80 | < 200 | 100.1622
| 1 | 5 | 102 | 112 | 105 | 80 | 110 | 100 | 120 | 87.4 | 80 | < 200 | 13
| Mean |  | 102.6 | 110 | 105 | 86 | 108 | 100 | 122 | 87.86 | 80
| Median |  | 102 | 110 | 105 | 90 | 110 | 100 | 120 | 87.9 | 80
| Std Dev. |  | 1.9494 | 2.1213 | 1.2247 | 5.4772 | 8.3666 | 0 | 4.4721 | 0.38471 | 0
| Rel.Std.Dev. |  | 1.9% | 1.928% | 1.166% | 6.369% | 7.747% | 0% | 3.666% | 0.4379% | 0%
| PDM3 |  | 2.434% | 9.822% | 4.83% | -14.14% | 7.825% | -0.162% | 21.8% | -12.28% | -20.13%
| Z-Score (Absolute) |  | 0.11 | 0.76 | 0.32 | 1.34 | 0.59 | 0.11 | 1.81 | 1.18 | 1.87 | 0.11
| Analytical results for Cr in OREAS 134a (Indicative Value 80 ppm) | [11 empty] | Indicative
Cr | Round | Replicate | Lab | [9 empty] | 1
| No | No | 1.09 | [9 empty] | ppm
|  |  | PF*OES | [9 empty] | 0
0
| 1 | 1 | 100 | [9 empty] | 1
| 1 | 2 | < 100 | [9 empty] | 20
| 1 | 3 | < 100 | [9 empty] | 16
| 1 | 4 | 100 | [9 empty] | 80
| 1 | 5 | 100 | [9 empty] | 36
| Mean |  | 100
| Median |  | 100
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 25%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 134a (Indicative Value 1278 ppm) | [11 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.16 |  | ppm
|  |  | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 0
0
| 1 | 1 | 1276 | 1439 | 1280 | 1300 | 1290 | 1200 | 1240 | 1250 | 1420 |  | 1
| 1 | 2 | 1255 | 1401 | 1240 | 1300 | 1310 | 1200 | 1280 | 1240 | 1490 |  | 22
| 1 | 3 | 1272 | 1397 | 1240 | 1300 | 1330 | 1200 | 1300 | 1310 | 1460 |  | 16
| 1 | 4 | 1246 | 1424 | 1220 | 1300 | 1310 | 1200 | 1320 | 1210 | 1450 |  | 1277.7
| 1 | 5 | 1262 | 1381 | 1260 | 1300 | 1430 | 1200 | 1280 | 1300 | 1340 |  | 37
| Mean |  | 1262.2 | 1408.4 | 1248 | 1300 | 1334 | 1200 | 1284 | 1262 | 1432
| Median |  | 1262 | 1401 | 1240 | 1300 | 1310 | 1200 | 1280 | 1250 | 1450
| Std Dev. |  | 12.2556 | 22.9957 | 22.8035 | 0 | 55.4977 | 0 | 29.6648 | 42.0714 | 57.1839
| Rel.Std.Dev. |  | 0.971% | 1.633% | 1.827% | 0% | 4.16% | 0% | 2.31% | 3.334% | 3.993%
| PDM3 |  | -1.213% | 10.23% | -2.324% | 1.745% | 4.406% | -6.081% | 0.4931% | -1.229% | 12.08%
| Z-Score (Absolute) |  | 0.41 | 2.33 | 0.67 | 0.3 | 0.94 | 1.57 | 0 | 0.41 | 2.77
| Analytical results for Fe in OREAS 134a (Certified Value 12.32 wt.%) | [11 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 12.262 | 13.03 | NVL | 12.6 | 11.6 | 12.14 | 12.6 | 11.95 | 10.8 | 13.2 | 1
| 1 | 2 | 12.335 | 12.98 | NVL | 12.55 | 11.2 | 12.4 | 12.7 | 11.988 | 10.7 | 13.8 | 0
| 1 | 3 | 12.543 | 13.25 | NVL | 12.45 | 11.45 | 12.43 | 13.5 | 12.032 | 11 | 13.6 | 16
| 1 | 4 | 12.543 | 13.51 | NVL | 12.15 | 11.7 | 12.62 | 12.95 | 11.904 | 10.9 | 13.2 | 12.3213
| 1 | 5 | 12.366 | 13.26 | NVL | 11.9 | 11.25 | 12.11 | 13.1 | 12.007 | 11.1 | 12.8 | 14
| Mean |  | 12.4098 | 13.206 |  | 12.33 | 11.44 | 12.34 | 12.97 | 11.9762 | 10.9 | 13.32
| Median |  | 12.366 | 13.25 |  | 12.45 | 11.45 | 12.4 | 12.95 | 11.988 | 10.9 | 13.2
| Std Dev. |  | 0.12732 | 0.21173 |  | 0.29707 | 0.21622 | 0.21389 | 0.35637 | 0.050251 | 0.15811 | 0.38987
| Rel.Std.Dev. |  | 1.026% | 1.603% |  | 2.409% | 1.89% | 1.733% | 2.748% | 0.4196% | 1.451% | 2.927%
| PDM3 |  | 0.718% | 7.18% |  | 0.07034% | -7.153% | 0.1515% | 5.265% | -2.801% | -11.54% | 8.105%
| Z-Score (Absolute) |  | 0.07 | 0.93 | NA | 0.01 | 0.96 | 0 | 0.67 | 0.39 | 1.54 | 1.05
| Analytical results for K in OREAS 134a (Indicative Value 0.929 wt.%) | [11 empty] | Indicative
K | Round | Replicate | Lab | Lab | [8 empty] | 1
| No | No | 1.09 | 1.12 | [8 empty] | wt.%
|  |  | PF*OES | BF*OES | [8 empty] | 3
3
| 1 | 1 | 1.01 | 0.855 | [8 empty] | 1
| 1 | 2 | 1.04 | 0.83 | [8 empty] | 32
| 1 | 3 | 1.02 | 0.847 | [8 empty] | 16
| 1 | 4 | 1 | 0.838 | [8 empty] | 0.92938
| 1 | 5 | 0.99 | 0.863 | [8 empty] | 38
| Mean |  | 1.012 | 0.8466
| Median |  | 1.01 | 0.847
| Std Dev. |  | 0.019235 | 0.013126
| Rel.Std.Dev. |  | 1.901% | 1.55%
| PDM3 |  | 8.89% | -8.907%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for LOI1000 in OREAS 134a (Indicative Value 15.12 wt.%) | [11 empty] | Indicative
LOI1000 | Round | Replicate | Lab | [9 empty] | 1
| No | No | 1.12 | [9 empty] | wt.%
|  |  | BF*OES | [9 empty] | 2
2
| 1 | 1 | 15.18 | [9 empty] | 1
| 1 | 2 | 14.96 | [9 empty] | 10
| 1 | 3 | 15.46 | [9 empty] | 16
| 1 | 4 | 14.95 | [9 empty] | 15.12
| 1 | 5 | 15.05 | [9 empty] | 39
| Mean |  | 15.12
| Median |  | 15.05
| Std Dev. |  | 0.21131
| Rel.Std.Dev. |  | 1.398%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 134a (Certified Value 3.55 wt.%) | [11 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 3.469 | 3.714 | NVL | 3.631 | 3.399 | 3.449 | 3.731 | 3.5021 | 3.582 | 3.897 | 1
| 1 | 2 | 3.469 | 3.681 | NVL | 3.598 | 3.333 | 3.499 | 3.748 | 3.5187 | 3.532 | 4.063 | 3
| 1 | 3 | 3.4275 | 3.631 | NVL | 3.549 | 3.466 | 3.515 | 3.88 | 3.5187 | 3.665 | 3.98 | 16
| 1 | 4 | 3.4109 | 3.731 | NVL | 3.499 | 3.449 | 3.549 | 3.665 | 3.5121 | 3.615 | 3.88 | 3.5548
| 1 | 5 | 3.4275 | 3.698 | NVL | 3.432 | 3.366 | 3.366 | 3.714 | 3.5253 | 3.731 | 3.814 | 15
| Mean |  | 3.4408 | 3.691 |  | 3.5418 | 3.4026 | 3.4756 | 3.7476 | 3.5154 | 3.625 | 3.9268
| Median |  | 3.4275 | 3.698 |  | 3.549 | 3.399 | 3.499 | 3.731 | 3.5187 | 3.615 | 3.897
| Std Dev. |  | 0.026638 | 0.038334 |  | 0.079144 | 0.055608 | 0.071076 | 0.080245 | 0.0087688 | 0.076541 | 0.096414
| Rel.Std.Dev. |  | 0.7742% | 1.039% |  | 2.235% | 1.634% | 2.045% | 2.141% | 0.2494% | 2.111% | 2.455%
| PDM3 |  | -3.207% | 3.832% |  | -0.3648% | -4.281% | -2.227% | 5.425% | -1.108% | 1.976% | 10.47%
| Z-Score (Absolute) |  | 0.67 | 0.99 | NA | 0 | 0.93 | 0.44 | 1.37 | 0.18 | 0.55 | 2.56
| Analytical results for Mn in OREAS 134a (Indicative Value 3371 ppm) | [11 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | [8 empty] | 1
| No | No | 1.09 | 1.12 | [8 empty] | ppm
|  |  | PF*OES | BF*OES | [8 empty] | 0
0
| 1 | 1 | 3400 | 3284 | [8 empty] | 1
| 1 | 2 | 3600 | 3261 | [8 empty] | 34
| 1 | 3 | 3500 | 3059 | [8 empty] | 16
| 1 | 4 | 3500 | 3477 | [8 empty] | 3371.0409
| 1 | 5 | 3400 | 3230 | [8 empty] | 40
| Mean |  | 3480 | 3262.2
| Median |  | 3500 | 3261
| Std Dev. |  | 83.666 | 149.1164
| Rel.Std.Dev. |  | 2.404% | 4.571%
| PDM3 |  | 3.232% | -3.229%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Na in OREAS 134a (Indicative Value 653 ppm) | [11 empty] | Indicative
Na | Round | Replicate | Lab | [9 empty] | 1
| No | No | 1.12 | [9 empty] | ppm
|  |  | BF*OES | [9 empty] | 0
0
| 1 | 1 | 670 | [9 empty] | 1
| 1 | 2 | 590 | [9 empty] | 29
| 1 | 3 | 740 | [9 empty] | 16
| 1 | 4 | 670 | [9 empty] | 652.819
| 1 | 5 | 590 | [9 empty] | 35
| Mean |  | 652
| Median |  | 670
| Std Dev. |  | 63.4035
| Rel.Std.Dev. |  | 9.724%
| PDM3 |  | -0.1255%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 134a (Indicative Value 297 ppm) | [11 empty] | Indicative
P | Round | Replicate | Lab | Lab | [8 empty] | 1
| No | No | 1.09 | 1.12 | [8 empty] | ppm
|  |  | PF*OES | BF*OES | [8 empty] | 0
0
| 1 | 1 | 400 | 260 | [8 empty] | 1
| 1 | 2 | 400 | 260 | [8 empty] | 30
| 1 | 3 | 300 | 220 | [8 empty] | 16
| 1 | 4 | 300 | 260 | [8 empty] | 296.5491
| 1 | 5 | 300 | 260 | [8 empty] | 36
| Mean |  | 340 | 252
| Median |  | 300 | 260
| Std Dev. |  | 54.7723 | 17.8885
| Rel.Std.Dev. |  | 16.11% | 7.099%
| PDM3 |  | 14.65% | -15.02%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pb in OREAS 134a (Certified Value 12.95 wt.%) | [11 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 11.711 | 13.856 | > 10 | 13.15 | 12.8 | 13.49 | 12.8 | 13.28 | 12.5 | 13.4 | 1
| 1 | 2 | 11.783 | 13.6669 | > 10 | 13 | 12.4 | 13.21 | 12.6 | 13.415 | 12.4 | 14.1 | 22
| 1 | 3 | 11.96 | 13.7065 | > 10 | 12.65 | 12.7 | 12.69 | 13.2 | 13.395 | 13 | 14 | 16
| 1 | 4 | 11.888 | 13.7092 | > 10 | 12.55 | 12.85 | 13.38 | 12.55 | 13.35 | 12.7 | 13.4 | 12.9451
| 1 | 5 | 11.87 | 13.7342 | > 10 | 12.3 | 12.4 | 12.73 | 12.5 | 13.44 | 13.5 | 13.4 | 16
| Mean |  | 11.8424 | 13.7346 |  | 12.73 | 12.63 | 13.1 | 12.73 | 13.376 | 12.82 | 13.66
| Median |  | 11.87 | 13.7092 |  | 12.65 | 12.7 | 13.21 | 12.6 | 13.395 | 12.7 | 13.4
| Std Dev. |  | 0.09678 | 0.072035 |  | 0.34387 | 0.21679 | 0.37 | 0.28636 | 0.062988 | 0.44385 | 0.35777
| Rel.Std.Dev. |  | 0.8172% | 0.5245% |  | 2.701% | 1.717% | 2.824% | 2.249% | 0.4709% | 3.462% | 2.619%
| PDM3 |  | -8.518% | 6.099% |  | -1.661% | -2.434% | 1.197% | -1.661% | 3.329% | -0.966% | 5.523%
| Z-Score (Absolute) |  | 2.35 | 2.2 | NA | 0.22 | 0.46 | 0.67 | 0.22 | 1.34 | 0 | 2.02
| Analytical results for S in OREAS 134a (Indicative Value 19.39 wt.%) | [11 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES |  |  | 2
2
| 1 | 1 | 18.43 | 15.25 | 19.05 | 19.76 | 20.3 | 6.764 | 18.3 | 73.3 |  |  | 1
| 1 | 2 | 19.1 | 14.35 | 18.65 | 20.89 | 20.1 | 7.61 | 18.2 | 39.6 |  |  | 31
| 1 | 3 | 20.01 | 15.15 | 19.3 | 18.46 | 21 | 7.45 | 19.3 | 40.3 |  |  | 16
| 1 | 4 | 20.48 | 14.8 | 19.3 | 18.66 | 19.75 | 7.688 | 18.8 | 39.3 |  |  | 19.3944
| 1 | 5 | 20.12 | 13.8 | 18.8 | 19.1 | 20.4 | 8.416 | 18.6 | 40.7 |  |  | 37
| Mean |  | 19.628 | 14.67 | 19.02 | 19.374 | 20.31 | 7.5856 | 18.64 | 46.64
| Median |  | 20.01 | 14.8 | 19.05 | 19.1 | 20.3 | 7.61 | 18.6 | 40.3
| Std Dev. |  | 0.84082 | 0.60062 | 0.29283 | 0.9833 | 0.4588 | 0.59035 | 0.43932 | 14.9137
| Rel.Std.Dev. |  | 4.284% | 4.094% | 1.54% | 5.075% | 2.259% | 7.783% | 2.357% | 31.98%
| PDM3 |  | 1.204% | -24.36% | -1.93% | -0.1052% | 4.721% | -60.89% | -3.89% | 140.5%
| Z-Score (Absolute) |  | 0.35 | 3.66 | 0.14 | 0.14 | 0.9 | 9.38 | 0.45 | 22.16
| Analytical results for Sb in OREAS 134a (Indicative Value 118 ppm) | [11 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.09 | 1.10 | 1.12 | 1.16 |  |  |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES |  |  |  | 0
0
| 1 | 1 | 92 | 120.1 | 121 | < 100 | 112 | 140 | < 400 |  |  |  | 1
| 1 | 2 | 92 | 120.3 | 117 | < 100 | 114 | 150 | < 400 |  |  |  | 26
| 1 | 3 | 94 | 118.1 | 119 | < 100 | 118 | 150 | < 400 |  |  |  | 16
| 1 | 4 | 96 | 118.7 | 120 | < 100 | 115 | 150 | < 400 |  |  |  | 118.1667
| 1 | 5 | 93 | 121.3 | 120 | < 100 | 118 | 150 | < 400 |  |  |  | 38
| Mean |  | 93.4 | 119.7 | 119.4 |  | 115.4 | 148
| Median |  | 93 | 120.1 | 120 |  | 115 | 150
| Std Dev. |  | 1.6733 | 1.2884 | 1.5166 |  | 2.6077 | 4.4721
| Rel.Std.Dev. |  | 1.792% | 1.076% | 1.27% |  | 2.26% | 3.022%
| PDM3 |  | -20.96% | 1.298% | 1.044% |  | -2.341% | 25.25%
| Z-Score (Absolute) |  | 0.67 | 0.01 | 0 | 1.8 | 0.1 | 0.74 | 2.09
| Analytical results for Sc in OREAS 134a (Indicative Value 2 ppm) | [11 empty] | Indicative
Sc | Round | Replicate | 1.04: Lab |  |  | 1
| No | No | 1.04: 1.12 |  |  | ppm
|  |  | 1.04: BF*OES |  |  | 2
2
| 1 | 1 | 1.04: 2 |  |  | 1
| 1 | 2 | 1.04: 2 |  |  | 27
| 1 | 3 | 1.04: 2 |  |  | 16
| 1 | 4 | 1.04: 2 |  |  | 2
| 1 | 5 | 1.04: 2 |  |  | 39
| Mean |  | 1.04: 2
| Median |  | 1.04: 2
| Std Dev. |  | 1.04: 0
Rel.Std.Dev. | bar: 1.04
PDM3 | bar: 1.04
Z-Score (Absolute) | bar: 1.04
| Analytical results for SiO2 in OREAS 134a (Certified Value 14.81 wt.%) | [11 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.16 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  | 2
3
| 1 | 1 | 13.788 | 16.26 | ****** | 15.573 | 14.718 | 14.7 | 14.523 | 13.76 | 15.745 |  | 1
| 1 | 2 | 13.86 | 16.04 | ****** | 15.445 | 14.375 | 14.7 | 14.5615 | 13.64 | 16.279 |  | 8
| 1 | 3 | 14.172 | 16.26 | ****** | 15.295 | 14.932 | 15.2 | 14.5936 | 12.8 | 16.472 |  | 16
| 1 | 4 | 13.908 | 16.04 | ****** | 15.017 | 14.867 | 14.4 | 14.5401 | 14.86 | 15.702 |  | 14.8131
| 1 | 5 | 13.438 | 16.26 | ****** | 14.76 | 14.525 | 14.5 | 14.5979 | 13.53 | 14.162 |  | 17
| Mean |  | 13.8332 | 16.172 |  | 15.218 | 14.6834 | 14.7 | 14.5632 | 13.718 | 15.672
| Median |  | 13.86 | 16.26 |  | 15.295 | 14.718 | 14.7 | 14.5615 | 13.64 | 15.745
| Std Dev. |  | 0.26426 | 0.1205 |  | 0.32926 | 0.23287 | 0.30822 | 0.032714 | 0.73995 | 0.90758
| Rel.Std.Dev. |  | 1.91% | 0.7451% |  | 2.164% | 1.586% | 2.097% | 0.2246% | 5.394% | 5.791%
| PDM3 |  | -6.615% | 9.174% |  | 2.734% | -0.8753% | -0.7632% | -1.687% | -7.393% | 5.799%
| Z-Score (Absolute) |  | 0.84 | 1.44 | NA | 0.51 | 0.01 | 0.01 | 0.13 | 0.95 | 0.95
| Analytical results for Sr in OREAS 134a (Indicative Value 28.8 ppm) | [11 empty] | Indicative
Sr | Round | Replicate | 1.04: Lab; 1.07: Lab |  |  | 1
| No | No | 1.04: 1.09; 1.07: 1.12 |  |  | ppm
|  |  | 1.04: PF*OES; 1.07: BF*OES |  |  | 1
1
| 1 | 1 | 1.04: 100; 1.07: 27 |  |  | 1
| 1 | 2 | 1.04: 100; 1.07: 28 |  |  | 34
| 1 | 3 | 1.04: 100; 1.07: 26 |  |  | 16
| 1 | 4 | 1.04: 100; 1.07: 35 |  |  | 28.8
| 1 | 5 | 1.04: 100; 1.07: 28 |  |  | 40
| Mean |  | 1.04: 100; 1.07: 28.8
| Median |  | 1.04: 100; 1.07: 28
| Std Dev. |  | 1.04: 0; 1.07: 3.5637
| Rel.Std.Dev. |  | 1.04: 0%; 1.07: 12.37%
| PDM3 |  | 1.04: 247.2%; 1.07: 0%
| Z-Score (Absolute) |  | 1.04: 0.67; 1.07: 0.67
| Analytical results for Ti in OREAS 134a (Indicative Value 613 ppm) | [11 empty] | Indicative
Ti | Round | Replicate | 1.04: Lab |  |  | 1
| No | No | 1.04: 1.12 |  |  | ppm
|  |  | 1.04: BF*OES |  |  | 0
0
| 1 | 1 | 1.04: 629 |  |  | 1
| 1 | 2 | 1.04: 593 |  |  | 29
| 1 | 3 | 1.04: 599 |  |  | 16
| 1 | 4 | 1.04: 623 |  |  | 612.6731
| 1 | 5 | 1.04: 617 |  |  | 35
| Mean |  | 1.04: 612.2
| Median |  | 1.04: 617
| Std Dev. |  | 1.04: 15.5306
| Rel.Std.Dev. |  | 1.04: 2.537%
| PDM3 |  | 1.04: -0.07722%
Z-Score (Absolute) | bar: 1.04
| Analytical results for V in OREAS 134a (Indicative Value 16.4 ppm) | [11 empty] | Indicative
V | Round | Replicate | 1.04: Lab |  |  | 1
| No | No | 1.04: 1.12 |  |  | ppm
|  |  | 1.04: BF*OES |  |  | 1
1
| 1 | 1 | 1.04: 16 |  |  | 1
| 1 | 2 | 1.04: 14 |  |  | 23
| 1 | 3 | 1.04: 22 |  |  | 16
| 1 | 4 | 1.04: 16 |  |  | 16.4
| 1 | 5 | 1.04: 14 |  |  | 36
| Mean |  | 1.04: 16.4
| Median |  | 1.04: 16
| Std Dev. |  | 1.04: 3.2863
| Rel.Std.Dev. |  | 1.04: 20.04%
PDM3 | bar: 1.04
Z-Score (Absolute) | bar: 1.04
| Analytical results for Y in OREAS 134a (Indicative Value 5.4 ppm) | [11 empty] | Indicative
Y | Round | Replicate | 1.04: Lab |  |  | 1
| No | No | 1.04: 1.12 |  |  | ppm
|  |  | 1.04: BF*OES |  |  | 2
2
| 1 | 1 | 1.04: 5 |  |  | 1
| 1 | 2 | 1.04: 5 |  |  | 25
| 1 | 3 | 1.04: 6 |  |  | 16
| 1 | 4 | 1.04: 5 |  |  | 5.4
| 1 | 5 | 1.04: 6 |  |  | 37
| Mean |  | 1.04: 5.4
| Median |  | 1.04: 5
| Std Dev. |  | 1.04: 0.54772
| Rel.Std.Dev. |  | 1.04: 10.14%
PDM3 | bar: 1.04
Z-Score (Absolute) | bar: 1.04
| Analytical results for Zn in OREAS 134a (Certified Value 17.49 wt.%) | [11 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 15.629 | 17.9631 | > 10 | 18 | 17.7 | 17.87 | 17.1 | 17.99 | 16.6 | 18 | 1
| 1 | 2 | 15.438 | 17.8522 | > 10 | 17.05 | 17.05 | 17.68 | 16.8 | 18.132 | 16.7 | 18.6 | 0
| 1 | 3 | 15.647 | 17.8473 | > 10 | 17.2 | 17.2 | 17.25 | 17.7 | 18.122 | 17 | 18 | 16
| 1 | 4 | 15.398 | 17.8586 | > 10 | 17.45 | 17.75 | 17.87 | 17.05 | 18.078 | 16.9 | 17.5 | 17.4916
| 1 | 5 | 15.773 | 17.8041 | > 10 | 16.75 | 17.1 | 17.55 | 16.9 | 18.136 | 17.9 | 17.5 | 18
| Mean |  | 15.577 | 17.8651 |  | 17.29 | 17.36 | 17.644 | 17.11 | 18.0916 | 17.02 | 17.92
| Median |  | 15.629 | 17.8522 |  | 17.2 | 17.2 | 17.68 | 17.05 | 18.122 | 16.9 | 18
| Std Dev. |  | 0.15603 | 0.058844 |  | 0.4709 | 0.33801 | 0.25861 | 0.35071 | 0.061309 | 0.51672 | 0.45497
| Rel.Std.Dev. |  | 1.002% | 0.3294% |  | 2.724% | 1.947% | 1.466% | 2.05% | 0.3389% | 3.036% | 2.539%
| PDM3 |  | -10.95% | 2.135% |  | -1.153% | -0.7526% | 0.871% | -2.182% | 3.43% | -2.696% | 2.449%
| Z-Score (Absolute) |  | 3.54 | 1 | NA | 0.14 | 0 | 0.56 | 0.5 | 1.45 | 0.67 | 1.11
| Analytical results for Zr in OREAS 134a (Indicative Value 30.2 ppm) | [11 empty] | Indicative
Zr | Round | Replicate | 1.04: Lab |  |  | 1
| No | No | 1.04: 1.12 |  |  | ppm
|  |  | 1.04: BF*OES |  |  | 1
1
| 1 | 1 | 1.04: 30 |  |  | 1
| 1 | 2 | 1.04: 30 |  |  | 27
| 1 | 3 | 1.04: 29 |  |  | 16
| 1 | 4 | 1.04: 31 |  |  | 30.2
| 1 | 5 | 1.04: 31 |  |  | 38
| Mean |  | 1.04: 30.2
| Median |  | 1.04: 30
| Std Dev. |  | 1.04: 0.83666
| Rel.Std.Dev. |  | 1.04: 2.77%
PDM3 | bar: 1.04
Z-Score (Absolute) | bar: 1.04
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C33: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C50: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C84: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C322: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C339: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C356: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C373: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C390: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C407: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C424: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 134a (Certified Value 201 ppm) | [16 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 215.8 | 149 | 200 | 191.1 | 203 | 201 | 196 | 200.3 | 196 | 218 | > 100 | 172 | 198 | 188 |  | 1
| 1 | 2 | 212.5 | 151 | 205 | 192.1 | 204 | 200 | 192 | 203.4 | 194 | 212 | > 100 | 178 | 201 | 190 |  | 12
| 1 | 3 | 210.4 | 144 | 201 | 193.1 | 203 | 202 | 201 | 201.1 | 195 | 211 | > 100 | 169 | 203 | 193 |  | 16
| 1 | 4 | 210.8 | 137 | 199 | 195.5 | 204 | 200 | 197 | 203.5 | 197 | 211 | > 100 | 176 | 204 | 191 |  | 200.8218
| 1 | 5 | 210.3 | 128 | 200 | 195.6 | 199 | 200 | 197 | 203.5 | 197 | 213 | > 100 | 169 | 205 | 190 |  | 20
| Mean |  | 211.96 | 141.8 | 201 | 193.48 | 202.6 | 200.6 | 196.6 | 202.36 | 195.8 | 213 |  | 172.8 | 202.2 | 190.4
| Median |  | 210.8 | 144 | 200 | 193.1 | 203 | 200 | 197 | 203.4 | 196 | 212 |  | 172 | 203 | 190
| Std Dev. |  | 2.3223 | 9.4181 | 2.3452 | 2.0179 | 2.0736 | 0.89443 | 3.2094 | 1.5421 | 1.3038 | 2.9155 |  | 4.0866 | 2.7749 | 1.8166
| Rel.Std.Dev. |  | 1.096% | 6.642% | 1.167% | 1.043% | 1.024% | 0.4459% | 1.632% | 0.762% | 0.6659% | 1.369% |  | 2.365% | 1.372% | 0.9541%
| PDM3 |  | 5.546% | -29.39% | 0.08873% | -3.656% | 0.8855% | -0.1105% | -2.102% | 0.7659% | -2.501% | 6.064% |  | -13.95% | 0.6863% | -5.19%
| Z-Score (Absolute) |  | 1.6 | 8.26 | 0.06 | 1 | 0.28 | 0 | 0.56 | 0.25 | 0.67 | 1.74 | NA | 3.91 | 0.22 | 1.43
| Analytical results for Al2O3 in OREAS 134a (Certified Value 2.53 wt.%) | [16 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.16 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  |  |  | 2
3
| 1 | 1 | 2.635 | 2.4 | 2.4615 | 2.26 | 2.66 | 2.66 | 2.55 | 2.532 | 2.57 | 2.475 | 2.6 | 2.475 |  |  |  | 1
| 1 | 2 | 2.675 | 2.4 | 2.4548 | 2.46 | 2.44 | 2.65 | 2.4 | 2.608 | 2.53 | 2.645 | 2.62 | 2.419 |  |  |  | 1
| 1 | 3 | 2.686 | 2.343 | 2.5153 | 2.3 | 2.66 | 2.82 | 2.48 | 2.589 | 2.65 | 2.589 | 2.58 | 2.532 |  |  |  | 16
| 1 | 4 | 2.667 | 2.419 | 2.4826 | 2.25 | 2.63 | 2.66 | 2.36 | 2.589 | 2.61 | 2.532 | 2.53 | 2.494 |  |  |  | 2.5326
| 1 | 5 | 2.635 | 2.078 | 2.4633 | 2.32 | 2.68 | 2.67 | 2.44 | 2.608 | 2.65 | 2.608 | 2.54 | 2.494 |  |  |  | 21
| Mean |  | 2.6596 | 2.328 | 2.4755 | 2.318 | 2.614 | 2.692 | 2.446 | 2.5852 | 2.602 | 2.5698 | 2.574 | 2.4828
| Median |  | 2.667 | 2.4 | 2.4633 | 2.3 | 2.66 | 2.66 | 2.44 | 2.589 | 2.61 | 2.589 | 2.58 | 2.494
| Std Dev. |  | 0.023448 | 0.14263 | 0.024529 | 0.08438 | 0.098894 | 0.071903 | 0.073348 | 0.03122 | 0.052154 | 0.066893 | 0.038471 | 0.04124
| Rel.Std.Dev. |  | 0.8816% | 6.127% | 0.9909% | 3.64% | 3.783% | 2.671% | 2.999% | 1.208% | 2.004% | 2.603% | 1.495% | 1.661%
| PDM3 |  | 5.016% | -8.077% | -2.253% | -8.472% | 3.215% | 6.295% | -3.418% | 2.078% | 2.742% | 1.47% | 1.636% | -1.965%
| Z-Score (Absolute) |  | 0.67 | 1.86 | 0.73 | 1.94 | 0.32 | 0.92 | 0.96 | 0.1 | 0.23 | 0.02 | 0.02 | 0.68
| Analytical results for As in OREAS 134a (Certified Value 228 ppm) | [16 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 198 | 236 | 243 | 227 | 250 | 229 | 241 | 270 | 223 | 247 | 219.4 | 221 | 280 | 190 | 148 | 1
| 1 | 2 | 200 | 207 | 242 | 227 | 255 | 208 | 238 | 230 | 230 | 245 | 215.6 | 214 | 310 | 180 | 171 | 2
| 1 | 3 | 199 | 223 | 239 | 226 | 247 | 226 | 255 | 250 | 224 | 224 | 211.5 | 220 | 230 | 180 | 97 | 16
| 1 | 4 | 201 | 216 | 245 | 228 | 252 | 228 | 241 | 240 | 228 | 254 | 211.6 | 215 | 240 | 190 | 187 | 227.7698
| 1 | 5 | 198 | 203 | 238 | 225 | 247 | 247 | 239 | 190 | 231 | 238 | 214.2 | 221 | 200 | 190 | 96 | 22
| Mean |  | 199.2 | 217 | 241.4 | 226.6 | 250.2 | 227.6 | 242.8 | 236 | 227.2 | 241.6 | 214.46 | 218.2 | 252 | 186 | 139.8
| Median |  | 199 | 216 | 242 | 227 | 250 | 228 | 241 | 240 | 228 | 245 | 214.2 | 220 | 240 | 190 | 148
| Std Dev. |  | 1.3038 | 13.1719 | 2.881 | 1.1402 | 3.4205 | 13.8311 | 6.9426 | 29.6648 | 3.5637 | 11.371 | 3.2677 | 3.4205 | 43.2435 | 5.4772 | 41.8891
| Rel.Std.Dev. |  | 0.6545% | 6.07% | 1.193% | 0.5032% | 1.367% | 6.077% | 2.859% | 12.57% | 1.569% | 4.707% | 1.524% | 1.568% | 17.16% | 2.945% | 29.96%
| PDM3 |  | -12.54% | -4.728% | 5.984% | -0.5136% | 9.848% | -0.07453% | 6.599% | 3.613% | -0.2501% | 6.072% | -5.844% | -4.202% | 10.64% | -18.34% | -38.62%
| Z-Score (Absolute) |  | 1.33 | 0.48 | 0.67 | 0.03 | 1.09 | 0.02 | 0.74 | 0.42 | 0 | 0.68 | 0.6 | 0.43 | 1.18 | 1.96 | 4.15
| Analytical results for B in OREAS 134a (Indicative Value 21 ppm) | [16 empty] | Indicative
B | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 1
1
| 1 | 1 | [15 empty] | 1 | bar: 1.01
| 1 | 2 | 1.01: 11 | 36
| 1 | 3 | 1.01: 22 | 16
| 1 | 4 | 1.01: 32 | 21
| 1 | 5 | 1.01: 39 | 42
| Mean |  | 1.01: 21
| Median |  | 1.01: 22
| Std Dev. |  | 1.01: 15.3786
| Rel.Std.Dev. |  | 1.01: 73.23%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ba in OREAS 134a (Indicative Value < 1400 ppm) | [16 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
0
| 1 | 1 | 990 | NA | < 5 | 430 | 110 | 940 | 94 | 1320 | 18 | 330 | 1350 | 290 |  |  |  | 1
| 1 | 2 | 999 | NA | 6 | 350 | 100 | 860 | 84 | 1170 | 9 | 380 | 1200 | 119 |  |  |  | 14
| 1 | 3 | 1011 | NA | 7 | 540 | 90 | 1030 | 102 | 1335 | 9 | 370 | 1200 | 280 |  |  |  | 16
| 1 | 4 | 1071 | NA | 9 | 430 | 130 | 1160 | 93 | 1330 | 19 | 570 | 1110 | 181 |  |  |  | < 1400
| 1 | 5 | 1042 | NA | 5 | 480 | 140 | 1060 | 192 | 1440 | 8 | 360 | 1260 | 188 |  |  |  | 43
| Mean |  | 1022.6 |  | 6.75 | 446 | 114 | 1010 | 113 | 1319 | 12.6 | 402 | 1224 | 211.6
| Median |  | 1011 |  | 6.5 | 430 | 110 | 1030 | 94 | 1330 | 9 | 370 | 1200 | 188
| Std Dev. |  | 33.441 |  | 1.7078 | 70.214 | 20.7364 | 114.8913 | 44.6206 | 96.4624 | 5.4129 | 95.7601 | 88.4873 | 72.2724
| Rel.Std.Dev. |  | 3.27% |  | 25.3% | 15.74% | 18.19% | 11.38% | 39.49% | 7.313% | 42.96% | 23.82% | 7.229% | 34.16%
PDM3
| Z-Score (Absolute) |  | 1.07 | NA | 0.69 | 0.08 | 0.5 | 1.05 | 0.5 | 1.59 | 0.67 | 0 | 1.42 | 0.33
| Analytical results for Be in OREAS 134a (Indicative Value 2 ppm) | [16 empty] | Indicative
Be | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 2
2
| 1 | 1 | 1.01: < 5; 1.02: 2 | 1
| 1 | 2 | 1.01: < 5; 1.02: 2 | 15
| 1 | 3 | 1.01: < 5; 1.02: 2 | 16
| 1 | 4 | 1.01: < 5; 1.02: 2 | 2
| 1 | 5 | 1.01: < 5; 1.02: 2 | 44
| Mean |  | 1.02: 2
| Median |  | 1.02: 2
| Std Dev. |  | 1.02: 0
Rel.Std.Dev. | bar: 1.02
PDM3 | bar: 1.02
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for CaO in OREAS 134a (Certified Value 6.04 wt.%) | [16 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.16 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  |  |  | 2
3
| 1 | 1 | 6.293 | 6.828 | 5.3361 | 5.33 | 5.77 | 6.34 | 6.32 | 6.128 | 6.32 | 5.975 | 5.74 | 6.142 |  |  |  | 1
| 1 | 2 | 6.374 | 6.576 | 5.6879 | 5.68 | 5.36 | 6.37 | 5.93 | 6.17 | 6.27 | 6.282 | 5.82 | 5.975 |  |  |  | 4
| 1 | 3 | 6.303 | 6.73 | 5.8856 | 5.32 | 5.8 | 6.74 | 5.8 | 6.184 | 6.53 | 6.156 | 5.79 | 6.087 |  |  |  | 16
| 1 | 4 | 6.344 | 6.8 | 5.7029 | 5.1 | 5.71 | 6.34 | 5.57 | 6.031 | 6.44 | 6.073 | 5.6 | 6.073 |  |  |  | 6.0361
| 1 | 5 | 6.371 | 6.59 | 5.6661 | 5.4 | 5.99 | 6.38 | 5.75 | 6.128 | 6.45 | 6.212 | 5.64 | 6.184 |  |  |  | 23
| Mean |  | 6.337 | 6.7048 | 5.6557 | 5.366 | 5.726 | 6.434 | 5.874 | 6.1282 | 6.402 | 6.1396 | 5.718 | 6.0922
| Median |  | 6.344 | 6.73 | 5.6879 | 5.33 | 5.77 | 6.37 | 5.8 | 6.128 | 6.44 | 6.156 | 5.74 | 6.087
| Std Dev. |  | 0.037636 | 0.11688 | 0.19896 | 0.20852 | 0.22985 | 0.17199 | 0.28077 | 0.05981 | 0.10521 | 0.11971 | 0.094974 | 0.079118
| Rel.Std.Dev. |  | 0.5939% | 1.743% | 3.518% | 3.886% | 4.014% | 2.673% | 4.78% | 0.976% | 1.643% | 1.95% | 1.661% | 1.299%
| PDM3 |  | 4.985% | 11.08% | -6.302% | -11.1% | -5.137% | 6.592% | -2.685% | 1.526% | 6.062% | 1.715% | -5.27% | 0.9299%
| Z-Score (Absolute) |  | 0.5 | 1.3 | 1 | 1.63 | 0.84 | 0.71 | 0.52 | 0.04 | 0.64 | 0.06 | 0.86 | 0.04
| Analytical results for Cd in OREAS 134a (Certified Value 536 ppm) | [16 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 546 | 606 | 548 | 573.7 | 537 | 561 | 558 | 648.6 | 572 | 505.4 | 462 | 493 | 482 | 525 |  | 1
| 1 | 2 | 555 | 546 | 537 | 578.2 | 502 | 564 | 549 | 686.2 | 564 | 502.1 | 485 | 513 | 479 | 536 |  | 17
| 1 | 3 | 542 | 585 | 543 | 582.7 | 542 | 589 | 568 | 673.4 | 567 | 492.8 | 470 | 493 | 483 | 545 |  | 16
| 1 | 4 | 549 | 579 | 545 | 595.9 | 535 | 563 | 561 | 664.5 | 567 | 486.1 | 470 | 516 | 487 | 538 |  | 535.6146
| 1 | 5 | 557 | 537 | 538 | 591.2 | 530 | 567 | 542 | 668.2 | 552 | 492.6 | 479 | 505 | 494 | 532 |  | 24
| Mean |  | 549.8 | 570.6 | 542.2 | 584.34 | 529.2 | 568.8 | 555.6 | 668.18 | 564.4 | 495.8 | 473.2 | 504 | 485 | 535.2
| Median |  | 549 | 579 | 543 | 582.7 | 535 | 564 | 558 | 668.2 | 567 | 492.8 | 470 | 505 | 483 | 536
| Std Dev. |  | 6.2209 | 28.571 | 4.6583 | 9.1418 | 15.8019 | 11.4978 | 10.2127 | 13.6829 | 7.5033 | 7.8291 | 8.9275 | 10.8167 | 5.7879 | 7.3959
| Rel.Std.Dev. |  | 1.131% | 5.007% | 0.8592% | 1.564% | 2.986% | 2.021% | 1.838% | 2.048% | 1.329% | 1.579% | 1.887% | 2.146% | 1.193% | 1.382%
| PDM3 |  | 2.648% | 6.532% | 1.23% | 9.097% | -1.198% | 6.196% | 3.731% | 24.75% | 5.374% | -7.433% | -11.65% | -5.902% | -9.45% | -0.07741%
| Z-Score (Absolute) |  | 0.11 | 0.7 | 0.11 | 1.09 | 0.48 | 0.65 | 0.27 | 3.48 | 0.52 | 1.43 | 2.07 | 1.19 | 1.74 | 0.31
| Analytical results for Ce in OREAS 134a (Indicative Value 13.2 ppm) | [16 empty] | Indicative
Ce | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 14; 1.02: 13.1 |  | 1
| 1 | 2 | 1.01: 16; 1.02: 12 |  | 18
| 1 | 3 | 1.01: 14; 1.02: 12.7 |  | 16
| 1 | 4 | 1.01: 15; 1.02: 10.5 |  | 13.15
| 1 | 5 | 1.01: 13; 1.02: 11.2 |  | 45
| Mean |  | 1.01: 14.4; 1.02: 11.9
| Median |  | 1.01: 14; 1.02: 12
| Std Dev. |  | 1.01: 1.1402; 1.02: 1.0654
| Rel.Std.Dev. |  | 1.01: 7.918%; 1.02: 8.953%
| PDM3 |  | 1.01: 9.506%; 1.02: -9.506%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Co in OREAS 134a (Certified Value 105 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 110 | 102 | 108 | 106.7 | 100 | 103 | 100 | 113 | 115 | 84.21 | 108 | 90 | 110 | 87 |  | 1
| 1 | 2 | 108 | 91 | 112 | 103.8 | 94 | 102 | 100 | 123 | 110 | 81.58 | 113 | 100 | 110 | 84 |  | 19
| 1 | 3 | 112 | 97.7 | 103 | 106.6 | 101 | 106 | 100 | 114 | 112 | 80.98 | 106 | 100 | 110 | 98 |  | 16
| 1 | 4 | 113 | 96.7 | 106 | 106.5 | 102 | 102 | 100 | 119 | 115 | 81.01 | 108 | 100 | 110 | 92 |  | 104.9669
| 1 | 5 | 110 | 89.7 | 116 | 105.4 | 104 | 102 | 100 | 118 | 110 | 80.93 | 111 | 100 | 120 | 89 |  | 25
| Mean |  | 110.6 | 95.42 | 109 | 105.8 | 100.2 | 103 | 100 | 117.4 | 112.4 | 81.742 | 109.2 | 98 | 112 | 90
| Median |  | 110 | 96.7 | 108 | 106.5 | 101 | 102 | 100 | 118 | 112 | 81.01 | 108 | 100 | 110 | 89
| Std Dev. |  | 1.9494 | 5.0593 | 5.099 | 1.2349 | 3.7683 | 1.7321 | 0 | 4.0373 | 2.51 | 1.4047 | 2.7749 | 4.4721 | 4.4721 | 5.3385
| Rel.Std.Dev. |  | 1.763% | 5.302% | 4.678% | 1.167% | 3.761% | 1.682% | 0% | 3.439% | 2.233% | 1.718% | 2.541% | 4.563% | 3.993% | 5.932%
| PDM3 |  | 5.367% | -9.095% | 3.842% | 0.7937% | -4.541% | -1.874% | -4.732% | 11.84% | 7.081% | -22.13% | 4.033% | -6.637% | 6.7% | -14.26%
| Z-Score (Absolute) |  | 0.66 | 0.96 | 0.49 | 0.15 | 0.45 | 0.15 | 0.47 | 1.39 | 0.86 | 2.43 | 0.51 | 0.69 | 0.81 | 1.54
| Analytical results for Cr in OREAS 134a (Indicative Value 16.6 ppm) | [16 empty] | Indicative
Cr | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 14.6; 1.02: 18.5 |  | 1
| 1 | 2 | 1.01: 16.8; 1.02: 16.6 |  | 20
| 1 | 3 | 1.01: 15.2; 1.02: 15.4 |  | 16
| 1 | 4 | 1.01: 16.4; 1.02: 16.9 |  | 16.57
| 1 | 5 | 1.01: 18; 1.02: 17.3 |  | 46
| Mean |  | 1.01: 16.2; 1.02: 16.94
| Median |  | 1.01: 16.4; 1.02: 16.9
| Std Dev. |  | 1.01: 1.3416; 1.02: 1.1238
| Rel.Std.Dev. |  | 1.01: 8.282%; 1.02: 6.634%
| PDM3 |  | 1.01: -2.233%; 1.02: 2.233%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Cs in OREAS 134a (Indicative Value 1.14 ppm) | [16 empty] | Indicative
Cs | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 1.2 |  | 1
| 1 | 2 | 1.01: 1.2 |  | 21
| 1 | 3 | 1.01: 1.2 |  | 16
| 1 | 4 | [15 empty] | 1.14 | bar: 1.01
| 1 | 5 | 1.01: 1.1 |  | 47
| Mean |  | 1.01: 1.14
| Median |  | 1.01: 1.2
| Std Dev. |  | 1.01: 0.089443
| Rel.Std.Dev. |  | 1.01: 7.846%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Cu in OREAS 134a (Certified Value 1291 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 1283 | 1380 | 1240 | 1296 | 1320 | 1360 | 1310 | 1284.9 | 1270 | 1261 | 1270 | 1190 | 1240 | 1230 |  | 1
| 1 | 2 | 1282 | 1420 | 1250 | 1293 | 1220 | 1370 | 1300 | 1316.3 | 1280 | 1243 | 1340 | 1280 | 1240 | 1210 |  | 22
| 1 | 3 | 1279 | 1340 | 1220 | 1328 | 1330 | 1450 | 1350 | 1306 | 1325 | 1242 | 1310 | 1230 | 1260 | 1260 |  | 16
| 1 | 4 | 1283 | 1260 | 1270 | 1302 | 1320 | 1380 | 1330 | 1305.8 | 1290 | 1249 | 1280 | 1280 | 1270 | 1260 |  | 1290.6657
| 1 | 5 | 1280 | 1180 | 1310 | 1293 | 1340 | 1380 | 1280 | 1317.6 | 1280 | 1239 | 1320 | 1270 | 1280 | 1260 |  | 26
| Mean |  | 1281.4 | 1316 | 1258 | 1302.4 | 1306 | 1388 | 1314 | 1306.12 | 1289 | 1246.8 | 1304 | 1250 | 1258 | 1244
| Median |  | 1282 | 1340 | 1250 | 1296 | 1320 | 1380 | 1310 | 1306 | 1280 | 1243 | 1310 | 1270 | 1260 | 1260
| Std Dev. |  | 1.8166 | 96.3328 | 34.2053 | 14.775 | 48.7852 | 35.6371 | 27.0185 | 13.0942 | 21.3307 | 8.7293 | 28.8097 | 39.37 | 17.8885 | 23.0217
| Rel.Std.Dev. |  | 0.1418% | 7.32% | 2.719% | 1.134% | 3.735% | 2.568% | 2.056% | 1.003% | 1.655% | 0.7001% | 2.209% | 3.15% | 1.422% | 1.851%
| PDM3 |  | -0.7179% | 1.963% | -2.531% | 0.9092% | 1.188% | 7.541% | 1.808% | 1.197% | -0.1291% | -3.399% | 1.033% | -3.151% | -2.531% | -3.616%
| Z-Score (Absolute) |  | 0.5 | 0.71 | 1.32 | 0.23 | 0.36 | 3.22 | 0.64 | 0.36 | 0.23 | 1.71 | 0.29 | 1.6 | 1.32 | 1.81
| Analytical results for Dy in OREAS 134a (Indicative Value 0.8 ppm) | [16 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.8 |  | 1
| 1 | 2 | 1.01: 0.8 |  | 23
| 1 | 3 | 1.01: 0.9 |  | 16
| 1 | 4 | 1.01: 0.7 |  | 0.8
| 1 | 5 | 1.01: 0.8 |  | 48
| Mean |  | 1.01: 0.8
| Median |  | 1.01: 0.8
| Std Dev. |  | 1.01: 0.070711
| Rel.Std.Dev. |  | 1.01: 8.839%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Er in OREAS 134a (Indicative Value 0.56 ppm) | [16 empty] | Indicative
Er | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.6 |  | 1
| 1 | 2 | 1.01: 0.6 |  | 24
| 1 | 3 | 1.01: 0.6 |  | 16
| 1 | 4 | 1.01: 0.5 |  | 0.56
| 1 | 5 | 1.01: 0.5 |  | 49
| Mean |  | 1.01: 0.56
| Median |  | 1.01: 0.6
| Std Dev. |  | 1.01: 0.054772
| Rel.Std.Dev. |  | 1.01: 9.781%
| PDM3 |  | 1.01: -2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Eu in OREAS 134a (Indicative Value 300 ppb) | [16 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppb
|  |  | 1.01: 4A*MS |  | 0
0
| 1 | 1 | 1.01: 300 |  | 1
| 1 | 2 | 1.01: 300 |  | 25
| 1 | 3 | 1.01: 300 |  | 16
| 1 | 4 | 1.01: 300 |  | 300
| 1 | 5 | 1.01: 300 |  | 50
| Mean |  | 1.01: 300
| Median |  | 1.01: 300
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Fe in OREAS 134a (Certified Value 12.04 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 12.004 | 12.7 | 11.8 | 12.32 | 11.1 | 12.05 | 11.8 | 11.8 | 12.04 | 12.8 | 11.988 | 11.6 | 10.5 | 12.8 | 12.1 | 1
| 1 | 2 | 12.223 | 12.9 | 11.9 | 12.8 | 11.9 | 12.1 | 11.65 | 11.6 | 12.42 | 12.6 | 12.002 | 12.5 | 11.7 | 12.8 | 11.8 | 0
| 1 | 3 | 12.058 | 12.4 | 11.4 | 13.22 | 11.1 | 12 | 11.9 | 12.05 | 12.29 | 13.25 | 11.924 | 12.2 | 11.3 | 13 | 12.3 | 16
| 1 | 4 | 12.053 | 11.6 | 11.6 | 12.86 | 10.6 | 12.05 | 11.75 | 11.9 | 12.31 | 13.1 | 11.952 | 11.9 | 11.85 | 13.2 | 12.2 | 12.0394
| 1 | 5 | 12.064 | 10.8 | 11.3 | 12.79 | 11.2 | 11.85 | 11.85 | 12 | 12.35 | 13.1 | 11.81 | 12.3 | 11.6 | 13.3 | 12.2 | 27
| Mean |  | 12.0804 | 12.08 | 11.6 | 12.798 | 11.18 | 12.01 | 11.79 | 11.87 | 12.282 | 12.97 | 11.9352 | 12.1 | 11.39 | 13.02 | 12.12
| Median |  | 12.058 | 12.4 | 11.6 | 12.8 | 11.1 | 12.05 | 11.8 | 11.9 | 12.31 | 13.1 | 11.952 | 12.2 | 11.6 | 13 | 12.2
| Std Dev. |  | 0.083206 | 0.87006 | 0.25495 | 0.32034 | 0.46583 | 0.096177 | 0.096177 | 0.17889 | 0.14412 | 0.26363 | 0.076375 | 0.35355 | 0.53666 | 0.22804 | 0.19235
| Rel.Std.Dev. |  | 0.6888% | 7.202% | 2.198% | 2.503% | 4.167% | 0.8008% | 0.8157% | 1.507% | 1.173% | 2.033% | 0.6399% | 2.922% | 4.712% | 1.751% | 1.587%
| PDM3 |  | 0.3405% | 0.3371% | -3.65% | 6.301% | -7.138% | -0.2443% | -2.072% | -1.407% | 2.015% | 7.73% | -0.8656% | 0.5033% | -5.394% | 8.145% | 0.6694%
| Z-Score (Absolute) |  | 0 | 0 | 1.54 | 2.31 | 2.89 | 0.22 | 0.93 | 0.67 | 0.65 | 2.86 | 0.46 | 0.06 | 2.22 | 3.02 | 0.13
| Analytical results for Ga in OREAS 134a (Indicative Value 4.88 ppm) | [16 empty] | Indicative
Ga | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 2
2
| 1 | 1 | 1.01: 5.4 | 1
| 1 | 2 | 1.01: 5.1 | 26
| 1 | 3 | 1.01: 4.3 | 16
| 1 | 4 | 1.01: 4.6 | 4.88
| 1 | 5 | 1.01: 5 | 51
| Mean |  | 1.01: 4.88
| Median |  | 1.01: 5
| Std Dev. |  | 1.01: 0.43243
| Rel.Std.Dev. |  | 1.01: 8.861%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Gd in OREAS 134a (Indicative Value 1.24 ppm) | [16 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.4 | 1
| 1 | 2 | 1.01: 1.3 | 27
| 1 | 3 | 1.01: 1.2 | 16
| 1 | 4 | 1.01: 1.1 | 1.24
| 1 | 5 | 1.01: 1.2 | 52
| Mean |  | 1.01: 1.24
| Median |  | 1.01: 1.2
| Std Dev. |  | 1.01: 0.11402
| Rel.Std.Dev. |  | 1.01: 9.195%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ge in OREAS 134a (Indicative Value 2080 ppb) | [16 empty] | Indicative
Ge | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 0
0
| 1 | 1 | 1.01: 1100 | 1
| 1 | 2 | 1.01: 3400 | 47
| 1 | 3 | 1.01: 2400 | 16
| 1 | 4 | 1.01: 1200 | 2080
| 1 | 5 | 1.01: 2300 | 53
| Mean |  | 1.01: 2080
| Median |  | 1.01: 2300
| Std Dev. |  | 1.01: 952.3655
| Rel.Std.Dev. |  | 1.01: 45.79%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Hf in OREAS 134a (Indicative Value 880 ppb) | [16 empty] | Indicative
Hf | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppb
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 0
0
| 1 | 1 | 1.01: 800; 1.02: 1100 | 1
| 1 | 2 | 1.01: 1000; 1.02: 600 | 48
| 1 | 3 | 1.01: 1100; 1.02: 700 | 16
| 1 | 4 | 1.01: 1200; 1.02: 600 | 880
| 1 | 5 | 1.01: 1000; 1.02: 700 | 54
| Mean |  | 1.01: 1020; 1.02: 740
| Median |  | 1.01: 1000; 1.02: 700
| Std Dev. |  | 1.01: 148.324; 1.02: 207.3644
| Rel.Std.Dev. |  | 1.01: 14.54%; 1.02: 28.02%
| PDM3 |  | 1.01: 15.91%; 1.02: -15.91%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Hg in OREAS 134a (Indicative Value 2000 ppb) | [16 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*OES | 0
0
| 1 | 1 | 1.01: 3000 | 1
| 1 | 2 | 1.01: 1000 | 49
| 1 | 3 | 1.01: 2000 | 16
| 1 | 4 | 1.01: 2000 | 2000
| 1 | 5 | 1.01: 2000 | 55
| Mean |  | 1.01: 2000
| Median |  | 1.01: 2000
| Std Dev. |  | 1.01: 707.1068
| Rel.Std.Dev. |  | 1.01: 35.36%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ho in OREAS 134a (Indicative Value 200 ppb) | [16 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 0
0
| 1 | 1 | 1.01: 200 | 1
| 1 | 2 | 1.01: 200 | 50
| 1 | 3 | 1.01: 200 | 16
| 1 | 4 | 1.01: 200 | 200
| 1 | 5 | 1.01: 200 | 56
| Mean |  | 1.01: 200
| Median |  | 1.01: 200
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for In in OREAS 134a (Indicative Value 1.16 ppm) | [16 empty] | Indicative
In | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.2 | 1
| 1 | 2 | 1.01: 1.1 | 51
| 1 | 3 | 1.01: 1.2 | 16
| 1 | 4 | [15 empty] | 1.16 | bar: 1.01
| 1 | 5 | 1.01: 1.3 | 57
| Mean |  | 1.01: 1.16
| Median |  | 1.01: 1.2
| Std Dev. |  | 1.01: 0.11402
| Rel.Std.Dev. |  | 1.01: 9.829%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for K in OREAS 134a (Indicative Value 0.795 wt.%) | [16 empty] | Indicative
K | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | wt.%
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 3
3
| 1 | 1 | 1.01: 0.75; 1.02: 0.8 | 1
| 1 | 2 | 1.01: 0.77; 1.02: 0.85 | 32
| 1 | 3 | 1.01: 0.76; 1.02: 0.83 | 16
| 1 | 4 | 1.01: 0.77; 1.02: 0.81 | 0.795
| 1 | 5 | 1.01: 0.77; 1.02: 0.84 | 42
| Mean |  | 1.01: 0.764; 1.02: 0.826
| Median |  | 1.01: 0.77; 1.02: 0.83
| Std Dev. |  | 1.01: 0.0089443; 1.02: 0.020736
| Rel.Std.Dev. |  | 1.01: 1.171%; 1.02: 2.51%
| PDM3 |  | 1.01: -3.899%; 1.02: 3.899%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for La in OREAS 134a (Indicative Value 6.1 ppm) | [16 empty] | Indicative
La | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 7.4; 1.02: 4.3 | 1
| 1 | 2 | 1.01: 8.4; 1.02: 4.5 | 16
| 1 | 3 | 1.01: 7.8; 1.02: 4.7 | 16
| 1 | 4 | 1.01: 7.7; 1.02: 4.2 | 6.1
| 1 | 5 | 1.01: 7.6; 1.02: 4.4 | 43
| Mean |  | 1.01: 7.78; 1.02: 4.42
| Median |  | 1.01: 7.7; 1.02: 4.4
| Std Dev. |  | 1.01: 0.37683; 1.02: 0.19235
| Rel.Std.Dev. |  | 1.01: 4.844%; 1.02: 4.352%
| PDM3 |  | 1.01: 27.54%; 1.02: -27.54%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Li in OREAS 134a (Indicative Value 17.8 ppm) | [16 empty] | Indicative
Li | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 1
1
| 1 | 1 | 1.01: 16.2; 1.02: 19.4 | 1
| 1 | 2 | 1.01: 21.4; 1.02: 16.5 | 38
| 1 | 3 | 1.01: 19.4; 1.02: 15.5 | 16
| 1 | 4 | 1.01: 22.7; 1.02: 15.6 | 17.84
| 1 | 5 | 1.01: 15.9; 1.02: 15.8 | 44
| Mean |  | 1.01: 19.12; 1.02: 16.56
| Median |  | 1.01: 19.4; 1.02: 15.8
| Std Dev. |  | 1.01: 3.0409; 1.02: 1.6349
| Rel.Std.Dev. |  | 1.01: 15.9%; 1.02: 9.873%
| PDM3 |  | 1.01: 7.175%; 1.02: -7.175%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Lu in OREAS 134a (Indicative Value 100 ppb) | [16 empty] | Indicative
Lu | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 0
0
| 1 | 1 | 1.01: 100 | 1
| 1 | 2 | 1.01: 100 | 17
| 1 | 3 | 1.01: 100 | 16
| 1 | 4 | 1.01: 100 | 100
| 1 | 5 | 1.01: 100 | 45
| Mean |  | 1.01: 100
| Median |  | 1.01: 100
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for MgO in OREAS 134a (Certified Value 3.49 wt.%) | [16 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.5369 | 3.6646 | 3.714 | 3.4495 | 2.8853 | 3.532 | 3.598 | 3.366 | 3.532 | 3.598 | 3.5618 | 2.67 | 3.2 | 3.217 | 3.615 | 1
| 1 | 2 | 3.5668 | 3.731 | 3.731 | 3.3738 | 3.1506 | 3.25 | 3.598 | 3.167 | 3.631 | 3.565 | 3.5751 | 2.753 | 3.233 | 3.217 | 3.449 | 3
| 1 | 3 | 3.5369 | 3.5485 | 3.698 | 3.4106 | 2.9848 | 3.532 | 3.83 | 3.233 | 3.598 | 3.714 | 3.522 | 2.703 | 3.233 | 3.267 | 3.681 | 16
| 1 | 4 | 3.5369 | 3.3661 | 3.814 | 3.4695 | 2.7029 | 3.482 | 3.615 | 3.084 | 3.598 | 3.648 | 3.5154 | 2.736 | 3.117 | 3.3 | 3.631 | 3.4938
| 1 | 5 | 3.585 | 3.134 | 3.698 | 3.4565 | 2.9184 | 3.549 | 3.631 | 3.217 | 3.615 | 3.648 | 3.4955 | 2.736 | 3.167 | 3.3 | 3.648 | 28
| Mean |  | 3.5525 | 3.4888 | 3.731 | 3.432 | 2.9284 | 3.469 | 3.6544 | 3.2134 | 3.5948 | 3.6346 | 3.534 | 2.7196 | 3.19 | 3.2602 | 3.6048
| Median |  | 3.5369 | 3.5485 | 3.714 | 3.4495 | 2.9184 | 3.532 | 3.615 | 3.217 | 3.598 | 3.648 | 3.522 | 2.736 | 3.2 | 3.267 | 3.631
| Std Dev. |  | 0.022309 | 0.2419 | 0.048363 | 0.039246 | 0.1623 | 0.12497 | 0.099118 | 0.10318 | 0.037692 | 0.056673 | 0.033295 | 0.033126 | 0.049132 | 0.041674 | 0.090467
| Rel.Std.Dev. |  | 0.628% | 6.934% | 1.296% | 1.144% | 5.542% | 3.602% | 2.712% | 3.211% | 1.049% | 1.559% | 0.9421% | 1.218% | 1.54% | 1.278% | 2.51%
| PDM3 |  | 1.679% | -0.143% | 6.788% | -1.771% | -16.18% | -0.7109% | 4.596% | -8.027% | 2.89% | 4.029% | 1.148% | -22.16% | -8.696% | -6.687% | 3.176%
| Z-Score (Absolute) |  | 0.29 | 0 | 1.12 | 0.26 | 2.59 | 0.09 | 0.77 | 1.27 | 0.49 | 0.67 | 0.21 | 3.56 | 1.38 | 1.06 | 0.54
| Analytical results for Mn in OREAS 134a (Indicative Value 3424 ppm) | [16 empty] | Indicative
Mn | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 0
0
| 1 | 1 | 1.01: 3532; 1.02: 3190 | 1
| 1 | 2 | 1.01: 3568; 1.02: 3400 | 34
| 1 | 3 | 1.01: 3589; 1.02: 3300 | 16
| 1 | 4 | 1.01: 3530; 1.02: 3260 | 3424.1
| 1 | 5 | 1.01: 3592; 1.02: 3280 | 46
| Mean |  | 1.01: 3562.2; 1.02: 3286
| Median |  | 1.01: 3568; 1.02: 3280
| Std Dev. |  | 1.01: 29.9533; 1.02: 76.0263
| Rel.Std.Dev. |  | 1.01: 0.8409%; 1.02: 2.314%
| PDM3 |  | 1.01: 4.033%; 1.02: -4.033%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Mo in OREAS 134a (Indicative Value 6.25 ppm) | [16 empty] | Indicative
Mo | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 2
2
| 1 | 1 | 1.01: 6.9; 1.02: 5 | 1
| 1 | 2 | 1.01: 8; 1.02: 6 | 18
| 1 | 3 | 1.01: 7.1; 1.02: 5 | 16
| 1 | 4 | 1.01: 6.9; 1.02: 5 | 6.25
| 1 | 5 | 1.01: 7.6; 1.02: 5 | 47
| Mean |  | 1.01: 7.3; 1.02: 5.2
| Median |  | 1.01: 7.1; 1.02: 5
| Std Dev. |  | 1.01: 0.48477; 1.02: 0.44721
| Rel.Std.Dev. |  | 1.01: 6.641%; 1.02: 8.6%
| PDM3 |  | 1.01: 16.8%; 1.02: -16.8%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Na in OREAS 134a (Indicative Value 303 ppm) | [16 empty] | Indicative
Na | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 0
0
| 1 | 1 | 1.01: 300; 1.02: 360 | 1
| 1 | 2 | 1.01: 300; 1.02: 310 | 29
| 1 | 3 | 1.01: 300; 1.02: 280 | 16
| 1 | 4 | 1.01: 300; 1.02: 290 | 303
| 1 | 5 | 1.01: 300; 1.02: 290 | 48
| Mean |  | 1.01: 300; 1.02: 306
| Median |  | 1.01: 300; 1.02: 290
| Std Dev. |  | 1.01: 0; 1.02: 32.0936
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 10.49%
| PDM3 |  | 1.01: -0.9901%; 1.02: 0.9901%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Nb in OREAS 134a (Indicative Value 1.86 ppm) | [16 empty] | Indicative
Nb | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.4; 1.02: 2.2 | 1
| 1 | 2 | 1.01: 1.7; 1.02: 2 | 19
| 1 | 3 | 1.01: 1.5; 1.02: 2.4 | 16
| 1 | 4 | 1.01: 2; 1.02: 1.8 | 1.86
| 1 | 5 | 1.01: 1.4; 1.02: 2.2 | 49
| Mean |  | 1.01: 1.6; 1.02: 2.12
| Median |  | 1.01: 1.5; 1.02: 2.2
| Std Dev. |  | 1.01: 0.25495; 1.02: 0.22804
| Rel.Std.Dev. |  | 1.01: 15.93%; 1.02: 10.76%
| PDM3 |  | 1.01: -13.98%; 1.02: 13.98%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Nd in OREAS 134a (Indicative Value 7.42 ppm) | [16 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 2
2
| 1 | 1 | 1.01: 8.1 | 1
| 1 | 2 | 1.01: 7.5 | 20
| 1 | 3 | 1.01: 7.5 | 16
| 1 | 4 | 1.01: 6.9 | 7.42
| 1 | 5 | 1.01: 7.1 | 50
| Mean |  | 1.01: 7.42
| Median |  | 1.01: 7.5
| Std Dev. |  | 1.01: 0.46043
| Rel.Std.Dev. |  | 1.01: 6.205%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ni in OREAS 134a (Indicative Value 21.3 ppm) | [16 empty] | Indicative
Ni | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 1
1
| 1 | 1 | 1.01: 20.7; 1.02: 21 | 1
| 1 | 2 | 1.01: 20.8; 1.02: 21 | 21
| 1 | 3 | 1.01: 20; 1.02: 21 | 16
| 1 | 4 | 1.01: 23.4; 1.02: 22 | 21.29
| 1 | 5 | 1.01: 21; 1.02: 22 | 51
| Mean |  | 1.01: 21.18; 1.02: 21.4
| Median |  | 1.01: 20.8; 1.02: 21
| Std Dev. |  | 1.01: 1.2969; 1.02: 0.54772
| Rel.Std.Dev. |  | 1.01: 6.123%; 1.02: 2.559%
| PDM3 |  | 1.01: -0.5167%; 1.02: 0.5167%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for P in OREAS 134a (Indicative Value 294 ppm) | [16 empty] | Indicative
P | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 0
0
| 1 | 1 | 1.01: 300; 1.02: 280 | 1
| 1 | 2 | 1.01: 300; 1.02: 300 | 30
| 1 | 3 | 1.01: 300; 1.02: 290 | 16
| 1 | 4 | 1.01: 300; 1.02: 280 | 294
| 1 | 5 | 1.01: 300; 1.02: 290 | 52
| Mean |  | 1.01: 300; 1.02: 288
| Median |  | 1.01: 300; 1.02: 290
| Std Dev. |  | 1.01: 0; 1.02: 8.3666
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 2.905%
| PDM3 |  | 1.01: 2.041%; 1.02: -2.041%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Pb in OREAS 134a (Certified Value 12.79 wt.%) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 11.97 | 5.33 | 12.8 | 13.9316 | 11.9 | 13 | 12.7 | 13.05 | 13.84 | 12.6 | 12.3 | 11.75 | 12.8 | 13.1 |  | 1
| 1 | 2 | 11.913 | 6.28 | 12.8 | 13.9731 | 11.9 | 13.15 | 12.3 | 12.85 | 14.2343 | 12.6 | 11.6 | 12.15 | 12.7 | 12.5 |  | 22
| 1 | 3 | 11.803 | 5.56 | 11.9 | 14.1516 | 11.8 | 13.05 | 12.5 | 13.35 | 14.1358 | 13 | 11.8 | 11.65 | 12.5 | 13.4 |  | 16
| 1 | 4 | 11.886 | 5.44 | 12.9 | 13.9479 | 12.2 | 13.1 | 12.75 | 13.15 | 14.7761 | 12.8 | 12.7 | 12.25 | 12.7 | 13.2 |  | 12.7865
| 1 | 5 | 11.874 | 5.34 | 12.1 | 13.8392 | 11.9 | 12.7 | 12.8 | 12.65 | 14.382 | 12.9 | 12.7 | 12 | 12.9 | 13.1 |  | 29
| Mean |  | 11.8892 | 5.59 | 12.5 | 13.9687 | 11.94 | 13 | 12.61 | 13.01 | 14.2736 | 12.78 | 12.22 | 11.96 | 12.72 | 13.06
| Median |  | 11.886 | 5.44 | 12.8 | 13.9479 | 11.9 | 13.05 | 12.7 | 13.05 | 14.2343 | 12.8 | 12.3 | 12 | 12.7 | 13.1
| Std Dev. |  | 0.060759 | 0.39674 | 0.46368 | 0.11407 | 0.15166 | 0.17678 | 0.20736 | 0.27019 | 0.3438 | 0.17889 | 0.50695 | 0.25593 | 0.14832 | 0.33615
| Rel.Std.Dev. |  | 0.511% | 7.097% | 3.709% | 0.8166% | 1.27% | 1.36% | 1.644% | 2.077% | 2.409% | 1.4% | 4.149% | 2.14% | 1.166% | 2.574%
| PDM3 |  | -7.017% | -56.28% | -2.24% | 9.246% | -6.62% | 1.67% | -1.38% | 1.748% | 11.63% | -0.05055% | -4.43% | -6.464% | -0.5198% | 2.139%
| Z-Score (Absolute) |  | 1.25 | 11.36 | 0.26 | 2.09 | 1.16 | 0.54 | 0.09 | 0.55 | 2.58 | 0.18 | 0.71 | 1.13 | 0.09 | 0.63
| Analytical results for Pr in OREAS 134a (Indicative Value 1.68 ppm) | [16 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 1.7 |  | 1
| 1 | 2 | 1.01: 1.7 |  | 23
| 1 | 3 | 1.01: 1.8 |  | 16
| 1 | 4 | 1.01: 1.5 |  | 1.68
| 1 | 5 | 1.01: 1.7 |  | 53
| Mean |  | 1.01: 1.68
| Median |  | 1.01: 1.7
| Std Dev. |  | 1.01: 0.10954
| Rel.Std.Dev. |  | 1.01: 6.521%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Rb in OREAS 134a (Indicative Value 27.4 ppm) | [16 empty] | Indicative
Rb | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 24.5; 1.02: 29.7 |  | 1
| 1 | 2 | 1.01: 27.6; 1.02: 27.6 |  | 24
| 1 | 3 | 1.01: 29; 1.02: 26.8 |  | 16
| 1 | 4 | 1.01: 28.4; 1.02: 25.7 |  | 27.38
| 1 | 5 | 1.01: 26.4; 1.02: 28.1 |  | 54
| Mean |  | 1.01: 27.18; 1.02: 27.58
| Median |  | 1.01: 27.6; 1.02: 27.6
| Std Dev. |  | 1.01: 1.7866; 1.02: 1.4923
| Rel.Std.Dev. |  | 1.01: 6.573%; 1.02: 5.411%
| PDM3 |  | 1.01: -0.7305%; 1.02: 0.7305%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Re in OREAS 134a (Indicative Value < 1 ppb) | [16 empty] | Indicative
Re | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppb
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: < 1 |  | 1
| 1 | 2 | 1.01: < 1 |  | 25
| 1 | 3 | 1.01: < 1 |  | 16
| 1 | 4 | 1.01: < 1 |  | < 1
| 1 | 5 | 1.01: < 1 |  | 55
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for S in OREAS 134a (Certified Value 19.27 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 | [5 empty] | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | [5 empty] | 2
3
| 1 | 1 | 18.6238 | > 10 | 19.05 | 20.9 | 19 | 19.3 | > 20 | 19.3 | > 10 | 16.7 | [5 empty] | 1
| 1 | 2 | 18.2809 | > 10 | 19.15 | 19.7 | 19.4 | 19.5 | > 20 | 19.5 | > 10 | 17.2 | [5 empty] | 31
| 1 | 3 | 18.4249 | > 10 | 19.95 | 19.8 | 19.4 | 20 | > 20 | 19.5 | > 10 | 18.3 | [5 empty] | 16
| 1 | 4 | 18.8223 | > 10 | 19.15 | 19.2 | 19.1 | 19.4 | > 20 | 18.9 | > 10 | 17.6 | [5 empty] | 19.2749
| 1 | 5 | 18.744 | > 10 | 19.25 | 19.2 | 19.6 | 19.7 | > 20 | 19.2 | > 10 | 17.3 | [5 empty] | 30
| Mean |  | 1.01: 18.5792; 1.03: 19.31; 1.04: 19.76; 1.05: 19.3; 1.06: 19.58; 1.08: 19.28; 1.10: 17.42
| Median |  | 1.01: 18.6238; 1.03: 19.15; 1.04: 19.7; 1.05: 19.4; 1.06: 19.5; 1.08: 19.3; 1.10: 17.3
| Std Dev. |  | 1.01: 0.22419; 1.03: 0.36469; 1.04: 0.69498; 1.05: 0.24495; 1.06: 0.27749; 1.08: 0.249; 1.10: 0.58907
| Rel.Std.Dev. |  | 1.01: 1.207%; 1.03: 1.889%; 1.04: 3.517%; 1.05: 1.269%; 1.06: 1.417%; 1.08: 1.291%; 1.10: 3.382%
| PDM3 |  | 1.01: -3.609%; 1.03: 0.1823%; 1.04: 2.517%; 1.05: 0.1304%; 1.06: 1.583%; 1.08: 0.02665%; 1.10: -9.623%
| Z-Score (Absolute) |  | 1.74 | NA | 0.02 | 1.11 | 0 | 0.67 | NA | 0.05 | NA | 4.53
| Analytical results for Sb in OREAS 134a (Certified Value 115 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 98 | 21.1 | 145 | 117.3 | 127 | 121 | 131 | 120 | 119.2 | 126 | 81.4 | 27 | 110 | 110 | 118 | 1
| 1 | 2 | 96 | 27.5 | 148 | 116.8 | 127 | 106 | 133 | 100 | 131.3 | 127 | 84.6 | 24 | 100 | 110 | 116 | 26
| 1 | 3 | 96 | 23.5 | 139 | 116.57 | 124 | 118 | 133 | 110 | 120.4 | 125 | 84.3 | 25 | 110 | 100 | 129 | 16
| 1 | 4 | 98 | 16.9 | 130 | 118.25 | 126 | 115 | 120 | 100 | 120.2 | 123 | 94.6 | 20 | 120 | 100 | 124 | 115.4795
| 1 | 5 | 96 | 15.7 | 131 | 119.15 | 124 | 121 | 119 | 110 | 120 | 131 | 79.7 | 26 | 100 | 120 | 120 | 31
| Mean |  | 96.8 | 20.94 | 138.6 | 117.614 | 125.6 | 116.2 | 127.2 | 108 | 122.22 | 126.4 | 84.92 | 24.4 | 108 | 108 | 121.4
| Median |  | 96 | 21.1 | 139 | 117.3 | 126 | 118 | 131 | 110 | 120.2 | 126 | 84.3 | 25 | 110 | 110 | 120
| Std Dev. |  | 1.0954 | 4.832 | 8.0808 | 1.0741 | 1.5166 | 6.2209 | 7.0852 | 8.3666 | 5.0963 | 2.9665 | 5.7842 | 2.7019 | 8.3666 | 8.3666 | 5.1769
| Rel.Std.Dev. |  | 1.132% | 23.08% | 5.83% | 0.9133% | 1.207% | 5.354% | 5.57% | 7.747% | 4.17% | 2.347% | 6.811% | 11.07% | 7.747% | 7.747% | 4.264%
| PDM3 |  | -16.18% | -81.87% | 20.02% | 1.848% | 8.764% | 0.6239% | 10.15% | -6.477% | 5.837% | 9.457% | -26.46% | -78.87% | -6.477% | -6.477% | 5.127%
| Z-Score (Absolute) |  | 1.39 | 6.83 | 1.61 | 0.1 | 0.67 | 0 | 0.79 | 0.59 | 0.43 | 0.73 | 2.24 | 6.59 | 0.59 | 0.59 | 0.37
| Analytical results for Sc in OREAS 134a (Indicative Value 1.4 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 2
2
| 1 | 1 | 1.01: 2; 1.02: < 4 | 1
| 1 | 2 | 1.01: 1; 1.02: < 4 | 27
| 1 | 3 | 1.01: 1; 1.02: < 4 | 16
| 1 | 4 | 1.01: 2; 1.02: < 4 | 1.4
| 1 | 5 | 1.01: 1; 1.02: < 4 | 56
| Mean |  | 1.01: 1.4
Median | bar: 1.01
| Std Dev. |  | 1.01: 0.54772
| Rel.Std.Dev. |  | 1.01: 39.12%
PDM3 | bar: 1.01
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Se in OREAS 134a (Indicative Value 7.44 ppm) | [16 empty] | Indicative
Se | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 2
2
| 1 | 1 | 1.01: 6.7 | 1
| 1 | 2 | 1.01: 6.5 | 51
| 1 | 3 | 1.01: 7.4 | 16
| 1 | 4 | 1.01: 6.5 | 7.44
| 1 | 5 | 1.01: 10.1 | 57
| Mean |  | 1.01: 7.44
| Median |  | 1.01: 6.7
| Std Dev. |  | 1.01: 1.5323
| Rel.Std.Dev. |  | 1.01: 20.6%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sm in OREAS 134a (Indicative Value 1.38 ppm) | [16 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.5 | 1
| 1 | 2 | 1.01: 1.5 | 13
| 1 | 3 | 1.01: 1.5 | 16
| 1 | 4 | 1.01: 1.2 | 1.38
| 1 | 5 | 1.01: 1.2 | 42
| Mean |  | 1.01: 1.38
| Median |  | 1.01: 1.5
| Std Dev. |  | 1.01: 0.16432
| Rel.Std.Dev. |  | 1.01: 11.91%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sn in OREAS 134a (Indicative Value 1.13 ppm) | [16 empty] | Indicative
Sn | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.4; 1.02: 1 | 1
| 1 | 2 | 1.01: 1.7; 1.02: 1 | 14
| 1 | 3 | 1.01: 1.9; 1.02: < 1 | 16
| 1 | 4 | 1.01: 0.5; 1.02: 1 | 1.13
| 1 | 5 | 1.01: 1.3; 1.02: 1 | 43
| Mean |  | 1.01: 1.36; 1.02: 1
| Median |  | 1.01: 1.4; 1.02: 1
| Std Dev. |  | 1.01: 0.53666; 1.02: 0
| Rel.Std.Dev. |  | 1.01: 39.46%; 1.02: 0%
| PDM3 |  | 1.01: 20.35%; 1.02: -11.5%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Sr in OREAS 134a (Indicative Value 18.3 ppm) | [16 empty] | Indicative
Sr | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 1
1
| 1 | 1 | 1.01: 21; 1.02: 14 | 1
| 1 | 2 | 1.01: 21; 1.02: 17 | 34
| 1 | 3 | 1.01: 21; 1.02: 16 | 16
| 1 | 4 | 1.01: 21; 1.02: 15 | 18.3
| 1 | 5 | 1.01: 22; 1.02: 15 | 44
| Mean |  | 1.01: 21.2; 1.02: 15.4
| Median |  | 1.01: 21; 1.02: 15
| Std Dev. |  | 1.01: 0.44721; 1.02: 1.1402
| Rel.Std.Dev. |  | 1.01: 2.109%; 1.02: 7.404%
| PDM3 |  | 1.01: 15.85%; 1.02: -15.85%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ta in OREAS 134a (Indicative Value < 500 ppb) | [16 empty] | Indicative
Ta | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppb
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 0
0
| 1 | 1 | 1.01: < 500; 1.02: 100 | 1
| 1 | 2 | 1.01: < 500; 1.02: < 100 | 16
| 1 | 3 | 1.01: < 500; 1.02: < 100 | 16
| 1 | 4 | 1.01: < 500; 1.02: < 100 | < 500
| 1 | 5 | 1.01: < 500; 1.02: < 100 | 45
| Mean |  | 1.02: 100
| Median |  | 1.02: 100
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Tb in OREAS 134a (Indicative Value 180 ppb) | [16 empty] | Indicative
Tb | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 0
0
| 1 | 1 | 1.01: 200 | 1
| 1 | 2 | 1.01: 200 | 17
| 1 | 3 | 1.01: 200 | 16
| 1 | 4 | 1.01: 100 | 180
| 1 | 5 | 1.01: 200 | 46
| Mean |  | 1.01: 180
| Median |  | 1.01: 200
| Std Dev. |  | 1.01: 44.7214
| Rel.Std.Dev. |  | 1.01: 24.85%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Te in OREAS 134a (Indicative Value 90 ppb) | [16 empty] | Indicative
Te | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 1
1
| 1 | 1 | 1.01: 100 | 1
| 1 | 2 | 1.01: 100 | 18
| 1 | 3 | 1.01: < 100 | 16
| 1 | 4 | 1.01: 100 | 90
| 1 | 5 | 1.01: 100 | 47
| Mean |  | 1.01: 100
| Median |  | 1.01: 100
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
| PDM3 |  | 1.01: 11.11%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Th in OREAS 134a (Indicative Value 1.44 ppm) | [16 empty] | Indicative
Th | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 2; 1.02: 1.5 | 1
| 1 | 2 | 1.01: 1.9; 1.02: 0.9 | 19
| 1 | 3 | 1.01: 1.8; 1.02: 1 | 16
| 1 | 4 | 1.01: 1.9; 1.02: 0.9 | 1.4425
| 1 | 5 | 1.01: 2.2; 1.02: 0.9 | 48
| Mean |  | 1.01: 1.96; 1.02: 1.04
| Median |  | 1.01: 1.9; 1.02: 0.9
| Std Dev. |  | 1.01: 0.15166; 1.02: 0.26077
| Rel.Std.Dev. |  | 1.01: 7.738%; 1.02: 25.07%
| PDM3 |  | 1.01: 35.88%; 1.02: -27.9%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ti in OREAS 134a (Indicative Value 479 ppm) | [16 empty] | Indicative
Ti | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*OES | 0
0
| 1 | 1 | 1.01: 350; 1.02: 600 | 1
| 1 | 2 | 1.01: 360; 1.02: 600 | 29
| 1 | 3 | 1.01: 360; 1.02: 600 | 16
| 1 | 4 | 1.01: 360; 1.02: 600 | 479
| 1 | 5 | 1.01: 360; 1.02: 600 | 49
| Mean |  | 1.01: 358; 1.02: 600
| Median |  | 1.01: 360; 1.02: 600
| Std Dev. |  | 1.01: 4.4721; 1.02: 0
| Rel.Std.Dev. |  | 1.01: 1.249%; 1.02: 0%
| PDM3 |  | 1.01: -25.26%; 1.02: 25.26%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Tl in OREAS 134a (Indicative Value 93 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: 4A*OES | 0
0
| 1 | 1 | 1.01: 90 | 1
| 1 | 2 | 1.01: 95 | 20
| 1 | 3 | 1.01: 97 | 16
| 1 | 4 | 1.01: 91 | 93
| 1 | 5 | 1.01: 92 | 50
| Mean |  | 1.01: 93
| Median |  | 1.01: 92
| Std Dev. |  | 1.01: 2.9155
| Rel.Std.Dev. |  | 1.01: 3.135%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Tm in OREAS 134a (Indicative Value 100 ppb) | [16 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 0
0
| 1 | 1 | 1.01: 100 | 1
| 1 | 2 | 1.01: 100 | 21
| 1 | 3 | 1.01: 100 | 16
| 1 | 4 | 1.01: 100 | 100
| 1 | 5 | 1.01: 100 | 51
| Mean |  | 1.01: 100
| Median |  | 1.01: 100
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for U in OREAS 134a (Indicative Value 1.47 ppm) | [16 empty] | Indicative
U | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.4; 1.02: 1.4 | 1
| 1 | 2 | 1.01: 1.5; 1.02: 1.4 | 22
| 1 | 3 | 1.01: 1.5; 1.02: 1.6 | 16
| 1 | 4 | 1.01: 1.5; 1.02: 1.3 | 1.47
| 1 | 5 | 1.01: 1.5; 1.02: 1.6 | 52
| Mean |  | 1.01: 1.48; 1.02: 1.46
| Median |  | 1.01: 1.5; 1.02: 1.4
| Std Dev. |  | 1.01: 0.044721; 1.02: 0.13416
| Rel.Std.Dev. |  | 1.01: 3.022%; 1.02: 9.189%
| PDM3 |  | 1.01: 0.6803%; 1.02: -0.6803%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for V in OREAS 134a (Indicative Value 12.2 ppm) | [16 empty] | Indicative
V | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 1
1
| 1 | 1 | 1.01: < 10; 1.02: 13 | 1
| 1 | 2 | 1.01: < 10; 1.02: 10 | 23
| 1 | 3 | 1.01: < 10; 1.02: 10 | 16
| 1 | 4 | 1.01: < 10; 1.02: 13 | 12.2
| 1 | 5 | 1.01: < 10; 1.02: 15 | 53
| Mean |  | 1.02: 12.2
| Median |  | 1.02: 13
| Std Dev. |  | 1.02: 2.1679
| Rel.Std.Dev. |  | 1.02: 17.77%
PDM3 | bar: 1.02
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for W in OREAS 134a (Indicative Value 0.52 ppm) | [16 empty] | Indicative
W | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 0.8; 1.02: 0.7 | 1
| 1 | 2 | 1.01: < 0.5; 1.02: 0.5 | 24
| 1 | 3 | 1.01: 0.6; 1.02: 0.6 | 16
| 1 | 4 | 1.01: < 0.5; 1.02: 0.5 | 0.52
| 1 | 5 | 1.01: 0.5; 1.02: 0.5 | 54
| Mean |  | 1.01: 0.63333; 1.02: 0.56
| Median |  | 1.01: 0.6; 1.02: 0.5
| Std Dev. |  | 1.01: 0.15275; 1.02: 0.089443
| Rel.Std.Dev. |  | 1.01: 24.12%; 1.02: 15.97%
| PDM3 |  | 1.01: 21.79%; 1.02: 7.692%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Y in OREAS 134a (Indicative Value 5.18 ppm) | [16 empty] | Indicative
Y | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS | 2
2
| 1 | 1 | 1.01: 5.2; 1.02: 5.3 | 1
| 1 | 2 | 1.01: 5.5; 1.02: 4.7 | 25
| 1 | 3 | 1.01: 5.8; 1.02: 5.1 | 16
| 1 | 4 | 1.01: 4.8; 1.02: 4.5 | 5.18
| 1 | 5 | 1.01: 5.4; 1.02: 5.5 | 55
| Mean |  | 1.01: 5.34; 1.02: 5.02
| Median |  | 1.01: 5.4; 1.02: 5.1
| Std Dev. |  | 1.01: 0.37148; 1.02: 0.41473
| Rel.Std.Dev. |  | 1.01: 6.957%; 1.02: 8.262%
| PDM3 |  | 1.01: 3.089%; 1.02: -3.089%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Yb in OREAS 134a (Indicative Value 520 ppb) | [16 empty] | Indicative
Yb | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: 4A*MS | 0
0
| 1 | 1 | 1.01: 600 | 1
| 1 | 2 | 1.01: 500 | 26
| 1 | 3 | 1.01: 500 | 16
| 1 | 4 | 1.01: 500 | 520
| 1 | 5 | 1.01: 500 | 56
| Mean |  | 1.01: 520
| Median |  | 1.01: 500
| Std Dev. |  | 1.01: 44.7214
| Rel.Std.Dev. |  | 1.01: 8.6%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Zn in OREAS 134a (Certified Value 17.27 wt.%) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 16.949 | 11.5 | 17.7 | 18.1444 | 17.4 | 17.3 | 16.9 | 16.85 | 18.018 | 16.6 | 17.5 | 16.15 | 17.1 | 17.1 |  | 1
| 1 | 2 | 16.43 | 11.9 | 17.6 | 18.4968 | 18.3 | 17.2 | 16.8 | 16.55 | 18.2506 | 16.2 | 17.2 | 16.85 | 16.8 | 17.3 |  | 0
| 1 | 3 | 16.977 | 11.1 | 17.5 | 18.0169 | 17.7 | 17.15 | 17.05 | 17.25 | 18.2353 | 17.1 | 17.2 | 16.4 | 16.7 | 17.4 |  | 16
| 1 | 4 | 16.648 | 11 | 17.6 | 17.9033 | 18.3 | 17.3 | 16.9 | 17 | 18.264 | 17.3 | 17.3 | 17 | 16.9 | 17.3 |  | 17.2671
| 1 | 5 | 16.658 | 10.4 | 17.5 | 17.828 | 18.6 | 17.3 | 17.1 | 16.8 | 18.066 | 16.8 | 17.3 | 16.75 | 17 | 17.1 |  | 32
| Mean |  | 16.7324 | 11.18 | 17.58 | 18.0779 | 18.06 | 17.25 | 16.95 | 16.89 | 18.1668 | 16.8 | 17.3 | 16.63 | 16.9 | 17.24
| Median |  | 16.658 | 11.1 | 17.6 | 18.0169 | 18.3 | 17.3 | 16.9 | 16.85 | 18.2353 | 16.8 | 17.3 | 16.75 | 16.9 | 17.3
| Std Dev. |  | 0.22959 | 0.56303 | 0.083666 | 0.26294 | 0.49295 | 0.070711 | 0.12247 | 0.25836 | 0.11561 | 0.43012 | 0.12247 | 0.34749 | 0.15811 | 0.13416
| Rel.Std.Dev. |  | 1.372% | 5.036% | 0.4759% | 1.455% | 2.73% | 0.4099% | 0.7226% | 1.53% | 0.6364% | 2.56% | 0.7079% | 2.09% | 0.9356% | 0.7782%
| PDM3 |  | -3.097% | -35.25% | 1.812% | 4.696% | 4.592% | -0.09905% | -1.836% | -2.184% | 5.21% | -2.705% | 0.1905% | -3.69% | -2.126% | -0.157%
| Z-Score (Absolute) |  | 0.74 | 12.13 | 0.99 | 2.02 | 1.98 | 0.32 | 0.3 | 0.42 | 2.2 | 0.6 | 0.42 | 0.95 | 0.4 | 0.3
| Analytical results for Zr in OREAS 134a (Indicative Value 27.8 ppm) | [16 empty] | Indicative
Zr | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 29.8; 1.02: 42 |  | 1
| 1 | 2 | 1.01: 32.5; 1.02: 22 |  | 27
| 1 | 3 | 1.01: 31.2; 1.02: 24 |  | 16
| 1 | 4 | 1.01: 33.7; 1.02: 24 |  | 27.75
| 1 | 5 | 1.01: 30.3; 1.02: 26 |  | 57
| Mean |  | 1.01: 31.5; 1.02: 27.6
| Median |  | 1.01: 31.2; 1.02: 24
| Std Dev. |  | 1.01: 1.6016; 1.02: 8.1731
| Rel.Std.Dev. |  | 1.01: 5.084%; 1.02: 29.61%
| PDM3 |  | 1.01: 13.51%; 1.02: -0.5405%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C33: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C50: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C84: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C322: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C339: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C356: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C373: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C390: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C407: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C424: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C628: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C645: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C866: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C883: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C934: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 0.15-1.0g
| Analytical results for Ag in OREAS 134a (Certified Value 194 ppm) | [16 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/AAS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 01g | 01g | 01g | 0.4g | 0.5g | 01g | 01g | 01g | 0.5g | 0.4g | 0.2g | 0.25g |  |  | 0
| 1 | 1 | 206.8 | 191 | 169 | 188.5 | 196 | 191 | 213.4 | 191 | 190.2 | > 100 | 180 | 214 | 184 |  |  | 1
| 1 | 2 | 205.9 | 198 | 168 | 184.08 | 197 | 190 | 203.7 | 190 | 189.1 | > 100 | 199 | 208 | 182 |  |  | 12
| 1 | 3 | 201.6 | 203 | 163 | 189.71 | 201 | 189 | 212.7 | 186 | 190.8 | > 100 | 190 | 204 | 182 |  |  | 16
| 1 | 4 | 197 | 199 | 163 | 191.99 | 202 | 189 | 209.9 | 186 | 194.5 | > 100 | 198 | 209 | 185 |  |  | 193.5817
| 1 | 5 | 204.8 | 199 | 169 | 102.15 | 197 | 192 | 219.1 | 189 | 188.2 | > 100 | 135 | 220 | 178 |  |  | 34
| Mean |  | 203.22 | 198 | 166.4 | 171.286 | 198.6 | 190.2 | 211.76 | 188.4 | 190.56 |  | 180.4 | 211 | 182.2
| Median |  | 204.8 | 199 | 168 | 188.5 | 197 | 190 | 212.7 | 189 | 190.2 |  | 190 | 209 | 182
| Std Dev. |  | 3.994 | 4.3589 | 3.1305 | 38.7553 | 2.7019 | 1.3038 | 5.6096 | 2.3022 | 2.4193 |  | 26.5009 | 6.1644 | 2.6833
| Rel.Std.Dev. |  | 1.965% | 2.201% | 1.881% | 22.63% | 1.36% | 0.6855% | 2.649% | 1.222% | 1.27% |  | 14.69% | 2.922% | 1.473%
| PDM3 |  | 4.979% | 2.282% | -14.04% | -11.52% | 2.592% | -1.747% | 9.391% | -2.677% | -1.561% |  | -6.809% | 8.998% | -5.88%
| Z-Score (Absolute) |  | 0.95 | 0.56 | 1.78 | 1.41 | 0.61 | 0.01 | 1.58 | 0.15 | 0.01 | NA | 0.74 | 1.53 | 0.61
| Analytical results for Al2O3 in OREAS 134a (Certified Value 0.774 wt.%) | [16 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.16 |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.747 | 0.624 | 0.80625 | 0.633 | 0.87 | 0.78 | 0.86 | 0.794 | 0.72 | 1.077 | 0.888 | 0.7105 |  |  |  | 1
| 1 | 2 | 0.747 | 0.68 | 0.81985 | 0.633 | 0.87 | 0.76 | 0.89 | 0.831 | 0.7 | 1.058 | 0.869 | 0.701 |  |  |  | 1
| 1 | 3 | 0.728 | 0.624 | 0.80701 | 0.6916 | 0.91 | 0.76 | 0.87 | 0.812 | 0.72 | 1.096 | 0.869 | 0.6953 |  |  |  | 16
| 1 | 4 | 0.706 | 0.68 | 0.83818 | 0.6557 | 0.89 | 0.74 | 0.9 | 0.831 | 0.72 | 1.096 | 0.888 | 0.6972 |  |  |  | 0.77352
| 1 | 5 | 0.747 | 0.68 | 0.83403 | 0.7426 | 0.89 | 0.76 | 0.87 | 0.85 | 0.74 | 1.304 | 0.869 | 0.6953 |  |  |  | 35
| Mean |  | 0.735 | 0.6576 | 0.82106 | 0.67118 | 0.886 | 0.76 | 0.878 | 0.8236 | 0.72 | 1.1262 | 0.8766 | 0.69986
| Median |  | 0.747 | 0.68 | 0.81985 | 0.6557 | 0.89 | 0.76 | 0.87 | 0.831 | 0.72 | 1.096 | 0.869 | 0.6972
| Std Dev. |  | 0.01818 | 0.030672 | 0.014828 | 0.046564 | 0.016733 | 0.014142 | 0.016432 | 0.021314 | 0.014142 | 0.10063 | 0.010407 | 0.0063869
| Rel.Std.Dev. |  | 2.473% | 4.664% | 1.806% | 6.938% | 1.889% | 1.861% | 1.871% | 2.588% | 1.964% | 8.936% | 1.187% | 0.9126%
| PDM3 |  | -4.979% | -14.99% | 6.147% | -13.23% | 14.54% | -1.747% | 13.51% | 6.475% | -6.919% | 45.59% | 13.33% | -9.522%
| Z-Score (Absolute) |  | 0.43 | 1.03 | 0.24 | 0.93 | 0.74 | 0.24 | 0.68 | 0.26 | 0.55 | 2.61 | 0.67 | 0.7
| Analytical results for As in OREAS 134a (Certified Value 226 ppm) | [16 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | 0
|  |  | 01g | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.2g | 0.25g | 0
| 1 | 1 | 211 | 208 | 195 | 224.7 | 248 | 270 | 250 | 277 | 235 | 222 | 217 | 205 | 210 | 200 | 224 | 1
| 1 | 2 | 216 | 220 | 197 | 221.9 | 239 | 268 | 250 | 277 | 234 | 228 | 210 | 211 | 170 | 200 | 212 | 2
| 1 | 3 | 230 | 218 | 201 | 223.7 | 239 | 274 | 238 | 270 | 222 | 220 | 210 | 200 | 140 | 200 | 221 | 16
| 1 | 4 | 219 | 227 | 202 | 227.8 | 238 | 272 | 238 | 272 | 236 | 219 | 215 | 199 | 220 | 200 | 217 | 225.505
| 1 | 5 | 224 | 228 | 199 | 239.8 | 255 | 271 | 241 | 262 | 236 | 224 | 214 | 200 | 200 | 200 | 212 | 36
| Mean |  | 220 | 220.2 | 198.8 | 227.58 | 243.8 | 271 | 243.4 | 271.6 | 232.6 | 222.6 | 213.2 | 203 | 188 | 200 | 217.2
| Median |  | 219 | 220 | 199 | 224.7 | 239 | 271 | 241 | 272 | 235 | 222 | 214 | 200 | 200 | 200 | 217
| Std Dev. |  | 7.3144 | 8.0747 | 2.8636 | 7.1587 | 7.4632 | 2.2361 | 6.1482 | 6.1887 | 5.9833 | 3.5777 | 3.1145 | 5.0498 | 32.7109 | 0 | 5.3572
| Rel.Std.Dev. |  | 3.325% | 3.667% | 1.44% | 3.146% | 3.061% | 0.8251% | 2.526% | 2.279% | 2.572% | 1.607% | 1.461% | 2.488% | 17.4% | 0% | 2.467%
| PDM3 |  | -2.441% | -2.352% | -11.84% | 0.9202% | 8.113% | 20.17% | 7.936% | 20.44% | 3.146% | -1.288% | -5.457% | -9.98% | -16.63% | -11.31% | -3.683%
| Z-Score (Absolute) |  | 0.01 | 0 | 0.84 | 0.29 | 0.93 | 1.99 | 0.91 | 2.02 | 0.49 | 0.09 | 0.27 | 0.67 | 1.26 | 0.79 | 0.12
| Analytical results for Au in OREAS 134a (Indicative Value < 0.5 ppb) | [16 empty] | Indicative
Au | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: AR*MS | 3
|  |  | 1.01: 0.5g | 3
| 1 | 1 | 1.01: < 0.5 | 1
| 1 | 2 | 1.01: < 0.5 | 53
| 1 | 3 | 1.01: < 0.5 | 16
| 1 | 4 | 1.01: < 0.5 | < 0.5
| 1 | 5 | 1.01: < 0.5 | 59
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for B in OREAS 134a (Indicative Value < 10 ppm) | [16 empty] | Indicative
B | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: AR*OES | 1
|  |  | 1.01: 0.5g | 1
| 1 | 1 | 1.01: < 10 | 1
| 1 | 2 | 1.01: < 10 | 36
| 1 | 3 | 1.01: < 10 | 16
| 1 | 4 | 1.01: < 10 | < 10
| 1 | 5 | 1.01: < 10 | 60
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ba in OREAS 134a (Indicative Value < 600 ppm) | [16 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.16 |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES |  |  |  | 0
|  |  | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.25g |  |  |  | 0
| 1 | 1 | 353 | 11 | 465 | 7 | 10 | 10 | 10 | 242 | 132 | < 10 | 570 | 110 |  |  |  | 1
| 1 | 2 | 360 | 12 | 511 | 9 | 10 | 10 | 10 | 256 | 139 | < 10 | 510 | 127 |  |  |  | 14
| 1 | 3 | 339 | 12 | 549 | 7 | 10 | 20 | 10 | 238 | 137 | < 10 | 550 | 124 |  |  |  | 16
| 1 | 4 | 376 | 12 | 588 | 7 | 10 | 10 | 10 | 248 | 117 | < 10 | 570 | 125 |  |  |  | < 600
| 1 | 5 | 342 | 10 | < 1 | 7 | 10 | 10 | 10 | 245 | 136 | < 10 | 570 | 118 |  |  |  | 61
| Mean |  | 354 | 11.4 | 528.25 | 7.4 | 10 | 12 | 10 | 245.8 | 132.2 |  | 554 | 120.8
| Median |  | 353 | 12 | 530 | 7 | 10 | 10 | 10 | 245 | 136 |  | 570 | 124
| Std Dev. |  | 14.9164 | 0.89443 | 52.5952 | 0.89443 | 0 | 4.4721 | 0 | 6.7971 | 8.8713 |  | 26.0768 | 6.9065
| Rel.Std.Dev. |  | 4.214% | 7.846% | 9.956% | 12.09% | 0% | 37.27% | 0% | 2.765% | 6.711% |  | 4.707% | 5.717%
PDM3
| Z-Score (Absolute) |  | 3.22 | 0.62 | 3.99 | 0.66 | 0.63 | 0.61 | 0.63 | 2.01 | 0.74 | 0.69 | 5.46 | 0.61
| Analytical results for Be in OREAS 134a (Indicative Value 0.7 ppm) | [16 empty] | Indicative
Be | Round | Replicate | 1.01: Lab |  |  |  | 1
| No | No | 1.01: 1.12 |  |  |  | ppm
|  |  | 1.01: AR*OES |  |  |  | 2
|  |  | 1.01: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.7 |  |  |  | 1
| 1 | 2 | 1.01: 0.7 |  |  |  | 15
| 1 | 3 | 1.01: 0.7 |  |  |  | 16
| 1 | 4 | 1.01: 0.7 |  |  |  | 0.7
| 1 | 5 | 1.01: 0.7 |  |  |  | 62
| Mean |  | 1.01: 0.7
| Median |  | 1.01: 0.7
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for CaO in OREAS 134a (Indicative Value 6.09 wt.%) | [16 empty] | Indicative
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.16 |  |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.25g |  |  |  |  | 2
| 1 | 1 | 6.248 | 6.017 | 6.017 | 4.8972 | 4.52 | 5.57 | 6.184 | 6.16 | 3.736 | 5.62 | 5.919 |  |  |  |  | 1
| 1 | 2 | 6.246 | 6.073 | 6.156 | 4.8832 | 4.39 | 5.44 | 6.226 | 6.18 | 4.156 | 5.65 | 6.059 |  |  |  |  | 4
| 1 | 3 | 6.263 | 5.821 | 6.17 | 4.9951 | 4.78 | 5.27 | 5.989 | 6.21 | 3.806 | 5.67 | 6.128 |  |  |  |  | 16
| 1 | 4 | 6.286 | 6.184 | 6.226 | 5.0091 | 4.67 | 5.23 | 6.352 | 6.2 | 4.03 | 5.64 | 6.101 |  |  |  |  | 6.0901
| 1 | 5 | 6.276 | 6.282 | 6.254 | 5.205 | 4.38 | 5.32 | 5.989 | 6.28 | 3.4 | 5.64 | 6.226 |  |  |  |  | 63
| Mean |  | 6.2638 | 6.0754 | 6.1646 | 4.9979 | 4.548 | 5.366 | 6.148 | 6.206 | 3.8256 | 5.644 | 6.0866
| Median |  | 6.263 | 6.073 | 6.17 | 4.9951 | 4.52 | 5.32 | 6.184 | 6.2 | 3.806 | 5.64 | 6.101
| Std Dev. |  | 0.017384 | 0.17508 | 0.091721 | 0.12876 | 0.17513 | 0.13867 | 0.15776 | 0.045607 | 0.29176 | 0.018166 | 0.11203
| Rel.Std.Dev. |  | 0.2775% | 2.882% | 1.488% | 2.576% | 3.851% | 2.584% | 2.566% | 0.7349% | 7.626% | 0.3219% | 1.841%
| PDM3 |  | 2.852% | -0.2411% | 1.224% | -17.93% | -25.32% | -11.89% | 0.951% | 1.903% | -37.18% | -7.325% | -0.05719%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.32 | 3.85 | 5.46 | 2.54 | 0.26 | 0.47 | 8.05 | 1.54 | 0.04
| Analytical results for Cd in OREAS 134a (Certified Value 547 ppm) | [16 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/AAS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 01g | 01g | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.2g | 0.25g |  | 0
| 1 | 1 | 546 | 486 | 486 | 581.79 | 570 | 600 | 535 | 633.1 | 530 | 584 | 551 | 550 | 480 | 524 |  | 1
| 1 | 2 | 553 | 500 | 492 | 580.71 | 567 | 600 | 533 | 633.9 | 534 | 567 | 555 | 551 | 460 | 514 |  | 17
| 1 | 3 | 542 | 514 | 491 | 575.64 | 587 | 600 | 528 | 625 | 547 | 571 | 522 | 550 | 440 | 515 |  | 16
| 1 | 4 | 550 | 509 | 499 | 583.31 | 584 | 600 | 532 | 662.6 | 525 | 583 | 532 | 552 | 450 | 526 |  | 546.7723
| 1 | 5 | 541 | 509 | 507 | 603.94 | 580 | 600 | 511 | 670.8 | 536 | 576 | 530 | 549 | 480 | 500 |  | 37
| Mean |  | 546.4 | 503.6 | 495 | 585.078 | 577.6 | 600 | 527.8 | 645.08 | 534.4 | 576.2 | 538 | 550.4 | 462 | 515.8
| Median |  | 546 | 509 | 492 | 581.79 | 580 | 600 | 532 | 633.9 | 534 | 576 | 532 | 550 | 460 | 515
| Std Dev. |  | 5.1284 | 11.0589 | 8.1548 | 10.9301 | 8.735 | 0 | 9.7314 | 20.2496 | 8.2037 | 7.3959 | 14.2653 | 1.1402 | 17.8885 | 10.3053
| Rel.Std.Dev. |  | 0.9386% | 2.196% | 1.647% | 1.868% | 1.512% | 0% | 1.844% | 3.139% | 1.535% | 1.284% | 2.652% | 0.2072% | 3.872% | 1.998%
| PDM3 |  | -0.06809% | -7.896% | -9.469% | 7.006% | 5.638% | 9.735% | -3.47% | 17.98% | -2.263% | 5.382% | -1.604% | 0.6635% | -15.5% | -5.665%
| Z-Score (Absolute) |  | 0.08 | 0.75 | 0.92 | 0.83 | 0.69 | 1.12 | 0.28 | 2 | 0.15 | 0.66 | 0.08 | 0.16 | 1.56 | 0.51
| Analytical results for Ce in OREAS 134a (Indicative Value 13 ppm) | [16 empty] | Indicative
Ce | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*MS |  | 1
|  |  | 1.01: 0.5g |  | 1
| 1 | 1 | 1.01: 13.4 |  | 1
| 1 | 2 | 1.01: 7.94 |  | 18
| 1 | 3 | 1.01: 14.4 |  | 16
| 1 | 4 | 1.01: 13.8 |  | 13.048
| 1 | 5 | 1.01: 15.7 |  | 64
| Mean |  | 1.01: 13.048
| Median |  | 1.01: 13.8
| Std Dev. |  | 1.01: 2.985
| Rel.Std.Dev. |  | 1.01: 22.88%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Co in OREAS 134a (Certified Value 101 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.2g | 0.25g |  | 0
| 1 | 1 | 112 | 88.6 | 97 | 102.3 | 110 | 100 | 101 | 118.9 | 102 | 89.84 | 90 | 108 | 100 | 101 |  | 1
| 1 | 2 | 112 | 87.5 | 98 | 99.7 | 108 | 100 | 100 | 110.1 | 100 | 88.57 | 93 | 115 | 100 | 101 |  | 19
| 1 | 3 | 111 | 85.2 | 94 | 102.4 | 112 | 96 | 99 | 112.5 | 100 | 88.79 | 89 | 116 | 100 | 101 |  | 16
| 1 | 4 | 107 | 93.6 | 98 | 104.4 | 110 | 95 | 99 | 113.2 | 104 | 88.99 | 90 | 115 | 100 | 104 |  | 101.1341
| 1 | 5 | 110 | 93.7 | 98 | 110.7 | 110 | 97 | 95 | 113 | 110 | 89.85 | 88 | 105 | 110 | 102 |  | 38
| Mean |  | 110.4 | 89.72 | 97 | 103.9 | 110 | 97.6 | 98.8 | 113.54 | 103.2 | 89.208 | 90 | 111.8 | 102 | 101.8
| Median |  | 111 | 88.6 | 98 | 102.4 | 110 | 97 | 99 | 113 | 102 | 88.99 | 90 | 115 | 100 | 101
| Std Dev. |  | 2.0736 | 3.7917 | 1.7321 | 4.1515 | 1.4142 | 2.3022 | 2.2804 | 3.2424 | 4.1473 | 0.60018 | 1.8708 | 4.9699 | 4.4721 | 1.3038
| Rel.Std.Dev. |  | 1.878% | 4.226% | 1.786% | 3.996% | 1.286% | 2.359% | 2.308% | 2.856% | 4.019% | 0.6728% | 2.079% | 4.445% | 4.384% | 1.281%
| PDM3 |  | 9.162% | -11.29% | -4.088% | 2.735% | 8.766% | -3.495% | -2.308% | 12.27% | 2.043% | -11.79% | -11.01% | 10.55% | 0.8561% | 0.6584%
| Z-Score (Absolute) |  | 0.88 | 1.26 | 0.51 | 0.21 | 0.84 | 0.45 | 0.32 | 1.21 | 0.13 | 1.32 | 1.23 | 1.03 | 0.01 | 0.01
| Analytical results for Cr in OREAS 134a (Indicative Value 12.9 ppm) | [16 empty] | Indicative
Cr | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*OES |  | 1
|  |  | 1.01: 01g; 1.02: 0.5g |  | 1
| 1 | 1 | 1.01: 14.3; 1.02: 12 |  | 1
| 1 | 2 | 1.01: 12.7; 1.02: 13 |  | 20
| 1 | 3 | 1.01: 14.5; 1.02: 11 |  | 16
| 1 | 4 | 1.01: 13.5; 1.02: 12 |  | 12.88
| 1 | 5 | 1.01: 13.8; 1.02: 12 |  | 65
| Mean |  | 1.01: 13.76; 1.02: 12
| Median |  | 1.01: 13.8; 1.02: 12
| Std Dev. |  | 1.01: 0.71274; 1.02: 0.70711
| Rel.Std.Dev. |  | 1.01: 5.18%; 1.02: 5.893%
| PDM3 |  | 1.01: 6.832%; 1.02: -6.832%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Cs in OREAS 134a (Indicative Value 0.68 ppm) | [16 empty] | Indicative
Cs | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*MS |  | 2
|  |  | 1.01: 0.5g |  | 2
| 1 | 1 | 1.01: 0.65 |  | 1
| 1 | 2 | 1.01: 0.67 |  | 21
| 1 | 3 | 1.01: 0.68 |  | 16
| 1 | 4 | 1.01: 0.69 |  | 0.684
| 1 | 5 | 1.01: 0.73 |  | 66
| Mean |  | 1.01: 0.684
| Median |  | 1.01: 0.68
| Std Dev. |  | 1.01: 0.029665
| Rel.Std.Dev. |  | 1.01: 4.337%
| PDM3 |  | 1.01: -1.11e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Cu in OREAS 134a (Certified Value 1279 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.2g | 0.25g |  | 0
| 1 | 1 | 1274 | 1270 | 1170 | 1360 | 1580 | 1430 | 1515 | 1362.2 | 1240 | 1296 | 1280 | 1330 | 1300 | 1230 |  | 1
| 1 | 2 | 1264 | 1250 | 1130 | 1383 | 1570 | 1440 | 1500 | 1318.1 | 1230 | 1258 | 1340 | 1330 | 1200 | 1220 |  | 22
| 1 | 3 | 1262 | 1200 | 1130 | 1389 | 1580 | 1400 | 1500 | 1349.3 | 1240 | 1256 | 1230 | 1310 | 1200 | 1230 |  | 16
| 1 | 4 | 1262 | 1250 | 1120 | 1405 | 1550 | 1390 | 1525 | 1362.7 | 1210 | 1276 | 1180 | 1330 | 1200 | 1230 |  | 1279.4917
| 1 | 5 | 1266 | 1290 | 1150 | 1399 | 1560 | 1410 | 1470 | 1352.2 | 1240 | 1275 | 1200 | 1320 | 1300 | 1250 |  | 39
| Mean |  | 1265.6 | 1252 | 1140 | 1387.2 | 1568 | 1414 | 1502 | 1348.9 | 1232 | 1272.2 | 1246 | 1324 | 1240 | 1232
| Median |  | 1264 | 1250 | 1130 | 1389 | 1570 | 1410 | 1500 | 1352.2 | 1240 | 1275 | 1230 | 1330 | 1200 | 1230
| Std Dev. |  | 4.98 | 33.4664 | 20 | 17.4413 | 13.0384 | 20.7364 | 20.7966 | 18.2141 | 13.0384 | 16.2234 | 64.6529 | 8.9443 | 54.7723 | 10.9545
| Rel.Std.Dev. |  | 0.3935% | 2.673% | 1.754% | 1.257% | 0.8315% | 1.467% | 1.385% | 1.35% | 1.058% | 1.275% | 5.189% | 0.6755% | 4.417% | 0.8892%
| PDM3 |  | -1.086% | -2.149% | -10.9% | 8.418% | 22.55% | 10.51% | 17.39% | 5.425% | -3.712% | -0.5699% | -2.618% | 3.479% | -3.087% | -3.712%
| Z-Score (Absolute) |  | 0.05 | 0.25 | 1.89 | 1.73 | 4.38 | 2.13 | 3.42 | 1.17 | 0.54 | 0.05 | 0.34 | 0.81 | 0.42 | 0.54
| Analytical results for Dy in OREAS 134a (Indicative Value 1.11 ppm) | [16 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*MS |  | 2
|  |  | 1.01: 0.5g |  | 2
| 1 | 1 | 1.01: 1.14 |  | 1
| 1 | 2 | 1.01: 1.05 |  | 23
| 1 | 3 | 1.01: 1.19 |  | 16
| 1 | 4 | 1.01: 1.13 |  | 1.11
| 1 | 5 | 1.01: 1.04 |  | 67
| Mean |  | 1.01: 1.11
| Median |  | 1.01: 1.13
| Std Dev. |  | 1.01: 0.06364
| Rel.Std.Dev. |  | 1.01: 5.733%
| PDM3 |  | 1.01: -2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Er in OREAS 134a (Indicative Value 0.56 ppm) | [16 empty] | Indicative
Er | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*MS |  | 2
|  |  | 1.01: 0.5g |  | 2
| 1 | 1 | 1.01: 0.6 |  | 1
| 1 | 2 | 1.01: 0.5 |  | 24
| 1 | 3 | 1.01: 0.6 |  | 16
| 1 | 4 | 1.01: 0.6 |  | 0.56
| 1 | 5 | 1.01: 0.5 |  | 68
| Mean |  | 1.01: 0.56
| Median |  | 1.01: 0.6
| Std Dev. |  | 1.01: 0.054772
| Rel.Std.Dev. |  | 1.01: 9.781%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Eu in OREAS 134a (Indicative Value 360 ppb) | [16 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppb
|  |  | 1.01: AR*MS |  | 0
|  |  | 1.01: 0.5g |  | 0
| 1 | 1 | 1.01: 400 |  | 1
| 1 | 2 | 1.01: 300 |  | 25
| 1 | 3 | 1.01: 400 |  | 16
| 1 | 4 | 1.01: 400 |  | 360
| 1 | 5 | 1.01: 300 |  | 69
| Mean |  | 1.01: 360
| Median |  | 1.01: 400
| Std Dev. |  | 1.01: 54.7723
| Rel.Std.Dev. |  | 1.01: 15.21%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Fe in OREAS 134a (Certified Value 11.89 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/AAS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/AAS | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 01g | 01g | 01g | 0.4g | 0.4g | 01g | 01g | 01g | 0.5g | 0.4g | 0.2g | 0.25g |  |  | 3
| 1 | 1 | 11.95 | 11.95 | 11.1 | 13.04 | 11.1 | 12.25 | 12.66 | 12.3 | 10.508 | 10.1 | 11.55 | 13.3 | 12.1 |  |  | 1
| 1 | 2 | 11.999 | 11.75 | 11.6 | 13.28 | 10.7 | 12.35 | 12.57 | 12.4 | 10.254 | 10.5 | 12.45 | 12.9 | 12.2 |  |  | 0
| 1 | 3 | 11.815 | 11.3 | 11.4 | 13.34 | 10.9 | 12.65 | 12.6 | 12.4 | 10.282 | 9.76 | 11.65 | 12.6 | 12.3 |  |  | 16
| 1 | 4 | 11.82 | 11.7 | 11.4 | 13.5 | 11.2 | 12.8 | 12.6 | 12.4 | 10.43 | 9.51 | 12.45 | 12.9 | 12.3 |  |  | 11.8925
| 1 | 5 | 11.902 | 12.05 | 11.1 | 13.46 | 11.7 | 12.4 | 12.7 | 12.4 | 10.48 | 10.2 | 11.65 | 13.5 | 12.6 |  |  | 40
| Mean |  | 11.8972 | 11.75 | 11.32 | 13.324 | 11.12 | 12.49 | 12.626 | 12.38 | 10.3908 | 10.014 | 11.95 | 13.04 | 12.3
| Median |  | 11.902 | 11.75 | 11.4 | 13.34 | 11.1 | 12.4 | 12.6 | 12.4 | 10.43 | 10.1 | 11.65 | 12.9 | 12.3
| Std Dev. |  | 0.080453 | 0.2894 | 0.21679 | 0.18188 | 0.37683 | 0.22749 | 0.052726 | 0.044721 | 0.11595 | 0.38624 | 0.45826 | 0.35777 | 0.18708
| Rel.Std.Dev. |  | 0.6762% | 2.463% | 1.915% | 1.365% | 3.389% | 1.821% | 0.4176% | 0.3612% | 1.116% | 3.857% | 3.835% | 2.744% | 1.521%
| PDM3 |  | 0.03984% | -1.198% | -4.814% | 12.04% | -6.495% | 5.025% | 6.168% | 4.1% | -12.63% | -15.8% | 0.4838% | 9.649% | 3.427%
| Z-Score (Absolute) |  | 0.06 | 0.21 | 0.67 | 1.47 | 0.89 | 0.58 | 0.72 | 0.46 | 1.67 | 2.07 | 0 | 1.17 | 0.37
| Analytical results for Ga in OREAS 134a (Indicative Value 3.13 ppm) | [16 empty] | Indicative
Ga | Round | Replicate | 1.01: Lab; 1.02: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS |  |  | 2
|  |  | 1.01: 01g; 1.02: 0.5g |  |  | 2
| 1 | 1 | 1.01: < 5; 1.02: 3.06 |  |  | 1
| 1 | 2 | 1.01: < 5; 1.02: 3.02 |  |  | 26
| 1 | 3 | 1.01: < 5; 1.02: 3.17 |  |  | 16
| 1 | 4 | 1.01: < 5; 1.02: 3.13 |  |  | 3.13
| 1 | 5 | 1.01: < 5; 1.02: 3.27 |  |  | 70
| Mean |  | 1.02: 3.13
| Median |  | 1.02: 3.13
| Std Dev. |  | 1.02: 0.097724
| Rel.Std.Dev. |  | 1.02: 3.122%
PDM3 | bar: 1.02
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Gd in OREAS 134a (Indicative Value 1.34 ppm) | [16 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 2
|  |  | 1.01: 0.5g |  |  | 2
| 1 | 1 | 1.01: 1.3 |  |  | 1
| 1 | 2 | 1.01: 1.3 |  |  | 27
| 1 | 3 | 1.01: 1.4 |  |  | 16
| 1 | 4 | 1.01: 1.4 |  |  | 1.34
| 1 | 5 | 1.01: 1.3 |  |  | 71
| Mean |  | 1.01: 1.34
| Median |  | 1.01: 1.3
| Std Dev. |  | 1.01: 0.054772
| Rel.Std.Dev. |  | 1.01: 4.087%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ge in OREAS 134a (Indicative Value 475 ppb) | [16 empty] | Indicative
Ge | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppb
|  |  | 1.01: AR*MS |  |  | 0
|  |  | 1.01: 0.5g |  |  | 0
| 1 | 1 | 1.01: 500 |  |  | 1
| 1 | 2 | 1.01: 400 |  |  | 47
| 1 | 3 | 1.01: 500 |  |  | 16
| 1 | 4 | 1.01: 500 |  |  | 475
| 1 | 5 | 1.01: 1300 |  |  | 72
| Mean |  | 1.01: 640
| Median |  | 1.01: 500
| Std Dev. |  | 1.01: 371.4835
| Rel.Std.Dev. |  | 1.01: 58.04%
| PDM3 |  | 1.01: 34.74%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Hf in OREAS 134a (Indicative Value 280 ppb) | [16 empty] | Indicative
Hf | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppb
|  |  | 1.01: AR*MS |  |  | 0
|  |  | 1.01: 0.5g |  |  | 0
| 1 | 1 | 1.01: 300 |  |  | 1
| 1 | 2 | 1.01: 300 |  |  | 48
| 1 | 3 | 1.01: 300 |  |  | 16
| 1 | 4 | 1.01: 300 |  |  | 280
| 1 | 5 | 1.01: 200 |  |  | 73
| Mean |  | 1.01: 280
| Median |  | 1.01: 300
| Std Dev. |  | 1.01: 44.7214
| Rel.Std.Dev. |  | 1.01: 15.97%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Hg in OREAS 134a (Indicative Value 2260 ppb) | [16 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.02: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  |  | ppb
|  |  | 1.01: AR*OES/MS; 1.02: AR*OES |  |  | 0
|  |  | 1.01: 01g; 1.02: 0.5g |  |  | 0
| 1 | 1 | 1.01: 2100; 1.02: 3000 |  |  | 1
| 1 | 2 | 1.01: 2210; 1.02: 2000 |  |  | 49
| 1 | 3 | 1.01: 2010; 1.02: 2000 |  |  | 16
| 1 | 4 | 1.01: 2120; 1.02: 3000 |  |  | 2260
| 1 | 5 | 1.01: 2160; 1.02: 2000 |  |  | 74
| Mean |  | 1.01: 2120; 1.02: 2400
| Median |  | 1.01: 2120; 1.02: 2000
| Std Dev. |  | 1.01: 74.4983; 1.02: 547.7226
| Rel.Std.Dev. |  | 1.01: 3.514%; 1.02: 22.82%
| PDM3 |  | 1.01: -6.195%; 1.02: 6.195%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ho in OREAS 134a (Indicative Value 200 ppb) | [16 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppb
|  |  | 1.01: AR*MS |  |  | 0
|  |  | 1.01: 0.5g |  |  | 0
| 1 | 1 | 1.01: 200 |  |  | 1
| 1 | 2 | 1.01: 200 |  |  | 50
| 1 | 3 | 1.01: 200 |  |  | 16
| 1 | 4 | 1.01: 200 |  |  | 200
| 1 | 5 | 1.01: 200 |  |  | 75
| Mean |  | 1.01: 200
| Median |  | 1.01: 200
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for In in OREAS 134a (Indicative Value 1.2 ppm) | [16 empty] | Indicative
In | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 2
|  |  | 1.01: 0.5g |  |  | 2
| 1 | 1 | 1.01: 1.19 |  |  | 1
| 1 | 2 | 1.01: 1.18 |  |  | 51
| 1 | 3 | 1.01: 1.19 |  |  | 16
| 1 | 4 | 1.01: 1.22 |  |  | 1.202
| 1 | 5 | 1.01: 1.23 |  |  | 59
| Mean |  | 1.01: 1.202
| Median |  | 1.01: 1.19
| Std Dev. |  | 1.01: 0.021679
| Rel.Std.Dev. |  | 1.01: 1.804%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for K in OREAS 134a (Indicative Value 0.15 wt.%) | [16 empty] | Indicative
K | Round | Replicate | 1.01: Lab; 1.02: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  |  | wt.%
|  |  | 1.01: AR*OES/MS; 1.02: AR*OES |  |  | 3
|  |  | 1.01: 01g; 1.02: 0.5g |  |  | 3
| 1 | 1 | 1.01: 0.15; 1.02: 0.15 |  |  | 1
| 1 | 2 | 1.01: 0.15; 1.02: 0.16 |  |  | 32
| 1 | 3 | 1.01: 0.15; 1.02: 0.14 |  |  | 16
| 1 | 4 | 1.01: 0.15; 1.02: 0.14 |  |  | 0.15
| 1 | 5 | 1.01: 0.16; 1.02: 0.15 |  |  | 60
| Mean |  | 1.01: 0.152; 1.02: 0.148
| Median |  | 1.01: 0.15; 1.02: 0.15
| Std Dev. |  | 1.01: 0.0044721; 1.02: 0.0083666
| Rel.Std.Dev. |  | 1.01: 2.942%; 1.02: 5.653%
| PDM3 |  | 1.01: 1.333%; 1.02: -1.333%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for La in OREAS 134a (Indicative Value 6.54 ppm) | [16 empty] | Indicative
La | Round | Replicate | 1.01: Lab; 1.02: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS |  |  | 2
|  |  | 1.01: 01g; 1.02: 0.5g |  |  | 2
| 1 | 1 | 1.01: 7.5; 1.02: 5.6 |  |  | 1
| 1 | 2 | 1.01: 7.6; 1.02: 1.6 |  |  | 16
| 1 | 3 | 1.01: 7.2; 1.02: 5.9 |  |  | 16
| 1 | 4 | 1.01: 7.3; 1.02: 5.7 |  |  | 6.54
| 1 | 5 | 1.01: 7.8; 1.02: 5.2 |  |  | 61
| Mean |  | 1.01: 7.48; 1.02: 4.8
| Median |  | 1.01: 7.5; 1.02: 5.6
| Std Dev. |  | 1.01: 0.23875; 1.02: 1.8069
| Rel.Std.Dev. |  | 1.01: 3.192%; 1.02: 37.64%
| PDM3 |  | 1.01: 14.37%; 1.02: -26.61%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Li in OREAS 134a (Indicative Value 15.3 ppm) | [16 empty] | Indicative
Li | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 1
|  |  | 1.01: 0.5g |  |  | 1
| 1 | 1 | 1.01: 14.6 |  |  | 1
| 1 | 2 | 1.01: 14.4 |  |  | 38
| 1 | 3 | 1.01: 15.1 |  |  | 16
| 1 | 4 | 1.01: 15.1 |  |  | 15.28
| 1 | 5 | 1.01: 17.2 |  |  | 62
| Mean |  | 1.01: 15.28
| Median |  | 1.01: 15.1
| Std Dev. |  | 1.01: 1.1167
| Rel.Std.Dev. |  | 1.01: 7.308%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Lu in OREAS 134a (Indicative Value < 100 ppb) | [16 empty] | Indicative
Lu | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppb
|  |  | 1.01: AR*MS |  |  | 0
|  |  | 1.01: 0.5g |  |  | 0
| 1 | 1 | 1.01: < 100 |  |  | 1
| 1 | 2 | 1.01: < 100 |  |  | 17
| 1 | 3 | 1.01: < 100 |  |  | 16
| 1 | 4 | 1.01: < 100 |  |  | < 100
| 1 | 5 | 1.01: < 100 |  |  | 63
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for MgO in OREAS 134a (Certified Value 3.43 wt.%) | [16 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.2g | 0.25g |  |  | 3
| 1 | 1 | 3.386 | 3.134 | 3.499 | 2.9682 | 3.515 | 3.416 | 3.366 | 3.515 | 3.3 | 3.615 | 3.001 | 3.482 | 3.416 |  |  | 1
| 1 | 2 | 3.4308 | 3.167 | 3.565 | 2.9184 | 3.482 | 3.399 | 3.35 | 3.449 | 3.316 | 3.565 | 3.001 | 3.383 | 3.432 |  |  | 3
| 1 | 3 | 3.386 | 3.051 | 3.582 | 2.9184 | 3.615 | 3.333 | 3.316 | 3.466 | 3.35 | 3.615 | 3.001 | 3.3 | 3.432 |  |  | 16
| 1 | 4 | 3.386 | 3.184 | 3.615 | 3.0179 | 3.598 | 3.3 | 3.35 | 3.499 | 3.333 | 3.648 | 2.985 | 3.366 | 3.449 |  |  | 3.425
| 1 | 5 | 3.4176 | 3.3 | 3.631 | 3.134 | 3.565 | 3.35 | 3.233 | 3.515 | 3.383 | 3.714 | 2.985 | 3.532 | 3.549 |  |  | 41
| Mean |  | 3.4013 | 3.1672 | 3.5784 | 2.9914 | 3.555 | 3.3596 | 3.323 | 3.4888 | 3.3364 | 3.6314 | 2.9946 | 3.4126 | 3.4556
| Median |  | 3.386 | 3.167 | 3.582 | 2.9682 | 3.565 | 3.35 | 3.35 | 3.499 | 3.333 | 3.615 | 3.001 | 3.383 | 3.432
| Std Dev. |  | 0.021437 | 0.09017 | 0.051486 | 0.089768 | 0.055852 | 0.047658 | 0.053516 | 0.02992 | 0.032052 | 0.054875 | 0.0087636 | 0.093273 | 0.0535
| Rel.Std.Dev. |  | 0.6303% | 2.847% | 1.439% | 3.001% | 1.571% | 1.419% | 1.61% | 0.8576% | 0.9607% | 1.511% | 0.2926% | 2.733% | 1.548%
| PDM3 |  | -0.6938% | -7.528% | 4.478% | -12.66% | 3.794% | -1.911% | -2.979% | 1.862% | -2.588% | 6.025% | -12.57% | -0.3633% | 0.8922%
| Z-Score (Absolute) |  | 0 | 1.8 | 1.36 | 3.16 | 1.19 | 0.32 | 0.6 | 0.67 | 0.5 | 1.77 | 3.13 | 0.09 | 0.42
| Analytical results for Mn in OREAS 134a (Indicative Value 3854 ppm) | [16 empty] | Indicative
Mn | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.03: AR*MS |  |  | 0
|  |  | 1.01: 01g; 1.03: 0.5g |  |  | 0
| 1 | 1 | 1.01: 3795; 1.03: 3880 |  |  | 1
| 1 | 2 | 1.01: 3678; 1.03: 3880 |  |  | 34
| 1 | 3 | 1.01: 3741; 1.03: 4000 |  |  | 16
| 1 | 4 | 1.01: 3828; 1.03: 3950 |  |  | 3854.1
| 1 | 5 | 1.01: 3729; 1.03: 4060 |  |  | 64
| Mean |  | 1.01: 3754.2; 1.03: 3954
| Median |  | 1.01: 3741; 1.03: 3950
| Std Dev. |  | 1.01: 58.5807; 1.03: 77.9744
| Rel.Std.Dev. |  | 1.01: 1.56%; 1.03: 1.972%
| PDM3 |  | 1.01: -2.592%; 1.03: 2.592%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for Mo in OREAS 134a (Indicative Value 6.8 ppm) | [16 empty] | Indicative
Mo | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.03: AR*MS |  |  | 2
|  |  | 1.01: 01g; 1.03: 0.5g |  |  | 2
| 1 | 1 | 1.01: 7.2; 1.03: 6.41 |  |  | 1
| 1 | 2 | 1.01: 7.3; 1.03: 6.36 |  |  | 18
| 1 | 3 | 1.01: 6.6; 1.03: 6.52 |  |  | 16
| 1 | 4 | 1.01: 7.2; 1.03: 6.5 |  |  | 6.8
| 1 | 5 | 1.01: 7.3; 1.03: 6.61 |  |  | 65
| Mean |  | 1.01: 7.12; 1.03: 6.48
| Median |  | 1.01: 7.2; 1.03: 6.5
| Std Dev. |  | 1.01: 0.29496; 1.03: 0.097724
| Rel.Std.Dev. |  | 1.01: 4.143%; 1.03: 1.508%
| PDM3 |  | 1.01: 4.706%; 1.03: -4.706%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for Na in OREAS 134a (Indicative Value 161 ppm) | [16 empty] | Indicative
Na | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.03: AR*OES |  |  | 0
|  |  | 1.01: 01g; 1.03: 0.5g |  |  | 0
| 1 | 1 | 1.01: 100; 1.03: 230 |  |  | 1
| 1 | 2 | 1.01: 100; 1.03: 230 |  |  | 29
| 1 | 3 | 1.01: 100; 1.03: 210 |  |  | 16
| 1 | 4 | 1.01: 100; 1.03: 210 |  |  | 161
| 1 | 5 | 1.01: 100; 1.03: 230 |  |  | 66
| Mean |  | 1.01: 100; 1.03: 222
| Median |  | 1.01: 100; 1.03: 230
| Std Dev. |  | 1.01: 0; 1.03: 10.9545
| Rel.Std.Dev. |  | 1.01: 0%; 1.03: 4.934%
| PDM3 |  | 1.01: -37.89%; 1.03: 37.89%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for Nb in OREAS 134a (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Nb | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 2
|  |  | 1.01: 0.5g |  |  | 2
| 1 | 1 | 1.01: < 0.1 |  |  | 1
| 1 | 2 | 1.01: < 0.1 |  |  | 19
| 1 | 3 | 1.01: < 0.1 |  |  | 16
| 1 | 4 | 1.01: < 0.1 |  |  | < 0.1
| 1 | 5 | 1.01: 0.1 |  |  | 67
| Mean |  | 1.01: 0.1
| Median |  | 1.01: 0.1
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for Nd in OREAS 134a (Indicative Value 6.37 ppm) | [16 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.12 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 2
|  |  | 1.01: 0.5g |  |  | 2
| 1 | 1 | 1.01: 6.37 |  |  | 1
| 1 | 2 | 1.01: 5.84 |  |  | 20
| 1 | 3 | 1.01: 6.75 |  |  | 16
| 1 | 4 | 1.01: 6.58 |  |  | 6.366
| 1 | 5 | 1.01: 6.29 |  |  | 68
| Mean |  | 1.01: 6.366
| Median |  | 1.01: 6.37
| Std Dev. |  | 1.01: 0.34486
| Rel.Std.Dev. |  | 1.01: 5.417%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ni in OREAS 134a (Indicative Value 21.6 ppm) | [16 empty] | Indicative
Ni | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.03: AR*MS |  |  | 1
|  |  | 1.01: 01g; 1.03: 0.5g |  |  | 1
| 1 | 1 | 1.01: 21.5; 1.03: 21.3 |  |  | 1
| 1 | 2 | 1.01: 21.3; 1.03: 21.1 |  |  | 21
| 1 | 3 | 1.01: 21; 1.03: 21.9 |  |  | 16
| 1 | 4 | 1.01: 22.7; 1.03: 21.9 |  |  | 21.64
| 1 | 5 | 1.01: 21.5; 1.03: 22.2 |  |  | 69
| Mean |  | 1.01: 21.6; 1.03: 21.68
| Median |  | 1.01: 21.5; 1.03: 21.9
| Std Dev. |  | 1.01: 0.64807; 1.03: 0.46043
| Rel.Std.Dev. |  | 1.01: 3%; 1.03: 2.124%
| PDM3 |  | 1.01: -0.1848%; 1.03: 0.1848%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for P in OREAS 134a (Indicative Value 281 ppm) | [16 empty] | Indicative
P | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.12 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.03: AR*OES |  |  | 0
|  |  | 1.01: 01g; 1.03: 0.5g |  |  | 0
| 1 | 1 | 1.01: 300; 1.03: 250 |  |  | 1
| 1 | 2 | 1.01: 320; 1.03: 260 |  |  | 30
| 1 | 3 | 1.01: 300; 1.03: 240 |  |  | 16
| 1 | 4 | 1.01: 300; 1.03: 240 |  |  | 281
| 1 | 5 | 1.01: 350; 1.03: 250 |  |  | 70
| Mean |  | 1.01: 314; 1.03: 248
| Median |  | 1.01: 300; 1.03: 250
| Std Dev. |  | 1.01: 21.9089; 1.03: 8.3666
| Rel.Std.Dev. |  | 1.01: 6.977%; 1.03: 3.374%
| PDM3 |  | 1.01: 11.74%; 1.03: -11.74%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for Pb in OREAS 134a (Certified Value 12.64 wt.%) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | wt.%
|  |  | AR*OES | MAR*OES | AR*OES | AR*OES | 3A*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*OES | MAR*OES | AR*MS | AR*OES/AAS | AR*OES | AR*OES |  | 2
|  |  | 01g |  | 01g | 01g |  | 0.4g | 0.5g | 01g | 01g |  | 0.5g | 0.4g | 0.2g | 0.25g |  | 3
| 1 | 1 | 11.676 | 14.84 | 2.44 | 11.7296 | > 10 | 12.6 | 12.55 | 14 | 11.2 | 12.8173 | > 1 | 12.3 | 12.3 | 13.1 |  | 1
| 1 | 2 | 11.652 | 14.77 | 2.46 | 12.5365 | > 10 | 12.75 | 12.45 | 14 | 11.9 | 13.2141 | > 1 | 13 | 12 | 13.1 |  | 22
| 1 | 3 | 11.882 | 14.94 | 2.4 | 12.547 | > 10 | 13.05 | 12.5 | 13.9757 | 11.7 | 13.3421 | > 1 | 12.4 | 11.7 | 13.3 |  | 16
| 1 | 4 | 11.529 | 14.32 | 2.42 | 13.2844 | > 10 | 13.2 | 12.4 | 14.105 | 11.5 | 13.2166 | > 1 | 13.1 | 11.9 | 13.4 |  | 12.6375
| 1 | 5 | 12.306 | 14.81 | 2.35 | 4.9735 | > 10 | 12.75 | 12.6 | 13.9 | 11.9 | 13.4086 | > 1 | 12.1 | 12.5 | 13.6 |  | 42
| Mean |  | 11.809 | 14.736 | 2.414 | 11.0142 |  | 12.87 | 12.5 | 13.9962 | 11.64 | 13.1997 |  | 12.58 | 12.08 | 13.3
| Median |  | 11.676 | 14.81 | 2.42 | 12.5365 |  | 12.75 | 12.5 | 14 | 11.7 | 13.2166 |  | 12.4 | 12 | 13.3
| Std Dev. |  | 0.3054 | 0.24089 | 0.04219 | 3.4213 |  | 0.24648 | 0.079057 | 0.073392 | 0.29665 | 0.22948 |  | 0.44385 | 0.31937 | 0.21213
| Rel.Std.Dev. |  | 2.586% | 1.635% | 1.748% | 31.06% |  | 1.915% | 0.6325% | 0.5244% | 2.549% | 1.739% |  | 3.528% | 2.644% | 1.595%
| PDM3 |  | -6.556% | 16.61% | -80.9% | -12.85% |  | 1.84% | -1.088% | 10.75% | -7.893% | 4.449% |  | -0.455% | -4.411% | 5.242%
| Z-Score (Absolute) |  | 0.66 | 1.99 | 9.16 | 1.38 | NA | 0.3 | 0.04 | 1.32 | 0.81 | 0.6 | NA | 0.04 | 0.42 | 0.69
| Analytical results for Pr in OREAS 134a (Indicative Value 1.54 ppm) | [16 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*MS |  | 2
|  |  | 1.01: 0.5g |  | 2
| 1 | 1 | 1.01: 1.6 |  | 1
| 1 | 2 | 1.01: 1.3 |  | 23
| 1 | 3 | 1.01: 1.6 |  | 16
| 1 | 4 | 1.01: 1.6 |  | 1.54
| 1 | 5 | 1.01: 1.6 |  | 71
| Mean |  | 1.01: 1.54
| Median |  | 1.01: 1.6
| Std Dev. |  | 1.01: 0.13416
| Rel.Std.Dev. |  | 1.01: 8.712%
| PDM3 |  | 1.01: -1.11e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Rb in OREAS 134a (Indicative Value 11.7 ppm) | [16 empty] | Indicative
Rb | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*MS |  | 1
|  |  | 1.01: 0.5g |  | 1
| 1 | 1 | 1.01: 11.1 |  | 1
| 1 | 2 | 1.01: 11.1 |  | 24
| 1 | 3 | 1.01: 11.3 |  | 16
| 1 | 4 | 1.01: 11.2 |  | 11.7
| 1 | 5 | 1.01: 13.8 |  | 72
| Mean |  | 1.01: 11.7
| Median |  | 1.01: 11.2
| Std Dev. |  | 1.01: 1.1769
| Rel.Std.Dev. |  | 1.01: 10.06%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Re in OREAS 134a (Indicative Value 2.6 ppb) | [16 empty] | Indicative
Re | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppb
|  |  | 1.01: AR*MS |  | 2
|  |  | 1.01: 0.5g |  | 2
| 1 | 1 | 1.01: 2 |  | 1
| 1 | 2 | 1.01: 2 |  | 25
| 1 | 3 | 1.01: 3 |  | 16
| 1 | 4 | 1.01: 2 |  | 2.6
| 1 | 5 | 1.01: 4 |  | 73
| Mean |  | 1.01: 2.6
| Median |  | 1.01: 2
| Std Dev. |  | 1.01: 0.89443
| Rel.Std.Dev. |  | 1.01: 34.4%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for S in OREAS 134a (Certified Value 18.49 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 |  |  |  |  | wt.%
|  |  | MAR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  |  | 2
|  |  |  | 01g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.2g | 0.25g |  |  |  |  | 3
| 1 | 1 | 20.05 | 17.9671 | > 10 | 17.55 | > 10 | 17.2 | 19.7 | 13.6 | 19 | 17 | 19 |  |  |  |  | 1
| 1 | 2 | 19.9 | 18.4424 | > 10 | 16.75 | > 10 | 17.5 | 20.1 | 14.2 | 18.95 | 16.5 | 19.4 |  |  |  |  | 31
| 1 | 3 | 20 | 18.782 | > 10 | 16.7 | > 10 | 16.8 | 20.2 | 13.4 | 18.7 | 16.2 | 19.1 |  |  |  |  | 16
| 1 | 4 | 19.25 | 19.2581 | > 10 | 16.05 | > 10 | 17.9 | 20.1 | 13.5 | 18.8 | 16.4 | 19.6 |  |  |  |  | 18.4925
| 1 | 5 | 19.75 | 18.2007 | > 10 | 17.05 | > 10 | 18 | 20.3 | 13.3 | 18.9 | 17.3 | 20.2 |  |  |  |  | 43
| Mean |  | 19.79 | 18.5301 |  | 16.82 |  | 17.48 | 20.08 | 13.6 | 18.87 | 16.68 | 19.46
| Median |  | 19.9 | 18.4424 |  | 16.75 |  | 17.5 | 20.1 | 13.5 | 18.9 | 16.5 | 19.4
| Std Dev. |  | 0.32288 | 0.50661 |  | 0.54727 |  | 0.49699 | 0.22804 | 0.35355 | 0.12042 | 0.45497 | 0.47749
| Rel.Std.Dev. |  | 1.632% | 2.734% |  | 3.254% |  | 2.843% | 1.136% | 2.6% | 0.6381% | 2.728% | 2.454%
| PDM3 |  | 7.016% | 0.2031% |  | -9.044% |  | -5.475% | 8.585% | -26.46% | 2.041% | -9.801% | 5.232%
| Z-Score (Absolute) |  | 0.67 | 0 | NA | 0.92 | NA | 0.56 | 0.83 | 2.64 | 0.18 | 0.99 | 0.5
| Analytical results for Sb in OREAS 134a (Indicative Value 101 ppm) | [16 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | 0
|  |  | 01g | 01g | 01g | 01g | 0.4g | 01g | 01g | 01g | 01g | 01g | 01g | 0.5g | 01g | 0.4g | 0.25g | 0
| 1 | 1 | 103 | 86.6 | 70 | 90.51 | 134 | 135 | 114 | 113 | 133 | 116 | 88 | 81 | 70 | 84 | 101 | 1
| 1 | 2 | 101 | 86.5 | 75 | 89.43 | 132 | 133 | 115 | 111 | 132.8 | 118 | 89.3 | 83 | 30 | 90 | 97 | 26
| 1 | 3 | 101 | 84.5 | 77 | 91.93 | 129 | 139 | 112 | 121 | 127.4 | 114 | 91.8 | 80 | 80 | 88 | 97 | 16
| 1 | 4 | 103 | 88.7 | 75 | 92.32 | 131 | 138 | 111 | 126 | 132.1 | 117 | 85.7 | 79 | 50 | 84 | 96 | 101.3958
| 1 | 5 | 101 | 90.2 | 67 | 92.52 | 136 | 139 | 116 | 114 | 136.2 | 120 | 88.6 | 104 | 70 | 90 | 88 | 74
| Mean |  | 101.8 | 87.3 | 72.8 | 91.342 | 132.4 | 136.8 | 113.6 | 117 | 132.3 | 117 | 88.68 | 85.4 | 60 | 87.2 | 95.8
| Median |  | 101 | 86.6 | 75 | 91.93 | 132 | 138 | 114 | 114 | 132.8 | 117 | 88.6 | 81 | 70 | 88 | 97
| Std Dev. |  | 1.0954 | 2.1989 | 4.1473 | 1.3265 | 2.7019 | 2.6833 | 2.0736 | 6.2849 | 3.1623 | 2.2361 | 2.2061 | 10.5024 | 20 | 3.0332 | 4.7645
| Rel.Std.Dev. |  | 1.076% | 2.519% | 5.697% | 1.452% | 2.041% | 1.961% | 1.825% | 5.372% | 2.39% | 1.911% | 2.488% | 12.3% | 33.33% | 3.478% | 4.973%
| PDM3 |  | 0.3987% | -13.9% | -28.2% | -9.915% | 30.58% | 34.92% | 12.04% | 15.39% | 30.48% | 15.39% | -12.54% | -15.78% | -40.83% | -14% | -5.519%
| Z-Score (Absolute) |  | 0.23 | 0.32 | 0.87 | 0.17 | 1.39 | 1.55 | 0.67 | 0.8 | 1.38 | 0.8 | 0.27 | 0.39 | 1.36 | 0.33 | 0
| Analytical results for Sc in OREAS 134a (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*OES | 2
|  |  | 1.01: 01g; 1.02: 0.5g | 2
| 1 | 1 | 1.01: < 0.5; 1.02: < 1 | 1
| 1 | 2 | 1.01: < 0.5; 1.02: < 1 | 27
| 1 | 3 | 1.01: < 0.5; 1.02: < 1 | 16
| 1 | 4 | 1.01: < 0.5; 1.02: < 1 | < 0.5
| 1 | 5 | 1.01: < 0.5; 1.02: < 1 | 75
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Se in OREAS 134a (Indicative Value 6.92 ppm) | [16 empty] | Indicative
Se | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS | 2
|  |  | 1.01: 01g; 1.02: 0.5g | 2
| 1 | 1 | 1.01: < 2; 1.02: 6.6 | 1
| 1 | 2 | 1.01: < 2; 1.02: 7 | 51
| 1 | 3 | 1.01: < 2; 1.02: 7.1 | 16
| 1 | 4 | 1.01: < 2; 1.02: 6.8 | 6.92
| 1 | 5 | 1.01: < 2; 1.02: 7.1 | 59
| Mean |  | 1.02: 6.92
| Median |  | 1.02: 7
| Std Dev. |  | 1.02: 0.21679
| Rel.Std.Dev. |  | 1.02: 3.133%
PDM3 | bar: 1.02
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Sm in OREAS 134a (Indicative Value 1.44 ppm) | [16 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: AR*MS | 2
|  |  | 1.01: 0.5g | 2
| 1 | 1 | 1.01: 1.5 | 1
| 1 | 2 | 1.01: 1.3 | 13
| 1 | 3 | 1.01: 1.5 | 16
| 1 | 4 | 1.01: 1.5 | 1.44
| 1 | 5 | 1.01: 1.4 | 60
| Mean |  | 1.01: 1.44
| Median |  | 1.01: 1.5
| Std Dev. |  | 1.01: 0.089443
| Rel.Std.Dev. |  | 1.01: 6.211%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sn in OREAS 134a (Indicative Value 0.93 ppm) | [16 empty] | Indicative
Sn | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: AR*MS | 2
|  |  | 1.01: 0.5g | 2
| 1 | 1 | 1.01: 0.95 | 1
| 1 | 2 | 1.01: 0.92 | 14
| 1 | 3 | 1.01: 0.92 | 16
| 1 | 4 | 1.01: 0.99 | 0.932
| 1 | 5 | 1.01: 0.88 | 61
| Mean |  | 1.01: 0.932
| Median |  | 1.01: 0.92
| Std Dev. |  | 1.01: 0.040866
| Rel.Std.Dev. |  | 1.01: 4.385%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sr in OREAS 134a (Indicative Value 23.8 ppm) | [16 empty] | Indicative
Sr | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS | 1
|  |  | 1.01: 01g; 1.02: 0.5g | 1
| 1 | 1 | 1.01: 31; 1.02: 23.8 | 1
| 1 | 2 | 1.01: 30; 1.02: < 0.5 | 34
| 1 | 3 | 1.01: 31; 1.02: 26.4 | 16
| 1 | 4 | 1.01: 31; 1.02: 23.2 | 23.835
| 1 | 5 | 1.01: 30; 1.02: 11.7 | 62
| Mean |  | 1.01: 30.6; 1.02: 21.275
| Median |  | 1.01: 31; 1.02: 23.5
| Std Dev. |  | 1.01: 0.54772; 1.02: 6.5327
| Rel.Std.Dev. |  | 1.01: 1.79%; 1.02: 30.71%
| PDM3 |  | 1.01: 28.38%; 1.02: -10.74%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ta in OREAS 134a (Indicative Value < 50 ppb) | [16 empty] | Indicative
Ta | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: AR*MS | 1
|  |  | 1.01: 0.5g | 1
| 1 | 1 | 1.01: < 50 | 1
| 1 | 2 | 1.01: < 50 | 16
| 1 | 3 | 1.01: < 50 | 16
| 1 | 4 | 1.01: < 50 | < 50
| 1 | 5 | 1.01: < 50 | 63
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for Tb in OREAS 134a (Indicative Value 200 ppb) | [16 empty] | Indicative
Tb | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: AR*MS | 0
|  |  | 1.01: 0.5g | 0
| 1 | 1 | 1.01: 200 | 1
| 1 | 2 | 1.01: 200 | 17
| 1 | 3 | 1.01: 200 | 16
| 1 | 4 | 1.01: 200 | 200
| 1 | 5 | 1.01: 200 | 64
| Mean |  | 1.01: 200
| Median |  | 1.01: 200
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Te in OREAS 134a (Indicative Value 100 ppb) | [16 empty] | Indicative
Te | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: AR*MS | 0
|  |  | 1.01: 0.5g | 0
| 1 | 1 | 1.01: 100 | 1
| 1 | 2 | 1.01: 100 | 18
| 1 | 3 | 1.01: 100 | 16
| 1 | 4 | 1.01: 90 | 100
| 1 | 5 | 1.01: 110 | 65
| Mean |  | 1.01: 100
| Median |  | 1.01: 100
| Std Dev. |  | 1.01: 7.0711
| Rel.Std.Dev. |  | 1.01: 7.071%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Th in OREAS 134a (Indicative Value 2.18 ppm) | [16 empty] | Indicative
Th | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS | 2
|  |  | 1.01: 01g; 1.02: 0.5g | 2
| 1 | 1 | 1.01: 2.1; 1.02: 2.2 | 1
| 1 | 2 | 1.01: 2.3; 1.02: 1.9 | 19
| 1 | 3 | 1.01: 2.1; 1.02: 2.4 | 16
| 1 | 4 | 1.01: 2.3; 1.02: 2.1 | 2.18
| 1 | 5 | 1.01: 2.2; 1.02: 2.2 | 66
| Mean |  | 1.01: 2.2; 1.02: 2.16
| Median |  | 1.01: 2.2; 1.02: 2.2
| Std Dev. |  | 1.01: 0.1; 1.02: 0.18166
| Rel.Std.Dev. |  | 1.01: 4.545%; 1.02: 8.41%
| PDM3 |  | 1.01: 0.9174%; 1.02: -0.9174%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ti in OREAS 134a (Indicative Value 60 ppm) | [16 empty] | Indicative
Ti | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*OES | 0
|  |  | 1.01: 01g; 1.02: 0.5g | 0
| 1 | 1 | 1.01: 60; 1.02: < 100 | 1
| 1 | 2 | 1.01: 60; 1.02: < 100 | 29
| 1 | 3 | 1.01: 60; 1.02: < 100 | 16
| 1 | 4 | 1.01: 60; 1.02: < 100 | 60
| 1 | 5 | 1.01: 60; 1.02: < 100 | 67
| Mean |  | 1.01: 60
| Median |  | 1.01: 60
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
PDM3 | bar: 1.01
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Tl in OREAS 134a (Indicative Value 67 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*OES | 0
|  |  | 1.01: 01g; 1.02: 0.5g | 0
| 1 | 1 | 1.01: 75.2; 1.02: 58 | 1
| 1 | 2 | 1.01: 80; 1.02: 58 | 20
| 1 | 3 | 1.01: 73; 1.02: 57 | 16
| 1 | 4 | 1.01: 77.6; 1.02: 57 | 66.81
| 1 | 5 | 1.01: 80.3; 1.02: 52 | 68
| Mean |  | 1.01: 77.22; 1.02: 56.4
| Median |  | 1.01: 77.6; 1.02: 57
| Std Dev. |  | 1.01: 3.1324; 1.02: 2.51
| Rel.Std.Dev. |  | 1.01: 4.056%; 1.02: 4.45%
| PDM3 |  | 1.01: 15.58%; 1.02: -15.58%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Tm in OREAS 134a (Indicative Value < 100 ppb) | [16 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: AR*MS | 0
|  |  | 1.01: 0.5g | 0
| 1 | 1 | 1.01: < 100 | 1
| 1 | 2 | 1.01: < 100 | 21
| 1 | 3 | 1.01: < 100 | 16
| 1 | 4 | 1.01: < 100 | < 100
| 1 | 5 | 1.01: < 100 | 69
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for U in OREAS 134a (Indicative Value 0.96 ppm) | [16 empty] | Indicative
U | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS | 2
|  |  | 1.01: 01g; 1.02: 0.5g | 2
| 1 | 1 | 1.01: 1; 1.02: 0.9 | 1
| 1 | 2 | 1.01: 1; 1.02: 0.9 | 22
| 1 | 3 | 1.01: 1; 1.02: 0.9 | 16
| 1 | 4 | 1.01: 1; 1.02: 0.9 | 0.96
| 1 | 5 | [15 empty] | 70 | bar: 1.01 to 1.02
| Mean |  | 1.01: 1; 1.02: 0.92
| Median |  | 1.01: 1; 1.02: 0.9
| Std Dev. |  | 1.01: 0; 1.02: 0.044721
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 4.861%
| PDM3 |  | 1.01: 4.167%; 1.02: -4.167%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for V in OREAS 134a (Indicative Value 7.4 ppm) | [16 empty] | Indicative
V | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS | 2
|  |  | 1.01: 01g; 1.02: 0.5g | 2
| 1 | 1 | 1.01: < 10; 1.02: 7 | 1
| 1 | 2 | 1.01: < 10; 1.02: 7 | 23
| 1 | 3 | 1.01: < 10; 1.02: 7 | 16
| 1 | 4 | 1.01: < 10; 1.02: 7 | 7.4
| 1 | 5 | 1.01: < 10; 1.02: 9 | 71
| Mean |  | 1.02: 7.4
| Median |  | 1.02: 7
| Std Dev. |  | 1.02: 0.89443
| Rel.Std.Dev. |  | 1.02: 12.09%
PDM3 | bar: 1.02
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for W in OREAS 134a (Indicative Value 0.24 ppm) | [16 empty] | Indicative
W | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.12 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS | 2
|  |  | 1.01: 01g; 1.02: 0.5g | 2
| 1 | 1 | 1.01: < 0.5; 1.02: 0.3 | 1
| 1 | 2 | 1.01: < 0.5; 1.02: 0.2 | 24
| 1 | 3 | 1.01: < 0.5; 1.02: 0.2 | 16
| 1 | 4 | 1.01: < 0.5; 1.02: 0.3 | 0.24
| 1 | 5 | 1.01: < 0.5; 1.02: 0.2 | 72
| Mean |  | 1.02: 0.24
| Median |  | 1.02: 0.2
| Std Dev. |  | 1.02: 0.054772
| Rel.Std.Dev. |  | 1.02: 22.82%
PDM3 | bar: 1.02
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Y in OREAS 134a (Indicative Value 3 ppm) | [16 empty] | Indicative
Y | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppm
|  |  | 1.01: AR*OES | 2
|  |  | 1.01: 0.5g | 2
| 1 | 1 | 1.01: 3 | 1
| 1 | 2 | 1.01: 3 | 25
| 1 | 3 | 1.01: 3 | 16
| 1 | 4 | 1.01: 3 | 3
| 1 | 5 | 1.01: 3 | 73
| Mean |  | 1.01: 3
| Median |  | 1.01: 3
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Yb in OREAS 134a (Indicative Value 480 ppb) | [16 empty] | Indicative
Yb | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.12 | ppb
|  |  | 1.01: AR*MS | 0
|  |  | 1.01: 0.5g | 0
| 1 | 1 | 1.01: 500 | 1
| 1 | 2 | 1.01: 500 | 26
| 1 | 3 | 1.01: 500 | 16
| 1 | 4 | 1.01: 500 | 480
| 1 | 5 | 1.01: 400 | 74
| Mean |  | 1.01: 480
| Median |  | 1.01: 500
| Std Dev. |  | 1.01: 44.7214
| Rel.Std.Dev. |  | 1.01: 9.317%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Zn in OREAS 134a (Certified Value 17 wt.%) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 |  | wt.%
|  |  | AR*OES | MAR*OES | AR*OES | AR*OES | 3A*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*OES | MAR*OES | AR*MS | AR*OES/AAS | AR*OES | AR*OES |  | 2
|  |  | 01g |  | 01g | 01g |  | 0.4g | 0.5g | 01g | 01g |  | 0.5g | 0.4g | 0.2g | 0.25g |  | 3
| 1 | 1 | 15.929 | 13.04 | 13 | 16.0493 | 17.1 | 17.45 | 17.15 | 18.15 | 15.9 | 16.6993 | > 1 | 16.75 | 16.5 | 17.5 |  | 1
| 1 | 2 | 15.984 | 12.76 | 13.2 | 16.3718 | 17.2 | 17.45 | 17 | 18.3484 | 16.1 | 17.0415 | > 1 | 17.5 | 16.1 | 17.4 |  | 0
| 1 | 3 | 15.792 | 13.01 | 13.4 | 16.3434 | 17.2 | 17.9 | 17 | 18.6654 | 16.1 | 17.272 | > 1 | 16.4 | 15.7 | 17.8 |  | 16
| 1 | 4 | 15.213 | 12.67 | 13.3 | 16.7174 | 17.7 | 17.95 | 16.9 | 18.9967 | 16.1 | 17.1605 | > 1 | 17.6 | 16 | 17.7 |  | 16.9999
| 1 | 5 | 15.434 | 12.79 | 13.1 | 16.8278 | 17.6 | 17.7 | 17.15 | 18.4531 | 16.2 | 17.461 | > 1 | 16.65 | 16.8 | 18.3 |  | 44
| Mean |  | 15.6704 | 12.854 | 13.2 | 16.4619 | 17.36 | 17.69 | 17.04 | 18.5227 | 16.08 | 17.1269 |  | 16.98 | 16.22 | 17.74
| Median |  | 15.792 | 12.79 | 13.2 | 16.3718 | 17.2 | 17.7 | 17 | 18.4531 | 16.1 | 17.1605 |  | 16.75 | 16.1 | 17.7
| Std Dev. |  | 0.33358 | 0.16257 | 0.15811 | 0.31287 | 0.27019 | 0.23822 | 0.1084 | 0.32372 | 0.10954 | 0.28458 |  | 0.53689 | 0.43243 | 0.35071
| Rel.Std.Dev. |  | 2.129% | 1.265% | 1.198% | 1.901% | 1.556% | 1.347% | 0.6361% | 1.748% | 0.6812% | 1.662% |  | 3.162% | 2.666% | 1.977%
| PDM3 |  | -7.821% | -24.39% | -22.35% | -3.164% | 2.118% | 4.059% | 0.2359% | 8.958% | -5.411% | 0.7469% |  | -0.117% | -4.588% | 4.354%
| Z-Score (Absolute) |  | 1.16 | 3.66 | 3.35 | 0.46 | 0.34 | 0.63 | 0.05 | 1.37 | 0.8 | 0.13 | NA | 0 | 0.67 | 0.67
| Analytical results for Zr in OREAS 134a (Indicative Value 10.4 ppm) | [16 empty] | Indicative
Zr | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.12 |  | ppm
|  |  | 1.01: AR*OES |  | 1
|  |  | 1.01: 0.5g |  | 1
| 1 | 1 | 1.01: 10 |  | 1
| 1 | 2 | 1.01: 11 |  | 27
| 1 | 3 | 1.01: 10 |  | 16
| 1 | 4 | 1.01: 10 |  | 10.4
| 1 | 5 | 1.01: 11 |  | 75
| Mean |  | 1.01: 10.4
| Median |  | 1.01: 10
| Std Dev. |  | 1.01: 0.54772
| Rel.Std.Dev. |  | 1.01: 5.267%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C33: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C50: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C84: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C322: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C339: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C356: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C373: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C390: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C407: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C424: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C628: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C645: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C866: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C883: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C934: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for S in OREAS 134a (Certified Value 19.11 wt.%) | [12 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 19 | 18.704 | 18.5 | 25.3 | 19.15 | 18.95 | 19.2 | 13.64 | 19.7 | 19 | 19.5 | 1
| 1 | 2 | 19.1 | 18.724 | 18.4 | 25.4 | 19.15 | 19.65 | 19.3 | 14.41 | 19.7 | 18.9 | 19.5 | 31
| 1 | 3 | 19.2 | 18.668 | 18.5 | 25.4 | 18.6 | 19.6 | 19.3 | 14.25 | 19.7 | 19.1 | 19.4 | 16
| 1 | 4 | 19.2 | 18.631 | 18.6 | 25.4 | 18.8 | 19 | 19.3 | 17.49 | 19.9 | 18.7 | 19.4 | 19.1121
| 1 | 5 | 19.5 | 18.618 | 18.6 | 25.4 | 18.7 | 19 | 19.2 | 17.38 | 19.9 | 19.1 | 19.7 | 46
| Mean |  | 19.2 | 18.669 | 18.52 | 25.38 | 18.88 | 19.24 | 19.26 | 15.434 | 19.78 | 18.96 | 19.5
| Median |  | 19.2 | 18.668 | 18.5 | 25.4 | 18.8 | 19 | 19.3 | 14.41 | 19.7 | 19 | 19.5
| Std Dev. |  | 0.18708 | 0.045541 | 0.083666 | 0.044721 | 0.25642 | 0.35249 | 0.054772 | 1.8495 | 0.10954 | 0.16733 | 0.12247
| Rel.Std.Dev. |  | 0.9744% | 0.2439% | 0.4518% | 0.1762% | 1.358% | 1.832% | 0.2844% | 11.98% | 0.5538% | 0.8826% | 0.6281%
| PDM3 |  | 0.4599% | -2.318% | -3.098% | 32.8% | -1.214% | 0.6692% | 0.7738% | -19.24% | 3.495% | -0.7959% | 2.03%
| Z-Score (Absolute) |  | 0 | 1.12 | 1.43 | 13.02 | 0.67 | 0.08 | 0.13 | 7.94 | 1.22 | 0.51 | 0.63
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.