## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 25b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 1.5954 | 0.85697 | 0 | 3.3093 | 0 | 4.1663 | 53.72% | 107.4% | 161.1% | 1.5156 | 1.6752
Pd, ppb | 1.4198 | 0.33092 | 0.75799 | 2.0817 | 0.42706 | 2.4126 | 23.31% | 46.61% | 69.92% | 1.3488 | 1.4908
Pt, ppb | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Borate Fusion XRF
Al2O3, wt.% | 16.8502 | 0.29852 | 16.2531 | 17.4472 | 15.9546 | 17.7457 | 1.772% | 3.543% | 5.315% | 16.0076 | 17.6927
BaO, ppm | 214.0126 | 28.1335 | 157.7457 | 270.2796 | 129.6122 | 298.4131 | 13.15% | 26.29% | 39.44% | 203.312 | 224.7133
CaO, wt.% | 0.35606 | 0.012496 | 0.33107 | 0.38105 | 0.31858 | 0.39355 | 3.509% | 7.019% | 10.53% | 0.33826 | 0.37387
Cr2O3, ppm | 295.8967 | 45.3592 | 205.1783 | 386.615 | 159.8192 | 431.9742 | 15.33% | 30.66% | 45.99% | 281.1018 | 310.6915
Fe2O3, wt.% | 12.2027 | 0.17156 | 11.8596 | 12.5459 | 11.6881 | 12.7174 | 1.406% | 2.812% | 4.218% | 11.5926 | 12.8129
K2O, wt.% | 0.93575 | 0.01858 | 0.89859 | 0.97291 | 0.88001 | 0.99149 | 1.986% | 3.971% | 5.957% | 0.88896 | 0.98254
MgO, wt.% | 0.62689 | 0.014029 | 0.59883 | 0.65495 | 0.58481 | 0.66898 | 2.238% | 4.476% | 6.714% | 0.59555 | 0.65824
MnO, wt.% | 0.030444 | 0.0036375 | 0.023169 | 0.037719 | 0.019531 | 0.041356 | 11.95% | 23.9% | 35.85% | 0.028922 | 0.031966
Na2O, wt.% | 0.093651 | 0.012665 | 0.068321 | 0.11898 | 0.055657 | 0.13164 | 13.52% | 27.05% | 40.57% | 0.088968 | 0.098333
P2O5, wt.% | 0.10042 | 0.0036337 | 0.093157 | 0.10769 | 0.089523 | 0.11133 | 3.618% | 7.237% | 10.85% | 0.095403 | 0.10545
SiO2, wt.% | 58.0934 | 0.30023 | 57.4929 | 58.6938 | 57.1927 | 58.9941 | 0.5168% | 1.034% | 1.55% | 55.1887 | 60.9981
SO3, wt.% | 0.058113 | 0.008483 | 0.041147 | 0.075079 | 0.032664 | 0.083562 | 14.6% | 29.2% | 43.79% | 0.055207 | 0.061019
SrO, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
TiO2, wt.% | 1.8783 | 0.019944 | 1.8384 | 1.9182 | 1.8185 | 1.9381 | 1.062% | 2.124% | 3.185% | 1.7844 | 1.9722
Thermogravimetry
LOI1000, wt.% | 8.4942 | 0.19474 | 8.1047 | 8.8837 | 7.91 | 9.0784 | 2.293% | 4.585% | 6.878% | 8.0695 | 8.9189
Borate / Peroxide Fusion ICP
Al, wt.% | 8.7314 | 0.22606 | 8.2793 | 9.1836 | 8.0533 | 9.4096 | 2.589% | 5.178% | 7.767% | 8.2949 | 9.168
Ba, ppm | 210.4738 | 8.3299 | 193.814 | 227.1336 | 185.4841 | 235.4636 | 3.958% | 7.915% | 11.87% | 199.9501 | 220.9975
Be, ppm | 1.5187 | 0.51557 | 0.48755 | 2.5498 | 0 | 3.0654 | 33.95% | 67.9% | 101.8% | 1.4428 | 1.5946
Ca, wt.% | 0.29034 | 0.044765 | 0.20081 | 0.37987 | 0.15604 | 0.42464 | 15.42% | 30.84% | 46.25% | 0.27582 | 0.30486
Ce, ppm | 52.6057 | 2.3281 | 47.9494 | 57.2619 | 45.6213 | 59.59 | 4.426% | 8.851% | 13.28% | 49.9754 | 55.2359
Co, ppm | 9.9147 | 0.87422 | 8.1663 | 11.6632 | 7.292 | 12.5374 | 8.817% | 17.63% | 26.45% | 9.419 | 10.4105
Cr, ppm | 219.1976 | 12.2273 | 194.7429 | 243.6523 | 182.5156 | 255.8796 | 5.578% | 11.16% | 16.73% | 208.2377 | 230.1575
Cs, ppm | 4.5651 | 0.17205 | 4.221 | 4.9092 | 4.0489 | 5.0812 | 3.769% | 7.537% | 11.31% | 4.3368 | 4.7933
Cu, ppm | 41.1978 | 9.0221 | 23.1535 | 59.2421 | 14.1314 | 68.2643 | 21.9% | 43.8% | 65.7% | 39.1379 | 43.2577
Dy, ppm | 3.6129 | 0.20411 | 3.2047 | 4.0211 | 3.0005 | 4.2252 | 5.65% | 11.3% | 16.95% | 3.4322 | 3.7935
Er, ppm | 2.16 | 0.14356 | 1.8729 | 2.4471 | 1.7293 | 2.5907 | 6.646% | 13.29% | 19.94% | 2.052 | 2.268
Eu, ppm | 0.88737 | 0.070722 | 0.74593 | 1.0288 | 0.6752 | 1.0995 | 7.97% | 15.94% | 23.91% | 0.843 | 0.93174
Fe, wt.% | 8.4388 | 0.26795 | 7.9029 | 8.9747 | 7.6349 | 9.2426 | 3.175% | 6.35% | 9.526% | 8.0168 | 8.8607
Ga, ppm | 25.8618 | 1.1363 | 23.5892 | 28.1343 | 22.4529 | 29.2706 | 4.394% | 8.787% | 13.18% | 24.5687 | 27.1549
Gd, ppm | 3.6567 | 0.13539 | 3.3859 | 3.9275 | 3.2505 | 4.0628 | 3.703% | 7.405% | 11.11% | 3.4738 | 3.8395
Ge, ppm | 2.0306 | 0.4355 | 1.1596 | 2.9016 | 0.72407 | 3.3371 | 21.45% | 42.89% | 64.34% | 1.9291 | 2.1321
Hf, ppm | 8.4652 | 0.39765 | 7.6699 | 9.2605 | 7.2723 | 9.6582 | 4.697% | 9.395% | 14.09% | 8.042 | 8.8885
Ho, ppm | 0.71738 | 0.038067 | 0.64125 | 0.79351 | 0.60318 | 0.83158 | 5.306% | 10.61% | 15.92% | 0.68151 | 0.75325
K, wt.% | 0.80533 | 0.058207 | 0.68891 | 0.92174 | 0.63071 | 0.97995 | 7.228% | 14.46% | 21.68% | 0.76506 | 0.84559
La, ppm | 27.9098 | 1.0655 | 25.7788 | 30.0409 | 24.7132 | 31.1065 | 3.818% | 7.636% | 11.45% | 26.5144 | 29.3053
Li, ppm | 30.202 | 4.5266 | 21.1488 | 39.2553 | 16.6221 | 43.7819 | 14.99% | 29.98% | 44.96% | 28.6919 | 31.7121
Lu, ppm | 0.3156 | 0.028911 | 0.25778 | 0.37342 | 0.22887 | 0.40233 | 9.161% | 18.32% | 27.48% | 0.29982 | 0.33138
Mg, wt.% | 0.37737 | 0.0071792 | 0.36301 | 0.39172 | 0.35583 | 0.3989 | 1.902% | 3.805% | 5.707% | 0.3585 | 0.39623
Mn, wt.% | 0.02419 | 0.0013902 | 0.02141 | 0.026971 | 0.02002 | 0.028361 | 5.747% | 11.49% | 17.24% | 0.022981 | 0.0254
Na, wt.% | 0.065559 | 0.016072 | 0.033415 | 0.097703 | 0.017343 | 0.11377 | 24.52% | 49.03% | 73.55% | 0.062281 | 0.068837
Nb, ppm | 27.6233 | 1.7253 | 24.1727 | 31.0739 | 22.4474 | 32.7991 | 6.246% | 12.49% | 18.74% | 26.2421 | 29.0044
Nd, ppm | 22.4515 | 1.5029 | 19.4456 | 25.4574 | 17.9427 | 26.9604 | 6.694% | 13.39% | 20.08% | 21.3289 | 23.5741
Ni, ppm | 68.5186 | 7.0551 | 54.4084 | 82.6288 | 47.3533 | 89.6839 | 10.3% | 20.59% | 30.89% | 65.0927 | 71.9445
P, wt.% | 0.040369 | 0.0026699 | 0.035029 | 0.045709 | 0.032359 | 0.048379 | 6.614% | 13.23% | 19.84% | 0.038351 | 0.042388
Pr, ppm | 6.0366 | 0.38622 | 5.2642 | 6.8091 | 4.878 | 7.1953 | 6.398% | 12.8% | 19.19% | 5.7348 | 6.3385
Rb, ppm | 60.8711 | 3.0555 | 54.76 | 66.9822 | 51.7045 | 70.0377 | 5.02% | 10.04% | 15.06% | 57.8276 | 63.9147
Sc, ppm | 16.1571 | 1.7457 | 12.6657 | 19.6484 | 10.9201 | 21.3941 | 10.8% | 21.61% | 32.41% | 15.3492 | 16.9649
Si, wt.% | 26.1678 | 0.65046 | 24.8669 | 27.4687 | 24.2165 | 28.1192 | 2.486% | 4.971% | 7.457% | 24.8594 | 27.4762
Sm, ppm | 4.1861 | 0.35807 | 3.4699 | 4.9022 | 3.1119 | 5.2603 | 8.554% | 17.11% | 25.66% | 3.9768 | 4.3954
Sn, ppm | 7.1343 | 0.47173 | 6.1909 | 8.0778 | 5.7191 | 8.5495 | 6.612% | 13.22% | 19.84% | 6.7776 | 7.491
Sr, ppm | 33.619 | 2.1936 | 29.2317 | 38.0062 | 27.0381 | 40.1998 | 6.525% | 13.05% | 19.57% | 31.938 | 35.2999
Ta, ppm | 1.9336 | 0.14892 | 1.6357 | 2.2314 | 1.4868 | 2.3803 | 7.702% | 15.4% | 23.11% | 1.8369 | 2.0302
Tb, ppm | 0.5953 | 0.025563 | 0.54417 | 0.64642 | 0.51861 | 0.67199 | 4.294% | 8.588% | 12.88% | 0.56553 | 0.62506
Th, ppm | 13.1662 | 0.59752 | 11.9711 | 14.3612 | 11.3736 | 14.9587 | 4.538% | 9.077% | 13.61% | 12.5079 | 13.8245
Ti, wt.% | 1.1067 | 0.040275 | 1.0262 | 1.1873 | 0.98588 | 1.2275 | 3.639% | 7.278% | 10.92% | 1.0514 | 1.162
Tl, ppm | < 0.5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tm, ppm | 0.31941 | 0.028667 | 0.26207 | 0.37674 | 0.23341 | 0.40541 | 8.975% | 17.95% | 26.92% | 0.30344 | 0.33538
U, ppm | 2.7267 | 0.16534 | 2.3961 | 3.0574 | 2.2307 | 3.2228 | 6.064% | 12.13% | 18.19% | 2.5904 | 2.8631
V, ppm | 188.2093 | 6.3975 | 175.4143 | 201.0042 | 169.0169 | 207.4017 | 3.399% | 6.798% | 10.2% | 178.7988 | 197.6197
W, ppm | 2.8978 | 0.57654 | 1.7447 | 4.0508 | 1.1682 | 4.6274 | 19.9% | 39.79% | 59.69% | 2.7529 | 3.0426
Y, ppm | 20.5914 | 1.027 | 18.5375 | 22.6454 | 17.5105 | 23.6723 | 4.987% | 9.975% | 14.96% | 19.5618 | 21.621
Yb, ppm | 2.1647 | 0.15262 | 1.8594 | 2.4699 | 1.7068 | 2.6225 | 7.051% | 14.1% | 21.15% | 2.0564 | 2.2729
Zn, ppm | 65.405 | 7.3653 | 50.6744 | 80.1356 | 43.3091 | 87.5009 | 11.26% | 22.52% | 33.78% | 62.1348 | 68.6753
Zr, ppm | 322.01 | 11.3374 | 299.3352 | 344.6848 | 287.9979 | 356.0221 | 3.521% | 7.042% | 10.56% | 305.9095 | 338.1105
4-Acid Digestion
Al, wt.% | 8.5964 | 0.25566 | 8.0851 | 9.1078 | 7.8295 | 9.3634 | 2.974% | 5.948% | 8.922% | 8.1666 | 9.0263
As, ppm | 14.5098 | 1.0267 | 12.4564 | 16.5632 | 11.4297 | 17.5899 | 7.076% | 14.15% | 21.23% | 13.7843 | 15.2353
Ba, ppm | 211.4159 | 8.2456 | 194.9248 | 227.9071 | 186.6792 | 236.1526 | 3.9% | 7.8% | 11.7% | 200.8451 | 221.9867
Be, ppm | 1.2972 | 0.079569 | 1.1381 | 1.4564 | 1.0585 | 1.5359 | 6.134% | 12.27% | 18.4% | 1.2324 | 1.3621
Bi, ppm | 0.40099 | 0.025888 | 0.34921 | 0.45277 | 0.32333 | 0.47866 | 6.456% | 12.91% | 19.37% | 0.38094 | 0.42104
Ca, wt.% | 0.25481 | 0.01862 | 0.21757 | 0.29205 | 0.19895 | 0.31067 | 7.307% | 14.61% | 21.92% | 0.24207 | 0.26755
Cd, ppm | 0.044389 | 0.015212 | 0.013964 | 0.074813 | 0 | 0.090026 | 34.27% | 68.54% | 102.8% | 0.042169 | 0.046608
Ce, ppm | 53.4083 | 5.7651 | 41.8781 | 64.9385 | 36.113 | 70.7036 | 10.79% | 21.59% | 32.38% | 50.7379 | 56.0787
Co, ppm | 9.2912 | 0.42585 | 8.4395 | 10.1429 | 8.0136 | 10.5687 | 4.583% | 9.167% | 13.75% | 8.8266 | 9.7558
Cr, ppm | 204.0091 | 9.937 | 184.1351 | 223.8831 | 174.198 | 233.8201 | 4.871% | 9.742% | 14.61% | 193.8086 | 214.2095
Cs, ppm | 4.6147 | 0.3003 | 4.0141 | 5.2153 | 3.7138 | 5.5156 | 6.508% | 13.02% | 19.52% | 4.3839 | 4.8454
Cu, ppm | 38.3398 | 3.7491 | 30.8415 | 45.838 | 27.0924 | 49.5872 | 9.779% | 19.56% | 29.34% | 36.4228 | 40.2567
Dy, ppm | 2.4689 | 0.28642 | 1.896 | 3.0417 | 1.6096 | 3.3282 | 11.6% | 23.2% | 34.8% | 2.3454 | 2.5923
Er, ppm | 1.2906 | 0.19367 | 0.90321 | 1.6779 | 0.70954 | 1.8716 | 15.01% | 30.01% | 45.02% | 1.226 | 1.3551
Eu, ppm | 0.83389 | 0.07334 | 0.68721 | 0.98057 | 0.61387 | 1.0539 | 8.795% | 17.59% | 26.38% | 0.79219 | 0.87558
Fe, wt.% | 8.1855 | 0.31794 | 7.5496 | 8.8214 | 7.2317 | 9.1393 | 3.884% | 7.768% | 11.65% | 7.7762 | 8.5948
Ga, ppm | 25.2889 | 1.1133 | 23.0624 | 27.5155 | 21.9491 | 28.6288 | 4.402% | 8.805% | 13.21% | 24.0245 | 26.5534
Gd, ppm | 3.1619 | 0.27815 | 2.6056 | 3.7183 | 2.3275 | 3.9964 | 8.797% | 17.59% | 26.39% | 3.0038 | 3.32
Hf, ppm | 4.6794 | 0.20376 | 4.2719 | 5.0869 | 4.0681 | 5.2907 | 4.354% | 8.709% | 13.06% | 4.4454 | 4.9134
Ho, ppm | 0.45722 | 0.080768 | 0.29569 | 0.61876 | 0.21492 | 0.69953 | 17.66% | 35.33% | 52.99% | 0.43436 | 0.48008
In, ppm | 0.095297 | 0.0047756 | 0.085746 | 0.10485 | 0.08097 | 0.10962 | 5.011% | 10.02% | 15.03% | 0.090532 | 0.10006
K, wt.% | 0.76199 | 0.018426 | 0.72514 | 0.79884 | 0.70671 | 0.81727 | 2.418% | 4.836% | 7.254% | 0.72389 | 0.80009
La, ppm | 28.32 | 2.5645 | 23.191 | 33.449 | 20.6265 | 36.0135 | 9.055% | 18.11% | 27.17% | 26.904 | 29.736
Li, ppm | 29.5442 | 1.2696 | 27.0049 | 32.0834 | 25.7353 | 33.3531 | 4.297% | 8.595% | 12.89% | 28.067 | 31.0214
Lu, ppm | 0.1769 | 0.032119 | 0.11267 | 0.24114 | 0.080547 | 0.27326 | 18.16% | 36.31% | 54.47% | 0.16806 | 0.18575
Mg, wt.% | 0.35468 | 0.012942 | 0.3288 | 0.38057 | 0.31586 | 0.39351 | 3.649% | 7.298% | 10.95% | 0.33695 | 0.37242
Mn, wt.% | 0.023602 | 0.0011818 | 0.021239 | 0.025966 | 0.020057 | 0.027148 | 5.007% | 10.01% | 15.02% | 0.022422 | 0.024782
Mo, ppm | 2.1151 | 0.1321 | 1.8509 | 2.3793 | 1.7188 | 2.5114 | 6.245% | 12.49% | 18.74% | 2.0094 | 2.2209
Na, wt.% | 0.070943 | 0.0033475 | 0.064248 | 0.077638 | 0.060901 | 0.080986 | 4.718% | 9.437% | 14.16% | 0.067396 | 0.07449
Nb, ppm | 23.7134 | 2.0283 | 19.6568 | 27.7699 | 17.6285 | 29.7982 | 8.553% | 17.11% | 25.66% | 22.5277 | 24.899
Nd, ppm | 21.1789 | 2.0623 | 17.0543 | 25.3035 | 14.992 | 27.3658 | 9.738% | 19.48% | 29.21% | 20.1199 | 22.2378
Ni, ppm | 68.178 | 3.0194 | 62.1393 | 74.2167 | 59.12 | 77.2361 | 4.429% | 8.857% | 13.29% | 64.7691 | 71.5869
P, wt.% | 0.041431 | 0.0016714 | 0.038088 | 0.044774 | 0.036417 | 0.046445 | 4.034% | 8.068% | 12.1% | 0.03936 | 0.043503
Pb, ppm | 20.9014 | 0.76324 | 19.375 | 22.4279 | 18.6117 | 23.1912 | 3.652% | 7.303% | 10.95% | 19.8564 | 21.9465
Pr, ppm | 5.7989 | 0.5713 | 4.6563 | 6.9415 | 4.085 | 7.5128 | 9.852% | 19.7% | 29.56% | 5.5089 | 6.0888
Rb, ppm | 62.8549 | 3.6247 | 55.6055 | 70.1043 | 51.9808 | 73.7291 | 5.767% | 11.53% | 17.3% | 59.7122 | 65.9977
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.02634 | 0.0046124 | 0.017115 | 0.035564 | 0.012503 | 0.040177 | 17.51% | 35.02% | 52.53% | 0.025023 | 0.027657
Sb, ppm | 0.98057 | 0.056156 | 0.86826 | 1.0929 | 0.8121 | 1.149 | 5.727% | 11.45% | 17.18% | 0.93154 | 1.0296
Sc, ppm | 15.288 | 0.95062 | 13.3867 | 17.1892 | 12.4361 | 18.1398 | 6.218% | 12.44% | 18.65% | 14.5236 | 16.0524
Sm, ppm | 3.9283 | 0.40604 | 3.1163 | 4.7404 | 2.7102 | 5.1464 | 10.34% | 20.67% | 31.01% | 3.7319 | 4.1247
Sn, ppm | 5.0739 | 0.30834 | 4.4573 | 5.6906 | 4.1489 | 5.999 | 6.077% | 12.15% | 18.23% | 4.8202 | 5.3276
Sr, ppm | 32.3363 | 2.0497 | 28.237 | 36.4356 | 26.1874 | 38.4853 | 6.339% | 12.68% | 19.02% | 30.7195 | 33.9532
Ta, ppm | 1.6509 | 0.19392 | 1.263 | 2.0387 | 1.0691 | 2.2326 | 11.75% | 23.49% | 35.24% | 1.5683 | 1.7334
Tb, ppm | 0.4268 | 0.03851 | 0.34978 | 0.50382 | 0.31127 | 0.54233 | 9.023% | 18.05% | 27.07% | 0.40546 | 0.44814
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 12.8301 | 0.80513 | 11.2198 | 14.4403 | 10.4147 | 15.2455 | 6.275% | 12.55% | 18.83% | 12.1886 | 13.4716
Ti, wt.% | 1.0119 | 0.034438 | 0.94301 | 1.0808 | 0.90857 | 1.1152 | 3.403% | 6.807% | 10.21% | 0.96129 | 1.0625
Tl, ppm | 0.39729 | 0.03041 | 0.33647 | 0.45811 | 0.30606 | 0.48852 | 7.654% | 15.31% | 22.96% | 0.37742 | 0.41715
Tm, ppm | 0.18972 | 0.023632 | 0.14246 | 0.23699 | 0.11883 | 0.26062 | 12.46% | 24.91% | 37.37% | 0.18024 | 0.19921
U, ppm | 2.3475 | 0.19254 | 1.9624 | 2.7325 | 1.7698 | 2.9251 | 8.202% | 16.4% | 24.61% | 2.2301 | 2.4648
V, ppm | 175.7634 | 8.0592 | 159.6451 | 191.8817 | 151.5859 | 199.9409 | 4.585% | 9.17% | 13.76% | 166.9752 | 184.5516
W, ppm | 2.0687 | 0.17528 | 1.7182 | 2.4193 | 1.5429 | 2.5946 | 8.473% | 16.95% | 25.42% | 1.9653 | 2.1722
Y, ppm | 11.7702 | 1.7823 | 8.2056 | 15.3348 | 6.4233 | 17.1171 | 15.14% | 30.29% | 45.43% | 11.1817 | 12.3587
Yb, ppm | 1.0964 | 0.17336 | 0.74968 | 1.4431 | 0.57632 | 1.6165 | 15.81% | 31.62% | 47.43% | 1.0416 | 1.1512
Zn, ppm | 65.5427 | 1.8499 | 61.8428 | 69.2425 | 59.9929 | 71.0925 | 2.822% | 5.645% | 8.467% | 62.2655 | 68.8198
Zr, ppm | 181.0367 | 9.7989 | 161.4389 | 200.6346 | 151.6399 | 210.4335 | 5.413% | 10.83% | 16.24% | 171.9849 | 190.0886
Aqua Regia Digestion
Ag, ppm | 0.047068 | 0.015104 | 0.01686 | 0.077276 | 0.0017563 | 0.09238 | 32.09% | 64.18% | 96.27% | 0.044715 | 0.049422
Al, wt.% | 4.2611 | 0.29363 | 3.6738 | 4.8483 | 3.3802 | 5.1419 | 6.891% | 13.78% | 20.67% | 4.048 | 4.4741
As, ppm | 2.1552 | 0.48571 | 1.1838 | 3.1267 | 0.69811 | 3.6124 | 22.54% | 45.07% | 67.61% | 2.0475 | 2.263
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 68.227 | 2.5267 | 63.1736 | 73.2805 | 60.6468 | 75.8072 | 3.703% | 7.407% | 11.11% | 64.8157 | 71.6384
Be, ppm | 0.76387 | 0.045017 | 0.67384 | 0.8539 | 0.62882 | 0.89892 | 5.893% | 11.79% | 17.68% | 0.72568 | 0.80206
Bi, ppm | 0.34354 | 0.027192 | 0.28916 | 0.39793 | 0.26196 | 0.42512 | 7.915% | 15.83% | 23.75% | 0.32636 | 0.36072
Ca, wt.% | 0.22081 | 0.010099 | 0.20061 | 0.241 | 0.19051 | 0.2511 | 4.574% | 9.147% | 13.72% | 0.20977 | 0.23185
Cd, ppm | 0.043228 | 0.010738 | 0.021752 | 0.064703 | 0.011015 | 0.075441 | 24.84% | 49.68% | 74.52% | 0.041066 | 0.045389
Ce, ppm | 32.6171 | 1.6342 | 29.3487 | 35.8856 | 27.7144 | 37.5199 | 5.01% | 10.02% | 15.03% | 30.9863 | 34.248
Co, ppm | 6.4047 | 1.1461 | 4.1125 | 8.6968 | 2.9664 | 9.8429 | 17.89% | 35.79% | 53.68% | 6.0844 | 6.7249
Cr, ppm | 140.4525 | 6.9784 | 126.4957 | 154.4094 | 119.5172 | 161.3879 | 4.969% | 9.937% | 14.91% | 133.4299 | 147.4752
Cs, ppm | 2.0263 | 0.24016 | 1.546 | 2.5066 | 1.3058 | 2.7468 | 11.85% | 23.7% | 35.56% | 1.925 | 2.1276
Cu, ppm | 29.1105 | 2.9915 | 23.1276 | 35.0935 | 20.1361 | 38.085 | 10.28% | 20.55% | 30.83% | 27.655 | 30.5661
Fe, wt.% | 7.6055 | 0.21793 | 7.1696 | 8.0414 | 6.9517 | 8.2593 | 2.865% | 5.731% | 8.596% | 7.2252 | 7.9858
Ga, ppm | 17.8782 | 0.99976 | 15.8787 | 19.8777 | 14.8789 | 20.8775 | 5.592% | 11.18% | 16.78% | 16.9843 | 18.7721
Hf, ppm | 0.7518 | 0.11879 | 0.51423 | 0.98938 | 0.39544 | 1.1082 | 15.8% | 31.6% | 47.4% | 0.71421 | 0.78939
Hg, ppm | 0.027625 | 0.0090535 | 0.009518 | 0.045732 | 0.00046451 | 0.054785 | 32.77% | 65.55% | 98.32% | 0.026244 | 0.029006
In, ppm | 0.064552 | 0.0054633 | 0.053626 | 0.075479 | 0.048162 | 0.080942 | 8.463% | 16.93% | 25.39% | 0.061325 | 0.06778
K, wt.% | 0.12534 | 0.0060999 | 0.11314 | 0.13754 | 0.10704 | 0.14364 | 4.867% | 9.733% | 14.6% | 0.11907 | 0.13161
La, ppm | 15.6887 | 0.73067 | 14.2274 | 17.15 | 13.4967 | 17.8807 | 4.657% | 9.315% | 13.97% | 14.9043 | 16.4731
Li, ppm | 14.0664 | 2.047 | 9.9725 | 18.1603 | 7.9255 | 20.2073 | 14.55% | 29.1% | 43.66% | 13.3631 | 14.7697
Mg, wt.% | 0.19251 | 0.026882 | 0.13874 | 0.24627 | 0.11186 | 0.27315 | 13.96% | 27.93% | 41.89% | 0.18288 | 0.20213
Mn, wt.% | 0.017504 | 0.0011909 | 0.015123 | 0.019886 | 0.013932 | 0.021077 | 6.803% | 13.61% | 20.41% | 0.016629 | 0.018379
Mo, ppm | 1.2431 | 0.26613 | 0.71083 | 1.7753 | 0.44471 | 2.0415 | 21.41% | 42.82% | 64.23% | 1.1809 | 1.3052
Na, wt.% | 0.018542 | 0.0029832 | 0.012575 | 0.024508 | 0.009592 | 0.027491 | 16.09% | 32.18% | 48.27% | 0.017615 | 0.019469
Nb, ppm | 0.58488 | 0.22728 | 0.13032 | 1.0394 | 0 | 1.2667 | 38.86% | 77.72% | 116.6% | 0.55563 | 0.61412
Ni, ppm | 34.4106 | 5.0305 | 24.3497 | 44.4715 | 19.3192 | 49.502 | 14.62% | 29.24% | 43.86% | 32.6901 | 36.1311
P, wt.% | 0.032405 | 0.001529 | 0.029347 | 0.035463 | 0.027818 | 0.036992 | 4.718% | 9.437% | 14.16% | 0.030785 | 0.034025
Pb, ppm | 18.8036 | 0.77112 | 17.2613 | 20.3458 | 16.4902 | 21.1169 | 4.101% | 8.202% | 12.3% | 17.8634 | 19.7437
Rb, ppm | 15.7742 | 1.7325 | 12.3092 | 19.2391 | 10.5768 | 20.9715 | 10.98% | 21.97% | 32.95% | 14.9854 | 16.5629
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.021894 | 0.0035087 | 0.014877 | 0.028911 | 0.011368 | 0.03242 | 16.03% | 32.05% | 48.08% | 0.020799 | 0.022989
Sb, ppm | 0.42152 | 0.095057 | 0.2314 | 0.61163 | 0.13635 | 0.70669 | 22.55% | 45.1% | 67.65% | 0.40044 | 0.44259
Sc, ppm | 10.0007 | 0.49663 | 9.0074 | 10.994 | 8.5108 | 11.4906 | 4.966% | 9.932% | 14.9% | 9.5007 | 10.5007
Se, ppm | 0.43208 | 0.10434 | 0.22341 | 0.64075 | 0.11908 | 0.74509 | 24.15% | 48.29% | 72.44% | 0.41048 | 0.45369
Sn, ppm | 2.2639 | 0.20501 | 1.8539 | 2.6739 | 1.6489 | 2.8789 | 9.055% | 18.11% | 27.17% | 2.1507 | 2.3771
Sr, ppm | 13.2966 | 1.1567 | 10.9832 | 15.6101 | 9.8264 | 16.7669 | 8.7% | 17.4% | 26.1% | 12.6318 | 13.9615
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.2408 | 0.029584 | 0.18163 | 0.29997 | 0.15205 | 0.32955 | 12.29% | 24.57% | 36.86% | 0.22876 | 0.25284
Th, ppm | 9.1163 | 0.55996 | 7.9964 | 10.2363 | 7.4364 | 10.7962 | 6.142% | 12.28% | 18.43% | 8.6605 | 9.5722
Ti, wt.% | 0.10499 | 0.011574 | 0.081839 | 0.12813 | 0.070265 | 0.13971 | 11.02% | 22.05% | 33.07% | 0.099737 | 0.11024
Tl, ppm | 0.13099 | 0.010847 | 0.1093 | 0.15269 | 0.098451 | 0.16354 | 8.281% | 16.56% | 24.84% | 0.12444 | 0.13754
U, ppm | 1.238 | 0.094686 | 1.0487 | 1.4274 | 0.95398 | 1.5221 | 7.648% | 15.3% | 22.94% | 1.1761 | 1.2999
V, ppm | 135.8591 | 4.5054 | 126.8483 | 144.8699 | 122.343 | 149.3753 | 3.316% | 6.632% | 9.949% | 129.0662 | 142.6521
W, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Y, ppm | 5.6277 | 0.17637 | 5.2749 | 5.9804 | 5.0986 | 6.1568 | 3.134% | 6.268% | 9.402% | 5.3463 | 5.9091
Yb, ppm | 0.43733 | 0.044577 | 0.34818 | 0.52649 | 0.3036 | 0.57106 | 10.19% | 20.39% | 30.58% | 0.41547 | 0.4592
Zn, ppm | 48.4367 | 4.2139 | 40.0089 | 56.8646 | 35.7949 | 61.0785 | 8.7% | 17.4% | 26.1% | 46.0149 | 50.8585
Zr, ppm | 29.2155 | 2.0053 | 25.2049 | 33.2261 | 23.1996 | 35.2314 | 6.864% | 13.73% | 20.59% | 27.7547 | 30.6763
Infrared Combustion
C, wt.% | 1.1443 | 0.027227 | 1.0899 | 1.1988 | 1.0627 | 1.226 | 2.379% | 4.759% | 7.138% | 1.0871 | 1.2016
S, wt.% | 0.021552 | 0.0074395 | 0.0066727 | 0.036431 | 0 | 0.04387 | 34.52% | 69.04% | 103.6% | 0.020474 | 0.022629
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 25b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 1.5954 | 0.64014 | 2.5506 | IND | IND
Pd, Palladium (ppb) | 1.4198 | 1.1484 | 1.6912 | IND | IND
Pt, Platinum (ppb) | < 5 | IND | IND | IND | IND
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 16.8502 | 16.648 | 17.0523 | 16.746 | 16.9543
BaO, Barium oxide (ppm) | 214.0126 | 137.9226 | 290.1026 | IND | IND
CaO, Calcium oxide (wt.%) | 0.35606 | 0.34368 | 0.36845 | 0.34742 | 0.36471
Cr2O3, Chromium(III) oxide (ppm) | 295.8967 | 271.5249 | 320.2684 | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 12.2027 | 12.0685 | 12.337 | 12.163 | 12.2425
K2O, Potassium oxide (wt.%) | 0.93575 | 0.92139 | 0.9501 | 0.92322 | 0.94828
MgO, Magnesium oxide (wt.%) | 0.62689 | 0.60808 | 0.6457 | 0.61545 | 0.63833
MnO, Manganese oxide (wt.%) | 0.030444 | 0.028786 | 0.032101 | IND | IND
Na2O, Sodium oxide (wt.%) | 0.093651 | 0.080425 | 0.10688 | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.10042 | 0.094138 | 0.10671 | IND | IND
SiO2, Silicon dioxide (wt.%) | 58.0934 | 57.7534 | 58.4334 | 57.8914 | 58.2953
SO3, Sulphur trioxide (wt.%) | 0.058113 | 0.048625 | 0.0676 | IND | IND
SrO, Strontium oxide (ppm) | < 100 | IND | IND | IND | IND
TiO2, Titanium dioxide (wt.%) | 1.8783 | 1.8555 | 1.9011 | 1.8592 | 1.8974
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 8.4942 | 8.3186 | 8.6698 | 8.4304 | 8.558
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 8.7314 | 8.4414 | 9.0215 | 8.5426 | 8.9203
Ba, Barium (ppm) | 210.4738 | 201.3181 | 219.6296 | 204.8938 | 216.0538
Be, Beryllium (ppm) | 1.5187 | 1.1314 | 1.906 | IND | IND
Ca, Calcium (wt.%) | 0.29034 | 0.26293 | 0.31775 | 0.27956 | 0.30112
Ce, Cerium (ppm) | 52.6057 | 50.6593 | 54.552 | 50.5814 | 54.63
Co, Cobalt (ppm) | 9.9147 | 8.7581 | 11.0713 | 9.2248 | 10.6047
Cr, Chromium (ppm) | 219.1976 | 208.0611 | 230.3341 | 213.5286 | 224.8666
Cs, Caesium (ppm) | 4.5651 | 4.3349 | 4.7953 | 4.3575 | 4.7727
Cu, Copper (ppm) | 41.1978 | 32.0866 | 50.309 | 36.4385 | 45.9572
Dy, Dysprosium (ppm) | 3.6129 | 3.3368 | 3.889 | 3.416 | 3.8098
Er, Erbium (ppm) | 2.16 | 1.9142 | 2.4058 | 1.9456 | 2.3744
Eu, Europium (ppm) | 0.88737 | 0.78918 | 0.98556 | 0.79137 | 0.98337
Fe, Iron (wt.%) | 8.4388 | 8.138 | 8.7395 | 8.2554 | 8.6221
Ga, Gallium (ppm) | 25.8618 | 24.5246 | 27.1989 | 24.8335 | 26.89
Gd, Gadolinium (ppm) | 3.6567 | 3.3951 | 3.9182 | 3.4792 | 3.8341
Ge, Germanium (ppm) | 2.0306 | 1.4196 | 2.6416 | IND | IND
Hf, Hafnium (ppm) | 8.4652 | 7.9869 | 8.9435 | 8.0409 | 8.8896
Ho, Holmium (ppm) | 0.71738 | 0.64783 | 0.78693 | 0.66142 | 0.77334
K, Potassium (wt.%) | 0.80533 | 0.74538 | 0.86528 | 0.78016 | 0.83049
La, Lanthanum (ppm) | 27.9098 | 26.8335 | 28.9862 | 26.8502 | 28.9695
Li, Lithium (ppm) | 30.202 | 24.9896 | 35.4144 | IND | IND
Lu, Lutetium (ppm) | 0.3156 | 0.26944 | 0.36176 | 0.281 | 0.3502
Mg, Magnesium (wt.%) | 0.37737 | 0.36795 | 0.38678 | 0.36782 | 0.38691
Mn, Manganese (wt.%) | 0.02419 | 0.02301 | 0.025371 | 0.023041 | 0.02534
Na, Sodium (wt.%) | 0.065559 | 0.056283 | 0.074835 | IND | IND
Nb, Niobium (ppm) | 27.6233 | 26.1039 | 29.1427 | 26.5969 | 28.6496
Nd, Neodymium (ppm) | 22.4515 | 20.8382 | 24.0649 | 21.1815 | 23.7215
Ni, Nickel (ppm) | 68.5186 | 55.5009 | 81.5363 | 63.6346 | 73.4025
P, Phosphorus (wt.%) | 0.040369 | 0.036542 | 0.044196 | IND | IND
Pr, Praseodymium (ppm) | 6.0366 | 5.6942 | 6.379 | 5.7863 | 6.2869
Rb, Rubidium (ppm) | 60.8711 | 57.6145 | 64.1277 | 57.679 | 64.0632
Sc, Scandium (ppm) | 16.1571 | 14.2984 | 18.0158 | 14.3138 | 18.0004
Si, Silicon (wt.%) | 26.1678 | 25.4021 | 26.9335 | 25.8319 | 26.5038
Sm, Samarium (ppm) | 4.1861 | 3.8062 | 4.5659 | 3.9402 | 4.432
Sn, Tin (ppm) | 7.1343 | 6.37 | 7.8987 | 6.1866 | 8.0821
Sr, Strontium (ppm) | 33.619 | 30.9932 | 36.2447 | 32.1819 | 35.0561
Ta, Tantalum (ppm) | 1.9336 | 1.798 | 2.0691 | 1.7735 | 2.0936
Tb, Terbium (ppm) | 0.5953 | 0.55485 | 0.63574 | 0.53992 | 0.65067
Th, Thorium (ppm) | 13.1662 | 12.5578 | 13.7745 | 12.7267 | 13.6056
Ti, Titanium (wt.%) | 1.1067 | 1.073 | 1.1404 | 1.0758 | 1.1376
Tl, Thallium (ppm) | < 0.5 | IND | IND | IND | IND
Tm, Thulium (ppm) | 0.31941 | 0.27998 | 0.35884 | 0.28989 | 0.34892
U, Uranium (ppm) | 2.7267 | 2.5732 | 2.8803 | 2.589 | 2.8644
V, Vanadium (ppm) | 188.2093 | 180.4635 | 195.955 | 182.5272 | 193.8913
W, Tungsten (ppm) | 2.8978 | 2.2596 | 3.5359 | 2.5089 | 3.2867
Y, Yttrium (ppm) | 20.5914 | 18.9065 | 22.2763 | 19.3216 | 21.8613
Yb, Ytterbium (ppm) | 2.1647 | 1.9418 | 2.3876 | 1.9661 | 2.3632
Zn, Zinc (ppm) | 65.405 | 53.5031 | 77.3069 | IND | IND
Zr, Zirconium (ppm) | 322.01 | 306.031 | 337.989 | 309.4804 | 334.5396
4-Acid Digestion
Al, Aluminium (wt.%) | 8.5964 | 8.3631 | 8.8298 | 8.4153 | 8.7776
As, Arsenic (ppm) | 14.5098 | 13.4224 | 15.5972 | 13.8784 | 15.1411
Ba, Barium (ppm) | 211.4159 | 203.6922 | 219.1397 | 204.4707 | 218.3612
Be, Beryllium (ppm) | 1.2972 | 1.171 | 1.4234 | 1.2518 | 1.3426
Bi, Bismuth (ppm) | 0.40099 | 0.37398 | 0.428 | 0.38092 | 0.42106
Ca, Calcium (wt.%) | 0.25481 | 0.24461 | 0.265 | 0.24438 | 0.26524
Cd, Cadmium (ppm) | 0.044389 | 0.025195 | 0.063583 | IND | IND
Ce, Cerium (ppm) | 53.4083 | 49.8733 | 56.9433 | 51.2848 | 55.5318
Co, Cobalt (ppm) | 9.2912 | 8.8949 | 9.6875 | 9.0365 | 9.5459
Cr, Chromium (ppm) | 204.0091 | 193.9728 | 214.0454 | 198.3505 | 209.6677
Cs, Caesium (ppm) | 4.6147 | 4.4117 | 4.8176 | 4.4509 | 4.7785
Cu, Copper (ppm) | 38.3398 | 35.9573 | 40.7222 | 36.5495 | 40.13
Dy, Dysprosium (ppm) | 2.4689 | 2.1721 | 2.7656 | 2.3311 | 2.6067
Er, Erbium (ppm) | 1.2906 | 1.0561 | 1.5251 | IND | IND
Eu, Europium (ppm) | 0.83389 | 0.74021 | 0.92757 | 0.76677 | 0.90101
Fe, Iron (wt.%) | 8.1855 | 7.9552 | 8.4158 | 8.0463 | 8.3247
Ga, Gallium (ppm) | 25.2889 | 24.4332 | 26.1447 | 24.7012 | 25.8766
Gd, Gadolinium (ppm) | 3.1619 | 2.8733 | 3.4506 | 2.9772 | 3.3467
Hf, Hafnium (ppm) | 4.6794 | 4.4464 | 4.9124 | 4.4682 | 4.8906
Ho, Holmium (ppm) | 0.45722 | 0.33889 | 0.57555 | IND | IND
In, Indium (ppm) | 0.095297 | 0.087403 | 0.10319 | 0.087043 | 0.10355
K, Potassium (wt.%) | 0.76199 | 0.7428 | 0.78118 | 0.74647 | 0.7775
La, Lanthanum (ppm) | 28.32 | 26.7061 | 29.9339 | 27.2887 | 29.3513
Li, Lithium (ppm) | 29.5442 | 28.6815 | 30.4069 | 28.6495 | 30.4389
Lu, Lutetium (ppm) | 0.1769 | 0.13332 | 0.22049 | IND | IND
Mg, Magnesium (wt.%) | 0.35468 | 0.34393 | 0.36543 | 0.34612 | 0.36324
Mn, Manganese (wt.%) | 0.023602 | 0.022577 | 0.024628 | 0.023103 | 0.024102
Mo, Molybdenum (ppm) | 2.1151 | 1.96 | 2.2702 | 2.0298 | 2.2004
Na, Sodium (wt.%) | 0.070943 | 0.069057 | 0.072829 | 0.067281 | 0.074605
Nb, Niobium (ppm) | 23.7134 | 21.9787 | 25.448 | 22.7484 | 24.6783
Nd, Neodymium (ppm) | 21.1789 | 19.1609 | 23.1969 | 20.2292 | 22.1286
Ni, Nickel (ppm) | 68.178 | 65.9576 | 70.3984 | 66.4656 | 69.8904
P, Phosphorus (wt.%) | 0.041431 | 0.039974 | 0.042888 | 0.040174 | 0.042689
Pb, Lead (ppm) | 20.9014 | 20.0079 | 21.795 | 20.1821 | 21.6208
Pr, Praseodymium (ppm) | 5.7989 | 5.2281 | 6.3697 | 5.4393 | 6.1585
Rb, Rubidium (ppm) | 62.8549 | 59.9387 | 65.7711 | 60.1383 | 65.5715
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.02634 | 0.022083 | 0.030596 | 0.023447 | 0.029233
Sb, Antimony (ppm) | 0.98057 | 0.90782 | 1.0533 | 0.92033 | 1.0408
Sc, Scandium (ppm) | 15.288 | 14.4354 | 16.1405 | 14.7845 | 15.7914
Sm, Samarium (ppm) | 3.9283 | 3.4948 | 4.3619 | 3.6588 | 4.1979
Sn, Tin (ppm) | 5.0739 | 4.7927 | 5.3552 | 4.8427 | 5.3052
Sr, Strontium (ppm) | 32.3363 | 30.791 | 33.8816 | 31.168 | 33.5047
Ta, Tantalum (ppm) | 1.6509 | 1.5007 | 1.8011 | 1.5824 | 1.7193
Tb, Terbium (ppm) | 0.4268 | 0.34931 | 0.50429 | 0.37804 | 0.47556
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 12.8301 | 12.1887 | 13.4715 | 12.4512 | 13.209
Ti, Titanium (wt.%) | 1.0119 | 0.98682 | 1.0369 | 0.99034 | 1.0334
Tl, Thallium (ppm) | 0.39729 | 0.37365 | 0.42093 | 0.38164 | 0.41293
Tm, Thulium (ppm) | 0.18972 | 0.16594 | 0.2135 | IND | IND
U, Uranium (ppm) | 2.3475 | 2.1655 | 2.5294 | 2.2389 | 2.456
V, Vanadium (ppm) | 175.7634 | 170.1458 | 181.381 | 172.1034 | 179.4235
W, Tungsten (ppm) | 2.0687 | 1.8783 | 2.2592 | 1.9315 | 2.206
Y, Yttrium (ppm) | 11.7702 | 10.6575 | 12.8829 | 11.1367 | 12.4037
Yb, Ytterbium (ppm) | 1.0964 | 0.8771 | 1.3157 | IND | IND
Zn, Zinc (ppm) | 65.5427 | 62.7154 | 68.3699 | 63.8749 | 67.2104
Zr, Zirconium (ppm) | 181.0367 | 173.1531 | 188.9204 | 175.0037 | 187.0697
Aqua Regia Digestion
Ag, Silver (ppm) | 0.047068 | 0.031403 | 0.062733 | IND | IND
Al, Aluminium (wt.%) | 4.2611 | 3.9919 | 4.5302 | 4.1563 | 4.3658
As, Arsenic (ppm) | 2.1552 | 1.6347 | 2.6757 | 1.9913 | 2.3191
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 68.227 | 65.9991 | 70.455 | 66.6469 | 69.8072
Be, Beryllium (ppm) | 0.76387 | 0.71408 | 0.81366 | 0.73206 | 0.79568
Bi, Bismuth (ppm) | 0.34354 | 0.31958 | 0.3675 | 0.32011 | 0.36698
Ca, Calcium (wt.%) | 0.22081 | 0.21051 | 0.2311 | 0.21172 | 0.22989
Cd, Cadmium (ppm) | 0.043228 | 0.031185 | 0.05527 | IND | IND
Ce, Cerium (ppm) | 32.6171 | 31.2832 | 33.9511 | 31.4116 | 33.8227
Co, Cobalt (ppm) | 6.4047 | 5.7022 | 7.1071 | 6.1683 | 6.641
Cr, Chromium (ppm) | 140.4525 | 135.753 | 145.1521 | 137.4835 | 143.4215
Cs, Caesium (ppm) | 2.0263 | 1.8471 | 2.2055 | 1.9281 | 2.1245
Cu, Copper (ppm) | 29.1105 | 27.0283 | 31.1928 | 28.168 | 30.0531
Fe, Iron (wt.%) | 7.6055 | 7.4241 | 7.7869 | 7.4687 | 7.7423
Ga, Gallium (ppm) | 17.8782 | 17.1526 | 18.6038 | 17.285 | 18.4714
Hf, Hafnium (ppm) | 0.7518 | 0.66056 | 0.84304 | 0.70309 | 0.80051
Hg, Mercury (ppm) | 0.027625 | 0.016083 | 0.039167 | IND | IND
In, Indium (ppm) | 0.064552 | 0.058851 | 0.070253 | 0.058478 | 0.070627
K, Potassium (wt.%) | 0.12534 | 0.119 | 0.13168 | 0.11982 | 0.13086
La, Lanthanum (ppm) | 15.6887 | 14.9673 | 16.4101 | 15.1478 | 16.2296
Li, Lithium (ppm) | 14.0664 | 12.3786 | 15.7543 | 13.5673 | 14.5655
Mg, Magnesium (wt.%) | 0.19251 | 0.17516 | 0.20985 | 0.18444 | 0.20057
Mn, Manganese (wt.%) | 0.017504 | 0.016762 | 0.018247 | 0.01708 | 0.017928
Mo, Molybdenum (ppm) | 1.2431 | 1.0472 | 1.4389 | 1.188 | 1.2982
Na, Sodium (wt.%) | 0.018542 | 0.016973 | 0.02011 | IND | IND
Nb, Niobium (ppm) | 0.58488 | 0.38067 | 0.78908 | 0.53745 | 0.6323
Ni, Nickel (ppm) | 34.4106 | 30.4604 | 38.3608 | 33.2204 | 35.6008
P, Phosphorus (wt.%) | 0.032405 | 0.031283 | 0.033528 | 0.031574 | 0.033237
Pb, Lead (ppm) | 18.8036 | 17.9736 | 19.6335 | 18.1296 | 19.4775
Rb, Rubidium (ppm) | 15.7742 | 14.4401 | 17.1082 | 15.0874 | 16.4609
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.021894 | 0.019793 | 0.023995 | IND | IND
Sb, Antimony (ppm) | 0.42152 | 0.33894 | 0.5041 | 0.38664 | 0.4564
Sc, Scandium (ppm) | 10.0007 | 9.6111 | 10.3903 | 9.7643 | 10.2371
Se, Selenium (ppm) | 0.43208 | 0.31396 | 0.55021 | IND | IND
Sn, Tin (ppm) | 2.2639 | 2.091 | 2.4368 | 2.1218 | 2.406
Sr, Strontium (ppm) | 13.2966 | 12.472 | 14.1212 | 12.9562 | 13.6371
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.2408 | 0.20125 | 0.28035 | 0.22901 | 0.25259
Th, Thorium (ppm) | 9.1163 | 8.7118 | 9.5208 | 8.7737 | 9.459
Ti, Titanium (wt.%) | 0.10499 | 0.094936 | 0.11504 | 0.098951 | 0.11102
Tl, Thallium (ppm) | 0.13099 | 0.11915 | 0.14284 | IND | IND
U, Uranium (ppm) | 1.238 | 1.1727 | 1.3033 | 1.1984 | 1.2776
V, Vanadium (ppm) | 135.8591 | 132.435 | 139.2832 | 133.2002 | 138.518
W, Tungsten (ppm) | < 0.05 | IND | IND | IND | IND
Y, Yttrium (ppm) | 5.6277 | 5.4496 | 5.8057 | 5.4769 | 5.7784
Yb, Ytterbium (ppm) | 0.43733 | 0.39384 | 0.48083 | IND | IND
Zn, Zinc (ppm) | 48.4367 | 45.0876 | 51.7859 | 46.5052 | 50.3682
Zr, Zirconium (ppm) | 29.2155 | 27.633 | 30.798 | 28.3744 | 30.0566
Infrared Combustion
C, Carbon (wt.%) | 1.1443 | 1.1204 | 1.1683 | 1.1295 | 1.1592
S, Sulphur (wt.%) | 0.021552 | 0.013902 | 0.029201 | IND | IND
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 25b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
As | ppm | < 100 | La | ppm | < 100 | Ta | ppm | < 100
Bi | ppm | < 100 | Mo | ppm | < 100 | Th | ppm | < 50
Ce | ppm | < 100 | Nb | ppm | < 100 | TOT_XRF | wt.% | 99.8961
Cl | ppm | 92.5 | Nd | ppm | < 100 | U | ppm | < 50
Co | ppm | < 37 | Ni | ppm | 52.3903 | V2O5 | ppm | 295.9621
Cu | ppm | < 40 | Pb | ppm | 46.4166 | W | ppm | < 100
Dy | ppm | < 90 | Pr | ppm | < 100 | Y | ppm | < 100
Eu | ppm | < 100 | Rb | ppm | < 100 | Zn | ppm | < 100
F | ppm | < 500 | Sb | ppm | < 100 | ZrO2 | ppm | 438.6159
Gd | ppm | < 100 | Sm | ppm | < 100
Hf | ppm | < 100 | Sn | ppm | < 100
Borate / Peroxide Fusion ICP
Ag | ppm | 5.4792 | In | ppm | 0.12131 | Sb | ppm | 1.0268
As | ppm | 13.7917 | Mo | ppm | 2.8522 | Se | ppm | < 20
B | ppm | 79.5858 | Pb | ppm | 23.3935 | Te | ppm | < 1
Bi | ppm | 0.43246 | Re | ppm | < 0.1
Cd | ppm | < 10 | S | wt.% | 0.032778
4-Acid Digestion
Ag | ppm | 0.062019 | Ge | ppm | 0.19114 | Se | ppm | 1.8033
Aqua Regia Digestion
Au | ppb | 1.339 | Ge | ppm | 0.0765 | Pr | ppm | 3.4852
Dy | ppm | 1.2974 | Ho | ppm | 0.21689 | Pt | ppb | < 1
Er | ppm | 0.56344 | Lu | ppm | 0.06025 | Sm | ppm | 2.5556
Eu | ppm | 0.51944 | Nd | ppm | 14.1352 | Te | ppm | 0.024306
Gd | ppm | 1.9875 | Pd | ppb | 3.6667 | Tm | ppm | 0.074833
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 25b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 25b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil

## Fire Assay
| Analytical results for Au in OREAS 25b (Certified Value 1.6 ppb) | [16 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppb
|  |  | FA*MS | FA*OES | FA*AAS | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*MS | 1
|  |  | 25g | 50g | 40g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 30g | 30g | 30g | 2
| 1 | 1 | 1 | 2.5 | 2 | < 5 | < 1 | 7 | 1 | < 1 | < 1 | 1 | 1.56 | 2 | < 2 | 9 | 3 | 1
| 1 | 2 | 2 | 2.8 | 2 | < 5 | 1 | 4 | 1 | 2 | < 1 | 1 | 1.5 | 1 | 3 | 5 | 3 | 21
| 1 | 3 | 2 | 2.6 | 3 | < 5 | 1 | 5 | 1 | < 1 | < 1 | 1 | 1.41 | 1 | < 2 | < 5 | 3 | 16
| 1 | 4 | 1 | 2.2 | 3 | < 5 | < 1 | 5 | 1 | 1 | < 1 | 1 | 1.28 | 1 | < 2 | 10 | 3 | 1.5954
| 1 | 5 | 2 | 2.8 | 2 | < 5 | 1 | 7 | 1 | < 1 | < 1 | 1 | 1.84 | 1 | < 2 | 23 | 3 | 7
| 1 | 6 | 2 | 2.5 | 2 | < 5 | 1 | 6 | 1 | < 1 | < 1 | 3 | 1.45 | 1 | 2 | < 5 | 3
| Mean |  | 1.6667 | 2.5667 | 2.3333 |  | 1 | 5.6667 | 1 | 1.5 |  | 1.3333 | 1.5067 | 1.1667 | 2.5 | 11.75 | 3
| Median |  | 2 | 2.55 | 2 |  | 1 | 5.5 | 1 | 1.5 |  | 1 | 1.475 | 1 | 2.5 | 9.5 | 3
| Std Dev. |  | 0.5164 | 0.22509 | 0.5164 |  | 0 | 1.2111 | 0 | 0.70711 |  | 0.8165 | 0.18864 | 0.40825 | 0.70711 | 7.8049 | 0
| Rel.Std.Dev. |  | 30.98% | 8.77% | 22.13% |  | 0% | 21.37% | 0% | 47.14% |  | 61.24% | 12.52% | 34.99% | 28.28% | 66.42% | 0%
| PDM3 |  | 4.468% | 60.88% | 46.26% |  | -37.32% | 255.2% | -37.32% | -5.979% |  | -16.43% | -5.561% | -26.87% | 56.7% | 636.5% | 88.04%
| Z-Score (Absolute) |  | 0.16 | 1.06 | 0.83 | 0.99 | 0.67 | 4.17 | 0.51 | 0.67 | 1.01 | 0.17 | 0 | 0.34 | 0.01 | 7.17 | 1.5
| Analytical results for Pd in OREAS 25b (Certified Value 1.42 ppb) | [16 empty] | Certified
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppb
|  |  | FA*MS | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*MS |  | 2
|  |  | 25g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 30g | 30g | 30g |  | 2
| 1 | 1 | 1.4 | 2.6 | < 5 | 1.3 | 4 | 1 | 1 | < 1 | 1.5 | 1.4 | 2 | < 2 | < 5 | 1.7 |  | 1
| 1 | 2 | 1.4 | 2.5 | < 5 | 1.2 | 3 | 1 | 1 | 1 | 1.5 | 1.9 | < 1 | 2 | < 5 | 1.7 |  | 25
| 1 | 3 | 1.5 | 2.4 | < 5 | 1.2 | < 2 | < 1 | 1 | < 1 | 1.6 | 1.56 | 2 | < 2 | < 5 | 1.8 |  | 16
| 1 | 4 | 1.5 | 2.2 | < 5 | 1.2 | < 2 | 1 | 1 | < 1 | 1.5 | 1.68 | 1 | < 2 | < 5 | 1.8 |  | 1.4198
| 1 | 5 | 1.7 | 1.9 | < 5 | 1.3 | < 2 | 2 | 1 | < 1 | 1.5 | 1.6 | 1 | < 2 | < 5 | 1.7 |  | 8
| 1 | 6 | 1.5 | 2.2 | < 5 | 1.4 | < 2 | 1 | 1 | 1 | 1.4 | 1.45 | 1 | < 2 | < 5 | 1.7
| Mean |  | 1.5 | 2.3 |  | 1.2667 | 3.5 | 1.2 | 1 | 1 | 1.5 | 1.5983 | 1.4 | 2 |  | 1.7333
| Median |  | 1.5 | 2.3 |  | 1.25 | 3.5 | 1 | 1 | 1 | 1.5 | 1.58 | 1 | 2 |  | 1.7
| Std Dev. |  | 0.10954 | 0.25298 |  | 0.08165 | 0.70711 | 0.44721 | 0 | 0 | 0.063246 | 0.17916 | 0.54772 |  |  | 0.05164
| Rel.Std.Dev. |  | 7.303% | 11% |  | 6.446% | 20.2% | 37.27% | 0% | 0% | 4.216% | 11.21% | 39.12% |  |  | 2.979%
| PDM3 |  | 5.646% | 61.99% |  | -10.79% | 146.5% | -15.48% | -29.57% | -29.57% | 5.646% | 12.57% | -1.397% | 40.86% |  | 22.08%
| Z-Score (Absolute) |  | 0 | 1.62 | 2.02 | 0.47 | 0.67 | 0.84 | 1.01 | 1.69 | 0 | 0.2 | 0.51 | 0.67 | 2.02 | 0.47
| Analytical results for Pt in OREAS 25b (Certified Value < 5 ppb) | [16 empty] | Certified
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppb
|  |  | FA*MS | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*MS |  | 2
|  |  | 25g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 30g | 30g | 30g |  | 2
| 1 | 1 | 1 | 1.4 | < 5 | 0.3 | 4 | < 5 | < 5 | < 5 | < 5 | < 5 | < 5 | < 3 | < 10 | 0.8 |  | 1
| 1 | 2 | 0.9 | 0.9 | < 5 | 0.4 | < 3 | < 5 | < 5 | < 5 | < 5 | < 5 | < 5 | < 3 | < 10 | 0.7 |  | 22
| 1 | 3 | 1 | 1.3 | < 5 | 0.4 | < 3 | < 5 | < 5 | < 5 | < 5 | < 5 | < 5 | < 3 | < 10 | 0.7 |  | 16
| 1 | 4 | 0.9 | 1.2 | < 5 | 0.4 | < 3 | < 5 | < 5 | < 5 | < 5 | < 5 | < 5 | < 3 | < 10 | 0.7 |  | < 5
| 1 | 5 | 1 | 1 | < 5 | 1.5 | < 3 | < 5 | < 5 | < 5 | < 5 | < 5 | < 5 | < 3 | < 10 | 0.7 |  | 9
| 1 | 6 | 1.1 | 1.4 | < 5 | 0.5 | 4 | < 5 | < 5 | < 5 | < 5 | < 5 | < 5 | < 3 | < 10 | 0.6
| Mean |  | 0.98333 | 1.2 |  | 0.58333 | 4 | [8 empty] | 0.7
| Median |  | 1 | 1.25 |  | 0.4 | 4 | [8 empty] | 0.7
| Std Dev. |  | 0.075277 | 0.20976 |  | 0.4535 | 0 | [8 empty] | 0.063246
| Rel.Std.Dev. |  | 7.655% | 17.48% |  | 77.74% | 0% | [8 empty] | 9.035%
PDM3
| Z-Score (Absolute) |  | 12.27 | 10.52 | 0 | 15.51 | 1.35 | 0 | 0 | 0 | 0 | 0 | 0 | 8.09 | 20.23 | 14.57
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 25b (Certified Value 16.85 wt.%) | [14 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 17.13 | 16.9 | 16.52 | 16.9344 | 16.48 | 17.18 | 17.02 | 17 | 17.09 | 16.98 | 16.48 | 16.3 | 17.1 | 1
| 1 | 2 | 17.12 | 17.1 | 16.58 | 16.9285 | 16.31 | 17.01 | 17.05 | 16.81 | 17.16 | 17.08 | 16.53 | 16.34 | 17.1 | #N/A
| 1 | 3 | 17.08 | 16.9 | 16.64 | 16.9559 | 16.5 | 17.11 | 17.03 | 16.8 | 17.14 | 17.08 | 16.57 | 16.22 | 17.1 | 16
| 1 | 4 | 17.1 | 16.8 | 16.54 | 16.9555 | 16.37 | 17.05 | 17.08 | 16.82 | 17.13 | 17.1 | 16.55 | 16.23 | 17.2 | 16.8502
| 1 | 5 | 17.14 | 16.5 | 16.56 | 16.9754 | 16.41 | 17.08 | 17.08 | 16.93 | 16.84 | 17 | 16.61 | 16.32 | 17.1 | 11
| 1 | 6 | 17.17 | 16.6 | 16.58 | 16.9081 | 16.49 | 17.04 | 17.14 | 16.98 | 17.25 | 17.02 | 16.54 | 16.25 | 17.2
| Mean |  | 17.1233 | 16.8 | 16.57 | 16.943 | 16.4267 | 17.0783 | 17.0667 | 16.89 | 17.1017 | 17.0433 | 16.5467 | 16.2767 | 17.1333
| Median |  | 17.125 | 16.85 | 16.57 | 16.9449 | 16.445 | 17.065 | 17.065 | 16.875 | 17.135 | 17.05 | 16.545 | 16.275 | 17.1
| Std Dev. |  | 0.031411 | 0.21909 | 0.041473 | 0.023965 | 0.076594 | 0.06047 | 0.043665 | 0.090774 | 0.13877 | 0.049666 | 0.043205 | 0.050067 | 0.05164
| Rel.Std.Dev. |  | 0.1834% | 1.304% | 0.2503% | 0.1414% | 0.4663% | 0.3541% | 0.2559% | 0.5374% | 0.8114% | 0.2914% | 0.2611% | 0.3076% | 0.3014%
| PDM3 |  | 1.621% | -0.2976% | -1.663% | 0.5509% | -2.513% | 1.354% | 1.285% | 0.2365% | 1.493% | 1.146% | -1.801% | -3.403% | 1.681%
| Z-Score (Absolute) |  | 0.77 | 0.61 | 1.58 | 0 | 2.19 | 0.58 | 0.53 | 0.23 | 0.67 | 0.43 | 1.68 | 2.83 | 0.81
| Analytical results for As in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
As | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 1
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 25b (Certified Value 214 ppm) | [14 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | 209.4866 | 220 | 200 | 300 | 100 | 200 | 300 | < 100 | 100 |  |  |  | 1
| 1 | 2 | < 100 | 206.4827 | 220 | 200 | 200 | 200 | 200 | 200 | < 100 | 200 |  |  |  | #N/A
| 1 | 3 | < 100 | 204.3661 | 220 | 200 | 300 | 200 | 200 | 200 | < 100 | 200 |  |  |  | 16
| 1 | 4 | < 100 | 207.7523 | 220 | 200 | 300 | 200 | 300 | 200 | < 100 | 300 |  |  |  | 214.0126
| 1 | 5 | < 100 | 210.5876 | 220 | 200 | 300 | 100 | 200 | 200 | < 100 | 200 |  |  |  | 12
| 1 | 6 | < 100 | 205.9796 | 220 | 200 | 300 | 200 | 200 | 200 | < 100 | 200
| Mean |  |  | 207.4425 | 220 | 200 | 283.3333 | 166.6667 | 216.6667 | 216.6667 |  | 200
| Median |  |  | 207.1175 | 220 | 200 | 300 | 200 | 200 | 200 |  | 200
| Std Dev. |  |  | 2.3105 | 0 | 0 | 40.8248 | 51.6398 | 40.8248 | 40.8248 |  | 63.2456
| Rel.Std.Dev. |  |  | 1.114% | 0% | 0% | 14.41% | 30.98% | 18.84% | 18.84% |  | 31.62%
| PDM3 |  |  | -3.07% | 2.798% | -6.548% | 32.39% | -22.12% | 1.24% | 1.24% |  | -6.548%
| Z-Score (Absolute) |  | 6.37 | 0.15 | 0.81 | 0.15 | 3.3 | 1.54 | 0.54 | 0.54 | 6.37 | 0.15
| Analytical results for Bi in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Bi | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 2
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 25b (Certified Value 0.356 wt.%) | [14 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.37 | 0.37 | 0.35 | 0.3666 | 0.36 | 0.36 | 0.36 | 0.35 | 0.35 | 0.37 | 0.35 | 0.33 | 0.35 | 1
| 1 | 2 | 0.36 | 0.38 | 0.35 | 0.36651 | 0.35 | 0.37 | 0.35 | 0.35 | 0.35 | 0.37 | 0.34 | 0.33 | 0.35 | #N/A
| 1 | 3 | 0.37 | 0.38 | 0.35 | 0.37092 | 0.35 | 0.37 | 0.35 | 0.34 | 0.35 | 0.37 | 0.34 | 0.33 | 0.36 | 16
| 1 | 4 | 0.36 | 0.37 | 0.34 | 0.36772 | 0.35 | 0.36 | 0.36 | 0.35 | 0.35 | 0.39 | 0.34 | 0.33 | 0.36 | 0.35606
| 1 | 5 | 0.36 | 0.36 | 0.35 | 0.36713 | 0.35 | 0.37 | 0.36 | 0.35 | 0.34 | 0.37 | 0.35 | 0.35 | 0.35 | 13
| 1 | 6 | 0.36 | 0.37 | 0.34 | 0.36401 | 0.36 | 0.36 | 0.36 | 0.35 | 0.35 | 0.37 | 0.34 | 0.34 | 0.37
| Mean |  | 0.36333 | 0.37167 | 0.34667 | 0.36715 | 0.35333 | 0.365 | 0.35667 | 0.34833 | 0.34833 | 0.37333 | 0.34333 | 0.335 | 0.35667
| Median |  | 0.36 | 0.37 | 0.35 | 0.36687 | 0.35 | 0.365 | 0.36 | 0.35 | 0.35 | 0.37 | 0.34 | 0.33 | 0.355
| Std Dev. |  | 0.005164 | 0.0075277 | 0.005164 | 0.0022438 | 0.005164 | 0.0054772 | 0.005164 | 0.0040825 | 0.0040825 | 0.008165 | 0.005164 | 0.0083666 | 0.008165
| Rel.Std.Dev. |  | 1.421% | 2.025% | 1.49% | 0.6111% | 1.462% | 1.501% | 1.448% | 1.172% | 1.172% | 2.187% | 1.504% | 2.497% | 2.289%
| PDM3 |  | 2.042% | 4.382% | -2.639% | 3.114% | -0.7666% | 2.51% | 0.1696% | -2.171% | -2.171% | 4.85% | -3.575% | -5.915% | 0.1696%
| Z-Score (Absolute) |  | 0.54 | 1.21 | 0.81 | 0.85 | 0.27 | 0.67 | 0 | 0.67 | 0.67 | 1.35 | 1.08 | 1.75 | 0
| Analytical results for Ce in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Ce | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 3
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 9
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 25b (Indicative Value 93 ppm) | [14 empty] | Indicative
Cl | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | [12 empty] | 1
| 1 | 2 | 150 | [12 empty] | 4
| 1 | 3 | 60 | [12 empty] | 16
| 1 | 4 | 110 | [12 empty] | 92.5
| 1 | 5 | 100 | [12 empty] | 10
| 1 | 6 | 110
| Mean |  | 106
| Median |  | 110
| Std Dev. |  | 32.0936
| Rel.Std.Dev. |  | 30.28%
| PDM3 |  | 14.59%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 25b (Indicative Value < 37 ppm) | [14 empty] | Indicative
Co | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 05 | 16 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 37 | < 70 | [11 empty] | 1
| 1 | 2 | < 37 | < 70 | [11 empty] | 5
| 1 | 3 | < 37 | < 70 | [11 empty] | 16
| 1 | 4 | < 37 | < 70 | [11 empty] | < 37
| 1 | 5 | < 37 | < 70 | [11 empty] | 11
| 1 | 6 | < 37 | < 70
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Cr2O3 in OREAS 25b (Certified Value 296 ppm) | [14 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 300 | 600 | 523.7164 | 200 | 300 | 300 | 300 | 300 | 300 | 300 | 200 | 290 |  | 1
| 1 | 2 | 300 | 600 | 381.993 | 200 | 300 | 300 | 300 | 300 | 300 | 400 | 300 | 290 |  | #N/A
| 1 | 3 | 300 | 500 | 572.2251 | 200 | 300 | 300 | 300 | 300 | 300 | 300 | 300 | 290 |  | 16
| 1 | 4 | 300 | 600 | 516.0039 | 200 | 300 | 400 | 300 | 300 | 300 | 400 | 200 | 150 |  | 295.8967
| 1 | 5 | 300 | 700 | 417.1968 | 300 | 300 | 400 | 300 | 300 | 300 | 300 | 200 | 290 |  | 14
| 1 | 6 | 300 | 600 | 473.7593 | 300 | 300 | 400 | 300 | 300 | 300 | 300 | 300 | 290
| Mean |  | 300 | 600 | 480.8158 | 233.3333 | 300 | 350 | 300 | 300 | 300 | 333.3333 | 250 | 266.6667
| Median |  | 300 | 600 | 494.8816 | 200 | 300 | 350 | 300 | 300 | 300 | 300 | 250 | 290
| Std Dev. |  | 0 | 63.2456 | 71.1301 | 51.6398 | 0 | 54.7723 | 0 | 0 | 0 | 51.6398 | 54.7723 | 57.1548
| Rel.Std.Dev. |  | 0% | 10.54% | 14.79% | 22.13% | 0% | 15.65% | 0% | 0% | 0% | 15.49% | 21.91% | 21.43%
| PDM3 |  | 1.387% | 102.8% | 62.49% | -21.14% | 1.387% | 18.28% | 1.387% | 1.387% | 1.387% | 12.65% | -15.51% | -9.878%
| Z-Score (Absolute) |  | 0 | 6.19 | 3.73 | 1.37 | 0 | 1.03 | 0 | 0 | 0 | 0.69 | 1.03 | 0.66
| Analytical results for Cu in OREAS 25b (Indicative Value < 40 ppm) | [14 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 05 | 16 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 40 | < 80 | [11 empty] | 1
| 1 | 2 | < 40 | < 80 | [11 empty] | 6
| 1 | 3 | < 40 | < 80 | [11 empty] | 16
| 1 | 4 | < 40 | < 80 | [11 empty] | < 40
| 1 | 5 | < 40 | < 80 | [11 empty] | 12
| 1 | 6 | < 40 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Dy in OREAS 25b (Indicative Value < 90 ppm) | [14 empty] | Indicative
Dy | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 90 | [12 empty] | 1
| 1 | 2 | < 90 | [12 empty] | 7
| 1 | 3 | < 90 | [12 empty] | 16
| 1 | 4 | < 90 | [12 empty] | < 90
| 1 | 5 | < 90 | [12 empty] | 13
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Eu | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 8
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for F in OREAS 25b (Indicative Value < 500 ppm) | [14 empty] | Indicative
F | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 500 | [12 empty] | 1
| 1 | 2 | < 500 | [12 empty] | 9
| 1 | 3 | < 500 | [12 empty] | 16
| 1 | 4 | 1000 | [12 empty] | < 500
| 1 | 5 | < 500 | [12 empty] | 15
| 1 | 6 | < 500
| Mean |  | 1000
| Median |  | 1000
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 25b (Certified Value 12.2 wt.%) | [14 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 12.22 | 12.2 | 12.09 | 12.1115 | 12.04 | 12.4 | 12.26 | 12.36 | 12.34 | 12.29 | 11.96 | 11.83 | 12.467 | 1
| 1 | 2 | 12.2 | 12.2 | 12.21 | 12.1188 | 11.99 | 12.28 | 12.28 | 12.29 | 12.45 | 12.36 | 11.97 | 11.9 | 12.453 | #N/A
| 1 | 3 | 12.21 | 12.2 | 12.17 | 12.1056 | 12.1 | 12.32 | 12.22 | 12.22 | 12.37 | 12.3 | 11.98 | 11.8 | 12.553 | 16
| 1 | 4 | 12.21 | 12.2 | 12.22 | 12.107 | 12.02 | 12.34 | 12.33 | 12.44 | 12.39 | 12.32 | 11.94 | 11.89 | 12.467 | 12.2027
| 1 | 5 | 12.22 | 12.1 | 12.13 | 12.1181 | 12.01 | 12.33 | 12.27 | 12.3 | 11.98 | 12.22 | 11.98 | 11.96 | 12.467 | 15
| 1 | 6 | 12.25 | 12.2 | 12.19 | 12.1104 | 12.03 | 12.22 | 12.35 | 12.32 | 12.48 | 12.22 | 11.96 | 11.87 | 12.41
| Mean |  | 12.2183 | 12.1833 | 12.1683 | 12.1119 | 12.0317 | 12.315 | 12.285 | 12.3217 | 12.335 | 12.285 | 11.965 | 11.875 | 12.4695
| Median |  | 12.215 | 12.2 | 12.18 | 12.1109 | 12.025 | 12.325 | 12.275 | 12.31 | 12.38 | 12.295 | 11.965 | 11.88 | 12.467
| Std Dev. |  | 0.017224 | 0.040825 | 0.049967 | 0.0055092 | 0.037639 | 0.060581 | 0.047645 | 0.073869 | 0.18141 | 0.055767 | 0.015166 | 0.056125 | 0.046483
| Rel.Std.Dev. |  | 0.141% | 0.3351% | 0.4106% | 0.04549% | 0.3128% | 0.4919% | 0.3878% | 0.5995% | 1.471% | 0.4539% | 0.1268% | 0.4726% | 0.3728%
| PDM3 |  | 0.1278% | -0.159% | -0.2819% | -0.7445% | -1.402% | 0.92% | 0.6741% | 0.9746% | 1.084% | 0.6741% | -1.948% | -2.686% | 2.186%
| Z-Score (Absolute) |  | 0 | 0.23 | 0.33 | 0.69 | 1.22 | 0.63 | 0.44 | 0.67 | 0.76 | 0.44 | 1.65 | 2.24 | 1.64
| Analytical results for Gd in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Gd | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 10
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 16
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Hf | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 11
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 25b (Certified Value 0.936 wt.%) | [14 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.94 | 0.94 | 0.92 | 0.89137 | 0.93 | 0.96 | 0.94 | 0.95 | 0.93 | 0.96 | 0.95 | 0.91 | 0.94 | 1
| 1 | 2 | 0.95 | 0.96 | 0.91 | 0.90749 | 0.92 | 0.95 | 0.93 | 0.95 | 0.94 | 0.96 | 0.95 | 0.91 | 0.928 | #N/A
| 1 | 3 | 0.94 | 0.95 | 0.91 | 0.91147 | 0.93 | 0.96 | 0.93 | 0.94 | 0.94 | 0.95 | 0.94 | 0.91 | 0.915 | 16
| 1 | 4 | 0.95 | 0.95 | 0.91 | 0.89749 | 0.92 | 0.94 | 0.94 | 0.95 | 0.94 | 0.96 | 0.95 | 0.91 | 0.94 | 0.93575
| 1 | 5 | 0.95 | 0.96 | 0.91 | 0.90219 | 0.92 | 0.95 | 0.94 | 0.94 | 0.91 | 0.96 | 0.94 | 0.91 | 0.94 | 16
| 1 | 6 | 0.95 | 0.96 | 0.92 | 0.88536 | 0.93 | 0.95 | 0.94 | 0.95 | 0.95 | 0.95 | 0.95 | 0.92 | 0.94
| Mean |  | 0.94667 | 0.95333 | 0.91333 | 0.89923 | 0.925 | 0.95167 | 0.93667 | 0.94667 | 0.935 | 0.95667 | 0.94667 | 0.91167 | 0.93383
| Median |  | 0.95 | 0.955 | 0.91 | 0.89984 | 0.925 | 0.95 | 0.94 | 0.95 | 0.94 | 0.96 | 0.95 | 0.91 | 0.94
| Std Dev. |  | 0.005164 | 0.008165 | 0.005164 | 0.0098393 | 0.0054772 | 0.0075277 | 0.005164 | 0.005164 | 0.013784 | 0.005164 | 0.005164 | 0.0040825 | 0.0104
| Rel.Std.Dev. |  | 0.5455% | 0.8565% | 0.5654% | 1.094% | 0.5921% | 0.791% | 0.5513% | 0.5455% | 1.474% | 0.5398% | 0.5455% | 0.4478% | 1.114%
| PDM3 |  | 1.167% | 1.879% | -2.396% | -3.903% | -1.149% | 1.701% | 0.09804% | 1.167% | -0.08007% | 2.235% | 1.167% | -2.574% | -0.2047%
| Z-Score (Absolute) |  | 0.58 | 0.96 | 1.35 | 2.16 | 0.67 | 0.87 | 0 | 0.58 | 0.1 | 1.16 | 0.58 | 1.44 | 0.18
| Analytical results for La in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
La | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 1
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 25b (Certified Value 0.627 wt.%) | [14 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.61 | 0.7 | 0.63 | 0.62532 | 0.58 | 0.64 | 0.65 | 0.63 | 0.63 | 0.65 | 0.62 | 0.61 | 0.61 | 1
| 1 | 2 | 0.61 | 0.72 | 0.62 | 0.62977 | 0.57 | 0.64 | 0.63 | 0.63 | 0.63 | 0.66 | 0.63 | 0.6 | 0.62 | #N/A
| 1 | 3 | 0.62 | 0.69 | 0.6 | 0.63149 | 0.57 | 0.64 | 0.64 | 0.62 | 0.63 | 0.64 | 0.64 | 0.6 | 0.63 | 16
| 1 | 4 | 0.61 | 0.69 | 0.62 | 0.62949 | 0.58 | 0.63 | 0.64 | 0.63 | 0.63 | 0.66 | 0.62 | 0.61 | 0.63 | 0.62689
| 1 | 5 | 0.61 | 0.71 | 0.62 | 0.63101 | 0.59 | 0.64 | 0.64 | 0.63 | 0.62 | 0.64 | 0.62 | 0.62 | 0.62 | 17
| 1 | 6 | 0.62 | 0.69 | 0.6 | 0.62784 | 0.57 | 0.64 | 0.62 | 0.65 | 0.64 | 0.64 | 0.63 | 0.6 | 0.61
| Mean |  | 0.61333 | 0.7 | 0.615 | 0.62915 | 0.57667 | 0.63833 | 0.63667 | 0.63167 | 0.63 | 0.64833 | 0.62667 | 0.60667 | 0.62
| Median |  | 0.61 | 0.695 | 0.62 | 0.62963 | 0.575 | 0.64 | 0.64 | 0.63 | 0.63 | 0.645 | 0.625 | 0.605 | 0.62
| Std Dev. |  | 0.005164 | 0.012649 | 0.012247 | 0.0022736 | 0.008165 | 0.0040825 | 0.010328 | 0.0098319 | 0.0063246 | 0.0098319 | 0.008165 | 0.008165 | 0.0089443
| Rel.Std.Dev. |  | 0.842% | 1.807% | 1.991% | 0.3614% | 1.416% | 0.6396% | 1.622% | 1.557% | 1.004% | 1.516% | 1.303% | 1.346% | 1.443%
| PDM3 |  | -2.163% | 11.66% | -1.897% | 0.3606% | -8.012% | 1.825% | 1.559% | 0.7615% | 0.4957% | 3.42% | -0.03606% | -3.226% | -1.1%
| Z-Score (Absolute) |  | 1.16 | 5.2 | 1.04 | 0 | 3.86 | 0.67 | 0.55 | 0.18 | 0.06 | 1.41 | 0.18 | 1.65 | 0.67
| Analytical results for MnO in OREAS 25b (Certified Value 0.03 wt.%) | [14 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.03 | 0.04 | 0.04 | 0.028281 | 0.035 | 0.04 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 1
| 1 | 2 | 0.03 | 0.03 | 0.04 | 0.027875 | 0.036 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | #N/A
| 1 | 3 | 0.03 | 0.03 | 0.04 | 0.028611 | 0.035 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 16
| 1 | 4 | 0.03 | 0.03 | 0.04 | 0.029085 | 0.035 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.030444
| 1 | 5 | 0.03 | 0.03 | 0.04 | 0.028161 | 0.034 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.02 | 0.03 | 18
| 1 | 6 | 0.03 | 0.03 | 0.04 | 0.027274 | 0.035 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03
| Mean |  | 0.03 | 0.031667 | 0.04 | 0.028215 | 0.035 | 0.033333 | 0.031667 | 0.03 | 0.03 | 0.03 | 0.036667 | 0.028333 | 0.026667
| Median |  | 0.03 | 0.03 | 0.04 | 0.028221 | 0.035 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03
| Std Dev. |  | 0 | 0.0040825 | 0 | 0.00062004 | 0.00063246 | 0.005164 | 0.0040825 | 0 | 0 | 0 | 0.005164 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 0% | 12.89% | 0% | 2.198% | 1.807% | 15.49% | 12.89% | 0% | 0% | 0% | 14.08% | 14.41% | 19.36%
| PDM3 |  | -1.458% | 4.017% | 31.39% | -7.322% | 14.97% | 9.492% | 4.017% | -1.458% | -1.458% | -1.458% | 20.44% | -6.932% | -12.41%
| Z-Score (Absolute) |  | 0 | 0.67 | 4.05 | 0.72 | 2.02 | 1.35 | 0.67 | 0 | 0 | 0 | 2.7 | 0.67 | 1.35
| Analytical results for Mo in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Mo | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 2
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 25b (Certified Value 0.094 wt.%) | [14 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.11 | < 0.1 | 0.09 | 0.086937 | 0.1 | 0.07 | 0.11 | 0.1 | 0.1 | < 0.01 | 0.04 | 0.07 | 0.081 | 1
| 1 | 2 | 0.1 | < 0.1 | 0.11 | 0.10221 | 0.09 | 0.08 | 0.1 | 0.1 | 0.1 | < 0.01 | 0.04 | 0.09 | 0.081 | #N/A
| 1 | 3 | 0.1 | 0.1 | 0.09 | 0.085834 | 0.1 | 0.08 | 0.1 | 0.1 | 0.1 | < 0.01 | 0.03 | 0.06 | 0.067 | 16
| 1 | 4 | 0.09 | 0.1 | 0.12 | 0.089333 | 0.09 | 0.07 | 0.1 | 0.09 | 0.1 | < 0.01 | 0.03 | 0.09 | 0.081 | 0.093651
| 1 | 5 | 0.11 | 0.1 | 0.11 | 0.087611 | 0.1 | 0.08 | 0.1 | 0.1 | 0.09 | < 0.01 | 0.04 | 0.1 | 0.081 | 19
| 1 | 6 | 0.11 | < 0.1 | 0.11 | 0.1007 | 0.09 | 0.08 | 0.1 | 0.1 | 0.11 | < 0.01 | 0.03 | 0.08 | 0.067
| Mean |  | 0.10333 | 0.1 | 0.105 | 0.092105 | 0.095 | 0.076667 | 0.10167 | 0.098333 | 0.1 |  | 0.035 | 0.081667 | 0.076333
| Median |  | 0.105 | 0.1 | 0.11 | 0.088472 | 0.095 | 0.08 | 0.1 | 0.1 | 0.1 |  | 0.035 | 0.085 | 0.081
| Std Dev. |  | 0.008165 | 1.6997e-17 | 0.012247 | 0.0073479 | 0.0054772 | 0.005164 | 0.0040825 | 0.0040825 | 0.0063246 |  | 0.0054772 | 0.01472 | 0.0072296
| Rel.Std.Dev. |  | 7.902% | 1.7e-14% | 11.66% | 7.978% | 5.766% | 6.736% | 4.016% | 4.152% | 6.325% |  | 15.65% | 18.02% | 9.471%
| PDM3 |  | 10.34% | 6.78% | 12.12% | -1.651% | 1.441% | -18.14% | 8.559% | 5% | 6.78% |  | -62.63% | -12.8% | -18.49%
| Z-Score (Absolute) |  | 0.67 | 1.03 | 0.77 | 0 | 0.17 | 0.93 | 0.57 | 0.37 | 0.47 | 5.23 | 3.43 | 0.63 | 0.94
| Analytical results for Nb in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Nb | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 3
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 9
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Nd | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 4
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 10
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 25b (Indicative Value 52 ppm) | [14 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 05 | 16 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | 55 | < 100 | [11 empty] | 1
| 1 | 2 | 39 | < 100 | [11 empty] | 5
| 1 | 3 | 55 | < 100 | [11 empty] | 16
| 1 | 4 | 71 | < 100 | [11 empty] | 52.3903
| 1 | 5 | 47 | < 100 | [11 empty] | 11
| 1 | 6 | 47 | < 100
| Mean |  | 52.3333
| Median |  | 51
| Std Dev. |  | 10.9301
| Rel.Std.Dev. |  | 20.89%
| PDM3 |  | -0.1088%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for P2O5 in OREAS 25b (Certified Value 0.1 wt.%) | [14 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.098 | 0.09 | 0.1 | 0.1016 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | 0.1 | 0.1 | 1
| 1 | 2 | 0.098 | 0.09 | 0.1 | 0.10427 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | 0.1 | 0.11 | #N/A
| 1 | 3 | 0.098 | 0.09 | 0.1 | 0.10423 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.09 | 0.09 | 0.1 | 0.1 | 16
| 1 | 4 | 0.099 | 0.09 | 0.1 | 0.10284 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | 0.11 | 0.10042
| 1 | 5 | 0.099 | 0.09 | 0.1 | 0.1023 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | 0.1 | 0.1 | 20
| 1 | 6 | 0.099 | 0.09 | 0.1 | 0.10175 | 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.09 | 0.09 | 0.1 | 0.11
| Mean |  | 0.0985 | 0.09 | 0.1 | 0.10283 | 0.1 | 0.1 | 0.10333 | 0.1 | 0.1 | 0.091667 | 0.095 | 0.1 | 0.105
| Median |  | 0.0985 | 0.09 | 0.1 | 0.10257 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.09 | 0.095 | 0.1 | 0.105
| Std Dev. |  | 0.00054772 | 1.5202e-17 | 1.5202e-17 | 0.0011823 | 1.5202e-17 | 1.5202e-17 | 0.005164 | 1.5202e-17 | 1.5202e-17 | 0.0040825 | 0.0054772 | 1.5202e-17 | 0.0054772
| Rel.Std.Dev. |  | 0.5561% | 1.689e-14% | 1.52e-14% | 1.15% | 1.52e-14% | 1.52e-14% | 4.997% | 1.52e-14% | 1.52e-14% | 4.454% | 5.766% | 1.52e-14% | 5.216%
| PDM3 |  | -1.916% | -10.38% | -0.4223% | 2.397% | -0.4223% | -0.4223% | 2.897% | -0.4223% | -0.4223% | -8.72% | -5.401% | -0.4223% | 4.557%
| Z-Score (Absolute) |  | 0.67 | 4.5 | 0 | 1.27 | 0 | 0 | 1.5 | 0 | 0 | 3.75 | 2.25 | 0 | 2.25
| Analytical results for Pb in OREAS 25b (Indicative Value 46.4 ppm) | [14 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 11 | 16 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | 37 | < 90 | [11 empty] | 1
| 1 | 2 | 46 | < 90 | [11 empty] | 6
| 1 | 3 | 56 | < 90 | [11 empty] | 16
| 1 | 4 | 37 | < 90 | [11 empty] | 46.4166
| 1 | 5 | 46 | < 90 | [11 empty] | 12
| 1 | 6 | 56 | < 90
| Mean |  | 46.3333
| Median |  | 46
| Std Dev. |  | 8.501
| Rel.Std.Dev. |  | 18.35%
| PDM3 |  | -0.1795%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Pr in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Pr | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 7
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Rb | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 8
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Sb | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 9
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 15
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 25b (Certified Value 58.09 wt.%) | [14 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 57.9 | 58.1 | 58.1 | 58.0827 | 56.13 | 58.52 | 57.39 | 58.15 | 57.84 | 58.38 | 56.52 | 56.07 | 58.2 | 1
| 1 | 2 | 58.01 | 58.3 | 58.17 | 58.0613 | 55.93 | 58.45 | 57.5 | 57.47 | 58.2 | 58.74 | 56.58 | 56.19 | 58.2 | #N/A
| 1 | 3 | 58.03 | 58.5 | 58.01 | 58.0573 | 56.46 | 58.63 | 57.6 | 57.37 | 58.03 | 58.53 | 56.66 | 55.82 | 58.1 | 16
| 1 | 4 | 57.9 | 57.9 | 58.04 | 58.0442 | 55.94 | 58.34 | 57.85 | 57.85 | 58.25 | 58.55 | 56.65 | 56.04 | 58 | 58.0934
| 1 | 5 | 58.19 | 58.1 | 58.09 | 58.0551 | 55.8 | 58.51 | 57.8 | 57.64 | 57.08 | 58.37 | 56.7 | 56.36 | 58.2 | 21
| 1 | 6 | 58.2 | 57.9 | 57.97 | 58.0629 | 56.14 | 58.27 | 57.79 | 57.94 | 58.58 | 58.31 | 56.74 | 56.21 | 58.1
| Mean |  | 58.0383 | 58.1333 | 58.0633 | 58.0606 | 56.0667 | 58.4533 | 57.655 | 57.7367 | 57.9967 | 58.48 | 56.6417 | 56.115 | 58.1333
| Median |  | 58.02 | 58.1 | 58.065 | 58.0593 | 56.035 | 58.48 | 57.695 | 57.745 | 58.115 | 58.455 | 56.655 | 56.13 | 58.15
| Std Dev. |  | 0.13288 | 0.23381 | 0.071461 | 0.012665 | 0.23235 | 0.13064 | 0.18684 | 0.29663 | 0.5122 | 0.15875 | 0.080104 | 0.18404 | 0.08165
| Rel.Std.Dev. |  | 0.2289% | 0.4022% | 0.1231% | 0.02181% | 0.4144% | 0.2235% | 0.3241% | 0.5138% | 0.8832% | 0.2715% | 0.1414% | 0.328% | 0.1405%
| PDM3 |  | -0.09477% | 0.06876% | -0.05174% | -0.0565% | -3.489% | 0.6196% | -0.7546% | -0.6141% | -0.1665% | 0.6655% | -2.499% | -3.406% | 0.06876%
| Z-Score (Absolute) |  | 0 | 0.21 | 0.06 | 0.05 | 4.41 | 0.93 | 0.86 | 0.67 | 0.09 | 0.99 | 3.12 | 4.3 | 0.21
| Analytical results for Sm in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Sm | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 10
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 16
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Sn | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 11
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 17
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 25b (Certified Value 0.058 wt.%) | [14 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 15 | 16 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.06 | < 0.01 | 0.068376 | 0.044 | 0.07 | 0.06 | 0.05 | 0.06 | < 0.01 | 0.06 |  |  |  | 1
| 1 | 2 | 0.06 | < 0.01 | 0.070555 | 0.04 | 0.06 | 0.07 | 0.05 | 0.06 | < 0.01 | 0.06 |  |  |  | #N/A
| 1 | 3 | 0.06 | < 0.01 | 0.072537 | 0.046 | 0.06 | 0.06 | 0.05 | 0.05 | < 0.01 | 0.06 |  |  |  | 16
| 1 | 4 | 0.06 | < 0.01 | 0.070989 | 0.044 | 0.06 | 0.06 | 0.05 | 0.05 | < 0.01 | 0.07 |  |  |  | 0.058113
| 1 | 5 | 0.06 | < 0.01 | 0.072342 | 0.033 | 0.06 | 0.06 | 0.05 | 0.05 | < 0.01 | 0.06 |  |  |  | 22
| 1 | 6 | 0.06 | < 0.01 | 0.070619 | 0.046 | 0.07 | 0.06 | 0.05 | 0.05 | < 0.01 | 0.06
| Mean |  | 0.06 |  | 0.070903 | 0.042167 | 0.063333 | 0.061667 | 0.05 | 0.053333 |  | 0.061667
| Median |  | 0.06 |  | 0.070804 | 0.044 | 0.06 | 0.06 | 0.05 | 0.05 |  | 0.06
| Std Dev. |  | 0 |  | 0.0015055 | 0.0049967 | 0.005164 | 0.0040825 | 7.6012e-18 | 0.005164 |  | 0.0040825
| Rel.Std.Dev. |  | 0% |  | 2.123% | 11.85% | 8.154% | 6.62% | 1.52e-14% | 9.682% |  | 6.62%
| PDM3 |  | 3.247% |  | 22.01% | -27.44% | 8.983% | 6.115% | -13.96% | -8.225% |  | 6.115%
| Z-Score (Absolute) |  | 0.34 | 5.23 | 1.44 | 1.47 | 0.67 | 0.51 | 0.67 | 0.34 | 5.23 | 0.51
| Analytical results for SrO in OREAS 25b (Certified Value < 100 ppm) | [14 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 06 | 07 | 08 | 09 | 10 | 11 | 16 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 100 | 95 | < 100 | 100 | < 100 | 100 | < 100 | < 120 | [5 empty] | 1
| 1 | 2 | 100 | 83 | < 100 | 100 | < 100 | 100 | < 100 | < 120 | [5 empty] | #N/A
| 1 | 3 | 100 | 95 | < 100 | 100 | < 100 | 100 | < 100 | < 120 | [5 empty] | 16
| 1 | 4 | 100 | 106 | < 100 | < 100 | < 100 | 100 | < 100 | < 120 | [5 empty] | < 100
| 1 | 5 | 100 | 95 | < 100 | 100 | < 100 | 100 | < 100 | < 120 | [5 empty] | 23
| 1 | 6 | 100 | 106 | < 100 | 200 | 100 | 100 | < 100 | < 120
| Mean |  | 100 | 96.6667 |  | 120 | 100 | 100
| Median |  | 100 | 95 |  | 100 | 100 | 100
| Std Dev. |  | 0 | 8.5946 |  | 44.7214 |  | 0
| Rel.Std.Dev. |  | 0% | 8.891% |  | 37.27% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.57 | 0.85 | 0.93 | 0.59 | 0.67 | 0.85 | 0.57
| Analytical results for Ta in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Ta | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 1
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 25b (Indicative Value < 50 ppm) | [14 empty] | Indicative
Th | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | [12 empty] | 2
| 1 | 3 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | [12 empty] | < 50
| 1 | 5 | < 50 | [12 empty] | 8
| 1 | 6 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 25b (Certified Value 1.88 wt.%) | [14 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.87 | 1.87 | 1.94 | 1.8654 | 1.88 | 1.91 | 1.88 | 1.88 | 1.84 | 1.88 | 1.89 | 1.85 | 1.87 | 1
| 1 | 2 | 1.87 | 1.89 | 1.93 | 1.865 | 1.87 | 1.9 | 1.88 | 1.86 | 1.86 | 1.88 | 1.91 | 1.86 | 1.91 | #N/A
| 1 | 3 | 1.86 | 1.88 | 1.93 | 1.8761 | 1.89 | 1.91 | 1.87 | 1.85 | 1.85 | 1.88 | 1.89 | 1.83 | 1.9 | 16
| 1 | 4 | 1.87 | 1.9 | 1.92 | 1.8754 | 1.88 | 1.9 | 1.89 | 1.89 | 1.86 | 1.93 | 1.9 | 1.83 | 1.9 | 1.8783
| 1 | 5 | 1.88 | 1.89 | 1.93 | 1.883 | 1.86 | 1.91 | 1.88 | 1.86 | 1.81 | 1.88 | 1.91 | 1.85 | 1.92 | 24
| 1 | 6 | 1.88 | 1.87 | 1.93 | 1.8832 | 1.87 | 1.9 | 1.9 | 1.87 | 1.89 | 1.88 | 1.9 | 1.84 | 1.87
| Mean |  | 1.8717 | 1.8833 | 1.93 | 1.8747 | 1.875 | 1.905 | 1.8833 | 1.8683 | 1.8517 | 1.8883 | 1.9 | 1.8433 | 1.895
| Median |  | 1.87 | 1.885 | 1.93 | 1.8757 | 1.875 | 1.905 | 1.88 | 1.865 | 1.855 | 1.88 | 1.9 | 1.845 | 1.9
| Std Dev. |  | 0.0075277 | 0.012111 | 0.0063246 | 0.0080583 | 0.010488 | 0.0054772 | 0.010328 | 0.01472 | 0.026394 | 0.020412 | 0.0089443 | 0.012111 | 0.020736
| Rel.Std.Dev. |  | 0.4022% | 0.643% | 0.3277% | 0.4299% | 0.5594% | 0.2875% | 0.5484% | 0.7878% | 1.425% | 1.081% | 0.4708% | 0.657% | 1.094%
| PDM3 |  | -0.3535% | 0.2677% | 2.752% | -0.1935% | -0.176% | 1.421% | 0.2677% | -0.5309% | -1.418% | 0.5339% | 1.155% | -1.862% | 0.8888%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.7 | 0.5 | 0.48 | 1.25 | 0 | 0.87 | 1.83 | 0.29 | 0.96 | 2.31 | 0.67
| Analytical results for TOT_XRF in OREAS 25b (Indicative Value 99.9 wt.%) | [14 empty] | Indicative
TOT_XRF | Round | Replicate | Lab | Lab | Lab | [10 empty] | 1
| No | No | 07 | 08 | 09 | [10 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 2
2
| 1 | 1 | 99.38 | 100.3 | 100.25 | [10 empty] | 1
| 1 | 2 | 99.57 | 99.2 | 100.8 | [10 empty] | 3
| 1 | 3 | 99.55 | 98.83 | 100.4 | [10 empty] | 16
| 1 | 4 | 99.89 | 99.82 | 100.75 | [10 empty] | 99.8961
| 1 | 5 | 99.79 | 99.45 | 98.78 | [10 empty] | 9
| 1 | 6 | 99.98 | 99.94 | 101.45
| Mean |  | 99.6933 | 99.59 | 100.405
| Median |  | 99.68 | 99.635 | 100.575
| Std Dev. |  | 0.22984 | 0.53457 | 0.89797
| Rel.Std.Dev. |  | 0.2305% | 0.5368% | 0.8943%
| PDM3 |  | -0.203% | -0.3064% | 0.5094%
| Z-Score (Absolute) |  | 0 | 0.67 | 4.64
| Analytical results for U in OREAS 25b (Indicative Value < 50 ppm) | [14 empty] | Indicative
U | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 50 | [12 empty] | 1
| 1 | 2 | < 50 | [12 empty] | 4
| 1 | 3 | < 50 | [12 empty] | 16
| 1 | 4 | < 50 | [12 empty] | < 50
| 1 | 5 | < 50 | [12 empty] | 10
| 1 | 6 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 25b (Indicative Value 296 ppm) | [14 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 05 | 11 | 15 | 16 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 309.8306 | 280 | 300 | 200 | 400 | [8 empty] | 1
| 1 | 2 | 305.3879 | 280 | 300 | 200 | 300 | [8 empty] | 5
| 1 | 3 | 291.8348 | 290 | 300 | 200 | 400 | [8 empty] | 16
| 1 | 4 | 317.861 | 270 | 300 | 200 | 300 | [8 empty] | 295.9621
| 1 | 5 | 297.8785 | 270 | 300 | 200 | 300 | [8 empty] | 11
| 1 | 6 | 310.2967 | 280 | 300 | 300 | 300
| Mean |  | 305.5149 | 278.3333 | 300 | 216.6667 | 333.3333
| Median |  | 307.6093 | 280 | 300 | 200 | 300
| Std Dev. |  | 9.3744 | 7.5277 | 0 | 40.8248 | 51.6398
| Rel.Std.Dev. |  | 3.068% | 2.705% | 0% | 18.84% | 15.49%
| PDM3 |  | 3.228% | -5.956% | 1.364% | -26.79% | 12.63%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 0 | 2.59 | 1.04
| Analytical results for W in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
W | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 6
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 12
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Y | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 7
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 25b (Indicative Value < 100 ppm) | [14 empty] | Indicative
Zn | Round | Replicate | Lab | [12 empty] | 1
| No | No | 16 | [12 empty] | ppm
|  |  | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | [12 empty] | 1
| 1 | 2 | < 100 | [12 empty] | 8
| 1 | 3 | < 100 | [12 empty] | 16
| 1 | 4 | < 100 | [12 empty] | < 100
| 1 | 5 | < 100 | [12 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for ZrO2 in OREAS 25b (Indicative Value 439 ppm) | [14 empty] | Indicative
ZrO2 | Round | Replicate | Lab | Lab | [11 empty] | 1
| No | No | 04 | 16 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | 471.3448 | 410 | [11 empty] | 1
| 1 | 2 | 474.9102 | 410 | [11 empty] | 9
| 1 | 3 | 470.042 | 410 | [11 empty] | 16
| 1 | 4 | 474.3192 | 410 | [11 empty] | 438.6159
| 1 | 5 | 469.3464 | 410 | [11 empty] | 15
| 1 | 6 | 472.0538 | 410
| Mean |  | 472.0027 | 410
| Median |  | 471.6993 | 410
| Std Dev. |  | 2.243 | 0
| Rel.Std.Dev. |  | 0.4752% | 0%
| PDM3 |  | 7.612% | -6.524%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 25b (Certified Value 8.49 wt.%) | [14 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 15 | 16 | wt.%
|  |  | LOI*TGA | LOI*Furnace | BF*XRF | LOI*Furnace | BF*XRF | BF*XRF | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | BF*OES | LOI*TGA | BF*XRF | 2
3
| 1 | 1 | 8.35 | 8.91 | 8.49 | 8.455 | 10.36 | 8.7 | 8.4 | 8.54 | 8.78 | 8.33 | 10.3 | 10.9869 | 8.15 | 1
| 1 | 2 | 8.34 | 8.66 | 8.64 | 8.5405 | 10.38 | 8.6 | 8.5 | 8.39 | 8.78 | 8.34 | 10.2 | 11.0066 | 8.22 | #N/A
| 1 | 3 | 8.31 | 8.63 | 8.66 | 8.5215 | 10.36 | 8.6 | 8.46 | 8.23 | 8.64 | 8.3 | 10.3 | 10.98 | 8.21 | 16
| 1 | 4 | 8.36 | 8.62 | 8.65 | 8.493 | 10.34 | 8.6 | 8.33 | 8.45 | 8.76 | 8.32 | 10.2 | 10.9326 | 8.2 | 8.4942
| 1 | 5 | 8.33 | 8.62 | 8.64 | 8.493 | 10.36 | 8.7 | 8.37 | 8.34 | 8.8 | 8.31 | 10.2 | 11.0484 | 8.23 | 26
| 1 | 6 | 8.36 | 8.99 | 8.67 | 8.569 | 10.35 | 8.6 | 8.4 | 8.42 | 8.83 | 8.31 | 10.2 | 11.0615 | 8.21
| Mean |  | 8.3417 | 8.7383 | 8.625 | 8.512 | 10.3583 | 8.6333 | 8.41 | 8.395 | 8.765 | 8.3183 | 10.2333 | 11.0027 | 8.2033
| Median |  | 8.345 | 8.645 | 8.645 | 8.5072 | 10.36 | 8.6 | 8.4 | 8.405 | 8.78 | 8.315 | 10.2 | 10.9968 | 8.21
| Std Dev. |  | 0.019408 | 0.16654 | 0.067157 | 0.040305 | 0.013292 | 0.05164 | 0.061319 | 0.10483 | 0.065651 | 0.01472 | 0.05164 | 0.047413 | 0.028048
| Rel.Std.Dev. |  | 0.2327% | 1.906% | 0.7786% | 0.4735% | 0.1283% | 0.5981% | 0.7291% | 1.249% | 0.749% | 0.177% | 0.5046% | 0.4309% | 0.3419%
| PDM3 |  | -1.796% | 2.874% | 1.54% | 0.2096% | 21.95% | 1.638% | -0.9913% | -1.168% | 3.188% | -2.07% | 20.47% | 29.53% | -3.424%
| Z-Score (Absolute) |  | 0.83 | 0.33 | 0 | 0.33 | 5.08 | 0.02 | 0.63 | 0.67 | 0.41 | 0.9 | 4.72 | 6.97 | 1.24
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 25b (Indicative Value 5.48 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | PF*MS | BF*OES/MS | PF*OES | [12 empty] | 2
2
| 1 | 1 | < 5 | 0.5 | 11 | [12 empty] | 1
| 1 | 2 | < 5 | < 0.5 | 12 | [12 empty] | 13
| 1 | 3 | < 5 | 0.5 | 6 | [12 empty] | 16
| 1 | 4 | < 5 | 0.5 | 12 | [12 empty] | 5.4792
| 1 | 5 | < 5 | 0.5 | 11 | [12 empty] | 19
| 1 | 6 | < 5 | 0.5 | 11
| Mean |  |  | 0.5 | 10.5
| Median |  |  | 0.5 | 11
| Std Dev. |  |  | 0 | 2.2583
| Rel.Std.Dev. |  |  | 0% | 21.51%
| PDM3 |  |  | -90.87% | 91.63%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.64
| Analytical results for Al in OREAS 25b (Certified Value 8.73 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 15 | 16 | [5 empty] | wt.%
|  |  | BF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 2
3
| 1 | 1 | 8.854 | 8.674 | 8.61 | 8.8821 | 8.669 | 8.912 | 8.4149 | 9 | 8.9 | 8.3 | [5 empty] | 1
| 1 | 2 | 8.558 | 8.484 | 8.65 | 8.882 | 8.891 | 8.859 | 8.362 | 8.93 | 8.93 | 8.44 | [5 empty] | #N/A
| 1 | 3 | 8.653 | 8.388 | 8.63 | 9.0229 | 8.918 | 8.796 | 8.6266 | 9.14 | 6.77 | 8.78 | [5 empty] | 16
| 1 | 4 | 8.627 | 9.267 | 8.54 | 8.9653 | 8.949 | 8.791 | 8.6795 | 9.12 | 8.76 | 8.53 | [5 empty] | 8.7314
| 1 | 5 | 8.722 | 8.568 | 8.65 | 8.9601 | 8.664 | 8.859 | 7.674 | 9.09 | 8.49 | 8.46 | [5 empty] | 28
| 1 | 6 | 8.457 | 8.235 | 8.66 | 9.0501 | 8.859 | 8.918 | 8.9442 | 9.08 | 8.84 | 8.4
| Mean |  | 8.6452 | 8.6027 | 8.6233 | 8.9604 | 8.825 | 8.8558 | 8.4502 | 9.06 | 8.4483 | 8.485
| Median |  | 8.64 | 8.526 | 8.64 | 8.9627 | 8.875 | 8.859 | 8.5207 | 9.085 | 8.8 | 8.45
| Std Dev. |  | 0.13632 | 0.3586 | 0.044572 | 0.06967 | 0.12633 | 0.054448 | 0.43342 | 0.07975 | 0.83727 | 0.16319
| Rel.Std.Dev. |  | 1.577% | 4.168% | 0.5169% | 0.7775% | 1.431% | 0.6148% | 5.129% | 0.8802% | 9.91% | 1.923%
| PDM3 |  | -0.988% | -1.475% | -1.238% | 2.623% | 1.072% | 1.425% | -3.221% | 3.763% | -3.242% | -2.822%
| Z-Score (Absolute) |  | 0.04 | 0.11 | 0.04 | 1.19 | 0.7 | 0.81 | 0.67 | 1.55 | 0.68 | 0.54
| Analytical results for As in OREAS 25b (Indicative Value 13.8 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 04 | 05 | 12 | 15 | 16 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | [9 empty] | 1
1
| 1 | 1 | < 20 | < 20 | 9 | < 20 | 13 | 10 | [9 empty] | 1
| 1 | 2 | 20 | < 20 | 9 | < 20 | 14 | 20 | [9 empty] | 1
| 1 | 3 | < 20 | < 20 | 11 | < 20 | 14 | 15 | [9 empty] | 16
| 1 | 4 | < 20 | < 20 | 10 | < 20 | 14 | 20 | [9 empty] | 13.7917
| 1 | 5 | < 20 | < 20 | 10 | < 20 | 15 | 30 | [9 empty] | 20
| 1 | 6 | < 20 | < 20 | 9 | < 20 | 13 | 25
| Mean |  | 20 |  | 9.6667 |  | 13.8333 | 20
| Median |  | 20 |  | 9.5 |  | 14 | 20
| Std Dev. |  |  |  | 0.8165 |  | 0.75277 | 7.0711
| Rel.Std.Dev. |  |  |  | 8.447% |  | 5.442% | 35.36%
| PDM3 |  | 45.02% |  | -29.91% |  | 0.3021% | 45.02%
| Z-Score (Absolute) |  | 0.56 | 0.56 | 0.79 | 0.56 | 2.02 | 6.18
| Analytical results for B in OREAS 25b (Indicative Value 80 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 04 | 12 | 16 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | [11 empty] | 0
0
| 1 | 1 | 86 | < 50 | 80 | 65 | [11 empty] | 1
| 1 | 2 | 81 | < 50 | 80 | 78 | [11 empty] | 15
| 1 | 3 | 90 | < 50 | 90 | 71 | [11 empty] | 16
| 1 | 4 | 88 | < 50 | 80 | 68 | [11 empty] | 79.5858
| 1 | 5 | 81 | < 50 | 80 | 75 | [11 empty] | 21
| 1 | 6 | 88 | < 50 | 80 | 71
| Mean |  | 85.6667 |  | 81.6667 | 71.3333
| Median |  | 87 |  | 80 | 71
| Std Dev. |  | 3.8297 |  | 4.0825 | 4.6762
| Rel.Std.Dev. |  | 4.47% |  | 4.999% | 6.555%
| PDM3 |  | 7.641% |  | 2.615% | -10.37%
| Z-Score (Absolute) |  | 0.86 | 4.88 | 0.48 | 0.48
| Analytical results for Ba in OREAS 25b (Certified Value 210 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | BF*MS | BF*OES | PF*OES/MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | PF*OES | 0
0
| 1 | 1 | 212.4 | 206 | 339 | 201.2315 | 219 | 204 | 212 | 213 | 198 | 201 | 223 | 270 | 211 | 199 | 242 | 1
| 1 | 2 | 222.4 | 224 | 341 | 199.4854 | 219 | 208 | 207 | 208 | 197.5 | 207 | 219 | 271 | 212 | 199 | 215 | #N/A
| 1 | 3 | 214.8 | 216 | 336 | 204.2545 | 223 | 207 | 223 | 213 | 203 | 211 | 216 | 267 | 209 | 159 | 206 | 16
| 1 | 4 | 217.4 | 242 | 337 | 198.8489 | 221 | 198 | 214 | 203 | 193.5 | 218 | 207 | 269 | 213 | 199 | 215 | 210.4738
| 1 | 5 | 220.1 | 206 | 335 | 200.8141 | 221 | 208 | 215 | 213 | 206 | 218 | 231 | 288 | 213 | 193 | 215 | 29
| 1 | 6 | 215.8 | 216 | 340 | 204.0264 | 222 | 205 | 216 | 212 | 202 | 211 | 223 | 264 | 211 | 199 | 215
| Mean |  | 217.15 | 218.3333 | 338 | 201.4435 | 220.8333 | 205 | 214.5 | 210.3333 | 200 | 211 | 219.8333 | 271.5 | 211.5 | 191.3333 | 218
| Median |  | 216.6 | 216 | 338 | 201.0228 | 221 | 206 | 214.5 | 212.5 | 200 | 211 | 221 | 269.5 | 211.5 | 199 | 215
| Std Dev. |  | 3.6396 | 13.471 | 2.3664 | 2.2619 | 1.6021 | 3.7947 | 5.244 | 4.0825 | 4.5056 | 6.5422 | 8.0602 | 8.4558 | 1.5166 | 16.0208 | 12.2963
| Rel.Std.Dev. |  | 1.676% | 6.17% | 0.7001% | 1.123% | 0.7255% | 1.851% | 2.445% | 1.941% | 2.253% | 3.101% | 3.667% | 3.114% | 0.7171% | 8.373% | 5.641%
| PDM3 |  | 3.172% | 3.734% | 60.59% | -4.29% | 4.922% | -2.601% | 1.913% | -0.06675% | -4.976% | 0.25% | 4.447% | 28.99% | 0.4876% | -9.094% | 3.576%
| Z-Score (Absolute) |  | 0.34 | 0.48 | 15.61 | 1.65 | 0.8 | 1.2 | 0 | 0.53 | 1.83 | 0.44 | 0.67 | 7.21 | 0.38 | 2.93 | 0.44
| Analytical results for Be in OREAS 25b (Certified Value 1.52 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 12 | 14 | 15 | 16 | [7 empty] | ppm
|  |  | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | PF*MS | BF*MS | PF*MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 1.4 | 1.3694 | 2 | 3 | 2 | 1 | 1 | 1.3 | [7 empty] | 1
| 1 | 2 | 1.5 | 1.3606 | 2 | < 1 | 2 | 1 | 1 | 1.3 | [7 empty] | #N/A
| 1 | 3 | 0.9 | 1.5159 | 2 | 3 | 1 | 1 | 1 | 1.2 | [7 empty] | 16
| 1 | 4 | 1.9 | 1.4863 | 2 | 2 | 2 | 1 | 1 | 1.2 | [7 empty] | 1.5187
| 1 | 5 | 1.4 | 1.4644 | 2 | 1 | 2 | 1 | 1 | 1.3 | [7 empty] | 30
| 1 | 6 | 1.4 | 1.5005 | 2 | 2 | 2 | 1 | 1 | 1.2
| Mean |  | 1.4167 | 1.4495 | 2 | 2.2 | 1.8333 | 1 | 1 | 1.25
| Median |  | 1.4 | 1.4753 | 2 | 2 | 2 | 1 | 1 | 1.25
| Std Dev. |  | 0.31885 | 0.067685 | 0 | 0.83666 | 0.40825 | 0 | 0 | 0.054772
| Rel.Std.Dev. |  | 22.51% | 4.67% | 0% | 38.03% | 22.27% | 0% | 0% | 4.382%
| PDM3 |  | -6.718% | -4.556% | 31.69% | 44.86% | 20.72% | -34.15% | -34.15% | -17.69%
| Z-Score (Absolute) |  | 0.03 | 0.03 | 0.92 | 0.78 | 0.65 | 0.7 | 0.7 | 0.3
| Analytical results for Bi in OREAS 25b (Indicative Value 0.43 ppm) | [16 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 04 | 05 | 12 | 16 | [10 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | [10 empty] | 2
2
| 1 | 1 | 0.5 | 0.35328 | < 0.4 | 10000 | < 0.5 | [10 empty] | 1
| 1 | 2 | 0.5 | 0.40486 | < 0.4 | 13000 | < 0.5 | [10 empty] | 2
| 1 | 3 | 0.7 | 0.22294 | < 0.4 | 10000 | < 0.5 | [10 empty] | 16
| 1 | 4 | 0.5 | 0.21271 | < 0.4 | 12000 | < 0.5 | [10 empty] | 0.43246
| 1 | 5 | 0.6 | 0.21997 | < 0.4 | 13000 | < 0.5 | [10 empty] | 22
| 1 | 6 | 0.5 | 0.47575 | < 0.4 | 13000 | < 0.5
| Mean |  | 0.55 | 0.31492 |  | 11833.3333
| Median |  | 0.5 | 0.28811 |  | 12500
| Std Dev. |  | 0.083666 | 0.11256 |  | 1471.9601
| Rel.Std.Dev. |  | 15.21% | 35.74% |  | 12.44%
| PDM3 |  | 27.18% | -27.18% |  | 2736190%
| Z-Score (Absolute) |  | 1.38 | 0 | 0.67 | 69432.99 | 0.38
| Analytical results for Ca in OREAS 25b (Certified Value 0.29 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 15 | 16 | [5 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 3
3
| 1 | 1 | 0.2 | 0.243 | 0.28 | 0.31585 | 0.257 | 0.25 | 0.2937 | 0.3 | 0.4 | 0.36 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.25 | 0.27 | 0.31344 | 0.25 | 0.25 | 0.2852 | 0.3 | 0.4 | 0.36 | [5 empty] | #N/A
| 1 | 3 | 0.3 | 0.243 | 0.27 | 0.31085 | 0.25 | 0.257 | 0.2995 | 0.3 | 0.2 | 0.36 | [5 empty] | 16
| 1 | 4 | 0.3 | 0.286 | 0.27 | 0.31422 | 0.25 | 0.25 | 0.2987 | 0.3 | 0.3 | 0.36 | [5 empty] | 0.29034
| 1 | 5 | 0.3 | 0.243 | 0.27 | 0.30248 | 0.25 | 0.25 | 0.3045 | 0.3 | 0.3 | 0.36 | [5 empty] | 31
| 1 | 6 | 0.3 | 0.243 | 0.28 | 0.3019 | 0.25 | 0.25 | 0.2995 | 0.3 | 0.4 | 0.36
| Mean |  | 0.28333 | 0.25133 | 0.27333 | 0.30979 | 0.25117 | 0.25117 | 0.29685 | 0.3 | 0.33333 | 0.36
| Median |  | 0.3 | 0.243 | 0.27 | 0.31215 | 0.25 | 0.25 | 0.2991 | 0.3 | 0.35 | 0.36
| Std Dev. |  | 0.040825 | 0.017212 | 0.005164 | 0.0061074 | 0.0028577 | 0.0028577 | 0.0066579 | 0 | 0.08165 | 6.0809e-17
| Rel.Std.Dev. |  | 14.41% | 6.848% | 1.889% | 1.971% | 1.138% | 1.138% | 2.243% | 0% | 24.49% | 1.689e-14%
| PDM3 |  | -2.413% | -13.43% | -5.858% | 6.699% | -13.49% | -13.49% | 2.242% | 3.327% | 14.81% | 23.99%
| Z-Score (Absolute) |  | 0.16 | 0.89 | 0.39 | 0.45 | 0.89 | 0.89 | 0.16 | 0.23 | 1 | 1.55
| Analytical results for Cd in OREAS 25b (Indicative Value < 10 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 04 | 16 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | [12 empty] | 1
1
| 1 | 1 | < 1 | < 10 | < 10 | [12 empty] | 1
| 1 | 2 | < 1 | < 10 | < 10 | [12 empty] | 17
| 1 | 3 | < 1 | < 10 | < 10 | [12 empty] | 16
| 1 | 4 | < 1 | < 10 | < 10 | [12 empty] | < 10
| 1 | 5 | < 1 | < 10 | < 10 | [12 empty] | 23
| 1 | 6 | < 1 | < 10 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Ce in OREAS 25b (Certified Value 53 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES |  |  |  | 0
1
| 1 | 1 | 53.8 | 46.6 | 53.2 | 51.3 | 53.1 | 56.8 | 54.9 | 50.9 | 63.5 | 49.2 | 47.1 | 45 |  |  |  | 1
| 1 | 2 | 53.3 | 54.9 | 53.9 | 51 | 52.8 | 55.4 | 55 | 51.9 | 62.6 | 50.1 | 50.1 | 45 |  |  |  | 3
| 1 | 3 | 53.5 | 53.4 | 55.5 | 52.5 | 54.5 | 56.7 | 54.6 | 53.8 | 57.8 | 49.3 | 50.3 | 40 |  |  |  | 16
| 1 | 4 | 52.7 | 51.9 | 52.2 | 49.7 | 54.4 | 54 | 53.7 | 53.9 | 57.6 | 52.2 | 54.2 | 40 |  |  |  | 52.6057
| 1 | 5 | 53.2 | 45.8 | 54 | 52.5 | 53.9 | 55.8 | 53.5 | 53.2 | 61.2 | 49.8 | 49.9 | 45 |  |  |  | 32
| 1 | 6 | 53.9 | 53.2 | 53.4 | 51.5 | 55.2 | 54.5 | 53.1 | 51.1 | 59.6 | 50.3 | 52.7 | 45
| Mean |  | 53.4 | 50.9667 | 53.7 | 51.4167 | 53.9833 | 55.5333 | 54.1333 | 52.4667 | 60.3833 | 50.15 | 50.7167 | 43.3333
| Median |  | 53.4 | 52.55 | 53.65 | 51.4 | 54.15 | 55.6 | 54.15 | 52.55 | 60.4 | 49.95 | 50.2 | 45
| Std Dev. |  | 0.43818 | 3.8213 | 1.0918 | 1.0477 | 0.90646 | 1.1378 | 0.80166 | 1.3426 | 2.4629 | 1.0932 | 2.4645 | 2.582
| Rel.Std.Dev. |  | 0.8206% | 7.498% | 2.033% | 2.038% | 1.679% | 2.049% | 1.481% | 2.559% | 4.079% | 2.18% | 4.859% | 5.958%
| PDM3 |  | 1.51% | -3.116% | 2.08% | -2.26% | 2.619% | 5.565% | 2.904% | -0.2642% | 14.78% | -4.668% | -3.591% | -17.63%
| Z-Score (Absolute) |  | 0.18 | 0.76 | 0.3 | 0.59 | 0.41 | 1.01 | 0.46 | 0.18 | 2.88 | 1.08 | 0.86 | 3.72
| Analytical results for Co in OREAS 25b (Certified Value 9.91 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS |  |  |  |  | 2
3
| 1 | 1 | 8.9 | 12.1 | 23 | 9.6128 | 10 | 9.2 | 11.9 | 12 | 8.3 | 8.6 | 10 |  |  |  |  | 1
| 1 | 2 | 9.8 | 9.9 | 23.7 | 9.2514 | 10 | 9.1 | 10.3 | 11 | 7.4 | 9.5 | 10 |  |  |  |  | 5
| 1 | 3 | 9.8 | 10 | 24.6 | 9.3714 | 10 | 9 | 10.1 | 12 | 7.7 | 8.8 | 10 |  |  |  |  | 16
| 1 | 4 | 10.3 | 9.4 | 23.1 | 9.5473 | 10 | 8.5 | 9.3 | 11 | 7.3 | 9.9 | 10 |  |  |  |  | 9.9147
| 1 | 5 | 9.5 | 8.2 | 23.2 | 9.2101 | 10 | 9.1 | 11 | 11 | 7.8 | 10.6 | 10 |  |  |  |  | 33
| 1 | 6 | 10 | 10.1 | 24.9 | 9.2417 | 9 | 9.4 | 10.8 | 11 | 7.1 | 9.4 | 9
| Mean |  | 9.7167 | 9.95 | 23.75 | 9.3724 | 9.8333 | 9.05 | 10.5667 | 11.3333 | 7.6 | 9.4667 | 9.8333
| Median |  | 9.8 | 9.95 | 23.45 | 9.3114 | 10 | 9.1 | 10.55 | 11 | 7.55 | 9.45 | 10
| Std Dev. |  | 0.47924 | 1.2661 | 0.8167 | 0.17114 | 0.40825 | 0.30166 | 0.88468 | 0.5164 | 0.42895 | 0.73121 | 0.40825
| Rel.Std.Dev. |  | 4.932% | 12.72% | 3.439% | 1.826% | 4.152% | 3.333% | 8.372% | 4.556% | 5.644% | 7.724% | 4.152%
| PDM3 |  | -1.998% | 0.3559% | 139.5% | -5.469% | -0.8208% | -8.722% | 6.576% | 14.31% | -23.35% | -4.519% | -0.8208%
| Z-Score (Absolute) |  | 0.17 | 0.17 | 20.36 | 0.67 | 0 | 1.15 | 1.07 | 2.19 | 3.27 | 0.54 | 0
| Analytical results for Cr in OREAS 25b (Certified Value 219 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  | ppm
|  |  | BF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES | PF*OES | PF*OES | PF*OES |  | 0
0
| 1 | 1 | 230 | 210 | 169 | 240.7326 | 220 | 205 | 225 | 235 | 228 | 204 | 205 | 210 | 209 | 230 |  | 1
| 1 | 2 | 223 | 270 | 170 | 235.9264 | 220 | 205 | 224 | 232 | 226 | 205 | 192 | 210 | 204 | 235 |  | #N/A
| 1 | 3 | 225 | 270 | 168 | 236.2525 | 220 | 205 | 229 | 236 | 227 | 211 | 226 | 210 | 164 | 220 |  | 16
| 1 | 4 | 230 | 270 | 167 | 238.1127 | 220 | 205 | 227 | 227 | 228 | 221 | 205 | 220 | 209 | 235 |  | 219.1976
| 1 | 5 | 225 | 210 | 167 | 239.0014 | 220 | 212 | 230 | 236 | 227 | 213 | 219 | 210 | 193 | 235 |  | 34
| 1 | 6 | 215 | 210 | 169 | 240.3416 | 210 | 205 | 233 | 232 | 210 | 207 | 212 | 210 | 214 | 230
| Mean |  | 224.6667 | 240 | 168.3333 | 238.3945 | 218.3333 | 206.1667 | 228 | 233 | 224.3333 | 210.1667 | 209.8333 | 211.6667 | 198.8333 | 230.8333
| Median |  | 225 | 240 | 168.5 | 238.557 | 220 | 205 | 228 | 233.5 | 227 | 209 | 208.5 | 210 | 206.5 | 232.5
| Std Dev. |  | 5.5377 | 32.8634 | 1.2111 | 2.0192 | 4.0825 | 2.8577 | 3.3466 | 3.4641 | 7.0616 | 6.3377 | 11.9569 | 4.0825 | 18.4977 | 5.8452
| Rel.Std.Dev. |  | 2.465% | 13.69% | 0.7194% | 0.847% | 1.87% | 1.386% | 1.468% | 1.487% | 3.148% | 3.016% | 5.698% | 1.929% | 9.303% | 2.532%
| PDM3 |  | 2.495% | 9.49% | -23.2% | 8.758% | -0.3943% | -5.945% | 4.016% | 6.297% | 2.343% | -4.12% | -4.272% | -3.436% | -9.29% | 5.308%
| Z-Score (Absolute) |  | 0.2 | 1.08 | 3.15 | 1.01 | 0.18 | 0.89 | 0.4 | 0.69 | 0.18 | 0.66 | 0.68 | 0.58 | 1.34 | 0.57
| Analytical results for Cs in OREAS 25b (Certified Value 4.57 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | PF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS |  | 2
3
| 1 | 1 | 4.7 | 4.31 | 4.7675 | 4.6 | 4.8 | 4.34 | 4.65 | 4.44 | 4.25 | 4.4 | 5.3 | 4.6 | 4.8 | 4.8 |  | 1
| 1 | 2 | 4.7 | 5.28 | 4.4455 | 4.6 | 4.7 | 4.34 | 4.51 | 4.62 | 4.54 | 4.1 | 5.1 | 4.6 | 4.5 | 4.4 |  | #N/A
| 1 | 3 | 4.6 | 5.16 | 4.401 | 4.7 | 4.4 | 4.49 | 4.64 | 4.57 | 4.62 | 4.7 | 4.7 | 4.4 | 4.6 | 4.8 |  | 16
| 1 | 4 | 4.4 | 5.23 | 4.6266 | 4.7 | 4.5 | 4.43 | 4.56 | 4.56 | 4.68 | 4.4 | 5.1 | 4.6 | 4.4 | 4.9 |  | 4.5651
| 1 | 5 | 4.5 | 4.28 | 4.775 | 4.7 | 4.3 | 4.41 | 4.64 | 4.52 | 4.46 | 5.1 | 4.8 | 4.6 | 4.2 | 4.7 |  | 35
| 1 | 6 | 4.8 | 5.1 | 4.7892 | 4.7 | 4.4 | 4.59 | 4.6 | 4.61 | 4.37 | 5.3 | 5.1 | 4.6 | 4.5 | 4.4
| Mean |  | 4.6167 | 4.8933 | 4.6341 | 4.6667 | 4.5167 | 4.4333 | 4.6 | 4.5533 | 4.4867 | 4.6667 | 5.0167 | 4.5667 | 4.5 | 4.6667
| Median |  | 4.65 | 5.13 | 4.6971 | 4.7 | 4.45 | 4.42 | 4.62 | 4.565 | 4.5 | 4.55 | 5.1 | 4.6 | 4.5 | 4.75
| Std Dev. |  | 0.1472 | 0.46758 | 0.17418 | 0.05164 | 0.19408 | 0.095638 | 0.055498 | 0.066232 | 0.16021 | 0.45898 | 0.22286 | 0.08165 | 0.2 | 0.21602
| Rel.Std.Dev. |  | 3.188% | 9.555% | 3.759% | 1.107% | 4.297% | 2.157% | 1.206% | 1.455% | 3.571% | 9.835% | 4.442% | 1.788% | 4.444% | 4.629%
| PDM3 |  | 1.13% | 7.191% | 1.513% | 2.226% | -1.06% | -2.886% | 0.7652% | -0.257% | -1.717% | 2.226% | 9.893% | 0.03506% | -1.425% | 2.226%
| Z-Score (Absolute) |  | 0.1 | 3.29 | 0.3 | 0.67 | 1.06 | 2.02 | 0.1 | 0.64 | 1.41 | 0.67 | 4.72 | 0.48 | 1.25 | 0.67
| Analytical results for Cu in OREAS 25b (Certified Value 41.2 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 11 | 12 | 15 | 16 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*MS | [8 empty] | 1
1
| 1 | 1 | 22 | 32.5822 | 40 | 40.6 | 60 | 95 | 45 | [8 empty] | 1
| 1 | 2 | 40 | 34.0061 | 40 | 39.9 | 70 | 95 | 45 | [8 empty] | 6
| 1 | 3 | 40 | 31.2216 | 40 | 38.1 | 50 | 75 | 45 | [8 empty] | 16
| 1 | 4 | 38 | 32.5001 | 40 | 36.2 | 80 | 99 | 50 | [8 empty] | 41.1978
| 1 | 5 | 23 | 32.0088 | 40 | 39.3 | 50 | 94 | 45 | [8 empty] | 36
| 1 | 6 | 41 | 33.6028 | 40 | 43.1 | 50 | 99 | 45
| Mean |  | 34 | 32.6536 | 40 | 39.5333 | 60 | 92.8333 | 45.8333
| Median |  | 39 | 32.5411 | 40 | 39.6 | 55 | 95 | 45
| Std Dev. |  | 8.9666 | 1.0223 | 0 | 2.333 | 12.6491 | 8.9981 | 2.0412
| Rel.Std.Dev. |  | 26.37% | 3.131% | 0% | 5.901% | 21.08% | 9.693% | 4.454%
| PDM3 |  | -17.47% | -20.74% | -2.907% | -4.04% | 45.64% | 125.3% | 11.25%
| Z-Score (Absolute) |  | 0.67 | 0.83 | 0 | 0.05 | 2.25 | 5.94 | 0.66
| Analytical results for Dy in OREAS 25b (Certified Value 3.61 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 4.1 | 3.42 | 3.8 | 3.57 | 3.58 | 3.7 | 3.71 | 3.49 | 4.55 | 3.41 | 3.18 | 3.5 |  |  |  | 1
| 1 | 2 | 3.8 | 3.92 | 3.8 | 3.54 | 3.59 | 3.89 | 3.45 | 3.68 | 4.47 | 3.41 | 3.41 | 3.6 |  |  |  | 7
| 1 | 3 | 3.9 | 3.92 | 3.9 | 3.66 | 3.78 | 3.81 | 3.51 | 3.65 | 3.98 | 3.22 | 3.46 | 3.7 |  |  |  | 16
| 1 | 4 | 3.7 | 3.77 | 3.8 | 3.24 | 3.72 | 3.33 | 3.69 | 3.61 | 4.01 | 3.46 | 3.53 | 3.7 |  |  |  | 3.6129
| 1 | 5 | 3.8 | 3.21 | 3.8 | 3.5 | 3.47 | 3.57 | 3.9 | 3.51 | 4.28 | 3.26 | 3.45 | 3.4 |  |  |  | 37
| 1 | 6 | 3.6 | 3.73 | 3.8 | 3.48 | 3.85 | 3.87 | 3.57 | 3.62 | 4.13 | 3.23 | 3.61 | 3.3
| Mean |  | 3.8167 | 3.6617 | 3.8167 | 3.4983 | 3.665 | 3.695 | 3.6383 | 3.5933 | 4.2367 | 3.3317 | 3.44 | 3.5333
| Median |  | 3.8 | 3.75 | 3.8 | 3.52 | 3.655 | 3.755 | 3.63 | 3.615 | 4.205 | 3.335 | 3.455 | 3.55
| Std Dev. |  | 0.17224 | 0.28715 | 0.040825 | 0.14148 | 0.14237 | 0.21483 | 0.16302 | 0.076594 | 0.23796 | 0.10647 | 0.1456 | 0.1633
| Rel.Std.Dev. |  | 4.513% | 7.842% | 1.07% | 4.044% | 3.885% | 5.814% | 4.481% | 2.132% | 5.617% | 3.196% | 4.233% | 4.622%
| PDM3 |  | 5.641% | 1.35% | 5.641% | -3.17% | 1.443% | 2.273% | 0.7046% | -0.541% | 17.27% | -7.784% | -4.785% | -2.202%
| Z-Score (Absolute) |  | 0.84 | 0.06 | 0.84 | 0.76 | 0.08 | 0.23 | 0.06 | 0.28 | 2.95 | 1.6 | 1.06 | 0.59
| Analytical results for Er in OREAS 25b (Certified Value 2.16 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 2.2 | 2.06 | 2.2 | 2.23 | 2.27 | 2.26 | 2.2 | 2.27 | 2.42 | 2.08 | 1.63 | 2 |  |  |  | 1
| 1 | 2 | 2.1 | 2.36 | 2.2 | 2.11 | 2.2 | 2.07 | 2.14 | 2.01 | 2.3 | 2.04 | 1.67 | 1.9 |  |  |  | 23
| 1 | 3 | 2.2 | 2.38 | 2.2 | 2.22 | 2.19 | 2.44 | 2.06 | 2.1 | 2.19 | 1.91 | 1.76 | 2.2 |  |  |  | 16
| 1 | 4 | 2.2 | 2.11 | 2.2 | 2.07 | 2.12 | 1.97 | 2.35 | 2.21 | 2.09 | 2.08 | 1.85 | 2.1 |  |  |  | 2.16
| 1 | 5 | 2 | 2 | 2.2 | 2.11 | 2.14 | 2.15 | 2.26 | 2.29 | 2.55 | 2.02 | 1.58 | 2 |  |  |  | 38
| 1 | 6 | 2.1 | 2.31 | 2.2 | 2.11 | 2.3 | 2.35 | 2 | 2.08 | 2.47 | 1.91 | 1.84 | 1.8
| Mean |  | 2.1333 | 2.2033 | 2.2 | 2.1417 | 2.2033 | 2.2067 | 2.1683 | 2.16 | 2.3367 | 2.0067 | 1.7217 | 2
| Median |  | 2.15 | 2.21 | 2.2 | 2.11 | 2.195 | 2.205 | 2.17 | 2.155 | 2.36 | 2.03 | 1.715 | 2
| Std Dev. |  | 0.08165 | 0.16597 | 4.8648e-16 | 0.066458 | 0.070616 | 0.17648 | 0.12906 | 0.11314 | 0.17546 | 0.078401 | 0.11232 | 0.14142
| Rel.Std.Dev. |  | 3.827% | 7.533% | 2.211e-14% | 3.103% | 3.205% | 7.998% | 5.952% | 5.238% | 7.509% | 3.907% | 6.524% | 7.071%
| PDM3 |  | -1.235% | 2.006% | 1.852% | -0.8488% | 2.006% | 2.16% | 0.3858% | 0% | 8.179% | -7.099% | -20.29% | -7.407%
| Z-Score (Absolute) |  | 0.53 | 0.67 | 0.62 | 0.39 | 0.67 | 0.73 | 0.07 | 0.07 | 2.97 | 2.71 | 7.62 | 2.83
| Analytical results for Eu in OREAS 25b (Certified Value 0.89 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 0.9 | 0.8 | 0.93 | 0.81 | 0.85 | 0.87 | 0.93 | 0.83 | 1.01 | 0.79 | 0.58 | 0.8 |  |  |  | 1
| 1 | 2 | 0.9 | 0.59 | 0.89 | 0.86 | 0.86 | 1.04 | 0.96 | 0.83 | 0.99 | 0.81 | 0.54 | 0.7 |  |  |  | 8
| 1 | 3 | 0.9 | 0.63 | 0.92 | 0.82 | 0.88 | 0.81 | 0.88 | 0.82 | 0.9 | 0.75 | 0.53 | 0.9 |  |  |  | 16
| 1 | 4 | 0.9 | 0.64 | 0.97 | 0.8 | 1.01 | 0.81 | 0.96 | 1 | 0.95 | 0.83 | 0.68 | 0.8 |  |  |  | 0.88737
| 1 | 5 | 0.8 | 0.78 | 0.96 | 0.88 | 0.9 | 0.83 | 0.82 | 0.97 | 1.19 | 0.81 | 0.5 | 0.9 |  |  |  | 39
| 1 | 6 | 0.8 | 0.64 | 0.94 | 0.85 | 0.95 | 0.96 | 0.87 | 0.8 | 0.99 | 0.85 | 0.71 | 0.7
| Mean |  | 0.86667 | 0.68 | 0.935 | 0.83667 | 0.90833 | 0.88667 | 0.90333 | 0.875 | 1.005 | 0.80667 | 0.59 | 0.8
| Median |  | 0.9 | 0.64 | 0.935 | 0.835 | 0.89 | 0.85 | 0.905 | 0.83 | 0.99 | 0.81 | 0.56 | 0.8
| Std Dev. |  | 0.05164 | 0.087407 | 0.02881 | 0.031411 | 0.061128 | 0.09395 | 0.056095 | 0.086429 | 0.098742 | 0.034448 | 0.08579 | 0.089443
| Rel.Std.Dev. |  | 5.958% | 12.85% | 3.081% | 3.754% | 6.73% | 10.6% | 6.21% | 9.878% | 9.825% | 4.27% | 14.54% | 11.18%
| PDM3 |  | -2.333% | -23.37% | 5.368% | -5.714% | 2.362% | -0.0793% | 1.799% | -1.394% | 13.26% | -9.095% | -33.51% | -9.846%
| Z-Score (Absolute) |  | 0.27 | NA | 1.02 | 0.83 | 0.52 | 0.11 | 0.42 | 0.11 | 2.34 | 1.4 | 5.47 | NA
Results from laboratories 2 and 16 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Fe in OREAS 25b (Certified Value 8.44 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 15 | 16 | [5 empty] | wt.%
|  |  | BF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 2
3
| 1 | 1 | 8.61 | 7.932 | 8.04 | 8.7297 | 8.498 | 8.365 | 8.883 | 8.4 | 8.33 | 8.55 | [5 empty] | 1
| 1 | 2 | 8.309 | 8.302 | 8.07 | 8.54 | 8.547 | 8.316 | 8.6032 | 8.41 | 8.4 | 8.71 | [5 empty] | #N/A
| 1 | 3 | 8.428 | 8.142 | 8.13 | 8.8944 | 8.659 | 8.351 | 8.9529 | 8.4 | 6.82 | 8.55 | [5 empty] | 16
| 1 | 4 | 8.386 | 9.044 | 8.11 | 8.5448 | 8.624 | 8.365 | 8.9529 | 8.58 | 8.25 | 8.81 | [5 empty] | 8.4388
| 1 | 5 | 8.463 | 7.785 | 8.03 | 8.4247 | 8.442 | 8.512 | 8.883 | 8.38 | 8.23 | 8.71 | [5 empty] | 40
| 1 | 6 | 8.205 | 8.044 | 8.1 | 8.9362 | 8.582 | 8.47 | 8.7431 | 8.37 | 8.32 | 8.67
| Mean |  | 8.4002 | 8.2082 | 8.08 | 8.6783 | 8.5587 | 8.3965 | 8.8363 | 8.4233 | 8.0583 | 8.6667
| Median |  | 8.407 | 8.093 | 8.085 | 8.6373 | 8.5645 | 8.365 | 8.883 | 8.4 | 8.285 | 8.69
| Std Dev. |  | 0.1381 | 0.44589 | 0.04 | 0.20836 | 0.080383 | 0.076521 | 0.13753 | 0.078145 | 0.6097 | 0.10152
| Rel.Std.Dev. |  | 1.644% | 5.432% | 0.495% | 2.401% | 0.9392% | 0.9113% | 1.556% | 0.9277% | 7.566% | 1.171%
| PDM3 |  | -0.4574% | -2.733% | -4.251% | 2.839% | 1.421% | -0.5008% | 4.711% | -0.1829% | -4.508% | 2.701%
| Z-Score (Absolute) |  | 0.03 | 0.6 | 0.98 | 0.78 | 0.43 | 0.04 | 1.25 | 0.03 | 1.04 | 0.75
| Analytical results for Ga in OREAS 25b (Certified Value 25.9 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS | 1
2
| 1 | 1 | 25 | 21.8 | 25.1 | 27.3564 | 25 | 22.1 | 25.2 | 27.3 | 25.1 | 25 | 28 | 31.3 | 26.3 | 22 | 25 | 1
| 1 | 2 | 26 | 24.4 | 26.8 | 26.9217 | 26 | 23.4 | 25.2 | 27.1 | 26.5 | 26.9 | 27.5 | 32.5 | 26.4 | 22 | 25 | #N/A
| 1 | 3 | 26 | 23.9 | 26.2 | 27.6992 | 26 | 22.2 | 26.2 | 27.9 | 25 | 27.4 | 24.3 | 31.6 | 24.6 | 21 | 25 | 16
| 1 | 4 | 25 | 23.1 | 26 | 27.4568 | 25 | 22.4 | 26 | 25.4 | 25.6 | 26.7 | 23.2 | 33.5 | 25.6 | 23 | 26 | 25.8618
| 1 | 5 | 26 | 21.5 | 25.8 | 26.9456 | 26 | 23 | 25 | 28 | 25.3 | 26 | 26.3 | 31.2 | 25.5 | 22 | 27 | 41
| 1 | 6 | 26 | 23.9 | 26.3 | 27.418 | 26 | 22.3 | 26.6 | 27.4 | 25.3 | 26.3 | 27.7 | 32.1 | 25.2 | 22 | 25
| Mean |  | 25.6667 | 23.1 | 26.0333 | 27.2996 | 25.6667 | 22.5667 | 25.7 | 27.1833 | 25.4667 | 26.3833 | 26.1667 | 32.0333 | 25.6 | 22 | 25.5
| Median |  | 26 | 23.5 | 26.1 | 27.3872 | 26 | 22.35 | 25.6 | 27.35 | 25.3 | 26.5 | 26.9 | 31.85 | 25.55 | 22 | 25
| Std Dev. |  | 0.5164 | 1.2017 | 0.56804 | 0.30652 | 0.5164 | 0.5164 | 0.65422 | 0.9411 | 0.5465 | 0.83287 | 1.9896 | 0.87101 | 0.67823 | 0.63246 | 0.83666
| Rel.Std.Dev. |  | 2.012% | 5.202% | 2.182% | 1.123% | 2.012% | 2.288% | 2.546% | 3.462% | 2.146% | 3.157% | 7.604% | 2.719% | 2.649% | 2.875% | 3.281%
| PDM3 |  | -0.7544% | -10.68% | 0.6634% | 5.56% | -0.7544% | -12.74% | -0.6255% | 5.11% | -1.528% | 2.017% | 1.179% | 23.86% | -1.012% | -14.93% | -1.399%
| Z-Score (Absolute) |  | 0 | 3.46 | 0.49 | 2.2 | 0 | 4.18 | 0.04 | 2.05 | 0.27 | 0.97 | 0.67 | 8.59 | 0.09 | 4.94 | 0.22
| Analytical results for Gd in OREAS 25b (Certified Value 3.66 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 3.8 | 3.46 | 3.7 | 3.65 | 3.52 | 3.53 | 3.62 | 3.7 | 4.41 | 3.51 | 3.4 | 3.8 |  |  |  | 1
| 1 | 2 | 3.7 | < 0.05 | 3.6 | 3.65 | 3.77 | 3.76 | 3.97 | 3.61 | 4.51 | 3.62 | 3.47 | 3.5 |  |  |  | 10
| 1 | 3 | 3.7 | < 0.05 | 3.8 | 3.72 | 3.69 | 3.74 | 3.65 | 3.94 | 4.41 | 3.53 | 3.66 | 3.7 |  |  |  | 16
| 1 | 4 | 3.8 | < 0.05 | 3.6 | 3.52 | 3.79 | 3.68 | 3.54 | 3.9 | 3.88 | 3.55 | 3.42 | 3.7 |  |  |  | 3.6567
| 1 | 5 | 3.4 | 3.41 | 3.8 | 3.74 | 3.64 | 3.63 | 3.91 | 3.77 | 4.4 | 3.49 | 3.64 | 3.6 |  |  |  | 42
| 1 | 6 | 3.6 | < 0.05 | 3.7 | 3.57 | 3.78 | 3.87 | 3.72 | 3.53 | 3.94 | 3.59 | 3.51 | 3.4
| Mean |  | 3.6667 | 3.435 | 3.7 | 3.6417 | 3.6983 | 3.7017 | 3.735 | 3.7417 | 4.2583 | 3.5483 | 3.5167 | 3.6167
| Median |  | 3.7 | 3.435 | 3.7 | 3.65 | 3.73 | 3.71 | 3.685 | 3.735 | 4.405 | 3.54 | 3.49 | 3.65
| Std Dev. |  | 0.15055 | 0.035355 | 0.089443 | 0.084715 | 0.10534 | 0.11686 | 0.17003 | 0.16068 | 0.27345 | 0.04916 | 0.11039 | 0.1472
| Rel.Std.Dev. |  | 4.106% | 1.029% | 2.417% | 2.326% | 2.848% | 3.157% | 4.552% | 4.294% | 6.422% | 1.385% | 3.139% | 4.07%
| PDM3 |  | 0.2735% | -6.062% | 1.185% | -0.4102% | 1.139% | 1.231% | 2.142% | 2.325% | 16.45% | -2.963% | -3.829% | -1.094%
| Z-Score (Absolute) |  | 0.19 | 30.44 | 0.21 | 0.49 | 0.19 | 0.23 | 0.63 | 0.71 | 6.95 | 1.62 | 2 | 0.8
| Analytical results for Ge in OREAS 25b (Certified Value 2.03 ppm) | [16 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 04 | 05 | 11 | 12 | 15 | 16 | [9 empty] | ppm
|  |  | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 1.4085 | 2 | 1.98 | 2.8 | 2 | 3 | [9 empty] | 1
| 1 | 2 | 1.3857 | 2 | 1.73 | 2.6 | 2 | 2 | [9 empty] | 19
| 1 | 3 | 1.2188 | 2 | 1.89 | 2.7 | 2 | 2 | [9 empty] | 16
| 1 | 4 | 1.4137 | 2 | 1.83 | 2.8 | 2 | 2 | [9 empty] | 2.0306
| 1 | 5 | 1.2185 | 2 | 2.4 | 2.6 | 2 | 2 | [9 empty] | 43
| 1 | 6 | 1.4057 | 2 | 2.32 | 2.4 | 2 | 2
| Mean |  | 1.3418 | 2 | 2.025 | 2.65 | 2 | 2.1667
| Median |  | 1.3957 | 2 | 1.935 | 2.65 | 2 | 2
| Std Dev. |  | 0.095873 | 0 | 0.27311 | 0.15166 | 0 | 0.40825
| Rel.Std.Dev. |  | 7.145% | 0% | 13.49% | 5.723% | 0% | 18.84%
| PDM3 |  | -33.92% | -1.506% | -0.2749% | 30.5% | -1.506% | 6.702%
| Z-Score (Absolute) |  | 5.43 | 0.1 | 0.1 | 5.16 | 0.1 | 1.25
| Analytical results for Hf in OREAS 25b (Certified Value 8.47 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | [5 empty] | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | [5 empty] | 2
3
| 1 | 1 | 8.6 | 7.17 | 8.3 | 8.4 | 8.29 | 8.45 | 8.35 | 7.69 | 9.02 | 8.2 | [5 empty] | 1
| 1 | 2 | 8.2 | 8.8 | 8.5 | 8.6 | 8.06 | 8.37 | 8.77 | 8.25 | 8.89 | 8.3 | [5 empty] | 11
| 1 | 3 | 8.2 | 9.13 | 8.8 | 8.3 | 8.5 | 9.04 | 8.84 | 8.29 | 8.07 | 8.1 | [5 empty] | 16
| 1 | 4 | 7.9 | 8.76 | 8.6 | 7.7 | 8.43 | 8.47 | 8.63 | 8.96 | 8.44 | 8.3 | [5 empty] | 8.4652
| 1 | 5 | 8.3 | 7.06 | 8.7 | 8.1 | 8.44 | 8.99 | 8.77 | 8.07 | 9.18 | 8.1 | [5 empty] | 44
| 1 | 6 | 8.4 | 8.81 | 8.9 | 8.2 | 8.43 | 9.06 | 7.97 | 7.97 | 9.06 | 8.2
| Mean |  | 8.2667 | 8.2883 | 8.6333 | 8.2167 | 8.3583 | 8.73 | 8.555 | 8.205 | 8.7767 | 8.2
| Median |  | 8.25 | 8.78 | 8.65 | 8.25 | 8.43 | 8.73 | 8.7 | 8.16 | 8.955 | 8.2
| Std Dev. |  | 0.23381 | 0.91916 | 0.21602 | 0.30605 | 0.16167 | 0.33112 | 0.33572 | 0.42847 | 0.43075 | 0.089443
| Rel.Std.Dev. |  | 2.828% | 11.09% | 2.502% | 3.725% | 1.934% | 3.793% | 3.924% | 5.222% | 4.908% | 1.091%
| PDM3 |  | -2.346% | -2.09% | 1.986% | -2.936% | -1.263% | 3.128% | 1.06% | -3.074% | 3.679% | -3.133%
| Z-Score (Absolute) |  | 0.32 | 0.2 | 1.73 | 0.6 | 0.2 | 2.27 | 1.29 | 0.66 | 2.53 | 0.69
| Analytical results for Ho in OREAS 25b (Certified Value 0.72 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 0.7 | 0.7 | 0.7 | 0.71 | 0.7 | 0.71 | 0.74 | 0.72 | 0.91 | 0.68 | 0.44 | 0.6 |  |  |  | 1
| 1 | 2 | 0.7 | 0.72 | 0.7 | 0.71 | 0.76 | 0.74 | 0.78 | 0.69 | 0.91 | 0.69 | 0.42 | 0.7 |  |  |  | 22
| 1 | 3 | 0.8 | 0.72 | 0.7 | 0.72 | 0.71 | 0.74 | 0.68 | 0.71 | 0.81 | 0.67 | 0.47 | 0.7 |  |  |  | 16
| 1 | 4 | 0.7 | 0.71 | 0.8 | 0.67 | 0.77 | 0.71 | 0.75 | 0.67 | 0.82 | 0.69 | 0.51 | 0.7 |  |  |  | 0.71738
| 1 | 5 | 0.8 | 0.66 | 0.8 | 0.72 | 0.7 | 0.74 | 0.77 | 0.72 | 0.95 | 0.68 | 0.49 | 0.7 |  |  |  | 45
| 1 | 6 | 0.8 | 0.77 | 0.8 | 0.73 | 0.8 | 0.81 | 0.75 | 0.65 | 0.87 | 0.66 | 0.52 | 0.7
| Mean |  | 0.75 | 0.71333 | 0.75 | 0.71 | 0.74 | 0.74167 | 0.745 | 0.69333 | 0.87833 | 0.67833 | 0.475 | 0.68333
| Median |  | 0.75 | 0.715 | 0.75 | 0.715 | 0.735 | 0.74 | 0.75 | 0.7 | 0.89 | 0.68 | 0.48 | 0.7
| Std Dev. |  | 0.054772 | 0.03559 | 0.054772 | 0.020976 | 0.042426 | 0.03656 | 0.035071 | 0.028752 | 0.055287 | 0.01169 | 0.03937 | 0.040825
| Rel.Std.Dev. |  | 7.303% | 4.989% | 7.303% | 2.954% | 5.733% | 4.929% | 4.708% | 4.147% | 6.295% | 1.723% | 8.288% | 5.974%
| PDM3 |  | 4.547% | -0.5642% | 4.547% | -1.029% | 3.153% | 3.385% | 3.85% | -3.352% | 22.44% | -5.443% | -33.79% | -4.746%
| Z-Score (Absolute) |  | NA | 0 | NA | 0.08 | 0.63 | 0.67 | 0.75 | 0.48 | 3.93 | 0.83 | 5.67 | NA
Results from laboratories 1, 5 & 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 25b (Indicative Value 0.12 ppm) | [16 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 04 | 05 | 11 | 12 | 16 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.1 | 0.13106 | < 0.2 | < 0.2 | < 0.05 | < 0.25 | [9 empty] | 1
| 1 | 2 | 0.1 | 0.11304 | < 0.2 | < 0.2 | < 0.05 | < 0.25 | [9 empty] | 13
| 1 | 3 | 0.1 | 0.11561 | < 0.2 | < 0.2 | < 0.05 | < 0.25 | [9 empty] | 16
| 1 | 4 | 0.1 | 0.12273 | < 0.2 | < 0.2 | < 0.05 | < 0.25 | [9 empty] | 0.12131
| 1 | 5 | 0.1 | 0.11435 | < 0.2 | 0.21 | < 0.05 | < 0.25 | [9 empty] | 19
| 1 | 6 | 0.1 | 0.12695 | < 0.2 | 0.22 | < 0.05 | < 0.25
| Mean |  | 0.1 | 0.12062 |  | 0.215
| Median |  | 0.1 | 0.11917 |  | 0.215
| Std Dev. |  | 1.5202e-17 | 0.0074214 |  | 0.0070711
| Rel.Std.Dev. |  | 1.52e-14% | 6.153% |  | 3.289%
| PDM3 |  | -17.57% | -0.5677% |  | 77.23%
| Z-Score (Absolute) |  | 0.56 | 0.56 | 0.56 | 1.51 | 4.6 | 0.79
| Analytical results for K in OREAS 25b (Certified Value 0.805 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 15 | 16 | [5 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 3
3
| 1 | 1 | 0.75 | 0.772 | 0.76 | 0.9808 | 0.747 | 0.772 | 0.7837 | 0.87 | 1 | 0.83 | [5 empty] | 1
| 1 | 2 | 0.81 | 0.772 | 0.75 | 0.83651 | 0.747 | 0.764 | 0.7886 | 0.9 | 1 | 0.83 | [5 empty] | #N/A
| 1 | 3 | 0.77 | 0.739 | 0.75 | 0.99917 | 0.755 | 0.764 | 0.7936 | 0.92 | 0.6 | 0.83 | [5 empty] | 16
| 1 | 4 | 0.77 | 0.805 | 0.74 | 0.8947 | 0.755 | 0.764 | 0.7969 | 0.91 | 0.9 | 0.83 | [5 empty] | 0.80533
| 1 | 5 | 0.84 | 0.755 | 0.74 | 0.83371 | 0.755 | 0.764 | 0.8285 | 0.89 | 0.9 | 0.83 | [5 empty] | 46
| 1 | 6 | 0.8 | 0.747 | 0.76 | 0.92673 | 0.755 | 0.772 | 0.8227 | 0.89 | 0.9 | 0.83
| Mean |  | 0.79 | 0.765 | 0.75 | 0.91194 | 0.75233 | 0.76667 | 0.80233 | 0.89667 | 0.88333 | 0.83
| Median |  | 0.785 | 0.7635 | 0.75 | 0.91071 | 0.755 | 0.764 | 0.79525 | 0.895 | 0.9 | 0.83
| Std Dev. |  | 0.032863 | 0.023656 | 0.0089443 | 0.070254 | 0.0041312 | 0.0041312 | 0.018661 | 0.017512 | 0.1472 | 0
| Rel.Std.Dev. |  | 4.16% | 3.092% | 1.193% | 7.704% | 0.5491% | 0.5388% | 2.326% | 1.953% | 16.66% | 0%
| PDM3 |  | -1.903% | -5.008% | -6.87% | 13.24% | -6.58% | -4.801% | -0.3718% | 11.34% | 9.686% | 3.064%
| Z-Score (Absolute) |  | 0 | 0.45 | 0.72 | 2.2 | 0.67 | 0.42 | 0.22 | 1.93 | NA | 0.73
Results from laboratory 15 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for La in OREAS 25b (Certified Value 27.9 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES |  |  | 1
2
| 1 | 1 | 27.5 | 23.8 | 46.3 | 28.2 | 28.5 | 26.8 | 28.6 | 27.9 | 27.2 | 33.7 | 26.3 | 28.1 | 30 |  |  | 1
| 1 | 2 | 27.7 | 29.7 | 47 | 28.4 | 28 | 26.1 | 28 | 27.9 | 27.3 | 32.3 | 26.7 | 28.5 | 25 |  |  | 1
| 1 | 3 | 27.7 | 28.4 | 46.8 | 29.3 | 27.7 | 27.8 | 28.8 | 27.9 | 29.8 | 30.3 | 26.4 | 28.8 | 30 |  |  | 16
| 1 | 4 | 27 | 27.6 | 46 | 27.7 | 27 | 27.3 | 27.3 | 27.5 | 28.8 | 29.4 | 27.6 | 28.6 | 30 |  |  | 27.9098
| 1 | 5 | 27.6 | 23.7 | 45.9 | 28.8 | 27.3 | 27 | 28.4 | 27.1 | 28.2 | 32.5 | 26.6 | 27.2 | 25 |  |  | 47
| 1 | 6 | 27.7 | 29 | 47.7 | 27.6 | 28.3 | 27.6 | 27.9 | 27.7 | 27.6 | 31.6 | 26.9 | 28.5 | 30
| Mean |  | 27.5333 | 27.0333 | 46.6167 | 28.3333 | 27.8 | 27.1 | 28.1667 | 27.6667 | 28.15 | 31.6333 | 26.75 | 28.2833 | 28.3333
| Median |  | 27.65 | 28 | 46.55 | 28.3 | 27.85 | 27.15 | 28.2 | 27.8 | 27.9 | 31.95 | 26.65 | 28.5 | 30
| Std Dev. |  | 0.27325 | 2.6357 | 0.68532 | 0.65013 | 0.57966 | 0.61319 | 0.5465 | 0.32042 | 1.0075 | 1.5642 | 0.46797 | 0.57764 | 2.582
| Rel.Std.Dev. |  | 0.9924% | 9.75% | 1.47% | 2.295% | 2.085% | 2.263% | 1.94% | 1.158% | 3.579% | 4.945% | 1.749% | 2.042% | 9.113%
| PDM3 |  | -1.349% | -3.141% | 67.03% | 1.517% | -0.3936% | -2.902% | 0.9202% | -0.8713% | 0.8605% | 13.34% | -4.156% | 1.338% | 1.517%
| Z-Score (Absolute) |  | 0.86 | 1.56 | 25.76 | 0.26 | 0.49 | 1.46 | 0.02 | 0.67 | 0 | 4.86 | 1.95 | 0.19 | 0.26
| Analytical results for Li in OREAS 25b (Certified Value 30.2 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 04 | 12 | 15 | 16 | [10 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | [10 empty] | 1
1
| 1 | 1 | 27 | 30.0139 | 36 | 35 | 30 | [10 empty] | 1
| 1 | 2 | 26 | 29.6856 | 38 | 35 | 25 | [10 empty] | #N/A
| 1 | 3 | 27 | 31.2587 | 37 | 20 | 30 | [10 empty] | 16
| 1 | 4 | 25 | 29.4472 | 37 | 25 | 25 | [10 empty] | 30.202
| 1 | 5 | 27 | 30.6181 | 36 | 30 | 30 | [10 empty] | 48
| 1 | 6 | 26 | 31.0372 | 37 | 30 | 30
| Mean |  | 26.3333 | 30.3434 | 36.8333 | 29.1667 | 28.3333
| Median |  | 26.5 | 30.316 | 37 | 30 | 30
| Std Dev. |  | 0.8165 | 0.74015 | 0.75277 | 5.8452 | 2.582
| Rel.Std.Dev. |  | 3.101% | 2.439% | 2.044% | 20.04% | 9.113%
| PDM3 |  | -12.81% | 0.4683% | 21.96% | -3.428% | -6.187%
| Z-Score (Absolute) |  | 1.62 | 0.67 | 4.39 | 0 | 0.48
| Analytical results for Lu in OREAS 25b (Certified Value 0.32 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 |  |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS |  |  |  |  | 2
3
| 1 | 1 | 0.4 | 0.26 | 0.33 | 0.34 | 0.31 | 0.3 | 0.28 | 0.34 | 0.35 | 0.31 | 0.3 |  |  |  |  | 1
| 1 | 2 | 0.3 | 0.32 | 0.35 | 0.31 | 0.33 | 0.35 | 0.34 | 0.3 | 0.33 | 0.31 | 0.3 |  |  |  |  | 23
| 1 | 3 | 0.3 | 0.26 | 0.35 | 0.31 | 0.33 | 0.35 | 0.28 | 0.28 | 0.32 | 0.28 | 0.3 |  |  |  |  | 16
| 1 | 4 | 0.3 | 0.31 | 0.34 | 0.31 | 0.34 | 0.3 | 0.35 | 0.29 | 0.29 | 0.29 | 0.3 |  |  |  |  | 0.3156
| 1 | 5 | 0.3 | 0.27 | 0.34 | 0.31 | 0.33 | 0.32 | 0.34 | 0.36 | 0.47 | 0.3 | 0.3 |  |  |  |  | 49
| 1 | 6 | 0.3 | 0.29 | 0.34 | 0.29 | 0.36 | 0.35 | 0.31 | 0.28 | 0.41 | 0.29 | 0.3
| Mean |  | 0.31667 | 0.285 | 0.34167 | 0.31167 | 0.33333 | 0.32833 | 0.31667 | 0.30833 | 0.36167 | 0.29667 | 0.3
| Median |  | 0.3 | 0.28 | 0.34 | 0.31 | 0.33 | 0.335 | 0.325 | 0.295 | 0.34 | 0.295 | 0.3
| Std Dev. |  | 0.040825 | 0.025884 | 0.0075277 | 0.016021 | 0.01633 | 0.024833 | 0.031411 | 0.033714 | 0.066458 | 0.012111 | 0
| Rel.Std.Dev. |  | 12.89% | 9.082% | 2.203% | 5.14% | 4.899% | 7.563% | 9.919% | 10.93% | 18.38% | 4.082% | 0%
| PDM3 |  | 0.338% | -9.696% | 8.259% | -1.246% | 5.619% | 4.035% | 0.338% | -2.302% | 14.6% | -5.999% | -4.943%
| Z-Score (Absolute) |  | NA | 1.24 | 1.17 | 0.11 | 0.82 | 0.6 | 0.11 | 0.25 | 2.02 | 0.75 | 0.6
Results from laboratory 1 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 25b (Certified Value 0.377 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 15 | 16 | [5 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 3
3
| 1 | 1 | 0.37 | 0.374 | 0.35 | 0.37832 | 0.374 | 0.38 | 0.3811 | 0.37 | 0.34 | 0.37 | [5 empty] | 1
| 1 | 2 | 0.37 | 0.38 | 0.36 | 0.37664 | 0.368 | 0.38 | 0.3673 | 0.38 | 0.35 | 0.38 | [5 empty] | #N/A
| 1 | 3 | 0.38 | 0.374 | 0.35 | 0.38091 | 0.374 | 0.386 | 0.3866 | 0.38 | 0.28 | 0.38 | [5 empty] | 16
| 1 | 4 | 0.37 | 0.416 | 0.34 | 0.37693 | 0.362 | 0.38 | 0.3854 | 0.38 | 0.35 | 0.38 | [5 empty] | 0.37737
| 1 | 5 | 0.38 | 0.374 | 0.35 | 0.3713 | 0.368 | 0.386 | 0.398 | 0.38 | 0.34 | 0.38 | [5 empty] | 50
| 1 | 6 | 0.37 | 0.374 | 0.36 | 0.38577 | 0.368 | 0.386 | 0.3938 | 0.38 | 0.34 | 0.37
| Mean |  | 0.37333 | 0.382 | 0.35167 | 0.37831 | 0.369 | 0.383 | 0.38537 | 0.37833 | 0.33333 | 0.37667
| Median |  | 0.37 | 0.374 | 0.35 | 0.37762 | 0.368 | 0.383 | 0.386 | 0.38 | 0.34 | 0.38
| Std Dev. |  | 0.005164 | 0.016829 | 0.0075277 | 0.0048236 | 0.0045166 | 0.0032863 | 0.010741 | 0.0040825 | 0.026583 | 0.005164
| Rel.Std.Dev. |  | 1.383% | 4.405% | 2.141% | 1.275% | 1.224% | 0.8581% | 2.787% | 1.079% | 7.975% | 1.371%
| PDM3 |  | -1.069% | 1.228% | -6.81% | 0.2504% | -2.217% | 1.493% | 2.12% | 0.2564% | -11.67% | -0.1853%
| Z-Score (Absolute) |  | 0.57 | 0.61 | 3.51 | 0.11 | 1.17 | 0.74 | 1.07 | 0.11 | 6.01 | 0.11
| Analytical results for Mn in OREAS 25b (Certified Value 0.024 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 15 | 16 | [5 empty] | wt.%
|  |  | BF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 3
3
| 1 | 1 | 0.023 | 0.023 | 0.0316 | 0.026944 | 0.0263 | 0.023 | 0.0232 | 0.023 | 0.0234 | 0.026 | [5 empty] | 1
| 1 | 2 | 0.023 | 0.023 | 0.0323 | 0.026285 | 0.0263 | 0.023 | 0.0225 | 0.024 | 0.0234 | 0.0255 | [5 empty] | #N/A
| 1 | 3 | 0.023 | 0.023 | 0.0321 | 0.023528 | 0.0263 | 0.023 | 0.024 | 0.024 | 0.0189 | 0.025 | [5 empty] | 16
| 1 | 4 | 0.023 | 0.031 | 0.0315 | 0.026365 | 0.0263 | 0.023 | 0.0232 | 0.024 | 0.0239 | 0.027 | [5 empty] | 0.02419
| 1 | 5 | 0.023 | 0.023 | 0.0316 | 0.026218 | 0.0263 | 0.023 | 0.024 | 0.023 | 0.0228 | 0.026 | [5 empty] | 51
| 1 | 6 | 0.023 | 0.023 | 0.0323 | 0.02378 | 0.0263 | 0.023 | 0.0232 | 0.023 | 0.0234 | 0.026
| Mean |  | 0.023 | 0.024333 | 0.0319 | 0.02552 | 0.0263 | 0.023 | 0.02335 | 0.0235 | 0.022633 | 0.025917
| Median |  | 0.023 | 0.023 | 0.03185 | 0.026251 | 0.0263 | 0.023 | 0.0232 | 0.0235 | 0.0234 | 0.026
| Std Dev. |  | 3.8006e-18 | 0.003266 | 0.00037417 | 0.0014702 | 0 | 3.8006e-18 | 0.00057184 | 0.00054772 | 0.0018619 | 0.00066458
| Rel.Std.Dev. |  | 1.652e-14% | 13.42% | 1.173% | 5.761% | 0% | 1.652e-14% | 2.449% | 2.331% | 8.226% | 2.564%
| PDM3 |  | -4.92% | 0.5916% | 31.87% | 5.497% | 8.722% | -4.92% | -3.473% | -2.853% | -6.436% | 7.137%
| Z-Score (Absolute) |  | 0.49 | 0.32 | 4.93 | 0.94 | 1.45 | 0.49 | 0.41 | 0.32 | 0.86 | 1.19
| Analytical results for Mo in OREAS 25b (Indicative Value 2.85 ppm) | [16 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 11 | 16 | [8 empty] | ppm
|  |  | BF*MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 3 | < 2 | 15.2 | 2.0056 | 2 | 4 | 2.7 | [8 empty] | 1
| 1 | 2 | 3 | 4 | 15.5 | 1.8541 | 2 | 3.4 | 2.6 | [8 empty] | 2
| 1 | 3 | 3 | 4 | 15.5 | 1.6378 | 2 | 3.88 | 2.6 | [8 empty] | 16
| 1 | 4 | 4 | 3 | 15.2 | 1.9455 | 2 | 3.38 | 2.5 | [8 empty] | 2.8522
| 1 | 5 | 3 | < 2 | 15.2 | 1.5726 | 2 | 4.29 | 2.6 | [8 empty] | 20
| 1 | 6 | 3 | 4 | 15.5 | 2.1821 | 2 | 3.53 | 2.5
| Mean |  | 3.1667 | 3.75 | 15.35 | 1.8663 | 2 | 3.7467 | 2.5833
| Median |  | 3 | 4 | 15.35 | 1.8998 | 2 | 3.705 | 2.6
| Std Dev. |  | 0.40825 | 0.5 | 0.16432 | 0.22979 | 0 | 0.36844 | 0.075277
| Rel.Std.Dev. |  | 12.89% | 13.33% | 1.07% | 12.31% | 0% | 9.834% | 2.914%
| PDM3 |  | 11.03% | 31.48% | 438.2% | -34.57% | -29.88% | 31.36% | -9.425%
| Z-Score (Absolute) |  | 0.27 | 0 | 10.13 | 0.78 | 0.67 | 0.74 | 0.2
| Analytical results for Na in OREAS 25b (Certified Value 0.066 wt.%) | [16 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 11 | [9 empty] | wt.%
|  |  | BF*OES | BF*OES | PF*OES/MS | BF*OES/MS | BF*OES/MS | BF*OES | [9 empty] | 3
3
| 1 | 1 | 0.067 | 0.141 | 0.1 | 0.067 | 0.067 | 0.0846 | [9 empty] | 1
| 1 | 2 | 0.067 | 0.045 | 0.08 | 0.067 | 0.067 | 0.0712 | [9 empty] | #N/A
| 1 | 3 | 0.067 | 0.037 | 0.08 | 0.067 | 0.067 | 0.0734 | [9 empty] | 16
| 1 | 4 | 0.067 | 0.037 | 0.09 | 0.067 | 0.067 | 0.0734 | [9 empty] | 0.065559
| 1 | 5 | 0.067 | 0.126 | 0.1 | 0.067 | 0.067 | 0.0742 | [9 empty] | 52
| 1 | 6 | 0.059 | 0.037 | 0.08 | 0.067 | 0.059 | 0.0757
| Mean |  | 0.065667 | 0.0705 | 0.088333 | 0.067 | 0.065667 | 0.075417
| Median |  | 0.067 | 0.041 | 0.085 | 0.067 | 0.067 | 0.0738
| Std Dev. |  | 0.003266 | 0.049127 | 0.0098319 | 0 | 0.003266 | 0.0047288
| Rel.Std.Dev. |  | 4.974% | 69.68% | 11.13% | 0% | 4.974% | 6.27%
| PDM3 |  | 0.1645% | 7.537% | 34.74% | 2.198% | 0.1645% | 15.04%
| Z-Score (Absolute) |  | 0.67 | 0.4 | 4.3 | 0.4 | 0.67 | 1.48
| Analytical results for Nb in OREAS 25b (Certified Value 27.6 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 |  |  | ppm
|  |  | BF*MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS |  |  | 1
2
| 1 | 1 | 27.3 | 23.63 | 34.7 | 30.3226 | 25 | 25.1 | 27.8 | 27.9 | 28.9 | 25.2 | 28.7 | 27.2 | 30 |  |  | 1
| 1 | 2 | 27.4 | 28.99 | 33.2 | 29.0463 | 26 | 25.3 | 27.6 | 28.3 | 30.1 | 26.2 | 28.2 | 27.4 | 30 |  |  | 3
| 1 | 3 | 27.6 | 29.16 | 33.8 | 28.2822 | 27 | 24.8 | 28 | 28.2 | 29.6 | 26.1 | 25.8 | 26.2 | 30 |  |  | 16
| 1 | 4 | 27.4 | 27.52 | 32.7 | 29.6627 | 27 | 24.5 | 27.7 | 26.9 | 30 | 26.6 | 26.2 | 26.4 | 30 |  |  | 27.6233
| 1 | 5 | 27.4 | 23.4 | 32.7 | 27.7502 | 26 | 25.6 | 28.4 | 28 | 29.5 | 26.9 | 29.4 | 27.2 | 30 |  |  | 53
| 1 | 6 | 27.7 | 29.32 | 34.7 | 29.7115 | 27 | 24.6 | 28.8 | 27.9 | 26.9 | 26 | 29 | 26.3 | 30
| Mean |  | 27.4667 | 27.0033 | 33.6333 | 29.1293 | 26.3333 | 24.9833 | 28.05 | 27.8667 | 29.1667 | 26.1667 | 27.8833 | 26.7833 | 30
| Median |  | 27.4 | 28.255 | 33.5 | 29.3545 | 26.5 | 24.95 | 27.9 | 27.95 | 29.55 | 26.15 | 28.45 | 26.8 | 30
| Std Dev. |  | 0.15055 | 2.7783 | 0.92014 | 0.96682 | 0.8165 | 0.42622 | 0.46368 | 0.50067 | 1.1894 | 0.58195 | 1.5158 | 0.53821 | 0
| Rel.Std.Dev. |  | 0.5481% | 10.29% | 2.736% | 3.319% | 3.101% | 1.706% | 1.653% | 1.797% | 4.078% | 2.224% | 5.436% | 2.009% | 0%
| PDM3 |  | -0.5669% | -2.244% | 21.76% | 5.452% | -4.67% | -9.557% | 1.545% | 0.8811% | 5.587% | -5.273% | 0.9415% | -3.041% | 8.604%
| Z-Score (Absolute) |  | 0.21 | 0.46 | 3.08 | 0.67 | 0.82 | 1.54 | 0.1 | 0 | 0.69 | 0.91 | 0.01 | 0.58 | 1.14
| Analytical results for Nd in OREAS 25b (Certified Value 22.5 ppm) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 1
2
| 1 | 1 | 21.9 | 19 | 22.3 | 21.6 | 22.9 | 23.7 | 22.9 | 21.4 | 25.3 | 20.2 | 18.4 | 24 |  |  |  | 1
| 1 | 2 | 23.1 | 22.2 | 22.9 | 21 | 22.7 | 24.1 | 22 | 22.1 | 26.5 | 20.2 | 19.1 | 21.6 |  |  |  | 4
| 1 | 3 | 23 | 21 | 23 | 21.1 | 23.7 | 24.3 | 21.4 | 23.8 | 23.5 | 19.8 | 19 | 23.2 |  |  |  | 16
| 1 | 4 | 23.3 | 20.9 | 22.5 | 20.5 | 22.8 | 23.1 | 22.3 | 24.2 | 23.4 | 21 | 19.8 | 23.7 |  |  |  | 22.4515
| 1 | 5 | 22.1 | 18.9 | 23.5 | 21.1 | 22.4 | 24.1 | 21.7 | 23.2 | 24.8 | 19.7 | 18.8 | 22.4 |  |  |  | 54
| 1 | 6 | 22.8 | 21.2 | 22.6 | 21.6 | 24 | 23.8 | 22.6 | 23.4 | 24.8 | 20.3 | 20.4 | 21.7
| Mean |  | 22.7 | 20.5333 | 22.8 | 21.15 | 23.0833 | 23.85 | 22.15 | 23.0167 | 24.7167 | 20.2 | 19.25 | 22.7667
| Median |  | 22.9 | 20.95 | 22.75 | 21.1 | 22.85 | 23.95 | 22.15 | 23.3 | 24.8 | 20.2 | 19.05 | 22.8
| Std Dev. |  | 0.56921 | 1.311 | 0.42895 | 0.41352 | 0.62423 | 0.42778 | 0.56125 | 1.0629 | 1.1618 | 0.46043 | 0.72595 | 1.0211
| Rel.Std.Dev. |  | 2.508% | 6.385% | 1.881% | 1.955% | 2.704% | 1.794% | 2.534% | 4.618% | 4.7% | 2.279% | 3.771% | 4.485%
| PDM3 |  | 1.107% | -8.544% | 1.552% | -5.797% | 2.814% | 6.229% | -1.343% | 2.517% | 10.09% | -10.03% | -14.26% | 1.404%
| Z-Score (Absolute) |  | 0.03 | 1.75 | 0.05 | 1.26 | 0.28 | 0.89 | 0.46 | 0.22 | 1.57 | 2.01 | 2.76 | 0.03
| Analytical results for Ni in OREAS 25b (Certified Value 69 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 11 | 12 | 16 | [7 empty] | ppm
|  |  | BF*MS | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 61 | 63.3 | 63.2009 | 80 | 66 | 71 | 70 | 70 | [7 empty] | 1
| 1 | 2 | 75 | 63.3 | 57.7195 | 70 | 63 | 63 | 70 | 75 | [7 empty] | 5
| 1 | 3 | 74 | 62 | 59.2231 | 80 | 66 | 86 | 80 | 70 | [7 empty] | 16
| 1 | 4 | 73 | 61.7 | 61.392 | 70 | 66 | 71 | 80 | 90 | [7 empty] | 68.5186
| 1 | 5 | 59 | 62.5 | 57.7849 | 70 | 63 | 71 | 70 | 65 | [7 empty] | 55
| 1 | 6 | 75 | 64.1 | 64.7344 | 70 | 65 | 86 | 70 | 65
| Mean |  | 69.5 | 62.8167 | 60.6758 | 73.3333 | 64.8333 | 74.6667 | 73.3333 | 72.5
| Median |  | 73.5 | 62.9 | 60.3076 | 70 | 65.5 | 71 | 70 | 70
| Std Dev. |  | 7.4229 | 0.90866 | 2.9178 | 5.164 | 1.472 | 9.3095 | 5.164 | 9.3541
| Rel.Std.Dev. |  | 10.68% | 1.447% | 4.809% | 7.042% | 2.27% | 12.47% | 7.042% | 12.9%
| PDM3 |  | 1.432% | -8.322% | -11.45% | 7.027% | -5.378% | 8.973% | 7.027% | 5.811%
| Z-Score (Absolute) |  | 0.34 | 1.84 | 2.32 | 0.53 | 1.39 | 0.82 | 0.53 | 0.34
| Analytical results for P in OREAS 25b (Certified Value 0.04 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 11 | 15 | 16 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | [7 empty] | 3
3
| 1 | 1 | 0.04 | 0.044 | 0.0443 | 0.039 | 0.039 | 0.0371 | 0.03 | 0.042 | [7 empty] | 1
| 1 | 2 | 0.04 | 0.057 | 0.0436 | 0.044 | 0.035 | 0.0371 | 0.03 | 0.044 | [7 empty] | #N/A
| 1 | 3 | 0.04 | 0.057 | 0.0433 | 0.039 | 0.039 | 0.0384 | 0.03 | 0.042 | [7 empty] | 16
| 1 | 4 | 0.04 | 0.065 | 0.0447 | 0.039 | 0.039 | 0.0384 | 0.03 | 0.041 | [7 empty] | 0.040369
| 1 | 5 | 0.04 | 0.031 | 0.0442 | 0.035 | 0.039 | 0.0367 | 0.03 | 0.043 | [7 empty] | 56
| 1 | 6 | 0.04 | 0.061 | 0.0443 | 0.044 | 0.039 | 0.038 | 0.03 | 0.04
| Mean |  | 0.04 | 0.0525 | 0.044067 | 0.04 | 0.038333 | 0.037617 | 0.03 | 0.042
| Median |  | 0.04 | 0.057 | 0.04425 | 0.039 | 0.039 | 0.03755 | 0.03 | 0.042
| Std Dev. |  | 0 | 0.012677 | 0.0005164 | 0.0034641 | 0.001633 | 0.0007414 | 0 | 0.0014142
| Rel.Std.Dev. |  | 0% | 24.15% | 1.172% | 8.66% | 4.26% | 1.971% | 0% | 3.367%
| PDM3 |  | -0.9147% | 30.05% | 9.159% | -0.9147% | -5.043% | -6.819% | -25.69% | 4.04%
| Z-Score (Absolute) |  | 0 | 3.79 | 1.25 | 0 | 0.45 | 0.74 | 3.07 | 0.61
| Analytical results for Pb in OREAS 25b (Indicative Value 23.4 ppm) | [16 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 04 | 05 | 12 | 16 | [10 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | [10 empty] | 1
1
| 1 | 1 | 24 | 20.316 | 21 | 35 | 15 | [10 empty] | 1
| 1 | 2 | 23 | 21.7855 | 20 | 37 | 15 | [10 empty] | 6
| 1 | 3 | 29 | < 20 | 21 | 37 | 20 | [10 empty] | 16
| 1 | 4 | 31 | < 20 | 22 | 36 | 15 | [10 empty] | 23.3935
| 1 | 5 | 31 | < 20 | 21 | 36 | 20 | [10 empty] | 21
| 1 | 6 | 26 | 21.7021 | 22 | 36 | 15
| Mean |  | 27.3333 | 21.2679 | 21.1667 | 36.1667 | 16.6667
| Median |  | 27.5 | 21.7021 | 21 | 36 | 15
| Std Dev. |  | 3.5024 | 0.82542 | 0.75277 | 0.75277 | 2.582
| Rel.Std.Dev. |  | 12.81% | 3.881% | 3.556% | 2.081% | 15.49%
| PDM3 |  | 16.84% | -9.086% | -9.519% | 54.6% | -28.75%
| Z-Score (Absolute) |  | 0.75 | 0.67 | 0 | 1.83 | 0.55
| Analytical results for Pr in OREAS 25b (Certified Value 6.04 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 6.2 | 5.22 | 6.06 | 6.01 | 5.93 | 6.44 | 5.89 | 5.81 | 6.77 | 5.46 | 4.6 | 5.9 |  |  |  | 1
| 1 | 2 | 6.2 | 6.37 | 6.1 | 5.89 | 6 | 6.43 | 6.06 | 6.06 | 6.78 | 5.44 | 4.87 | 5.5 |  |  |  | 7
| 1 | 3 | 5.9 | 6.07 | 6.14 | 5.92 | 5.99 | 6.54 | 5.97 | 6.4 | 6.23 | 5.39 | 4.92 | 6.1 |  |  |  | 16
| 1 | 4 | 6 | 6.03 | 5.83 | 5.66 | 5.99 | 6.3 | 5.93 | 6.51 | 6.18 | 5.74 | 5.38 | 6 |  |  |  | 6.0366
| 1 | 5 | 6.1 | 5.11 | 6.33 | 5.92 | 6.11 | 6.44 | 6.03 | 6.25 | 7.15 | 5.45 | 4.77 | 5.6 |  |  |  | 57
| 1 | 6 | 6.2 | 6.15 | 6.25 | 5.92 | 6.16 | 6.61 | 5.64 | 6.09 | 6.7 | 5.52 | 5.23 | 5.3
| Mean |  | 6.1 | 5.825 | 6.1183 | 5.8867 | 6.03 | 6.46 | 5.92 | 6.1867 | 6.635 | 5.5 | 4.9617 | 5.7333
| Median |  | 6.15 | 6.05 | 6.12 | 5.92 | 5.995 | 6.44 | 5.95 | 6.17 | 6.735 | 5.455 | 4.895 | 5.75
| Std Dev. |  | 0.12649 | 0.52573 | 0.17291 | 0.11827 | 0.086487 | 0.10602 | 0.15073 | 0.25351 | 0.36871 | 0.12474 | 0.29144 | 0.31411
| Rel.Std.Dev. |  | 2.074% | 9.025% | 2.826% | 2.009% | 1.434% | 1.641% | 2.546% | 4.098% | 5.557% | 2.268% | 5.874% | 5.479%
| PDM3 |  | 1.05% | -3.506% | 1.353% | -2.484% | -0.1099% | 7.013% | -1.932% | 2.485% | 9.912% | -8.89% | -17.81% | -5.024%
| Z-Score (Absolute) |  | 0.47 | 0.56 | 0.53 | 0.33 | 0.21 | 1.81 | 0.21 | 0.79 | 2.46 | 1.77 | 3.78 | 0.9
| Analytical results for Rb in OREAS 25b (Certified Value 61 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | PF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS |  | 0
1
| 1 | 1 | 58.6 | 54.3 | 67.4386 | 60 | 62.6 | 62.1 | 61 | 59.8 | 53.4 | 60 | 55.1 | 62.4 | 67 | 63.4 |  | 1
| 1 | 2 | 61.1 | 68.4 | 65.4433 | 62 | 61.1 | 61 | 60.8 | 59.6 | 60.6 | 57.2 | 55.6 | 60.3 | 67 | 60.3 |  | 8
| 1 | 3 | 60.6 | 67.8 | 67.1677 | 62 | 61.4 | 63.2 | 62.6 | 58.8 | 58.3 | 56.4 | 59.6 | 58.2 | 67 | 63.4 |  | 16
| 1 | 4 | 59.7 | 65.4 | 63.496 | 62 | 61.2 | 63.7 | 58.7 | 61.3 | 63.8 | 56.8 | 58.6 | 60.5 | 67 | 65.7 |  | 60.8711
| 1 | 5 | 61.4 | 54.4 | 64.3806 | 62 | 61.2 | 64.4 | 60.5 | 57.2 | 60.5 | 66.5 | 57.9 | 61 | 67 | 64.5 |  | 58
| 1 | 6 | 61.5 | 67.6 | 67.6541 | 61 | 59.6 | 62.8 | 60.6 | 54.9 | 60.4 | 66.4 | 57.3 | 58.9 | 70 | 62.5
| Mean |  | 60.4833 | 62.9833 | 65.93 | 61.5 | 61.1833 | 62.8667 | 60.7 | 58.6 | 59.5 | 60.55 | 57.35 | 60.2167 | 67.5 | 63.3
| Median |  | 60.85 | 66.5 | 66.3055 | 62 | 61.2 | 63 | 60.7 | 59.2 | 60.45 | 58.6 | 57.6 | 60.4 | 67 | 63.4
| Std Dev. |  | 1.1339 | 6.7641 | 1.7517 | 0.83666 | 0.95586 | 1.2028 | 1.2458 | 2.2548 | 3.4687 | 4.742 | 1.7352 | 1.5012 | 1.2247 | 1.8363
| Rel.Std.Dev. |  | 1.875% | 10.74% | 2.657% | 1.36% | 1.562% | 1.913% | 2.052% | 3.848% | 5.83% | 7.832% | 3.026% | 2.493% | 1.814% | 2.901%
| PDM3 |  | -0.637% | 3.47% | 8.311% | 1.033% | 0.5129% | 3.278% | -0.2811% | -3.731% | -2.252% | -0.5275% | -5.785% | -1.075% | 10.89% | 3.99%
| Z-Score (Absolute) |  | 0.18 | 0.82 | 2 | 0.22 | 0.1 | 0.77 | 0.1 | 0.94 | 0.58 | 0.16 | 1.44 | 0.29 | 2.63 | 0.94
| Analytical results for Re in OREAS 25b (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 04 | 12 | 16 | [11 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | [11 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 0.1 | < 1 | [11 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.1 | < 1 | [11 empty] | 16
| 1 | 3 | < 0.1 | < 0.1 | < 0.1 | < 1 | [11 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.1 | < 1 | [11 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 0.1 | < 1 | [11 empty] | 22
| 1 | 6 | < 0.1 | < 0.1 | < 0.1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for S in OREAS 25b (Indicative Value 0.033 wt.%) | [16 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 12 | 15 | 16 | [9 empty] | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 3
3
| 1 | 1 | < 0.05 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.04 | [9 empty] | 1
| 1 | 2 | < 0.05 | 0.03 | < 0.05 | < 0.05 | 0.03 | 0.03 | [9 empty] | 17
| 1 | 3 | < 0.05 | 0.03 | < 0.05 | < 0.05 | 0.03 | 0.04 | [9 empty] | 16
| 1 | 4 | < 0.05 | 0.02 | < 0.05 | < 0.05 | 0.04 | 0.03 | [9 empty] | 0.032778
| 1 | 5 | < 0.05 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.03 | [9 empty] | 23
| 1 | 6 | < 0.05 | 0.03 | < 0.05 | < 0.05 | 0.04 | 0.03
| Mean |  |  | 0.028333 |  |  | 0.036667 | 0.033333
| Median |  |  | 0.03 |  |  | 0.04 | 0.03
| Std Dev. |  |  | 0.0040825 |  |  | 0.005164 | 0.005164
| Rel.Std.Dev. |  |  | 14.41% |  |  | 14.08% | 15.49%
| PDM3 |  |  | -13.56% |  |  | 11.86% | 1.695%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 0.67 | 4.05 | 2.7
| Analytical results for Sb in OREAS 25b (Indicative Value 1.03 ppm) | [16 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 04 | 05 | 16 | [11 empty] | ppm
|  |  | PF*MS | PF*MS | BF*OES/MS | PF*MS | [11 empty] | 2
2
| 1 | 1 | 1 | 0.80303 | 0.8 | 1.4 | [11 empty] | 1
| 1 | 2 | 1.2 | 0.53557 | 0.9 | 1.4 | [11 empty] | 9
| 1 | 3 | 1.3 | 0.59683 | 1 | 1.1 | [11 empty] | 16
| 1 | 4 | 1.3 | 0.82425 | 1 | 1.3 | [11 empty] | 1.0268
| 1 | 5 | 1.2 | 0.54258 | 0.9 | 1.5 | [11 empty] | 19
| 1 | 6 | 1.2 | 0.84108 | 0.8 | 1.2
| Mean |  | 1.2 | 0.69056 | 0.9 | 1.3167
| Median |  | 1.2 | 0.69993 | 0.9 | 1.35
| Std Dev. |  | 0.10954 | 0.14689 | 0.089443 | 0.1472
| Rel.Std.Dev. |  | 9.129% | 21.27% | 9.938% | 11.18%
| PDM3 |  | 16.87% | -32.75% | -12.35% | 28.23%
| Z-Score (Absolute) |  | 0.49 | 1.16 | 0.49 | 0.86
| Analytical results for Sc in OREAS 25b (Certified Value 16.2 ppm) | [16 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 16 |  |  |  |  | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | PF*OES |  |  |  |  | 1
1
| 1 | 1 | 15 | < 20 | 15 | 15 | 15.4 | 18 | 18 | 17 | 18 | < 20 | 8 |  |  |  |  | 1
| 1 | 2 | 14 | < 20 | 16 | 14 | 14.9 | 18.2 | 17.9 | 16.6 | 16 | < 20 | 11 |  |  |  |  | #N/A
| 1 | 3 | 13 | < 20 | 16 | 15 | 16 | 17.6 | 20 | 17.2 | 16 | < 20 | < 5 |  |  |  |  | 16
| 1 | 4 | 13 | < 20 | 16 | 14 | 15 | 16.3 | 19 | 17.4 | 14 | < 20 | 11 |  |  |  |  | 16.1571
| 1 | 5 | 13 | < 20 | 16 | 15 | 15.5 | 16.4 | 20.3 | 15.8 | 17 | < 20 | 10 |  |  |  |  | 59
| 1 | 6 | 14 | < 20 | 16 | 15 | 15.2 | 19.2 | 22.7 | 16.6 | 17 | < 20 | 11
| Mean |  | 13.6667 |  | 15.8333 | 14.6667 | 15.3333 | 17.6167 | 19.65 | 16.7667 | 16.3333 |  | 10.2
| Median |  | 13.5 |  | 16 | 15 | 15.3 | 17.8 | 19.5 | 16.8 | 16.5 |  | 11
| Std Dev. |  | 0.8165 |  | 0.40825 | 0.5164 | 0.39833 | 1.1143 | 1.7919 | 0.57155 | 1.3663 |  | 1.3038
| Rel.Std.Dev. |  | 5.974% |  | 2.578% | 3.521% | 2.598% | 6.325% | 9.119% | 3.409% | 8.365% |  | 12.78%
| PDM3 |  | -15.41% |  | -2.004% | -9.225% | -5.098% | 9.034% | 21.62% | 3.773% | 1.091% |  | -36.87%
| Z-Score (Absolute) |  | 0.67 | 2.16 | 0.2 | 0.27 | 0 | 0.92 | 1.75 | 0.58 | 0.4 | 2.16 | 2.6
| Analytical results for Se in OREAS 25b (Indicative Value < 20 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 04 | 16 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 20 | < 20 | < 50 | [12 empty] | 1
| 1 | 2 | < 20 | < 20 | < 50 | [12 empty] | 14
| 1 | 3 | < 20 | < 20 | < 50 | [12 empty] | 16
| 1 | 4 | < 20 | < 20 | < 50 | [12 empty] | < 20
| 1 | 5 | < 20 | < 20 | < 50 | [12 empty] | 20
| 1 | 6 | < 20 | < 20 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Si in OREAS 25b (Certified Value 26.17 wt.%) | [16 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 11 | 15 | 16 | [6 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | [6 empty] | 2
3
| 1 | 1 | 26 | 26.272 | 24.75 | 25.223 | 26.379 | 26.337 | 27.5337 | 25.7 | 27.53 | [6 empty] | 1
| 1 | 2 | 26.7 | 26.627 | 24.97 | 25.3374 | 26.253 | 26.458 | 26.7857 | 26.4 | 28.28 | [6 empty] | #N/A
| 1 | 3 | 26.8 | 26.145 | 24.88 | 25.3393 | 26.622 | 26.482 | 27.2064 | 19.7 | 28.05 | [6 empty] | 16
| 1 | 4 | 26.4 | 28.039 | 24.55 | 25.2675 | 26.547 | 26.435 | 26.0378 | 25 | 28.56 | [6 empty] | 26.1678
| 1 | 5 | 27.2 | 25.958 | 24.92 | 25.1093 | 26.468 | 26.379 | 26.926 | 25 | 28.52 | [6 empty] | 60
| 1 | 6 | 26.7 | 25.813 | 24.96 | 25.1126 | 26.589 | 26.29 | 26.2715 | 25.3 | 28.28
| Mean |  | 26.6333 | 26.4757 | 24.8383 | 25.2315 | 26.4763 | 26.3968 | 26.7935 | 24.5167 | 28.2033
| Median |  | 26.7 | 26.2085 | 24.9 | 25.2453 | 26.5075 | 26.407 | 26.8558 | 25.15 | 28.28
| Std Dev. |  | 0.40332 | 0.81546 | 0.16216 | 0.10323 | 0.14023 | 0.074516 | 0.56206 | 2.4178 | 0.37845
| Rel.Std.Dev. |  | 1.514% | 3.08% | 0.6529% | 0.4091% | 0.5297% | 0.2823% | 2.098% | 9.862% | 1.342%
| PDM3 |  | 1.779% | 1.176% | -5.081% | -3.578% | 1.179% | 0.8752% | 2.391% | -6.31% | 7.779%
| Z-Score (Absolute) |  | 0.33 | 0 | 3.47 | 2.64 | 0 | 0.17 | 0.67 | 4.16 | 3.67
| Analytical results for Sm in OREAS 25b (Certified Value 4.19 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 4.4 | 3.8 | 4.4 | 3.93 | 4.43 | 4.9 | 4.39 | 3.86 | 4.58 | 3.73 | 3 | < 100 |  |  |  | 1
| 1 | 2 | 4.2 | 4 | 4.4 | 3.99 | 4.33 | 4.42 | 4.6 | 4.51 | 4.84 | 3.87 | 3.3 | < 100 |  |  |  | 10
| 1 | 3 | 4.4 | 4.7 | 4.5 | 3.94 | 4.36 | 4.68 | 4.29 | 4.65 | 4.63 | 3.66 | 3.9 | < 100 |  |  |  | 16
| 1 | 4 | 3.8 | 4.3 | 4.3 | 3.93 | 4.14 | 4.32 | 3.86 | 4.04 | 4.25 | 3.88 | 3.6 | < 100 |  |  |  | 4.1861
| 1 | 5 | 4.1 | 3.4 | 4.4 | 3.71 | 4.51 | 4.45 | 4.5 | 3.97 | 4.53 | 3.66 | 3.2 | < 100 |  |  |  | 61
| 1 | 6 | 4.3 | 4.4 | 4.3 | 3.99 | 4.43 | 4.43 | 4.34 | 4.28 | 4.56 | 3.84 | 4 | < 100
| Mean |  | 4.2 | 4.1 | 4.3833 | 3.915 | 4.3667 | 4.5333 | 4.33 | 4.2183 | 4.565 | 3.7733 | 3.5
| Median |  | 4.25 | 4.15 | 4.4 | 3.935 | 4.395 | 4.44 | 4.365 | 4.16 | 4.57 | 3.785 | 3.45
| Std Dev. |  | 0.22804 | 0.46476 | 0.075277 | 0.10426 | 0.12754 | 0.21538 | 0.25612 | 0.3153 | 0.19003 | 0.1027 | 0.4
| Rel.Std.Dev. |  | 5.429% | 11.34% | 1.717% | 2.663% | 2.921% | 4.751% | 5.915% | 7.475% | 4.163% | 2.722% | 11.43%
| PDM3 |  | 0.333% | -2.056% | 4.713% | -6.475% | 4.314% | 8.296% | 3.439% | 0.771% | 9.052% | -9.86% | -16.39%
| Z-Score (Absolute) |  | 0.23 | 0.54 | 0.34 | 1.12 | 0.29 | 0.81 | 0.17 | 0.17 | 0.91 | 1.56 | 2.41 | 142.31
| Analytical results for Sn in OREAS 25b (Certified Value 7.13 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 14 | 16 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS |  |  |  |  | 2
3
| 1 | 1 | 7 | 7.1 | < 100 | 7 | 7 | 6.9 | 8.8 | 6.6 | 7.1 | 7 | 7 |  |  |  |  | 1
| 1 | 2 | 7 | < 0.3 | < 100 | 7 | 7 | 6.8 | 8.2 | 7.4 | 7.3 | 6.7 | 7 |  |  |  |  | 11
| 1 | 3 | 8 | < 0.3 | < 100 | 8 | 7 | 7.7 | 7.8 | 7 | 7.3 | 6.7 | 8 |  |  |  |  | 16
| 1 | 4 | 7 | < 0.3 | < 100 | 7 | 7 | 7.5 | 8.2 | 6.3 | 7.5 | 6.4 | 8 |  |  |  |  | 7.1343
| 1 | 5 | 8 | 6.6 | < 100 | 7 | 7 | 6.6 | 7 | 6.6 | 7 | 7.3 | 7 |  |  |  |  | 62
| 1 | 6 | 7 | < 0.3 | < 100 | 7 | 7 | 7.1 | 7.5 | 6.6 | 6.7 | 7.2 | 8
| Mean |  | 7.3333 | 6.85 |  | 7.1667 | 7 | 7.1 | 7.9167 | 6.75 | 7.15 | 6.8833 | 7.5
| Median |  | 7 | 6.85 |  | 7 | 7 | 7 | 8 | 6.6 | 7.2 | 6.85 | 7.5
| Std Dev. |  | 0.5164 | 0.35355 |  | 0.40825 | 0 | 0.42426 | 0.62743 | 0.38859 | 0.28107 | 0.34303 | 0.54772
| Rel.Std.Dev. |  | 7.042% | 5.161% |  | 5.696% | 0% | 5.976% | 7.925% | 5.757% | 3.931% | 4.983% | 7.303%
| PDM3 |  | 2.789% | -3.985% |  | 0.4532% | -1.883% | -0.4812% | 10.97% | -5.387% | 0.2196% | -3.518% | 5.125%
| Z-Score (Absolute) |  | 0.46 | 12.05 | 108.35 | 0.04 | 0.38 | 0.13 | 1.94 | 1.01 | 0 | 0.67 | 0.89
| Analytical results for Sr in OREAS 25b (Certified Value 33.6 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | BF*MS | BF*OES | PF*OES/MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | PF*OES | 1
2
| 1 | 1 | 32.3 | 31 | 38 | 30.9314 | 33 | 32.6 | 31.1 | 32.6 | 31.9 | 30.9 | 35.1 | 40 | 33.7 | 40 | 35 | 1
| 1 | 2 | 32.2 | 35 | 36.7 | 26.2387 | 34 | 32.8 | 31.9 | 32.1 | 32.9 | 29.9 | 33.4 | 40 | 33.5 | 40 | 35 | #N/A
| 1 | 3 | 32.3 | 31 | 36.7 | 31.3884 | 34 | 33 | 38.4 | 32.5 | 29.4 | 32.4 | 31.4 | 40 | 32.5 | 25 | 35 | 16
| 1 | 4 | 33.3 | 39 | 36.2 | 31.3295 | 33 | 30.5 | 32.9 | 31.3 | 30 | 33.3 | 31.7 | 50 | 33.9 | 35 | 35 | 33.619
| 1 | 5 | 33.3 | 31 | 36.7 | 26.9079 | 33 | 32.4 | 34 | 32.9 | 31.5 | 32.3 | 35.9 | 40 | 33.9 | 35 | 35 | 63
| 1 | 6 | 33.2 | 33 | 37.3 | 30.9742 | 34 | 31.1 | 34.4 | 32.7 | 33.2 | 31.8 | 35.5 | 40 | 31.9 | 35 | 35
| Mean |  | 32.7667 | 33.3333 | 36.9333 | 29.6284 | 33.5 | 32.0667 | 33.7833 | 32.35 | 31.4833 | 31.7667 | 33.8333 | 41.6667 | 33.2333 | 35 | 35
| Median |  | 32.75 | 32 | 36.7 | 30.9528 | 33.5 | 32.5 | 33.45 | 32.55 | 31.7 | 32.05 | 34.25 | 40 | 33.6 | 35 | 35
| Std Dev. |  | 0.55015 | 3.2042 | 0.62823 | 2.3829 | 0.54772 | 1.0191 | 2.5795 | 0.57879 | 1.5276 | 1.2061 | 1.9654 | 4.0825 | 0.83586 | 5.4772 | 0
| Rel.Std.Dev. |  | 1.679% | 9.612% | 1.701% | 8.043% | 1.635% | 3.178% | 7.635% | 1.789% | 4.852% | 3.797% | 5.809% | 9.798% | 2.515% | 15.65% | 0%
| PDM3 |  | -2.535% | -0.8496% | 9.859% | -11.87% | -0.3539% | -4.617% | 0.4889% | -3.775% | -6.352% | -5.51% | 0.6376% | 23.94% | -1.147% | 4.108% | 4.108%
| Z-Score (Absolute) |  | 0.32 | 0.03 | 2.29 | 2.29 | 0.14 | 0.76 | 0.31 | 0.59 | 1.13 | 0.95 | 0.34 | NA | 0.03 | 1.08 | 1.08
Results from laboratory 12 were removed due to their 10 ppm reading resolution.
| Analytical results for Ta in OREAS 25b (Certified Value 1.93 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 2 | 1.81 | 1.8686 | 2.1 | 1.8 | 1.8 | 1.9 | 2 | 1.7 | 2.1 | 2.1 | < 10 |  |  |  | 1
| 1 | 2 | 1.9 | 2.14 | 1.7469 | 2.2 | 1.8 | 1.8 | 1.9 | 2 | 1.8 | 2.1 | 2 | < 10 |  |  |  | 1
| 1 | 3 | 2 | 2.21 | 1.6483 | 2.1 | 1.7 | 1.8 | 1.8 | 2 | 1.8 | 1.9 | 2 | < 10 |  |  |  | 16
| 1 | 4 | 2.1 | 2.09 | 1.7601 | 2.1 | 1.8 | 1.8 | 1.8 | 2 | 1.9 | 2 | 2 | < 10 |  |  |  | 1.9336
| 1 | 5 | 2.1 | 1.73 | 1.6557 | 2.2 | 1.8 | 1.8 | 1.8 | 2 | 1.8 | 2.2 | 2 | < 10 |  |  |  | 64
| 1 | 6 | 2 | 2 | 1.8551 | 2.2 | 2 | 1.8 | 1.9 | 1.9 | 1.9 | 2.1 | 2 | < 10
| Mean |  | 2.0167 | 1.9967 | 1.7558 | 2.15 | 1.8167 | 1.8 | 1.85 | 1.9833 | 1.8167 | 2.0667 | 2.0167
| Median |  | 2 | 2.045 | 1.7535 | 2.15 | 1.8 | 1.8 | 1.85 | 2 | 1.8 | 2.1 | 2
| Std Dev. |  | 0.075277 | 0.19012 | 0.094104 | 0.054772 | 0.098319 | 0 | 0.054772 | 0.040825 | 0.075277 | 0.10328 | 0.040825
| Rel.Std.Dev. |  | 3.733% | 9.522% | 5.36% | 2.548% | 5.412% | 0% | 2.961% | 2.058% | 4.144% | 4.997% | 2.024%
| PDM3 |  | 4.298% | 3.264% | -9.193% | 11.19% | -6.045% | -6.907% | -4.321% | 2.574% | -6.045% | 6.884% | 4.298%
| Z-Score (Absolute) |  | 0.12 | 0.03 | 1.05 | 0.72 | 0.78 | 0.85 | 0.63 | 0.03 | 0.78 | 0.34 | 0.12 | 13.53
| Analytical results for Tb in OREAS 25b (Certified Value 0.6 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 0.6 | 0.53 | 0.6 | 0.58 | 0.57 | 0.62 | 0.66 | 0.59 | 0.73 | 0.56 | 0.28 | 0.6 |  |  |  | 1
| 1 | 2 | 0.6 | 0.49 | 0.6 | 0.6 | 0.57 | 0.59 | 0.63 | 0.53 | 0.71 | 0.57 | 0.29 | 0.6 |  |  |  | #N/A
| 1 | 3 | 0.6 | 0.46 | 0.6 | 0.6 | 0.6 | 0.64 | 0.65 | 0.64 | 0.61 | 0.56 | 0.29 | 0.6 |  |  |  | 16
| 1 | 4 | 0.5 | 0.43 | 0.6 | 0.56 | 0.63 | 0.59 | 0.66 | 0.59 | 0.66 | 0.59 | 0.36 | 0.6 |  |  |  | 0.5953
| 1 | 5 | 0.6 | 0.5 | 0.6 | 0.58 | 0.6 | 0.59 | 0.58 | 0.58 | 0.81 | 0.55 | 0.33 | 0.6 |  |  |  | 65
| 1 | 6 | 0.6 | 0.45 | 0.6 | 0.57 | 0.61 | 0.59 | 0.62 | 0.57 | 0.73 | 0.57 | 0.34 | 0.5
| Mean |  | 0.58333 | 0.47667 | 0.6 | 0.58167 | 0.59667 | 0.60333 | 0.63333 | 0.58333 | 0.70833 | 0.56667 | 0.315 | 0.58333
| Median |  | 0.6 | 0.475 | 0.6 | 0.58 | 0.6 | 0.59 | 0.64 | 0.585 | 0.72 | 0.565 | 0.31 | 0.6
| Std Dev. |  | 0.040825 | 0.036697 | 0 | 0.016021 | 0.023381 | 0.021602 | 0.030768 | 0.03559 | 0.06824 | 0.013663 | 0.032711 | 0.040825
| Rel.Std.Dev. |  | 6.999% | 7.699% | 0% | 2.754% | 3.919% | 3.581% | 4.858% | 6.101% | 9.634% | 2.411% | 10.38% | 6.999%
| PDM3 |  | -2.01% | -19.93% | 0.7901% | -2.29% | 0.2302% | 1.35% | 6.39% | -2.01% | 18.99% | -4.809% | -47.09% | -2.01%
| Z-Score (Absolute) |  | 0 | 4.32 | 0.67 | 0.07 | 0.54 | 0.81 | 2.02 | 0 | 5.06 | 0.67 | 10.86 | 0
| Analytical results for Te in OREAS 25b (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 04 | 16 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 1 | < 1 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | < 1 | < 1 | [12 empty] | 15
| 1 | 3 | < 1 | < 1 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | < 1 | < 1 | [12 empty] | < 1
| 1 | 5 | < 1 | < 1 | < 1 | [12 empty] | 21
| 1 | 6 | < 1 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Th in OREAS 25b (Certified Value 13.2 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | BF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS |  | 1
2
| 1 | 1 | 13.3 | 12.9 | 12.3654 | 13.3 | 12.7 | 12.65 | 13.7 | 12.75 | 12.4 | 15.3 | 15.5 | 14.1 | 11.1 | 12.6 |  | 1
| 1 | 2 | 13.2 | 13.6 | 12.5817 | 13.3 | 12.9 | 13.45 | 13.15 | 12.75 | 12.8 | 14.6 | 15.5 | 13.4 | 11.6 | 12.6 |  | 2
| 1 | 3 | 13.5 | 13.5 | 12.3745 | 13.6 | 13.1 | 13.25 | 13.8 | 12.65 | 12.95 | 13.3 | 15.4 | 13 | 11.2 | 13.8 |  | 16
| 1 | 4 | 13.3 | 13.2 | 12.2802 | 13.3 | 12.4 | 13.6 | 13.15 | 12.5 | 13.15 | 13.3 | 15.9 | 13.4 | 12.2 | 13.5 |  | 13.1662
| 1 | 5 | 13.3 | 12.8 | 11.5494 | 13.5 | 12.7 | 13.5 | 13.35 | 12.65 | 13.15 | 15.1 | 15.5 | 13.3 | 11.7 | 13.2 |  | 66
| 1 | 6 | 13.4 | 13.5 | 11.6924 | 13.4 | 12.9 | 13.55 | 13.75 | 13 | 12.85 | 14.7 | 15.1 | 12.8 | 12.4 | 12.3
| Mean |  | 13.3333 | 13.25 | 12.1406 | 13.4 | 12.7833 | 13.3333 | 13.4833 | 12.7167 | 12.8833 | 14.3833 | 15.4833 | 13.3333 | 11.7 | 13
| Median |  | 13.3 | 13.35 | 12.3228 | 13.35 | 12.8 | 13.475 | 13.525 | 12.7 | 12.9 | 14.65 | 15.5 | 13.35 | 11.65 | 12.9
| Std Dev. |  | 0.10328 | 0.33912 | 0.41706 | 0.12649 | 0.24014 | 0.3559 | 0.30277 | 0.16633 | 0.27869 | 0.87731 | 0.25626 | 0.44572 | 0.52154 | 0.58992
| Rel.Std.Dev. |  | 0.7746% | 2.559% | 3.435% | 0.944% | 1.879% | 2.669% | 2.245% | 1.308% | 2.163% | 6.099% | 1.655% | 3.343% | 4.458% | 4.538%
| PDM3 |  | 1.27% | 0.6368% | -7.789% | 1.776% | -2.908% | 1.27% | 2.409% | -3.414% | -2.148% | 9.245% | 17.6% | 1.27% | -11.14% | -1.262%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 2.22 | 0.21 | 0.98 | 0.08 | 0.37 | 1.11 | 0.79 | 2.1 | 4.22 | 0.08 | 3.07 | 0.56
| Analytical results for Ti in OREAS 25b (Certified Value 1.11 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  | wt.%
|  |  | BF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 1.157 | 1.061 | 1.1 | 1.0805 | 1.0785 | 1.103 | 1.13 | 1.17 | 1.17 | 1.04 | 1.0491 | 1.1 | 1.06 | 1.06 |  | 1
| 1 | 2 | 1.115 | 1.091 | 1.13 | 1.0923 | 1.0887 | 1.103 | 1.12 | 1.18 | 1.17 | 1.07 | 0.8393 | 1.1 | 1.07 | 1.08 |  | #N/A
| 1 | 3 | 1.127 | 1.079 | 1.12 | 1.106 | 1.0935 | 1.103 | 1.17 | 1.18 | 1.15 | 1.08 | 0.9652 | 1.1 | 0.85 | 1.07 |  | 16
| 1 | 4 | 1.121 | 1.193 | 1.09 | 1.1058 | 1.0773 | 1.097 | 1.15 | 1.13 | 1.16 | 1.13 | 1.115 | 1.2 | 1.07 | 1.09 |  | 1.1067
| 1 | 5 | 1.133 | 1.049 | 1.11 | 1.0881 | 1.0887 | 1.103 | 1.14 | 1.19 | 1.15 | 1.11 | 1.139 | 1.1 | 1.04 | 1.09 |  | 67
| 1 | 6 | 1.097 | 1.067 | 1.13 | 1.1148 | 1.0857 | 1.109 | 1.16 | 1.17 | 1.02 | 1.08 | 1.145 | 1.1 | 1.07 | 1.08
| Mean |  | 1.125 | 1.09 | 1.1133 | 1.0979 | 1.0854 | 1.103 | 1.145 | 1.17 | 1.1367 | 1.085 | 1.0421 | 1.1167 | 1.0267 | 1.0783
| Median |  | 1.124 | 1.073 | 1.115 | 1.0991 | 1.0872 | 1.103 | 1.145 | 1.175 | 1.155 | 1.08 | 1.082 | 1.1 | 1.065 | 1.08
| Std Dev. |  | 0.01996 | 0.052501 | 0.01633 | 0.012978 | 0.0063356 | 0.0037947 | 0.018708 | 0.020976 | 0.05785 | 0.031464 | 0.12032 | 0.040825 | 0.087331 | 0.01169
| Rel.Std.Dev. |  | 1.774% | 4.817% | 1.467% | 1.182% | 0.5837% | 0.344% | 1.634% | 1.793% | 5.089% | 2.9% | 11.55% | 3.656% | 8.506% | 1.084%
| PDM3 |  | 1.653% | -1.509% | 0.5991% | -0.7947% | -1.925% | -0.3346% | 3.46% | 5.719% | 2.707% | -1.961% | -5.837% | 0.9003% | -7.232% | -2.563%
| Z-Score (Absolute) |  | 0.86 | 0.37 | 0.45 | 0.09 | 0.53 | 0.09 | 1.57 | 2.45 | 1.27 | 0.54 | 2.05 | 0.57 | 2.6 | 0.78
| Analytical results for Tl in OREAS 25b (Certified Value < 0.5 ppm) | [16 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 11 | 12 | 16 | [8 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | < 0.5 | 0.4 | < 0.5 | < 0.5 | 0.4 | [8 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | < 0.5 | 0.4 | < 0.5 | < 0.5 | 0.4 | [8 empty] | #N/A
| 1 | 3 | < 0.5 | < 0.5 | < 0.5 | 0.4 | < 0.5 | < 0.5 | 0.3 | [8 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | < 0.5 | 0.4 | < 0.5 | < 0.5 | 0.4 | [8 empty] | < 0.5
| 1 | 5 | < 0.5 | < 0.5 | < 0.5 | 0.4 | < 0.5 | < 0.5 | 0.4 | [8 empty] | 68
| 1 | 6 | < 0.5 | < 0.5 | < 0.5 | 0.4 | < 0.5 | < 0.5 | 0.3
| Mean |  |  |  |  | 0.4 |  |  | 0.36667
| Median |  |  |  |  | 0.4 |  |  | 0.4
| Std Dev. |  |  |  |  | 6.0809e-17 |  |  | 0.05164
| Rel.Std.Dev. |  |  |  |  | 1.52e-14% |  |  | 14.08%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tm in OREAS 25b (Certified Value 0.32 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 |  |  |  |  | ppm
|  |  | PF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS |  |  |  |  | 2
3
| 1 | 1 | 0.3 | 0.31 | 0.33 | 0.31 | 0.33 | 0.31 | 0.35 | 0.28 | 0.38 | 0.29 | 0.3 |  |  |  |  | 1
| 1 | 2 | 0.3 | 0.33 | 0.32 | 0.32 | 0.31 | 0.37 | 0.3 | 0.3 | 0.33 | 0.3 | 0.3 |  |  |  |  | 25
| 1 | 3 | 0.3 | 0.33 | 0.32 | 0.32 | 0.3 | 0.38 | 0.37 | 0.32 | 0.29 | 0.28 | 0.3 |  |  |  |  | 16
| 1 | 4 | 0.3 | 0.3 | 0.34 | 0.3 | 0.31 | 0.33 | 0.31 | 0.32 | 0.32 | 0.29 | 0.3 |  |  |  |  | 0.31941
| 1 | 5 | 0.3 | 0.28 | 0.32 | 0.31 | 0.3 | 0.31 | 0.31 | 0.31 | 0.48 | 0.3 | 0.3 |  |  |  |  | 69
| 1 | 6 | 0.3 | 0.28 | 0.33 | 0.31 | 0.35 | 0.35 | 0.35 | 0.28 | 0.42 | 0.29 | 0.3
| Mean |  | 0.3 | 0.305 | 0.32667 | 0.31167 | 0.31667 | 0.34167 | 0.33167 | 0.30167 | 0.37 | 0.29167 | 0.3
| Median |  | 0.3 | 0.305 | 0.325 | 0.31 | 0.31 | 0.34 | 0.33 | 0.305 | 0.355 | 0.29 | 0.3
| Std Dev. |  | 0 | 0.022583 | 0.008165 | 0.0075277 | 0.019664 | 0.029944 | 0.028577 | 0.018348 | 0.070993 | 0.0075277 | 0
| Rel.Std.Dev. |  | 0% | 7.404% | 2.499% | 2.415% | 6.21% | 8.764% | 8.616% | 6.082% | 19.19% | 2.581% | 0%
| PDM3 |  | -6.076% | -4.511% | 2.273% | -2.423% | -0.8581% | 6.969% | 3.838% | -5.554% | 15.84% | -8.685% | -6.076%
| Z-Score (Absolute) |  | NA | 0.52 | 0.45 | 0.22 | 0 | 1.12 | 0.67 | 0.67 | 2.4 | 1.12 | NA
Results from laboratory 1 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 25b (Certified Value 2.73 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | PF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS |  | 2
3
| 1 | 1 | 2.7 | 2.8 | 2.4462 | 2.8 | 2.8 | 2.6 | 2.69 | 2.6 | 2.52 | 3.08 | 3 | 3 | 2 | 2.7 |  | 1
| 1 | 2 | 2.8 | 2.86 | 2.5017 | 2.8 | 2.7 | 2.6 | 2.77 | 2.56 | 2.48 | 2.87 | 3 | 2.9 | 2 | 2.5 |  | 4
| 1 | 3 | 2.7 | 2.83 | 2.431 | 2.8 | 2.6 | 2.67 | 2.72 | 2.49 | 2.67 | 2.73 | 3 | 2.8 | 2 | 2.9 |  | 16
| 1 | 4 | 2.8 | 2.78 | 2.4798 | 2.8 | 2.7 | 2.62 | 2.67 | 2.54 | 2.7 | 2.8 | 3.1 | 2.9 | 2 | 2.8 |  | 2.7267
| 1 | 5 | 2.8 | 2.67 | 2.4005 | 2.8 | 2.7 | 2.65 | 2.81 | 2.6 | 2.62 | 3.05 | 2.8 | 2.8 | 2 | 2.6 |  | 70
| 1 | 6 | 2.7 | 2.87 | 2.4067 | 2.8 | 2.7 | 2.67 | 2.81 | 2.69 | 2.56 | 3.07 | 2.8 | 2.8 | 2 | 2.4
| Mean |  | 2.75 | 2.8017 | 2.4443 | 2.8 | 2.7 | 2.635 | 2.745 | 2.58 | 2.5917 | 2.9333 | 2.95 | 2.8667 | 2 | 2.65
| Median |  | 2.75 | 2.815 | 2.4386 | 2.8 | 2.7 | 2.635 | 2.745 | 2.58 | 2.59 | 2.96 | 3 | 2.85 | 2 | 2.65
| Std Dev. |  | 0.054772 | 0.073052 | 0.040174 | 4.8648e-16 | 0.063246 | 0.032711 | 0.060581 | 0.067823 | 0.086352 | 0.15293 | 0.12247 | 0.08165 | 0 | 0.18708
| Rel.Std.Dev. |  | 1.992% | 2.607% | 1.644% | 1.737e-14% | 2.342% | 1.241% | 2.207% | 2.629% | 3.332% | 5.213% | 4.152% | 2.848% | 0% | 7.06%
| PDM3 |  | 0.8529% | 2.748% | -10.36% | 2.687% | -0.9808% | -3.365% | 0.6695% | -5.382% | -4.954% | 7.576% | 8.188% | 5.132% | -26.65% | -2.814%
| Z-Score (Absolute) |  | 0.03 | 0.35 | 1.84 | 0.34 | 0.28 | 0.67 | 0 | 1.01 | 0.94 | 1.15 | 1.26 | 0.75 | NA | 0.58
| Analytical results for V in OREAS 25b (Certified Value 188 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | BF*OES | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES | PF*OES | 0
0
| 1 | 1 | 192 | 181 | 184 | 184.266 | 186 | 191 | 194 | 190 | 196 | 184 | 198 | 190 | 196 | 165 | 185 | 1
| 1 | 2 | 181 | 195 | 188 | 181.2763 | 185 | 189 | 192 | 190 | 202 | 179 | 192 | 190 | 186 | 165 | 185 | #N/A
| 1 | 3 | 189 | 196 | 189 | 181.8858 | 188 | 193 | 199 | 193 | 193 | 221 | 178 | 180 | 178 | 134 | 185 | 16
| 1 | 4 | 185 | 217 | 184 | 181.05 | 187 | 191 | 195 | 183 | 190 | 191 | 174 | 190 | 183 | 164 | 195 | 188.2093
| 1 | 5 | 187 | 184 | 187 | 183.5672 | 187 | 192 | 197 | 190 | 188 | 182 | 204 | 190 | 190 | 159 | 190 | 71
| 1 | 6 | 179 | 210 | 187 | 182.7328 | 188 | 190 | 200 | 190 | 179 | 181 | 213 | 190 | 181 | 164 | 185
| Mean |  | 185.5 | 197.1667 | 186.5 | 182.463 | 186.8333 | 191 | 196.1667 | 189.3333 | 191.3333 | 189.6667 | 193.1667 | 188.3333 | 185.6667 | 158.5 | 187.5
| Median |  | 186 | 195.5 | 187 | 182.3093 | 187 | 191 | 196 | 190 | 191.5 | 183 | 195 | 190 | 184.5 | 164 | 185
| Std Dev. |  | 4.8888 | 14.1339 | 2.0736 | 1.2869 | 1.169 | 1.4142 | 3.0605 | 3.3267 | 7.7889 | 15.8955 | 15.0521 | 4.0825 | 6.532 | 12.2107 | 4.1833
| Rel.Std.Dev. |  | 2.635% | 7.168% | 1.112% | 0.7053% | 0.6257% | 0.7404% | 1.56% | 1.757% | 4.071% | 8.381% | 7.792% | 2.168% | 3.518% | 7.704% | 2.231%
| PDM3 |  | -1.439% | 4.759% | -0.9082% | -3.053% | -0.7311% | 1.483% | 4.228% | 0.5972% | 1.66% | 0.7744% | 2.634% | 0.06592% | -1.351% | -15.79% | -0.3768%
| Z-Score (Absolute) |  | 0.72 | 2.23 | 0.46 | 1.48 | 0.38 | 0.67 | 1.98 | 0.25 | 0.76 | 0.34 | 1.22 | 0 | 0.67 | 7.54 | 0.21
| Analytical results for W in OREAS 25b (Certified Value 2.9 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 |  |  |  | ppm
|  |  | PF*MS | BF*MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS |  |  |  | 2
2
| 1 | 1 | 3 | < 0.1 | 2.4325 | 2 | 2.3 | 2.6 | 3 | 3.8 | 5.1 | 2.83 | 3.2 | 3 |  |  |  | 1
| 1 | 2 | 2 | 2.3 | 2.2456 | 4 | 2.5 | 2.5 | 2.9 | 4 | 3.2 | 2.66 | 3.6 | 3 |  |  |  | 6
| 1 | 3 | 3 | 3.7 | 2.4679 | 10 | 2.4 | 2.6 | 3 | 3.8 | 4.4 | 2.63 | 2 | 3 |  |  |  | 16
| 1 | 4 | 3 | 5 | 2.2488 | 3 | 2 | 3 | 2.8 | 3 | 2.7 | 2.54 | 2.6 | 2 |  |  |  | 2.8978
| 1 | 5 | 5 | < 0.1 | 2.3785 | 3 | 1.9 | 2.4 | 2.7 | 3.6 | 3.2 | 3.13 | 2.7 | 3 |  |  |  | 72
| 1 | 6 | 4 | 3.5 | 2.3657 | 3 | 2 | 3 | 2.8 | 3.7 | 3.1 | 2.85 | 3.6 | 3
| Mean |  | 3.3333 | 3.625 | 2.3565 | 4.1667 | 2.1833 | 2.6833 | 2.8667 | 3.65 | 3.6167 | 2.7733 | 2.95 | 2.8333
| Median |  | 3 | 3.6 | 2.3721 | 3 | 2.15 | 2.6 | 2.85 | 3.75 | 3.2 | 2.745 | 2.95 | 3
| Std Dev. |  | 1.0328 | 1.1057 | 0.092368 | 2.9269 | 0.24833 | 0.25626 | 0.12111 | 0.34496 | 0.92394 | 0.21172 | 0.63166 | 0.40825
| Rel.Std.Dev. |  | 30.98% | 30.5% | 3.92% | 70.25% | 11.37% | 9.55% | 4.225% | 9.451% | 25.55% | 7.634% | 21.41% | 14.41%
| PDM3 |  | 15.03% | 25.1% | -18.68% | 43.79% | -24.65% | -7.4% | -1.073% | 25.96% | 24.81% | -4.294% | 1.803% | -2.223%
| Z-Score (Absolute) |  | NA | 0.67 | 0.83 | NA | 1.17 | 0.18 | 0.19 | 1.74 | 1.67 | 0 | 0.35 | NA
Results from laboratories 1, 5 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 25b (Certified Value 20.6 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | BF*MS | BF*MS | PF*OES/MS | PF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*MS | 1
2
| 1 | 1 | 19.9 | 18.69 | 20.6 | 20.2925 | 21 | 20.5 | 21.5 | 20.8 | 20.3 | 19.1 | 23 | 15.3 | 21.1 | 18.6 | 20.5 | 1
| 1 | 2 | 19.9 | 22.14 | 20.6 | 19.4318 | 21 | 20.8 | 20.5 | 21.5 | 20.4 | 20 | 22 | 14.9 | 20.9 | 19.9 | 20.2 | 7
| 1 | 3 | 20.8 | 20.94 | 20.9 | 19.776 | 21 | 20.7 | 22.6 | 21.8 | 19.4 | 20.1 | 20.5 | 16 | 19.1 | 19.2 | 21.6 | 16
| 1 | 4 | 20.7 | 20.12 | 20.3 | 19.9304 | 21 | 19 | 21.4 | 19.5 | 20 | 20.6 | 20.6 | 15.3 | 20 | 21.9 | 21.3 | 20.5914
| 1 | 5 | 20.5 | 17.13 | 20.3 | 19.0093 | 21 | 19.7 | 20.5 | 20.7 | 19.5 | 20.7 | 22.5 | 15.5 | 20.6 | 21.6 | 21.5 | 73
| 1 | 6 | 21.1 | 22.12 | 21.1 | 19.1265 | 21 | 19.4 | 23.2 | 22.2 | 19.4 | 19.5 | 22.9 | 15.4 | 18.9 | 21.4 | 19.5
| Mean |  | 20.4833 | 20.19 | 20.6333 | 19.5944 | 21 | 20.0167 | 21.6167 | 21.0833 | 19.8333 | 20 | 21.9167 | 15.4 | 20.1 | 20.4333 | 20.7667
| Median |  | 20.6 | 20.53 | 20.6 | 19.6039 | 21 | 20.1 | 21.45 | 21.15 | 19.75 | 20.05 | 22.25 | 15.35 | 20.3 | 20.65 | 20.9
| Std Dev. |  | 0.4916 | 1.9854 | 0.32042 | 0.49415 | 0 | 0.75211 | 1.098 | 0.96626 | 0.45898 | 0.61968 | 1.1161 | 0.35777 | 0.93167 | 1.3866 | 0.83825
| Rel.Std.Dev. |  | 2.4% | 9.833% | 1.553% | 2.522% | 0% | 3.757% | 5.08% | 4.583% | 2.314% | 3.098% | 5.092% | 2.323% | 4.635% | 6.786% | 4.037%
| PDM3 |  | -0.5249% | -1.949% | 0.2036% | -4.842% | 1.984% | -2.791% | 4.979% | 2.389% | -3.682% | -2.872% | 6.436% | -25.21% | -2.386% | -0.7677% | 0.8511%
| Z-Score (Absolute) |  | 0.08 | 0.38 | 0.31 | 1.31 | 0.88 | 0.65 | 1.84 | 1.01 | 0.93 | 0.67 | 2.31 | 7.83 | 0.52 | 0 | 0.52
| Analytical results for Yb in OREAS 25b (Certified Value 2.16 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS |  |  |  | 2
3
| 1 | 1 | 2.1 | 2.1 | 2.2 | 1.96 | 2.03 | 2.21 | 2.4 | 2.14 | 2.52 | 2.39 | 1.7 | 1.7 |  |  |  | 1
| 1 | 2 | 1.8 | 2.3 | 2.1 | 2 | 2.02 | 2.13 | 2.36 | 2 | 2.36 | 2.38 | 1.9 | 1.7 |  |  |  | #N/A
| 1 | 3 | 2.1 | 2.1 | 2.1 | 2.17 | 2.24 | 2.24 | 2.17 | 2.05 | 2.13 | 2.27 | 1.8 | 2.1 |  |  |  | 16
| 1 | 4 | 2.2 | 2.2 | 2.2 | 1.94 | 2.18 | 2.17 | 2.05 | 2.25 | 2.25 | 2.34 | 2 | 1.8 |  |  |  | 2.1647
| 1 | 5 | 2.4 | 2 | 2.1 | 2 | 1.92 | 2.15 | 2.29 | 1.96 | 2.41 | 2.26 | 1.8 | 1.7 |  |  |  | 74
| 1 | 6 | 2.2 | 2.4 | 2.2 | 2.07 | 2.05 | 2.06 | 1.95 | 1.97 | 2.44 | 2.2 | 1.7 | 1.5
| Mean |  | 2.1333 | 2.1833 | 2.15 | 2.0233 | 2.0733 | 2.16 | 2.2033 | 2.0617 | 2.3517 | 2.3067 | 1.8167 | 1.75
| Median |  | 2.15 | 2.15 | 2.15 | 2 | 2.04 | 2.16 | 2.23 | 2.025 | 2.385 | 2.305 | 1.8 | 1.7
| Std Dev. |  | 0.19664 | 0.1472 | 0.054772 | 0.084538 | 0.11656 | 0.063246 | 0.17862 | 0.11339 | 0.14077 | 0.075277 | 0.1169 | 0.19748
| Rel.Std.Dev. |  | 9.217% | 6.742% | 2.548% | 4.178% | 5.622% | 2.928% | 8.107% | 5.5% | 5.986% | 3.263% | 6.435% | 11.28%
| PDM3 |  | -1.447% | 0.8623% | -0.6775% | -6.529% | -4.219% | -0.2156% | 1.786% | -4.758% | 8.639% | 6.56% | -16.08% | -19.16%
| Z-Score (Absolute) |  | 0.08 | 0.38 | 0.08 | 1.08 | 0.62 | 0.17 | 0.56 | 0.73 | 1.91 | 1.5 | 2.95 | 3.56
| Analytical results for Zn in OREAS 25b (Certified Value 65 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 11 | 12 | 15 | 16 | [7 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 73 | 93.2 | 73.5205 | 60 | < 80 | 70 | 60 | 65 | [7 empty] | 1
| 1 | 2 | 63 | 96.8 | 69.4393 | 70 | < 80 | 70 | 45 | 60 | [7 empty] | 8
| 1 | 3 | 63 | 95.5 | 70.3172 | 70 | < 80 | 70 | 55 | 65 | [7 empty] | 16
| 1 | 4 | 93 | 94.2 | 73.361 | 70 | < 80 | 70 | 55 | 60 | [7 empty] | 65.405
| 1 | 5 | 71 | 93.9 | 75.4613 | 70 | < 80 | 70 | 50 | 60 | [7 empty] | 75
| 1 | 6 | 62 | 95.1 | 74.0811 | 70 | < 80 | 70 | 55 | 60
| Mean |  | 70.8333 | 94.7833 | 72.6967 | 68.3333 |  | 70 | 53.3333 | 61.6667
| Median |  | 67 | 94.65 | 73.4407 | 70 |  | 70 | 55 | 60
| Std Dev. |  | 11.8054 | 1.2891 | 2.3217 | 4.0825 |  | 0 | 5.164 | 2.582
| Rel.Std.Dev. |  | 16.67% | 1.36% | 3.194% | 5.974% |  | 0% | 9.682% | 4.187%
| PDM3 |  | 8.3% | 44.92% | 11.15% | 4.477% |  | 7.025% | -18.46% | -5.716%
| Z-Score (Absolute) |  | 0.2 | 3.13 | 0.43 | 0.1 | 3.55 | 0.1 | 1.94 | 0.92
| Analytical results for Zr in OREAS 25b (Certified Value 322 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 |  |  |  |  | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | BF*OES/MS | BF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 327 | 310 | 318 | 324 | 308.6 | 338 | 325 | 326 | 306 | 358 | 313.3 |  |  |  |  | 1
| 1 | 2 | 327 | 316 | 319 | 315 | 310.5 | 325 | 331 | 338 | 321 | 360 | 311.2 |  |  |  |  | #N/A
| 1 | 3 | 320 | 334 | 322 | 327 | 314.5 | 345 | 342 | 318 | 331 | 330 | 302.3 |  |  |  |  | 16
| 1 | 4 | 313 | 349 | 313 | 321 | 302.7 | 338 | 318 | 332 | 341 | 328 | 312.7 |  |  |  |  | 322.01
| 1 | 5 | 321 | 302 | 319 | 318 | 316.7 | 335 | 328 | 324 | 332 | 358 | 308.8 |  |  |  |  | 76
| 1 | 6 | 331 | 315 | 321 | 325 | 310.1 | 343 | 323 | 304 | 328 | 348 | 310.2
| Mean |  | 323.1667 | 321 | 318.6667 | 321.6667 | 310.5167 | 337.3333 | 327.8333 | 323.6667 | 326.5 | 347 | 309.75
| Median |  | 324 | 315.5 | 319 | 322.5 | 310.3 | 338 | 326.5 | 325 | 329.5 | 353 | 310.7
| Std Dev. |  | 6.4627 | 17.2974 | 3.1411 | 4.5461 | 4.8795 | 7.0616 | 8.2321 | 11.8265 | 11.9457 | 14.5739 | 4.0004
| Rel.Std.Dev. |  | 2% | 5.389% | 0.9857% | 1.413% | 1.571% | 2.093% | 2.511% | 3.654% | 3.659% | 4.2% | 1.291%
| PDM3 |  | 0.3592% | -0.3137% | -1.038% | -0.1066% | -3.569% | 4.759% | 1.808% | 0.5145% | 1.394% | 7.761% | -3.807%
| Z-Score (Absolute) |  | 0 | 0.32 | 0.67 | 0.22 | 1.9 | 2.12 | 0.7 | 0.07 | 0.5 | 3.57 | 2.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C813: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C831: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C867: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C904: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C922: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C940: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C958: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C976: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C994: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1013: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1031: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1068: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1086: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1104: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1122: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 25b (Indicative Value 0.062 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 3
3
| 1 | 1 | < 0.05 | 0.03 | < 0.5 | < 0.1 | < 0.3 | 0.046 | < 0.01 | 0.13 | 0.03 | < 0.01 | 0.1 | < 0.1 | < 0.02 | < 0.05 |  | 1
| 1 | 2 | < 0.05 | 0.04 | < 0.5 | < 0.1 | 0.4 | 0.04 | 0.04 | 0.09 | 0.07 | 0.03 | 0.11 | < 0.1 | 0.2 | < 0.05 |  | 13
| 1 | 3 | < 0.05 | 0.04 | < 0.5 | < 0.1 | < 0.3 | 0.054 | < 0.01 | 0.06 | 0.03 | 0.04 | 0.09 | < 0.1 | 0.24 | < 0.05 |  | 16
| 1 | 4 | < 0.05 | 0.04 | < 0.5 | < 0.1 | < 0.3 | 0.039 | 0.01 | 0.06 | 0.07 | 0.05 | 0.1 | < 0.1 | 0.2 | < 0.05 |  | 0.062019
| 1 | 5 | < 0.05 | 0.05 | < 0.5 | < 0.1 | < 0.3 | 0.044 | 0.02 | 0.08 | 0.03 | 0.05 | 0.09 | < 0.1 | 0.21 | < 0.05 |  | 25
| 1 | 6 | < 0.05 | 0.04 | 0.5 | < 0.1 | < 0.3 | 0.041 | 0.01 | 0.07 | 0.05 | 0.01 | 0.09 | 0.1 | 0.19 | < 0.05
| Mean |  |  | 0.04 | 0.5 |  | 0.4 | 0.044 | 0.02 | 0.081667 | 0.046667 | 0.036 | 0.096667 | 0.1 | 0.208
| Median |  |  | 0.04 | 0.5 |  | 0.4 | 0.0425 | 0.015 | 0.075 | 0.04 | 0.04 | 0.095 | 0.1 | 0.2
| Std Dev. |  |  | 0.0063246 |  |  |  | 0.0055498 | 0.014142 | 0.026394 | 0.019664 | 0.016733 | 0.008165 |  | 0.019235
| Rel.Std.Dev. |  |  | 15.81% |  |  |  | 12.61% | 70.71% | 32.32% | 42.14% | 46.48% | 8.447% |  | 9.248%
| PDM3 |  |  | -35.5% | 706.2% |  | 545% | -29.05% | -67.75% | 31.68% | -24.75% | -41.95% | 55.87% | 61.24% | 235.4%
| Z-Score (Absolute) |  | 0.67 | 0.24 | 7.03 | 0.05 | 4.14 | 0.13 | 0.96 | 0.96 | 0.05 | 0.51 | 1.4 | 0.29 | 3.66 | 0.67
| Analytical results for Al in OREAS 25b (Certified Value 8.6 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 8.5272 | 4.08 | 8.6 | 8.6696 | 9.11 | 8.59 | 8.81 | 8.76 | 8.56 | 8.56 | 7.49 | 8.54 | 8.15 | 9.85 | 7.89 | 1
| 1 | 2 | 8.6414 | 4.45 | 8.6 | 8.6341 | 9.02 | 8.57 | 8.98 | 8.69 | 8.42 | 8.79 | 7.58 | 8.62 | 8.46 | 9.87 | 6.48 | #N/A
| 1 | 3 | 8.439 | 5.13 | 8.6 | 8.7158 | 9.08 | 8.69 | 8.95 | 8.88 | 8.53 | 8.68 | 7.63 | 8.52 | 8.48 | 10.04 | 8.28 | 16
| 1 | 4 | 8.6004 | 4.92 | 8.6 | 8.6125 | 9.01 | 8.66 | 9.07 | 8.78 | 8.42 | 8.47 | 7.31 | 8.68 | 8.44 | 9.54 | 7.9 | 8.5964
| 1 | 5 | 8.6169 | 5.18 | 8.6 | 8.7233 | 8.87 | 8.52 | 9.06 | 8.71 | 8.5 | 8.62 | 7.43 | 8.54 | 8.23 | 9.79 | 7.72 | 78
| 1 | 6 | 8.4466 | 4.23 | 8.5 | 8.5778 | 8.89 | 8.62 | 9.1 | 8.43 | 8.33 | 8.39 | 7.47 | 8.62 | 8.43 | 8.81 | 7.98
| Mean |  | 8.5453 | 4.665 | 8.5833 | 8.6555 | 8.9967 | 8.6083 | 8.995 | 8.7083 | 8.46 | 8.585 | 7.485 | 8.5867 | 8.365 | 9.65 | 7.7083
| Median |  | 8.5638 | 4.685 | 8.6 | 8.6518 | 9.015 | 8.605 | 9.02 | 8.735 | 8.46 | 8.59 | 7.48 | 8.58 | 8.435 | 9.82 | 7.895
| Std Dev. |  | 0.088053 | 0.47416 | 0.040825 | 0.057917 | 0.097912 | 0.061779 | 0.1071 | 0.15171 | 0.085557 | 0.14433 | 0.11309 | 0.062823 | 0.13896 | 0.44222 | 0.62917
| Rel.Std.Dev. |  | 1.03% | 10.16% | 0.4756% | 0.6691% | 1.088% | 0.7177% | 1.191% | 1.742% | 1.011% | 1.681% | 1.511% | 0.7316% | 1.661% | 4.583% | 8.162%
| PDM3 |  | -0.5955% | -45.73% | -0.1525% | 0.6872% | 4.656% | 0.1384% | 4.636% | 1.302% | -1.587% | -0.1331% | -12.93% | -0.1137% | -2.692% | 12.26% | -10.33%
| Z-Score (Absolute) |  | 0.21 | 21.15 | 0.01 | 0.38 | 2.22 | 0.13 | 2.21 | 0.67 | 0.67 | 0 | 5.93 | 0.01 | 1.19 | 5.75 | 4.73
| Analytical results for As in OREAS 25b (Certified Value 14.5 ppm) | [16 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 15.3 | 26 | 12.7 | 13.8881 | 7 | 13.3 | 15.6 | 16 | 16.2 | 15.1 | 13.6 | 16 | 13.2 | 19 | 15 | 1
| 1 | 2 | 15.2 | 27 | 13 | 15.1074 | 6.5 | 13.8 | 16.2 | 15.3 | 15.3 | 14.2 | 13.6 | 16 | 14.3 | 18 | 13 | 1
| 1 | 3 | 15.2 | 30 | 13.2 | 13.8943 | 6.8 | 13.5 | 16.2 | 15.3 | 16.1 | 14.2 | 13.7 | 16 | 14 | 18 | 14 | 16
| 1 | 4 | 15.2 | 28 | 12.8 | 15.3477 | 6.7 | 13.9 | 15.6 | 15.3 | 15.7 | 14.2 | 13 | 16 | 13.9 | 19 | 12 | 14.5098
| 1 | 5 | 14.8 | 30 | 13.2 | 15.0336 | 6.9 | 13.6 | 15.2 | 15 | 16.4 | 15.4 | 13.4 | 16 | 14.7 | 19 | 12 | 79
| 1 | 6 | 15.3 | 25 | 13.1 | 14.0162 | 7.9 | 13.7 | 15.6 | 14.6 | 15.5 | 13.8 | 13.2 | 16 | 13.9 | 17 | 13
| Mean |  | 15.1667 | 27.6667 | 13 | 14.5478 | 6.9667 | 13.6333 | 15.7333 | 15.25 | 15.8667 | 14.4833 | 13.4167 | 16 | 14 | 18.3333 | 13.1667
| Median |  | 15.2 | 27.5 | 13.05 | 14.5249 | 6.85 | 13.65 | 15.6 | 15.3 | 15.9 | 14.2 | 13.5 | 16 | 13.95 | 18.5 | 13
| Std Dev. |  | 0.18619 | 2.0656 | 0.20976 | 0.68321 | 0.48854 | 0.21602 | 0.39328 | 0.45935 | 0.43205 | 0.62102 | 0.27142 | 0 | 0.498 | 0.8165 | 1.169
| Rel.Std.Dev. |  | 1.228% | 7.466% | 1.614% | 4.696% | 7.012% | 1.585% | 2.5% | 3.012% | 2.723% | 4.288% | 2.023% | 0% | 3.557% | 4.454% | 8.879%
| PDM3 |  | 4.527% | 90.68% | -10.41% | 0.2623% | -51.99% | -6.04% | 8.433% | 5.101% | 9.351% | -0.1823% | -7.534% | 10.27% | -3.513% | 26.35% | -9.257%
| Z-Score (Absolute) |  | 0.54 | NA | 1.18 | 0.05 | 5.96 | 0.67 | 0.99 | 0.61 | 1.1 | 0 | 0.85 | NA | 0.38 | NA | NA
Results from laboratories 2, 12, 15 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Ba in OREAS 25b (Certified Value 211 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 208.2 | 180 | 208 | 216.9776 | 241 | 196 | 220 | 220 | 220 | 210 | 222 | 203 | 213 | 345 | 200 | 1
| 1 | 2 | 214 | 199 | 209 | 208.8873 | 241 | 207 | 220 | 220 | 210 | 220 | 223 | 207 | 210 | 317 | 187 | #N/A
| 1 | 3 | 211.9 | 205 | 209 | 220.3297 | 242 | 209 | 220 | 230 | 210 | 210 | 228 | 203 | 213 | 361 | 198 | 16
| 1 | 4 | 209.9 | 195 | 210 | 210.3205 | 240 | 210 | 230 | 220 | 210 | 210 | 218 | 199 | 215 | 339 | 198 | 211.4159
| 1 | 5 | 205.8 | 205 | 207 | 206.4375 | 234 | 210 | 230 | 220 | 210 | 210 | 217 | 206 | 217 | 379 | 197 | 80
| 1 | 6 | 211.7 | 274 | 206 | 205.5904 | 242 | 214 | 230 | 220 | 210 | 210 | 223 | 207 | 211 | 317 | 199
| Mean |  | 210.25 | 209.6667 | 208.1667 | 211.4238 | 240 | 207.6667 | 225 | 221.6667 | 211.6667 | 211.6667 | 221.8333 | 204.1667 | 213.1667 | 343 | 196.5
| Median |  | 210.8 | 202 | 208.5 | 209.6039 | 241 | 209.5 | 225 | 220 | 210 | 210 | 222.5 | 204.5 | 213 | 342 | 198
| Std Dev. |  | 2.9331 | 32.837 | 1.472 | 5.9448 | 3.0332 | 6.1536 | 5.4772 | 4.0825 | 4.0825 | 4.0825 | 3.9707 | 3.1252 | 2.5626 | 24.4622 | 4.7645
| Rel.Std.Dev. |  | 1.395% | 15.66% | 0.7071% | 2.812% | 1.264% | 2.963% | 2.434% | 1.842% | 1.929% | 1.929% | 1.79% | 1.531% | 1.202% | 7.132% | 2.425%
| PDM3 |  | -0.5515% | -0.8274% | -1.537% | 0.003739% | 13.52% | -1.773% | 6.425% | 4.849% | 0.1186% | 0.1186% | 4.927% | -3.429% | 0.8281% | 62.24% | -7.055%
| Z-Score (Absolute) |  | 0.24 | 0.34 | 0.59 | 0.04 | 4.78 | 0.67 | 2.25 | 1.69 | 0 | 0 | 1.71 | 1.26 | 0.25 | 22.14 | 2.56
| Analytical results for Be in OREAS 25b (Certified Value 1.3 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 1.29 | 1.1 | < 5 | 1.3421 | 1.3 | 1 | 1.33 | 1.3 | 1.26 | 1.27 | 1.35 | 1.3 | 2 | 1.6 | 2 | 1
| 1 | 2 | 1.35 | 1.1 | < 5 | 1.4008 | 1.3 | 1 | 1.35 | 1.3 | 1.25 | 1.36 | 1.22 | 1.3 | 2 | 1.53 | 2 | #N/A
| 1 | 3 | 1.45 | 1.4 | < 5 | 1.4188 | 1.2 | 1 | 1.34 | 1.32 | 1.26 | 1.34 | 1.5 | 1.3 | 2 | 1.58 | 2 | 16
| 1 | 4 | 1.35 | 1.2 | < 5 | 1.3953 | 1.1 | 1 | 1.35 | 1.31 | 1.25 | 1.25 | 1.31 | 1.2 | 2 | 1.5 | 2.1 | 1.2972
| 1 | 5 | 1.39 | 1.2 | < 5 | 1.3323 | 1.2 | 1 | 1.35 | 1.29 | 1.27 | 1.27 | 1.34 | 1.3 | 1 | 1.68 | 2 | 81
| 1 | 6 | 1.4 | 1.3 | < 5 | 1.3541 | 1.2 | 1 | 1.36 | 1.26 | 1.25 | 1.24 | 1.21 | 1.3 | 2 | 1.51 | 2
| Mean |  | 1.3717 | 1.2167 |  | 1.3739 | 1.2167 | 1 | 1.3467 | 1.2967 | 1.2567 | 1.2883 | 1.3217 | 1.2833 | 1.8333 | 1.5667 | 2.0167
| Median |  | 1.37 | 1.2 |  | 1.3747 | 1.2 | 1 | 1.35 | 1.3 | 1.255 | 1.27 | 1.325 | 1.3 | 2 | 1.555 | 2
| Std Dev. |  | 0.054559 | 0.1169 |  | 0.035601 | 0.075277 | 0 | 0.010328 | 0.020656 | 0.008165 | 0.049565 | 0.10572 | 0.040825 | 0.40825 | 0.06802 | 0.040825
| Rel.Std.Dev. |  | 3.978% | 9.609% |  | 2.591% | 6.187% | 0% | 0.7669% | 1.593% | 0.6497% | 3.847% | 7.999% | 3.181% | 22.27% | 4.342% | 2.024%
| PDM3 |  | 5.739% | -6.21% |  | 5.911% | -6.21% | -22.91% | 3.811% | -0.04294% | -3.126% | -0.6853% | 1.884% | -1.071% | 41.33% | 20.77% | 55.46%
| Z-Score (Absolute) |  | 0.52 | 1.09 | 12.22 | 0.54 | 1.09 | 3.34 | 0.26 | 0.26 | 0.67 | 0.35 | 0 | 0.4 | 5.31 | 2.54 | 7.21
| Analytical results for Bi in OREAS 25b (Certified Value 0.4 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.41 | 0.36 | 0.37395 | 0.44 | 0.4 | 0.37 | 0.4 | 0.38 | 0.42 | 0.41 | 0.45 | 0.43 | 0.68 | 0.4 |  | 1
| 1 | 2 | 0.44 | 0.38 | 0.3709 | 0.44 | 0.41 | 0.37 | 0.37 | 0.4 | 0.39 | 0.4 | 0.42 | 0.42 | 0.64 | 0.3 |  | 2
| 1 | 3 | 0.42 | 0.39 | 0.3676 | 0.43 | 0.42 | 0.37 | 0.38 | 0.38 | 0.42 | 0.4 | 0.44 | 0.44 | 0.66 | 0.3 |  | 16
| 1 | 4 | 0.42 | 0.36 | 0.38034 | 0.42 | 0.43 | 0.35 | 0.39 | 0.38 | 0.38 | 0.4 | 0.43 | 0.43 | 0.61 | 0.3 |  | 0.40099
| 1 | 5 | 0.4 | 0.37 | 0.38015 | 0.43 | 0.41 | 0.36 | 0.37 | 0.38 | 0.42 | 0.4 | 0.41 | 0.44 | 0.66 | 0.3 |  | 82
| 1 | 6 | 0.42 | 0.4 | 0.36839 | 0.44 | 0.41 | 0.38 | 0.37 | 0.38 | 0.4 | 0.39 | 0.44 | 0.42 | 0.61 | 0.3
| Mean |  | 0.41833 | 0.37667 | 0.37356 | 0.43333 | 0.41333 | 0.36667 | 0.38 | 0.38333 | 0.405 | 0.4 | 0.43167 | 0.43 | 0.64333 | 0.31667
| Median |  | 0.42 | 0.375 | 0.37243 | 0.435 | 0.41 | 0.37 | 0.375 | 0.38 | 0.41 | 0.4 | 0.435 | 0.43 | 0.65 | 0.3
| Std Dev. |  | 0.013292 | 0.01633 | 0.0056387 | 0.008165 | 0.010328 | 0.010328 | 0.012649 | 0.008165 | 0.017607 | 0.0063246 | 0.01472 | 0.0089443 | 0.028752 | 0.040825
| Rel.Std.Dev. |  | 3.177% | 4.335% | 1.509% | 1.884% | 2.499% | 2.817% | 3.329% | 2.13% | 4.347% | 1.581% | 3.41% | 2.08% | 4.469% | 12.89%
| PDM3 |  | 4.325% | -6.066% | -6.842% | 8.066% | 3.078% | -8.56% | -5.235% | -4.403% | 0.9998% | -0.2471% | 7.65% | 7.234% | 60.44% | -21.03%
| Z-Score (Absolute) |  | 0.36 | 0.76 | 0.85 | 0.76 | 0.22 | 1.03 | 0.67 | 0.58 | 0 | 0.13 | 0.72 | 0.67 | 6.43 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 25b (Certified Value 0.255 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.2465 | 0.15 | 0.3 | 0.26073 | 0.25 | 0.25 | 0.27 | 0.26 | 0.26 | 0.25 | 0.27 | 0.25 | 0.28 | 0.263 | 0.226 | 1
| 1 | 2 | 0.2524 | 0.18 | 0.3 | 0.22331 | 0.25 | 0.25 | 0.27 | 0.26 | 0.25 | 0.26 | 0.27 | 0.248 | 0.26 | 0.255 | 0.176 | #N/A
| 1 | 3 | 0.2486 | 0.2 | 0.3 | 0.20457 | 0.25 | 0.25 | 0.27 | 0.26 | 0.26 | 0.25 | 0.27 | 0.245 | 0.26 | 0.266 | 0.239 | 16
| 1 | 4 | 0.2523 | 0.18 | 0.2 | 0.25248 | 0.25 | 0.25 | 0.27 | 0.26 | 0.25 | 0.25 | 0.27 | 0.245 | 0.26 | 0.254 | 0.233 | 0.25481
| 1 | 5 | 0.2469 | 0.19 | 0.3 | 0.2759 | 0.24 | 0.25 | 0.27 | 0.26 | 0.26 | 0.26 | 0.27 | 0.247 | 0.28 | 0.274 | 0.223 | 83
| 1 | 6 | 0.2455 | 0.18 | 0.2 | 0.22866 | 0.25 | 0.25 | 0.27 | 0.25 | 0.25 | 0.25 | 0.27 | 0.245 | 0.27 | 0.248 | 0.234
| Mean |  | 0.2487 | 0.18 | 0.26667 | 0.24094 | 0.24833 | 0.25 | 0.27 | 0.25833 | 0.255 | 0.25333 | 0.27 | 0.24667 | 0.26833 | 0.26 | 0.22183
| Median |  | 0.24775 | 0.18 | 0.3 | 0.24057 | 0.25 | 0.25 | 0.27 | 0.26 | 0.255 | 0.25 | 0.27 | 0.246 | 0.265 | 0.259 | 0.2295
| Std Dev. |  | 0.0029993 | 0.016733 | 0.05164 | 0.026576 | 0.0040825 | 0 | 0 | 0.0040825 | 0.0054772 | 0.005164 | 0 | 0.0020656 | 0.0098319 | 0.0094446 | 0.023181
| Rel.Std.Dev. |  | 1.206% | 9.296% | 19.36% | 11.03% | 1.644% | 0% | 0% | 1.58% | 2.148% | 2.038% | 0% | 0.8374% | 3.664% | 3.633% | 10.45%
| PDM3 |  | -2.397% | -29.36% | 4.654% | -5.442% | -2.541% | -1.887% | 5.962% | 1.384% | 0.07548% | -0.5786% | 5.962% | -3.195% | 5.308% | 2.038% | -12.94%
| Z-Score (Absolute) |  | 0.47 | 7.42 | 1.35 | 1.25 | 0.51 | 0.34 | 1.69 | 0.51 | 0.17 | 0 | 1.69 | 0.67 | 1.52 | 0.67 | 3.19
| Analytical results for Cd in OREAS 25b (Certified Value 0.044 ppm) | [16 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 3
3
| 1 | 1 | 0.05 | 0.08 | 0.98 | < 0.05 | < 0.3 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.06 | < 0.05 | 0.24 | 0.03 | 0.1 | 1
| 1 | 2 | 0.05 | 0.06 | 0.94 | 0.052581 | < 0.3 | 0.03 | 0.02 | 0.04 | 0.06 | 0.02 | 0.06 | 0.05 | 0.13 | 0.04 | 0.1 | 17
| 1 | 3 | 0.04 | 0.07 | 0.96 | 0.056569 | < 0.3 | 0.04 | 0.02 | 0.02 | 0.02 | 0.06 | 0.07 | 0.08 | 0.12 | 0.05 | 0.1 | 16
| 1 | 4 | 0.04 | 0.06 | 0.98 | 0.05568 | < 0.3 | 0.05 | 0.04 | 0.04 | 0.03 | 0.03 | 0.07 | < 0.05 | 0.12 | 0.05 | 0.1 | 0.044389
| 1 | 5 | 0.04 | 0.06 | 0.96 | < 0.05 | < 0.3 | 0.03 | 0.03 | 0.04 | 0.03 | 0.05 | 0.07 | 0.05 | 0.23 | 0.04 | 0.1 | 84
| 1 | 6 | 0.05 | 0.06 | 1.01 | < 0.05 | < 0.3 | 0.03 | 0.02 | 0.02 | 0.04 | 0.05 | 0.06 | 0.05 | 0.13 | 0.05 | 0.1
| Mean |  | 0.045 | 0.065 | 0.97167 | 0.054943 |  | 0.035 | 0.026667 | 0.033333 | 0.035 | 0.04 | 0.065 | 0.0575 | 0.16167 | 0.043333 | 0.1
| Median |  | 0.045 | 0.06 | 0.97 | 0.05568 |  | 0.03 | 0.025 | 0.04 | 0.03 | 0.04 | 0.065 | 0.05 | 0.13 | 0.045 | 0.1
| Std Dev. |  | 0.0054772 | 0.0083666 | 0.024014 | 0.0020934 |  | 0.0083666 | 0.008165 | 0.010328 | 0.013784 | 0.015492 | 0.0054772 | 0.015 | 0.057067 | 0.008165 | 1.5202e-17
| Rel.Std.Dev. |  | 12.17% | 12.87% | 2.471% | 3.81% |  | 23.9% | 30.62% | 30.98% | 39.38% | 38.73% | 8.427% | 26.09% | 35.3% | 18.84% | 1.52e-14%
| PDM3 |  | 1.377% | 46.43% | 2089% | 23.78% |  | -21.15% | -39.92% | -24.91% | -21.15% | -9.887% | 46.43% | 29.54% | 264.2% | -2.378% | 125.3%
| Z-Score (Absolute) |  | 0.06 | 1.4 | 62.54 | 0.28 | 7.14 | 0.62 | 1.18 | 0.73 | 0.62 | 0.28 | 1.4 | 0.17 | 7.92 | 0.06 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 25b (Certified Value 53 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 53.58 | 39.63 | 59.2 | 47.22 | 53.5 | 59.7 | 57.8 | 55.2 | 58.3 | 53.42 | 44.45 | 59 |  |  |  | 1
| 1 | 2 | 53.98 | 42.7 | 60.6 | 50.66 | 54.1 | 60.6 | 56.5 | 52.5 | 56.8 | 46.32 | 52.74 | 39.4 |  |  |  | 3
| 1 | 3 | 55.07 | 46.09 | 59.3 | 50.06 | 54.6 | 59.8 | 56.4 | 52.6 | 58.5 | 48.48 | 60.33 | 60.1 |  |  |  | 16
| 1 | 4 | 54.56 | 44.59 | 58.7 | 50.7 | 53.4 | 60.9 | 56.9 | 54.9 | 56.9 | 46.12 | 44.22 | 48.4 |  |  |  | 53.4083
| 1 | 5 | 52.86 | 45.66 | 58.9 | 49.51 | 53.6 | 60.1 | 59 | 55.3 | 55.5 | 52.25 | 64.29 | 50.6 |  |  |  | 85
| 1 | 6 | 54.75 | 41.71 | 60 | 49.66 | 54.1 | 57.7 | 55.5 | 55.4 | 56.7 | 48.35 | 43.22 | 45.8
| Mean |  | 54.1333 | 43.3967 | 59.45 | 49.635 | 53.8833 | 59.8 | 57.0167 | 54.3167 | 57.1167 | 49.1567 | 51.5417 | 50.55
| Median |  | 54.27 | 43.645 | 59.25 | 49.86 | 53.85 | 59.95 | 56.7 | 55.05 | 56.85 | 48.415 | 48.595 | 49.5
| Std Dev. |  | 0.8235 | 2.5027 | 0.71764 | 1.2817 | 0.46224 | 1.1278 | 1.2254 | 1.379 | 1.1179 | 3.0371 | 9.1032 | 7.9261
| Rel.Std.Dev. |  | 1.521% | 5.767% | 1.207% | 2.582% | 0.8579% | 1.886% | 2.149% | 2.539% | 1.957% | 6.179% | 17.66% | 15.68%
| PDM3 |  | 1.358% | -18.75% | 11.31% | -7.065% | 0.8894% | 11.97% | 6.756% | 1.701% | 6.943% | -7.961% | -3.495% | -5.352%
| Z-Score (Absolute) |  | 0.03 | 2.18 | 1.12 | 0.9 | 0.03 | 1.19 | 0.62 | 0.06 | 0.64 | 1 | 0.51 | 0.71
| Analytical results for Co in OREAS 25b (Certified Value 9.29 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 10.1 | 9.7 | 21.3 | 9.4976 | 9 | 9.4 | 9 | 9.9 | 8.9 | 8.9 | 8.93 | 9 | 9.2 | 9.7 | 9 | 1
| 1 | 2 | 10 | 9.6 | 21.2 | 9.4955 | 9 | 9.7 | 8.9 | 9.5 | 8.8 | 9 | 8.88 | 9 | 9.1 | 9.4 | 9 | 5
| 1 | 3 | 9.9 | 9.9 | 21.3 | 9.3794 | 9 | 9.8 | 9 | 9.6 | 9 | 8.9 | 9.19 | 9 | 9.6 | 9.8 | 6 | 16
| 1 | 4 | 9.8 | 9.5 | 21.2 | 8.8675 | 9 | 9.6 | 8.8 | 9.6 | 8.8 | 8.9 | 8.69 | 9 | 9.3 | 9.4 | 9 | 9.2912
| 1 | 5 | 9.8 | 9.7 | 21.1 | 9.1269 | 9 | 9.5 | 8.7 | 9.3 | 9 | 8.9 | 8.67 | 9 | 9.2 | 10.4 | 8 | 86
| 1 | 6 | 9.7 | 9.1 | 20.9 | 8.8719 | 9 | 9.5 | 8.9 | 9.2 | 8.6 | 8.8 | 8.42 | 9 | 9.4 | 9.5 | 9
| Mean |  | 9.8833 | 9.5833 | 21.1667 | 9.2065 | 9 | 9.5833 | 8.8833 | 9.5167 | 8.85 | 8.9 | 8.7967 | 9 | 9.3 | 9.7 | 8.3333
| Median |  | 9.85 | 9.65 | 21.2 | 9.2531 | 9 | 9.55 | 8.9 | 9.55 | 8.85 | 8.9 | 8.785 | 9 | 9.25 | 9.6 | 9
| Std Dev. |  | 0.1472 | 0.27142 | 0.15055 | 0.29373 | 0 | 0.1472 | 0.1169 | 0.24833 | 0.15166 | 0.063246 | 0.2641 | 0 | 0.17889 | 0.37947 | 1.2111
| Rel.Std.Dev. |  | 1.489% | 2.832% | 0.7113% | 3.19% | 0% | 1.536% | 1.316% | 2.609% | 1.714% | 0.7106% | 3.002% | 0% | 1.923% | 3.912% | 14.53%
| PDM3 |  | 6.373% | 3.144% | 127.8% | -0.9118% | -3.134% | 3.144% | -4.39% | 2.427% | -4.749% | -4.21% | -5.323% | -3.134% | 0.09477% | 4.4% | -10.31%
| Z-Score (Absolute) |  | 0.84 | 0.31 | 20.68 | 0.36 | NA | 0.31 | 0.92 | 0.19 | 0.98 | 0.89 | 1.08 | NA | 0.19 | 0.51 | NA
Results from laboratories 5, 12 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 25b (Certified Value 204 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 227 | 178 | 206.8 | 216.6259 | 154 | 203 | 203 | 197 | 200 | 192 | 203 | 187 | 205 | 212 | 200 | 1
| 1 | 2 | 229 | 176 | 207.7 | 219.4178 | 180 | 202 | 209 | 199 | 198 | 195 | 204 | 184 | 211 | 203 | 175 | #N/A
| 1 | 3 | 223 | 170 | 206.4 | 205.0494 | 157 | 204 | 206 | 200 | 200 | 189 | 204 | 187 | 214 | 214 | 165 | 16
| 1 | 4 | 229 | 154 | 206 | 207.2517 | 179 | 206 | 207 | 195 | 202 | 194 | 200 | 189 | 206 | 201 | 160 | 204.0091
| 1 | 5 | 223 | 161 | 205.7 | 202.3306 | 170 | 203 | 210 | 200 | 199 | 194 | 199 | 195 | 200 | 215 | 160 | 87
| 1 | 6 | 220 | 161 | 204.1 | 204.2795 | 172 | 202 | 208 | 191 | 196 | 188 | 202 | 197 | 224 | 198 | 165
| Mean |  | 225.1667 | 166.6667 | 206.1167 | 209.1591 | 168.6667 | 203.3333 | 207.1667 | 197 | 199.1667 | 192 | 202 | 189.8333 | 210 | 207.1667 | 170.8333
| Median |  | 225 | 165.5 | 206.2 | 206.1505 | 171 | 203 | 207.5 | 198 | 199.5 | 193 | 202.5 | 188 | 208.5 | 207.5 | 165
| Std Dev. |  | 3.7103 | 9.5009 | 1.209 | 7.0988 | 10.9484 | 1.5055 | 2.4833 | 3.5214 | 2.0412 | 2.8983 | 2.0976 | 5.0761 | 8.4143 | 7.3598 | 15.3025
| Rel.Std.Dev. |  | 1.648% | 5.701% | 0.5866% | 3.394% | 6.491% | 0.7404% | 1.199% | 1.787% | 1.025% | 1.51% | 1.038% | 2.674% | 4.007% | 3.553% | 8.958%
| PDM3 |  | 10.37% | -18.3% | 1.033% | 2.524% | -17.32% | -0.3312% | 1.548% | -3.436% | -2.374% | -5.887% | -0.9848% | -6.949% | 2.937% | 1.548% | -16.26%
| Z-Score (Absolute) |  | 2.18 | 3.33 | 0.39 | 0.67 | 3.14 | 0.13 | 0.49 | 0.47 | 0.27 | 0.94 | 0 | 1.15 | 0.75 | 0.49 | 2.94
| Analytical results for Cs in OREAS 25b (Certified Value 4.61 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 4.84 | < 5 | 4.6293 | 5.11 | 4.4 | 4.57 | 4.96 | 4.4 | 4.57 | 4.89 | 4.7 | 4.7 | 4.72 | 4.98 |  | 1
| 1 | 2 | 4.87 | < 5 | 3.7966 | 5.06 | 4.5 | 4.58 | 4.95 | 4.31 | 4.34 | 4.91 | 4.8 | 4.1 | 4.77 | 4.19 |  | #N/A
| 1 | 3 | 4.83 | < 5 | 3.5925 | 5.08 | 4.5 | 4.55 | 4.94 | 4.36 | 4.3 | 5 | 4.7 | 4.2 | 4.94 | 4.72 |  | 16
| 1 | 4 | 4.83 | < 5 | 4.4071 | 5.01 | 4.6 | 4.51 | 5.02 | 4.24 | 4.49 | 4.79 | 4.6 | 4 | 4.58 | 4.27 |  | 4.6147
| 1 | 5 | 4.76 | < 5 | 4.5684 | 5.07 | 4.5 | 4.53 | 4.82 | 4.36 | 4.49 | 4.8 | 4.6 | 4.6 | 5.12 | 4.33 |  | 88
| 1 | 6 | 4.69 | < 5 | 4.0687 | 5.11 | 4.5 | 4.64 | 4.8 | 4.27 | 4.36 | 4.9 | 4.7 | 4.1 | 4.77 | 4.08
| Mean |  | 4.8033 |  | 4.1771 | 5.0733 | 4.5 | 4.5633 | 4.915 | 4.3233 | 4.425 | 4.8817 | 4.6833 | 4.2833 | 4.8167 | 4.4283
| Median |  | 4.83 |  | 4.2379 | 5.075 | 4.5 | 4.56 | 4.945 | 4.335 | 4.425 | 4.895 | 4.7 | 4.15 | 4.77 | 4.3
| Std Dev. |  | 0.066232 |  | 0.42634 | 0.037238 | 0.063246 | 0.045461 | 0.086197 | 0.060882 | 0.10635 | 0.077825 | 0.075277 | 0.29269 | 0.18811 | 0.34707
| Rel.Std.Dev. |  | 1.379% |  | 10.21% | 0.734% | 1.405% | 0.9962% | 1.754% | 1.408% | 2.403% | 1.594% | 1.607% | 6.833% | 3.905% | 7.837%
| PDM3 |  | 4.088% |  | -9.482% | 9.939% | -2.485% | -1.112% | 6.508% | -6.313% | -4.11% | 5.786% | 1.488% | -7.18% | 4.377% | -4.038%
| Z-Score (Absolute) |  | 0.7 | 5.27 | 0.92 | 1.4 | 0.08 | 0.08 | 0.99 | 0.54 | 0.28 | 0.91 | 0.39 | 0.64 | 0.74 | 0.27
| Analytical results for Cu in OREAS 25b (Certified Value 38.3 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 41.3 | 39.1 | 34.1 | 33.6778 | 41 | 37 | 36.2 | 37.7 | 38.3 | 36.4 | 38 | 33 | 38.3 | 40.4 | 50 | 1
| 1 | 2 | 45.8 | 39.6 | 34.2 | 34.8292 | 41 | 38.7 | 36.2 | 35.2 | 38.3 | 37.5 | 45.7 | 33 | 39.9 | 39.6 | 42 | 6
| 1 | 3 | 40.3 | 46 | 34.3 | 33.9997 | 42 | 38.8 | 35.8 | 35.7 | 39.7 | 36.1 | 48.4 | 33 | 38.3 | 41.3 | 41 | 16
| 1 | 4 | 38.8 | 40.6 | 34.2 | 37.6397 | 40 | 37.8 | 35.6 | 35.8 | 38 | 35.8 | 55.7 | 33 | 38.8 | 53 | 47 | 38.3398
| 1 | 5 | 38.9 | 45.1 | 33.8 | 38.4086 | 41 | 41.9 | 35.9 | 34.1 | 38.2 | 35.9 | 51.9 | 32 | 38.7 | 40.9 | 40 | 89
| 1 | 6 | 39.9 | 48.6 | 34 | 34.673 | 41 | 39 | 36.9 | 34.3 | 36.8 | 35 | 37 | 33 | 38.5 | 38.3 | 40
| Mean |  | 40.8333 | 43.1667 | 34.1 | 35.538 | 41 | 38.8667 | 36.1 | 35.4667 | 38.2167 | 36.1167 | 46.1167 | 32.8333 | 38.75 | 42.25 | 43.3333
| Median |  | 40.1 | 42.85 | 34.15 | 34.7511 | 41 | 38.75 | 36.05 | 35.45 | 38.25 | 36 | 47.05 | 33 | 38.6 | 40.65 | 41.5
| Std Dev. |  | 2.6044 | 3.9277 | 0.17889 | 1.9867 | 0.63246 | 1.6657 | 0.45607 | 1.3003 | 0.92394 | 0.82321 | 7.4789 | 0.40825 | 0.59917 | 5.3728 | 4.1793
| Rel.Std.Dev. |  | 6.378% | 9.099% | 0.5246% | 5.59% | 1.543% | 4.286% | 1.263% | 3.666% | 2.418% | 2.279% | 16.22% | 1.243% | 1.546% | 12.72% | 9.645%
| PDM3 |  | 6.504% | 12.59% | -11.06% | -7.308% | 6.939% | 1.374% | -5.842% | -7.494% | -0.321% | -5.798% | 20.28% | -14.36% | 1.07% | 10.2% | 13.02%
| Z-Score (Absolute) |  | 0.44 | 0.93 | 0.98 | 0.67 | 0.47 | 0.02 | 0.56 | 0.69 | 0.11 | 0.55 | 1.55 | 1.24 | 0 | 0.73 | 0.96
| Analytical results for Dy in OREAS 25b (Certified Value 2.47 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 2.66 | 2.9 | 2.2 | 2.8 | 2.16 | 2.7 | [9 empty] | 1
| 1 | 2 | 2.69 | 3 | 2.3 | 2.3 | 2.21 | 1.8 | [9 empty] | 7
| 1 | 3 | 2.67 | 2.9 | 2.3 | 2.4 | 2.26 | 2.6 | [9 empty] | 16
| 1 | 4 | 2.64 | 2.9 | 2.3 | 2.3 | 2.15 | 2.2 | [9 empty] | 2.4689
| 1 | 5 | 2.71 | 2.8 | 2.3 | 2.5 | 2.28 | 2.3 | [9 empty] | 90
| 1 | 6 | 2.64 | 2.8 | 2.3 | 2.6 | 2.21 | 2.1
| Mean |  | 2.6683 | 2.8833 | 2.2833 | 2.4833 | 2.2117 | 2.2833
| Median |  | 2.665 | 2.9 | 2.3 | 2.45 | 2.21 | 2.25
| Std Dev. |  | 0.027869 | 0.075277 | 0.040825 | 0.19408 | 0.051929 | 0.33116
| Rel.Std.Dev. |  | 1.044% | 2.611% | 1.788% | 7.815% | 2.348% | 14.5%
| PDM3 |  | 8.078% | 16.79% | -7.516% | 0.5851% | -10.42% | -7.516%
| Z-Score (Absolute) |  | 1.41 | 2.48 | 0.5 | 0.5 | 0.85 | 0.5
| Analytical results for Er in OREAS 25b (Certified Value 1.29 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 1.36 | 1.7 | 1.1 | 1.4 | 1.06 | 1.4 | [9 empty] | 1
| 1 | 2 | 1.33 | 1.6 | 1.1 | 1.3 | 1.11 | 1 | [9 empty] | 23
| 1 | 3 | 1.49 | 1.6 | 1.2 | 1.2 | 1.08 | 1.4 | [9 empty] | 16
| 1 | 4 | 1.4 | 1.6 | 1.2 | 1.2 | 1.16 | 1.2 | [9 empty] | 1.2906
| 1 | 5 | 1.43 | 1.6 | 1.2 | 1.2 | 1.12 | 1.2 | [9 empty] | 91
| 1 | 6 | 1.42 | 1.6 | 1.2 | 1.2 | 1 | 1.1
| Mean |  | 1.405 | 1.6167 | 1.1667 | 1.25 | 1.0883 | 1.2167
| Median |  | 1.41 | 1.6 | 1.2 | 1.2 | 1.095 | 1.2
| Std Dev. |  | 0.056125 | 0.040825 | 0.05164 | 0.083666 | 0.055287 | 0.16021
| Rel.Std.Dev. |  | 3.995% | 2.525% | 4.426% | 6.693% | 5.08% | 13.17%
| PDM3 |  | 8.868% | 25.27% | -9.6% | -3.142% | -15.67% | -5.725%
| Z-Score (Absolute) |  | 1.09 | 2.44 | 0.42 | 0.11 | 0.92 | 0.11
| Analytical results for Eu in OREAS 25b (Certified Value 0.83 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 0.88 | 0.89 | 0.7 | 0.9 | 0.77 | 0.95 | [9 empty] | 1
| 1 | 2 | 0.87 | 0.94 | 0.8 | 0.9 | 0.78 | 0.62 | [9 empty] | 8
| 1 | 3 | 0.86 | 0.91 | 0.8 | 0.9 | 0.79 | 0.92 | [9 empty] | 16
| 1 | 4 | 0.85 | 0.88 | 0.8 | 0.9 | 0.73 | 0.77 | [9 empty] | 0.83389
| 1 | 5 | 0.86 | 0.88 | 0.8 | 0.8 | 0.85 | 0.8 | [9 empty] | 92
| 1 | 6 | 0.87 | 0.87 | 0.8 | 0.9 | 0.74 | 0.74
| Mean |  | 0.865 | 0.895 | 0.78333 | 0.88333 | 0.77667 | 0.8
| Median |  | 0.865 | 0.885 | 0.8 | 0.9 | 0.775 | 0.785
| Std Dev. |  | 0.010488 | 0.025884 | 0.040825 | 0.040825 | 0.04274 | 0.12149
| Rel.Std.Dev. |  | 1.212% | 2.892% | 5.212% | 4.622% | 5.503% | 15.19%
| PDM3 |  | 3.731% | 7.328% | -6.063% | 5.929% | -6.862% | -4.064%
| Z-Score (Absolute) |  | 0.44 | 0.84 | 0.66 | 0.69 | 0.75 | 0.44
| Analytical results for Fe in OREAS 25b (Certified Value 8.19 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 8.41 | 7.55 | 8.14 | 8.6616 | 8.68 | 8.09 | 8.32 | 8.38 | 8.11 | 8.07 | 6.02 | 8.12 | 7.81 | 8.54 | 7.81 | 1
| 1 | 2 | 8.31 | 7.76 | 8.1 | 8.4146 | 8.7 | 8.12 | 8.48 | 8.32 | 7.98 | 8.09 | 6.08 | 8.11 | 7.91 | 8.4 | 7.62 | #N/A
| 1 | 3 | 8.28 | 8.05 | 8.1 | 8.6839 | 8.8 | 8.16 | 8.46 | 8.47 | 8.04 | 8.02 | 6.07 | 8.08 | 7.95 | 8.48 | 7.7 | 16
| 1 | 4 | 8.21 | 7.58 | 8.07 | 8.712 | 8.92 | 8.11 | 8.56 | 8.35 | 8 | 7.98 | 6.05 | 8.02 | 8.01 | 8.1 | 7.73 | 8.1855
| 1 | 5 | 8.31 | 7.88 | 8.08 | 8.4193 | 8.79 | 8.04 | 8.56 | 8.37 | 8.06 | 8.11 | 6.01 | 8.15 | 7.93 | 8.61 | 7.78 | 93
| 1 | 6 | 8.29 | 7.56 | 8.14 | 8.64 | 8.81 | 8.09 | 8.58 | 8.11 | 7.89 | 7.89 | 5.93 | 8.14 | 7.95 | 8.12 | 7.65
| Mean |  | 8.3017 | 7.73 | 8.105 | 8.5886 | 8.7833 | 8.1017 | 8.4933 | 8.3333 | 8.0133 | 8.0267 | 6.0267 | 8.1033 | 7.9267 | 8.375 | 7.715
| Median |  | 8.3 | 7.67 | 8.1 | 8.6508 | 8.795 | 8.1 | 8.52 | 8.36 | 8.02 | 8.045 | 6.035 | 8.115 | 7.94 | 8.44 | 7.715
| Std Dev. |  | 0.064627 | 0.20474 | 0.029496 | 0.13507 | 0.08641 | 0.039707 | 0.097707 | 0.12044 | 0.075807 | 0.082138 | 0.05465 | 0.04761 | 0.066232 | 0.21668 | 0.073417
| Rel.Std.Dev. |  | 0.7785% | 2.649% | 0.3639% | 1.573% | 0.9838% | 0.4901% | 1.15% | 1.445% | 0.946% | 1.023% | 0.9068% | 0.5875% | 0.8356% | 2.587% | 0.9516%
| PDM3 |  | 1.419% | -5.565% | -0.9834% | 4.924% | 7.304% | -1.024% | 3.761% | 1.806% | -2.103% | -1.94% | -26.37% | -1.004% | -3.162% | 2.315% | -5.748%
| Z-Score (Absolute) |  | 0.58 | 1.09 | 0 | 1.42 | 1.99 | 0 | 1.14 | 0.67 | 0.26 | 0.22 | 6.09 | 0 | 0.52 | 0.8 | 1.14
| Analytical results for Ga in OREAS 25b (Certified Value 25.3 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 25.21 | 24.8 | 29.6 | 26.0848 | 26.6 | 24.56 | 24.5 | 26.5 | 25.1 | 24.1 | 25.6 | 26.4 | 25.63 | 26.21 | 23.6 | 1
| 1 | 2 | 25.87 | 25.1 | 29.2 | 25.6708 | 27.1 | 24.93 | 24.4 | 25.8 | 24.7 | 24.7 | 25.2 | 26.4 | 24.41 | 26.17 | 23.3 | #N/A
| 1 | 3 | 25.21 | 26.1 | 29.4 | 25.5623 | 27.4 | 25.53 | 24.3 | 26.4 | 25.1 | 24.6 | 25.8 | 26.4 | 26.29 | 27.17 | 23.2 | 16
| 1 | 4 | 25.29 | 24.6 | 29.1 | 26.0281 | 27.5 | 25.27 | 23.9 | 26.1 | 24.5 | 24.1 | 24.5 | 27 | 25.46 | 26.08 | 22.3 | 25.2889
| 1 | 5 | 24.81 | 25.2 | 29.1 | 25.4631 | 27.5 | 25.07 | 24.2 | 25.6 | 25.3 | 24 | 24.8 | 26.8 | 25.07 | 29.36 | 22.9 | 94
| 1 | 6 | 25.27 | 24 | 29 | 25.8054 | 27.5 | 24.71 | 24.3 | 25.4 | 24.3 | 23.1 | 24.9 | 26.3 | 24.93 | 26.15 | 23.4
| Mean |  | 25.2767 | 24.9667 | 29.2333 | 25.7691 | 27.2667 | 25.0117 | 24.2667 | 25.9667 | 24.8333 | 24.1 | 25.1333 | 26.55 | 25.2983 | 26.8567 | 23.1167
| Median |  | 25.24 | 24.95 | 29.15 | 25.7381 | 27.45 | 25 | 24.3 | 25.95 | 24.9 | 24.1 | 25.05 | 26.4 | 25.265 | 26.19 | 23.25
| Std Dev. |  | 0.34027 | 0.70048 | 0.22509 | 0.25067 | 0.36148 | 0.35791 | 0.20656 | 0.44121 | 0.39328 | 0.56921 | 0.49666 | 0.28107 | 0.64753 | 1.2928 | 0.46224
| Rel.Std.Dev. |  | 1.346% | 2.806% | 0.77% | 0.9727% | 1.326% | 1.431% | 0.8512% | 1.699% | 1.584% | 2.362% | 1.976% | 1.059% | 2.56% | 4.814% | 2%
| PDM3 |  | -0.04851% | -1.274% | 15.6% | 1.899% | 7.821% | -1.096% | -4.042% | 2.68% | -1.802% | -4.701% | -0.6153% | 4.987% | 0.03717% | 6.199% | -8.59%
| Z-Score (Absolute) |  | 0 | 0.3 | 3.87 | 0.48 | 1.94 | 0.26 | 0.99 | 0.67 | 0.43 | 1.15 | 0.14 | 1.24 | 0.02 | 1.54 | 2.11
| Analytical results for Gd in OREAS 25b (Certified Value 3.16 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 3.19 | 3.5 | 2.8 | 3.3 | 2.92 | 3.6 | [9 empty] | 1
| 1 | 2 | 3.2 | 3.7 | 3 | 3.3 | 2.89 | 2.4 | [9 empty] | 10
| 1 | 3 | 3.2 | 3.5 | 3 | 3.4 | 2.98 | 3.5 | [9 empty] | 16
| 1 | 4 | 3.18 | 3.5 | 3 | 3.2 | 2.78 | 3 | [9 empty] | 3.1619
| 1 | 5 | 3.18 | 3.4 | 3.1 | 3 | 3.1 | 3.1 | [9 empty] | 95
| 1 | 6 | 3.2 | 3.5 | 2.9 | 3.3 | 2.81 | 2.9
| Mean |  | 3.1917 | 3.5167 | 2.9667 | 3.25 | 2.9133 | 3.0833
| Median |  | 3.195 | 3.5 | 3 | 3.3 | 2.905 | 3.05
| Std Dev. |  | 0.0098319 | 0.098319 | 0.10328 | 0.13784 | 0.1169 | 0.43551
| Rel.Std.Dev. |  | 0.308% | 2.796% | 3.481% | 4.241% | 4.013% | 14.12%
| PDM3 |  | 0.94% | 11.22% | -6.176% | 2.785% | -7.863% | -2.486%
| Z-Score (Absolute) |  | 0.26 | 1.8 | 0.81 | 0.54 | 1.07 | 0.26
| Analytical results for Ge in OREAS 25b (Indicative Value 0.19 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 14 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 1.8 | 0.2 | 1.2335 | < 0.1 | 0.19 | 0.17 | 0.15 | 0.25 | 0.37 | 1.8 | 0.08 |  |  |  |  | 1
| 1 | 2 | 1.6 | 0.3 | 1.2182 | < 0.1 | 0.17 | 0.16 | 0.15 | 0.1 | 0.34 | 1.7 | 0.89 |  |  |  |  | 19
| 1 | 3 | 1.7 | 0.2 | 1.2057 | 0.1 | 0.17 | 0.14 | 0.15 | 0.1 | 0.44 | 1.9 | 1.09 |  |  |  |  | 16
| 1 | 4 | 1.7 | 0.2 | 1.2013 | 0.2 | 0.18 | 0.18 | 0.14 | 0.21 | 0.33 | 1.8 | 0.97 |  |  |  |  | 0.19114
| 1 | 5 | 1.8 | 0.2 | 1.2341 | 0.1 | 0.18 | 0.15 | 0.14 | 0.25 | 0.29 | 1.8 | 0.1 |  |  |  |  | 25
| 1 | 6 | 1.8 | 0.2 | 1.2131 | 0.1 | 0.19 | 0.13 | 0.14 | 0.22 | 0.31 | 1.8 | 0.84
| Mean |  | 1.7333 | 0.21667 | 1.2176 | 0.125 | 0.18 | 0.155 | 0.145 | 0.18833 | 0.34667 | 1.8 | 0.66167
| Median |  | 1.75 | 0.2 | 1.2156 | 0.1 | 0.18 | 0.155 | 0.145 | 0.215 | 0.335 | 1.8 | 0.865
| Std Dev. |  | 0.08165 | 0.040825 | 0.013792 | 0.05 | 0.0089443 | 0.018708 | 0.0054772 | 0.070261 | 0.053166 | 0.063246 | 0.45084
| Rel.Std.Dev. |  | 4.711% | 18.84% | 1.133% | 40% | 4.969% | 12.07% | 3.777% | 37.31% | 15.34% | 3.514% | 68.14%
| PDM3 |  | 806.8% | 13.35% | 537% | -34.6% | -5.83% | -18.91% | -24.14% | -1.47% | 81.37% | 841.7% | 246.2%
| Z-Score (Absolute) |  | 8.77 | 0 | 5.79 | 0.67 | 0.21 | 0.36 | 0.41 | 0.16 | 0.75 | 9.15 | 2.57
| Analytical results for Hf in OREAS 25b (Certified Value 4.68 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 4.62 | 3.45 | 4.9252 | 2.1 | 3.92 | 4.5 | 5 | 4.7 | 4.5 | 4.91 | 4.8 | 4.83 | 4.34 | 4.89 |  | 1
| 1 | 2 | 4.65 | 3.6 | 4.6887 | 2.8 | 4.25 | 4.7 | 4.9 | 4.6 | 5.2 | 4.69 | 4.8 | 4.45 | 4.3 | 4.69 |  | 11
| 1 | 3 | 4.77 | 3.55 | 4.8844 | 1.9 | 4.22 | 4.6 | 4.9 | 4.7 | 5 | 4.73 | 4.8 | 4.61 | 4.32 | 4.87 |  | 16
| 1 | 4 | 4.74 | 3.49 | 4.7217 | 2.3 | 4.22 | 4.5 | 4.9 | 4.5 | 4.7 | 4.66 | 4.7 | 4.41 | 4.07 | 4.69 |  | 4.6794
| 1 | 5 | 4.81 | 3.6 | 4.7391 | 1.8 | 4.16 | 4.6 | 4.8 | 4.6 | 4.6 | 4.57 | 5 | 4.75 | 4.52 | 4.68 |  | 96
| 1 | 6 | 4.68 | 3.47 | 4.8107 | 2.7 | 4.25 | 4.5 | 4.8 | 4.9 | 4.5 | 4.24 | 4.7 | 4.32 | 4.25 | 4.54
| Mean |  | 4.7117 | 3.5267 | 4.7949 | 2.2667 | 4.17 | 4.5667 | 4.8833 | 4.6667 | 4.75 | 4.6333 | 4.8 | 4.5617 | 4.3 | 4.7267
| Median |  | 4.71 | 3.52 | 4.7749 | 2.2 | 4.22 | 4.55 | 4.9 | 4.65 | 4.65 | 4.675 | 4.8 | 4.53 | 4.31 | 4.69
| Std Dev. |  | 0.073598 | 0.065929 | 0.094884 | 0.41312 | 0.12681 | 0.08165 | 0.075277 | 0.13663 | 0.2881 | 0.22295 | 0.10954 | 0.20183 | 0.14546 | 0.13186
| Rel.Std.Dev. |  | 1.562% | 1.869% | 1.979% | 18.23% | 3.041% | 1.788% | 1.542% | 2.928% | 6.065% | 4.812% | 2.282% | 4.425% | 3.383% | 2.79%
| PDM3 |  | 0.6891% | -24.63% | 2.469% | -51.56% | -10.89% | -2.41% | 4.358% | -0.2725% | 1.508% | -0.9849% | 2.577% | -2.516% | -8.108% | 1.01%
| Z-Score (Absolute) |  | 0.34 | 6.18 | 0.8 | 13.12 | 2.64 | 0.46 | 1.28 | 0.09 | 0.55 | 0.09 | 0.83 | 0.49 | 1.93 | 0.42
| Analytical results for Ho in OREAS 25b (Certified Value 0.46 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 0.5 | 0.6 | 0.4 | 0.6 | 0.38 | 0.49 | [9 empty] | 1
| 1 | 2 | 0.49 | 0.6 | 0.4 | 0.4 | 0.37 | 0.34 | [9 empty] | 22
| 1 | 3 | 0.5 | 0.6 | 0.4 | 0.4 | 0.42 | 0.48 | [9 empty] | 16
| 1 | 4 | 0.51 | 0.6 | 0.4 | 0.4 | 0.38 | 0.42 | [9 empty] | 0.45722
| 1 | 5 | 0.48 | 0.5 | 0.4 | 0.5 | 0.4 | 0.42 | [9 empty] | 97
| 1 | 6 | 0.53 | 0.6 | 0.4 | 0.4 | 0.36 | 0.39
| Mean |  | 0.50167 | 0.58333 | 0.4 | 0.45 | 0.385 | 0.42333
| Median |  | 0.5 | 0.6 | 0.4 | 0.4 | 0.38 | 0.42
| Std Dev. |  | 0.017224 | 0.040825 | 6.0809e-17 | 0.083666 | 0.021679 | 0.056095
| Rel.Std.Dev. |  | 3.433% | 6.999% | 1.52e-14% | 18.59% | 5.631% | 13.25%
| PDM3 |  | 9.721% | 27.58% | -12.52% | -1.58% | -15.8% | -7.412%
| Z-Score (Absolute) |  | 0.99 | 2.24 | 0.56 | 0.2 | 0.79 | 0.2
| Analytical results for In in OREAS 25b (Certified Value 0.095 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 3
3
| 1 | 1 | 0.1 | 0.09 | 0.092273 | 0.1 | 0.09 | 0.104 | 0.094 | 0.096 | 0.098 | 0.095 | 0.09 | 0.11 | 0.1 | 0.09 |  | 1
| 1 | 2 | 0.1 | 0.09 | 0.095397 | 0.1 | 0.1 | 0.097 | 0.103 | 0.093 | 0.098 | 0.092 | 0.09 | 0.1 | 0.096 | 0.07 |  | 13
| 1 | 3 | 0.09 | 0.1 | 0.09106 | 0.1 | 0.1 | 0.104 | 0.1 | 0.096 | 0.092 | 0.1 | 0.1 | 0.1 | 0.096 | 0.08 |  | 16
| 1 | 4 | 0.09 | 0.09 | 0.086072 | 0.1 | 0.1 | 0.099 | 0.104 | 0.097 | 0.097 | 0.093 | 0.09 | 0.1 | 0.094 | 0.08 |  | 0.095297
| 1 | 5 | 0.1 | 0.1 | 0.094113 | 0.1 | 0.09 | 0.1 | 0.097 | 0.091 | 0.096 | 0.09 | 0.1 | 0.1 | 0.103 | 0.08 |  | 98
| 1 | 6 | 0.09 | 0.09 | 0.084503 | 0.1 | 0.09 | 0.104 | 0.102 | 0.095 | 0.087 | 0.095 | 0.09 | 0.1 | 0.087 | 0.08
| Mean |  | 0.095 | 0.093333 | 0.09057 | 0.1 | 0.095 | 0.10133 | 0.1 | 0.094667 | 0.094667 | 0.094167 | 0.093333 | 0.10167 | 0.096 | 0.08
| Median |  | 0.095 | 0.09 | 0.091666 | 0.1 | 0.095 | 0.102 | 0.101 | 0.0955 | 0.0965 | 0.094 | 0.09 | 0.1 | 0.096 | 0.08
| Std Dev. |  | 0.0054772 | 0.005164 | 0.0043828 | 1.5202e-17 | 0.0054772 | 0.0030768 | 0.0038471 | 0.0022509 | 0.0043665 | 0.0034303 | 0.005164 | 0.0040825 | 0.0054772 | 0.0063246
| Rel.Std.Dev. |  | 5.766% | 5.533% | 4.839% | 1.52e-14% | 5.766% | 3.036% | 3.847% | 2.378% | 4.613% | 3.643% | 5.533% | 4.016% | 5.705% | 7.906%
| PDM3 |  | -0.3119% | -2.061% | -4.961% | 4.935% | -0.3119% | 6.334% | 4.935% | -0.6617% | -0.6617% | -1.186% | -2.061% | 6.684% | 0.7374% | -16.05%
| Z-Score (Absolute) |  | 0.07 | 0.67 | 1.92 | 2.32 | 0.07 | 2.92 | 2.32 | 0.07 | 0.07 | 0.3 | 0.67 | 3.07 | 0.52 | 6.67
| Analytical results for K in OREAS 25b (Certified Value 0.762 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.7702 | 0.73 | 0.78 | 0.79317 | 0.87 | 0.76 | 0.78 | 0.77 | 0.77 | 0.75 | 0.73 | 0.737 | 0.73 | 1.22 | 0.772 | 1
| 1 | 2 | 0.7684 | 0.77 | 0.78 | 0.77484 | 0.81 | 0.78 | 0.79 | 0.76 | 0.75 | 0.77 | 0.74 | 0.741 | 0.74 | 0.98 | 0.751 | #N/A
| 1 | 3 | 0.7537 | 0.79 | 0.78 | 0.76484 | 0.84 | 0.78 | 0.79 | 0.79 | 0.76 | 0.76 | 0.74 | 0.736 | 0.76 | 1 | 0.766 | 16
| 1 | 4 | 0.7651 | 0.74 | 0.78 | 0.7933 | 0.85 | 0.78 | 0.79 | 0.76 | 0.76 | 0.75 | 0.74 | 0.735 | 0.76 | 0.95 | 0.761 | 0.76199
| 1 | 5 | 0.7638 | 0.77 | 0.78 | 0.76329 | 0.86 | 0.77 | 0.8 | 0.76 | 0.76 | 0.76 | 0.73 | 0.738 | 0.75 | 0.99 | 0.77 | 99
| 1 | 6 | 0.7567 | 0.76 | 0.78 | 0.76774 | 0.84 | 0.77 | 0.8 | 0.74 | 0.74 | 0.74 | 0.74 | 0.736 | 0.75 | 0.93 | 0.767
| Mean |  | 0.76298 | 0.76 | 0.78 | 0.7762 | 0.845 | 0.77333 | 0.79167 | 0.76333 | 0.75667 | 0.755 | 0.73667 | 0.73717 | 0.74833 | 1.0117 | 0.7645
| Median |  | 0.76445 | 0.765 | 0.78 | 0.77129 | 0.845 | 0.775 | 0.79 | 0.76 | 0.76 | 0.755 | 0.74 | 0.7365 | 0.75 | 0.985 | 0.7665
| Std Dev. |  | 0.006515 | 0.021909 | 1.2162e-16 | 0.01378 | 0.020736 | 0.008165 | 0.0075277 | 0.01633 | 0.010328 | 0.010488 | 0.005164 | 0.002137 | 0.01169 | 0.10534 | 0.0076092
| Rel.Std.Dev. |  | 0.8539% | 2.883% | 1.559e-14% | 1.775% | 2.454% | 1.056% | 0.9509% | 2.139% | 1.365% | 1.389% | 0.701% | 0.2899% | 1.562% | 10.41% | 0.9953%
| PDM3 |  | 0.1306% | -0.2609% | 2.364% | 1.865% | 10.89% | 1.489% | 3.895% | 0.1765% | -0.6984% | -0.9171% | -3.323% | -3.257% | -1.792% | 32.77% | 0.3296%
| Z-Score (Absolute) |  | 0.02 | 0.17 | 0.87 | 0.67 | 4.28 | 0.52 | 1.49 | 0 | 0.35 | 0.44 | 1.4 | 1.37 | 0.79 | 13.02 | 0.06
| Analytical results for La in OREAS 25b (Certified Value 28.3 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 28.31 | 19.1 | 24.8 | 30.4 | 25.3 | 28.7 | 31.2 | 30.1 | 28.3 | 30.7 | 27.6 | 22.96 | 30.7 |  |  | 1
| 1 | 2 | 28.42 | 19.7 | 24.8 | 31.2 | 26.4 | 29.3 | 30.8 | 29.5 | 28.9 | 31.4 | 28.7 | 21.31 | 19.9 |  |  | 1
| 1 | 3 | 28.94 | 21.5 | 24.5 | 30.6 | 27 | 29 | 31.3 | 29.4 | 28 | 32 | 28.5 | 29.82 | 30.8 |  |  | 16
| 1 | 4 | 28.67 | 20.6 | 23.8 | 30.1 | 26.9 | 28.3 | 32 | 29.5 | 29 | 30.6 | 27.1 | 22.49 | 24.8 |  |  | 28.32
| 1 | 5 | 27.84 | 21.7 | 24.7 | 30.3 | 26.9 | 29.1 | 31.1 | 30.1 | 28.3 | 30.1 | 28.7 | 31.95 | 25.8 |  |  | 100
| 1 | 6 | 28.08 | 19.4 | 24.6 | 30.7 | 26.5 | 28.4 | 30.1 | 28.8 | 28.3 | 31.1 | 28.7 | 22.03 | 23.2
| Mean |  | 28.3767 | 20.3333 | 24.5333 | 30.55 | 26.5 | 28.8 | 31.0833 | 29.5667 | 28.4667 | 30.9833 | 28.2167 | 25.0933 | 25.8667
| Median |  | 28.365 | 20.15 | 24.65 | 30.5 | 26.7 | 28.85 | 31.15 | 29.5 | 28.3 | 30.9 | 28.6 | 22.725 | 25.3
| Std Dev. |  | 0.39642 | 1.1039 | 0.37771 | 0.38341 | 0.63561 | 0.4 | 0.62423 | 0.48854 | 0.39328 | 0.66758 | 0.69402 | 4.5691 | 4.2791
| Rel.Std.Dev. |  | 1.397% | 5.429% | 1.54% | 1.255% | 2.399% | 1.389% | 2.008% | 1.652% | 1.382% | 2.155% | 2.46% | 18.21% | 16.54%
| PDM3 |  | 0.2001% | -28.2% | -13.37% | 7.874% | -6.427% | 1.695% | 9.758% | 4.402% | 0.5179% | 9.404% | -0.3649% | -11.39% | -8.663%
| Z-Score (Absolute) |  | 0 | 2.5 | 1.19 | 0.67 | 0.58 | 0.13 | 0.84 | 0.37 | 0.03 | 0.81 | 0.05 | 1.02 | 0.78
| Analytical results for Li in OREAS 25b (Certified Value 29.5 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 29.3 | 33 | 30.3 | 31.335 | 29 | 29.6 | 31.1 | 29.4 | 30.5 | 29.8 | 29.5 | 28.8 | 29.2 | 50 | 32.1 | 1
| 1 | 2 | 29.9 | 34 | 30 | 31.1073 | 29 | 30.3 | 31.7 | 29.3 | 29.7 | 30.4 | 29.4 | 29.4 | 27 | 44 | 26.1 | #N/A
| 1 | 3 | 30 | 36 | 29.9 | 31.4107 | 29 | 30.6 | 31.2 | 30.1 | 29.9 | 29.9 | 30.5 | 28.8 | 27.4 | 43 | 26.7 | 16
| 1 | 4 | 29.9 | 33 | 29.7 | 30.8499 | 29 | 29.8 | 31.4 | 29.4 | 29.7 | 29.9 | 28.2 | 29.5 | 27.7 | 40 | 38.3 | 29.5442
| 1 | 5 | 29.5 | 35 | 29.7 | 30.8032 | 28 | 29.6 | 31.8 | 29.3 | 30 | 30.3 | 29.5 | 28.6 | 29.5 | 41 | 26.2 | 101
| 1 | 6 | 29.7 | 32 | 30 | 30.8994 | 29 | 29.2 | 31.9 | 28.6 | 29.5 | 31.8 | 28.8 | 28.6 | 27.3 | 38 | 26.8
| Mean |  | 29.7167 | 33.8333 | 29.9333 | 31.0676 | 28.8333 | 29.85 | 31.5167 | 29.35 | 29.8833 | 30.35 | 29.3167 | 28.95 | 28.0167 | 42.6667 | 29.3667
| Median |  | 29.8 | 33.5 | 29.95 | 31.0034 | 29 | 29.7 | 31.55 | 29.35 | 29.8 | 30.1 | 29.45 | 28.8 | 27.55 | 42 | 26.75
| Std Dev. |  | 0.27142 | 1.472 | 0.22509 | 0.25936 | 0.40825 | 0.51284 | 0.33116 | 0.47645 | 0.34881 | 0.75033 | 0.77309 | 0.39875 | 1.061 | 4.1793 | 4.933
| Rel.Std.Dev. |  | 0.9133% | 4.351% | 0.752% | 0.8348% | 1.416% | 1.718% | 1.051% | 1.623% | 1.167% | 2.472% | 2.637% | 1.377% | 3.787% | 9.795% | 16.8%
| PDM3 |  | 0.5838% | 14.52% | 1.317% | 5.156% | -2.406% | 1.035% | 6.676% | -0.6572% | 1.148% | 2.728% | -0.7701% | -2.011% | -5.17% | 44.42% | -0.6008%
| Z-Score (Absolute) |  | 0.17 | 5.04 | 0.11 | 1.54 | 1.29 | 0 | 2.11 | 0.63 | 0.04 | 0.63 | 0.67 | 1.14 | 2.32 | 16.2 | 0.61
| Analytical results for Lu in OREAS 25b (Certified Value 0.18 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 14 | 15 | 16 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [8 empty] | 2
2
| 1 | 1 | 0.21 | 0.12 | 0.2 | 0.1 | 0.2 | 0.16 | 0.2 | [8 empty] | 1
| 1 | 2 | 0.21 | 0.15 | 0.2 | 0.1 | 0.2 | 0.15 | 0.14 | [8 empty] | 23
| 1 | 3 | 0.21 | 0.15 | 0.2 | 0.2 | 0.2 | 0.16 | 0.2 | [8 empty] | 16
| 1 | 4 | 0.2 | 0.15 | 0.2 | 0.2 | 0.2 | 0.14 | 0.16 | [8 empty] | 0.1769
| 1 | 5 | 0.2 | 0.14 | 0.2 | 0.2 | 0.2 | 0.18 | 0.17 | [8 empty] | 102
| 1 | 6 | 0.21 | 0.13 | 0.2 | 0.2 | 0.2 | 0.14 | 0.15
| Mean |  | 0.20667 | 0.14 | 0.2 | 0.16667 | 0.2 | 0.155 | 0.17
| Median |  | 0.21 | 0.145 | 0.2 | 0.2 | 0.2 | 0.155 | 0.165
| Std Dev. |  | 0.005164 | 0.012649 | 3.0405e-17 | 0.05164 | 3.0405e-17 | 0.015166 | 0.025298
| Rel.Std.Dev. |  | 2.499% | 9.035% | 1.52e-14% | 30.98% | 1.52e-14% | 9.784% | 14.88%
| PDM3 |  | 16.82% | -20.86% | 13.06% | -5.787% | 13.06% | -12.38% | -3.903%
| Z-Score (Absolute) |  | 0.82 | 0.67 | 0.67 | 0.07 | 0.67 | 0.34 | 0
| Analytical results for Mg in OREAS 25b (Certified Value 0.355 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.356 | 0.27 | 0.36 | 0.38876 | 0.37 | 0.36 | 0.35 | 0.34 | 0.34 | 0.34 | 0.4 | 0.357 | 0.35 | 0.375 | 0.345 | 1
| 1 | 2 | 0.359 | 0.27 | 0.36 | 0.35535 | 0.36 | 0.37 | 0.36 | 0.34 | 0.33 | 0.35 | 0.4 | 0.355 | 0.35 | 0.367 | 0.309 | #N/A
| 1 | 3 | 0.3522 | 0.3 | 0.36 | 0.36706 | 0.36 | 0.37 | 0.36 | 0.34 | 0.34 | 0.34 | 0.4 | 0.356 | 0.36 | 0.372 | 0.352 | 16
| 1 | 4 | 0.3569 | 0.28 | 0.36 | 0.36449 | 0.36 | 0.37 | 0.36 | 0.34 | 0.33 | 0.33 | 0.4 | 0.357 | 0.36 | 0.357 | 0.343 | 0.35468
| 1 | 5 | 0.3558 | 0.3 | 0.36 | 0.38778 | 0.36 | 0.37 | 0.36 | 0.34 | 0.34 | 0.34 | 0.4 | 0.358 | 0.36 | 0.377 | 0.341 | 103
| 1 | 6 | 0.3531 | 0.27 | 0.35 | 0.36681 | 0.37 | 0.36 | 0.36 | 0.33 | 0.33 | 0.33 | 0.39 | 0.355 | 0.36 | 0.356 | 0.344
| Mean |  | 0.3555 | 0.28167 | 0.35833 | 0.37171 | 0.36333 | 0.36667 | 0.35833 | 0.33833 | 0.335 | 0.33833 | 0.39833 | 0.35633 | 0.35667 | 0.36733 | 0.339
| Median |  | 0.3559 | 0.275 | 0.36 | 0.36694 | 0.36 | 0.37 | 0.36 | 0.34 | 0.335 | 0.34 | 0.4 | 0.3565 | 0.36 | 0.3695 | 0.3435
| Std Dev. |  | 0.002498 | 0.01472 | 0.0040825 | 0.013525 | 0.005164 | 0.005164 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0075277 | 0.0040825 | 0.0012111 | 0.005164 | 0.009048 | 0.015166
| Rel.Std.Dev. |  | 0.7027% | 5.226% | 1.139% | 3.639% | 1.421% | 1.408% | 1.139% | 1.207% | 1.635% | 2.225% | 1.025% | 0.3399% | 1.448% | 2.463% | 4.474%
| PDM3 |  | 0.2304% | -20.59% | 1.029% | 4.8% | 2.439% | 3.379% | 1.029% | -4.61% | -5.549% | -4.61% | 12.31% | 0.4654% | 0.5593% | 3.567% | -4.422%
| Z-Score (Absolute) |  | 0.07 | 4.74 | 0.11 | 0.95 | 0.42 | 0.63 | 0.11 | 1.16 | 1.37 | 1.16 | 2.63 | 0.02 | 0 | 0.67 | 1.12
| Analytical results for Mn in OREAS 25b (Certified Value 0.024 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0245 | 0.02 | 0.0304 | 0.024657 | 0.0252 | 0.0244 | 0.0232 | 0.024 | 0.0241 | 0.0229 | 0.0215 | 0.0224 | 0.0243 | 0.0242 | 0.0262 | 1
| 1 | 2 | 0.0245 | 0.02 | 0.0305 | 0.025354 | 0.0251 | 0.0243 | 0.0238 | 0.0237 | 0.0238 | 0.0236 | 0.0217 | 0.0227 | 0.0235 | 0.0235 | 0.0216 | #N/A
| 1 | 3 | 0.0238 | 0.02 | 0.0304 | 0.026033 | 0.0249 | 0.0248 | 0.0237 | 0.0243 | 0.0241 | 0.0232 | 0.0218 | 0.0223 | 0.0236 | 0.0239 | 0.0235 | 16
| 1 | 4 | 0.0242 | 0.02 | 0.0304 | 0.024516 | 0.025 | 0.0245 | 0.0237 | 0.024 | 0.0237 | 0.0229 | 0.0214 | 0.0225 | 0.0236 | 0.023 | 0.0209 | 0.023602
| 1 | 5 | 0.0241 | 0.02 | 0.0303 | 0.02526 | 0.0246 | 0.0241 | 0.0239 | 0.0238 | 0.0236 | 0.0234 | 0.0214 | 0.0225 | 0.0249 | 0.0244 | 0.021 | 104
| 1 | 6 | 0.0236 | 0.02 | 0.0302 | 0.025821 | 0.0258 | 0.0243 | 0.0242 | 0.0234 | 0.0241 | 0.0227 | 0.0208 | 0.0224 | 0.0241 | 0.023 | 0.0217
| Mean |  | 0.024117 | 0.02 | 0.030367 | 0.025273 | 0.0251 | 0.0244 | 0.02375 | 0.023867 | 0.0239 | 0.023117 | 0.021433 | 0.022467 | 0.024 | 0.023667 | 0.022483
| Median |  | 0.02415 | 0.02 | 0.0304 | 0.025307 | 0.02505 | 0.02435 | 0.02375 | 0.0239 | 0.02395 | 0.02305 | 0.02145 | 0.02245 | 0.02385 | 0.0237 | 0.02165
| Std Dev. |  | 0.0003656 | 0 | 0.00010328 | 0.00060601 | 0.0004 | 0.00023664 | 0.00032711 | 0.00030768 | 0.00022804 | 0.00034303 | 0.00035024 | 0.00013663 | 0.00054406 | 0.00059889 | 0.0020469
| Rel.Std.Dev. |  | 1.516% | 0% | 0.3401% | 2.398% | 1.594% | 0.9698% | 1.377% | 1.289% | 0.9541% | 1.484% | 1.634% | 0.6081% | 2.267% | 2.531% | 9.104%
| PDM3 |  | 2.179% | -15.26% | 28.66% | 7.08% | 6.346% | 3.38% | 0.6258% | 1.12% | 1.261% | -2.058% | -9.19% | -4.812% | 1.685% | 0.2727% | -4.741%
| Z-Score (Absolute) |  | 0.22 | 3.48 | 5.84 | 1.26 | 1.11 | 0.48 | 0.1 | 0 | 0.03 | 0.67 | 2.19 | 1.26 | 0.12 | 0.18 | 1.24
| Analytical results for Mo in OREAS 25b (Certified Value 2.12 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 2.3 | 1.82 | 2 | 2.1489 | < 1 | 2.03 | 2.21 | 2.35 | 2.1 | 2.27 | 2.13 | 2.2 | 2.18 | 2 |  | 1
| 1 | 2 | 2.5 | 1.93 | 2 | 2.1224 | < 1 | 2.1 | 2.17 | 2.21 | 2.06 | 2.26 | 2.13 | 2.2 | 2.12 | 1.9 |  | 2
| 1 | 3 | 2.4 | 1.86 | 2.1 | 2.1403 | < 1 | 2.19 | 2.19 | 2.2 | 2.08 | 2.17 | 2.07 | 2.1 | 2.16 | 2 |  | 16
| 1 | 4 | 2.2 | 1.84 | 2.3 | 2.1339 | < 1 | 2.15 | 2.15 | 2.27 | 1.97 | 2.2 | 2.02 | 2.3 | 2.13 | 1.9 |  | 2.1151
| 1 | 5 | 2.4 | 1.9 | < 2 | 2.1377 | < 1 | 2.16 | 2.27 | 2.19 | 2.08 | 2.1 | 1.91 | 2.1 | 2.01 | 1.9 |  | 105
| 1 | 6 | 2.2 | 2.08 | 2 | 2.1236 | < 1 | 2.23 | 2.18 | 2.23 | 2.04 | 2.11 | 2.47 | 2.1 | 2.06 | 2
| Mean |  | 2.3333 | 1.905 | 2.08 | 2.1345 |  | 2.1433 | 2.195 | 2.2417 | 2.055 | 2.185 | 2.1217 | 2.1667 | 2.11 | 1.95
| Median |  | 2.35 | 1.88 | 2 | 2.1358 |  | 2.155 | 2.185 | 2.22 | 2.07 | 2.185 | 2.1 | 2.15 | 2.125 | 1.95
| Std Dev. |  | 0.12111 | 0.094604 | 0.13038 | 0.010152 |  | 0.070333 | 0.041833 | 0.060139 | 0.046368 | 0.072319 | 0.18936 | 0.08165 | 0.063875 | 0.054772
| Rel.Std.Dev. |  | 5.19% | 4.966% | 6.268% | 0.4756% |  | 3.281% | 1.906% | 2.683% | 2.256% | 3.31% | 8.925% | 3.768% | 3.027% | 2.809%
| PDM3 |  | 10.32% | -9.934% | -1.66% | 0.9154% |  | 1.334% | 3.777% | 5.983% | -2.842% | 3.304% | 0.3098% | 2.437% | -0.2418% | -7.806%
| Z-Score (Absolute) |  | 1.98 | 2.15 | 2.2 | 0.06 | 15.68 | 0.15 | 0.64 | 1.09 | 0.7 | 0.55 | 0.06 | 0.37 | 0.17 | 1.72
| Analytical results for Na in OREAS 25b (Certified Value 0.071 wt.%) | [16 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.0666 | 0.14 | 0.08 | 0.076781 | 0.07 | 0.069 | 0.07 | 0.07 | 0.08 | 0.07 | 0.07 | 0.07 | 0.066 | 0.151 | 0.078 | 1
| 1 | 2 | 0.0676 | 0.11 | 0.08 | 0.070873 | 0.07 | 0.072 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.07 | 0.073 | 0.134 | 0.0775 | #N/A
| 1 | 3 | 0.0653 | 0.12 | 0.08 | 0.079406 | 0.07 | 0.072 | 0.07 | 0.07 | 0.08 | 0.07 | 0.07 | 0.07 | 0.074 | 0.16 | 0.079 | 16
| 1 | 4 | 0.0662 | 0.11 | 0.08 | 0.076676 | 0.07 | 0.071 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.069 | 0.075 | 0.131 | 0.0805 | 0.070943
| 1 | 5 | 0.067 | 0.12 | 0.08 | 0.072059 | 0.07 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.07 | 0.07 | 0.067 | 0.163 | 0.08 | 106
| 1 | 6 | 0.0666 | 0.13 | 0.08 | 0.075161 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.074 | 0.115 | 0.0795
| Mean |  | 0.06655 | 0.12167 | 0.08 | 0.075159 | 0.07 | 0.070667 | 0.07 | 0.07 | 0.075 | 0.071667 | 0.07 | 0.069833 | 0.0715 | 0.14233 | 0.079083
| Median |  | 0.0666 | 0.12 | 0.08 | 0.075918 | 0.07 | 0.0705 | 0.07 | 0.07 | 0.075 | 0.07 | 0.07 | 0.07 | 0.0735 | 0.1425 | 0.07925
| Std Dev. |  | 0.00077395 | 0.01169 | 0 | 0.0031922 | 0 | 0.0012111 | 0 | 0 | 0.0054772 | 0.0040825 | 0 | 0.00040825 | 0.003937 | 0.018758 | 0.0011583
| Rel.Std.Dev. |  | 1.163% | 9.609% | 0% | 4.247% | 0% | 1.714% | 0% | 0% | 7.303% | 5.696% | 0% | 0.5846% | 5.506% | 13.18% | 1.465%
| PDM3 |  | -6.193% | 71.5% | 12.77% | 5.943% | -1.33% | -0.3899% | -1.33% | -1.33% | 5.718% | 1.02% | -1.33% | -1.565% | 0.7847% | 100.6% | 11.47%
| Z-Score (Absolute) |  | 2 | 20.3 | 3.44 | 1.48 | 0.61 | 0.34 | 0.61 | 0.61 | 1.42 | 0.07 | 0.61 | 0.67 | 0 | 28.66 | 3.07
| Analytical results for Nb in OREAS 25b (Certified Value 23.7 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 24.64 | 13.4 | 21.6 | 24.6847 | < 0.1 | 21.05 | 25.8 | 25.9 | 25.8 | 25 | 20.2 | 24.5 | 25.39 | 22.8 | 25.7 | 1
| 1 | 2 | 25.22 | 12.4 | 21 | 24.3957 | 0.1 | 22.37 | 25.7 | 25.2 | 25.5 | 24.4 | 19.6 | 24.4 | 25.03 | 22 | 23.5 | 3
| 1 | 3 | 25.36 | 11.4 | 22.1 | 22.5135 | 0.1 | 22.86 | 26.2 | 25 | 26.7 | 24.1 | 18.3 | 24.4 | 26.21 | 22.9 | 24.9 | 16
| 1 | 4 | 24.78 | 12.5 | 21.6 | 24.4385 | 0.2 | 22.35 | 25.2 | 25.5 | 26.1 | 25 | 17.7 | 24 | 25.25 | 21.4 | 22.2 | 23.7134
| 1 | 5 | 26.27 | 13.5 | 21.2 | 22.201 | 0.3 | 21.74 | 25.4 | 24.8 | 26.2 | 25 | 16.1 | 24.3 | 24.52 | 23.9 | 22.8 | 107
| 1 | 6 | 25.13 | 13.1 | 21.6 | 22.618 | 0.3 | 21.78 | 26.1 | 24.5 | 25.8 | 24.5 | 18.6 | 24.3 | 25.06 | 21.3 | 22.7
| Mean |  | 25.2333 | 12.7167 | 21.5167 | 23.4752 | 0.2 | 22.025 | 25.7333 | 25.15 | 26.0167 | 24.6667 | 18.4167 | 24.3167 | 25.2433 | 22.3833 | 23.6333
| Median |  | 25.175 | 12.8 | 21.6 | 23.5069 | 0.2 | 22.065 | 25.75 | 25.1 | 25.95 | 24.75 | 18.45 | 24.35 | 25.155 | 22.4 | 23.15
| Std Dev. |  | 0.57604 | 0.78846 | 0.38166 | 1.1421 | 0.1 | 0.63463 | 0.38816 | 0.501 | 0.41673 | 0.38816 | 1.4497 | 0.17224 | 0.55827 | 1.0028 | 1.3794
| Rel.Std.Dev. |  | 2.283% | 6.2% | 1.774% | 4.865% | 50% | 2.881% | 1.508% | 1.992% | 1.602% | 1.574% | 7.872% | 0.7083% | 2.212% | 4.48% | 5.837%
| PDM3 |  | 6.41% | -46.37% | -9.263% | -1.004% | -99.16% | -7.12% | 8.518% | 6.058% | 9.713% | 4.02% | -22.34% | 2.544% | 6.452% | -5.609% | -0.3374%
| Z-Score (Absolute) |  | 0.67 | 4.58 | 0.89 | 0.07 | 9.84 | 0.67 | 0.88 | 0.64 | 1 | 0.43 | 2.19 | 0.29 | 0.68 | 0.52 | 0
| Analytical results for Nd in OREAS 25b (Certified Value 21.2 ppm) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 1
2
| 1 | 1 | 22.8 | 23.3 | 20.1 | 21.7 | 18.3 | 23.8 | [9 empty] | 1
| 1 | 2 | 22.71 | 23.7 | 21.1 | 20.1 | 18.1 | 16.3 | [9 empty] | 4
| 1 | 3 | 23.43 | 23.3 | 21.3 | 21 | 18.6 | 23.6 | [9 empty] | 16
| 1 | 4 | 23.05 | 23.1 | 21.5 | 20.3 | 17.9 | 19.9 | [9 empty] | 21.1789
| 1 | 5 | 22.81 | 23.3 | 21.5 | 21.1 | 19.3 | 20.2 | [9 empty] | 108
| 1 | 6 | 22.84 | 23 | 21.1 | 20.7 | 17.7 | 18.7
| Mean |  | 22.94 | 23.2833 | 21.1 | 20.8167 | 18.3167 | 20.4167
| Median |  | 22.825 | 23.3 | 21.2 | 20.85 | 18.2 | 20.05
| Std Dev. |  | 0.26518 | 0.24014 | 0.52154 | 0.58109 | 0.57417 | 2.891
| Rel.Std.Dev. |  | 1.156% | 1.031% | 2.472% | 2.791% | 3.135% | 14.16%
| PDM3 |  | 8.315% | 9.937% | -0.3725% | -1.71% | -13.51% | -3.599%
| Z-Score (Absolute) |  | 1.06 | 1.24 | 0.08 | 0.08 | 1.41 | 0.29
| Analytical results for Ni in OREAS 25b (Certified Value 68 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 0
1
| 1 | 1 | 70.5 | 65.9 | 70.2 | 68.1252 | 73 | 68.7 | 73.7 | 69.7 | 68.5 | 68.1 | 68.2 | 66 | 65 | 71 | 64 | 1
| 1 | 2 | 72.2 | 65.7 | 69.9 | 67.9692 | 74 | 69.5 | 71.2 | 67.4 | 67.4 | 66.8 | 69 | 67 | 63.4 | 69 | 60 | 5
| 1 | 3 | 70.6 | 67.8 | 70.7 | 65.7901 | 74 | 70.3 | 70.8 | 67.7 | 69 | 66.8 | 73.3 | 66 | 63.7 | 72 | 63 | 16
| 1 | 4 | 69.9 | 64 | 70.4 | 65.8009 | 72 | 69.3 | 70.1 | 67.8 | 67.8 | 66.8 | 70.2 | 67 | 64.3 | 69 | 61 | 68.178
| 1 | 5 | 71 | 68 | 69.6 | 65.7886 | 72 | 69.2 | 69.9 | 66.7 | 68.3 | 66.8 | 69 | 66 | 65.8 | 76 | 64 | 109
| 1 | 6 | 70 | 64.3 | 69.9 | 66.8086 | 74 | 68.3 | 71.5 | 64.9 | 66.3 | 65.3 | 68.8 | 67 | 64.1 | 70 | 63
| Mean |  | 70.7 | 65.95 | 70.1167 | 66.7138 | 73.1667 | 69.2167 | 71.2 | 67.3667 | 67.8833 | 66.7667 | 69.75 | 66.5 | 64.3833 | 71.1667 | 62.5
| Median |  | 70.55 | 65.8 | 70.05 | 66.3047 | 73.5 | 69.25 | 71 | 67.55 | 68.05 | 66.8 | 69 | 66.5 | 64.2 | 70.5 | 63
| Std Dev. |  | 0.83905 | 1.6861 | 0.39707 | 1.1063 | 0.98319 | 0.68823 | 1.3711 | 1.5667 | 0.95376 | 0.88694 | 1.8566 | 0.54772 | 0.88412 | 2.6394 | 1.6432
| Rel.Std.Dev. |  | 1.187% | 2.557% | 0.5663% | 1.658% | 1.344% | 0.9943% | 1.926% | 2.326% | 1.405% | 1.328% | 2.662% | 0.8236% | 1.373% | 3.709% | 2.629%
| PDM3 |  | 3.699% | -3.268% | 2.843% | -2.148% | 7.317% | 1.523% | 4.432% | -1.19% | -0.4322% | -2.07% | 2.306% | -2.461% | -5.566% | 4.384% | -8.328%
| Z-Score (Absolute) |  | 0.98 | 0.67 | 0.78 | 0.41 | 1.84 | 0.47 | 1.16 | 0.18 | 0 | 0.39 | 0.65 | 0.48 | 1.22 | 1.15 | 1.88
| Analytical results for P in OREAS 25b (Certified Value 0.041 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.0411 | 0.0309 | 0.0443 | 0.04273 | 0.032 | 0.04 | 0.043 | 0.041 | 0.042 | 0.041 | 0.0386 | 0.041 | 0.043 | 0.049 | 0.042 | 1
| 1 | 2 | 0.0417 | 0.0301 | 0.044 | 0.043963 | 0.033 | 0.041 | 0.043 | 0.042 | 0.041 | 0.042 | 0.0391 | 0.042 | 0.039 | 0.045 | 0.0405 | #N/A
| 1 | 3 | 0.0417 | 0.0314 | 0.0437 | 0.042883 | 0.033 | 0.041 | 0.042 | 0.042 | 0.041 | 0.042 | 0.0382 | 0.042 | 0.039 | 0.046 | 0.0415 | 16
| 1 | 4 | 0.0408 | 0.0304 | 0.0448 | 0.043982 | 0.033 | 0.04 | 0.043 | 0.042 | 0.041 | 0.039 | 0.0384 | 0.042 | 0.039 | 0.046 | 0.04 | 0.041431
| 1 | 5 | 0.0403 | 0.0331 | 0.0445 | 0.04261 | 0.033 | 0.039 | 0.043 | 0.042 | 0.042 | 0.041 | 0.0385 | 0.042 | 0.043 | 0.048 | 0.04 | 110
| 1 | 6 | 0.0406 | 0.031 | 0.044 | 0.044178 | 0.034 | 0.039 | 0.043 | 0.04 | 0.041 | 0.041 | 0.0394 | 0.042 | 0.039 | 0.041 | 0.042
| Mean |  | 0.041033 | 0.03115 | 0.044217 | 0.043391 | 0.033 | 0.04 | 0.042833 | 0.0415 | 0.041333 | 0.041 | 0.0387 | 0.041833 | 0.040333 | 0.045833 | 0.041
| Median |  | 0.04095 | 0.03095 | 0.04415 | 0.043423 | 0.033 | 0.04 | 0.043 | 0.042 | 0.041 | 0.041 | 0.03855 | 0.042 | 0.039 | 0.046 | 0.041
| Std Dev. |  | 0.0005785 | 0.0010597 | 0.00039707 | 0.00072112 | 0.00063246 | 0.00089443 | 0.00040825 | 0.00083666 | 0.0005164 | 0.0010954 | 0.00045607 | 0.00040825 | 0.0020656 | 0.0027869 | 0.00094868
| Rel.Std.Dev. |  | 1.41% | 3.402% | 0.898% | 1.662% | 1.917% | 2.236% | 0.9531% | 2.016% | 1.249% | 2.672% | 1.178% | 0.9759% | 5.121% | 6.08% | 2.314%
| PDM3 |  | -0.9603% | -24.82% | 6.723% | 4.73% | -20.35% | -3.454% | 3.384% | 0.1661% | -0.2362% | -1.041% | -6.592% | 0.9706% | -2.65% | 10.63% | -1.041%
| Z-Score (Absolute) |  | 0 | 6.45 | 2.08 | 1.54 | 5.24 | 0.67 | 1.17 | 0.3 | 0.2 | 0.02 | 1.52 | 0.52 | 0.46 | 3.13 | 0.02
| Analytical results for Pb in OREAS 25b (Certified Value 20.9 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 22.3 | 20.7 | 21.3 | 21.0081 | 20 | 20.43 | 20.9 | 22.5 | 21.3 | 20.4 | 26.1 | 21 | 21.06 | 27.3 | 21 | 1
| 1 | 2 | 22.5 | 21 | 21 | 20.4955 | 19 | 21.48 | 21.6 | 20.8 | 20.9 | 22.3 | 22.5 | 21 | 20.12 | 25.6 | 18 | 6
| 1 | 3 | 21.8 | 21 | 20.9 | 20.8514 | 20 | 21.53 | 22.3 | 20.9 | 21.7 | 21.6 | 21.8 | 21 | 20.6 | 25.3 | 19 | 16
| 1 | 4 | 22.1 | 20.3 | 21.4 | 20.6977 | 20 | 21.52 | 20.8 | 21.1 | 21.3 | 20.3 | 24.8 | 21 | 22.13 | 27.6 | 19 | 20.9014
| 1 | 5 | 22.1 | 20.4 | 21.3 | 20.4563 | 20 | 21.62 | 20.9 | 21.1 | 21.9 | 20.5 | 24.9 | 21 | 20.34 | 27.8 | 20 | 111
| 1 | 6 | 21.7 | 21.7 | 21.2 | 20.7837 | 20 | 22.13 | 21.2 | 20.9 | 21.6 | 20.5 | 20.9 | 21 | 20.1 | 25.3 | 19
| Mean |  | 22.0833 | 20.85 | 21.1833 | 20.7155 | 19.8333 | 21.4517 | 21.2833 | 21.2167 | 21.45 | 20.9333 | 23.5 | 21 | 20.725 | 26.4833 | 19.3333
| Median |  | 22.1 | 20.85 | 21.25 | 20.7407 | 20 | 21.525 | 21.05 | 21 | 21.45 | 20.5 | 23.65 | 21 | 20.47 | 26.45 | 19
| Std Dev. |  | 0.29944 | 0.50892 | 0.19408 | 0.21196 | 0.40825 | 0.55568 | 0.57764 | 0.64005 | 0.35637 | 0.82138 | 2.0523 | 0 | 0.77547 | 1.2024 | 1.0328
| Rel.Std.Dev. |  | 1.356% | 2.441% | 0.9162% | 1.023% | 2.058% | 2.59% | 2.714% | 3.017% | 1.661% | 3.924% | 8.733% | 0% | 3.742% | 4.54% | 5.342%
| PDM3 |  | 5.655% | -0.2461% | 1.349% | -0.8898% | -5.11% | 2.632% | 1.827% | 1.508% | 2.624% | 0.1526% | 12.43% | 0.4715% | -0.8442% | 26.71% | -7.502%
| Z-Score (Absolute) |  | 1.82 | 0.67 | 0 | 0.95 | 2.73 | 0.54 | 0.2 | 0.07 | 0.54 | 0.51 | 4.69 | 0.37 | 0.93 | 10.72 | 3.74
| Analytical results for Pr in OREAS 25b (Certified Value 5.8 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
3
| 1 | 1 | 6.33 | 6.4 | 5.5 | 6.1 | 4.98 | 6.61 | [9 empty] | 1
| 1 | 2 | 6.29 | 6.5 | 5.8 | 5.5 | 4.91 | 4.44 | [9 empty] | 7
| 1 | 3 | 6.39 | 6.3 | 5.8 | 5.7 | 5.28 | 6.49 | [9 empty] | 16
| 1 | 4 | 6.42 | 6.3 | 5.9 | 5.4 | 4.88 | 5.36 | [9 empty] | 5.7989
| 1 | 5 | 6.18 | 6.3 | 5.8 | 6.1 | 5.46 | 5.59 | [9 empty] | 112
| 1 | 6 | 6.22 | 5.6 | 5.7 | 5.5 | 4.83 | 5.14
| Mean |  | 6.305 | 6.2333 | 5.75 | 5.7167 | 5.0567 | 5.605
| Median |  | 6.31 | 6.3 | 5.8 | 5.6 | 4.945 | 5.475
| Std Dev. |  | 0.093968 | 0.32042 | 0.13784 | 0.31252 | 0.25398 | 0.82788
| Rel.Std.Dev. |  | 1.49% | 5.14% | 2.397% | 5.467% | 5.023% | 14.77%
| PDM3 |  | 8.728% | 7.492% | -0.8431% | -1.418% | -12.8% | -3.344%
| Z-Score (Absolute) |  | 1.23 | 1.07 | 0.04 | 0.04 | 1.45 | 0.28
| Analytical results for Rb in OREAS 25b (Certified Value 63 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 0
1
| 1 | 1 | 62.81 | 48 | 62.9539 | 71 | 59.1 | 62.6 | 70.5 | 63.9 | 62.5 | 62.1 | 59.9 | 62.5 | 65 | 65.9 |  | 1
| 1 | 2 | 64.48 | 57 | 69.3553 | 69.8 | 59.9 | 62.4 | 68.8 | 62.9 | 64.7 | 61.7 | 61 | 67.6 | 63.7 | 61 |  | 8
| 1 | 3 | 63.53 | 59 | 69.6149 | 68.5 | 61.4 | 62.3 | 69 | 64 | 62.2 | 62.6 | 60.2 | 69.5 | 64.7 | 63.9 |  | 16
| 1 | 4 | 63.74 | 55.9 | 58.2924 | 69.3 | 60.9 | 61.5 | 69.6 | 61.7 | 61.4 | 60.1 | 60 | 65.8 | 61.9 | 59 |  | 62.8549
| 1 | 5 | 61.95 | 57.8 | 62.4591 | 62.9 | 61.3 | 60.7 | 68 | 64.5 | 63 | 59.6 | 61.3 | 60.1 | 68.9 | 57.9 |  | 113
| 1 | 6 | 62.59 | 55.4 | 58.8168 | 63.7 | 60.3 | 63.3 | 66.6 | 61.3 | 59.3 | 60.4 | 61.2 | 66.1 | 62 | 56.7
| Mean |  | 63.1833 | 55.5167 | 63.582 | 67.5333 | 60.4833 | 62.1333 | 68.75 | 63.05 | 62.1833 | 61.0833 | 60.6 | 65.2667 | 64.3667 | 60.7333
| Median |  | 63.17 | 56.45 | 62.7065 | 68.9 | 60.6 | 62.35 | 68.9 | 63.4 | 62.35 | 61.05 | 60.6 | 65.95 | 64.2 | 60
| Std Dev. |  | 0.90818 | 3.9041 | 4.9411 | 3.3874 | 0.89088 | 0.90921 | 1.3442 | 1.3142 | 1.7904 | 1.2123 | 0.63561 | 3.4262 | 2.5766 | 3.5803
| Rel.Std.Dev. |  | 1.437% | 7.032% | 7.771% | 5.016% | 1.473% | 1.463% | 1.955% | 2.084% | 2.879% | 1.985% | 1.049% | 5.25% | 4.003% | 5.895%
| PDM3 |  | 0.5225% | -11.67% | 1.157% | 7.443% | -3.773% | -1.148% | 9.379% | 0.3104% | -1.068% | -2.819% | -3.587% | 3.837% | 2.405% | -3.375%
| Z-Score (Absolute) |  | 0.21 | 2.64 | 0.36 | 1.82 | 0.79 | 0.18 | 2.28 | 0.16 | 0.16 | 0.57 | 0.75 | 0.98 | 0.65 | 0.7
| Analytical results for Re in OREAS 25b (Certified Value < 0.002 ppm) | [16 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.059 | < 0.05 | < 0.002 |  |  |  |  | 1
| 1 | 2 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.06 | < 0.05 | < 0.002 |  |  |  |  | 16
| 1 | 3 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.058 | < 0.05 | < 0.002 |  |  |  |  | 16
| 1 | 4 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.06 | < 0.05 | < 0.002 |  |  |  |  | < 0.002
| 1 | 5 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.049 | < 0.05 | 0.002 |  |  |  |  | 114
| 1 | 6 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.063 | < 0.05 | < 0.002
| Mean | [9 empty] | 0.058167 |  | 0.002
| Median | [9 empty] | 0.0595 |  | 0.002
| Std Dev. | [9 empty] | 0.0047924
| Rel.Std.Dev. | [9 empty] | 8.239%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 25b (Certified Value 0.026 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | < 0.05 | 0.02 | 0.03 | 0.025466 | 0.02 | < 0.04 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.023 | < 0.04 | 0.031 | 0.025 | 1
| 1 | 2 | < 0.05 | 0.02 | 0.03 | 0.023154 | 0.02 | < 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.023 | < 0.04 | 0.03 | 0.021 | 17
| 1 | 3 | < 0.05 | 0.02 | 0.03 | 0.023366 | 0.02 | < 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.022 | < 0.04 | 0.031 | 0.0265 | 16
| 1 | 4 | < 0.05 | 0.02 | 0.03 | 0.025227 | 0.02 | < 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.022 | < 0.04 | 0.029 | 0.0245 | 0.02634
| 1 | 5 | < 0.05 | 0.02 | 0.03 | 0.023713 | 0.02 | < 0.04 | 0.03 | 0.03 | 0.03 | 0.02 | 0.03 | 0.022 | < 0.04 | 0.032 | 0.0235 | 115
| 1 | 6 | < 0.05 | 0.04 | 0.03 | 0.024032 | 0.02 | < 0.04 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.023 | < 0.04 | 0.028 | 0.025
| Mean |  |  | 0.023333 | 0.03 | 0.02416 | 0.02 |  | 0.03 | 0.028333 | 0.028333 | 0.025 | 0.03 | 0.0225 |  | 0.030167 | 0.02425
| Median |  |  | 0.02 | 0.03 | 0.023873 | 0.02 |  | 0.03 | 0.03 | 0.03 | 0.025 | 0.03 | 0.0225 |  | 0.0305 | 0.02475
| Std Dev. |  |  | 0.008165 | 0 | 0.00096975 | 0 |  | 0 | 0.0040825 | 0.0040825 | 0.0054772 | 0 | 0.00054772 |  | 0.001472 | 0.0018641
| Rel.Std.Dev. |  |  | 34.99% | 0% | 4.014% | 0% |  | 0% | 14.41% | 14.41% | 21.91% | 0% | 2.434% |  | 4.879% | 7.687%
| PDM3 |  |  | -11.41% | 13.9% | -8.277% | -24.07% |  | 13.9% | 7.569% | 7.569% | -5.086% | 13.9% | -14.58% |  | 14.53% | -7.934%
| Z-Score (Absolute) |  | 0 | 0.34 | 1.01 | 0.17 | 1.01 | 1.01 | 1.01 | 0.67 | 0.67 | 0 | 1.01 | 0.51 | 1.01 | 1.05 | 0.15
| Analytical results for Sb in OREAS 25b (Certified Value 0.98 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 1.02 | 0.55 | < 5 | 0.9596 | < 0.1 | 0.92 | 0.96 | 1.12 | 0.98 | 0.99 | 0.79 | 1 | 1 | 1.02 | 1.1 | 1
| 1 | 2 | 1.03 | 0.52 | < 5 | 0.87013 | < 0.1 | 0.95 | 0.97 | 1.04 | 1 | 1.05 | 0.76 | 1 | 0.92 | 0.99 | 0.9 | 9
| 1 | 3 | 0.97 | 0.46 | < 5 | 0.96167 | < 0.1 | 0.98 | 1 | 1.04 | 0.95 | 1 | 0.75 | 1.1 | 0.93 | 1.12 | 1 | 16
| 1 | 4 | 1.01 | 0.51 | < 5 | 0.87008 | < 0.1 | 0.96 | 0.94 | 1.03 | 0.93 | 0.93 | 0.72 | 1.1 | 0.93 | 1.01 | 0.9 | 0.98057
| 1 | 5 | 1.07 | 0.5 | < 5 | 0.94585 | < 0.1 | 0.95 | 0.98 | 1.06 | 0.96 | 0.92 | 0.71 | 1.1 | 0.89 | 1.07 | 0.9 | 116
| 1 | 6 | 1.08 | 0.54 | < 5 | 1.0255 | < 0.1 | 0.96 | 0.98 | 0.99 | 0.94 | 0.98 | 0.6 | 1.1 | 0.92 | 0.96 | 0.9
| Mean |  | 1.03 | 0.51333 |  | 0.93881 |  | 0.95333 | 0.97167 | 1.0467 | 0.96 | 0.97833 | 0.72167 | 1.0667 | 0.93167 | 1.0283 | 0.95
| Median |  | 1.025 | 0.515 |  | 0.95272 |  | 0.955 | 0.975 | 1.04 | 0.955 | 0.985 | 0.735 | 1.1 | 0.925 | 1.015 | 0.9
| Std Dev. |  | 0.040497 | 0.032042 |  | 0.059943 |  | 0.019664 | 0.020412 | 0.04274 | 0.026077 | 0.047924 | 0.066156 | 0.05164 | 0.03656 | 0.057764 | 0.083666
| Rel.Std.Dev. |  | 3.932% | 6.242% |  | 6.385% |  | 2.063% | 2.101% | 4.083% | 2.716% | 4.898% | 9.167% | 4.841% | 3.924% | 5.617% | 8.807%
| PDM3 |  | 5.041% | -47.65% |  | -4.259% |  | -2.778% | -0.9081% | 6.741% | -2.098% | -0.2282% | -26.4% | 8.78% | -4.987% | 4.871% | -3.118%
| Z-Score (Absolute) |  | 0.69 | 4.41 | 15.2 | 0.21 | 8.98 | 0.07 | 0.12 | 0.86 | 0 | 0.18 | 2.35 | NA | 0.28 | 0.67 | NA
Results from laboratories 12 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 25b (Certified Value 15.3 ppm) | [16 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 15.9 | 6.9 | 14.9 | 17.0386 | 17 | 14.7 | 15.5 | 16.2 | 15.7 | 14.7 | 16 | 14 | 14.8 | 16.6 | 13.4 | 1
| 1 | 2 | 16.6 | 8.3 | 14.8 | 16.6807 | 17 | 14.9 | 15.4 | 15.5 | 15.6 | 15.2 | 15.8 | 14 | 14.8 | 16.1 | 9.6 | #N/A
| 1 | 3 | 16.2 | 9.1 | 14.8 | 16.8613 | 16 | 15.1 | 15.5 | 15.6 | 15.8 | 15 | 16 | 14 | 15 | 17.1 | 14.2 | 16
| 1 | 4 | 16.2 | 9.1 | 14.7 | 14.4675 | 17 | 14.9 | 15 | 15.8 | 15.4 | 14.6 | 15.3 | 14 | 14.8 | 16.4 | 14.1 | 15.288
| 1 | 5 | 15.9 | 8.9 | 14.8 | 14.7507 | 16 | 14.8 | 15.2 | 15 | 15.8 | 14.7 | 15.1 | 14 | 15 | 18.1 | 12.8 | 117
| 1 | 6 | 16.1 | 7.1 | 14.9 | 14.0299 | 16 | 14.6 | 15.2 | 14.9 | 15.3 | 14.1 | 15.4 | 14 | 14.9 | 16.1 | 14.3
| Mean |  | 16.15 | 8.2333 | 14.8167 | 15.6381 | 16.5 | 14.8333 | 15.3 | 15.5 | 15.6 | 14.7167 | 15.6 | 14 | 14.8833 | 16.7333 | 13.0667
| Median |  | 16.15 | 8.6 | 14.8 | 15.7157 | 16.5 | 14.85 | 15.3 | 15.55 | 15.65 | 14.7 | 15.6 | 14 | 14.85 | 16.5 | 13.75
| Std Dev. |  | 0.25884 | 1.0013 | 0.075277 | 1.363 | 0.54772 | 0.17512 | 0.2 | 0.4899 | 0.20976 | 0.37639 | 0.38471 | 0 | 0.098319 | 0.76594 | 1.7929
| Rel.Std.Dev. |  | 1.603% | 12.16% | 0.5081% | 8.716% | 3.32% | 1.181% | 1.307% | 3.161% | 1.345% | 2.558% | 2.466% | 0% | 0.6606% | 4.577% | 13.72%
| PDM3 |  | 5.639% | -46.14% | -3.083% | 2.29% | 7.928% | -2.974% | 0.07875% | 1.387% | 2.041% | -3.737% | 2.041% | -8.425% | -2.647% | 9.454% | -14.53%
| Z-Score (Absolute) |  | 1.19 | 9.86 | 0.67 | 0.47 | 1.67 | 0.65 | 0 | 0.28 | 0.42 | 0.81 | 0.42 | 1.81 | 0.58 | 2 | 3.12
| Analytical results for Se in OREAS 25b (Indicative Value 1.8 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | 1.8 | < 2 | < 5 | < 2 | 0.5 | 2 | 2 | 2 | 2 | 2 | 1.4 | 2 | 2.1 | < 2 |  | 1
| 1 | 2 | 1.9 | < 2 | < 5 | < 2 | 0.4 | 2 | 2 | < 1 | 2 | 2 | 1.4 | 3 | 2.7 | < 2 |  | 14
| 1 | 3 | 1.5 | < 2 | < 5 | < 2 | 0.6 | 1.9 | 2 | < 1 | 2 | 2 | 1.4 | 3 | 2.9 | < 2 |  | 16
| 1 | 4 | 1.7 | < 2 | < 5 | < 2 | 0.5 | 2 | 2 | 1 | 2 | 2 | 1.4 | 3 | 2.7 | < 2 |  | 1.8033
| 1 | 5 | 1.6 | < 2 | < 5 | < 2 | 0.8 | 2 | 2 | 1 | 2 | 2 | 1.3 | 2 | 2.2 | < 2 |  | 25
| 1 | 6 | 1.6 | < 2 | < 5 | < 2 | 0.9 | 1.9 | 2 | < 1 | 2 | 2 | 1.3 | 2 | 2.8 | < 2
| Mean |  | 1.6833 |  |  |  | 0.61667 | 1.9667 | 2 | 1.3333 | 2 | 2 | 1.3667 | 2.5 | 2.5667
| Median |  | 1.65 |  |  |  | 0.55 | 2 | 2 | 1 | 2 | 2 | 1.4 | 2.5 | 2.7
| Std Dev. |  | 0.1472 |  |  |  | 0.19408 | 0.05164 | 0 | 0.57735 | 0 | 0 | 0.05164 | 0.54772 | 0.33267
| Rel.Std.Dev. |  | 8.744% |  |  |  | 31.47% | 2.626% | 0% | 43.3% | 0% | 0% | 3.779% | 21.91% | 12.96%
| PDM3 |  | -6.654% |  |  |  | -65.8% | 9.057% | 10.91% | -26.06% | 10.91% | 10.91% | -24.21% | 38.63% | 42.33%
| Z-Score (Absolute) |  | 0.14 | 0.82 | 0.67 | 0.82 | 1.21 | 0.14 | 0.17 | 0.91 | 0.17 | 0.17 | 0.46 | 0.67 | 0.74 | 0.82
| Analytical results for Sm in OREAS 25b (Certified Value 3.93 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 4.09 | 4.7 | 3.7 | 4.2 | 3.41 | 4.3 | [9 empty] | 1
| 1 | 2 | 4.09 | 4.4 | 3.8 | 3.9 | 3.35 | 2.9 | [9 empty] | 10
| 1 | 3 | 4.11 | 4.3 | 3.9 | 4.1 | 3.61 | 4.2 | [9 empty] | 16
| 1 | 4 | 4.14 | 4.5 | 3.9 | 4 | 3.34 | 3.5 | [9 empty] | 3.9283
| 1 | 5 | 4.06 | 4.4 | 3.9 | 4.1 | 3.65 | 3.6 | [9 empty] | 118
| 1 | 6 | 4 | 4.7 | 3.9 | 3.9 | 3.27 | 3.5
| Mean |  | 4.0817 | 4.5 | 3.85 | 4.0333 | 3.4383 | 3.6667
| Median |  | 4.09 | 4.45 | 3.9 | 4.05 | 3.38 | 3.55
| Std Dev. |  | 0.047924 | 0.16733 | 0.083666 | 0.12111 | 0.15549 | 0.5164
| Rel.Std.Dev. |  | 1.174% | 3.718% | 2.173% | 3.003% | 4.522% | 14.08%
| PDM3 |  | 3.903% | 14.55% | -1.994% | 2.673% | -12.47% | -6.661%
| Z-Score (Absolute) |  | 0.45 | 1.81 | 0.3 | 0.3 | 1.64 | 0.89
| Analytical results for Sn in OREAS 25b (Certified Value 5.07 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 5.4 | 4.7 | < 10 | 5.0249 | < 1 | 4.8 | 5.1 | 5.5 | 4.8 | 5.1 | 5 | 5.6 | 5.1 | 5.3 | 5.3 | 1
| 1 | 2 | 5.5 | 4.8 | < 10 | 5.3053 | < 1 | 4.9 | 5.1 | 5.4 | 5 | 5.2 | 5 | 5.4 | 5.2 | 5.3 | 4.2 | 11
| 1 | 3 | 5.6 | 4.5 | < 10 | 5.0067 | < 1 | 5.1 | 5.1 | 5.5 | 4.9 | 5 | 5 | 5.2 | 5.2 | 5.6 | 4.6 | 16
| 1 | 4 | 5.3 | 4.4 | < 10 | 5.2578 | < 1 | 5.2 | 4.9 | 5.5 | 4.8 | 5 | 4.8 | 5.3 | 5 | 5.1 | 4.8 | 5.0739
| 1 | 5 | 5.5 | 4.5 | < 10 | 4.9521 | < 1 | 5.1 | 5 | 5.5 | 4.9 | 5 | 4.7 | 5 | 4.9 | 5.8 | 4.8 | 119
| 1 | 6 | 5.2 | 4.5 | < 10 | 5.1209 | < 1 | 5 | 5.2 | 5.6 | 4.9 | 5 | 4.7 | 5.1 | 5.2 | 5.2 | 4.7
| Mean |  | 5.4167 | 4.5667 |  | 5.1113 |  | 5.0167 | 5.0667 | 5.5 | 4.8833 | 5.05 | 4.8667 | 5.2667 | 5.1 | 5.3833 | 4.7333
| Median |  | 5.45 | 4.5 |  | 5.0729 |  | 5.05 | 5.1 | 5.5 | 4.9 | 5 | 4.9 | 5.25 | 5.15 | 5.3 | 4.75
| Std Dev. |  | 0.1472 | 0.15055 |  | 0.14349 |  | 0.1472 | 0.10328 | 0.063246 | 0.075277 | 0.083666 | 0.15055 | 0.21602 | 0.12649 | 0.26394 | 0.3559
| Rel.Std.Dev. |  | 2.717% | 3.297% |  | 2.807% |  | 2.934% | 2.038% | 1.15% | 1.542% | 1.657% | 3.094% | 4.102% | 2.48% | 4.903% | 7.519%
| PDM3 |  | 6.755% | -9.998% |  | 0.7358% |  | -1.129% | -0.1434% | 8.397% | -3.757% | -0.4719% | -4.085% | 3.798% | 0.5135% | 6.098% | -6.713%
| Z-Score (Absolute) |  | 1.35 | 1.78 | 0.18 | 0.23 | 16.74 | 0.12 | 0.06 | 1.66 | 0.61 | 0 | 0.67 | 0.8 | 0.18 | 1.23 | 1.16
| Analytical results for Sr in OREAS 25b (Certified Value 32.3 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 33.7 | 22.1 | 46.7 | 31.8041 | 36.4 | 31 | 34.2 | 36 | 32.9 | 33.2 | 31 | 30.5 | 32 | 33.4 | 33.2 | 1
| 1 | 2 | 33.03 | 24.5 | 45.5 | 28.6601 | 35.6 | 32 | 33.8 | 35.2 | 31.6 | 30.6 | 30.8 | 30.7 | 33 | 32.9 | 28.5 | #N/A
| 1 | 3 | 33.18 | 26.7 | 46.4 | 26.8383 | 36.2 | 32 | 34.5 | 35.3 | 32.3 | 30.1 | 31.1 | 30.5 | 34 | 32.5 | 32.4 | 16
| 1 | 4 | 33.57 | 24.1 | 45.7 | 31.5845 | 35.1 | 32 | 32.8 | 36 | 32.1 | 32.9 | 29.9 | 30.2 | 33 | 32.1 | 28.6 | 32.3363
| 1 | 5 | 33.21 | 25.9 | 45.5 | 33.7539 | 35.3 | 32 | 34.2 | 34.3 | 32.2 | 33.2 | 29.8 | 30.8 | 33 | 33 | 28.7 | 120
| 1 | 6 | 32.35 | 26.3 | 44.7 | 29.6134 | 34.7 | 31 | 34.7 | 34.2 | 32.2 | 32.4 | 30.1 | 30.5 | 33 | 31.3 | 28.6
| Mean |  | 33.1733 | 24.9333 | 45.75 | 30.3757 | 35.55 | 31.6667 | 34.0333 | 35.1667 | 32.2167 | 32.0667 | 30.45 | 30.5333 | 33 | 32.5333 | 30
| Median |  | 33.195 | 25.2 | 45.6 | 30.599 | 35.45 | 32 | 34.2 | 35.25 | 32.2 | 32.65 | 30.45 | 30.5 | 33 | 32.7 | 28.65
| Std Dev. |  | 0.47635 | 1.7224 | 0.71484 | 2.488 | 0.65345 | 0.5164 | 0.67725 | 0.78655 | 0.41673 | 1.3706 | 0.58224 | 0.20656 | 0.63246 | 0.75011 | 2.1845
| Rel.Std.Dev. |  | 1.436% | 6.908% | 1.562% | 8.191% | 1.838% | 1.631% | 1.99% | 2.237% | 1.294% | 4.274% | 1.912% | 0.6765% | 1.917% | 2.306% | 7.282%
| PDM3 |  | 2.588% | -22.89% | 41.48% | -6.063% | 9.938% | -2.071% | 5.248% | 8.753% | -0.3701% | -0.834% | -5.833% | -5.576% | 2.052% | 0.6092% | -7.225%
| Z-Score (Absolute) |  | 0.37 | 2.78 | 5.17 | 0.7 | 1.27 | 0.21 | 0.69 | 1.13 | 0 | 0.06 | 0.67 | 0.64 | 0.3 | 0.12 | 0.85
| Analytical results for Ta in OREAS 25b (Certified Value 1.65 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 2
2
| 1 | 1 | 1.67 | 0.49 | 12.3 | 1.6746 | < 0.1 | 1.4 | 1.59 | 1.82 | 1.64 | 1.55 | 1.33 | 1.66 | 1.8 | 1.89 | 1.94 | 1
| 1 | 2 | 1.72 | 0.48 | 12.4 | 1.6492 | < 0.1 | 1.4 | 1.62 | 1.82 | 1.61 | 1.66 | 1.23 | 1.61 | 1.9 | 1.78 | 2.07 | 1
| 1 | 3 | 1.82 | 0.4 | 12.4 | 1.6331 | < 0.1 | 1.5 | 1.67 | 1.79 | 1.63 | 1.62 | 1.14 | 1.57 | 2 | 1.86 | 1.96 | 16
| 1 | 4 | 1.73 | 0.52 | 12.1 | 1.618 | < 0.1 | 1.4 | 1.62 | 1.82 | 1.56 | 1.52 | 1.11 | 1.57 | 2 | 1.72 | 1.71 | 1.6509
| 1 | 5 | 1.75 | 0.49 | 12.1 | 1.6606 | < 0.1 | 1.5 | 1.61 | 1.84 | 1.62 | 1.57 | 1.02 | 1.6 | 1.8 | 1.9 | 1.63 | 121
| 1 | 6 | 1.78 | 0.47 | 12.1 | 1.6963 | < 0.1 | 1.5 | 1.65 | 1.8 | 1.6 | 1.52 | 1.19 | 1.58 | 1.9 | 1.83 | 1.53
| Mean |  | 1.745 | 0.475 | 12.2333 | 1.6553 |  | 1.45 | 1.6267 | 1.815 | 1.61 | 1.5733 | 1.17 | 1.5983 | 1.9 | 1.83 | 1.8067
| Median |  | 1.74 | 0.485 | 12.2 | 1.6549 |  | 1.45 | 1.62 | 1.82 | 1.615 | 1.56 | 1.165 | 1.59 | 1.9 | 1.845 | 1.825
| Std Dev. |  | 0.051672 | 0.040373 | 0.15055 | 0.028291 |  | 0.054772 | 0.028752 | 0.017607 | 0.028284 | 0.056451 | 0.1064 | 0.034303 | 0.089443 | 0.069282 | 0.21342
| Rel.Std.Dev. |  | 2.961% | 8.5% | 1.231% | 1.709% |  | 3.777% | 1.768% | 0.9701% | 1.757% | 3.588% | 9.094% | 2.146% | 4.708% | 3.786% | 11.81%
| PDM3 |  | 5.703% | -71.23% | 641% | 0.2691% |  | -12.17% | -1.465% | 9.943% | -2.475% | -4.696% | -29.13% | -3.182% | 15.09% | 10.85% | 9.438%
| Z-Score (Absolute) |  | 0.44 | 4.31 | 39.73 | 0.11 | 5.91 | 0.66 | 0 | 0.71 | 0.06 | 0.2 | 1.71 | 0.11 | 1.02 | 0.76 | 0.67
| Analytical results for Tb in OREAS 25b (Certified Value 0.43 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 14 | 15 | 16 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [8 empty] | 2
3
| 1 | 1 | 0.43 | 0.3 | 0.5 | 0.4 | 0.5 | 0.42 | 0.47 | [8 empty] | 1
| 1 | 2 | 0.48 | 0.35 | 0.5 | 0.4 | 0.4 | 0.4 | 0.34 | [8 empty] | #N/A
| 1 | 3 | 0.47 | 0.38 | 0.5 | 0.4 | 0.4 | 0.43 | 0.49 | [8 empty] | 16
| 1 | 4 | 0.46 | 0.36 | 0.5 | 0.4 | 0.4 | 0.39 | 0.43 | [8 empty] | 0.4268
| 1 | 5 | 0.48 | 0.38 | 0.5 | 0.4 | 0.4 | 0.44 | 0.44 | [8 empty] | 122
| 1 | 6 | 0.48 | 0.33 | 0.5 | 0.4 | 0.4 | 0.4 | 0.41
| Mean |  | 0.46667 | 0.35 | 0.5 | 0.4 | 0.41667 | 0.41333 | 0.43
| Median |  | 0.475 | 0.355 | 0.5 | 0.4 | 0.4 | 0.41 | 0.435
| Std Dev. |  | 0.019664 | 0.030984 | 0 | 6.0809e-17 | 0.040825 | 0.019664 | 0.052536
| Rel.Std.Dev. |  | 4.214% | 8.853% | 0% | 1.52e-14% | 9.798% | 4.757% | 12.22%
| PDM3 |  | 9.341% | -17.99% | 17.15% | -6.279% | -2.374% | -3.155% | 0.7498%
| Z-Score (Absolute) |  | 2.02 | 2.7 | 3.37 | 0.67 | 0 | 0.13 | 0.54
| Analytical results for Te in OREAS 25b (Certified Value < 0.05 ppm) | [16 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 3
3
| 1 | 1 | < 0.2 | < 0.05 | 17.9 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.04 | 0.1 | 0.06 | < 0.1 |  | 1
| 1 | 2 | < 0.2 | < 0.05 | 17.5 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | 0.04 | 0.1 | < 0.05 | < 0.1 |  | 15
| 1 | 3 | < 0.2 | < 0.05 | 17.7 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.1 | 0.04 | 0.1 | < 0.05 | < 0.1 |  | 16
| 1 | 4 | < 0.2 | < 0.05 | 18.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.04 | 0.1 | < 0.05 | < 0.1 |  | < 0.05
| 1 | 5 | < 0.2 | < 0.05 | 18.3 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.02 | 0.1 | 0.11 | < 0.1 |  | 123
| 1 | 6 | < 0.2 | < 0.05 | 18.3 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.03 | 0.1 | < 0.05 | < 0.1
| Mean |  |  |  | 17.9667 |  |  |  |  | 0.055 |  | 0.1 | 0.035 | 0.1 | 0.085
| Median |  |  |  | 18 |  |  |  |  | 0.055 |  | 0.1 | 0.04 | 0.1 | 0.085
| Std Dev. |  |  |  | 0.3266 |  |  |  |  | 0.0070711 |  |  | 0.0083666 | 1.5202e-17 | 0.035355
| Rel.Std.Dev. |  |  |  | 1.818% |  |  |  |  | 12.86% |  |  | 23.9% | 1.52e-14% | 41.59%
PDM3
| Z-Score (Absolute) |  | 3.17 | 0.88 | 966.98 | 0.47 | 0.47 | 0.88 | 0.88 | 0.34 | 0.88 | 0.2 | 0.34 | 3.17 | 0.2 | 0.47
| Analytical results for Th in OREAS 25b (Certified Value 12.8 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 13.18 | 7.2 | 12.9255 | 15.4 | 11.8 | 13.2 | 13.15 | 13.35 | 12.3 | 13.8 | 12.7 | 12.6 | 12.13 | 13.5 |  | 1
| 1 | 2 | 13.23 | 8.6 | 14.8314 | 15.5 | 12.4 | 13.5 | 13.25 | 13.25 | 13.6 | 13.2 | 12.6 | 11.8 | 11.8 | 10 |  | 2
| 1 | 3 | 12.84 | 9 | 14.6928 | 15.2 | 12.5 | 13.65 | 13.65 | 13.35 | 13.15 | 13.5 | 12.6 | 11.9 | 12.27 | 12.7 |  | 16
| 1 | 4 | 12.89 | 8.6 | 14.2584 | 15 | 12.6 | 13 | 13.35 | 13.5 | 11.85 | 13 | 12.4 | 11.2 | 11.52 | 11.1 |  | 12.8301
| 1 | 5 | 12.8 | 9.2 | 12.657 | 15.3 | 12.5 | 13.4 | 13.2 | 13.5 | 12.2 | 13.4 | 12.5 | 12.2 | 12.94 | 11.3 |  | 124
| 1 | 6 | 12.98 | 8.2 | 14.7418 | 15.3 | 12.5 | 13.4 | 13.35 | 13.35 | 12 | 13.1 | 12.7 | 12.2 | 11.9 | 10.8
| Mean |  | 12.9867 | 8.4667 | 14.0178 | 15.2833 | 12.3833 | 13.3583 | 13.325 | 13.3833 | 12.5167 | 13.3333 | 12.5833 | 11.9833 | 12.0933 | 11.5667
| Median |  | 12.935 | 8.6 | 14.4756 | 15.3 | 12.5 | 13.4 | 13.3 | 13.35 | 12.25 | 13.3 | 12.6 | 12.05 | 12.015 | 11.2
| Std Dev. |  | 0.18019 | 0.71181 | 0.97411 | 0.17224 | 0.29269 | 0.22895 | 0.17819 | 0.098319 | 0.69761 | 0.29439 | 0.1169 | 0.47504 | 0.49021 | 1.2925
| Rel.Std.Dev. |  | 1.387% | 8.407% | 6.949% | 1.127% | 2.364% | 1.714% | 1.337% | 0.7346% | 5.573% | 2.208% | 0.929% | 3.964% | 4.054% | 11.17%
| PDM3 |  | 1.22% | -34.01% | 9.257% | 19.12% | -3.482% | 4.117% | 3.857% | 4.312% | -2.443% | 3.922% | -1.923% | -6.6% | -5.742% | -9.847%
| Z-Score (Absolute) |  | 0.23 | 4.97 | 1.42 | 2.88 | 0.46 | 0.66 | 0.62 | 0.69 | 0.31 | 0.63 | 0.23 | 0.92 | 0.8 | 1.4
| Analytical results for Ti in OREAS 25b (Certified Value 1.01 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.0192 | 0.82 | 0.98 | 1.0648 | 0.103 | 1.01 | 1.05 | 1.035 | 1.01 | 1.01 | 0.83 | 0.955 | 1.044 | 0.985 | 1.03 | 1
| 1 | 2 | 1.0058 | 0.88 | 0.97 | 1.0662 | 0.142 | 1.001 | 1.06 | 1.015 | 1.005 | 1.05 | 0.83 | 0.957 | 1.022 | 0.977 | 1.02 | #N/A
| 1 | 3 | 1.0208 | 0.85 | 0.99 | 1.0604 | 0.0758 | 1.025 | 1.05 | 1.03 | 1.005 | 1.04 | 0.81 | 0.959 | 1.031 | 0.986 | 1.03 | 16
| 1 | 4 | 1.0266 | 0.82 | 0.97 | 1.0433 | 0.139 | 1.005 | 1.055 | 1.035 | 1.005 | 0.998 | 0.8 | 0.959 | 1.034 | 0.939 | 1.01 | 1.0119
| 1 | 5 | 1.0103 | 0.85 | 0.97 | 1.0378 | 0.101 | 1.002 | 1.08 | 1.03 | 0.992 | 1.015 | 0.78 | 0.948 | 1.052 | 0.996 | 1.03 | 125
| 1 | 6 | 1.008 | 0.84 | 0.97 | 1.0432 | 0.12 | 1.004 | 1.08 | 0.991 | 1 | 0.981 | 0.78 | 0.945 | 1.042 | 0.931 | 1.01
| Mean |  | 1.0151 | 0.84333 | 0.975 | 1.0526 | 0.11347 | 1.0078 | 1.0625 | 1.0227 | 1.0028 | 1.0157 | 0.805 | 0.95383 | 1.0375 | 0.969 | 1.0217
| Median |  | 1.0147 | 0.845 | 0.97 | 1.0518 | 0.1115 | 1.0045 | 1.0575 | 1.03 | 1.005 | 1.0125 | 0.805 | 0.956 | 1.038 | 0.981 | 1.025
| Std Dev. |  | 0.0082642 | 0.022509 | 0.0083666 | 0.012569 | 0.025269 | 0.0089759 | 0.014053 | 0.017166 | 0.0061779 | 0.025758 | 0.022583 | 0.005947 | 0.010654 | 0.027137 | 0.0098319
| Rel.Std.Dev. |  | 0.8141% | 2.669% | 0.8581% | 1.194% | 22.27% | 0.8906% | 1.323% | 1.679% | 0.616% | 2.536% | 2.805% | 0.6235% | 1.027% | 2.8% | 0.9623%
| PDM3 |  | 0.3198% | -16.66% | -3.645% | 4.026% | -88.79% | -0.4% | 5.002% | 1.066% | -0.8942% | 0.3741% | -20.45% | -5.737% | 2.532% | -4.238% | 0.9671%
| Z-Score (Absolute) |  | 0.15 | 3.38 | 0.67 | 0.92 | 18.37 | 0 | 1.12 | 0.3 | 0.1 | 0.16 | 4.17 | 1.11 | 0.61 | 0.8 | 0.28
| Analytical results for Tl in OREAS 25b (Certified Value 0.4 ppm) | [16 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 0.43 | 0.35 | 0.42654 | 0.42 | 0.38 | 0.35 | 0.38 | 0.39 | 0.35 | 0.47 | 0.41 | 0.4 | 0.39 | 0.4 |  | 1
| 1 | 2 | 0.42 | 0.38 | 0.40492 | 0.43 | 0.39 | 0.36 | 0.4 | 0.38 | 0.37 | 0.45 | 0.41 | 0.38 | 0.39 | 0.3 |  | #N/A
| 1 | 3 | 0.42 | 0.38 | 0.42952 | 0.42 | 0.4 | 0.35 | 0.4 | 0.37 | 0.41 | 0.45 | 0.41 | 0.4 | 0.44 | 0.4 |  | 16
| 1 | 4 | 0.42 | 0.37 | 0.42944 | 0.42 | 0.4 | 0.35 | 0.41 | 0.37 | 0.33 | 0.45 | 0.41 | 0.39 | 0.39 | 0.4 |  | 0.39729
| 1 | 5 | 0.42 | 0.37 | 0.42461 | 0.41 | 0.4 | 0.35 | 0.4 | 0.38 | 0.33 | 0.44 | 0.42 | 0.38 | 0.44 | 0.4 |  | 126
| 1 | 6 | 0.41 | 0.36 | 0.40326 | 0.43 | 0.4 | 0.36 | 0.39 | 0.37 | 0.33 | 0.43 | 0.41 | 0.4 | 0.4 | 0.4
| Mean |  | 0.42 | 0.36833 | 0.41971 | 0.42167 | 0.395 | 0.35333 | 0.39667 | 0.37667 | 0.35333 | 0.44833 | 0.41167 | 0.39167 | 0.40833 | 0.38333
| Median |  | 0.42 | 0.37 | 0.42558 | 0.42 | 0.4 | 0.35 | 0.4 | 0.375 | 0.34 | 0.45 | 0.41 | 0.395 | 0.395 | 0.4
| Std Dev. |  | 0.0063246 | 0.01169 | 0.012255 | 0.0075277 | 0.0083666 | 0.005164 | 0.010328 | 0.008165 | 0.032042 | 0.013292 | 0.0040825 | 0.0098319 | 0.024833 | 0.040825
| Rel.Std.Dev. |  | 1.506% | 3.174% | 2.92% | 1.785% | 2.118% | 1.462% | 2.604% | 2.168% | 9.068% | 2.965% | 0.9917% | 2.51% | 6.081% | 10.65%
| PDM3 |  | 5.717% | -7.288% | 5.645% | 6.137% | -0.5753% | -11.06% | -0.1558% | -5.19% | -11.06% | 12.85% | 3.62% | -1.414% | 2.781% | -3.512%
| Z-Score (Absolute) |  | 0.68 | 0.83 | 0.67 | 0.73 | 0.05 | 1.27 | 0 | 0.59 | 1.27 | 1.51 | 0.44 | 0.15 | 0.34 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 25b (Certified Value 0.19 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 05 | 06 | 14 | 15 | 16 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 2
2
| 1 | 1 | 0.23 | 0.2 | 0.1 | 0.2 | 0.17 | 0.2 | [9 empty] | 1
| 1 | 2 | 0.23 | 0.2 | 0.2 | 0.2 | 0.15 | 0.13 | [9 empty] | 25
| 1 | 3 | 0.22 | 0.2 | 0.2 | 0.2 | 0.16 | 0.19 | [9 empty] | 16
| 1 | 4 | 0.2 | 0.2 | 0.2 | 0.2 | 0.15 | 0.17 | [9 empty] | 0.18972
| 1 | 5 | 0.2 | 0.2 | 0.2 | 0.2 | 0.16 | 0.16 | [9 empty] | 127
| 1 | 6 | 0.2 | 0.2 | 0.2 | 0.2 | 0.15 | 0.16
| Mean |  | 0.21333 | 0.2 | 0.18333 | 0.2 | 0.15667 | 0.16833
| Median |  | 0.21 | 0.2 | 0.2 | 0.2 | 0.155 | 0.165
| Std Dev. |  | 0.015055 | 3.0405e-17 | 0.040825 | 3.0405e-17 | 0.008165 | 0.024833
| Rel.Std.Dev. |  | 7.057% | 1.52e-14% | 22.27% | 1.52e-14% | 5.212% | 14.75%
| PDM3 |  | 12.45% | 5.417% | -3.367% | 5.417% | -17.42% | -11.27%
| Z-Score (Absolute) |  | 0.97 | 0.37 | 0.37 | 0.37 | 1.57 | 1.05
| Analytical results for U in OREAS 25b (Certified Value 2.35 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 2.41 | 2 | 2.8474 | 2.7 | 2.1 | 2.1 | 2.4 | 2.3 | 2.3 | 2.47 | 2.34 | 2.3 | 2.19 | 2.31 |  | 1
| 1 | 2 | 2.47 | 2.2 | 2.3149 | 2.7 | 2.2 | 2.1 | 2.5 | 2.3 | 2.5 | 2.4 | 2.39 | 2.3 | 2.22 | 2.06 |  | 4
| 1 | 3 | 2.54 | 2.3 | 2.2368 | 2.8 | 2.3 | 2.1 | 2.5 | 2.3 | 2.5 | 2.45 | 2.38 | 2.4 | 2.2 | 2.13 |  | 16
| 1 | 4 | 2.35 | 2.2 | 2.7028 | 2.6 | 2.3 | 2.1 | 2.5 | 2.3 | 2.3 | 2.41 | 2.38 | 2.3 | 2.07 | 2.04 |  | 2.3475
| 1 | 5 | 2.45 | 2.3 | 2.6243 | 2.6 | 2.3 | 2.1 | 2.5 | 2.3 | 2.3 | 2.4 | 2.38 | 2.2 | 2.29 | 2.11 |  | 128
| 1 | 6 | 2.63 | 2.3 | 2.744 | 2.8 | 2.3 | 2.1 | 2.5 | 2.3 | 2.3 | 2.27 | 2.37 | 2.3 | 2.17 | 1.98
| Mean |  | 2.475 | 2.2167 | 2.5784 | 2.7 | 2.25 | 2.1 | 2.4833 | 2.3 | 2.3667 | 2.4 | 2.3733 | 2.3 | 2.19 | 2.105
| Median |  | 2.46 | 2.25 | 2.6635 | 2.7 | 2.3 | 2.1 | 2.5 | 2.3 | 2.3 | 2.405 | 2.38 | 2.3 | 2.195 | 2.085
| Std Dev. |  | 0.098742 | 0.1169 | 0.24639 | 0.089443 | 0.083666 | 0 | 0.040825 | 4.8648e-16 | 0.10328 | 0.069857 | 0.017512 | 0.063246 | 0.071833 | 0.11362
| Rel.Std.Dev. |  | 3.99% | 5.274% | 9.556% | 3.313% | 3.718% | 0% | 1.644% | 2.115e-14% | 4.364% | 2.911% | 0.7379% | 2.75% | 3.28% | 5.398%
| PDM3 |  | 5.433% | -5.571% | 9.837% | 15.02% | -4.152% | -10.54% | 5.788% | -2.022% | 0.8184% | 2.238% | 1.102% | -2.022% | -6.707% | -10.33%
| Z-Score (Absolute) |  | 0.74 | 0.61 | 1.28 | 1.91 | 0.44 | 1.22 | 0.78 | 0.17 | 0.17 | 0.35 | 0.21 | 0.17 | 0.75 | 1.19
| Analytical results for V in OREAS 25b (Certified Value 176 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 183 | 123 | 186 | 191.6781 | 43 | 177 | 177 | 177 | 175 | 168 | 169 | 166 | 180 | 184 | 170 | 1
| 1 | 2 | 177 | 144 | 186 | 191.2292 | 56 | 177 | 179 | 176 | 172 | 176 | 167 | 170 | 178 | 176 | 146 | #N/A
| 1 | 3 | 177 | 139 | 185 | 189.9701 | 34 | 177 | 180 | 179 | 174 | 175 | 169 | 166 | 179 | 185 | 158 | 16
| 1 | 4 | 180 | 138 | 183 | 189.5298 | 40 | 176 | 181 | 176 | 171 | 167 | 162 | 165 | 181 | 172 | 156 | 175.7634
| 1 | 5 | 178 | 142 | 186 | 190.2393 | 31 | 175 | 181 | 178 | 174 | 170 | 159 | 167 | 181 | 185 | 168 | 129
| 1 | 6 | 178 | 132 | 184 | 191.0999 | 44 | 175 | 182 | 174 | 171 | 164 | 165 | 168 | 180 | 172 | 167
| Mean |  | 178.8333 | 136.3333 | 185 | 190.6244 | 41.3333 | 176.1667 | 180 | 176.6667 | 172.8333 | 170 | 165.1667 | 167 | 179.8333 | 179 | 160.8333
| Median |  | 178 | 138.5 | 185.5 | 190.6696 | 41.5 | 176.5 | 180.5 | 176.5 | 173 | 169 | 166 | 166.5 | 180 | 180 | 162.5
| Std Dev. |  | 2.3166 | 7.7115 | 1.2649 | 0.83386 | 8.8015 | 0.98319 | 1.7889 | 1.7512 | 1.7224 | 4.6904 | 4.0208 | 1.7889 | 1.169 | 6.3875 | 9.2177
| Rel.Std.Dev. |  | 1.295% | 5.656% | 0.6837% | 0.4374% | 21.29% | 0.5581% | 0.9938% | 0.9912% | 0.9966% | 2.759% | 2.434% | 1.071% | 0.6501% | 3.568% | 5.731%
| PDM3 |  | 1.747% | -22.43% | 5.255% | 8.455% | -76.48% | 0.2294% | 2.41% | 0.5139% | -1.667% | -3.279% | -6.029% | -4.986% | 2.316% | 1.841% | -8.494%
| Z-Score (Absolute) |  | 0.29 | 4.36 | 0.97 | 1.58 | 14.74 | 0 | 0.42 | 0.05 | 0.36 | 0.67 | 1.2 | 1 | 0.4 | 0.31 | 1.68
| Analytical results for W in OREAS 25b (Certified Value 2.07 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 2.1 | 1.1 | < 10 | 2.1423 | < 0.1 | 1.7 | 1.9 | 2.2 | 2.1 | 2 | 1.5 | 2.1 | 2.2 | 2.4 | 2.2 | 1
| 1 | 2 | 2.1 | 1.1 | < 10 | 2.0062 | < 0.1 | 1.7 | 1.9 | 2.2 | 1.9 | 2.3 | 1.3 | 2.1 | 1.9 | 2.2 | 2.2 | 6
| 1 | 3 | 2.1 | 0.9 | < 10 | 2.0374 | < 0.1 | 1.8 | 1.9 | 2.2 | 2 | 2.2 | 1.2 | 1.9 | 1.9 | 2.4 | 2.4 | 16
| 1 | 4 | 2.1 | 1 | < 10 | 2.1392 | < 0.1 | 1.7 | 1.9 | 2.2 | 1.9 | 2 | 1.1 | 2.2 | 2 | 2.2 | 2.2 | 2.0687
| 1 | 5 | 2.1 | 1 | < 10 | 2.1673 | < 0.1 | 1.8 | 1.9 | 2.2 | 2 | 2.1 | 1 | 1.9 | 2.1 | 2.4 | 2.2 | 130
| 1 | 6 | 2.3 | 1.2 | < 10 | 2.2445 | < 0.1 | 1.8 | 1.9 | 2.3 | 2 | 2 | 1.1 | 2 | 1.9 | 2.3 | 2
| Mean |  | 2.1333 | 1.05 |  | 2.1228 |  | 1.75 | 1.9 | 2.2167 | 1.9833 | 2.1 | 1.2 | 2.0333 | 2 | 2.3167 | 2.2
| Median |  | 2.1 | 1.05 |  | 2.1407 |  | 1.75 | 1.9 | 2.2 | 2 | 2.05 | 1.15 | 2.05 | 1.95 | 2.35 | 2.2
| Std Dev. |  | 0.08165 | 0.10488 |  | 0.087575 |  | 0.054772 | 2.4324e-16 | 0.040825 | 0.075277 | 0.12649 | 0.17889 | 0.12111 | 0.12649 | 0.098319 | 0.12649
| Rel.Std.Dev. |  | 3.827% | 9.989% |  | 4.125% |  | 3.13% | 1.28e-14% | 1.842% | 3.795% | 6.023% | 14.91% | 5.956% | 6.325% | 4.244% | 5.75%
| PDM3 |  | 3.122% | -49.24% |  | 2.614% |  | -15.41% | -8.157% | 7.151% | -4.128% | 1.511% | -41.99% | -1.712% | -3.323% | 11.98% | 6.345%
| Z-Score (Absolute) |  | 0.4 | 3.98 | 12 | 0.36 | 8.02 | 1.15 | 0.54 | 0.74 | 0.2 | 0.27 | 3.37 | 0 | 0.13 | 1.15 | 0.67
| Analytical results for Y in OREAS 25b (Certified Value 11.8 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 1
1
| 1 | 1 | 12.08 | 7.4 | 11.1 | 9.9621 | 14.6 | 10.5 | 12.6 | 12.9 | 13.6 | 12.4 | 12.2 | 12.9 | 13.2 | 9.9 | 9.81 | 1
| 1 | 2 | 12.28 | 8.8 | 11.2 | 8.5762 | 14.7 | 10.9 | 13 | 12.8 | 13.3 | 11.6 | 12.1 | 13.1 | 13.7 | 9.7 | 8.3 | 7
| 1 | 3 | 15.11 | 9.3 | 11.5 | 8.9721 | 14.8 | 11.2 | 12.6 | 13.6 | 13.4 | 11.7 | 11.8 | 12.8 | 12.5 | 9.6 | 13.2 | 16
| 1 | 4 | 13.11 | 9 | 11.6 | 9.5693 | 14.8 | 11.1 | 12.2 | 13 | 28.9 | 11.8 | 11.5 | 12.5 | 11.5 | 9.7 | 10.5 | 11.7702
| 1 | 5 | 13.06 | 9.2 | 11.5 | 10.2926 | 14.3 | 10.9 | 12.2 | 12.4 | 13.7 | 11.9 | 11.9 | 12.8 | 13.9 | 10.3 | 10.7 | 131
| 1 | 6 | 14.83 | 7.8 | 11.6 | 8.7843 | 14.5 | 10.8 | 12.5 | 12.4 | 14.2 | 11.7 | 11.3 | 12.4 | 11.9 | 9.2 | 10.2
| Mean |  | 13.4117 | 8.5833 | 11.4167 | 9.3594 | 14.6167 | 10.9 | 12.5167 | 12.85 | 16.1833 | 11.85 | 11.8 | 12.75 | 12.7833 | 9.7333 | 10.4517
| Median |  | 13.085 | 8.9 | 11.5 | 9.2707 | 14.65 | 10.9 | 12.55 | 12.85 | 13.65 | 11.75 | 11.85 | 12.8 | 12.85 | 9.7 | 10.35
| Std Dev. |  | 1.2779 | 0.79099 | 0.2137 | 0.68883 | 0.19408 | 0.24495 | 0.29944 | 0.44609 | 6.2378 | 0.2881 | 0.34641 | 0.25884 | 0.97656 | 0.36148 | 1.5951
| Rel.Std.Dev. |  | 9.528% | 9.215% | 1.872% | 7.36% | 1.328% | 2.247% | 2.392% | 3.472% | 38.54% | 2.431% | 2.936% | 2.03% | 7.639% | 3.714% | 15.26%
| PDM3 |  | 13.95% | -27.08% | -3.003% | -20.48% | 24.18% | -7.393% | 6.342% | 9.174% | 37.49% | 0.6781% | 0.2533% | 8.325% | 8.608% | -17.31% | -11.2%
| Z-Score (Absolute) |  | 1.05 | 2.2 | 0.29 | 1.68 | 1.87 | 0.64 | 0.45 | 0.67 | 2.92 | 0 | 0.03 | 0.61 | 0.63 | 1.43 | 0.94
| Analytical results for Yb in OREAS 25b (Certified Value 1.1 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 14 | 15 | 16 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [8 empty] | 2
2
| 1 | 1 | 1.45 | 0.8 | 1.7 | 1 | 0.9 | 1.1 | 1.3 | [8 empty] | 1
| 1 | 2 | 1.19 | 1 | 1.6 | 1.1 | 1.2 | 1 | 0.9 | [8 empty] | #N/A
| 1 | 3 | 1.37 | 1 | 1.6 | 1.1 | 1.1 | 1 | 1.3 | [8 empty] | 16
| 1 | 4 | 1.3 | 1 | 1.6 | 1.1 | 1 | 1 | 1.1 | [8 empty] | 1.0964
| 1 | 5 | 1.48 | 1 | 1.6 | 1.1 | 0.8 | 1.1 | 1.2 | [8 empty] | 132
| 1 | 6 | 1.48 | 0.9 | 1.6 | 1.1 | 1 | 1 | 1
| Mean |  | 1.3783 | 0.95 | 1.6167 | 1.0833 | 1 | 1.0333 | 1.1333
| Median |  | 1.41 | 1 | 1.6 | 1.1 | 1 | 1 | 1.15
| Std Dev. |  | 0.11618 | 0.083666 | 0.040825 | 0.040825 | 0.14142 | 0.05164 | 0.1633
| Rel.Std.Dev. |  | 8.429% | 8.807% | 2.525% | 3.768% | 14.14% | 4.997% | 14.41%
| PDM3 |  | 25.72% | -13.35% | 47.45% | -1.191% | -8.791% | -5.751% | 3.37%
| Z-Score (Absolute) |  | 2.39 | 1.08 | 4.32 | 0 | 0.67 | 0.4 | 0.4
| Analytical results for Zn in OREAS 25b (Certified Value 66 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 66 | 65 | 85.4 | 69.343 | 70 | 62 | 66 | 68 | 66 | 64 | 62.9 | 66 | 64.1 | 64 | 58 | 1
| 1 | 2 | 66 | 67 | 84.8 | 66.2117 | 71 | 64.5 | 69 | 67 | 68 | 66 | 63.5 | 65 | 65.3 | 63 | 57 | 8
| 1 | 3 | 65 | 74 | 84.7 | 70.1114 | 71 | 66.4 | 65 | 68 | 65 | 66 | 65.2 | 66 | 66 | 63 | 55 | 16
| 1 | 4 | 65 | 63 | 84.3 | 66.6345 | 72 | 65.4 | 66 | 68 | 65 | 64 | 62.2 | 66 | 66.9 | 64 | 60 | 65.5427
| 1 | 5 | 69 | 65 | 83.8 | 66.4444 | 73 | 66.4 | 67 | 67 | 65 | 65 | 62.3 | 65 | 66.2 | 67 | 57 | 133
| 1 | 6 | 67 | 66 | 84.3 | 68.9275 | 70 | 64.3 | 66 | 65 | 64 | 64 | 62.2 | 65 | 67.4 | 61 | 57
| Mean |  | 66.3333 | 66.6667 | 84.55 | 67.9454 | 71.1667 | 64.8333 | 66.5 | 67.1667 | 65.5 | 64.8333 | 63.05 | 65.5 | 65.9833 | 63.6667 | 57.3333
| Median |  | 66 | 65.5 | 84.5 | 67.781 | 71 | 64.95 | 66 | 67.5 | 65 | 64.5 | 62.6 | 65.5 | 66.1 | 63.5 | 57
| Std Dev. |  | 1.5055 | 3.8297 | 0.54681 | 1.708 | 1.169 | 1.6525 | 1.3784 | 1.169 | 1.3784 | 0.98319 | 1.1709 | 0.54772 | 1.1754 | 1.9664 | 1.633
| Rel.Std.Dev. |  | 2.27% | 5.745% | 0.6467% | 2.514% | 1.643% | 2.549% | 2.073% | 1.741% | 2.104% | 1.516% | 1.857% | 0.8362% | 1.781% | 3.089% | 2.848%
| PDM3 |  | 1.206% | 1.715% | 29% | 3.666% | 8.581% | -1.082% | 1.461% | 2.478% | -0.06511% | -1.082% | -3.803% | -0.06511% | 0.6723% | -2.862% | -12.53%
| Z-Score (Absolute) |  | 0.21 | 0.4 | 10.89 | 1.15 | 3.04 | 0.67 | 0.3 | 0.69 | 0.28 | 0.67 | 1.72 | 0.28 | 0 | 1.36 | 5.07
| Analytical results for Zr in OREAS 25b (Certified Value 181 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES |  | 0
0
| 1 | 1 | 186.4 | 137.8 | 168 | 182.3943 | 77 | 160.5 | 184.5 | 189.5 | 199.5 | 179.5 | 181 | 181 | 176.9 | 183 |  | 1
| 1 | 2 | 180.7 | 145.1 | 168 | 187.9211 | 100 | 164.4 | 187 | 185 | 199.5 | 184 | 174 | 184 | 188.5 | 180 |  | #N/A
| 1 | 3 | 183.4 | 146.4 | 168 | 180.3347 | 69 | 167.8 | 190 | 188.5 | 197 | 179.5 | 170 | 181 | 181.2 | 182 |  | 16
| 1 | 4 | 183.9 | 138.8 | 167 | 190.0926 | 84 | 166.9 | 184.5 | 188.5 | 194 | 182 | 162 | 181 | 176.7 | 181 |  | 181.0367
| 1 | 5 | 189.8 | 147.5 | 167 | 181.94 | 65 | 162.5 | 189 | 189.5 | 202 | 183.5 | 220 | 181 | 180.3 | 183 |  | 134
| 1 | 6 | 189 | 139.3 | 171 | 188.2613 | 100 | 162 | 189 | 179.5 | 192.5 | 195 | 160 | 184 | 180.5 | 188
| Mean |  | 185.5333 | 142.4833 | 168.1667 | 185.1573 | 82.5 | 164.0167 | 187.3333 | 186.75 | 197.4167 | 183.9167 | 177.8333 | 182 | 180.6833 | 182.8333
| Median |  | 185.15 | 142.2 | 168 | 185.1577 | 80.5 | 163.45 | 188 | 188.5 | 198.25 | 182.75 | 172 | 181 | 180.4 | 182.5
| Std Dev. |  | 3.5087 | 4.3125 | 1.472 | 4.0712 | 15.0566 | 2.8813 | 2.4014 | 3.9211 | 3.6251 | 5.7569 | 22.0583 | 1.5492 | 4.2832 | 2.7869
| Rel.Std.Dev. |  | 1.891% | 3.027% | 0.8753% | 2.199% | 18.25% | 1.757% | 1.282% | 2.1% | 1.836% | 3.13% | 12.4% | 0.8512% | 2.371% | 1.524%
| PDM3 |  | 2.484% | -21.3% | -7.109% | 2.276% | -54.43% | -9.401% | 3.478% | 3.156% | 9.048% | 1.591% | -1.769% | 0.5321% | -0.1952% | 0.9924%
| Z-Score (Absolute) |  | 0.47 | 6.04 | 2.16 | 0.41 | 15.11 | 2.78 | 0.74 | 0.66 | 2.27 | 0.23 | 0.69 | 0.06 | 0.26 | 0.06
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C904: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C922: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C940: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C959: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C977: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C995: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1013: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1031: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 25b (Certified Value 0.047 ppm) | [16 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.04 | 0.04 | < 0.5 | < 0.05 | 0.024 | < 0.1 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.05 | 0.058 | 0.25 | 0.1 | 1
| 1 | 2 | 0.06 | 0.04 | < 0.5 | < 0.05 | 0.131 | < 0.1 | 0.02 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | 0.069 | 0.23 | < 0.1 | 13
| 1 | 3 | 0.04 | 0.05 | < 0.5 | < 0.05 | 0.07 | < 0.1 | 0.03 | 0.04 | 0.03 | 0.04 | 0.04 | < 0.05 | 0.058 | 0.27 | < 0.1 | 16
| 1 | 4 | 0.04 | 0.05 | < 0.5 | < 0.05 | 0.236 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.06 | 0.04 | < 0.05 | 0.069 | 0.27 | < 0.1 | 0.047068
| 1 | 5 | 0.06 | 0.05 | < 0.5 | < 0.05 | 0.027 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | 0.069 | 0.24 | < 0.1 | 136
| 1 | 6 | 0.04 | 0.05 | < 0.5 | < 0.05 | 0.054 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | 0.069 | 0.49 | < 0.1
| Mean |  | 0.046667 | 0.046667 |  |  | 0.090333 |  | 0.03 | 0.031667 | 0.031667 | 0.036667 | 0.04 | 0.05 | 0.065333 | 0.29167 | 0.1
| Median |  | 0.04 | 0.05 |  |  | 0.062 |  | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | 0.069 | 0.26 | 0.1
| Std Dev. |  | 0.010328 | 0.005164 |  |  | 0.081262 |  | 0.0063246 | 0.0040825 | 0.0040825 | 0.012111 | 0 | 0 | 0.0056804 | 0.098472
| Rel.Std.Dev. |  | 22.13% | 11.07% |  |  | 89.96% |  | 21.08% | 12.89% | 12.89% | 33.03% | 0% | 0% | 8.694% | 33.76%
| PDM3 |  | -0.8531% | -0.8531% |  |  | 91.92% |  | -36.26% | -32.72% | -32.72% | -22.1% | -15.02% | 6.229% | 38.81% | 519.7% | 112.5%
| Z-Score (Absolute) |  | 0 | 0 | 9.14 | 0.97 | 1.96 | 0.15 | 0.75 | 0.67 | 0.67 | 0.45 | 0.3 | 0.22 | 0.84 | 11.01 | 0.52
| Analytical results for Al in OREAS 25b (Certified Value 4.26 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 4.496 | 3.32 | 4.28 | 5.431 | 4.92 | 4.47 | 4.2 | 3.97 | 4.31 | 4.06 | 3 | 5.48 | 4.34 | 3.78 | 4.04 | 1
| 1 | 2 | 4.497 | 3.19 | 4.26 | 5.8369 | 4.85 | 4.39 | 3.91 | 3.83 | 4.27 | 4.18 | 3.05 | 5.42 | 4.33 | 3.778 | 3.93 | #N/A
| 1 | 3 | 4.418 | 3.21 | 4.26 | 5.3915 | 4.9 | 4.48 | 4.08 | 4.03 | 4.34 | 4.1 | 2.83 | 5.55 | 4.04 | 3.748 | 3.95 | 16
| 1 | 4 | 4.42 | 3.2 | 4.29 | 5.7602 | 4.6 | 4.66 | 4.03 | 4.09 | 4.29 | 4.18 | 2.95 | 5.64 | 4.55 | 3.739 | 4.05 | 4.2611
| 1 | 5 | 4.527 | 3.1 | 4.27 | 5.8871 | 4.88 | 4.52 | 4.08 | 4.22 | 4.38 | 3.97 | 3.08 | 5.53 | 4.69 | 3.69 | 4.39 | 137
| 1 | 6 | 4.428 | 3.05 | 4.26 | 5.5828 | 4.86 | 4.33 | 4.04 | 4.08 | 4.31 | 4.16 | 3.05 | 5.45 | 4.29 | 3.897 | 4.34
| Mean |  | 4.4643 | 3.1783 | 4.27 | 5.6483 | 4.835 | 4.475 | 4.0567 | 4.0367 | 4.3167 | 4.1083 | 2.9933 | 5.5117 | 4.3733 | 3.772 | 4.1167
| Median |  | 4.462 | 3.195 | 4.265 | 5.6715 | 4.87 | 4.475 | 4.06 | 4.055 | 4.31 | 4.13 | 3.025 | 5.505 | 4.335 | 3.763 | 4.045
| Std Dev. |  | 0.047811 | 0.09411 | 0.012649 | 0.21101 | 0.11794 | 0.11362 | 0.09395 | 0.13079 | 0.038816 | 0.083046 | 0.092232 | 0.079352 | 0.22474 | 0.069446 | 0.19876
| Rel.Std.Dev. |  | 1.071% | 2.961% | 0.2962% | 3.736% | 2.439% | 2.539% | 2.316% | 3.24% | 0.8992% | 2.021% | 3.081% | 1.44% | 5.139% | 1.841% | 4.828%
| PDM3 |  | 4.77% | -25.41% | 0.2098% | 32.56% | 13.47% | 5.021% | -4.797% | -5.266% | 1.305% | -3.584% | -29.75% | 29.35% | 2.635% | -11.48% | -3.389%
| Z-Score (Absolute) |  | 0.61 | 3.45 | 0 | 4.36 | 1.79 | 0.65 | 0.67 | 0.74 | 0.15 | 0.51 | 4.04 | 3.92 | 0.33 | 1.57 | 0.48
| Analytical results for As in OREAS 25b (Certified Value 2.16 ppm) | [16 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 2.34 | 3 | < 2 | 3.6819 | 1.3 | 2 | 2.4 | 2.7 | 2.8 | 2.3 | 4.5 | 3 | 0.9 | 3 | 2 | 1
| 1 | 2 | 2.34 | 3 | < 2 | 4.0192 | 1.2 | 1.8 | 2.5 | 2.1 | 2.7 | 2.1 | 4.4 | 3 | 0.9 | 2 | 2 | 1
| 1 | 3 | 2.13 | 3 | < 2 | 3.5513 | 1.2 | 2 | 2.4 | 2.6 | 2.7 | 1.9 | 4.4 | 3 | 1.1 | 2 | 3 | 16
| 1 | 4 | 2.24 | 3 | < 2 | 3.4818 | 1.2 | 2 | 2.5 | 2.3 | 2.8 | 2.1 | 4.5 | 3 | 1.2 | 3 | 3 | 2.1552
| 1 | 5 | 2.3 | 3 | < 2 | 3.8916 | 1.2 | 1.9 | 2.6 | 2.5 | 2.6 | 1.9 | 4.4 | 3 | 0.6 | 2 | 2 | 138
| 1 | 6 | 2.17 | 3 | < 2 | 3.2859 | 1 | 1.8 | 2.5 | 2.4 | 2.7 | 2.3 | 4.3 | 3 | 0.5 | 2 | 2
| Mean |  | 2.2533 | 3 |  | 3.6519 | 1.1833 | 1.9167 | 2.4833 | 2.4333 | 2.7167 | 2.1 | 4.4167 | 3 | 0.86667 | 2.3333 | 2.3333
| Median |  | 2.27 | 3 |  | 3.6166 | 1.2 | 1.95 | 2.5 | 2.45 | 2.7 | 2.1 | 4.4 | 3 | 0.9 | 2 | 2
| Std Dev. |  | 0.088919 | 0 |  | 0.27068 | 0.098319 | 0.098319 | 0.075277 | 0.21602 | 0.075277 | 0.17889 | 0.075277 | 0 | 0.27325 | 0.5164 | 0.5164
| Rel.Std.Dev. |  | 3.946% | 0% |  | 7.412% | 8.309% | 5.13% | 3.031% | 8.878% | 2.771% | 8.518% | 1.704% | 0% | 31.53% | 22.13% | 22.13%
| PDM3 |  | 4.551% | 39.2% |  | 69.44% | -45.1% | -11.07% | 15.22% | 12.9% | 26.05% | -2.563% | 104.9% | 39.2% | -59.79% | 8.263% | 8.263%
| Z-Score (Absolute) |  | 0 | NA | 1.82 | 2.04 | 1.56 | 0.49 | 0.33 | 0.26 | 0.67 | 0.22 | 3.15 | NA | 2.02 | NA | NA
Results from laboratories 2, 12, 15 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Au in OREAS 25b (Indicative Value 1.3 ppb) | [16 empty] | Indicative
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 14 | [6 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 0.9 | 2.0334 | < 0.5 | 1.8 | < 20 | < 20 | < 20 | < 20 | < 0.2 | [6 empty] | 1
| 1 | 2 | 1.2 | 1.5282 | < 0.5 | 0.8 | < 20 | < 20 | < 20 | < 20 | < 0.2 | [6 empty] | 21
| 1 | 3 | 0.9 | 1.5783 | < 0.5 | 0.7 | < 20 | < 20 | < 20 | < 20 | 9 | [6 empty] | 16
| 1 | 4 | 0.8 | 1.4482 | < 0.5 | 1.7 | < 20 | < 20 | < 20 | < 20 | < 0.2 | [6 empty] | 1.339
| 1 | 5 | 1 | 2.1658 | < 0.5 | 1.8 | < 20 | < 20 | < 20 | < 20 | < 0.2 | [6 empty] | 27
| 1 | 6 | 1 | 1.8474 | < 0.5 | 0.9 | < 20 | < 20 | < 20 | < 20 | < 0.2
| Mean |  | 0.96667 | 1.7669 |  | 1.2833 |  |  |  |  | 9
| Median |  | 0.95 | 1.7129 |  | 1.3 |  |  |  |  | 9
| Std Dev. |  | 0.13663 | 0.2935 |  | 0.53448
| Rel.Std.Dev. |  | 14.13% | 16.61% |  | 41.65%
| PDM3 |  | -27.8% | 31.96% |  | -4.155% |  |  |  |  | 572.2%
| Z-Score (Absolute) |  | 0.36 | 0 | 0.67 | 0.21 | 3.66 | 3.66 | 3.66 | 3.66 | 0.08
| Analytical results for B in OREAS 25b (Certified Value < 10 ppm) | [16 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | < 5 | < 10 | 49.3 | < 10 | 5 | < 20 | < 10 | 10 | < 10 | < 10 | 2.6 | < 20 |  |  |  | 1
| 1 | 2 | < 5 | < 10 | 49.1 | < 10 | 6 | < 20 | < 10 | 10 | < 10 | < 10 | 2.5 | < 20 |  |  |  | 15
| 1 | 3 | < 5 | < 10 | 49.9 | < 10 | 6 | < 20 | < 10 | 10 | < 10 | < 10 | 2.3 | < 20 |  |  |  | 16
| 1 | 4 | < 5 | < 10 | 49.4 | < 10 | 6 | < 20 | < 10 | 10 | < 10 | < 10 | 2.6 | < 20 |  |  |  | < 10
| 1 | 5 | < 5 | < 10 | 46.9 | < 10 | 6 | < 20 | < 10 | 10 | < 10 | < 10 | 2.6 | < 20 |  |  |  | 139
| 1 | 6 | < 5 | < 10 | 49.5 | < 10 | 6 | < 20 | < 10 | 10 | < 10 | < 10 | 2.2 | < 20
| Mean |  |  |  | 49.0167 |  | 5.8333 |  |  | 10 |  |  | 2.4667
| Median |  |  |  | 49.35 |  | 6 |  |  | 10 |  |  | 2.55
| Std Dev. |  |  |  | 1.0704 |  | 0.40825 |  |  | 0 |  |  | 0.17512
| Rel.Std.Dev. |  |  |  | 2.184% |  | 6.999% |  |  | 0% |  |  | 7.099%
PDM3
| Z-Score (Absolute) |  | 1.01 | 0 | 17.81 | 0 | 0.34 | 2.02 | 0 | 2.02 | 0 | 0 | 1.02 | 2.02
| Analytical results for Ba in OREAS 25b (Certified Value 68 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 1
| 1 | 1 | 68.4 | 58 | 66.6 | 72.3067 | 63.4 | 67.6 | 70 | 70 | 70 | 70 | 64 | 82 | 65.3 | 67 | 68 | 1
| 1 | 2 | 69.46 | 57 | 66.6 | 67.5536 | 65.1 | 69.3 | 70 | 70 | 70 | 70 | 64 | 81 | 62.7 | 70 | 67 | #N/A
| 1 | 3 | 68.26 | 57 | 66.2 | 70.3418 | 63.3 | 69.3 | 70 | 70 | 70 | 70 | 62 | 81 | 60.7 | 68 | 65 | 16
| 1 | 4 | 68.8 | 58 | 66.3 | 68.186 | 61 | 71.5 | 70 | 70 | 70 | 70 | 65 | 81 | 66 | 65 | 68 | 68.227
| 1 | 5 | 68.82 | 56 | 65.1 | 70.643 | 63.1 | 71.9 | 70 | 70 | 70 | 70 | 67 | 81 | 67.4 | 63 | 70 | 140
| 1 | 6 | 68.99 | 56 | 65.6 | 71.9852 | 64.9 | 70.4 | 70 | 70 | 70 | 70 | 67 | 81 | 64.1 | 66 | 70
| Mean |  | 68.7883 | 57 | 66.0667 | 70.1694 | 63.4667 | 70 | 70 | 70 | 70 | 70 | 64.8333 | 81.1667 | 64.3667 | 66.5 | 68
| Median |  | 68.81 | 57 | 66.25 | 70.4924 | 63.35 | 69.85 | 70 | 70 | 70 | 70 | 64.5 | 81 | 64.7 | 66.5 | 68
| Std Dev. |  | 0.42962 | 0.89443 | 0.59889 | 1.944 | 1.4814 | 1.5975 | 0 | 0 | 0 | 0 | 1.9408 | 0.40825 | 2.4097 | 2.429 | 1.8974
| Rel.Std.Dev. |  | 0.6246% | 1.569% | 0.9065% | 2.77% | 2.334% | 2.282% | 0% | 0% | 0% | 0% | 2.994% | 0.503% | 3.744% | 3.653% | 2.79%
| PDM3 |  | 0.8227% | -16.46% | -3.166% | 2.847% | -6.977% | 2.599% | 2.599% | 2.599% | 2.599% | 2.599% | -4.974% | 18.97% | -5.658% | -2.531% | -0.3328%
| Z-Score (Absolute) |  | 0 | 5.76 | 1.33 | 0.67 | 2.6 | 0.59 | 0.59 | 0.59 | 0.59 | 0.59 | 1.93 | 6.04 | 2.16 | 1.12 | 0.38
| Analytical results for Be in OREAS 25b (Certified Value 0.76 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  |  | 3
| 1 | 1 | 0.82 | 0.8 | 0.80869 | 0.7 | 0.74 | 0.75 | 0.79 | 0.72 | 0.79 | 0.85 | 0.66 | 0.8 |  |  |  | 1
| 1 | 2 | 0.8 | 0.8 | 0.75871 | 0.8 | 0.7 | 0.73 | 0.77 | 0.72 | 0.67 | 0.81 | 0.75 | 0.8 |  |  |  | #N/A
| 1 | 3 | 0.79 | 0.8 | 0.80956 | 0.7 | 0.71 | 0.74 | 0.78 | 0.73 | 0.74 | 0.83 | 0.75 | 0.8 |  |  |  | 16
| 1 | 4 | 0.82 | 0.7 | 0.74771 | 0.7 | 0.72 | 0.75 | 0.78 | 0.74 | 0.79 | 0.82 | 0.75 | 0.8 |  |  |  | 0.76387
| 1 | 5 | 0.81 | 0.7 | 0.77873 | 0.7 | 0.72 | 0.77 | 0.79 | 0.71 | 0.84 | 0.8 | 0.79 | 0.8 |  |  |  | 141
| 1 | 6 | 0.79 | 0.8 | 0.78681 | 0.7 | 0.72 | 0.75 | 0.77 | 0.73 | 0.58 | 0.88 | 0.79 | 0.9
| Mean |  | 0.805 | 0.76667 | 0.7817 | 0.71667 | 0.71833 | 0.74833 | 0.78 | 0.725 | 0.735 | 0.83167 | 0.74833 | 0.81667
| Median |  | 0.805 | 0.8 | 0.78277 | 0.7 | 0.72 | 0.75 | 0.78 | 0.725 | 0.765 | 0.825 | 0.75 | 0.8
| Std Dev. |  | 0.013784 | 0.05164 | 0.025391 | 0.040825 | 0.013292 | 0.013292 | 0.0089443 | 0.010488 | 0.095237 | 0.029269 | 0.047504 | 0.040825
| Rel.Std.Dev. |  | 1.712% | 6.736% | 3.248% | 5.696% | 1.85% | 1.776% | 1.147% | 1.447% | 12.96% | 3.519% | 6.348% | 4.999%
| PDM3 |  | 5.384% | 0.3661% | 2.335% | -6.18% | -5.961% | -2.034% | 2.112% | -5.089% | -3.779% | 8.875% | -2.034% | 6.912%
| Z-Score (Absolute) |  | 1.24 | NA | 0.73 | 0.69 | 0.66 | 0 | 0.69 | 0.51 | 0.29 | 1.82 | 0 | NA
Results from laboratories 2 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 25b (Certified Value 0.34 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.4g |  |  | 3
| 1 | 1 | 0.38 | 0.33 | 0.3595 | 0.35 | 0.37 | 0.32 | 0.35 | 0.32 | 0.33 | 0.35 | 0.37 | 0.3 | 0.3 |  |  | 1
| 1 | 2 | 0.41 | 0.33 | 0.36581 | 0.32 | 0.37 | 0.31 | 0.34 | 0.32 | 0.34 | 0.35 | 0.36 | 0.3 | 0.3 |  |  | 2
| 1 | 3 | 0.38 | 0.33 | 0.34974 | 0.33 | 0.37 | 0.31 | 0.35 | 0.33 | 0.35 | 0.33 | 0.36 | 0.3 | 0.3 |  |  | 16
| 1 | 4 | 0.39 | 0.3 | 0.33853 | 0.32 | 0.39 | 0.31 | 0.36 | 0.32 | 0.34 | 0.35 | 0.36 | 0.32 | 0.4 |  |  | 0.34354
| 1 | 5 | 0.38 | 0.38 | 0.32776 | 0.33 | 0.39 | 0.29 | 0.36 | 0.3 | 0.34 | 0.36 | 0.35 | 0.35 | 0.4 |  |  | 142
| 1 | 6 | 0.4 | 0.33 | 0.33963 | 0.34 | 0.38 | 0.31 | 0.39 | 0.32 | 0.35 | 0.35 | 0.37 | 0.32 | 0.4
| Mean |  | 0.39 | 0.33333 | 0.34683 | 0.33167 | 0.37833 | 0.30833 | 0.35833 | 0.31833 | 0.34167 | 0.34833 | 0.36167 | 0.315 | 0.35
| Median |  | 0.385 | 0.33 | 0.34469 | 0.33 | 0.375 | 0.31 | 0.355 | 0.32 | 0.34 | 0.35 | 0.36 | 0.31 | 0.35
| Std Dev. |  | 0.012649 | 0.02582 | 0.014237 | 0.01169 | 0.0098319 | 0.0098319 | 0.017224 | 0.0098319 | 0.0075277 | 0.0098319 | 0.0075277 | 0.019748 | 0.054772
| Rel.Std.Dev. |  | 3.243% | 7.746% | 4.105% | 3.525% | 2.599% | 3.189% | 4.807% | 3.089% | 2.203% | 2.823% | 2.081% | 6.269% | 15.65%
| PDM3 |  | 13.52% | -2.971% | 0.9571% | -3.457% | 10.13% | -10.25% | 4.306% | -7.338% | -0.5457% | 1.395% | 5.276% | -8.308% | 1.88%
| Z-Score (Absolute) |  | 1.96 | 0.47 | 0.11 | 0.54 | 1.46 | 1.54 | 0.6 | 1.11 | 0.11 | 0.17 | 0.75 | 1.25 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 25b (Certified Value 0.221 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.217 | 0.21 | 0.22 | 0.21233 | 0.2 | 0.23 | 0.23 | 0.23 | 0.22 | 0.22 | 0.23 | 0.2 | 0.23 | 0.149 | 0.224 | 1
| 1 | 2 | 0.216 | 0.21 | 0.22 | 0.2183 | 0.21 | 0.23 | 0.22 | 0.23 | 0.22 | 0.23 | 0.22 | 0.2 | 0.21 | 0.153 | 0.223 | #N/A
| 1 | 3 | 0.211 | 0.21 | 0.22 | 0.20625 | 0.19 | 0.23 | 0.23 | 0.23 | 0.23 | 0.23 | 0.22 | 0.2 | 0.21 | 0.153 | 0.222 | 16
| 1 | 4 | 0.211 | 0.21 | 0.23 | 0.22132 | 0.19 | 0.23 | 0.23 | 0.23 | 0.23 | 0.23 | 0.22 | 0.2 | 0.23 | 0.147 | 0.226 | 0.22081
| 1 | 5 | 0.218 | 0.2 | 0.23 | 0.22034 | 0.2 | 0.23 | 0.23 | 0.23 | 0.23 | 0.22 | 0.23 | 0.2 | 0.23 | 0.146 | 0.228 | 143
| 1 | 6 | 0.214 | 0.2 | 0.22 | 0.20537 | 0.2 | 0.23 | 0.23 | 0.23 | 0.22 | 0.23 | 0.23 | 0.2 | 0.22 | 0.151 | 0.229
| Mean |  | 0.2145 | 0.20667 | 0.22333 | 0.21398 | 0.19833 | 0.23 | 0.22833 | 0.23 | 0.225 | 0.22667 | 0.225 | 0.2 | 0.22167 | 0.14983 | 0.22533
| Median |  | 0.215 | 0.21 | 0.22 | 0.21532 | 0.2 | 0.23 | 0.23 | 0.23 | 0.225 | 0.23 | 0.225 | 0.2 | 0.225 | 0.15 | 0.225
| Std Dev. |  | 0.0030166 | 0.005164 | 0.005164 | 0.0070642 | 0.0075277 | 0 | 0.0040825 | 0 | 0.0054772 | 0.005164 | 0.0054772 | 3.0405e-17 | 0.0098319 | 0.0029944 | 0.0028048
| Rel.Std.Dev. |  | 1.406% | 2.499% | 2.312% | 3.301% | 3.795% | 0% | 1.788% | 0% | 2.434% | 2.278% | 2.434% | 1.52e-14% | 4.435% | 1.999% | 1.245%
| PDM3 |  | -2.856% | -6.404% | 1.144% | -3.089% | -10.18% | 4.164% | 3.409% | 4.164% | 1.899% | 2.654% | 1.899% | -9.423% | 0.3895% | -32.14% | 2.05%
| Z-Score (Absolute) |  | 0.89 | 1.69 | 0 | 0.95 | 2.53 | 0.67 | 0.51 | 0.67 | 0.17 | 0.34 | 0.17 | 2.36 | 0.17 | 7.43 | 0.2
| Analytical results for Cd in OREAS 25b (Certified Value 0.043 ppm) | [16 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.4g |  | 3
| 1 | 1 | 0.041 | 0.04 | 0.92 | 0.060744 | 0.05 | < 0.1 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.06 | 0.09 | 0.1 |  | 1
| 1 | 2 | 0.044 | 0.03 | 0.9 | 0.051613 | 0.06 | < 0.1 | 0.04 | 0.03 | 0.03 | 0.03 | 0.04 | 0.06 | 0.09 | < 0.1 |  | 17
| 1 | 3 | 0.04 | 0.05 | 0.9 | 0.063869 | 0.03 | < 0.1 | 0.05 | 0.03 | 0.04 | 0.04 | 0.03 | 0.05 | 0.09 | 0.1 |  | 16
| 1 | 4 | 0.044 | 0.05 | 0.9 | 0.053555 | 0.04 | < 0.1 | 0.04 | 0.03 | 0.04 | 0.03 | 0.05 | 0.06 | 0.09 | < 0.1 |  | 0.043228
| 1 | 5 | 0.042 | 0.05 | 0.9 | < 0.05 | 0.09 | < 0.1 | 0.04 | 0.02 | 0.04 | 0.03 | 0.05 | 0.05 | 0.09 | 0.1 |  | 144
| 1 | 6 | 0.04 | 0.05 | 0.89 | < 0.05 | 0.06 | < 0.1 | 0.03 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06 | 0.1 | < 0.1
| Mean |  | 0.041833 | 0.045 | 0.90167 | 0.057445 | 0.055 |  | 0.04 | 0.028333 | 0.036667 | 0.035 | 0.043333 | 0.056667 | 0.091667 | 0.1
| Median |  | 0.0415 | 0.05 | 0.9 | 0.05715 | 0.055 |  | 0.04 | 0.03 | 0.04 | 0.035 | 0.045 | 0.06 | 0.09 | 0.1
| Std Dev. |  | 0.0018348 | 0.0083666 | 0.0098319 | 0.0058105 | 0.020736 |  | 0.0063246 | 0.0075277 | 0.005164 | 0.0054772 | 0.008165 | 0.005164 | 0.0040825 | 1.6997e-17
| Rel.Std.Dev. |  | 4.386% | 18.59% | 1.09% | 10.11% | 37.7% |  | 15.81% | 26.57% | 14.08% | 15.65% | 18.84% | 9.113% | 4.454% | 1.7e-14%
| PDM3 |  | -3.226% | 4.1% | 1986% | 32.89% | 27.23% |  | -7.467% | -34.46% | -15.18% | -19.03% | 0.244% | 31.09% | 112.1% | 131.3%
| Z-Score (Absolute) |  | 0.26 | 0 | 69.32 | 0.13 | 0.81 | 0.4 | 0.4 | 1.35 | 0.67 | 0.81 | 0.13 | 0.94 | 3.78 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 25b (Certified Value 32.6 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  |  | 2
| 1 | 1 | 34.606 | 27.2 | 32.8567 | 31 | 31.8 | 31.3 | 33.8 | 33.2 | 31.8 | 34.5 | 33.44 | 28.1 |  |  |  | 1
| 1 | 2 | 34.347 | 26.29 | 33.9319 | 30.9 | 31.2 | 31.2 | 33.2 | 33.6 | 31.3 | 34.4 | 35.08 | 28.4 |  |  |  | 3
| 1 | 3 | 33.51 | 26.33 | 33.6581 | 30.2 | 31 | 31.3 | 35.5 | 33.1 | 30.6 | 34 | 33.07 | 26.6 |  |  |  | 16
| 1 | 4 | 33.861 | 25.64 | 32.6945 | 29.7 | 31.5 | 32.2 | 34.9 | 33.9 | 31.8 | 34.4 | 33.89 | 32.4 |  |  |  | 32.6171
| 1 | 5 | 34.225 | 25.07 | 33.2251 | 30.4 | 31.5 | 32.3 | 34.6 | 31.9 | 32.7 | 35 | 33.29 | 30 |  |  |  | 145
| 1 | 6 | 33.921 | 24.31 | 32.1547 | 30.7 | 31.8 | 32.2 | 33.8 | 34.5 | 32.4 | 33.9 | 34.29 | 32.5
| Mean |  | 34.0783 | 25.8067 | 33.0868 | 30.4833 | 31.4667 | 31.75 | 34.3 | 33.3667 | 31.7667 | 34.3667 | 33.8433 | 29.6667
| Median |  | 34.073 | 25.965 | 33.0409 | 30.55 | 31.5 | 31.75 | 34.2 | 33.4 | 31.8 | 34.4 | 33.665 | 29.2
| Std Dev. |  | 0.39169 | 1.0256 | 0.65351 | 0.48751 | 0.32042 | 0.53198 | 0.84853 | 0.88015 | 0.75542 | 0.39328 | 0.74806 | 2.4114
| Rel.Std.Dev. |  | 1.149% | 3.974% | 1.975% | 1.599% | 1.018% | 1.676% | 2.474% | 2.638% | 2.378% | 1.144% | 2.21% | 8.128%
| PDM3 |  | 4.48% | -20.88% | 1.44% | -6.542% | -3.527% | -2.659% | 5.159% | 2.298% | -2.607% | 5.364% | 3.759% | -9.046%
| Z-Score (Absolute) |  | 0.73 | 2.91 | 0.29 | 0.85 | 0.42 | 0.3 | 0.82 | 0.41 | 0.29 | 0.85 | 0.62 | 1.21
| Analytical results for Co in OREAS 25b (Certified Value 6.4 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 6.44 | 5.1 | 7.9 | 8.4642 | 7.5 | 6.3 | 5.8 | 5.4 | 6 | 5.9 | 5 | 8.7 | 6.6 | 5.5 | 5 | 1
| 1 | 2 | 6.37 | 4.9 | 7.9 | 8.3774 | 7.5 | 6.2 | 5.5 | 5.3 | 5.9 | 6.2 | 4.9 | 9 | 6.7 | 5.7 | 5 | 5
| 1 | 3 | 6.34 | 4.9 | 7.9 | 8.076 | 7 | 6.3 | 5.7 | 5.4 | 6.1 | 6.2 | 4.5 | 8.6 | 6.5 | 5.2 | 4 | 16
| 1 | 4 | 6.26 | 4.9 | 7.9 | 7.6097 | 6.9 | 6.5 | 5.6 | 5.6 | 6.1 | 6.4 | 4.9 | 8.6 | 6.7 | 5.6 | 7 | 6.4047
| 1 | 5 | 6.31 | 4.7 | 7.6 | 7.8793 | 7 | 6.4 | 5.7 | 5.6 | 5.9 | 5.9 | 5 | 8.7 | 7.3 | 5.3 | 6 | 146
| 1 | 6 | 6.29 | 4.6 | 7.7 | 7.4963 | 7.1 | 6.2 | 5.6 | 5.5 | 6 | 6.3 | 5 | 9 | 6.8 | 5.7 | 5
| Mean |  | 6.335 | 4.85 | 7.8167 | 7.9838 | 7.1667 | 6.3167 | 5.65 | 5.4667 | 6 | 6.15 | 4.8833 | 8.7667 | 6.7667 | 5.5 | 5.3333
| Median |  | 6.325 | 4.9 | 7.9 | 7.9777 | 7.05 | 6.3 | 5.65 | 5.45 | 6 | 6.2 | 4.95 | 8.7 | 6.7 | 5.55 | 5
| Std Dev. |  | 0.064109 | 0.17607 | 0.13292 | 0.39567 | 0.26583 | 0.1169 | 0.10488 | 0.12111 | 0.089443 | 0.20736 | 0.19408 | 0.18619 | 0.28048 | 0.20976 | 1.0328
| Rel.Std.Dev. |  | 1.012% | 3.63% | 1.7% | 4.956% | 3.709% | 1.851% | 1.856% | 2.215% | 1.491% | 3.372% | 3.974% | 2.124% | 4.145% | 3.814% | 19.36%
| PDM3 |  | -1.088% | -24.27% | 22.05% | 24.66% | 11.9% | -1.374% | -11.78% | -14.65% | -6.318% | -3.976% | -23.75% | 36.88% | 5.652% | -14.13% | -16.73%
| Z-Score (Absolute) |  | 0.09 | 1.24 | 1.42 | 1.57 | 0.84 | 0.07 | 0.52 | 0.69 | 0.21 | 0.07 | 1.21 | 2.28 | 0.48 | 0.66 | NA
Results from laboratory 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 25b (Certified Value 140 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 0
| 1 | 1 | 165.3 | 131 | 139 | 148.0595 | 137 | 153 | 143 | 147 | 151 | 138 | 135 | 132 | 144.3 | 140 | 139 | 1
| 1 | 2 | 160.9 | 131 | 139 | 145.3433 | 134 | 151 | 136 | 143 | 149 | 140 | 134 | 131 | 136.8 | 141 | 138 | #N/A
| 1 | 3 | 160 | 131 | 141 | 146.3556 | 135 | 152.9 | 139 | 145 | 149 | 140 | 130 | 131 | 134 | 140 | 136 | 16
| 1 | 4 | 158.4 | 129 | 142 | 141.9844 | 130 | 157.8 | 139 | 147 | 150 | 141 | 136 | 134 | 143.2 | 138 | 139 | 140.4525
| 1 | 5 | 161.2 | 128 | 140 | 145.0629 | 130 | 159.7 | 139 | 149 | 150 | 138 | 135 | 135 | 147.9 | 139 | 142 | 147
| 1 | 6 | 162.2 | 128 | 141 | 141.8083 | 137 | 154.4 | 139 | 146 | 149 | 140 | 135 | 131 | 142.4 | 142 | 144
| Mean |  | 161.3333 | 129.6667 | 140.3333 | 144.769 | 133.8333 | 154.8 | 139.1667 | 146.1667 | 149.6667 | 139.5 | 134.1667 | 132.3333 | 141.4333 | 140 | 139.6667
| Median |  | 161.05 | 130 | 140.5 | 145.2031 | 134.5 | 153.7 | 139 | 146.5 | 149.5 | 140 | 135 | 131.5 | 142.8 | 140 | 139
| Std Dev. |  | 2.3269 | 1.5055 | 1.2111 | 2.4609 | 3.1885 | 3.3003 | 2.2286 | 2.0412 | 0.8165 | 1.2247 | 2.137 | 1.7512 | 5.1149 | 1.4142 | 2.8752
| Rel.Std.Dev. |  | 1.442% | 1.161% | 0.863% | 1.7% | 2.382% | 2.132% | 1.601% | 1.397% | 0.5455% | 0.878% | 1.593% | 1.323% | 3.617% | 1.01% | 2.059%
| PDM3 |  | 14.87% | -7.679% | -0.08488% | 3.073% | -4.713% | 10.22% | -0.9155% | 4.068% | 6.56% | -0.6782% | -4.475% | -5.781% | 0.6983% | -0.3222% | -0.5595%
| Z-Score (Absolute) |  | 2.47 | 1.19 | 0.04 | 0.55 | 0.71 | 1.71 | 0.1 | 0.71 | 1.12 | 0.06 | 0.67 | 0.89 | 0.17 | 0 | 0.04
| Analytical results for Cs in OREAS 25b (Certified Value 2.03 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  |  | 2
| 1 | 1 | 2.14 | 1.76 | 2.6194 | 1.78 | 1.82 | 1.94 | 2.14 | 2 | 2.25 | 3.02 | 1.97 | 2.08 |  |  |  | 1
| 1 | 2 | 2.12 | 1.68 | 2.5612 | 1.79 | 1.74 | 1.9 | 2.09 | 2.13 | 2.21 | 2.95 | 2.02 | 2.1 |  |  |  | #N/A
| 1 | 3 | 2.12 | 1.72 | 2.6031 | 1.74 | 1.76 | 1.94 | 2.17 | 2.04 | 2.11 | 2.82 | 1.86 | 1.96 |  |  |  | 16
| 1 | 4 | 2.1 | 1.68 | 2.4575 | 1.65 | 1.78 | 1.94 | 2.23 | 2.1 | 2.18 | 2.8 | 1.91 | 2.25 |  |  |  | 2.0263
| 1 | 5 | 2.16 | 1.67 | 2.4499 | 1.71 | 1.78 | 2 | 2.2 | 1.95 | 2.25 | 2.82 | 1.89 | 2.22 |  |  |  | 148
| 1 | 6 | 2.09 | 1.71 | 2.4846 | 1.72 | 1.72 | 1.98 | 2.14 | 2.11 | 2.19 | 3.06 | 2 | 2.17
| Mean |  | 2.1217 | 1.7033 | 2.5293 | 1.7317 | 1.7667 | 1.95 | 2.1617 | 2.055 | 2.1983 | 2.9117 | 1.9417 | 2.13
| Median |  | 2.12 | 1.695 | 2.5229 | 1.73 | 1.77 | 1.94 | 2.155 | 2.07 | 2.2 | 2.885 | 1.94 | 2.135
| Std Dev. |  | 0.025626 | 0.033862 | 0.074884 | 0.051153 | 0.035024 | 0.035214 | 0.049565 | 0.07064 | 0.052313 | 0.11356 | 0.064317 | 0.10621
| Rel.Std.Dev. |  | 1.208% | 1.988% | 2.961% | 2.954% | 1.982% | 1.806% | 2.293% | 3.437% | 2.38% | 3.9% | 3.312% | 4.986%
| PDM3 |  | 4.707% | -15.94% | 24.82% | -14.54% | -12.81% | -3.765% | 6.681% | 1.416% | 8.49% | 43.69% | -4.177% | 5.118%
| Z-Score (Absolute) |  | 0.16 | 1.82 | 2.09 | 1.69 | 1.52 | 0.65 | 0.35 | 0.16 | 0.52 | 3.9 | 0.69 | 0.2
| Analytical results for Cu in OREAS 25b (Certified Value 29.1 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 1
| 1 | 1 | 29.94 | 30.5 | 30.2 | 34.0991 | 36.4 | 27.8 | 28.7 | 26.2 | 29.6 | 28.7 | 26.7 | 31 | 30.98 | 21.8 | 29 | 1
| 1 | 2 | 29.09 | 26.7 | 30.1 | 32.6226 | 34.6 | 28 | 27.2 | 26.2 | 28 | 28.8 | 23.3 | 31 | 29.24 | 22.1 | 28 | 6
| 1 | 3 | 28.97 | 26.7 | 29.9 | 34.3896 | 35.3 | 27.8 | 26.9 | 26.3 | 29.3 | 29.5 | 21.6 | 31 | 28.8 | 22.5 | 27 | 16
| 1 | 4 | 28.81 | 26.5 | 30.1 | 31.8581 | 35.1 | 28.4 | 27.1 | 26.7 | 29.7 | 29.7 | 23.4 | 32 | 31.77 | 21.5 | 29 | 29.1105
| 1 | 5 | 29.14 | 28.4 | 29.5 | 31.3694 | 35.6 | 27.1 | 27.3 | 26.5 | 28.8 | 29.8 | 22.7 | 32 | 33.02 | 21.6 | 30 | 149
| 1 | 6 | 28.86 | 26.2 | 29.5 | 33.467 | 35.3 | 27 | 27.6 | 26.4 | 28.6 | 29.5 | 22.8 | 31 | 30.58 | 22.6 | 30
| Mean |  | 29.135 | 27.5 | 29.8833 | 32.9676 | 35.3833 | 27.6833 | 27.4667 | 26.3833 | 29 | 29.3333 | 23.4167 | 31.3333 | 30.7317 | 22.0167 | 28.8333
| Median |  | 29.03 | 26.7 | 30 | 33.0448 | 35.3 | 27.8 | 27.25 | 26.35 | 29.05 | 29.5 | 23.05 | 31 | 30.78 | 21.95 | 29
| Std Dev. |  | 0.41438 | 1.6601 | 0.31252 | 1.2211 | 0.59805 | 0.53821 | 0.64704 | 0.19408 | 0.65422 | 0.46762 | 1.7314 | 0.5164 | 1.572 | 0.46224 | 1.169
| Rel.Std.Dev. |  | 1.422% | 6.037% | 1.046% | 3.704% | 1.69% | 1.944% | 2.356% | 0.7356% | 2.256% | 1.594% | 7.394% | 1.648% | 5.115% | 2.1% | 4.054%
| PDM3 |  | 0.08401% | -5.533% | 2.655% | 13.25% | 21.55% | -4.903% | -5.647% | -9.368% | -0.3797% | 0.7653% | -19.56% | 7.636% | 5.569% | -24.37% | -0.9523%
| Z-Score (Absolute) |  | 0.06 | 0.66 | 0.39 | 1.74 | 2.81 | 0.58 | 0.67 | 1.15 | 0 | 0.15 | 2.46 | 1.03 | 0.76 | 3.07 | 0.07
| Analytical results for Dy in OREAS 25b (Indicative Value 1.3 ppm) | [16 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 1.353 | 1.2 | 1.2 | [12 empty] | 1
| 1 | 2 | 1.357 | 1.2 | 1.27 | [12 empty] | 7
| 1 | 3 | 1.33 | 1.3 | 1.24 | [12 empty] | 16
| 1 | 4 | 1.338 | 1.2 | 1.42 | [12 empty] | 1.2974
| 1 | 5 | 1.358 | 1.3 | 1.3 | [12 empty] | 28
| 1 | 6 | 1.338 | 1.3 | 1.35
| Mean |  | 1.3457 | 1.25 | 1.2967
| Median |  | 1.3455 | 1.25 | 1.285
| Std Dev. |  | 0.01181 | 0.054772 | 0.079162
| Rel.Std.Dev. |  | 0.8776% | 4.382% | 6.105%
| PDM3 |  | 3.717% | -3.657% | -0.05995%
| Z-Score (Absolute) |  | 0.71 | 0.67 | 0
| Analytical results for Er in OREAS 25b (Indicative Value 0.56 ppm) | [16 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 0.588 | 0.5 | 0.57 | [12 empty] | 1
| 1 | 2 | 0.6 | 0.5 | 0.59 | [12 empty] | 23
| 1 | 3 | 0.58 | 0.5 | 0.57 | [12 empty] | 16
| 1 | 4 | 0.579 | 0.5 | 0.66 | [12 empty] | 0.56344
| 1 | 5 | 0.596 | 0.5 | 0.6 | [12 empty] | 29
| 1 | 6 | 0.589 | 0.5 | 0.62
| Mean |  | 0.58867 | 0.5 | 0.60167
| Median |  | 0.5885 | 0.5 | 0.595
| Std Dev. |  | 0.0083825 | 0 | 0.034303
| Rel.Std.Dev. |  | 1.424% | 0% | 5.701%
| PDM3 |  | 4.476% | -11.26% | 6.784%
| Z-Score (Absolute) |  | 0 | 4.6 | 0.67
| Analytical results for Eu in OREAS 25b (Indicative Value 0.52 ppm) | [16 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 0.56 | 0.5 | 0.47 | [12 empty] | 1
| 1 | 2 | 0.564 | 0.5 | 0.5 | [12 empty] | 8
| 1 | 3 | 0.559 | 0.5 | 0.47 | [12 empty] | 16
| 1 | 4 | 0.556 | 0.4 | 0.58 | [12 empty] | 0.51944
| 1 | 5 | 0.564 | 0.5 | 0.52 | [12 empty] | 30
| 1 | 6 | 0.557 | 0.5 | 0.55
| Mean |  | 0.56 | 0.48333 | 0.515
| Median |  | 0.5595 | 0.5 | 0.51
| Std Dev. |  | 0.0034059 | 0.040825 | 0.044159
| Rel.Std.Dev. |  | 0.6082% | 8.447% | 8.575%
| PDM3 |  | 7.807% | -6.952% | -0.8556%
| Z-Score (Absolute) |  | 0.96 | 0.67 | 0
| Analytical results for Fe in OREAS 25b (Certified Value 7.61 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 7.837 | 7.03 | 7.59 | 7.6084 | 7.45 | 8.23 | 7.68 | 7.75 | 7.61 | 7.59 | 7.42 | 7.52 | 7.01 | 6.9 | 8 | 1
| 1 | 2 | 7.595 | 6.96 | 7.6 | 7.8676 | 7.43 | 8.18 | 7.36 | 7.54 | 7.58 | 7.66 | 7.4 | 7.5 | 6.58 | 7.09 | 7.97 | #N/A
| 1 | 3 | 7.531 | 6.98 | 7.6 | 7.4117 | 7.15 | 8.24 | 7.5 | 7.69 | 7.64 | 7.7 | 7.16 | 7.5 | 6.51 | 7.06 | 7.91 | 16
| 1 | 4 | 7.501 | 7.03 | 7.68 | 7.9511 | 7.14 | 8.39 | 7.46 | 7.82 | 7.58 | 7.75 | 7.39 | 7.44 | 6.96 | 6.86 | 8.09 | 7.6055
| 1 | 5 | 7.719 | 6.88 | 7.57 | 8.0128 | 7.35 | 8.44 | 7.49 | 7.82 | 7.68 | 7.6 | 7.47 | 7.45 | 7.07 | 6.75 | 8.17 | 150
| 1 | 6 | 7.547 | 6.94 | 7.58 | 7.4236 | 7.34 | 8.36 | 7.47 | 7.66 | 7.54 | 7.67 | 7.41 | 7.38 | 6.89 | 7.05 | 8.2
| Mean |  | 7.6217 | 6.97 | 7.6033 | 7.7125 | 7.31 | 8.3067 | 7.4933 | 7.7133 | 7.605 | 7.6617 | 7.375 | 7.465 | 6.8367 | 6.9517 | 8.0567
| Median |  | 7.571 | 6.97 | 7.595 | 7.738 | 7.345 | 8.3 | 7.48 | 7.72 | 7.595 | 7.665 | 7.405 | 7.475 | 6.925 | 6.975 | 8.045
| Std Dev. |  | 0.13031 | 0.057271 | 0.039328 | 0.2668 | 0.13491 | 0.10386 | 0.10424 | 0.10727 | 0.049699 | 0.06047 | 0.10895 | 0.052058 | 0.23458 | 0.13586 | 0.11553
| Rel.Std.Dev. |  | 1.71% | 0.8217% | 0.5172% | 3.459% | 1.846% | 1.25% | 1.391% | 1.391% | 0.6535% | 0.7893% | 1.477% | 0.6974% | 3.431% | 1.954% | 1.434%
| PDM3 |  | 0.2125% | -8.356% | -0.02851% | 1.407% | -3.885% | 9.219% | -1.475% | 1.418% | -0.006597% | 0.7385% | -3.031% | -1.847% | -10.11% | -8.597% | 5.932%
| Z-Score (Absolute) |  | 0.09 | 3.09 | 0 | 0.53 | 1.43 | 3.43 | 0.54 | 0.54 | 0.01 | 0.28 | 1.11 | 0.67 | 3.74 | 3.18 | 2.21
| Analytical results for Ga in OREAS 25b (Certified Value 17.9 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 18.9 | 15.9 | 17.8 | 18.5274 | 18 | 18.1 | 17.75 | 17.05 | 18.7 | 17.3 | 15.7 | 18.2 | 18.6 | 14.6 | 17.8 | 1
| 1 | 2 | 18.8 | 15.2 | 17.8 | 18.3323 | 18 | 18.2 | 17 | 16.9 | 18.9 | 17.75 | 15.4 | 17.9 | 18.6 | 15 | 16.4 | #N/A
| 1 | 3 | 18.6 | 15.3 | 17.6 | 18.3906 | 18 | 18.6 | 17.25 | 16.9 | 19.7 | 17.3 | 15 | 18.3 | 17.8 | 14.2 | 17.8 | 16
| 1 | 4 | 18.5 | 15.3 | 17.6 | 17.8675 | 17.7 | 19 | 17.15 | 17.4 | 19.35 | 18 | 15.5 | 17.8 | 18.4 | 14.8 | 19.7 | 17.8782
| 1 | 5 | 18.8 | 14.9 | 17.6 | 18.1165 | 19.1 | 18.9 | 17.3 | 17.15 | 18.8 | 17.05 | 15.9 | 17.3 | 19.6 | 14 | 17.4 | 151
| 1 | 6 | 18.6 | 14.6 | 17.5 | 17.8756 | 18.5 | 18.4 | 17.5 | 17.15 | 18.75 | 18.25 | 15.8 | 17.6 | 18.7 | 14.9 | 20.3
| Mean |  | 18.7 | 15.2 | 17.65 | 18.185 | 18.2167 | 18.5333 | 17.325 | 17.0917 | 19.0333 | 17.6083 | 15.55 | 17.85 | 18.6167 | 14.5833 | 18.2333
| Median |  | 18.7 | 15.25 | 17.6 | 18.2244 | 18 | 18.5 | 17.275 | 17.1 | 18.85 | 17.525 | 15.6 | 17.85 | 18.6 | 14.7 | 17.8
| Std Dev. |  | 0.15492 | 0.43818 | 0.12247 | 0.27655 | 0.50365 | 0.36697 | 0.26599 | 0.18819 | 0.40208 | 0.46628 | 0.32711 | 0.37283 | 0.58109 | 0.40208 | 1.4733
| Rel.Std.Dev. |  | 0.8284% | 2.883% | 0.6939% | 1.521% | 2.765% | 1.98% | 1.535% | 1.101% | 2.112% | 2.648% | 2.104% | 2.089% | 3.121% | 2.757% | 8.08%
| PDM3 |  | 4.597% | -14.98% | -1.276% | 1.716% | 1.893% | 3.664% | -3.094% | -4.399% | 6.461% | -1.509% | -13.02% | -0.1578% | 4.131% | -18.43% | 1.986%
| Z-Score (Absolute) |  | 0.84 | 2.62 | 0.2 | 0.33 | 0.36 | 0.67 | 0.52 | 0.75 | 1.17 | 0.24 | 2.27 | 0 | 0.76 | 3.22 | 0.38
| Analytical results for Gd in OREAS 25b (Indicative Value 1.99 ppm) | [16 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 2.021 | 1.9 | 1.84 | [12 empty] | 1
| 1 | 2 | 2.035 | 2 | 1.94 | [12 empty] | 10
| 1 | 3 | 1.975 | 2 | 1.82 | [12 empty] | 16
| 1 | 4 | 2.008 | 1.9 | 2.17 | [12 empty] | 1.9875
| 1 | 5 | 1.997 | 2.1 | 1.98 | [12 empty] | 31
| 1 | 6 | 1.989 | 2 | 2.1
| Mean |  | 2.0042 | 1.9833 | 1.975
| Median |  | 2.0025 | 2 | 1.96
| Std Dev. |  | 0.021821 | 0.075277 | 0.13939
| Rel.Std.Dev. |  | 1.089% | 3.795% | 7.058%
| PDM3 |  | 0.8386% | -0.2096% | -0.6289%
| Z-Score (Absolute) |  | 1.69 | 0 | 0.67
| Analytical results for Ge in OREAS 25b (Indicative Value 0.077 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 07 | 08 | 09 | 10 | 11 | 15 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.05 | < 0.1 | < 0.1 | 0.07 | 0.08 | 0.11 | 0.07 | 0.06 | < 0.1 | [6 empty] | 1
| 1 | 2 | < 0.05 | < 0.1 | < 0.1 | 0.06 | 0.09 | 0.1 | 0.07 | 0.06 | < 0.1 | [6 empty] | 19
| 1 | 3 | < 0.05 | < 0.1 | < 0.1 | 0.06 | 0.09 | 0.1 | 0.07 | < 0.05 | < 0.1 | [6 empty] | 16
| 1 | 4 | < 0.05 | < 0.1 | < 0.1 | 0.06 | 0.09 | 0.12 | 0.09 | 0.06 | < 0.1 | [6 empty] | 0.0765
| 1 | 5 | < 0.05 | < 0.1 | < 0.1 | 0.07 | 0.11 | 0.11 | 0.08 | < 0.05 | < 0.1 | [6 empty] | 27
| 1 | 6 | < 0.05 | < 0.1 | < 0.1 | 0.07 | 0.08 | 0.11 | 0.08 | < 0.05 | < 0.1
| Mean |  |  |  |  | 0.065 | 0.09 | 0.10833 | 0.076667 | 0.06
| Median |  |  |  |  | 0.065 | 0.09 | 0.11 | 0.075 | 0.06
| Std Dev. |  |  |  |  | 0.0054772 | 0.010954 | 0.0075277 | 0.008165 | 0
| Rel.Std.Dev. |  |  |  |  | 8.427% | 12.17% | 6.949% | 10.65% | 0%
| PDM3 |  |  |  |  | -15.03% | 17.65% | 41.61% | 0.2179% | -21.57%
| Z-Score (Absolute) |  | 1.12 | 0 | 0 | 0.67 | 1.8 | 2.62 | 1.2 | 0.34 | 0
| Analytical results for Hf in OREAS 25b (Certified Value 0.75 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  |  | 2
| 1 | 1 | 0.71 | 0.56 | 0.86349 | 0.3 | 0.74 | 0.68 | 0.66 | 0.88 | 0.71 | 0.95 | 0.85 | 0.71 |  |  |  | 1
| 1 | 2 | 0.74 | 0.55 | 0.77787 | 0.3 | 0.71 | 0.65 | 0.63 | 0.9 | 0.73 | 0.91 | 0.93 | 0.7 |  |  |  | 11
| 1 | 3 | 0.72 | 0.53 | 0.83479 | 0.3 | 0.71 | 0.65 | 0.61 | 0.92 | 0.68 | 0.88 | 0.86 | 0.67 |  |  |  | 16
| 1 | 4 | 0.74 | 0.53 | 0.80969 | 0.3 | 0.7 | 0.64 | 0.64 | 0.96 | 0.74 | 0.9 | 0.87 | 0.77 |  |  |  | 0.7518
| 1 | 5 | 0.73 | 0.53 | 0.78248 | 0.3 | 0.69 | 0.68 | 0.61 | 0.9 | 0.79 | 0.89 | 0.84 | 0.77 |  |  |  | 152
| 1 | 6 | 0.72 | 0.52 | 0.84067 | 0.3 | 0.69 | 0.66 | 0.67 | 0.91 | 0.83 | 0.92 | 0.96 | 0.78
| Mean |  | 0.72667 | 0.53667 | 0.81817 | 0.3 | 0.70667 | 0.66 | 0.63667 | 0.91167 | 0.74667 | 0.90833 | 0.885 | 0.73333
| Median |  | 0.725 | 0.53 | 0.82224 | 0.3 | 0.705 | 0.655 | 0.635 | 0.905 | 0.735 | 0.905 | 0.865 | 0.74
| Std Dev. |  | 0.012111 | 0.015055 | 0.034076 | 0 | 0.018619 | 0.016733 | 0.025033 | 0.027142 | 0.05465 | 0.024833 | 0.048477 | 0.045898
| Rel.Std.Dev. |  | 1.667% | 2.805% | 4.165% | 0% | 2.635% | 2.535% | 3.932% | 2.977% | 7.319% | 2.734% | 5.478% | 6.259%
| PDM3 |  | -3.343% | -28.62% | 8.827% | -60.1% | -6.004% | -12.21% | -15.31% | 21.26% | -0.6832% | 20.82% | 17.72% | -2.457%
| Z-Score (Absolute) |  | 0.05 | 1.56 | 0.67 | NA | 0.21 | 0.58 | 0.77 | 1.42 | 0.11 | 1.39 | 1.21 | 0
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 25b (Certified Value 0.028 ppm) | [16 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g |  |  |  | 3
| 1 | 1 | 0.03 | 0.02 | < 0.05 | 0.05 | 0.03 | 0.03 | 0.01 | 0.03 | 0.05 | 0.14 | 0.021 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.03 | 0.02 | < 0.05 | 0.04 | 0.03 | 0.03 | 0.01 | 0.04 | 0.04 | 0.09 | 0.021 | < 0.1 |  |  |  | #N/A
| 1 | 3 | 0.03 | 0.02 | < 0.05 | 0.07 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.08 | 0.032 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.03 | 0.03 | < 0.05 | 0.06 | 0.03 | 0.03 | 0.01 | 0.02 | 0.03 | 0.07 | 0.021 | < 0.1 |  |  |  | 0.027625
| 1 | 5 | 0.03 | 0.04 | < 0.05 | 0.06 | 0.03 | 0.03 | 0.01 | 0.03 | 0.04 | 0.06 | 0.021 | < 0.1 |  |  |  | 153
| 1 | 6 | 0.03 | 0.02 | < 0.05 | 0.06 | 0.03 | 0.03 | 0.01 | 0.04 | 0.04 | 0.06 | 0.021 | < 0.1
| Mean |  | 0.03 | 0.025 |  | 0.056667 | 0.03 | 0.03 | 0.013333 | 0.031667 | 0.04 | 0.083333 | 0.022833
| Median |  | 0.03 | 0.02 |  | 0.06 | 0.03 | 0.03 | 0.01 | 0.03 | 0.04 | 0.075 | 0.021
| Std Dev. |  | 0 | 0.0083666 |  | 0.010328 | 0 | 0 | 0.008165 | 0.0075277 | 0.0063246 | 0.030111 | 0.0044907
| Rel.Std.Dev. |  | 0% | 33.47% |  | 18.23% | 0% | 0% | 61.24% | 23.77% | 15.81% | 36.13% | 19.67%
| PDM3 |  | 8.597% | -9.502% |  | 105.1% | 8.597% | 8.597% | -51.73% | 14.63% | 44.8% | 201.7% | -17.35%
| Z-Score (Absolute) |  | 0 | 0.55 | 0.55 | 2.96 | 0 | 0 | 1.85 | 0.18 | 1.11 | 5.91 | 0.79 | 2.22
| Analytical results for Ho in OREAS 25b (Indicative Value 0.22 ppm) | [16 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 0.233 | 0.2 | 0.21 | [12 empty] | 1
| 1 | 2 | 0.231 | 0.2 | 0.22 | [12 empty] | 22
| 1 | 3 | 0.222 | 0.2 | 0.21 | [12 empty] | 16
| 1 | 4 | 0.23 | 0.2 | 0.24 | [12 empty] | 0.21689
| 1 | 5 | 0.232 | 0.2 | 0.22 | [12 empty] | 28
| 1 | 6 | 0.226 | 0.2 | 0.23
| Mean |  | 0.229 | 0.2 | 0.22167
| Median |  | 0.2305 | 0.2 | 0.22
| Std Dev. |  | 0.0041952 | 3.0405e-17 | 0.01169
| Rel.Std.Dev. |  | 1.832% | 1.52e-14% | 5.274%
| PDM3 |  | 5.584% | -7.787% | 2.203%
| Z-Score (Absolute) |  | 0.67 | 1.99 | 0
| Analytical results for In in OREAS 25b (Certified Value 0.065 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  |  | 3
| 1 | 1 | 0.06 | 0.06 | 0.071598 | 0.07 | 0.065 | 0.06 | 0.071 | 0.065 | 0.061 | 0.07 | 0.06 | 0.06 |  |  |  | 1
| 1 | 2 | 0.06 | 0.06 | 0.071955 | 0.07 | 0.065 | 0.06 | 0.065 | 0.068 | 0.06 | 0.07 | 0.063 | 0.05 |  |  |  | 13
| 1 | 3 | 0.07 | 0.06 | 0.072933 | 0.06 | 0.067 | 0.063 | 0.069 | 0.064 | 0.062 | 0.07 | 0.06 | 0.06 |  |  |  | 16
| 1 | 4 | 0.07 | 0.06 | 0.069289 | 0.06 | 0.058 | 0.067 | 0.073 | 0.067 | 0.057 | 0.07 | 0.059 | 0.06 |  |  |  | 0.064552
| 1 | 5 | 0.07 | 0.06 | 0.063087 | 0.06 | 0.065 | 0.06 | 0.074 | 0.064 | 0.063 | 0.07 | 0.057 | 0.07 |  |  |  | 154
| 1 | 6 | 0.07 | 0.06 | 0.064101 | 0.08 | 0.071 | 0.058 | 0.064 | 0.077 | 0.061 | 0.07 | 0.06 | 0.07
| Mean |  | 0.066667 | 0.06 | 0.068827 | 0.066667 | 0.065167 | 0.061333 | 0.069333 | 0.0675 | 0.060667 | 0.07 | 0.059833 | 0.061667
| Median |  | 0.07 | 0.06 | 0.070443 | 0.065 | 0.065 | 0.06 | 0.07 | 0.066 | 0.061 | 0.07 | 0.06 | 0.06
| Std Dev. |  | 0.005164 | 0 | 0.0042382 | 0.008165 | 0.0042151 | 0.0032042 | 0.0041312 | 0.0049295 | 0.0020656 | 0 | 0.0019408 | 0.0075277
| Rel.Std.Dev. |  | 7.746% | 0% | 6.158% | 12.25% | 6.468% | 5.224% | 5.958% | 7.303% | 3.405% | 0% | 3.244% | 12.21%
| PDM3 |  | 3.275% | -7.052% | 6.622% | 3.275% | 0.9518% | -4.987% | 7.407% | 4.566% | -6.019% | 8.439% | -7.31% | -4.47%
| Z-Score (Absolute) |  | 0.13 | 1.06 | 0.52 | 0.13 | 0.13 | 0.82 | 0.61 | 0.28 | 0.94 | 0.73 | 1.09 | 0.76
| Analytical results for K in OREAS 25b (Certified Value 0.125 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.126 | 0.1 | 0.13 | 0.13709 | 0.13 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12 | 0.1 | 0.164 | 0.12 | 0.213 | 0.129 | 1
| 1 | 2 | 0.126 | 0.1 | 0.13 | 0.13251 | 0.12 | 0.13 | 0.11 | 0.12 | 0.12 | 0.12 | 0.1 | 0.16 | 0.12 | 0.214 | 0.126 | #N/A
| 1 | 3 | 0.124 | 0.1 | 0.13 | 0.13808 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.1 | 0.165 | 0.12 | 0.212 | 0.12 | 16
| 1 | 4 | 0.125 | 0.1 | 0.13 | 0.13732 | 0.11 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12 | 0.1 | 0.166 | 0.13 | 0.209 | 0.127 | 0.12534
| 1 | 5 | 0.128 | 0.1 | 0.13 | 0.13118 | 0.12 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12 | 0.11 | 0.17 | 0.13 | 0.205 | 0.135 | 155
| 1 | 6 | 0.126 | 0.09 | 0.13 | 0.13129 | 0.13 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12 | 0.11 | 0.163 | 0.12 | 0.217 | 0.133
| Mean |  | 0.12583 | 0.098333 | 0.13 | 0.13458 | 0.12167 | 0.13 | 0.11833 | 0.12 | 0.12667 | 0.12 | 0.10333 | 0.16467 | 0.12333 | 0.21167 | 0.12833
| Median |  | 0.126 | 0.1 | 0.13 | 0.1348 | 0.12 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12 | 0.1 | 0.1645 | 0.12 | 0.2125 | 0.128
| Std Dev. |  | 0.0013292 | 0.0040825 | 0 | 0.0032482 | 0.0075277 | 0 | 0.0040825 | 0 | 0.005164 | 0 | 0.005164 | 0.0033267 | 0.005164 | 0.0041793 | 0.0053541
| Rel.Std.Dev. |  | 1.056% | 4.152% | 0% | 2.414% | 6.187% | 0% | 3.45% | 0% | 4.077% | 0% | 4.997% | 2.02% | 4.187% | 1.974% | 4.172%
| PDM3 |  | 0.3932% | -21.55% | 3.718% | 7.37% | -2.931% | 3.718% | -5.59% | -4.261% | 1.058% | -4.261% | -17.56% | 31.38% | -1.601% | 68.87% | 2.388%
| Z-Score (Absolute) |  | 0 | 3.18 | 0.48 | 1.01 | 0.48 | 0.48 | 0.87 | 0.67 | 0.1 | 0.67 | 2.6 | 4.49 | 0.29 | 9.92 | 0.29
| Analytical results for La in OREAS 25b (Certified Value 15.7 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 16.284 | 12.8 | 15.1 | 16.8795 | 15.6 | 16 | 16 | 15.2 | 16.2 | 15.7 | 15.2 | 18.3 | 15.5 | 15.7 | 13.9 | 1
| 1 | 2 | 15.991 | 12.3 | 15 | 16.4801 | 15.7 | 15.9 | 15.3 | 15.1 | 15.4 | 16.7 | 15 | 18.3 | 14.8 | 16.7 | 14.2 | 1
| 1 | 3 | 15.619 | 12.4 | 14.9 | 16.9446 | 15.1 | 16 | 15.3 | 15.3 | 17.1 | 16 | 14.6 | 18.3 | 14.4 | 15.8 | 13.3 | 16
| 1 | 4 | 15.749 | 12 | 15 | 16.7724 | 14.4 | 16.8 | 15.1 | 15.7 | 17 | 16.8 | 15.2 | 18.3 | 15.6 | 16.1 | 16.1 | 15.6887
| 1 | 5 | 15.678 | 11.7 | 14.6 | 16.4735 | 15.1 | 16.5 | 15.6 | 15.6 | 16.8 | 15.6 | 15.7 | 18 | 15.9 | 15.4 | 15.2 | 156
| 1 | 6 | 15.791 | 11.3 | 14.7 | 16.3946 | 15.1 | 16.3 | 15.6 | 15.5 | 16.4 | 17.2 | 15.6 | 18.1 | 15.1 | 16.4 | 16
| Mean |  | 15.852 | 12.0833 | 14.8833 | 16.6574 | 15.1667 | 16.25 | 15.4833 | 15.4 | 16.4833 | 16.3333 | 15.2167 | 18.2167 | 15.2167 | 16.0167 | 14.7833
| Median |  | 15.77 | 12.15 | 14.95 | 16.6262 | 15.1 | 16.15 | 15.45 | 15.4 | 16.6 | 16.35 | 15.2 | 18.3 | 15.3 | 15.95 | 14.7
| Std Dev. |  | 0.24692 | 0.53448 | 0.19408 | 0.23638 | 0.46332 | 0.35071 | 0.31885 | 0.23664 | 0.63377 | 0.65625 | 0.40208 | 0.13292 | 0.55648 | 0.47924 | 1.1583
| Rel.Std.Dev. |  | 1.558% | 4.423% | 1.304% | 1.419% | 3.055% | 2.158% | 2.059% | 1.537% | 3.845% | 4.018% | 2.642% | 0.7296% | 3.657% | 2.992% | 7.835%
| PDM3 |  | 1.041% | -22.98% | -5.133% | 6.175% | -3.327% | 3.578% | -1.309% | -1.84% | 5.065% | 4.109% | -3.009% | 16.11% | -3.009% | 2.091% | -5.771%
| Z-Score (Absolute) |  | 0.41 | 3.82 | 0.67 | 1.32 | 0.36 | 0.86 | 0 | 0.09 | 1.12 | 0.96 | 0.3 | 3.07 | 0.3 | 0.6 | 0.79
| Analytical results for Li in OREAS 25b (Certified Value 14.1 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 1
| 1 | 1 | 14.76 | 9 | 15.1 | 16.9166 | 16.2 | 13.8 | 12.5 | 14.5 | 13.3 | 9.6 | 21.3 | 14 |  |  |  | 1
| 1 | 2 | 15.23 | 8 | 15.1 | 17.2393 | 16.1 | 12.7 | 12.1 | 14.4 | 13.6 | 9.8 | 21.1 | 14 |  |  |  | #N/A
| 1 | 3 | 14.99 | 9 | 15.4 | 17.2405 | 16.4 | 13.3 | 12.8 | 14.7 | 13.3 | 8.7 | 21.3 | 14 |  |  |  | 16
| 1 | 4 | 14.94 | 8 | 15.4 | 16.4451 | 15.6 | 13.2 | 12.9 | 14.4 | 13.7 | 9.2 | 21.1 | 13 |  |  |  | 14.0664
| 1 | 5 | 15.37 | 8 | 14.8 | 16.6495 | 15.8 | 13.4 | 13.4 | 14.8 | 13 | 10.4 | 21.5 | 13 |  |  |  | 157
| 1 | 6 | 14.97 | 8 | 14.8 | 16.3348 | 16.1 | 13.3 | 12.8 | 14.5 | 13.5 | 10.1 | 21.6 | 14
| Mean |  | 15.0433 | 8.3333 | 15.1 | 16.8043 | 16.0333 | 13.2833 | 12.75 | 14.55 | 13.4 | 9.6333 | 21.3167 | 13.6667
| Median |  | 14.98 | 8 | 15.1 | 16.7831 | 16.1 | 13.3 | 12.8 | 14.5 | 13.4 | 9.7 | 21.3 | 14
| Std Dev. |  | 0.21942 | 0.5164 | 0.26833 | 0.39129 | 0.28752 | 0.35449 | 0.43243 | 0.16432 | 0.25298 | 0.61536 | 0.20412 | 0.5164
| Rel.Std.Dev. |  | 1.459% | 6.197% | 1.777% | 2.329% | 1.793% | 2.669% | 3.392% | 1.129% | 1.888% | 6.388% | 0.9576% | 3.779%
| PDM3 |  | 6.945% | -40.76% | 7.348% | 19.46% | 13.98% | -5.567% | -9.358% | 3.438% | -4.738% | -31.52% | 51.54% | -2.842%
| Z-Score (Absolute) |  | 0.11 | NA | 0.14 | 0.93 | 0.57 | 0.7 | 0.95 | 0.11 | 0.65 | 2.39 | 3.02 | NA
Results from laboratories 2 and 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 25b (Indicative Value 0.06 ppm) | [16 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 05 | 15 | 16 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.25g | 0.5g | 0.25g | 0.4g | [10 empty] | 3
| 1 | 1 | 0.062 | 0.05 | < 0.1 | 0.06 | 0.06 | [10 empty] | 1
| 1 | 2 | 0.062 | 0.05 | < 0.1 | 0.07 | 0.07 | [10 empty] | 23
| 1 | 3 | 0.063 | 0.05 | < 0.1 | 0.09 | 0.06 | [10 empty] | 16
| 1 | 4 | 0.062 | 0.06 | < 0.1 | 0.06 | 0.07 | [10 empty] | 0.06025
| 1 | 5 | 0.062 | 0.05 | < 0.1 | 0.06 | 0.06 | [10 empty] | 29
| 1 | 6 | 0.063 | 0.05 | < 0.1 | 0.06 | 0.07
| Mean |  | 0.062333 | 0.051667 |  | 0.066667 | 0.065
| Median |  | 0.062 | 0.05 |  | 0.06 | 0.065
| Std Dev. |  | 0.0005164 | 0.0040825 |  | 0.012111 | 0.0054772
| Rel.Std.Dev. |  | 0.8284% | 7.902% |  | 18.17% | 8.427%
| PDM3 |  | 3.458% | -14.25% |  | 10.65% | 7.884%
| Z-Score (Absolute) |  | 0 | 1.66 | 1.92 | 0.67 | 0.41
| Analytical results for Mg in OREAS 25b (Certified Value 0.193 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.202 | 0.16 | 0.2 | 0.22435 | 0.24 | 0.19 | 0.18 | 0.17 | 0.17 | 0.17 | 0.16 | 0.24 | 0.19 | 0.199 | 0.196 | 1
| 1 | 2 | 0.203 | 0.16 | 0.2 | 0.23826 | 0.27 | 0.19 | 0.17 | 0.16 | 0.17 | 0.17 | 0.16 | 0.24 | 0.18 | 0.2 | 0.192 | #N/A
| 1 | 3 | 0.202 | 0.16 | 0.2 | 0.22012 | 0.24 | 0.2 | 0.17 | 0.17 | 0.17 | 0.17 | 0.15 | 0.24 | 0.18 | 0.199 | 0.185 | 16
| 1 | 4 | 0.2 | 0.16 | 0.2 | 0.23046 | 0.23 | 0.19 | 0.17 | 0.17 | 0.17 | 0.18 | 0.16 | 0.25 | 0.19 | 0.196 | 0.196 | 0.19251
| 1 | 5 | 0.206 | 0.15 | 0.2 | 0.235 | 0.24 | 0.19 | 0.17 | 0.17 | 0.18 | 0.17 | 0.17 | 0.25 | 0.2 | 0.192 | 0.207 | 158
| 1 | 6 | 0.202 | 0.15 | 0.2 | 0.22535 | 0.24 | 0.19 | 0.17 | 0.17 | 0.17 | 0.17 | 0.16 | 0.25 | 0.18 | 0.202 | 0.205
| Mean |  | 0.2025 | 0.15667 | 0.2 | 0.22892 | 0.24333 | 0.19167 | 0.17167 | 0.16833 | 0.17167 | 0.17167 | 0.16 | 0.245 | 0.18667 | 0.198 | 0.19683
| Median |  | 0.202 | 0.16 | 0.2 | 0.22791 | 0.24 | 0.19 | 0.17 | 0.17 | 0.17 | 0.17 | 0.16 | 0.245 | 0.185 | 0.199 | 0.196
| Std Dev. |  | 0.0019748 | 0.005164 | 3.0405e-17 | 0.0068939 | 0.013663 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0063246 | 0.0054772 | 0.008165 | 0.0035214 | 0.0081833
| Rel.Std.Dev. |  | 0.9752% | 3.296% | 1.52e-14% | 3.011% | 5.615% | 2.13% | 2.378% | 2.425% | 2.378% | 2.378% | 3.953% | 2.236% | 4.374% | 1.778% | 4.157%
| PDM3 |  | 5.192% | -18.62% | 3.893% | 18.92% | 26.4% | -0.436% | -10.83% | -12.56% | -10.83% | -10.83% | -16.89% | 27.27% | -3.033% | 2.854% | 2.248%
| Z-Score (Absolute) |  | 0.37 | 1.18 | 0.28 | 1.26 | 1.74 | 0 | 0.67 | 0.79 | 0.67 | 0.67 | 1.07 | 1.8 | 0.17 | 0.21 | 0.17
| Analytical results for Mn in OREAS 25b (Certified Value 0.018 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.01825 | 0.0158 | 0.0247 | 0.020762 | 0.0184 | 0.0176 | 0.0171 | 0.0172 | 0.0175 | 0.0167 | 0.0165 | 0.0182 | 0.0188 | 0.0161 | 0.0194 | 1
| 1 | 2 | 0.01786 | 0.0155 | 0.0247 | 0.019715 | 0.0182 | 0.0174 | 0.0163 | 0.017 | 0.0174 | 0.0167 | 0.0165 | 0.0181 | 0.018 | 0.0161 | 0.0178 | #N/A
| 1 | 3 | 0.0176 | 0.0154 | 0.0248 | 0.020711 | 0.0177 | 0.0176 | 0.0167 | 0.0174 | 0.0175 | 0.0169 | 0.0155 | 0.018 | 0.0175 | 0.0162 | 0.018 | 16
| 1 | 4 | 0.01763 | 0.0155 | 0.0247 | 0.019323 | 0.0172 | 0.0178 | 0.0166 | 0.0174 | 0.0175 | 0.017 | 0.0169 | 0.0182 | 0.019 | 0.0159 | 0.0191 | 0.017504
| 1 | 5 | 0.01814 | 0.0151 | 0.0244 | 0.019347 | 0.018 | 0.0179 | 0.0167 | 0.0176 | 0.0179 | 0.0165 | 0.0167 | 0.0184 | 0.0194 | 0.0154 | 0.017 | 159
| 1 | 6 | 0.01776 | 0.0151 | 0.0248 | 0.020641 | 0.0185 | 0.0177 | 0.0166 | 0.0173 | 0.0175 | 0.0169 | 0.0166 | 0.0183 | 0.0184 | 0.0161 | 0.0184
| Mean |  | 0.017873 | 0.0154 | 0.024683 | 0.020083 | 0.018 | 0.017667 | 0.016667 | 0.017317 | 0.01755 | 0.016783 | 0.01645 | 0.0182 | 0.018517 | 0.015967 | 0.018283
| Median |  | 0.01781 | 0.01545 | 0.0247 | 0.020178 | 0.0181 | 0.01765 | 0.01665 | 0.01735 | 0.0175 | 0.0168 | 0.01655 | 0.0182 | 0.0186 | 0.0161 | 0.0182
| Std Dev. |  | 0.0002683 | 0.00026833 | 0.0001472 | 0.0006959 | 0.0004858 | 0.00017512 | 0.0002582 | 0.00020412 | 0.00017607 | 0.00018348 | 0.00048888 | 0.00014142 | 0.00069402 | 0.00029439 | 0.00088185
| Rel.Std.Dev. |  | 1.501% | 1.742% | 0.5963% | 3.465% | 2.699% | 0.9912% | 1.549% | 1.179% | 1.003% | 1.093% | 2.972% | 0.777% | 3.748% | 1.844% | 4.823%
| PDM3 |  | 2.108% | -12.02% | 41.01% | 14.73% | 2.832% | 0.9278% | -4.785% | -1.072% | 0.2613% | -4.119% | -6.023% | 3.975% | 5.784% | -8.784% | 4.451%
| Z-Score (Absolute) |  | 0.16 | 1.8 | 5.57 | 1.92 | 0.26 | 0 | 0.79 | 0.28 | 0.09 | 0.7 | 0.97 | 0.42 | 0.67 | 1.35 | 0.49
| Analytical results for Mo in OREAS 25b (Certified Value 1.24 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 1.43 | 1.51 | < 2 | 1.2735 | 0.36 | 0.8 | 1.29 | 1.08 | 1.18 | 0.9 | 1.74 | 1.4 | 1.5 | 1.07 | 1.4 | 1
| 1 | 2 | 1.42 | 1.46 | < 2 | 1.2395 | 0.36 | 0.84 | 1.19 | 1.12 | 1.17 | 0.97 | 1.64 | 1.3 | 1.47 | 1.1 | 1.4 | 2
| 1 | 3 | 1.38 | 1.48 | < 2 | 1.3131 | 0.35 | 0.78 | 1.16 | 1.1 | 1.17 | 0.99 | 1.64 | 1.4 | 1.52 | 0.97 | 1.4 | 16
| 1 | 4 | 1.45 | 1.51 | < 2 | 1.3418 | 0.39 | 0.78 | 1.14 | 1.07 | 1.17 | 0.96 | 1.73 | 1.3 | 1.51 | 0.97 | 1.4 | 1.2431
| 1 | 5 | 1.49 | 1.46 | < 2 | 1.2668 | 0.37 | 0.9 | 1.34 | 1.1 | 1.17 | 0.91 | 1.7 | 1.3 | 1.57 | 0.87 | 1.4 | 160
| 1 | 6 | 1.51 | 1.48 | < 2 | 1.3269 | 0.35 | 0.85 | 1.17 | 1.16 | 1.12 | 0.94 | 1.67 | 1.3 | 1.56 | 0.9 | 1.3
| Mean |  | 1.4467 | 1.4833 |  | 1.2936 | 0.36333 | 0.825 | 1.215 | 1.105 | 1.1633 | 0.945 | 1.6867 | 1.3333 | 1.5217 | 0.98 | 1.3833
| Median |  | 1.44 | 1.48 |  | 1.2933 | 0.36 | 0.82 | 1.18 | 1.1 | 1.17 | 0.95 | 1.685 | 1.3 | 1.515 | 0.97 | 1.4
| Std Dev. |  | 0.04761 | 0.022509 |  | 0.039656 | 0.015055 | 0.047223 | 0.080685 | 0.032094 | 0.021602 | 0.035071 | 0.043665 | 0.05164 | 0.037639 | 0.090774 | 0.040825
| Rel.Std.Dev. |  | 3.291% | 1.517% |  | 3.066% | 4.144% | 5.724% | 6.641% | 2.904% | 1.857% | 3.711% | 2.589% | 3.873% | 2.474% | 9.263% | 2.951%
| PDM3 |  | 16.38% | 19.33% |  | 4.064% | -70.77% | -33.63% | -2.259% | -11.11% | -6.416% | -23.98% | 35.68% | 7.26% | 22.41% | -21.16% | 11.28%
| Z-Score (Absolute) |  | 0.88 | 0.99 | 0.5 | 0.4 | 2.47 | 1.04 | 0.16 | 0.18 | 0 | 0.67 | 1.62 | NA | 1.11 | 0.57 | NA
Results from laboratories 12 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 25b (Certified Value 0.019 wt.%) | [16 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.019 | 0.03 | 0.02 | 0.026121 | 0.023 | 0.017 | 0.02 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.017 | 0.019 | 0.0185 | 1
| 1 | 2 | 0.019 | 0.03 | 0.02 | 0.02768 | 0.023 | 0.017 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.02 | 0.016 | 0.021 | 0.0185 | #N/A
| 1 | 3 | 0.019 | 0.03 | 0.02 | 0.025749 | 0.023 | 0.017 | 0.01 | 0.02 | 0.01 | 0.02 | 0.01 | 0.02 | 0.016 | 0.019 | 0.018 | 16
| 1 | 4 | 0.019 | 0.03 | 0.02 | 0.02773 | 0.021 | 0.017 | 0.02 | 0.02 | 0.01 | 0.02 | 0.01 | 0.02 | 0.017 | 0.017 | 0.0185 | 0.018542
| 1 | 5 | 0.019 | 0.03 | 0.02 | 0.027421 | 0.022 | 0.017 | 0.02 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.018 | 0.019 | 0.019 | 161
| 1 | 6 | 0.019 | 0.03 | 0.02 | 0.025921 | 0.021 | 0.017 | 0.01 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.017 | 0.018 | 0.0195
| Mean |  | 0.019 | 0.03 | 0.02 | 0.02677 | 0.022167 | 0.017 | 0.015 | 0.018333 | 0.01 | 0.02 | 0.016667 | 0.02 | 0.016833 | 0.018833 | 0.018667
| Median |  | 0.019 | 0.03 | 0.02 | 0.026771 | 0.0225 | 0.017 | 0.015 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.017 | 0.019 | 0.0185
| Std Dev. |  | 0 | 0 | 0 | 0.00093373 | 0.00098319 | 0 | 0.0054772 | 0.0040825 | 0 | 0 | 0.005164 | 0 | 0.00075277 | 0.0013292 | 0.0005164
| Rel.Std.Dev. |  | 0% | 0% | 0% | 3.488% | 4.435% | 0% | 36.51% | 22.27% | 0% | 0% | 30.98% | 0% | 4.472% | 7.057% | 2.766%
| PDM3 |  | 2.472% | 61.8% | 7.865% | 44.38% | 19.55% | -8.315% | -19.1% | -1.124% | -46.07% | 7.865% | -10.11% | 7.865% | -9.213% | 1.573% | 0.6742%
| Z-Score (Absolute) |  | 0.06 | 4.11 | 0.43 | 2.92 | 1.23 | 0.67 | 1.41 | 0.18 | 3.25 | 0.43 | 0.8 | 0.43 | 0.74 | 0 | 0.06
| Analytical results for Nb in OREAS 25b (Certified Value 0.58 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  | 2
| 1 | 1 | 0.51 | 0.63 | < 5 | 1.6856 | < 0.1 | 0.51 | 0.78 | 0.38 | 0.62 | 1.06 | 0.7 | 0.32 | 0.9 |  |  | 1
| 1 | 2 | 0.52 | 0.56 | < 5 | 1.7903 | < 0.1 | 0.53 | 0.79 | 0.33 | 0.63 | 1.03 | 0.5 | 0.28 | 0.9 |  |  | 3
| 1 | 3 | 0.5 | 0.59 | < 5 | 1.8548 | < 0.1 | 0.48 | 0.77 | 0.33 | 0.69 | 0.97 | 0.5 | 0.31 | 0.8 |  |  | 16
| 1 | 4 | 0.51 | 0.57 | < 5 | 1.5567 | < 0.1 | 0.52 | 0.73 | 0.38 | 0.67 | 1.04 | 0.5 | 0.29 | 0.9 |  |  | 0.58488
| 1 | 5 | 0.5 | 0.58 | < 5 | 1.8912 | < 0.1 | 0.51 | 0.71 | 0.32 | 0.62 | 1.07 | 0.5 | 0.25 | 0.9 |  |  | 162
| 1 | 6 | 0.52 | 0.57 | < 5 | 1.5854 | < 0.1 | 0.5 | 0.74 | 0.37 | 0.67 | 1.07 | 0.7 | 0.3 | 0.9
| Mean |  | 0.51 | 0.58333 |  | 1.7273 |  | 0.50833 | 0.75333 | 0.35167 | 0.65 | 1.04 | 0.56667 | 0.29167 | 0.88333
| Median |  | 0.51 | 0.575 |  | 1.738 |  | 0.51 | 0.755 | 0.35 | 0.65 | 1.05 | 0.5 | 0.295 | 0.9
| Std Dev. |  | 0.0089443 | 0.025033 |  | 0.14005 |  | 0.017224 | 0.031411 | 0.027869 | 0.030332 | 0.037947 | 0.10328 | 0.024833 | 0.040825
| Rel.Std.Dev. |  | 1.754% | 4.291% |  | 8.108% |  | 3.388% | 4.17% | 7.925% | 4.666% | 3.649% | 18.23% | 8.514% | 4.622%
| PDM3 |  | -12.8% | -0.2636% |  | 195.3% |  | -13.09% | 28.8% | -39.87% | 11.13% | 77.82% | -3.113% | -50.13% | 51.03%
| Z-Score (Absolute) |  | 0.21 | 0 | 5.58 | 3.33 | 1.55 | 0.22 | 0.49 | 0.67 | 0.19 | 1.33 | NA | 0.85 | NA
| Analytical results for Nd in OREAS 25b (Indicative Value 14.1 ppm) | [16 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 1
| 1 | 1 | 14.032 | 14.9 | 13.4 | [12 empty] | 1
| 1 | 2 | 14.03 | 14.5 | 13.6 | [12 empty] | 4
| 1 | 3 | 13.719 | 13.8 | 12.7 | [12 empty] | 16
| 1 | 4 | 14.081 | 13.8 | 15.8 | [12 empty] | 14.1352
| 1 | 5 | 13.925 | 14.1 | 14 | [12 empty] | 30
| 1 | 6 | 13.846 | 14.8 | 15.4
| Mean |  | 13.9388 | 14.3167 | 14.15
| Median |  | 13.9775 | 14.3 | 13.8
| Std Dev. |  | 0.13736 | 0.48751 | 1.2062
| Rel.Std.Dev. |  | 0.9854% | 3.405% | 8.525%
| PDM3 |  | -1.389% | 1.284% | 0.1049%
| Z-Score (Absolute) |  | 0.85 | 0.67 | 0
| Analytical results for Ni in OREAS 25b (Certified Value 34.4 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 1
| 1 | 1 | 38.9 | 25.6 | 39.8 | 52.728 | 60.2 | 37.8 | 34.2 | 31.3 | 36.5 | 34.9 | 27.1 | 52 | 35.9 | 25.8 | 36 | 1
| 1 | 2 | 38.9 | 24.5 | 39.8 | 50.5257 | 60.9 | 36.5 | 32.1 | 30.4 | 35.9 | 36.9 | 24.1 | 51 | 37.2 | 25.8 | 34 | 5
| 1 | 3 | 38.8 | 24.6 | 39.6 | 53.0431 | 62.9 | 37.4 | 33.1 | 32.1 | 37.4 | 36.9 | 22.1 | 52 | 34.1 | 24.6 | 32 | 16
| 1 | 4 | 38.3 | 24.1 | 39.9 | 50.4308 | 61.7 | 39.4 | 33.6 | 32.7 | 37.5 | 38.1 | 23.2 | 52 | 38.6 | 26.2 | 37 | 34.4106
| 1 | 5 | 39 | 23.8 | 39.4 | 50.5531 | 62.3 | 38 | 34.1 | 33.8 | 36.5 | 34.8 | 24.9 | 54 | 40.7 | 25.8 | 40 | 163
| 1 | 6 | 38.4 | 23.1 | 39.6 | 50.0441 | 61.4 | 35.8 | 33.1 | 32.6 | 36.3 | 37.9 | 25.6 | 52 | 36.3 | 27.1 | 39
| Mean |  | 38.7167 | 24.2833 | 39.6833 | 51.2208 | 61.5667 | 37.4833 | 33.3667 | 32.15 | 36.6833 | 36.5833 | 24.5 | 52.1667 | 37.1333 | 25.8833 | 36.3333
| Median |  | 38.85 | 24.3 | 39.7 | 50.5394 | 61.55 | 37.6 | 33.35 | 32.35 | 36.5 | 36.9 | 24.5 | 52 | 36.75 | 25.8 | 36.5
| Std Dev. |  | 0.29269 | 0.84242 | 0.18348 | 1.3062 | 0.96678 | 1.2529 | 0.77889 | 1.1845 | 0.63377 | 1.4317 | 1.7743 | 0.98319 | 2.2932 | 0.80602 | 3.0111
| Rel.Std.Dev. |  | 0.756% | 3.469% | 0.4624% | 2.55% | 1.57% | 3.342% | 2.334% | 3.684% | 1.728% | 3.913% | 7.242% | 1.885% | 6.176% | 3.114% | 8.287%
| PDM3 |  | 12.51% | -29.43% | 15.32% | 48.85% | 78.92% | 8.93% | -3.034% | -6.57% | 6.605% | 6.314% | -28.8% | 51.6% | 7.912% | -24.78% | 5.588%
| Z-Score (Absolute) |  | 0.41 | 2.52 | 0.61 | 2.96 | 5.06 | 0.16 | 0.67 | 0.92 | 0 | 0.02 | 2.48 | 3.15 | 0.09 | 2.2 | 0.07
| Analytical results for P in OREAS 25b (Certified Value 0.032 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.0316 | 0.0293 | 0.0324 | 0.032112 | 0.031 | 0.034 | 0.032 | 0.032 | 0.034 | 0.031 | 0.035 | 0.0337 | 0.034 | 0.041 | 0.033 | 1
| 1 | 2 | 0.0329 | 0.029 | 0.0324 | 0.033081 | 0.031 | 0.034 | 0.031 | 0.031 | 0.033 | 0.032 | 0.034 | 0.034 | 0.032 | 0.042 | 0.0325 | #N/A
| 1 | 3 | 0.0317 | 0.0291 | 0.0322 | 0.031117 | 0.032 | 0.034 | 0.031 | 0.032 | 0.034 | 0.032 | 0.034 | 0.0333 | 0.032 | 0.042 | 0.0325 | 16
| 1 | 4 | 0.0311 | 0.0291 | 0.032 | 0.033004 | 0.032 | 0.035 | 0.031 | 0.032 | 0.033 | 0.032 | 0.035 | 0.0333 | 0.034 | 0.041 | 0.0335 | 0.032405
| 1 | 5 | 0.0314 | 0.0285 | 0.032 | 0.033216 | 0.032 | 0.035 | 0.031 | 0.033 | 0.034 | 0.031 | 0.035 | 0.0333 | 0.034 | 0.041 | 0.033 | 164
| 1 | 6 | 0.0302 | 0.0288 | 0.0319 | 0.031004 | 0.032 | 0.035 | 0.031 | 0.032 | 0.033 | 0.032 | 0.034 | 0.0338 | 0.033 | 0.042 | 0.034
| Mean |  | 0.031483 | 0.028967 | 0.03215 | 0.032256 | 0.031667 | 0.0345 | 0.031167 | 0.032 | 0.0335 | 0.031667 | 0.0345 | 0.033567 | 0.033167 | 0.0415 | 0.033083
| Median |  | 0.0315 | 0.02905 | 0.0321 | 0.032558 | 0.032 | 0.0345 | 0.031 | 0.032 | 0.0335 | 0.032 | 0.0345 | 0.0335 | 0.0335 | 0.0415 | 0.033
| Std Dev. |  | 0.00087958 | 0.00028048 | 0.00021679 | 0.0010046 | 0.0005164 | 0.00054772 | 0.00040825 | 0.00063246 | 0.00054772 | 0.0005164 | 0.00054772 | 0.00030768 | 0.00098319 | 0.00054772 | 0.00058452
| Rel.Std.Dev. |  | 2.794% | 0.9683% | 0.6743% | 3.115% | 1.631% | 1.588% | 1.31% | 1.976% | 1.635% | 1.631% | 1.588% | 0.9166% | 2.964% | 1.32% | 1.767%
| PDM3 |  | -2.845% | -10.61% | -0.7874% | -0.4615% | -2.279% | 6.465% | -3.822% | -1.25% | 3.379% | -2.279% | 6.465% | 3.584% | 2.35% | 28.07% | 2.093%
| Z-Score (Absolute) |  | 0.57 | 2.43 | 0.08 | 0 | 0.44 | 1.66 | 0.81 | 0.19 | 0.92 | 0.44 | 1.66 | 0.97 | 0.67 | 6.84 | 0.61
| Analytical results for Pb in OREAS 25b (Certified Value 18.8 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 19.1 | 18 | 19.2 | 19.6581 | 19.2 | 19.26 | 18.2 | 17.6 | 19.9 | 18.6 | 18.6 | 19.2 | 15.18 | 24.8 | 15 | 1
| 1 | 2 | 18.8 | 17.6 | 19.4 | 19.023 | 19.2 | 19.64 | 17.6 | 17.6 | 19.2 | 19 | 18.7 | 18.5 | 15.38 | 24.1 | 16 | 6
| 1 | 3 | 18.5 | 17.4 | 18.9 | 19.3821 | 18.4 | 19.52 | 18 | 17.4 | 20.2 | 20.3 | 18.2 | 18.5 | 14.42 | 23 | 15 | 16
| 1 | 4 | 18.7 | 17.4 | 18.9 | 18.9599 | 18.6 | 19.96 | 17.9 | 18 | 19.1 | 19.4 | 18.7 | 19.2 | 16.43 | 23 | 17 | 18.8036
| 1 | 5 | 19.1 | 17.9 | 18.8 | 19.0416 | 19.7 | 19.91 | 18 | 17.8 | 19 | 19.4 | 19.2 | 19.2 | 16.49 | 22.9 | 16 | 165
| 1 | 6 | 19.1 | 16.6 | 19.5 | 19.3519 | 19.1 | 19.65 | 17.8 | 17.9 | 19.6 | 19.6 | 19 | 19.3 | 15.54 | 24.3 | 16
| Mean |  | 18.8833 | 17.4833 | 19.1167 | 19.2361 | 19.0333 | 19.6567 | 17.9167 | 17.7167 | 19.5 | 19.3833 | 18.7333 | 18.9833 | 15.5733 | 23.6833 | 15.8333
| Median |  | 18.95 | 17.5 | 19.05 | 19.1967 | 19.15 | 19.645 | 17.95 | 17.7 | 19.4 | 19.4 | 18.7 | 19.2 | 15.46 | 23.55 | 16
| Std Dev. |  | 0.25626 | 0.49967 | 0.29269 | 0.27285 | 0.46762 | 0.25789 | 0.20412 | 0.22286 | 0.48166 | 0.57417 | 0.34448 | 0.37639 | 0.78709 | 0.81833 | 0.75277
| Rel.Std.Dev. |  | 1.357% | 2.858% | 1.531% | 1.418% | 2.457% | 1.312% | 1.139% | 1.258% | 2.47% | 2.962% | 1.839% | 1.983% | 5.054% | 3.455% | 4.754%
| PDM3 |  | 0.4242% | -7.021% | 1.665% | 2.3% | 1.222% | 4.537% | -4.717% | -5.78% | 3.704% | 3.083% | -0.3735% | 0.956% | -17.18% | 25.95% | -15.8%
| Z-Score (Absolute) |  | 0.13 | 1.96 | 0.17 | 0.33 | 0.07 | 0.88 | 1.39 | 1.65 | 0.67 | 0.52 | 0.33 | 0 | 4.45 | 6.13 | 4.11
| Analytical results for Pd in OREAS 25b (Indicative Value 3.67 ppb) | [16 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 01 | 04 | [13 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.5g | 01g | [13 empty] | 2
| 1 | 1 | 4 | < 10 | [13 empty] | 1
| 1 | 2 | 3 | < 10 | [13 empty] | 25
| 1 | 3 | 4 | < 10 | [13 empty] | 16
| 1 | 4 | 4 | < 10 | [13 empty] | 3.6667
| 1 | 5 | 3 | < 10 | [13 empty] | 31
| 1 | 6 | 4 | < 10
| Mean |  | 3.6667
| Median |  | 4
| Std Dev. |  | 0.5164
| Rel.Std.Dev. |  | 14.08%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pr in OREAS 25b (Indicative Value 3.49 ppm) | [16 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 3.744 | 3.3 | 3.26 | [12 empty] | 1
| 1 | 2 | 3.734 | 3.4 | 3.22 | [12 empty] | 7
| 1 | 3 | 3.689 | 3.3 | 3.08 | [12 empty] | 16
| 1 | 4 | 3.698 | 3.3 | 3.71 | [12 empty] | 3.4852
| 1 | 5 | 3.734 | 3.3 | 3.53 | [12 empty] | 27
| 1 | 6 | 3.615 | 3.4 | 3.72
| Mean |  | 3.7023 | 3.3333 | 3.42
| Median |  | 3.716 | 3.3 | 3.395
| Std Dev. |  | 0.048094 | 0.05164 | 0.27107
| Rel.Std.Dev. |  | 1.299% | 1.549% | 7.926%
| PDM3 |  | 6.229% | -4.358% | -1.871%
| Z-Score (Absolute) |  | 2.2 | 0.67 | 0
| Analytical results for Pt in OREAS 25b (Indicative Value < 1 ppb) | [16 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 01 | 04 | [13 empty] | ppb
|  |  | AR*MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.5g | 01g | [13 empty] | 2
| 1 | 1 | < 1 | < 5 | [13 empty] | 1
| 1 | 2 | < 1 | < 5 | [13 empty] | 22
| 1 | 3 | < 1 | < 5 | [13 empty] | 16
| 1 | 4 | < 1 | < 5 | [13 empty] | < 1
| 1 | 5 | < 1 | < 5 | [13 empty] | 28
| 1 | 6 | < 1 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Rb in OREAS 25b (Certified Value 15.8 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  |  | 1
| 1 | 1 | 16.95 | 13.5 | 19.9121 | 16.1 | 15 | 14.4 | 17.2 | 15.9 | 14.3 | 24.6 | 15.52 | 16 |  |  |  | 1
| 1 | 2 | 16.79 | 13.1 | 18.7669 | 16 | 14.4 | 14.2 | 16.7 | 16.8 | 14.1 | 23.5 | 15.53 | 16.9 |  |  |  | 8
| 1 | 3 | 16.4 | 12.6 | 19.871 | 16 | 14.6 | 14.6 | 17.6 | 16.2 | 13.4 | 23.9 | 14.96 | 15 |  |  |  | 16
| 1 | 4 | 16.67 | 12.8 | 18.2177 | 15.2 | 14.7 | 14.7 | 17.4 | 17 | 13.9 | 24 | 15.34 | 17.5 |  |  |  | 15.7742
| 1 | 5 | 16.63 | 12.4 | 18.3445 | 15.5 | 14.6 | 14.8 | 17.4 | 15.6 | 14.6 | 25 | 15.15 | 17.1 |  |  |  | 166
| 1 | 6 | 16.47 | 12.4 | 19.4818 | 15.9 | 14.6 | 14.8 | 17.2 | 16.6 | 14.4 | 23.6 | 15.89 | 18.5
| Mean |  | 16.6517 | 12.8 | 19.099 | 15.7833 | 14.65 | 14.5833 | 17.25 | 16.35 | 14.1167 | 24.1 | 15.3983 | 16.8333
| Median |  | 16.65 | 12.7 | 19.1244 | 15.95 | 14.6 | 14.65 | 17.3 | 16.4 | 14.2 | 23.95 | 15.43 | 17
| Std Dev. |  | 0.20262 | 0.43359 | 0.75631 | 0.35449 | 0.19748 | 0.24014 | 0.30822 | 0.54314 | 0.42622 | 0.58652 | 0.32591 | 1.2127
| Rel.Std.Dev. |  | 1.217% | 3.387% | 3.96% | 2.246% | 1.348% | 1.647% | 1.787% | 3.322% | 3.019% | 2.434% | 2.117% | 7.204%
| PDM3 |  | 5.563% | -18.85% | 21.08% | 0.05821% | -7.127% | -7.549% | 9.356% | 3.651% | -10.51% | 52.78% | -2.382% | 6.715%
| Z-Score (Absolute) |  | 0.3 | 1.69 | 1.57 | 0.15 | 0.73 | 0.77 | 0.61 | 0.15 | 1.01 | 4.17 | 0.35 | 0.4
| Analytical results for Re in OREAS 25b (Certified Value < 0.001 ppm) | [16 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 16 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.4g |  |  |  |  | 3
| 1 | 1 | < 0.001 | < 0.1 | < 0.05 | 0.004 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 |  |  |  |  | 1
| 1 | 2 | < 0.001 | < 0.1 | < 0.05 | < 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 |  |  |  |  | 16
| 1 | 3 | < 0.001 | < 0.1 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 |  |  |  |  | 16
| 1 | 4 | < 0.001 | < 0.1 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 |  |  |  |  | < 0.001
| 1 | 5 | < 0.001 | < 0.1 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 |  |  |  |  | 167
| 1 | 6 | < 0.001 | < 0.1 | < 0.05 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05
| Mean |  |  |  |  | 0.004 |  | 0.001
| Median |  |  |  |  | 0.004 |  | 0.001
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 399.87 | 197.57 | 4.05 | 0.67 | 0 | 0.67 | 0.67 | 0.67 | 197.57 | 197.57
| Analytical results for S in OREAS 25b (Certified Value 0.022 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | < 0.05 | 0.02 | 0.02 | 0.024556 | 0.021 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0225 | 1
| 1 | 2 | < 0.05 | 0.02 | 0.02 | 0.020544 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0215 | 17
| 1 | 3 | < 0.05 | 0.02 | 0.02 | 0.023008 | 0.021 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0215 | 16
| 1 | 4 | < 0.05 | 0.02 | 0.02 | 0.021878 | 0.021 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0225 | 0.021894
| 1 | 5 | < 0.05 | 0.02 | 0.02 | 0.02604 | 0.021 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.023 | 168
| 1 | 6 | < 0.05 | 0.02 | 0.02 | 0.023572 | 0.022 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0225
| Mean |  |  | 0.02 | 0.02 | 0.023266 | 0.021 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02225
| Median |  |  | 0.02 | 0.02 | 0.02329 | 0.021 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | 0.0225
| Std Dev. |  |  | 0 | 0 | 0.001942 | 0.00063246 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.00061237
| Rel.Std.Dev. |  |  | 0% | 0% | 8.347% | 3.012% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 2.752%
| PDM3 |  |  | -8.651% | -8.651% | 6.268% | -4.083% | -8.651% | 37.02% | -8.651% | -8.651% | -8.651% | 37.02% | -8.651% | -8.651% | -8.651% | 1.626%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Sb in OREAS 25b (Certified Value 0.42 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.4g |  | 2
| 1 | 1 | 0.51 | 0.47 | < 5 | 0.55591 | 0.05 | 0.1 | 0.38 | 0.39 | 0.36 | 0.23 | 0.71 | 0.47 | 0.49 | 0.7 |  | 1
| 1 | 2 | 0.51 | 0.47 | < 5 | 0.52503 | 0.06 | < 0.1 | 0.37 | 0.37 | 0.31 | 0.24 | 0.7 | 0.47 | 0.48 | 0.7 |  | 9
| 1 | 3 | 0.47 | 0.49 | < 5 | 0.58968 | 0.05 | < 0.1 | 0.35 | 0.35 | 0.32 | 0.27 | 0.98 | 0.48 | 0.48 | 0.7 |  | 16
| 1 | 4 | 0.49 | 0.49 | < 5 | 0.57915 | 0.06 | 0.1 | 0.37 | 0.35 | 0.34 | 0.25 | 0.8 | 0.4 | 0.47 | 0.7 |  | 0.42152
| 1 | 5 | 0.51 | 0.49 | < 5 | 0.52908 | 0.05 | 0.1 | 0.37 | 0.35 | 0.32 | 0.26 | 0.82 | 0.47 | 0.48 | 0.7 |  | 169
| 1 | 6 | 0.5 | 0.49 | < 5 | 0.54911 | 0.06 | 0.1 | 0.36 | 0.36 | 0.35 | 0.25 | 0.82 | 0.43 | 0.49 | 0.7
| Mean |  | 0.49833 | 0.48333 |  | 0.55466 | 0.055 | 0.1 | 0.36667 | 0.36167 | 0.33333 | 0.25 | 0.805 | 0.45333 | 0.48167 | 0.7
| Median |  | 0.505 | 0.49 |  | 0.55251 | 0.055 | 0.1 | 0.37 | 0.355 | 0.33 | 0.25 | 0.81 | 0.47 | 0.48 | 0.7
| Std Dev. |  | 0.016021 | 0.010328 |  | 0.026041 | 0.0054772 | 0 | 0.010328 | 0.016021 | 0.019664 | 0.014142 | 0.10114 | 0.031411 | 0.0075277 | 1.2162e-16
| Rel.Std.Dev. |  | 3.215% | 2.137% |  | 4.695% | 9.959% | 0% | 2.817% | 4.43% | 5.899% | 5.657% | 12.56% | 6.929% | 1.563% | 1.737e-14%
| PDM3 |  | 18.22% | 14.67% |  | 31.59% | -86.95% | -76.28% | -13.01% | -14.2% | -20.92% | -40.69% | 90.98% | 7.548% | 14.27% | 66.07%
| Z-Score (Absolute) |  | 0.2 | 0.1 | 13.26 | 0.57 | 2.69 | NA | 0.66 | 0.69 | 0.88 | 1.42 | 2.2 | 0.09 | 0.09 | NA
Results from laboratories 6 and 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 25b (Certified Value 10 ppm) | [16 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 10.37 | 9.5 | 10 | 9.5485 | 9.8 | 10.9 | 10.6 | 10 | 10.5 | 10.2 | 9.5 | 10 | 10.5 | 10 | 10.1 | 1
| 1 | 2 | 10.2 | 9.5 | 9.9 | 9.3266 | 9.8 | 10.8 | 10.2 | 10 | 10.3 | 10.4 | 9.3 | 10 | 9.9 | 10.2 | 9.9 | #N/A
| 1 | 3 | 10 | 9.3 | 9.9 | 9.4011 | 9.3 | 10.8 | 10.2 | 9.9 | 10.4 | 10.3 | 8.9 | 10 | 9.8 | 9.7 | 9.9 | 16
| 1 | 4 | 10.06 | 9.3 | 9.9 | 9.4145 | 9.5 | 11.3 | 10.3 | 10.2 | 10.7 | 10.7 | 9.3 | 10 | 10.5 | 10.1 | 10.1 | 10.0007
| 1 | 5 | 10 | 9.1 | 9.8 | 9.327 | 9.2 | 11 | 10.3 | 10 | 10.2 | 10.2 | 9.7 | 10 | 10.8 | 9.6 | 10.5 | 170
| 1 | 6 | 9.97 | 8.9 | 9.8 | 9.5282 | 9.4 | 11.2 | 10.2 | 10.1 | 10.4 | 10.8 | 9.6 | 10 | 10.3 | 10 | 10.4
| Mean |  | 10.1 | 9.2667 | 9.8833 | 9.4243 | 9.5 | 11 | 10.3 | 10.0333 | 10.4167 | 10.4333 | 9.3833 | 10 | 10.3 | 9.9333 | 10.15
| Median |  | 10.03 | 9.3 | 9.9 | 9.4078 | 9.45 | 10.95 | 10.25 | 10 | 10.4 | 10.35 | 9.4 | 10 | 10.4 | 10 | 10.1
| Std Dev. |  | 0.15582 | 0.23381 | 0.075277 | 0.095772 | 0.25298 | 0.20976 | 0.15492 | 0.10328 | 0.17224 | 0.2582 | 0.28577 | 0 | 0.38471 | 0.23381 | 0.251
| Rel.Std.Dev. |  | 1.543% | 2.523% | 0.7617% | 1.016% | 2.663% | 1.907% | 1.504% | 1.029% | 1.654% | 2.475% | 3.046% | 0% | 3.735% | 2.354% | 2.473%
| PDM3 |  | 0.993% | -7.34% | -1.173% | -5.763% | -5.007% | 9.992% | 2.993% | 0.3264% | 4.159% | 4.326% | -6.173% | -0.006893% | 2.993% | -0.6735% | 1.493%
| Z-Score (Absolute) |  | 0.17 | 1.94 | 0.38 | 1.54 | 1.35 | 2.44 | 0.67 | 0 | 0.97 | 1.01 | 1.64 | 0.08 | 0.67 | 0.25 | 0.3
| Analytical results for Se in OREAS 25b (Certified Value 0.43 ppm) | [16 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.4g |  | 2
| 1 | 1 | 0.37 | < 1 | < 5 | < 1 | 0.6 | 0.5 | 0.5 | 0.5 | 0.5 | 0.3 | 0.5 | 1 | 0.8 | < 1 |  | 1
| 1 | 2 | 0.34 | < 1 | < 5 | < 1 | 0.5 | 0.5 | 0.4 | 0.6 | 0.3 | 0.3 | 0.6 | 1 | 0.8 | < 1 |  | 14
| 1 | 3 | 0.38 | < 1 | < 5 | < 1 | 0.5 | 0.4 | 0.4 | 0.3 | 0.5 | < 0.2 | 0.6 | 1 | 0.7 | < 1 |  | 16
| 1 | 4 | 0.33 | < 1 | < 5 | < 1 | 0.5 | 0.5 | 0.4 | 0.3 | 0.4 | 0.2 | 0.5 | 1 | 0.8 | < 1 |  | 0.43208
| 1 | 5 | 0.39 | < 1 | < 5 | < 1 | 0.6 | 0.5 | 0.5 | 0.4 | 0.5 | 0.2 | 0.5 | 1 | 0.9 | < 1 |  | 171
| 1 | 6 | 0.37 | < 1 | < 5 | < 1 | 0.5 | 0.5 | 0.3 | 0.5 | 0.4 | 0.3 | 0.5 | 1 | 0.8 | < 1
| Mean |  | 0.36333 |  |  |  | 0.53333 | 0.48333 | 0.41667 | 0.43333 | 0.43333 | 0.26 | 0.53333 | 1 | 0.8
| Median |  | 0.37 |  |  |  | 0.5 | 0.5 | 0.4 | 0.45 | 0.45 | 0.3 | 0.5 | 1 | 0.8
| Std Dev. |  | 0.023381 |  |  |  | 0.05164 | 0.040825 | 0.075277 | 0.12111 | 0.08165 | 0.054772 | 0.05164 | 0 | 0.063246
| Rel.Std.Dev. |  | 6.435% |  |  |  | 9.682% | 8.447% | 18.07% | 27.95% | 18.84% | 21.07% | 9.682% | 0% | 7.906%
| PDM3 |  | -15.91% |  |  |  | 23.43% | 11.86% | -3.568% | 0.2893% | 0.2893% | -39.83% | 23.43% | 131.4% | 85.15%
| Z-Score (Absolute) |  | 1.38 | 0 | 20.23 | 0 | 0.34 | 0.17 | 0.84 | 0.67 | 0.67 | 2.7 | 0.34 | NA | 3.03 | 0
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 25b (Indicative Value 2.56 ppm) | [16 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 05 | 16 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.4g | [12 empty] | 2
| 1 | 1 | 2.644 | 2.3 | 2.44 | [12 empty] | 1
| 1 | 2 | 2.592 | 2.5 | 2.53 | [12 empty] | 10
| 1 | 3 | 2.534 | 2.6 | 2.38 | [12 empty] | 16
| 1 | 4 | 2.594 | 2.6 | 2.85 | [12 empty] | 2.5556
| 1 | 5 | 2.559 | 2.6 | 2.59 | [12 empty] | 29
| 1 | 6 | 2.558 | 2.3 | 2.83
| Mean |  | 2.5802 | 2.4833 | 2.6033
| Median |  | 2.5755 | 2.55 | 2.56
| Std Dev. |  | 0.03867 | 0.1472 | 0.19715
| Rel.Std.Dev. |  | 1.499% | 5.927% | 7.573%
| PDM3 |  | 0.9608% | -2.828% | 1.867%
| Z-Score (Absolute) |  | 0 | 2.82 | 0.67
| Analytical results for Sn in OREAS 25b (Certified Value 2.26 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  | 3
| 1 | 1 | 2.4 | 2.3 | < 10 | 2.555 | 2.05 | 2.4 | 2.2 | 2.4 | 2.3 | 2.1 | 3 | 1.9 | 2.3 |  |  | 1
| 1 | 2 | 2.4 | 2.2 | < 10 | 2.5578 | 2.14 | 2.3 | 2.3 | 2.4 | 2.4 | 2.1 | 3 | 1.9 | 2.3 |  |  | 11
| 1 | 3 | 2.4 | 2.1 | < 10 | 2.5571 | 2.04 | 2.2 | 2.3 | 2.5 | 2.3 | 2 | 2.7 | 1.8 | 2.3 |  |  | 16
| 1 | 4 | 2.36 | 2.2 | < 10 | 2.4494 | 1.98 | 2.3 | 2.4 | 2.5 | 2.4 | 2.1 | 2.7 | 2 | 2.5 |  |  | 2.2639
| 1 | 5 | 2.47 | 2.1 | < 10 | 2.4314 | 1.97 | 2.4 | 2.4 | 2.5 | 2.2 | 2.2 | 3 | 1.7 | 2.5 |  |  | 172
| 1 | 6 | 2.42 | 2.1 | < 10 | 2.4667 | 1.97 | 2.4 | 2.4 | 2.4 | 2.4 | 2.2 | 2.9 | 1.8 | 2.4
| Mean |  | 2.4083 | 2.1667 |  | 2.5029 | 2.025 | 2.3333 | 2.3333 | 2.45 | 2.3333 | 2.1167 | 2.8833 | 1.85 | 2.3833
| Median |  | 2.4 | 2.15 |  | 2.5109 | 2.01 | 2.35 | 2.35 | 2.45 | 2.35 | 2.1 | 2.95 | 1.85 | 2.35
| Std Dev. |  | 0.036009 | 0.08165 |  | 0.059906 | 0.066558 | 0.08165 | 0.08165 | 0.054772 | 0.08165 | 0.075277 | 0.1472 | 0.10488 | 0.098319
| Rel.Std.Dev. |  | 1.495% | 3.768% |  | 2.393% | 3.287% | 3.499% | 3.499% | 2.236% | 3.499% | 3.556% | 5.105% | 5.669% | 4.125%
| PDM3 |  | 6.38% | -4.295% |  | 10.56% | -10.55% | 3.067% | 3.067% | 8.22% | 3.067% | -6.504% | 27.36% | -18.28% | 5.275%
| Z-Score (Absolute) |  | 0.3 | 0.67 | 10.79 | 0.69 | 1.25 | 0 | 0 | 0.47 | 0 | 0.88 | 2.23 | 1.96 | 0.2
| Analytical results for Sr in OREAS 25b (Certified Value 13.3 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 13.88 | 11.5 | 22.5 | 15.2122 | 14.3 | 13.9 | 12.4 | 12.6 | 14.1 | 12.6 | 11.8 | 15.1 | 13.6 | 12.9 | 13.1 | 1
| 1 | 2 | 13.62 | 11.1 | 22.7 | 15.2411 | 14.5 | 13.8 | 12 | 12.6 | 13.5 | 13 | 11.5 | 14.9 | 12.9 | 13.5 | 12.9 | #N/A
| 1 | 3 | 13.38 | 11 | 22.8 | 15.2757 | 14.8 | 13.5 | 12.1 | 12.7 | 14.2 | 12.5 | 11.2 | 15.1 | 12.5 | 13 | 12.5 | 16
| 1 | 4 | 13.41 | 11.2 | 23 | 14.9363 | 13.7 | 14.5 | 12.2 | 13.2 | 14.5 | 13.1 | 11.7 | 15 | 13.8 | 12.9 | 13.2 | 13.2966
| 1 | 5 | 13.59 | 11 | 22.7 | 14.9515 | 14 | 14.3 | 12.2 | 12.9 | 14.4 | 12.3 | 11.9 | 15.1 | 14.1 | 12.7 | 13.8 | 173
| 1 | 6 | 13.62 | 11 | 23 | 15.0407 | 14.1 | 13.8 | 12.3 | 13 | 13.8 | 12.9 | 11.9 | 15.2 | 13.1 | 13.4 | 13.8
| Mean |  | 13.5833 | 11.1333 | 22.7833 | 15.1096 | 14.2333 | 13.9667 | 12.2 | 12.8333 | 14.0833 | 12.7333 | 11.6667 | 15.0667 | 13.3333 | 13.0667 | 13.2167
| Median |  | 13.605 | 11.05 | 22.75 | 15.1264 | 14.2 | 13.85 | 12.2 | 12.8 | 14.15 | 12.75 | 11.75 | 15.1 | 13.35 | 12.95 | 13.15
| Std Dev. |  | 0.18007 | 0.19664 | 0.19408 | 0.15177 | 0.38816 | 0.36697 | 0.14142 | 0.24221 | 0.37639 | 0.31411 | 0.27325 | 0.10328 | 0.60222 | 0.31411 | 0.51153
| Rel.Std.Dev. |  | 1.326% | 1.766% | 0.8518% | 1.004% | 2.727% | 2.627% | 1.159% | 1.887% | 2.673% | 2.467% | 2.342% | 0.6855% | 4.517% | 2.404% | 3.87%
| PDM3 |  | 2.156% | -16.27% | 71.35% | 13.63% | 7.045% | 5.039% | -8.247% | -3.484% | 5.917% | -4.236% | -12.26% | 13.31% | 0.276% | -1.73% | -0.6014%
| Z-Score (Absolute) |  | 0.22 | 1.98 | 8.5 | 1.6 | 0.81 | 0.57 | 1.02 | 0.45 | 0.67 | 0.54 | 1.5 | 1.56 | 0 | 0.24 | 0.1
| Analytical results for Ta in OREAS 25b (Certified Value < 0.01 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  | 3
| 1 | 1 | < 0.01 | < 0.05 | 5.4 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.13 | < 0.05 |  |  | 1
| 1 | 2 | < 0.01 | < 0.05 | 5.4 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.14 | < 0.05 |  |  | 1
| 1 | 3 | < 0.01 | < 0.05 | 5.1 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.12 | < 0.05 |  |  | 16
| 1 | 4 | < 0.01 | < 0.05 | 5.2 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.11 | < 0.05 |  |  | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | 5 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.1 | < 0.05 |  |  | 174
| 1 | 6 | < 0.01 | < 0.05 | 5 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.12 | < 0.05
| Mean |  |  |  | 5.1833 | [8 empty] | 0.12
| Median |  |  |  | 5.15 | [8 empty] | 0.12
| Std Dev. |  |  |  | 0.18348 | [8 empty] | 0.014142
| Rel.Std.Dev. |  |  |  | 3.54% | [8 empty] | 11.79%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 173.92 | 0 | 0 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0 | 3.2 | 0
| Analytical results for Tb in OREAS 25b (Certified Value 0.24 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 05 | 15 | 16 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 0.25g | 0.5g | 0.25g | 0.4g | [10 empty] | 2
| 1 | 1 | 0.255 | 0.23 | 0.2 | 0.28 | 0.24 | [10 empty] | 1
| 1 | 2 | 0.255 | 0.22 | 0.2 | 0.28 | 0.25 | [10 empty] | #N/A
| 1 | 3 | 0.255 | 0.23 | 0.2 | 0.27 | 0.24 | [10 empty] | 16
| 1 | 4 | 0.25 | 0.21 | 0.2 | 0.27 | 0.29 | [10 empty] | 0.2408
| 1 | 5 | 0.256 | 0.21 | 0.2 | 0.25 | 0.26 | [10 empty] | 175
| 1 | 6 | 0.253 | 0.21 | 0.2 | 0.28 | 0.28
| Mean |  | 0.254 | 0.21833 | 0.2 | 0.27167 | 0.26
| Median |  | 0.255 | 0.215 | 0.2 | 0.275 | 0.255
| Std Dev. |  | 0.0021909 | 0.0098319 | 3.0405e-17 | 0.01169 | 0.020976
| Rel.Std.Dev. |  | 0.8626% | 4.503% | 1.52e-14% | 4.303% | 8.068%
| PDM3 |  | 5.482% | -9.33% | -16.94% | 12.82% | 7.973%
| Z-Score (Absolute) |  | 0 | 1.36 | 2.06 | 0.67 | 0.23
| Analytical results for Te in OREAS 25b (Indicative Value 0.024 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.4g |  | 3
| 1 | 1 | < 0.02 | < 0.05 | 8.5 | < 0.05 | 0.03 | < 0.02 | 0.02 | < 0.01 | 0.04 | 0.14 | 0.01 | < 0.05 | 0.03 | < 0.1 |  | 1
| 1 | 2 | < 0.02 | < 0.05 | 8.8 | < 0.05 | < 0.02 | < 0.02 | 0.04 | 0.02 | 0.05 | 0.13 | < 0.01 | < 0.05 | 0.02 | < 0.1 |  | 15
| 1 | 3 | < 0.02 | < 0.05 | 9.1 | < 0.05 | 0.02 | < 0.02 | 0.03 | 0.03 | 0.05 | 0.13 | 0.02 | < 0.05 | 0.02 | < 0.1 |  | 16
| 1 | 4 | < 0.02 | < 0.05 | 8.7 | < 0.05 | < 0.02 | < 0.02 | 0.02 | 0.04 | 0.05 | 0.11 | 0.01 | < 0.05 | 0.02 | < 0.1 |  | 0.024306
| 1 | 5 | < 0.02 | < 0.05 | 8.5 | < 0.05 | < 0.02 | < 0.02 | 0.02 | 0.02 | 0.05 | 0.16 | 0.03 | < 0.05 | 0.02 | < 0.1 |  | 30
| 1 | 6 | < 0.02 | < 0.05 | 8.8 | < 0.05 | < 0.02 | < 0.02 | 0.03 | 0.02 | 0.04 | 0.16 | < 0.01 | < 0.05 | 0.02 | < 0.1
| Mean |  |  |  | 8.7333 |  | 0.025 |  | 0.026667 | 0.026 | 0.046667 | 0.13833 | 0.0175 |  | 0.021667
| Median |  |  |  | 8.75 |  | 0.025 |  | 0.025 | 0.02 | 0.05 | 0.135 | 0.015 |  | 0.02
| Std Dev. |  |  |  | 0.22509 |  | 0.0070711 |  | 0.008165 | 0.0089443 | 0.005164 | 0.019408 | 0.0095743 |  | 0.0040825
| Rel.Std.Dev. |  |  |  | 2.577% |  | 28.28% |  | 30.62% | 34.4% | 11.07% | 14.03% | 54.71% |  | 18.84%
| PDM3 |  |  |  | 35831% |  | 2.857% |  | 9.714% | 6.971% | 92% | 469.1% | -28% |  | -10.86%
| Z-Score (Absolute) |  | 0.93 | 0 | 542.04 | 0 | 0.62 | 0.93 | 0.1 | 0.16 | 1.35 | 7.05 | 0.73 | 0 | 0.21 | 1.56
| Analytical results for Th in OREAS 25b (Certified Value 9.12 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g |  | 3
| 1 | 1 | 9.842 | 8.2 | 9.41 | 9 | 9.8 | 8.3 | 9.1 | 9.2 | 8.9 | 9.4 | 10.9 | 8.7 | 9.11 | 9.14 |  | 1
| 1 | 2 | 9.811 | 8.1 | 9.6141 | 9 | 9.7 | 8 | 9.2 | 9.1 | 9 | 9.5 | 10.9 | 8.9 | 9.74 | 9.09 |  | 2
| 1 | 3 | 9.46 | 8.2 | 9.516 | 8.9 | 9.8 | 8.2 | 9.2 | 9.4 | 9 | 9.3 | 11.1 | 8.7 | 9.17 | 8.97 |  | 16
| 1 | 4 | 9.661 | 8.1 | 9.2691 | 8.6 | 10 | 8.2 | 9.3 | 9 | 9.4 | 9.6 | 11 | 9 | 9.26 | 9.93 |  | 9.1163
| 1 | 5 | 9.713 | 8 | 9.463 | 8.6 | 10.2 | 8.1 | 9.4 | 8.8 | 8.8 | 9.6 | 11.3 | 9.9 | 8.97 | 9.65 |  | 176
| 1 | 6 | 9.71 | 7.9 | 9.2249 | 8.9 | 10.1 | 8 | 9.5 | 9.1 | 9.2 | 9.6 | 11.1 | 8.8 | 9.48 | 9.45
| Mean |  | 9.6995 | 8.0833 | 9.4162 | 8.8333 | 9.9333 | 8.1333 | 9.2833 | 9.1 | 9.05 | 9.5 | 11.05 | 9 | 9.2883 | 9.3717
| Median |  | 9.7115 | 8.1 | 9.4365 | 8.9 | 9.9 | 8.15 | 9.25 | 9.1 | 9 | 9.55 | 11.05 | 8.85 | 9.215 | 9.295
| Std Dev. |  | 0.13555 | 0.1169 | 0.14805 | 0.18619 | 0.19664 | 0.12111 | 0.1472 | 0.2 | 0.21679 | 0.12649 | 0.15166 | 0.45607 | 0.27881 | 0.3714
| Rel.Std.Dev. |  | 1.398% | 1.446% | 1.572% | 2.108% | 1.98% | 1.489% | 1.586% | 2.198% | 2.396% | 1.331% | 1.372% | 5.067% | 3.002% | 3.963%
| PDM3 |  | 6.397% | -11.33% | 3.289% | -3.104% | 8.962% | -10.78% | 1.832% | -0.1792% | -0.7277% | 4.209% | 21.21% | -1.276% | 1.887% | 2.801%
| Z-Score (Absolute) |  | 1.07 | 3.11 | 0.34 | 1.17 | 1.67 | 2.98 | 0.01 | 0.48 | 0.61 | 0.55 | 4.56 | 0.74 | 0.01 | 0.22
| Analytical results for Ti in OREAS 25b (Certified Value 0.105 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.1122 | 0.09 | 0.1 | 0.23401 | 0.11 | 0.086 | 0.11 | 0.101 | 0.125 | 0.11 | 0.1 | 0.182 | 0.099 | 0.097 | 0.129 | 1
| 1 | 2 | 0.1106 | 0.09 | 0.1 | 0.25016 | 0.12 | 0.086 | 0.105 | 0.099 | 0.119 | 0.115 | 0.099 | 0.171 | 0.101 | 0.102 | 0.122 | #N/A
| 1 | 3 | 0.1104 | 0.09 | 0.1 | 0.23158 | 0.12 | 0.089 | 0.106 | 0.104 | 0.12 | 0.115 | 0.09 | 0.169 | 0.095 | 0.095 | 0.113 | 16
| 1 | 4 | 0.1118 | 0.09 | 0.1 | 0.21445 | 0.11 | 0.096 | 0.108 | 0.101 | 0.125 | 0.115 | 0.098 | 0.165 | 0.107 | 0.096 | 0.132 | 0.10499
| 1 | 5 | 0.1125 | 0.08 | 0.1 | 0.24452 | 0.11 | 0.088 | 0.11 | 0.106 | 0.124 | 0.105 | 0.104 | 0.168 | 0.109 | 0.092 | 0.15 | 177
| 1 | 6 | 0.1108 | 0.08 | 0.1 | 0.21979 | 0.11 | 0.089 | 0.105 | 0.099 | 0.126 | 0.112 | 0.108 | 0.173 | 0.099 | 0.095 | 0.146
| Mean |  | 0.11138 | 0.086667 | 0.1 | 0.23242 | 0.11333 | 0.089 | 0.10733 | 0.10167 | 0.12317 | 0.112 | 0.099833 | 0.17133 | 0.10167 | 0.096167 | 0.132
| Median |  | 0.1113 | 0.09 | 0.1 | 0.2328 | 0.11 | 0.0885 | 0.107 | 0.101 | 0.1245 | 0.1135 | 0.0995 | 0.17 | 0.1 | 0.0955 | 0.1305
| Std Dev. |  | 0.00089536 | 0.005164 | 1.5202e-17 | 0.013761 | 0.005164 | 0.0036878 | 0.0023381 | 0.0028048 | 0.0029269 | 0.004 | 0.00608 | 0.0058878 | 0.0053166 | 0.0033116 | 0.014071
| Rel.Std.Dev. |  | 0.8039% | 5.958% | 1.52e-14% | 5.921% | 4.556% | 4.144% | 2.178% | 2.759% | 2.376% | 3.571% | 6.09% | 3.436% | 5.229% | 3.444% | 10.66%
| PDM3 |  | 6.094% | -17.45% | -4.749% | 121.4% | 7.951% | -15.23% | 2.236% | -3.162% | 17.32% | 6.681% | -4.908% | 63.2% | -3.162% | -8.4% | 25.73%
| Z-Score (Absolute) |  | 0.36 | 1.86 | 0.66 | 11.25 | 0.54 | 1.65 | 0 | 0.51 | 1.42 | 0.42 | 0.67 | 5.75 | 0.51 | 1 | 2.22
| Analytical results for Tl in OREAS 25b (Certified Value 0.13 ppm) | [16 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 3
| 1 | 1 | 0.14 | 0.1 | < 10 | 0.14774 | 0.12 | 0.13 | 0.11 | 0.13 | 0.13 | 0.13 | 0.14 | 0.17 | 0.13 | 0.14 | 0.1 | 1
| 1 | 2 | 0.14 | 0.1 | < 10 | 0.15659 | 0.13 | 0.14 | 0.11 | 0.11 | 0.11 | 0.13 | 0.14 | 0.18 | 0.13 | 0.13 | 0.1 | #N/A
| 1 | 3 | 0.14 | 0.1 | < 10 | 0.16118 | 0.12 | 0.14 | 0.11 | 0.12 | 0.13 | 0.13 | 0.13 | 0.17 | 0.12 | 0.13 | 0.1 | 16
| 1 | 4 | 0.14 | 0.1 | < 10 | 0.14036 | 0.12 | 0.14 | 0.12 | 0.13 | 0.12 | 0.13 | 0.13 | 0.18 | 0.13 | 0.13 | 0.1 | 0.13099
| 1 | 5 | 0.14 | 0.1 | < 10 | 0.14285 | 0.12 | 0.14 | 0.12 | 0.13 | 0.12 | 0.13 | 0.14 | 0.17 | 0.15 | 0.13 | 0.1 | 178
| 1 | 6 | 0.14 | 0.1 | < 10 | 0.13683 | 0.12 | 0.14 | 0.12 | 0.13 | 0.12 | 0.13 | 0.14 | 0.18 | 0.13 | 0.14 | 0.1
| Mean |  | 0.14 | 0.1 |  | 0.14759 | 0.12167 | 0.13833 | 0.115 | 0.125 | 0.12167 | 0.13 | 0.13667 | 0.175 | 0.13167 | 0.13333 | 0.1
| Median |  | 0.14 | 0.1 |  | 0.1453 | 0.12 | 0.14 | 0.115 | 0.13 | 0.12 | 0.13 | 0.14 | 0.175 | 0.13 | 0.13 | 0.1
| Std Dev. |  | 0 | 1.5202e-17 |  | 0.0095537 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0083666 | 0.0075277 | 0 | 0.005164 | 0.0054772 | 0.0098319 | 0.005164 | 1.5202e-17
| Rel.Std.Dev. |  | 0% | 1.52e-14% |  | 6.473% | 3.355% | 2.951% | 4.763% | 6.693% | 6.187% | 0% | 3.779% | 3.13% | 7.467% | 3.873% | 1.52e-14%
| PDM3 |  | 6.876% | -23.66% |  | 12.67% | -7.12% | 5.603% | -12.21% | -4.575% | -7.12% | -0.7583% | 4.331% | 33.59% | 0.5141% | 1.786% | -23.66%
| Z-Score (Absolute) |  | 0.55 | 2.39 | 358.06 | 1.11 | 0.8 | 0.43 | 1.29 | 0.55 | 0.8 | 0.18 | 0.31 | 3.13 | 0.06 | 0.06 | NA
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 25b (Indicative Value 0.075 ppm) | [16 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 01 | 05 | [13 empty] | ppm
|  |  | AR*MS | AR*OES/MS | [13 empty] | 3
|  |  | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | 0.075 | < 0.1 | [13 empty] | 1
| 1 | 2 | 0.076 | < 0.1 | [13 empty] | 25
| 1 | 3 | 0.074 | < 0.1 | [13 empty] | 16
| 1 | 4 | 0.072 | < 0.1 | [13 empty] | 0.074833
| 1 | 5 | 0.076 | < 0.1 | [13 empty] | 31
| 1 | 6 | 0.076 | < 0.1
| Mean |  | 0.074833
| Median |  | 0.0755
| Std Dev. |  | 0.0016021
| Rel.Std.Dev. |  | 2.141%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 25b (Certified Value 1.24 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g |  | 3
| 1 | 1 | 1.31 | 1.19 | 1.2249 | 1.1 | 1.3 | 1.12 | 1.16 | 1.19 | 1.23 | 1.3 | 1.42 | 1.17 | 1.23 | 1.29 |  | 1
| 1 | 2 | 1.332 | 1.17 | 1.2727 | 1.1 | 1.3 | 1.08 | 1.18 | 1.16 | 1.25 | 1.31 | 1.42 | 1.17 | 1.34 | 1.29 |  | 4
| 1 | 3 | 1.298 | 1.17 | 1.2339 | 1.1 | 1.3 | 1.1 | 1.17 | 1.2 | 1.26 | 1.28 | 1.42 | 1.14 | 1.3 | 1.29 |  | 16
| 1 | 4 | 1.318 | 1.19 | 1.2079 | 1.1 | 1.3 | 1.1 | 1.17 | 1.16 | 1.26 | 1.35 | 1.39 | 1.2 | 1.23 | 1.41 |  | 1.238
| 1 | 5 | 1.316 | 1.18 | 1.2603 | 1.1 | 1.3 | 1.11 | 1.19 | 1.13 | 1.22 | 1.33 | 1.42 | 1.28 | 1.2 | 1.37 |  | 179
| 1 | 6 | 1.299 | 1.15 | 1.2722 | 1.1 | 1.3 | 1.1 | 1.18 | 1.16 | 1.27 | 1.33 | 1.46 | 1.17 | 1.27 | 1.38
| Mean |  | 1.3122 | 1.175 | 1.2453 | 1.1 | 1.3 | 1.1017 | 1.175 | 1.1667 | 1.2483 | 1.3167 | 1.4217 | 1.1883 | 1.2617 | 1.3383
| Median |  | 1.313 | 1.175 | 1.2471 | 1.1 | 1.3 | 1.1 | 1.175 | 1.16 | 1.255 | 1.32 | 1.42 | 1.17 | 1.25 | 1.33
| Std Dev. |  | 0.012813 | 0.015166 | 0.026973 | 2.4324e-16 | 0 | 0.013292 | 0.010488 | 0.025033 | 0.019408 | 0.025033 | 0.022286 | 0.048751 | 0.051929 | 0.054559
| Rel.Std.Dev. |  | 0.9765% | 1.291% | 2.166% | 2.211e-14% | 0% | 1.206% | 0.8926% | 2.146% | 1.555% | 1.901% | 1.568% | 4.102% | 4.116% | 4.077%
| PDM3 |  | 5.988% | -5.091% | 0.5877% | -11.15% | 5.005% | -11.01% | -5.091% | -5.765% | 0.8319% | 6.351% | 14.83% | -4.014% | 1.909% | 8.102%
| Z-Score (Absolute) |  | 0.62 | 0.68 | 0.01 | 1.4 | 0.51 | 1.38 | 0.68 | 0.76 | 0.01 | 0.66 | 1.66 | 0.56 | 0.14 | 0.87
| Analytical results for V in OREAS 25b (Certified Value 136 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 0
| 1 | 1 | 142.3 | 130 | 131 | 140.5002 | 107 | 136 | 137 | 141 | 141 | 132 | 133 | 143 | 138 | 133 | 136 | 1
| 1 | 2 | 141 | 129 | 130 | 144.7622 | 106 | 136 | 131 | 137 | 139 | 131 | 129 | 141 | 130 | 138 | 133 | #N/A
| 1 | 3 | 137.5 | 129 | 131 | 136.7455 | 108 | 137 | 133 | 138 | 140 | 135 | 127 | 139 | 129 | 135 | 135 | 16
| 1 | 4 | 139 | 129 | 130 | 147.1098 | 101 | 139 | 133 | 140 | 140 | 135 | 132 | 140 | 137 | 134 | 138 | 135.8591
| 1 | 5 | 138.8 | 128 | 130 | 144.7832 | 101 | 140 | 133 | 141 | 141 | 132 | 132 | 141 | 139 | 133 | 140 | 180
| 1 | 6 | 138.2 | 128 | 130 | 137.2643 | 106 | 139 | 133 | 138 | 139 | 133 | 132 | 139 | 136 | 137 | 140
| Mean |  | 139.4667 | 128.8333 | 130.3333 | 141.8609 | 104.8333 | 137.8333 | 133.3333 | 139.1667 | 140 | 133 | 130.8333 | 140.5 | 134.8333 | 135 | 137
| Median |  | 138.9 | 129 | 130 | 142.6312 | 106 | 138 | 133 | 139 | 140 | 132.5 | 132 | 140.5 | 136.5 | 134.5 | 137
| Std Dev. |  | 1.8173 | 0.75277 | 0.5164 | 4.3279 | 3.0605 | 1.7224 | 1.9664 | 1.7224 | 0.89443 | 1.6733 | 2.3166 | 1.5166 | 4.2622 | 2.0976 | 2.8284
| Rel.Std.Dev. |  | 1.303% | 0.5843% | 0.3962% | 3.051% | 2.919% | 1.25% | 1.475% | 1.238% | 0.6389% | 1.258% | 1.771% | 1.079% | 3.161% | 1.554% | 2.065%
| PDM3 |  | 2.655% | -5.171% | -4.067% | 4.418% | -22.84% | 1.453% | -1.859% | 2.435% | 3.048% | -2.104% | -3.699% | 3.416% | -0.755% | -0.6324% | 0.8398%
| Z-Score (Absolute) |  | 0.72 | 1 | 0.76 | 1.11 | 4.88 | 0.46 | 0.27 | 0.67 | 0.81 | 0.32 | 0.67 | 0.89 | 0.03 | 0 | 0.32
| Analytical results for W in OREAS 25b (Certified Value < 0.05 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 16 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  | 3
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.4g |  | 3
| 1 | 1 | 0.04 | 0.1 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.24 | 0.1 | < 0.05 | 0.2 |  | 1
| 1 | 2 | 0.02 | < 0.1 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.23 | < 0.1 | < 0.05 | 0.1 |  | 6
| 1 | 3 | < 0.02 | 0.1 | < 5 | < 0.1 | 0.4 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.23 | < 0.1 | < 0.05 | 0.2 |  | 16
| 1 | 4 | 0.02 | < 0.1 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.24 | < 0.1 | < 0.05 | 0.1 |  | < 0.05
| 1 | 5 | 0.03 | 0.1 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.22 | 0.1 | < 0.05 | 0.1 |  | 181
| 1 | 6 | 0.02 | < 0.1 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.22 | < 0.1 | < 0.05 | < 0.1
| Mean |  | 0.026 | 0.1 |  |  | 0.4 | [5 empty] | 0.23 | 0.1 |  | 0.14
| Median |  | 0.02 | 0.1 |  |  | 0.4 | [5 empty] | 0.23 | 0.1 |  | 0.1
| Std Dev. |  | 0.0089443 | 1.6997e-17 | [8 empty] | 0.0089443 | 0 |  | 0.054772
| Rel.Std.Dev. |  | 34.4% | 1.7e-14% | [8 empty] | 3.889% | 0% |  | 39.12%
PDM3
| Z-Score (Absolute) |  | 0.72 | 0.67 | 66.08 | 0 | 1.57 | 0 | 0.67 | 0.67 | 0.67 | 0.67 | 4.86 | 0.45 | 0.67 | 2.02
| Analytical results for Y in OREAS 25b (Certified Value 5.63 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  | 3
| 1 | 1 | 5.73 | 5.3 | 5.7 | 6.0951 | 4.88 | 5.62 | 5.69 | 5.93 | 5.52 | 5.44 | 5.92 | 5.59 | 5.44 |  |  | 1
| 1 | 2 | 5.62 | 5.28 | 5.6 | 6.0676 | 4.74 | 5.39 | 5.66 | 5.94 | 5.66 | 5.34 | 5.71 | 5.58 | 5.54 |  |  | 7
| 1 | 3 | 5.62 | 5.24 | 5.6 | 6.0643 | 4.76 | 5.51 | 5.65 | 6 | 5.59 | 5.21 | 5.72 | 5.25 | 5.28 |  |  | 16
| 1 | 4 | 5.61 | 5.21 | 5.6 | 5.9171 | 4.57 | 5.5 | 5.78 | 5.98 | 5.87 | 5.5 | 5.67 | 5.52 | 5.8 |  |  | 5.6277
| 1 | 5 | 5.61 | 5.06 | 5.5 | 6.0543 | 4.76 | 5.42 | 5.69 | 5.91 | 5.49 | 5.61 | 6 | 5.31 | 5.68 |  |  | 182
| 1 | 6 | 5.55 | 4.98 | 5.6 | 5.9329 | 4.72 | 5.53 | 5.72 | 5.74 | 5.75 | 5.51 | 5.72 | 5.71 | 5.77
| Mean |  | 5.6233 | 5.1783 | 5.6 | 6.0219 | 4.7383 | 5.495 | 5.6983 | 5.9167 | 5.6467 | 5.435 | 5.79 | 5.4933 | 5.585
| Median |  | 5.615 | 5.225 | 5.6 | 6.0593 | 4.75 | 5.505 | 5.69 | 5.935 | 5.625 | 5.47 | 5.72 | 5.55 | 5.61
| Std Dev. |  | 0.058538 | 0.12906 | 0.063246 | 0.076431 | 0.099683 | 0.082158 | 0.047081 | 0.092664 | 0.14459 | 0.14153 | 0.13535 | 0.17739 | 0.20276
| Rel.Std.Dev. |  | 1.041% | 2.492% | 1.129% | 1.269% | 2.104% | 1.495% | 0.8262% | 1.566% | 2.561% | 2.604% | 2.338% | 3.229% | 3.63%
| PDM3 |  | -0.077% | -7.984% | -0.4916% | 7.005% | -15.8% | -2.357% | 1.256% | 5.135% | 0.3376% | -3.424% | 2.885% | -2.387% | -0.7582%
| Z-Score (Absolute) |  | 0.15 | 2.67 | 0 | 2.67 | 5.45 | 0.66 | 0.62 | 2 | 0.3 | 1.04 | 1.2 | 0.67 | 0.09
| Analytical results for Yb in OREAS 25b (Certified Value 0.44 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 05 | 15 | 16 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 0.25g | 0.5g | 0.25g | 0.4g | [10 empty] | 2
| 1 | 1 | 0.469 | 0.4 | 0.4 | 0.5 | 0.4 | [10 empty] | 1
| 1 | 2 | 0.474 | 0.4 | 0.4 | 0.5 | 0.5 | [10 empty] | #N/A
| 1 | 3 | 0.472 | 0.4 | 0.4 | 0.4 | 0.4 | [10 empty] | 16
| 1 | 4 | 0.467 | 0.4 | 0.4 | 0.4 | 0.5 | [10 empty] | 0.43733
| 1 | 5 | 0.467 | 0.4 | 0.4 | 0.4 | 0.5 | [10 empty] | 183
| 1 | 6 | 0.471 | 0.4 | 0.4 | 0.5 | 0.5
| Mean |  | 0.47 | 0.4 | 0.4 | 0.45 | 0.46667
| Median |  | 0.47 | 0.4 | 0.4 | 0.45 | 0.5
| Std Dev. |  | 0.0028284 | 6.0809e-17 | 6.0809e-17 | 0.054772 | 0.05164
| Rel.Std.Dev. |  | 0.6018% | 1.52e-14% | 1.52e-14% | 12.17% | 11.07%
| PDM3 |  | 7.47% | -8.537% | -8.537% | 2.896% | 6.707%
| Z-Score (Absolute) |  | 0.67 | 1.69 | 1.69 | 0 | 0.56
| Analytical results for Zn in OREAS 25b (Certified Value 48.4 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.4g | 2
| 1 | 1 | 49.4 | 42 | 64.4 | 56.415 | 50.5 | 48 | 50 | 50 | 51 | 47 | 42.8 | 53 | 54.3 | 46 | 62 | 1
| 1 | 2 | 48.9 | 40 | 65.1 | 53.1082 | 52.3 | 49 | 46 | 49 | 50 | 46 | 39.3 | 53 | 50.3 | 48 | 60 | 8
| 1 | 3 | 48.3 | 42 | 63.7 | 54.8992 | 50 | 47 | 46 | 47 | 48 | 51 | 38.5 | 54 | 51.5 | 47 | 60 | 16
| 1 | 4 | 48.4 | 42 | 63.7 | 50.6385 | 47.1 | 48 | 48 | 50 | 49 | 49 | 40.4 | 54 | 54.1 | 46 | 63 | 48.4367
| 1 | 5 | 50.7 | 43 | 64 | 51.0526 | 54.7 | 48 | 48 | 50 | 48 | 47 | 39.5 | 55 | 54.9 | 45 | 62 | 184
| 1 | 6 | 48.8 | 42 | 64.6 | 53.9502 | 52.3 | 46 | 47 | 49 | 48 | 47 | 40.9 | 54 | 53.9 | 47 | 63
| Mean |  | 49.0833 | 41.8333 | 64.25 | 53.3439 | 51.15 | 47.6667 | 47.5 | 49.1667 | 49 | 47.8333 | 40.2333 | 53.8333 | 53.1667 | 46.5 | 61.6667
| Median |  | 48.85 | 42 | 64.2 | 53.5292 | 51.4 | 48 | 47.5 | 49.5 | 48.5 | 47 | 39.95 | 54 | 54 | 46.5 | 62
| Std Dev. |  | 0.88412 | 0.98319 | 0.55408 | 2.229 | 2.5844 | 1.0328 | 1.5166 | 1.169 | 1.2649 | 1.8348 | 1.5148 | 0.75277 | 1.8272 | 1.0488 | 1.3663
| Rel.Std.Dev. |  | 1.801% | 2.35% | 0.8624% | 4.179% | 5.053% | 2.167% | 3.193% | 2.378% | 2.581% | 3.836% | 3.765% | 1.398% | 3.437% | 2.256% | 2.216%
| PDM3 |  | 1.335% | -13.63% | 32.65% | 10.13% | 5.602% | -1.59% | -1.934% | 1.507% | 1.163% | -1.246% | -16.94% | 11.14% | 9.765% | -3.998% | 27.31%
| Z-Score (Absolute) |  | 0 | 1.89 | 3.96 | 1.11 | 0.54 | 0.37 | 0.41 | 0.02 | 0.02 | 0.33 | 2.31 | 1.24 | 1.07 | 0.67 | 3.28
| Analytical results for Zr in OREAS 25b (Certified Value 29.2 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.4g |  |  | 2
| 1 | 1 | 29.9 | 27.5 | 26.9 | 30.1695 | 11.9 | 29.8 | 29.5 | 29.7 | 35.8 | 31.3 | 33.9 | 29.2 | 27.5 |  |  | 1
| 1 | 2 | 29.93 | 26.3 | 26.7 | 29.6855 | 11.3 | 27.4 | 29.4 | 28.1 | 38 | 30.2 | 31.8 | 29.7 | 25.6 |  |  | #N/A
| 1 | 3 | 29.02 | 26.3 | 26.6 | 30.394 | 11.5 | 28.1 | 29.4 | 30 | 36.9 | 29.6 | 35 | 28.7 | 25.5 |  |  | 16
| 1 | 4 | 29.11 | 26.5 | 26.9 | 29.5471 | 11.6 | 27.8 | 29.6 | 30.2 | 37.8 | 31.5 | 32.2 | 28.4 | 28.3 |  |  | 29.2155
| 1 | 5 | 29.5 | 26.3 | 27.4 | 30.6563 | 11.1 | 29.2 | 29.5 | 28.9 | 36.7 | 32.2 | 33.8 | 28.4 | 29.7 |  |  | 185
| 1 | 6 | 29.22 | 26.7 | 27.7 | 29.2713 | 11.1 | 28.1 | 29.6 | 30.5 | 38.7 | 33.3 | 32.1 | 28.9 | 29.5
| Mean |  | 29.4467 | 26.6 | 27.0333 | 29.954 | 11.4167 | 28.4 | 29.5 | 29.5667 | 37.3167 | 31.35 | 33.1333 | 28.8833 | 27.6833
| Median |  | 29.36 | 26.4 | 26.9 | 29.9275 | 11.4 | 28.1 | 29.5 | 29.85 | 37.35 | 31.4 | 33 | 28.8 | 27.9
| Std Dev. |  | 0.39717 | 0.46904 | 0.4274 | 0.53609 | 0.31252 | 0.90995 | 0.089443 | 0.90259 | 1.0458 | 1.3368 | 1.2832 | 0.50365 | 1.8378
| Rel.Std.Dev. |  | 1.349% | 1.763% | 1.581% | 1.79% | 2.737% | 3.204% | 0.3032% | 3.053% | 2.802% | 4.264% | 3.873% | 1.744% | 6.639%
| PDM3 |  | 0.7912% | -8.952% | -7.469% | 2.528% | -60.92% | -2.791% | 0.9738% | 1.202% | 27.73% | 7.306% | 13.41% | -1.137% | -5.244%
| Z-Score (Absolute) |  | 0 | 1.09 | 0.92 | 0.19 | 6.89 | 0.4 | 0.02 | 0.05 | 3.01 | 0.73 | 1.41 | 0.22 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 25b (Certified Value 1.14 wt.%) | [16 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 1.14 | 1.1 | 1.14 | 1.0594 | 1.14 | 1.21 | 1.18 | 1.14 | 1.18 | 1.15 | 1.12 | 1.27 | 1.16 | 1.178 | 1.12 | 1
| 1 | 2 | 1.14 | 1.07 | 1.14 | 1.063 | 1.13 | 1.21 | 1.17 | 1.14 | 1.17 | 1.15 | 1.18 | 1.27 | 1.14 | 1.193 | 1.13 | #N/A
| 1 | 3 | 1.13 | 1.11 | 1.15 | 1.055 | 1.09 | 1.21 | 1.15 | 1.13 | 1.16 | 1.15 | 1.14 | 1.27 | 1.17 | 1.177 | 1.14 | 16
| 1 | 4 | 1.15 | 1.08 | 1.13 | 1.0581 | 1.14 | 1.2 | 1.18 | 1.14 | 1.17 | 1.16 | 1.14 | 1.26 | 1.16 | 1.188 | 1.13 | 1.1443
| 1 | 5 | 1.15 | 1.09 | 1.14 | 1.0617 | 1.14 | 1.21 | 1.17 | 1.14 | 1.16 | 1.16 | 1.11 | 1.26 | 1.15 | 1.202 | 1.13 | 187
| 1 | 6 | 1.14 | 1.08 | 1.15 | 1.0574 | 1.11 | 1.19 | 1.17 | 1.14 | 1.17 | 1.14 | 1.13 | 1.28 | 1.13 | 1.195 | 1.13
| Mean |  | 1.1417 | 1.0883 | 1.1417 | 1.0591 | 1.125 | 1.205 | 1.17 | 1.1383 | 1.1683 | 1.1517 | 1.1367 | 1.2683 | 1.1517 | 1.1888 | 1.13
| Median |  | 1.14 | 1.085 | 1.14 | 1.0588 | 1.135 | 1.21 | 1.17 | 1.14 | 1.17 | 1.15 | 1.135 | 1.27 | 1.155 | 1.1905 | 1.13
| Std Dev. |  | 0.0075277 | 0.01472 | 0.0075277 | 0.0029341 | 0.020736 | 0.0083666 | 0.010954 | 0.0040825 | 0.0075277 | 0.0075277 | 0.024221 | 0.0075277 | 0.01472 | 0.0098675 | 0.0063246
| Rel.Std.Dev. |  | 0.6594% | 1.352% | 0.6594% | 0.277% | 1.843% | 0.6943% | 0.9363% | 0.3586% | 0.6443% | 0.6536% | 2.131% | 0.5935% | 1.278% | 0.83% | 0.5597%
| PDM3 |  | -0.2342% | -4.895% | -0.2342% | -7.45% | -1.691% | 5.3% | 2.242% | -0.5255% | 2.096% | 0.6396% | -0.6712% | 10.83% | 0.6396% | 3.887% | -1.254%
| Z-Score (Absolute) |  | 0 | 2.16 | 0 | 3.34 | 0.67 | 2.56 | 1.15 | 0.13 | 1.08 | 0.4 | 0.2 | 5.12 | 0.4 | 1.91 | 0.47
| Analytical results for S in OREAS 25b (Certified Value 0.022 wt.%) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.02 | 0.03 | 0.08 | 0.029161 | 0.02 | 0.04 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 | 0.01 | 0.04 | 0.036 | 0.011 | 1
| 1 | 2 | 0.02 | 0.03 | 0.09 | 0.027745 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 | 0.02 | 0.04 | 0.005 | 17
| 1 | 3 | 0.02 | 0.03 | 0.09 | 0.021465 | 0.01 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 | 0.04 | 0.005 | 16
| 1 | 4 | 0.02 | 0.03 | 0.09 | 0.025233 | 0.01 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.058 | < 0.005 | 0.021552
| 1 | 5 | 0.01 | 0.03 | 0.09 | 0.019723 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.04 | 0.01 | 0.02 | 0.048 | < 0.005 | 188
| 1 | 6 | 0.02 | 0.03 | 0.09 | 0.025701 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.051 | < 0.005
| Mean |  | 0.018333 | 0.03 | 0.088333 | 0.024838 | 0.016667 | 0.031667 | 0.03 | 0.02 | 0.02 | 0.021667 | 0.021667 | 0.015 | 0.023333 | 0.0455 | 0.007
| Median |  | 0.02 | 0.03 | 0.09 | 0.025467 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015 | 0.02 | 0.044 | 0.005
| Std Dev. |  | 0.0040825 | 0 | 0.0040825 | 0.0036216 | 0.005164 | 0.0040825 | 0 | 0 | 0 | 0.0040825 | 0.0098319 | 0.0054772 | 0.008165 | 0.0082885 | 0.0034641
| Rel.Std.Dev. |  | 22.27% | 0% | 4.622% | 14.58% | 30.98% | 12.89% | 0% | 0% | 0% | 18.84% | 45.38% | 36.51% | 34.99% | 18.22% | 49.49%
| PDM3 |  | -14.93% | 39.2% | 309.9% | 15.25% | -22.67% | 46.93% | 39.2% | -7.2% | -7.2% | 0.5337% | 0.5337% | -30.4% | 8.267% | 111.1% | -67.52%
| Z-Score (Absolute) |  | 0.45 | 1.12 | 8.99 | 0.43 | 0.67 | 1.35 | 1.12 | 0.22 | 0.22 | 0 | 0 | 0.9 | 0.22 | 3.21 | 2.28
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.