## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 625
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.66675 | 0.65941 | 0.67408 | 0.6651 | 0.6684
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.66659 | 0.6558 | 0.67738 | 0.66478 | 0.6684
4-Acid Digestion
Ag, Silver (ppm) | 11.6934 | 11.2121 | 12.1748 | 11.3179 | 12.069
Al, Aluminium (wt.%) | 6.6781 | 6.4553 | 6.901 | 6.5562 | 6.8001
As, Arsenic (ppm) | 143.3644 | 138.4476 | 148.2813 | 138.9234 | 147.8055
Be, Beryllium (ppm) | 2.5671 | 2.4463 | 2.6878 | 2.4705 | 2.6637
Bi, Bismuth (ppm) | 8.6125 | 8.2582 | 8.9669 | 8.3878 | 8.8373
Ca, Calcium (wt.%) | 1.1444 | 1.1126 | 1.1762 | 1.1161 | 1.1727
Cd, Cadmium (ppm) | 82.9767 | 80.0351 | 85.9183 | 80.9974 | 84.956
Ce, Cerium (ppm) | 72.4712 | 66.832 | 78.1103 | 69.7367 | 75.2056
Co, Cobalt (ppm) | 4.2137 | 3.9922 | 4.4352 | 4.0256 | 4.4019
Cr, Chromium (ppm) | 13.276 | 11.6488 | 14.9032 | 12.3335 | 14.2185
Cs, Caesium (ppm) | 5.9425 | 5.676 | 6.209 | 5.7649 | 6.1201
Cu, Copper (wt.%) | 0.171 | 0.16581 | 0.17618 | 0.16854 | 0.17346
Dy, Dysprosium (ppm) | 3.5182 | 3.217 | 3.8195 | 3.3621 | 3.6744
Er, Erbium (ppm) | 1.2423 | 1.1388 | 1.3458 | 1.1712 | 1.3135
Eu, Europium (ppm) | 1.3626 | 1.19 | 1.5352 | 1.3117 | 1.4134
Fe, Iron (wt.%) | 3.9776 | 3.8734 | 4.0818 | 3.9054 | 4.0499
Ga, Gallium (ppm) | 20.5577 | 19.8102 | 21.3053 | 20.0685 | 21.047
Gd, Gadolinium (ppm) | 5.463 | 5.1341 | 5.792 | 5.268 | 5.658
Hf, Hafnium (ppm) | 6.3553 | 6.0332 | 6.6775 | 6.185 | 6.5257
Ho, Holmium (ppm) | 0.53594 | 0.4897 | 0.58219 | 0.50135 | 0.57054
In, Indium (ppm) | 2.1179 | 2.0135 | 2.2224 | 2.0528 | 2.183
K, Potassium (wt.%) | 2.8649 | 2.7975 | 2.9322 | 2.8054 | 2.9243
La, Lanthanum (ppm) | 34.7157 | 30.9589 | 38.4725 | 33.427 | 36.0044
Li, Lithium (ppm) | 17.8834 | 17.1662 | 18.6005 | 17.3783 | 18.3885
Lu, Lutetium (ppm) | 0.13519 | 0.12412 | 0.14626 | IND | IND
Mg, Magnesium (wt.%) | 0.26131 | 0.25106 | 0.27157 | 0.25685 | 0.26577
Mn, Manganese (wt.%) | 0.44334 | 0.43337 | 0.45331 | 0.436 | 0.45068
Mo, Molybdenum (ppm) | 12.295 | 11.6288 | 12.9612 | 11.9844 | 12.6056
Na, Sodium (wt.%) | 1.9132 | 1.8615 | 1.9649 | 1.8791 | 1.9473
Nb, Niobium (ppm) | 14.7794 | 13.9081 | 15.6506 | 14.3643 | 15.1945
Nd, Neodymium (ppm) | 37.3294 | 32.9775 | 41.6812 | 36.2975 | 38.3612
Ni, Nickel (ppm) | 5.9846 | 5.5499 | 6.4194 | 5.5982 | 6.3711
P, Phosphorus (wt.%) | 0.029731 | 0.028794 | 0.030668 | 0.028747 | 0.030716
Pb, Lead (wt.%) | 0.82207 | 0.80395 | 0.8402 | 0.8112 | 0.83294
Pr, Praseodymium (ppm) | 9.738 | 8.5671 | 10.9089 | 9.4133 | 10.0627
Rb, Rubidium (ppm) | 135.319 | 130.2657 | 140.3724 | 131.7023 | 138.9357
Re, Rhenium (ppm) | 0.0039861 | 0.0020213 | 0.005951 | IND | IND
S, Sulphur (wt.%) | 3.8039 | 3.6897 | 3.9181 | 3.7346 | 3.8733
Sb, Antimony (ppm) | 35.1259 | 33.3798 | 36.872 | 34.1094 | 36.1423
Sc, Scandium (ppm) | 4.6011 | 4.2852 | 4.917 | 4.4463 | 4.7559
Sm, Samarium (ppm) | 6.801 | 6.2215 | 7.3805 | 6.4039 | 7.1981
Sn, Tin (ppm) | 3.5668 | 3.3744 | 3.7593 | 3.4263 | 3.7074
Sr, Strontium (ppm) | 130.2398 | 125.4472 | 135.0324 | 125.8789 | 134.6006
Ta, Tantalum (ppm) | 1.1074 | 1.0191 | 1.1956 | 1.0532 | 1.1615
Tb, Terbium (ppm) | 0.72399 | 0.65218 | 0.7958 | 0.68979 | 0.75819
Te, Tellurium (ppm) | 0.063198 | 0.016542 | 0.10985 | IND | IND
Th, Thorium (ppm) | 12.6735 | 11.6587 | 13.6883 | 12.2763 | 13.0707
Ti, Titanium (wt.%) | 0.12026 | 0.11542 | 0.12509 | 0.11713 | 0.12339
Tl, Thallium (ppm) | 9.7907 | 9.3978 | 10.1836 | 9.5478 | 10.0336
Tm, Thulium (ppm) | 0.16596 | 0.14105 | 0.19086 | IND | IND
U, Uranium (ppm) | 5.5198 | 5.269 | 5.7706 | 5.3672 | 5.6725
V, Vanadium (ppm) | 11.4585 | 10.6144 | 12.3025 | 10.9622 | 11.9547
W, Tungsten (ppm) | 3.9863 | 3.6708 | 4.3018 | 3.8353 | 4.1373
Y, Yttrium (ppm) | 14.8882 | 13.9272 | 15.8493 | 14.4407 | 15.3358
Yb, Ytterbium (ppm) | 0.9574 | 0.84772 | 1.0671 | 0.91896 | 0.99584
Zn, Zinc (wt.%) | 3.1686 | 3.1015 | 3.2358 | 3.1236 | 3.2137
Zr, Zirconium (ppm) | 229.2116 | 221.0077 | 237.4155 | 224.0437 | 234.3795
Aqua Regia Digestion
Ag, Silver (ppm) | 11.6562 | 11.2876 | 12.0248 | 11.3832 | 11.9292
Al, Aluminium (wt.%) | 0.56512 | 0.53123 | 0.59901 | 0.54959 | 0.58065
As, Arsenic (ppm) | 133.6096 | 127.3002 | 139.919 | 130.4199 | 136.7992
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.61007 | 0.57906 | 0.64108 | 0.58515 | 0.63499
Bi, Bismuth (ppm) | 8.8116 | 8.4761 | 9.1472 | 8.5329 | 9.0903
Ca, Calcium (wt.%) | 0.9658 | 0.93591 | 0.99569 | 0.94779 | 0.98381
Cd, Cadmium (ppm) | 83.2787 | 80.5538 | 86.0036 | 81.9959 | 84.5615
Ce, Cerium (ppm) | 36.7836 | 34.1761 | 39.3911 | 35.9122 | 37.655
Co, Cobalt (ppm) | 3.5367 | 3.3496 | 3.7239 | 3.3907 | 3.6828
Cr, Chromium (ppm) | 13.0762 | 12.2049 | 13.9475 | 12.484 | 13.6683
Cs, Caesium (ppm) | 1.1424 | 1.0669 | 1.2178 | 1.1093 | 1.1754
Cu, Copper (wt.%) | 0.17359 | 0.16843 | 0.17874 | 0.17085 | 0.17632
Fe, Iron (wt.%) | 3.2729 | 3.1465 | 3.3994 | 3.2105 | 3.3353
Ga, Gallium (ppm) | 2.6263 | 2.3385 | 2.9142 | 2.551 | 2.7017
Hf, Hafnium (ppm) | 1.461 | 1.3783 | 1.5436 | 1.4069 | 1.515
Hg, Mercury (ppm) | 1.7164 | 1.5796 | 1.8533 | 1.6616 | 1.7712
In, Indium (ppm) | 2.0651 | 1.952 | 2.1782 | 1.9723 | 2.1579
K, Potassium (wt.%) | 0.27558 | 0.2607 | 0.29046 | 0.26717 | 0.28399
La, Lanthanum (ppm) | 17.6507 | 16.1717 | 19.1297 | 17.2219 | 18.0795
Li, Lithium (ppm) | 3.4013 | 3.1012 | 3.7013 | 3.2128 | 3.5898
Lu, Lutetium (ppm) | 0.045083 | 0.036047 | 0.05412 | IND | IND
Mg, Magnesium (wt.%) | 0.11907 | 0.11268 | 0.12547 | 0.11477 | 0.12338
Mn, Manganese (wt.%) | 0.43395 | 0.42053 | 0.44737 | 0.42658 | 0.44133
Mo, Molybdenum (ppm) | 11.6409 | 11.0827 | 12.1991 | 11.3256 | 11.9563
Na, Sodium (wt.%) | 0.058867 | 0.05593 | 0.061805 | 0.056977 | 0.060758
Nb, Niobium (ppm) | 0.69694 | 0.60304 | 0.79083 | 0.65651 | 0.73736
Ni, Nickel (ppm) | 5.3202 | 4.9864 | 5.654 | 5.1014 | 5.539
P, Phosphorus (wt.%) | 0.025122 | 0.023911 | 0.026333 | 0.024336 | 0.025908
Pb, Lead (wt.%) | 0.82622 | 0.80731 | 0.84513 | 0.81601 | 0.83642
Rb, Rubidium (ppm) | 14.2798 | 13.3194 | 15.2402 | 13.728 | 14.8316
Re, Rhenium (ppm) | 0.0040833 | 0.0026885 | 0.0054782 | IND | IND
S, Sulphur (wt.%) | 3.7662 | 3.6395 | 3.8929 | 3.6869 | 3.8456
Sb, Antimony (ppm) | 28.5046 | 26.981 | 30.0281 | 27.5201 | 29.489
Sc, Scandium (ppm) | 1.1189 | 0.9984 | 1.2395 | IND | IND
Sn, Tin (ppm) | 0.94654 | 0.82572 | 1.0674 | 0.87337 | 1.0197
Sr, Strontium (ppm) | 20.7281 | 19.8368 | 21.6195 | 20.2416 | 21.2147
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.38144 | 0.31875 | 0.44413 | 0.36382 | 0.39906
Te, Tellurium (ppm) | 0.050324 | 0.027683 | 0.072965 | IND | IND
Th, Thorium (ppm) | 7.4911 | 7.0384 | 7.9438 | 7.2773 | 7.7049
Ti, Titanium (wt.%) | 0.02342 | 0.021145 | 0.025696 | 0.022704 | 0.024137
Tl, Thallium (ppm) | 3.0305 | 2.8865 | 3.1745 | 2.9248 | 3.1363
U, Uranium (ppm) | 2.77 | 2.5965 | 2.9436 | 2.7047 | 2.8354
V, Vanadium (ppm) | 3.1105 | 2.8753 | 3.3457 | 2.8206 | 3.4004
W, Tungsten (ppm) | 1.4702 | 1.3534 | 1.5869 | 1.4173 | 1.523
Y, Yttrium (ppm) | 6.3963 | 6.0393 | 6.7533 | 6.2162 | 6.5765
Yb, Ytterbium (ppm) | 0.3408 | 0.30651 | 0.37508 | IND | IND
Zn, Zinc (wt.%) | 3.1792 | 3.1134 | 3.245 | 3.1322 | 3.2263
Zr, Zirconium (ppm) | 54.7422 | 49.6983 | 59.7862 | 52.6277 | 56.8568
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.8273 | 6.6508 | 7.0038 | 6.7342 | 6.9204
As, Arsenic (ppm) | 144.1385 | 133.5928 | 154.6842 | 139.403 | 148.8739
Ba, Barium (ppm) | 3415.3709 | 3249.3142 | 3581.4276 | 3350.5006 | 3480.2412
Bi, Bismuth (ppm) | 8.6117 | 7.9285 | 9.2948 | 8.3225 | 8.9008
Ca, Calcium (wt.%) | 1.1477 | 1.0876 | 1.2077 | 1.1095 | 1.1858
Cd, Cadmium (ppm) | 82.6872 | 77.2221 | 88.1523 | 80.2531 | 85.1213
Ce, Cerium (ppm) | 86.0175 | 79.9662 | 92.0688 | 82.5329 | 89.5022
Co, Cobalt (ppm) | 4.5042 | 3.796 | 5.2123 | 4.068 | 4.9404
Cs, Caesium (ppm) | 5.9997 | 5.3966 | 6.6029 | 5.7754 | 6.2241
Cu, Copper (wt.%) | 0.17029 | 0.16575 | 0.17484 | 0.16658 | 0.17401
Dy, Dysprosium (ppm) | 3.6586 | 3.224 | 4.0932 | 3.5001 | 3.8171
Er, Erbium (ppm) | 1.4143 | 1.2487 | 1.5799 | 1.3502 | 1.4784
Eu, Europium (ppm) | 1.5774 | 1.2257 | 1.9291 | 1.4872 | 1.6676
Fe, Iron (wt.%) | 4.0112 | 3.9144 | 4.108 | 3.9473 | 4.0751
Ga, Gallium (ppm) | 20.4465 | 18.8279 | 22.065 | 19.6833 | 21.2097
Gd, Gadolinium (ppm) | 5.5297 | 4.9882 | 6.0712 | 5.2839 | 5.7755
Ge, Germanium (ppm) | 1.3052 | 0.42714 | 2.1832 | IND | IND
Ho, Holmium (ppm) | 0.60395 | 0.57856 | 0.62934 | 0.58184 | 0.62606
In, Indium (ppm) | 2.1052 | 1.8472 | 2.3631 | 1.8508 | 2.3596
K, Potassium (wt.%) | 2.8869 | 2.7975 | 2.9763 | 2.8269 | 2.947
La, Lanthanum (ppm) | 43.3584 | 40.8598 | 45.8571 | 41.6802 | 45.0367
Li, Lithium (ppm) | 18.1905 | 15.9366 | 20.4443 | IND | IND
Lu, Lutetium (ppm) | 0.1433 | 0.11345 | 0.17315 | IND | IND
Mg, Magnesium (wt.%) | 0.26668 | 0.25712 | 0.27623 | 0.25587 | 0.27748
Mn, Manganese (wt.%) | 0.44679 | 0.43786 | 0.45572 | 0.43963 | 0.45396
Mo, Molybdenum (ppm) | 13.6828 | 12.0667 | 15.2989 | IND | IND
Nb, Niobium (ppm) | 14.7113 | 12.3941 | 17.0284 | IND | IND
Nd, Neodymium (ppm) | 37.7173 | 35.9174 | 39.5172 | 36.9452 | 38.4894
P, Phosphorus (wt.%) | 0.030419 | 0.028373 | 0.032464 | IND | IND
Pb, Lead (wt.%) | 0.80589 | 0.78322 | 0.82855 | 0.79237 | 0.8194
Pr, Praseodymium (ppm) | 9.9529 | 8.9775 | 10.9284 | 9.5874 | 10.3185
Rb, Rubidium (ppm) | 136.3857 | 127.1761 | 145.5954 | 130.8241 | 141.9474
S, Sulphur (wt.%) | 3.851 | 3.7268 | 3.9752 | 3.76 | 3.942
Sb, Antimony (ppm) | 36.1703 | 33.0799 | 39.2607 | 35.1944 | 37.1462
Si, Silicon (wt.%) | 29.0185 | 28.2464 | 29.7906 | 28.5762 | 29.4609
Sm, Samarium (ppm) | 7.0366 | 6.2101 | 7.863 | 6.6617 | 7.4115
Sr, Strontium (ppm) | 142.3775 | 135.2878 | 149.4673 | 139.0672 | 145.6879
Tb, Terbium (ppm) | 0.77226 | 0.67589 | 0.86864 | 0.75422 | 0.7903
Th, Thorium (ppm) | 13.7295 | 12.9332 | 14.5258 | 13.3662 | 14.0928
Ti, Titanium (wt.%) | 0.12943 | 0.12248 | 0.13638 | 0.12354 | 0.13532
Tl, Thallium (ppm) | 9.9868 | 9.2094 | 10.7643 | 9.7177 | 10.256
Tm, Thulium (ppm) | 0.18296 | 0.16344 | 0.20248 | IND | IND
U, Uranium (ppm) | 5.7822 | 5.2789 | 6.2855 | 5.4915 | 6.073
V, Vanadium (ppm) | 11.5024 | 8.2576 | 14.7471 | IND | IND
W, Tungsten (ppm) | 4.3931 | 3.0442 | 5.7421 | 4.2514 | 4.5349
Y, Yttrium (ppm) | 17.5649 | 16.4675 | 18.6622 | 17.3759 | 17.7539
Yb, Ytterbium (ppm) | 1.0549 | 0.9327 | 1.1772 | IND | IND
Zn, Zinc (wt.%) | 3.2007 | 3.1149 | 3.2866 | 3.1546 | 3.2469
Infrared Combustion
S, Sulphur (wt.%) | 3.9116 | 3.828 | 3.9952 | 3.8717 | 3.9515
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 625
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Ag | ppm | 6.6667 | Pd | ppb | 0.75 | Pt | ppb | < 10
4-Acid Digestion
B | ppm | 20.8333 | Ge | ppm | 0.22872 | Se | ppm | 2.0667
Ba | ppm | 527.7437 | Hg | ppm | 1.348
Aqua Regia Digestion
Ba | ppm | 76.9419 | Ge | ppm | 0.084259 | Pt | ppb | < 5
Dy | ppm | 1.6888 | Ho | ppm | 0.24104 | Se | ppm | 1.5503
Er | ppm | 0.55165 | Nd | ppm | 17.1554 | Sm | ppm | 3.3006
Eu | ppm | 0.67319 | Pd | ppb | 11.5 | Tm | ppm | 0.059004
Gd | ppm | 2.9003 | Pr | ppm | 4.5233
Peroxide Fusion ICP
Ag | ppm | 12.5554 | Hg | ppm | < 5 | Se | ppm | < 5
B | ppm | 25.3333 | Na | wt.% | 1.8833 | Sn | ppm | 3.4754
Be | ppm | 2.9078 | Ni | ppm | 20.7019 | Ta | ppm | 1.4266
Cr | ppm | 56.7123 | Re | ppm | < 0.05 | Te | ppm | < 1
Hf | ppm | 6.292 | Sc | ppm | 4.8388 | Zr | ppm | 248.0719
Infrared Combustion
C | wt.% | 0.38389
Borate Fusion XRF
Al2O3 | wt.% | 13.1626 | MgO | wt.% | 0.46655 | Sn | ppm | 27.5059
As | ppm | 3156.1728 | Mn | wt.% | 0.45362 | Sr | ppm | 99.1256
Ba | ppm | 3587.7583 | Na2O | wt.% | 2.6942 | TiO2 | wt.% | 0.21066
CaO | wt.% | 1.6087 | Ni | ppm | 10.5991 | V | ppm | 17.9829
Cr | ppm | 27.3813 | P | wt.% | 0.02906 | Zn | wt.% | 3.3177
Cu | wt.% | 0.16943 | Pb | wt.% | 0.8709 | Zr | ppm | 295.8355
Fe | wt.% | 4.0512 | S | wt.% | 3.3264
K2O | wt.% | 3.5187 | SiO2 | wt.% | 62.8471
Thermogravimetry
LOI1000 | wt.% | 4.3667
Laser Ablation ICP-MS
Ag | ppm | 10.85 | Hf | ppm | 6.795 | Se | ppm | < 5
As | ppm | 141.5 | Ho | ppm | 0.595 | Sm | ppm | 6.89
Ba | ppm | 3450 | In | ppm | 1.95 | Sn | ppm | 3.4
Be | ppm | 2.6 | La | ppm | 42.7 | Sr | ppm | 134
Bi | ppm | 8.9 | Lu | ppm | 0.13 | Ta | ppm | 1.175
Ce | ppm | 81.6 | Mn | wt.% | 0.454 | Tb | ppm | 0.74
Co | ppm | 5.05 | Mo | ppm | 11.6 | Te | ppm | 0.15
Cr | ppm | 16 | Nb | ppm | 15.6 | Th | ppm | 14.25
Cs | ppm | 6.07 | Nd | ppm | 36.2 | Ti | wt.% | 0.1335
Cu | wt.% | 0.2065 | Ni | ppm | 8 | Tm | ppm | 0.16
Dy | ppm | 3.825 | Pb | wt.% | 0.7365 | U | ppm | 5.595
Er | ppm | 1.365 | Pr | ppm | 10.2 | V | ppm | 11.1
Eu | ppm | 1.425 | Rb | ppm | 136 | W | ppm | 4.5
Ga | ppm | 19.75 | Re | ppm | < 0.01 | Y | ppm | 16.55
Gd | ppm | 5.36 | Sb | ppm | 36.35 | Yb | ppm | 1.065
Ge | ppm | 1.075 | Sc | ppm | 5.1 | Zr | ppm | 240.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 625
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.66675 | 0.019853 | 0.62704 | 0.70645 | 0.60719 | 0.72631 | 2.978% | 5.955% | 8.933% | 0.63341 | 0.70008
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.66659 | 0.023663 | 0.61926 | 0.71392 | 0.5956 | 0.73758 | 3.55% | 7.1% | 10.65% | 0.63326 | 0.69992
4-Acid Digestion
Ag, ppm | 11.6934 | 0.60652 | 10.4804 | 12.9065 | 9.8739 | 13.513 | 5.187% | 10.37% | 15.56% | 11.1088 | 12.2781
Al, wt.% | 6.6781 | 0.27313 | 6.1319 | 7.2244 | 5.8587 | 7.4975 | 4.09% | 8.18% | 12.27% | 6.3442 | 7.012
As, ppm | 143.3644 | 4.0765 | 135.2114 | 151.5175 | 131.1349 | 155.594 | 2.843% | 5.687% | 8.53% | 136.1962 | 150.5327
Be, ppm | 2.5671 | 0.15691 | 2.2533 | 2.8809 | 2.0964 | 3.0378 | 6.112% | 12.22% | 18.34% | 2.4387 | 2.6954
Bi, ppm | 8.6125 | 0.48267 | 7.6472 | 9.5779 | 7.1645 | 10.0605 | 5.604% | 11.21% | 16.81% | 8.1819 | 9.0432
Ca, wt.% | 1.1444 | 0.052834 | 1.0387 | 1.2501 | 0.9859 | 1.3029 | 4.617% | 9.233% | 13.85% | 1.0872 | 1.2016
Cd, ppm | 82.9767 | 5.2832 | 72.4103 | 93.5431 | 67.1271 | 98.8263 | 6.367% | 12.73% | 19.1% | 78.8279 | 87.1255
Ce, ppm | 72.4712 | 9.8843 | 52.7026 | 92.2397 | 42.8184 | 102.1239 | 13.64% | 27.28% | 40.92% | 68.8476 | 76.0947
Co, ppm | 4.2137 | 0.23594 | 3.7418 | 4.6856 | 3.5059 | 4.9215 | 5.599% | 11.2% | 16.8% | 4.003 | 4.4244
Cr, ppm | 13.276 | 1.6182 | 10.0396 | 16.5124 | 8.4214 | 18.1305 | 12.19% | 24.38% | 36.57% | 12.6122 | 13.9398
Cs, ppm | 5.9425 | 0.20948 | 5.5235 | 6.3615 | 5.3141 | 6.5709 | 3.525% | 7.05% | 10.58% | 5.6454 | 6.2396
Cu, wt.% | 0.171 | 0.0063526 | 0.15829 | 0.1837 | 0.15194 | 0.19005 | 3.715% | 7.43% | 11.15% | 0.16245 | 0.17955
Dy, ppm | 3.5182 | 0.17465 | 3.1689 | 3.8675 | 2.9943 | 4.0422 | 4.964% | 9.928% | 14.89% | 3.3423 | 3.6942
Er, ppm | 1.2423 | 0.048676 | 1.145 | 1.3397 | 1.0963 | 1.3883 | 3.918% | 7.836% | 11.75% | 1.1802 | 1.3044
Eu, ppm | 1.3626 | 0.1089 | 1.1448 | 1.5804 | 1.0359 | 1.6893 | 7.993% | 15.99% | 23.98% | 1.2944 | 1.4307
Fe, wt.% | 3.9776 | 0.13117 | 3.7153 | 4.24 | 3.5841 | 4.3711 | 3.298% | 6.595% | 9.893% | 3.7787 | 4.1765
Ga, ppm | 20.5577 | 0.63475 | 19.2882 | 21.8272 | 18.6535 | 22.462 | 3.088% | 6.175% | 9.263% | 19.5298 | 21.5856
Gd, ppm | 5.463 | 0.20624 | 5.0506 | 5.8755 | 4.8443 | 6.0818 | 3.775% | 7.55% | 11.33% | 5.1899 | 5.7362
Hf, ppm | 6.3553 | 0.37716 | 5.601 | 7.1097 | 5.2239 | 7.4868 | 5.935% | 11.87% | 17.8% | 6.0376 | 6.6731
Ho, ppm | 0.53594 | 0.024348 | 0.48725 | 0.58464 | 0.4629 | 0.60899 | 4.543% | 9.086% | 13.63% | 0.50915 | 0.56274
In, ppm | 2.1179 | 0.097986 | 1.9219 | 2.3139 | 1.824 | 2.4119 | 4.627% | 9.253% | 13.88% | 2.012 | 2.2238
K, wt.% | 2.8649 | 0.086911 | 2.691 | 3.0387 | 2.6041 | 3.1256 | 3.034% | 6.067% | 9.101% | 2.7216 | 3.0081
La, ppm | 34.7157 | 5.9044 | 22.9069 | 46.5245 | 17.0024 | 52.4289 | 17.01% | 34.02% | 51.02% | 32.9799 | 36.4515
Li, ppm | 17.8834 | 0.71956 | 16.4442 | 19.3225 | 15.7247 | 20.042 | 4.024% | 8.047% | 12.07% | 16.9892 | 18.7775
Lu, ppm | 0.13519 | 0.0064478 | 0.1223 | 0.14809 | 0.11585 | 0.15454 | 4.769% | 9.539% | 14.31% | 0.12843 | 0.14195
Mg, wt.% | 0.26131 | 0.012333 | 0.23665 | 0.28598 | 0.22431 | 0.29831 | 4.72% | 9.439% | 14.16% | 0.24825 | 0.27438
Mn, wt.% | 0.44334 | 0.015405 | 0.41253 | 0.47415 | 0.39713 | 0.48955 | 3.475% | 6.949% | 10.42% | 0.42117 | 0.46551
Mo, ppm | 12.295 | 0.8956 | 10.5038 | 14.0862 | 9.6082 | 14.9818 | 7.284% | 14.57% | 21.85% | 11.6803 | 12.9098
Na, wt.% | 1.9132 | 0.058061 | 1.7971 | 2.0293 | 1.739 | 2.0874 | 3.035% | 6.07% | 9.104% | 1.8176 | 2.0089
Nb, ppm | 14.7794 | 1.3387 | 12.102 | 17.4567 | 10.7633 | 18.7954 | 9.058% | 18.12% | 27.17% | 14.0404 | 15.5183
Nd, ppm | 37.3294 | 3.7216 | 29.8861 | 44.7726 | 26.1644 | 48.4943 | 9.97% | 19.94% | 29.91% | 35.4629 | 39.1958
Ni, ppm | 5.9846 | 0.53363 | 4.9174 | 7.0519 | 4.3838 | 7.5855 | 8.917% | 17.83% | 26.75% | 5.6854 | 6.2839
P, wt.% | 0.029731 | 0.0010252 | 0.027681 | 0.031782 | 0.026656 | 0.032807 | 3.448% | 6.896% | 10.34% | 0.028245 | 0.031218
Pb, wt.% | 0.82207 | 0.01886 | 0.78435 | 0.85979 | 0.76549 | 0.87865 | 2.294% | 4.588% | 6.882% | 0.78097 | 0.86317
Pr, ppm | 9.738 | 0.98711 | 7.7638 | 11.7122 | 6.7767 | 12.6993 | 10.14% | 20.27% | 30.41% | 9.2511 | 10.2249
Rb, ppm | 135.319 | 6.1469 | 123.0253 | 147.6127 | 116.8785 | 153.7596 | 4.542% | 9.085% | 13.63% | 128.5531 | 142.085
Re, ppm | 0.0039861 | 0.0010998 | 0.0017865 | 0.0061857 | 0.00068672 | 0.0072855 | 27.59% | 55.18% | 82.77% | 0.0037868 | 0.0041854
S, wt.% | 3.8039 | 0.11963 | 3.5647 | 4.0432 | 3.445 | 4.1628 | 3.145% | 6.29% | 9.435% | 3.6137 | 3.9941
Sb, ppm | 35.1259 | 2.3038 | 30.5183 | 39.7335 | 28.2145 | 42.0373 | 6.559% | 13.12% | 19.68% | 33.3696 | 36.8822
Sc, ppm | 4.6011 | 0.3242 | 3.9527 | 5.2495 | 3.6285 | 5.5737 | 7.046% | 14.09% | 21.14% | 4.3711 | 4.8312
Sm, ppm | 6.801 | 0.35697 | 6.0871 | 7.515 | 5.7301 | 7.8719 | 5.249% | 10.5% | 15.75% | 6.461 | 7.1411
Sn, ppm | 3.5668 | 0.17069 | 3.2254 | 3.9082 | 3.0548 | 4.0789 | 4.785% | 9.571% | 14.36% | 3.3885 | 3.7452
Sr, ppm | 130.2398 | 6.9942 | 116.2513 | 144.2283 | 109.2571 | 151.2225 | 5.37% | 10.74% | 16.11% | 123.7278 | 136.7518
Ta, ppm | 1.1074 | 0.081243 | 0.94487 | 1.2698 | 0.86362 | 1.3511 | 7.337% | 14.67% | 22.01% | 1.052 | 1.1627
Tb, ppm | 0.72399 | 0.063945 | 0.5961 | 0.85188 | 0.53215 | 0.91582 | 8.832% | 17.66% | 26.5% | 0.68779 | 0.76019
Te, ppm | 0.063198 | 0.030203 | 0.0027914 | 0.1236 | 0 | 0.15381 | 47.79% | 95.58% | 143.4% | 0.060038 | 0.066358
Th, ppm | 12.6735 | 1.2716 | 10.1303 | 15.2167 | 8.8587 | 16.4882 | 10.03% | 20.07% | 30.1% | 12.0398 | 13.3072
Ti, wt.% | 0.12026 | 0.003881 | 0.11249 | 0.12802 | 0.10861 | 0.1319 | 3.227% | 6.455% | 9.682% | 0.11424 | 0.12627
Tl, ppm | 9.7907 | 0.4134 | 8.9639 | 10.6175 | 8.5505 | 11.0309 | 4.222% | 8.445% | 12.67% | 9.3012 | 10.2802
Tm, ppm | 0.16596 | 0.022578 | 0.1208 | 0.21111 | 0.098222 | 0.23369 | 13.6% | 27.21% | 40.81% | 0.15766 | 0.17425
U, ppm | 5.5198 | 0.26335 | 4.9931 | 6.0465 | 4.7298 | 6.3099 | 4.771% | 9.542% | 14.31% | 5.2438 | 5.7958
V, ppm | 11.4585 | 1.0561 | 9.3463 | 13.5706 | 8.2903 | 14.6266 | 9.216% | 18.43% | 27.65% | 10.8855 | 12.0314
W, ppm | 3.9863 | 0.40664 | 3.173 | 4.7996 | 2.7664 | 5.2062 | 10.2% | 20.4% | 30.6% | 3.787 | 4.1856
Y, ppm | 14.8882 | 1.0324 | 12.8235 | 16.953 | 11.7911 | 17.9853 | 6.934% | 13.87% | 20.8% | 14.1438 | 15.6326
Yb, ppm | 0.9574 | 0.055039 | 0.84732 | 1.0675 | 0.79228 | 1.1225 | 5.749% | 11.5% | 17.25% | 0.90953 | 1.0053
Zn, wt.% | 3.1686 | 0.097836 | 2.973 | 3.3643 | 2.8751 | 3.4621 | 3.088% | 6.175% | 9.263% | 3.0102 | 3.3271
Zr, ppm | 229.2116 | 9.781 | 209.6496 | 248.7736 | 199.8686 | 258.5547 | 4.267% | 8.534% | 12.8% | 217.751 | 240.6722
Aqua Regia Digestion
Ag, ppm | 11.6562 | 0.39662 | 10.863 | 12.4494 | 10.4663 | 12.8461 | 3.403% | 6.805% | 10.21% | 11.0734 | 12.239
Al, wt.% | 0.56512 | 0.063952 | 0.43722 | 0.69302 | 0.37327 | 0.75698 | 11.32% | 22.63% | 33.95% | 0.53686 | 0.59338
As, ppm | 133.6096 | 9.5461 | 114.5174 | 152.7017 | 104.9714 | 162.2478 | 7.145% | 14.29% | 21.43% | 126.9291 | 140.2901
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.61007 | 0.035102 | 0.53987 | 0.68027 | 0.50477 | 0.71537 | 5.754% | 11.51% | 17.26% | 0.57957 | 0.64057
Bi, ppm | 8.8116 | 0.29499 | 8.2216 | 9.4016 | 7.9267 | 9.6966 | 3.348% | 6.696% | 10.04% | 8.3711 | 9.2522
Ca, wt.% | 0.9658 | 0.035231 | 0.89534 | 1.0363 | 0.8601 | 1.0715 | 3.648% | 7.296% | 10.94% | 0.91751 | 1.0141
Cd, ppm | 83.2787 | 3.7735 | 75.7316 | 90.8258 | 71.9581 | 94.5993 | 4.531% | 9.062% | 13.59% | 79.1148 | 87.4426
Ce, ppm | 36.7836 | 3.8418 | 29.1 | 44.4673 | 25.2581 | 48.3091 | 10.44% | 20.89% | 31.33% | 34.9444 | 38.6228
Co, ppm | 3.5367 | 0.30189 | 2.933 | 4.1405 | 2.6311 | 4.4424 | 8.536% | 17.07% | 25.61% | 3.3599 | 3.7136
Cr, ppm | 13.0762 | 0.77576 | 11.5247 | 14.6277 | 10.7489 | 15.4034 | 5.933% | 11.87% | 17.8% | 12.4224 | 13.73
Cs, ppm | 1.1424 | 0.10781 | 0.92672 | 1.358 | 0.81891 | 1.4658 | 9.438% | 18.88% | 28.31% | 1.0852 | 1.1995
Cu, wt.% | 0.17359 | 0.005702 | 0.16218 | 0.18499 | 0.15648 | 0.19069 | 3.285% | 6.57% | 9.855% | 0.16491 | 0.18226
Fe, wt.% | 3.2729 | 0.22758 | 2.8178 | 3.7281 | 2.5902 | 3.9557 | 6.954% | 13.91% | 20.86% | 3.1093 | 3.4366
Ga, ppm | 2.6263 | 0.43894 | 1.7484 | 3.5042 | 1.3095 | 3.9431 | 16.71% | 33.43% | 50.14% | 2.495 | 2.7576
Hf, ppm | 1.461 | 0.094639 | 1.2717 | 1.6503 | 1.1771 | 1.7449 | 6.478% | 12.96% | 19.43% | 1.3879 | 1.534
Hg, ppm | 1.7164 | 0.13749 | 1.4415 | 1.9914 | 1.304 | 2.1289 | 8.01% | 16.02% | 24.03% | 1.6306 | 1.8022
In, ppm | 2.0651 | 0.1352 | 1.7947 | 2.3355 | 1.6595 | 2.4707 | 6.547% | 13.09% | 19.64% | 1.9618 | 2.1684
K, wt.% | 0.27558 | 0.026013 | 0.22355 | 0.32761 | 0.19754 | 0.35362 | 9.439% | 18.88% | 28.32% | 0.2618 | 0.28936
La, ppm | 17.6507 | 2.5286 | 12.5934 | 22.708 | 10.0648 | 25.2366 | 14.33% | 28.65% | 42.98% | 16.7682 | 18.5332
Li, ppm | 3.4013 | 0.46214 | 2.477 | 4.3256 | 2.0149 | 4.7877 | 13.59% | 27.17% | 40.76% | 3.2312 | 3.5713
Lu, ppm | 0.045083 | 0.0050383 | 0.035007 | 0.05516 | 0.029969 | 0.060198 | 11.18% | 22.35% | 33.53% | 0.042829 | 0.047337
Mg, wt.% | 0.11907 | 0.0091627 | 0.10075 | 0.1374 | 0.091585 | 0.14656 | 7.695% | 15.39% | 23.09% | 0.11312 | 0.12503
Mn, wt.% | 0.43395 | 0.012287 | 0.40938 | 0.45852 | 0.39709 | 0.47081 | 2.831% | 5.663% | 8.494% | 0.41225 | 0.45565
Mo, ppm | 11.6409 | 0.66266 | 10.3156 | 12.9663 | 9.6529 | 13.6289 | 5.693% | 11.39% | 17.08% | 11.0589 | 12.223
Na, wt.% | 0.058867 | 0.0064544 | 0.045959 | 0.071776 | 0.039504 | 0.078231 | 10.96% | 21.93% | 32.89% | 0.055924 | 0.061811
Nb, ppm | 0.69694 | 0.14621 | 0.40452 | 0.98935 | 0.25831 | 1.1356 | 20.98% | 41.96% | 62.94% | 0.66209 | 0.73178
Ni, ppm | 5.3202 | 0.45718 | 4.4058 | 6.2346 | 3.9487 | 6.6917 | 8.593% | 17.19% | 25.78% | 5.0542 | 5.5862
P, wt.% | 0.025122 | 0.0017283 | 0.021665 | 0.028579 | 0.019937 | 0.030307 | 6.88% | 13.76% | 20.64% | 0.023866 | 0.026378
Pb, wt.% | 0.82622 | 0.020949 | 0.78432 | 0.86811 | 0.76337 | 0.88906 | 2.536% | 5.071% | 7.607% | 0.78491 | 0.86753
Rb, ppm | 14.2798 | 1.4551 | 11.3695 | 17.1901 | 9.9144 | 18.6452 | 10.19% | 20.38% | 30.57% | 13.5658 | 14.9938
Re, ppm | 0.0040833 | 0.0006455 | 0.0027923 | 0.0053743 | 0.0021468 | 0.0060198 | 15.81% | 31.62% | 47.42% | 0.0038792 | 0.0042875
S, wt.% | 3.7662 | 0.1836 | 3.399 | 4.1334 | 3.2154 | 4.317 | 4.875% | 9.75% | 14.62% | 3.5779 | 3.9545
Sb, ppm | 28.5046 | 2.5774 | 23.3497 | 33.6594 | 20.7723 | 36.2368 | 9.042% | 18.08% | 27.13% | 27.0793 | 29.9298
Sc, ppm | 1.1189 | 0.15474 | 0.80944 | 1.4284 | 0.6547 | 1.5832 | 13.83% | 27.66% | 41.49% | 1.063 | 1.1749
Sn, ppm | 0.94654 | 0.16631 | 0.61392 | 1.2792 | 0.44762 | 1.4455 | 17.57% | 35.14% | 52.71% | 0.89921 | 0.99387
Sr, ppm | 20.7281 | 1.614 | 17.5001 | 23.9562 | 15.8861 | 25.5702 | 7.787% | 15.57% | 23.36% | 19.6917 | 21.7645
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.38144 | 0.05048 | 0.28048 | 0.4824 | 0.23 | 0.53288 | 13.23% | 26.47% | 39.7% | 0.36237 | 0.40051
Te, ppm | 0.050324 | 0.014578 | 0.021169 | 0.07948 | 0.0065907 | 0.094058 | 28.97% | 57.94% | 86.9% | 0.047808 | 0.052841
Th, ppm | 7.4911 | 0.60745 | 6.2762 | 8.706 | 5.6687 | 9.3134 | 8.109% | 16.22% | 24.33% | 7.1165 | 7.8656
Ti, wt.% | 0.02342 | 0.0045155 | 0.014389 | 0.032452 | 0.0098739 | 0.036967 | 19.28% | 38.56% | 57.84% | 0.022249 | 0.024592
Tl, ppm | 3.0305 | 0.14398 | 2.7426 | 3.3185 | 2.5986 | 3.4625 | 4.751% | 9.502% | 14.25% | 2.879 | 3.182
U, ppm | 2.77 | 0.29062 | 2.1888 | 3.3513 | 1.8981 | 3.6419 | 10.49% | 20.98% | 31.48% | 2.6315 | 2.9085
V, ppm | 3.1105 | 0.48173 | 2.147 | 4.0739 | 1.6653 | 4.5557 | 15.49% | 30.97% | 46.46% | 2.955 | 3.266
W, ppm | 1.4702 | 0.12962 | 1.2109 | 1.7294 | 1.0813 | 1.859 | 8.817% | 17.63% | 26.45% | 1.3966 | 1.5437
Y, ppm | 6.3963 | 0.50478 | 5.3868 | 7.4059 | 4.882 | 7.9107 | 7.892% | 15.78% | 23.68% | 6.0765 | 6.7161
Yb, ppm | 0.3408 | 0.040717 | 0.25936 | 0.42223 | 0.21865 | 0.46295 | 11.95% | 23.9% | 35.84% | 0.32376 | 0.35784
Zn, wt.% | 3.1792 | 0.089572 | 3.0001 | 3.3584 | 2.9105 | 3.4479 | 2.817% | 5.635% | 8.452% | 3.0203 | 3.3382
Zr, ppm | 54.7422 | 9.0395 | 36.6633 | 72.8212 | 27.6238 | 81.8607 | 16.51% | 33.03% | 49.54% | 52.0051 | 57.4794
Peroxide Fusion ICP
Al, wt.% | 6.8273 | 0.17089 | 6.4855 | 7.1691 | 6.3146 | 7.3399 | 2.503% | 5.006% | 7.509% | 6.4859 | 7.1686
As, ppm | 144.1385 | 22.4799 | 99.1786 | 189.0984 | 76.6987 | 211.5783 | 15.6% | 31.19% | 46.79% | 136.9316 | 151.3454
Ba, ppm | 3415.3709 | 242.8932 | 2929.5845 | 3901.1573 | 2686.6913 | 4144.0505 | 7.112% | 14.22% | 21.34% | 3244.6024 | 3586.1395
Bi, ppm | 8.6117 | 0.68374 | 7.2442 | 9.9791 | 6.5604 | 10.6629 | 7.94% | 15.88% | 23.82% | 8.1811 | 9.0422
Ca, wt.% | 1.1477 | 0.055343 | 1.037 | 1.2584 | 0.98165 | 1.3137 | 4.822% | 9.644% | 14.47% | 1.0903 | 1.2051
Cd, ppm | 82.6872 | 4.5916 | 73.5041 | 91.8703 | 68.9125 | 96.4619 | 5.553% | 11.11% | 16.66% | 78.5528 | 86.8216
Ce, ppm | 86.0175 | 2.4755 | 81.0664 | 90.9686 | 78.5909 | 93.4441 | 2.878% | 5.756% | 8.634% | 81.7166 | 90.3184
Co, ppm | 4.5042 | 0.40567 | 3.6929 | 5.3155 | 3.2872 | 5.7212 | 9.007% | 18.01% | 27.02% | 4.279 | 4.7294
Cs, ppm | 5.9997 | 0.28074 | 5.4383 | 6.5612 | 5.1575 | 6.842 | 4.679% | 9.358% | 14.04% | 5.6998 | 6.2997
Cu, wt.% | 0.17029 | 0.0048333 | 0.16063 | 0.17996 | 0.15579 | 0.18479 | 2.838% | 5.676% | 8.515% | 0.16178 | 0.17881
Dy, ppm | 3.6586 | 0.28396 | 3.0907 | 4.2265 | 2.8067 | 4.5105 | 7.761% | 15.52% | 23.28% | 3.4757 | 3.8415
Er, ppm | 1.4143 | 0.092332 | 1.2297 | 1.599 | 1.1373 | 1.6913 | 6.528% | 13.06% | 19.59% | 1.3436 | 1.485
Eu, ppm | 1.5774 | 0.30422 | 0.96895 | 2.1858 | 0.66474 | 2.49 | 19.29% | 38.57% | 57.86% | 1.4985 | 1.6563
Fe, wt.% | 4.0112 | 0.10938 | 3.7925 | 4.23 | 3.6831 | 4.3394 | 2.727% | 5.454% | 8.181% | 3.8107 | 4.2118
Ga, ppm | 20.4465 | 1.3367 | 17.7732 | 23.1198 | 16.4365 | 24.4565 | 6.537% | 13.07% | 19.61% | 19.4242 | 21.4688
Gd, ppm | 5.5297 | 0.24535 | 5.039 | 6.0204 | 4.7936 | 6.2657 | 4.437% | 8.874% | 13.31% | 5.2532 | 5.8062
Ge, ppm | 1.3052 | 0.44998 | 0.4052 | 2.2051 | 0 | 2.6551 | 34.48% | 68.95% | 103.4% | 1.2399 | 1.3704
Ho, ppm | 0.60395 | 0.023322 | 0.5573 | 0.65059 | 0.53398 | 0.67392 | 3.862% | 7.723% | 11.58% | 0.57375 | 0.63415
In, ppm | 2.1052 | 0.18696 | 1.7312 | 2.4791 | 1.5443 | 2.6661 | 8.881% | 17.76% | 26.64% | 1.9999 | 2.2104
K, wt.% | 2.8869 | 0.075335 | 2.7363 | 3.0376 | 2.6609 | 3.1129 | 2.61% | 5.219% | 7.829% | 2.7426 | 3.0313
La, ppm | 43.3584 | 1.722 | 39.9145 | 46.8024 | 38.1926 | 48.5243 | 3.971% | 7.943% | 11.91% | 41.1905 | 45.5264
Li, ppm | 18.1905 | 1.7275 | 14.7354 | 21.6456 | 13.0078 | 23.3731 | 9.497% | 18.99% | 28.49% | 17.281 | 19.1
Lu, ppm | 0.1433 | 0.027658 | 0.087985 | 0.19862 | 0.060327 | 0.22627 | 19.3% | 38.6% | 57.9% | 0.13614 | 0.15047
Mg, wt.% | 0.26668 | 0.012644 | 0.24139 | 0.29196 | 0.22874 | 0.30461 | 4.741% | 9.482% | 14.22% | 0.25334 | 0.28001
Mn, wt.% | 0.44679 | 0.0094931 | 0.42781 | 0.46578 | 0.41831 | 0.47527 | 2.125% | 4.249% | 6.374% | 0.42445 | 0.46913
Mo, ppm | 13.6828 | 1.6549 | 10.3729 | 16.9927 | 8.718 | 18.6476 | 12.09% | 24.19% | 36.28% | 12.9987 | 14.3669
Nb, ppm | 14.7113 | 2.2449 | 10.2216 | 19.201 | 7.9767 | 21.4459 | 15.26% | 30.52% | 45.78% | 13.9757 | 15.4468
Nd, ppm | 37.7173 | 1.0882 | 35.541 | 39.8936 | 34.4528 | 40.9818 | 2.885% | 5.77% | 8.655% | 35.8314 | 39.6032
P, wt.% | 0.030419 | 0.0013782 | 0.027662 | 0.033175 | 0.026284 | 0.034553 | 4.531% | 9.061% | 13.59% | 0.028898 | 0.03194
Pb, wt.% | 0.80589 | 0.033691 | 0.7385 | 0.87327 | 0.70481 | 0.90696 | 4.181% | 8.361% | 12.54% | 0.76559 | 0.84618
Pr, ppm | 9.9529 | 0.9413 | 8.0703 | 11.8355 | 7.129 | 12.7768 | 9.458% | 18.92% | 28.37% | 9.4553 | 10.4506
Rb, ppm | 136.3857 | 9.1931 | 117.9996 | 154.7719 | 108.8066 | 163.9649 | 6.74% | 13.48% | 20.22% | 129.5665 | 143.205
S, wt.% | 3.851 | 0.14978 | 3.5514 | 4.1505 | 3.4017 | 4.3003 | 3.889% | 7.779% | 11.67% | 3.6584 | 4.0435
Sb, ppm | 36.1703 | 1.3941 | 33.382 | 38.9586 | 31.9879 | 40.3528 | 3.854% | 7.709% | 11.56% | 34.3618 | 37.9788
Si, wt.% | 29.0185 | 0.7563 | 27.506 | 30.5311 | 26.7497 | 31.2874 | 2.606% | 5.212% | 7.819% | 27.5676 | 30.4695
Sm, ppm | 7.0366 | 0.45383 | 6.1289 | 7.9442 | 5.6751 | 8.3981 | 6.45% | 12.9% | 19.35% | 6.6847 | 7.3884
Sr, ppm | 142.3775 | 9.39 | 123.5975 | 161.1576 | 114.2074 | 170.5476 | 6.595% | 13.19% | 19.79% | 135.2586 | 149.4964
Tb, ppm | 0.77226 | 0.064536 | 0.64319 | 0.90133 | 0.57865 | 0.96587 | 8.357% | 16.71% | 25.07% | 0.73365 | 0.81088
Th, ppm | 13.7295 | 0.94042 | 11.8487 | 15.6103 | 10.9083 | 16.5508 | 6.85% | 13.7% | 20.55% | 13.043 | 14.416
Ti, wt.% | 0.12943 | 0.0059215 | 0.11758 | 0.14127 | 0.11166 | 0.14719 | 4.575% | 9.15% | 13.73% | 0.12296 | 0.1359
Tl, ppm | 9.9868 | 0.70175 | 8.5833 | 11.3903 | 7.8816 | 12.0921 | 7.027% | 14.05% | 21.08% | 9.4875 | 10.4862
Tm, ppm | 0.18296 | 0.017491 | 0.14798 | 0.21794 | 0.13049 | 0.23543 | 9.56% | 19.12% | 28.68% | 0.17381 | 0.19211
U, ppm | 5.7822 | 0.50618 | 4.7699 | 6.7946 | 4.2637 | 7.3008 | 8.754% | 17.51% | 26.26% | 5.4931 | 6.0713
V, ppm | 11.5024 | 2.2196 | 7.0632 | 15.9416 | 4.8436 | 18.1612 | 19.3% | 38.59% | 57.89% | 10.9273 | 12.0775
W, ppm | 4.3931 | 0.56329 | 3.2666 | 5.5197 | 2.7033 | 6.083 | 12.82% | 25.64% | 38.47% | 4.1735 | 4.6128
Y, ppm | 17.5649 | 0.95341 | 15.6581 | 19.4717 | 14.7046 | 20.4251 | 5.428% | 10.86% | 16.28% | 16.6866 | 18.4431
Yb, ppm | 1.0549 | 0.065278 | 0.92438 | 1.1855 | 0.8591 | 1.2508 | 6.188% | 12.38% | 18.56% | 1.0022 | 1.1077
Zn, wt.% | 3.2007 | 0.081931 | 3.0369 | 3.3646 | 2.9549 | 3.4465 | 2.56% | 5.12% | 7.679% | 3.0407 | 3.3608
Infrared Combustion
S, wt.% | 3.9116 | 0.11155 | 3.6885 | 4.1347 | 3.577 | 4.2463 | 2.852% | 5.704% | 8.556% | 3.716 | 4.1072
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 625
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Text Values:
NA | Unable to report due to QC failure (Lab 1.20)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 625
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSALABS, Vancouver, BC, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 625 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 16 | 8 | 255580 | 0.082507 | 0.70455 | 0.69159
2 | 1 | ANSLu | 1 | 3 | 2 | 255581 | 0.087379 | 0.69355 | 0.691
2 | 1 | ANSLu | 1 | 19 | 10 | 255582 | 0.083864 | 0.70368 | 0.69154
2 | 1 | ANSLu | 1 | 1 | 1 | 255583 | 0.08347 | 0.67581 | 0.69005
2 | 1 | ANSLu | 1 | 8 | 4 | 255584 | 0.085497 | 0.69589 | 0.69112
2 | 1 | ANSLu | 1 | 12 | 6 | 255585 | 0.083364 | 0.69021 | 0.69082
2 | 1 | ANSLu | 1 | 17 | 9 | 255586 | 0.086204 | 0.68808 | 0.69071
2 | 1 | ANSLu | 1 | 18 | 9 | 255587 | 0.086524 | 0.67684 | 0.69011
2 | 1 | ANSLu | 1 | 9 | 5 | 255588 | 0.087395 | 0.69661 | 0.69116
2 | 1 | ANSLu | 1 | 4 | 2 | 255589 | 0.085792 | 0.69175 | 0.6909
2 | 1 | ANSLu | 1 | 7 | 4 | 255590 | 0.085364 | 0.69385 | 0.69102
2 | 1 | ANSLu | 1 | 2 | 1 | 255591 | 0.084136 | 0.69784 | 0.69123
2 | 1 | ANSLu | 1 | 15 | 8 | 255592 | 0.083922 | 0.68474 | 0.69053
2 | 1 | ANSLu | 1 | 5 | 3 | 255593 | 0.085467 | 0.68202 | 0.69038
2 | 1 | ANSLu | 1 | 20 | 10 | 255594 | 0.08736 | 0.68767 | 0.69069
2 | 1 | ANSLu | 1 | 10 | 5 | 255595 | 0.088445 | 0.68941 | 0.69078
2 | 1 | ANSLu | 1 | 11 | 6 | 255596 | 0.084475 | 0.6859 | 0.69059
2 | 1 | ANSLu | 1 | 13 | 7 | 255597 | 0.086187 | 0.70854 | 0.6918
2 | 1 | ANSLu | 1 | 6 | 3 | 255598 | 0.087449 | 0.70097 | 0.6914
2 | 1 | ANSLu | 1 | 14 | 7 | 255599 | 0.085174 | 0.6692 | 0.6897
Mean | [6 empty] | 0.085499 | 0.69086 | 0.69086
Median | [7 empty] | 0.69098 | 0.69086
SD | [7 empty] | 0.010113 | 0.00053989
RSD | [7 empty] | 1.464% | 0.07815%
RSD Upscale Factor | [7 empty] | 0.053385
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Ag in OREAS 625 (Indicative Value 6.67 ppm) | [31 empty] | Indicative
Ag | Round | Replicate | Lab | [29 empty] | 1
| No | No | 1.28 | [29 empty] | ppm
|  |  | FA*GRAV | [29 empty] | 2
|  |  | 30g | [29 empty] | 2
| 1 | 1 | < 10 | [29 empty] | 1
| 1 | 2 | < 10 | [29 empty] | 1
| 1 | 3 | 10 | [29 empty] | 16
| 1 | 4 | < 10 | [29 empty] | 6.6667
| 1 | 5 | < 10 | [29 empty] | 7
| 1 | 6 | 10
| Mean |  | 10
| Median |  | 10
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 50%
| Z-Score (Absolute) |  | NA
| Analytical results for Au in OREAS 625 (Certified Value 0.667 ppm) | [31 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 30g | 50g | 30g | 15g | 30g | 30g | 50g | 30g | 15g | 30g | 30g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 0.67581 | 0.75 | 0.702 | 0.66 | 0.665 | 0.602 | 0.646 | 0.695 | 0.659 | 0.585 | 0.669 | 0.684 | 0.669 | 0.68265 | 0.67 | 0.65057 | 0.61 | 0.652 | 0.657 | 0.68 | 0.661 | 0.693 | 0.58835 | 0.648 | 0.677 | 0.669 | 0.669 | 0.697 | 0.67 | 0.64 | 1
| 1 | 2 | 0.69784 | 0.74 | 0.689 | 0.66 | 0.67 | 0.666 | 0.668 | 0.68 | 0.654 | 0.471 | 0.691 | 0.664 | 0.646 | 0.67866 | 0.675 | 0.6575 | 0.61 | 0.647 | 0.642 | 0.67 | 0.659 | 0.687 | 0.58825 | 0.642 | 0.692 | 0.657 | 0.678 | 0.684 | 0.64 | 0.63 | 0
| 1 | 3 | 0.69355 | 0.75 | 0.712 | 0.648 | 0.675 | 0.651 | 0.636 | 0.676 | 0.651 | 0.625 | 0.663 | 0.664 | 0.684 | 0.68085 | 0.669 | 0.653 | 0.58 | 0.649 | 0.655 | 0.7 | 0.613 | 0.692 | 0.59072 | 0.636 | 0.67 | 0.659 | 0.669 | 0.708 | 0.68 | 0.63 | 16
| 1 | 4 | 0.69175 | 0.74 | 0.704 | 0.664 | 0.672 | 0.669 | 0.648 | 0.665 | 0.666 | 0.554 | 0.667 | 0.664 | 0.67 | 0.68233 | 0.654 | 0.67105 | 0.6 | 0.64 | 0.642 | 0.69 | 0.597 | 0.688 | 0.57726 | 0.641 | 0.684 | 0.654 | 0.663 | 0.68 | 0.66 | 0.65 | 0.66675
| 1 | 5 | 0.68202 | 0.74 | 0.717 | 0.65 | 0.675 | 0.664 | 0.625 | 0.677 | 0.624 | 0.64 | 0.656 | 0.678 | 0.679 | 0.68114 | 0.674 | 0.65878 | 0.6 | 0.652 | 0.642 | 0.71 | 0.638 | 0.694 | 0.59676 | 0.634 | 0.677 | 0.661 | 0.66 | 0.663 | 0.68 | 0.65 | 7
| 1 | 6 | 0.70097 | 0.74 | 0.712 | 0.675 | 0.685 | 0.655 | 0.647 | 0.676 | 0.664 | 0.606 | 0.679 | 0.705 | 0.681 | 0.68314 | 0.671 | 0.67378 | 0.61 | 0.646 | 0.656 | 0.7 | 0.633 | 0.686 | 0.58904 | 0.629 | 0.666 | 0.659 | 0.671 | 0.698 | 0.66 | 0.64
|  | 7 | 00: 0.69385
|  | 8 | 00: 0.69589
|  | 9 | 00: 0.69661
|  | 10 | 00: 0.68941
|  | 11 | 00: 0.6859
|  | 12 | 00: 0.69021
|  | 13 | 00: 0.70854
|  | 14 | 00: 0.6692
|  | 15 | 00: 0.68474
|  | 16 | 00: 0.70455
|  | 17 | 00: 0.68808
|  | 18 | 00: 0.67684
|  | 19 | 00: 0.70368
|  | 20 | 00: 0.68767
| Mean |  | 0.69086 | 0.74333 | 0.706 | 0.6595 | 0.67367 | 0.65117 | 0.645 | 0.67817 | 0.653 | 0.58017 | 0.67083 | 0.6765 | 0.6715 | 0.68146 | 0.66883 | 0.66078 | 0.60167 | 0.64767 | 0.649 | 0.69167 | 0.6335 | 0.69 | 0.5884 | 0.63833 | 0.67767 | 0.65983 | 0.66833 | 0.68833 | 0.665 | 0.64
| Median |  | 0.69098 | 0.74 | 0.708 | 0.66 | 0.6735 | 0.6595 | 0.6465 | 0.6765 | 0.6565 | 0.5955 | 0.668 | 0.671 | 0.6745 | 0.68173 | 0.6705 | 0.65814 | 0.605 | 0.648 | 0.6485 | 0.695 | 0.6355 | 0.69 | 0.58869 | 0.6385 | 0.677 | 0.659 | 0.669 | 0.6905 | 0.665 | 0.64
| Std Dev. |  | 0.010113 | 0.005164 | 0.01002 | 0.0098336 | 0.0066833 | 0.025039 | 0.014311 | 0.009704 | 0.01531 | 0.061454 | 0.012432 | 0.016368 | 0.013867 | 0.0016321 | 0.0076267 | 0.009529 | 0.01169 | 0.0045019 | 0.0076942 | 0.01472 | 0.025217 | 0.0034059 | 0.0063215 | 0.0067132 | 0.009395 | 0.0050761 | 0.006314 | 0.016033 | 0.015166 | 0.0089443
| Rel.Std.Dev. |  | 1.464% | 0.6947% | 1.419% | 1.491% | 0.9921% | 3.845% | 2.219% | 1.431% | 2.345% | 10.59% | 1.853% | 2.419% | 2.065% | 0.2395% | 1.14% | 1.442% | 1.943% | 0.6951% | 1.186% | 2.128% | 3.981% | 0.4936% | 1.074% | 1.052% | 1.386% | 0.7693% | 0.9447% | 2.329% | 2.281% | 1.398%
| PDM3 |  | 3.616% | 11.49% | 5.887% | -1.087% | 1.038% | -2.337% | -3.262% | 1.713% | -2.062% | -12.99% | 0.6129% | 1.463% | 0.7129% | 2.207% | 0.3129% | -0.895% | -9.761% | -2.862% | -2.662% | 3.738% | -4.986% | 3.488% | -11.75% | -4.262% | 1.638% | -1.037% | 0.2379% | 3.238% | -0.262% | -4.012%
| Z-Score (Absolute) |  | NA | 3.3 | 1.73 | 0.23 | 0.37 | 0.58 | 0.84 | 0.55 | 0.51 | 3.58 | 0.25 | 0.48 | 0.27 | 0.69 | 0.16 | 0.18 | 2.67 | 0.73 | 0.67 | 1.12 | 1.33 | 1.05 | 3.23 | 1.12 | 0.53 | 0.22 | 0.14 | 0.98 | 0 | 1.05
| Analytical results for Pd in OREAS 625 (Indicative Value 0.75 ppb) | [31 empty] | Indicative
Pd | Round | Replicate | 00: Lab | 1
| No | No | 00: 1.23 | ppb
|  |  | 00: FA*OES | 3
|  |  | 00: 30g | 3
| 1 | 1 | [30 empty] | 1 | bar: 00
| 1 | 2 | [30 empty] | 1 | bar: 00
| 1 | 3 | 00: < 1 | 16
| 1 | 4 | 00: < 1 | 0.75
| 1 | 5 | [30 empty] | 7 | bar: 00
| 1 | 6 | 00: < 1
Mean | bar: 00
Median | bar: 00
| Std Dev. |  | 00: 0
Rel.Std.Dev. | bar: 00
| PDM3 |  | 00: 33.33%
Z-Score (Absolute) | bar: 00
| Analytical results for Pt in OREAS 625 (Indicative Value < 10 ppb) | [31 empty] | Indicative
Pt | Round | Replicate | 00: Lab | 1
| No | No | 00: 1.23 | ppb
|  |  | 00: FA*OES | 1
|  |  | 00: 30g | 1
| 1 | 1 | 00: < 10 | 1
| 1 | 2 | 00: < 10 | 1
| 1 | 3 | 00: < 10 | 16
| 1 | 4 | 00: < 10 | < 10
| 1 | 5 | 00: < 10 | 7
| 1 | 6 | 00: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 625 (Certified Value 0.667 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.23 | 1.24 | 1.26 | 1.27 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*OES/MS | 3
|  |  | 0.085g | 30g | 30g | 15g | 25g | 25g | 25g | 25g | 25g | 25g | 25g | 15g | 10g | 50g | 25g | 10g | 30g | 50g | 25g | 30g | 25g | 20g | 3
| 1 | 1 | 0.67581 | 0.74 | 0.693 | 0.6984 | 0.671 | 0.671 | 0.681 | 0.683 | 0.635 | 0.5 | 0.646 | 0.6976 | 0.66 | 0.666 | 0.6962 | 0.59 | 0.6 | 0.65 | > 0.5 | 0.661 | 0.636 | 0.6471 | 1
| 1 | 2 | 0.69784 | 0.74 | 0.667 | 0.7566 | 0.667 | 0.682 | 0.67 | 0.677 | 0.611 | 0.54 | 0.67 | 0.6609 | 0.69 | 0.657 | 0.69241 | 0.59 | 0.61 | 0.7 | > 0.5 | 0.662 | 0.621 | 0.6706 | 0
| 1 | 3 | 0.69355 | 0.75 | 0.639 | 0.7275 | 0.679 | 0.652 | 0.673 | 0.667 | 0.635 | 0.51 | 0.655 | 0.7049 | 0.69 | 0.69 | 0.67151 | 0.56 | 0.6 | 0.65 | > 0.5 | 0.682 | 0.654 | 0.651 | 16
| 1 | 4 | 0.69175 | 0.74 | 0.647 | 0.6693 | 0.68 | 0.678 | 0.661 | 0.673 | 0.637 | 0.49 | 0.665 | 0.707 | 0.69 | 0.63 | 0.67927 | 0.57 | 0.6 | 0.64 | > 0.5 | 0.675 | 0.632 | 0.6856 | 0.66659
| 1 | 5 | 0.68202 | 0.74 | 0.667 | 0.7128 | 0.675 | 0.679 | 0.66 | 0.669 | 0.608 | 0.54 | 0.663 | 0.6889 | 0.71 | 0.647 | 0.68615 | 0.58 | 0.6 | 0.63 | > 0.5 | 0.679 | 0.628 | 0.6814 | 9
| 1 | 6 | 0.70097 | 0.74 | 0.631 | 0.6984 | 0.689 | 0.651 | 0.669 | 0.668 | 0.645 | 0.54 | 0.638 | 0.6672 | 0.71 | 0.636 | 0.66657 | 0.6 | 0.64 | 0.64 | > 0.5 | 0.672 | 0.643 | 0.6806
|  | 7 | 00: 0.69385
|  | 8 | 00: 0.69589
|  | 9 | 00: 0.69661
|  | 10 | 00: 0.68941
|  | 11 | 00: 0.6859
|  | 12 | 00: 0.69021
|  | 13 | 00: 0.70854
|  | 14 | 00: 0.6692
|  | 15 | 00: 0.68474
|  | 16 | 00: 0.70455
|  | 17 | 00: 0.68808
|  | 18 | 00: 0.67684
|  | 19 | 00: 0.70368
|  | 20 | 00: 0.68767
| Mean |  | 0.69086 | 0.74167 | 0.65733 | 0.7105 | 0.67683 | 0.66883 | 0.669 | 0.67283 | 0.6285 | 0.52 | 0.65617 | 0.68775 | 0.69167 | 0.65433 | 0.68202 | 0.58167 | 0.60833 | 0.65167 |  | 0.67183 | 0.63567 | 0.66938
| Median |  | 0.69098 | 0.74 | 0.657 | 0.7056 | 0.677 | 0.6745 | 0.6695 | 0.671 | 0.635 | 0.525 | 0.659 | 0.69325 | 0.69 | 0.652 | 0.68271 | 0.585 | 0.6 | 0.645 |  | 0.6735 | 0.634 | 0.6756
| Std Dev. |  | 0.010113 | 0.0040825 | 0.022783 | 0.029698 | 0.007705 | 0.013906 | 0.007823 | 0.0062102 | 0.015202 | 0.022804 | 0.012254 | 0.019526 | 0.018348 | 0.021897 | 0.011688 | 0.01472 | 0.016021 | 0.024833 |  | 0.0087044 | 0.011639 | 0.016548
| Rel.Std.Dev. |  | 1.464% | 0.5504% | 3.466% | 4.18% | 1.138% | 2.079% | 1.169% | 0.923% | 2.419% | 4.385% | 1.868% | 2.839% | 2.653% | 3.346% | 1.714% | 2.531% | 2.634% | 3.811% |  | 1.296% | 1.831% | 2.472%
| PDM3 |  | 3.64% | 11.26% | -1.389% | 6.587% | 1.537% | 0.3366% | 0.3616% | 0.9367% | -5.714% | -21.99% | -1.564% | 3.174% | 3.762% | -1.839% | 2.315% | -12.74% | -8.739% | -2.239% |  | 0.7867% | -4.639% | 0.4191%
| Z-Score (Absolute) |  | NA | 3.08 | 0.49 | 1.76 | 0.34 | 0 | 0 | 0.17 | 1.71 | 6.31 | 0.54 | 0.8 | 0.96 | 0.62 | 0.56 | 3.7 | 2.57 | 0.73 | NA | 0.12 | 1.41 | 0.02
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 625 (Certified Value 11.7 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 13.2 | 11.2 | 11.9 | 11.7 | 11.9 | 11.1 | 11.2 | 11.45 | 11.1 | 11.75 | 12 | 11.52 | 11.8124 | 10.6556 | 12.1337 | 10.64 | 12.4 | 10.9 | 12.67 | 11.2 | 11.4 | 12.88 | 11.86 | 12 | 12.2 | 11.9 | 11.78 | 12 | 11.58 | 1
| 1 | 2 | 13.3 | 11.1 | 11.1 | 11.9 | 10.5 | 10.85 | 11 | 12.25 | 11.4 | 11.89 | 12.1 | 12.27 | 11.6454 | 10.5588 | 12.0211 | 11.09 | 12.8 | 10.2 | 12.95 | 11.6 | 11.4 | 12.73 | 11.67 | 12.6 | 12.1 | 11.8 | 11.42 | 12 | 11.2 | 1
| 1 | 3 | 13.6 | 11 | 11.4 | 12.15 | 11.75 | 10.65 | 11.15 | 12.65 | 11.9 | 12.14 | 11.8 | 12.09 | 11.643 | 12.1239 | 12.2284 | 10.46 | 12.6 | 10.3 | 12.27 | 11.4 | 11.4 | 12.95 | 11.68 | 12.5 | 12.3 | 11.7 | 11.44 | 12 | 11.62 | 16
| 1 | 4 | 13.8 | 11 | 11.5 | 12.4 | 11.85 | 11.3 | 11.25 | 11.95 | 12.1 | 11.98 | 11.8 | 11.92 | 11.4011 | 11.3211 | 11.9039 | 10.65 | 12.6 | 10.9 | 12.42 | 11.4 | 11.4 | 12.22 | 11.87 | 12.3 | 12.6 | 11.8 | 11.34 | 12 | 11.73 | 11.6934
| 1 | 5 | 13.5 | 11.4 | 11.2 | 11.9 | 11.8 | 11.1 | 11.1 | 12.75 | 11.5 | 11.92 | 11.7 | 11.72 | 11.3055 | 12.01 | 11.6635 | 11.57 | 12.3 | 10.5 | 12.75 | 11.7 | 11.7 | 12.76 | 11.59 | 12.3 | 12.3 | 11.8 | 11.45 | 12 | 11.76 | 11
| 1 | 6 | 13 | 10.8 | 11 | 12.1 | 11.9 | 10.7 | 11.7 | 11.6 | 11.6 | 11.95 | 11.8 | 11.85 | 11.5943 | 10.9644 | 11.7283 | 10.41 | 12.5 | 10 | 12.35 | 11.5 | 11.2 | 12.93 | 11.54 | 12.5 | 12.4 | 11.7 | 11.49 | 12 | 11.8
| Mean |  | 13.4 | 11.0833 | 11.35 | 12.025 | 11.6167 | 10.95 | 11.2333 | 12.1083 | 11.6 | 11.9383 | 11.8667 | 11.895 | 11.5669 | 11.2723 | 11.9465 | 10.8033 | 12.5333 | 10.4667 | 12.5683 | 11.4667 | 11.4167 | 12.745 | 11.7017 | 12.3667 | 12.3167 | 11.7833 | 11.4867 | 12 | 11.615
| Median |  | 13.4 | 11.05 | 11.3 | 12 | 11.825 | 10.975 | 11.175 | 12.1 | 11.55 | 11.935 | 11.8 | 11.885 | 11.6186 | 11.1427 | 11.9625 | 10.645 | 12.55 | 10.4 | 12.545 | 11.45 | 11.4 | 12.82 | 11.675 | 12.4 | 12.3 | 11.8 | 11.445 | 12 | 11.675
| Std Dev. |  | 0.28983 | 0.20412 | 0.32711 | 0.24444 | 0.55015 | 0.2569 | 0.24427 | 0.53704 | 0.35777 | 0.12703 | 0.15055 | 0.26554 | 0.18372 | 0.67172 | 0.22345 | 0.44559 | 0.17512 | 0.37238 | 0.2637 | 0.17512 | 0.16021 | 0.27223 | 0.13674 | 0.21602 | 0.17224 | 0.075277 | 0.15201 | 0 | 0.21998
| Rel.Std.Dev. |  | 2.163% | 1.842% | 2.882% | 2.033% | 4.736% | 2.346% | 2.174% | 4.435% | 3.084% | 1.064% | 1.269% | 2.232% | 1.588% | 5.959% | 1.87% | 4.125% | 1.397% | 3.558% | 2.098% | 1.527% | 1.403% | 2.136% | 1.169% | 1.747% | 1.398% | 0.6388% | 1.323% | 0% | 1.894%
| PDM3 |  | 14.59% | -5.217% | -2.937% | 2.836% | -0.6564% | -6.358% | -3.935% | 3.548% | -0.7989% | 2.094% | 1.482% | 1.724% | -1.082% | -3.601% | 2.164% | -7.612% | 7.183% | -10.49% | 7.482% | -1.939% | -2.367% | 8.993% | 0.07051% | 5.757% | 5.33% | 0.7689% | -1.768% | 2.622% | -0.6706%
| Z-Score (Absolute) |  | 3.48 | 1.15 | 0.62 | 0.73 | 0.08 | 1.42 | 0.85 | 0.9 | 0.12 | 0.56 | 0.41 | 0.47 | 0.18 | 0.77 | 0.57 | 1.71 | 1.75 | 2.38 | 1.82 | 0.38 | 0.48 | 2.17 | 0.08 | 1.41 | 1.31 | 0.25 | 0.34 | NA | 0.09
Results from laboratory 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Al in OREAS 625 (Certified Value 6.68 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 6.5 | 6.51 | 6.59 | 6.82 | 6.26 | 6.4 | 6.89 | 6.97 | 6.62 | 6.87 | 7.0253 | 6.7023 | 6.653 | 6.39 | 6.69 | 6.1 | 6.9868 | 6.2236 | 6.91 | 6.7721 | 6.62 | 6.6 | 6.88 | 6.86 | 7.18 | 6.99 |  |  |  | 1
| 1 | 2 | 6.35 | 6.87 | 6.51 | 7.15 | 6.29 | 6.56 | 6.58 | 6.88 | 6.41 | 6.85 | 7.0362 | 6.5481 | 6.7206 | 6.44 | 6.66 | 6.07 | 6.4808 | 6.3531 | 6.85 | 6.7121 | 6.66 | 6.56 | 6.32 | 7.04 | 7.22 | 6.87 |  |  |  | 0
| 1 | 3 | 6.61 | 6.59 | 6.76 | 7.01 | 6.25 | 6.09 | 6.51 | 6.95 | 6.48 | 6.61 | 6.9266 | 6.5826 | 6.8849 | 6.37 | 6.75 | 6.12 | 6.6485 | 6.2698 | 6.88 | 6.7531 | 6.61 | 6.55 | 6.68 | 6.82 | 7.17 | 6.97 |  |  |  | 16
| 1 | 4 | 6.69 | 7.38 | 6.64 | 7.05 | 6.5 | 6.44 | 6.63 | 6.94 | 6.49 | 6.77 | 6.963 | 6.4562 | 6.7043 | 6.42 | 6.68 | 6.02 | 6.8534 | 6.3522 | 7.05 | 6.7568 | 6.64 | 6.55 | 6.77 | 6.96 | 7.15 | 6.86 |  |  |  | 6.6781
| 1 | 5 | 6.62 | 7.14 | 6.71 | 6.83 | 6.44 | 6.25 | 6.51 | 7.43 | 6.48 | 6.68 | 6.792 | 6.3818 | 6.6016 | 6.37 | 6.71 | 6.2 | 6.8986 | 6.4182 | 7.12 | 6.7606 | 6.63 | 6.6 | 6.74 | 6.77 | 7.17 | 7.11 |  |  |  | 12
| 1 | 6 | 6.57 | 6.12 | 6.36 | 6.92 | 6.64 | 6.35 | 6.8 | 6.81 | 6.53 | 6.63 | 6.928 | 6.5077 | 6.8915 | 6.38 | 6.68 | 6.16 | 6.7863 | 6.3395 | 6.9 | 6.7564 | 6.49 | 6.57 | 6.62 | 6.96 | 7.13 | 7.13
| Mean |  | 6.5567 | 6.7683 | 6.595 | 6.9633 | 6.3967 | 6.3483 | 6.6533 | 6.9967 | 6.5017 | 6.735 | 6.9452 | 6.5298 | 6.7427 | 6.395 | 6.695 | 6.1117 | 6.7757 | 6.3261 | 6.9517 | 6.7519 | 6.6083 | 6.5717 | 6.6683 | 6.9017 | 7.17 | 6.9883
| Median |  | 6.59 | 6.73 | 6.615 | 6.965 | 6.365 | 6.375 | 6.605 | 6.945 | 6.485 | 6.725 | 6.9455 | 6.5279 | 6.7125 | 6.385 | 6.685 | 6.11 | 6.8198 | 6.3459 | 6.905 | 6.7566 | 6.625 | 6.565 | 6.71 | 6.91 | 7.17 | 6.98
| Std Dev. |  | 0.11894 | 0.4564 | 0.14488 | 0.13018 | 0.15706 | 0.16266 | 0.15782 | 0.22015 | 0.06969 | 0.11167 | 0.08846 | 0.1102 | 0.12025 | 0.02881 | 0.031464 | 0.064005 | 0.18383 | 0.068918 | 0.10759 | 0.020559 | 0.06047 | 0.023166 | 0.19188 | 0.10167 | 0.030332 | 0.11462
| Rel.Std.Dev. |  | 1.814% | 6.743% | 2.197% | 1.869% | 2.455% | 2.562% | 2.372% | 3.147% | 1.072% | 1.658% | 1.274% | 1.688% | 1.783% | 0.4505% | 0.47% | 1.047% | 2.713% | 1.089% | 1.548% | 0.3045% | 0.9151% | 0.3525% | 2.877% | 1.473% | 0.423% | 1.64%
| PDM3 |  | -1.819% | 1.351% | -1.245% | 4.271% | -4.214% | -4.938% | -0.3711% | 4.77% | -2.642% | 0.8518% | 3.999% | -2.221% | 0.9666% | -4.239% | 0.2529% | -8.482% | 1.462% | -5.272% | 4.096% | 1.104% | -1.045% | -1.594% | -0.1464% | 3.348% | 7.366% | 4.645%
| Z-Score (Absolute) |  | 0.51 | 0.35 | 0.35 | 1.14 | 1.16 | 1.35 | 0.12 | 1.28 | 0.73 | 0.22 | 1.07 | 0.62 | 0.25 | 1.16 | 0.05 | 2.32 | 0.38 | 1.44 | 1.1 | 0.29 | 0.3 | 0.45 | 0.05 | 0.89 | 1.98 | 1.25
| Analytical results for As in OREAS 625 (Certified Value 143 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 139 | 143 | 141 | 135.5 | 142.5 | 147.5 | 140 | 148.5 | 144 | 141.6 | 127 | 136.9 | 145.0962 | 126.101 | 142.5154 | 109.6 | 149 | 41.6 | 126 | 128.5 | 142 | 149 | 166 | 147 | 146 | 145.3 | 150 | 140 |  | 1
| 1 | 2 | 145 | 143 | 137 | 139 | 133 | 146 | 136.5 | 148 | 144 | 140.1 | 129 | 148.4 | 141.5722 | 138.022 | 145.0265 | 104.6 | 147 | 39.9 | 123 | 132.8 | 139 | 144 | 162 | 147 | 146 | 143.9 | 144 | 140 |  | 21
| 1 | 3 | 143 | 143 | 147 | 143 | 144 | 144 | 136.5 | 150 | 145 | 143.8 | 128 | 145.3 | 140.53 | 134.589 | 146.1817 | 115.1 | 148 | 38.8 | 125 | 131.2 | 141 | 142 | 176 | 147 | 154 | 143.4 | 149 | 140 |  | 16
| 1 | 4 | 147 | 141 | 146 | 142.5 | 144.5 | 147.5 | 139.5 | 149.5 | 144 | 142.8 | 130 | 144.6 | 138.2634 | 121.801 | 141.1059 | 104.5 | 150 | 41.7 | 130 | 133.7 | 141 | 143 | 176 | 148 | 149 | 143.8 | 147 | 130 |  | 143.3644
| 1 | 5 | 142 | 146 | 141 | 140 | 138.5 | 149.5 | 138.5 | 161.5 | 146 | 139 | 124 | 145 | 133.8365 | 126.348 | 141.4672 | 114.3 | 154 | 38.5 | 128 | 139.5 | 145 | 139 | 178 | 145 | 143 | 141.8 | 146 | 140 |  | 13
| 1 | 6 | 144 | 144 | 145 | 141.5 | 143.5 | 144 | 145 | 144.5 | 141 | 140.6 | 125 | 147.7 | 139.401 | 125.991 | 143.7308 | 104.9 | 151 | 38.7 | 128 | 133.2 | 139 | 145 | 165 | 148 | 145 | 147.6 | 148 | 140
| Mean |  | 143.3333 | 143.3333 | 142.8333 | 140.25 | 141 | 146.4167 | 139.3333 | 150.3333 | 144 | 141.3167 | 127.1667 | 144.65 | 139.7832 | 128.8087 | 143.3379 | 108.8333 | 149.8333 | 39.8667 | 126.6667 | 133.15 | 141.1667 | 143.6667 | 170.5 | 147 | 147.1667 | 144.3 | 147.3333 | 138.3333
| Median |  | 143.5 | 143 | 143 | 140.75 | 143 | 146.75 | 139 | 149 | 144 | 141.1 | 127.5 | 145.15 | 139.9655 | 126.2245 | 143.1231 | 107.25 | 149.5 | 39.35 | 127 | 133 | 141 | 143.5 | 171 | 147 | 146 | 143.85 | 147.5 | 140
| Std Dev. |  | 2.7325 | 1.633 | 3.8166 | 2.7704 | 4.4721 | 2.1775 | 3.1411 | 5.8023 | 1.6733 | 1.7781 | 2.3166 | 4.0982 | 3.7364 | 6.1438 | 2.0125 | 4.9379 | 2.4833 | 1.4652 | 2.5033 | 3.638 | 2.2286 | 3.3267 | 6.921 | 1.0954 | 3.8687 | 1.9677 | 2.1602 | 4.0825
| Rel.Std.Dev. |  | 1.906% | 1.139% | 2.672% | 1.975% | 3.172% | 1.487% | 2.254% | 3.86% | 1.162% | 1.258% | 1.822% | 2.833% | 2.673% | 4.77% | 1.404% | 4.537% | 1.657% | 3.675% | 1.976% | 2.732% | 1.579% | 2.316% | 4.059% | 0.7452% | 2.629% | 1.364% | 1.466% | 2.951%
| PDM3 |  | -0.0217% | -0.0217% | -0.3705% | -2.172% | -1.649% | 2.129% | -2.812% | 4.861% | 0.4433% | -1.428% | -11.3% | 0.8967% | -2.498% | -10.15% | -0.01848% | -24.09% | 4.512% | -72.19% | -11.65% | -7.125% | -1.533% | 0.2108% | 18.93% | 2.536% | 2.652% | 0.6526% | 2.768% | -3.509%
| Z-Score (Absolute) |  | 0.05 | 0.05 | 0.05 | 0.54 | 0.4 | 0.63 | 0.71 | 1.38 | 0.17 | 0.34 | 3.03 | 0.3 | 0.63 | 2.72 | 0.05 | 6.52 | 1.29 | 19.65 | 3.13 | 1.89 | 0.36 | 0.11 | 5.22 | 0.75 | 0.78 | 0.23 | 0.81 | 0.9
| Analytical results for B in OREAS 625 (Indicative Value 20.8 ppm) | [30 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02 |  | ppm
|  |  | 1.01: 4A*OES; 1.02: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 21; 1.02: < 20 |  | 1
| 1 | 2 | 1.01: 20; 1.02: < 20 |  | 3
| 1 | 3 | 1.01: 22; 1.02: < 20 |  | 16
| 1 | 4 | 1.01: 21; 1.02: < 20 |  | 20.8333
| 1 | 5 | 1.01: 20; 1.02: < 20 |  | 9
| 1 | 6 | 1.01: 21; 1.02: < 20
| Mean |  | 1.01: 20.8333
| Median |  | 1.01: 21
| Std Dev. |  | 1.01: 0.75277
| Rel.Std.Dev. |  | 1.01: 3.613%
| PDM3 |  | 1.01: 1.554e-13%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ba in OREAS 625 (Indicative Value 528 ppm) | [30 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | [7 empty] | 0
0
| 1 | 1 | 358 | 118 | 50 | 540 | 410 | 180 | 300 | 160 | 1970 | 37 | 3029.9864 | 3266 | 1020 | > 1000 | 3145 | 3214 | 179 | 442 | 2735 | 770 | 197 | 936 | [7 empty] | 1
| 1 | 2 | 348 | 110 | 46 | 1860 | 480 | 180 | 230 | 230 | 1750 | 39 | 2985.6442 | 3295 | 1030 | > 1000 | 2942 | 3291 | 203 | 446 | 2129 | 1563 | 319 | 1884 | [7 empty] | 4
| 1 | 3 | 263 | 139 | 45 | 820 | 270 | 180 | 240 | 170 | 1900 | 25 | 3090.9398 | 3292 | 1070 | > 1000 | 3124 | 3236 | 180 | 472 | 2686 | 707 | 238 | 1074 | [7 empty] | 16
| 1 | 4 | 317 | 127 | 50 | 1100 | 470 | 180 | 290 | 140 | 1850 | 33 | 3301.2764 | 3187 | 1060 | > 1000 | 3134 | 3254 | 194 | 466 | 2711 | 1819 | 224 | 1758 | [7 empty] | 527.7437
| 1 | 5 | 241 | 140 | 46 | 410 | 720 | 170 | 240 | 150 | 1670 | 33 | 3193.207 | 3257 | 1100 | > 1000 | 3138 | 3292 | 168 | 441 | 2697 | 1033 | 230 | 1083 | [7 empty] | 10
| 1 | 6 | 335 | 144 | 50 | 790 | 390 | 180 | 330 | 210 | 2150 | 39 | 2805.7023 | 3237 | 1100 | > 1000 | 3161 | 3279 | 230 | 465 | 2648 | 207 | 223 | 914
| Mean |  | 310.3333 | 129.6667 | 47.8333 | 920 | 456.6667 | 178.3333 | 271.6667 | 176.6667 | 1881.6667 | 34.3333 | 3067.7927 | 3255.6667 | 1063.3333 |  | 3107.3333 | 3261 | 192.3333 | 455.3333 | 2601 | 1016.5 | 238.5 | 1274.8333
| Median |  | 326 | 133 | 48 | 805 | 440 | 180 | 265 | 165 | 1875 | 35 | 3060.4631 | 3261.5 | 1065 |  | 3136 | 3266.5 | 187 | 455.5 | 2691.5 | 901.5 | 227 | 1078.5
| Std Dev. |  | 47.7312 | 13.6333 | 2.4014 | 519.1146 | 149.3542 | 4.0825 | 40.7022 | 35.5903 | 169.2828 | 5.3166 | 171.8419 | 40.0982 | 33.8625 |  | 81.9309 | 31.8371 | 22.169 | 13.8227 | 233.0193 | 592.416 | 41.7744 | 430.5059
| Rel.Std.Dev. |  | 15.38% | 10.51% | 5.02% | 56.43% | 32.71% | 2.289% | 14.98% | 20.15% | 8.996% | 15.49% | 5.601% | 1.232% | 3.185% |  | 2.637% | 0.9763% | 11.53% | 3.036% | 8.959% | 58.28% | 17.52% | 33.77%
| PDM3 |  | -41.2% | -75.43% | -90.94% | 74.33% | -13.47% | -66.21% | -48.52% | -66.52% | 256.5% | -93.49% | 481.3% | 516.9% | 101.5% |  | 488.8% | 517.9% | -63.56% | -13.72% | 392.9% | 92.61% | -54.81% | 141.6%
| Z-Score (Absolute) |  | 0.23 | 0.52 | 0.65 | 0.74 | 0 | 0.44 | 0.3 | 0.45 | 2.28 | 0.67 | 4.17 | 4.47 | 0.97 | NA | 4.23 | 4.48 | 0.42 | 0 | 3.42 | 0.89 | 0.35 | 1.31
| Analytical results for Be in OREAS 625 (Certified Value 2.57 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | < 5 | 2.9 | 2 | 2.55 | 2.42 | 2.43 | 2.66 | 2.54 | 2.51 | 2.54 | 3 | 2.58 | 2.5209 | 2.6702 | 2.41 | 2.4 | 2 | < 5 | 2.49 | 2.6 | 3.5709 | 2.95 | 3 | 3 | 2.8 | 2.34 | 2.8 | 2.52 |  | 1
| 1 | 2 | < 5 | 2.7 | 2 | 2.62 | 2.38 | 2.45 | 2.55 | 2.52 | 2.55 | 2.6 | 3 | 2.56 | 2.4782 | 2.4496 | 2.42 | 2.4 | 2 | < 5 | 2.56 | 2.7 | 3.403 | 2.81 | 3 | 3 | 2.7 | 2.59 | 2.6 | 2.43 |  | 14
| 1 | 3 | < 5 | 2.6 | 2 | 2.59 | 2.41 | 2.33 | 2.55 | 2.54 | 2.92 | 2.62 | 3 | 2.65 | 2.502 | 2.5684 | 2.42 | 2.4 | 2 | < 5 | 2.52 | 2.8 | 3.3848 | 2.77 | 3 | 3 | 2.9 | 2.51 | 2.3 | 2.41 |  | 16
| 1 | 4 | < 5 | 2.7 | 2 | 2.6 | 2.44 | 2.44 | 2.59 | 2.54 | 2.91 | 2.63 | 3 | 2.58 | 2.4342 | 2.4266 | 2.38 | 2.3 | 2 | < 5 | 2.54 | 2.7 | 3.5396 | 2.76 | 3 | 3 | 2.9 | 2.59 | 2.5 | 2.45 |  | 2.5671
| 1 | 5 | < 5 | 2.7 | 2 | 2.53 | 2.41 | 2.38 | 2.56 | 2.7 | 2.85 | 2.56 | 3 | 2.69 | 2.4255 | 2.6008 | 2.37 | 2.4 | 2 | < 5 | 2.57 | 2.7 | 3.3395 | 3.03 | 3 | 3 | 2.7 | 2.62 | 2.3 | 2.45 |  | 14
| 1 | 6 | < 5 | 2.5 | 2 | 2.57 | 2.49 | 2.41 | 2.64 | 2.48 | 2.57 | 2.56 | 3 | 2.69 | 2.4582 | 2.56 | 2.35 | 2.3 | 2 | < 5 | 2.55 | 2.7 | 3.5043 | 3 | 3 | 3 | 2.7 | 2.55 | 2.4 | 2.55
| Mean |  |  | 2.6833 | 2 | 2.5767 | 2.425 | 2.4067 | 2.5917 | 2.5533 | 2.7183 | 2.585 | 3 | 2.625 | 2.4698 | 2.5459 | 2.3917 | 2.3667 | 2 |  | 2.5383 | 2.7 | 3.457 | 2.8867 | 3 | 3 | 2.7833 | 2.5333 | 2.4833 | 2.4683
| Median |  |  | 2.7 | 2 | 2.58 | 2.415 | 2.42 | 2.575 | 2.54 | 2.71 | 2.58 | 3 | 2.615 | 2.4682 | 2.5642 | 2.395 | 2.4 | 2 |  | 2.545 | 2.7 | 3.4536 | 2.88 | 3 | 3 | 2.75 | 2.57 | 2.45 | 2.45
| Std Dev. |  |  | 0.13292 | 0 | 0.033267 | 0.037283 | 0.045019 | 0.047924 | 0.075542 | 0.19416 | 0.036742 | 0 | 0.058907 | 0.037637 | 0.092369 | 0.029269 | 0.05164 | 0 |  | 0.029269 | 0.063246 | 0.093756 | 0.12078 | 0 | 0 | 0.098319 | 0.10211 | 0.19408 | 0.054559
| Rel.Std.Dev. |  |  | 4.953% | 0% | 1.291% | 1.537% | 1.871% | 1.849% | 2.959% | 7.142% | 1.421% | 0% | 2.244% | 1.524% | 3.628% | 1.224% | 2.182% | 0% |  | 1.153% | 2.342% | 2.712% | 4.184% | 0% | 0% | 3.532% | 4.031% | 7.815% | 2.21%
| PDM3 |  |  | 4.528% | -22.09% | 0.3731% | -5.535% | -6.249% | 0.9575% | -0.5358% | 5.892% | 0.6978% | 16.86% | 2.256% | -3.789% | -0.8241% | -6.833% | -7.807% | -22.09% |  | -1.12% | 5.178% | 34.67% | 12.45% | 16.86% | 16.86% | 8.424% | -1.315% | -3.263% | -3.847%
| Z-Score (Absolute) |  | 0.4 | 1.19 | NA | 0.27 | 1.05 | 1.21 | 0.4 | 0.06 | 1.5 | 0.34 | NA | 0.69 | 0.66 | 0 | 1.34 | 1.56 | NA | 0.4 | 0.07 | 1.34 | 7.92 | 2.96 | NA | NA | 2.06 | 0.11 | 0.54 | 0.67
Results from laboratories 1.03, 1.11, 1.18, 1.24 and 1.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 625 (Certified Value 8.61 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 10 | 8.8 | 9.3 | 8.55 | 9.04 | 9.1 | 8.25 | 7.53 | 8.93 | 9.37 | 9 | 8.74 | 8.9145 | 4.931 | 8.2581 | 5.03 | 9.75 | 8.3 | < 5 | 8.4 | 8.9 | 7.99 | 11 | 8.66 | 8.16 | 8.65 | 8 | 7.81 |  | 1
| 1 | 2 | 10 | 8.63 | 8.7 | 8.27 | 8.12 | 8.23 | 8.04 | 7.83 | 8.84 | 9.27 | 8.9 | 9.08 | 8.6162 | 5.064 | 8.5177 | < 5 | 9.61 | 8.2 | < 5 | 8.57 | 9.02 | 8.28 | 10 | 8.75 | 7.92 | 8.55 | 6 | 7.59 |  | 35
| 1 | 3 | 9 | 8.82 | 8.7 | 8.3 | 9.31 | 8.45 | 8.1 | 8.13 | 8.78 | 9.36 | 8.8 | 9.25 | 8.6595 | 3.85 | 8.3276 | < 5 | 9.5 | 8.2 | < 5 | 8.54 | 8.74 | 8.21 | 10 | 9.18 | 8.52 | 8.63 | 8 | 7.8 |  | 16
| 1 | 4 | 11 | 8.74 | 9.3 | 8.24 | 9.12 | 8.67 | 8.26 | 7.87 | 8.69 | 9.32 | 8.9 | 9.11 | 8.9696 | 4.542 | 8.1232 | < 5 | 9.48 | 8.2 | < 5 | 8.43 | 8.84 | 8.32 | 10 | 9.41 | 8.14 | 8.66 | 8 | 7.84 |  | 8.6125
| 1 | 5 | 9 | 8.81 | 8.2 | 8.26 | 8.94 | 8.28 | 8.19 | 8.56 | 9.09 | 9.21 | 8.8 | 8.86 | 8.7004 | 5.503 | 8.5788 | 5.62 | 9.54 | 8.3 | < 5 | 8.65 | 8.88 | 8.76 | 10 | 9.07 | 7.93 | 8.63 | 11 | 7.81 |  | 15
| 1 | 6 | 8 | 8.61 | 8.7 | 8.16 | 9.23 | 8.3 | 8.28 | 7.79 | 8.53 | 9.33 | 8.7 | 9.11 | 9.0158 | 5.986 | 8.1944 | 6.3 | 9.61 | 8.4 | < 5 | 8.4 | 8.68 | 8.25 | 10 | 9.28 | 8.01 | 8.69 | 11 | 7.95
| Mean |  | 9.5 | 8.735 | 8.8167 | 8.2967 | 8.96 | 8.505 | 8.1867 | 7.9517 | 8.81 | 9.31 | 8.85 | 9.025 | 8.8127 | 4.9793 | 8.3333 | 5.65 | 9.5817 | 8.2667 |  | 8.4983 | 8.8433 | 8.3017 | 10.1667 | 9.0583 | 8.1133 | 8.635 | 8.6667 | 7.8
| Median |  | 9.5 | 8.77 | 8.7 | 8.265 | 9.08 | 8.375 | 8.22 | 7.85 | 8.81 | 9.325 | 8.85 | 9.095 | 8.8075 | 4.9975 | 8.2929 | 5.62 | 9.575 | 8.25 |  | 8.485 | 8.86 | 8.265 | 10 | 9.125 | 8.075 | 8.64 | 8 | 7.81
| Std Dev. |  | 1.0488 | 0.093541 | 0.42151 | 0.13277 | 0.43206 | 0.33243 | 0.097091 | 0.3542 | 0.1936 | 0.060332 | 0.10488 | 0.18791 | 0.17371 | 0.74347 | 0.18078 | 0.63553 | 0.098675 | 0.08165 |  | 0.10381 | 0.12094 | 0.25262 | 0.40825 | 0.29715 | 0.22358 | 0.047223 | 1.9664 | 0.11696
| Rel.Std.Dev. |  | 11.04% | 1.071% | 4.781% | 1.6% | 4.822% | 3.909% | 1.186% | 4.454% | 2.197% | 0.648% | 1.185% | 2.082% | 1.971% | 14.93% | 2.169% | 11.25% | 1.03% | 0.9877% |  | 1.222% | 1.368% | 3.043% | 4.016% | 3.28% | 2.756% | 0.5469% | 22.69% | 1.5%
| PDM3 |  | 10.3% | 1.422% | 2.37% | -3.668% | 4.034% | -1.249% | -4.945% | -7.673% | 2.293% | 8.098% | 2.757% | 4.789% | 2.324% | -42.19% | -3.242% | -34.4% | 11.25% | -4.016% |  | -1.326% | 2.68% | -3.61% | 18.04% | 5.176% | -5.796% | 0.2607% | 0.6284% | -9.434%
| Z-Score (Absolute) |  | NA | 0.46 | 0.62 | 0.42 | 0.91 | 0 | 0.63 | 1.1 | 0.61 | 1.6 | 0.69 | 1.04 | 0.61 | 7.03 | 0.34 | 8.83 | 2.15 | 0.48 | 11.97 | 0.01 | 0.67 | 0.41 | NA | 1.1 | 0.78 | 0.26 | NA | 1.41
Results from laboratories 1.01, 1.24 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 625 (Certified Value 1.14 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 1.08 | 1.2 | 0.98 | 1.18 | 1.15 | 1.12 | 1.16 | 1.18 | 1.07 | 1.1408 | 1.15 | 1.1293 | 1.1169 | 0.79 | 1.1427 | 1.13 | 1.21 | 1.1 | 1.1761 | 1.16 | 1.14 | 1.1394 | 1.166 | 1.27 | 1.22 | 1.12 | 1.14 | 1.223 | 1.09 | 1
| 1 | 2 | 1.1 | 1.16 | 0.98 | 1.21 | 1.14 | 1.1 | 1.12 | 1.15 | 1.04 | 1.1516 | 1.15 | 1.1712 | 1.113 | 1.257 | 1.1401 | 1.13 | 1.25 | 1.11 | 1.1022 | 1.1925 | 1.13 | 1.1256 | 1.173 | 1.3 | 1.22 | 1.11 | 1.13 | 1.238 | 1.06 | 0
| 1 | 3 | 1.1 | 1.17 | 1.11 | 1.2 | 1.15 | 1.1 | 1.12 | 1.17 | 1.06 | 1.1446 | 1.12 | 1.198 | 1.1002 | 1.268 | 1.18 | 1.12 | 1.24 | 1.1 | 1.1554 | 1.1728 | 1.13 | 1.1683 | 1.162 | 1.28 | 1.22 | 1.11 | 1.11 | 1.218 | 1.09 | 16
| 1 | 4 | 1.12 | 1.18 | 1.07 | 1.2 | 1.16 | 1.12 | 1.13 | 1.17 | 1.05 | 1.141 | 1.14 | 1.1849 | 1.0818 | 1.061 | 1.1523 | 1.13 | 1.23 | 1.13 | 1.1687 | 1.178 | 1.17 | 1.118 | 1.167 | 1.27 | 1.22 | 1.1 | 1.13 | 1.219 | 1.09 | 1.1444
| 1 | 5 | 1.1 | 1.2 | 1.13 | 1.18 | 1.15 | 1.12 | 1.12 | 1.26 | 1.07 | 1.1537 | 1.13 | 1.19 | 1.0718 | 1.04 | 1.1708 | 1.12 | 1.25 | 1.17 | 1.1693 | 1.1992 | 1.17 | 1.1437 | 1.171 | 1.29 | 1.23 | 1.09 | 1.1 | 1.22 | 1.11 | 16
| 1 | 6 | 1.1 | 1.11 | 1.05 | 1.2 | 1.17 | 1.11 | 1.15 | 1.16 | 1.01 | 1.1458 | 1.12 | 1.2008 | 1.1059 | 1.012 | 1.1728 | 1.11 | 1.23 | 1.17 | 1.1635 | 1.1871 | 1.14 | 1.1626 | 1.147 | 1.33 | 1.22 | 1.09 | 1.15 | 1.217 | 1.1
| Mean |  | 1.1 | 1.17 | 1.0533 | 1.195 | 1.1533 | 1.1117 | 1.1333 | 1.1817 | 1.05 | 1.1462 | 1.135 | 1.179 | 1.0983 | 1.0713 | 1.1598 | 1.1233 | 1.235 | 1.13 | 1.1559 | 1.1816 | 1.1467 | 1.1429 | 1.1643 | 1.29 | 1.2217 | 1.1033 | 1.1267 | 1.2225 | 1.09
| Median |  | 1.1 | 1.175 | 1.06 | 1.2 | 1.15 | 1.115 | 1.125 | 1.17 | 1.055 | 1.1452 | 1.135 | 1.1874 | 1.103 | 1.0505 | 1.1615 | 1.125 | 1.235 | 1.12 | 1.1661 | 1.1825 | 1.14 | 1.1416 | 1.1665 | 1.285 | 1.22 | 1.105 | 1.13 | 1.2195 | 1.09
| Std Dev. |  | 0.012649 | 0.033466 | 0.063456 | 0.012247 | 0.010328 | 0.0098319 | 0.017512 | 0.039707 | 0.022804 | 0.0053724 | 0.013784 | 0.026544 | 0.017859 | 0.17715 | 0.016938 | 0.008165 | 0.015166 | 0.032863 | 0.027178 | 0.014253 | 0.018619 | 0.019826 | 0.009331 | 0.022804 | 0.0040825 | 0.012111 | 0.018619 | 0.0078677 | 0.016733
| Rel.Std.Dev. |  | 1.15% | 2.86% | 6.024% | 1.025% | 0.8955% | 0.8844% | 1.545% | 3.36% | 2.172% | 0.4687% | 1.214% | 2.251% | 1.626% | 16.54% | 1.461% | 0.7269% | 1.228% | 2.908% | 2.351% | 1.206% | 1.624% | 1.735% | 0.8014% | 1.768% | 0.3342% | 1.098% | 1.653% | 0.6436% | 1.535%
| PDM3 |  | -3.88% | 2.237% | -7.958% | 4.422% | 0.7807% | -2.86% | -0.967% | 3.257% | -8.249% | 0.1617% | -0.8213% | 3.026% | -4.03% | -6.385% | 1.343% | -1.841% | 7.917% | -1.258% | 1.002% | 3.251% | 0.1981% | -0.1281% | 1.742% | 12.72% | 6.752% | -3.588% | -1.55% | 6.825% | -4.754%
| Z-Score (Absolute) |  | 0.9 | 0.46 | 1.81 | 0.95 | 0.14 | 0.67 | 0.25 | 0.69 | 1.88 | 0 | 0.22 | 0.64 | 0.94 | 1.46 | 0.26 | 0.45 | 1.73 | 0.32 | 0.19 | 0.69 | 0.01 | 0.06 | 0.35 | 2.8 | 1.47 | 0.84 | 0.38 | 1.49 | 1.1
| Analytical results for Cd in OREAS 625 (Certified Value 83 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 91 | 85.7 | 76.5 | 84.5 | 81.8 | 78.6 | 82.4 | 85.1 | 84.2 | 87.78 | 84.2 | 81.48 | 84.343 | 78.834 | 84.5624 | 81.91 | 85.6 | 78.6 | 72.71 | 79.71 | 84.29 | 87.18 | 88.91 | 95.5 | 73.4 | 82.6 | 83.42 | 88 | 80 | 1
| 1 | 2 | 93 | 84.1 | 70.9 | 86.6 | 75.4 | 76 | 80.8 | 87.4 | 83.5 | 86.63 | 83.3 | 88.07 | 82.2151 | 80.621 | 88.2466 | 81.6 | 85.7 | 79.3 | 69.93 | 81.31 | 82 | 87.68 | 84.27 | 91.9 | 72.3 | 79.5 | 82.08 | 88 | 80 | 0
| 1 | 3 | 92 | 83.5 | 72.3 | 87.2 | 80.4 | 73.7 | 82.3 | 90 | 84.3 | 89.39 | 83.7 | 86.63 | 82.2803 | 77.765 | 87.0301 | 82.39 | 86.8 | 78 | 71.56 | 80.52 | 82.67 | 89.33 | 85.58 | 88.8 | 75.2 | 82.4 | 82.93 | 87 | 80 | 16
| 1 | 4 | 96 | 84.6 | 73.8 | 88 | 81.6 | 79.9 | 82.2 | 87.7 | 82.8 | 87.76 | 84.6 | 85.59 | 81.836 | 79.618 | 84.9558 | 80.11 | 86.5 | 77.7 | 71.77 | 79.33 | 82.78 | 87.05 | 87.16 | 91.9 | 76.6 | 81.3 | 82.85 | 88 | 80 | 82.9767
| 1 | 5 | 92 | 86.1 | 71.8 | 85.8 | 79.6 | 78.3 | 82.1 | 93.5 | 84.8 | 87.72 | 81.8 | 83.49 | 80.9899 | 79.407 | 86.3738 | 81.17 | 88 | 77.9 | 71.89 | 82.78 | 84.73 | 86.87 | 86.68 | 93.3 | 71.6 | 81 | 83.31 | 88 | 80 | 17
| 1 | 6 | 95 | 83.1 | 70.7 | 87.6 | 82.3 | 75.2 | 85.6 | 85.5 | 82.4 | 86.13 | 82.1 | 86.21 | 83.0222 | 79.018 | 87.2769 | 81.31 | 86.1 | 78.4 | 71.11 | 81.4 | 82.04 | 88.64 | 83.54 | 89.5 | 75.6 | 80.1 | 84.07 | 87 | 80
| Mean |  | 93.1667 | 84.5167 | 72.6667 | 86.6167 | 80.1833 | 76.95 | 82.5667 | 88.2 | 83.6667 | 87.5683 | 83.2833 | 85.245 | 82.4477 | 79.2105 | 86.4076 | 81.415 | 86.45 | 78.3167 | 71.495 | 80.8417 | 83.085 | 87.7917 | 86.0233 | 91.8167 | 74.1167 | 81.15 | 83.11 | 87.6667 | 80
| Median |  | 92.5 | 84.35 | 72.05 | 86.9 | 81 | 77.15 | 82.25 | 87.55 | 83.85 | 87.74 | 83.5 | 85.9 | 82.2477 | 79.2125 | 86.7019 | 81.455 | 86.3 | 78.2 | 71.665 | 80.915 | 82.725 | 87.43 | 86.13 | 91.9 | 74.3 | 81.15 | 83.12 | 88 | 80
| Std Dev. |  | 1.9408 | 1.194 | 2.1842 | 1.2937 | 2.5443 | 2.3569 | 1.5983 | 3.1343 | 0.9331 | 1.1289 | 1.1268 | 2.3751 | 1.1405 | 0.94472 | 1.4169 | 0.77511 | 0.88713 | 0.58452 | 0.92839 | 1.2615 | 1.1569 | 0.98715 | 1.9745 | 2.4596 | 1.9844 | 1.2276 | 0.66612 | 0.5164 | 0
| Rel.Std.Dev. |  | 2.083% | 1.413% | 3.006% | 1.494% | 3.173% | 3.063% | 1.936% | 3.554% | 1.115% | 1.289% | 1.353% | 2.786% | 1.383% | 1.193% | 1.64% | 0.952% | 1.026% | 0.7464% | 1.299% | 1.561% | 1.392% | 1.124% | 2.295% | 2.679% | 2.677% | 1.513% | 0.8015% | 0.589% | 0%
| PDM3 |  | 12.28% | 1.856% | -12.43% | 4.387% | -3.366% | -7.263% | -0.4941% | 6.295% | 0.8315% | 5.534% | 0.3695% | 2.734% | -0.6375% | -4.539% | 4.135% | -1.882% | 4.186% | -5.616% | -13.84% | -2.573% | 0.1305% | 5.803% | 3.672% | 10.65% | -10.68% | -2.201% | 0.1607% | 5.652% | -3.587%
| Z-Score (Absolute) |  | 2.06 | 0.29 | 2.14 | 0.72 | 0.6 | 1.26 | 0.11 | 1.04 | 0.11 | 0.91 | 0.04 | 0.44 | 0.14 | 0.8 | 0.67 | 0.35 | 0.68 | 0.98 | 2.38 | 0.46 | 0.01 | 0.96 | 0.6 | 1.78 | 1.84 | 0.4 | 0 | 0.93 | 0.64
| Analytical results for Ce in OREAS 625 (Certified Value 72 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 0
0
| 1 | 1 | 63.9 | 71 | 75.6 | 62.9 | 67.6 | 71.9 | 71.1 | 82.6 | 86.5 | 64 | 83 | 71.7259 | 82 | 48.43 | 59 | 76 | 83.46 | 78.2 | 62.36 | 74.31 | 81 | [8 empty] | 1
| 1 | 2 | 66.3 | 74 | 81.4 | 59.9 | 69.9 | 71.4 | 74.9 | 81.8 | 82.76 | 59 | 80 | 69.887 | 80 | 47.84 | 61.1 | 74.6 | 81.88 | 79.9 | 52.77 | 78.31 | 82 | [8 empty] | 36
| 1 | 3 | 66.9 | 76 | 79.5 | 62.9 | 54 | 71.1 | 73 | 81.9 | 86.74 | 66 | 79 | 72.1975 | 80 | 48.76 | 59.7 | 73.2 | 81.17 | 94.5 | 65.13 | 73.49 | 82 | [8 empty] | 16
| 1 | 4 | 65.1 | 75 | 83.2 | 70.3 | 65.6 | 74 | 70.6 | 81.8 | 85.15 | 65 | 79 | 70.1658 | 83 | 48.62 | 60.5 | 75.8 | 77.97 | 86.6 | 58.97 | 72.22 | 79 | [8 empty] | 72.4712
| 1 | 5 | 67.9 | 77 | 75.7 | 71.6 | 61.8 | 71.6 | 76 | 83.5 | 83.34 | 63 | 76 | 68.7153 | 80 | 48.53 | 60.9 | 76.4 | 81.66 | 84.4 | 58.51 | 78.64 | 81 | [8 empty] | 18
| 1 | 6 | 55.1 | 74 | 80 | 70 | 59.5 | 75.7 | 76.4 | 80 | 82.84 | 61 | 77 | 71.162 | 81 | 48.72 | 60 | 74.7 | 82.14 | 84.3 | 62.42 | 77.86 | 79
| Mean |  | 64.2 | 74.5 | 79.2333 | 66.2667 | 63.0667 | 72.6167 | 73.6667 | 81.9333 | 84.555 | 63 | 79 | 70.6422 | 81 | 48.4833 | 60.2 | 75.1167 | 81.38 | 84.65 | 60.0267 | 75.805 | 80.6667
| Median |  | 65.7 | 74.5 | 79.75 | 66.45 | 63.7 | 71.75 | 73.95 | 81.85 | 84.245 | 63.5 | 79 | 70.6639 | 80.5 | 48.575 | 60.25 | 75.25 | 81.77 | 84.35 | 60.665 | 76.085 | 81
| Std Dev. |  | 4.6712 | 2.0736 | 3.0585 | 4.9367 | 5.8343 | 1.8302 | 2.4849 | 1.1553 | 1.8194 | 2.6077 | 2.4495 | 1.2941 | 1.2649 | 0.33768 | 0.78994 | 1.1839 | 1.8387 | 5.7476 | 4.3195 | 2.7922 | 1.3663
| Rel.Std.Dev. |  | 7.276% | 2.783% | 3.86% | 7.45% | 9.251% | 2.52% | 3.373% | 1.41% | 2.152% | 4.139% | 3.101% | 1.832% | 1.562% | 0.6965% | 1.312% | 1.576% | 2.259% | 6.79% | 7.196% | 3.683% | 1.694%
| PDM3 |  | -11.41% | 2.8% | 9.331% | -8.561% | -12.98% | 0.2008% | 1.65% | 13.06% | 16.67% | -13.07% | 9.009% | -2.524% | 11.77% | -33.1% | -16.93% | 3.65% | 12.29% | 16.81% | -17.17% | 4.6% | 11.31%
| Z-Score (Absolute) |  | 1.01 | 0 | 0.46 | 0.81 | 1.12 | 0.18 | 0.08 | 0.73 | 0.99 | 1.13 | 0.44 | 0.38 | 0.64 | 2.55 | 1.4 | 0.06 | 0.67 | 0.99 | 1.42 | 0.13 | 0.6
| Analytical results for Co in OREAS 625 (Certified Value 4.21 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 4 | 4.4 | 4.1 | 4.3 | 4.8 | 4.2 | 4 | 4.4 | 4.1 | 4.1 | 4.5 | 4.3 | 4.0691 | 4.336 | 3.925 | 8.94 | 4 | 4.1 | < 5 | 4 | 4.22 | 6.6 | 4.4 | 6 | 4.3 | 4.4 | 4.1 | 4 | 3.8 | 1
| 1 | 2 | 4 | 4.5 | 3.9 | 4.3 | 4.1 | 4.1 | 4 | 4.3 | 4.15 | 4 | 4.5 | 4.5 | 3.8686 | 4.814 | 4.1629 | 9.6 | 4 | 3.8 | < 5 | 4.1 | 4.13 | 6 | 4.4 | 6 | 4.3 | 4.2 | 4.3 | 4 | 3.6 | 37
| 1 | 3 | 4 | 4.4 | 4 | 4.5 | 4.7 | 4 | 4.1 | 4.2 | 4.19 | 4.1 | 4.3 | 4.5 | 3.8931 | 4.185 | 4.212 | 9 | 4 | 3.8 | < 5 | 4.3 | 4.16 | 6.58 | 4.3 | 6 | 4.4 | 4.4 | 4.2 | 4 | 3.7 | 16
| 1 | 4 | 4 | 4.4 | 4 | 4.4 | 4.7 | 4.4 | 4.1 | 3.9 | 4.19 | 4.2 | 4.5 | 4.6 | 3.8249 | 4.116 | 4.1232 | 9.02 | 4 | 3.9 | < 5 | 4.6 | 4.36 | 6.27 | 4.3 | 6 | 4.4 | 4.3 | 4.3 | 4 | 3.8 | 4.2137
| 1 | 5 | 4 | 4.3 | 3.9 | 4.2 | 4.5 | 4.2 | 4.1 | 4.5 | 4.22 | 4 | 4.4 | 4.5 | 3.7492 | 4.215 | 3.9817 | 9.39 | 4 | 4 | < 5 | 4.5 | 4.3 | 6.63 | 4.5 | 6 | 4.4 | 4 | 4.2 | 4 | 3.8 | 19
| 1 | 6 | 4 | 4.1 | 3.9 | 4.3 | 4.6 | 4 | 4.3 | 4.5 | 4.07 | 4 | 4.4 | 4.5 | 3.8679 | 4.228 | 4.2753 | 9.26 | 4 | 4.2 | < 5 | 4 | 4.14 | 5.92 | 4.6 | 6 | 4.5 | 4.3 | 4.2 | 4 | 3.9
| Mean |  | 4 | 4.35 | 3.9667 | 4.3333 | 4.5667 | 4.15 | 4.1 | 4.3 | 4.1533 | 4.0667 | 4.4333 | 4.4833 | 3.8788 | 4.3157 | 4.1133 | 9.2017 | 4 | 3.9667 |  | 4.25 | 4.2183 | 6.3333 | 4.4167 | 6 | 4.3833 | 4.2667 | 4.2167 | 4 | 3.7667
| Median |  | 4 | 4.4 | 3.95 | 4.3 | 4.65 | 4.15 | 4.1 | 4.35 | 4.17 | 4.05 | 4.45 | 4.5 | 3.8682 | 4.2215 | 4.143 | 9.14 | 4 | 3.95 |  | 4.2 | 4.19 | 6.425 | 4.4 | 6 | 4.4 | 4.3 | 4.2 | 4 | 3.8
| Std Dev. |  | 0 | 0.13784 | 0.08165 | 0.10328 | 0.25033 | 0.15166 | 0.10954 | 0.22804 | 0.058195 | 0.08165 | 0.08165 | 0.098319 | 0.10615 | 0.25437 | 0.13515 | 0.26065 | 0 | 0.1633 |  | 0.25884 | 0.093897 | 0.3181 | 0.1169 | 0 | 0.075277 | 0.15055 | 0.075277 | 0 | 0.10328
| Rel.Std.Dev. |  | 0% | 3.169% | 2.058% | 2.383% | 5.482% | 3.654% | 2.672% | 5.303% | 1.401% | 2.008% | 1.842% | 2.193% | 2.737% | 5.894% | 3.286% | 2.833% | 0% | 4.117% |  | 6.09% | 2.226% | 5.023% | 2.647% | 0% | 1.717% | 3.529% | 1.785% | 0% | 2.742%
| PDM3 |  | -5.072% | 3.234% | -5.863% | 2.839% | 8.376% | -1.512% | -2.699% | 2.048% | -1.433% | -3.49% | 5.212% | 6.399% | -7.948% | 2.419% | -2.382% | 118.4% | -5.072% | -5.863% |  | 0.8611% | 0.1096% | 50.3% | 4.816% | 42.39% | 4.025% | 1.257% | 0.07003% | -5.072% | -10.61%
| Z-Score (Absolute) |  | NA | 0.45 | 1.27 | 0.37 | 1.42 | 0.45 | 0.67 | 0.22 | 0.43 | 0.82 | 0.82 | 1.05 | 1.67 | 0.3 | 0.61 | 22.26 | NA | 1.27 | 7.87 | 0 | 0.14 | 9.37 | 0.75 | NA | 0.6 | 0.07 | 0.15 | NA | 2.17
Results from laboratories 1.01, 1.17, 1.24 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 625 (Certified Value 13.3 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 15 | 15 | 12 | 12 | 13 | 15 | 11 | 13 | 13.8 | 17 | 10 | 14 | 12.8373 | 20.987 | 18.7138 | 15.97 | 14 | 13 | 15.24 | 13.5 | 12 | 12 | 14 | 14 | 13 | 11 | 14 |  |  | 1
| 1 | 2 | 14 | 16 | 9 | 13 | 14 | 14 | 11 | 13 | 14.2 | 14 | 8 | 15 | 12.9137 | 20.55 | 19.1526 | 20.76 | 14 | 14 | 14.94 | 14.3 | 12 | 10 | 14 | 14 | 11 | 11 | 14 |  |  | 15
| 1 | 3 | 14 | 16 | 11 | 15 | 13 | 13 | 11 | 13 | 15 | 16 | 8 | 13 | 13.0475 | 21.815 | 25.359 | 14.46 | 14 | 14 | 15.16 | 14 | 11 | 9 | 14 | 14 | 13 | 11 | 13 |  |  | 16
| 1 | 4 | 14 | 16 | 12 | 12 | 14 | 13 | 11 | 12 | 13.2 | 15 | 8 | 14 | 12.2234 | 22.842 | 18.3151 | 23.89 | 14 | 13 | 15.14 | 13.9 | 12 | 10 | 15 | 13 | 11 | 11 | 13 |  |  | 13.276
| 1 | 5 | 15 | 15 | 10 | 12 | 12 | 15 | 11 | 14 | 13.4 | 15 | 8 | 13 | 12.6594 | 22.657 | 22.4435 | 15.15 | 14 | 17 | 15.15 | 13.9 | 12 | 12 | 15 | 14 | 11 | 11 | 13 |  |  | 20
| 1 | 6 | 14 | 13 | 11 | 12 | 14 | 13 | 11 | 12 | 13.8 | 15 | 7 | 12 | 13.3823 | 22.361 | 24.5736 | 16.24 | 14 | 15 | 15.12 | 14 | 12 | 10 | 13 | 14 | 11 | 12 | 14
| Mean |  | 14.3333 | 15.1667 | 10.8333 | 12.6667 | 13.3333 | 13.8333 | 11 | 12.8333 | 13.9 | 15.3333 | 8.1667 | 13.5 | 12.8439 | 21.8687 | 21.4263 | 17.745 | 14 | 14.3333 | 15.125 | 13.9333 | 11.8333 | 10.5 | 14.1667 | 13.8333 | 11.6667 | 11.1667 | 13.5
| Median |  | 14 | 15.5 | 11 | 12 | 13.5 | 13.5 | 11 | 13 | 13.8 | 15 | 8 | 13.5 | 12.8755 | 22.088 | 20.7981 | 16.105 | 14 | 14 | 15.145 | 13.95 | 12 | 10 | 14 | 14 | 11 | 11 | 13.5
| Std Dev. |  | 0.5164 | 1.169 | 1.169 | 1.2111 | 0.8165 | 0.98319 | 0 | 0.75277 | 0.64187 | 1.0328 | 0.98319 | 1.0488 | 0.38869 | 0.93063 | 3.1181 | 3.7361 | 0 | 1.5055 | 0.099549 | 0.2582 | 0.40825 | 1.2247 | 0.75277 | 0.40825 | 1.0328 | 0.40825 | 0.54772
| Rel.Std.Dev. |  | 3.603% | 7.708% | 10.79% | 9.561% | 6.124% | 7.107% | 0% | 5.866% | 4.618% | 6.736% | 12.04% | 7.769% | 3.026% | 4.256% | 14.55% | 21.05% | 0% | 10.5% | 0.6582% | 1.853% | 3.45% | 11.66% | 5.314% | 2.951% | 8.853% | 3.656% | 4.057%
| PDM3 |  | 7.964% | 14.24% | -18.4% | -4.59% | 0.4319% | 4.198% | -17.14% | -3.334% | 4.7% | 15.5% | -38.49% | 1.687% | -3.254% | 64.72% | 61.39% | 33.66% | 5.453% | 7.964% | 13.93% | 4.951% | -10.87% | -20.91% | 6.709% | 4.198% | -12.12% | -15.89% | 1.687%
| Z-Score (Absolute) |  | 0.29 | 0.77 | 1.73 | 0.67 | 0.29 | 0 | 1.64 | 0.58 | 0.04 | 0.87 | 3.28 | 0.19 | 0.57 | 4.64 | 4.39 | 2.26 | 0.1 | 0.29 | 0.75 | 0.06 | 1.16 | 1.93 | 0.19 | 0 | 1.25 | 1.54 | 0.19
| Analytical results for Cs in OREAS 625 (Certified Value 5.94 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 2
3
| 1 | 1 | 6.02 | 5.78 | 6.17 | 5.91 | 5.77 | 5.81 | 5.79 | 6.03 | 6 | 6.1918 | 6.1386 | 6.1 | 5.6 | 6.3 | 5.78 | 7.1614 | 6.04 | 5 | 5.77 | 5.96 | 5.99 | [8 empty] | 1
| 1 | 2 | 5.94 | 5.84 | 5.41 | 5.71 | 5.63 | 6.01 | 5.78 | 5.86 | 6.29 | 5.936 | 6.1036 | 5.9 | 5.5 | 6.4 | 5.66 | 7.3712 | 6.06 | 5 | 5.06 | 5.9 | 5.83 | [8 empty] | 39
| 1 | 3 | 5.86 | 5.92 | 6.11 | 5.5 | 5.8 | 6.18 | 5.93 | 6.14 | 6.35 | 5.9695 | 5.8107 | 6 | 5.2 | 6.5 | 5.7 | 7.1204 | 6.01 | 6 | 6.03 | 5.96 | 5.98 | [8 empty] | 16
| 1 | 4 | 6 | 6.08 | 6.19 | 5.81 | 5.78 | 6.04 | 5.79 | 5.92 | 6.2 | 5.9375 | 6.1442 | 6 | 5.6 | 6.3 | 5.71 | 7.1167 | 6.11 | 6 | 5.68 | 5.88 | 6.06 | [8 empty] | 5.9425
| 1 | 5 | 6.01 | 5.77 | 6.04 | 6.07 | 5.76 | 6.33 | 6.01 | 5.94 | 6.18 | 5.6951 | 5.783 | 6 | 5.9 | 6.3 | 5.89 | 7.6065 | 6.03 | 5 | 5.59 | 5.92 | 6.13 | [8 empty] | 21
| 1 | 6 | 5.72 | 5.9 | 6.27 | 5.51 | 6.12 | 5.78 | 5.66 | 5.96 | 6.19 | 5.983 | 5.8923 | 6 | 5.8 | 6.3 | 5.63 | 7.4878 | 6.12 | 6 | 5.51 | 5.96 | 6.13
| Mean |  | 5.925 | 5.8817 | 6.0317 | 5.7517 | 5.81 | 6.025 | 5.8267 | 5.975 | 6.2017 | 5.9521 | 5.9787 | 6 | 5.6 | 6.35 | 5.7283 | 7.3107 | 6.0617 | 5.5 | 5.6067 | 5.93 | 6.02
| Median |  | 5.97 | 5.87 | 6.14 | 5.76 | 5.775 | 6.025 | 5.79 | 5.95 | 6.195 | 5.9535 | 5.9979 | 6 | 5.6 | 6.3 | 5.705 | 7.2663 | 6.05 | 5.5 | 5.635 | 5.94 | 6.025
| Std Dev. |  | 0.11692 | 0.11462 | 0.31422 | 0.22507 | 0.16346 | 0.21154 | 0.12404 | 0.097929 | 0.11923 | 0.15829 | 0.16884 | 0.063246 | 0.24495 | 0.083666 | 0.09411 | 0.20913 | 0.04446 | 0.54772 | 0.32241 | 0.035214 | 0.11349
| Rel.Std.Dev. |  | 1.973% | 1.949% | 5.21% | 3.913% | 2.813% | 3.511% | 2.129% | 1.639% | 1.923% | 2.659% | 2.824% | 1.054% | 4.374% | 1.318% | 1.643% | 2.861% | 0.7335% | 9.959% | 5.75% | 0.5938% | 1.885%
| PDM3 |  | -0.2945% | -1.024% | 1.5% | -3.211% | -2.23% | 1.388% | -1.949% | 0.5469% | 4.361% | 0.1623% | 0.6095% | 0.9676% | -5.764% | 6.857% | -3.604% | 23.02% | 2.005% | -7.446% | -5.651% | -0.2104% | 1.304%
| Z-Score (Absolute) |  | 0.29 | 0.61 | 0.51 | 1.59 | 1.15 | 0.46 | 1.03 | 0.09 | 1.78 | 0.09 | 0.11 | 0.27 | 2.72 | 2.9 | 1.76 | 10.09 | 0.73 | NA | 2.67 | 0.25 | 0.42
Results from laboratory 1.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 625 (Certified Value 0.171 wt.%) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.1763 | 0.183 | 0.17673 | 0.1725 | 0.169 | 0.1625 | 0.1695 | 0.1835 | 0.171 | 0.16401 | 0.17429 | 0.15437 | 0.17009 | 0.16924 | 0.16977 | 0.1944 | 0.172 | 0.16208 | 0.1762 | 0.17071 | 0.1756 | 0.19001 | 0.1683 | 0.1735 | 0.1625 | 0.1783 | 0.17266 | 0.1627 | 0.169 | 1
| 1 | 2 | 0.1782 | 0.183 | 0.16987 | 0.1745 | 0.1665 | 0.1625 | 0.1645 | 0.183 | 0.167 | 0.16334 | 0.17579 | 0.1658 | 0.16734 | 0.1758 | 0.16951 | 0.1944 | 0.172 | 0.15963 | 0.1767 | 0.17495 | 0.1774 | 0.18759 | 0.1696 | 0.1747 | 0.1614 | 0.1788 | 0.17436 | 0.1635 | 0.169 | 25
| 1 | 3 | 0.181 | 0.179 | 0.18225 | 0.176 | 0.168 | 0.16 | 0.1645 | 0.187 | 0.174 | 0.16334 | 0.17016 | 0.162 | 0.16715 | 0.17273 | 0.17073 | 0.1953 | 0.176 | 0.1603 | 0.1758 | 0.17331 | 0.1765 | 0.19389 | 0.1677 | 0.1783 | 0.1636 | 0.1788 | 0.17282 | 0.1628 | 0.169 | 16
| 1 | 4 | 0.1785 | 0.177 | 0.18132 | 0.1735 | 0.167 | 0.1645 | 0.167 | 0.187 | 0.177 | 0.16337 | 0.17386 | 0.16309 | 0.16292 | 0.17112 | 0.1704 | 0.1952 | 0.172 | 0.16134 | 0.1768 | 0.19015 | 0.1758 | 0.19337 | 0.1695 | 0.1705 | 0.1687 | 0.1792 | 0.16893 | 0.1639 | 0.169 | 0.171
| 1 | 5 | 0.1734 | 0.184 | 0.17685 | 0.176 | 0.1675 | 0.165 | 0.164 | 0.1995 | 0.17 | 0.16121 | 0.17143 | 0.16027 | 0.16148 | 0.17145 | 0.17186 | 0.1942 | 0.169 | 0.16159 | 0.1786 | 0.17734 | 0.1777 | 0.19145 | 0.1682 | 0.1764 | 0.163 | 0.1726 | 0.17417 | 0.163 | 0.169 | 22
| 1 | 6 | 0.1761 | 0.176 | 0.1766 | 0.1745 | 0.171 | 0.1635 | 0.1665 | 0.1785 | 0.169 | 0.16429 | 0.17272 | 0.16109 | 0.16639 | 0.17003 | 0.17208 | 0.1933 | 0.166 | 0.16375 | 0.1764 | 0.17563 | 0.1761 | 0.19323 | 0.1669 | 0.1752 | 0.1667 | 0.1743 | 0.17446 | 0.1624 | 0.168
| Mean |  | 0.17725 | 0.18033 | 0.17727 | 0.1745 | 0.16817 | 0.163 | 0.166 | 0.18642 | 0.17133 | 0.16326 | 0.17304 | 0.1611 | 0.16589 | 0.17173 | 0.17072 | 0.19447 | 0.17117 | 0.16145 | 0.17675 | 0.17702 | 0.17652 | 0.19159 | 0.16837 | 0.17477 | 0.16432 | 0.177 | 0.1729 | 0.16305 | 0.16883
| Median |  | 0.17725 | 0.181 | 0.17679 | 0.1745 | 0.16775 | 0.163 | 0.1655 | 0.18525 | 0.1705 | 0.16336 | 0.17329 | 0.16154 | 0.16677 | 0.17128 | 0.17056 | 0.1944 | 0.172 | 0.16146 | 0.17655 | 0.17529 | 0.1763 | 0.19234 | 0.16825 | 0.17495 | 0.1633 | 0.17855 | 0.1735 | 0.1629 | 0.169
| Std Dev. |  | 0.0025913 | 0.0034448 | 0.0044018 | 0.0013784 | 0.001633 | 0.0017889 | 0.0020976 | 0.0071373 | 0.0036148 | 0.0010817 | 0.0020389 | 0.0038169 | 0.0031576 | 0.0023301 | 0.0010623 | 0.00073121 | 0.0033714 | 0.001439 | 0.00097519 | 0.0068161 | 0.0008612 | 0.0024373 | 0.0010424 | 0.0026515 | 0.002791 | 0.0028164 | 0.0020978 | 0.00055408 | 0.00040825
| Rel.Std.Dev. |  | 1.462% | 1.91% | 2.483% | 0.7899% | 0.9711% | 1.097% | 1.264% | 3.829% | 2.11% | 0.6626% | 1.178% | 2.369% | 1.903% | 1.357% | 0.6222% | 0.376% | 1.97% | 0.8913% | 0.5517% | 3.851% | 0.4879% | 1.272% | 0.6191% | 1.517% | 1.699% | 1.591% | 1.213% | 0.3398% | 0.2418%
| PDM3 |  | 3.657% | 5.46% | 3.669% | 2.049% | -1.655% | -4.676% | -2.922% | 9.018% | 0.197% | -4.524% | 1.196% | -5.786% | -2.984% | 0.4281% | -0.1603% | 13.73% | 0.09949% | -5.584% | 3.365% | 3.52% | 3.228% | 12.04% | -1.538% | 2.205% | -3.906% | 3.511% | 1.113% | -4.647% | -1.265%
| Z-Score (Absolute) |  | 0.7 | 1.1 | 0.71 | 0.35 | 0.45 | 1.11 | 0.73 | 1.87 | 0.05 | 1.08 | 0.17 | 1.36 | 0.74 | 0 | 0.13 | 2.9 | 0.07 | 1.31 | 0.64 | 0.67 | 0.61 | 2.53 | 0.43 | 0.39 | 0.95 | 0.67 | 0.15 | 1.11 | 0.37
| Analytical results for Dy in OREAS 625 (Certified Value 3.52 ppm) | [30 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | 2
3
| 1 | 1 | 1.01: 3.5; 1.02: 3.61; 1.03: 3.59; 1.04: 3.399; 1.05: 4.5862; 1.06: 3.5 | 1
| 1 | 2 | 1.01: 3.4; 1.02: 3.49; 1.03: 3.76; 1.04: 3.2605; 1.05: 4.2478; 1.06: 3.58 | 7
| 1 | 3 | 1.01: 3.5; 1.02: 3.56; 1.03: 3.86; 1.04: 3.2845; 1.05: 4.1817; 1.06: 3.53 | 16
| 1 | 4 | 1.01: 3.5; 1.02: 3.55; 1.03: 3.82; 1.04: 3.2949; 1.05: 4.6938; 1.06: 3.46 | 3.5182
| 1 | 5 | 1.01: 3.5; 1.02: 3.5; 1.03: 3.8; 1.04: 3.2752; 1.05: 4.6159; 1.06: 3.55 | 23
| 1 | 6 | 1.01: 3.2; 1.02: 3.47; 1.03: 3.88; 1.04: 3.3832; 1.05: 4.5627; 1.06: 3.54
| Mean |  | 1.01: 3.4333; 1.02: 3.53; 1.03: 3.785; 1.04: 3.3162; 1.05: 4.4814; 1.06: 3.5267
| Median |  | 1.01: 3.5; 1.02: 3.525; 1.03: 3.81; 1.04: 3.2897; 1.05: 4.5745; 1.06: 3.535
| Std Dev. |  | 1.01: 0.12111; 1.02: 0.052536; 1.03: 0.10464; 1.04: 0.059298; 1.05: 0.21222; 1.06: 0.041793
| Rel.Std.Dev. |  | 1.01: 3.527%; 1.02: 1.488%; 1.03: 2.765%; 1.04: 1.788%; 1.05: 4.736%; 1.06: 1.185%
| PDM3 |  | 1.01: -2.413%; 1.02: 0.3342%; 1.03: 7.582%; 1.04: -5.742%; 1.05: 27.37%; 1.06: 0.2394%
| Z-Score (Absolute) |  | 1.01: 0.42; 1.02: 0.01; 1.03: 1.13; 1.04: 0.93; 1.05: 4.18; 1.06: 0.01
| Analytical results for Er in OREAS 625 (Certified Value 1.24 ppm) | [30 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | 2
3
| 1 | 1 | 1.01: 1.2; 1.02: 1.25; 1.03: 1.27; 1.04: 1.2532; 1.05: 1.788; 1.06: 1.19 | 1
| 1 | 2 | 1.01: 1.2; 1.02: 1.32; 1.03: 1.31; 1.04: 1.1767; 1.05: 1.5969; 1.06: 1.17 | 8
| 1 | 3 | 1.01: 1.3; 1.02: 1.25; 1.03: 1.31; 1.04: 1.2201; 1.05: 1.5854; 1.06: 1.19 | 16
| 1 | 4 | 1.01: 1.2; 1.02: 1.25; 1.03: 1.31; 1.04: 1.2398; 1.05: 1.8209; 1.06: 1.21 | 1.2423
| 1 | 5 | 1.01: 1.3; 1.02: 1.19; 1.03: 1.29; 1.04: 1.2072; 1.05: 1.6438; 1.06: 1.22 | 24
| 1 | 6 | 1.01: 1.2; 1.02: 1.27; 1.03: 1.33; 1.04: 1.2461; 1.05: 1.7558; 1.06: 1.35
| Mean |  | 1.01: 1.2333; 1.02: 1.255; 1.03: 1.3033; 1.04: 1.2239; 1.05: 1.6985; 1.06: 1.2217
| Median |  | 1.01: 1.2; 1.02: 1.25; 1.03: 1.31; 1.04: 1.23; 1.05: 1.6998; 1.06: 1.2
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.041833; 1.03: 0.020656; 1.04: 0.028704; 1.05: 0.10236; 1.06: 0.065243
| Rel.Std.Dev. |  | 1.01: 4.187%; 1.02: 3.333%; 1.03: 1.585%; 1.04: 2.345%; 1.05: 6.026%; 1.06: 5.341%
| PDM3 |  | 1.01: -0.7222%; 1.02: 1.022%; 1.03: 4.913%; 1.04: -1.485%; 1.05: 36.72%; 1.06: -1.661%
| Z-Score (Absolute) |  | 1.01: 0.34; 1.02: 0.34; 1.03: 1.86; 1.04: 0.64; 1.05: 14.31; 1.06: 0.71
| Analytical results for Eu in OREAS 625 (Certified Value 1.36 ppm) | [30 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS | 2
3
| 1 | 1 | 1.01: 1.25; 1.02: 1.4; 1.03: 1.39; 1.04: 1.2272; 1.05: 1.48 | 1
| 1 | 2 | 1.01: 1.32; 1.02: 1.36; 1.03: 1.47; 1.04: 1.1642; 1.05: 1.47 | 9
| 1 | 3 | 1.01: 1.32; 1.02: 1.43; 1.03: 1.5; 1.04: 1.2252; 1.05: 1.49 | 16
| 1 | 4 | 1.01: 1.29; 1.02: 1.36; 1.03: 1.46; 1.04: 1.2197; 1.05: 1.5 | 1.3626
| 1 | 5 | 1.01: 1.33; 1.02: 1.31; 1.03: 1.45; 1.04: 1.1987; 1.05: 1.51 | 25
| 1 | 6 | 1.01: 1.19; 1.02: 1.33; 1.03: 1.45; 1.04: 1.2101; 1.05: 1.52
| Mean |  | 1.01: 1.2833; 1.02: 1.365; 1.03: 1.4533; 1.04: 1.2075; 1.05: 1.495
| Median |  | 1.01: 1.305; 1.02: 1.36; 1.03: 1.455; 1.04: 1.2149; 1.05: 1.495
| Std Dev. |  | 1.01: 0.054283; 1.02: 0.044159; 1.03: 0.036148; 1.04: 0.023724; 1.05: 0.018708
| Rel.Std.Dev. |  | 1.01: 4.23%; 1.02: 3.235%; 1.03: 2.487%; 1.04: 1.965%; 1.05: 1.251%
| PDM3 |  | 1.01: -5.815%; 1.02: 0.1782%; 1.03: 6.661%; 1.04: -11.38%; 1.05: 9.719%
| Z-Score (Absolute) |  | 1.01: 0.62; 1.02: 0; 1.03: 0.67; 1.04: 1.2; 1.05: 0.99
| Analytical results for Fe in OREAS 625 (Certified Value 3.98 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 3.85 | 4.1 | 3.76 | 3.99 | 3.91 | 3.76 | 4.02 | 3.88 | 4.01 | 3.95 | 4 | 4.15 | 4.0014 | 3.616 | 4.0005 | 3.99 | 4.03 | 3.68 | 4.0559 | 3.9822 | 4.13 | 4.0073 | 4.02 | 4.17 | 4.01 | 3.87 | 4.07 | 4.15 | 3.73 | 1
| 1 | 2 | 3.96 | 4.12 | 3.68 | 4.09 | 3.84 | 3.75 | 3.88 | 3.85 | 3.95 | 3.97 | 4.04 | 4.19 | 3.9771 | 3.885 | 4.0158 | 4.03 | 4.02 | 3.7 | 3.9426 | 4.0497 | 4.09 | 4.0072 | 4.05 | 4.24 | 4.07 | 3.72 | 4.06 | 4.18 | 3.67 | 26
| 1 | 3 | 3.92 | 3.96 | 3.88 | 4.99 | 3.88 | 3.72 | 3.89 | 3.91 | 4.06 | 4.04 | 3.96 | 4.09 | 3.9262 | 3.736 | 4.1257 | 4.01 | 4.1 | 3.59 | 4.0264 | 3.9972 | 4.11 | 4.0132 | 3.99 | 4.21 | 4.04 | 3.9 | 3.98 | 4.11 | 3.74 | 16
| 1 | 4 | 3.93 | 4.07 | 3.86 | 4.07 | 3.89 | 3.8 | 3.93 | 3.9 | 3.99 | 3.95 | 4.01 | 4.1 | 3.8712 | 3.742 | 4.0517 | 4.05 | 4.04 | 3.71 | 4.0447 | 4.0102 | 4.22 | 4.0938 | 4.04 | 4.24 | 4.1 | 3.79 | 4.03 | 4.15 | 3.72 | 3.9776
| 1 | 5 | 3.85 | 4.1 | 3.8 | 4.01 | 3.87 | 3.78 | 3.9 | 4.17 | 4.05 | 3.99 | 3.95 | 4.02 | 3.8362 | 3.739 | 4.021 | 4 | 4.04 | 3.82 | 3.9865 | 4.0612 | 4.27 | 4.0324 | 4.04 | 4.17 | 4.03 | 3.88 | 3.94 | 4.14 | 3.78 | 26
| 1 | 6 | 3.82 | 3.86 | 3.78 | 4.02 | 3.98 | 3.75 | 4 | 3.81 | 3.84 | 3.97 | 3.95 | 4.15 | 3.9249 | 3.856 | 4.1073 | 4.02 | 4.03 | 3.81 | 4.0382 | 4.086 | 4.12 | 4.086 | 3.99 | 4.22 | 4.05 | 3.84 | 4.09 | 4.13 | 3.76
| Mean |  | 3.8883 | 4.035 | 3.7933 | 4.195 | 3.895 | 3.76 | 3.9367 | 3.92 | 3.9833 | 3.9783 | 3.985 | 4.1167 | 3.9228 | 3.7623 | 4.0537 | 4.0167 | 4.0433 | 3.7183 | 4.0157 | 4.0311 | 4.1567 | 4.04 | 4.0217 | 4.2083 | 4.05 | 3.8333 | 4.0283 | 4.1433 | 3.7333
| Median |  | 3.885 | 4.085 | 3.79 | 4.045 | 3.885 | 3.755 | 3.915 | 3.89 | 4 | 3.97 | 3.98 | 4.125 | 3.9256 | 3.7405 | 4.0363 | 4.015 | 4.035 | 3.705 | 4.0323 | 4.03 | 4.125 | 4.0228 | 4.03 | 4.215 | 4.045 | 3.855 | 4.045 | 4.145 | 3.735
| Std Dev. |  | 0.055648 | 0.1031 | 0.072296 | 0.39129 | 0.047645 | 0.027568 | 0.059554 | 0.12775 | 0.080911 | 0.033714 | 0.037283 | 0.059889 | 0.06216 | 0.096835 | 0.051763 | 0.021602 | 0.028752 | 0.08612 | 0.043067 | 0.040603 | 0.071461 | 0.039829 | 0.026394 | 0.031885 | 0.031623 | 0.067429 | 0.057764 | 0.023381 | 0.037771
| Rel.Std.Dev. |  | 1.431% | 2.555% | 1.906% | 9.328% | 1.223% | 0.7332% | 1.513% | 3.259% | 2.031% | 0.8475% | 0.9356% | 1.455% | 1.585% | 2.574% | 1.277% | 0.5378% | 0.7111% | 2.316% | 1.072% | 1.007% | 1.719% | 0.9859% | 0.6563% | 0.7577% | 0.7808% | 1.759% | 1.434% | 0.5643% | 1.012%
| PDM3 |  | -2.245% | 1.443% | -4.633% | 5.465% | -2.077% | -5.471% | -1.03% | -1.449% | 0.1437% | 0.018% | 0.1856% | 3.496% | -1.377% | -5.412% | 1.912% | 0.9817% | 1.652% | -6.519% | 0.9578% | 1.344% | 4.501% | 1.568% | 1.107% | 5.8% | 1.82% | -3.627% | 1.275% | 4.166% | -6.141%
| Z-Score (Absolute) |  | 0.92 | 0.14 | 1.61 | 1.3 | 0.88 | 1.86 | 0.57 | 0.69 | 0.24 | 0.27 | 0.22 | 0.73 | 0.67 | 1.84 | 0.28 | 0.01 | 0.2 | 2.16 | 0 | 0.11 | 1.02 | 0.18 | 0.04 | 1.4 | 0.25 | 1.32 | 0.09 | 0.93 | 2.05
| Analytical results for Ga in OREAS 625 (Certified Value 20.6 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 18 | 22.6 | 20.2 | 20.9 | 21.5 | 20.3 | 21.3 | 21.2 | 20.89 | 19.85 | 22.643 | 18.766 | 20.126 | 21.2 | 21 | 18.14 | 17 | 20.78 | 21.05 | 19.8 | 20.4 | 19.57 | 18 | [6 empty] | 1
| 1 | 2 | 19 | 22.6 | 21.2 | 17.8 | 20 | 20 | 22.7 | 21.5 | 20.78 | 20.23 | 21.7758 | 20.429 | 19.476 | 20.5 | 21 | 17.13 | 17.72 | 20.43 | 20.69 | 19.7 | 19.8 | 20.38 | 18 | [6 empty] | 44
| 1 | 3 | 19 | 21.6 | 21.2 | 20.7 | 19.75 | 20.8 | 23 | 21.5 | 20.93 | 20.57 | 21.6718 | 19.347 | 20.6473 | 20.9 | 20 | 18.05 | 17.28 | 20.5 | 20.96 | 20.1 | 21.2 | 19.75 | 18 | [6 empty] | 16
| 1 | 4 | 19 | 21.9 | 21.2 | 20.7 | 20.6 | 20.7 | 22.4 | 21.3 | 20.78 | 20.49 | 21.3171 | 19.149 | 19.5716 | 20.5 | 21 | 18.1 | 17.25 | 20.74 | 20.86 | 20.6 | 20.4 | 20.11 | 18 | [6 empty] | 20.5577
| 1 | 5 | 19 | 22 | 20.8 | 20 | 21.1 | 20.5 | 23.4 | 21.7 | 20.73 | 20.63 | 20.9795 | 19.317 | 20.2332 | 20.6 | 21 | 18.11 | 17.69 | 21.11 | 20.67 | 19.9 | 19.8 | 20.03 | 18 | [6 empty] | 27
| 1 | 6 | 19 | 21.7 | 21.4 | 20 | 19.8 | 21.4 | 21.5 | 20.8 | 21.23 | 20.82 | 21.5962 | 19.173 | 20.2138 | 21.1 | 21 | 18.06 | 17.38 | 20.26 | 20.62 | 20.8 | 20.1 | 19.9 | 19
| Mean |  | 18.8333 | 22.0667 | 21 | 20.0167 | 20.4583 | 20.6167 | 22.3833 | 21.3333 | 20.89 | 20.4317 | 21.6639 | 19.3635 | 20.0447 | 20.8 | 20.8333 | 17.9317 | 17.3867 | 20.6367 | 20.8083 | 20.15 | 20.2833 | 19.9567 | 18.1667
| Median |  | 19 | 21.95 | 21.2 | 20.35 | 20.3 | 20.6 | 22.55 | 21.4 | 20.835 | 20.53 | 21.634 | 19.245 | 20.1699 | 20.75 | 21 | 18.08 | 17.33 | 20.62 | 20.775 | 20 | 20.25 | 19.965 | 18
| Std Dev. |  | 0.40825 | 0.43665 | 0.43818 | 1.1514 | 0.72967 | 0.47924 | 0.83287 | 0.31411 | 0.18276 | 0.344 | 0.55936 | 0.56161 | 0.44297 | 0.30984 | 0.40825 | 0.39413 | 0.27667 | 0.30283 | 0.17475 | 0.45056 | 0.52313 | 0.28409 | 0.40825
| Rel.Std.Dev. |  | 2.168% | 1.979% | 2.087% | 5.752% | 3.567% | 2.325% | 3.721% | 1.472% | 0.8749% | 1.684% | 2.582% | 2.9% | 2.21% | 1.49% | 1.96% | 2.198% | 1.591% | 1.467% | 0.8398% | 2.236% | 2.579% | 1.424% | 2.247%
| PDM3 |  | -8.388% | 7.34% | 2.151% | -2.632% | -0.4835% | 0.2867% | 8.88% | 3.773% | 1.616% | -0.6132% | 5.381% | -5.809% | -2.496% | 1.178% | 1.341% | -12.77% | -15.43% | 0.3839% | 1.219% | -1.983% | -1.335% | -2.924% | -11.63%
| Z-Score (Absolute) |  | 2.48 | 2.46 | 0.83 | 0.67 | 0 | 0.24 | 2.94 | 1.34 | 0.66 | 0.04 | 1.84 | 1.67 | 0.63 | 0.52 | 0.57 | 3.86 | 4.69 | 0.27 | 0.53 | 0.47 | 0.27 | 0.77 | 3.5
| Analytical results for Gd in OREAS 625 (Certified Value 5.46 ppm) | [30 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.22; 1.07: 1.30 | [6 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 5.2; 1.02: 5.32; 1.04: 5.48; 1.05: 5.5609; 1.06: 8.4448; 1.07: 5.7 | [6 empty] | 1
| 1 | 2 | 1.01: 5.1; 1.02: 5.35; 1.04: 5.76; 1.05: 5.3129; 1.06: 7.9661; 1.07: 5.55 | [6 empty] | 10
| 1 | 3 | 1.01: 5.2; 1.02: 5.37; 1.04: 5.84; 1.05: 5.4482; 1.06: 7.8922; 1.07: 5.6 | [6 empty] | 16
| 1 | 4 | 1.01: 5.3; 1.02: 5.37; 1.04: 5.72; 1.05: 5.4023; 1.06: 7.9391; 1.07: 5.49 | [6 empty] | 5.463
| 1 | 5 | 1.01: 5.3; 1.02: 5.29; 1.04: 5.66; 1.05: 5.3107; 1.06: 7.7602; 1.07: 5.69 | [6 empty] | 28
| 1 | 6 | 1.01: 4.7; 1.02: 5.29; 1.04: 5.89; 1.05: 5.4864; 1.06: 8.281; 1.07: 5.68
| Mean |  | 1.01: 5.1333; 1.02: 5.3317; 1.04: 5.725; 1.05: 5.4202; 1.06: 8.0472; 1.07: 5.6183
| Median |  | 1.01: 5.2; 1.02: 5.335; 1.04: 5.74; 1.05: 5.4253; 1.06: 7.9526; 1.07: 5.64
| Std Dev. |  | 1.01: 0.22509; 1.02: 0.037103; 1.04: 0.14557; 1.05: 0.098794; 1.06: 0.25977; 1.07: 0.08612
| Rel.Std.Dev. |  | 1.01: 4.385%; 1.02: 0.6959%; 1.04: 2.543%; 1.05: 1.823%; 1.06: 3.228%; 1.07: 1.533%
| PDM3 |  | 1.01: -6.035%; 1.02: -2.405%; 1.04: 4.795%; 1.05: -0.7836%; 1.06: 47.3%; 1.07: 2.842%
| Z-Score (Absolute) |  | 1.01: 1.32; 1.02: 0.64; 1.04: 0.71; 1.05: 0.34; 1.06: 8.67; 1.07: 0.34
| Analytical results for Ge in OREAS 625 (Indicative Value 0.23 ppm) | [30 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | < 0.1 | 0.21 | 0.14 | 0.14 | 0.13 | 0.22 | 0.37 | 0.39 | 1.2 | 1 | < 0.2 | 0.26926 | 1.4 | 3.2 | 0.15 | 0.7 | 0.08 | 0.09 | [11 empty] | 1
| 1 | 2 | < 0.1 | 0.13 | 0.18 | 0.11 | 0.14 | 0.23 | 0.3 | 0.41 | 1.3 | 1 | < 0.2 | 0.27659 | 1.4 | 3.2 | 0.13 | 0.5 | 0.1 | 0.09 | [11 empty] | 46
| 1 | 3 | < 0.1 | 0.11 | 0.18 | 0.13 | 0.12 | 0.24 | 0.43 | 0.39 | 1.4 | 1 | < 0.2 | 0.27216 | 1.3 | 2.9 | 0.13 | 0.8 | 0.09 | 0.09 | [11 empty] | 16
| 1 | 4 | < 0.1 | 0.08 | 0.19 | 0.13 | 0.12 | 0.24 | 0.32 | 0.41 | 1.2 | 1.1 | < 0.2 | 0.2789 | 1.4 | 3.1 | 0.09 | 0.7 | 0.09 | 0.13 | [11 empty] | 0.22872
| 1 | 5 | < 0.1 | 0.11 | 0.18 | 0.15 | 0.12 | 0.25 | 0.3 | 0.4 | 1.3 | 1.1 | < 0.2 | 0.27423 | 1.4 | 3.4 | 0.12 | 0.8 | 0.09 | 0.09 | [11 empty] | 9
| 1 | 6 | < 0.1 | 0.13 | 0.16 | 0.15 | 0.12 | 0.24 | 0.36 | 0.41 | 1.2 | 1 | < 0.2 | 0.26701 | 1.4 | 3.3 | 0.16 | 0.6 | 0.1 | 0.09
| Mean |  |  | 0.12833 | 0.17167 | 0.135 | 0.125 | 0.23667 | 0.34667 | 0.40167 | 1.2667 | 1.0333 |  | 0.27303 | 1.3833 | 3.1833 | 0.13 | 0.68333 | 0.091667 | 0.096667
| Median |  |  | 0.12 | 0.18 | 0.135 | 0.12 | 0.24 | 0.34 | 0.405 | 1.25 | 1 |  | 0.27319 | 1.4 | 3.2 | 0.13 | 0.7 | 0.09 | 0.09
| Std Dev. |  |  | 0.044008 | 0.018348 | 0.015166 | 0.0083666 | 0.010328 | 0.050465 | 0.0098319 | 0.08165 | 0.05164 |  | 0.0044676 | 0.040825 | 0.17224 | 0.024495 | 0.1169 | 0.0075277 | 0.01633
| Rel.Std.Dev. |  |  | 34.29% | 10.69% | 11.23% | 6.693% | 4.364% | 14.56% | 2.448% | 6.446% | 4.997% |  | 1.636% | 2.951% | 5.411% | 18.84% | 17.11% | 8.212% | 16.89%
| PDM3 |  |  | -43.89% | -24.95% | -40.98% | -45.35% | 3.472% | 51.57% | 75.61% | 453.8% | 351.8% |  | 19.37% | 504.8% | 1292% | -43.16% | 198.8% | -59.92% | -57.74%
| Z-Score (Absolute) |  | 0.95 | 0.46 | 0.2 | 0.42 | 0.49 | 0.2 | 0.87 | 1.21 | 6.51 | 5.08 | 0.64 | 0.42 | 7.23 | 18.26 | 0.45 | 2.94 | 0.69 | 0.66
| Analytical results for Hf in OREAS 625 (Certified Value 6.36 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 5.9 | 6.3 | 7 | 6.5 | 6 | 5.9 | 5.9 | 6.7 | 6.57 | 6.2 | 6 | 6.4172 | 6.39 | 6.5 | 6.2 | 6.83 | 6.25 | 8.0187 | 6.41 | 6.81 | 6.93 | 6.1 | 5.8 | [6 empty] | 1
| 1 | 2 | 6.1 | 6.3 | 6.8 | 5.7 | 5.7 | 5.8 | 6.1 | 6.7 | 6.11 | 6.2 | 6.29 | 6.1721 | 6.4287 | 6.5 | 6.4 | 6.92 | 6.29 | 7.7904 | 6.66 | 6.89 | 6.8 | 5.8 | 6 | [6 empty] | 16
| 1 | 3 | 6.3 | 6.3 | 6.7 | 6.4 | 5.5 | 5.6 | 6.2 | 7 | 6.41 | 6.1 | 6.44 | 6.3083 | 6.3633 | 6.7 | 6.3 | 6.87 | 6.39 | 7.8637 | 6.3 | 7.18 | 7.07 | 6.1 | 5.8 | [6 empty] | 16
| 1 | 4 | 6.1 | 6.4 | 6.8 | 6.5 | 5.9 | 5.8 | 6.1 | 7.2 | 6.31 | 6.2 | 6.24 | 6.4214 | 6.4341 | 6.4 | 6.2 | 6.8 | 6.39 | 7.8127 | 6.36 | 7.22 | 6.65 | 6 | 5.8 | [6 empty] | 6.3553
| 1 | 5 | 6.3 | 5.9 | 6.6 | 6.5 | 5.7 | 5.6 | 6.4 | 6.9 | 6.23 | 6.1 | 6.2 | 6.6659 | 6.1743 | 6.6 | 6.3 | 6.5 | 6.49 | 7.8392 | 6.59 | 6.89 | 6.64 | 6 | 5.9 | [6 empty] | 29
| 1 | 6 | 5.8 | 6.4 | 6.5 | 6.6 | 5.5 | 6 | 6.1 | 6.8 | 6.01 | 6.1 | 6.29 | 6.4654 | 6.2091 | 6.5 | 6.3 | 6.86 | 6.31 | 7.8766 | 6.39 | 7.37 | 6.7 | 6.1 | 6
| Mean |  | 6.0833 | 6.2667 | 6.7333 | 6.3667 | 5.7167 | 5.7833 | 6.1333 | 6.8833 | 6.2733 | 6.15 | 6.2433 | 6.4084 | 6.3333 | 6.5333 | 6.2833 | 6.7967 | 6.3533 | 7.8669 | 6.4517 | 7.06 | 6.7983 | 6.0167 | 5.8833
| Median |  | 6.1 | 6.3 | 6.75 | 6.5 | 5.7 | 5.8 | 6.1 | 6.85 | 6.27 | 6.15 | 6.265 | 6.4193 | 6.3767 | 6.5 | 6.3 | 6.845 | 6.35 | 7.8515 | 6.4 | 7.035 | 6.75 | 6.05 | 5.85
| Std Dev. |  | 0.20412 | 0.18619 | 0.17512 | 0.33267 | 0.20412 | 0.16021 | 0.1633 | 0.19408 | 0.20294 | 0.054772 | 0.14431 | 0.16471 | 0.11324 | 0.10328 | 0.075277 | 0.15082 | 0.087101 | 0.080889 | 0.14106 | 0.22645 | 0.17198 | 0.1169 | 0.098319
| Rel.Std.Dev. |  | 3.355% | 2.971% | 2.601% | 5.225% | 3.571% | 2.77% | 2.662% | 2.82% | 3.235% | 0.8906% | 2.311% | 2.57% | 1.788% | 1.581% | 1.198% | 2.219% | 1.371% | 1.028% | 2.186% | 3.208% | 2.53% | 1.943% | 1.671%
| PDM3 |  | -4.28% | -1.395% | 5.948% | 0.1781% | -10.05% | -9.001% | -3.493% | 8.308% | -1.29% | -3.231% | -1.763% | 0.8346% | -0.3477% | 2.801% | -1.133% | 6.944% | -0.03169% | 23.78% | 1.516% | 11.09% | 6.97% | -5.329% | -7.427%
| Z-Score (Absolute) |  | 0.84 | 0.22 | 1.35 | 0.11 | 2.08 | 1.85 | 0.67 | 1.85 | 0.2 | 0.62 | 0.3 | 0.25 | 0 | 0.67 | 0.17 | 1.56 | 0.07 | 5.17 | 0.4 | 2.45 | 1.57 | 1.07 | 1.52
| Analytical results for Hg in OREAS 625 (Indicative Value 1.35 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab | [6 empty] | 1
| No | No | 1.02: 1.13; 1.03: 1.14; 1.04: 1.20; 1.05: 1.30 | [6 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*OES/MS; 1.04: 4A*OES/MS; 1.05: 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.02: 1.4848; 1.03: 1.0626; 1.04: 1.33; 1.05: < 2 | [6 empty] | 1
| 1 | 2 | 1.02: 1.4282; 1.03: 0.8265; 1.04: 1.31; 1.05: < 2 | [6 empty] | 4
| 1 | 3 | 1.02: 1.4102; 1.03: 1.8901; 1.04: 1.31; 1.05: < 2 | [6 empty] | 16
| 1 | 4 | 1.02: 1.4879; 1.03: 1.4279; 1.04: 1.12; 1.05: < 2 | [6 empty] | 1.348
| 1 | 5 | 1.02: 1.4244; 1.03: 1.4327; 1.04: 1.24; 1.05: < 2 | [6 empty] | 10
| 1 | 6 | 1.02: 1.559; 1.03: 0.9932; 1.04: 1.34; 1.05: < 2
| Mean |  | 1.02: 1.4658; 1.03: 1.2722; 1.04: 1.275
| Median |  | 1.02: 1.4565; 1.03: 1.2452; 1.04: 1.31
| Std Dev. |  | 1.02: 0.056131; 1.03: 0.38787; 1.04: 0.083606
| Rel.Std.Dev. |  | 1.02: 3.829%; 1.03: 30.49%; 1.04: 6.557%
| PDM3 |  | 1.02: 8.738%; 1.03: -5.624%; 1.04: -5.414%
| Z-Score (Absolute) |  | 1.02: 1.34; 1.03: 0.01; 1.04: 0.01; 1.05: 1.91
| Analytical results for Ho in OREAS 625 (Certified Value 0.54 ppm) | [30 empty] | Certified
Ho | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.02: 1.02; 1.03: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.22; 1.07: 1.30 | [6 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.02: 0.5; 1.03: 0.56; 1.04: 0.53; 1.05: 0.54771; 1.06: 0.72676; 1.07: 0.54 | [6 empty] | 1
| 1 | 2 | 1.02: 0.5; 1.03: 0.51; 1.04: 0.58; 1.05: 0.512; 1.06: 0.64529; 1.07: 0.54 | [6 empty] | 7
| 1 | 3 | 1.02: 0.5; 1.03: 0.56; 1.04: 0.57; 1.05: 0.53178; 1.06: 0.63957; 1.07: 0.55 | [6 empty] | 16
| 1 | 4 | 1.02: 0.5; 1.03: 0.54; 1.04: 0.55; 1.05: 0.54572; 1.06: 0.72913; 1.07: 0.54 | [6 empty] | 0.53594
| 1 | 5 | 1.02: 0.5; 1.03: 0.55; 1.04: 0.56; 1.05: 0.52151; 1.06: 0.62975; 1.07: 0.55 | [6 empty] | 30
| 1 | 6 | 1.02: 0.5; 1.03: 0.52; 1.04: 0.57; 1.05: 0.53964; 1.06: 0.67956; 1.07: 0.56
| Mean |  | 1.02: 0.5; 1.03: 0.54; 1.04: 0.56; 1.05: 0.53306; 1.06: 0.67501; 1.07: 0.54667
| Median |  | 1.02: 0.5; 1.03: 0.545; 1.04: 0.565; 1.05: 0.53571; 1.06: 0.66243; 1.07: 0.545
| Std Dev. |  | 1.02: 0; 1.03: 0.020976; 1.04: 0.017889; 1.05: 0.014117; 1.06: 0.044307; 1.07: 0.008165
| Rel.Std.Dev. |  | 1.02: 0%; 1.03: 3.884%; 1.04: 3.194%; 1.05: 2.648%; 1.06: 6.564%; 1.07: 1.494%
| PDM3 |  | 1.02: -6.707%; 1.03: 0.7566%; 1.04: 4.488%; 1.05: -0.5387%; 1.06: 25.95%; 1.07: 2.001%
| Z-Score (Absolute) |  | 1.02: 2.17; 1.03: 0.17; 1.04: 0.83; 1.05: 0.51; 1.06: 6.59; 1.07: 0.17
| Analytical results for In in OREAS 625 (Certified Value 2.12 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 2
3
| 1 | 1 | 2 | 2.1 | 2.16 | 2.12 | 2.18 | 1.925 | 2.06 | 2.11 | 2.12 | 2.24 | 2.13 | 2.08 | 2.0703 | 2.15 | 1.99 | 2.229 | 2.3 | 2.078 | 2.06 | 2 | 2.136 | 2.693 | [7 empty] | 1
| 1 | 2 | < 2 | 2 | 2 | 2.21 | 1.935 | 1.915 | 2.04 | 2.24 | 2.11 | 2.22 | 2.17 | 2.19 | 2.082 | 2.19 | 2.05 | 2.194 | 2.2 | 2.151 | 2.02 | 1.89 | 2.109 | 2.078 | [7 empty] | 33
| 1 | 3 | < 2 | 2.1 | 2.05 | 2.2 | 2.15 | 1.83 | 2.09 | 2.32 | 2.14 | 2.27 | 2.14 | 2.15 | 2.0884 | 2.18 | 2.01 | 2.249 | 2.2 | 2.082 | 2.15 | 2.11 | 2.095 | 2.091 | [7 empty] | 16
| 1 | 4 | 2 | 2 | 2.17 | 2.23 | 2.2 | 2.02 | 2.12 | 2.25 | 2.11 | 2.23 | 2.14 | 2.19 | 2.0369 | 2.2 | 2 | 2.145 | 2.2 | 2.136 | 2.16 | 1.99 | 2.109 | 2.126 | [7 empty] | 2.1179
| 1 | 5 | < 2 | 2.1 | 1.99 | 2.16 | 2.12 | 1.945 | 2.11 | 2.39 | 2.13 | 2.26 | 2.1 | 2.15 | 2.0617 | 2.18 | 2.01 | 2.421 | 2.2 | 2.035 | 2.03 | 1.91 | 2.126 | 2.139 | [7 empty] | 31
| 1 | 6 | 2 | 2 | 2.04 | 2.25 | 2.22 | 1.845 | 2.2 | 2.19 | 2.13 | 2.28 | 2.11 | 2.19 | 2.0292 | 2.16 | 2.02 | 2.201 | 2.2 | 2.063 | 2.18 | 1.93 | 2.123 | 2.843
| Mean |  | 2 | 2.05 | 2.0683 | 2.195 | 2.1342 | 1.9133 | 2.1033 | 2.25 | 2.1233 | 2.25 | 2.1317 | 2.1583 | 2.0614 | 2.1767 | 2.0133 | 2.2398 | 2.2167 | 2.0908 | 2.1 | 1.9717 | 2.1163 | 2.3283
| Median |  | 2 | 2.05 | 2.045 | 2.205 | 2.165 | 1.92 | 2.1 | 2.245 | 2.125 | 2.25 | 2.135 | 2.17 | 2.066 | 2.18 | 2.01 | 2.215 | 2.2 | 2.08 | 2.105 | 1.96 | 2.116 | 2.1325
| Std Dev. |  | 0 | 0.054772 | 0.078337 | 0.047645 | 0.10385 | 0.069474 | 0.056095 | 0.097775 | 0.012111 | 0.023664 | 0.024833 | 0.043089 | 0.023988 | 0.018619 | 0.020656 | 0.095525 | 0.040825 | 0.04426 | 0.071274 | 0.080602 | 0.014747 | 0.34457
| Rel.Std.Dev. |  | 0% | 2.672% | 3.787% | 2.171% | 4.866% | 3.631% | 2.667% | 4.346% | 0.5704% | 1.052% | 1.165% | 1.996% | 1.164% | 0.8554% | 1.026% | 4.265% | 1.842% | 2.117% | 3.394% | 4.088% | 0.6968% | 14.8%
| PDM3 |  | -5.568% | -3.207% | -2.341% | 3.639% | 0.767% | -9.66% | -0.6888% | 6.236% | 0.2555% | 6.236% | 0.649% | 1.908% | -2.668% | 2.774% | -4.938% | 5.756% | 4.662% | -1.279% | -0.8462% | -6.906% | -0.07499% | 9.935%
| Z-Score (Absolute) |  | NA | 0.8 | 0.6 | 0.78 | 0.12 | 2.29 | 0.22 | 1.38 | 0 | 1.38 | 0.09 | 0.38 | 0.67 | 0.58 | 1.2 | 1.27 | 1.02 | 0.35 | 0.25 | 1.65 | 0.08 | 2.23
| Analytical results for K in OREAS 625 (Certified Value 2.86 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 2.85 | 2.68 | 2.81 | 2.94 | 2.84 | 2.81 | 2.94 | 3 | 2.68 | 2.807 | 3.04 | 2.9396 | 2.8477 | 2.291 | 2.8224 | 2.89 | 3 | 2.78 | 3.051 | 2.6788 | 2.89 | 2.6356 | 2.83 | 3.04 | 2.9 | 2.8 | 2.94 | 2.89 | 2.87 | 1
| 1 | 2 | 2.89 | 3.2 | 2.67 | 2.95 | 2.79 | 2.77 | 2.82 | 3.05 | 2.59 | 2.8187 | 3.05 | 2.958 | 2.7978 | 2.855 | 2.8239 | 2.89 | 3.09 | 2.77 | 2.789 | 2.7328 | 2.86 | 2.6317 | 2.84 | 3.07 | 2.91 | 2.8 | 2.91 | 2.9 | 2.83 | 0
| 1 | 3 | 2.88 | 3.03 | 2.86 | 2.96 | 2.82 | 2.77 | 2.78 | 3.09 | 2.64 | 2.8147 | 3.03 | 2.8971 | 2.7958 | 2.672 | 2.8753 | 2.88 | 3.05 | 2.78 | 2.8654 | 2.6996 | 2.87 | 2.6048 | 2.82 | 3.05 | 2.96 | 2.79 | 2.92 | 2.87 | 2.86 | 16
| 1 | 4 | 2.9 | 2.63 | 2.86 | 2.96 | 2.81 | 2.83 | 2.83 | 3.01 | 2.63 | 2.7921 | 3.07 | 2.8961 | 2.7051 | 2.304 | 2.8585 | 2.9 | 3.04 | 2.82 | 2.8635 | 2.7197 | 2.93 | 2.65 | 2.86 | 3.06 | 2.96 | 2.75 | 2.89 | 2.84 | 2.9 | 2.8649
| 1 | 5 | 2.85 | 3.3 | 2.82 | 2.98 | 2.81 | 2.84 | 2.76 | 3.2 | 2.66 | 2.8219 | 3 | 2.8463 | 2.6884 | 2.373 | 2.7927 | 2.88 | 3.03 | 2.81 | 2.8457 | 2.7646 | 2.97 | 2.6471 | 2.83 | 3 | 2.92 | 2.78 | 2.91 | 2.88 | 2.89 | 32
| 1 | 6 | 2.9 | 3.07 | 2.81 | 2.95 | 2.84 | 2.79 | 2.88 | 2.98 | 2.53 | 2.8087 | 2.98 | 2.8812 | 2.7907 | 2.581 | 2.8914 | 2.89 | 2.97 | 2.84 | 2.8976 | 2.752 | 2.88 | 2.6821 | 2.79 | 3.11 | 3 | 2.78 | 2.93 | 2.91 | 2.92
| Mean |  | 2.8783 | 2.985 | 2.805 | 2.9567 | 2.8183 | 2.8017 | 2.835 | 3.055 | 2.6217 | 2.8105 | 3.0283 | 2.903 | 2.7709 | 2.5127 | 2.844 | 2.8883 | 3.03 | 2.8 | 2.8854 | 2.7246 | 2.9 | 2.6419 | 2.8283 | 3.055 | 2.9417 | 2.7833 | 2.9167 | 2.8817 | 2.8783
| Median |  | 2.885 | 3.05 | 2.815 | 2.955 | 2.815 | 2.8 | 2.825 | 3.03 | 2.635 | 2.8117 | 3.035 | 2.8966 | 2.7933 | 2.477 | 2.8412 | 2.89 | 3.035 | 2.795 | 2.8645 | 2.7262 | 2.885 | 2.6414 | 2.83 | 3.055 | 2.94 | 2.785 | 2.915 | 2.885 | 2.88
| Std Dev. |  | 0.023166 | 0.2734 | 0.070071 | 0.013663 | 0.019408 | 0.029944 | 0.066257 | 0.081179 | 0.054191 | 0.010667 | 0.033116 | 0.040341 | 0.061252 | 0.22779 | 0.037272 | 0.0075277 | 0.041473 | 0.027568 | 0.08868 | 0.032139 | 0.041952 | 0.025412 | 0.023166 | 0.036194 | 0.038166 | 0.018619 | 0.017512 | 0.024833 | 0.031885
| Rel.Std.Dev. |  | 0.8048% | 9.159% | 2.498% | 0.4621% | 0.6886% | 1.069% | 2.337% | 2.657% | 2.067% | 0.3795% | 1.094% | 1.39% | 2.211% | 9.066% | 1.311% | 0.2606% | 1.369% | 0.9846% | 3.073% | 1.18% | 1.447% | 0.9619% | 0.8191% | 1.185% | 1.297% | 0.6689% | 0.6004% | 0.8618% | 1.108%
| PDM3 |  | 0.4705% | 4.194% | -2.089% | 3.205% | -1.624% | -2.206% | -1.042% | 6.637% | -8.489% | -1.897% | 5.706% | 1.333% | -3.279% | -12.29% | -0.7275% | 0.8196% | 5.765% | -2.264% | 0.716% | -4.896% | 1.227% | -7.783% | -1.275% | 6.637% | 2.681% | -2.846% | 1.809% | 0.5869% | 0.4705%
| Z-Score (Absolute) |  | 0 | 0.98 | 0.67 | 0.72 | 0.55 | 0.7 | 0.4 | 1.62 | 2.36 | 0.62 | 1.38 | 0.23 | 0.99 | 3.36 | 0.32 | 0.09 | 1.39 | 0.72 | 0.06 | 1.41 | 0.2 | 2.17 | 0.46 | 1.62 | 0.58 | 0.87 | 0.35 | 0.03 | 0
| Analytical results for La in OREAS 625 (Certified Value 34.7 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 37 | 25.3 | 27.2 | 34.5 | 27.8 | 29.9 | 34.1 | 31.2 | 40.7 | 43.41 | 27.6 | 42 | 31.0824 | 37.72 | 41 | 25.68 | 28.9 | 34 | 38.03 | 39 | 44.1 | 30.3 | 37.2 | 34.6 | 38 |  |  |  |  | 1
| 1 | 2 | 38 | 28.6 | 29.8 | 37.5 | 26.8 | 32.5 | 33.8 | 34.3 | 41.8 | 41.08 | 23.9 | 42 | 30.3582 | 38.74 | 40 | 25.36 | 29.7 | 34.2 | 38.18 | 38 | 45 | 24.3 | 39.6 | 36.5 | 39 |  |  |  |  | 34
| 1 | 3 | 38 | 27.4 | 30 | 36.1 | 27 | 23.4 | 34.1 | 34.1 | 40.7 | 43.55 | 27.5 | 41 | 31.1865 | 38.14 | 40 | 25.45 | 29.2 | 33 | 37.62 | 39 | 48.9 | 32.7 | 37.5 | 35.6 | 38 |  |  |  |  | 16
| 1 | 4 | 37 | 24.7 | 28.8 | 38.7 | 32.5 | 29.4 | 34.8 | 31.1 | 41.4 | 42.11 | 26.6 | 40 | 30.2435 | 37.22 | 43 | 25.51 | 29.6 | 34.3 | 34.63 | 38 | 49.2 | 28.8 | 36.4 | 36.3 | 38 |  |  |  |  | 34.7157
| 1 | 5 | 38 | 28.3 | 30.5 | 33.5 | 33.1 | 27 | 33.7 | 32.9 | 42.6 | 41.66 | 26.7 | 40 | 29.7594 | 38 | 44 | 25.65 | 29.6 | 35.3 | 38.44 | 38 | 45.9 | 28.6 | 38.3 | 33.4 | 39 |  |  |  |  | 33
| 1 | 6 | 37 | 21.8 | 28.9 | 36 | 30.7 | 27.6 | 36.4 | 34.9 | 40.8 | 41.28 | 25.2 | 40 | 31.1186 | 37.98 | 44 | 25.46 | 29.4 | 34 | 37.4 | 39 | 46.9 | 30.6 | 38.2 | 36 | 37
| Mean |  | 37.5 | 26.0167 | 29.2 | 36.05 | 29.65 | 28.3 | 34.4833 | 33.0833 | 41.3333 | 42.1817 | 26.25 | 40.8333 | 30.6248 | 37.9667 | 42 | 25.5183 | 29.4 | 34.1333 | 37.3833 | 38.5 | 46.6667 | 29.2167 | 37.8667 | 35.4 | 38.1667
| Median |  | 37.5 | 26.35 | 29.35 | 36.05 | 29.25 | 28.5 | 34.1 | 33.5 | 41.1 | 41.885 | 26.65 | 40.5 | 30.7203 | 37.99 | 42 | 25.485 | 29.5 | 34.1 | 37.825 | 38.5 | 46.4 | 29.55 | 37.85 | 35.8 | 38
| Std Dev. |  | 0.54772 | 2.6011 | 1.1781 | 1.8992 | 2.8176 | 3.0854 | 1.0147 | 1.6327 | 0.76333 | 1.0663 | 1.4377 | 0.98319 | 0.58891 | 0.49955 | 1.8974 | 0.12384 | 0.30332 | 0.73666 | 1.4005 | 0.54772 | 2.0695 | 2.8266 | 1.0985 | 1.1883 | 0.75277
| Rel.Std.Dev. |  | 1.461% | 9.998% | 4.035% | 5.268% | 9.503% | 10.9% | 2.943% | 4.935% | 1.847% | 2.528% | 5.477% | 2.408% | 1.923% | 1.316% | 4.518% | 0.4853% | 1.032% | 2.158% | 3.746% | 1.423% | 4.435% | 9.675% | 2.901% | 3.357% | 1.972%
| PDM3 |  | 8.02% | -25.06% | -15.89% | 3.844% | -14.59% | -18.48% | -0.6693% | -4.702% | 19.06% | 21.51% | -24.39% | 17.62% | -11.78% | 9.365% | 20.98% | -26.49% | -15.31% | -1.677% | 7.684% | 10.9% | 34.43% | -15.84% | 9.077% | 1.971% | 9.941%
| Z-Score (Absolute) |  | 0.26 | 1.16 | 0.77 | 0.08 | 0.71 | 0.88 | 0.11 | 0.29 | 0.74 | 0.84 | 1.14 | 0.67 | 0.59 | 0.32 | 0.82 | 1.23 | 0.74 | 0.16 | 0.25 | 0.38 | 1.4 | 0.77 | 0.31 | 0 | 0.34
| Analytical results for Li in OREAS 625 (Certified Value 17.9 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 18 | 18.4 | 18.2 | 18.3 | 17.8 | 17 | 18.2 | 18.8 | 17.1 | 17.3 | 18.1 | 17.5 | 18.0607 | 17.432 | 17.4077 | 15.86 | 18.1 | 14.2 | 18.47 | 17 | 17.3 | 24.9742 | 18.4 | 18 | 18 | 15.8 | 18 | 17 |  | 1
| 1 | 2 | 18 | 18.3 | 17.3 | 18.7 | 17.2 | 16.9 | 17.4 | 20.1 | 17.1 | 16.9 | 18.2 | 18.3 | 17.5114 | 21.583 | 17.3804 | 15.86 | 19.1 | 15.1 | 17.92 | 17 | 17.3 | 24.4036 | 16.6 | 19 | 17 | 17.4 | 19 | 18 |  | 0
| 1 | 3 | 18 | 17.7 | 18.5 | 18.6 | 17.6 | 16.9 | 17.4 | 21.3 | 18 | 17.5 | 17 | 18.6 | 17.5737 | 23.881 | 17.2666 | 15.5 | 18.1 | 15.1 | 18.22 | 17 | 17.6 | 24.2565 | 18.6 | 19 | 18 | 17.3 | 18 | 18 |  | 16
| 1 | 4 | 19 | 18.3 | 18.5 | 18.4 | 17.6 | 17.2 | 17.4 | 20.2 | 18.5 | 17 | 18.1 | 18.1 | 17.4563 | 19.734 | 17.293 | 15.35 | 18.5 | 14.6 | 18.14 | 17 | 17.1 | 24.2084 | 17.8 | 20 | 17 | 17.5 | 18 | 18 |  | 17.8834
| 1 | 5 | 18 | 18.1 | 18.2 | 18.4 | 17.7 | 17.4 | 17.4 | 20.8 | 17.8 | 17.3 | 17.6 | 18.4 | 17.2856 | 20.225 | 17.3847 | 15.17 | 17.8 | 16 | 18.19 | 18 | 17 | 24.3629 | 18.9 | 18 | 17 | 17.3 | 18 | 19 |  | 34
| 1 | 6 | 18 | 17.2 | 18.2 | 18.6 | 18.1 | 17 | 17.8 | 18.2 | 16.8 | 17.2 | 17.8 | 18.4 | 17.6776 | 20.223 | 17.3058 | 15.46 | 18.2 | 17.1 | 18.31 | 17 | 16.7 | 24.6287 | 17.8 | 20 | 17 | 17 | 20 | 18
| Mean |  | 18.1667 | 18 | 18.15 | 18.5 | 17.6667 | 17.0667 | 17.6 | 19.9 | 17.55 | 17.2 | 17.8 | 18.2167 | 17.5942 | 20.513 | 17.3397 | 15.5333 | 18.3 | 15.35 | 18.2083 | 17.1667 | 17.1667 | 24.4724 | 18.0167 | 19 | 17.3333 | 17.05 | 18.5 | 18
| Median |  | 18 | 18.2 | 18.2 | 18.5 | 17.65 | 17 | 17.4 | 20.15 | 17.45 | 17.25 | 17.95 | 18.35 | 17.5425 | 20.224 | 17.3431 | 15.48 | 18.15 | 15.1 | 18.205 | 17 | 17.2 | 24.3833 | 18.1 | 19 | 17 | 17.3 | 18 | 18
| Std Dev. |  | 0.40825 | 0.46476 | 0.44159 | 0.15492 | 0.29439 | 0.19664 | 0.33466 | 1.1832 | 0.65345 | 0.21909 | 0.45166 | 0.38687 | 0.26309 | 2.1333 | 0.058255 | 0.27768 | 0.45166 | 1.0483 | 0.1828 | 0.40825 | 0.30768 | 0.28608 | 0.82077 | 0.89443 | 0.5164 | 0.63482 | 0.83666 | 0.63246
| Rel.Std.Dev. |  | 2.247% | 2.582% | 2.433% | 0.8374% | 1.666% | 1.152% | 1.902% | 5.946% | 3.723% | 1.274% | 2.537% | 2.124% | 1.495% | 10.4% | 0.336% | 1.788% | 2.468% | 6.83% | 1.004% | 2.378% | 1.792% | 1.169% | 4.556% | 4.708% | 2.979% | 3.723% | 4.522% | 3.514%
| PDM3 |  | 1.584% | 0.6522% | 1.491% | 3.448% | -1.212% | -4.567% | -1.584% | 11.28% | -1.864% | -3.821% | -0.4661% | 1.864% | -1.617% | 14.7% | -3.04% | -13.14% | 2.33% | -14.17% | 1.817% | -4.008% | -4.008% | 36.84% | 0.7454% | 6.244% | -3.076% | -4.66% | 3.448% | 0.6522%
| Z-Score (Absolute) |  | 0.32 | 0.12 | 0.3 | 0.72 | 0.28 | 1 | 0.36 | 2.39 | 0.42 | 0.84 | 0.12 | 0.38 | 0.37 | 3.13 | 0.67 | 2.83 | 0.48 | 3.05 | 0.37 | 0.88 | 0.88 | 7.87 | 0.14 | 1.32 | 0.68 | 1.02 | 0.72 | 0.12
| Analytical results for Lu in OREAS 625 (Certified Value 0.14 ppm) | [30 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.23; 1.06: 1.25; 1.07: 1.26; 1.08: 1.30 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS; 1.08: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 0.1; 1.02: 0.14; 1.03: 0.13; 1.04: 0.14156; 1.05: 0.14; 1.06: 0.15; 1.07: 0.13; 1.08: 0.14 |  | 1
| 1 | 2 | 1.01: 0.1; 1.02: 0.13; 1.03: 0.14; 1.04: 0.13123; 1.05: 0.13; 1.06: 0.15; 1.07: 0.11; 1.08: 0.14 |  | 35
| 1 | 3 | 1.01: 0.1; 1.02: 0.14; 1.03: 0.15; 1.04: 0.12994; 1.05: 0.13; 1.06: 0.17; 1.07: 0.14; 1.08: 0.14 |  | 16
| 1 | 4 | 1.01: 0.1; 1.02: 0.13; 1.03: 0.14; 1.04: 0.13516; 1.05: 0.13; 1.06: 0.17; 1.07: 0.13; 1.08: 0.14 |  | 0.13519
| 1 | 5 | 1.01: 0.1; 1.02: 0.12; 1.03: 0.14; 1.04: 0.135; 1.05: 0.13; 1.06: 0.16; 1.07: 0.13; 1.08: 0.15 |  | 35
| 1 | 6 | 1.01: 0.1; 1.02: 0.13; 1.03: 0.14; 1.04: 0.13209; 1.05: 0.13; 1.06: 0.17; 1.07: 0.13; 1.08: 0.14
| Mean |  | 1.01: 0.1; 1.02: 0.13167; 1.03: 0.14; 1.04: 0.13416; 1.05: 0.13167; 1.06: 0.16167; 1.07: 0.12833; 1.08: 0.14167
| Median |  | 1.01: 0.1; 1.02: 0.13; 1.03: 0.14; 1.04: 0.13354; 1.05: 0.13; 1.06: 0.165; 1.07: 0.13; 1.08: 0.14
| Std Dev. |  | 1.01: 1.5202e-17; 1.02: 0.0075277; 1.03: 0.0063246; 1.04: 0.0041753; 1.05: 0.0040825; 1.06: 0.0098319; 1.07: 0.0098319; 1.08: 0.0040825
| Rel.Std.Dev. |  | 1.01: 1.52e-14%; 1.02: 5.717%; 1.03: 4.518%; 1.04: 3.112%; 1.05: 3.101%; 1.06: 6.082%; 1.07: 7.661%; 1.08: 2.882%
| PDM3 |  | 1.01: -26.03%; 1.02: -2.609%; 1.03: 3.555%; 1.04: -0.7624%; 1.05: -2.609%; 1.06: 19.58%; 1.07: -5.075%; 1.08: 4.788%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.29; 1.03: 0.68; 1.04: 0; 1.05: 0.29; 1.06: 3.18; 1.07: 0.67; 1.08: 0.87
Results from laboratory 1.02 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 625 (Certified Value 0.261 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.25 | 0.25 | 0.25 | 0.25 | 0.22 | 0.23 | 0.25 | 0.24 | 0.26 | 0.2631 | 0.28 | 0.2683 | 0.25312 | 0.26331 | 0.28 | 0.262 | 0.26 | 0.2653 | 0.263 | 0.26 | 0.28062 | 0.26 | 0.27 | 0.26 | 0.24 | 0.25 | 0.286 | 0.25 |  | 1
| 1 | 2 | 0.27 | 0.25 | 0.25 | 0.26 | 0.23 | 0.23 | 0.24 | 0.23 | 0.25 | 0.2646 | 0.28 | 0.2702 | 0.24749 | 0.26595 | 0.28 | 0.262 | 0.25 | 0.2509 | 0.2697 | 0.25 | 0.27637 | 0.261 | 0.28 | 0.26 | 0.25 | 0.26 | 0.288 | 0.25 |  | 0
| 1 | 3 | 0.26 | 0.25 | 0.26 | 0.25 | 0.23 | 0.21 | 0.24 | 0.24 | 0.26 | 0.2639 | 0.27 | 0.2624 | 0.24816 | 0.27181 | 0.28 | 0.265 | 0.25 | 0.2576 | 0.2664 | 0.25 | 0.27721 | 0.259 | 0.27 | 0.26 | 0.25 | 0.25 | 0.287 | 0.26 |  | 16
| 1 | 4 | 0.28 | 0.27 | 0.26 | 0.25 | 0.23 | 0.23 | 0.25 | 0.24 | 0.26 | 0.2625 | 0.27 | 0.2648 | 0.24371 | 0.2615 | 0.28 | 0.263 | 0.26 | 0.2588 | 0.2687 | 0.26 | 0.27868 | 0.261 | 0.28 | 0.27 | 0.25 | 0.26 | 0.285 | 0.26 |  | 0.26131
| 1 | 5 | 0.26 | 0.27 | 0.26 | 0.24 | 0.23 | 0.22 | 0.24 | 0.25 | 0.26 | 0.2649 | 0.27 | 0.2599 | 0.24076 | 0.25937 | 0.28 | 0.263 | 0.27 | 0.2587 | 0.2715 | 0.26 | 0.27825 | 0.261 | 0.27 | 0.26 | 0.25 | 0.25 | 0.287 | 0.27 |  | 36
| 1 | 6 | 0.28 | 0.22 | 0.26 | 0.25 | 0.24 | 0.22 | 0.25 | 0.23 | 0.25 | 0.2629 | 0.28 | 0.2619 | 0.24617 | 0.27079 | 0.28 | 0.263 | 0.27 | 0.2573 | 0.269 | 0.25 | 0.28043 | 0.258 | 0.28 | 0.26 | 0.25 | 0.26 | 0.286 | 0.26
| Mean |  | 0.26667 | 0.25167 | 0.25667 | 0.25 | 0.23 | 0.22333 | 0.245 | 0.23833 | 0.25667 | 0.26365 | 0.275 | 0.26458 | 0.24657 | 0.26545 | 0.28 | 0.263 | 0.26 | 0.2581 | 0.26805 | 0.255 | 0.27859 | 0.26 | 0.275 | 0.26167 | 0.24833 | 0.255 | 0.2865 | 0.25833
| Median |  | 0.265 | 0.25 | 0.26 | 0.25 | 0.23 | 0.225 | 0.245 | 0.24 | 0.26 | 0.2635 | 0.275 | 0.2636 | 0.24683 | 0.26463 | 0.28 | 0.263 | 0.26 | 0.25815 | 0.26885 | 0.255 | 0.27847 | 0.2605 | 0.275 | 0.26 | 0.25 | 0.255 | 0.2865 | 0.26
| Std Dev. |  | 0.012111 | 0.018348 | 0.005164 | 0.0063246 | 0.0063246 | 0.008165 | 0.0054772 | 0.0075277 | 0.005164 | 0.00097108 | 0.0054772 | 0.0039827 | 0.0042033 | 0.0050251 | 0 | 0.0010954 | 0.0089443 | 0.0045917 | 0.0029724 | 0.0054772 | 0.0017008 | 0.0012649 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0010488 | 0.0075277
| Rel.Std.Dev. |  | 4.541% | 7.291% | 2.012% | 2.53% | 2.75% | 3.656% | 2.236% | 3.158% | 2.012% | 0.3683% | 1.992% | 1.505% | 1.705% | 1.893% | 0% | 0.4165% | 3.44% | 1.779% | 1.109% | 2.148% | 0.6105% | 0.4865% | 1.992% | 1.56% | 1.644% | 2.148% | 0.3661% | 2.914%
| PDM3 |  | 2.048% | -3.692% | -1.778% | -4.33% | -11.98% | -14.53% | -6.243% | -8.794% | -1.778% | 0.8939% | 5.237% | 1.251% | -5.644% | 1.584% | 7.151% | 0.6452% | -0.5028% | -1.23% | 2.578% | -2.416% | 6.612% | -0.5028% | 5.237% | 0.135% | -4.967% | -2.416% | 9.638% | -1.141%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.22 | 0.82 | 2.62 | 3.22 | 1.27 | 1.87 | 0.22 | 0.4 | 1.42 | 0.49 | 1.13 | 0.57 | 1.87 | 0.34 | 0.07 | 0.1 | 0.8 | 0.37 | 1.75 | 0.07 | 1.42 | 0.22 | 0.97 | 0.37 | 2.46 | 0.07
| Analytical results for Mn in OREAS 625 (Certified Value 0.443 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.4262 | 0.478 | 0.3961 | 0.452 | 0.431 | 0.441 | 0.448 | 0.463 | 0.477 | 0.4379 | 0.4576 | 0.4714 | 0.43369 | 0.29 | 0.44934 | 0.4293 | 0.434 | 0.4392 | 0.4432 | 0.4065 | 0.4435 | 0.47057 | 0.4456 | 0.457 | 0.4424 | 0.439 | 0.4399 | 0.4717 | 0.4295 | 1
| 1 | 2 | 0.4296 | 0.471 | 0.3834 | 0.456 | 0.424 | 0.433 | 0.432 | 0.456 | 0.453 | 0.4407 | 0.4568 | 0.4748 | 0.43126 | 0.319 | 0.44084 | 0.4329 | 0.431 | 0.4194 | 0.4195 | 0.4165 | 0.4449 | 0.47933 | 0.447 | 0.4665 | 0.4428 | 0.4238 | 0.4318 | 0.4733 | 0.4313 | 21
| 1 | 3 | 0.4298 | 0.444 | 0.4073 | 0.455 | 0.429 | 0.439 | 0.431 | 0.466 | 0.477 | 0.4383 | 0.4481 | 0.4639 | 0.42829 | 0.327 | 0.45098 | 0.4295 | 0.439 | 0.4113 | 0.4367 | 0.4121 | 0.4377 | 0.47572 | 0.4439 | 0.4551 | 0.4439 | 0.4608 | 0.4291 | 0.4672 | 0.437 | 16
| 1 | 4 | 0.4357 | 0.463 | 0.4065 | 0.456 | 0.427 | 0.444 | 0.432 | 0.464 | 0.462 | 0.4366 | 0.4556 | 0.467 | 0.42301 | 0.301 | 0.44243 | 0.4317 | 0.435 | 0.4426 | 0.4388 | 0.4089 | 0.4408 | 0.47719 | 0.4474 | 0.4563 | 0.4446 | 0.4399 | 0.431 | 0.4718 | 0.4369 | 0.44334
| 1 | 5 | 0.4346 | 0.468 | 0.3981 | 0.45 | 0.425 | 0.449 | 0.435 | 0.499 | 0.473 | 0.443 | 0.4481 | 0.4564 | 0.42307 | 0.311 | 0.44569 | 0.4292 | 0.432 | 0.4584 | 0.4396 | 0.419 | 0.4466 | 0.46931 | 0.4458 | 0.4555 | 0.4414 | 0.4261 | 0.4233 | 0.4707 | 0.4398 | 37
| 1 | 6 | 0.4329 | 0.438 | 0.3994 | 0.458 | 0.435 | 0.441 | 0.44 | 0.448 | 0.45 | 0.4394 | 0.4425 | 0.4634 | 0.4316 | 0.307 | 0.44518 | 0.4303 | 0.435 | 0.4655 | 0.4386 | 0.4154 | 0.4392 | 0.47605 | 0.4407 | 0.4694 | 0.4409 | 0.4264 | 0.4408 | 0.4668 | 0.4393
| Mean |  | 0.43147 | 0.46033 | 0.39847 | 0.4545 | 0.4285 | 0.44117 | 0.43633 | 0.466 | 0.46533 | 0.43932 | 0.45145 | 0.46615 | 0.42849 | 0.30917 | 0.44574 | 0.43048 | 0.43433 | 0.4394 | 0.43607 | 0.41307 | 0.44212 | 0.47469 | 0.44507 | 0.45997 | 0.44267 | 0.436 | 0.43265 | 0.47025 | 0.43563
| Median |  | 0.43135 | 0.4655 | 0.39875 | 0.4555 | 0.428 | 0.441 | 0.4335 | 0.4635 | 0.4675 | 0.43885 | 0.45185 | 0.46545 | 0.42978 | 0.309 | 0.44543 | 0.4299 | 0.4345 | 0.4409 | 0.4387 | 0.41375 | 0.44215 | 0.47589 | 0.4457 | 0.45665 | 0.4426 | 0.4327 | 0.4314 | 0.4712 | 0.43695
| Std Dev. |  | 0.0035719 | 0.015858 | 0.0086738 | 0.0029496 | 0.0040866 | 0.0053072 | 0.0065929 | 0.01747 | 0.012078 | 0.0022763 | 0.0061027 | 0.006493 | 0.0045564 | 0.013091 | 0.0038905 | 0.0015079 | 0.0028048 | 0.021176 | 0.0083918 | 0.0047719 | 0.0034499 | 0.0039143 | 0.0024671 | 0.0062858 | 0.0014194 | 0.013988 | 0.0066702 | 0.0026539 | 0.0042585
| Rel.Std.Dev. |  | 0.8279% | 3.445% | 2.177% | 0.649% | 0.9537% | 1.203% | 1.511% | 3.749% | 2.595% | 0.5182% | 1.352% | 1.393% | 1.063% | 4.234% | 0.8728% | 0.3503% | 0.6458% | 4.819% | 1.924% | 1.155% | 0.7803% | 0.8246% | 0.5543% | 1.367% | 0.3206% | 3.208% | 1.542% | 0.5644% | 0.9775%
| PDM3 |  | -2.678% | 3.833% | -10.12% | 2.517% | -3.347% | -0.4904% | -1.581% | 5.111% | 4.961% | -0.9077% | 1.829% | 5.145% | -3.35% | -30.26% | 0.5418% | -2.9% | -2.032% | -0.8889% | -1.641% | -6.829% | -0.2761% | 7.072% | 0.3893% | 3.75% | -0.1521% | -1.656% | -2.411% | 6.07% | -1.738%
| Z-Score (Absolute) |  | 0.6 | 1.58 | 3.1 | 1.14 | 0.82 | 0.13 | 0.23 | 2.01 | 1.96 | 0.01 | 0.91 | 2.02 | 0.83 | 9.85 | 0.48 | 0.67 | 0.38 | 0 | 0.25 | 1.99 | 0.21 | 2.67 | 0.43 | 1.56 | 0.25 | 0.26 | 0.51 | 2.33 | 0.28
| Analytical results for Mo in OREAS 625 (Certified Value 12.3 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 13 | 14.1 | 11.4 | 12.05 | 13.3 | 12.1 | 11.55 | 12.6 | 11.9 | 12.5 | 12 | 12.8 | 11.762 | 10.783 | 11.6688 | 9.99 | 13 | 15.8 | 11.27 | 13.8 | 12.41 | 13.53 | 12 | 12.2 | 12.94 | 12.5 | 12 | 11 |  | 1
| 1 | 2 | 13 | 13.5 | 11 | 12.55 | 11.55 | 11.6 | 11.2 | 12.65 | 12.4 | 12.1 | 12 | 13.4 | 11.5209 | 11.395 | 11.7082 | 10.59 | 13.3 | 15.9 | 10.98 | 14 | 12.47 | 13.41 | 13 | 12.8 | 12.93 | 12.19 | 12 | 11 |  | 37
| 1 | 3 | 12 | 13.8 | 11 | 12.9 | 13.1 | 11.4 | 11.75 | 13.65 | 12.4 | 12.9 | 11.9 | 13.1 | 11.4241 | 11.093 | 11.851 | 10.44 | 13 | 15.7 | 11.06 | 14.1 | 12.35 | 12.61 | 13 | 12.1 | 12.9 | 12.2 | 13 | 11 |  | 16
| 1 | 4 | 13 | 13.4 | 11.4 | 12.7 | 13 | 12 | 11.6 | 12.85 | 12.4 | 12.1 | 12 | 13 | 11.4177 | 10.918 | 11.8908 | 10.28 | 13.1 | 15.5 | 11.12 | 13.6 | 12.6 | 12.88 | 13 | 13 | 12.86 | 12.35 | 13 | 11 |  | 12.295
| 1 | 5 | 13 | 13.6 | 11 | 12.65 | 12.5 | 11.9 | 11.75 | 14.2 | 12.3 | 12.6 | 11.7 | 13.3 | 11.3291 | 11.241 | 11.9436 | 11 | 13.1 | 16.2 | 11.21 | 13.7 | 12.64 | 12.84 | 13 | 12 | 12.81 | 12.2 | 12 | 12 |  | 38
| 1 | 6 | 12 | 13 | 11.4 | 12.4 | 13.5 | 11.4 | 12.25 | 12.95 | 12.1 | 12.5 | 11.5 | 13.7 | 11.532 | 11.409 | 11.6876 | 10.66 | 13.1 | 15.8 | 11.08 | 13.6 | 12.32 | 12.85 | 13 | 12.6 | 12.95 | 12.29 | 12 | 11
| Mean |  | 12.6667 | 13.5667 | 11.2 | 12.5417 | 12.825 | 11.7333 | 11.6833 | 13.15 | 12.25 | 12.45 | 11.85 | 13.2167 | 11.4976 | 11.1398 | 11.7917 | 10.4933 | 13.1 | 15.8167 | 11.12 | 13.8 | 12.465 | 13.02 | 12.8333 | 12.45 | 12.8983 | 12.2883 | 12.3333 | 11.1667
| Median |  | 13 | 13.55 | 11.2 | 12.6 | 13.05 | 11.75 | 11.675 | 12.9 | 12.35 | 12.5 | 11.95 | 13.2 | 11.4725 | 11.167 | 11.7796 | 10.515 | 13.1 | 15.8 | 11.1 | 13.75 | 12.44 | 12.865 | 13 | 12.4 | 12.915 | 12.245 | 12 | 11
| Std Dev. |  | 0.5164 | 0.37238 | 0.21909 | 0.29226 | 0.70975 | 0.30768 | 0.34303 | 0.63797 | 0.20736 | 0.30822 | 0.20736 | 0.31885 | 0.14951 | 0.25555 | 0.11777 | 0.34512 | 0.10954 | 0.23166 | 0.10526 | 0.20976 | 0.13126 | 0.36365 | 0.40825 | 0.40866 | 0.054191 | 0.12156 | 0.5164 | 0.40825
| Rel.Std.Dev. |  | 4.077% | 2.745% | 1.956% | 2.33% | 5.534% | 2.622% | 2.936% | 4.851% | 1.693% | 2.476% | 1.75% | 2.413% | 1.3% | 2.294% | 0.9987% | 3.289% | 0.8362% | 1.465% | 0.9466% | 1.52% | 1.053% | 2.793% | 3.181% | 3.282% | 0.4201% | 0.9892% | 4.187% | 3.656%
| PDM3 |  | 3.023% | 10.34% | -8.906% | 2.006% | 4.31% | -4.569% | -4.975% | 6.954% | -0.3663% | 1.26% | -3.62% | 7.496% | -6.486% | -9.396% | -4.094% | -14.65% | 6.547% | 28.64% | -9.557% | 12.24% | 1.382% | 5.896% | 4.378% | 1.26% | 4.907% | -0.05451% | 0.3115% | -9.177%
| Z-Score (Absolute) |  | NA | 1.11 | 1.24 | 0.09 | 0.37 | 0.71 | 0.76 | 0.69 | 0.2 | 0 | 0.6 | 0.76 | 0.95 | 1.3 | 0.65 | 1.94 | 0.65 | 3.34 | 1.32 | 1.34 | 0.01 | 0.57 | NA | 0 | 0.45 | 0.16 | NA | NA
| Analytical results for Na in OREAS 625 (Certified Value 1.91 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.8 | 2.11 | 1.913 | 1.93 | 1.88 | 1.88 | 2 | 2 | 1.84 | 1.8681 | 1.978 | 1.9731 | 1.9075 | 1.8498 | 1.86 | 1.87 | 1.87 | 1.8833 | 1.8632 | 2.03 | 1.8798 | 1.895 | 1.86 | 1.9 | 1.97 | 2.02 | 1.989 | 1.96 |  | 1
| 1 | 2 | 1.82 | 2.09 | 1.804 | 1.97 | 1.83 | 1.86 | 1.91 | 2.07 | 1.83 | 1.8677 | 1.971 | 1.9879 | 1.8632 | 1.8542 | 1.87 | 1.87 | 1.86 | 1.7283 | 1.9138 | 2 | 1.8564 | 1.909 | 1.96 | 1.91 | 1.87 | 1.99 | 1.986 | 1.94 |  | 13
| 1 | 3 | 1.82 | 2 | 1.938 | 1.96 | 1.87 | 1.86 | 1.91 | 2.13 | 1.92 | 1.8683 | 1.93 | 1.9469 | 1.856 | 1.8641 | 1.86 | 1.92 | 1.89 | 1.8524 | 1.9069 | 1.99 | 1.9186 | 1.893 | 1.91 | 1.9 | 1.94 | 1.96 | 1.966 | 1.95 |  | 16
| 1 | 4 | 1.83 | 2.07 | 1.934 | 1.97 | 1.88 | 1.9 | 1.94 | 2.07 | 1.95 | 1.8508 | 1.968 | 1.95 | 1.8295 | 1.8743 | 1.85 | 1.88 | 1.9 | 1.8762 | 1.9129 | 2.08 | 1.8234 | 1.903 | 1.93 | 1.92 | 1.94 | 1.96 | 1.963 | 1.98 |  | 1.9132
| 1 | 5 | 1.79 | 2.15 | 1.898 | 1.95 | 1.86 | 1.9 | 1.9 | 2.18 | 1.78 | 1.8735 | 1.931 | 1.9139 | 1.8054 | 1.8797 | 1.85 | 1.95 | 1.9 | 1.8554 | 1.943 | 2.11 | 1.8797 | 1.9 | 1.89 | 1.9 | 1.94 | 1.95 | 1.982 | 1.99 |  | 39
| 1 | 6 | 1.81 | 1.92 | 1.909 | 1.95 | 1.9 | 1.86 | 1.97 | 1.96 | 1.85 | 1.8666 | 1.907 | 1.9343 | 1.8407 | 1.8795 | 1.85 | 1.92 | 1.92 | 1.8645 | 1.8956 | 2.03 | 1.915 | 1.868 | 1.97 | 1.9 | 1.97 | 2.05 | 1.962 | 1.96
| Mean |  | 1.8117 | 2.0567 | 1.8993 | 1.955 | 1.87 | 1.8767 | 1.9383 | 2.0683 | 1.8617 | 1.8658 | 1.9475 | 1.951 | 1.8504 | 1.8669 | 1.8567 | 1.9017 | 1.89 | 1.8434 | 1.9059 | 2.04 | 1.8788 | 1.8947 | 1.92 | 1.905 | 1.9383 | 1.9883 | 1.9747 | 1.9633
| Median |  | 1.815 | 2.08 | 1.911 | 1.955 | 1.875 | 1.87 | 1.925 | 2.07 | 1.845 | 1.8679 | 1.9495 | 1.9485 | 1.8483 | 1.8692 | 1.855 | 1.9 | 1.895 | 1.8599 | 1.9099 | 2.03 | 1.8798 | 1.8975 | 1.92 | 1.9 | 1.94 | 1.975 | 1.974 | 1.96
| Std Dev. |  | 0.01472 | 0.083347 | 0.049127 | 0.015166 | 0.023664 | 0.019664 | 0.039707 | 0.08085 | 0.062423 | 0.0077469 | 0.028711 | 0.026545 | 0.034676 | 0.012948 | 0.008165 | 0.033116 | 0.021909 | 0.057595 | 0.026147 | 0.046476 | 0.035964 | 0.014264 | 0.041952 | 0.0083666 | 0.03656 | 0.039707 | 0.012323 | 0.018619
| Rel.Std.Dev. |  | 0.8125% | 4.053% | 2.587% | 0.7757% | 1.265% | 1.048% | 2.049% | 3.909% | 3.353% | 0.4152% | 1.474% | 1.361% | 1.874% | 0.6935% | 0.4398% | 1.741% | 1.159% | 3.124% | 1.372% | 2.278% | 1.914% | 0.7529% | 2.185% | 0.4392% | 1.886% | 1.997% | 0.6241% | 0.9483%
| PDM3 |  | -5.308% | 7.498% | -0.7256% | 2.184% | -2.259% | -1.91% | 1.313% | 8.108% | -2.694% | -2.477% | 1.792% | 1.976% | -3.284% | -2.42% | -2.956% | -0.6036% | -1.213% | -3.652% | -0.3823% | 6.627% | -1.798% | -0.9695% | 0.3547% | -0.4294% | 1.313% | 3.926% | 3.212% | 2.62%
| Z-Score (Absolute) |  | 1.56 | 2.61 | 0.07 | 0.88 | 0.57 | 0.45 | 0.6 | 2.81 | 0.71 | 0.64 | 0.75 | 0.81 | 0.9 | 0.62 | 0.79 | 0.03 | 0.23 | 1.02 | 0.04 | 2.33 | 0.42 | 0.15 | 0.28 | 0.03 | 0.6 | 1.45 | 1.22 | 1.02
| Analytical results for Nb in OREAS 625 (Certified Value 14.8 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 9 | 16.6 | 15 | 14.4 | 17.1 | 16.1 | 14.2 | 14.8 | 12.9 | 15.35 | 12.4 | 15.62 | 14.0698 | 15.2995 | 14.5 | 15 | 11.72 | 15 | 15.7 | 12 | 13.7 | 13.6 | 15.1 | 16.8 | [5 empty] | 1
| 1 | 2 | 9 | 16.5 | 14.2 | 14.4 | 15.3 | 15.5 | 13.8 | 16 | 13.4 | 14.32 | 12.3 | 16.16 | 13.6883 | 15.2344 | 14.6 | 15.4 | 11.86 | 14.9 | 15.9 | 11 | 13.9 | 12.1 | 15.5 | 16.6 | [5 empty] | 38
| 1 | 3 | 9 | 16.3 | 14.5 | 14.6 | 17.1 | 15.8 | 14.2 | 16.2 | 13.4 | 15.64 | 13 | 15.74 | 13.8106 | 15.2442 | 15.1 | 14.9 | 11.95 | 14.8 | 15.5 | 13 | 14.2 | 14.4 | 15.1 | 16.6 | [5 empty] | 16
| 1 | 4 | 9 | 16.1 | 14.7 | 15.2 | 17.6 | 16.1 | 13.9 | 15.6 | 13.3 | 14.85 | 11.9 | 16.03 | 14.2171 | 15.2567 | 14.8 | 15.5 | 11.86 | 14.5 | 14.8 | 12 | 14.8 | 13.8 | 15.2 | 17 | [5 empty] | 14.7794
| 1 | 5 | 8 | 16.5 | 13.8 | 14.1 | 16.9 | 16.2 | 13.9 | 16.7 | 13 | 14.89 | 11.7 | 16.06 | 13.6584 | 15.2907 | 14.5 | 15.4 | 11 | 14.9 | 14.8 | 12 | 13.3 | 13.4 | 15.4 | 17.3 | [5 empty] | 40
| 1 | 6 | 9 | 16.2 | 14.2 | 14.3 | 17.4 | 15.1 | 14.9 | 15.2 | 12.9 | 14.47 | 11.8 | 15.8 | 13.6346 | 15.2992 | 14.7 | 15.6 | 11.53 | 14.9 | 15.4 | 12 | 14.4 | 13.6 | 15.3 | 16.4
| Mean |  | 8.8333 | 16.3667 | 14.4 | 14.5 | 16.9 | 15.8 | 14.15 | 15.75 | 13.15 | 14.92 | 12.1833 | 15.9017 | 13.8465 | 15.2708 | 14.7 | 15.3 | 11.6533 | 14.8333 | 15.35 | 12 | 14.05 | 13.4833 | 15.2667 | 16.7833
| Median |  | 9 | 16.4 | 14.35 | 14.4 | 17.1 | 15.95 | 14.05 | 15.8 | 13.15 | 14.87 | 12.1 | 15.915 | 13.7494 | 15.2737 | 14.65 | 15.4 | 11.79 | 14.9 | 15.45 | 12 | 14.05 | 13.6 | 15.25 | 16.7
| Std Dev. |  | 0.40825 | 0.19664 | 0.42426 | 0.37947 | 0.82219 | 0.42895 | 0.40373 | 0.6921 | 0.2429 | 0.5043 | 0.48751 | 0.2117 | 0.24239 | 0.029193 | 0.22804 | 0.28284 | 0.35212 | 0.17512 | 0.45935 | 0.63246 | 0.53198 | 0.76004 | 0.1633 | 0.32506
| Rel.Std.Dev. |  | 4.622% | 1.201% | 2.946% | 2.617% | 4.865% | 2.715% | 2.853% | 4.394% | 1.847% | 3.38% | 4.001% | 1.331% | 1.751% | 0.1912% | 1.551% | 1.849% | 3.022% | 1.181% | 2.992% | 5.27% | 3.786% | 5.637% | 1.07% | 1.937%
| PDM3 |  | -40.23% | 10.74% | -2.567% | -1.89% | 14.35% | 6.906% | -4.258% | 6.567% | -11.02% | 0.9515% | -17.57% | 7.594% | -6.312% | 3.325% | -0.5371% | 3.523% | -21.15% | 0.3651% | 3.861% | -18.81% | -4.935% | -8.769% | 3.297% | 13.56%
| Z-Score (Absolute) |  | NA | 1.18 | 0.38 | 0.3 | 1.61 | 0.73 | 0.58 | 0.69 | 1.37 | 0.03 | 2.14 | 0.81 | 0.82 | 0.31 | 0.14 | 0.34 | 2.56 | 0.03 | 0.38 | NA | 0.66 | 1.11 | 0.31 | 1.51
Results from laboratories 1.01 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 625 (Certified Value 37.3 ppm) | [30 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | [5 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | [5 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | [5 empty] | 1
2
| 1 | 1 | 1.01: 34.4; 1.02: 37.21; 1.03: 38.27; 1.04: 34.9769; 1.05: 44.561; 1.06: 35.9 | [5 empty] | 1
| 1 | 2 | 1.01: 33.6; 1.02: 36.5; 1.03: 38.57; 1.04: 32.9256; 1.05: 44.2012; 1.06: 35.3 | [5 empty] | 8
| 1 | 3 | 1.01: 33.9; 1.02: 37.04; 1.03: 39.18; 1.04: 34.2939; 1.05: 44.2124; 1.06: 35.6 | [5 empty] | 16
| 1 | 4 | 1.01: 34.2; 1.02: 36.43; 1.03: 39.52; 1.04: 33.6865; 1.05: 44.8676; 1.06: 34.9 | [5 empty] | 37.3294
| 1 | 5 | 1.01: 34.6; 1.02: 36.78; 1.03: 39.39; 1.04: 33.0582; 1.05: 44.1601; 1.06: 36.5 | [5 empty] | 41
| 1 | 6 | 1.01: 29.4; 1.02: 36.3; 1.03: 39.57; 1.04: 34.2246; 1.05: 44.6888; 1.06: 36.2
| Mean |  | 1.01: 33.35; 1.02: 36.71; 1.03: 39.0833; 1.04: 33.861; 1.05: 44.4485; 1.06: 35.7333
| Median |  | 1.01: 34.05; 1.02: 36.64; 1.03: 39.285; 1.04: 33.9555; 1.05: 44.3867; 1.06: 35.75
| Std Dev. |  | 1.01: 1.9675; 1.02: 0.36177; 1.03: 0.53962; 1.04: 0.78925; 1.05: 0.29869; 1.06: 0.58878
| Rel.Std.Dev. |  | 1.01: 5.9%; 1.02: 0.9855%; 1.03: 1.381%; 1.04: 2.331%; 1.05: 0.672%; 1.06: 1.648%
| PDM3 |  | 1.01: -10.66%; 1.02: -1.659%; 1.03: 4.699%; 1.04: -9.291%; 1.05: 19.07%; 1.06: -4.276%
| Z-Score (Absolute) |  | 1.01: 0.74; 1.02: 0.13; 1.03: 0.74; 1.04: 0.61; 1.05: 2.12; 1.06: 0.13
| Analytical results for Ni in OREAS 625 (Certified Value 5.98 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 7 | 6.1 | 4.5 | 5.3 | 6.5 | 6 | 5.7 | 5.8 | 6.4 | 6 | 6.6 | 5.8 | 5.7306 | 5.215 | 7.0493 | 6.42 | 5 | 6.8 | 6.34 | 6.3 | 5.8 | 8.48 | 6 | 6.9 | 6.2 | 5 | 5 | 5 |  | 1
| 1 | 2 | 7 | 6 | 4.7 | 5.1 | 5.8 | 5.8 | 5.3 | 6.3 | 6.6 | 5.7 | 6 | 6.1 | 5.6966 | 5.212 | 6.3765 | 8.58 | 5 | 5.4 | 5.82 | 5.9 | 5.8 | 8.84 | 6 | 6.8 | 6.5 | 5.3 | 5 | 7 |  | 14
| 1 | 3 | 6 | 6 | 4.7 | 7.1 | 6.6 | 5.6 | 5.5 | 6.5 | 6.8 | 6 | 5.8 | 6.2 | 5.6411 | 4.926 | 6.4602 | 7.1 | 5 | 5.8 | 6.25 | 5.9 | 5.8 | 8.67 | 6 | 7 | 6.3 | 5.4 | 5 | 5 |  | 16
| 1 | 4 | 6 | 5.8 | 4.7 | 5.7 | 6.6 | 6 | 5.6 | 5.9 | 6.3 | 5.9 | 6 | 6.4 | 5.292 | 5.02 | 6.1154 | 7.99 | 5 | 6.5 | 6.18 | 5.7 | 6.1 | 8 | 6 | 7.3 | 6.1 | 5.3 | 6 | 6 |  | 5.9846
| 1 | 5 | 7 | 5.9 | 4.6 | 5.3 | 6.2 | 5.9 | 5.5 | 6.7 | 6.3 | 5.6 | 6.6 | 6.1 | 5.3396 | 4.786 | 6.7345 | 9.13 | 5 | 6.6 | 6.14 | 5.9 | 6.1 | 7.82 | 6 | 7.2 | 5.8 | 5.1 | 5 | 6 |  | 42
| 1 | 6 | 7 | 5.7 | 4.7 | 5.5 | 6.5 | 5.6 | 5.7 | 6 | 6.2 | 5.7 | 5.8 | 6.2 | 5.7975 | 4.99 | 6.7511 | 6.82 | 5 | 5.5 | 6.31 | 5.7 | 5.7 | 8.04 | 6 | 7.1 | 6.7 | 5.1 | 5 | 6
| Mean |  | 6.6667 | 5.9167 | 4.65 | 5.6667 | 6.3667 | 5.8167 | 5.55 | 6.2 | 6.4333 | 5.8167 | 6.1333 | 6.1333 | 5.5829 | 5.0248 | 6.5812 | 7.6733 | 5 | 6.1 | 6.1733 | 5.9 | 5.8833 | 8.3083 | 6 | 7.05 | 6.2667 | 5.2 | 5.1667 | 5.8333
| Median |  | 7 | 5.95 | 4.7 | 5.4 | 6.5 | 5.85 | 5.55 | 6.15 | 6.35 | 5.8 | 6 | 6.15 | 5.6689 | 5.005 | 6.5973 | 7.545 | 5 | 6.15 | 6.215 | 5.9 | 5.8 | 8.26 | 6 | 7.05 | 6.25 | 5.2 | 5 | 6
| Std Dev. |  | 0.5164 | 0.1472 | 0.083666 | 0.73121 | 0.31411 | 0.18348 | 0.15166 | 0.35777 | 0.22509 | 0.17224 | 0.37238 | 0.19664 | 0.21357 | 0.16688 | 0.33013 | 1.0651 | 0 | 0.60663 | 0.18886 | 0.21909 | 0.17224 | 0.41194 | 0 | 0.18708 | 0.31411 | 0.15492 | 0.40825 | 0.75277
| Rel.Std.Dev. |  | 7.746% | 2.488% | 1.799% | 12.9% | 4.934% | 3.154% | 2.733% | 5.77% | 3.499% | 2.961% | 6.071% | 3.206% | 3.826% | 3.321% | 5.016% | 13.88% | 0% | 9.945% | 3.059% | 3.713% | 2.928% | 4.958% | 0% | 2.654% | 5.012% | 2.979% | 7.902% | 12.9%
| PDM3 |  | 11.4% | -1.136% | -22.3% | -5.313% | 6.383% | -2.807% | -7.263% | 3.598% | 7.497% | -2.807% | 2.485% | 2.485% | -6.713% | -16.04% | 9.967% | 28.22% | -16.45% | 1.928% | 3.153% | -1.414% | -1.693% | 38.83% | 0.2566% | 17.8% | 4.712% | -13.11% | -13.67% | -2.528%
| Z-Score (Absolute) |  | NA | 0.44 | 3.45 | 1.03 | 0.63 | 0.67 | 1.31 | 0.24 | 0.79 | 0.67 | 0.08 | 0.08 | 1.23 | 2.56 | 1.15 | 3.74 | NA | 0 | 0.17 | 0.48 | 0.52 | 5.26 | NA | 2.26 | 0.4 | 2.14 | NA | NA
Results from laboratories 1.01, 1.17, 1.23 and 1.28 were removed due to their 1 ppm reading resolution. Results from laboratory 1.30 were removed due to their 10 ppm reading resolution.
| Analytical results for P in OREAS 625 (Certified Value 0.03 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.0342 | 0.031 | 0.031 | 0.027 | 0.029 | 0.031 | 0.031 | 0.0305 | 0.0279 | 0.03 | 0.0302 | 0.029935 | 0.029562 | 0.03 | 0.031 | 0.029 | 0.0261 | 0.03 | 0.03 | 0.03 | 0.026 | 0.03 | 0.0291 | 0.0305 | 0.028 | 0.03 |  |  |  | 1
| 1 | 2 | 0.0349 | 0.029 | 0.031 | 0.028 | 0.028 | 0.03 | 0.031 | 0.0289 | 0.0282 | 0.03 | 0.0299 | 0.029334 | 0.029796 | 0.03 | 0.031 | 0.029 | 0.0254 | 0.0309 | 0.03 | 0.03 | 0.028 | 0.03 | 0.0284 | 0.0303 | 0.026 | 0.03 |  |  |  | 24
| 1 | 3 | 0.0341 | 0.032 | 0.031 | 0.028 | 0.028 | 0.03 | 0.032 | 0.0313 | 0.0286 | 0.029 | 0.0295 | 0.029198 | 0.030126 | 0.03 | 0.03 | 0.029 | 0.0266 | 0.0306 | 0.03 | 0.03 | 0.026 | 0.03 | 0.029 | 0.0299 | 0.027 | 0.03 |  |  |  | 16
| 1 | 4 | 0.0365 | 0.032 | 0.031 | 0.028 | 0.029 | 0.03 | 0.032 | 0.0305 | 0.0282 | 0.03 | 0.0297 | 0.028637 | 0.029434 | 0.03 | 0.031 | 0.029 | 0.0298 | 0.0302 | 0.03 | 0.03 | 0.027 | 0.03 | 0.0279 | 0.0299 | 0.027 | 0.03 |  |  |  | 0.029731
| 1 | 5 | 0.0356 | 0.031 | 0.03 | 0.028 | 0.029 | 0.03 | 0.034 | 0.0298 | 0.0283 | 0.029 | 0.0295 | 0.028337 | 0.030185 | 0.03 | 0.03 | 0.029 | 0.0288 | 0.0314 | 0.03 | 0.03 | 0.027 | 0.03 | 0.0285 | 0.0294 | 0.027 | 0.03 |  |  |  | 43
| 1 | 6 | 0.0362 | 0.031 | 0.031 | 0.029 | 0.028 | 0.031 | 0.031 | 0.0298 | 0.0282 | 0.029 | 0.0301 | 0.029015 | 0.030249 | 0.03 | 0.03 | 0.03 | 0.0278 | 0.0304 | 0.03 | 0.03 | 0.027 | 0.03 | 0.028 | 0.0307 | 0.029 | 0.03
| Mean |  | 0.03525 | 0.031 | 0.030833 | 0.028 | 0.0285 | 0.030333 | 0.031833 | 0.030133 | 0.028233 | 0.0295 | 0.029817 | 0.029076 | 0.029892 | 0.03 | 0.0305 | 0.029167 | 0.027417 | 0.030583 | 0.03 | 0.03 | 0.026833 | 0.03 | 0.028483 | 0.030117 | 0.027333 | 0.03
| Median |  | 0.03525 | 0.031 | 0.031 | 0.028 | 0.0285 | 0.03 | 0.0315 | 0.03015 | 0.0282 | 0.0295 | 0.0298 | 0.029106 | 0.029961 | 0.03 | 0.0305 | 0.029 | 0.0272 | 0.0305 | 0.03 | 0.03 | 0.027 | 0.03 | 0.02845 | 0.0301 | 0.027 | 0.03
| Std Dev. |  | 0.0010134 | 0.0010954 | 0.00040825 | 0.00063246 | 0.00054772 | 0.0005164 | 0.001169 | 0.00082138 | 0.00022509 | 0.00054772 | 0.00029944 | 0.00055879 | 0.00034506 | 0 | 0.00054772 | 0.00040825 | 0.0016857 | 0.00050761 | 0 | 0 | 0.00075277 | 0 | 0.00049565 | 0.00047504 | 0.0010328 | 0
| Rel.Std.Dev. |  | 2.875% | 3.534% | 1.324% | 2.259% | 1.922% | 1.702% | 3.672% | 2.726% | 0.7973% | 1.857% | 1.004% | 1.922% | 1.154% | 0% | 1.796% | 1.4% | 6.149% | 1.66% | 0% | 0% | 2.805% | 0% | 1.74% | 1.577% | 3.779% | 0%
| PDM3 |  | 18.56% | 4.268% | 3.707% | -5.823% | -4.141% | 2.025% | 7.07% | 1.353% | -5.038% | -0.7777% | 0.2874% | -2.204% | 0.5411% | 0.904% | 2.586% | -1.899% | -7.785% | 2.866% | 0.904% | 0.904% | -9.747% | 0.904% | -4.197% | 1.296% | -8.065% | 0.904%
| Z-Score (Absolute) |  | 5 | 0.95 | 0.79 | 1.9 | 1.43 | 0.32 | 1.75 | 0.13 | 1.68 | 0.48 | 0.17 | 0.88 | 0.1 | 0 | 0.48 | 0.79 | 2.46 | 0.56 | 0 | 0 | 3.01 | 0 | 1.44 | 0.11 | 2.54 | 0
| Analytical results for Pb in OREAS 625 (Certified Value 0.822 wt.%) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.7958 | 0.841 | 0.78795 | 0.825 | 0.822 | 0.798 | 0.822 | 0.868 | 0.743 | 0.86095 | 0.83719 | 0.8313 | 0.76892 | 0.78597 | 0.8275 | 0.829 | 0.83 | 0.713 | 0.8534 | 0.8295 | 0.8394 | 0.80186 | 0.8617 | 0.827 | 0.7565 | 0.8239 | 0.81136 | 0.802 | 0.81261 | 1
| 1 | 2 | 0.7943 | 0.843 | 0.75728 | 0.842 | 0.809 | 0.784 | 0.793 | 0.859 | 0.732 | 0.81739 | 0.83912 | 0.83 | 0.74385 | 0.84272 | 0.81361 | 0.8243 | 0.81 | 0.701 | 0.8513 | 0.8425 | 0.8383 | 0.80559 | 0.8327 | 0.8289 | 0.7565 | 0.8338 | 0.82163 | 0.8195 | 0.80636 | 25
| 1 | 3 | 0.7914 | 0.838 | 0.81847 | 0.836 | 0.819 | 0.795 | 0.796 | 0.872 | 0.724 | 0.84535 | 0.8215 | 0.8114 | 0.75636 | 0.80981 | 0.82118 | 0.8261 | 0.82 | 0.697 | 0.8502 | 0.8256 | 0.8354 | 0.80154 | 0.8628 | 0.8234 | 0.7532 | 0.8297 | 0.80714 | 0.8099 | 0.81544 | 16
| 1 | 4 | 0.7869 | 0.827 | 0.81046 | 0.837 | 0.819 | 0.803 | 0.802 | 0.877 | 0.722 | 0.83428 | 0.82868 | 0.8211 | 0.76375 | 0.81727 | 0.83439 | 0.8125 | 0.83 | 0.715 | 0.8505 | 0.8284 | 0.8343 | 0.80495 | 0.7949 | 0.8215 | 0.7565 | 0.8201 | 0.79971 | 0.8141 | 0.82651 | 0.82207
| 1 | 5 | 0.7921 | 0.826 | 0.79273 | 0.831 | 0.814 | 0.809 | 0.792 | 0.935 | 0.758 | 0.82287 | 0.82214 | 0.813 | 0.74915 | 0.80355 | 0.80491 | 0.8257 | 0.8 | 0.72 | 0.8511 | 0.7994 | 0.8428 | 0.80801 | 0.851 | 0.8234 | 0.7526 | 0.8234 | 0.80584 | 0.8081 | 0.82053 | 44
| 1 | 6 | 0.7992 | 0.832 | 0.79829 | 0.831 | 0.829 | 0.794 | 0.815 | 0.848 | 0.704 | 0.8256 | 0.81888 | 0.8246 | 0.76701 | 0.83593 | 0.81455 | 0.8252 | 0.83 | 0.719 | 0.8506 | 0.8345 | 0.8318 | 0.80911 | 0.802 | 0.8277 | 0.7569 | 0.8218 | 0.80627 | 0.8012 | 0.83115
| Mean |  | 0.79328 | 0.8345 | 0.7942 | 0.83367 | 0.81867 | 0.79717 | 0.80333 | 0.8765 | 0.7305 | 0.83441 | 0.82792 | 0.8219 | 0.75817 | 0.81587 | 0.81936 | 0.8238 | 0.82 | 0.71083 | 0.85118 | 0.82665 | 0.837 | 0.80518 | 0.83418 | 0.82532 | 0.75537 | 0.82545 | 0.80866 | 0.80913 | 0.81877
| Median |  | 0.7932 | 0.835 | 0.79551 | 0.8335 | 0.819 | 0.7965 | 0.799 | 0.87 | 0.728 | 0.82994 | 0.82541 | 0.82285 | 0.76005 | 0.81354 | 0.81786 | 0.82545 | 0.825 | 0.714 | 0.85085 | 0.82895 | 0.83685 | 0.80527 | 0.84185 | 0.8252 | 0.7565 | 0.82365 | 0.80671 | 0.809 | 0.81798
| Std Dev. |  | 0.0041969 | 0.0072319 | 0.021328 | 0.0059217 | 0.0068313 | 0.0085186 | 0.012453 | 0.030442 | 0.018609 | 0.016276 | 0.0085825 | 0.0083828 | 0.010145 | 0.021014 | 0.010588 | 0.0057592 | 0.012649 | 0.0096003 | 0.0011583 | 0.014613 | 0.0039492 | 0.0030967 | 0.029799 | 0.0029417 | 0.0019263 | 0.0052218 | 0.0073701 | 0.0070372 | 0.0091576
| Rel.Std.Dev. |  | 0.529% | 0.8666% | 2.686% | 0.7103% | 0.8344% | 1.069% | 1.55% | 3.473% | 2.547% | 1.951% | 1.037% | 1.02% | 1.338% | 2.576% | 1.292% | 0.6991% | 1.543% | 1.351% | 0.1361% | 1.768% | 0.4718% | 0.3846% | 3.572% | 0.3564% | 0.255% | 0.6326% | 0.9114% | 0.8697% | 1.118%
| PDM3 |  | -3.502% | 1.512% | -3.391% | 1.411% | -0.4141% | -3.029% | -2.279% | 6.621% | -11.14% | 1.501% | 0.7113% | -0.02079% | -7.773% | -0.7538% | -0.3299% | 0.2103% | -0.2519% | -13.53% | 3.541% | 0.557% | 1.816% | -2.055% | 1.473% | 0.3948% | -8.114% | 0.411% | -1.632% | -1.574% | -0.4019%
| Z-Score (Absolute) |  | 1.23 | 0.71 | 1.19 | 0.67 | 0.03 | 1.05 | 0.76 | 2.69 | 4.19 | 0.71 | 0.4 | 0.12 | 2.88 | 0.16 | 0 | 0.21 | 0.03 | 5.11 | 1.5 | 0.34 | 0.83 | 0.67 | 0.7 | 0.28 | 3.02 | 0.29 | 0.5 | 0.48 | 0.03
| Analytical results for Pr in OREAS 625 (Certified Value 9.74 ppm) | [30 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | 2
2
| 1 | 1 | 1.01: 8.5; 1.02: 9.37; 1.03: 10.08; 1.04: 9.2297; 1.05: 11.3371; 1.06: 9.83 | 1
| 1 | 2 | 1.01: 8.6; 1.02: 9.01; 1.03: 10.43; 1.04: 8.8454; 1.05: 11.4294; 1.06: 9.74 | 10
| 1 | 3 | 1.01: 8.6; 1.02: 9.35; 1.03: 10.55; 1.04: 9.1634; 1.05: 11.2564; 1.06: 9.76 | 16
| 1 | 4 | 1.01: 8.6; 1.02: 9.13; 1.03: 10.43; 1.04: 8.8834; 1.05: 11.0569; 1.06: 9.63 | 9.738
| 1 | 5 | 1.01: 8.7; 1.02: 9.17; 1.03: 10.48; 1.04: 8.8006; 1.05: 11.7624; 1.06: 10.06 | 45
| 1 | 6 | 1.01: 7.3; 1.02: 8.94; 1.03: 10.56; 1.04: 8.937; 1.05: 11.786; 1.06: 9.96
| Mean |  | 1.01: 8.3833; 1.02: 9.1617; 1.03: 10.4217; 1.04: 8.9766; 1.05: 11.4381; 1.06: 9.83
| Median |  | 1.01: 8.6; 1.02: 9.15; 1.03: 10.455; 1.04: 8.9102; 1.05: 11.3833; 1.06: 9.795
| Std Dev. |  | 1.01: 0.53448; 1.02: 0.1744; 1.03: 0.17657; 1.04: 0.17741; 1.05: 0.28805; 1.06: 0.15672
| Rel.Std.Dev. |  | 1.01: 6.375%; 1.02: 1.904%; 1.03: 1.694%; 1.04: 1.976%; 1.05: 2.518%; 1.06: 1.594%
| PDM3 |  | 1.01: -13.91%; 1.02: -5.918%; 1.03: 7.021%; 1.04: -7.819%; 1.05: 17.46%; 1.06: 0.9448%
| Z-Score (Absolute) |  | 1.01: 1.04; 1.02: 0.31; 1.03: 0.86; 1.04: 0.48; 1.05: 1.81; 1.06: 0.31
| Analytical results for Rb in OREAS 625 (Certified Value 135 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 0
0
| 1 | 1 | 130 | 128.9 | 136.5 | 140 | 134 | 124.5 | 131.5 | 137 | 138.65 | 139.5 | 133.96 | 135.0873 | 136.9952 | 143 | 138.2 | 145.6 | 134 | 135 | 163 | 137.8 | 125.3 | 135.6 | [7 empty] | 1
| 1 | 2 | 139 | 122.7 | 139.5 | 124 | 134 | 123 | 136.5 | 136 | 134.15 | 136.4 | 143.38 | 132.1887 | 137.8558 | 142 | 136.9 | 148.4 | 131 | 137 | 162 | 121.6 | 132.2 | 134.4 | [7 empty] | 43
| 1 | 3 | 136 | 123.7 | 139.5 | 140.5 | 116 | 122.5 | 138.5 | 139 | 139.9 | 136.5 | 144.11 | 132.2613 | 135.5363 | 143 | 132.7 | 149.3 | 132 | 133 | 177 | 141.4 | 126.3 | 136.4 | [7 empty] | 16
| 1 | 4 | 129 | 124.4 | 141.5 | 146 | 132.5 | 124 | 135.5 | 137 | 137.99 | 136.4 | 141.9 | 133.0503 | 135.2817 | 143 | 139.2 | 143.3 | 131 | 137 | 175 | 137.1 | 131.1 | 138.6 | [7 empty] | 135.319
| 1 | 5 | 141 | 124.5 | 135 | 140.5 | 126 | 123.5 | 140.5 | 139 | 134.55 | 134.6 | 141.93 | 129.9536 | 135.8834 | 141 | 142.1 | 145.5 | 135 | 135 | 174 | 130.1 | 127.2 | 137.9 | [7 empty] | 46
| 1 | 6 | 131 | 120.2 | 137.5 | 138.5 | 124 | 134 | 133 | 133 | 135.46 | 133.1 | 143.15 | 129.219 | 137.1191 | 143 | 142.5 | 144 | 130 | 135 | 175 | 132.7 | 128.3 | 138.9
| Mean |  | 134.3333 | 124.0667 | 138.25 | 138.25 | 127.75 | 125.25 | 135.9167 | 136.8333 | 136.7833 | 136.0833 | 141.405 | 131.96 | 136.4453 | 142.5 | 138.6 | 146.0167 | 132.1667 | 135.3333 | 171 | 133.45 | 128.4 | 136.9667
| Median |  | 133.5 | 124.05 | 138.5 | 140.25 | 129.25 | 123.75 | 136 | 137 | 136.725 | 136.4 | 142.54 | 132.225 | 136.4393 | 143 | 138.7 | 145.55 | 131.5 | 135 | 174.5 | 134.9 | 127.75 | 137.15
| Std Dev. |  | 5.0465 | 2.85 | 2.3611 | 7.4347 | 7.1537 | 4.3445 | 3.3529 | 2.2286 | 2.3802 | 2.1517 | 3.7473 | 2.1284 | 1.0239 | 0.83666 | 3.6244 | 2.3811 | 1.9408 | 1.5055 | 6.6633 | 7.0378 | 2.7276 | 1.7918
| Rel.Std.Dev. |  | 3.757% | 2.297% | 1.708% | 5.378% | 5.6% | 3.469% | 2.467% | 1.629% | 1.74% | 1.581% | 2.65% | 1.613% | 0.7504% | 0.5871% | 2.615% | 1.631% | 1.468% | 1.112% | 3.897% | 5.274% | 2.124% | 1.308%
| PDM3 |  | -0.7284% | -8.315% | 2.166% | 2.166% | -5.593% | -7.441% | 0.4417% | 1.119% | 1.082% | 0.5648% | 4.498% | -2.482% | 0.8323% | 5.307% | 2.425% | 7.906% | -2.33% | 0.01058% | 26.37% | -1.381% | -5.113% | 1.218%
| Z-Score (Absolute) |  | 0.51 | 3.19 | 0.52 | 0.52 | 2.23 | 2.88 | 0.09 | 0.15 | 0.14 | 0.05 | 1.35 | 1.13 | 0.05 | 1.63 | 0.61 | 2.55 | 1.07 | 0.24 | 9.1 | 0.74 | 2.06 | 0.18
| Analytical results for Re in OREAS 625 (Certified Value 0.004 ppm) | [30 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 3
3
| 1 | 1 | 0.004 | < 0.005 | 0.003 | 0.003 | 0.005 | 0.005 | 0.004 | 0.004 | 0.003 | < 0.005 | 0.004 | < 0.01 | < 0.05 | < 0.05 | 0.006 | 0.004 | 0.004 | < 0.002 | [11 empty] | 1
| 1 | 2 | 0.003 | < 0.005 | 0.005 | 0.003 | 0.003 | 0.003 | 0.006 | 0.005 | 0.003 | < 0.005 | 0.006 | < 0.01 | < 0.05 | < 0.05 | 0.008 | 0.004 | 0.003 | < 0.002 | [11 empty] | 15
| 1 | 3 | 0.005 | < 0.005 | 0.004 | 0.004 | 0.004 | 0.002 | 0.005 | 0.005 | 0.006 | < 0.005 | 0.006 | < 0.01 | < 0.05 | < 0.05 | 0.007 | 0.005 | 0.004 | < 0.002 | [11 empty] | 16
| 1 | 4 | 0.003 | < 0.005 | 0.003 | 0.004 | 0.004 | 0.003 | 0.004 | 0.005 | 0.004 | < 0.005 | 0.006 | < 0.01 | < 0.05 | < 0.05 | 0.009 | 0.005 | 0.004 | 0.003 | [11 empty] | 0.0039861
| 1 | 5 | 0.003 | < 0.005 | 0.003 | 0.004 | 0.004 | 0.004 | 0.006 | 0.005 | 0.002 | < 0.005 | 0.002 | < 0.01 | < 0.05 | < 0.05 | 0.008 | 0.005 | 0.003 | < 0.002 | [11 empty] | 47
| 1 | 6 | 0.003 | < 0.005 | 0.007 | 0.004 | 0.004 | 0.003 | 0.005 | 0.005 | 0.003 | < 0.005 | 0.005 | < 0.01 | < 0.05 | < 0.05 | 0.008 | 0.006 | 0.004 | < 0.002
| Mean |  | 0.0035 |  | 0.0041667 | 0.0036667 | 0.004 | 0.0033333 | 0.005 | 0.0048333 | 0.0035 |  | 0.0048333 |  |  |  | 0.0076667 | 0.0048333 | 0.0036667 | 0.003
| Median |  | 0.003 |  | 0.0035 | 0.004 | 0.004 | 0.003 | 0.005 | 0.005 | 0.003 |  | 0.0055 |  |  |  | 0.008 | 0.005 | 0.004 | 0.003
| Std Dev. |  | 0.00083666 |  | 0.0016021 | 0.0005164 | 0.00063246 | 0.0010328 | 0.00089443 | 0.00040825 | 0.0013784 |  | 0.0016021 |  |  |  | 0.0010328 | 0.00075277 | 0.0005164
| Rel.Std.Dev. |  | 23.9% |  | 38.45% | 14.08% | 15.81% | 30.98% | 17.89% | 8.447% | 39.38% |  | 33.15% |  |  |  | 13.47% | 15.57% | 14.08%
| PDM3 |  | -12.2% |  | 4.53% | -8.014% | 0.3484% | -16.38% | 25.44% | 21.25% | -12.2% |  | 21.25% |  |  |  | 92.33% | 21.25% | -8.014% | -24.74%
| Z-Score (Absolute) |  | 0.52 | 1.42 | 0.07 | 0.37 | 0.07 | 0.67 | 0.82 | 0.67 | 0.52 | 1.42 | 0.67 | 0.82 | 18.81 | 18.81 | 3.22 | 0.67 | 0.37 | 2.47
| Analytical results for S in OREAS 625 (Certified Value 3.8 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 2
3
| 1 | 1 | 4.04 | 3.8 | 3.88 | 3.78 | 3.74 | 3.95 | 4.01 | 3.62 | 3.696 | 3.8 | 3.61 | > 1 | 3.718 | 3.76 | 3.67 | 3.73 | 3.856 | 3.89 | 3.8536 | 3.738 | 3.4 | 3.74 | 3.81 | 3.88 | 3.925 | 4.2 |  |  |  | 1
| 1 | 2 | 4.2 | 3.7 | 4 | 3.75 | 3.68 | 3.8 | 3.93 | 3.47 | 3.6531 | 3.8 | 3.86 | > 1 | 3.7015 | 4.13 | 3.64 | 3.69 | 4.0629 | 3.82 | 3.8215 | 3.751 | 3.5 | 3.72 | 3.77 | 3.87 | 3.907 | 4.17 |  |  |  | 26
| 1 | 3 | 4.11 | 4 | 3.95 | 3.78 | 3.67 | 3.8 | 3.92 | 3.67 | 3.6267 | 3.6 | 3.74 | > 1 | 3.7993 | 3.87 | 3.69 | 3.72 | 3.8629 | 3.84 | 3.9214 | 3.753 | 3.5 | 3.73 | 3.85 | 3.89 | 3.879 | 4.19 |  |  |  | 16
| 1 | 4 | 4.2 | 3.9 | 3.99 | 3.79 | 3.75 | 3.83 | 3.98 | 3.61 | 3.6328 | 3.7 | 3.87 | > 1 | 3.7336 | 3.89 | 3.66 | 3.79 | 3.9663 | 3.94 | 3.8933 | 3.78 | 3.5 | 3.8 | 3.82 | 3.83 | 3.904 | 4.21 |  |  |  | 3.8039
| 1 | 5 | 4.05 | 3.9 | 3.9 | 3.78 | 3.76 | 3.78 | 4.3 | 3.54 | 3.6579 | 3.7 | 3.81 | > 1 | 3.6589 | 3.71 | 3.65 | 3.72 | 4.1022 | 4.01 | 3.8934 | 3.736 | 3.5 | 3.76 | 3.86 | 3.88 | 3.9 | 4.2 |  |  |  | 48
| 1 | 6 | 4.16 | 3.9 | 3.92 | 3.87 | 3.71 | 3.9 | 3.93 | 3.55 | 3.6296 | 3.7 | 3.85 | > 1 | 3.834 | 3.57 | 3.65 | 3.81 | 3.988 | 3.88 | 3.8945 | 3.719 | 3.5 | 3.74 | 3.82 | 3.91 | 3.858 | 4.16
| Mean |  | 4.1267 | 3.8667 | 3.94 | 3.7917 | 3.7183 | 3.8433 | 4.0117 | 3.5767 | 3.6494 | 3.7167 | 3.79 |  | 3.7409 | 3.8217 | 3.66 | 3.7433 | 3.9731 | 3.8967 | 3.8796 | 3.7462 | 3.4833 | 3.7483 | 3.8217 | 3.8767 | 3.8955 | 4.1883
| Median |  | 4.135 | 3.9 | 3.935 | 3.78 | 3.725 | 3.815 | 3.955 | 3.58 | 3.6429 | 3.7 | 3.83 |  | 3.7258 | 3.815 | 3.655 | 3.725 | 3.9771 | 3.885 | 3.8933 | 3.7445 | 3.5 | 3.74 | 3.82 | 3.88 | 3.902 | 4.195
| Std Dev. |  | 0.071461 | 0.10328 | 0.04858 | 0.040702 | 0.037639 | 0.067132 | 0.14552 | 0.070899 | 0.026231 | 0.075277 | 0.1002 |  | 0.064709 | 0.19062 | 0.017889 | 0.046332 | 0.10085 | 0.069474 | 0.03578 | 0.020605 | 0.040825 | 0.028577 | 0.031885 | 0.026583 | 0.023552 | 0.019408
| Rel.Std.Dev. |  | 1.732% | 2.671% | 1.233% | 1.073% | 1.012% | 1.747% | 3.627% | 1.982% | 0.7188% | 2.025% | 2.644% |  | 1.73% | 4.988% | 0.4888% | 1.238% | 2.538% | 1.783% | 0.9222% | 0.55% | 1.172% | 0.7624% | 0.8343% | 0.6857% | 0.6046% | 0.4634%
| PDM3 |  | 8.485% | 1.65% | 3.577% | -0.3221% | -2.25% | 1.036% | 5.461% | -5.974% | -4.063% | -2.294% | -0.3659% |  | -1.657% | 0.4665% | -3.783% | -1.593% | 4.446% | 2.438% | 1.99% | -1.518% | -8.428% | -1.461% | 0.4665% | 1.912% | 2.408% | 10.11%
| Z-Score (Absolute) |  | 2.63 | 0.39 | 1.02 | 0.26 | 0.89 | 0.19 | 1.64 | 2.11 | 1.48 | 0.9 | 0.27 | NA | 0.7 | 0 | 1.39 | 0.67 | 1.3 | 0.65 | 0.5 | 0.65 | 2.91 | 0.63 | 0 | 0.47 | 0.64 | 3.16
| Analytical results for Sb in OREAS 625 (Certified Value 35.1 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 32 | 38.2 | 35.4 | 34.9 | 36.4 | 34.8 | 32.4 | 35.7 | 34.3 | 36.93 | 31.7 | 35.09 | 35.5575 | 32.945 | 36.3153 | 29.62 | 36.6 | 34.6 | 33.37 | 39.06 | 35.2 | 36.31 | 36 | 29.2 | 34.7 | 35.2 | 37 | 37 |  | 1
| 1 | 2 | 32 | 38.1 | 32.6 | 35.6 | 32.4 | 33.8 | 31.5 | 37.8 | 35 | 36.02 | 28.3 | 37.15 | 34.5422 | 33.785 | 36.4298 | 28.59 | 36.5 | 34.7 | 33.98 | 39.38 | 34.6 | 36.12 | 38 | 30.3 | 33.6 | 34.4 | 38 | 41 |  | 45
| 1 | 3 | 32 | 36.9 | 32.6 | 36.1 | 37.6 | 32.4 | 32.2 | 38.7 | 36.9 | 37.38 | 32.7 | 37.02 | 34.6623 | 33.981 | 36.6301 | 27.68 | 36 | 34.6 | 33.26 | 39.91 | 34.94 | 35.91 | 39 | 28.6 | 36.3 | 34.7 | 37 | 36 |  | 16
| 1 | 4 | 33 | 37.9 | 34.4 | 36.2 | 36.9 | 34.4 | 32.1 | 37.9 | 37.4 | 37.75 | 32.9 | 36.91 | 34.292 | 33.552 | 36.0687 | 30.09 | 36.4 | 34.7 | 33.24 | 38.13 | 34.81 | 35.43 | 36 | 32.5 | 34.6 | 34.7 | 34 | 34 |  | 35.1259
| 1 | 5 | 32 | 38.1 | 32.1 | 35.4 | 36.5 | 34.8 | 32 | 39.7 | 35.4 | 37.03 | 32.1 | 36.38 | 33.493 | 32.002 | 34.8825 | 25.93 | 36.5 | 35.2 | 33.51 | 38.49 | 35.6 | 34.65 | 38 | 22.7 | 33.2 | 34.5 | 35 | 39 |  | 49
| 1 | 6 | 33 | 37.1 | 32 | 36.3 | 36.7 | 32.8 | 34.2 | 36.6 | 35.7 | 36.93 | 29.3 | 36.59 | 35.0327 | 34.047 | 34.9634 | 28.06 | 36 | 35.2 | 33.41 | 39.14 | 34.28 | 34.66 | 38 | 30.4 | 34 | 35 | 33 | 38
| Mean |  | 32.3333 | 37.7167 | 33.1833 | 35.75 | 36.0833 | 33.8333 | 32.4 | 37.7333 | 35.7833 | 37.0067 | 31.1667 | 36.5233 | 34.5966 | 33.3853 | 35.8816 | 28.3283 | 36.3333 | 34.8333 | 33.4617 | 39.0183 | 34.905 | 35.5133 | 37.5 | 28.95 | 34.4 | 34.75 | 35.6667 | 37.5
| Median |  | 32 | 38 | 32.6 | 35.85 | 36.6 | 34.1 | 32.15 | 37.85 | 35.55 | 36.98 | 31.9 | 36.75 | 34.6022 | 33.6685 | 36.192 | 28.325 | 36.45 | 34.7 | 33.39 | 39.1 | 34.875 | 35.67 | 38 | 29.75 | 34.3 | 34.7 | 36 | 37.5
| Std Dev. |  | 0.5164 | 0.56716 | 1.3891 | 0.54681 | 1.8541 | 1.0309 | 0.93167 | 1.432 | 1.1686 | 0.57926 | 1.9086 | 0.75741 | 0.69737 | 0.78583 | 0.76487 | 1.4889 | 0.26583 | 0.28752 | 0.27272 | 0.63414 | 0.46129 | 0.72671 | 1.2247 | 3.3399 | 1.0936 | 0.30166 | 1.9664 | 2.429
| Rel.Std.Dev. |  | 1.597% | 1.504% | 4.186% | 1.53% | 5.138% | 3.047% | 2.876% | 3.795% | 3.266% | 1.565% | 6.124% | 2.074% | 2.016% | 2.354% | 2.132% | 5.256% | 0.7316% | 0.8254% | 0.815% | 1.625% | 1.322% | 2.046% | 3.266% | 11.54% | 3.179% | 0.8681% | 5.513% | 6.477%
| PDM3 |  | -7.95% | 7.376% | -5.53% | 1.777% | 2.726% | -3.68% | -7.76% | 7.423% | 1.872% | 5.354% | -11.27% | 3.978% | -1.507% | -4.955% | 2.152% | -19.35% | 3.438% | -0.8328% | -4.738% | 11.08% | -0.6288% | 1.103% | 6.759% | -17.58% | -2.066% | -1.07% | 1.54% | 6.759%
| Z-Score (Absolute) |  | 1.24 | 1.08 | 0.87 | 0.23 | 0.38 | 0.59 | 1.21 | 1.09 | 0.25 | 0.78 | 1.75 | 0.57 | 0.26 | 0.79 | 0.29 | 2.97 | 0.49 | 0.16 | 0.75 | 1.64 | 0.13 | 0.13 | 0.99 | 2.7 | 0.35 | 0.2 | 0.2 | 0.99
| Analytical results for Sc in OREAS 625 (Certified Value 4.6 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | < 5 | 4 | 4.3 | 4.6 | 4.5 | 4.1 | 4.2 | 5.3 | 4.4 | 5 | 4.8 | 4.5358 | 4.5098 | 4 | 4.1 | 4.64 | 5.7 | 5.5666 | 4.9 | 5 | 4.8 | 4.7 | 4.2 | 4.5 | 4.9 |  |  |  |  | 1
| 1 | 2 | < 5 | 4 | 4.5 | 4 | 4.6 | 4 | 4.3 | 5.3 | 4.7 | 5 | 4.8 | 4.3959 | 4.9093 | 4 | 4.1 | 4.77 | 5.7 | 5.8583 | 4.8 | 5 | 4.9 | 4 | 4.6 | 4.5 | 4.8 |  |  |  |  | 16
| 1 | 3 | < 5 | 4 | 4.6 | 4.6 | 4.1 | 4 | 4.5 | 5.4 | 4.6 | 5 | 4.7 | 4.4626 | 4.8824 | 4 | 4 | 4.7 | 5.8 | 5.6772 | 4.9 | 5 | 5.2 | 4.7 | 4.3 | 4.6 | 4.8 |  |  |  |  | 16
| 1 | 4 | < 5 | 4 | 4.6 | 4.7 | 4.6 | 4 | 4.4 | 5.4 | 4.7 | 5 | 4.7 | 4.3248 | 4.6622 | 4 | 4.4 | 4.72 | 5.6 | 5.8256 | 4.7 | 5 | 5.3 | 4.6 | 4.5 | 4.3 | 4.9 |  |  |  |  | 4.6011
| 1 | 5 | < 5 | 4 | 4.4 | 4.5 | 4.3 | 4 | 4.6 | 5.3 | 4.3 | 5 | 4.6 | 4.2458 | 4.4514 | 4 | 4.6 | 4.79 | 5.7 | 5.8046 | 4.9 | 5 | 4.9 | 4.6 | 4.4 | 4.6 | 5 |  |  |  |  | 50
| 1 | 6 | < 5 | 4 | 4.5 | 4.6 | 4.2 | 4.3 | 4.2 | 5.3 | 4.4 | 5 | 4.7 | 4.3898 | 4.8409 | 4 | 4.6 | 4.73 | 6 | 5.7212 | 4.9 | 5 | 5.2 | 4.6 | 4.4 | 4.2 | 4.9
| Mean |  |  | 4 | 4.4833 | 4.5 | 4.3833 | 4.0667 | 4.3667 | 5.3333 | 4.5167 | 5 | 4.7167 | 4.3924 | 4.7093 | 4 | 4.3 | 4.725 | 5.75 | 5.7423 | 4.85 | 5 | 5.05 | 4.5333 | 4.4 | 4.45 | 4.8833
| Median |  |  | 4 | 4.5 | 4.6 | 4.4 | 4 | 4.35 | 5.3 | 4.5 | 5 | 4.7 | 4.3928 | 4.7516 | 4 | 4.25 | 4.725 | 5.7 | 5.7629 | 4.9 | 5 | 5.05 | 4.6 | 4.4 | 4.5 | 4.9
| Std Dev. |  |  | 0 | 0.1169 | 0.25298 | 0.2137 | 0.12111 | 0.1633 | 0.05164 | 0.17224 | 0 | 0.075277 | 0.10154 | 0.19791 | 0 | 0.26833 | 0.053198 | 0.13784 | 0.10937 | 0.083666 | 0 | 0.20736 | 0.26583 | 0.14142 | 0.16432 | 0.075277
| Rel.Std.Dev. |  |  | 0% | 2.608% | 5.622% | 4.875% | 2.978% | 3.74% | 0.9682% | 3.813% | 0% | 1.596% | 2.312% | 4.202% | 0% | 6.24% | 1.126% | 2.397% | 1.905% | 1.725% | 0% | 4.106% | 5.864% | 3.214% | 3.693% | 1.542%
| PDM3 |  |  | -13.06% | -2.56% | -2.198% | -4.733% | -11.62% | -5.095% | 15.91% | -1.835% | 8.669% | 2.511% | -4.535% | 2.352% | -13.06% | -6.544% | 2.692% | 24.97% | 24.8% | 5.409% | 8.669% | 9.756% | -1.473% | -4.371% | -3.284% | 6.134%
| Z-Score (Absolute) |  | 6.8 | NA | 0.11 | 0.06 | 0.45 | 1.52 | 0.51 | 2.75 | 0 | NA | 0.67 | 0.42 | 0.65 | NA | 0.73 | 0.7 | 4.16 | 4.13 | 1.12 | NA | 1.8 | 0.06 | 0.39 | 0.22 | 1.24
| Analytical results for Se in OREAS 625 (Indicative Value 2.07 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | < 5 | 1.8 | < 1 | 3 | 2 | 3 | 2 | 3 | 0.7 | 0.7 | 1 | 0.8 | 1.6827 | < 2 | < 1 | < 10 | 0.56 | 2 | 2 | 4 | 4 | < 1 | 3 | [6 empty] | 1
| 1 | 2 | < 5 | 1.8 | < 1 | 3 | 2 | 3 | 2 | 4 | 0.7 | 0.7 | 1 | 0.9 | 1.9582 | < 2 | < 1 | < 10 | 0.56 | 2.3 | 2 | 3 | 2 | < 1 | 3 | [6 empty] | 3
| 1 | 3 | < 5 | 2 | < 1 | 4 | 2 | 3 | 2 | 2 | 0.7 | 0.5 | 1 | 0.8 | 1.7538 | < 2 | < 1 | < 10 | 0.6 | 2.2 | 1 | 3 | 4 | < 1 | 3 | [6 empty] | 16
| 1 | 4 | < 5 | 1.7 | < 1 | 4 | 2 | 3 | 2 | 3 | 0.7 | < 0.5 | < 1 | 0.6 | 1.605 | < 2 | < 1 | < 10 | 0.74 | 2.1 | 2 | 3 | 4 | < 1 | 3 | [6 empty] | 2.0667
| 1 | 5 | < 5 | 1.9 | < 1 | 4 | 2 | 3 | 2 | 3 | 0.6 | 0.8 | < 1 | 1 | 1.8229 | < 2 | < 1 | < 10 | 0.61 | 2.2 | 1 | 4 | 4 | < 1 | 3 | [6 empty] | 9
| 1 | 6 | < 5 | 1.9 | < 1 | 4 | 2 | 3 | 3 | 3 | 0.7 | 0.9 | 1 | 0.7 | 1.9735 | < 2 | < 1 | < 10 | 0.58 | 2.4 | 2 | 4 | 3 | < 1 | 3
| Mean |  |  | 1.85 |  | 3.6667 | 2 | 3 | 2.1667 | 3 | 0.68333 | 0.72 | 1 | 0.8 | 1.7994 |  |  |  | 0.60833 | 2.2 | 1.6667 | 3.5 | 3.5 |  | 3
| Median |  |  | 1.85 |  | 4 | 2 | 3 | 2 | 3 | 0.7 | 0.7 | 1 | 0.8 | 1.7884 |  |  |  | 0.59 | 2.2 | 2 | 3.5 | 4 |  | 3
| Std Dev. |  |  | 0.10488 |  | 0.5164 | 0 | 0 | 0.40825 | 0.63246 | 0.040825 | 0.14832 | 0 | 0.14142 | 0.14804 |  |  |  | 0.067651 | 0.14142 | 0.5164 | 0.54772 | 0.83666 |  | 0
| Rel.Std.Dev. |  |  | 5.669% |  | 14.08% | 0% | 0% | 18.84% | 21.08% | 5.974% | 20.6% | 0% | 17.68% | 8.227% |  |  |  | 11.12% | 6.428% | 30.98% | 15.65% | 23.9% |  | 0%
| PDM3 |  |  | -10.49% |  | 77.41% | -3.23% | 45.16% | 4.835% | 45.16% | -66.94% | -65.16% | -51.61% | -61.29% | -12.94% |  |  |  | -70.57% | 6.447% | -19.36% | 69.35% | 69.35% |  | 45.16%
| Z-Score (Absolute) |  | 0.38 | 0 | 0.79 | 1.07 | 0.09 | 0.67 | 0.19 | 0.67 | 0.68 | 0.71 | 0.6 | 0.62 | 0.03 | 0.5 | 0.79 | 1.85 | 0.73 | 0.21 | 0.11 | 0.97 | 0.97 | 0.79 | 0.67
| Analytical results for Sm in OREAS 625 (Certified Value 6.8 ppm) | [30 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | [6 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 6.9; 1.02: 7; 1.03: 7.14; 1.04: 6.5741; 1.05: 8.335; 1.06: 6.86 | [6 empty] | 1
| 1 | 2 | 1.01: 7.1; 1.02: 6.73; 1.03: 7.13; 1.04: 6.3038; 1.05: 8.802; 1.06: 6.72 | [6 empty] | 8
| 1 | 3 | 1.01: 6.8; 1.02: 7.03; 1.03: 7.23; 1.04: 6.7524; 1.05: 8.3161; 1.06: 6.8 | [6 empty] | 16
| 1 | 4 | 1.01: 6.6; 1.02: 6.71; 1.03: 7.32; 1.04: 6.6549; 1.05: 8.7972; 1.06: 6.65 | [6 empty] | 6.801
| 1 | 5 | 1.01: 6.6; 1.02: 6.57; 1.03: 7.35; 1.04: 6.1812; 1.05: 8.3051; 1.06: 6.91 | [6 empty] | 51
| 1 | 6 | 1.01: 5.8; 1.02: 6.62; 1.03: 7.55; 1.04: 6.624; 1.05: 8.5837; 1.06: 6.82
| Mean |  | 1.01: 6.6333; 1.02: 6.7767; 1.03: 7.2867; 1.04: 6.5151; 1.05: 8.5232; 1.06: 6.7933
| Median |  | 1.01: 6.7; 1.02: 6.72; 1.03: 7.275; 1.04: 6.5991; 1.05: 8.4593; 1.06: 6.81
| Std Dev. |  | 1.01: 0.45019; 1.02: 0.19387; 1.03: 0.15731; 1.04: 0.22241; 1.05: 0.23761; 1.06: 0.094587
| Rel.Std.Dev. |  | 1.01: 6.787%; 1.02: 2.861%; 1.03: 2.159%; 1.04: 3.414%; 1.05: 2.788%; 1.06: 1.392%
| PDM3 |  | 1.01: -2.466%; 1.02: -0.358%; 1.03: 7.141%; 1.04: -4.204%; 1.05: 25.32%; 1.06: -0.113%
| Z-Score (Absolute) |  | 1.01: 0.49; 1.02: 0.03; 1.03: 1.6; 1.04: 0.86; 1.05: 5.56; 1.06: 0.03
| Analytical results for Sn in OREAS 625 (Certified Value 3.57 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | < 10 | 4 | 3.6 | 3.6 | 3.7 | 3.4 | 3.4 | 3.6 | 3.4 | 3.5 | 3.4 | 3.5 | 3.5744 | 2.51 | 3.2721 | < 5 | 3.4 | 3.7 | < 10 | 3.65 | 3.5 | 3.8 | 4 | 3.2 | 3.6 | 3.5 | < 10 | 3.7 |  | 1
| 1 | 2 | < 10 | 4 | 3.6 | 3.9 | 3.2 | 3.4 | 3.3 | 3.7 | 3.5 | 3.5 | 3.5 | 3.8 | 3.5924 | 2.41 | 3.513 | < 5 | 3.3 | 3.7 | < 10 | 3.73 | 3.5 | 3.8 | 4 | 3.3 | 3.6 | 3.5 | < 10 | 3.6 |  | 13
| 1 | 3 | < 10 | 4 | 3.5 | 4.3 | 3.6 | 3.3 | 3.3 | 3.7 | 3.6 | 3.8 | 3.4 | 3.8 | 3.4956 | 2.06 | 3.6274 | < 5 | 3.5 | 3.5 | < 10 | 3.78 | 3.5 | 3.8 | 4 | 3.4 | 3.8 | 3.5 | < 10 | 3.7 |  | 16
| 1 | 4 | < 10 | 4 | 3.5 | 3.8 | 3.8 | 3.6 | 3.4 | 3.9 | 3.7 | 3.7 | 3.4 | 3.7 | 3.5189 | 2.53 | 3.4332 | < 5 | 3.4 | 3.6 | < 10 | 3.66 | 3.5 | 3.8 | 4 | 3.5 | 3.7 | 3.6 | < 10 | 3.7 |  | 3.5668
| 1 | 5 | < 10 | 4 | 3.3 | 3.8 | 3.5 | 3.5 | 3.4 | 3.9 | 3.5 | 3.7 | 3.4 | 3.7 | 3.4021 | 2.32 | 3.3129 | < 5 | 3.4 | 3.7 | < 10 | 4.59 | 3.6 | 3.8 | 4 | 3.2 | 3.7 | 3.6 | < 10 | 3.8 |  | 52
| 1 | 6 | < 10 | 4 | 3.3 | 3.8 | 3.8 | 3.2 | 3.4 | 3.7 | 3.5 | 3.6 | 3.3 | 3.7 | 3.5714 | 2.05 | 3.4872 | < 5 | 3.4 | 3.7 | < 10 | 3.63 | 3.5 | 3.7 | 4 | 3.4 | 3.7 | 3.5 | < 10 | 3.8
| Mean |  |  | 4 | 3.4667 | 3.8667 | 3.6 | 3.4 | 3.3667 | 3.75 | 3.5333 | 3.6333 | 3.4 | 3.7 | 3.5258 | 2.3133 | 3.441 |  | 3.4 | 3.65 |  | 3.84 | 3.5167 | 3.7833 | 4 | 3.3333 | 3.6833 | 3.5333 |  | 3.7167
| Median |  |  | 4 | 3.5 | 3.8 | 3.65 | 3.4 | 3.4 | 3.7 | 3.5 | 3.65 | 3.4 | 3.7 | 3.5452 | 2.365 | 3.4602 |  | 3.4 | 3.7 |  | 3.695 | 3.5 | 3.8 | 4 | 3.35 | 3.7 | 3.5 |  | 3.7
| Std Dev. |  |  | 0 | 0.13663 | 0.23381 | 0.22804 | 0.14142 | 0.05164 | 0.12247 | 0.10328 | 0.12111 | 0.063246 | 0.10954 | 0.070897 | 0.21379 | 0.13196 |  | 0.063246 | 0.083666 |  | 0.3717 | 0.040825 | 0.040825 | 0 | 0.12111 | 0.075277 | 0.05164 |  | 0.075277
| Rel.Std.Dev. |  |  | 0% | 3.941% | 6.047% | 6.334% | 4.159% | 1.534% | 3.266% | 2.923% | 3.333% | 1.86% | 2.961% | 2.011% | 9.242% | 3.835% |  | 1.86% | 2.292% |  | 9.68% | 1.161% | 1.079% | 0% | 3.633% | 2.044% | 1.462% |  | 2.025%
| PDM3 |  |  | 12.14% | -2.808% | 8.406% | 0.9299% | -4.677% | -5.612% | 5.135% | -0.9391% | 1.864% | -4.677% | 3.734% | -1.15% | -35.14% | -3.529% |  | -4.677% | 2.332% |  | 7.659% | -1.406% | 6.07% | 12.14% | -6.546% | 3.266% | -0.9391% |  | 4.201%
| Z-Score (Absolute) |  | 5.8 | NA | 0.4 | 1.21 | 0.13 | 0.67 | 0.81 | 0.74 | 0.13 | 0.27 | 0.67 | 0.54 | 0.17 | 5.07 | 0.51 | 4.32 | 0.67 | 0.34 | 5.8 | 1.11 | 0.2 | 0.88 | NA | 0.94 | 0.47 | 0.13 | 5.8 | 0.61
Results from laboratories 1.02 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 625 (Certified Value 130 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 129 | 85.9 | 114 | 134.5 | 131.5 | 125.5 | 126.5 | 134.5 | 140 | 129.95 | 105 | 140 | 125.9452 | 83.2 | 137.1746 | 130.8 | 136 | 128 | 119 | 143 | 127 | 128 | 133 | 131 | 128.1 | 131.9 | 130 | 132 |  | 1
| 1 | 2 | 128 | 109 | 118 | 136.5 | 119.5 | 128.5 | 123 | 144 | 139 | 123.93 | 93 | 141 | 123.8181 | 106.52 | 132.6164 | 132.3 | 135 | 127 | 115 | 144 | 125 | 131 | 131 | 134 | 133.9 | 133.7 | 132 | 130 |  | 21
| 1 | 3 | 129 | 113 | 120 | 135 | 127.5 | 117.5 | 123.5 | 145.5 | 141 | 125.49 | 106 | 138 | 124.2495 | 117.18 | 132.9989 | 130.3 | 136 | 122 | 116 | 146 | 123 | 128 | 132 | 135 | 130.1 | 127.6 | 128 | 129 |  | 16
| 1 | 4 | 131 | 87.2 | 118 | 138.5 | 131.5 | 126.5 | 126 | 140 | 140 | 127.83 | 103 | 138 | 120.425 | 92.36 | 134.7367 | 131.8 | 135 | 129 | 114 | 141 | 124 | 128 | 128 | 137 | 124.1 | 132.4 | 127 | 133 |  | 130.2398
| 1 | 5 | 128 | 112 | 121 | 131.5 | 132 | 126 | 123.5 | 146 | 142 | 124.05 | 101 | 135 | 121.1027 | 91.69 | 135.2722 | 130.2 | 133 | 134 | 112 | 143 | 128 | 131 | 136 | 132 | 125.4 | 128.8 | 128 | 133 |  | 53
| 1 | 6 | 129 | 100 | 115 | 136 | 128 | 121.5 | 126 | 133 | 136 | 125.35 | 97 | 138 | 123.592 | 96.51 | 137.8364 | 130.8 | 133 | 132 | 118 | 142 | 122 | 125 | 134 | 134 | 123.5 | 131.3 | 131 | 132
| Mean |  | 129 | 101.1833 | 117.6667 | 135.3333 | 128.3333 | 124.25 | 124.75 | 140.5 | 139.6667 | 126.1 | 100.8333 | 138.3333 | 123.1888 | 97.91 | 135.1059 | 131.0333 | 134.6667 | 128.6667 | 115.6667 | 143.1667 | 124.8333 | 128.5 | 132.3333 | 133.8333 | 127.5167 | 130.95 | 129.3333 | 131.5
| Median |  | 129 | 104.5 | 118 | 135.5 | 129.75 | 125.75 | 124.75 | 142 | 140 | 125.42 | 102 | 138 | 123.705 | 94.435 | 135.0044 | 130.8 | 135 | 128.5 | 115.5 | 143 | 124.5 | 128 | 132.5 | 134 | 126.75 | 131.6 | 129 | 132
| Std Dev. |  | 1.0954 | 12.2331 | 2.7325 | 2.3381 | 4.7399 | 4.0218 | 1.5732 | 5.6569 | 2.0656 | 2.3521 | 4.9967 | 2.0656 | 2.0626 | 12.1103 | 2.123 | 0.84063 | 1.3663 | 4.1793 | 2.582 | 1.7224 | 2.3166 | 2.2583 | 2.7325 | 2.137 | 4.0012 | 2.3037 | 1.9664 | 1.6432
| Rel.Std.Dev. |  | 0.8492% | 12.09% | 2.322% | 1.728% | 3.693% | 3.237% | 1.261% | 4.026% | 1.479% | 1.865% | 4.955% | 1.493% | 1.674% | 12.37% | 1.571% | 0.6415% | 1.015% | 3.248% | 2.232% | 1.203% | 1.856% | 1.757% | 2.065% | 1.597% | 3.138% | 1.759% | 1.52% | 1.25%
| PDM3 |  | -0.9519% | -22.31% | -9.654% | 3.911% | -1.464% | -4.599% | -4.215% | 7.878% | 7.238% | -3.179% | -22.58% | 6.214% | -5.414% | -24.82% | 3.736% | 0.6093% | 3.399% | -1.208% | -11.19% | 9.925% | -4.151% | -1.336% | 1.607% | 2.759% | -2.091% | 0.5453% | -0.696% | 0.9676%
| Z-Score (Absolute) |  | 0.02 | 3.89 | 1.57 | 0.91 | 0.07 | 0.64 | 0.57 | 1.64 | 1.52 | 0.38 | 3.94 | 1.34 | 0.79 | 4.35 | 0.88 | 0.31 | 0.82 | 0.02 | 1.85 | 2.02 | 0.56 | 0.05 | 0.49 | 0.7 | 0.19 | 0.3 | 0.07 | 0.38
| Analytical results for Ta in OREAS 625 (Certified Value 1.11 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 2
3
| 1 | 1 | 5 | 1 | 1.4 | 1.17 | 1.17 | 1.22 | 1.02 | 1.08 | 1.01 | 1.13 | 1 | 1.14 | 1.1616 | 1.07 | 1.4 | 0.75 | 1.4 | 1.14 | 1.14 | 0.6 | 1.23 | 2.47 | [7 empty] | 1
| 1 | 2 | 5 | 1.1 | 1.5 | 1.13 | 1.05 | 1.12 | 0.98 | 1.12 | 1.02 | 1.1 | 0.9 | 1.19 | 1.1192 | 1.11 | 1.7 | 0.75 | 1.4 | 1.15 | 1.18 | 0.79 | 1.17 | 2.13 | [7 empty] | 14
| 1 | 3 | 5 | 1.1 | 1.3 | 1.1 | 1.19 | 1.1 | 1 | 1.14 | 1 | 1.13 | 1 | 1.19 | 1.2607 | 1.15 | 2.2 | 0.71 | 1.4 | 1.09 | 1.18 | 0.63 | 1.16 | 2.29 | [7 empty] | 16
| 1 | 4 | 5 | 1.1 | 1.4 | 1.14 | 1.18 | 1.21 | 1 | 1.06 | 1.01 | 1.11 | 0.9 | 1.15 | 1.2152 | 1.07 | 1.5 | 0.78 | 1.5 | 1.09 | 1.23 | 0.68 | 1.16 | 2.72 | [7 empty] | 1.1074
| 1 | 5 | 5 | 1.1 | 1.2 | 1.1 | 1.16 | 1.18 | 0.98 | 1.2 | 1 | 1.1 | 0.9 | 1.16 | 1.0789 | 1.1 | 1.9 | 0.75 | 1.4 | 1.11 | 1.11 | 0.81 | 1.18 | 2.5 | [7 empty] | 54
| 1 | 6 | 5 | 1.1 | 1.3 | 1.11 | 1.2 | 1.09 | 1.04 | 1.13 | 0.99 | 1.11 | 0.9 | 1.15 | 1.2681 | 1.05 | 2.2 | 0.74 | 1.5 | 1.06 | 1.23 | 0.54 | 1.17 | 2.27
| Mean |  | 5 | 1.0833 | 1.35 | 1.125 | 1.1583 | 1.1533 | 1.0033 | 1.1217 | 1.005 | 1.1133 | 0.93333 | 1.1633 | 1.1839 | 1.0917 | 1.8167 | 0.74667 | 1.4333 | 1.1067 | 1.1783 | 0.675 | 1.1783 | 2.3967
| Median |  | 5 | 1.1 | 1.35 | 1.12 | 1.175 | 1.15 | 1 | 1.125 | 1.005 | 1.11 | 0.9 | 1.155 | 1.1884 | 1.085 | 1.8 | 0.75 | 1.4 | 1.1 | 1.18 | 0.655 | 1.17 | 2.38
| Std Dev. |  | 0 | 0.040825 | 0.10488 | 0.027386 | 0.054924 | 0.057155 | 0.023381 | 0.04916 | 0.010488 | 0.013663 | 0.05164 | 0.021602 | 0.077017 | 0.036009 | 0.34303 | 0.022509 | 0.05164 | 0.033862 | 0.047924 | 0.1071 | 0.026394 | 0.20935
| Rel.Std.Dev. |  | 0% | 3.768% | 7.769% | 2.434% | 4.742% | 4.956% | 2.33% | 4.383% | 1.044% | 1.227% | 5.533% | 1.857% | 6.505% | 3.299% | 18.88% | 3.015% | 3.603% | 3.06% | 4.067% | 15.87% | 2.24% | 8.735%
| PDM3 |  | 351.5% | -2.169% | 21.91% | 1.594% | 4.604% | 4.152% | -9.393% | 1.293% | -9.243% | 0.5401% | -15.71% | 5.055% | 6.917% | -1.416% | 64.06% | -32.57% | 29.44% | -0.06189% | 6.41% | -39.04% | 6.41% | 116.4%
| Z-Score (Absolute) |  | NA | 0.53 | 2.84 | 0 | 0.42 | 0.36 | 1.54 | 0.04 | 1.52 | 0.15 | 2.42 | 0.48 | 0.75 | 0.42 | 8.74 | 4.78 | 3.9 | 0.23 | 0.67 | 5.69 | 0.67 | 16.08
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 625 (Certified Value 0.72 ppm) | [30 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab | [7 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.21; 1.06: 1.22; 1.07: 1.23; 1.08: 1.25; 1.09: 1.26; 1.10: 1.30 | [7 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*MS | [7 empty] | 2
3
| 1 | 1 | 1.01: 0.7; 1.02: 0.71; 1.03: 0.74; 1.04: 0.72869; 1.05: 0.8; 1.06: 1.0187; 1.07: 0.73; 1.08: 0.72; 1.09: 0.64; 1.10: 0.71 | [7 empty] | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.68; 1.03: 0.76; 1.04: 0.69266; 1.05: 0.9; 1.06: 0.99079; 1.07: 0.71; 1.08: 0.72; 1.09: 0.55; 1.10: 0.7 | [7 empty] | 36
| 1 | 3 | 1.01: 0.7; 1.02: 0.71; 1.03: 0.76; 1.04: 0.71443; 1.05: 0.8; 1.06: 0.96662; 1.07: 0.72; 1.08: 0.81; 1.09: 0.66; 1.10: 0.71 | [7 empty] | 16
| 1 | 4 | 1.01: 0.7; 1.02: 0.7; 1.03: 0.77; 1.04: 0.70893; 1.05: 0.9; 1.06: 1.0785; 1.07: 0.72; 1.08: 0.77; 1.09: 0.62; 1.10: 0.71 | [7 empty] | 0.72399
| 1 | 5 | 1.01: 0.7; 1.02: 0.7; 1.03: 0.74; 1.04: 0.68467; 1.05: 0.8; 1.06: 0.96325; 1.07: 0.69; 1.08: 0.74; 1.09: 0.61; 1.10: 0.72 | [7 empty] | 55
| 1 | 6 | 1.01: 0.7; 1.02: 0.68; 1.03: 0.78; 1.04: 0.70211; 1.05: 0.8; 1.06: 1.0323; 1.07: 0.71; 1.08: 0.78; 1.09: 0.63; 1.10: 0.71
| Mean |  | 1.01: 0.7; 1.02: 0.69667; 1.03: 0.75833; 1.04: 0.70525; 1.05: 0.83333; 1.06: 1.0084; 1.07: 0.71333; 1.08: 0.75667; 1.09: 0.61833; 1.10: 0.71
| Median |  | 1.01: 0.7; 1.02: 0.7; 1.03: 0.76; 1.04: 0.70552; 1.05: 0.8; 1.06: 1.0048; 1.07: 0.715; 1.08: 0.755; 1.09: 0.625; 1.10: 0.71
| Std Dev. |  | 1.01: 1.2162e-16; 1.02: 0.013663; 1.03: 0.016021; 1.04: 0.015743; 1.05: 0.05164; 1.06: 0.044013; 1.07: 0.013663; 1.08: 0.036148; 1.09: 0.037639; 1.10: 0.0063246
| Rel.Std.Dev. |  | 1.01: 1.737e-14%; 1.02: 1.961%; 1.03: 2.113%; 1.04: 2.232%; 1.05: 6.197%; 1.06: 4.365%; 1.07: 1.915%; 1.08: 4.777%; 1.09: 6.087%; 1.10: 0.8908%
| PDM3 |  | 1.01: -3.314%; 1.02: -3.774%; 1.03: 4.744%; 1.04: -2.588%; 1.05: 15.1%; 1.06: 39.28%; 1.07: -1.472%; 1.08: 4.513%; 1.09: -14.59%; 1.10: -1.932%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.26; 1.03: 0.7; 1.04: 0.13; 1.05: 1.87; 1.06: 4.59; 1.07: 0; 1.08: 0.67; 1.09: 1.48; 1.10: 0.05
Results from laboratory 1.02 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 625 (Certified Value 0.063 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | 8 | < 0.1 | < 0.5 | 0.08 | < 0.05 | 0.05 | < 0.05 | 0.07 | 0.1 | < 0.5 | < 0.2 | < 5 | 0.12049 | < 0.1 | 0.1 | < 10 | 0.07 | < 0.05 | 0.05 | 0.08 | 0.1 | < 0.05 | [7 empty] | 1
| 1 | 2 | 8 | < 0.1 | < 0.5 | 0.06 | < 0.05 | < 0.05 | 0.08 | < 0.05 | 0.09 | < 0.5 | < 0.2 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 10 | 0.08 | 0.06 | 0.06 | 0.08 | 0.09 | 0.11 | [7 empty] | 15
| 1 | 3 | 8 | < 0.1 | < 0.5 | 0.06 | < 0.05 | < 0.05 | 0.05 | 0.07 | 0.08 | < 0.5 | < 0.2 | < 5 | 0.11821 | < 0.1 | 0.2 | < 10 | 0.09 | < 0.05 | 0.08 | 0.07 | 0.09 | 0.06 | [7 empty] | 16
| 1 | 4 | 8 | < 0.1 | < 0.5 | 0.11 | 0.07 | 0.05 | 0.07 | 0.07 | 0.08 | < 0.5 | < 0.2 | < 5 | < 0.1 | < 0.1 | < 0.1 | < 10 | 0.13 | < 0.05 | 0.07 | 0.06 | 0.08 | 0.07 | [7 empty] | 0.063198
| 1 | 5 | 8 | < 0.1 | < 0.5 | 0.06 | < 0.05 | 0.05 | 0.05 | 0.08 | 0.11 | < 0.5 | < 0.2 | < 5 | 0.13126 | < 0.1 | 0.2 | < 10 | 0.04 | < 0.05 | 0.06 | 0.06 | 0.09 | < 0.05 | [7 empty] | 56
| 1 | 6 | 8 | < 0.1 | < 0.5 | < 0.05 | < 0.05 | < 0.05 | 0.08 | < 0.05 | 0.1 | < 0.5 | < 0.2 | < 5 | 0.10366 | < 0.1 | 0.1 | < 10 | 0.06 | < 0.05 | 0.08 | < 0.05 | 0.1 | < 0.05
| Mean |  | 8 |  |  | 0.074 | 0.07 | 0.05 | 0.066 | 0.0725 | 0.093333 |  |  |  | 0.11841 |  | 0.15 |  | 0.078333 | 0.06 | 0.066667 | 0.07 | 0.091667 | 0.08
| Median |  | 8 |  |  | 0.06 | 0.07 | 0.05 | 0.07 | 0.07 | 0.095 |  |  |  | 0.11935 |  | 0.15 |  | 0.075 | 0.06 | 0.065 | 0.07 | 0.09 | 0.07
| Std Dev. |  | 1.01: 0; 1.04: 0.021909; 1.06: 8.4984e-18; 1.07: 0.015166; 1.08: 0.005; 1.09: 0.012111; 1.15: 0.011359; 1.18: 0.057735; 1.20: 0.030605; 1.25: 0.012111; 1.26: 0.01; 1.27: 0.0075277; 1.30: 0.026458
| Rel.Std.Dev. |  | 1.01: 0%; 1.04: 29.61%; 1.06: 1.7e-14%; 1.07: 22.98%; 1.08: 6.897%; 1.09: 12.98%; 1.15: 9.593%; 1.18: 38.49%; 1.20: 39.07%; 1.25: 18.17%; 1.26: 14.29%; 1.27: 8.212%; 1.30: 33.07%
| PDM3 |  | 12559% |  |  | 17.09% | 10.76% | -20.88% | 4.434% | 14.72% | 47.68% |  |  |  | 87.36% |  | 137.3% |  | 23.95% | -5.06% | 5.489% | 10.76% | 45.05% | 26.59%
| Z-Score (Absolute) |  | NA | 0.42 | 4.64 | 0.02 | 0.86 | 0.74 | 0.19 | 0.25 | 0.67 | 4.64 | 0.84 | 61.53 | 0.73 | 0.42 | 1.26 | 124.75 | 0.3 | 0.91 | 0 | 0.11 | 0.63 | 0.36
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 625 (Certified Value 12.7 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 1
1
| 1 | 1 | 6.4 | 10.2 | 13.1 | 11.8 | 13.3 | 12.25 | 10.85 | 14.3 | 13.52 | 10.7 | 13.93 | 12.6066 | 13.5511 | 13.1 | 13.7 | 13.5 | 12.96 | 14.9914 | 12.92 | 16 | 12.3 | 13.9 | 10.9 | 11.19 | [5 empty] | 1
| 1 | 2 | 7.4 | 9.8 | 12.8 | 10.8 | 12.85 | 11.95 | 11.2 | 14.1 | 13.26 | 9.5 | 14.56 | 11.9857 | 12.5922 | 12.7 | 13.6 | 13.77 | 12.92 | 15.0451 | 13.52 | 15 | 12.8 | 12.2 | 11.9 | 11.8 | [5 empty] | 38
| 1 | 3 | 8.2 | 10 | 12.45 | 11.8 | 11.4 | 11.7 | 11.45 | 14.1 | 13.54 | 11.4 | 14.72 | 12.57 | 12.7157 | 13.2 | 13 | 13.47 | 13.11 | 14.9531 | 12.72 | 17 | 14.3 | 13.8 | 11.5 | 11.16 | [5 empty] | 16
| 1 | 4 | 6.6 | 10.6 | 12.85 | 12.6 | 12.9 | 12.2 | 10.65 | 13.9 | 13.42 | 11 | 14.15 | 12.1502 | 13.7166 | 13.1 | 13.8 | 13.45 | 14.8 | 15.0337 | 12.23 | 16 | 13.8 | 12.8 | 11.7 | 11.14 | [5 empty] | 12.6735
| 1 | 5 | 7.7 | 10.3 | 12.15 | 12.6 | 11.65 | 11.75 | 11.65 | 14.5 | 13.33 | 10.4 | 14.27 | 12.1366 | 12.4769 | 12.8 | 14.4 | 12.88 | 13.3 | 14.9576 | 13.18 | 17 | 13.1 | 12.9 | 11.6 | 11.37 | [5 empty] | 57
| 1 | 6 | 7.5 | 10.6 | 12.3 | 12.5 | 12.1 | 12.3 | 11.15 | 13.6 | 13.41 | 9.9 | 14.62 | 12.0749 | 14.1255 | 13.1 | 14.5 | 13.55 | 13.36 | 14.9748 | 12.48 | 17 | 13.1 | 13.2 | 11.6 | 11.21
| Mean |  | 7.3 | 10.25 | 12.6083 | 12.0167 | 12.3667 | 12.025 | 11.1583 | 14.0833 | 13.4133 | 10.4833 | 14.375 | 12.254 | 13.1964 | 13 | 13.8333 | 13.4367 | 13.4083 | 14.9926 | 12.8417 | 16.3333 | 13.2333 | 13.1333 | 11.5333 | 11.3117
| Median |  | 7.45 | 10.25 | 12.625 | 12.15 | 12.475 | 12.075 | 11.175 | 14.1 | 13.415 | 10.55 | 14.415 | 12.1434 | 13.1334 | 13.1 | 13.75 | 13.485 | 13.205 | 14.9831 | 12.82 | 16.5 | 13.1 | 13.05 | 11.6 | 11.2
| Std Dev. |  | 0.68118 | 0.32094 | 0.36526 | 0.70545 | 0.76267 | 0.2622 | 0.36935 | 0.31252 | 0.10764 | 0.70261 | 0.30755 | 0.26562 | 0.68899 | 0.2 | 0.55377 | 0.29636 | 0.70417 | 0.038857 | 0.46914 | 0.8165 | 0.71461 | 0.64395 | 0.33862 | 0.25278
| Rel.Std.Dev. |  | 9.331% | 3.131% | 2.897% | 5.871% | 6.167% | 2.18% | 3.31% | 2.219% | 0.8025% | 6.702% | 2.14% | 2.168% | 5.221% | 1.538% | 4.003% | 2.206% | 5.252% | 0.2592% | 3.653% | 4.999% | 5.4% | 4.903% | 2.936% | 2.235%
| PDM3 |  | -42.4% | -19.12% | -0.5141% | -5.183% | -2.421% | -5.117% | -11.96% | 11.12% | 5.838% | -17.28% | 13.43% | -3.31% | 4.126% | 2.576% | 9.152% | 6.022% | 5.798% | 18.3% | 1.327% | 28.88% | 4.417% | 3.628% | -8.996% | -10.75%
| Z-Score (Absolute) |  | 4.58 | 2.14 | 0.19 | 0.68 | 0.39 | 0.67 | 1.39 | 1.03 | 0.47 | 1.95 | 1.27 | 0.49 | 0.29 | 0.13 | 0.82 | 0.49 | 0.47 | 1.78 | 0 | NA | 0.32 | 0.24 | 1.08 | 1.26
Results from laboratory 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 625 (Certified Value 0.12 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 3
3
| 1 | 1 | 0.11 | 0.119 | 0.124 | 0.123 | 0.121 | 0.125 | 0.128 | 0.12 | 0.1199 | 0.113 | 0.1222 | 0.12699 | < 0.1 | 0.12285 | 0.12 | 0.119 | 0.12 | 0.1004 | 0.12 | 0.11758 | 0.115 | 0.11 | 0.12 | 0.1 | 0.12 | 0.111 | 0.121 |  |  | 1
| 1 | 2 | 0.11 | 0.114 | 0.127 | 0.121 | 0.12 | 0.119 | 0.128 | 0.12 | 0.1163 | 0.112 | 0.1249 | 0.12617 | < 0.1 | 0.12237 | 0.12 | 0.118 | 0.118 | 0.1012 | 0.12 | 0.11672 | 0.116 | 0.11 | 0.12 | 0.1 | 0.13 | 0.117 | 0.122 |  |  | 39
| 1 | 3 | 0.11 | 0.123 | 0.125 | 0.123 | 0.118 | 0.119 | 0.131 | 0.13 | 0.1231 | 0.111 | 0.1198 | 0.12459 | < 0.1 | 0.12526 | 0.12 | 0.123 | 0.118 | 0.1005 | 0.12 | 0.12054 | 0.115 | 0.12 | 0.12 | 0.11 | 0.12 | 0.112 | 0.122 |  |  | 16
| 1 | 4 | 0.11 | 0.124 | 0.126 | 0.121 | 0.122 | 0.119 | 0.129 | 0.13 | 0.1169 | 0.11 | 0.1186 | 0.12249 | < 0.1 | 0.12708 | 0.12 | 0.122 | 0.122 | 0.1038 | 0.12 | 0.11292 | 0.115 | 0.12 | 0.12 | 0.11 | 0.12 | 0.114 | 0.119 |  |  | 0.12026
| 1 | 5 | 0.11 | 0.119 | 0.125 | 0.121 | 0.122 | 0.118 | 0.139 | 0.12 | 0.1187 | 0.109 | 0.1175 | 0.12186 | < 0.1 | 0.12362 | 0.12 | 0.12 | 0.121 | 0.101 | 0.12 | 0.11875 | 0.117 | 0.12 | 0.12 | 0.11 | 0.12 | 0.113 | 0.122 |  |  | 58
| 1 | 6 | 0.1 | 0.121 | 0.125 | 0.124 | 0.121 | 0.122 | 0.128 | 0.13 | 0.1163 | 0.109 | 0.1164 | 0.12425 | < 0.1 | 0.12624 | 0.12 | 0.114 | 0.122 | 0.1018 | 0.12 | 0.11993 | 0.114 | 0.12 | 0.12 | 0.11 | 0.12 | 0.112 | 0.122
| Mean |  | 0.10833 | 0.12 | 0.12533 | 0.12217 | 0.12067 | 0.12033 | 0.1305 | 0.125 | 0.11853 | 0.11067 | 0.1199 | 0.12439 |  | 0.12457 | 0.12 | 0.11933 | 0.12017 | 0.10145 | 0.12 | 0.11774 | 0.11533 | 0.11667 | 0.12 | 0.10667 | 0.12167 | 0.11317 | 0.12133
| Median |  | 0.11 | 0.12 | 0.125 | 0.122 | 0.121 | 0.119 | 0.1285 | 0.125 | 0.1178 | 0.1105 | 0.1192 | 0.12442 |  | 0.12444 | 0.12 | 0.1195 | 0.1205 | 0.1011 | 0.12 | 0.11817 | 0.115 | 0.12 | 0.12 | 0.11 | 0.12 | 0.1125 | 0.122
| Std Dev. |  | 0.0040825 | 0.0035777 | 0.0010328 | 0.0013292 | 0.0015055 | 0.0026583 | 0.0043243 | 0.0054772 | 0.0026606 | 0.001633 | 0.0031623 | 0.0019999 |  | 0.0019133 | 0 | 0.0032042 | 0.0018348 | 0.0012582 | 0 | 0.0027552 | 0.0010328 | 0.005164 | 0 | 0.005164 | 0.0040825 | 0.002137 | 0.0012111
| Rel.Std.Dev. |  | 3.768% | 2.981% | 0.824% | 1.088% | 1.248% | 2.209% | 3.314% | 4.382% | 2.245% | 1.476% | 2.637% | 1.608% |  | 1.536% | 0% | 2.685% | 1.527% | 1.24% | 0% | 2.34% | 0.8955% | 4.426% | 0% | 4.841% | 3.355% | 1.888% | 0.9981%
| PDM3 |  | -9.914% | -0.2126% | 4.222% | 1.589% | 0.3418% | 0.06458% | 8.519% | 3.945% | -1.432% | -7.974% | -0.2958% | 3.438% |  | 3.589% | -0.2126% | -0.767% | -0.07401% | -15.64% | -0.2126% | -2.092% | -4.093% | -2.984% | -0.2126% | -11.3% | 1.173% | -5.895% | 0.8961%
| Z-Score (Absolute) |  | 3.48 | 0 | 1.59 | 0.65 | 0.2 | 0.1 | 3.13 | 1.49 | 0.44 | 2.78 | 0.03 | 1.31 | 20.89 | 1.36 | 0 | 0.2 | 0.05 | 5.53 | 0 | 0.67 | 1.39 | 0.99 | 0 | 3.98 | 0.5 | 2.04 | 0.4
| Analytical results for Tl in OREAS 625 (Certified Value 9.79 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 9.99 | 10 | 9.78 | 10.4 | 10.7 | 9.03 | 10.15 | 9.8 | 9.89 | 10.5 | 9.91 | 10.0538 | 10.0216 | 10.2 | 9.5 | 9.722 | 9.83 | 9.02 | 9.46 | 9.37 | 9.8 | 7 | [7 empty] | 1
| 1 | 2 | 9.84 | 9.2 | 9.61 | 9.2 | 9.43 | 8.95 | 10.4 | 9.75 | 9.65 | 10.4 | 10.41 | 9.8888 | 10.2271 | 10.2 | 9.6 | 9.941 | 9.7 | 9.37 | 10 | 9.21 | 9.85 | 8 | [7 empty] | 0
| 1 | 3 | 9.96 | 9.3 | 9.53 | 10.4 | 9.6 | 8.8 | 10.7 | 9.74 | 9.86 | 10.3 | 10.5 | 9.4706 | 10.1726 | 10.2 | 9.5 | 9.712 | 9.48 | 9.27 | 10.2 | 9.8 | 9.75 | 8 | [7 empty] | 16
| 1 | 4 | 9.92 | 9.9 | 9.52 | 10.4 | 10 | 9.08 | 10.6 | 9.64 | 9.7 | 10.5 | 10.34 | 9.7759 | 10.1267 | 10.3 | 9.4 | 9.452 | 9.8 | 9.26 | 10.4 | 9.27 | 10.08 | 8 | [7 empty] | 9.7907
| 1 | 5 | 10 | 9.3 | 9.17 | 10.2 | 9.64 | 8.86 | 11.25 | 10 | 9.69 | 10.4 | 10.33 | 9.21 | 10.0558 | 10.5 | 9.5 | 9.322 | 9.74 | 9.55 | 9.85 | 20.77 | 9.75 | 7 | [7 empty] | 59
| 1 | 6 | 9.93 | 9.5 | 9.49 | 10.4 | 9.39 | 9.32 | 10.15 | 9.41 | 9.72 | 10.2 | 10.59 | 9.9169 | 10.0045 | 10.4 | 9.7 | 9.808 | 9.65 | 9.2 | 10.2 | 9.34 | 9.84 | 9
| Mean |  | 9.94 | 9.5333 | 9.5167 | 10.1667 | 9.7933 | 9.0067 | 10.5417 | 9.7233 | 9.7517 | 10.3833 | 10.3467 | 9.7193 | 10.1014 | 10.3 | 9.5333 | 9.6595 | 9.7 | 9.2783 | 10.0183 | 11.2933 | 9.845 | 7.8333
| Median |  | 9.945 | 9.4 | 9.525 | 10.4 | 9.62 | 8.99 | 10.5 | 9.745 | 9.71 | 10.4 | 10.375 | 9.8324 | 10.0912 | 10.25 | 9.5 | 9.717 | 9.72 | 9.265 | 10.1 | 9.355 | 9.82 | 8
| Std Dev. |  | 0.05831 | 0.33862 | 0.19957 | 0.48028 | 0.49403 | 0.18522 | 0.41402 | 0.19418 | 0.098675 | 0.1169 | 0.23568 | 0.31752 | 0.088712 | 0.12649 | 0.10328 | 0.23018 | 0.12602 | 0.17634 | 0.33229 | 4.6473 | 0.12276 | 0.75277
| Rel.Std.Dev. |  | 0.5866% | 3.552% | 2.097% | 4.724% | 5.045% | 2.056% | 3.928% | 1.997% | 1.012% | 1.126% | 2.278% | 3.267% | 0.8782% | 1.228% | 1.083% | 2.383% | 1.299% | 1.901% | 3.317% | 41.15% | 1.247% | 9.61%
| PDM3 |  | 1.525% | -2.629% | -2.799% | 3.84% | 0.02678% | -8.008% | 7.67% | -0.6882% | -0.3988% | 6.053% | 5.678% | -0.7291% | 3.173% | 5.202% | -2.629% | -1.34% | -0.9265% | -5.233% | 2.325% | 15.35% | 0.5545% | -19.99%
| Z-Score (Absolute) |  | 0.38 | 0.67 | 0.72 | 0.97 | 0 | 2.04 | 1.94 | 0.18 | 0.11 | 1.53 | 1.44 | 0.19 | 0.8 | 1.31 | 0.67 | 0.35 | 0.24 | 1.34 | 0.58 | 3.89 | 0.13 | NA
Results from laboratory 1.30 were removed due to their 10 ppm reading resolution.
| Analytical results for Tm in OREAS 625 (Certified Value 0.17 ppm) | [30 empty] | Certified
Tm | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [7 empty] | 1
| No | No | 1.02: 1.02; 1.03: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.22; 1.07: 1.30 | [7 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*MS | [7 empty] | 2
2
| 1 | 1 | 1.02: 0.1; 1.03: 0.17; 1.04: 0.15; 1.05: 0.15119; 1.06: 0.19562; 1.07: 0.16 | [7 empty] | 1
| 1 | 2 | 1.02: 0.2; 1.03: 0.16; 1.04: 0.15; 1.05: 0.14196; 1.06: 0.18294; 1.07: 0.15 | [7 empty] | 9
| 1 | 3 | 1.02: 0.2; 1.03: 0.17; 1.04: 0.16; 1.05: 0.14782; 1.06: 0.17779; 1.07: 0.16 | [7 empty] | 16
| 1 | 4 | 1.02: 0.2; 1.03: 0.16; 1.04: 0.16; 1.05: 0.1501; 1.06: 0.19998; 1.07: 0.16 | [7 empty] | 0.16596
| 1 | 5 | 1.02: 0.2; 1.03: 0.16; 1.04: 0.15; 1.05: 0.14767; 1.06: 0.17317; 1.07: 0.16 | [7 empty] | 60
| 1 | 6 | 1.02: 0.2; 1.03: 0.15; 1.04: 0.16; 1.05: 0.15088; 1.06: 0.20528; 1.07: 0.16
| Mean |  | 1.02: 0.18333; 1.03: 0.16167; 1.04: 0.155; 1.05: 0.14827; 1.06: 0.18913; 1.07: 0.15833
| Median |  | 1.02: 0.2; 1.03: 0.16; 1.04: 0.155; 1.05: 0.14896; 1.06: 0.18928; 1.07: 0.16
| Std Dev. |  | 1.02: 0.040825; 1.03: 0.0075277; 1.04: 0.0054772; 1.05: 0.0034364; 1.06: 0.01298; 1.07: 0.0040825
| Rel.Std.Dev. |  | 1.02: 22.27%; 1.03: 4.656%; 1.04: 3.534%; 1.05: 2.318%; 1.06: 6.863%; 1.07: 2.578%
| PDM3 |  | 1.02: 10.47%; 1.03: -2.584%; 1.04: -6.602%; 1.05: -10.66%; 1.06: 13.96%; 1.07: -4.593%
| Z-Score (Absolute) |  | 1.02: 1.88; 1.03: 0.13; 1.04: 0.4; 1.05: 0.95; 1.06: 2.35; 1.07: 0.13
| Analytical results for U in OREAS 625 (Certified Value 5.52 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 5.5 | 5.2 | 5.4 | 5.1 | 5.7 | 5.3 | 4.7 | 5.54 | 5.71 | 5.7 | 5.74 | 5.7947 | 5.6673 | 5.77 | 5 | 5.84 | 5.68 | 5.9606 | 5.22 | < 20 | 5.23 | 5.6 | 5.2 | 5.5 | [5 empty] | 1
| 1 | 2 | 5.5 | 5.1 | 5.4 | 4.9 | 5.4 | 5.2 | 4.8 | 5.49 | 5.62 | 5.5 | 5.86 | 5.4458 | 5.6058 | 5.55 | 5 | 6.08 | 5.54 | 5.9146 | 5.45 | < 20 | 5.46 | 5.3 | 5.4 | 5.4 | [5 empty] | 41
| 1 | 3 | 5.6 | 5.1 | 5.3 | 5.4 | 4.9 | 5.1 | 5 | 5.45 | 5.85 | 5.6 | 5.89 | 5.6034 | 5.7807 | 5.74 | 4.9 | 5.89 | 5.37 | 5.9938 | 5.37 | < 20 | 5.69 | 5.8 | 5.4 | 5.5 | [5 empty] | 16
| 1 | 4 | 5.7 | 5.1 | 5.4 | 5.6 | 5.7 | 5.2 | 4.8 | 5.42 | 5.75 | 5.6 | 5.74 | 5.9279 | 5.6812 | 5.75 | 5.1 | 5.85 | 5.75 | 6.2391 | 5.29 | < 20 | 5.65 | 5.5 | 5.4 | 5.5 | [5 empty] | 5.5198
| 1 | 5 | 5.7 | 5.1 | 5.1 | 5.4 | 5.4 | 5.1 | 5.2 | 5.63 | 5.62 | 5.6 | 5.75 | 5.6953 | 5.5651 | 5.75 | 5.2 | 5.61 | 5.62 | 6.0135 | 5.68 | < 20 | 5.6 | 5.4 | 5.4 | 5.6 | [5 empty] | 61
| 1 | 6 | 5.2 | 5.1 | 5.2 | 5.7 | 5.3 | 5.3 | 4.9 | 5.25 | 5.73 | 5.5 | 5.68 | 5.6981 | 5.7451 | 5.8 | 5.3 | 5.86 | 5.52 | 5.9612 | 5.18 | < 20 | 5.6 | 5.5 | 5.4 | 5.7
| Mean |  | 5.5333 | 5.1167 | 5.3 | 5.35 | 5.4 | 5.2 | 4.9 | 5.4633 | 5.7133 | 5.5833 | 5.7767 | 5.6942 | 5.6742 | 5.7267 | 5.0833 | 5.855 | 5.58 | 6.0138 | 5.365 |  | 5.5383 | 5.5167 | 5.3667 | 5.5333
| Median |  | 5.55 | 5.1 | 5.35 | 5.4 | 5.4 | 5.2 | 4.85 | 5.47 | 5.72 | 5.6 | 5.745 | 5.6967 | 5.6742 | 5.75 | 5.05 | 5.855 | 5.58 | 5.9775 | 5.33 |  | 5.6 | 5.5 | 5.4 | 5.5
| Std Dev. |  | 0.18619 | 0.040825 | 0.12649 | 0.30166 | 0.29665 | 0.089443 | 0.17889 | 0.12801 | 0.086872 | 0.075277 | 0.080664 | 0.1641 | 0.081287 | 0.089144 | 0.1472 | 0.14977 | 0.13401 | 0.11541 | 0.18295 |  | 0.16987 | 0.17224 | 0.08165 | 0.10328
| Rel.Std.Dev. |  | 3.365% | 0.7979% | 2.387% | 5.639% | 5.493% | 1.72% | 3.651% | 2.343% | 1.521% | 1.348% | 1.396% | 2.882% | 1.433% | 1.557% | 2.896% | 2.558% | 2.402% | 1.919% | 3.41% |  | 3.067% | 3.122% | 1.521% | 1.866%
| PDM3 |  | 0.245% | -7.304% | -3.982% | -3.076% | -2.171% | -5.794% | -11.23% | -1.023% | 3.506% | 1.151% | 4.653% | 3.16% | 2.797% | 3.748% | -7.907% | 6.072% | 1.09% | 8.95% | -2.805% |  | 0.3356% | -0.05695% | -2.774% | 0.245%
| Z-Score (Absolute) |  | 0 | 1.61 | 0.9 | 0.71 | 0.52 | 1.29 | 2.45 | 0.27 | 0.7 | 0.19 | 0.94 | 0.62 | 0.55 | 0.75 | 1.74 | 1.25 | 0.18 | 1.86 | 0.65 | 17.29 | 0.02 | 0.06 | 0.65 | 0
| Analytical results for V in OREAS 625 (Certified Value 11.5 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 12 | 12 | 11 | 11 | 10 | 10 | 11 | 11 | 11.7 | 10 | 14 | 11 | 11.4595 | 12.2 | 10.1148 | 25.85 | 11 | 10 | 13.1 | 11 | 11 | 13.2 | 11 | 12 | 12 | 13 | 11 | 12 | 9 | 1
| 1 | 2 | 12 | 12 | 11 | 11 | 10 | 10 | 11 | 11 | 12 | 10 | 14 | 11 | 11.3966 | 12.456 | 11.4142 | 25.21 | 11 | 11 | 12.58 | 11 | 11 | 12.95 | 11 | 13 | 12 | 12 | 12 | 12 | 9 | 23
| 1 | 3 | 12 | 12 | 11 | 11 | 10 | 10 | 10 | 11 | 12.3 | 10 | 13 | 11 | 11.5938 | 11.981 | 12.2545 | 24.52 | 11 | 11 | 12.76 | 11 | 10 | 13.51 | 11 | 12 | 12 | 13 | 12 | 12 | 9 | 16
| 1 | 4 | 13 | 13 | 11 | 11 | 10 | 10 | 11 | 11 | 12.4 | 10 | 14 | 10 | 11.2023 | 12.177 | 12.3998 | 23.62 | 11 | 11 | 12.95 | 11 | 11 | 12.63 | 11 | 13 | 12 | 12 | 12 | 12 | 10 | 11.4585
| 1 | 5 | 12 | 13 | 11 | 11 | 10 | 10 | 11 | 12 | 12.1 | 10 | 13 | 11 | 11.3842 | 12.195 | 11.2448 | 25.62 | 11 | 11 | 12.96 | 11 | 12 | 13.27 | 11 | 12 | 12 | 13 | 12 | 12 | 10 | 62
| 1 | 6 | 12 | 11 | 11 | 11 | 10 | 10 | 11 | 10 | 12.2 | 10 | 14 | 11 | 11.4475 | 12.041 | 10.4584 | 25.54 | 11 | 11 | 12.86 | 11 | 11 | 13.13 | 11 | 13 | 12 | 12 | 12 | 12 | 10
| Mean |  | 12.1667 | 12.1667 | 11 | 11 | 10 | 10 | 10.8333 | 11 | 12.1167 | 10 | 13.6667 | 10.8333 | 11.414 | 12.175 | 11.3144 | 25.06 | 11 | 10.8333 | 12.8683 | 11 | 11 | 13.115 | 11 | 12.5 | 12 | 12.5 | 11.8333 | 12 | 9.5
| Median |  | 12 | 12 | 11 | 11 | 10 | 10 | 11 | 11 | 12.15 | 10 | 14 | 11 | 11.4221 | 12.186 | 11.3295 | 25.375 | 11 | 11 | 12.905 | 11 | 11 | 13.165 | 11 | 12.5 | 12 | 12.5 | 12 | 12 | 9.5
| Std Dev. |  | 0.40825 | 0.75277 | 0 | 0 | 0 | 0 | 0.40825 | 0.63246 | 0.24833 | 0 | 0.5164 | 0.40825 | 0.12769 | 0.16467 | 0.92179 | 0.8433 | 0 | 0.40825 | 0.18093 | 0 | 0.63246 | 0.29998 | 0 | 0.54772 | 0 | 0.54772 | 0.40825 | 0 | 0.54772
| Rel.Std.Dev. |  | 3.355% | 6.187% | 0% | 0% | 0% | 0% | 3.768% | 5.75% | 2.049% | 0% | 3.779% | 3.768% | 1.119% | 1.353% | 8.147% | 3.365% | 0% | 3.768% | 1.406% | 0% | 5.75% | 2.287% | 0% | 4.382% | 0% | 4.382% | 3.45% | 0% | 5.766%
| PDM3 |  | 6.181% | 6.181% | -4.001% | -4.001% | -12.73% | -12.73% | -5.456% | -4.001% | 5.744% | -12.73% | 19.27% | -5.456% | -0.3881% | 6.253% | -1.257% | 118.7% | -4.001% | -5.456% | 12.3% | -4.001% | -4.001% | 14.46% | -4.001% | 9.09% | 4.726% | 9.09% | 3.272% | 4.726% | -17.09%
| Z-Score (Absolute) |  | 0.84 | 0.84 | 0.31 | 0.31 | 1.29 | 1.29 | 0.47 | 0.31 | 0.79 | 1.29 | 2.31 | 0.47 | 0.1 | 0.85 | 0 | 13.52 | 0.31 | 0.47 | 1.53 | 0.31 | 0.31 | 1.77 | 0.31 | 1.17 | 0.67 | 1.17 | 0.51 | 0.67 | 1.78
| Analytical results for W in OREAS 625 (Certified Value 3.99 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | < 10 | 4.5 | 3.9 | 4.5 | 4.2 | 4.2 | 3.9 | 3.9 | 3.4 | 4.2 | 3.7 | 4.1 | < 10 | 3.8556 | < 5 | 4.1 | 3.7 | 3.53 | 4.1 | 4.1 | < 4 | 3 | 3.5 | 4.3 | 104 | 4.8 |  |  |  | 1
| 1 | 2 | < 10 | 4.5 | 4.4 | 4.4 | 3.7 | 3.8 | 3.9 | 4.1 | 3.5 | 3.9 | 3.5 | 4.3 | < 10 | 4.1093 | < 5 | 3.8 | 3.7 | 3.46 | 4.1 | 3.7 | < 4 | 3 | 3.5 | 4.2 | 102 | 4.8 |  |  |  | 43
| 1 | 3 | < 10 | 4.6 | 3.7 | 4.2 | 4.3 | 4.1 | 3.9 | 4.2 | 3.6 | 4.4 | 4.7 | 4.5 | < 10 | 3.8843 | < 5 | 4.1 | 3.8 | 3.53 | 4 | 3.8 | < 4 | 3 | 3.8 | 4.2 | 103 | 5 |  |  |  | 16
| 1 | 4 | < 10 | 4.5 | 4 | 4.4 | 4.2 | 4.1 | 3.9 | 3.9 | 3.4 | 4.4 | 3.5 | 4.4 | < 10 | 4.044 | < 5 | 3.8 | 4 | 3.51 | 4.1 | 3.8 | < 4 | 3 | 3.6 | 4.2 | 105 | 4.8 |  |  |  | 3.9863
| 1 | 5 | < 10 | 4.5 | 3.7 | 4 | 4.2 | 4 | 4.1 | 4.3 | 3.5 | 4.1 | 3.6 | 4.2 | < 10 | 3.7659 | < 5 | 3.8 | 3.7 | 3.87 | 4.1 | 4 | < 4 | 4 | 3.5 | 4.1 | 110 | 5 |  |  |  | 63
| 1 | 6 | < 10 | 4.7 | 3.7 | 4.2 | 4.4 | 3.8 | 4 | 4 | 3.4 | 4 | 3.4 | 4.2 | < 10 | 4.0786 | < 5 | 4 | 3.9 | 3.5 | 4 | 3.8 | < 4 | 5 | 3.6 | 4.2 | 104 | 4.6
| Mean |  |  | 4.55 | 3.9 | 4.2833 | 4.1667 | 4 | 3.95 | 4.0667 | 3.4667 | 4.1667 | 3.7333 | 4.2833 |  | 3.9563 |  | 3.9333 | 3.8 | 3.5667 | 4.0667 | 3.8667 |  | 3.5 | 3.5833 | 4.2 | 104.6667 | 4.8333
| Median |  |  | 4.5 | 3.8 | 4.3 | 4.2 | 4.05 | 3.9 | 4.05 | 3.45 | 4.15 | 3.55 | 4.25 |  | 3.9641 |  | 3.9 | 3.75 | 3.52 | 4.1 | 3.8 |  | 3 | 3.55 | 4.2 | 104 | 4.8
| Std Dev. |  |  | 0.083666 | 0.27568 | 0.18348 | 0.24221 | 0.16733 | 0.083666 | 0.1633 | 0.08165 | 0.20656 | 0.48442 | 0.1472 |  | 0.13971 |  | 0.15055 | 0.12649 | 0.15082 | 0.05164 | 0.15055 |  | 0.83666 | 0.1169 | 0.063246 | 2.8048 | 0.15055
| Rel.Std.Dev. |  |  | 1.839% | 7.069% | 4.284% | 5.813% | 4.183% | 2.118% | 4.016% | 2.355% | 4.957% | 12.98% | 3.436% |  | 3.531% |  | 3.828% | 3.329% | 4.229% | 1.27% | 3.894% |  | 23.9% | 3.262% | 1.506% | 2.68% | 3.115%
| PDM3 |  |  | 14.14% | -2.165% | 7.451% | 4.525% | 0.3437% | -0.9106% | 2.016% | -13.04% | 4.525% | -6.346% | 7.451% |  | -0.7534% |  | -1.329% | -4.673% | -10.53% | 2.016% | -3.001% |  | -12.2% | -10.11% | 5.361% | 2526% | 21.25%
| Z-Score (Absolute) |  | 2.51 | 1.4 | 0.19 | 0.75 | 0.46 | 0.05 | 0.07 | 0.22 | 1.25 | 0.46 | 0.6 | 0.75 | 2.51 | 0.05 | 3.62 | 0.11 | 0.44 | 1.01 | 0.22 | 0.27 | 4.85 | 1.17 | 0.97 | 0.54 | 246.89 | 2.1
| Analytical results for Y in OREAS 625 (Certified Value 14.9 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 15 | 13.5 | 14.3 | 14.3 | 14.5 | 14.2 | 14.1 | 12.9 | 15.4 | 15.48 | 15.1 | 14.9 | 14.5501 | 14.832 | 16.2 | 14.4 | 12.26 | 15.33 | 15.1 | 15.1 | 17 | 13.3 | 14.3 | 17.6 | 17.3 | 14.3 |  |  |  | 1
| 1 | 2 | 15 | 14.8 | 13.5 | 15.3 | 13.3 | 14 | 13.9 | 13.7 | 15.5 | 15.15 | 15.4 | 16.05 | 14.1709 | 15.0673 | 15.8 | 14.2 | 11.95 | 15.58 | 14.7 | 14.4 | 16 | 13.4 | 11.3 | 17.4 | 17.3 | 14 |  |  |  | 44
| 1 | 3 | 15 | 14.6 | 14.1 | 14.9 | 14.7 | 12.7 | 14 | 14.1 | 15.5 | 15.39 | 15.2 | 15.92 | 14.3558 | 14.9937 | 16.1 | 14.3 | 12.16 | 15.7 | 15 | 15.1 | 17 | 15.6 | 14.6 | 16.8 | 16.7 | 14.1 |  |  |  | 16
| 1 | 4 | 16 | 14.1 | 14.1 | 15.5 | 15.7 | 14.2 | 14.5 | 13.2 | 15.4 | 15.36 | 15.1 | 15.41 | 13.9648 | 14.8203 | 16 | 14.8 | 12.24 | 15.15 | 14.9 | 14.8 | 16 | 14.6 | 14 | 15.9 | 17.2 | 14.1 |  |  |  | 14.8882
| 1 | 5 | 15 | 14.9 | 13.5 | 14.4 | 15.4 | 12.9 | 14.3 | 14.1 | 15.5 | 15.16 | 14.8 | 15.5 | 13.7706 | 14.6756 | 16 | 15.6 | 12.06 | 15.33 | 15.3 | 15 | 17 | 14.3 | 13.5 | 17.4 | 16.7 | 14.5 |  |  |  | 64
| 1 | 6 | 16 | 12.8 | 13.2 | 15.3 | 15.3 | 13.1 | 15.4 | 13.6 | 15.9 | 15.43 | 14.8 | 15.96 | 14.0928 | 14.8224 | 16.1 | 15.3 | 12.21 | 15.12 | 14.7 | 15 | 17 | 14.5 | 13.6 | 17.2 | 16.8 | 14.4
| Mean |  | 15.3333 | 14.1167 | 13.7833 | 14.95 | 14.8167 | 13.5167 | 14.3667 | 13.6 | 15.5333 | 15.3283 | 15.0667 | 15.6233 | 14.1509 | 14.8685 | 16.0333 | 14.7667 | 12.1467 | 15.3683 | 14.95 | 14.9 | 16.6667 | 14.2833 | 13.55 | 17.05 | 17 | 14.2333
| Median |  | 15 | 14.35 | 13.8 | 15.1 | 15 | 13.55 | 14.2 | 13.65 | 15.5 | 15.375 | 15.1 | 15.71 | 14.1319 | 14.8272 | 16.05 | 14.6 | 12.185 | 15.33 | 14.95 | 15 | 17 | 14.4 | 13.8 | 17.3 | 17 | 14.2
| Std Dev. |  | 0.5164 | 0.82805 | 0.44008 | 0.50498 | 0.86814 | 0.69113 | 0.55015 | 0.48166 | 0.18619 | 0.1402 | 0.23381 | 0.43985 | 0.27727 | 0.14014 | 0.13663 | 0.57504 | 0.11994 | 0.23112 | 0.23452 | 0.26833 | 0.5164 | 0.85186 | 1.1777 | 0.6253 | 0.29665 | 0.19664
| Rel.Std.Dev. |  | 3.368% | 5.866% | 3.193% | 3.378% | 5.859% | 5.113% | 3.829% | 3.542% | 1.199% | 0.9147% | 1.552% | 2.815% | 1.959% | 0.9425% | 0.8521% | 3.894% | 0.9875% | 1.504% | 1.569% | 1.801% | 3.098% | 5.964% | 8.692% | 3.667% | 1.745% | 1.382%
| PDM3 |  | 2.99% | -5.182% | -7.421% | 0.4149% | -0.4807% | -9.212% | -3.503% | -8.653% | 4.333% | 2.956% | 1.198% | 4.937% | -4.953% | -0.1323% | 7.691% | -0.8165% | -18.41% | 3.225% | 0.4149% | 0.07903% | 11.95% | -4.063% | -8.989% | 14.52% | 14.18% | -4.399%
| Z-Score (Absolute) |  | NA | 0.75 | 1.1 | 0.11 | 0.03 | 1.38 | 0.49 | 1.29 | 0.72 | 0.5 | 0.23 | 0.81 | 0.72 | 0.03 | 1.24 | 0.08 | 2.8 | 0.55 | 0.11 | 0.06 | NA | 0.58 | 1.34 | 2.29 | 2.24 | 0.63
Results from laboratories 1.01 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 625 (Certified Value 0.96 ppm) | [30 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.21; 1.06: 1.22; 1.07: 1.23; 1.08: 1.25; 1.09: 1.26; 1.10: 1.30 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.01: 0.8; 1.02: 0.96; 1.03: 0.96; 1.04: 0.94269; 1.05: 1; 1.06: 1.2613; 1.07: 1; 1.08: 1.1; 1.09: 1; 1.10: 0.95 |  |  |  | 1
| 1 | 2 | 1.01: 0.9; 1.02: 0.93; 1.03: 1.02; 1.04: 0.96756; 1.05: 1; 1.06: 1.2068; 1.07: 1; 1.08: 1.1; 1.09: 0.8; 1.10: 0.93 |  |  |  | 45
| 1 | 3 | 1.01: 1; 1.02: 1.02; 1.03: 1.04; 1.04: 0.94865; 1.05: 0.9; 1.06: 1.1644; 1.07: 1; 1.08: 1.2; 1.09: 1; 1.10: 0.94 |  |  |  | 16
| 1 | 4 | 1.01: 0.9; 1.02: 0.98; 1.03: 1.03; 1.04: 0.93958; 1.05: 1; 1.06: 1.4187; 1.07: 1; 1.08: 1.2; 1.09: 0.9; 1.10: 0.95 |  |  |  | 0.9574
| 1 | 5 | 1.01: 1; 1.02: 1.02; 1.03: 1.02; 1.04: 0.93697; 1.05: 1; 1.06: 1.2208; 1.07: 0.9; 1.08: 1.1; 1.09: 0.9; 1.10: 0.96 |  |  |  | 65
| 1 | 6 | 1.01: 0.9; 1.02: 0.96; 1.03: 1.01; 1.04: 0.92578; 1.05: 0.9; 1.06: 1.254; 1.07: 0.9; 1.08: 1.2; 1.09: 0.9; 1.10: 0.95
| Mean |  | 1.01: 0.91667; 1.02: 0.97833; 1.03: 1.0133; 1.04: 0.94354; 1.05: 0.96667; 1.06: 1.2544; 1.07: 0.96667; 1.08: 1.15; 1.09: 0.91667; 1.10: 0.94667
| Median |  | 1.01: 0.9; 1.02: 0.97; 1.03: 1.02; 1.04: 0.94113; 1.05: 1; 1.06: 1.2374; 1.07: 1; 1.08: 1.15; 1.09: 0.9; 1.10: 0.95
| Std Dev. |  | 1.01: 0.075277; 1.02: 0.036009; 1.03: 0.028048; 1.04: 0.013986; 1.05: 0.05164; 1.06: 0.087781; 1.07: 0.05164; 1.08: 0.054772; 1.09: 0.075277; 1.10: 0.010328
| Rel.Std.Dev. |  | 1.01: 8.212%; 1.02: 3.681%; 1.03: 2.768%; 1.04: 1.482%; 1.05: 5.342%; 1.06: 6.998%; 1.07: 5.342%; 1.08: 4.763%; 1.09: 8.212%; 1.10: 1.091%
| PDM3 |  | 1.01: -4.255%; 1.02: 2.186%; 1.03: 5.842%; 1.04: -1.448%; 1.05: 0.9678%; 1.06: 31.02%; 1.07: 0.9678%; 1.08: 20.12%; 1.09: -4.255%; 1.10: -1.121%
| Z-Score (Absolute) |  | 1.01: 0.97; 1.02: 0.23; 1.03: 0.9; 1.04: 0.45; 1.05: 0; 1.06: 5.56; 1.07: 0; 1.08: 3.54; 1.09: 0.97; 1.10: 0.39
| Analytical results for Zn in OREAS 625 (Certified Value 3.17 wt.%) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.1875 | 3.2 | 3.091 | 3.15 | 2.88 | 3.02 | 3.1 | 3.13 | 3.03 | 3.1847 | 3.28 | 2.9705 | 3.1835 | 3.0172 | 3.2726 | 3.34 | 3.05 | 3.38 | 3.3038 | 3.2665 | 3.047 | 3.1629 | 3.13 | 3.1 | 3.25 | 3.14 | 3.14 | > 1 | 3.18 | 1
| 1 | 2 | 2.9835 | 3.21 | 3.134 | 3.17 | 3.13 | 3 | 3.15 | 3.19 | 3.01 | 3.1603 | 3.3 | 3.0359 | 3.2008 | 3.2523 | 3.1509 | 3.33 | 3.02 | 3.27 | 3.2577 | 3.3205 | 3.059 | 3.171 | 3.14 | 3.28 | 3.23 | 3.222 | 3.17 | > 1 | 3.16 | 24
| 1 | 3 | 2.9155 | 3.18 | 3.033 | 3.16 | 3.13 | 3.04 | 3.14 | 3.14 | 3.1 | 3.1649 | 3.17 | 3.382 | 3.207 | 3.1019 | 3.3062 | 3.32 | 3.07 | 3.38 | 3.3025 | 3.2601 | 3.042 | 3.1526 | 3.16 | 3.17 | 3.21 | 3.138 | 3.15 | > 1 | 3.21 | 16
| 1 | 4 | 3.2385 | 3.13 | 3.091 | 3.14 | 3.14 | 3 | 3.12 | 3.16 | 3.24 | 3.1752 | 3.31 | 3.2726 | 3.1988 | 3.135 | 3.5352 | 3.33 | 3.03 | 3.26 | 3.3005 | 3.3064 | 3.031 | 3.1412 | 3.16 | 3.19 | 3.22 | 3.153 | 3.16 | > 1 | 3.16 | 3.1686
| 1 | 5 | 3.026 | 3.13 | 3.077 | 3.15 | 3.05 | 3.04 | 3.11 | 3.05 | 3.1 | 3.1623 | 3.3 | 3.3078 | 3.2476 | 3.1159 | 3.3977 | 3.32 | 3.04 | 3.29 | 3.3025 | 3.3177 | 3.029 | 3.1603 | 3.17 | 3.15 | 3.25 | 3.166 | 3.16 | > 1 | 3.21 | 66
| 1 | 6 | 3.077 | 3.16 | 3.122 | 3.16 | 3.12 | 3.01 | 3.12 | 3.18 | 3.03 | 3.1644 | 3.31 | 3.2207 | 3.2165 | 3.2084 | 3.0308 | 3.33 | 3.03 | 3.29 | 3.3014 | 3.2764 | 3.048 | 3.1078 | 3.14 | 3.25 | 3.27 | 3.136 | 3.16 | > 1 | 3.14
| Mean |  | 3.0713 | 3.1683 | 3.0913 | 3.155 | 3.075 | 3.0183 | 3.1233 | 3.1417 | 3.085 | 3.1686 | 3.2783 | 3.1982 | 3.209 | 3.1385 | 3.2822 | 3.3283 | 3.04 | 3.3117 | 3.2947 | 3.2913 | 3.0427 | 3.1493 | 3.15 | 3.19 | 3.2383 | 3.1592 | 3.1567 |  | 3.1767
| Median |  | 3.0515 | 3.17 | 3.091 | 3.155 | 3.125 | 3.015 | 3.12 | 3.15 | 3.065 | 3.1646 | 3.3 | 3.2466 | 3.2039 | 3.1255 | 3.2894 | 3.33 | 3.035 | 3.29 | 3.302 | 3.2914 | 3.0445 | 3.1564 | 3.15 | 3.18 | 3.24 | 3.1465 | 3.16 |  | 3.17
| Std Dev. |  | 0.1229 | 0.034303 | 0.035691 | 0.010488 | 0.10095 | 0.018348 | 0.018619 | 0.050365 | 0.085029 | 0.0094093 | 0.054191 | 0.16126 | 0.021757 | 0.082962 | 0.17802 | 0.0075277 | 0.017889 | 0.054191 | 0.018177 | 0.026788 | 0.01129 | 0.022696 | 0.015492 | 0.06603 | 0.022286 | 0.032817 | 0.010328 |  | 0.028752
| Rel.Std.Dev. |  | 4.002% | 1.083% | 1.155% | 0.3324% | 3.283% | 0.6079% | 0.5961% | 1.603% | 2.756% | 0.297% | 1.653% | 5.042% | 0.678% | 2.643% | 5.424% | 0.2262% | 0.5884% | 1.636% | 0.5517% | 0.8139% | 0.3711% | 0.7207% | 0.4918% | 2.07% | 0.6882% | 1.039% | 0.3272% |  | 0.9051%
| PDM3 |  | -3.07% | -0.009152% | -2.439% | -0.4299% | -2.955% | -4.743% | -1.429% | -0.8507% | -2.639% | 0.0003157% | 3.462% | 0.935% | 1.275% | -0.9522% | 3.585% | 5.04% | -4.059% | 4.514% | 3.98% | 3.871% | -3.975% | -0.6098% | -0.5877% | 0.6746% | 2.2% | -0.2984% | -0.3773% |  | 0.2538%
| Z-Score (Absolute) |  | 0.99 | 0.12 | 0.76 | 0.03 | 0.95 | 1.6 | 0.4 | 0.19 | 0.84 | 0.12 | 1.38 | 0.46 | 0.59 | 0.22 | 1.42 | 1.95 | 1.35 | 1.76 | 1.57 | 1.53 | 1.32 | 0.1 | 0.09 | 0.37 | 0.92 | 0.01 | 0.01 | NA | 0.21
| Analytical results for Zr in OREAS 625 (Certified Value 229 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 226 | 213 | 232.5 | 249 | 232 | 224 | 221 | 224 | 235 | 232.4 | 221.6 | 210.2 | 221.1598 | 231.3358 | 227.9 | 236 | 229.9 | 197 | 245 | 234 | 224.46 | 236 | 247 | 245 | 236.3 | 221.2 | 224 | 228.6 |  | 1
| 1 | 2 | 237 | 228 | 224.4 | 250 | 213 | 226 | 215 | 230 | 236 | 223.1 | 220.8 | 226.8 | 216.9532 | 228.1223 | 219.9 | 231 | 231 | 193 | 245.2 | 230 | 224.6 | 237 | 256 | 244 | 221.3 | 224.4 | 224 | 230.5 |  | 16
| 1 | 3 | 222 | 222 | 239.1 | 251 | 230 | 211 | 212 | 231 | 239 | 240.1 | 222 | 219.2 | 217.0746 | 230.946 | 218.7 | 239 | 227.4 | 196 | 249.3 | 239 | 231.52 | 236 | 258 | 246 | 234.1 | 217.8 | 221 | 238.9 |  | 16
| 1 | 4 | 234 | 223 | 237.3 | 256 | 232 | 223 | 215 | 228 | 236 | 230.6 | 222 | 220.3 | 212.3138 | 222.1512 | 218.4 | 235 | 234 | 194 | 238.3 | 231 | 227.09 | 237 | 260 | 246 | 228.5 | 224.6 | 222 | 225.4 |  | 229.2116
| 1 | 5 | 223 | 227 | 233.4 | 248 | 230 | 218 | 210 | 245 | 237 | 231.4 | 221 | 221.1 | 210.151 | 222.5567 | 220.7 | 241 | 231.8 | 196 | 254.4 | 235 | 226.68 | 236 | 256 | 242 | 234.3 | 218.2 | 224 | 231.3 |  | 67
| 1 | 6 | 231 | 210 | 234.4 | 249 | 234 | 217 | 227 | 226 | 241 | 227.6 | 215.4 | 222.2 | 210.4291 | 229.4799 | 216.7 | 236 | 233.1 | 195 | 241.2 | 231 | 230.17 | 232 | 262 | 249 | 228.9 | 223 | 220 | 230.3
| Mean |  | 228.8333 | 220.5 | 233.5167 | 250.5 | 228.5 | 219.8333 | 216.6667 | 230.6667 | 237.3333 | 230.8667 | 220.4667 | 219.9667 | 214.6803 | 227.432 | 220.3833 | 236.3333 | 231.2 | 195.1667 | 245.5667 | 233.3333 | 227.42 | 235.6667 | 256.5 | 245.3333 | 230.5667 | 221.5333 | 222.5 | 230.8333
| Median |  | 228.5 | 222.5 | 233.9 | 249.5 | 231 | 220.5 | 215 | 229 | 236.5 | 231 | 221.3 | 220.7 | 214.6335 | 228.8011 | 219.3 | 236 | 231.4 | 195.5 | 245.1 | 232.5 | 226.885 | 236 | 257 | 245.5 | 231.5 | 222.1 | 223 | 230.4
| Std Dev. |  | 6.1128 | 7.3959 | 5.1074 | 2.881 | 7.7395 | 5.5648 | 6.2823 | 7.4744 | 2.2509 | 5.6383 | 2.5319 | 5.4599 | 4.4061 | 4.0967 | 3.9286 | 3.4448 | 2.3656 | 1.472 | 5.7322 | 3.3862 | 2.8894 | 1.8619 | 5.2058 | 2.3381 | 5.5175 | 2.9978 | 1.7607 | 4.4765
| Rel.Std.Dev. |  | 2.671% | 3.354% | 2.187% | 1.15% | 3.387% | 2.531% | 2.9% | 3.24% | 0.9484% | 2.442% | 1.148% | 2.482% | 2.052% | 1.801% | 1.783% | 1.458% | 1.023% | 0.7542% | 2.334% | 1.451% | 1.271% | 0.7901% | 2.03% | 0.953% | 2.393% | 1.353% | 0.7913% | 1.939%
| PDM3 |  | -0.165% | -3.801% | 1.878% | 9.288% | -0.3105% | -4.092% | -5.473% | 0.6348% | 3.543% | 0.7221% | -3.815% | -4.033% | -6.34% | -0.7764% | -3.852% | 3.107% | 0.8675% | -14.85% | 7.135% | 1.798% | -0.7816% | 2.816% | 11.91% | 7.034% | 0.5912% | -3.35% | -2.928% | 0.7075%
| Z-Score (Absolute) |  | 0.08 | 0.84 | 0.35 | 1.89 | 0.11 | 0.9 | 1.18 | 0.09 | 0.69 | 0.11 | 0.84 | 0.88 | 1.37 | 0.21 | 0.85 | 0.6 | 0.14 | 3.14 | 1.44 | 0.33 | 0.21 | 0.54 | 2.44 | 1.42 | 0.08 | 0.74 | 0.65 | 0.1
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 625 (Certified Value 11.7 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | 2
| 1 | 1 | 11.9 | 10.8 | 10.9 | 13.05 | 11.5 | 13.4 | 10.85 | 11.7 | 11.5 | 11.8 | 11.923 | 10.6 | 11.8855 | 11.38 | 11.5302 | 10.59 | 11.8 | 10.4 | 16.08 | 11.8 | 12.1 | 13.0572 | 11.8 | 11.7 | 11.8 | > 10 | 10.39 | 12 | 11.5 | 1
| 1 | 2 | 11.5 | 10.8 | 11.2 | 13.05 | 11.9 | 12.95 | 11.1 | 12.1 | 12.3 | 11.77 | 11.789 | 10.8 | 11.6029 | 12.6 | 11.765 | 10.86 | 11.7 | 10.4 | 15.97 | 12 | 12.1 | 12.7548 | 11.84 | 11.6 | 11.7 | > 10 | 11.01 | 12 | 11.5 | 1
| 1 | 3 | 11.7 | 10.9 | 11.4 | 12.75 | 12.1 | 13.3 | 10.85 | 12 | 12.1 | 11.71 | 12.078 | 10.2 | 11.5249 | 10.7 | 12.1592 | 10.76 | 11.8 | 10.8 | 16.05 | 11.7 | 12.2 | 12.6792 | 12.03 | 12 | 11.8 | > 10 | 10.47 | 12 | 11.4 | 16
| 1 | 4 | 11.7 | 11.1 | 11.1 | 13.05 | 11.8 | 13.2 | 10.8 | 11.8 | 12 | 11.99 | 12.03 | 11 | 11.5135 | 10.65 | 11.7334 | 10.91 | 11.6 | 10.5 | 16.01 | 11.6 | 11.8 | 12.6576 | 11.57 | 11.8 | 11.7 | > 10 | 10.99 | 12 | 11.5 | 11.6562
| 1 | 5 | 12 | 11 | 11 | 12.85 | 11.75 | 13.15 | 10.95 | 11.85 | 11.7 | 12.14 | 11.886 | 10.8 | 11.5697 | 12.1735 | 11.3368 | 10.44 | 11.9 | 10.7 | 16.04 | 11.7 | 12 | 13.1436 | 11.98 | 11.7 | 11.8 | > 10 | 10.62 | 12 | 11.3 | 69
| 1 | 6 | 11.7 | 11.3 | 11 | 12.9 | 12 | 12.95 | 11 | 11.65 | 12.4 | 11.38 | 11.79 | 10.8 | 11.3237 | 12.07 | 11.9629 | 10.89 | 11.9 | 10.5 | 16.08 | 11.7 | 12.1 | 12.7224 | 11.92 | 11.5 | 11.9 | > 10 | 10.54 | 12 | 11.5
| Mean |  | 11.75 | 10.9833 | 11.1 | 12.9417 | 11.8417 | 13.1583 | 10.925 | 11.85 | 12 | 11.7983 | 11.916 | 10.7 | 11.57 | 11.5956 | 11.7479 | 10.7417 | 11.7833 | 10.55 | 16.0383 | 11.75 | 12.05 | 12.8358 | 11.8567 | 11.7167 | 11.7833 |  | 10.67 | 12 | 11.45
| Median |  | 11.7 | 10.95 | 11.05 | 12.975 | 11.85 | 13.175 | 10.9 | 11.825 | 12.05 | 11.785 | 11.9045 | 10.8 | 11.5473 | 11.725 | 11.7492 | 10.81 | 11.8 | 10.5 | 16.045 | 11.7 | 12.1 | 12.7386 | 11.88 | 11.7 | 11.8 |  | 10.58 | 12 | 11.5
| Std Dev. |  | 0.17607 | 0.19408 | 0.17889 | 0.12813 | 0.21075 | 0.1828 | 0.11292 | 0.17321 | 0.34641 | 0.25965 | 0.12016 | 0.27568 | 0.18251 | 0.81371 | 0.29402 | 0.18904 | 0.1169 | 0.16432 | 0.042622 | 0.13784 | 0.13784 | 0.2095 | 0.16428 | 0.17224 | 0.075277 |  | 0.26676 | 0 | 0.083666
| Rel.Std.Dev. |  | 1.498% | 1.767% | 1.612% | 0.99% | 1.78% | 1.389% | 1.034% | 1.462% | 2.887% | 2.201% | 1.008% | 2.576% | 1.577% | 7.017% | 2.503% | 1.76% | 0.9921% | 1.558% | 0.2658% | 1.173% | 1.144% | 1.632% | 1.386% | 1.47% | 0.6388% |  | 2.5% | 0% | 0.7307%
| PDM3 |  | 0.8047% | -5.773% | -4.772% | 11.03% | 1.591% | 12.89% | -6.273% | 1.663% | 2.949% | 1.219% | 2.229% | -8.203% | -0.7392% | -0.5201% | 0.7868% | -7.846% | 1.091% | -9.49% | 37.59% | 0.8047% | 3.378% | 10.12% | 1.72% | 0.5187% | 1.091% |  | -8.461% | 2.949% | -1.769%
| Z-Score (Absolute) |  | 0 | 2.07 | 1.75 | 3.21 | 0.25 | 3.8 | 2.23 | 0.27 | 0.67 | 0.13 | 0.45 | 2.83 | 0.49 | 0.42 | 0.01 | 2.72 | 0.09 | 3.24 | 11.57 | 0 | 0.81 | 2.93 | 0.29 | 0.09 | 0.09 | NA | 2.91 | NA | 0.81
| Analytical results for Al in OREAS 625 (Certified Value 0.565 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.63 | 0.53 | 0.54 | 0.54 | 0.58 | 0.53 | 0.51 | 0.52 | 0.41 | 0.5574 | 0.57 | 0.7913 | 0.585 | 0.72627 | 0.53 | 0.728 | 0.57 | 0.5841 | 0.96 | 1.0861 | 0.59 | 0.5 | 0.51 | 0.51 | 0.61 | 0.608 | 0.57 |  |  | 1
| 1 | 2 | 0.63 | 0.52 | 0.5 | 0.55 | 0.59 | 0.53 | 0.52 | 0.53 | 0.43 | 0.5891 | 0.56 | 0.7694 | 0.5721 | 0.74418 | 0.54 | 0.725 | 0.57 | 0.5621 | 0.98 | 1.0182 | 0.57 | 0.52 | 0.5 | 0.51 | 0.57 | 0.588 | 0.56 |  |  | 0
| 1 | 3 | 0.62 | 0.52 | 0.55 | 0.55 | 0.6 | 0.53 | 0.5 | 0.53 | 0.43 | 0.5807 | 0.58 | 0.7475 | 0.5741 | 0.77114 | 0.59 | 0.738 | 0.59 | 0.5736 | 0.98 | 1.0212 | 0.59 | 0.52 | 0.5 | 0.51 | 0.58 | 0.588 | 0.53 |  |  | 16
| 1 | 4 | 0.64 | 0.52 | 0.51 | 0.56 | 0.59 | 0.53 | 0.5 | 0.53 | 0.44 | 0.6015 | 0.58 | 0.758 | 0.57099 | 0.77568 | 0.55 | 0.733 | 0.57 | 0.5744 | 0.98 | 1.0025 | 0.57 | 0.52 | 0.48 | 0.51 | 0.61 | 0.592 | 0.58 |  |  | 0.56512
| 1 | 5 | 0.63 | 0.54 | 0.54 | 0.54 | 0.59 | 0.52 | 0.5 | 0.53 | 0.4 | 0.5885 | 0.57 | 0.7423 | 0.57872 | 0.73288 | 0.55 | 0.733 | 0.59 | 0.5796 | 0.99 | 1.0271 | 0.59 | 0.52 | 0.51 | 0.5 | 0.59 | 0.6 | 0.54 |  |  | 70
| 1 | 6 | 0.63 | 0.54 | 0.55 | 0.53 | 0.59 | 0.52 | 0.5 | 0.52 | 0.43 | 0.5614 | 0.55 | 0.7899 | 0.56801 | 0.73166 | 0.61 | 0.723 | 0.57 | 0.5786 | 0.95 | 1.0093 | 0.58 | 0.51 | 0.5 | 0.52 | 0.61 | 0.599 | 0.57
| Mean |  | 0.63 | 0.52833 | 0.53167 | 0.545 | 0.59 | 0.52667 | 0.505 | 0.52667 | 0.42333 | 0.57977 | 0.56833 | 0.7664 | 0.57482 | 0.74697 | 0.56167 | 0.73 | 0.57667 | 0.5754 | 0.97333 | 1.0274 | 0.58167 | 0.515 | 0.5 | 0.51 | 0.595 | 0.59583 | 0.55833
| Median |  | 0.63 | 0.525 | 0.54 | 0.545 | 0.59 | 0.53 | 0.5 | 0.53 | 0.43 | 0.5846 | 0.57 | 0.7637 | 0.5731 | 0.73853 | 0.55 | 0.7305 | 0.57 | 0.5765 | 0.98 | 1.0197 | 0.585 | 0.52 | 0.5 | 0.51 | 0.6 | 0.5955 | 0.565
| Std Dev. |  | 0.0063246 | 0.0098319 | 0.02137 | 0.010488 | 0.0063246 | 0.005164 | 0.0083666 | 0.005164 | 0.015055 | 0.017171 | 0.01169 | 0.020928 | 0.0061285 | 0.021343 | 0.031252 | 0.0056569 | 0.010328 | 0.0075459 | 0.015055 | 0.030058 | 0.0098319 | 0.0083666 | 0.010954 | 0.0063246 | 0.017607 | 0.0079099 | 0.019408
| Rel.Std.Dev. |  | 1.004% | 1.861% | 4.019% | 1.924% | 1.072% | 0.9805% | 1.657% | 0.9805% | 3.556% | 2.962% | 2.057% | 2.731% | 1.066% | 2.857% | 5.564% | 0.7749% | 1.791% | 1.311% | 1.547% | 2.926% | 1.69% | 1.625% | 2.191% | 1.24% | 2.959% | 1.328% | 3.476%
| PDM3 |  | 11.48% | -6.51% | -5.92% | -3.56% | 4.402% | -6.805% | -10.64% | -6.805% | -25.09% | 2.592% | 0.5684% | 35.62% | 1.717% | 32.18% | -0.6112% | 29.18% | 2.043% | 1.819% | 72.23% | 81.8% | 2.928% | -8.869% | -11.52% | -9.754% | 5.287% | 5.435% | -1.201%
| Z-Score (Absolute) |  | 0.8 | 0.67 | 0.63 | 0.43 | 0.22 | 0.7 | 1.01 | 0.7 | 2.2 | 0.07 | 0.09 | 2.78 | 0 | 2.5 | 0.19 | 2.25 | 0.03 | 0.01 | 5.78 | 6.56 | 0.1 | 0.87 | 1.09 | 0.94 | 0.29 | 0.3 | 0.24
| Analytical results for As in OREAS 625 (Certified Value 134 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | 0
| 1 | 1 | 137 | 126 | 145.4 | 139.5 | 137 | 139 | 137.5 | 144.5 | 123 | 130 | 141.9 | 130 | 140.8418 | 102.995 | 132.7621 | 134.8 | 147 | 122.5 | 114 | 118.9 | 137 | 128.26 | 137 | 120 | 131 | 120 | 136.7 | 129 | 152 | 1
| 1 | 2 | 136 | 123 | 143.6 | 143 | 142.5 | 134 | 140.5 | 151 | 133 | 134 | 141 | 132 | 136.764 | 105.142 | 136.1661 | 133.7 | 146 | 119.7 | 103 | 118.1 | 145 | 125.85 | 134 | 126 | 132 | 123 | 136.3 | 129 | 150.2 | 21
| 1 | 3 | 134 | 121 | 141.6 | 140 | 144 | 137.5 | 136 | 141 | 130 | 131 | 149.8 | 125 | 136.0942 | 99.982 | 140.2996 | 132.8 | 144 | 124.2 | 108 | 118.7 | 145 | 121.9 | 139 | 121 | 132 | 124 | 131.6 | 123 | 159.6 | 16
| 1 | 4 | 136 | 126 | 140.5 | 143.5 | 143 | 138 | 134.5 | 147 | 123 | 134 | 145.7 | 135 | 137.9845 | 101.053 | 137.6176 | 131.6 | 146 | 122 | 112 | 120.8 | 139 | 125.55 | 129 | 120 | 131 | 124 | 141.6 | 125 | 149 | 133.6096
| 1 | 5 | 137 | 123 | 145.5 | 138 | 141 | 134.5 | 136 | 146 | 127 | 135 | 144.8 | 133 | 136.6012 | 104.537 | 131.7999 | 136.1 | 148 | 125.1 | 118 | 120.3 | 140 | 125.92 | 133 | 119 | 131 | 119 | 136.6 | 128 | 154.2 | 71
| 1 | 6 | 136 | 126 | 143.5 | 137.5 | 146.5 | 134.5 | 136 | 138 | 131 | 135 | 139.7 | 133 | 137.4379 | 101.309 | 136.8822 | 133.4 | 147 | 119.4 | 109 | 119.6 | 143 | 123.78 | 136 | 130 | 132 | 120 | 138.2 | 133 | 152.7
| Mean |  | 136 | 124.1667 | 143.35 | 140.25 | 142.3333 | 136.25 | 136.75 | 144.5833 | 127.8333 | 133.1667 | 143.8167 | 131.3333 | 137.6206 | 102.503 | 135.9212 | 133.7333 | 146.3333 | 122.15 | 110.6667 | 119.4 | 141.5 | 125.21 | 134.6667 | 122.6667 | 131.5 | 121.6667 | 136.8333 | 127.8333 | 152.95
| Median |  | 136 | 124.5 | 143.55 | 139.75 | 142.75 | 136 | 136 | 145.25 | 128.5 | 134 | 143.35 | 132.5 | 137.101 | 102.152 | 136.5241 | 133.55 | 146.5 | 122.25 | 110.5 | 119.25 | 141.5 | 125.7 | 135 | 120.5 | 131.5 | 121.5 | 136.65 | 128.5 | 152.35
| Std Dev. |  | 1.0954 | 2.137 | 2.0047 | 2.505 | 3.1885 | 2.1622 | 2.0676 | 4.5871 | 4.2151 | 2.137 | 3.7113 | 3.5024 | 1.7114 | 2.0604 | 3.1621 | 1.5667 | 1.3663 | 2.3072 | 5.2026 | 1.0237 | 3.3317 | 2.1602 | 3.5024 | 4.3665 | 0.54772 | 2.2509 | 3.2352 | 3.4881 | 3.7394
| Rel.Std.Dev. |  | 0.8055% | 1.721% | 1.398% | 1.786% | 2.24% | 1.587% | 1.512% | 3.173% | 3.297% | 1.605% | 2.581% | 2.667% | 1.244% | 2.01% | 2.326% | 1.172% | 0.9337% | 1.889% | 4.701% | 0.8574% | 2.355% | 1.725% | 2.601% | 3.56% | 0.4165% | 1.85% | 2.364% | 2.729% | 2.445%
| PDM3 |  | 1.789% | -7.068% | 7.29% | 4.97% | 6.529% | 1.976% | 2.35% | 8.213% | -4.323% | -0.3315% | 7.639% | -1.704% | 3.002% | -23.28% | 1.73% | 0.09262% | 9.523% | -8.577% | -17.17% | -10.64% | 5.906% | -6.287% | 0.7912% | -8.19% | -1.579% | -8.939% | 2.413% | -4.323% | 14.48%
| Z-Score (Absolute) |  | 0.13 | 1.04 | 0.86 | 0.55 | 0.76 | 0.16 | 0.21 | 0.98 | 0.67 | 0.15 | 0.9 | 0.33 | 0.29 | 3.17 | 0.12 | 0.09 | 1.15 | 1.24 | 2.37 | 1.51 | 0.67 | 0.93 | 0 | 1.18 | 0.31 | 1.28 | 0.21 | 0.67 | 1.8
| Analytical results for B in OREAS 625 (Certified Value < 10 ppm) | [30 empty] | Certified
B | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.26 | 1.27 | [14 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [14 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 15g | 0.2g | 0.5g | 01g | 0.25g | 0.5g | [14 empty] | 1
| 1 | 1 | 16 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 28.82 | < 10 | < 10 | < 10 | [14 empty] | 1
| 1 | 2 | 15 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 32.84 | < 10 | < 10 | < 10 | [14 empty] | 3
| 1 | 3 | 15 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 24.22 | < 10 | < 10 | 11 | [14 empty] | 16
| 1 | 4 | 15 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 33.43 | < 10 | < 10 | 11 | [14 empty] | < 10
| 1 | 5 | 14 | 2 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 24.78 | < 10 | < 10 | < 10 | [14 empty] | 72
| 1 | 6 | 15 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 31.14 | < 10 | < 10 | < 10
| Mean |  | 1.01: 15; 1.02: 2.8333; 1.11: 2; 1.14: 29.205; 1.27: 11
| Median |  | 1.01: 15; 1.02: 3; 1.11: 2; 1.14: 29.98; 1.27: 11
| Std Dev. |  | 1.01: 0.63246; 1.02: 0.40825; 1.11: 0; 1.14: 3.9846; 1.27: 0
| Rel.Std.Dev. |  | 1.01: 4.216%; 1.02: 14.41%; 1.11: 0%; 1.14: 13.64%; 1.27: 0%
PDM3
| Z-Score (Absolute) |  | 3.37 | 0.73 | 1.69 | 0 | 0 | 0 | 0 | 0 | 0.84 | 1.01 | 1.35 | 8.16 | 0 | 0 | 0.67
| Analytical results for Ba in OREAS 625 (Indicative Value 77 ppm) | [30 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 0
| 1 | 1 | 16 | 4 | 57 | 50 | 100 | 30 | 80 | 40 | 182 | 24 | 23.6 | 477 | 220.5751 | 337.46 | 45.334 | 397.1 | 470 | 32 | 377 | 134 | 74 | 121 | 150 | 99 | 374 | 376 |  |  |  | 1
| 1 | 2 | 15 | 3.5 | 54 | 50 | 60 | 30 | 60 | 30 | 216 | 22 | 22.3 | 468 | 222.1778 | 338.53 | 44.2509 | 403.1 | 466 | 30 | 369 | 101 | 74 | 120 | 151 | 100 | 355 | 382 |  |  |  | 4
| 1 | 3 | 15 | 5.4 | 59 | 40 | 70 | 30 | 70 | 60 | 221 | 21 | 31.3 | 456 | 233.2856 | 336.64 | 46.8032 | 420.9 | 471 | 31 | 372 | 112 | 77 | 124 | 153 | 102 | 362 | 359 |  |  |  | 16
| 1 | 4 | 15 | 4 | 52 | 40 | 70 | 20 | 70 | 40 | 202 | 21 | 24.3 | 467 | 225.3622 | 346.75 | 44.17 | 406.5 | 470 | 31 | 375 | 132 | 76 | 121 | 151 | 110 | 350 | 378 |  |  |  | 76.9419
| 1 | 5 | 16 | 4.1 | 56 | 40 | 70 | 30 | 70 | 50 | 187 | 22 | 23.9 | 456 | 233.8578 | 347.7 | 45.1929 | 397.7 | 470 | 31 | 372 | 119 | 76 | 123 | 152 | 102 | 367 | 368 |  |  |  | 12
| 1 | 6 | 16 | 3.8 | 58 | 40 | 90 | 40 | 80 | 60 | 215 | 21 | 23.3 | 483 | 225.7745 | 343.62 | 46.4335 | 432.2 | 471 | 33 | 374 | 95 | 73 | 122 | 156 | 102 | 364 | 382
| Mean |  | 15.5 | 4.1333 | 56 | 43.3333 | 76.6667 | 30 | 71.6667 | 46.6667 | 203.8333 | 21.8333 | 24.7833 | 467.8333 | 226.8388 | 341.7833 | 45.3641 | 409.5833 | 469.6667 | 31.3333 | 373.1667 | 115.5 | 75 | 121.8333 | 152.1667 | 102.5 | 362 | 374.1667
| Median |  | 15.5 | 4 | 56.5 | 40 | 70 | 30 | 70 | 45 | 208.5 | 21.5 | 23.75 | 467.5 | 225.5684 | 341.075 | 45.2634 | 404.8 | 470 | 31 | 373 | 115.5 | 75 | 121.5 | 151.5 | 102 | 363 | 377
| Std Dev. |  | 0.54772 | 0.65625 | 2.6077 | 5.164 | 15.0555 | 6.3246 | 7.5277 | 12.1106 | 16.3146 | 1.169 | 3.2634 | 10.9072 | 5.5696 | 4.874 | 1.0872 | 14.0582 | 1.8619 | 1.0328 | 2.7869 | 15.9342 | 1.5492 | 1.472 | 2.137 | 3.8859 | 8.5557 | 9.0425
| Rel.Std.Dev. |  | 3.534% | 15.88% | 4.657% | 11.92% | 19.64% | 21.08% | 10.5% | 25.95% | 8.004% | 5.354% | 13.17% | 2.331% | 2.455% | 1.426% | 2.397% | 3.432% | 0.3964% | 3.296% | 0.7468% | 13.8% | 2.066% | 1.208% | 1.404% | 3.791% | 2.363% | 2.417%
| PDM3 |  | -79.85% | -94.63% | -27.22% | -43.68% | -0.3577% | -61.01% | -6.856% | -39.35% | 164.9% | -71.62% | -67.79% | 508% | 194.8% | 344.2% | -41.04% | 432.3% | 510.4% | -59.28% | 385% | 50.11% | -2.524% | 58.34% | 97.77% | 33.22% | 370.5% | 386.3%
| Z-Score (Absolute) |  | 0.78 | 0.9 | 0.36 | 0.49 | 0.14 | 0.63 | 0.19 | 0.45 | 1.21 | 0.72 | 0.69 | 4 | 1.45 | 2.67 | 0.47 | 3.39 | 4.02 | 0.62 | 3 | 0.27 | 0.15 | 0.34 | 0.66 | 0.14 | 2.88 | 3.01
| Analytical results for Be in OREAS 625 (Certified Value 0.61 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [7 empty] | 3
| 1 | 1 | < 5 | 0.6 | 0.6 | 0.6 | 0.62 | 0.59 | 0.59 | 0.53 | 0.59 | 0.65708 | 0.61314 | < 2 | 0.71 | < 1 | < 5 | 0.8 | 0.8 | 0.6 | 0.7 | 0.65 | 1.3 | 0.84 | [7 empty] | 1
| 1 | 2 | < 5 | 0.6 | 0.61 | 0.62 | 0.62 | 0.59 | 0.6 | 0.48 | 0.64 | 0.62887 | 0.68893 | < 2 | 0.75 | < 1 | < 5 | 0.9 | 0.8 | 0.7 | 0.8 | 0.65 | 1.3 | 0.83 | [7 empty] | 14
| 1 | 3 | < 5 | 0.6 | 0.6 | 0.63 | 0.63 | 0.58 | 0.59 | 0.59 | 0.58 | 0.63133 | 0.69143 | < 2 | 0.72 | < 1 | < 5 | 0.9 | 0.7 | 0.6 | 0.7 | 0.63 | < 0.5 | 0.87 | [7 empty] | 16
| 1 | 4 | < 5 | 0.7 | 0.61 | 0.61 | 0.63 | 0.58 | 0.59 | 0.54 | 0.57 | 0.64728 | 0.59325 | < 2 | 0.76 | < 1 | < 5 | 0.8 | 0.7 | 0.6 | 0.8 | 0.69 | 1.3 | 0.82 | [7 empty] | 0.61007
| 1 | 5 | < 5 | 0.6 | 0.59 | 0.61 | 0.6 | 0.58 | 0.59 | 0.54 | 0.6 | 0.64746 | 0.6413 | < 2 | 0.74 | < 1 | < 5 | 0.8 | 0.7 | 0.7 | 0.7 | 0.66 | 1.3 | 0.72 | [7 empty] | 73
| 1 | 6 | < 5 | 0.6 | 0.58 | 0.63 | 0.62 | 0.58 | 0.58 | 0.6 | 0.59 | 0.64259 | 0.70754 | < 2 | 0.74 | < 1 | < 5 | 0.8 | 0.7 | 0.7 | 0.7 | 0.65 | 1.3 | 0.84
| Mean |  |  | 0.61667 | 0.59833 | 0.61667 | 0.62 | 0.58333 | 0.59 | 0.54667 | 0.595 | 0.64243 | 0.65593 |  | 0.73667 |  |  | 0.83333 | 0.73333 | 0.65 | 0.73333 | 0.655 | 1.3 | 0.82
| Median |  |  | 0.6 | 0.6 | 0.615 | 0.62 | 0.58 | 0.59 | 0.54 | 0.59 | 0.64493 | 0.66512 |  | 0.74 |  |  | 0.8 | 0.7 | 0.65 | 0.7 | 0.65 | 1.3 | 0.835
| Std Dev. |  |  | 0.040825 | 0.01169 | 0.012111 | 0.010954 | 0.005164 | 0.0063246 | 0.043665 | 0.02429 | 0.010683 | 0.046875 |  | 0.018619 |  |  | 0.05164 | 0.05164 | 0.054772 | 0.05164 | 0.019748 | 0 | 0.051769
| Rel.Std.Dev. |  |  | 6.62% | 1.954% | 1.964% | 1.767% | 0.8853% | 1.072% | 7.988% | 4.082% | 1.663% | 7.146% |  | 2.527% |  |  | 6.197% | 7.042% | 8.427% | 7.042% | 3.015% | 0% | 6.313%
| PDM3 |  |  | 1.081% | -1.924% | 1.081% | 1.628% | -4.383% | -3.29% | -10.39% | -2.47% | 5.305% | 7.517% |  | 20.75% |  |  | 36.6% | 20.2% | 6.545% | 20.2% | 7.365% | 113.1% | 34.41%
| Z-Score (Absolute) |  | 28.29 | NA | 0.5 | 0.22 | 0.17 | 0.72 | 0.62 | 1.28 | 0.55 | 0.17 | 0.37 | 5.58 | 1.6 | 1.99 | 28.29 | NA | NA | NA | NA | 0.36 | NA | 2.86
| Analytical results for Bi in OREAS 625 (Certified Value 8.81 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 8 | 8.47 | 8.7 | 9.41 | 8.41 | 8.44 | 8.28 | 7.56 | 9.21 | 8.64 | 9.02 | 8 | 8.4367 | 6.18 | 8.7717 | 5.8 | 9.55 | 7.5 | 7.17 | 8.83 | 9.23 | 8.81 | 9 | 9 | 8.49 | 8.75 | 11 | 10 |  | 1
| 1 | 2 | 9 | 8.61 | 9 | 9.46 | 8.69 | 8.52 | 9.12 | 7.75 | 9.25 | 8.96 | 8.76 | 9 | 8.5103 | 6.2 | 8.8867 | 8.3 | 9.49 | 7.4 | 6.98 | 8.75 | 8.87 | 9.24 | 9 | 8.74 | 8.84 | 8.96 | 12 | 8 |  | 35
| 1 | 3 | 8 | 8.86 | 9 | 9.35 | 8.9 | 8.95 | 8.02 | 7.62 | 9.27 | 8.82 | 9.22 | 6 | 8.5279 | 6.16 | 8.7914 | < 5 | 9.56 | 7.9 | 7.15 | 8.8 | 9 | 9.21 | 9 | 8.86 | 8.64 | 8.54 | < 5 | 7 |  | 16
| 1 | 4 | 8 | 8.8 | 8.8 | 9.74 | 8.69 | 8.78 | 8.16 | 7.7 | 9.26 | 8.85 | 9.08 | 7 | 8.6107 | 5.48 | 8.8625 | 6.6 | 9.45 | 7.5 | 7.09 | 8.67 | 8.78 | 9.45 | 9 | 8.68 | 8.73 | 8.97 | 8 | 8 |  | 8.8116
| 1 | 5 | 9 | 8.68 | 8.9 | 9.34 | 8.56 | 8.83 | 8.31 | 7.56 | 8.71 | 8.98 | 8.95 | 7 | 8.6437 | 6.89 | 8.7296 | 8.4 | 9.56 | 7.7 | 7.11 | 8.34 | 8.8 | 9.05 | 9 | 8.56 | 8.39 | 8.89 | 14 | 10 |  | 74
| 1 | 6 | 8 | 8.83 | 8.7 | 9.37 | 9.39 | 8.71 | 8.25 | 7.43 | 9.26 | 8.61 | 8.76 | 5 | 8.5009 | 6.19 | 8.8044 | 8.3 | 9.59 | 7.6 | 7.08 | 8.64 | 8.86 | 8.95 | 9 | 8.78 | 8.34 | 9.1 | 9 | 9
| Mean |  | 8.3333 | 8.7083 | 8.85 | 9.445 | 8.7733 | 8.705 | 8.3567 | 7.6033 | 9.16 | 8.81 | 8.965 | 7 | 8.5384 | 6.1833 | 8.8077 | 7.48 | 9.5333 | 7.6 | 7.0967 | 8.6717 | 8.9233 | 9.1183 | 9 | 8.77 | 8.5717 | 8.8683 | 10.8 | 8.6667
| Median |  | 8 | 8.74 | 8.85 | 9.39 | 8.69 | 8.745 | 8.265 | 7.59 | 9.255 | 8.835 | 8.985 | 7 | 8.5191 | 6.185 | 8.7979 | 8.3 | 9.555 | 7.55 | 7.1 | 8.71 | 8.865 | 9.13 | 9 | 8.76 | 8.565 | 8.925 | 11 | 8.5
| Std Dev. |  | 0.5164 | 0.15065 | 0.13784 | 0.1511 | 0.34285 | 0.19274 | 0.38836 | 0.11396 | 0.22145 | 0.1562 | 0.18196 | 1.4142 | 0.076111 | 0.44608 | 0.058179 | 1.2029 | 0.052409 | 0.17889 | 0.066833 | 0.17815 | 0.16884 | 0.2286 | 0 | 0.15113 | 0.19753 | 0.19732 | 2.3875 | 1.2111
| Rel.Std.Dev. |  | 6.197% | 1.73% | 1.558% | 1.6% | 3.908% | 2.214% | 4.647% | 1.499% | 2.418% | 1.773% | 2.03% | 20.2% | 0.8914% | 7.214% | 0.6605% | 16.08% | 0.5497% | 2.354% | 0.9418% | 2.054% | 1.892% | 2.507% | 0% | 1.723% | 2.304% | 2.225% | 22.11% | 13.97%
| PDM3 |  | -5.428% | -1.172% | 0.4354% | 7.188% | -0.4347% | -1.21% | -5.163% | -13.71% | 3.953% | -0.01854% | 1.74% | -20.56% | -3.101% | -29.83% | -0.04435% | -15.11% | 8.19% | -13.75% | -19.46% | -1.588% | 1.268% | 3.481% | 2.138% | -0.4725% | -2.723% | 0.6435% | 22.57% | -1.645%
| Z-Score (Absolute) |  | NA | 0.21 | 0.27 | 2.29 | 0.01 | 0.22 | 1.41 | 3.97 | 1.33 | 0.14 | 0.66 | NA | 0.79 | 8.79 | 0.13 | 7.21 | 2.6 | 3.98 | 5.69 | 0.33 | 0.52 | 1.18 | NA | 0 | 0.67 | 0.33 | NA | NA
Results from laboratories 1.01, 1.12, 1.24, 1.28 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 625 (Certified Value 0.966 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.99 | 0.88 | 0.82 | 1.02 | 0.96 | 0.99 | 0.95 | 1 | 0.85 | 0.92 | 0.96 | 0.98 | 0.96849 | 1.0219 | 0.94 | 0.97 | 0.92 | 1.0119 | 0.9061 | 0.94 | 0.95 | 1.01 | 0.99 | 0.99 | 0.97 | 0.941 | 0.92 |  |  | 1
| 1 | 2 | 0.94 | 0.83 | 0.81 | 1.01 | 0.97 | 0.99 | 0.96 | 1.02 | 0.93 | 0.93 | 0.94 | 0.99 | 0.95654 | 1.0287 | 0.94 | 0.97 | 0.91 | 0.9907 | 0.9254 | 0.92 | 0.96 | 1.03 | 0.97 | 1 | 0.96 | 0.903 | 0.95 |  |  | 0
| 1 | 3 | 0.97 | 0.82 | 0.79 | 1.01 | 1 | 0.99 | 0.95 | 1.01 | 0.91 | 0.94 | 0.97 | 0.95 | 0.96031 | 1.0236 | 0.95 | 0.99 | 0.93 | 1.0025 | 0.9127 | 0.93 | 0.97 | 1.04 | 0.99 | 0.99 | 0.94 | 0.903 | 0.98 |  |  | 16
| 1 | 4 | 0.96 | 0.85 | 0.81 | 1.04 | 1 | 1 | 0.95 | 1.02 | 0.94 | 0.95 | 0.96 | 1.01 | 0.96142 | 0.9849 | 0.94 | 0.97 | 0.92 | 1.0036 | 0.931 | 0.94 | 0.94 | 1.04 | 0.97 | 0.99 | 0.99 | 0.915 | 0.91 |  |  | 0.9658
| 1 | 5 | 0.99 | 0.82 | 0.83 | 1.01 | 0.99 | 0.97 | 0.95 | 1.01 | 0.86 | 0.95 | 0.92 | 0.97 | 0.95146 | 0.98618 | 0.94 | 0.97 | 0.93 | 1.0088 | 0.9051 | 0.94 | 0.96 | 1.02 | 0.97 | 0.99 | 0.97 | 0.935 | 0.95 |  |  | 75
| 1 | 6 | 0.97 | 0.86 | 0.79 | 1 | 0.99 | 0.98 | 0.95 | 1 | 0.92 | 0.93 | 0.95 | 1.01 | 0.95425 | 1.0242 | 0.95 | 0.98 | 0.92 | 1.0035 | 0.9145 | 0.91 | 0.95 | 1.01 | 0.98 | 1 | 0.97 | 0.953 | 0.93
| Mean |  | 0.97 | 0.84333 | 0.80833 | 1.015 | 0.985 | 0.98667 | 0.95167 | 1.01 | 0.90167 | 0.93667 | 0.95 | 0.985 | 0.95874 | 1.0116 | 0.94333 | 0.975 | 0.92167 | 1.0035 | 0.9158 | 0.93 | 0.955 | 1.025 | 0.97833 | 0.99333 | 0.96667 | 0.925 | 0.94
| Median |  | 0.97 | 0.84 | 0.81 | 1.01 | 0.99 | 0.99 | 0.95 | 1.01 | 0.915 | 0.935 | 0.955 | 0.985 | 0.95842 | 1.0227 | 0.94 | 0.97 | 0.92 | 1.0035 | 0.9136 | 0.935 | 0.955 | 1.025 | 0.975 | 0.99 | 0.97 | 0.925 | 0.94
| Std Dev. |  | 0.018974 | 0.024221 | 0.016021 | 0.013784 | 0.016432 | 0.010328 | 0.0040825 | 0.0089443 | 0.037639 | 0.012111 | 0.017889 | 0.023452 | 0.0060445 | 0.0203 | 0.005164 | 0.0083666 | 0.0075277 | 0.0072595 | 0.010422 | 0.012649 | 0.010488 | 0.013784 | 0.0098319 | 0.005164 | 0.01633 | 0.021014 | 0.025298
| Rel.Std.Dev. |  | 1.956% | 2.872% | 1.982% | 1.358% | 1.668% | 1.047% | 0.429% | 0.8856% | 4.174% | 1.293% | 1.883% | 2.381% | 0.6305% | 2.007% | 0.5474% | 0.8581% | 0.8168% | 0.7234% | 1.138% | 1.36% | 1.098% | 1.345% | 1.005% | 0.5199% | 1.689% | 2.272% | 2.691%
| PDM3 |  | 0.435% | -12.68% | -16.3% | 5.094% | 1.988% | 2.161% | -1.463% | 4.577% | -6.64% | -3.016% | -1.636% | 1.988% | -0.7303% | 4.74% | -2.326% | 0.9528% | -4.569% | 3.904% | -5.177% | -3.707% | -1.118% | 6.13% | 1.298% | 2.851% | 0.08991% | -4.224% | -2.671%
| Z-Score (Absolute) |  | 0.27 | 2.79 | 3.63 | 1.36 | 0.63 | 0.67 | 0.17 | 1.24 | 1.38 | 0.53 | 0.21 | 0.63 | 0 | 1.28 | 0.37 | 0.39 | 0.9 | 1.08 | 1.04 | 0.69 | 0.09 | 1.6 | 0.47 | 0.84 | 0.19 | 0.81 | 0.45
| Analytical results for Cd in OREAS 625 (Certified Value 83 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 1
| 1 | 1 | 90 | 74.6 | 79.1 | 82.8 | 79.6 | 84.2 | 81.3 | 85.4 | 81.2 | 84.66 | 84.47 | 81.1 | 85.2038 | 95.13 | 81.9623 | 85.9 | 86.2 | 70 | 73.75 | 81.55 | 81.37 | 79.1 | 89.7 | 81.73 | 76.3 | 84 | 84.89 |  |  | 1
| 1 | 2 | 90 | 74.8 | 78.9 | 83 | 81.4 | 84.3 | 81.9 | 87.3 | 87.4 | 84.63 | 81.42 | 83.4 | 84.0075 | 89.54 | 82.7622 | 85.3 | 84.8 | 68.7 | 72.99 | 83.54 | 77.75 | 80.3 | 89.6 | 83.44 | 78.5 | 81 | 85.76 |  |  | 0
| 1 | 3 | 88 | 77 | 76.2 | 82.4 | 84.3 | 86.1 | 80.2 | 86.2 | 86.3 | 84.41 | 85.73 | 77.6 | 83.1853 | 90.67 | 82.2309 | 85.7 | 85.8 | 71.5 | 73.26 | 83.81 | 78.65 | 81 | 91.3 | 82.76 | 75.45 | 81 | 88.51 |  |  | 16
| 1 | 4 | 90 | 75.5 | 76.1 | 85.3 | 83.5 | 86.9 | 79.7 | 85.1 | 83.9 | 84.41 | 85.97 | 84 | 83.7249 | 89.77 | 82.4641 | 85 | 84.5 | 70.3 | 73.67 | 81.5 | 77.15 | 79.9 | 88.2 | 81.81 | 78.62 | 81 | 84.16 |  |  | 83.2787
| 1 | 5 | 90 | 75.6 | 79.9 | 82.9 | 82.4 | 85.4 | 81.1 | 86.2 | 83.3 | 84.96 | 83.41 | 81.8 | 83.7432 | 91.57 | 82.5481 | 85.1 | 85.9 | 71.1 | 73.48 | 83.1 | 80.49 | 78.8 | 87.6 | 81.95 | 76.56 | 83 | 85.35 |  |  | 76
| 1 | 6 | 89 | 77.1 | 77.7 | 82.7 | 82.6 | 84.1 | 80.5 | 85.6 | 87.1 | 84.44 | 83.43 | 82.4 | 83.7185 | 89.59 | 82.9912 | 85.6 | 85.3 | 69.2 | 73.29 | 83.49 | 80 | 79.5 | 93.1 | 82.89 | 76.55 | 84 | 85.17
| Mean |  | 89.5 | 75.7667 | 77.9833 | 83.1833 | 82.3 | 85.1667 | 80.7833 | 85.9667 | 84.8667 | 84.585 | 84.0717 | 81.7167 | 83.9305 | 91.045 | 82.4931 | 85.4333 | 85.4167 | 70.1333 | 73.4067 | 82.8317 | 79.235 | 79.7667 | 89.9167 | 82.43 | 76.9967 | 82.3333 | 85.64
| Median |  | 90 | 75.55 | 78.3 | 82.85 | 82.5 | 84.85 | 80.8 | 85.9 | 85.1 | 84.535 | 83.95 | 82.1 | 83.734 | 90.22 | 82.5061 | 85.45 | 85.55 | 70.15 | 73.385 | 83.295 | 79.325 | 79.7 | 89.65 | 82.355 | 76.555 | 82 | 85.26
| Std Dev. |  | 0.83666 | 1.0671 | 1.5855 | 1.0572 | 1.6517 | 1.1622 | 0.80104 | 0.78655 | 2.4614 | 0.21473 | 1.6967 | 2.2737 | 0.67895 | 2.1499 | 0.36739 | 0.3559 | 0.66758 | 1.0746 | 0.28331 | 1.0373 | 1.65 | 0.80911 | 2.0253 | 0.69934 | 1.2777 | 1.5055 | 1.5039
| Rel.Std.Dev. |  | 0.9348% | 1.408% | 2.033% | 1.271% | 2.007% | 1.365% | 0.9916% | 0.915% | 2.9% | 0.2539% | 2.018% | 2.782% | 0.8089% | 2.361% | 0.4454% | 0.4166% | 0.7816% | 1.532% | 0.386% | 1.252% | 2.082% | 1.014% | 2.252% | 0.8484% | 1.659% | 1.829% | 1.756%
| PDM3 |  | 7.47% | -9.02% | -6.359% | -0.1145% | -1.175% | 2.267% | -2.996% | 3.228% | 1.907% | 1.569% | 0.9522% | -1.876% | 0.7827% | 9.326% | -0.9433% | 2.587% | 2.567% | -15.78% | -11.85% | -0.5368% | -4.856% | -4.217% | 7.971% | -1.019% | -7.543% | -1.135% | 2.835%
| Z-Score (Absolute) |  | 1.74 | 1.84 | 1.26 | 0.09 | 0.14 | 0.61 | 0.53 | 0.82 | 0.53 | 0.46 | 0.32 | 0.29 | 0.29 | 2.14 | 0.09 | 0.68 | 0.67 | 3.31 | 2.46 | 0 | 0.94 | 0.8 | 1.85 | 0.1 | 1.52 | 0.13 | 0.73
| Analytical results for Ce in OREAS 625 (Certified Value 36.8 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [9 empty] | 1
| 1 | 1 | 30.5 | 36.9 | 36.7 | 35.7 | 35.5 | 34.4 | 32.9 | 33.635 | 63 | 43.0771 | 51.9315 | 61.9 | 34 | 44.34 | 58.2 | 41.03 | 45.4 | 32.81 | 38.2 | 34 | [9 empty] | 1
| 1 | 2 | 29.5 | 37.5 | 36.8 | 35 | 35.5 | 34.8 | 36 | 37.371 | 62 | 42.7995 | 51.6503 | 60.4 | 33 | 43.51 | 59.8 | 36.25 | 42.8 | 34.33 | 39.32 | 33 | [9 empty] | 36
| 1 | 3 | 30.6 | 36.4 | 38.1 | 35.2 | 34 | 36.1 | 35.5 | 35.389 | 60 | 42.3159 | 52.1896 | 60.7 | 35 | 43.25 | 61.5 | 40.31 | 44.4 | 33.32 | 36.87 | 31 | [9 empty] | 16
| 1 | 4 | 30.4 | 37.7 | 36.7 | 35.6 | 33.9 | 35.3 | 35.5 | 36.564 | 62 | 43.1755 | 51.7313 | 61.5 | 34 | 43.45 | 60.2 | 36.88 | 41.8 | 33.74 | 40 | 34 | [9 empty] | 36.7836
| 1 | 5 | 30.7 | 36.5 | 36.6 | 35.9 | 34.6 | 35.3 | 34.8 | 35.856 | 59 | 42.305 | 52.0523 | 59.8 | 35 | 43.51 | 60.4 | 40.18 | 43.4 | 31.95 | 38.61 | 33 | [9 empty] | 77
| 1 | 6 | 30.6 | 36.1 | 39.7 | 34.7 | 34.7 | 35.2 | 35.7 | 34.018 | 63 | 42.4805 | 51.9879 | 60.7 | 34 | 43.31 | 60 | 39.47 | 44.3 | 32.18 | 38.58 | 34
| Mean |  | 30.3833 | 36.85 | 37.4333 | 35.35 | 34.7 | 35.1833 | 35.0667 | 35.4722 | 61.5 | 42.6922 | 51.9238 | 60.8333 | 34.1667 | 43.5617 | 60.0167 | 39.02 | 43.6833 | 33.055 | 38.5967 | 33.1667
| Median |  | 30.55 | 36.7 | 36.75 | 35.4 | 34.65 | 35.25 | 35.5 | 35.6225 | 62 | 42.64 | 51.9597 | 60.7 | 34 | 43.48 | 60.1 | 39.825 | 43.85 | 33.065 | 38.595 | 33.5
| Std Dev. |  | 0.4446 | 0.63797 | 1.2453 | 0.45935 | 0.6957 | 0.57067 | 1.1325 | 1.4451 | 1.6432 | 0.38199 | 0.2016 | 0.75807 | 0.75277 | 0.396 | 1.0704 | 1.9751 | 1.2844 | 0.9178 | 1.0613 | 1.169
| Rel.Std.Dev. |  | 1.463% | 1.731% | 3.327% | 1.299% | 2.005% | 1.622% | 3.23% | 4.074% | 2.672% | 0.8948% | 0.3883% | 1.246% | 2.203% | 0.9091% | 1.783% | 5.062% | 2.94% | 2.777% | 2.75% | 3.525%
| PDM3 |  | -17.4% | 0.1805% | 1.766% | -3.897% | -5.665% | -4.351% | -4.668% | -3.565% | 67.19% | 16.06% | 41.16% | 65.38% | -7.114% | 18.43% | 63.16% | 6.08% | 18.76% | -10.14% | 4.929% | -9.833%
| Z-Score (Absolute) |  | 1.31 | 0.06 | 0.06 | 0.35 | 0.47 | 0.38 | 0.4 | 0.32 | 4.73 | 1.08 | 2.87 | 4.6 | 0.58 | 1.25 | 4.44 | 0.36 | 1.27 | 0.79 | 0.28 | 0.77
| Analytical results for Co in OREAS 625 (Certified Value 3.54 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 3 | 3 | 3.3 | 3.5 | 3.5 | 3.1 | 3.2 | 3.5 | 3.3 | 3.2 | 3.3 | 4 | 3.4579 | 1.08 | 3.6397 | 4.1 | 4 | 2.2 | < 5 | 3.82 | 3.8 | 3 | 3.6 | 3.7 | 3.4 | 4 | 3.9 |  |  | 1
| 1 | 2 | 3 | 3.1 | 3.2 | 3.4 | 3.7 | 3.2 | 3.3 | 3.6 | 3.6 | 3.2 | 3.3 | 4 | 3.3519 |  | 3.634 | 4.1 | 3.8 | 2.1 | < 5 | 3.97 | 3.8 | 4 | 3.6 | 3.8 | 3.4 | 4 | 3.9 |  |  | 37
| 1 | 3 | 3 | 3.2 | 3.2 | 3.4 | 3.7 | 3.3 | 3.2 | 3.5 | 3.5 | 3.2 | 3.5 | 3 | 3.265 | 1.7 | 3.79 | 4.1 | 4.2 | 2.2 | < 5 | 3.86 | 3.8 | 4 | 3.5 | 3.7 | 3.3 | 4 | 4.1 |  |  | 16
| 1 | 4 | 3 | 3.2 | 3.2 | 3.5 | 3.7 | 3.2 | 3.1 | 3.5 | 3.5 | 3.2 | 3.4 | 4 | 3.3217 | 1.08 | 3.7618 | 4 | 4 | 2.2 | < 5 | 3.86 | 3.7 | 4 | 3.5 | 3.8 | 3.4 | 4 | 3.8 |  |  | 3.5367
| 1 | 5 | 3 | 3.2 | 3.2 | 3.4 | 3.6 | 3.2 | 3.2 | 3.6 | 3.3 | 3.3 | 3.4 | 4 | 3.279 | 1.6869 | 3.826 | 4.1 | 4.3 | 2.3 | < 5 | 3.91 | 3.7 | 4 | 3.5 | 3.7 | 3.4 | 4 | 3.9 |  |  | 78
| 1 | 6 | 3 | 3.2 | 3.1 | 3.4 | 3.6 | 3.2 | 3.2 | 3.4 | 3.6 | 3.1 | 3.3 | 4 | 3.2794 | 2.4032 | 3.7099 | 4 | 4.2 | 2.2 | < 5 | 3.93 | 3.7 | 4 | 3.6 | 3.6 | 3.4 | 4 | 3.8
| Mean |  | 3 | 3.15 | 3.2 | 3.4333 | 3.6333 | 3.2 | 3.2 | 3.5167 | 3.4667 | 3.2 | 3.3667 | 3.8333 | 3.3258 | 1.59 | 3.7269 | 4.0667 | 4.0833 | 2.2 |  | 3.8917 | 3.75 | 3.8333 | 3.55 | 3.7167 | 3.3833 | 4 | 3.9
| Median |  | 3 | 3.2 | 3.2 | 3.4 | 3.65 | 3.2 | 3.2 | 3.5 | 3.5 | 3.2 | 3.35 | 4 | 3.3006 | 1.6869 | 3.7359 | 4.1 | 4.1 | 2.2 |  | 3.885 | 3.75 | 4 | 3.55 | 3.7 | 3.4 | 4 | 3.9
| Std Dev. |  | 0 | 0.083666 | 0.063246 | 0.05164 | 0.08165 | 0.063246 | 0.063246 | 0.075277 | 0.13663 | 0.063246 | 0.08165 | 0.40825 | 0.072371 | 0.5484 | 0.07943 | 0.05164 | 0.18348 | 0.063246 |  | 0.054924 | 0.054772 | 0.40825 | 0.054772 | 0.075277 | 0.040825 | 0 | 0.10954
| Rel.Std.Dev. |  | 0% | 2.656% | 1.976% | 1.504% | 2.247% | 1.976% | 1.976% | 2.141% | 3.941% | 1.976% | 2.425% | 10.65% | 2.176% | 34.49% | 2.131% | 1.27% | 4.494% | 2.875% |  | 1.411% | 1.461% | 10.65% | 1.543% | 2.025% | 1.207% | 0% | 2.809%
| PDM3 |  | -15.18% | -10.93% | -9.521% | -2.924% | 2.731% | -9.521% | -9.521% | -0.5674% | -1.981% | -9.521% | -4.809% | 8.386% | -5.964% | -55.04% | 5.378% | 14.98% | 15.45% | -37.8% |  | 10.04% | 6.03% | 8.386% | 0.3751% | 5.088% | -4.337% | 13.1% | 10.27%
| Z-Score (Absolute) |  | NA | 0.82 | 0.67 | 0 | 0.58 | 0.67 | 0.67 | 0.24 | 0.1 | 0.67 | 0.19 | NA | 0.31 | 5.33 | 0.85 | 1.83 | 1.88 | 3.56 | 2.7 | 1.32 | 0.92 | NA | 0.34 | 0.82 | 0.14 | NA | 1.35
| Analytical results for Cr in OREAS 625 (Certified Value 13.1 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 2
| 1 | 1 | 14 | 6 | 12 | 13 | 14 | 13 | 13 | 13 | 13.1 | 12 | 13.3 | 13 | 13.4384 | 44.54 | 14.5558 | 15 | 14 | 12.15 | 13 | 12.8 | 13 | 22 | 14 | 13 | 14 | 14 | 13 |  |  | 1
| 1 | 2 | 13 | 6 | 12 | 13 | 13 | 12 | 12 | 13 | 14 | 12 | 13 | 13 | 13.05 | 35.2 | 13.9476 | 15 | 12 | 12.13 | 14 | 12.25 | 13 | 24 | 13 | 13 | 14 | 13 | 14 |  |  | 15
| 1 | 3 | 13 | 6 | 12 | 13 | 14 | 13 | 12 | 13 | 14 | 13 | 13.6 | 12 | 13.136 | 55.87 | 14.4203 | 16 | 12 | 12.21 | 14 | 12.35 | 13 | 24 | 13 | 13 | 14 | 15 | 13 |  |  | 16
| 1 | 4 | 14 | 6 | 11 | 13 | 13 | 15 | 13 | 13 | 14.4 | 13 | 13 | 13 | 12.8283 | 57.19 | 14.6798 | 15 | 12 | 12.34 | 13 | 12.27 | 13 | 24 | 13 | 13 | 14 | 14 | 13 |  |  | 13.0762
| 1 | 5 | 14 | 7 | 12 | 13 | 14 | 12 | 13 | 13 | 13.6 | 13 | 12.9 | 12 | 12.9865 |  | 13.95 | 16 | 13 | 12.19 | 13 | 12.19 | 14 | 22 | 14 | 12 | 14 | 14 | 13 |  |  | 79
| 1 | 6 | 14 | 7 | 12 | 12 | 13 | 12 | 12 | 13 | 14 | 12 | 13 | 13 | 12.879 | 43.83 | 14.721 | 15 | 14 | 12.22 | 13 | 12.14 | 14 | 23 | 13 | 13 | 13 | 13 | 13
| Mean |  | 13.6667 | 6.3333 | 11.8333 | 12.8333 | 13.5 | 12.8333 | 12.5 | 13 | 13.85 | 12.5 | 13.1333 | 12.6667 | 13.053 | 47.326 | 14.3791 | 15.3333 | 12.8333 | 12.2067 | 13.3333 | 12.3333 | 13.3333 | 23.1667 | 13.3333 | 12.8333 | 13.8333 | 13.8333 | 13.1667
| Median |  | 14 | 6 | 12 | 13 | 13.5 | 12.5 | 12.5 | 13 | 14 | 12.5 | 13 | 13 | 13.0182 | 44.54 | 14.4881 | 15 | 12.5 | 12.2 | 13 | 12.26 | 13 | 23.5 | 13 | 13 | 14 | 14 | 13
| Std Dev. |  | 0.5164 | 0.5164 | 0.40825 | 0.40825 | 0.54772 | 1.169 | 0.54772 | 0 | 0.44609 | 0.54772 | 0.26583 | 0.5164 | 0.21937 | 9.1832 | 0.34944 | 0.5164 | 0.98319 | 0.073937 | 0.5164 | 0.23956 | 0.5164 | 0.98319 | 0.5164 | 0.40825 | 0.40825 | 0.75277 | 0.40825
| Rel.Std.Dev. |  | 3.779% | 8.154% | 3.45% | 3.181% | 4.057% | 9.109% | 4.382% | 0% | 3.221% | 4.382% | 2.024% | 4.077% | 1.681% | 19.4% | 2.43% | 3.368% | 7.661% | 0.6057% | 3.873% | 1.942% | 3.873% | 4.244% | 3.873% | 3.181% | 2.951% | 5.442% | 3.101%
| PDM3 |  | 4.516% | -51.57% | -9.505% | -1.857% | 3.241% | -1.857% | -4.406% | -0.5826% | 5.918% | -4.406% | 0.4371% | -3.132% | -0.1769% | 261.9% | 9.964% | 17.26% | -1.857% | -6.65% | 1.967% | -5.681% | 1.967% | 77.17% | 1.967% | -1.857% | 5.79% | 5.79% | 0.692%
| Z-Score (Absolute) |  | 0.67 | 8.6 | 1.64 | 0.38 | 0.46 | 0.38 | 0.8 | 0.17 | 0.91 | 0.8 | 0 | 0.59 | 0.1 | 43.23 | 1.58 | 2.78 | 0.38 | 1.17 | 0.25 | 1.01 | 0.25 | 12.69 | 0.25 | 0.38 | 0.89 | 0.89 | 0.04
| Analytical results for Cs in OREAS 625 (Certified Value 1.14 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 3
| 1 | 1 | 0.96 | 1.02 | 1.1 | 1.18 | 1.03 | 1.04 | 1.08 | 1.38 | 1.2077 | 1.219 | 1.53 | 1 | 1.7 | 1.25 | 1.19 | 1.09 | 1.15 | 1.25 | [11 empty] | 1
| 1 | 2 | 0.94 | 1.01 | 1.15 | 1.2 | 1.04 | 1.04 | 1.13 | 1.4 | 1.1706 | 1.2453 | 1.5 | 1 | 1.76 | 1.15 | 1.18 | 1.18 | 1.13 | 1.24 | [11 empty] | 39
| 1 | 3 | 0.97 | 1 | 1.16 | 1.2 | 1.01 | 1.08 | 1.1 | 1.42 | 1.1687 | 1.2363 | 1.58 | 1 | 1.76 | 1.21 | 1.19 | 1.1 | 1.15 | 1.24 | [11 empty] | 16
| 1 | 4 | 0.94 | 1.05 | 1.14 | 1.18 | 0.99 | 1.04 | 1.12 | 1.35 | 1.1967 | 1.2361 | 1.52 | 1 | 1.71 | 1.12 | 1.15 | 1.14 | 1.19 | 1.25 | [11 empty] | 1.1424
| 1 | 5 | 0.94 | 1 | 1.11 | 1.2 | 1 | 1 | 1.14 | 1.4 | 1.1885 | 1.2247 | 1.52 | 1 | 1.75 | 1.17 | 1.14 | 1.1 | 1.19 | 1.21 | [11 empty] | 80
| 1 | 6 | 0.97 | 0.99 | 1.14 | 1.17 | 1.04 | 1.02 | 1.14 | 1.32 | 1.1606 | 1.2194 | 1.52 | 1 | 1.76 | 1.2 | 1.22 | 1.11 | 1.21 | 1.28
| Mean |  | 0.95333 | 1.0117 | 1.1333 | 1.1883 | 1.0183 | 1.0367 | 1.1183 | 1.3783 | 1.1821 | 1.2301 | 1.5283 | 1 | 1.74 | 1.1833 | 1.1783 | 1.12 | 1.17 | 1.245
| Median |  | 0.95 | 1.005 | 1.14 | 1.19 | 1.02 | 1.04 | 1.125 | 1.39 | 1.1796 | 1.2304 | 1.52 | 1 | 1.755 | 1.185 | 1.185 | 1.105 | 1.17 | 1.245
| Std Dev. |  | 0.015055 | 0.02137 | 0.023381 | 0.013292 | 0.02137 | 0.026583 | 0.024014 | 0.037103 | 0.01836 | 0.010697 | 0.027142 | 0 | 0.027568 | 0.046332 | 0.029269 | 0.034059 | 0.030984 | 0.022583
| Rel.Std.Dev. |  | 1.579% | 2.112% | 2.063% | 1.119% | 2.099% | 2.564% | 2.147% | 2.692% | 1.553% | 0.8696% | 1.776% | 0% | 1.584% | 3.915% | 2.484% | 3.041% | 2.648% | 1.814%
| PDM3 |  | -16.55% | -11.44% | -0.7894% | 4.025% | -10.86% | -9.252% | -2.103% | 20.66% | 3.484% | 7.684% | 33.79% | -12.46% | 52.32% | 3.587% | 3.15% | -1.957% | 2.42% | 8.986%
| Z-Score (Absolute) |  | 2.53 | 1.87 | 0.51 | 0.11 | 1.8 | 1.59 | 0.67 | 2.25 | 0.04 | 0.58 | 3.93 | NA | 6.31 | 0.06 | 0 | 0.66 | 0.09 | 0.75
| Analytical results for Cu in OREAS 625 (Certified Value 0.174 wt.%) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | 3
| 1 | 1 | 0.17 | 0.161 | 0.17494 | 0.175 | 0.174 | 0.1655 | 0.171 | 0.1735 | 0.163 | 0.15225 | 0.172 | 0.1692 | 0.17852 | 0.18027 | 0.16969 | 0.1907 | 0.172 | 0.16604 | 0.169 | 0.1845 | 0.174 | 0.19541 | 0.1734 | 0.1733 | 0.1749 | 0.17422 | 0.17793 | 0.1818 | 0.17842 | 1
| 1 | 2 | 0.171 | 0.159 | 0.17564 | 0.175 | 0.1775 | 0.168 | 0.1725 | 0.1785 | 0.175 | 0.15724 | 0.172 | 0.1689 | 0.17445 | 0.17621 | 0.17229 | 0.1915 | 0.17 | 0.16448 | 0.1664 | 0.1887 | 0.1738 | 0.19285 | 0.1751 | 0.1766 | 0.1718 | 0.17715 | 0.17543 | 0.1742 | 0.18326 | 25
| 1 | 3 | 0.1704 | 0.157 | 0.176 | 0.1765 | 0.18 | 0.168 | 0.1695 | 0.1795 | 0.172 | 0.15347 | 0.172 | 0.1629 | 0.17253 | 0.17863 | 0.16379 | 0.1907 | 0.173 | 0.17261 | 0.1685 | 0.1863 | 0.1738 | 0.18782 | 0.1771 | 0.1785 | 0.1757 | 0.17791 | 0.17092 | 0.1747 | 0.18895 | 16
| 1 | 4 | 0.1721 | 0.162 | 0.17428 | 0.179 | 0.1805 | 0.1695 | 0.1725 | 0.1775 | 0.169 | 0.15553 | 0.174 | 0.1719 | 0.17397 | 0.17618 | 0.16772 | 0.1928 | 0.169 | 0.1659 | 0.1699 | 0.1899 | 0.1737 | 0.19242 | 0.1718 | 0.1784 | 0.1717 | 0.17427 | 0.17936 | 0.1776 | 0.17575 | 0.17359
| 1 | 5 | 0.1707 | 0.162 | 0.17749 | 0.173 | 0.1795 | 0.1645 | 0.171 | 0.1785 | 0.168 | 0.1564 | 0.173 | 0.166 | 0.17339 | 0.18318 | 0.16874 | 0.1904 | 0.171 | 0.17159 | 0.1698 | 0.1853 | 0.1728 | 0.19437 | 0.1744 | 0.1751 | 0.1726 | 0.17662 | 0.17475 | 0.1805 | 0.18232 | 81
| 1 | 6 | 0.171 | 0.163 | 0.17765 | 0.1765 | 0.1775 | 0.1665 | 0.171 | 0.177 | 0.173 | 0.15039 | 0.172 | 0.1733 | 0.17317 | 0.17909 | 0.16403 | 0.1893 | 0.172 | 0.16731 | 0.1695 | 0.1857 | 0.1741 | 0.19144 | 0.173 | 0.1732 | 0.1735 | 0.17764 | 0.1768 | 0.183 | 0.17878
| Mean |  | 0.17087 | 0.16067 | 0.176 | 0.17583 | 0.17817 | 0.167 | 0.17125 | 0.17742 | 0.17 | 0.15421 | 0.1725 | 0.1687 | 0.17434 | 0.17893 | 0.16771 | 0.1909 | 0.17117 | 0.16799 | 0.16885 | 0.18673 | 0.1737 | 0.19238 | 0.17413 | 0.17585 | 0.17337 | 0.1763 | 0.17586 | 0.17863 | 0.18125
| Median |  | 0.17085 | 0.1615 | 0.17582 | 0.17575 | 0.1785 | 0.16725 | 0.171 | 0.178 | 0.1705 | 0.1545 | 0.172 | 0.16905 | 0.17368 | 0.17886 | 0.16823 | 0.1907 | 0.1715 | 0.16667 | 0.16925 | 0.186 | 0.1738 | 0.19263 | 0.1739 | 0.17585 | 0.17305 | 0.17688 | 0.17611 | 0.17905 | 0.18055
| Std Dev. |  | 0.00071461 | 0.0022509 | 0.0013531 | 0.0020166 | 0.0024014 | 0.0018439 | 0.0011292 | 0.0021075 | 0.0042895 | 0.0026344 | 0.00083666 | 0.003811 | 0.0021543 | 0.0026435 | 0.0033112 | 0.0011713 | 0.001472 | 0.0033244 | 0.0013096 | 0.0021068 | 0.00046476 | 0.0026439 | 0.0018479 | 0.0023738 | 0.0016513 | 0.0016529 | 0.0029407 | 0.0037087 | 0.0046655
| Rel.Std.Dev. |  | 0.4182% | 1.401% | 0.7688% | 1.147% | 1.348% | 1.104% | 0.6594% | 1.188% | 2.523% | 1.708% | 0.485% | 2.259% | 1.236% | 1.477% | 1.974% | 0.6136% | 0.86% | 1.979% | 0.7756% | 1.128% | 0.2676% | 1.374% | 1.061% | 1.35% | 0.9525% | 0.9375% | 1.672% | 2.076% | 2.574%
| PDM3 |  | -1.566% | -7.442% | 1.391% | 1.295% | 2.639% | -3.794% | -1.345% | 2.207% | -2.065% | -11.16% | -0.6251% | -2.814% | 0.4344% | 3.077% | -3.384% | 9.975% | -1.393% | -3.224% | -2.728% | 7.575% | 0.06624% | 10.83% | 0.3159% | 1.305% | -0.1258% | 1.565% | 1.313% | 2.908% | 4.414%
| Z-Score (Absolute) |  | 0.55 | 2.25 | 0.31 | 0.28 | 0.67 | 1.19 | 0.48 | 0.55 | 0.69 | 3.33 | 0.27 | 0.91 | 0.03 | 0.8 | 1.07 | 2.8 | 0.5 | 1.03 | 0.88 | 2.11 | 0.07 | 3.05 | 0 | 0.29 | 0.13 | 0.36 | 0.29 | 0.75 | 1.19
| Analytical results for Dy in OREAS 625 (Indicative Value 1.69 ppm) | [30 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | 2
| 1 | 1 | 1.01: 1.7; 1.02: 1.612; 1.03: 1.7028; 1.04: 1.6758 | 1
| 1 | 2 | 1.01: 1.8; 1.02: 1.706; 1.03: 1.7164; 1.04: 1.5172 | 7
| 1 | 3 | 1.01: 1.7; 1.02: 1.717; 1.03: 1.6863; 1.04: 1.6853 | 16
| 1 | 4 | 1.01: 1.7; 1.02: 1.723; 1.03: 1.6742; 1.04: 1.6229 | 1.6888
| 1 | 5 | 1.01: 1.8; 1.02: 1.649; 1.03: 1.6674; 1.04: 1.6675 | 13
| 1 | 6 | 1.01: 1.7; 1.02: 1.581; 1.03: 1.735; 1.04: 1.6503
| Mean |  | 1.01: 1.7333; 1.02: 1.6647; 1.03: 1.697; 1.04: 1.6365
| Median |  | 1.01: 1.7; 1.02: 1.6775; 1.03: 1.6945; 1.04: 1.6589
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.059782; 1.03: 0.025935; 1.04: 0.062448
| Rel.Std.Dev. |  | 1.01: 2.979%; 1.02: 3.591%; 1.03: 1.528%; 1.04: 3.816%
| PDM3 |  | 1.01: 2.635%; 1.02: -1.431%; 1.03: 0.4839%; 1.04: -3.1%
| Z-Score (Absolute) |  | 1.01: 1.17; 1.02: 0.36; 1.03: 0.36; 1.04: 0.99
| Analytical results for Er in OREAS 625 (Indicative Value 0.55 ppm) | [30 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | 2
| 1 | 1 | 1.01: 0.6; 1.02: 0.487; 1.03: 0.54653; 1.04: 0.57121 | 1
| 1 | 2 | 1.01: 0.6; 1.02: 0.563; 1.03: 0.54532; 1.04: 0.54228 | 8
| 1 | 3 | 1.01: 0.5; 1.02: 0.512; 1.03: 0.54458; 1.04: 0.56971 | 16
| 1 | 4 | 1.01: 0.6; 1.02: 0.52; 1.03: 0.54246; 1.04: 0.53825 | 0.55165
| 1 | 5 | 1.01: 0.6; 1.02: 0.537; 1.03: 0.57655; 1.04: 0.56166 | 14
| 1 | 6 | 1.01: 0.6; 1.02: 0.515; 1.03: 0.53896; 1.04: 0.56114
| Mean |  | 1.01: 0.58333; 1.02: 0.52233; 1.03: 0.54907; 1.04: 0.55737
| Median |  | 1.01: 0.6; 1.02: 0.5175; 1.03: 0.54495; 1.04: 0.5614
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.025626; 1.03: 0.013722; 1.04: 0.013926
| Rel.Std.Dev. |  | 1.01: 6.999%; 1.02: 4.906%; 1.03: 2.499%; 1.04: 2.499%
| PDM3 |  | 1.01: 5.743%; 1.02: -5.315%; 1.03: -0.469%; 1.04: 1.037%
| Z-Score (Absolute) |  | 1.01: 1.19; 1.02: 1.22; 1.03: 0.16; 1.04: 0.16
| Analytical results for Eu in OREAS 625 (Indicative Value 0.67 ppm) | [30 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | 2
| 1 | 1 | 1.01: 0.6; 1.02: 0.633; 1.03: 0.76798; 1.04: 0.72626 | 1
| 1 | 2 | 1.01: 0.6; 1.02: 0.652; 1.03: 0.77361; 1.04: 0.65268 | 9
| 1 | 3 | 1.01: 0.6; 1.02: 0.656; 1.03: 0.78624; 1.04: 0.71714 | 16
| 1 | 4 | 1.01: 0.5; 1.02: 0.649; 1.03: 0.75714; 1.04: 0.67021 | 0.67319
| 1 | 5 | 1.01: 0.6; 1.02: 0.656; 1.03: 0.76739; 1.04: 0.70894 | 15
| 1 | 6 | 1.01: 0.6; 1.02: 0.625; 1.03: 0.76853; 1.04: 0.68945
| Mean |  | 1.01: 0.58333; 1.02: 0.64517; 1.03: 0.77015; 1.04: 0.69411
| Median |  | 1.01: 0.6; 1.02: 0.6505; 1.03: 0.76825; 1.04: 0.6992
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.013045; 1.03: 0.0095372; 1.04: 0.028612
| Rel.Std.Dev. |  | 1.01: 6.999%; 1.02: 2.022%; 1.03: 1.238%; 1.04: 4.122%
| PDM3 |  | 1.01: -13.35%; 1.02: -4.163%; 1.03: 14.4%; 1.04: 3.108%
| Z-Score (Absolute) |  | 1.01: 1.05; 1.02: 0.3; 1.03: 1.22; 1.04: 0.3
| Analytical results for Fe in OREAS 625 (Certified Value 3.27 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 3.5 | 3.06 | 3.11 | 3.24 | 3.27 | 3.12 | 3.09 | 3.16 | 2.94 | 3.25 | 3.09 | 3.53 | 3.4375 | 3.9 | 3.7092 | 3.55 | 3.55 | 3.08 | 3.2644 | 2.885 | 3.49 | 3.37 | 2.79 | 3.14 | 3.24 | 3.3 | 3.42 | 3.12 |  | 1
| 1 | 2 | 3.45 | 2.98 | 3.05 | 3.3 | 3.33 | 3.11 | 3.08 | 3.21 | 3.17 | 3.33 | 3.04 | 3.55 | 3.3705 | 3.65 | 3.595 | 3.49 | 3.58 | 3.01 | 3.2542 | 2.9698 | 3.45 | 3.34 | 2.85 | 3.14 | 3.28 | 3.3 | 3.26 | 3.18 |  | 26
| 1 | 3 | 3.47 | 3.01 | 3.06 | 3.26 | 3.37 | 3.14 | 3.04 | 3.19 | 3.09 | 3.32 | 3.19 | 3.41 | 3.3268 | 3.6 | 3.6159 | 3.58 | 3.66 | 3.15 | 3.2497 | 2.9183 | 3.49 | 3.41 | 2.86 | 3.14 | 3.26 | 3.24 | 3.27 | 3.24 |  | 16
| 1 | 4 | 3.45 | 3.08 | 2.99 | 3.34 | 3.33 | 3.16 | 3.06 | 3.19 | 3.19 | 3.42 | 3.12 | 3.61 | 3.3889 | 3.69 | 3.7309 | 3.57 | 3.59 | 3.08 | 3.2556 | 2.9712 | 3.49 | 3.31 | 2.85 | 3.18 | 3.24 | 3.45 | 3.32 | 3.07 |  | 3.2729
| 1 | 5 | 3.53 | 3 | 3.12 | 3.25 | 3.31 | 3.06 | 3.04 | 3.2 | 2.96 | 3.41 | 3.12 | 3.48 | 3.3451 | 3.94 | 3.7368 | 3.54 | 3.77 | 3.13 | 3.2631 | 2.8666 | 3.49 | 3.37 | 2.83 | 3.14 | 3.29 | 3.36 | 3.38 | 3.15 |  | 82
| 1 | 6 | 3.45 | 3.1 | 3.11 | 3.18 | 3.22 | 3.08 | 3.04 | 3.14 | 3.14 | 3.28 | 3.03 | 3.62 | 3.3372 | 3.58 | 3.7722 | 3.59 | 3.57 | 3.08 | 3.2531 | 2.9339 | 3.4 | 3.35 | 2.77 | 3.16 | 3.25 | 3.35 | 3.43 | 3.12
| Mean |  | 3.475 | 3.0383 | 3.0733 | 3.2617 | 3.305 | 3.1117 | 3.0583 | 3.1817 | 3.0817 | 3.335 | 3.0983 | 3.5333 | 3.3677 | 3.7267 | 3.6933 | 3.5533 | 3.62 | 3.0883 | 3.2567 | 2.9241 | 3.4683 | 3.3583 | 2.825 | 3.15 | 3.26 | 3.3333 | 3.3467 | 3.1467
| Median |  | 3.46 | 3.035 | 3.085 | 3.255 | 3.32 | 3.115 | 3.05 | 3.19 | 3.115 | 3.325 | 3.105 | 3.54 | 3.3578 | 3.67 | 3.72 | 3.56 | 3.585 | 3.08 | 3.2549 | 2.9261 | 3.49 | 3.36 | 2.84 | 3.14 | 3.255 | 3.325 | 3.35 | 3.135
| Std Dev. |  | 0.033317 | 0.048339 | 0.050067 | 0.054559 | 0.05282 | 0.037103 | 0.022286 | 0.026394 | 0.10759 | 0.068337 | 0.059133 | 0.079666 | 0.041096 | 0.15513 | 0.071341 | 0.036148 | 0.082462 | 0.048751 | 0.0058253 | 0.043062 | 0.037103 | 0.033714 | 0.036742 | 0.016733 | 0.020976 | 0.071461 | 0.074207 | 0.058538
| Rel.Std.Dev. |  | 0.9588% | 1.591% | 1.629% | 1.673% | 1.598% | 1.192% | 0.7287% | 0.8296% | 3.491% | 2.049% | 1.909% | 2.255% | 1.22% | 4.163% | 1.932% | 1.017% | 2.278% | 1.579% | 0.1789% | 1.473% | 1.07% | 1.004% | 1.301% | 0.5312% | 0.6434% | 2.144% | 2.217% | 1.86%
| PDM3 |  | 6.174% | -7.168% | -6.098% | -0.3439% | 0.9801% | -4.927% | -6.556% | -2.788% | -5.844% | 1.897% | -5.334% | 7.957% | 2.895% | 13.86% | 12.85% | 8.568% | 10.6% | -5.64% | -0.4962% | -10.66% | 5.971% | 2.61% | -13.69% | -3.756% | -0.3948% | 1.846% | 2.253% | -3.858%
| Z-Score (Absolute) |  | 0.86 | 0.9 | 0.75 | 0 | 0.18 | 0.6 | 0.82 | 0.32 | 0.72 | 0.3 | 0.65 | 1.1 | 0.43 | 1.88 | 1.74 | 1.18 | 1.45 | 0.69 | 0.02 | 1.36 | 0.84 | 0.39 | 1.75 | 0.45 | 0 | 0.29 | 0.35 | 0.46
| Analytical results for Ga in OREAS 625 (Certified Value 2.63 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g |  |  |  |  | 2
| 1 | 1 | 3 | 4.26 | 7 | 2.46 | 2.43 | 2.27 | 2.12 | 2.86 | 2.52 | 2.55 | 2.1 | 2.9429 | 3.01 | 2.8427 | 109.2 | 3.8 | 3 | < 5 | 5.13 | 3 | < 5 | 2.3 | 2.3 | 2.63 | < 2 |  |  |  |  | 1
| 1 | 2 | 3 | 4.24 | 6 | 2.42 | 2.54 | 2.16 | 2.24 | 2.86 | 2.61 | 2.5 | 2.1 | 2.8856 | 3.02 | 3.1261 | 109.5 | 3.6 | 2 | < 5 | 5.45 | 2.6 | < 5 | 2.3 | 2.5 | 2.56 | < 2 |  |  |  |  | 44
| 1 | 3 | 2 | 3.94 | 7 | 2.39 | 2.53 | 2.3 | 2.14 | 2.86 | 2.61 | 2.4 | 2.2 | 2.8011 | 3.16 | 3.0543 | 109.1 | 3.8 | 3 | < 5 | 5.53 | 2.9 | < 5 | 2.1 | 2.3 | 2.61 | < 2 |  |  |  |  | 16
| 1 | 4 | 2 | 4.26 | 6 | 2.48 | 2.53 | 2.3 | 2.1 | 2.7 | 2.51 | 2.33 | 2.2 | 2.8951 | 3.39 | 3.0089 | 107.9 | 3.7 | 3 | < 5 | 5.15 | 2.8 | < 5 | 2.1 | 2.3 | 2.71 | < 2 |  |  |  |  | 2.6263
| 1 | 5 | 3 | 4.24 | 7 | 2.39 | 2.48 | 2.37 | 2.12 | 2.78 | 2.55 | 2.23 | 2.1 | 2.9164 | 3.38 | 2.8127 | 107.1 | 3.7 | 3 | < 5 | 5.4 | 2.9 | < 5 | 2.2 | 2.4 | 2.55 | < 2 |  |  |  |  | 83
| 1 | 6 | 3 | 4.2 | 7 | 2.32 | 2.71 | 2.27 | 2.11 | 2.7 | 2.66 | 2.24 | 2.1 | 2.8758 | 3.48 | 3.0144 | 107.7 | 3.9 | 2 | < 5 | 5.44 | 2.8 | < 5 | 2.2 | 2.3 | 2.53 | < 2
| Mean |  | 2.6667 | 4.19 | 6.6667 | 2.41 | 2.5367 | 2.2783 | 2.1383 | 2.7933 | 2.5767 | 2.375 | 2.1333 | 2.8861 | 3.24 | 2.9765 | 108.4167 | 3.75 | 2.6667 |  | 5.35 | 2.8333 |  | 2.2 | 2.35 | 2.5983
| Median |  | 3 | 4.24 | 7 | 2.405 | 2.53 | 2.285 | 2.12 | 2.82 | 2.58 | 2.365 | 2.1 | 2.8904 | 3.27 | 3.0117 | 108.5 | 3.75 | 3 |  | 5.42 | 2.85 |  | 2.2 | 2.3 | 2.585
| Std Dev. |  | 0.5164 | 0.12442 | 0.5164 | 0.057271 | 0.094587 | 0.068532 | 0.051543 | 0.078655 | 0.059217 | 0.13278 | 0.05164 | 0.048069 | 0.20367 | 0.12303 | 0.97656 | 0.10488 | 0.5164 |  | 0.16817 | 0.13663 |  | 0.089443 | 0.083666 | 0.066458
| Rel.Std.Dev. |  | 19.36% | 2.969% | 7.746% | 2.376% | 3.729% | 3.008% | 2.41% | 2.816% | 2.298% | 5.591% | 2.421% | 1.666% | 6.286% | 4.133% | 0.9007% | 2.797% | 19.36% |  | 3.143% | 4.822% |  | 4.066% | 3.56% | 2.558%
| PDM3 |  | 1.537% | 59.54% | 153.8% | -8.236% | -3.413% | -13.25% | -18.58% | 6.36% | -1.89% | -9.569% | -18.77% | 9.893% | 23.37% | 13.33% | 4028% | 42.79% | 1.537% |  | 103.7% | 7.883% |  | -16.23% | -10.52% | -1.065%
| Z-Score (Absolute) |  | NA | 3.62 | NA | 0.33 | 0.04 | 0.62 | 0.93 | 0.53 | 0.04 | 0.4 | 0.94 | 0.73 | 1.52 | 0.93 | 234.89 | 2.65 | NA | 0.13 | 6.2 | 0.61 | 0.13 | 0.79 | 0.46 | 0.09 | 3.45
Results from laboratories 1.01, 1.03 and 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 625 (Indicative Value 2.9 ppm) | [30 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  |  |  |  | 2
| 1 | 1 | 1.01: 2.4; 1.02: 2.465; 1.03: 3.0931; 1.04: 3.7608 |  |  |  |  | 1
| 1 | 2 | 1.01: 2.3; 1.02: 2.667; 1.03: 3.0267; 1.04: 3.3816 |  |  |  |  | 10
| 1 | 3 | 1.01: 2.5; 1.02: 2.54; 1.03: 2.9693; 1.04: 3.7897 |  |  |  |  | 16
| 1 | 4 | 1.01: 2.5; 1.02: 2.598; 1.03: 2.937; 1.04: 3.5058 |  |  |  |  | 2.9003
| 1 | 5 | 1.01: 2.5; 1.02: 2.554; 1.03: 2.9311; 1.04: 3.6867 |  |  |  |  | 16
| 1 | 6 | 1.01: 2.4; 1.02: 2.458; 1.03: 3.0276; 1.04: 3.6169
| Mean |  | 1.01: 2.4333; 1.02: 2.547; 1.03: 2.9975; 1.04: 3.6236
| Median |  | 1.01: 2.45; 1.02: 2.547; 1.03: 2.998; 1.04: 3.6518
| Std Dev. |  | 1.01: 0.08165; 1.02: 0.079679; 1.03: 0.062875; 1.04: 0.15676
| Rel.Std.Dev. |  | 1.01: 3.355%; 1.02: 3.128%; 1.03: 2.098%; 1.04: 4.326%
| PDM3 |  | 1.01: -16.1%; 1.02: -12.18%; 1.03: 3.349%; 1.04: 24.94%
| Z-Score (Absolute) |  | 1.01: 0.81; 1.02: 0.54; 1.03: 0.54; 1.04: 2.04
| Analytical results for Ge in OREAS 625 (Indicative Value 0.084 ppm) | [30 empty] | Indicative
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.04; 1.03: 1.05; 1.04: 1.06; 1.05: 1.07; 1.06: 1.08; 1.07: 1.09; 1.08: 1.13; 1.09: 1.18; 1.10: 1.23; 1.11: 1.25; 1.13: 1.26; 1.14: 1.27; 1.15: 1.30 |  |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*OES/MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS; 1.15: AR*MS |  |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.2g; 1.09: 0.25g; 1.10: 0.25g; 1.11: 0.25g; 1.13: 0.25g; 1.14: 0.5g; 1.15: 0.2g |  |  |  |  | 3
| 1 | 1 | 1.01: < 0.1; 1.02: 0.07; 1.03: 0.07; 1.04: < 0.05; 1.05: 0.05; 1.06: 0.11; 1.07: 0.1; 1.08: < 0.2; 1.09: 0.3; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.05; 1.15: 0.14 |  |  |  |  | 1
| 1 | 2 | 1.01: < 0.1; 1.02: 0.07; 1.03: 0.08; 1.04: < 0.05; 1.05: < 0.05; 1.06: 0.11; 1.07: 0.12; 1.08: < 0.2; 1.09: 0.2; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: < 0.05; 1.15: 0.14 |  |  |  |  | 46
| 1 | 3 | 1.01: < 0.1; 1.02: 0.06; 1.03: 0.07; 1.04: < 0.05; 1.05: < 0.05; 1.06: 0.13; 1.07: 0.13; 1.08: < 0.2; 1.09: 0.2; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: < 0.05; 1.15: 0.15 |  |  |  |  | 16
| 1 | 4 | 1.01: < 0.1; 1.02: 0.07; 1.03: 0.08; 1.04: < 0.05; 1.05: < 0.05; 1.06: 0.1; 1.07: 0.12; 1.08: < 0.2; 1.09: 0.2; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.05; 1.15: 0.13 |  |  |  |  | 0.084259
| 1 | 5 | 1.01: < 0.1; 1.02: 0.06; 1.03: 0.06; 1.04: < 0.05; 1.05: 0.05; 1.06: 0.11; 1.07: 0.12; 1.08: < 0.2; 1.09: 0.2; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: < 0.05; 1.15: 0.14 |  |  |  |  | 12
| 1 | 6 | 1.01: < 0.1; 1.02: 0.06; 1.03: 0.08; 1.04: 0.05; 1.05: 0.05; 1.06: 0.12; 1.07: 0.11; 1.08: < 0.2; 1.09: 0.2; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: < 0.05; 1.15: 0.14
| Mean |  | 1.02: 0.065; 1.03: 0.073333; 1.04: 0.05; 1.05: 0.05; 1.06: 0.11333; 1.07: 0.11667; 1.09: 0.21667; 1.13: 0.1; 1.14: 0.05; 1.15: 0.14
| Median |  | 1.02: 0.065; 1.03: 0.075; 1.04: 0.05; 1.05: 0.05; 1.06: 0.11; 1.07: 0.12; 1.09: 0.2; 1.13: 0.1; 1.14: 0.05; 1.15: 0.14
| Std Dev. |  | 1.02: 0.0054772; 1.03: 0.008165; 1.05: 8.4984e-18; 1.06: 0.010328; 1.07: 0.010328; 1.09: 0.040825; 1.13: 1.5202e-17; 1.14: 0; 1.15: 0.0063246
| Rel.Std.Dev. |  | 1.02: 8.427%; 1.03: 11.13%; 1.05: 1.7e-14%; 1.06: 9.113%; 1.07: 8.853%; 1.09: 18.84%; 1.13: 1.52e-14%; 1.14: 0%; 1.15: 4.518%
| PDM3 |  | 1.02: -22.86%; 1.03: -12.97%; 1.04: -40.66%; 1.05: -40.66%; 1.06: 34.51%; 1.07: 38.46%; 1.09: 157.1%; 1.13: 18.68%; 1.14: -40.66%; 1.15: 66.15%
| Z-Score (Absolute) |  | 1.01: 0.41; 1.02: 0.09; 1.03: 0.09; 1.04: 0.86; 1.05: 0.68; 1.06: 0.95; 1.07: 1.02; 1.08: 0.67; 1.09: 3.18; 1.10: 0.41; 1.11: 0.41; 1.13: 0.67; 1.14: 0.77; 1.15: 1.53
| Analytical results for Hf in OREAS 625 (Certified Value 1.46 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 3
| 1 | 1 | 1.2 | 1.43 | 1.44 | 1.34 | 1.38 | 1.49 | 1.09 | 1.46 | 1.4363 | 1.6012 | 1.91 | 1.7 | 1.73 | 1.65 | 1.48 | 1.12 | 1.38 | 1.45 | [11 empty] | 1
| 1 | 2 | 0.8 | 1.43 | 1.49 | 1.34 | 1.43 | 1.47 | 1.12 | 1.57 | 1.409 | 1.6127 | 1.88 | 1.7 | 1.82 | 1.63 | 1.43 | 1.15 | 1.42 | 1.15 | [11 empty] | 16
| 1 | 3 | 1.4 | 1.46 | 1.51 | 1.37 | 1.34 | 1.45 | 1.11 | 1.51 | 1.3973 | 1.6071 | 1.87 | 1.6 | 1.88 | 1.51 | 1.42 | 1.17 | 1.37 | 1.28 | [11 empty] | 16
| 1 | 4 | 1.1 | 1.52 | 1.49 | 1.4 | 1.34 | 1.45 | 1.07 | 1.6 | 1.3826 | 1.582 | 1.87 | 1.7 | 1.7 | 1.65 | 1.32 | 1.16 | 1.41 | 1.51 | [11 empty] | 1.461
| 1 | 5 | 1.2 | 1.47 | 1.46 | 1.4 | 1.38 | 1.45 | 1.04 | 1.53 | 1.3885 | 1.55 | 1.87 | 1.7 | 1.81 | 1.64 | 1.52 | 1.14 | 1.43 | 1.25 | [11 empty] | 84
| 1 | 6 | 1.1 | 1.41 | 1.51 | 1.38 | 1.36 | 1.42 | 1.12 | 1.47 | 1.3866 | 1.5547 | 1.88 | 1.7 | 1.77 | 1.62 | 1.37 | 1.13 | 1.43 | 1.65
| Mean |  | 1.1333 | 1.4533 | 1.4833 | 1.3717 | 1.3717 | 1.455 | 1.0917 | 1.5233 | 1.4001 | 1.5846 | 1.88 | 1.6833 | 1.785 | 1.6167 | 1.4233 | 1.145 | 1.4067 | 1.3817
| Median |  | 1.15 | 1.445 | 1.49 | 1.375 | 1.37 | 1.45 | 1.1 | 1.52 | 1.3929 | 1.5916 | 1.875 | 1.7 | 1.79 | 1.635 | 1.425 | 1.145 | 1.415 | 1.365
| Std Dev. |  | 0.19664 | 0.039328 | 0.028048 | 0.027142 | 0.033714 | 0.023452 | 0.031885 | 0.055015 | 0.02007 | 0.027084 | 0.015492 | 0.040825 | 0.065345 | 0.053541 | 0.072296 | 0.018708 | 0.02582 | 0.18681
| Rel.Std.Dev. |  | 17.35% | 2.706% | 1.891% | 1.979% | 2.458% | 1.612% | 2.921% | 3.611% | 1.433% | 1.709% | 0.824% | 2.425% | 3.661% | 3.312% | 5.079% | 1.634% | 1.836% | 13.52%
| PDM3 |  | -22.43% | -0.5234% | 1.53% | -6.113% | -6.113% | -0.4093% | -25.28% | 4.268% | -4.17% | 8.462% | 28.68% | 15.22% | 22.18% | 10.66% | -2.577% | -21.63% | -3.718% | -5.429%
| Z-Score (Absolute) |  | 3.26 | 0.34 | 0.67 | 0.58 | 0.58 | 0.36 | 3.73 | 1.12 | 0.26 | 1.81 | 5.13 | NA | 4.06 | 2.17 | 0 | 3.13 | 0.19 | 0.47
Results from laboratory 1.18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 625 (Certified Value 1.72 ppm) | [30 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.18 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 01g | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [6 empty] | 3
| 1 | 1 | < 5 | 1.71 | 1.91 | 1.66 | 1.51 | 1.6 | 1.65 | 1.66 | 1.71 | 1.7 | 1.769 | 1.8514 | 1.6323 | < 5 | 0.964 | 1.46 | 1.59 | 2 | 1.9 | 2 | 1.591 | 3 | < 2 | [6 empty] | 1
| 1 | 2 | < 5 | 1.66 | 2.11 | 1.67 | 1.56 | 1.78 | 1.74 | 1.73 | 1.73 | 1.8 | 1.635 | 1.9386 | 1.6198 | < 5 | 0.995 | 1.48 | 1.5 | 1 | 1.9 | 1.99 | 1.658 | 3 | < 2 | [6 empty] | 4
| 1 | 3 | < 5 | 1.72 | 1.92 | 1.72 | 1.63 | 1.67 | 1.58 | 1.74 | 1.73 | 1.7 | 1.757 | 1.6769 | 1.6496 | < 5 | 0.981 | 1.46 | 1.57 | 2 | 1.9 | 2 | 1.568 | 3 | < 2 | [6 empty] | 16
| 1 | 4 | < 5 | 1.74 | 1.77 | 1.69 | 1.57 | 1.61 | 1.58 | 1.8 | 1.77 | 1.8 | 1.766 | 1.8665 | 1.6227 | < 5 | 0.976 | 1.49 | 1.53 | 2 | 1.8 | 1.99 | 1.586 | 3 | < 2 | [6 empty] | 1.7164
| 1 | 5 | < 5 | 1.73 | 1.82 | 1.7 | 1.53 | 1.88 | 1.7 | 1.72 | 1.65 | 1.9 | 1.73 | 1.7968 | 1.6514 | < 5 | 0.994 | 1.4 | 1.58 | 2 | 1.9 | 1.92 | 1.646 | 3 | < 2 | [6 empty] | 85
| 1 | 6 | < 5 | 1.74 | 1.84 | 1.69 | 1.75 | 1.8 | 1.64 | 1.78 | 1.73 | 1.8 | 1.639 | 1.7776 | 1.6266 | < 5 | 0.99 | 1.47 | 1.66 | 2 | 1.9 | 1.93 | 1.636 | 1 | < 2
| Mean |  |  | 1.7167 | 1.895 | 1.6883 | 1.5917 | 1.7233 | 1.6483 | 1.7383 | 1.72 | 1.7833 | 1.716 | 1.818 | 1.6337 |  | 0.98333 | 1.46 | 1.5717 | 1.8333 | 1.8833 | 1.9717 | 1.6142 | 2.6667
| Median |  |  | 1.725 | 1.875 | 1.69 | 1.565 | 1.725 | 1.645 | 1.735 | 1.73 | 1.8 | 1.7435 | 1.8241 | 1.6294 |  | 0.9855 | 1.465 | 1.575 | 2 | 1.9 | 1.99 | 1.6135 | 3
| Std Dev. |  |  | 0.030111 | 0.11946 | 0.02137 | 0.087731 | 0.11361 | 0.064005 | 0.04916 | 0.039497 | 0.075277 | 0.062731 | 0.089458 | 0.013653 |  | 0.012061 | 0.031623 | 0.054924 | 0.40825 | 0.040825 | 0.03656 | 0.037075 | 0.8165
| Rel.Std.Dev. |  |  | 1.754% | 6.304% | 1.266% | 5.512% | 6.592% | 3.883% | 2.828% | 2.296% | 4.221% | 3.656% | 4.921% | 0.8357% |  | 1.227% | 2.166% | 3.495% | 22.27% | 2.168% | 1.854% | 2.297% | 30.62%
| PDM3 |  |  | 0.01416% | 10.4% | -1.637% | -7.268% | 0.4026% | -3.967% | 1.276% | 0.2084% | 3.898% | -0.02468% | 5.916% | -4.817% |  | -42.71% | -14.94% | -8.434% | 6.811% | 9.724% | 14.87% | -5.958% | 55.36%
| Z-Score (Absolute) |  | 5.15 | 0 | 1.17 | 0.19 | 0.82 | 0.04 | 0.45 | 0.14 | 0.02 | 0.44 | 0 | 0.67 | 0.55 | 5.15 | 4.82 | 1.69 | 0.95 | NA | 1.1 | 1.68 | 0.67 | NA | 4.71
Results from laboratories 1.24 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 625 (Indicative Value 0.24 ppm) | [30 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | [6 empty] | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | [6 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | [6 empty] | 2
| 1 | 1 | 1.01: 0.2; 1.02: 0.241; 1.03: 0.25661; 1.04: 0.23328 | [6 empty] | 1
| 1 | 2 | 1.01: 0.2; 1.02: 0.25; 1.03: 0.25222; 1.04: 0.22098 | [6 empty] | 7
| 1 | 3 | 1.01: 0.2; 1.02: 0.247; 1.03: 0.25512; 1.04: 0.23927 | [6 empty] | 16
| 1 | 4 | 1.01: 0.2; 1.02: 0.249; 1.03: 0.25229; 1.04: 0.23406 | [6 empty] | 0.24104
| 1 | 5 | 1.01: 0.3; 1.02: 0.249; 1.03: 0.24923; 1.04: 0.23223 | [6 empty] | 13
| 1 | 6 | 1.01: 0.3; 1.02: 0.247; 1.03: 0.25022; 1.04: 0.22648
| Mean |  | 1.01: 0.23333; 1.02: 0.24717; 1.03: 0.25262; 1.04: 0.23105
| Median |  | 1.01: 0.2; 1.02: 0.248; 1.03: 0.25226; 1.04: 0.23275
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.0032506; 1.03: 0.0028166; 1.04: 0.0064071
| Rel.Std.Dev. |  | 1.01: 22.13%; 1.02: 1.315%; 1.03: 1.115%; 1.04: 2.773%
| PDM3 |  | 1.01: -3.198%; 1.02: 2.541%; 1.03: 4.802%; 1.04: -4.145%
| Z-Score (Absolute) |  | 1.01: 0.58; 1.02: 0.58; 1.03: 1.03; 1.04: 0.77
| Analytical results for In in OREAS 625 (Certified Value 2.07 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.14 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.5g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [10 empty] | 3
| 1 | 1 | < 5 | 1.83 | 2.15 | 1.915 | 2.37 | 1.965 | 2.15 | 1.97 | 2.06 | 3.57 | 1.9156 | 2.18 | 1.35 | 2.2 | 2.12 | 2.08 | 1.93 | 1.98 | 4.102 | [10 empty] | 1
| 1 | 2 | < 5 | 1.81 | 2.11 | 1.985 | 2.28 | 1.99 | 2.2 | 2.14 | 2.14 | 6.04 | 1.9197 | 2.01 | 1.34 | 2.2 | 2 | 2.08 | 2.06 | 1.962 | 3.24 | [10 empty] | 33
| 1 | 3 | < 5 | 1.83 | 2.14 | 2.05 | 2.41 | 1.96 | 2.15 | 2.1 | 2.1 | 4.7 | 1.9473 | 2.17 | 1.35 | 2.3 | 2.09 | 2.11 | 2 | 1.871 | 9.273 | [10 empty] | 16
| 1 | 4 | < 5 | 1.91 | 2.23 | 2.02 | 2.42 | 1.93 | 2.12 | 2.04 | 2.11 | 2.15 | 1.9462 | 2.1 | 1.39 | 2.2 | 2 | 2.02 | 1.98 | 1.962 | 2.484 | [10 empty] | 2.0651
| 1 | 5 | < 5 | 1.79 | 2.15 | 1.97 | 2.42 | 1.96 | 2.15 | 2.03 | 2.1 | 4.91 | 1.9326 | 2.17 | 1.4 | 2.2 | 2.08 | 2.01 | 1.98 | 1.883 | 2.731 | [10 empty] | 86
| 1 | 6 | < 5 | 1.88 | 2.14 | 2.11 | 2.3 | 1.965 | 2.13 | 2.11 | 2.1 | 4.88 | 1.9208 | 2.16 | 1.35 | 2.2 | 2.08 | 2.12 | 1.93 | 1.915 | 2.625
| Mean |  |  | 1.8417 | 2.1533 | 2.0083 | 2.3667 | 1.9617 | 2.15 | 2.065 | 2.1017 | 4.375 | 1.9304 | 2.1317 | 1.3633 | 2.2167 | 2.0617 | 2.07 | 1.98 | 1.9288 | 4.0758
| Median |  |  | 1.83 | 2.145 | 2.0025 | 2.39 | 1.9625 | 2.15 | 2.07 | 2.1 | 4.79 | 1.9267 | 2.165 | 1.35 | 2.2 | 2.08 | 2.08 | 1.98 | 1.9385 | 2.9855
| Std Dev. |  |  | 0.044907 | 0.040332 | 0.067725 | 0.062503 | 0.019149 | 0.027568 | 0.062849 | 0.025626 | 1.343 | 0.013877 | 0.066156 | 0.025033 | 0.040825 | 0.049967 | 0.045607 | 0.04858 | 0.045727 | 2.6137
| Rel.Std.Dev. |  |  | 2.438% | 1.873% | 3.372% | 2.641% | 0.9761% | 1.282% | 3.044% | 1.219% | 30.7% | 0.7189% | 3.104% | 1.836% | 1.842% | 2.424% | 2.203% | 2.454% | 2.371% | 64.13%
| PDM3 |  |  | -10.82% | 4.272% | -2.749% | 14.6% | -5.009% | 4.111% | -0.00497% | 1.771% | 111.9% | -6.524% | 3.223% | -33.98% | 7.339% | -0.1664% | 0.2371% | -4.121% | -6.599% | 97.37%
| Z-Score (Absolute) |  | 2.68 | 1.42 | 0.52 | 0.38 | 1.85 | 0.67 | 0.5 | 0.03 | 0.2 | 14.35 | 0.87 | 0.38 | 4.4 | 0.91 | 0.05 | 0 | 0.56 | 0.88 | 12.49
| Analytical results for K in OREAS 625 (Certified Value 0.276 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.29 | 0.26 | 0.26 | 0.26 | 0.28 | 0.25 | 0.25 | 0.26 | 0.23 | 0.2748 | 0.27 | 0.3852 | 0.28707 | 2.7e-05 | 0.36512 | 0.31 | 0.344 | 0.27 | 0.2824 | 0.42 | 0.3 | 0.25 | 0.26 | 0.27 | 0.29 | 0.32 | 0.26 |  |  | 1
| 1 | 2 | 0.28 | 0.26 | 0.24 | 0.26 | 0.28 | 0.26 | 0.25 | 0.26 | 0.25 | 0.2882 | 0.27 | 0.3778 | 0.28069 | 2.7e-05 | 0.37661 | 0.32 | 0.34 | 0.27 | 0.2731 | 0.43 | 0.28 | 0.25 | 0.25 | 0.26 | 0.29 | 0.31 | 0.27 |  |  | 0
| 1 | 3 | 0.29 | 0.25 | 0.26 | 0.26 | 0.28 | 0.26 | 0.25 | 0.26 | 0.24 | 0.2845 | 0.28 | 0.3639 | 0.27967 | 2.7e-05 | 0.37946 | 0.33 | 0.349 | 0.27 | 0.2821 | 0.44 | 0.3 | 0.24 | 0.26 | 0.27 | 0.29 | 0.31 | 0.26 |  |  | 16
| 1 | 4 | 0.29 | 0.26 | 0.24 | 0.27 | 0.28 | 0.26 | 0.25 | 0.26 | 0.26 | 0.2943 | 0.28 | 0.3703 | 0.27988 | 2.8e-05 | 0.38551 | 0.31 | 0.346 | 0.27 | 0.2822 | 0.43 | 0.29 | 0.24 | 0.25 | 0.26 | 0.31 | 0.31 | 0.26 |  |  | 0.27558
| 1 | 5 | 0.29 | 0.26 | 0.26 | 0.26 | 0.28 | 0.25 | 0.24 | 0.26 | 0.23 | 0.2883 | 0.28 | 0.3613 | 0.27975 | 2.8e-05 | 0.3668 | 0.31 | 0.348 | 0.28 | 0.2795 | 0.42 | 0.29 | 0.25 | 0.26 | 0.26 | 0.3 | 0.32 | 0.26 |  |  | 87
| 1 | 6 | 0.29 | 0.25 | 0.27 | 0.25 | 0.29 | 0.25 | 0.25 | 0.26 | 0.24 | 0.2781 | 0.27 | 0.3842 | 0.2759 | 2.7e-05 | 0.36629 | 0.34 | 0.342 | 0.27 | 0.282 | 0.42 | 0.29 | 0.24 | 0.26 | 0.27 | 0.3 | 0.32 | 0.27
| Mean |  | 0.28833 | 0.25667 | 0.255 | 0.26 | 0.28167 | 0.255 | 0.24833 | 0.26 | 0.24167 | 0.2847 | 0.275 | 0.37378 | 0.28049 | 2.7333e-05 | 0.3733 | 0.32 | 0.34483 | 0.27167 | 0.28022 | 0.42667 | 0.29167 | 0.245 | 0.25667 | 0.265 | 0.29667 | 0.315 | 0.26333
| Median |  | 0.29 | 0.26 | 0.26 | 0.26 | 0.28 | 0.255 | 0.25 | 0.26 | 0.24 | 0.28635 | 0.275 | 0.37405 | 0.27982 | 2.7e-05 | 0.3717 | 0.315 | 0.345 | 0.27 | 0.28205 | 0.425 | 0.29 | 0.245 | 0.26 | 0.265 | 0.295 | 0.315 | 0.26
| Std Dev. |  | 0.0040825 | 0.005164 | 0.012247 | 0.0063246 | 0.0040825 | 0.0054772 | 0.0040825 | 0 | 0.01169 | 0.0071969 | 0.0054772 | 0.010204 | 0.0036309 | 5.164e-07 | 0.0084414 | 0.012649 | 0.0034881 | 0.0040825 | 0.0036493 | 0.008165 | 0.0075277 | 0.0054772 | 0.005164 | 0.0054772 | 0.008165 | 0.0054772 | 0.005164
| Rel.Std.Dev. |  | 1.416% | 2.012% | 4.803% | 2.433% | 1.449% | 2.148% | 1.644% | 0% | 4.837% | 2.528% | 1.992% | 2.73% | 1.294% | 1.889% | 2.261% | 3.953% | 1.012% | 1.503% | 1.302% | 1.914% | 2.581% | 2.236% | 2.012% | 2.067% | 2.752% | 1.739% | 1.961%
| PDM3 |  | 4.628% | -6.863% | -7.468% | -5.653% | 2.209% | -7.468% | -9.887% | -5.653% | -12.31% | 3.309% | -0.2104% | 35.64% | 1.783% | -99.99% | 35.46% | 16.12% | 25.13% | -1.42% | 1.683% | 54.83% | 5.837% | -11.1% | -6.863% | -3.839% | 7.652% | 14.3% | -4.444%
| Z-Score (Absolute) |  | 0.36 | 0.71 | 0.76 | 0.59 | 0.14 | 0.76 | 0.99 | 0.59 | 1.21 | 0.24 | 0.09 | 3.24 | 0.1 | NA | 3.23 | 1.43 | 2.27 | 0.2 | 0.09 | 5.03 | 0.47 | 1.1 | 0.71 | 0.43 | 0.64 | 1.26 | 0.48
| Analytical results for La in OREAS 625 (Certified Value 17.7 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 1
| 1 | 1 | 15 | 13.2 | 16 | 18.1 | 17.9 | 17.2 | 17.4 | 17.6 | 16.1 | 16.13 | 15.4 | 31 | 19.9714 | 23.4768 | 16.4 | 30.2 | 15 | 21.87 | 29.2 | 21 | 15 | 18.1 | 16 | 17.1 | 22.4 | 18 |  |  |  | 1
| 1 | 2 | 16 | 12.9 | 15 | 18.4 | 17.8 | 16.8 | 17.4 | 16.8 | 17.7 | 18.047 | 14.4 | 30 | 19.951 | 23.477 | 17.5 | 30.3 | 14 | 21.48 | 29.9 | 18.6 | 15 | 17.3 | 16.6 | 17.5 | 22.2 | 18 |  |  |  | 34
| 1 | 3 | 15 | 13.6 | 16 | 17.8 | 18.1 | 17.2 | 17 | 18.1 | 17.4 | 16.946 | 16.2 | 29 | 19.9715 | 23.5401 | 18.2 | 30.5 | 15 | 21.54 | 30.7 | 20.6 | 16 | 18.2 | 16.7 | 17.5 | 22.5 | 17 |  |  |  | 16
| 1 | 4 | 16 | 13.2 | 15 | 19 | 17.8 | 17.4 | 16.4 | 16.7 | 16.3 | 17.43 | 15.7 | 30 | 19.8647 | 23.5195 | 18.3 | 29.2 | 15 | 21.55 | 28.9 | 19.3 | 15 | 17 | 17 | 18.7 | 23.3 | 18 |  |  |  | 17.6507
| 1 | 5 | 16 | 13.3 | 13 | 17.9 | 17.8 | 17.3 | 17.2 | 17.6 | 16.6 | 17.265 | 15.4 | 29 | 19.619 | 23.4978 | 17.4 | 29.7 | 15 | 21.66 | 29.5 | 20 | 16 | 17.6 | 15.9 | 17.4 | 22.3 | 17 |  |  |  | 88
| 1 | 6 | 16 | 13.3 | 16 | 17.7 | 19.6 | 17 | 16.9 | 17.9 | 17.5 | 16.621 | 14.5 | 32 | 19.7015 | 23.5474 | 20.1 | 29.1 | 15 | 21.56 | 29.3 | 20.1 | 15 | 18.1 | 16.3 | 17.4 | 23.5 | 18
| Mean |  | 15.6667 | 13.25 | 15.1667 | 18.15 | 18.1667 | 17.15 | 17.05 | 17.45 | 16.9333 | 17.0732 | 15.2667 | 30.1667 | 19.8465 | 23.5098 | 17.9833 | 29.8333 | 14.8333 | 21.61 | 29.5833 | 19.9333 | 15.3333 | 17.7167 | 16.4167 | 17.6 | 22.7 | 17.6667
| Median |  | 16 | 13.25 | 15.5 | 18 | 17.85 | 17.2 | 17.1 | 17.6 | 17 | 17.1055 | 15.4 | 30 | 19.9079 | 23.5087 | 17.85 | 29.95 | 15 | 21.555 | 29.4 | 20.05 | 15 | 17.85 | 16.45 | 17.45 | 22.45 | 18
| Std Dev. |  | 0.5164 | 0.22583 | 1.169 | 0.48477 | 0.71181 | 0.21679 | 0.37815 | 0.57533 | 0.68313 | 0.66673 | 0.69761 | 1.169 | 0.15182 | 0.030772 | 1.2416 | 0.59217 | 0.40825 | 0.14 | 0.64005 | 0.87101 | 0.5164 | 0.49565 | 0.42622 | 0.55857 | 0.55498 | 0.5164
| Rel.Std.Dev. |  | 3.296% | 1.704% | 7.708% | 2.671% | 3.918% | 1.264% | 2.218% | 3.297% | 4.034% | 3.905% | 4.57% | 3.875% | 0.765% | 0.1309% | 6.904% | 1.985% | 2.752% | 0.6478% | 2.164% | 4.37% | 3.368% | 2.798% | 2.596% | 3.174% | 2.445% | 2.923%
| PDM3 |  | -11.24% | -24.93% | -14.07% | 2.829% | 2.923% | -2.837% | -3.403% | -1.137% | -4.064% | -3.272% | -13.51% | 70.91% | 12.44% | 33.19% | 1.885% | 69.02% | -15.96% | 22.43% | 67.6% | 12.93% | -13.13% | 0.3737% | -6.991% | -0.2873% | 28.61% | 0.09045%
| Z-Score (Absolute) |  | 0.63 | 1.41 | 0.8 | 0.17 | 0.17 | 0.16 | 0.19 | 0.06 | 0.23 | 0.18 | 0.76 | 4.04 | 0.71 | 1.9 | 0.11 | 3.94 | 0.9 | 1.28 | 3.86 | 0.74 | 0.74 | 0.03 | 0.39 | 0.01 | 1.63 | 0.01
| Analytical results for Li in OREAS 625 (Certified Value 3.4 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.2g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [6 empty] | 2
| 1 | 1 | 4 | 3.9 | 3.3 | 3.5 | 3.3 | 3 | 3.3 | 2.5 | 3 | 5 | 3.3273 | 3.2 | 3.6059 | 4.3 | 3.9 | < 5 | 4.9 | 5 | 3 | 3 | 3 | 4 | 4 | [6 empty] | 1
| 1 | 2 | 4 | 3.7 | 3.2 | 3.6 | 3.4 | 3.1 | 3.3 | 2.7 | 3.2 | 5 | 3.2971 | 3.13 | 3.6751 | 4.3 | 3.5 | < 5 | 5.3 | 5 | 3 | 3 | 3 | < 1 | 3 | [6 empty] | 0
| 1 | 3 | 4 | 3.7 | 3.3 | 3.6 | 3.3 | 2.9 | 3.4 | 2.7 | 3.1 | 5 | 3.2345 | 3.22 | 3.8263 | 4.2 | 4.5 | < 5 | 5.4 | 5 | 3 | 3 | 3.1 | 4 | 3 | [6 empty] | 16
| 1 | 4 | 4 | 3.8 | 3.3 | 3.6 | 3.3 | 3 | 3.3 | 2.6 | 3.2 | 5 | 3.2363 | 4.01 | 3.5292 | 4.4 | 4 | < 5 | 5.4 | 5 | 3 | 3 | 3.2 | 1 | 3 | [6 empty] | 3.4013
| 1 | 5 | 4 | 4 | 3.2 | 3.6 | 3.2 | 3 | 3.3 | 2.7 | 3 | 5 | 3.3043 | 2.79 | 3.8132 | 4.4 | 3.3 | < 5 | 5.7 | 5 | 3 | 3 | 3.1 | 5 | 3 | [6 empty] | 89
| 1 | 6 | 4 | 3.8 | 3.1 | 3.6 | 3.3 | 3.1 | 3.2 | 2.5 | 2.9 | 5 | 3.2441 | 4.51 | 3.6538 | 4.4 | 3.8 | < 5 | 5.5 | 5 | 3 | 3 | 3.1 | 7 | 3
| Mean |  | 4 | 3.8167 | 3.2333 | 3.5833 | 3.3 | 3.0167 | 3.3 | 2.6167 | 3.0667 | 5 | 3.2739 | 3.4767 | 3.6839 | 4.3333 | 3.8333 |  | 5.3667 | 5 | 3 | 3 | 3.0833 | 4.2 | 3.1667
| Median |  | 4 | 3.8 | 3.25 | 3.6 | 3.3 | 3 | 3.3 | 2.65 | 3.05 | 5 | 3.2706 | 3.21 | 3.6644 | 4.35 | 3.85 |  | 5.4 | 5 | 3 | 3 | 3.1 | 4 | 3
| Std Dev. |  | 0 | 0.1169 | 0.08165 | 0.040825 | 0.063246 | 0.075277 | 0.063246 | 0.098319 | 0.12111 | 0 | 0.040413 | 0.64596 | 0.1166 | 0.08165 | 0.41793 |  | 0.26583 | 0 | 0 | 0 | 0.075277 | 2.1679 | 0.40825
| Rel.Std.Dev. |  | 0% | 3.063% | 2.525% | 1.139% | 1.917% | 2.495% | 1.917% | 3.757% | 3.949% | 0% | 1.234% | 18.58% | 3.165% | 1.884% | 10.9% |  | 4.953% | 0% | 0% | 0% | 2.441% | 51.62% | 12.89%
| PDM3 |  | 17.6% | 12.21% | -4.938% | 5.353% | -2.978% | -11.31% | -2.978% | -23.07% | -9.838% | 47% | -3.744% | 2.217% | 8.31% | 27.4% | 12.7% |  | 57.78% | 47% | -11.8% | -11.8% | -9.348% | 23.48% | -6.898%
| Z-Score (Absolute) |  | NA | 1.23 | 0.16 | 0.67 | 0 | 0.67 | 0 | 1.63 | 0.56 | NA | 0.06 | 0.42 | 0.91 | 2.46 | 1.27 | 1.9 | 4.92 | NA | NA | NA | 0.52 | NA | NA
Results from laboratories 1.01, 1.12, 1.23, 1.25, 1.26, 1.28 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 625 (Certified Value 0.045 ppm) | [30 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.23; 1.05: 1.25; 1.06: 1.26 |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.06: AR*OES/MS |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.25g; 1.05: 0.25g; 1.06: 0.25g |  |  |  | 3
| 1 | 1 | 1.01: < 0.1; 1.02: 0.04; 1.03: < 0.05; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04 |  |  |  | 1
| 1 | 2 | 1.01: < 0.1; 1.02: 0.041; 1.03: < 0.05; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04 |  |  |  | 35
| 1 | 3 | 1.01: < 0.1; 1.02: 0.041; 1.03: < 0.05; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04 |  |  |  | 16
| 1 | 4 | 1.01: < 0.1; 1.02: 0.04; 1.03: < 0.05; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04 |  |  |  | 0.045083
| 1 | 5 | 1.01: < 0.1; 1.02: 0.041; 1.03: < 0.05; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04 |  |  |  | 90
| 1 | 6 | 1.01: < 0.1; 1.02: 0.039; 1.03: < 0.05; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04
| Mean |  | 1.02: 0.040333; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04
| Median |  | 1.02: 0.0405; 1.04: 0.05; 1.05: 0.05; 1.06: 0.04
| Std Dev. |  | 1.02: 0.0008165; 1.04: 7.6012e-18; 1.05: 7.6012e-18; 1.06: 0
| Rel.Std.Dev. |  | 1.02: 2.024%; 1.04: 1.52e-14%; 1.05: 1.52e-14%; 1.06: 0%
| PDM3 |  | 1.02: -10.54%; 1.04: 10.91%; 1.05: 10.91%; 1.06: -11.28%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67; 1.03: 2.81; 1.04: 0.67; 1.05: 0.67; 1.06: 0.72
| Analytical results for Mg in OREAS 625 (Certified Value 0.119 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.12 | 0.11 | 0.12 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.12 | 0.12 | 0.13 | 0.12172 | 0.12384 | 0.14198 | 0.1 | 0.13 | 0.12 | 0.1264 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.124 | 0.12 |  |  | 1
| 1 | 2 | 0.11 | 0.11 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12029 | 0.12527 | 0.12386 | 0.11 | 0.13 | 0.12 | 0.124 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.119 | 0.12 |  |  | 0
| 1 | 3 | 0.12 | 0.11 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.13 | 0.11883 | 0.12075 | 0.13996 | 0.11 | 0.13 | 0.12 | 0.1254 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.118 | 0.12 |  |  | 16
| 1 | 4 | 0.12 | 0.11 | 0.12 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12013 | 0.12422 | 0.1455 | 0.11 | 0.13 | 0.12 | 0.1256 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.122 | 0.12 |  |  | 0.11907
| 1 | 5 | 0.12 | 0.12 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.12 | 0.12 | 0.13 | 0.11785 | 0.12639 | 0.1258 | 0.1 | 0.13 | 0.12 | 0.1253 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.123 | 0.12 |  |  | 91
| 1 | 6 | 0.12 | 0.12 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12015 | 0.12619 | 0.13384 | 0.11 | 0.13 | 0.12 | 0.1255 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.125 | 0.12
| Mean |  | 0.11833 | 0.11333 | 0.12 | 0.11333 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12667 | 0.12 | 0.12 | 0.13 | 0.11983 | 0.12444 | 0.13516 | 0.10667 | 0.13 | 0.12 | 0.12537 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.12183 | 0.12
| Median |  | 0.12 | 0.11 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12014 | 0.12475 | 0.1369 | 0.11 | 0.13 | 0.12 | 0.12545 | 0.14 | 0.12 | 0.1 | 0.11 | 0.12 | 0.12 | 0.1225 | 0.12
| Std Dev. |  | 0.0040825 | 0.005164 | 0 | 0.005164 | 0 | 0 | 0 | 0 | 0.005164 | 0 | 0 | 0 | 0.0013354 | 0.0020795 | 0.0088718 | 0.005164 | 0 | 0 | 0.00077632 | 0 | 0 | 1.5202e-17 | 0 | 0 | 0 | 0.0027869 | 0
| Rel.Std.Dev. |  | 3.45% | 4.556% | 0% | 4.556% | 0% | 0% | 0% | 0% | 4.077% | 0% | 0% | 0% | 1.114% | 1.671% | 6.564% | 4.841% | 0% | 0% | 0.6192% | 0% | 0% | 1.52e-14% | 0% | 0% | 0% | 2.287% | 0%
| PDM3 |  | -0.6209% | -4.82% | 0.7788% | -4.82% | -7.619% | -7.619% | -7.619% | -7.619% | 6.378% | 0.7788% | 0.7788% | 9.177% | 0.6335% | 4.511% | 13.51% | -10.42% | 9.177% | 0.7788% | 5.286% | 17.58% | 0.7788% | -16.02% | -7.619% | 0.7788% | 0.7788% | 2.318% | 0.7788%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 0 | 0.67 | 1.01 | 1.01 | 1.01 | 1.01 | 0.67 | 0 | 0 | 1.01 | 0.02 | 0.45 | 1.53 | 1.35 | 1.01 | 0 | 0.54 | 2.02 | 0 | 2.02 | 1.01 | 0 | 0 | 0.19 | 0
| Analytical results for Mn in OREAS 625 (Certified Value 0.434 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.3057 | 0.397 | 0.3615 | 0.438 | 0.423 | 0.441 | 0.425 | 0.434 | 0.394 | 0.4249 | 0.4406 | 0.4346 | 0.040687 | 0.40716 | 0.41402 | 0.4258 | 0.448 | 0.4068 | 0.4366 | 0.4651 | 0.4432 | 0.4492 | 0.4281 | 0.4283 | 0.4413 | 0.4557 | 0.4396 |  |  | 1
| 1 | 2 | 0.3051 | 0.385 | 0.3587 | 0.431 | 0.432 | 0.445 | 0.418 | 0.442 | 0.42 | 0.4314 | 0.4288 | 0.4338 | 0.040207 | 0.40143 | 0.4212 | 0.4243 | 0.444 | 0.3945 | 0.4306 | 0.4287 | 0.4425 | 0.4578 | 0.4223 | 0.4544 | 0.4388 | 0.4333 | 0.4387 |  |  | 21
| 1 | 3 | 0.3037 | 0.39 | 0.3493 | 0.442 | 0.437 | 0.447 | 0.414 | 0.435 | 0.411 | 0.4341 | 0.4417 | 0.4192 | 0.039719 | 0.40906 | 0.43005 | 0.429 | 0.446 | 0.4037 | 0.4352 | 0.4307 | 0.4501 | 0.4615 | 0.4286 | 0.4524 | 0.43 | 0.4344 | 0.4324 |  |  | 16
| 1 | 4 | 0.3075 | 0.397 | 0.3462 | 0.448 | 0.438 | 0.453 | 0.424 | 0.44 | 0.424 | 0.4427 | 0.4387 | 0.4452 | 0.040151 | 0.4027 | 0.42602 | 0.4281 | 0.444 | 0.4073 | 0.4355 | 0.438 | 0.4375 | 0.4614 | 0.4231 | 0.4435 | 0.4562 | 0.4419 | 0.4347 |  |  | 0.43395
| 1 | 5 | 0.3073 | 0.389 | 0.3617 | 0.436 | 0.438 | 0.434 | 0.419 | 0.437 | 0.397 | 0.4425 | 0.4193 | 0.4281 | 0.039394 | 0.42008 | 0.43231 | 0.4245 | 0.446 | 0.4123 | 0.4364 | 0.4204 | 0.4424 | 0.4542 | 0.426 | 0.4377 | 0.4406 | 0.448 | 0.4357 |  |  | 92
| 1 | 6 | 0.306 | 0.395 | 0.3551 | 0.429 | 0.431 | 0.445 | 0.419 | 0.426 | 0.417 | 0.4299 | 0.4381 | 0.4468 | 0.040159 | 0.41332 | 0.41951 | 0.4301 | 0.442 | 0.3889 | 0.4356 | 0.4227 | 0.4415 | 0.4477 | 0.4272 | 0.4384 | 0.4428 | 0.4559 | 0.433
| Mean |  | 0.30588 | 0.39217 | 0.35542 | 0.43733 | 0.43317 | 0.44417 | 0.41983 | 0.43567 | 0.4105 | 0.43425 | 0.43453 | 0.43462 | 0.040053 | 0.40896 | 0.42385 | 0.42697 | 0.445 | 0.40225 | 0.43498 | 0.43427 | 0.44287 | 0.4553 | 0.42588 | 0.44245 | 0.44162 | 0.44487 | 0.43568
| Median |  | 0.30585 | 0.3925 | 0.3569 | 0.437 | 0.4345 | 0.445 | 0.419 | 0.436 | 0.414 | 0.43275 | 0.4384 | 0.4342 | 0.040155 | 0.40811 | 0.42361 | 0.42695 | 0.445 | 0.40525 | 0.43555 | 0.4297 | 0.44245 | 0.456 | 0.4266 | 0.44095 | 0.44095 | 0.44495 | 0.4352
| Std Dev. |  | 0.0014176 | 0.004916 | 0.0064765 | 0.0070333 | 0.0058452 | 0.0063377 | 0.0040702 | 0.0056095 | 0.012406 | 0.0071268 | 0.0087541 | 0.010404 | 0.0004455 | 0.0069535 | 0.0068897 | 0.002441 | 0.0020976 | 0.0088045 | 0.0022149 | 0.016337 | 0.0040854 | 0.0059679 | 0.0026317 | 0.0098193 | 0.0084627 | 0.010011 | 0.0029458
| Rel.Std.Dev. |  | 0.4635% | 1.254% | 1.822% | 1.608% | 1.349% | 1.427% | 0.9695% | 1.288% | 3.022% | 1.641% | 2.015% | 2.394% | 1.112% | 1.7% | 1.625% | 0.5717% | 0.4714% | 2.189% | 0.5092% | 3.762% | 0.9225% | 1.311% | 0.6179% | 2.219% | 1.916% | 2.25% | 0.6761%
| PDM3 |  | -29.51% | -9.629% | -18.1% | 0.7792% | -0.1809% | 2.354% | -3.253% | 0.3952% | -5.404% | 0.06872% | 0.134% | 0.1532% | -90.77% | -5.76% | -2.328% | -1.61% | 2.546% | -7.305% | 0.2377% | 0.07256% | 2.054% | 4.919% | -1.859% | 1.958% | 1.766% | 2.515% | 0.399%
| Z-Score (Absolute) |  | 10.07 | 3.3 | 6.18 | 0.24 | 0.09 | 0.78 | 1.13 | 0.11 | 1.86 | 0 | 0.02 | 0.03 | 30.91 | 1.98 | 0.82 | 0.57 | 0.84 | 2.51 | 0.06 | 0 | 0.67 | 1.65 | 0.66 | 0.64 | 0.58 | 0.83 | 0.11
| Analytical results for Mo in OREAS 625 (Certified Value 11.6 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 2
| 1 | 1 | 11 | 10.5 | 10.6 | 11.55 | 11.25 | 11.8 | 11.5 | 12.5 | 11.8 | 11.8 | 11.38 | 12 | 12.1471 | 10.97 | 11.6985 | 10.7 | 12.2 | 11.2 | 10.75 | 11.3 | 12.45 | 11.95 | 12 | 12.5 | 12.69 | 11.1 | 11 | 9 |  | 1
| 1 | 2 | 11 | 10.7 | 10.7 | 11.5 | 11.85 | 11.55 | 11.3 | 12.15 | 12.7 | 12.1 | 11.12 | 12 | 11.7389 | 10.55 | 11.6562 | 11.8 | 12.3 | 10.9 | 10.62 | 11.5 | 12.5 | 11.8 | 12 | 12.7 | 12.91 | 11.47 | 12 | 9 |  | 37
| 1 | 3 | 11 | 10.8 | 10.4 | 12.4 | 12.2 | 11.95 | 11.1 | 12.2 | 12.1 | 12.2 | 11.66 | 11 | 11.6877 | 11.32 | 12.1923 | 10.7 | 12.4 | 11.4 | 10.72 | 11.5 | 12.76 | 11.91 | 12 | 12.5 | 12.79 | 10.92 | 13 | 9 |  | 16
| 1 | 4 | 11 | 10.7 | 10.7 | 13.15 | 11.95 | 12.05 | 10.8 | 11.9 | 12 | 12.3 | 11.43 | 12 | 11.637 | 10.76 | 12.1936 | 11.6 | 12.5 | 11.2 | 10.74 | 11.4 | 12.3 | 11.22 | 13 | 12 | 12.7 | 11.51 | 13 | 9 |  | 11.6409
| 1 | 5 | 11 | 10.7 | 10.4 | 12.25 | 11.85 | 11.65 | 10.85 | 12 | 11.8 | 12.2 | 11.23 | 12 | 11.8639 | 11.65 | 11.3766 | 11 | 12.4 | 11.4 | 10.68 | 11.4 | 12.45 | 11.47 | 12 | 12.2 | 12.5 | 10.96 | 13 | 8 |  | 93
| 1 | 6 | 11 | 10.8 | 10.4 | 12.1 | 12.25 | 11.6 | 10.85 | 11.45 | 12.4 | 12 | 11.1 | 12 | 11.9878 | 11.04 | 11.9489 | 10.9 | 12.2 | 11.2 | 10.77 | 11.3 | 12.5 | 11.81 | 12 | 12.7 | 12.7 | 11.16 | 14 | 9
| Mean |  | 11 | 10.7 | 10.5333 | 12.1583 | 11.8917 | 11.7667 | 11.0667 | 12.0333 | 12.1333 | 12.1 | 11.32 | 11.8333 | 11.8438 | 11.0483 | 11.8444 | 11.1167 | 12.3333 | 11.2167 | 10.7133 | 11.4 | 12.4933 | 11.6933 | 12.1667 | 12.4333 | 12.715 | 11.1867 | 12.6667 | 8.8333
| Median |  | 11 | 10.7 | 10.5 | 12.175 | 11.9 | 11.725 | 10.975 | 12.075 | 12.05 | 12.15 | 11.305 | 12 | 11.8014 | 11.005 | 11.8237 | 10.95 | 12.35 | 11.2 | 10.73 | 11.4 | 12.475 | 11.805 | 12 | 12.5 | 12.7 | 11.13 | 13 | 9
| Std Dev. |  | 0 | 0.10954 | 0.15055 | 0.60944 | 0.35835 | 0.20166 | 0.28577 | 0.35166 | 0.3559 | 0.17889 | 0.21345 | 0.40825 | 0.19559 | 0.39331 | 0.32539 | 0.47081 | 0.12111 | 0.18348 | 0.055015 | 0.089443 | 0.14989 | 0.28696 | 0.40825 | 0.28048 | 0.13487 | 0.25121 | 1.0328 | 0.40825
| Rel.Std.Dev. |  | 0% | 1.024% | 1.429% | 5.013% | 3.013% | 1.714% | 2.582% | 2.922% | 2.933% | 1.478% | 1.886% | 3.45% | 1.651% | 3.56% | 2.747% | 4.235% | 0.9819% | 1.636% | 0.5135% | 0.7846% | 1.2% | 2.454% | 3.355% | 2.256% | 1.061% | 2.246% | 8.154% | 4.622%
| PDM3 |  | -5.506% | -8.083% | -9.515% | 4.445% | 2.154% | 1.08% | -4.933% | 3.371% | 4.23% | 3.944% | -2.757% | 1.653% | 1.742% | -5.091% | 1.748% | -4.504% | 5.948% | -3.645% | -7.968% | -2.07% | 7.322% | 0.4501% | 4.516% | 6.807% | 9.227% | -3.902% | 8.811% | -24.12%
| Z-Score (Absolute) |  | NA | 1.31 | 1.51 | 0.48 | 0.15 | 0 | 0.86 | 0.33 | 0.45 | 0.41 | 0.55 | NA | 0.09 | 0.88 | 0.1 | 0.8 | 0.69 | 0.67 | 1.29 | 0.45 | 0.89 | 0.09 | NA | 0.82 | 1.16 | 0.71 | NA | NA
Results from laboratories 1.01, 1.12, 1.24, 1.28 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 625 (Certified Value 0.059 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.06 | 0.049 | 0.056 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 0.05 | 0.06 | 0.06 | 0.0642 | < 0.1 | 0.059066 | 0.09 | 0.06 | 0.05 | 0.0655 | 0.09 | 0.07 | 0.07 | 0.06 | 0.05 | 0.06 | 0.068 | 0.07 |  |  | 1
| 1 | 2 | 0.06 | 0.048 | 0.054 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 0.05 | 0.06 | 0.06 | 0.062839 | < 0.1 | 0.061529 | 0.09 | 0.06 | 0.06 | 0.0654 | 0.09 | 0.06 | 0.07 | 0.06 | 0.05 | 0.06 | 0.069 | 0.07 |  |  | 13
| 1 | 3 | 0.06 | 0.049 | 0.057 | 0.05 | 0.06 | 0.06 | 0.05 | 0.07 | 0.05 | 0.05 | 0.062 | 0.06 | 0.062932 | < 0.1 | 0.060164 | 0.1 | 0.06 | 0.06 | 0.0653 | 0.09 | 0.06 | 0.07 | 0.06 | 0.05 | 0.06 | 0.07 | 0.07 |  |  | 16
| 1 | 4 | 0.06 | 0.05 | 0.054 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 0.06 | 0.062 | 0.06 | 0.063171 | < 0.1 | 0.060388 | 0.09 | 0.06 | 0.05 | 0.0656 | 0.09 | 0.06 | 0.07 | 0.06 | 0.05 | 0.06 | 0.067 | 0.07 |  |  | 0.058867
| 1 | 5 | 0.06 | 0.049 | 0.056 | 0.05 | 0.06 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.062 | 0.06 | 0.063174 | < 0.1 | 0.058299 | 0.09 | 0.06 | 0.05 | 0.0666 | 0.09 | 0.06 | 0.07 | 0.06 | 0.05 | 0.06 | 0.068 | 0.07 |  |  | 94
| 1 | 6 | 0.06 | 0.051 | 0.057 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 0.05 | 0.059 | 0.06 | 0.062576 | < 0.1 | 0.05797 | 0.1 | 0.06 | 0.06 | 0.0652 | 0.09 | 0.06 | 0.07 | 0.06 | 0.06 | 0.06 | 0.07 | 0.07
| Mean |  | 0.06 | 0.049333 | 0.055667 | 0.05 | 0.06 | 0.058333 | 0.05 | 0.061667 | 0.05 | 0.051667 | 0.060833 | 0.06 | 0.063149 |  | 0.059569 | 0.093333 | 0.06 | 0.055 | 0.0656 | 0.09 | 0.061667 | 0.07 | 0.06 | 0.051667 | 0.06 | 0.068667 | 0.07
| Median |  | 0.06 | 0.049 | 0.056 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 0.05 | 0.061 | 0.06 | 0.063051 |  | 0.059615 | 0.09 | 0.06 | 0.055 | 0.06545 | 0.09 | 0.06 | 0.07 | 0.06 | 0.05 | 0.06 | 0.0685 | 0.07
| Std Dev. |  | 0 | 0.0010328 | 0.0013663 | 7.6012e-18 | 0 | 0.0040825 | 7.6012e-18 | 0.0040825 | 7.6012e-18 | 0.0040825 | 0.0013292 | 0 | 0.0005617 |  | 0.0013631 | 0.005164 | 0 | 0.0054772 | 0.0005099 | 1.5202e-17 | 0.0040825 | 0 | 0 | 0.0040825 | 0 | 0.0012111 | 0
| Rel.Std.Dev. |  | 0% | 2.094% | 2.454% | 1.52e-14% | 0% | 6.999% | 1.52e-14% | 6.62% | 1.52e-14% | 7.902% | 2.185% | 0% | 0.8895% |  | 2.288% | 5.533% | 0% | 9.959% | 0.7773% | 1.689e-14% | 6.62% | 0% | 0% | 7.902% | 0% | 1.764% | 0%
| PDM3 |  | 1.924% | -16.2% | -5.437% | -15.06% | 1.924% | -0.9073% | -15.06% | 4.755% | -15.06% | -12.23% | 3.34% | 1.924% | 7.273% |  | 1.192% | 58.55% | 1.924% | -6.57% | 11.44% | 52.89% | 4.755% | 18.91% | 1.924% | -12.23% | 1.924% | 16.65% | 18.91%
| Z-Score (Absolute) |  | 0 | 1.44 | 0.58 | 1.35 | 0 | 0.22 | 1.35 | 0.22 | 1.35 | 1.12 | 0.11 | 0 | 0.42 | 1.35 | 0.06 | 4.5 | 0 | 0.67 | 0.76 | 4.05 | 0.22 | 1.35 | 0 | 1.12 | 0 | 1.17 | 1.35
| Analytical results for Nb in OREAS 625 (Certified Value 0.7 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 2
| 1 | 1 | < 5 | 0.5 | 0.63 | 0.71 | 0.62 | 0.59 | 0.48 | 1 | 1 | 0.58146 | 0.97709 | 1 | 0.3 | 0.82 | 0.67 | 0.8 | 0.6 | 0.75 | [11 empty] | 1
| 1 | 2 | < 5 | 0.5 | 0.62 | 0.68 | 0.63 | 0.57 | 0.46 | 0.93 | 1.01 | 0.57151 | 0.97389 | 1 | 0.3 | 0.7 | 0.65 | 0.79 | 0.56 | 0.82 | [11 empty] | 38
| 1 | 3 | < 5 | 0.6 | 0.59 | 0.7 | 0.64 | 0.55 | 0.51 | 0.87 | 0.96 | 0.55359 | 0.97219 | 1.1 | 0.3 | 0.77 | 0.68 | 0.79 | 0.57 | 0.93 | [11 empty] | 16
| 1 | 4 | < 5 | 0.5 | 0.62 | 0.69 | 0.63 | 0.55 | 0.49 | 0.98 | 1 | 0.5851 | 0.9244 | 1.1 | 0.3 | 0.75 | 0.67 | 0.8 | 0.66 | 0.66 | [11 empty] | 0.69694
| 1 | 5 | < 5 | 0.6 | 0.6 | 0.69 | 0.67 | 0.58 | 0.48 | 1.03 | 0.97 | 0.60073 | 0.92799 | 1.1 | 0.3 | 0.78 | 0.67 | 0.8 | 0.58 | 0.66 | [11 empty] | 95
| 1 | 6 | < 5 | 0.6 | 0.63 | 0.75 | 0.65 | 0.57 | 0.51 | 0.94 | 1.01 | 0.57928 | 0.92585 | 1.1 | 0.3 | 0.75 | 0.71 | 0.83 | 0.57 | 0.74
| Mean |  |  | 0.55 | 0.615 | 0.70333 | 0.64 | 0.56833 | 0.48833 | 0.95833 | 0.99167 | 0.57861 | 0.95023 | 1.0667 | 0.3 | 0.76167 | 0.675 | 0.80167 | 0.59 | 0.76
| Median |  |  | 0.55 | 0.62 | 0.695 | 0.635 | 0.57 | 0.485 | 0.96 | 1 | 0.58037 | 0.95009 | 1.1 | 0.3 | 0.76 | 0.67 | 0.8 | 0.575 | 0.745
| Std Dev. |  |  | 0.054772 | 0.016432 | 0.025033 | 0.017889 | 0.016021 | 0.019408 | 0.057067 | 0.02137 | 0.015598 | 0.026531 | 0.05164 | 0 | 0.039707 | 0.019748 | 0.01472 | 0.036878 | 0.10296
| Rel.Std.Dev. |  |  | 9.959% | 2.672% | 3.559% | 2.795% | 2.819% | 3.974% | 5.955% | 2.155% | 2.696% | 2.792% | 4.841% | 0% | 5.213% | 2.926% | 1.836% | 6.251% | 13.55%
| PDM3 |  |  | -21.08% | -11.76% | 0.9178% | -8.17% | -18.45% | -29.93% | 37.51% | 42.29% | -16.98% | 36.34% | 53.05% | -56.95% | 9.288% | -3.148% | 15.03% | -15.34% | 9.049%
| Z-Score (Absolute) |  | 10.69 | NA | 0.53 | 0 | 0.38 | 0.8 | 1.28 | 1.52 | 1.72 | 0.74 | 1.47 | NA | NA | 0.35 | 0.17 | 0.59 | 0.67 | 0.34
Results from laboratories 1.02, 1.17 and 1.18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 625 (Indicative Value 17.2 ppm) | [30 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  |  | 1
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  |  | 1
| 1 | 1 | 1.01: 14.1; 1.02: 15.211; 1.03: 18.4925; 1.04: 20.7026 |  |  | 1
| 1 | 2 | 1.01: 13.5; 1.02: 16.797; 1.03: 18.3754; 1.04: 20.2406 |  |  | 8
| 1 | 3 | 1.01: 14; 1.02: 15.657; 1.03: 17.888; 1.04: 20.8046 |  |  | 16
| 1 | 4 | 1.01: 13.8; 1.02: 16.352; 1.03: 18.3918; 1.04: 20.865 |  |  | 17.1554
| 1 | 5 | 1.01: 14; 1.02: 16.337; 1.03: 18.1409; 1.04: 20.4572 |  |  | 14
| 1 | 6 | 1.01: 14; 1.02: 15.508; 1.03: 18.0392; 1.04: 20.0706
| Mean |  | 1.01: 13.9; 1.02: 15.977; 1.03: 18.2213; 1.04: 20.5234
| Median |  | 1.01: 14; 1.02: 15.997; 1.03: 18.2582; 1.04: 20.5799
| Std Dev. |  | 1.01: 0.21909; 1.02: 0.60856; 1.03: 0.2354; 1.04: 0.32164
| Rel.Std.Dev. |  | 1.01: 1.576%; 1.02: 3.809%; 1.03: 1.292%; 1.04: 1.567%
| PDM3 |  | 1.01: -18.98%; 1.02: -6.869%; 1.03: 6.213%; 1.04: 19.63%
| Z-Score (Absolute) |  | 1.01: 1; 1.02: 0.35; 1.03: 0.35; 1.04: 1.07
| Analytical results for Ni in OREAS 625 (Certified Value 5.32 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 6 | 4.4 | 4.9 | 4.9 | 5.4 | 5.1 | 5.3 | 5.4 | 5.5 | 5.3 | 5.4 | 5 | 5.8004 | 3.2017 | 5.587 | 5.68 | 2 | 4.4 | 10.54 | 5.8 | 5.09 | 5.7 | 8 | 5.8 | 5.7 | 5.2 | 5 | 6 |  | 1
| 1 | 2 | 6 | 4.8 | 4.7 | 4.8 | 5.5 | 5 | 5.4 | 5.5 | 6 | 5.4 | 5.3 | 5 | 5.7163 | 3.0498 | 5.0703 | 5.42 | 2 | 4.1 | 10.01 | 6.1 | 4.93 | 5.7 | 8 | 5.8 | 6.2 | 5.3 | 5 | 6 |  | 14
| 1 | 3 | 6 | 4.4 | 4.6 | 4.9 | 5.7 | 5.3 | 5.3 | 5.4 | 5.8 | 5.6 | 5.5 | 5 | 5.6844 | 3.3078 | 5.0241 | 5.21 | 3 | 4.4 | 10.25 | 6.1 | 4.78 | 5.8 | 8 | 5.8 | 5.9 | 5.2 | 5 | 6 |  | 16
| 1 | 4 | 6 | 4.6 | 4.5 | 5 | 5.6 | 5.5 | 5.3 | 5.2 | 5.7 | 5.5 | 5.4 | 5 | 5.4858 | 3.4661 | 5.6074 | 5 | 2 | 4.3 | 10.51 | 6.2 | 4.87 | 5.3 | 8 | 5.5 | 6 | 5.2 | 5 | 5 |  | 5.3202
| 1 | 5 | 6 | 4.3 | 4.9 | 4.9 | 5.6 | 5.2 | 5.4 | 5.4 | 5.6 | 5.6 | 5.5 | 5 | 5.6703 | 3.7134 | 6.0237 | 5.01 | 3 | 4.4 | 10.55 | 5.8 | 4.93 | 5.5 | 9 | 5.6 | 5.9 | 5.3 | 5 | 5 |  | 96
| 1 | 6 | 6 | 5.2 | 4.7 | 4.8 | 5.8 | 5 | 5.3 | 5.3 | 6 | 5.4 | 5.3 | 5 | 5.6253 | 3.4001 | 5.6281 | 4.49 | 2 | 4.5 | 10.34 | 6 | 4.79 | 5.6 | 9 | 5.9 | 5.7 | 5.3 | 4 | 6
| Mean |  | 6 | 4.6167 | 4.7167 | 4.8833 | 5.6 | 5.1833 | 5.3333 | 5.3667 | 5.7667 | 5.4667 | 5.4 | 5 | 5.6638 | 3.3565 | 5.4901 | 5.135 | 2.3333 | 4.35 | 10.3667 | 6 | 4.8983 | 5.6 | 8.3333 | 5.7333 | 5.9 | 5.25 | 4.8333 | 5.6667
| Median |  | 6 | 4.5 | 4.7 | 4.9 | 5.6 | 5.15 | 5.3 | 5.4 | 5.75 | 5.45 | 5.4 | 5 | 5.6774 | 3.354 | 5.5972 | 5.11 | 2 | 4.4 | 10.425 | 6.05 | 4.9 | 5.65 | 8 | 5.8 | 5.9 | 5.25 | 5 | 6
| Std Dev. |  | 0 | 0.33714 | 0.16021 | 0.075277 | 0.14142 | 0.19408 | 0.05164 | 0.10328 | 0.20656 | 0.12111 | 0.089443 | 0 | 0.10495 | 0.22878 | 0.37954 | 0.40835 | 0.5164 | 0.13784 | 0.21248 | 0.16733 | 0.11427 | 0.17889 | 0.5164 | 0.15055 | 0.18974 | 0.054772 | 0.40825 | 0.5164
| Rel.Std.Dev. |  | 0% | 7.303% | 3.397% | 1.542% | 2.525% | 3.744% | 0.9682% | 1.924% | 3.582% | 2.215% | 1.656% | 0% | 1.853% | 6.816% | 6.913% | 7.952% | 22.13% | 3.169% | 2.05% | 2.789% | 2.333% | 3.194% | 6.197% | 2.626% | 3.216% | 1.043% | 8.447% | 9.113%
| PDM3 |  | 12.78% | -13.22% | -11.34% | -8.211% | 5.259% | -2.572% | 0.247% | 0.8735% | 8.392% | 2.753% | 1.5% | -6.018% | 6.458% | -36.91% | 3.194% | -3.481% | -56.14% | -18.24% | 94.86% | 12.78% | -7.929% | 5.259% | 56.64% | 7.766% | 10.9% | -1.319% | -9.151% | 6.512%
| Z-Score (Absolute) |  | NA | 1.64 | 1.43 | 1.07 | 0.46 | 0.43 | 0.11 | 0.04 | 0.82 | 0.18 | 0.04 | NA | 0.6 | 4.34 | 0.23 | 0.53 | NA | 2.21 | 10.66 | 1.32 | 1.04 | 0.46 | NA | 0.75 | 1.11 | 0.29 | NA | NA
Results from laboratories 1.01, 1.12, 1.17, 1.24, 1.28 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 625 (Certified Value 0.025 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 3
| 1 | 1 | 0.0268 | 0.026 | 0.024 | 0.025 | 0.026 | 0.026 | 0.025 | 0.025 | 0.0225 | 0.0238 | 0.024 | 0.0245 | 0.026949 | 0.027431 | 0.03 | 0.0261 | 0.022 | 0.0242 | 0.03 | 0.03 | 0.02 | 0.03 | 0.0256 | 0.0259 | 0.024 | 0.0235 |  |  |  | 1
| 1 | 2 | 0.0254 | 0.027 | 0.024 | 0.026 | 0.026 | 0.025 | 0.024 | 0.026 | 0.0244 | 0.0239 | 0.023 | 0.025 | 0.026344 | 0.026016 | 0.03 | 0.0259 | 0.022 | 0.0239 | 0.03 | 0.03 | 0.021 | 0.03 | 0.0264 | 0.0258 | 0.025 | 0.0236 |  |  |  | 24
| 1 | 3 | 0.0266 | 0.026 | 0.024 | 0.025 | 0.027 | 0.026 | 0.024 | 0.025 | 0.0232 | 0.0243 | 0.024 | 0.0237 | 0.025913 | 0.027199 | 0.03 | 0.0261 | 0.021 | 0.0241 | 0.03 | 0.03 | 0.021 | 0.03 | 0.0253 | 0.0251 | 0.025 | 0.0234 |  |  |  | 16
| 1 | 4 | 0.0263 | 0.027 | 0.023 | 0.026 | 0.026 | 0.025 | 0.024 | 0.025 | 0.0245 | 0.0245 | 0.024 | 0.0254 | 0.026922 | 0.027815 | 0.03 | 0.0257 | 0.022 | 0.0246 | 0.03 | 0.02 | 0.019 | 0.03 | 0.0257 | 0.0267 | 0.025 | 0.0237 |  |  |  | 0.025122
| 1 | 5 | 0.024 | 0.026 | 0.024 | 0.025 | 0.027 | 0.025 | 0.025 | 0.025 | 0.0228 | 0.0247 | 0.024 | 0.025 | 0.026397 | 0.02569 | 0.03 | 0.0263 | 0.022 | 0.0244 | 0.03 | 0.03 | 0.021 | 0.03 | 0.0252 | 0.0258 | 0.026 | 0.0235 |  |  |  | 97
| 1 | 6 | 0.0252 | 0.027 | 0.024 | 0.025 | 0.026 | 0.025 | 0.024 | 0.025 | 0.0239 | 0.0236 | 0.023 | 0.0252 | 0.026227 | 0.027165 | 0.03 | 0.0265 | 0.02 | 0.0245 | 0.03 | 0.03 | 0.019 | 0.03 | 0.0257 | 0.026 | 0.026 | 0.0233
| Mean |  | 0.025717 | 0.0265 | 0.023833 | 0.025333 | 0.026333 | 0.025333 | 0.024333 | 0.025167 | 0.02355 | 0.024133 | 0.023667 | 0.0248 | 0.026458 | 0.026886 | 0.03 | 0.0261 | 0.0215 | 0.024283 | 0.03 | 0.028333 | 0.020167 | 0.03 | 0.02565 | 0.025883 | 0.025167 | 0.0235
| Median |  | 0.02585 | 0.0265 | 0.024 | 0.025 | 0.026 | 0.025 | 0.024 | 0.025 | 0.02355 | 0.0241 | 0.024 | 0.025 | 0.02637 | 0.027182 | 0.03 | 0.0261 | 0.022 | 0.0243 | 0.03 | 0.03 | 0.0205 | 0.03 | 0.02565 | 0.02585 | 0.025 | 0.0235
| Std Dev. |  | 0.0010591 | 0.00054772 | 0.00040825 | 0.0005164 | 0.0005164 | 0.0005164 | 0.0005164 | 0.00040825 | 0.00084083 | 0.00043205 | 0.0005164 | 0.00061644 | 0.0004059 | 0.0008396 | 0 | 0.00028284 | 0.00083666 | 0.00026394 | 0 | 0.0040825 | 0.00098319 | 0 | 0.00042308 | 0.00051153 | 0.00075277 | 0.00014142
| Rel.Std.Dev. |  | 4.118% | 2.067% | 1.713% | 2.038% | 1.961% | 2.038% | 2.122% | 1.622% | 3.57% | 1.79% | 2.182% | 2.486% | 1.534% | 3.123% | 0% | 1.084% | 3.891% | 1.087% | 0% | 14.41% | 4.875% | 0% | 1.649% | 1.976% | 2.991% | 0.6018%
| PDM3 |  | 2.368% | 5.486% | -5.129% | 0.8417% | 4.822% | 0.8417% | -3.139% | 0.1783% | -6.257% | -3.935% | -5.793% | -1.281% | 5.321% | 7.023% | 19.42% | 3.894% | -14.42% | -3.338% | 19.42% | 12.78% | -19.72% | 19.42% | 2.102% | 3.031% | 0.1783% | -6.456%
| Z-Score (Absolute) |  | 0.23 | 0.69 | 0.88 | 0 | 0.59 | 0 | 0.59 | 0.1 | 1.05 | 0.71 | 0.98 | 0.31 | 0.66 | 0.91 | 2.75 | 0.45 | 2.26 | 0.62 | 2.75 | 1.77 | 3.04 | 2.75 | 0.19 | 0.32 | 0.1 | 1.08
| Analytical results for Pb in OREAS 625 (Certified Value 0.826 wt.%) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 0.7546 | 0.815 | 0.78738 | 0.841 | 0.821 | 0.807 | 0.803 | 0.815 | 0.84 | > 0.5 | 0.8065 | > 0.5 | 0.84667 | 0.851 | 0.86162 | 0.8422 | > 0.5 | 0.70236 | 0.821 | 0.7418 | 0.8335 | 0.8715 | 0.812 | 0.8167 | 0.82546 | 0.82271 | 0.8245 | 0.81735 |  | 1
| 1 | 2 | 0.7504 | 0.821 | 0.7872 | 0.852 | 0.842 | 0.807 | 0.812 | 0.834 | 0.89 | > 0.5 | 0.78705 | > 0.5 | 0.82792 | 0.82837 | 0.86547 | 0.8344 | > 0.5 | 0.70502 | 0.8274 | 0.7661 | 0.8364 | 0.8654 | 0.8165 | 0.8131 | 0.83019 | 0.81831 | 0.8016 | 0.81927 |  | 25
| 1 | 3 | 0.7418 | 0.837 | 0.76929 | 0.85 | 0.855 | 0.815 | 0.802 | 0.826 | 0.871 | > 0.5 | 0.81031 | > 0.5 | 0.8339 | 0.83992 | 0.86295 | 0.8528 | > 0.5 | 0.73026 | 0.8254 | 0.7519 | 0.8378 | 0.8397 | 0.8317 | 0.8137 | 0.82644 | 0.79872 | 0.7938 | 0.84439 |  | 16
| 1 | 4 | 0.7554 | 0.827 | 0.76483 | 0.858 | 0.846 | 0.825 | 0.801 | 0.83 | 0.919 | > 0.5 | 0.8382 | > 0.5 | 0.82888 | 0.83466 | 0.8604 | 0.8404 | > 0.5 | 0.7148 | 0.8267 | 0.7709 | 0.8386 | 0.8698 | 0.8239 | 0.8161 | 0.81967 | 0.84258 | 0.8054 | 0.80999 |  | 0.82622
| 1 | 5 | 0.7538 | 0.828 | 0.79781 | 0.837 | 0.84 | 0.799 | 0.798 | 0.831 | 0.853 | > 0.5 | 0.82205 | > 0.5 | 0.8266 | 0.85324 | 0.86238 | 0.8463 | > 0.5 | 0.72396 | 0.8269 | 0.7397 | 0.833 | 0.8702 | 0.8209 | 0.8178 | 0.80445 | 0.81961 | 0.8132 | 0.81574 |  | 98
| 1 | 6 | 0.7513 | 0.835 | 0.77978 | 0.825 | 0.818 | 0.801 | 0.8 | 0.818 | 0.884 | > 0.5 | 0.79116 | > 0.5 | 0.83005 | 0.8393 | 0.86489 | 0.8428 | > 0.5 | 0.7034 | 0.8259 | 0.7468 | 0.837 | 0.8553 | 0.8211 | 0.8154 | 0.82692 | 0.82331 | 0.8241 | 0.82471
| Mean |  | 0.75122 | 0.82717 | 0.78105 | 0.84383 | 0.837 | 0.809 | 0.80267 | 0.82567 | 0.87617 |  | 0.80921 |  | 0.83234 | 0.84108 | 0.86295 | 0.84315 |  | 0.7133 | 0.82555 | 0.75287 | 0.83605 | 0.86198 | 0.82102 | 0.81547 | 0.82219 | 0.82087 | 0.81043 | 0.82191
| Median |  | 0.75255 | 0.8275 | 0.78349 | 0.8455 | 0.841 | 0.807 | 0.8015 | 0.828 | 0.8775 |  | 0.8084 |  | 0.82946 | 0.83961 | 0.86266 | 0.8425 |  | 0.70991 | 0.8263 | 0.74935 | 0.8367 | 0.8676 | 0.821 | 0.81575 | 0.82595 | 0.82116 | 0.8093 | 0.81831
| Std Dev. |  | 0.0049994 | 0.0083046 | 0.012343 | 0.011957 | 0.014533 | 0.0096333 | 0.0048854 | 0.007607 | 0.028138 |  | 0.019133 |  | 0.0074501 | 0.0095242 | 0.001936 | 0.0061351 |  | 0.011745 | 0.0023416 | 0.012914 | 0.0022976 | 0.012426 | 0.0066958 | 0.0017941 | 0.0093401 | 0.013996 | 0.012428 | 0.012009
| Rel.Std.Dev. |  | 0.6655% | 1.004% | 1.58% | 1.417% | 1.736% | 1.191% | 0.6086% | 0.9213% | 3.212% |  | 2.364% |  | 0.8951% | 1.132% | 0.2243% | 0.7276% |  | 1.647% | 0.2836% | 1.715% | 0.2748% | 1.442% | 0.8155% | 0.22% | 1.136% | 1.705% | 1.534% | 1.461%
| PDM3 |  | -9.077% | 0.115% | -5.467% | 2.132% | 1.305% | -2.084% | -2.85% | -0.06652% | 6.046% |  | -2.058% |  | 0.7408% | 1.799% | 4.446% | 2.05% |  | -13.67% | -0.08064% | -8.878% | 1.19% | 4.329% | -0.6293% | -1.301% | -0.4875% | -0.6467% | -1.91% | -0.5214%
| Z-Score (Absolute) |  | 3.45 | 0.24 | 2 | 1.05 | 0.72 | 0.64 | 0.95 | 0.17 | 2.63 | NA | 0.63 | NA | 0.49 | 0.92 | 1.98 | 1.02 | NA | 5.3 | 0.16 | 3.37 | 0.67 | 1.94 | 0.06 | 0.33 | 0 | 0.06 | 0.57 | 0.01
| Analytical results for Pd in OREAS 625 (Indicative Value 11.5 ppb) | [30 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.15 |  | ppb
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*OES/MS |  | 1
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 01g |  | 1
| 1 | 1 | 1.01: 12; 1.02: < 10; 1.03: < 10 |  | 1
| 1 | 2 | 1.01: 10; 1.02: < 10; 1.03: < 10 |  | 1
| 1 | 3 | 1.01: 11; 1.02: < 10; 1.03: < 10 |  | 16
| 1 | 4 | 1.01: 14; 1.02: < 10; 1.03: < 10 |  | 11.5
| 1 | 5 | 1.01: 11; 1.02: < 10; 1.03: < 10 |  | 15
| 1 | 6 | 1.01: 11; 1.02: < 10; 1.03: < 10
| Mean |  | 1.01: 11.5
| Median |  | 1.01: 11
| Std Dev. |  | 1.01: 1.3784
| Rel.Std.Dev. |  | 1.01: 11.99%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Pr in OREAS 625 (Indicative Value 4.52 ppm) | [30 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  | 2
| 1 | 1 | 1.01: 3.5; 1.02: 3.966; 1.03: 4.9214; 1.04: 5.7233 |  | 1
| 1 | 2 | 1.01: 3.3; 1.02: 4.363; 1.03: 4.9223; 1.04: 5.2744 |  | 10
| 1 | 3 | 1.01: 3.4; 1.02: 4.13; 1.03: 4.9215; 1.04: 5.7152 |  | 16
| 1 | 4 | 1.01: 3.5; 1.02: 4.328; 1.03: 4.8964; 1.04: 5.4428 |  | 4.5233
| 1 | 5 | 1.01: 3.6; 1.02: 4.209; 1.03: 4.8626; 1.04: 5.634 |  | 16
| 1 | 6 | 1.01: 3.5; 1.02: 4.007; 1.03: 4.8907; 1.04: 5.5507
| Mean |  | 1.01: 3.4667; 1.02: 4.1672; 1.03: 4.9025; 1.04: 5.5567
| Median |  | 1.01: 3.5; 1.02: 4.1695; 1.03: 4.9089; 1.04: 5.5923
| Std Dev. |  | 1.01: 0.10328; 1.02: 0.16338; 1.03: 0.023998; 1.04: 0.17411
| Rel.Std.Dev. |  | 1.01: 2.979%; 1.02: 3.921%; 1.03: 0.4895%; 1.04: 3.133%
| PDM3 |  | 1.01: -23.36%; 1.02: -7.872%; 1.03: 8.384%; 1.04: 22.85%
| Z-Score (Absolute) |  | 1.01: 1.04; 1.02: 0.36; 1.03: 0.36; 1.04: 0.99
| Analytical results for Pt in OREAS 625 (Indicative Value < 5 ppb) | [30 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.15 |  | ppb
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*OES/MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 01g |  | 2
| 1 | 1 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 1
| 1 | 2 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 1
| 1 | 3 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 16
| 1 | 4 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | < 5
| 1 | 5 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 12
| 1 | 6 | 1.01: < 10; 1.02: < 5; 1.03: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Rb in OREAS 625 (Certified Value 14.3 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [12 empty] | 1
| 1 | 1 | 12.8 | 14.6 | 15.3 | 12.9 | 13.5 | 13 | 11.7 | 14.79 | 16.0345 | 16.3581 | 20.6 | 12.5 | 16.4 | 14.3 | 13.5 | 15.2 | 16 | [12 empty] | 1
| 1 | 2 | 12.6 | 14.7 | 15.5 | 12.4 | 14 | 13.3 | 12.5 | 15.89 | 15.7956 | 16.972 | 20.6 | 12.2 | 14.9 | 13.7 | 14 | 15.3 | 16 | [12 empty] | 43
| 1 | 3 | 12.4 | 14.5 | 15.8 | 12.8 | 13.2 | 13.2 | 12.3 | 15.14 | 15.5117 | 16.4195 | 20.2 | 12.6 | 16.3 | 13.6 | 13.9 | 14.9 | 15.6 | [12 empty] | 16
| 1 | 4 | 12.7 | 15.1 | 15.3 | 12.9 | 13.3 | 13 | 11.5 | 15.87 | 15.7551 | 17.1185 | 20.8 | 12.4 | 15.2 | 13.3 | 14 | 15.5 | 15.8 | [12 empty] | 14.2798
| 1 | 5 | 12.9 | 14.7 | 15.3 | 12.7 | 13.6 | 13 | 12 | 15.5 | 15.5905 | 16.6511 | 21.8 | 12.9 | 15.6 | 13.5 | 13.6 | 15.8 | 15.4 | [12 empty] | 99
| 1 | 6 | 12.8 | 14.2 | 16.6 | 12.2 | 13.4 | 12.7 | 12.5 | 14.74 | 15.2904 | 16.6938 | 20.9 | 12.1 | 15.5 | 13.8 | 13.5 | 15.4 | 16.3
| Mean |  | 12.7 | 14.6333 | 15.6333 | 12.65 | 13.5 | 13.0333 | 12.0833 | 15.3217 | 15.663 | 16.7022 | 20.8167 | 12.45 | 15.65 | 13.7 | 13.75 | 15.35 | 15.85
| Median |  | 12.75 | 14.65 | 15.4 | 12.75 | 13.45 | 13 | 12.15 | 15.32 | 15.6728 | 16.6724 | 20.7 | 12.45 | 15.55 | 13.65 | 13.75 | 15.35 | 15.9
| Std Dev. |  | 0.17889 | 0.29439 | 0.51251 | 0.2881 | 0.28284 | 0.20656 | 0.42151 | 0.51176 | 0.25731 | 0.29903 | 0.53821 | 0.2881 | 0.59582 | 0.34059 | 0.2429 | 0.30166 | 0.32094
| Rel.Std.Dev. |  | 1.409% | 2.012% | 3.278% | 2.277% | 2.095% | 1.585% | 3.488% | 3.34% | 1.643% | 1.79% | 2.585% | 2.314% | 3.807% | 2.486% | 1.767% | 1.965% | 2.025%
| PDM3 |  | -11.06% | 2.476% | 9.479% | -11.41% | -5.461% | -8.729% | -15.38% | 7.296% | 9.686% | 16.96% | 45.78% | -12.81% | 9.595% | -4.06% | -3.71% | 7.495% | 11%
| Z-Score (Absolute) |  | 1.15 | 0 | 0.59 | 1.18 | 0.67 | 0.95 | 1.52 | 0.41 | 0.61 | 1.23 | 3.68 | 1.3 | 0.6 | 0.56 | 0.53 | 0.43 | 0.72
| Analytical results for Re in OREAS 625 (Certified Value 0.004 ppm) | [30 empty] | Certified
Re | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.04; 1.03: 1.05; 1.04: 1.06; 1.05: 1.07; 1.06: 1.08; 1.07: 1.09; 1.08: 1.10; 1.09: 1.13; 1.10: 1.15; 1.11: 1.17; 1.12: 1.18; 1.13: 1.25; 1.14: 1.26; 1.15: 1.27; 1.16: 1.30 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.09: AR*MS; 1.10: AR*OES/MS; 1.11: AR*OES/MS; 1.12: AR*OES/MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS; 1.15: AR*OES/MS; 1.16: AR*MS |  | 3
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 25g; 1.09: 0.2g; 1.10: 01g; 1.11: 01g; 1.12: 0.25g; 1.13: 0.25g; 1.14: 0.25g; 1.15: 0.5g; 1.16: 0.2g |  | 3
| 1 | 1 | 1.01: 0.003; 1.02: 0.003; 1.03: 0.004; 1.04: 0.005; 1.05: 0.004; 1.06: 0.003; 1.07: 0.004; 1.08: 0.004; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.004; 1.15: 0.004; 1.16: 0.003 |  | 1
| 1 | 2 | 1.01: 0.003; 1.02: 0.004; 1.03: 0.004; 1.04: 0.005; 1.05: 0.005; 1.06: 0.005; 1.07: 0.004; 1.08: 0.006; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.005; 1.15: 0.004; 1.16: 0.005 |  | 15
| 1 | 3 | 1.01: 0.004; 1.02: 0.005; 1.03: 0.004; 1.04: 0.005; 1.05: 0.004; 1.06: 0.004; 1.07: 0.005; 1.08: 0.004; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.005; 1.15: 0.004; 1.16: 0.004 |  | 16
| 1 | 4 | 1.01: 0.004; 1.02: 0.004; 1.03: 0.004; 1.04: 0.004; 1.05: 0.004; 1.06: 0.003; 1.07: 0.005; 1.08: 0.004; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.004; 1.15: 0.004; 1.16: 0.004 |  | 0.0040833
| 1 | 5 | 1.01: 0.004; 1.02: 0.004; 1.03: 0.004; 1.04: 0.004; 1.05: 0.004; 1.06: 0.005; 1.07: 0.004; 1.08: 0.003; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.005; 1.15: 0.004; 1.16: 0.004 |  | 100
| 1 | 6 | 1.01: 0.004; 1.02: 0.004; 1.03: 0.004; 1.04: 0.003; 1.05: 0.005; 1.06: 0.004; 1.07: 0.004; 1.08: 0.003; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.005; 1.15: 0.004; 1.16: 0.005
| Mean |  | 1.01: 0.0036667; 1.02: 0.004; 1.03: 0.004; 1.04: 0.0043333; 1.05: 0.0043333; 1.06: 0.004; 1.07: 0.0043333; 1.08: 0.004; 1.14: 0.0046667; 1.15: 0.004; 1.16: 0.0041667
| Median |  | 1.01: 0.004; 1.02: 0.004; 1.03: 0.004; 1.04: 0.0045; 1.05: 0.004; 1.06: 0.004; 1.07: 0.004; 1.08: 0.004; 1.14: 0.005; 1.15: 0.004; 1.16: 0.004
| Std Dev. |  | 1.01: 0.0005164; 1.02: 0.00063246; 1.03: 0; 1.04: 0.0008165; 1.05: 0.0005164; 1.06: 0.00089443; 1.07: 0.0005164; 1.08: 0.0010954; 1.14: 0.0005164; 1.15: 0; 1.16: 0.00075277
| Rel.Std.Dev. |  | 1.01: 14.08%; 1.02: 15.81%; 1.03: 0%; 1.04: 18.84%; 1.05: 11.92%; 1.06: 22.36%; 1.07: 11.92%; 1.08: 27.39%; 1.14: 11.07%; 1.15: 0%; 1.16: 18.07%
| PDM3 |  | 1.01: -10.2%; 1.02: -2.041%; 1.03: -2.041%; 1.04: 6.122%; 1.05: 6.122%; 1.06: -2.041%; 1.07: 6.122%; 1.08: -2.041%; 1.14: 14.29%; 1.15: -2.041%; 1.16: 2.041%
| Z-Score (Absolute) |  | 1.01: 1.57; 1.02: 0.67; 1.03: 0.67; 1.04: 0.22; 1.05: 0.22; 1.06: 0.67; 1.07: 0.22; 1.08: 0.67; 1.09: 2.02; 1.10: 55.97; 1.11: 55.97; 1.12: 4.72; 1.13: 15.51; 1.14: 1.12; 1.15: 0.67; 1.16: 0.22
| Analytical results for S in OREAS 625 (Certified Value 3.77 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 3
| 1 | 1 | 4 | 3.917 | 3.79 | 3.81 | 3.91 | 3.87 | 3.67 | 3.8 | 3.34 | 3.5982 | 3.8 | 3.4396 | 3.9903 | 3.9121 | 3.29 | 3.74 | 3.47 | 3.7419 | 3.8 | 3.8 | 3.36 | 3.97 | 3.79 | 3.9 | 3.99 | 3.27 |  |  |  | 1
| 1 | 2 | 3.86 | 4.098 | 3.76 | 3.89 | 3.98 | 3.81 | 3.59 | 3.89 | 3.59 | 3.6431 | 3.71 | 3.5302 | 3.8238 | 3.9354 | 3.29 | 3.49 | 3.42 | 3.8421 | 3.74 | 3.78 | 3.38 | 3.92 | 3.76 | 3.97 | 3.86 | 3.24 |  |  |  | 26
| 1 | 3 | 3.93 | 4.036 | 3.75 | 3.83 | 4.02 | 3.86 | 3.6 | 3.83 | 3.5 | 3.6929 | 3.96 | 3.2837 | 3.8619 | 3.9342 | 3.12 | 3.67 | 3.58 | 3.7943 | 3.79 | 3.85 | 3.45 | 3.95 | 3.7 | 3.92 | 3.83 | 3.16 |  |  |  | 16
| 1 | 4 | 3.93 | 4.086 | 3.68 | 3.92 | 3.97 | 3.83 | 3.59 | 3.84 | 3.6 | 3.7314 | 3.84 | 3.4668 | 3.87 | 3.9457 | 3.16 | 3.65 | 3.47 | 3.8311 | 3.8 | 3.72 | 3.49 | 3.92 | 3.78 | 3.95 | 3.9 | 3.41 |  |  |  | 3.7662
| 1 | 5 | 3.99 | 4.044 | 3.79 | 3.79 | 3.95 | 3.8 | 3.61 | 3.85 | 3.36 | 3.7468 | 3.78 | 3.3779 | 3.8422 | 3.7872 | 3.18 | 3.68 | 3.53 | 3.7515 | 3.81 | 3.78 | 3.39 | 3.99 | 3.65 | 3.95 | 3.93 | 3.09 |  |  |  | 101
| 1 | 6 | 3.92 | 4.088 | 3.79 | 3.73 | 4 | 3.77 | 3.56 | 3.79 | 3.54 | 3.6232 | 3.72 | 3.4042 | 3.844 | 3.7608 | 3.21 | 3.67 | 3.49 | 3.7791 | 3.7 | 3.77 | 3.31 | 3.94 | 3.67 | 3.99 | 3.98 | 3.39
| Mean |  | 3.9383 | 4.0448 | 3.76 | 3.8283 | 3.9717 | 3.8233 | 3.6033 | 3.8333 | 3.4883 | 3.6726 | 3.8017 | 3.4171 | 3.872 | 3.8792 | 3.2083 | 3.65 | 3.4933 | 3.79 | 3.7733 | 3.7833 | 3.3967 | 3.9483 | 3.725 | 3.9467 | 3.915 | 3.26
| Median |  | 3.93 | 4.065 | 3.775 | 3.82 | 3.975 | 3.82 | 3.595 | 3.835 | 3.52 | 3.668 | 3.79 | 3.4219 | 3.853 | 3.9231 | 3.195 | 3.67 | 3.48 | 3.7867 | 3.795 | 3.78 | 3.385 | 3.945 | 3.73 | 3.95 | 3.915 | 3.255
| Std Dev. |  | 0.051153 | 0.067535 | 0.042895 | 0.068823 | 0.038687 | 0.037771 | 0.036697 | 0.036148 | 0.11321 | 0.060367 | 0.091742 | 0.083984 | 0.060158 | 0.082665 | 0.06969 | 0.084143 | 0.055377 | 0.040833 | 0.043665 | 0.042269 | 0.064395 | 0.027869 | 0.059582 | 0.03266 | 0.064109 | 0.12554
| Rel.Std.Dev. |  | 1.299% | 1.67% | 1.141% | 1.798% | 0.9741% | 0.9879% | 1.018% | 0.943% | 3.245% | 1.644% | 2.413% | 2.458% | 1.554% | 2.131% | 2.172% | 2.305% | 1.585% | 1.077% | 1.157% | 1.117% | 1.896% | 0.7058% | 1.6% | 0.8275% | 1.638% | 3.851%
| PDM3 |  | 4.569% | 7.397% | -0.1656% | 1.649% | 5.455% | 1.516% | -4.325% | 1.782% | -7.379% | -2.486% | 0.9407% | -9.271% | 2.809% | 3% | -14.81% | -3.086% | -7.246% | 0.6309% | 0.1884% | 0.4539% | -9.813% | 4.835% | -1.095% | 4.791% | 3.95% | -13.44%
| Z-Score (Absolute) |  | 0.77 | 1.31 | 0.14 | 0.21 | 0.94 | 0.19 | 0.93 | 0.24 | 1.52 | 0.58 | 0.08 | 1.88 | 0.43 | 0.47 | 2.94 | 0.7 | 1.49 | 0.02 | 0.07 | 0.02 | 1.98 | 0.82 | 0.31 | 0.81 | 0.65 | 2.68
| Analytical results for Sb in OREAS 625 (Certified Value 28.5 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 2
| 1 | 1 | 28 | 20.3 | 32.2 | 26.4 | 28.6 | 28.3 | 28.8 | 28.7 | 30.7 | 32.19 | 27.74 | 26 | 27.6637 | 27.004 | 26.8299 | 31.5 | 28.5 | 22.86 | 27.43 | 29.35 | 28.27 | 29.16 | 31 | 30.3 | 22.66 | 29.42 | 32 | 29.75 |  | 1
| 1 | 2 | 27 | 18.2 | 32 | 27.2 | 29.7 | 28.2 | 29.6 | 28.5 | 32.6 | 31.92 | 27.01 | 27 | 27.2018 | 28.589 | 27.0092 | 33.8 | 26.5 | 23.37 | 25.17 | 29.63 | 29.25 | 27.21 | 33 | 30.7 | 23.79 | 28.89 | 27 | 29.18 |  | 45
| 1 | 3 | 27 | 20.6 | 31.8 | 26 | 30.6 | 28.8 | 28.4 | 29.7 | 32 | 32.53 | 30.01 | 24 | 26.8987 | 25.288 | 26.9364 | 31.7 | 28.6 | 22.69 | 26.53 | 29.15 | 29.6 | 28.34 | 34 | 30.5 | 23.2 | 28.8 | 32 | 30.46 |  | 16
| 1 | 4 | 28 | 17.8 | 30.1 | 26.9 | 29.6 | 29 | 28.2 | 28.9 | 31.4 | 31.53 | 27.97 | 27 | 27.0575 | 22.477 | 26.7964 | 33 | 27.5 | 22.87 | 26.84 | 29.17 | 27.53 | 27.02 | 31 | 29.4 | 23.32 | 29.96 | 26 | 29.14 |  | 28.5046
| 1 | 5 | 28 | 20.9 | 31.4 | 26.8 | 29.5 | 28.8 | 29.4 | 29.1 | 31.2 | 32.94 | 27.61 | 26 | 26.7972 | 25.628 | 27.0539 | 32.6 | 29.6 | 22.95 | 26.73 | 29.91 | 28.77 | 27.81 | 30 | 30.1 | 22.56 | 28.5 | 30 | 29.72 |  | 102
| 1 | 6 | 27 | 20.4 | 31.7 | 26.8 | 32.8 | 27.6 | 28.6 | 29.4 | 32.4 | 30.75 | 28.03 | 26 | 27.4052 | 25.102 | 26.9737 | 32.3 | 29.3 | 22.96 | 26.44 | 29.3 | 29.35 | 27.23 | 31 | 31.2 | 22.98 | 29.03 | 33 | 30.24
| Mean |  | 27.5 | 19.7 | 31.5333 | 26.6833 | 30.1333 | 28.45 | 28.8333 | 29.05 | 31.7167 | 31.9767 | 28.0617 | 26 | 27.1707 | 25.6813 | 26.9332 | 32.4833 | 28.3333 | 22.95 | 26.5233 | 29.4183 | 28.795 | 27.795 | 31.6667 | 30.3667 | 23.085 | 29.1 | 30 | 29.7483
| Median |  | 27.5 | 20.35 | 31.75 | 26.8 | 29.65 | 28.55 | 28.7 | 29 | 31.7 | 32.055 | 27.855 | 26 | 27.1296 | 25.458 | 26.955 | 32.45 | 28.55 | 22.91 | 26.63 | 29.325 | 29.01 | 27.52 | 31 | 30.4 | 23.09 | 28.96 | 31 | 29.735
| Std Dev. |  | 0.54772 | 1.3387 | 0.75277 | 0.42151 | 1.4528 | 0.52058 | 0.55737 | 0.44609 | 0.73869 | 0.77293 | 1.0217 | 1.0954 | 0.32412 | 2.0477 | 0.10136 | 0.85186 | 1.157 | 0.22742 | 0.7488 | 0.29627 | 0.77969 | 0.82735 | 1.5055 | 0.60553 | 0.45448 | 0.51788 | 2.8983 | 0.53704
| Rel.Std.Dev. |  | 1.992% | 6.795% | 2.387% | 1.58% | 4.821% | 1.83% | 1.933% | 1.536% | 2.329% | 2.417% | 3.641% | 4.213% | 1.193% | 7.973% | 0.3763% | 2.622% | 4.084% | 0.9909% | 2.823% | 1.007% | 2.708% | 2.977% | 4.754% | 1.994% | 1.969% | 1.78% | 9.661% | 1.805%
| PDM3 |  | -3.524% | -30.89% | 10.63% | -6.389% | 5.714% | -0.1914% | 1.153% | 1.914% | 11.27% | 12.18% | -1.554% | -8.786% | -4.68% | -9.904% | -5.513% | 13.96% | -0.6007% | -19.49% | -6.951% | 3.206% | 1.019% | -2.489% | 11.09% | 6.533% | -19.01% | 2.089% | 5.246% | 4.363%
| Z-Score (Absolute) |  | 0.47 | 3.76 | 1.23 | 0.82 | 0.64 | 0.07 | 0.09 | 0.18 | 1.3 | 1.41 | 0.24 | 1.11 | 0.61 | 1.24 | 0.71 | 1.63 | 0.12 | 2.39 | 0.88 | 0.34 | 0.07 | 0.35 | 1.28 | 0.74 | 2.33 | 0.2 | 0.58 | 0.47
| Analytical results for Sc in OREAS 625 (Certified Value 1.12 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  |  | 2
| 1 | 1 | < 5 | 1.2 | 1 | 1.2 | 1.3 | 1.1 | 1 | 1.3 | 0.9 | 0.9 | 0.9 | 1 | 1.2004 | 1.2197 | < 1 | 1 | 0.5 | 1.6 | 1.4 | < 5 | 1.2 | 0.9 | 1.1 | 1.3 | 1.1 |  |  |  |  | 1
| 1 | 2 | < 5 | 1.1 | 0.9 | 1.2 | 1.4 | 1.1 | 1.1 | 1.3 | 1 | 1 | 0.9 | 1 | 1.2016 | 1.273 | < 1 | 1 | 0.5 | 1.9 | 1.3 | < 5 | 1.2 | 0.9 | 1.1 | 1.3 | 1.1 |  |  |  |  | 16
| 1 | 3 | < 5 | 1 | 1 | 1.1 | 1.4 | 1.1 | 1 | 1.4 | 1 | 0.9 | 0.9 | 1 | 1.1976 | 1.2717 | 1.1 | 1 | 0.5 | 1.7 | 1.3 | < 5 | 1.2 | 0.9 | 1.1 | 1.3 | 1 |  |  |  |  | 16
| 1 | 4 | < 5 | 1.1 | 0.9 | 1.2 | 1.4 | 1.1 | 1 | 1.3 | 0.8 | 1 | 0.9 | 1 | 1.2126 | 1.2839 | < 1 | 1 | 0.5 | 1.7 | 1.2 | < 5 | 1.1 | 0.9 | 1.2 | 1.3 | 1.1 |  |  |  |  | 1.1189
| 1 | 5 | < 5 | 1.3 | 1 | 1.1 | 1.4 | 1.1 | 1 | 1.3 | 1 | 1 | 0.9 | 1 | 1.1856 | 1.2033 | < 1 | 1 | 0.5 | 1.8 | 1.3 | < 5 | 1.2 | 0.9 | 1.1 | 1.3 | 1 |  |  |  |  | 103
| 1 | 6 | < 5 | 1.2 | 1 | 1.1 | 1.4 | 1.1 | 1 | 1.3 | 1 | 0.9 | 0.9 | 1 | 1.1863 | 1.2229 | 1 | 1 | 0.5 | 1.8 | 1.3 | < 5 | 1.3 | 0.9 | 1.1 | 1.3 | 1.1
| Mean |  |  | 1.15 | 0.96667 | 1.15 | 1.3833 | 1.1 | 1.0167 | 1.3167 | 0.95 | 0.95 | 0.9 | 1 | 1.1973 | 1.2457 | 1.05 | 1 | 0.5 | 1.75 | 1.3 |  | 1.2 | 0.9 | 1.1167 | 1.3 | 1.0667
| Median |  |  | 1.15 | 1 | 1.15 | 1.4 | 1.1 | 1 | 1.3 | 1 | 0.95 | 0.9 | 1 | 1.199 | 1.2473 | 1.05 | 1 | 0.5 | 1.75 | 1.3 |  | 1.2 | 0.9 | 1.1 | 1.3 | 1.1
| Std Dev. |  |  | 0.10488 | 0.05164 | 0.054772 | 0.040825 | 2.4324e-16 | 0.040825 | 0.040825 | 0.083666 | 0.054772 | 0 | 0 | 0.010191 | 0.034262 | 0.070711 | 0 | 0 | 0.10488 | 0.063246 |  | 0.063246 | 0 | 0.040825 | 0 | 0.05164
| Rel.Std.Dev. |  |  | 9.12% | 5.342% | 4.763% | 2.951% | 2.211e-14% | 4.016% | 3.101% | 8.807% | 5.766% | 0% | 0% | 0.8511% | 2.75% | 6.734% | 0% | 0% | 5.993% | 4.865% |  | 5.27% | 0% | 3.656% | 0% | 4.841%
| PDM3 |  |  | 2.776% | -13.61% | 2.776% | 23.63% | -1.692% | -9.14% | 17.67% | -15.1% | -15.1% | -19.57% | -10.63% | 7.008% | 11.33% | -6.161% | -10.63% | -55.31% | 56.4% | 16.18% |  | 7.245% | -19.57% | -0.2026% | 16.18% | -4.671%
| Z-Score (Absolute) |  | 5.46 | 0 | 0.74 | 0 | 0.94 | 0.2 | 0.54 | 0.67 | 0.81 | 0.81 | 1.01 | NA | 0.19 | 0.39 | 1.89 | NA | 2.63 | 2.43 | 0.61 | 5.46 | 0.2 | 1.01 | 0.13 | 0.61 | 0.34
Results from laboratories 1.12 and 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 625 (Indicative Value 1.55 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [5 empty] | 2
| 1 | 1 | < 5 | 3 | 1 | 3 | 1.8 | 0.9 | 1.6 | 3.7 | 0.5 | < 1 | 0.9 | < 1.5 | 2.56 | 1.3459 | < 1 | 1.2 | < 10 | 0.52 | 1.1 | < 1 | 1 | 2 | 0.5 | 2.4 | [5 empty] | 1
| 1 | 2 | < 5 | 2.6 | 1 | 3 | 1.6 | 1.4 | 1.2 | 4 | 0.6 | < 1 | 0.9 | < 1.5 | 2.52 | 1.3588 | < 1 | 1.3 | < 10 | 0.54 | 1.4 | 2 | 2 | 2 | 0.5 | 2.5 | [5 empty] | 3
| 1 | 3 | < 5 | 2.7 | 1 | 2.9 | 1.7 | 1.1 | 1.4 | 4 | 0.6 | < 1 | 1 | < 1.5 | 2.86 | 1.3414 | < 1 | 1.3 | < 10 | 0.54 | 1.9 | 2 | 2 | 2 | 0.5 | 2.8 | [5 empty] | 16
| 1 | 4 | < 5 | 2.7 | 1 | 3 | 1.8 | 1.1 | 1.2 | 3.6 | 0.6 | < 1 | 0.9 | < 1.5 | 2.47 | 1.3298 | < 1 | 1.1 | < 10 | 0.52 | 1.6 | 1 | 3 | 2 | 0.5 | 2.5 | [5 empty] | 1.5503
| 1 | 5 | < 5 | 3 | 0.8 | 2.7 | 1.7 | 1.3 | 1.4 | 3.9 | 0.5 | < 1 | 0.9 | < 1.5 | 2.38 | 1.302 | < 1 | 1.2 | < 10 | 0.5 | 1.6 | 1 | 2 | 2 | 0.5 | 2.5 | [5 empty] | 13
| 1 | 6 | < 5 | 2.6 | 0.8 | 2.8 | 1.5 | 1.2 | 1.4 | 3.8 | 0.6 | < 1 | 1 | < 1.5 | 2.97 | 1.2979 | < 1 | 1.2 | < 10 | 0.52 | 1.6 | 1 | 2 | 2 | 0.5 | 2.4
| Mean |  |  | 2.7667 | 0.93333 | 2.9 | 1.6833 | 1.1667 | 1.3667 | 3.8333 | 0.56667 |  | 0.93333 |  | 2.6267 | 1.3293 |  | 1.2167 |  | 0.52333 | 1.5333 | 1.4 | 2 | 2 | 0.5 | 2.5167
| Median |  |  | 2.7 | 1 | 2.95 | 1.7 | 1.15 | 1.4 | 3.85 | 0.6 |  | 0.9 |  | 2.54 | 1.3356 |  | 1.2 |  | 0.52 | 1.6 | 1 | 2 | 2 | 0.5 | 2.5
| Std Dev. |  |  | 0.18619 | 0.10328 | 0.12649 | 0.1169 | 0.17512 | 0.15055 | 0.1633 | 0.05164 |  | 0.05164 |  | 0.23389 | 0.024581 |  | 0.075277 |  | 0.015055 | 0.26583 | 0.54772 | 0.63246 | 0 | 0 | 0.1472
| Rel.Std.Dev. |  |  | 6.73% | 11.07% | 4.362% | 6.945% | 15.01% | 11.02% | 4.26% | 9.113% |  | 5.533% |  | 8.905% | 1.849% |  | 6.187% |  | 2.877% | 17.34% | 39.12% | 31.62% | 0% | 0% | 5.849%
| PDM3 |  |  | 78.46% | -39.8% | 87.06% | 8.579% | -24.75% | -11.85% | 147.3% | -63.45% |  | -39.8% |  | 69.43% | -14.26% |  | -21.52% |  | -66.24% | -1.096% | -9.697% | 29% | 29% | -67.75% | 62.33%
| Z-Score (Absolute) |  | 1.08 | 1.33 | 0.39 | 1.46 | 0.32 | 0.17 | 0.02 | 2.34 | 0.74 | 0.8 | 0.39 | 0.56 | 1.2 | 0.02 | 0.8 | 0.12 | 3.44 | 0.78 | 0.17 | 0.09 | 0.61 | 0.61 | 0.8 | 1.1
| Analytical results for Sm in OREAS 625 (Indicative Value 3.3 ppm) | [30 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | [5 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | [5 empty] | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | [5 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | [5 empty] | 2
| 1 | 1 | 1.01: 2.3; 1.02: 2.952; 1.03: 3.5402; 1.04: 3.8646 | [5 empty] | 1
| 1 | 2 | 1.01: 2.9; 1.02: 3.172; 1.03: 3.4968; 1.04: 3.6461 | [5 empty] | 8
| 1 | 3 | 1.01: 2.7; 1.02: 3.149; 1.03: 3.49; 1.04: 3.8649 | [5 empty] | 16
| 1 | 4 | 1.01: 2.8; 1.02: 3.214; 1.03: 3.4676; 1.04: 3.6881 | [5 empty] | 3.3006
| 1 | 5 | 1.01: 3.3; 1.02: 3.065; 1.03: 3.3843; 1.04: 3.9052 | [5 empty] | 14
| 1 | 6 | 1.01: 3; 1.02: 3.034; 1.03: 3.4971; 1.04: 3.7847
| Mean |  | 1.01: 2.8333; 1.02: 3.0977; 1.03: 3.4793; 1.04: 3.7923
| Median |  | 1.01: 2.85; 1.02: 3.107; 1.03: 3.4934; 1.04: 3.8247
| Std Dev. |  | 1.01: 0.33267; 1.02: 0.098014; 1.03: 0.052177; 1.04: 0.10538
| Rel.Std.Dev. |  | 1.01: 11.74%; 1.02: 3.164%; 1.03: 1.5%; 1.04: 2.779%
| PDM3 |  | 1.01: -14.16%; 1.02: -6.15%; 1.03: 5.413%; 1.04: 14.89%
| Z-Score (Absolute) |  | 1.01: 0.95; 1.02: 0.4; 1.03: 0.4; 1.04: 1.05
| Analytical results for Sn in OREAS 625 (Certified Value 0.95 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [7 empty] | 2
| 1 | 1 | < 10 | 1.1 | 0.9 | 0.9 | 1 | 0.8 | 0.9 | 0.7 | 0.93 | 1.0569 | 0.97682 | < 5 | 1.3 | 1 | < 10 | 1.4 | 1.2 | 0.7 | 0.8 | 0.9 | < 10 | 0.7 | [7 empty] | 1
| 1 | 2 | < 10 | 1.04 | 0.9 | 0.9 | 1.1 | 0.8 | 0.9 | 0.8 | 1.09 | 1.0707 | 1.0164 | < 5 | 1.3 | 1 | < 10 | 1.5 | 1.1 | 0.7 | 0.8 | 0.9 | < 10 | 0.7 | [7 empty] | 13
| 1 | 3 | < 10 | 1.06 | 0.9 | 0.9 | 1 | 0.8 | 0.9 | 0.8 | 1.1 | 1.0313 | 1.0691 | < 5 | 1.4 | 1.1 | < 10 | 1.6 | 1.1 | 0.7 | 0.9 | 0.9 | < 10 | 0.7 | [7 empty] | 16
| 1 | 4 | < 10 | 1.07 | 0.9 | 0.9 | 1.1 | 0.8 | 0.9 | 0.7 | 1.04 | 1.0402 | 0.98416 | < 5 | 1.4 | 1 | < 10 | 1.4 | 1.1 | 0.7 | 0.8 | 1 | < 10 | 0.7 | [7 empty] | 0.94654
| 1 | 5 | < 10 | 1.1 | 0.9 | 0.9 | 1.1 | 0.8 | 0.8 | 0.7 | 1.06 | 0.98272 | 1.066 | < 5 | 1.2 | 1 | < 10 | 1.4 | 1.2 | 0.7 | 0.8 | 0.9 | < 10 | 0.7 | [7 empty] | 104
| 1 | 6 | < 10 | 1.13 | 0.9 | 1 | 1 | 0.8 | 0.8 | 0.8 | 0.95 | 1.0485 | 1.0674 | < 5 | 1.3 | 1 | < 10 | 1.4 | 1.1 | 0.7 | 0.9 | 0.9 | < 10 | 0.7
| Mean |  |  | 1.0833 | 0.9 | 0.91667 | 1.05 | 0.8 | 0.86667 | 0.75 | 1.0283 | 1.0384 | 1.03 |  | 1.3167 | 1.0167 |  | 1.45 | 1.1333 | 0.7 | 0.83333 | 0.91667 |  | 0.7
| Median |  |  | 1.085 | 0.9 | 0.9 | 1.05 | 0.8 | 0.9 | 0.75 | 1.05 | 1.0444 | 1.0412 |  | 1.3 | 1 |  | 1.4 | 1.1 | 0.7 | 0.8 | 0.9 |  | 0.7
| Std Dev. |  |  | 0.03266 | 0 | 0.040825 | 0.054772 | 1.2162e-16 | 0.05164 | 0.054772 | 0.071949 | 0.030466 | 0.043212 |  | 0.075277 | 0.040825 |  | 0.083666 | 0.05164 | 1.2162e-16 | 0.05164 | 0.040825 |  | 1.2162e-16
| Rel.Std.Dev. |  |  | 3.015% | 0% | 4.454% | 5.216% | 1.52e-14% | 5.958% | 7.303% | 6.997% | 2.934% | 4.195% |  | 5.717% | 4.016% |  | 5.77% | 4.556% | 1.737e-14% | 6.197% | 4.454% |  | 1.737e-14%
| PDM3 |  |  | 14.45% | -4.917% | -3.156% | 10.93% | -15.48% | -8.438% | -20.76% | 8.641% | 9.704% | 8.815% |  | 39.1% | 7.409% |  | 53.19% | 19.73% | -26.05% | -11.96% | -3.156% |  | -26.05%
| Z-Score (Absolute) |  | 14.95 | 0.2 | 0.49 | 0.42 | 0.08 | 0.86 | 0.61 | 1.05 | 0 | 0.03 | 0 | 5.54 | 1.08 | 0.05 | 14.95 | 1.58 | 0.39 | 1.24 | 0.74 | 0.42 | 14.95 | 1.24
| Analytical results for Sr in OREAS 625 (Certified Value 20.7 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 2
| 1 | 1 | 22 | 14.3 | 19.1 | 21.3 | 19.9 | 18.6 | 19.4 | 20.1 | 17.8 | 19.38 | 18.5 | 22 | 20.5092 | 21.66 | 19.3846 | 22.8 | 23.2 | 18 | 20.48 | 20.4 | 19.9 | 24 | 22.9 | 21.6 | 20.1 | 22.4 | 23 |  |  | 1
| 1 | 2 | 20 | 13.3 | 18.9 | 21.2 | 20.1 | 18 | 20.4 | 21 | 19.5 | 19.45 | 18 | 22 | 20.1455 | 21 | 19.4806 | 22.7 | 23 | 18 | 20.28 | 21.5 | 19.7 | 24 | 22.3 | 21.9 | 20.2 | 22.2 | 23 |  |  | 21
| 1 | 3 | 22 | 14.8 | 19.4 | 20.7 | 20.7 | 18.8 | 19.6 | 21.1 | 19.1 | 19.27 | 19.2 | 21 | 19.8865 | 20.13 | 19.8892 | 23.6 | 23.6 | 19 | 20.35 | 21.4 | 20.1 | 25 | 22.1 | 21.9 | 19.5 | 22.5 | 23 |  |  | 16
| 1 | 4 | 21 | 13.5 | 18.8 | 21.6 | 20.2 | 18.5 | 19.5 | 20.8 | 18.6 | 19.53 | 18.5 | 22 | 20.3179 | 20.78 | 19.323 | 22.9 | 23.2 | 18 | 20.41 | 20.6 | 19.5 | 24 | 22.5 | 22.3 | 20.5 | 22 | 23 |  |  | 20.7281
| 1 | 5 | 22 | 14.1 | 19.2 | 20.6 | 20.1 | 18.5 | 19.4 | 20.8 | 18.3 | 19.86 | 18.3 | 22 | 19.8809 | 20.8 | 19.6253 | 22.8 | 23.1 | 19 | 20.44 | 21 | 19.9 | 24 | 22.2 | 21.5 | 19.8 | 22.3 | 23 |  |  | 105
| 1 | 6 | 21 | 13.5 | 19.4 | 20.9 | 20 | 18.4 | 19.8 | 20.5 | 19.3 | 18.94 | 18.3 | 22 | 20.3291 | 22.17 | 19.7983 | 23.2 | 23.1 | 18 | 20.37 | 21.3 | 19.5 | 23 | 22.8 | 21.8 | 20.1 | 22.2 | 23
| Mean |  | 21.3333 | 13.9167 | 19.1333 | 21.05 | 20.1667 | 18.4667 | 19.6833 | 20.7167 | 18.7667 | 19.405 | 18.4667 | 21.8333 | 20.1782 | 21.09 | 19.5835 | 23 | 23.2 | 18.3333 | 20.3883 | 21.0333 | 19.7667 | 24 | 22.4667 | 21.8333 | 20.0333 | 22.2667 | 23
| Median |  | 21.5 | 13.8 | 19.15 | 21.05 | 20.1 | 18.5 | 19.55 | 20.8 | 18.85 | 19.415 | 18.4 | 22 | 20.2317 | 20.9 | 19.5529 | 22.85 | 23.15 | 18 | 20.39 | 21.15 | 19.8 | 24 | 22.4 | 21.85 | 20.1 | 22.25 | 23
| Std Dev. |  | 0.8165 | 0.58109 | 0.25033 | 0.38341 | 0.28048 | 0.26583 | 0.38166 | 0.3656 | 0.65013 | 0.30323 | 0.40332 | 0.40825 | 0.25549 | 0.72128 | 0.22775 | 0.34059 | 0.20976 | 0.5164 | 0.070828 | 0.45019 | 0.24221 | 0.63246 | 0.3266 | 0.28048 | 0.34448 | 0.17512 | 0
| Rel.Std.Dev. |  | 3.827% | 4.176% | 1.308% | 1.821% | 1.391% | 1.44% | 1.939% | 1.765% | 3.464% | 1.563% | 2.184% | 1.87% | 1.266% | 3.42% | 1.163% | 1.481% | 0.9041% | 2.817% | 0.3474% | 2.14% | 1.225% | 2.635% | 1.454% | 1.285% | 1.72% | 0.7865% | 0%
| PDM3 |  | 2.92% | -32.86% | -7.694% | 1.553% | -2.709% | -10.91% | -5.041% | -0.05536% | -9.463% | -6.383% | -10.91% | 5.332% | -2.653% | 1.746% | -5.522% | 10.96% | 11.93% | -11.55% | -1.639% | 1.472% | -4.638% | 15.78% | 8.387% | 5.332% | -3.352% | 7.422% | 10.96%
| Z-Score (Absolute) |  | 0.51 | 3.48 | 0.67 | 0.36 | 0.12 | 1.03 | 0.38 | 0.18 | 0.87 | 0.53 | 1.03 | 0.78 | 0.11 | 0.38 | 0.43 | 1.4 | 1.51 | 1.1 | 0 | 0.35 | 0.33 | 1.94 | 1.12 | 0.78 | 0.19 | 1.01 | 1.4
| Analytical results for Ta in OREAS 625 (Certified Value < 0.01 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 3
| 1 | 1 | < 5 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | 0.05 | < 0.01 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 5 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | [11 empty] | 14
| 1 | 3 | < 5 | < 0.05 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 5 | < 0.05 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | 0.11 | < 0.01 | < 0.01 | [11 empty] | < 0.01
| 1 | 5 | < 5 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | 0.09 | < 0.01 | 0.01 | [11 empty] | 106
| 1 | 6 | < 5 | < 0.05 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | 0.07 | < 0.01 | < 0.01
| Mean |  | 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.07: 0.01; 1.16: 0.08; 1.18: 0.01
| Median |  | 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.07: 0.01; 1.16: 0.08; 1.18: 0.01
| Std Dev. |  | 1.04: 0; 1.05: 0; 1.07: 0; 1.16: 0.02582
| Rel.Std.Dev. |  | 1.04: 0%; 1.05: 0%; 1.07: 0%; 1.16: 32.27%
PDM3
| Z-Score (Absolute) |  | 143.48 | 0.43 | 0.67 | 0.63 | 0.43 | 0.72 | 0.58 | 0.72 | 0.72 | 0.43 | 0.43 | 1.88 | 7.66 | 0.43 | 0.43 | 2.55 | 0.72 | 0.67
| Analytical results for Tb in OREAS 625 (Certified Value 0.38 ppm) | [30 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.21; 1.05: 1.22; 1.06: 1.23; 1.07: 1.25; 1.08: 1.26 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.2g; 1.05: 0.5g; 1.06: 0.25g; 1.07: 0.25g; 1.08: 0.25g |  |  | 2
| 1 | 1 | 1.01: 0.3; 1.02: 0.329; 1.03: 0.39662; 1.04: 0.6; 1.05: 0.42264; 1.06: 0.38; 1.07: 0.49; 1.08: 0.32 |  |  | 1
| 1 | 2 | 1.01: 0.4; 1.02: 0.354; 1.03: 0.39522; 1.04: 0.6; 1.05: 0.38433; 1.06: 0.35; 1.07: 0.47; 1.08: 0.34 |  |  | 36
| 1 | 3 | 1.01: 0.3; 1.02: 0.342; 1.03: 0.39777; 1.04: 0.6; 1.05: 0.4291; 1.06: 0.34; 1.07: 0.48; 1.08: 0.33 |  |  | 16
| 1 | 4 | 1.01: 0.4; 1.02: 0.354; 1.03: 0.40025; 1.04: 0.5; 1.05: 0.40071; 1.06: 0.35; 1.07: 0.43; 1.08: 0.33 |  |  | 0.38144
| 1 | 5 | 1.01: 0.4; 1.02: 0.344; 1.03: 0.39378; 1.04: 0.5; 1.05: 0.41178; 1.06: 0.35; 1.07: 0.47; 1.08: 0.32 |  |  | 107
| 1 | 6 | 1.01: 0.4; 1.02: 0.332; 1.03: 0.3869; 1.04: 0.5; 1.05: 0.40976; 1.06: 0.35; 1.07: 0.46; 1.08: 0.32
| Mean |  | 1.01: 0.36667; 1.02: 0.3425; 1.03: 0.39509; 1.04: 0.55; 1.05: 0.40972; 1.06: 0.35333; 1.07: 0.46667; 1.08: 0.32667
| Median |  | 1.01: 0.4; 1.02: 0.343; 1.03: 0.39592; 1.04: 0.55; 1.05: 0.41077; 1.06: 0.35; 1.07: 0.47; 1.08: 0.325
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.010578; 1.03: 0.0045825; 1.04: 0.054772; 1.05: 0.015951; 1.06: 0.013663; 1.07: 0.020656; 1.08: 0.008165
| Rel.Std.Dev. |  | 1.01: 14.08%; 1.02: 3.089%; 1.03: 1.16%; 1.04: 9.959%; 1.05: 3.893%; 1.06: 3.867%; 1.07: 4.426%; 1.08: 2.499%
| PDM3 |  | 1.01: -3.873%; 1.02: -10.21%; 1.03: 3.578%; 1.04: 44.19%; 1.05: 7.414%; 1.06: -7.369%; 1.07: 22.34%; 1.08: -14.36%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.64; 1.03: 0.42; 1.04: NA; 1.05: 0.71; 1.06: 0.42; 1.07: 1.85; 1.08: 0.95
Results from laboratories 1.02 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 625 (Certified Value 0.05 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 15g | 01g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 3
| 1 | 1 | < 5 | 0.04 | 1.6 | 0.05 | 0.03 | 0.05 | 0.02 | 0.04 | 0.06 | 0.03 | < 2 | < 5 | 0.057979 | < 0.05 | < 0.1 | < 10 | 0.05 | 0.07 | < 0.05 | < 0.05 | 0.07 | 0.06 | 0.03 | [6 empty] | 1
| 1 | 2 | < 5 | 0.04 | 2 | 0.05 | 0.03 | 0.03 | 0.03 | 0.04 | 0.05 | 0.03 | < 2 | < 5 | 0.059758 | < 0.05 | < 0.1 | < 10 | 0.08 | 0.09 | < 0.05 | 0.06 | 0.06 | 0.06 | 0.05 | [6 empty] | 15
| 1 | 3 | < 5 | 0.07 | 1.8 | 0.03 | 0.03 | 0.06 | 0.03 | 0.06 | 0.03 | 0.04 | < 2 | < 5 | 0.06031 | < 0.05 | < 0.1 | < 10 | 0.06 | 0.09 | < 0.05 | < 0.05 | 0.05 | 0.05 | 0.03 | [6 empty] | 16
| 1 | 4 | < 5 | 0.07 | 1 | 0.04 | 0.04 | 0.07 | 0.03 | 0.03 | 0.06 | 0.06 | < 2 | < 5 | 0.050711 | 0.05 | < 0.1 | < 10 | < 0.03 | 0.07 | < 0.05 | < 0.05 | 0.06 | 0.05 | 0.05 | [6 empty] | 0.050324
| 1 | 5 | < 5 | 0.04 | 1.1 | 0.04 | 0.04 | 0.06 | 0.03 | 0.05 | 0.03 | 0.05 | < 2 | < 5 | < 0.05 | < 0.05 | < 0.1 | < 10 | 0.08 | 0.12 | < 0.05 | < 0.05 | 0.06 | 0.05 | 0.05 | [6 empty] | 108
| 1 | 6 | < 5 | 0.05 | 1.2 | 0.05 | 0.06 | 0.05 | 0.02 | 0.03 | 0.08 | 0.04 | < 2 | < 5 | < 0.05 | 0.05 | < 0.1 | < 10 | 0.07 | 0.09 | 0.07 | < 0.05 | 0.05 | 0.05 | 0.03
| Mean |  |  | 0.051667 | 1.45 | 0.043333 | 0.038333 | 0.053333 | 0.026667 | 0.041667 | 0.051667 | 0.041667 |  |  | 0.05719 | 0.05 |  |  | 0.068 | 0.088333 | 0.07 | 0.06 | 0.058333 | 0.053333 | 0.04
| Median |  |  | 0.045 | 1.4 | 0.045 | 0.035 | 0.055 | 0.03 | 0.04 | 0.055 | 0.04 |  |  | 0.058869 | 0.05 |  |  | 0.07 | 0.09 | 0.07 | 0.06 | 0.06 | 0.05 | 0.04
| Std Dev. |  |  | 0.01472 | 0.40866 | 0.008165 | 0.01169 | 0.013663 | 0.005164 | 0.01169 | 0.019408 | 0.01169 |  |  | 0.0044321 | 0 |  |  | 0.013038 | 0.018348 |  |  | 0.0075277 | 0.005164 | 0.010954
| Rel.Std.Dev. |  |  | 28.49% | 28.18% | 18.84% | 30.5% | 25.62% | 19.36% | 28.06% | 37.56% | 28.06% |  |  | 7.75% | 0% |  |  | 19.17% | 20.77% |  |  | 12.9% | 9.682% | 27.39%
| PDM3 |  |  | 2.667% | 2781% | -13.89% | -23.83% | 5.979% | -47.01% | -17.2% | 2.667% | -17.2% |  |  | 13.64% | -0.6445% |  |  | 35.12% | 75.53% | 39.1% | 19.23% | 15.91% | 5.979% | -20.52%
| Z-Score (Absolute) |  | 141.51 | 0 | 80.82 | 0.48 | 0.77 | 0.1 | 1.44 | 0.58 | 0 | 0.58 | 54.81 | 141.51 | 0.3 | 1.06 | 0.1 | 286 | 0.43 | 2.12 | 1.11 | 1.2 | 0.39 | 0.1 | 0.67
| Analytical results for Th in OREAS 625 (Certified Value 7.49 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 01g | 01g | 01g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 3
| 1 | 1 | 6.5 | 7.3 | 7.9 | 7.7 | 7.4 | 7.3 | 7 | 6.8 | 7.3 | 7 | 8.2307 | 8.8648 | 5 | 9.52 | 7.5 | 7.58 | 4.6646 | 8.2 | 5 | 7.6 | 6.3 | 7.8 | 10 | [6 empty] | 1
| 1 | 2 | 6.5 | 7.2 | 7.9 | 7.5 | 7.4 | 7.8 | 7 | 7 | 7.53 | 6.6 | 8.2769 | 8.5464 | < 5 | 9.4 | 7.5 | 7.89 | 4.2439 | 8.2 | 5 | 7.2 | 6.3 | 8.2 | 8.7 | [6 empty] | 38
| 1 | 3 | 6.9 | 7.6 | 7.8 | 7.9 | 7.7 | 7.2 | 7.3 | 7 | 7.42 | 6.9 | 8.1646 | 8.9839 | 5.8 | 9.4 | 7.9 | 7.89 | 4.6704 | 8 | 5 | 7.4 | 6.6 | 7.8 | 9.6 | [6 empty] | 16
| 1 | 4 | 6.7 | 7 | 8 | 7.6 | 7.4 | 7.2 | 7.1 | 7 | 7.59 | 7.1 | 8.3698 | 8.4822 | 5.3 | 9.45 | 7.6 | 7.16 | 4.1947 | 8.2 | 5 | 7 | 6.6 | 8.6 | 10 | [6 empty] | 7.4911
| 1 | 5 | 6.6 | 7.1 | 7.9 | 7.7 | 7.6 | 7.3 | 7 | 6.5 | 7.54 | 7.1 | 8.247 | 8.704 | 5.2 | 9.53 | 7.8 | 7.49 | 4.7099 | 8.3 | 5 | 7.4 | 6.2 | 8.1 | 9.5 | [6 empty] | 109
| 1 | 6 | 6.7 | 7.3 | 7.7 | 8.2 | 7.6 | 7.3 | 7 | 7 | 7.22 | 6.8 | 8.2271 | 8.6805 | 6.1 | 9.5 | 7.6 | 7.63 | 4.4816 | 8.6 | 5 | 7.4 | 6.3 | 8 | 10.4
| Mean |  | 6.65 | 7.25 | 7.8667 | 7.7667 | 7.5167 | 7.35 | 7.0667 | 6.8833 | 7.4333 | 6.9167 | 8.2527 | 8.7103 | 5.48 | 9.4667 | 7.65 | 7.6067 | 4.4942 | 8.25 | 5 | 7.3333 | 6.3833 | 8.0833 | 9.7
| Median |  | 6.65 | 7.25 | 7.9 | 7.7 | 7.5 | 7.3 | 7 | 7 | 7.475 | 6.95 | 8.2388 | 8.6922 | 5.3 | 9.475 | 7.6 | 7.605 | 4.5731 | 8.2 | 5 | 7.4 | 6.3 | 8.05 | 9.8
| Std Dev. |  | 0.15166 | 0.20736 | 0.10328 | 0.25033 | 0.13292 | 0.22583 | 0.12111 | 0.20412 | 0.14747 | 0.19408 | 0.06818 | 0.18896 | 0.45497 | 0.058538 | 0.16432 | 0.27384 | 0.22766 | 0.19748 | 0 | 0.20656 | 0.17224 | 0.29944 | 0.58652
| Rel.Std.Dev. |  | 2.281% | 2.86% | 1.313% | 3.223% | 1.768% | 3.073% | 1.714% | 2.965% | 1.984% | 2.806% | 0.8262% | 2.169% | 8.302% | 0.6184% | 2.148% | 3.6% | 5.066% | 2.394% | 0% | 2.817% | 2.698% | 3.704% | 6.047%
| PDM3 |  | -11.23% | -3.218% | 5.014% | 3.679% | 0.3414% | -1.883% | -5.666% | -8.113% | -0.771% | -7.668% | 10.17% | 16.28% | -26.85% | 26.37% | 2.121% | 1.543% | -40.01% | 10.13% | -33.25% | -2.106% | -14.79% | 7.906% | 29.49%
| Z-Score (Absolute) |  | 0.97 | 0.26 | 0.46 | 0.34 | 0.05 | 0.15 | 0.48 | 0.69 | 0.05 | 0.65 | 0.91 | 1.45 | 2.92 | 2.34 | 0.21 | 0.15 | 3.5 | 0.91 | NA | 0.17 | 1.28 | 0.71 | 2.61
Results from laboratory 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 625 (Certified Value 0.023 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  |  | 3
| 1 | 1 | 0.02 | 0.03 | 0.017 | 0.023 | 0.023 | 0.021 | 0.02 | 0.02 | 0.02 | 0.0202 | 0.018 | 0.0313 | 0.024649 | < 0.1 | 0.038356 | 0.03 | 0.0303 | 0.02 | 0.04 | 0.03 | 0.02 | 0.02 | 0.019 | 0.029 | 0.013 |  |  |  |  | 1
| 1 | 2 | 0.02 | 0.03 | 0.015 | 0.024 | 0.024 | 0.021 | 0.02 | 0.02 | 0.02 | 0.022 | 0.018 | 0.0309 | 0.02409 | < 0.1 | 0.038442 | 0.03 | 0.0302 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.028 | 0.013 |  |  |  |  | 39
| 1 | 3 | 0.03 | 0.03 | 0.018 | 0.024 | 0.024 | 0.02 | 0.019 | 0.021 | 0.02 | 0.0213 | 0.019 | 0.0304 | 0.023921 | < 0.1 | 0.038155 | 0.03 | 0.0303 | 0.021 | 0.04 | 0.03 | 0.02 | 0.02 | 0.02 | 0.028 | 0.012 |  |  |  |  | 16
| 1 | 4 | 0.02 | 0.03 | 0.015 | 0.024 | 0.024 | 0.02 | 0.02 | 0.02 | 0.02 | 0.0224 | 0.019 | 0.031 | 0.024241 | < 0.1 | 0.038091 | 0.03 | 0.0307 | 0.021 | 0.04 | 0.03 | 0.02 | 0.02 | 0.024 | 0.028 | 0.013 |  |  |  |  | 0.02342
| 1 | 5 | 0.03 | 0.03 | 0.018 | 0.023 | 0.024 | 0.02 | 0.02 | 0.02 | 0.02 | 0.0218 | 0.019 | 0.0302 | 0.024054 | < 0.1 | 0.037779 | 0.03 | 0.0314 | 0.021 | 0.04 | 0.03 | 0.02 | 0.02 | 0.022 | 0.029 | 0.012 |  |  |  |  | 110
| 1 | 6 | 0.02 | 0.03 | 0.018 | 0.023 | 0.025 | 0.021 | 0.02 | 0.02 | 0.02 | 0.0205 | 0.018 | 0.0329 | 0.023986 | < 0.1 | 0.037617 | 0.03 | 0.0304 | 0.021 | 0.04 | 0.03 | 0.02 | 0.02 | 0.021 | 0.029 | 0.013
| Mean |  | 0.023333 | 0.03 | 0.016833 | 0.0235 | 0.024 | 0.0205 | 0.019833 | 0.020167 | 0.02 | 0.021367 | 0.0185 | 0.031117 | 0.024157 |  | 0.038073 | 0.03 | 0.03055 | 0.020667 | 0.04 | 0.028333 | 0.02 | 0.02 | 0.021 | 0.0285 | 0.012667
| Median |  | 0.02 | 0.03 | 0.0175 | 0.0235 | 0.024 | 0.0205 | 0.02 | 0.02 | 0.02 | 0.02155 | 0.0185 | 0.03095 | 0.024072 |  | 0.038123 | 0.03 | 0.03035 | 0.021 | 0.04 | 0.03 | 0.02 | 0.02 | 0.0205 | 0.0285 | 0.013
| Std Dev. |  | 0.005164 | 0 | 0.001472 | 0.00054772 | 0.00063246 | 0.00054772 | 0.00040825 | 0.00040825 | 0 | 0.00086872 | 0.00054772 | 0.00096212 | 0.00026419 |  | 0.00032181 | 0 | 0.00045056 | 0.0005164 | 0 | 0.0040825 | 0 | 0 | 0.0017889 | 0.00054772 | 0.0005164
| Rel.Std.Dev. |  | 22.13% | 0% | 8.744% | 2.331% | 2.635% | 2.672% | 2.058% | 2.024% | 0% | 4.066% | 2.961% | 3.092% | 1.094% |  | 0.8452% | 0% | 1.475% | 2.499% | 0% | 14.41% | 0% | 0% | 8.518% | 1.922% | 4.077%
| PDM3 |  | -0.3722% | 28.09% | -28.13% | 0.3395% | 2.474% | -12.47% | -15.32% | -13.89% | -14.6% | -8.769% | -21.01% | 32.86% | 3.145% |  | 62.56% | 28.09% | 30.44% | -11.76% | 70.79% | 20.98% | -14.6% | -14.6% | -10.33% | 21.69% | -45.92%
| Z-Score (Absolute) |  | 0 | 1.28 | 1.25 | 0.03 | 0.13 | 0.55 | 0.67 | 0.61 | 0.64 | 0.38 | 0.93 | 1.5 | 0.16 | 5.14 | 2.84 | 1.28 | 1.39 | 0.51 | 3.21 | 0.96 | 0.64 | 0.64 | 0.45 | 1 | 2.06
| Analytical results for Tl in OREAS 625 (Certified Value 3.03 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g |  |  |  | 3
| 1 | 1 | < 10 | 3.31 | 3 | 2.78 | 2.99 | 3.01 | 2.91 | 2.84 | 3.12 | 3.12 | 3.21 | < 5 | 2.9888 | 5.99 | 3.0533 | < 5 | 3.81 | 1.89 | 2.808 | 3.83 | 2.99 | < 5 | 2.94 | 2.96 | 3.01 | < 5 |  |  |  | 1
| 1 | 2 | < 10 | 3.29 | 3 | 2.82 | 3.16 | 2.91 | 3.05 | 2.95 | 3.19 | 3.26 | 3.07 | < 5 | 2.9154 | 6.3 | 3.0628 | < 5 | 3.69 | 1.9 | 2.85 | 3.84 | 3.17 | < 5 | 2.88 | 3.03 | 3.15 | < 5 |  |  |  | 0
| 1 | 3 | < 10 | 3.3 | 3.2 | 2.81 | 3.2 | 3.05 | 2.86 | 2.95 | 3.15 | 3.2 | 3.26 | < 5 | 2.9288 | 6.25 | 3.0846 | < 5 | 3.85 | 1.98 | 2.809 | 3.87 | 2.95 | < 5 | 2.89 | 3 | 2.97 | < 5 |  |  |  | 16
| 1 | 4 | < 10 | 3.35 | 3 | 2.92 | 3.14 | 3.06 | 2.88 | 3.02 | 3.13 | 3.22 | 3.27 | < 5 | 3.0219 | 5.34 | 3.0306 | < 5 | 3.72 | 1.93 | 2.813 | 3.77 | 3.15 | < 5 | 2.85 | 3 | 3.16 | < 5 |  |  |  | 3.0305
| 1 | 5 | < 10 | 3.35 | 3 | 2.85 | 3.06 | 3.13 | 2.91 | 2.97 | 2.97 | 3.2 | 3.19 | < 5 | 2.9938 | 8.47 | 3.0343 | < 5 | 3.76 | 1.98 | 2.709 | 3.84 | 2.99 | < 5 | 2.9 | 2.92 | 3.04 | < 5 |  |  |  | 111
| 1 | 6 | < 10 | 3.31 | 3.1 | 2.82 | 2.9 | 3.05 | 2.94 | 2.89 | 3.16 | 3.18 | 3.11 | < 5 | 2.9093 | 6.48 | 3.0952 | < 5 | 3.75 | 1.95 | 2.825 | 3.8 | 3.06 | < 5 | 2.99 | 2.88 | 3.07 | < 5
| Mean |  |  | 3.3183 | 3.05 | 2.8333 | 3.075 | 3.035 | 2.925 | 2.9367 | 3.12 | 3.1967 | 3.185 |  | 2.9597 | 6.4717 | 3.0602 |  | 3.7633 | 1.9383 | 2.8023 | 3.825 | 3.0517 |  | 2.9083 | 2.965 | 3.0667
| Median |  |  | 3.31 | 3 | 2.82 | 3.1 | 3.05 | 2.91 | 2.95 | 3.14 | 3.2 | 3.2 |  | 2.9588 | 6.275 | 3.0581 |  | 3.755 | 1.94 | 2.811 | 3.835 | 3.025 |  | 2.895 | 2.98 | 3.055
| Std Dev. |  |  | 0.025626 | 0.083666 | 0.048028 | 0.11415 | 0.072595 | 0.067157 | 0.06314 | 0.07746 | 0.046332 | 0.080436 |  | 0.047629 | 1.0569 | 0.026181 |  | 0.058538 | 0.038687 | 0.048347 | 0.035071 | 0.091305 |  | 0.049565 | 0.05648 | 0.07607
| Rel.Std.Dev. |  |  | 0.7722% | 2.743% | 1.695% | 3.712% | 2.392% | 2.296% | 2.15% | 2.483% | 1.449% | 2.525% |  | 1.609% | 16.33% | 0.8555% |  | 1.555% | 1.996% | 1.725% | 0.9169% | 2.992% |  | 1.704% | 1.905% | 2.481%
| PDM3 |  |  | 9.497% | 0.6428% | -6.507% | 1.468% | 0.1479% | -3.482% | -3.097% | 2.953% | 5.482% | 5.098% |  | -2.338% | 113.5% | 0.9779% |  | 24.18% | -36.04% | -7.53% | 26.22% | 0.6978% |  | -4.032% | -2.162% | 1.193%
| Z-Score (Absolute) |  | 8.9 | 1.25 | 0.03 | 0.95 | 0.15 | 0.03 | 0.53 | 0.48 | 0.35 | 0.7 | 0.65 | 2.47 | 0.38 | 15.59 | 0.08 | 2.47 | 3.28 | 5.02 | 1.09 | 3.56 | 0.04 | 2.47 | 0.61 | 0.35 | 0.11 | 2.47
| Analytical results for Tm in OREAS 625 (Indicative Value 0.059 ppm) | [30 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13 |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g |  |  |  | 3
| 1 | 1 | 1.01: < 0.1; 1.02: 0.058; 1.03: 0.057322 |  |  |  | 1
| 1 | 2 | 1.01: < 0.1; 1.02: 0.062; 1.03: 0.059452 |  |  |  | 9
| 1 | 3 | 1.01: < 0.1; 1.02: 0.06; 1.03: 0.058357 |  |  |  | 16
| 1 | 4 | 1.01: < 0.1; 1.02: 0.06; 1.03: 0.058381 |  |  |  | 0.059004
| 1 | 5 | 1.01: < 0.1; 1.02: 0.061; 1.03: 0.052244 |  |  |  | 15
| 1 | 6 | 1.01: < 0.1; 1.02: 0.063; 1.03: 0.058292
| Mean |  | 1.02: 0.060667; 1.03: 0.057342
| Median |  | 1.02: 0.0605; 1.03: 0.058325
| Std Dev. |  | 1.02: 0.0017512; 1.03: 0.0025868
| Rel.Std.Dev. |  | 1.02: 2.887%; 1.03: 4.511%
| PDM3 |  | 1.02: 2.818%; 1.03: -2.818%
| Z-Score (Absolute) |  | 1.01: 1.49; 1.02: 0.67; 1.03: 0
| Analytical results for U in OREAS 625 (Certified Value 2.77 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 01g | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 2
| 1 | 1 | 3 | 2.5 | 2.79 | 2.61 | 2.58 | 2.58 | 2.53 | 2.4 | 2.94 | 2.7 | 2.8273 | 3.2406 | 3.33 | 2.9 | 2.57 | 3.23 | 1.6561 | 2.97 | < 8 | 2.62 | 2.3 | 2.66 | 2.53 | [6 empty] | 1
| 1 | 2 | 2.9 | 2.5 | 2.75 | 2.84 | 2.6 | 2.69 | 2.63 | 2.43 | 3.12 | 2.6 | 2.7699 | 3.2373 | 3.31 | 3 | 2.61 | 3.24 | 1.4969 | 2.91 | < 8 | 2.6 | 2.37 | 2.79 | 2.39 | [6 empty] | 41
| 1 | 3 | 3 | 2.7 | 2.77 | 2.89 | 2.72 | 2.58 | 2.53 | 2.42 | 3.06 | 2.7 | 2.8027 | 3.3017 | 3.34 | 3.1 | 2.51 | 3.24 | 1.6411 | 2.96 | < 8 | 2.59 | 2.33 | 2.61 | 2.39 | [6 empty] | 16
| 1 | 4 | 3 | 2.5 | 2.85 | 2.82 | 2.63 | 2.56 | 2.57 | 2.37 | 3.08 | 2.7 | 2.8571 | 3.2988 | 3.33 | 3 | 2.57 | 3.11 | 1.5023 | 3.02 | < 8 | 2.45 | 2.34 | 2.74 | 2.51 | [6 empty] | 2.77
| 1 | 5 | 2.9 | 2.6 | 2.77 | 2.86 | 2.72 | 2.59 | 2.54 | 2.25 | 3.07 | 2.6 | 2.8695 | 3.2448 | 3.39 | 3 | 2.37 | 3.23 | 1.6783 | 3.05 | < 8 | 2.65 | 2.23 | 2.73 | 2.39 | [6 empty] | 112
| 1 | 6 | 3 | 2.5 | 2.78 | 2.65 | 2.73 | 2.58 | 2.52 | 2.41 | 2.93 | 2.6 | 2.9415 | 3.2438 | 3.35 | 3.1 | 2.98 | 3.13 | 1.5864 | 3.01 | < 8 | 2.54 | 2.26 | 2.74 | 2.53
| Mean |  | 2.9667 | 2.55 | 2.785 | 2.7783 | 2.6633 | 2.5967 | 2.5533 | 2.38 | 3.0333 | 2.65 | 2.8447 | 3.2612 | 3.3417 | 3.0167 | 2.6017 | 3.1967 | 1.5935 | 2.9867 |  | 2.575 | 2.305 | 2.7117 | 2.4567
| Median |  | 3 | 2.5 | 2.775 | 2.83 | 2.675 | 2.58 | 2.535 | 2.405 | 3.065 | 2.65 | 2.8422 | 3.2443 | 3.335 | 3 | 2.57 | 3.23 | 1.6137 | 2.99 |  | 2.595 | 2.315 | 2.735 | 2.45
| Std Dev. |  | 0.05164 | 0.083666 | 0.034496 | 0.11788 | 0.067725 | 0.046762 | 0.041312 | 0.066933 | 0.078909 | 0.054772 | 0.059687 | 0.030421 | 0.027142 | 0.075277 | 0.20361 | 0.059889 | 0.078834 | 0.050067 |  | 0.071204 | 0.05244 | 0.064936 | 0.073394
| Rel.Std.Dev. |  | 1.741% | 3.281% | 1.239% | 4.243% | 2.543% | 1.801% | 1.618% | 2.812% | 2.601% | 2.067% | 2.098% | 0.9328% | 0.8122% | 2.495% | 7.826% | 1.873% | 4.947% | 1.676% |  | 2.765% | 2.275% | 2.395% | 2.988%
| PDM3 |  | 7.099% | -7.943% | 0.5406% | 0.3% | -3.852% | -6.258% | -7.823% | -14.08% | 9.506% | -4.333% | 2.695% | 17.73% | 20.64% | 8.904% | -6.078% | 15.4% | -42.47% | 7.821% |  | -7.041% | -16.79% | -2.107% | -11.31%
| Z-Score (Absolute) |  | 0.67 | 0.43 | 0.19 | 0.18 | 0.13 | 0.3 | 0.42 | 0.88 | 0.85 | 0.16 | 0.35 | 1.45 | 1.67 | 0.81 | 0.29 | 1.28 | 2.96 | 0.73 | 3.41 | 0.36 | 1.08 | 0 | 0.67
| Analytical results for V in OREAS 625 (Certified Value 3.11 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 2
| 1 | 1 | < 5 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2.7 | 3 | 3 | 4 | 3.5377 | 3.13 | 3.1679 | 12.6 | 5 | 3 | < 10 | 5 | 4 | 4 | 4 | 6 | 3 | 1 | 3 |  |  | 1
| 1 | 2 | < 5 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2.8 | 3 | 3 | 4 | 3.5103 | 2.75 | 3.4896 | 12.7 | 5 | 2 | < 10 | 5 | 4 | 3 | 3 | 6 | 3 | 1 | 3 |  |  | 23
| 1 | 3 | < 5 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2.8 | 3 | 3 | 4 | 3.4638 | 2.43 | 3.3579 | 12.9 | 5 | 2 | < 10 | 5 | 4 | 3 | 3 | 5 | 3 | 1 | 2 |  |  | 16
| 1 | 4 | < 5 | 3 | 4 | 3 | 3 | 3 | 3 | 3 | 2.8 | 3 | 3 | 4 | 3.5258 | 2.46 | 3.484 | 13.2 | 5 | 2 | < 10 | 5 | 4 | 3 | 3 | 6 | 4 | 1 | 3 |  |  | 3.1105
| 1 | 5 | < 5 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2.6 | 3 | 3 | 4 | 3.6325 | 2.51 | 3.0236 | 11.8 | 5 | 3 | < 10 | 5 | 3 | 4 | 4 | 6 | 4 | 1 | 2 |  |  | 113
| 1 | 6 | < 5 | 3 | 4 | 3 | 3 | 3 | 3 | 3 | 2.8 | 3 | 3 | 4 | 3.3906 | 2.37 | 3.1496 | 13.3 | 5 | 2 | < 10 | 5 | 4 | 3 | 3 | 6 | 3 | 1 | 3
| Mean |  |  | 3 | 3.3333 | 3 | 3 | 3 | 3 | 3 | 2.75 | 3 | 3 | 4 | 3.5101 | 2.6083 | 3.2788 | 12.75 | 5 | 2.3333 |  | 5 | 3.8333 | 3.3333 | 3.3333 | 5.8333 | 3.3333 | 1 | 2.6667
| Median |  |  | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2.8 | 3 | 3 | 4 | 3.5181 | 2.485 | 3.2629 | 12.8 | 5 | 2 |  | 5 | 4 | 3 | 3 | 6 | 3 | 1 | 3
| Std Dev. |  |  | 0 | 0.5164 | 0 | 0 | 0 | 0 | 0 | 0.083666 | 0 | 0 | 0 | 0.080504 | 0.28722 | 0.19335 | 0.53944 | 0 | 0.5164 |  | 0 | 0.40825 | 0.5164 | 0.5164 | 0.40825 | 0.5164 | 0 | 0.5164
| Rel.Std.Dev. |  |  | 0% | 15.49% | 0% | 0% | 0% | 0% | 0% | 3.042% | 0% | 0% | 0% | 2.293% | 11.01% | 5.897% | 4.231% | 0% | 22.13% |  | 0% | 10.65% | 15.49% | 15.49% | 6.999% | 15.49% | 0% | 19.36%
| PDM3 |  |  | -3.552% | 7.165% | -3.552% | -3.552% | -3.552% | -3.552% | -3.552% | -11.59% | -3.552% | -3.552% | 28.6% | 12.85% | -16.14% | 5.41% | 309.9% | 60.75% | -24.98% |  | 60.75% | 23.24% | 7.165% | 7.165% | 87.54% | 7.165% | -67.85% | -14.27%
| Z-Score (Absolute) |  | 1.01 | 0 | 0.67 | 0 | 0 | 0 | 0 | 0 | 0.51 | 0 | 0 | 2.02 | 1.03 | 0.79 | 0.56 | 19.72 | 4.05 | 1.35 | 4.05 | 4.05 | 1.69 | 0.67 | 0.67 | 5.73 | 0.67 | 4.05 | 0.67
| Analytical results for W in OREAS 625 (Certified Value 1.47 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  |  | 3
| 1 | 1 | < 5 | 1.2 | 1.5 | 1.4 | 1.49 | 1.4 | 1.38 | 1.36 | 1.6 | 1.44 | 1.2 | 3 | 1.3376 | 1.3296 | < 5 | 1.5 | 1.7 | 1.6 | 1.6 | < 2 | 1.7 | 1.6 | 1.43 | < 10 | 1.45 |  |  |  |  | 1
| 1 | 2 | < 5 | 1.2 | 1.7 | 1.38 | 1.51 | 1.38 | 1.38 | 1.37 | 1.62 | 1.49 | 1.2 | 6 | 1.284 | 1.3838 | < 5 | 1.5 | 1.7 | 1.5 | 1.7 | < 2 | 1.6 | 1.6 | 1.5 | < 10 | 1.32 |  |  |  |  | 43
| 1 | 3 | < 5 | 1.4 | 1.6 | 1.36 | 1.52 | 1.44 | 1.37 | 1.46 | 1.56 | 1.48 | 1.3 | 3 | 1.275 | 1.423 | < 5 | 1.5 | 1.8 | 1.6 | 1.6 | < 2 | 1.7 | 1.6 | 1.41 | < 10 | 1.52 |  |  |  |  | 16
| 1 | 4 | < 5 | 1.3 | 1.4 | 1.42 | 1.54 | 1.41 | 1.38 | 1.34 | 1.61 | 1.54 | 1.3 | 9 | 1.3051 | 1.4311 | < 5 | 1.5 | 1.7 | 1.5 | 1.6 | < 2 | 1.6 | 1.6 | 1.44 | < 10 | 1.46 |  |  |  |  | 1.4702
| 1 | 5 | < 5 | 1.4 | 1.6 | 1.38 | 1.51 | 1.52 | 1.45 | 1.36 | 1.52 | 1.48 | 1.2 | 12 | 1.2666 | 1.5347 | < 5 | 1.5 | 1.6 | 1.5 | 1.6 | < 2 | 1.5 | 1.5 | 1.48 | < 10 | 1.5 |  |  |  |  | 114
| 1 | 6 | < 5 | 1.3 | 1.5 | 1.38 | 1.46 | 1.43 | 1.4 | 1.31 | 1.66 | 1.45 | 1.2 | 3 | 1.28 | 1.5131 | < 5 | 1.5 | 1.7 | 1.5 | 1.6 | < 2 | 1.6 | 1.5 | 1.45 | < 10 | 1.52
| Mean |  |  | 1.3 | 1.55 | 1.3867 | 1.505 | 1.43 | 1.3933 | 1.3667 | 1.595 | 1.48 | 1.2333 | 6 | 1.2914 | 1.4359 |  | 1.5 | 1.7 | 1.5333 | 1.6167 |  | 1.6167 | 1.5667 | 1.4517 |  | 1.4617
| Median |  |  | 1.3 | 1.55 | 1.38 | 1.51 | 1.42 | 1.38 | 1.36 | 1.605 | 1.48 | 1.2 | 4.5 | 1.282 | 1.427 |  | 1.5 | 1.7 | 1.5 | 1.6 |  | 1.6 | 1.6 | 1.445 |  | 1.48
| Std Dev. |  |  | 0.089443 | 0.10488 | 0.020656 | 0.027386 | 0.04899 | 0.029439 | 0.050465 | 0.048888 | 0.035214 | 0.05164 | 3.7947 | 0.026039 | 0.077352 |  | 0 | 0.063246 | 0.05164 | 0.040825 |  | 0.075277 | 0.05164 | 0.033116 |  | 0.075476
| Rel.Std.Dev. |  |  | 6.88% | 6.767% | 1.49% | 1.82% | 3.426% | 2.113% | 3.693% | 3.065% | 2.379% | 4.187% | 63.25% | 2.016% | 5.387% |  | 0% | 3.72% | 3.368% | 2.525% |  | 4.656% | 3.296% | 2.281% |  | 5.164%
| PDM3 |  |  | -11.57% | 5.431% | -5.679% | 2.37% | -2.731% | -5.225% | -7.039% | 8.492% | 0.6699% | -16.11% | 308.1% | -12.16% | -2.331% |  | 2.03% | 15.63% | 4.298% | 9.966% |  | 9.966% | 6.565% | -1.257% |  | -0.5771%
| Z-Score (Absolute) |  | 5.95 | 1.19 | 0.3 | 0.67 | 0.03 | 0.42 | 0.63 | 0.79 | 0.57 | 0.12 | 1.59 | 26.77 | 1.24 | 0.38 | 5.95 | 0 | 1.19 | 0.2 | 0.69 | 2.97 | 0.69 | 0.4 | 0.29 | 20.82 | 0.23
| Analytical results for Y in OREAS 625 (Certified Value 6.4 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [8 empty] | 3
| 1 | 1 | 7 | 6.14 | 6.86 | 6.53 | 6.58 | 6.06 | 5.77 | 5.24 | 6.36 | 7.016 | 7.245 | 8.32 | 6.1 | < 10 | 8.1 | 6.96 | 6.13 | 5.9 | 6.56 | 6.6 | 6.31 | [8 empty] | 1
| 1 | 2 | 7 | 6.1 | 6.82 | 6.96 | 6.46 | 6.14 | 6.07 | 5.69 | 6.87 | 6.8404 | 7.2101 | 8.18 | 6.1 | < 10 | 8.3 | 6.42 | 5.83 | 6.27 | 6.62 | 6.5 | 6.3 | [8 empty] | 44
| 1 | 3 | 7 | 6.17 | 6.85 | 6.94 | 6.68 | 5.94 | 5.82 | 5.56 | 6.73 | 6.7878 | 7.6495 | 8.36 | 6.3 | < 10 | 8.4 | 7.05 | 5.83 | 6.21 | 6.53 | 6.6 | 6.23 | [8 empty] | 16
| 1 | 4 | 7 | 6.15 | 6.99 | 6.92 | 6.67 | 5.88 | 5.79 | 5.18 | 6.79 | 6.8502 | 7.6045 | 8.2 | 6.2 | < 10 | 8.1 | 6.27 | 5.58 | 6.41 | 6.87 | 6.5 | 6.23 | [8 empty] | 6.3963
| 1 | 5 | 7 | 6.38 | 6.91 | 6.87 | 6.65 | 5.97 | 5.76 | 5.33 | 6.71 | 6.7957 | 7.3239 | 8.33 | 6.3 | < 10 | 8.3 | 6.67 | 5.82 | 5.92 | 6.39 | 6.6 | 6.13 | [8 empty] | 115
| 1 | 6 | 7 | 6.5 | 6.67 | 6.98 | 6.46 | 5.99 | 5.64 | 5.62 | 6.35 | 6.8395 | 7.3665 | 8.35 | 6.1 | < 10 | 8.3 | 6.63 | 5.72 | 5.94 | 6.49 | 6.5 | 6.35
| Mean |  | 7 | 6.24 | 6.85 | 6.8667 | 6.5833 | 5.9967 | 5.8083 | 5.4367 | 6.635 | 6.8549 | 7.3999 | 8.29 | 6.1833 |  | 8.25 | 6.6667 | 5.8183 | 6.1083 | 6.5767 | 6.55 | 6.2583
| Median |  | 7 | 6.16 | 6.855 | 6.93 | 6.615 | 5.98 | 5.78 | 5.445 | 6.72 | 6.8399 | 7.3452 | 8.325 | 6.15 |  | 8.3 | 6.65 | 5.825 | 6.075 | 6.545 | 6.55 | 6.265
| Std Dev. |  | 0 | 0.16112 | 0.1064 | 0.16919 | 0.10172 | 0.091797 | 0.14219 | 0.21398 | 0.22394 | 0.082971 | 0.18498 | 0.078994 | 0.098319 |  | 0.12247 | 0.30084 | 0.18104 | 0.21665 | 0.16293 | 0.054772 | 0.078592
| Rel.Std.Dev. |  | 0% | 2.582% | 1.553% | 2.464% | 1.545% | 1.531% | 2.448% | 3.936% | 3.375% | 1.21% | 2.5% | 0.9529% | 1.59% |  | 1.485% | 4.513% | 3.112% | 3.547% | 2.477% | 0.8362% | 1.256%
| PDM3 |  | 9.438% | -2.444% | 7.093% | 7.353% | 2.924% | -6.248% | -9.193% | -15% | 3.731% | 7.17% | 15.69% | 29.61% | -3.33% |  | 28.98% | 4.227% | -9.036% | -4.502% | 2.819% | 2.403% | -2.157%
| Z-Score (Absolute) |  | NA | 0.62 | 0.55 | 0.58 | 0.04 | 1.09 | 1.45 | 2.16 | 0.14 | 0.56 | 1.6 | 3.31 | 0.73 | 3 | 3.23 | 0.2 | 1.43 | 0.87 | 0.03 | 0.03 | 0.58
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 625 (Certified Value 0.34 ppm) | [30 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.21; 1.05: 1.22; 1.06: 1.23; 1.07: 1.25; 1.08: 1.26 |  |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS |  |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.2g; 1.05: 0.5g; 1.06: 0.25g; 1.07: 0.25g; 1.08: 0.25g |  |  |  |  | 2
| 1 | 1 | 1.01: 0.4; 1.02: 0.326; 1.03: 0.35265; 1.04: 0.4; 1.05: 0.3512; 1.06: 0.3; 1.07: 0.3; 1.08: 0.3 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.4; 1.02: 0.337; 1.03: 0.34731; 1.04: 0.4; 1.05: 0.306; 1.06: 0.3; 1.07: 0.4; 1.08: 0.3 |  |  |  |  | 45
| 1 | 3 | 1.01: 0.4; 1.02: 0.33; 1.03: 0.3327; 1.04: 0.4; 1.05: 0.35136; 1.06: 0.3; 1.07: 0.3; 1.08: 0.3 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.4; 1.02: 0.334; 1.03: 0.33563; 1.04: 0.4; 1.05: 0.3307; 1.06: 0.3; 1.07: 0.3; 1.08: 0.3 |  |  |  |  | 0.3408
| 1 | 5 | 1.01: 0.4; 1.02: 0.328; 1.03: 0.30344; 1.04: 0.4; 1.05: 0.34021; 1.06: 0.3; 1.07: 0.3; 1.08: 0.3 |  |  |  |  | 116
| 1 | 6 | 1.01: 0.4; 1.02: 0.341; 1.03: 0.33562; 1.04: 0.4; 1.05: 0.33811; 1.06: 0.3; 1.07: 0.3; 1.08: 0.3
| Mean |  | 1.01: 0.4; 1.02: 0.33267; 1.03: 0.33456; 1.04: 0.4; 1.05: 0.33626; 1.06: 0.3; 1.07: 0.31667; 1.08: 0.3
| Median |  | 1.01: 0.4; 1.02: 0.332; 1.03: 0.33562; 1.04: 0.4; 1.05: 0.33916; 1.06: 0.3; 1.07: 0.3; 1.08: 0.3
| Std Dev. |  | 1.01: 6.0809e-17; 1.02: 0.0057155; 1.03: 0.017111; 1.04: 6.0809e-17; 1.05: 0.016835; 1.06: 0; 1.07: 0.040825; 1.08: 0
| Rel.Std.Dev. |  | 1.01: 1.52e-14%; 1.02: 1.718%; 1.03: 5.115%; 1.04: 1.52e-14%; 1.05: 5.006%; 1.06: 0%; 1.07: 12.89%; 1.08: 0%
| PDM3 |  | 1.01: 17.37%; 1.02: -2.386%; 1.03: -1.831%; 1.04: 17.37%; 1.05: -1.331%; 1.06: -11.97%; 1.07: -7.081%; 1.08: -11.97%
| Z-Score (Absolute) |  | 1.01: 1.77; 1.02: 0.03; 1.03: 0.03; 1.04: 1.77; 1.05: 0.07; 1.06: 0.9; 1.07: 0.45; 1.08: 0.9
| Analytical results for Zn in OREAS 625 (Certified Value 3.18 wt.%) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*AAS | AR*OES | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*AAS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*AAS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*OES | AR*OES | AR*OES/MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.2g | 0.5g | 0.1g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 0.25g | 01g | 0.25g | 0.25g | 0.4g | 0.25g | 0.2g |  | 3
| 1 | 1 | 3.123 | 3.19 | 3.035 | 3.11 | 3.03 | 3.1 | 3.15 | 3.21 | 3.18 | > 2 | 3.19 | > 1 | 3.216 | 3.6181 | 3.2937 | 3.38 | > 2 | > 1 | 3.1392 | 3.295 | 3.047 | 3.22 | 3.21 | > 1 | 3.14 | 3.1 | > 1 | 3.2768 |  | 1
| 1 | 2 | 3.042 | 3.18 | 2.934 | 3.16 | 3.01 | 3.13 | 3.15 | 3.19 | 3.36 | > 2 | 3.18 | > 1 | 3.1596 | 3.3079 | 3.276 | 3.36 | > 2 | > 1 | 3.1546 | 3.3502 | 3.059 | 3.14 | 3.16 | > 1 | 3.21 | 3.091 | > 1 | 3.2624 |  | 24
| 1 | 3 | 3.06 | 3.26 | 3.01 | 3.13 | 3.06 | 3.15 | 3.15 | 3.21 | 3.28 | > 2 | 3.18 | > 1 | 3.1914 | 3.3254 | 3.2579 | 3.37 | > 2 | > 1 | 3.1526 | 3.2885 | 3.042 | 3.2 | 3.21 | > 1 | 3.14 | 3.131 | > 1 | 3.2217 |  | 16
| 1 | 4 | 3.159 | 3.23 | 3.112 | 3.14 | 3.03 | 3.11 | 3.16 | 3.22 | 3.47 | > 2 | 3.22 | > 1 | 3.228 | 3.3546 | 3.2434 | 3.36 | > 2 | > 1 | 3.1536 | 3.3357 | 3.031 | 3.22 | 3.18 | > 1 | 3.13 | 3.134 | > 1 | 3.2511 |  | 3.1792
| 1 | 5 | 3.177 | 3.24 | 3.109 | 3.17 | 3.07 | 3.16 | 3.16 | 3.49 | 3.23 | > 2 | 3.21 | > 1 | 3.2213 | 3.5289 | 3.2835 | 3.34 | > 2 | > 1 | 3.1547 | 3.3472 | 3.029 | 3.18 | 3.17 | > 1 | 3.14 | 3.059 | > 1 | 3.2153 |  | 117
| 1 | 6 | 3.096 | 3.27 | 3.078 | 3.16 | 3.11 | 3.15 | 3.18 | 3.14 | 3.34 | > 2 | 3.18 | > 1 | 3.2366 | 3.2816 | 3.2348 | 3.37 | > 2 | > 1 | 3.1537 | 3.3051 | 3.048 | 3.22 | 3.17 | > 1 | 3.13 | 3.131 | > 1 | 3.2046
| Mean |  | 3.1095 | 3.2283 | 3.0463 | 3.145 | 3.0517 | 3.1333 | 3.1583 | 3.2433 | 3.31 |  | 3.1933 |  | 3.2088 | 3.4028 | 3.2649 | 3.3633 |  |  | 3.1514 | 3.3203 | 3.0427 | 3.1967 | 3.1833 |  | 3.1483 | 3.1077 |  | 3.2386
| Median |  | 3.1095 | 3.235 | 3.0565 | 3.15 | 3.045 | 3.14 | 3.155 | 3.21 | 3.31 |  | 3.185 |  | 3.2187 | 3.34 | 3.2669 | 3.365 |  |  | 3.1536 | 3.3204 | 3.0445 | 3.21 | 3.175 |  | 3.14 | 3.1155 |  | 3.2364
| Std Dev. |  | 0.053624 | 0.03656 | 0.068278 | 0.022583 | 0.036009 | 0.024221 | 0.01169 | 0.1242 | 0.10315 |  | 0.017512 |  | 0.028535 | 0.1373 | 0.023287 | 0.013663 |  |  | 0.0060256 | 0.027339 | 0.01129 | 0.032042 | 0.021602 |  | 0.030605 | 0.029958 |  | 0.028867
| Rel.Std.Dev. |  | 1.725% | 1.132% | 2.241% | 0.7181% | 1.18% | 0.773% | 0.3701% | 3.83% | 3.116% |  | 0.5484% |  | 0.8893% | 4.035% | 0.7133% | 0.4062% |  |  | 0.1912% | 0.8234% | 0.3711% | 1.002% | 0.6786% |  | 0.9721% | 0.964% |  | 0.8913%
| PDM3 |  | -2.193% | 1.544% | -4.18% | -1.077% | -4.012% | -1.444% | -0.6574% | 2.016% | 4.113% |  | 0.4435% |  | 0.9312% | 7.031% | 2.694% | 5.791% |  |  | -0.8755% | 4.437% | -4.296% | 0.5484% | 0.129% |  | -0.9719% | -2.251% |  | 1.869%
| Z-Score (Absolute) |  | 0.97 | 0.49 | 1.74 | 0.53 | 1.68 | 0.67 | 0.37 | 0.67 | 1.49 | NA | 0.06 | NA | 0.25 | 2.63 | 0.94 | 2.15 | NA | NA | 0.45 | 1.62 | 1.79 | 0.1 | 0.06 | NA | 0.49 | 0.99 | NA | 0.62
| Analytical results for Zr in OREAS 625 (Certified Value 55 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [7 empty] | 0
| 1 | 1 | 62 | 48.1 | 58.3 | 54 | 61 | 50.7 | 55.2 | 37.7 | 54.2 | 58.397 | 52.4938 | 30.5 | 69.5 | 61.6 | 54.33 | 61.4 | 61 | 45.8 | 44.8 | 51 | 71.1 | 50.2 | [7 empty] | 1
| 1 | 2 | 62 | 35.3 | 58.3 | 57.1 | 59.7 | 51.6 | 57.6 | 41.1 | 60 | 57.6064 | 52.7165 | 31.5 | 69.8 | 60 | 53.12 | 63.9 | 58.1 | 43.7 | 46.2 | 53.5 | 70.5 | 38.4 | [7 empty] | 16
| 1 | 3 | 61 | 52.6 | 57.7 | 58 | 60.1 | 49.5 | 55.5 | 40.2 | 57.8 | 56.6996 | 54.3949 | 39.8 | 75.6 | 62 | 54.16 | 65.4 | 61.5 | 42.6 | 47.9 | 51 | 69.8 | 43.5 | [7 empty] | 16
| 1 | 4 | 63 | 44.6 | 59.1 | 57.1 | 63.1 | 48.5 | 56.1 | 36.2 | 60 | 57.6552 | 54.0843 | 39.1 | 74.6 | 61.9 | 54.74 | 56.5 | 59 | 44.5 | 49.4 | 58.1 | 71.4 | 51.8 | [7 empty] | 54.7422
| 1 | 5 | 62 | 49.6 | 58.3 | 56.1 | 60.9 | 49.5 | 55.4 | 38 | 59 | 56.8984 | 52.6108 | 38.3 | 70.1 | 63 | 54.38 | 61.8 | 60.7 | 45.8 | 46.3 | 54.2 | 72.3 | 43.8 | [7 empty] | 118
| 1 | 6 | 62 | 45.7 | 57.3 | 60.8 | 58 | 48.7 | 54.7 | 40.2 | 54.7 | 57.06 | 52.9699 | 43.2 | 70 | 60.7 | 54.18 | 62.8 | 60.4 | 42.9 | 46.5 | 53.5 | 72.2 | 53.2
| Mean |  | 62 | 45.9833 | 58.1667 | 57.1833 | 60.4667 | 49.75 | 55.75 | 38.9 | 57.6167 | 57.3861 | 53.2117 | 37.0667 | 71.6 | 61.5333 | 54.1517 | 61.9667 | 60.1167 | 44.2167 | 46.85 | 53.55 | 71.2167 | 46.8167
| Median |  | 62 | 46.9 | 58.3 | 57.1 | 60.5 | 49.5 | 55.45 | 39.1 | 58.4 | 57.3332 | 52.8432 | 38.7 | 70.05 | 61.75 | 54.255 | 62.3 | 60.55 | 44.1 | 46.4 | 53.5 | 71.25 | 47
| Std Dev. |  | 0.63246 | 5.9577 | 0.61536 | 2.2391 | 1.686 | 1.1929 | 1.0134 | 1.8847 | 2.588 | 0.62584 | 0.81752 | 4.9971 | 2.7372 | 1.0539 | 0.54686 | 3.0481 | 1.2983 | 1.3934 | 1.5909 | 2.6129 | 0.9704 | 5.796
| Rel.Std.Dev. |  | 1.02% | 12.96% | 1.058% | 3.916% | 2.788% | 2.398% | 1.818% | 4.845% | 4.492% | 1.091% | 1.536% | 13.48% | 3.823% | 1.713% | 1.01% | 4.919% | 2.16% | 3.151% | 3.396% | 4.879% | 1.363% | 12.38%
| PDM3 |  | 13.26% | -16% | 6.256% | 4.459% | 10.46% | -9.12% | 1.841% | -28.94% | 5.251% | 4.83% | -2.796% | -32.29% | 30.79% | 12.41% | -1.079% | 13.2% | 9.818% | -19.23% | -14.42% | -2.178% | 30.09% | -14.48%
| Z-Score (Absolute) |  | 0.71 | 1.34 | 0.22 | 0.09 | 0.51 | 0.86 | 0.09 | 2.24 | 0.15 | 0.12 | 0.42 | 2.48 | 1.93 | 0.65 | 0.3 | 0.7 | 0.47 | 1.56 | 1.23 | 0.37 | 1.88 | 1.23
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C817: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C835: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 625 (Indicative Value 12.6 ppm) | [24 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.10 | 1.12 | 1.16 | 1.23 | [19 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*MS | [19 empty] | 1
1
| 1 | 1 | 13 | 15 | 11.13 | 12 | [19 empty] | 1
| 1 | 2 | 13 | 15 | 10.99 | 11 | [19 empty] | 1
| 1 | 3 | 12 | 15 | 11.06 | 12 | [19 empty] | 16
| 1 | 4 | 12 | 13 | 11.15 | 12 | [19 empty] | 12.5554
| 1 | 5 | 12 | 17 | 10.98 | 12 | [19 empty] | 18
| 1 | 6 | 12 | 15 | 11.02 | 12
| Mean |  | 12.3333 | 15 | 11.055 | 11.8333
| Median |  | 12 | 15 | 11.04 | 12
| Std Dev. |  | 0.5164 | 1.2649 | 0.071764 | 0.40825
| Rel.Std.Dev. |  | 4.187% | 8.433% | 0.6491% | 3.45%
| PDM3 |  | -1.769% | 19.47% | -11.95% | -5.751%
| Z-Score (Absolute) |  | 0.26 | 3.08 | 1.08 | 0.26
| Analytical results for Al in OREAS 625 (Certified Value 6.83 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 7.71 | 6.81 | 7.19 | 6.774 | 6.854 | 6.986 | 6.748 | 6.616 | 6.96 | 6.98 | 6.57 | 6.75 | 6.5911 | 6.8881 | 6.49 | 7.09 | 7.3976 | 5.9972 | 6.95 | 6.82 | 6.67 | 6.83 |  | 1
| 1 | 2 | 7.7 | 6.86 | 6.66 | 6.748 | 6.933 | 6.801 | 6.721 | 6.536 | 7 | 7.13 | 6.8 | 6.6 | 6.6758 | 6.8868 | 6.47 | 6.98 | 7.7769 | 5.9615 | 7.02 | 6.88 | 6.37 | 6.93 |  | 0
| 1 | 3 | 7.78 | 6.88 | 6.94 | 6.801 | 6.774 | 6.854 | 6.774 | 6.748 | 6.96 | 6.86 | 6.76 | 6.61 | 6.5632 | 6.9402 | 6.53 | 7.07 | 7.5256 | 5.9192 | 7.02 | 6.83 | 6.71 | 6.93 |  | 16
| 1 | 4 | 7.76 | 6.86 | 6.87 | 6.748 | 6.933 | 6.774 | 6.96 | 6.774 | 7.01 | 6.85 | 6.61 | 6.66 | 6.6582 | 7.0245 | 6.53 | 7.02 | 7.4109 | 5.9276 | 7.06 | 6.85 | 6.88 | 6.94 |  | 6.8273
| 1 | 5 | 7.79 | 6.83 | 6.9 | 6.774 | 6.933 | 6.801 | 6.907 | 6.616 | 6.99 | 7.06 | 6.5 | 6.69 | 6.8971 | 7.0757 | 6.52 | 7.03 | 7.6158 | 5.9961 | 6.88 | 6.84 | 6.65 | 6.95 |  | 120
| 1 | 6 | 7.72 | 6.88 | 7.22 | 6.854 | 6.854 | 6.88 | 6.827 | 6.642 | 6.95 | 6.87 | 6.76 | 6.78 | 6.7203 | 7.0861 | 6.44 | 7.05 | 7.3498 | 5.9752 | 6.98 | 6.88 | 6.63 | 7.03
| Mean |  | 7.7433 | 6.8533 | 6.9633 | 6.7832 | 6.8802 | 6.8493 | 6.8228 | 6.6553 | 6.9783 | 6.9583 | 6.6667 | 6.6817 | 6.6843 | 6.9836 | 6.4967 | 7.04 | 7.5128 | 5.9628 | 6.985 | 6.85 | 6.6517 | 6.935
| Median |  | 7.74 | 6.86 | 6.92 | 6.774 | 6.8935 | 6.8275 | 6.8005 | 6.629 | 6.975 | 6.925 | 6.685 | 6.675 | 6.667 | 6.9823 | 6.505 | 7.04 | 7.4683 | 5.9683 | 7 | 6.845 | 6.66 | 6.935
| Std Dev. |  | 0.038297 | 0.028048 | 0.21097 | 0.039942 | 0.06483 | 0.077443 | 0.094097 | 0.089645 | 0.024833 | 0.11788 | 0.12291 | 0.073052 | 0.1189 | 0.090606 | 0.036697 | 0.038987 | 0.16171 | 0.033424 | 0.063797 | 0.025298 | 0.16473 | 0.063797
| Rel.Std.Dev. |  | 0.4946% | 0.4093% | 3.03% | 0.5888% | 0.9423% | 1.131% | 1.379% | 1.347% | 0.3559% | 1.694% | 1.844% | 1.093% | 1.779% | 1.297% | 0.5649% | 0.5538% | 2.152% | 0.5605% | 0.9133% | 0.3693% | 2.477% | 0.9199%
| PDM3 |  | 13.42% | 0.3816% | 1.993% | -0.6461% | 0.7746% | 0.323% | -0.06513% | -2.519% | 2.212% | 1.92% | -2.353% | -2.133% | -2.095% | 2.289% | -4.843% | 3.116% | 10.04% | -12.66% | 2.31% | 0.3328% | -2.572% | 1.578%
| Z-Score (Absolute) |  | 4.53 | 0.01 | 0.57 | 0.35 | 0.14 | 0.01 | 0.15 | 1 | 0.64 | 0.54 | 0.94 | 0.86 | 0.85 | 0.67 | 1.81 | 0.96 | 3.36 | 4.52 | 0.68 | 0.01 | 1.02 | 0.42
| Analytical results for As in OREAS 625 (Certified Value 144 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 145 | 144 | 145 | 100 | 200 | 200 | 200 | 200 | 136 | 158 | 200 | 146 | 152.2877 | 135.2 | 150 | 132 | 156 | 140 | 200 | 100 |  |  |  | 1
| 1 | 2 | 145 | 150 | 131 | 100 | 100 | 200 | 200 | 100 | 141 | 149 | 100 | 135 | 150.2324 | 131.1 | 150 | 134 | 156 | 150 | 100 | 100 |  |  |  | 21
| 1 | 3 | 145 | 155 | 164 | 200 | 200 | 200 | 200 | 200 | 134 | 151 | 200 | 126 | 142.7054 | 134.5 | 150 | 137 | 154 | 130 | 200 | < 100 |  |  |  | 16
| 1 | 4 | 145 | 145 | 139 | 100 | 200 | 200 | 200 | 200 | 131 | 154 | < 100 | 134 | 142.2413 | 129.7 | 150 | 138 | 158 | 140 | 200 | < 100 |  |  |  | 144.1385
| 1 | 5 | 144 | 144 | 152 | 200 | 200 | 200 | 100 | 200 | 136 | 155 | < 100 | 135 | 152.9902 | 131.2 | 150 | 131 | 156 | 140 | 100 | 100 |  |  |  | 121
| 1 | 6 | 145 | 144 | 151 | < 100 | 200 | 200 | 200 | 200 | 129 | 150 | 100 | 136 | 149.3062 | 137.6 | 150 | 135 | 157 | 140 | 100 | < 100
| Mean |  | 144.8333 | 147 | 147 | 140 | 183.3333 | 200 | 183.3333 | 183.3333 | 134.5 | 152.8333 | 150 | 135.3333 | 148.2939 | 133.2167 | 150 | 134.5 | 156.1667 | 140 | 150 | 100
| Median |  | 145 | 144.5 | 148 | 100 | 200 | 200 | 200 | 200 | 135 | 152.5 | 150 | 135 | 149.7693 | 132.85 | 150 | 134.5 | 156 | 140 | 150 | 100
| Std Dev. |  | 0.40825 | 4.5607 | 11.4368 | 54.7723 | 40.8248 | 0 | 40.8248 | 40.8248 | 4.2308 | 3.4303 | 57.735 | 6.377 | 4.7043 | 3.0235 | 0 | 2.7386 | 1.3292 | 6.3246 | 54.7723 | 0
| Rel.Std.Dev. |  | 0.2819% | 3.103% | 7.78% | 39.12% | 22.27% | 0% | 22.27% | 22.27% | 3.146% | 2.244% | 38.49% | 4.712% | 3.172% | 2.27% | 0% | 2.036% | 0.8511% | 4.518% | 36.51% | 0%
| PDM3 |  | 0.4821% | 1.985% | 1.985% | -2.871% | 27.19% | 38.76% | 27.19% | 27.19% | -6.687% | 6.032% | 4.067% | -6.109% | 2.883% | -7.577% | 4.067% | -6.687% | 8.345% | -2.871% | 4.067% | -30.62%
| Z-Score (Absolute) |  | 0.12 | 0 | 0 | 1.23 | 2.03 | 2.96 | 2.03 | 2.03 | 0.7 | 0.33 | 1.69 | 0.65 | 0.07 | 0.77 | 0.17 | 0.7 | 0.51 | 0.39 | 0.17 | 4.02
| Analytical results for B in OREAS 625 (Indicative Value 25.3 ppm) | [24 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.20; 1.06: 1.26 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES; 1.05: PF*MS; 1.06: PF*OES |  |  |  | 1
1
| 1 | 1 | 1.01: 10; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 35; 1.06: 30 |  |  |  | 1
| 1 | 2 | 1.01: 10; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 35; 1.06: 30 |  |  |  | 3
| 1 | 3 | 1.01: 20; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 32; 1.06: 30 |  |  |  | 16
| 1 | 4 | 1.01: 10; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 31; 1.06: 30 |  |  |  | 25.3333
| 1 | 5 | 1.01: 10; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 33; 1.06: 30 |  |  |  | 19
| 1 | 6 | 1.01: 30; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 30; 1.06: 20
| Mean |  | 1.01: 15; 1.05: 32.6667; 1.06: 28.3333
| Median |  | 1.01: 10; 1.05: 32.5; 1.06: 30
| Std Dev. |  | 1.01: 8.3666; 1.05: 2.0656; 1.06: 4.0825
| Rel.Std.Dev. |  | 1.01: 55.78%; 1.05: 6.323%; 1.06: 14.41%
| PDM3 |  | 1.01: -40.79%; 1.05: 28.95%; 1.06: 11.84%
| Z-Score (Absolute) |  | 1.01: 1.23; 1.02: 1.84; 1.03: 0; 1.04: 0; 1.05: 0.94; 1.06: 0.41
| Analytical results for Ba in OREAS 625 (Certified Value 3415 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | [9 empty] | 0
0
| 1 | 1 | 3635 | 3380 | 3382 | 3490 | 3504 | 3277 | 3248.3089 | 3439.8585 | 3565 | 3870 | 3772 | 3302 | 2973 | 3000 | [9 empty] | 1
| 1 | 2 | 3615 | 3300 | 3204 | 3520 | 3484 | 3197 | 3226.1395 | 3451.1412 | 3444.6 | 3810 | 3860 | 3360 | 3012 | 3030 | [9 empty] | 4
| 1 | 3 | 3611 | 3660 | 3322 | 3460 | 3515 | 3188 | 3201.3948 | 3359.3777 | 3537.1 | 3750 | 3800 | 3310 | 3028 | 2990 | [9 empty] | 16
| 1 | 4 | 3661 | 3410 | 3278 | 3510 | 3533 | 3218 | 3213.523 | 3381.2751 | 3495.4 | 3810 | 3750 | 3326 | 3055 | 3070 | [9 empty] | 3415.3709
| 1 | 5 | 3646 | 3450 | 3290 | 3500 | 3527 | 3235 | 3290.6892 | 3411.6627 | 3528 | 3890 | 3904 | 3280 | 3032 | 3050 | [9 empty] | 122
| 1 | 6 | 3630 | 3460 | 3450 | 3470 | 3515 | 3284 | 3311.6652 | 3454.6207 | 3554.4 | 3750 | 3829 | 3313 | 3020 | 3020
| Mean |  | 3633 | 3443.3333 | 3321 | 3491.6667 | 3513 | 3233.1667 | 3248.6201 | 3416.3226 | 3520.75 | 3813.3333 | 3819.1667 | 3315.1667 | 3020 | 3026.6667
| Median |  | 3632.5 | 3430 | 3306 | 3495 | 3515 | 3226.5 | 3237.2242 | 3425.7606 | 3532.55 | 3810 | 3814.5 | 3311.5 | 3024 | 3025
| Std Dev. |  | 18.8361 | 120.7753 | 85.8627 | 23.1661 | 17.47 | 40.2165 | 44.0727 | 39.3085 | 44.3885 | 58.5377 | 57.1609 | 26.7164 | 27.225 | 30.1109
| Rel.Std.Dev. |  | 0.5185% | 3.508% | 2.585% | 0.6635% | 0.4973% | 1.244% | 1.357% | 1.151% | 1.261% | 1.535% | 1.497% | 0.8059% | 0.9015% | 0.9949%
| PDM3 |  | 6.372% | 0.8187% | -2.763% | 2.234% | 2.859% | -5.335% | -4.882% | 0.02787% | 3.085% | 11.65% | 11.82% | -2.934% | -11.58% | -11.38%
| Z-Score (Absolute) |  | 0.93 | 0.06 | 0.5 | 0.28 | 0.38 | 0.9 | 0.83 | 0.06 | 0.41 | 1.75 | 1.77 | 0.52 | 1.87 | 1.84
| Analytical results for Be in OREAS 625 (Indicative Value 2.91 ppm) | [24 empty] | Indicative
Be | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.16; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23; 1.09: 1.25; 1.10: 1.26 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES; 1.09: PF*OES; 1.10: PF*OES |  |  |  | 2
2
| 1 | 1 | 1.01: < 3; 1.02: < 5; 1.03: 3; 1.04: 2.5367; 1.05: 2.9; 1.06: 3; 1.07: 0.6; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 1
| 1 | 2 | 1.01: < 3; 1.02: < 5; 1.03: 3; 1.04: 2.5915; 1.05: 2.8; 1.06: 3; 1.07: 2.1; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 14
| 1 | 3 | 1.01: < 3; 1.02: < 5; 1.03: 3; 1.04: 2.5805; 1.05: 2.8; 1.06: 3; 1.07: 1.8; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 16
| 1 | 4 | 1.01: < 3; 1.02: < 5; 1.03: 3; 1.04: 3.3068; 1.05: 2.7; 1.06: 3; 1.07: 2.7; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 2.9078
| 1 | 5 | 1.01: < 3; 1.02: < 5; 1.03: 3; 1.04: 3.0839; 1.05: 2.7; 1.06: 3; 1.07: 1.1; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 20
| 1 | 6 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 3.0881; 1.05: 2.7; 1.06: 3; 1.07: 2.2; 1.08: < 5; 1.09: < 5; 1.10: < 5
| Mean |  | 1.03: 2.8333; 1.04: 2.8646; 1.05: 2.7667; 1.06: 3; 1.07: 1.75
| Median |  | 1.03: 3; 1.04: 2.8377; 1.05: 2.75; 1.06: 3; 1.07: 1.95
| Std Dev. |  | 1.03: 0.40825; 1.04: 0.33358; 1.05: 0.08165; 1.06: 0; 1.07: 0.77136
| Rel.Std.Dev. |  | 1.03: 14.41%; 1.04: 11.64%; 1.05: 2.951%; 1.06: 0%; 1.07: 44.08%
| PDM3 |  | 1.03: -2.562%; 1.04: -1.486%; 1.05: -4.854%; 1.06: 3.17%; 1.07: -39.82%
| Z-Score (Absolute) |  | 1.01: 2.25; 1.02: 0; 1.03: 0.75; 1.04: 0.82; 1.05: 0.6; 1.06: 1.12; 1.07: 1.69; 1.08: 0; 1.09: 0; 1.10: 0
| Analytical results for Bi in OREAS 625 (Certified Value 8.61 ppm) | [24 empty] | Certified
Bi | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.16; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*MS |  |  |  | 2
3
| 1 | 1 | 1.01: 8; 1.02: 8; 1.03: 9.4; 1.04: 9.2756; 1.05: 8.9; 1.06: 100000; 1.07: 7; 1.08: 9.3 |  |  |  | 1
| 1 | 2 | 1.01: 8; 1.02: 8.2; 1.03: 9.4; 1.04: 9.2659; 1.05: 8.8; 1.06: 103000; 1.07: 7; 1.08: 9.4 |  |  |  | 35
| 1 | 3 | 1.01: 8; 1.02: 8; 1.03: 9.4; 1.04: 8.6065; 1.05: 9; 1.06: 99600; 1.07: 8; 1.08: 9.3 |  |  |  | 16
| 1 | 4 | 1.01: 8; 1.02: 8.1; 1.03: 9.2; 1.04: 8.5319; 1.05: 8.7; 1.06: 97200; 1.07: 8; 1.08: 9.7 |  |  |  | 8.6117
| 1 | 5 | 1.01: 8; 1.02: 8.3; 1.03: 9.3; 1.04: 9.1376; 1.05: 9; 1.06: 102000; 1.07: 8; 1.08: 9.2 |  |  |  | 123
| 1 | 6 | 1.01: 8; 1.02: 7.8; 1.03: 9.5; 1.04: 9.1723; 1.05: 8.9; 1.06: 95800; 1.07: 8; 1.08: 9.3
| Mean |  | 1.01: 8; 1.02: 8.0667; 1.03: 9.3667; 1.04: 8.9983; 1.05: 8.8833; 1.06: 99600; 1.07: 7.6667; 1.08: 9.3667
| Median |  | 1.01: 8; 1.02: 8.05; 1.03: 9.4; 1.04: 9.1549; 1.05: 8.9; 1.06: 99800; 1.07: 8; 1.08: 9.3
| Std Dev. |  | 1.01: 0; 1.02: 0.17512; 1.03: 0.10328; 1.04: 0.3374; 1.05: 0.1169; 1.06: 2745.1776; 1.07: 0.5164; 1.08: 0.17512
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 2.171%; 1.03: 1.103%; 1.04: 3.75%; 1.05: 1.316%; 1.06: 2.756%; 1.07: 6.736%; 1.08: 1.87%
| PDM3 |  | 1.01: -7.103%; 1.02: -6.329%; 1.03: 8.767%; 1.04: 4.49%; 1.05: 3.155%; 1.06: 1156471%; 1.07: -10.97%; 1.08: 8.767%
| Z-Score (Absolute) |  | 1.01: 1.23; 1.02: 1.14; 1.03: 0.67; 1.04: 0.16; 1.05: 0; 1.06: NA; 1.07: 1.7; 1.08: 0.67
| Analytical results for Ca in OREAS 625 (Certified Value 1.15 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 1.18 | 1.14 | 1.12 | 0.943 | 1.158 | 1.101 | 1.144 | 1.172 | 1.12 | 1.2 | 1.04 | 1.2 | 1.1848 | 1.13 | 1.16 | 1.286 | 1.1 | 1.14 | 1 | 1.18 |  |  |  | 1
| 1 | 2 | 1.18 | 1.16 | 0.96 | 0.879 | 1.201 | 1.072 | 1.122 | 1.029 | 1.13 | 1.3 | 1.12 | 1.2 | 1.1652 | 1.13 | 1.14 | 1.422 | 1.2 | 1.14 | 1.06 | 1.19 |  |  |  | 0
| 1 | 3 | 1.17 | 1.16 | 1.36 | 0.936 | 0.915 | 1.058 | 1.129 | 1.158 | 1.15 | 1.2 | 1.12 | 1.2 | 1.1265 | 1.14 | 1.15 | 1.5058 | 1.2 | 1.16 | 1.02 | 1.2 |  |  |  | 16
| 1 | 4 | 1.17 | 1.21 | 1 | 0.929 | 1.258 | 1.065 | 1.108 | 1.172 | 1.15 | 1.2 | 1.04 | 1.2 | 1.12 | 1.14 | 1.14 | 1.3913 | 1.2 | 1.14 | 1 | 1.22 |  |  |  | 1.1477
| 1 | 5 | 1.18 | 1.18 | 1.21 | 0.972 | 1.043 | 1.072 | 1.115 | 1.101 | 1.14 | 1.3 | 1.03 | 1.2 | 1.1753 | 1.14 | 1.15 | 1.2059 | 1.1 | 1.15 | 1.01 | 1.19 |  |  |  | 124
| 1 | 6 | 1.17 | 1.23 | 1.36 | 0.958 | 1.144 | 1.072 | 1.101 | 1.144 | 1.13 | 1.2 | 1.14 | 1.2 | 1.1893 | 1.12 | 1.15 | 1.1273 | 1.1 | 1.15 | 1 | 1.16
| Mean |  | 1.175 | 1.18 | 1.1683 | 0.93617 | 1.1198 | 1.0733 | 1.1198 | 1.1293 | 1.1367 | 1.2333 | 1.0817 | 1.2 | 1.1602 | 1.1333 | 1.1483 | 1.3231 | 1.15 | 1.1467 | 1.015 | 1.19
| Median |  | 1.175 | 1.17 | 1.165 | 0.9395 | 1.151 | 1.072 | 1.1185 | 1.151 | 1.135 | 1.2 | 1.08 | 1.2 | 1.1703 | 1.135 | 1.15 | 1.3386 | 1.15 | 1.145 | 1.005 | 1.19
| Std Dev. |  | 0.0054772 | 0.034059 | 0.17279 | 0.032021 | 0.12295 | 0.014665 | 0.015433 | 0.055763 | 0.012111 | 0.05164 | 0.049967 | 0 | 0.029856 | 0.008165 | 0.0075277 | 0.14231 | 0.054772 | 0.008165 | 0.023452 | 0.02
| Rel.Std.Dev. |  | 0.4661% | 2.886% | 14.79% | 3.42% | 10.98% | 1.366% | 1.378% | 4.938% | 1.065% | 4.187% | 4.619% | 0% | 2.573% | 0.7204% | 0.6555% | 10.76% | 4.763% | 0.7121% | 2.311% | 1.681%
| PDM3 |  | 2.381% | 2.816% | 1.8% | -18.43% | -2.426% | -6.478% | -2.426% | -1.598% | -0.9593% | 7.463% | -5.752% | 4.559% | 1.09% | -1.25% | 0.05722% | 15.28% | 0.2024% | -0.088% | -11.56% | 3.688%
| Z-Score (Absolute) |  | 0.67 | 0.79 | 0.51 | 5.12 | 0.67 | 1.8 | 0.67 | 0.44 | 0.26 | 2.08 | 1.6 | 1.27 | 0.31 | 0.34 | 0.02 | 4.26 | 0.06 | 0.02 | 3.21 | 1.03
| Analytical results for Cd in OREAS 625 (Certified Value 83 ppm) | [24 empty] | Certified
Cd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.09; 1.05: 1.10; 1.06: 1.15; 1.07: 1.16; 1.08: 1.23; 1.09: 1.25; 1.10: 1.26 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*OES; 1.10: PF*OES |  |  |  | 0
1
| 1 | 1 | 1.01: 84; 1.02: 87; 1.03: < 100; 1.04: 81; 1.05: 90; 1.06: 72.7143; 1.07: 83.7; 1.08: 82.4; 1.09: 76; 1.10: 60 |  |  |  | 1
| 1 | 2 | 1.01: 84; 1.02: 89; 1.03: < 100; 1.04: 84; 1.05: 87; 1.06: 78.3229; 1.07: 85.6; 1.08: 86; 1.09: 79; 1.10: 60 |  |  |  | 0
| 1 | 3 | 1.01: 85; 1.02: 83; 1.03: < 100; 1.04: 80; 1.05: 88; 1.06: 76.1377; 1.07: 84.9; 1.08: 83.4; 1.09: 78; 1.10: 60 |  |  |  | 16
| 1 | 4 | 1.01: 84; 1.02: 93; 1.03: < 100; 1.04: 80; 1.05: 85; 1.06: 73.3254; 1.07: 84; 1.08: 86; 1.09: 75; 1.10: 60 |  |  |  | 82.6872
| 1 | 5 | 1.01: 84; 1.02: 91; 1.03: < 100; 1.04: 85; 1.05: 89; 1.06: 75.6621; 1.07: 84.8; 1.08: 80.4; 1.09: 78; 1.10: 60 |  |  |  | 125
| 1 | 6 | 1.01: 85; 1.02: 79; 1.03: < 100; 1.04: 80; 1.05: 84; 1.06: 78.9235; 1.07: 85.6; 1.08: 81.1; 1.09: 79; 1.10: 60
| Mean |  | 1.01: 84.3333; 1.02: 87; 1.04: 81.6667; 1.05: 87.1667; 1.06: 75.8476; 1.07: 84.7667; 1.08: 83.2167; 1.09: 77.5; 1.10: 60
| Median |  | 1.01: 84; 1.02: 88; 1.04: 80.5; 1.05: 87.5; 1.06: 75.8999; 1.07: 84.85; 1.08: 82.9; 1.09: 78; 1.10: 60
| Std Dev. |  | 1.01: 0.5164; 1.02: 5.2154; 1.04: 2.2509; 1.05: 2.3166; 1.06: 2.5253; 1.07: 0.79162; 1.08: 2.392; 1.09: 1.6432; 1.10: 0
| Rel.Std.Dev. |  | 1.01: 0.6123%; 1.02: 5.995%; 1.04: 2.756%; 1.05: 2.658%; 1.06: 3.329%; 1.07: 0.9339%; 1.08: 2.874%; 1.09: 2.12%; 1.10: 0%
| PDM3 |  | 1.01: 1.991%; 1.02: 5.216%; 1.04: -1.234%; 1.05: 5.417%; 1.06: -8.272%; 1.07: 2.515%; 1.08: 0.6403%; 1.09: -6.273%; 1.10: -27.44%
| Z-Score (Absolute) |  | 1.01: 0.27; 1.02: 0.66; 1.03: 4.71; 1.04: 0.11; 1.05: 0.69; 1.06: 0.96; 1.07: 0.34; 1.08: 0.11; 1.09: 0.72; 1.10: 3.26
| Analytical results for Ce in OREAS 625 (Certified Value 86 ppm) | [24 empty] | Certified
Ce | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.23 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS |  |  |  | 0
1
| 1 | 1 | 1.01: 81.9; 1.02: 72.4; 1.03: 88.1; 1.04: 84.2693; 1.05: 87.134; 1.06: 120; 1.07: 84.8 |  |  |  | 1
| 1 | 2 | 1.01: 83.2; 1.02: 75.1; 1.03: 88.1; 1.04: 86.1651; 1.05: 86.389; 1.06: 127; 1.07: 89.5 |  |  |  | 36
| 1 | 3 | 1.01: 79.6; 1.02: 73.8; 1.03: 85.3; 1.04: 83.2524; 1.05: 86.748; 1.06: 117; 1.07: 86.7 |  |  |  | 16
| 1 | 4 | 1.01: 90.5; 1.02: 73; 1.03: 86.3; 1.04: 83.1799; 1.05: 86.853; 1.06: 117; 1.07: 91.9 |  |  |  | 86.0175
| 1 | 5 | 1.01: 85.7; 1.02: 76.3; 1.03: 88.1; 1.04: 84.5637; 1.05: 87.025; 1.06: 103; 1.07: 85.7 |  |  |  | 126
| 1 | 6 | 1.01: 84.1; 1.02: 71.9; 1.03: 86.1; 1.04: 86.6868; 1.05: 86.659; 1.06: 104; 1.07: 86
| Mean |  | 1.01: 84.1667; 1.02: 73.75; 1.03: 87; 1.04: 84.6862; 1.05: 86.8013; 1.06: 114.6667; 1.07: 87.4333
| Median |  | 1.01: 83.65; 1.02: 73.4; 1.03: 87.2; 1.04: 84.4165; 1.05: 86.8005; 1.06: 117; 1.07: 86.35
| Std Dev. |  | 1.01: 3.7254; 1.02: 1.6814; 1.03: 1.2506; 1.04: 1.4631; 1.05: 0.26702; 1.06: 9.395; 1.07: 2.7112
| Rel.Std.Dev. |  | 1.01: 4.426%; 1.02: 2.28%; 1.03: 1.437%; 1.04: 1.728%; 1.05: 0.3076%; 1.06: 8.193%; 1.07: 3.101%
| PDM3 |  | 1.01: -2.152%; 1.02: -14.26%; 1.03: 1.142%; 1.04: -1.548%; 1.05: 0.9112%; 1.06: 33.31%; 1.07: 1.646%
| Z-Score (Absolute) |  | 1.01: 0.84; 1.02: 4.16; 1.03: 0.06; 1.04: 0.67; 1.05: 0; 1.06: 8.88; 1.07: 0.2
| Analytical results for Co in OREAS 625 (Certified Value 4.5 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 7 | 5 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 4 | 4 | < 20 | < 20 | < 10 | 4.1835 | 4.5 | 5 | 6 | 4.2 | < 10 | < 10 | < 20 |  |  | 1
| 1 | 2 | 7 | 4.6 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 4 | 4 | < 20 | < 20 | < 10 | 4.074 | 4.7 | 6 | 11 | 4.2 | < 10 | 10 | < 20 |  |  | 37
| 1 | 3 | 7 | 4.9 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 4 | 4 | < 20 | < 20 | < 10 | 3.9258 | 4.4 | 6 | 59 | 4.3 | < 10 | 10 | < 20 |  |  | 16
| 1 | 4 | 7 | 5.2 | 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 4 | 4 | < 20 | < 20 | < 10 | 3.8105 | 4.6 | 6 | 76 | 4.3 | < 10 | 10 | < 20 |  |  | 4.5042
| 1 | 5 | 7 | 4.8 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 4 | 5 | < 20 | < 20 | < 10 | 3.8668 | 4.5 | 6 | 8 | 4 | < 10 | 10 | < 20 |  |  | 127
| 1 | 6 | 7 | 4.2 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 4 | 4 | < 20 | < 20 | < 10 | 4.2653 | 4.5 | 6 | 8 | 4.1 | < 10 | 10 | < 20
| Mean |  | 1.01: 7; 1.02: 4.7833; 1.03: 10; 1.09: 4; 1.10: 4.1667; 1.15: 4.021; 1.16: 4.5333; 1.17: 5.8333; 1.20: 28; 1.23: 4.1833; 1.26: 10
| Median |  | 1.01: 7; 1.02: 4.85; 1.03: 10; 1.09: 4; 1.10: 4; 1.15: 3.9999; 1.16: 4.5; 1.17: 6; 1.20: 9.5; 1.23: 4.2; 1.26: 10
| Std Dev. |  | 1.01: 0; 1.02: 0.34881; 1.09: 0; 1.10: 0.40825; 1.15: 0.18224; 1.16: 0.10328; 1.17: 0.40825; 1.20: 31.1063; 1.23: 0.1169; 1.26: 0
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 7.292%; 1.09: 0%; 1.10: 9.798%; 1.15: 4.532%; 1.16: 2.278%; 1.17: 6.999%; 1.20: 111.1%; 1.23: 2.795%; 1.26: 0%
| PDM3 |  | 1.01: 55.41%; 1.02: 6.197%; 1.03: 122%; 1.09: -11.19%; 1.10: -7.494%; 1.15: -10.73%; 1.16: 0.6469%; 1.17: 29.51%; 1.20: 521.6%; 1.23: -7.124%; 1.26: 122%
| Z-Score (Absolute) |  | NA | 4.22 | 3.37 | 0 | 0 | 0 | 0 | 0 | NA | NA | 0 | 0 | 4.05 | 4.84 | 4.42 | NA | 14.57 | 4.71 | 4.05 | 0.67 | 0
Results from laboratories 1.01, 1.09, 1.10 and 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 625 (Indicative Value 57 ppm) | [24 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.15 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [6 empty] | 0
0
| 1 | 1 | 50 | 106 | < 70 | 140 | < 70 | 70 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 16 | < 10 | 21 | 20 | < 100 | [6 empty] | 1
| 1 | 2 | 70 | < 100 | < 70 | 70 | < 70 | 210 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 17 | < 10 | 21 | 20 | < 100 | [6 empty] | 15
| 1 | 3 | 50 | < 100 | < 70 | 70 | < 70 | 340 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 20 | 15 | 22 | 30 | < 100 | [6 empty] | 16
| 1 | 4 | 50 | 116 | < 70 | 70 | < 70 | 140 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 29 | < 10 | 20 | 20 | < 100 | [6 empty] | 56.7123
| 1 | 5 | 50 | < 100 | < 70 | 70 | < 70 | 140 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 18 | 12 | 20 | 20 | < 100 | [6 empty] | 21
| 1 | 6 | 40 | < 100 | < 70 | 70 | < 70 | 70 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 22 | < 10 | 21 | 20 | < 100
| Mean |  | 1.01: 51.6667; 1.02: 111; 1.04: 81.6667; 1.06: 161.6667; 1.07: 70; 1.13: 20.3333; 1.15: 13.5; 1.16: 20.8333; 1.17: 21.6667
| Median |  | 1.01: 50; 1.02: 111; 1.04: 70; 1.06: 140; 1.07: 70; 1.13: 19; 1.15: 13.5; 1.16: 21; 1.17: 20
| Std Dev. |  | 1.01: 9.8319; 1.02: 7.0711; 1.04: 28.5774; 1.06: 101.8659; 1.07: 0; 1.13: 4.761; 1.15: 2.1213; 1.16: 0.75277; 1.17: 4.0825
| Rel.Std.Dev. |  | 1.01: 19.03%; 1.02: 6.37%; 1.04: 34.99%; 1.06: 63.01%; 1.07: 0%; 1.13: 23.41%; 1.15: 15.71%; 1.16: 3.613%; 1.17: 18.84%
| PDM3 |  | 1.01: -8.897%; 1.02: 95.72%; 1.04: 44%; 1.06: 185.1%; 1.07: 23.43%; 1.13: -64.15%; 1.15: -76.2%; 1.16: -63.26%; 1.17: -61.8%
| Z-Score (Absolute) |  | 0.75 | 1.54 | 0 | 1.95 | 0 | 5.36 | 1.46 | 1.03 | 0.67 | 0.39 | 0.39 | 0.39 | 0.59 | 1.13 | 0.57 | 0.54 | 0.67
| Analytical results for Cs in OREAS 625 (Certified Value 6 ppm) | [24 empty] | Certified
Cs | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.15; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 5.8; 1.02: 5.2; 1.03: 6.4; 1.04: 5.8268; 1.05: 6.062; 1.06: 6.4; 1.07: 7; 1.08: 5.8 |  |  | 1
| 1 | 2 | 1.01: 5.6; 1.02: 5.4; 1.03: 6.4; 1.04: 5.8745; 1.05: 5.9348; 1.06: 6.5; 1.07: 7; 1.08: 6 |  |  | 39
| 1 | 3 | 1.01: 6.6; 1.02: 5.1; 1.03: 6.2; 1.04: 5.7725; 1.05: 6.174; 1.06: 6.6; 1.07: 7; 1.08: 5.8 |  |  | 16
| 1 | 4 | 1.01: 6.1; 1.02: 5.2; 1.03: 6.5; 1.04: 5.5812; 1.05: 5.6393; 1.06: 6.5; 1.07: 7; 1.08: 6.1 |  |  | 5.9997
| 1 | 5 | 1.01: 6.2; 1.02: 5.2; 1.03: 6.4; 1.04: 5.9139; 1.05: 5.6967; 1.06: 6.5; 1.07: 7; 1.08: 5.6 |  |  | 128
| 1 | 6 | 1.01: 6; 1.02: 5.1; 1.03: 6.3; 1.04: 5.8568; 1.05: 5.8922; 1.06: 6.5; 1.07: 7; 1.08: 5.7
| Mean |  | 1.01: 6.05; 1.02: 5.2; 1.03: 6.3667; 1.04: 5.8043; 1.05: 5.8998; 1.06: 6.5; 1.07: 7; 1.08: 5.8333
| Median |  | 1.01: 6.05; 1.02: 5.2; 1.03: 6.4; 1.04: 5.8418; 1.05: 5.9135; 1.06: 6.5; 1.07: 7; 1.08: 5.8
| Std Dev. |  | 1.01: 0.34496; 1.02: 0.10954; 1.03: 0.10328; 1.04: 0.11915; 1.05: 0.20585; 1.06: 0.063246; 1.07: 0; 1.08: 0.18619
| Rel.Std.Dev. |  | 1.01: 5.702%; 1.02: 2.107%; 1.03: 1.622%; 1.04: 2.053%; 1.05: 3.489%; 1.06: 0.973%; 1.07: 0%; 1.08: 3.192%
| PDM3 |  | 1.01: 0.8377%; 1.02: -13.33%; 1.03: 6.116%; 1.04: -3.258%; 1.05: -1.666%; 1.06: 8.338%; 1.07: 16.67%; 1.08: -2.774%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 3.14; 1.03: 2.1; 1.04: 0.43; 1.05: 0; 1.06: 2.69; 1.07: NA; 1.08: 0.3
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 625 (Certified Value 0.17 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.1745 | 0.18 | 0.1736 | 0.165 | 0.169 | 0.173 | 0.164 | 0.164 | 0.17 | 0.1635 | 0.17 | 0.1691 | 0.17964 | 0.16704 | 0.174 | 0.1623 | 0.1398 | 0.1671 | 0.1623 | 0.172 | 0.171 | 0.17 |  | 1
| 1 | 2 | 0.1717 | 0.183 | 0.1678 | 0.165 | 0.175 | 0.172 | 0.157 | 0.161 | 0.17 | 0.1688 | 0.17 | 0.165 | 0.17951 | 0.16933 | 0.17 | 0.1685 | 0.134 | 0.1703 | 0.1681 | 0.178 | 0.17 | 0.18 |  | 25
| 1 | 3 | 0.1758 | 0.199 | 0.1735 | 0.163 | 0.17 | 0.17 | 0.164 | 0.166 | 0.167 | 0.1646 | 0.173 | 0.1646 | 0.1779 | 0.16747 | 0.173 | 0.1688 | 0.1424 | 0.1696 | 0.168 | 0.175 | 0.172 | 0.18 |  | 16
| 1 | 4 | 0.1773 | 0.181 | 0.173 | 0.164 | 0.173 | 0.17 | 0.168 | 0.166 | 0.171 | 0.1677 | 0.173 | 0.1653 | 0.18013 | 0.1679 | 0.171 | 0.1696 | 0.1277 | 0.1681 | 0.1612 | 0.175 | 0.174 | 0.17 |  | 0.17029
| 1 | 5 | 0.1764 | 0.178 | 0.1705 | 0.163 | 0.169 | 0.169 | 0.173 | 0.164 | 0.171 | 0.1676 | 0.172 | 0.1665 | 0.17673 | 0.16707 | 0.171 | 0.1989 | 0.1379 | 0.166 | 0.1686 | 0.173 | 0.172 | 0.17 |  | 129
| 1 | 6 | 0.1777 | 0.17 | 0.1791 | 0.167 | 0.171 | 0.17 | 0.167 | 0.164 | 0.168 | 0.169 | 0.173 | 0.169 | 0.17713 | 0.16659 | 0.174 | 0.1681 | 0.1428 | 0.1669 | 0.1667 | 0.17 | 0.172 | 0.17
| Mean |  | 0.17557 | 0.18183 | 0.17292 | 0.1645 | 0.17117 | 0.17067 | 0.1655 | 0.16417 | 0.1695 | 0.16687 | 0.17183 | 0.16658 | 0.17851 | 0.16757 | 0.17217 | 0.1727 | 0.13743 | 0.168 | 0.16582 | 0.17383 | 0.17183 | 0.17333
| Median |  | 0.1761 | 0.1805 | 0.17325 | 0.1645 | 0.1705 | 0.17 | 0.1655 | 0.164 | 0.17 | 0.16765 | 0.1725 | 0.1659 | 0.1787 | 0.16727 | 0.172 | 0.16865 | 0.13885 | 0.1676 | 0.16735 | 0.174 | 0.172 | 0.17
| Std Dev. |  | 0.0022088 | 0.0095376 | 0.0037701 | 0.0015166 | 0.0024014 | 0.0015055 | 0.0053198 | 0.0018348 | 0.0016432 | 0.0022801 | 0.001472 | 0.0020134 | 0.0014391 | 0.00096937 | 0.0017224 | 0.013101 | 0.0057587 | 0.0016661 | 0.0032307 | 0.0027869 | 0.0013292 | 0.005164
| Rel.Std.Dev. |  | 1.258% | 5.245% | 2.18% | 0.9219% | 1.403% | 0.8822% | 3.214% | 1.118% | 0.9694% | 1.366% | 0.8566% | 1.209% | 0.8062% | 0.5785% | 1% | 7.586% | 4.19% | 0.9917% | 1.948% | 1.603% | 0.7735% | 2.979%
| PDM3 |  | 3.096% | 6.776% | 1.54% | -3.403% | 0.5122% | 0.2186% | -2.815% | -3.598% | -0.4665% | -2.013% | 0.9037% | -2.179% | 4.822% | -1.603% | 1.099% | 1.413% | -19.3% | -1.347% | -2.629% | 2.078% | 0.9037% | 1.785%
| Z-Score (Absolute) |  | 1 | 2.35 | 0.43 | 1.38 | 0.05 | 0.05 | 1.17 | 1.45 | 0.3 | 0.87 | 0.2 | 0.93 | 1.63 | 0.72 | 0.27 | 0.38 | 7.2 | 0.63 | 1.1 | 0.63 | 0.2 | 0.52
| Analytical results for Dy in OREAS 625 (Certified Value 3.66 ppm) | [24 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 2
3
| 1 | 1 | 1.01: 3.4; 1.02: 3.3; 1.03: 4; 1.04: 4.0803; 1.05: 3.467; 1.06: 7; 1.07: 3.7521; 1.08: 3.95 |  | 1
| 1 | 2 | 1.01: 3.3; 1.02: 3.55; 1.03: 4.1; 1.04: 3.7266; 1.05: 3.434; 1.06: 6; 1.07: 3.4657; 1.08: 4.16 |  | 7
| 1 | 3 | 1.01: 3.5; 1.02: 3.41; 1.03: 3.7; 1.04: 3.3476; 1.05: 3.459; 1.06: 7; 1.07: 3.5923; 1.08: 3.87 |  | 16
| 1 | 4 | 1.01: 3.5; 1.02: 3.5; 1.03: 3.9; 1.04: 3.4895; 1.05: 3.453; 1.06: 6; 1.07: 3.399; 1.08: 4.28 |  | 3.6586
| 1 | 5 | 1.01: 3.5; 1.02: 3.5; 1.03: 4.1; 1.04: 3.2545; 1.05: 3.461; 1.06: 5; 1.07: 3.7386; 1.08: 4.12 |  | 130
| 1 | 6 | 1.01: 3.5; 1.02: 3.42; 1.03: 3.9; 1.04: 3.6694; 1.05: 3.447; 1.06: 5; 1.07: 3.8445; 1.08: 4.12
| Mean |  | 1.01: 3.45; 1.02: 3.4467; 1.03: 3.95; 1.04: 3.5947; 1.05: 3.4535; 1.06: 6; 1.07: 3.632; 1.08: 4.0833
| Median |  | 1.01: 3.5; 1.02: 3.46; 1.03: 3.95; 1.04: 3.5794; 1.05: 3.456; 1.06: 6; 1.07: 3.6655; 1.08: 4.12
| Std Dev. |  | 1.01: 0.083666; 1.02: 0.089368; 1.03: 0.15166; 1.04: 0.2989; 1.05: 0.01176; 1.06: 0.89443; 1.07: 0.17575; 1.08: 0.14868
| Rel.Std.Dev. |  | 1.01: 2.425%; 1.02: 2.593%; 1.03: 3.839%; 1.04: 8.315%; 1.05: 0.3405%; 1.06: 14.91%; 1.07: 4.839%; 1.08: 3.641%
| PDM3 |  | 1.01: -5.702%; 1.02: -5.793%; 1.03: 7.965%; 1.04: -1.748%; 1.05: -5.606%; 1.06: 64%; 1.07: -0.7262%; 1.08: 11.61%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.69; 1.03: 1.66; 1.04: 0; 1.05: 0.66; 1.06: NA; 1.07: 0.17; 1.08: 2.28
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 625 (Certified Value 1.41 ppm) | [24 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 2
3
| 1 | 1 | 1.01: 1.4; 1.02: 1.29; 1.03: 1.3; 1.04: 1.4817; 1.05: 1.434; 1.06: 3; 1.07: 1.4632; 1.08: 1.42 |  | 1
| 1 | 2 | 1.01: 1.4; 1.02: 1.3; 1.03: 1.6; 1.04: 1.4319; 1.05: 1.394; 1.06: 2; 1.07: 1.6879; 1.08: 1.5 |  | 8
| 1 | 3 | 1.01: 1.4; 1.02: 1.25; 1.03: 1.5; 1.04: 1.3428; 1.05: 1.421; 1.06: 3; 1.07: 1.4189; 1.08: 1.37 |  | 16
| 1 | 4 | 1.01: 1.2; 1.02: 1.25; 1.03: 1.5; 1.04: 1.4834; 1.05: 1.428; 1.06: 3; 1.07: 1.325; 1.08: 1.51 |  | 1.4143
| 1 | 5 | 1.01: 1.5; 1.02: 1.28; 1.03: 1.5; 1.04: 1.4804; 1.05: 1.431; 1.06: < 2; 1.07: 1.4116; 1.08: 1.5 |  | 131
| 1 | 6 | 1.01: 1.3; 1.02: 1.29; 1.03: 1.4; 1.04: 1.4793; 1.05: 1.427; 1.06: < 2; 1.07: 1.4886; 1.08: 1.55
| Mean |  | 1.01: 1.3667; 1.02: 1.2767; 1.03: 1.4667; 1.04: 1.4499; 1.05: 1.4225; 1.06: 2.75; 1.07: 1.4658; 1.08: 1.475
| Median |  | 1.01: 1.4; 1.02: 1.285; 1.03: 1.5; 1.04: 1.4799; 1.05: 1.4275; 1.06: 3; 1.07: 1.441; 1.08: 1.5
| Std Dev. |  | 1.01: 0.10328; 1.02: 0.021602; 1.03: 0.10328; 1.04: 0.056081; 1.05: 0.014625; 1.06: 0.5; 1.07: 0.12231; 1.08: 0.066558
| Rel.Std.Dev. |  | 1.01: 7.557%; 1.02: 1.692%; 1.03: 7.042%; 1.04: 3.868%; 1.05: 1.028%; 1.06: 18.18%; 1.07: 8.344%; 1.08: 4.512%
| PDM3 |  | 1.01: -3.37%; 1.02: -9.733%; 1.03: 3.701%; 1.04: 2.516%; 1.05: 0.578%; 1.06: 94.44%; 1.07: 3.643%; 1.08: 4.29%
| Z-Score (Absolute) |  | 1.01: 2.24; 1.02: 4.66; 1.03: 0.45; 1.04: 0; 1.05: 0.74; 1.06: NA; 1.07: 0.43; 1.08: 0.67
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 625 (Certified Value 1.58 ppm) | [24 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS |  | 2
2
| 1 | 1 | 1.01: 2.1; 1.02: 1.26; 1.03: 1.3; 1.04: 1.6703; 1.05: 1.581; 1.06: < 2; 1.07: 1.42 |  | 1
| 1 | 2 | 1.01: 2.1; 1.02: 1.3; 1.03: 1.3; 1.04: 1.7264; 1.05: 1.565; 1.06: < 2; 1.07: 1.48 |  | 9
| 1 | 3 | 1.01: 2; 1.02: 1.28; 1.03: 1.4; 1.04: 1.6412; 1.05: 1.572; 1.06: 2; 1.07: 1.44 |  | 16
| 1 | 4 | 1.01: 2.1; 1.02: 1.22; 1.03: 1.3; 1.04: 1.734; 1.05: 1.534; 1.06: 2; 1.07: 1.5 |  | 1.5774
| 1 | 5 | 1.01: 2.3; 1.02: 1.3; 1.03: 1.3; 1.04: 1.7578; 1.05: 1.566; 1.06: 2; 1.07: 1.38 |  | 132
| 1 | 6 | 1.01: 2.3; 1.02: 1.23; 1.03: 1.4; 1.04: 1.7212; 1.05: 1.577; 1.06: 2; 1.07: 1.43
| Mean |  | 1.01: 2.15; 1.02: 1.265; 1.03: 1.3333; 1.04: 1.7085; 1.05: 1.5658; 1.06: 2; 1.07: 1.4417
| Median |  | 1.01: 2.1; 1.02: 1.27; 1.03: 1.3; 1.04: 1.7238; 1.05: 1.569; 1.06: 2; 1.07: 1.435
| Std Dev. |  | 1.01: 0.12247; 1.02: 0.034496; 1.03: 0.05164; 1.04: 0.043697; 1.05: 0.016774; 1.06: 0; 1.07: 0.043089
| Rel.Std.Dev. |  | 1.01: 5.696%; 1.02: 2.727%; 1.03: 3.873%; 1.04: 2.558%; 1.05: 1.071%; 1.06: 0%; 1.07: 2.989%
| PDM3 |  | 1.01: 36.3%; 1.02: -19.8%; 1.03: -15.47%; 1.04: 8.31%; 1.05: -0.7323%; 1.06: 26.79%; 1.07: -8.604%
| Z-Score (Absolute) |  | 1.01: 2.32; 1.02: 0.86; 1.03: 0.61; 1.04: 0.74; 1.05: 0.22; 1.06: NA; 1.07: 0.22
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Fe in OREAS 625 (Certified Value 4.01 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 4.05 | 4.02 | 3.51 | 3.91 | 3.917 | 4.001 | 3.91 | 3.938 | 4.25 | 3.99 | 3.87 | 3.92 | 3.9023 | 4.1331 | 4.01 | 4.09 | 4.5024 | 4.1793 | 3.97 | 4.03 | 3.9 | 4.02 | 4.29 | 1
| 1 | 2 | 4.03 | 4.01 | 3.24 | 3.924 | 3.98 | 3.973 | 3.882 | 3.84 | 4.22 | 4.09 | 3.9 | 3.82 | 3.9266 | 4.0921 | 4.01 | 4.02 | 4.6568 | 4.15 | 4 | 4.08 | 3.73 | 4.03 | 4.27 | 26
| 1 | 3 | 4.08 | 4.06 | 3.39 | 3.903 | 3.861 | 4.008 | 3.847 | 3.931 | 4.19 | 3.95 | 3.97 | 3.81 | 3.8908 | 3.9537 | 4.05 | 3.98 | 4.5387 | 4.1327 | 4 | 4.01 | 3.94 | 4.07 | 4.28 | 16
| 1 | 4 | 4.09 | 4.09 | 3.44 | 3.903 | 4.057 | 4.008 | 4.071 | 3.98 | 4.23 | 3.94 | 3.92 | 3.85 | 3.9489 | 3.9924 | 4.06 | 4.03 | 4.5193 | 4.1139 | 3.99 | 4 | 3.93 | 4.12 | 4.28 | 4.0112
| 1 | 5 | 4.08 | 4.05 | 3.32 | 3.903 | 3.868 | 3.882 | 3.973 | 3.917 | 4.26 | 4.03 | 3.87 | 3.87 | 3.9715 | 4.0839 | 4.06 | 4.02 | 4.586 | 4.1209 | 3.91 | 4.06 | 3.93 | 4.01 | 4.15 | 133
| 1 | 6 | 4.08 | 4.06 | 3.56 | 3.959 | 3.917 | 3.98 | 3.91 | 3.952 | 4.21 | 3.96 | 3.89 | 3.93 | 3.8857 | 4.1682 | 4.01 | 4.03 | 4.4253 | 4.1639 | 3.98 | 4.13 | 3.91 | 4.09 | 4.19
| Mean |  | 4.0683 | 4.0483 | 3.41 | 3.917 | 3.9333 | 3.9753 | 3.9322 | 3.9263 | 4.2267 | 3.9933 | 3.9033 | 3.8667 | 3.921 | 4.0706 | 4.0333 | 4.0283 | 4.5381 | 4.1435 | 3.975 | 4.0517 | 3.89 | 4.0567 | 4.2433
| Median |  | 4.08 | 4.055 | 3.415 | 3.9065 | 3.917 | 3.9905 | 3.91 | 3.9345 | 4.225 | 3.975 | 3.895 | 3.86 | 3.9145 | 4.088 | 4.03 | 4.025 | 4.529 | 4.1414 | 3.985 | 4.045 | 3.92 | 4.05 | 4.275
| Std Dev. |  | 0.023166 | 0.029269 | 0.119 | 0.022136 | 0.074169 | 0.048024 | 0.079618 | 0.047407 | 0.02582 | 0.057504 | 0.037771 | 0.050067 | 0.034264 | 0.082281 | 0.02582 | 0.035449 | 0.078404 | 0.025469 | 0.033912 | 0.048751 | 0.07975 | 0.043665 | 0.058538
| Rel.Std.Dev. |  | 0.5694% | 0.723% | 3.49% | 0.5651% | 1.886% | 1.208% | 2.025% | 1.207% | 0.6109% | 1.44% | 0.9677% | 1.295% | 0.8739% | 2.021% | 0.6402% | 0.88% | 1.728% | 0.6147% | 0.8531% | 1.203% | 2.05% | 1.076% | 1.38%
| PDM3 |  | 1.424% | 0.9252% | -14.99% | -2.349% | -1.942% | -0.8947% | -1.971% | -2.116% | 5.371% | -0.4459% | -2.69% | -3.604% | -2.25% | 1.48% | 0.5513% | 0.4266% | 13.13% | 3.297% | -0.903% | 1.008% | -3.022% | 1.133% | 5.787%
| Z-Score (Absolute) |  | 0.7 | 0.51 | 5.44 | 0.71 | 0.56 | 0.17 | 0.57 | 0.63 | 2.17 | 0 | 0.84 | 1.18 | 0.67 | 0.72 | 0.37 | 0.33 | 5.08 | 1.4 | 0.17 | 0.54 | 0.96 | 0.59 | 2.33
| Analytical results for Ga in OREAS 625 (Certified Value 20.4 ppm) | [24 empty] | Certified
Ga | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.16; 1.08: 1.17; 1.09: 1.20; 1.10: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*OES/MS; 1.09: PF*MS; 1.10: PF*MS | 1
2
| 1 | 1 | 1.01: 22; 1.02: 22.2; 1.03: 18.3; 1.04: 19; 1.05: 20.7654; 1.06: 21.6027; 1.07: 112.6; 1.08: 21; 1.09: 19; 1.10: 20 | 1
| 1 | 2 | 1.01: 20; 1.02: 19; 1.03: 19.4; 1.04: 19; 1.05: 21.6685; 1.06: 21.4383; 1.07: 114.1; 1.08: 22.1; 1.09: 19; 1.10: 20 | 44
| 1 | 3 | 1.01: 21; 1.02: 23.5; 1.03: 18.5; 1.04: 18; 1.05: 21.4087; 1.06: 20.038; 1.07: 112.5; 1.08: 21; 1.09: 19; 1.10: 21 | 16
| 1 | 4 | 1.01: 21; 1.02: 18.9; 1.03: 18.2; 1.04: 19; 1.05: 21.0693; 1.06: 20.2914; 1.07: 112.4; 1.08: 22.9; 1.09: 18; 1.10: 21 | 20.4465
| 1 | 5 | 1.01: 21; 1.02: 19.8; 1.03: 18.7; 1.04: 18; 1.05: 21.3556; 1.06: 21.1132; 1.07: 111.6; 1.08: 21.4; 1.09: 19; 1.10: 20 | 134
| 1 | 6 | 1.01: 20; 1.02: 21.9; 1.03: 18.7; 1.04: 17; 1.05: 21.6429; 1.06: 22.2373; 1.07: 113.4; 1.08: 21.3; 1.09: 19; 1.10: 20
| Mean |  | 1.01: 20.8333; 1.02: 20.8833; 1.03: 18.6333; 1.04: 18.3333; 1.05: 21.3184; 1.06: 21.1202; 1.07: 112.7667; 1.08: 21.6167; 1.09: 18.8333; 1.10: 20.3333
| Median |  | 1.01: 21; 1.02: 20.85; 1.03: 18.6; 1.04: 18.5; 1.05: 21.3821; 1.06: 21.2758; 1.07: 112.55; 1.08: 21.35; 1.09: 19; 1.10: 20
| Std Dev. |  | 1.01: 0.75277; 1.02: 1.9115; 1.03: 0.4274; 1.04: 0.8165; 1.05: 0.34793; 1.06: 0.82945; 1.07: 0.86872; 1.08: 0.74677; 1.09: 0.40825; 1.10: 0.5164
| Rel.Std.Dev. |  | 1.01: 3.613%; 1.02: 9.153%; 1.03: 2.294%; 1.04: 4.454%; 1.05: 1.632%; 1.06: 3.927%; 1.07: 0.7704%; 1.08: 3.455%; 1.09: 2.168%; 1.10: 2.54%
| PDM3 |  | 1.01: 1.892%; 1.02: 2.137%; 1.03: -8.868%; 1.04: -10.34%; 1.05: 4.264%; 1.06: 3.295%; 1.07: 451.5%; 1.08: 5.723%; 1.09: -7.89%; 1.10: -0.5534%
| Z-Score (Absolute) |  | 1.01: 0.03; 1.02: 0.03; 1.03: 2.34; 1.04: 2.65; 1.05: 0.48; 1.06: 0.28; 1.07: 96.58; 1.08: 0.8; 1.09: 2.13; 1.10: 0.55
| Analytical results for Gd in OREAS 625 (Certified Value 5.53 ppm) | [24 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | 2
3
| 1 | 1 | 1.01: 5.9; 1.02: 5.31; 1.03: 5.8; 1.04: 5.3764; 1.05: 5.39; 1.06: 6; 1.07: 6.7732; 1.08: 5.7 | 1
| 1 | 2 | 1.01: 7; 1.02: 5.46; 1.03: 5.3; 1.04: 5.4304; 1.05: 5.409; 1.06: 5; 1.07: 6.4034; 1.08: 5.94 | 10
| 1 | 3 | 1.01: 6.8; 1.02: 5.23; 1.03: 5.3; 1.04: 5.5185; 1.05: 5.395; 1.06: 6; 1.07: 6.6535; 1.08: 5.89 | 16
| 1 | 4 | 1.01: 5.7; 1.02: 5.36; 1.03: 5.7; 1.04: 5.1354; 1.05: 5.397; 1.06: 6; 1.07: 6.1823; 1.08: 6.18 | 5.5297
| 1 | 5 | 1.01: 5.8; 1.02: 5.6; 1.03: 5.8; 1.04: 5.9239; 1.05: 5.39; 1.06: 5; 1.07: 6.2133; 1.08: 5.88 | 135
| 1 | 6 | 1.01: 7.2; 1.02: 5.17; 1.03: 5.3; 1.04: 5.6309; 1.05: 5.392; 1.06: 5; 1.07: 6.6145; 1.08: 5.9
| Mean |  | 1.01: 6.4; 1.02: 5.355; 1.03: 5.5333; 1.04: 5.5026; 1.05: 5.3955; 1.06: 5.5; 1.07: 6.4734; 1.08: 5.915
| Median |  | 1.01: 6.35; 1.02: 5.335; 1.03: 5.5; 1.04: 5.4744; 1.05: 5.3935; 1.06: 5.5; 1.07: 6.509; 1.08: 5.895
| Std Dev. |  | 1.01: 0.67231; 1.02: 0.15681; 1.03: 0.2582; 1.04: 0.26463; 1.05: 0.0071764; 1.06: 0.54772; 1.07: 0.24475; 1.08: 0.15437
| Rel.Std.Dev. |  | 1.01: 10.5%; 1.02: 2.928%; 1.03: 4.666%; 1.04: 4.809%; 1.05: 0.133%; 1.06: 9.959%; 1.07: 3.781%; 1.08: 2.61%
| PDM3 |  | 1.01: 15.74%; 1.02: -3.159%; 1.03: 0.06598%; 1.04: -0.49%; 1.05: -2.427%; 1.06: -0.5368%; 1.07: 17.07%; 1.08: 6.968%
| Z-Score (Absolute) |  | 1.01: 3.28; 1.02: 0.67; 1.03: 0; 1.04: 0.12; 1.05: 0.52; 1.06: NA; 1.07: 3.55; 1.08: 1.44
| Analytical results for Ge in OREAS 625 (Certified Value 1.31 ppm) | [24 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.17; 1.06: 1.20; 1.07: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS | 2
2
| 1 | 1 | 1.01: 1.8; 1.02: 1; 1.03: 1; 1.04: 1.3306; 1.05: 2.6; 1.06: < 2; 1.07: 1 | 1
| 1 | 2 | 1.01: 1.5; 1.02: 2; 1.03: 1; 1.04: < 1; 1.05: 2.5; 1.06: < 2; 1.07: 1 | 46
| 1 | 3 | 1.01: 1.7; 1.02: 2; 1.03: 1; 1.04: 1.3174; 1.05: 2.2; 1.06: < 2; 1.07: 1 | 16
| 1 | 4 | 1.01: 1.7; 1.02: 2; 1.03: 1; 1.04: < 1; 1.05: 2.9; 1.06: < 2; 1.07: 1 | 1.3052
| 1 | 5 | 1.01: 1.9; 1.02: 2; 1.03: 1; 1.04: 1.2759; 1.05: 2.7; 1.06: < 2; 1.07: 1 | 136
| 1 | 6 | 1.01: 1.3; 1.02: 2; 1.03: 1; 1.04: 1.331; 1.05: 2.4; 1.06: < 2; 1.07: 1
| Mean |  | 1.01: 1.65; 1.02: 1.8333; 1.03: 1; 1.04: 1.3137; 1.05: 2.55; 1.07: 1
| Median |  | 1.01: 1.7; 1.02: 2; 1.03: 1; 1.04: 1.324; 1.05: 2.55; 1.07: 1
| Std Dev. |  | 1.01: 0.21679; 1.02: 0.40825; 1.03: 0; 1.04: 0.025979; 1.05: 0.2429; 1.07: 0
| Rel.Std.Dev. |  | 1.01: 13.14%; 1.02: 22.27%; 1.03: 0%; 1.04: 1.977%; 1.05: 9.525%; 1.07: 0%
| PDM3 |  | 1.01: 26.42%; 1.02: 40.47%; 1.03: -23.38%; 1.04: 0.6558%; 1.05: 95.38%; 1.07: -23.38%
| Z-Score (Absolute) |  | 1.01: 9.64; 1.02: 12.55; 1.03: 0.67; 1.04: 0; 1.05: 23.93; 1.06: 0.67; 1.07: 0.67
| Analytical results for Hf in OREAS 625 (Indicative Value 6.29 ppm) | [24 empty] | Indicative
Hf | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.20; 1.05: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*MS | 2
2
| 1 | 1 | 1.01: < 10; 1.02: 6.2; 1.03: 6.9336; 1.04: 6; 1.05: 6 | 1
| 1 | 2 | 1.01: < 10; 1.02: 6.6; 1.03: 6.2702; 1.04: 6; 1.05: 6 | 16
| 1 | 3 | 1.01: < 10; 1.02: 6.2; 1.03: 7.1588; 1.04: 6; 1.05: 6 | 16
| 1 | 4 | 1.01: 20; 1.02: 6.6; 1.03: 6.6137; 1.04: 6; 1.05: 6 | 6.292
| 1 | 5 | 1.01: 10; 1.02: 6.6; 1.03: 6.603; 1.04: 6; 1.05: 6 | 22
| 1 | 6 | 1.01: < 10; 1.02: 6.5; 1.03: 6.728; 1.04: 6; 1.05: 6
| Mean |  | 1.01: 15; 1.02: 6.45; 1.03: 6.7179; 1.04: 6; 1.05: 6
| Median |  | 1.01: 15; 1.02: 6.55; 1.03: 6.6709; 1.04: 6; 1.05: 6
| Std Dev. |  | 1.01: 7.0711; 1.02: 0.19748; 1.03: 0.30514; 1.04: 0; 1.05: 0
| Rel.Std.Dev. |  | 1.01: 47.14%; 1.02: 3.062%; 1.03: 4.542%; 1.04: 0%; 1.05: 0%
| PDM3 |  | 1.01: 138.4%; 1.02: 2.512%; 1.03: 6.769%; 1.04: -4.64%; 1.05: -4.64%
| Z-Score (Absolute) |  | 1.01: 2.82; 1.02: 0; 1.03: 0.4; 1.04: 0.67; 1.05: 0.67
| Analytical results for Hg in OREAS 625 (Indicative Value < 5 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.16 | ppm
|  |  | 1.01: PF*OES | 2
2
| 1 | 1 | 1.01: < 5 | 1
| 1 | 2 | 1.01: < 5 | 4
| 1 | 3 | 1.01: < 5 | 16
| 1 | 4 | 1.01: < 5 | < 5
| 1 | 5 | 1.01: < 5 | 18
| 1 | 6 | 1.01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ho in OREAS 625 (Certified Value 0.6 ppm) | [24 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | 2
3
| 1 | 1 | 1.01: 0.6; 1.02: 0.52; 1.03: 0.6; 1.04: 0.61204; 1.05: 0.59; 1.06: < 2; 1.07: 0.56877; 1.08: 0.6 | 1
| 1 | 2 | 1.01: 0.6; 1.02: 0.54; 1.03: 0.6; 1.04: 0.60394; 1.05: 0.573; 1.06: < 2; 1.07: 0.63092; 1.08: 0.64 | 7
| 1 | 3 | 1.01: 0.6; 1.02: 0.52; 1.03: 0.6; 1.04: 0.61117; 1.05: 0.583; 1.06: < 2; 1.07: 0.62547; 1.08: 0.63 | 16
| 1 | 4 | 1.01: 0.6; 1.02: 0.52; 1.03: 0.6; 1.04: 0.60189; 1.05: 0.569; 1.06: < 2; 1.07: 0.5553; 1.08: 0.65 | 0.60395
| 1 | 5 | 1.01: 0.7; 1.02: 0.54; 1.03: 0.6; 1.04: 0.67451; 1.05: 0.584; 1.06: < 2; 1.07: 0.56849; 1.08: 0.64 | 137
| 1 | 6 | 1.01: 0.6; 1.02: 0.51; 1.03: 0.6; 1.04: 0.60627; 1.05: 0.587; 1.06: < 2; 1.07: 0.61285; 1.08: 0.65
| Mean |  | 1.01: 0.61667; 1.02: 0.525; 1.03: 0.6; 1.04: 0.6183; 1.05: 0.581; 1.07: 0.59363; 1.08: 0.635
| Median |  | 1.01: 0.6; 1.02: 0.52; 1.03: 0.6; 1.04: 0.60872; 1.05: 0.5835; 1.07: 0.59081; 1.08: 0.64
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.012247; 1.03: 0; 1.04: 0.02782; 1.05: 0.0082219; 1.07: 0.033146; 1.08: 0.018708
| Rel.Std.Dev. |  | 1.01: 6.62%; 1.02: 2.333%; 1.03: 0%; 1.04: 4.499%; 1.05: 1.415%; 1.07: 5.584%; 1.08: 2.946%
| PDM3 |  | 1.01: 2.106%; 1.02: -13.07%; 1.03: -0.6539%; 1.04: 2.377%; 1.05: -3.8%; 1.07: -1.708%; 1.08: 5.141%
| Z-Score (Absolute) |  | 1.01: 0.27; 1.02: 2.72; 1.03: 0.27; 1.04: 0.33; 1.05: 0.89; 1.06: 12.77; 1.07: 0.48; 1.08: 0.87
| Analytical results for In in OREAS 625 (Certified Value 2.11 ppm) | [24 empty] | Certified
In | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.17; 1.06: 1.20; 1.07: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS | 2
3
| 1 | 1 | 1.01: 1.9; 1.02: 1.9; 1.03: 2.1; 1.04: 2.0672; 1.05: 2.4; 1.06: 2.3; 1.07: 2 | 1
| 1 | 2 | 1.01: 2; 1.02: 2; 1.03: 2.3; 1.04: 2.079; 1.05: 2.39; 1.06: 2.2; 1.07: 2.1 | 33
| 1 | 3 | 1.01: 2.4; 1.02: 2; 1.03: 2.2; 1.04: 2.0078; 1.05: 2.19; 1.06: 2.4; 1.07: 2 | 16
| 1 | 4 | 1.01: 1.9; 1.02: 1.9; 1.03: 2.1; 1.04: 2.0242; 1.05: 2.35; 1.06: 2.4; 1.07: 2.1 | 2.1052
| 1 | 5 | 1.01: 1.9; 1.02: 2; 1.03: 2.2; 1.04: 1.9741; 1.05: 2.38; 1.06: 2.5; 1.07: 1.9 | 138
| 1 | 6 | 1.01: 1.7; 1.02: 1.9; 1.03: 2.3; 1.04: 2.0247; 1.05: 2.15; 1.06: 2.3; 1.07: 2
| Mean |  | 1.01: 1.9667; 1.02: 1.95; 1.03: 2.2; 1.04: 2.0295; 1.05: 2.31; 1.06: 2.35; 1.07: 2.0167
| Median |  | 1.01: 1.9; 1.02: 1.95; 1.03: 2.2; 1.04: 2.0245; 1.05: 2.365; 1.06: 2.35; 1.07: 2
| Std Dev. |  | 1.01: 0.23381; 1.02: 0.054772; 1.03: 0.089443; 1.04: 0.038621; 1.05: 0.11045; 1.06: 0.10488; 1.07: 0.075277
| Rel.Std.Dev. |  | 1.01: 11.89%; 1.02: 2.809%; 1.03: 4.066%; 1.04: 1.903%; 1.05: 4.782%; 1.06: 4.463%; 1.07: 3.733%
| PDM3 |  | 1.01: -6.579%; 1.02: -7.371%; 1.03: 4.505%; 1.04: -3.595%; 1.05: 9.73%; 1.06: 11.63%; 1.07: -4.204%
| Z-Score (Absolute) |  | 1.01: 0.53; 1.02: 0.67; 1.03: 1.45; 1.04: 0; 1.05: 2.38; 1.06: 2.72; 1.07: 0.11
| Analytical results for K in OREAS 625 (Certified Value 2.89 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 2.89 | 2.8 | 0.5 | 2.831 | 2.881 | 2.947 | 2.847 | 2.897 | 2.86 | 2.96 | 2.7 | 2.62 | 2.8013 | 3.0241 | 2.88 | 3.29 | 2.9173 | 2.9 | 2.9 | 2.77 | 2.9 |  |  | 1
| 1 | 2 | 2.9 | 2.8 | 0.5 | 2.847 | 2.972 | 2.864 | 2.806 | 2.847 | 2.93 | 3.02 | 2.95 | 2.58 | 2.8935 | 2.9907 | 2.84 | 3.2 | 3.0456 | 2.9 | 2.8 | 2.88 | 2.9 |  |  | 0
| 1 | 3 | 2.89 | 2.9 | 0.5 | 2.881 | 2.864 | 2.914 | 2.823 | 2.922 | 2.92 | 2.94 | 2.84 | 2.54 | 2.6988 | 3.0461 | 2.88 | 3.22 | 2.9618 | 2.9 | 2.8 | 2.91 | 2.9 |  |  | 16
| 1 | 4 | 2.89 | 2.8 | 0.5 | 2.847 | 2.939 | 2.939 | 2.922 | 2.964 | 2.92 | 2.92 | 2.69 | 2.55 | 2.7354 | 2.9878 | 2.88 | 3.23 | 2.913 | 2.9 | 2.8 | 2.88 | 3 |  |  | 2.8869
| 1 | 5 | 2.88 | 2.9 | 0.5 | 2.856 | 2.872 | 2.847 | 2.889 | 2.889 | 2.91 | 2.98 | 2.72 | 2.61 | 2.6928 | 3.0027 | 2.88 | 3.22 | 3.0259 | 2.9 | 2.9 | 2.79 | 2.9 |  |  | 139
| 1 | 6 | 2.88 | 2.9 | 0.5 | 2.881 | 2.897 | 2.922 | 2.889 | 2.872 | 2.89 | 2.94 | 3.06 | 2.63 | 2.7646 | 3.0481 | 2.86 | 3.22 | 2.8613 | 2.9 | 2.9 | 2.78 | 2.9
| Mean |  | 2.8883 | 2.85 | 0.5 | 2.8572 | 2.9042 | 2.9055 | 2.8627 | 2.8985 | 2.905 | 2.96 | 2.8267 | 2.5883 | 2.7644 | 3.0166 | 2.87 | 3.23 | 2.9541 | 2.9 | 2.85 | 2.835 | 2.9167
| Median |  | 2.89 | 2.85 | 0.5 | 2.8515 | 2.889 | 2.918 | 2.868 | 2.893 | 2.915 | 2.95 | 2.78 | 2.595 | 2.75 | 3.0134 | 2.88 | 3.22 | 2.9396 | 2.9 | 2.85 | 2.835 | 2.9
| Std Dev. |  | 0.0075277 | 0.054772 | 0 | 0.020144 | 0.042546 | 0.040825 | 0.044581 | 0.040717 | 0.025884 | 0.035777 | 0.15227 | 0.037639 | 0.075249 | 0.026866 | 0.016733 | 0.030984 | 0.071052 | 0 | 0.054772 | 0.061563 | 0.040825
| Rel.Std.Dev. |  | 0.2606% | 1.922% | 0% | 0.705% | 1.465% | 1.405% | 1.557% | 1.405% | 0.891% | 1.209% | 5.387% | 1.454% | 2.722% | 0.8906% | 0.583% | 0.9593% | 2.405% | 0% | 1.922% | 2.172% | 1.4%
| PDM3 |  | 0.0484% | -1.279% | -82.68% | -1.031% | 0.5968% | 0.643% | -0.8407% | 0.4006% | 0.6257% | 2.531% | -2.088% | -10.34% | -4.244% | 4.491% | -0.5866% | 11.88% | 2.328% | 0.4525% | -1.279% | -1.799% | 1.03%
| Z-Score (Absolute) |  | 0 | 0.67 | 42.01 | 0.55 | 0.28 | 0.3 | 0.45 | 0.18 | 0.29 | 1.26 | 1.08 | 5.28 | 2.18 | 2.26 | 0.32 | 6.01 | 1.16 | 0.21 | 0.67 | 0.94 | 0.5
| Analytical results for La in OREAS 625 (Certified Value 43.4 ppm) | [24 empty] | Certified
La | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.23; 1.10: 1.25; 1.11: 1.26 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*OES; 1.07: BF*MS; 1.08: PF*MS; 1.09: PF*MS; 1.10: PF*OES; 1.11: PF*OES |  |  | 1
2
| 1 | 1 | 1.01: 43; 1.02: 43.9; 1.03: 37.6; 1.04: 44.9; 1.05: 43.966; 1.06: 43.1; 1.07: 44.957; 1.08: 54; 1.09: 41.4; 1.10: 38; 1.11: 40 |  |  | 1
| 1 | 2 | 1.01: 43; 1.02: 44.5; 1.03: 38.9; 1.04: 45.2; 1.05: 43.7512; 1.06: 43.8; 1.07: 44.456; 1.08: 61; 1.09: 43.4; 1.10: 39; 1.11: 40 |  |  | 34
| 1 | 3 | 1.01: 44; 1.02: 46.9; 1.03: 38.1; 1.04: 43.4; 1.05: 42.7452; 1.06: 45.1; 1.07: 44.512; 1.08: 59; 1.09: 42.3; 1.10: 39; 1.11: 40 |  |  | 16
| 1 | 4 | 1.01: 44; 1.02: 45.3; 1.03: 37.5; 1.04: 44.2; 1.05: 41.9939; 1.06: 45.1; 1.07: 44.534; 1.08: 58; 1.09: 44.9; 1.10: 39; 1.11: 40 |  |  | 43.3584
| 1 | 5 | 1.01: 44; 1.02: 44.7; 1.03: 39.3; 1.04: 45; 1.05: 42.2493; 1.06: 45.1; 1.07: 44.853; 1.08: 48; 1.09: 41.5; 1.10: 39; 1.11: 40 |  |  | 140
| 1 | 6 | 1.01: 43; 1.02: 41.4; 1.03: 36.7; 1.04: 43.3; 1.05: 43.6605; 1.06: 43.9; 1.07: 44.527; 1.08: 50; 1.09: 41.7; 1.10: 39; 1.11: 40
| Mean |  | 1.01: 43.5; 1.02: 44.45; 1.03: 38.0167; 1.04: 44.3333; 1.05: 43.061; 1.06: 44.35; 1.07: 44.6398; 1.08: 55; 1.09: 42.5333; 1.10: 38.8333; 1.11: 40
| Median |  | 1.01: 43.5; 1.02: 44.6; 1.03: 37.85; 1.04: 44.55; 1.05: 43.2028; 1.06: 44.5; 1.07: 44.5305; 1.08: 56; 1.09: 42; 1.10: 39; 1.11: 40
| Std Dev. |  | 1.01: 0.54772; 1.02: 1.8108; 1.03: 0.96003; 1.04: 0.83347; 1.05: 0.84285; 1.06: 0.8666; 1.07: 0.20981; 1.08: 5.2154; 1.09: 1.375; 1.10: 0.40825; 1.11: 0
| Rel.Std.Dev. |  | 1.01: 1.259%; 1.02: 4.074%; 1.03: 2.525%; 1.04: 1.88%; 1.05: 1.957%; 1.06: 1.954%; 1.07: 0.47%; 1.08: 9.482%; 1.09: 3.233%; 1.10: 1.051%; 1.11: 0%
| PDM3 |  | 1.01: 0.3265%; 1.02: 2.518%; 1.03: -12.32%; 1.04: 2.248%; 1.05: -0.686%; 1.06: 2.287%; 1.07: 2.955%; 1.08: 26.85%; 1.09: -1.903%; 1.10: -10.44%; 1.11: -7.746%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.66; 1.03: 3.82; 1.04: 0.58; 1.05: 0.31; 1.06: 0.59; 1.07: 0.8; 1.08: 8.02; 1.09: 0.67; 1.10: 3.26; 1.11: 2.44
| Analytical results for Li in OREAS 625 (Certified Value 18.2 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 1
2
| 1 | 1 | 17 | 18 | 10 | 20 | 17 | 20 | < 100 | 35.8447 | 19.4 | 14 | 20 | 37 | 30 | < 100 | [9 empty] | 1
| 1 | 2 | 18 | 16 | 40 | 20 | 18 | 10 | < 100 | 20.1433 | 19.7 | 15 | 18 | 37 | 20 | < 100 | [9 empty] | 0
| 1 | 3 | 17 | 24 | 40 | 19 | 19 | 20 | < 100 | 27.4584 | 19.4 | 16 | 20 | 38 | 30 | < 100 | [9 empty] | 16
| 1 | 4 | 17 | 18 | < 10 | 20 | 18 | 10 | < 100 | 36.7893 | 19.2 | 18 | 18 | 38 | 30 | < 100 | [9 empty] | 18.1905
| 1 | 5 | 17 | 18 | 40 | 20 | 18 | 20 | < 100 | 25.1833 | 19.2 | 15 | 23 | 38 | 30 | < 100 | [9 empty] | 141
| 1 | 6 | 18 | 17 | 20 | 20 | 18 | 20 | < 100 | 24.2616 | 18.7 | 16 | 20 | 37 | 30 | < 100
| Mean |  | 1.01: 17.3333; 1.02: 18.5; 1.03: 30; 1.04: 19.8333; 1.05: 18; 1.06: 16.6667; 1.08: 28.2801; 1.09: 19.2667; 1.10: 15.6667; 1.11: 19.8333; 1.12: 37.5; 1.13: 28.3333
| Median |  | 1.01: 17; 1.02: 18; 1.03: 40; 1.04: 20; 1.05: 18; 1.06: 20; 1.08: 26.3208; 1.09: 19.3; 1.10: 15.5; 1.11: 20; 1.12: 37.5; 1.13: 30
| Std Dev. |  | 1.01: 0.5164; 1.02: 2.8107; 1.03: 14.1421; 1.04: 0.40825; 1.05: 0.63246; 1.06: 5.164; 1.08: 6.6671; 1.09: 0.33267; 1.10: 1.3663; 1.11: 1.8348; 1.12: 0.54772; 1.13: 4.0825
| Rel.Std.Dev. |  | 1.01: 2.979%; 1.02: 15.19%; 1.03: 47.14%; 1.04: 2.058%; 1.05: 3.514%; 1.06: 30.98%; 1.08: 23.58%; 1.09: 1.727%; 1.10: 8.721%; 1.11: 9.251%; 1.12: 1.461%; 1.13: 14.41%
| PDM3 |  | 1.01: -4.712%; 1.02: 1.702%; 1.03: 64.92%; 1.04: 9.031%; 1.05: -1.047%; 1.06: -8.377%; 1.08: 55.47%; 1.09: 5.916%; 1.10: -13.87%; 1.11: 9.031%; 1.12: 106.2%; 1.13: 55.76%
| Z-Score (Absolute) |  | 0.51 | 0.27 | 1.21 | 0 | 0.37 | NA | 6.1 | 1.71 | 0.11 | 0.84 | 0 | 3.57 | NA | 6.1
Results from laboratories 1.11 and 1.26 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 625 (Certified Value 0.14 ppm) | [24 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.13; 1.04: 1.19; 1.05: 1.20; 1.06: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: BF*MS; 1.05: PF*MS; 1.06: PF*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.11; 1.02: 0.1; 1.03: 0.13755; 1.04: 0.135; 1.05: < 2; 1.06: 0.14 |  |  | 1
| 1 | 2 | 1.01: 0.13; 1.02: 0.2; 1.03: 0.14659; 1.04: 0.128; 1.05: < 2; 1.06: 0.15 |  |  | 35
| 1 | 3 | 1.01: 0.11; 1.02: 0.2; 1.03: 0.13985; 1.04: 0.132; 1.05: < 2; 1.06: 0.15 |  |  | 16
| 1 | 4 | 1.01: 0.11; 1.02: 0.2; 1.03: 0.1421; 1.04: 0.136; 1.05: < 2; 1.06: 0.16 |  |  | 0.1433
| 1 | 5 | 1.01: 0.13; 1.02: 0.1; 1.03: 0.15279; 1.04: 0.138; 1.05: < 2; 1.06: 0.16 |  |  | 142
| 1 | 6 | 1.01: 0.13; 1.02: 0.2; 1.03: 0.13613; 1.04: 0.135; 1.05: < 2; 1.06: 0.16
| Mean |  | 1.01: 0.12; 1.02: 0.16667; 1.03: 0.1425; 1.04: 0.134; 1.06: 0.15333
| Median |  | 1.01: 0.12; 1.02: 0.2; 1.03: 0.14097; 1.04: 0.135; 1.06: 0.155
| Std Dev. |  | 1.01: 0.010954; 1.02: 0.05164; 1.03: 0.0062446; 1.04: 0.0035214; 1.06: 0.008165
| Rel.Std.Dev. |  | 1.01: 9.129%; 1.02: 30.98%; 1.03: 4.382%; 1.04: 2.628%; 1.06: 5.325%
| PDM3 |  | 1.01: -16.26%; 1.02: 16.31%; 1.03: -0.5575%; 1.04: -6.49%; 1.06: 7.001%
| Z-Score (Absolute) |  | 1.01: 1.15; 1.02: 0.77; 1.03: 0.22; 1.04: 0.57; 1.05: 35.18; 1.06: 0.22
| Analytical results for Mg in OREAS 625 (Certified Value 0.267 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.48 | 0.26 | 0.28 | 0.265 | 0.259 | 0.241 | 0.259 | 0.247 | 0.26 | 0.27 | 0.27 | 0.26 | 0.27011 | 0.27671 | 0.28 | 0.27 | 0.2835 | 0.27305 | 0.27 | 0.26 | 0.25 | 0.29 |  | 1
| 1 | 2 | 0.38 | 0.26 | 0.26 | 0.253 | 0.265 | 0.259 | 0.265 | 0.241 | 0.26 | 0.28 | 0.27 | 0.26 | 0.27796 | 0.26905 | 0.28 | 0.26 | 0.2881 | 0.26708 | 0.27 | 0.26 | 0.25 | 0.29 |  | 0
| 1 | 3 | 0.42 | 0.26 | 0.28 | 0.265 | 0.259 | 0.259 | 0.259 | 0.253 | 0.26 | 0.27 | 0.28 | 0.26 | 0.26641 | 0.26304 | 0.28 | 0.26 | 0.2834 | 0.27771 | 0.27 | 0.26 | 0.24 | 0.29 |  | 16
| 1 | 4 | 0.42 | 0.27 | 0.28 | 0.259 | 0.271 | 0.241 | 0.265 | 0.247 | 0.27 | 0.27 | 0.27 | 0.26 | 0.29591 | 0.25986 | 0.28 | 0.26 | 0.2811 | 0.27248 | 0.27 | 0.25 | 0.25 | 0.29 |  | 0.26668
| 1 | 5 | 0.48 | 0.26 | 0.27 | 0.259 | 0.265 | 0.259 | 0.259 | 0.247 | 0.27 | 0.28 | 0.27 | 0.26 | 0.28309 | 0.27351 | 0.28 | 0.26 | 0.2885 | 0.27461 | 0.26 | 0.26 | 0.25 | 0.29 |  | 143
| 1 | 6 | 0.43 | 0.26 | 0.29 | 0.265 | 0.265 | 0.229 | 0.253 | 0.247 | 0.26 | 0.27 | 0.27 | 0.26 | 0.27856 | 0.27735 | 0.28 | 0.26 | 0.279 | 0.27202 | 0.27 | 0.27 | 0.24 | 0.29
| Mean |  | 0.435 | 0.26167 | 0.27667 | 0.261 | 0.264 | 0.248 | 0.26 | 0.247 | 0.26333 | 0.27333 | 0.27167 | 0.26 | 0.27867 | 0.26992 | 0.28 | 0.26167 | 0.28393 | 0.27283 | 0.26833 | 0.26 | 0.24667 | 0.29
| Median |  | 0.425 | 0.26 | 0.28 | 0.262 | 0.265 | 0.25 | 0.259 | 0.247 | 0.26 | 0.27 | 0.27 | 0.26 | 0.27826 | 0.27128 | 0.28 | 0.26 | 0.28345 | 0.27277 | 0.27 | 0.26 | 0.25 | 0.29
| Std Dev. |  | 0.038859 | 0.0040825 | 0.010328 | 0.004899 | 0.0045166 | 0.012822 | 0.0045166 | 0.0037947 | 0.005164 | 0.005164 | 0.0040825 | 0 | 0.010399 | 0.0072604 | 0 | 0.0040825 | 0.0037697 | 0.0034903 | 0.0040825 | 0.0063246 | 0.005164 | 0
| Rel.Std.Dev. |  | 8.933% | 1.56% | 3.733% | 1.877% | 1.711% | 5.17% | 1.737% | 1.536% | 1.961% | 1.889% | 1.503% | 0% | 3.732% | 2.69% | 0% | 1.56% | 1.328% | 1.279% | 1.521% | 2.433% | 2.094% | 0%
| PDM3 |  | 63.12% | -1.878% | 3.746% | -2.128% | -1.003% | -7.003% | -2.503% | -7.378% | -1.253% | 2.496% | 1.871% | -2.503% | 4.499% | 1.216% | 4.996% | -1.878% | 6.471% | 2.306% | 0.6215% | -2.503% | -7.503% | 8.746%
| Z-Score (Absolute) |  | 17.74 | 0.49 | 1.09 | 0.53 | 0.21 | 1.9 | 0.63 | 2.01 | 0.32 | 0.74 | 0.56 | 0.67 | 1.3 | 0.38 | 1.44 | 0.49 | 1.85 | 0.68 | 0.21 | 0.67 | 2.07 | 2.49
| Analytical results for Mn in OREAS 625 (Certified Value 0.447 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.4472 | 0.457 | 0.4547 | 0.441 | 0.449 | 0.457 | 0.441 | 0.441 | 0.456 | 0.4438 | 0.43 | 0.411 | 0.44026 | 0.49214 | 0.4315 | 0.461 | 0.4643 | 0.4462 | 0.4426 | 0.404 | 0.47 | 0.45 |  | 1
| 1 | 2 | 0.4472 | 0.464 | 0.4245 | 0.441 | 0.457 | 0.434 | 0.449 | 0.434 | 0.456 | 0.4592 | 0.44 | 0.3997 | 0.44168 | 0.49482 | 0.4302 | 0.45 | 0.4676 | 0.4503 | 0.4449 | 0.403 | 0.47 | 0.45 |  | 21
| 1 | 3 | 0.4412 | 0.518 | 0.4424 | 0.441 | 0.449 | 0.441 | 0.441 | 0.441 | 0.454 | 0.4476 | 0.44 | 0.3998 | 0.43961 | 0.47369 | 0.4341 | 0.446 | 0.4605 | 0.4487 | 0.447 | 0.406 | 0.47 | 0.45 |  | 16
| 1 | 4 | 0.4413 | 0.466 | 0.4425 | 0.441 | 0.457 | 0.449 | 0.449 | 0.441 | 0.459 | 0.4541 | 0.44 | 0.4026 | 0.44129 | 0.47138 | 0.4337 | 0.45 | 0.4605 | 0.449 | 0.4432 | 0.421 | 0.48 | 0.46 |  | 0.44679
| 1 | 5 | 0.4427 | 0.461 | 0.4305 | 0.449 | 0.449 | 0.441 | 0.441 | 0.434 | 0.459 | 0.456 | 0.43 | 0.4074 | 0.43393 | 0.49124 | 0.4349 | 0.453 | 0.462 | 0.4406 | 0.4431 | 0.412 | 0.47 | 0.44 |  | 144
| 1 | 6 | 0.4404 | 0.463 | 0.4607 | 0.449 | 0.449 | 0.441 | 0.441 | 0.434 | 0.457 | 0.4597 | 0.44 | 0.4093 | 0.43677 | 0.48087 | 0.4297 | 0.452 | 0.4581 | 0.446 | 0.441 | 0.405 | 0.48 | 0.45
| Mean |  | 0.44333 | 0.4715 | 0.44255 | 0.44367 | 0.45167 | 0.44383 | 0.44367 | 0.4375 | 0.45683 | 0.4534 | 0.43667 | 0.40497 | 0.43892 | 0.48402 | 0.43235 | 0.452 | 0.46217 | 0.4468 | 0.44363 | 0.4085 | 0.47333 | 0.45
| Median |  | 0.442 | 0.4635 | 0.44245 | 0.441 | 0.449 | 0.441 | 0.441 | 0.4375 | 0.4565 | 0.45505 | 0.44 | 0.405 | 0.43993 | 0.48605 | 0.4326 | 0.451 | 0.46125 | 0.44745 | 0.44315 | 0.4055 | 0.47 | 0.45
| Std Dev. |  | 0.0030852 | 0.022985 | 0.01377 | 0.0041312 | 0.0041312 | 0.0080104 | 0.0041312 | 0.0038341 | 0.0019408 | 0.0064246 | 0.005164 | 0.0049221 | 0.0030014 | 0.010111 | 0.0021797 | 0.00502 | 0.0033524 | 0.0034681 | 0.0020695 | 0.006892 | 0.005164 | 0.0063246
| Rel.Std.Dev. |  | 0.6959% | 4.875% | 3.112% | 0.9311% | 0.9147% | 1.805% | 0.9311% | 0.8764% | 0.4248% | 1.417% | 1.183% | 1.215% | 0.6838% | 2.089% | 0.5041% | 1.111% | 0.7254% | 0.7762% | 0.4665% | 1.687% | 1.091% | 1.405%
| PDM3 |  | -0.7741% | 5.53% | -0.9494% | -0.6995% | 1.091% | -0.6622% | -0.6995% | -2.08% | 2.247% | 1.479% | -2.266% | -9.361% | -1.761% | 8.333% | -3.232% | 1.166% | 3.441% | 0.001783% | -0.707% | -8.57% | 5.94% | 0.718%
| Z-Score (Absolute) |  | 0.06 | 2.45 | 0.13 | 0 | 0.69 | 0 | 0 | 0.58 | 1.14 | 0.84 | 0.66 | 3.49 | 0.46 | 3.57 | 1.04 | 0.71 | 1.62 | 0.25 | 0.03 | 3.17 | 2.62 | 0.53
| Analytical results for Mo in OREAS 625 (Certified Value 13.7 ppm) | [24 empty] | Certified
Mo | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.12; 1.06: 1.13; 1.07: 1.15; 1.08: 1.16; 1.09: 1.17; 1.10: 1.20; 1.11: 1.23; 1.12: 1.25; 1.13: 1.26 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES; 1.09: PF*OES/MS; 1.10: PF*MS; 1.11: PF*MS; 1.12: PF*OES; 1.13: PF*OES |  | 1
1
| 1 | 1 | 1.01: 15; 1.02: 14; 1.03: 10; 1.04: 12; 1.05: < 20; 1.06: 14.3926; 1.07: 13.7679; 1.08: 12.6; 1.09: 16; 1.10: 24; 1.11: 13; 1.12: 16; 1.13: 20 |  | 1
| 1 | 2 | 1.01: 14; 1.02: 14; 1.03: 10; 1.04: 13; 1.05: < 20; 1.06: 14.932; 1.07: 14.0794; 1.08: 12.2; 1.09: 16; 1.10: 18; 1.11: 13; 1.12: 15; 1.13: 20 |  | 37
| 1 | 3 | 1.01: 15; 1.02: 14; 1.03: 10; 1.04: 13; 1.05: < 20; 1.06: 13.9888; 1.07: 14.0517; 1.08: 13.1; 1.09: 16; 1.10: 18; 1.11: 13; 1.12: 16; 1.13: 10 |  | 16
| 1 | 4 | 1.01: 16; 1.02: 13; 1.03: 10; 1.04: 13; 1.05: < 20; 1.06: 14.364; 1.07: 13.8388; 1.08: 12.8; 1.09: 15; 1.10: 19; 1.11: 13; 1.12: 16; 1.13: 20 |  | 13.6828
| 1 | 5 | 1.01: 14; 1.02: 14; 1.03: 11; 1.04: 13; 1.05: < 20; 1.06: 14.6699; 1.07: 14.0699; 1.08: 13.2; 1.09: 15; 1.10: 19; 1.11: 12; 1.12: 15; 1.13: 20 |  | 145
| 1 | 6 | 1.01: 14; 1.02: 10; 1.03: 10; 1.04: 14; 1.05: < 20; 1.06: 14.369; 1.07: 13.6443; 1.08: 13.1; 1.09: 15; 1.10: 18; 1.11: 12; 1.12: 17; 1.13: 20
| Mean |  | 1.01: 14.6667; 1.02: 13.1667; 1.03: 10.1667; 1.04: 13; 1.06: 14.4527; 1.07: 13.9087; 1.08: 12.8333; 1.09: 15.5; 1.10: 19.3333; 1.11: 12.6667; 1.12: 15.8333; 1.13: 18.3333
| Median |  | 1.01: 14.5; 1.02: 14; 1.03: 10; 1.04: 13; 1.06: 14.3808; 1.07: 13.9453; 1.08: 12.95; 1.09: 15.5; 1.10: 18.5; 1.11: 13; 1.12: 16; 1.13: 20
| Std Dev. |  | 1.01: 0.8165; 1.02: 1.6021; 1.03: 0.40825; 1.04: 0.63246; 1.06: 0.31958; 1.07: 0.1845; 1.08: 0.38297; 1.09: 0.54772; 1.10: 2.3381; 1.11: 0.5164; 1.12: 0.75277; 1.13: 4.0825
| Rel.Std.Dev. |  | 1.01: 5.567%; 1.02: 12.17%; 1.03: 4.016%; 1.04: 4.865%; 1.06: 2.211%; 1.07: 1.326%; 1.08: 2.984%; 1.09: 3.534%; 1.10: 12.09%; 1.11: 4.077%; 1.12: 4.754%; 1.13: 22.27%
| PDM3 |  | 1.01: 7.19%; 1.02: -3.772%; 1.03: -25.7%; 1.04: -4.99%; 1.06: 5.627%; 1.07: 1.651%; 1.08: -6.208%; 1.09: 13.28%; 1.10: 41.3%; 1.11: -7.426%; 1.12: 15.72%; 1.13: 33.99%
| Z-Score (Absolute) |  | 1.01: 0.74; 1.02: 0.24; 1.03: 2.22; 1.04: 0.35; 1.05: 2.33; 1.06: 0.6; 1.07: 0.24; 1.08: 0.46; 1.09: 1.29; 1.10: 3.82; 1.11: 0.57; 1.12: 1.51; 1.13: NA
Results from laboratory 1.26 were removed due to their 10 ppm reading resolution.
| Analytical results for Na in OREAS 625 (Indicative Value 1.88 wt.%) | [24 empty] | Indicative
Na | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.16 |  | wt.%
|  |  | 1.01: PF*OES |  | 2
2
| 1 | 1 | 1.01: 1.9 |  | 1
| 1 | 2 | 1.01: 1.87 |  | 13
| 1 | 3 | 1.01: 1.89 |  | 16
| 1 | 4 | 1.01: 1.88 |  | 1.8833
| 1 | 5 | 1.01: 1.89 |  | 19
| 1 | 6 | 1.01: 1.87
| Mean |  | 1.01: 1.8833
| Median |  | 1.01: 1.885
| Std Dev. |  | 1.01: 0.012111
| Rel.Std.Dev. |  | 1.01: 0.643%
| PDM3 |  | 1.01: 1.776e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Nb in OREAS 625 (Certified Value 14.7 ppm) | [24 empty] | Certified
Nb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.20; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 1
1
| 1 | 1 | 1.01: 11; 1.02: 14.9; 1.03: 11; 1.04: 15; 1.05: 14.8797; 1.06: 17.9986; 1.07: 17; 1.08: 14 |  | 1
| 1 | 2 | 1.01: 12; 1.02: 15.7; 1.03: 12; 1.04: 16; 1.05: 15.8688; 1.06: 16.9096; 1.07: 17; 1.08: 15 |  | 38
| 1 | 3 | 1.01: 12; 1.02: 16.7; 1.03: 11; 1.04: 15; 1.05: 14.6104; 1.06: 17.8576; 1.07: 16; 1.08: 14 |  | 16
| 1 | 4 | 1.01: 12; 1.02: 15.9; 1.03: 10; 1.04: 16; 1.05: 14.9952; 1.06: 17.5479; 1.07: 16; 1.08: 15 |  | 14.7113
| 1 | 5 | 1.01: 12; 1.02: 16.7; 1.03: 11; 1.04: 14; 1.05: 15.4439; 1.06: 16.9368; 1.07: 17; 1.08: 14 |  | 146
| 1 | 6 | 1.01: 12; 1.02: 16.5; 1.03: 10; 1.04: 15; 1.05: 14.9825; 1.06: 17.7108; 1.07: 17; 1.08: 15
| Mean |  | 1.01: 11.8333; 1.02: 16.0667; 1.03: 10.8333; 1.04: 15.1667; 1.05: 15.1301; 1.06: 17.4935; 1.07: 16.6667; 1.08: 14.5
| Median |  | 1.01: 12; 1.02: 16.2; 1.03: 11; 1.04: 15; 1.05: 14.9888; 1.06: 17.6294; 1.07: 17; 1.08: 14.5
| Std Dev. |  | 1.01: 0.40825; 1.02: 0.70899; 1.03: 0.75277; 1.04: 0.75277; 1.05: 0.45098; 1.06: 0.46664; 1.07: 0.5164; 1.08: 0.54772
| Rel.Std.Dev. |  | 1.01: 3.45%; 1.02: 4.413%; 1.03: 6.949%; 1.04: 4.963%; 1.05: 2.981%; 1.06: 2.668%; 1.07: 3.098%; 1.08: 3.777%
| PDM3 |  | 1.01: -19.56%; 1.02: 9.213%; 1.03: -26.36%; 1.04: 3.095%; 1.05: 2.847%; 1.06: 18.91%; 1.07: 13.29%; 1.08: -1.436%
| Z-Score (Absolute) |  | 1.01: 1.83; 1.02: 0.51; 1.03: 2.39; 1.04: 0.01; 1.05: 0.01; 1.06: 1.3; 1.07: 0.84; 1.08: 0.36
| Analytical results for Nd in OREAS 625 (Certified Value 37.7 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 1
2
| 1 | 1 | 1.01: 37.9; 1.02: 31; 1.03: 36.9; 1.04: 37.7671; 1.05: 37.508; 1.06: 50; 1.07: 37.7246; 1.08: 36.8 |  | 1
| 1 | 2 | 1.01: 39.6; 1.02: 33; 1.03: 37.9; 1.04: 36.8202; 1.05: 37.099; 1.06: 42; 1.07: 37.5417; 1.08: 38.7 |  | 8
| 1 | 3 | 1.01: 40.7; 1.02: 32; 1.03: 37.1; 1.04: 37.6771; 1.05: 37.248; 1.06: 44; 1.07: 37.373; 1.08: 37.2 |  | 16
| 1 | 4 | 1.01: 40.6; 1.02: 31; 1.03: 37.2; 1.04: 36.1207; 1.05: 37.356; 1.06: 41; 1.07: 37.3046; 1.08: 39.8 |  | 37.7173
| 1 | 5 | 1.01: 39.7; 1.02: 33; 1.03: 36.7; 1.04: 36.9831; 1.05: 37.447; 1.06: 41; 1.07: 37.7074; 1.08: 37.1 |  | 147
| 1 | 6 | 1.01: 39.7; 1.02: 31; 1.03: 36.7; 1.04: 37.9359; 1.05: 37.164; 1.06: 40; 1.07: 37.9455; 1.08: 37.2
| Mean |  | 1.01: 39.7; 1.02: 31.8333; 1.03: 37.0833; 1.04: 37.2174; 1.05: 37.3037; 1.06: 43; 1.07: 37.5995; 1.08: 37.8
| Median |  | 1.01: 39.7; 1.02: 31.5; 1.03: 37; 1.04: 37.3301; 1.05: 37.302; 1.06: 41.5; 1.07: 37.6245; 1.08: 37.2
| Std Dev. |  | 1.01: 1.006; 1.02: 0.98319; 1.03: 0.44907; 1.04: 0.69931; 1.05: 0.16095; 1.06: 3.6878; 1.07: 0.24026; 1.08: 1.1849
| Rel.Std.Dev. |  | 1.01: 2.534%; 1.02: 3.089%; 1.03: 1.211%; 1.04: 1.879%; 1.05: 0.4315%; 1.06: 8.576%; 1.07: 0.639%; 1.08: 3.135%
| PDM3 |  | 1.01: 5.257%; 1.02: -15.6%; 1.03: -1.681%; 1.04: -1.326%; 1.05: -1.097%; 1.06: 14.01%; 1.07: -0.3124%; 1.08: 0.2193%
| Z-Score (Absolute) |  | 1.01: 4.23; 1.02: 10.57; 1.03: 0.69; 1.04: 0.44; 1.05: 0.28; 1.06: 10.44; 1.07: 0.28; 1.08: 0.66
| Analytical results for Ni in OREAS 625 (Indicative Value 20.7 ppm) | [24 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 1
1
| 1 | 1 | 11 | 20 | 89 | < 20 | 120 | < 20 | < 20 | < 20 | < 10 | < 20 | < 50 | < 20 | 31.7968 | < 20 | 7.16 | 30 | NA | 5 | < 10 | 10 | < 50 | < 100 |  | 1
| 1 | 2 | 11 | 10 | < 50 | < 20 | < 20 | 30 | 30 | < 20 | < 10 | < 20 | < 50 | < 20 | 29.6852 | < 20 | 8.22 | 50 | NA | 6 | < 10 | 20 | < 50 | < 100 |  | 14
| 1 | 3 | 11 | 20 | < 50 | < 20 | 40 | < 20 | 90 | < 20 | 12 | < 20 | < 50 | < 20 | 30.5916 | < 20 | 7.61 | 40 | NA | 14 | < 10 | 30 | < 50 | < 100 |  | 16
| 1 | 4 | 11 | 10 | 176 | < 20 | 20 | < 20 | 30 | 80 | < 10 | < 20 | < 50 | < 20 | 29.4528 | < 20 | 8.12 | 40 | NA | 8 | < 10 | 10 | < 50 | < 100 |  | 20.7019
| 1 | 5 | 11 | 10 | < 50 | < 20 | < 20 | < 20 | < 20 | < 20 | 11 | < 20 | < 50 | < 20 | 37.192 | < 20 | 8.08 | 40 | NA | 7 | < 10 | 10 | < 50 | < 100 |  | 20
| 1 | 6 | 11 | 10 | < 50 | < 20 | 20 | < 20 | 40 | < 20 | 11 | < 20 | < 50 | 21 | 36.4106 | < 20 | 7.22 | 40 | NA | < 5 | < 10 | 10 | < 50 | < 100
| Mean |  | 11 | 13.3333 | 132.5 |  | 50 | 30 | 47.5 | 80 | 11.3333 |  |  | 21 | 32.5215 |  | 7.735 | 40 |  | 8 |  | 15
| Median |  | 11 | 10 | 132.5 |  | 30 | 30 | 35 | 80 | 11 |  |  | 21 | 31.1942 |  | 7.845 | 40 |  | 7 |  | 10
| Std Dev. |  | 1.01: 0; 1.02: 5.164; 1.03: 61.5183; 1.05: 47.6095; 1.07: 28.7228; 1.09: 0.57735; 1.13: 3.4249; 1.16: 0.47201; 1.17: 6.3246; 1.23: 3.5355; 1.26: 8.3666
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 38.73%; 1.03: 46.43%; 1.05: 95.22%; 1.07: 60.47%; 1.09: 5.094%; 1.13: 10.53%; 1.16: 6.102%; 1.17: 15.81%; 1.23: 44.19%; 1.26: 55.78%
| PDM3 |  | -46.86% | -35.59% | 540% |  | 141.5% | 44.91% | 129.4% | 286.4% | -45.25% |  |  | 1.44% | 57.09% |  | -62.64% | 93.22% |  | -61.36% |  | -27.54%
| Z-Score (Absolute) |  | 0.25 | 0 | 5.12 | 0.36 | 2.52 | 0 | 2.34 | 0.9 | 0.56 | 0.36 | 1.26 | 0.16 | 2.07 | 0.36 | 0.6 | 2.88 | NA | 0.67 | 0.9 | 0.18 | 1.26 | 3.96
| Analytical results for P in OREAS 625 (Certified Value 0.03 wt.%) | [24 empty] | Certified
P | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.09; 1.03: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.16; 1.07: 1.17; 1.08: 1.20; 1.09: 1.23; 1.10: 1.25; 1.11: 1.26 |  | wt.%
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES; 1.05: PF*OES/MS; 1.06: PF*OES; 1.07: PF*OES/MS; 1.08: PF*MS; 1.09: PF*OES; 1.10: PF*OES; 1.11: PF*OES |  | 3
3
| 1 | 1 | 1.01: 0.0197; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.03; 1.07: 0.03; 1.08: 0.0363; 1.09: 0.03; 1.10: 0.03; 1.11: 0.02 |  | 1
| 1 | 2 | 1.01: 0.0196; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.03; 1.07: 0.03; 1.08: 0.0353; 1.09: 0.03; 1.10: 0.04; 1.11: 0.03 |  | 24
| 1 | 3 | 1.01: 0.0196; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.03; 1.07: 0.03; 1.08: 0.03; 1.09: 0.03; 1.10: 0.03; 1.11: 0.03 |  | 16
| 1 | 4 | 1.01: 0.0214; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.03; 1.07: 0.03; 1.08: 0.0326; 1.09: 0.03; 1.10: 0.03; 1.11: 0.02 |  | 0.030419
| 1 | 5 | 1.01: 0.0196; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.03; 1.07: 0.03; 1.08: 0.0324; 1.09: 0.03; 1.10: 0.03; 1.11: 0.03 |  | 148
| 1 | 6 | 1.01: 0.0196; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.03; 1.07: 0.03; 1.08: 0.0335; 1.09: 0.03; 1.10: 0.03; 1.11: 0.03
| Mean |  | 1.01: 0.019917; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.06: 0.03; 1.07: 0.03; 1.08: 0.03335; 1.09: 0.03; 1.10: 0.031667; 1.11: 0.026667
| Median |  | 1.01: 0.0196; 1.02: 0.05; 1.03: 0.03; 1.04: 0.03; 1.06: 0.03; 1.07: 0.03; 1.08: 0.03305; 1.09: 0.03; 1.10: 0.03; 1.11: 0.03
| Std Dev. |  | 1.01: 0.00072778; 1.02: 7.6012e-18; 1.03: 0; 1.04: 0; 1.06: 0; 1.07: 0; 1.08: 0.0022457; 1.09: 0; 1.10: 0.0040825; 1.11: 0.005164
| Rel.Std.Dev. |  | 1.01: 3.654%; 1.02: 1.52e-14%; 1.03: 0%; 1.04: 0%; 1.06: 0%; 1.07: 0%; 1.08: 6.734%; 1.09: 0%; 1.10: 12.89%; 1.11: 19.36%
| PDM3 |  | 1.01: -34.53%; 1.02: 64.37%; 1.03: -1.377%; 1.04: -1.377%; 1.06: -1.377%; 1.07: -1.377%; 1.08: 9.636%; 1.09: -1.377%; 1.10: 4.102%; 1.11: -12.33%
| Z-Score (Absolute) |  | 1.01: 4.08; 1.02: 8.09; 1.03: 0; 1.04: 0; 1.05: 2.02; 1.06: 0; 1.07: 0; 1.08: 1.36; 1.09: 0; 1.10: 0.67; 1.11: 1.35
| Analytical results for Pb in OREAS 625 (Certified Value 0.806 wt.%) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | 3
3
| 1 | 1 | 0.7826 | 0.82 | 0.7452 | 0.77 | 0.8 | 0.8 | 0.78 | 0.77 | 0.761 | 0.8741 | 0.81 | 0.8034 | 0.81711 | 0.81408 | 0.8 | 0.8433 | 0.8197 | 0.78706 | 0.8569 | 0.804 | 0.834 | 0.83 | 0.84 | 1
| 1 | 2 | 0.7855 | 0.81 | 0.6969 | 0.76 | 0.82 | 0.75 | 0.76 | 0.75 | 0.809 | 0.8743 | 0.81 | 0.783 | 0.81434 | 0.80093 | 0.806 | 0.8186 | 0.8352 | 0.79718 | 0.8572 | 0.8036 | 0.834 | 0.85 | 0.83 | 25
| 1 | 3 | 0.7832 | 0.82 | 0.7292 | 0.77 | 0.79 | 0.75 | 0.75 | 0.78 | 0.766 | 0.867 | 0.82 | 0.7813 | 0.80255 | 0.81563 | 0.801 | 0.8365 | 0.8368 | 0.7874 | 0.8535 | 0.8004 | 0.849 | 0.85 | 0.83 | 16
| 1 | 4 | 0.783 | 0.82 | 0.7287 | 0.77 | 0.81 | 0.81 | 0.81 | 0.77 | 0.786 | 0.8774 | 0.8 | 0.7884 | 0.80278 | 0.80032 | 0.822 | 0.8356 | 0.8182 | 0.78045 | 0.882 | 0.8052 | 0.831 | 0.87 | 0.83 | 0.80589
| 1 | 5 | 0.7844 | 0.81 | 0.7013 | 0.77 | 0.8 | 0.73 | 0.8 | 0.77 | 0.797 | 0.8794 | 0.8 | 0.7965 | 0.80781 | 0.80817 | 0.804 | 0.5368 | 0.8288 | 0.79561 | 0.8505 | 0.8028 | 0.832 | 0.84 | 0.81 | 149
| 1 | 6 | 0.7866 | 0.82 | 0.7596 | 0.78 | 0.81 | 0.8 | 0.79 | 0.77 | 0.772 | 0.8716 | 0.79 | 0.8071 | 0.80393 | 0.81588 | 0.814 | 0.8347 | 0.8132 | 0.78205 | 0.8631 | 0.8058 | 0.836 | 0.86 | 0.82
| Mean |  | 0.78422 | 0.81667 | 0.72682 | 0.77 | 0.805 | 0.77333 | 0.78167 | 0.76833 | 0.78183 | 0.87397 | 0.805 | 0.79328 | 0.80809 | 0.80917 | 0.80783 | 0.78425 | 0.82532 | 0.78829 | 0.86053 | 0.80363 | 0.836 | 0.85 | 0.82667
| Median |  | 0.7838 | 0.82 | 0.72895 | 0.77 | 0.805 | 0.775 | 0.785 | 0.77 | 0.779 | 0.8742 | 0.805 | 0.79245 | 0.80587 | 0.81112 | 0.805 | 0.83515 | 0.82425 | 0.78723 | 0.85705 | 0.8038 | 0.834 | 0.85 | 0.83
| Std Dev. |  | 0.0015804 | 0.005164 | 0.024364 | 0.0063246 | 0.010488 | 0.033862 | 0.023166 | 0.0098319 | 0.018798 | 0.004372 | 0.010488 | 0.01074 | 0.0062713 | 0.0071845 | 0.0085421 | 0.1215 | 0.0097025 | 0.0068628 | 0.01133 | 0.0019201 | 0.006603 | 0.014142 | 0.010328
| Rel.Std.Dev. |  | 0.2015% | 0.6323% | 3.352% | 0.8214% | 1.303% | 4.379% | 2.964% | 1.28% | 2.404% | 0.5003% | 1.303% | 1.354% | 0.7761% | 0.8879% | 1.057% | 15.49% | 1.176% | 0.8706% | 1.317% | 0.2389% | 0.7898% | 1.664% | 1.249%
| PDM3 |  | -2.689% | 1.338% | -9.811% | -4.453% | -0.1099% | -4.039% | -3.005% | -4.66% | -2.985% | 8.448% | -0.1099% | -1.564% | 0.2731% | 0.4073% | 0.2417% | -2.685% | 2.411% | -2.183% | 6.781% | -0.2795% | 3.737% | 5.474% | 2.579%
| Z-Score (Absolute) |  | 0.67 | 0.38 | 2.54 | 1.14 | 0 | 1.03 | 0.76 | 1.19 | 0.75 | 2.24 | 0 | 0.38 | 0.1 | 0.14 | 0.09 | 0.67 | 0.66 | 0.54 | 1.8 | 0.04 | 1.01 | 1.46 | 0.7
| Analytical results for Pr in OREAS 625 (Certified Value 9.95 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | 2
3
| 1 | 1 | 1.01: 9.5; 1.02: 8.18; 1.03: 10; 1.04: 10.0052; 1.05: 9.978; 1.06: 15; 1.07: 11.3256; 1.08: 9.66 | 1
| 1 | 2 | 1.01: 10.1; 1.02: 8.5; 1.03: 10; 1.04: 10.279; 1.05: 9.84; 1.06: 12; 1.07: 11.5135; 1.08: 10.21 | 10
| 1 | 3 | 1.01: 10.1; 1.02: 8.28; 1.03: 9.6; 1.04: 10.045; 1.05: 9.854; 1.06: 13; 1.07: 11.8292; 1.08: 9.82 | 16
| 1 | 4 | 1.01: 10.5; 1.02: 8.08; 1.03: 9.9; 1.04: 9.7179; 1.05: 9.924; 1.06: 12; 1.07: 11.1842; 1.08: 10.56 | 9.9529
| 1 | 5 | 1.01: 9.8; 1.02: 8.34; 1.03: 9.9; 1.04: 10.2507; 1.05: 9.916; 1.06: 12; 1.07: 11.7242; 1.08: 9.82 | 150
| 1 | 6 | 1.01: 11.2; 1.02: 7.99; 1.03: 9.9; 1.04: 10.253; 1.05: 9.883; 1.06: 12; 1.07: 11.3986; 1.08: 9.85
| Mean |  | 1.01: 10.2; 1.02: 8.2283; 1.03: 9.8833; 1.04: 10.0918; 1.05: 9.8992; 1.06: 12.6667; 1.07: 11.4959; 1.08: 9.9867
| Median |  | 1.01: 10.1; 1.02: 8.23; 1.03: 9.9; 1.04: 10.1479; 1.05: 9.8995; 1.06: 12; 1.07: 11.4561; 1.08: 9.835
| Std Dev. |  | 1.01: 0.5933; 1.02: 0.18444; 1.03: 0.1472; 1.04: 0.21714; 1.05: 0.050827; 1.06: 1.2111; 1.07: 0.24458; 1.08: 0.33452
| Rel.Std.Dev. |  | 1.01: 5.817%; 1.02: 2.241%; 1.03: 1.489%; 1.04: 2.152%; 1.05: 0.5134%; 1.06: 9.561%; 1.07: 2.128%; 1.08: 3.35%
| PDM3 |  | 1.01: 2.482%; 1.02: -17.33%; 1.03: -0.6993%; 1.04: 1.395%; 1.05: -0.5402%; 1.06: 27.27%; 1.07: 15.5%; 1.08: 0.3389%
| Z-Score (Absolute) |  | 1.01: 0.68; 1.02: 7.71; 1.03: 0.66; 1.04: 0.22; 1.05: 0.6; 1.06: 11.19; 1.07: 6.2; 1.08: 0.22
| Analytical results for Rb in OREAS 625 (Certified Value 136 ppm) | [24 empty] | Certified
Rb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.15; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*MS | 0
0
| 1 | 1 | 1.01: 139; 1.02: 120; 1.03: 147.1; 1.04: 133.538; 1.05: 141.7184; 1.06: 152; 1.07: 133; 1.08: 126 | 1
| 1 | 2 | 1.01: 139; 1.02: 125; 1.03: 144.2; 1.04: 138.2853; 1.05: 137.8595; 1.06: 149; 1.07: 130; 1.08: 128 | 43
| 1 | 3 | 1.01: 154; 1.02: 119; 1.03: 141.8; 1.04: 134.3329; 1.05: 137.4056; 1.06: 150; 1.07: 139; 1.08: 128 | 16
| 1 | 4 | 1.01: 141; 1.02: 119; 1.03: 143.6; 1.04: 134.3176; 1.05: 136.4841; 1.06: 147; 1.07: 140; 1.08: 130 | 136.3857
| 1 | 5 | 1.01: 138; 1.02: 121; 1.03: 145; 1.04: 132.7046; 1.05: 140.5276; 1.06: 154; 1.07: 134; 1.08: 126 | 151
| 1 | 6 | 1.01: 145; 1.02: 118; 1.03: 142.5; 1.04: 136.8141; 1.05: 139.928; 1.06: 148; 1.07: 134; 1.08: 126
| Mean |  | 1.01: 142.6667; 1.02: 120.3333; 1.03: 144.0333; 1.04: 134.9988; 1.05: 138.9872; 1.06: 150; 1.07: 135; 1.08: 127.3333
| Median |  | 1.01: 140; 1.02: 119.5; 1.03: 143.9; 1.04: 134.3253; 1.05: 138.8937; 1.06: 149.5; 1.07: 134; 1.08: 127
| Std Dev. |  | 1.01: 6.0882; 1.02: 2.5033; 1.03: 1.8917; 1.04: 2.1166; 1.05: 2.0374; 1.06: 2.6077; 1.07: 3.7947; 1.08: 1.633
| Rel.Std.Dev. |  | 1.01: 4.267%; 1.02: 2.08%; 1.03: 1.313%; 1.04: 1.568%; 1.05: 1.466%; 1.06: 1.738%; 1.07: 2.811%; 1.08: 1.282%
| PDM3 |  | 1.01: 4.605%; 1.02: -11.77%; 1.03: 5.607%; 1.04: -1.017%; 1.05: 1.907%; 1.06: 9.982%; 1.07: -1.016%; 1.08: -6.637%
| Z-Score (Absolute) |  | 1.01: 0.6; 1.02: 1.77; 1.03: 0.75; 1.04: 0.21; 1.05: 0.21; 1.06: 1.38; 1.07: 0.21; 1.08: 1.02
| Analytical results for Re in OREAS 625 (Indicative Value < 0.05 ppm) | [24 empty] | Indicative
Re | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.15; 1.02: 1.17; 1.03: 1.20 | ppm
|  |  | 1.01: PF*MS; 1.02: PF*OES/MS; 1.03: PF*MS | 3
3
| 1 | 1 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 1
| 1 | 2 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 15
| 1 | 3 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 16
| 1 | 4 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | < 0.05
| 1 | 5 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 21
| 1 | 6 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for S in OREAS 625 (Certified Value 3.85 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.20 | 1.26 | 1.27 | [6 empty] | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | [6 empty] | 2
3
| 1 | 1 | 3.92 | 3.69 | 3.79 | 3.92 | 3.93 | 3.74 | 3.78 | 3.98 | 3.84 | 3.67 | 3.63 | 4.3082 | 3.87 | 4 | 4.0291 | 3.23 | 3.9 | [6 empty] | 1
| 1 | 2 | 3.89 | 3.68 | 3.79 | 3.89 | 3.81 | 3.67 | 3.71 | 3.96 | 3.94 | 3.66 | 3.59 | 4.2574 | 4.18 | 3.9 | 4.1735 | 3.31 | 3.95 | [6 empty] | 26
| 1 | 3 | 3.98 | 3.75 | 3.81 | 3.79 | 3.79 | 3.78 | 3.88 | 3.86 | 3.85 | 3.71 | 3.56 | 4.1737 | 4.21 | 3.93 | 4.1771 | 3.25 | 4.04 | [6 empty] | 16
| 1 | 4 | 3.95 | 3.74 | 3.79 | 3.9 | 3.9 | 3.83 | 3.83 | 3.98 | 3.83 | 3.68 | 3.54 | 4.1772 | 3.79 | 3.95 | 4.0486 | 3.23 | 4.07 | [6 empty] | 3.851
| 1 | 5 | 3.95 | 3.63 | 3.78 | 3.86 | 3.83 | 3.8 | 3.76 | 4 | 3.91 | 3.65 | 3.6 | 4.2406 | 3.74 | 3.95 | 4.2088 | 3.31 | 3.94 | [6 empty] | 152
| 1 | 6 | 3.98 | 3.72 | 3.82 | 3.81 | 3.9 | 3.72 | 3.8 | 3.91 | 3.91 | 3.72 | 3.56 | 4.3069 | 3.67 | 3.98 | 4.0418 | 3.12 | 4.01
| Mean |  | 3.945 | 3.7017 | 3.7967 | 3.8617 | 3.86 | 3.7567 | 3.7933 | 3.9483 | 3.88 | 3.6817 | 3.58 | 4.244 | 3.91 | 3.9517 | 4.1132 | 3.2417 | 3.985
| Median |  | 3.95 | 3.705 | 3.79 | 3.875 | 3.865 | 3.76 | 3.79 | 3.97 | 3.88 | 3.675 | 3.575 | 4.249 | 3.83 | 3.95 | 4.111 | 3.24 | 3.98
| Std Dev. |  | 0.035071 | 0.04446 | 0.015055 | 0.051929 | 0.057271 | 0.058195 | 0.058538 | 0.053072 | 0.045607 | 0.027869 | 0.032863 | 0.059451 | 0.23039 | 0.035449 | 0.081489 | 0.069976 | 0.065345
| Rel.Std.Dev. |  | 0.889% | 1.201% | 0.3965% | 1.345% | 1.484% | 1.549% | 1.543% | 1.344% | 1.175% | 0.757% | 0.918% | 1.401% | 5.892% | 0.8971% | 1.981% | 2.159% | 1.64%
| PDM3 |  | 2.441% | -3.877% | -1.411% | 0.2773% | 0.234% | -2.449% | -1.497% | 2.528% | 0.7534% | -4.397% | -7.037% | 10.21% | 1.532% | 2.614% | 6.808% | -15.82% | 3.48%
| Z-Score (Absolute) |  | 0.62 | 1.2 | 0.49 | 0 | 0.01 | 0.79 | 0.51 | 0.65 | 0.14 | 1.35 | 2.11 | 2.86 | 0.36 | 0.67 | 1.88 | 4.65 | 0.92
| Analytical results for Sb in OREAS 625 (Certified Value 36.2 ppm) | [24 empty] | Certified
Sb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | [6 empty] | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.04: 1.09; 1.05: 1.10; 1.06: 1.15; 1.07: 1.16; 1.08: 1.23; 1.09: 1.25; 1.10: 1.26; 1.11: 1.28 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*OES; 1.10: PF*OES; 1.11: PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 1.01: 262; 1.02: 36; 1.04: 35; 1.05: 37.3; 1.06: 37.6199; 1.07: 48.8; 1.08: 33.7; 1.09: < 50; 1.10: < 50; 1.11: < 100 | [6 empty] | 1
| 1 | 2 | 1.01: 265; 1.02: 34; 1.04: 37; 1.05: 34.7; 1.06: 38.2345; 1.07: 48.9; 1.08: 36.4; 1.09: < 50; 1.10: < 50; 1.11: < 100 | [6 empty] | 45
| 1 | 3 | 1.01: 262; 1.02: 39; 1.04: 35; 1.05: 35.8; 1.06: 35.9516; 1.07: 48.5; 1.08: 35.2; 1.09: < 50; 1.10: < 50; 1.11: < 100 | [6 empty] | 16
| 1 | 4 | 1.01: 261; 1.02: 36; 1.04: 35; 1.05: 35.3; 1.06: 36.0405; 1.07: 49.8; 1.08: 37.3; 1.09: < 50; 1.10: < 50; 1.11: < 100 | [6 empty] | 36.1703
| 1 | 5 | 1.01: 262; 1.02: 37; 1.04: 36; 1.05: 37.3; 1.06: 38.9746; 1.07: 48.9; 1.08: 34.3; 1.09: < 50; 1.10: < 50; 1.11: < 100 | [6 empty] | 153
| 1 | 6 | 1.01: 263; 1.02: 38; 1.04: 35; 1.05: 35.4; 1.06: 36.9885; 1.07: 48.4; 1.08: 35.6; 1.09: < 50; 1.10: < 50; 1.11: < 100
| Mean |  | 1.01: 262.5; 1.02: 36.6667; 1.04: 35.5; 1.05: 35.9667; 1.06: 37.3016; 1.07: 48.8833; 1.08: 35.4167
| Median |  | 1.01: 262; 1.02: 36.5; 1.04: 35; 1.05: 35.6; 1.06: 37.3042; 1.07: 48.85; 1.08: 35.4
| Std Dev. |  | 1.01: 1.3784; 1.02: 1.7512; 1.04: 0.83666; 1.05: 1.0912; 1.06: 1.2067; 1.07: 0.49565; 1.08: 1.3258
| Rel.Std.Dev. |  | 1.01: 0.5251%; 1.02: 4.776%; 1.04: 2.357%; 1.05: 3.034%; 1.06: 3.235%; 1.07: 1.014%; 1.08: 3.743%
| PDM3 |  | 1.01: 625.7%; 1.02: 1.372%; 1.04: -1.853%; 1.05: -0.563%; 1.06: 3.128%; 1.07: 35.15%; 1.08: -2.084%
| Z-Score (Absolute) |  | 1.01: 24.8; 1.02: 0.04; 1.04: 0.09; 1.05: 0.04; 1.06: 0.11; 1.07: 1.38; 1.08: 0.1; 1.09: 1.24; 1.10: 1.24; 1.11: 1.5
| Analytical results for Sc in OREAS 625 (Indicative Value 4.84 ppm) | [24 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.09 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | 261 | < 10 | < 10 | < 20 | 4.3 | 3 | 8.176 | 10 | < 5 | < 5 | 5 | [12 empty] | 1
| 1 | 2 | 261 | < 10 | < 10 | < 20 | 4.4 | 3 | 8.459 | 6 | < 5 | < 5 | 5 | [12 empty] | 16
| 1 | 3 | 261 | < 10 | < 10 | < 20 | 4.3 | 3 | 8.331 | 6 | < 5 | < 5 | 6 | [12 empty] | 16
| 1 | 4 | 261 | < 10 | < 10 | < 20 | 4.4 | 3 | 8.263 | 6 | 5 | < 5 | 6 | [12 empty] | 4.8388
| 1 | 5 | 261 | < 10 | < 10 | < 20 | 4.3 | 3 | 8.214 | 6 | < 5 | < 5 | 5 | [12 empty] | 22
| 1 | 6 | 260 | < 10 | < 10 | < 20 | 4.3 | 3 | 8.254 | 6 | < 5 | < 5 | < 5
| Mean |  | 1.01: 260.8333; 1.06: 4.3333; 1.07: 3; 1.08: 8.2828; 1.09: 6.6667; 1.10: 5; 1.12: 5.4
| Median |  | 1.01: 261; 1.06: 4.3; 1.07: 3; 1.08: 8.2585; 1.09: 6; 1.10: 5; 1.12: 5
| Std Dev. |  | 1.01: 0.40825; 1.06: 0.05164; 1.07: 0; 1.08: 0.10073; 1.09: 1.633; 1.12: 0.54772
| Rel.Std.Dev. |  | 1.01: 0.1565%; 1.06: 1.192%; 1.07: 0%; 1.08: 1.216%; 1.09: 24.49%; 1.12: 10.14%
| PDM3 |  | 1.01: 5290%; 1.06: -10.45%; 1.07: -38%; 1.08: 71.18%; 1.09: 37.78%; 1.10: 3.331%; 1.12: 11.6%
| Z-Score (Absolute) |  | 86.26 | 0 | 0 | 1.69 | 0.22 | 0.67 | 1.11 | 0.56 | 0.7 | 0.84 | 0.03
| Analytical results for Se in OREAS 625 (Indicative Value < 5 ppm) | [24 empty] | Indicative
Se | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.15 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: < 8; 1.02: < 5; 1.04: < 20 | [6 empty] | 1
| 1 | 2 | 1.01: < 8; 1.02: < 5; 1.04: < 20 | [6 empty] | 3
| 1 | 3 | 1.01: 10; 1.02: < 5; 1.04: < 20 | [6 empty] | 16
| 1 | 4 | 1.01: < 8; 1.02: < 5; 1.04: < 20 | [6 empty] | < 5
| 1 | 5 | 1.01: 11; 1.02: < 5; 1.04: < 20 | [6 empty] | 18
| 1 | 6 | 1.01: 9; 1.02: < 5; 1.04: < 20
| Mean |  | 1.01: 10
| Median |  | 1.01: 10
Std Dev. | bar: 1.01
| Rel.Std.Dev. |  | 1.01: 10%
PDM3
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 1.01; 1.04: 0.67
| Analytical results for Si in OREAS 625 (Certified Value 29.02 wt.%) | [24 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.15 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [7 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 28.8 | > 15 | 27.91 | 29.68 | 29.4 | 29.31 | 29.78 | 29.1 | 29.5 | 29.2 | 29.3812 | 28.7157 | 28.1 | 28.2 | 27.53 | 29.9 | [7 empty] | 1
| 1 | 2 | 28.6 | > 15 | 27.91 | 30.2 | 28.61 | 29.31 | 29.5 | 29 | 30.1 | 28.4 | 28.878 | 29.1052 | 28.9 | 28.6 | 27.62 | 30 | [7 empty] | 0
| 1 | 3 | 29.1 | > 15 | 27.91 | 29.59 | 29.22 | 29.22 | 30.2 | 28.5 | 29 | 28.6 | 28.7749 | 29.2551 | 28.5 | 28 | 27.7 | 30.3 | [7 empty] | 16
| 1 | 4 | 29.1 | > 15 | 27.81 | 30.2 | 28.98 | 30.2 | 30.62 | 29.2 | 29 | 28.8 | 28.705 | 28.4922 | 29 | 28.2 | 28.34 | 30.1 | [7 empty] | 29.0185
| 1 | 5 | 29.1 | > 15 | 27.81 | 29.92 | 28.8 | 29.78 | 30.01 | 29.2 | 29.9 | 28.9 | 29.2761 | 29.3646 | 28.4 | 28.7 | 27.4 | 29.9 | [7 empty] | 154
| 1 | 6 | 29.2 | > 15 | 28 | 29.78 | 29.12 | 29.59 | 29.68 | 28.8 | 29.2 | 29.3 | 29.2237 | 28.1485 | 29.4 | 28.4 | 27.22 | 30.3
| Mean |  | 28.9833 |  | 27.8917 | 29.895 | 29.0217 | 29.5683 | 29.965 | 28.9667 | 29.45 | 28.8667 | 29.0398 | 28.8469 | 28.7167 | 28.35 | 27.635 | 30.0833
| Median |  | 29.1 |  | 27.91 | 29.85 | 29.05 | 29.45 | 29.895 | 29.05 | 29.35 | 28.85 | 29.0508 | 28.9105 | 28.7 | 28.3 | 27.575 | 30.05
| Std Dev. |  | 0.23166 |  | 0.072226 | 0.26044 | 0.28722 | 0.37392 | 0.40446 | 0.27325 | 0.46797 | 0.34448 | 0.28797 | 0.47585 | 0.47081 | 0.26646 | 0.38469 | 0.18348
| Rel.Std.Dev. |  | 0.7993% |  | 0.259% | 0.8712% | 0.9897% | 1.265% | 1.35% | 0.9433% | 1.589% | 1.193% | 0.9916% | 1.65% | 1.64% | 0.9399% | 1.392% | 0.6099%
| PDM3 |  | -0.1213% |  | -3.883% | 3.02% | 0.01077% | 1.895% | 3.262% | -0.1788% | 1.487% | -0.5234% | 0.07333% | -0.5915% | -1.04% | -2.304% | -4.768% | 3.669%
| Z-Score (Absolute) |  | 0 | NA | 1.58 | 1.32 | 0.06 | 0.84 | 1.42 | 0.02 | 0.67 | 0.17 | 0.08 | 0.2 | 0.39 | 0.92 | 1.95 | 1.59
| Analytical results for Sm in OREAS 625 (Certified Value 7.04 ppm) | [24 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.19; 1.07: 1.20; 1.08: 1.22; 1.09: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: BF*MS; 1.07: PF*MS; 1.08: PF*MS; 1.09: PF*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 6.2; 1.02: 6.2; 1.04: 7.1; 1.05: 7.431; 1.06: 7.091; 1.07: 7; 1.08: 7.1548; 1.09: 6.9 | [6 empty] | 1
| 1 | 2 | 1.01: 5.9; 1.02: 6.2; 1.04: 6.8; 1.05: 7.0467; 1.06: 6.97; 1.07: 8; 1.08: 8.1903; 1.09: 7.4 | [6 empty] | 8
| 1 | 3 | 1.01: 6.4; 1.02: 6; 1.04: 6.9; 1.05: 6.7945; 1.06: 6.854; 1.07: 8; 1.08: 8.0761; 1.09: 7 | [6 empty] | 16
| 1 | 4 | 1.01: 6.7; 1.02: 5.9; 1.04: 6.9; 1.05: 6.7706; 1.06: 6.973; 1.07: 8; 1.08: 7.4457; 1.09: 7.6 | [6 empty] | 7.0366
| 1 | 5 | 1.01: 7.2; 1.02: 6.2; 1.04: 6.6; 1.05: 7.0134; 1.06: 6.967; 1.07: 10; 1.08: 7.7178; 1.09: 7.1 | [6 empty] | 155
| 1 | 6 | 1.01: 6.5; 1.02: 6; 1.04: 6.9; 1.05: 7.3636; 1.06: 6.924; 1.07: 9; 1.08: 7.3331; 1.09: 7.1
| Mean |  | 1.01: 6.4833; 1.02: 6.0833; 1.04: 6.8667; 1.05: 7.07; 1.06: 6.9632; 1.07: 8.3333; 1.08: 7.653; 1.09: 7.1833
| Median |  | 1.01: 6.45; 1.02: 6.1; 1.04: 6.9; 1.05: 7.03; 1.06: 6.9685; 1.07: 8; 1.08: 7.5817; 1.09: 7.1
| Std Dev. |  | 1.01: 0.4446; 1.02: 0.13292; 1.04: 0.1633; 1.05: 0.27776; 1.06: 0.077396; 1.07: 1.0328; 1.08: 0.41604; 1.09: 0.26394
| Rel.Std.Dev. |  | 1.01: 6.858%; 1.02: 2.185%; 1.04: 2.378%; 1.05: 3.929%; 1.06: 1.112%; 1.07: 12.39%; 1.08: 5.436%; 1.09: 3.674%
| PDM3 |  | 1.01: -7.862%; 1.02: -13.55%; 1.04: -2.415%; 1.05: 0.4746%; 1.06: -1.043%; 1.07: 18.43%; 1.08: 8.76%; 1.09: 2.086%
| Z-Score (Absolute) |  | 1.01: 1.47; 1.02: 2.69; 1.04: 0.3; 1.05: 0.33; 1.06: 0; 1.07: NA; 1.08: 2.11; 1.09: 0.67
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 625 (Indicative Value 3.48 ppm) | [24 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | [10 empty] | 2
2
| 1 | 1 | 14 | < 100 | 5 | 5 | < 50 | 2.1754 | < 100 | < 5 | NA | 3 | < 50 | < 50 | < 100 | [10 empty] | 1
| 1 | 2 | 11.4 | < 100 | 5 | 5 | < 50 | < 2 | < 100 | < 5 | NA | 3 | < 50 | < 50 | < 100 | [10 empty] | 13
| 1 | 3 | 7.3 | < 100 | 5 | 5 | < 50 | < 2 | < 100 | < 5 | NA | 3 | < 50 | < 50 | < 100 | [10 empty] | 16
| 1 | 4 | 14.2 | < 100 | 5 | 4 | < 50 | < 2 | < 100 | < 5 | NA | 3 | < 50 | < 50 | < 100 | [10 empty] | 3.4754
| 1 | 5 | 13.7 | < 100 | 5 | 4 | < 50 | 2.3556 | < 100 | < 5 | NA | 3 | < 50 | < 50 | < 100 | [10 empty] | 19
| 1 | 6 | 13.3 | < 100 | 5 | 5 | < 50 | < 2 | < 100 | < 5 | NA | 3 | < 50 | < 50 | < 100
| Mean |  | 1.01: 12.3167; 1.04: 5; 1.05: 4.6667; 1.07: 2.2655; 1.11: 3
| Median |  | 1.01: 13.5; 1.04: 5; 1.05: 5; 1.07: 2.2655; 1.11: 3
| Std Dev. |  | 1.01: 2.6559; 1.04: 0; 1.05: 0.5164; 1.07: 0.12743; 1.11: 0
| Rel.Std.Dev. |  | 1.01: 21.56%; 1.04: 0%; 1.05: 11.07%; 1.07: 5.625%; 1.11: 0%
| PDM3 |  | 1.01: 254.4%; 1.04: 43.87%; 1.05: 34.28%; 1.07: -34.81%; 1.11: -13.68%
| Z-Score (Absolute) |  | 0.29 | 1.43 | 0.62 | 0.64 | 0.29 | 0.78 | 1.43 | 0.73 | NA | 0.71 | 0.29 | 0.29 | 1.43
| Analytical results for Sr in OREAS 625 (Certified Value 142 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 0
0
| 1 | 1 | 138 | 140 | 134 | 141 | 157 | 153.86 | 143.2576 | 134.9 | 169 | 160 | 158.869 | 138 | 134 | 130 | [9 empty] | 1
| 1 | 2 | 139 | 139 | 128 | 142 | 155 | 162.4372 | 143.4861 | 136.9 | 161 | 130 | 158.793 | 141 | 135 | 130 | [9 empty] | 21
| 1 | 3 | 138 | 159 | 144 | 143 | 153 | 153.8133 | 138.0521 | 136 | 163 | 140 | 155.53 | 141 | 137 | 130 | [9 empty] | 16
| 1 | 4 | 138 | 145 | 128 | 143 | 156 | 154.8939 | 135.7818 | 135.1 | 166 | 143 | 155.942 | 140 | 135 | 130 | [9 empty] | 142.3775
| 1 | 5 | 138 | 140 | 133 | 144 | 153 | 151.8858 | 135.7436 | 136.9 | 171 | 136 | 158.55 | 137 | 134 | 130 | [9 empty] | 156
| 1 | 6 | 139 | 147 | 148 | 142 | 154 | 160.7218 | 145.1007 | 136.1 | 161 | 147 | 161.629 | 138 | 134 | 130
| Mean |  | 138.3333 | 145 | 135.8333 | 142.5 | 154.6667 | 156.2687 | 140.237 | 135.9833 | 165.1667 | 142.6667 | 158.2188 | 139.1667 | 134.8333 | 130
| Median |  | 138 | 142.5 | 133.5 | 142.5 | 154.5 | 154.3769 | 140.6549 | 136.05 | 164.5 | 141.5 | 158.6715 | 139 | 134.5 | 130
| Std Dev. |  | 0.5164 | 7.5631 | 8.3526 | 1.0488 | 1.633 | 4.2617 | 4.1988 | 0.85421 | 4.2151 | 10.3086 | 2.2318 | 1.7224 | 1.169 | 0
| Rel.Std.Dev. |  | 0.3733% | 5.216% | 6.149% | 0.736% | 1.056% | 2.727% | 2.994% | 0.6282% | 2.552% | 7.226% | 1.411% | 1.238% | 0.867% | 0%
| PDM3 |  | -2.84% | 1.842% | -4.596% | 0.08602% | 8.631% | 9.757% | -1.503% | -4.491% | 16.01% | 0.2031% | 11.13% | -2.255% | -5.299% | -8.693%
| Z-Score (Absolute) |  | 0.37 | 0.45 | 0.68 | 0.14 | 1.64 | 1.84 | 0.14 | 0.67 | 2.94 | 0.16 | 2.08 | 0.27 | 0.81 | 1.4
| Analytical results for Ta in OREAS 625 (Indicative Value 1.43 ppm) | [24 empty] | Indicative
Ta | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.20; 1.08: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: 1.6; 1.02: < 0.5; 1.04: 1.1; 1.05: 1.4934; 1.06: 1.9495; 1.07: NA; 1.08: 1.1 | [6 empty] | 1
| 1 | 2 | 1.01: 1.8; 1.02: < 0.5; 1.04: 1.1; 1.05: 1.3667; 1.06: 2.0419; 1.07: NA; 1.08: 1.2 | [6 empty] | 14
| 1 | 3 | 1.01: 1.3; 1.02: < 0.5; 1.04: 1; 1.05: 1.4005; 1.06: 1.9457; 1.07: NA; 1.08: 1.1 | [6 empty] | 16
| 1 | 4 | 1.01: 1.1; 1.02: < 0.5; 1.04: 1; 1.05: 1.4849; 1.06: 2.0656; 1.07: NA; 1.08: 1.2 | [6 empty] | 1.4266
| 1 | 5 | 1.01: 1.8; 1.02: < 0.5; 1.04: 1.1; 1.05: 1.4756; 1.06: 2.0957; 1.07: NA; 1.08: 1.1 | [6 empty] | 20
| 1 | 6 | 1.01: 1.3; 1.02: < 0.5; 1.04: 0.9; 1.05: 1.5136; 1.06: 1.9646; 1.07: NA; 1.08: 1.2
| Mean |  | 1.01: 1.4833; 1.04: 1.0333; 1.05: 1.4558; 1.06: 2.0105; 1.08: 1.15
| Median |  | 1.01: 1.45; 1.04: 1.05; 1.05: 1.4802; 1.06: 2.0032; 1.08: 1.15
| Std Dev. |  | 1.01: 0.29269; 1.04: 0.08165; 1.05: 0.058292; 1.06: 0.065303; 1.08: 0.054772
| Rel.Std.Dev. |  | 1.01: 19.73%; 1.04: 7.902%; 1.05: 4.004%; 1.06: 3.248%; 1.08: 4.763%
| PDM3 |  | 1.01: 3.978%; 1.04: -27.57%; 1.05: 2.046%; 1.06: 40.93%; 1.08: -19.39%
| Z-Score (Absolute) |  | 1.01: 0.54; 1.02: 3.16; 1.04: 0.81; 1.05: 0.46; 1.06: 2.12; 1.07: NA; 1.08: 0.46
| Analytical results for Tb in OREAS 625 (Certified Value 0.77 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.19; 1.07: 1.20; 1.08: 1.22; 1.09: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: BF*MS; 1.07: PF*MS; 1.08: PF*MS; 1.09: PF*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 0.7; 1.02: 0.64; 1.04: 0.8; 1.05: 0.76393; 1.06: 0.761; 1.07: < 2; 1.08: 0.80713; 1.09: 0.8 | [6 empty] | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.68; 1.04: 0.7; 1.05: 0.78983; 1.06: 0.757; 1.07: < 2; 1.08: 0.83509; 1.09: 0.86 | [6 empty] | 36
| 1 | 3 | 1.01: 0.9; 1.02: 0.66; 1.04: 0.8; 1.05: 0.74154; 1.06: 0.754; 1.07: < 2; 1.08: 0.79767; 1.09: 0.82 | [6 empty] | 16
| 1 | 4 | 1.01: 1; 1.02: 0.65; 1.04: 0.7; 1.05: 0.80305; 1.06: 0.758; 1.07: < 2; 1.08: 0.84479; 1.09: 0.88 | [6 empty] | 0.77226
| 1 | 5 | 1.01: 0.8; 1.02: 0.7; 1.04: 0.7; 1.05: 0.79755; 1.06: 0.754; 1.07: < 2; 1.08: 0.81518; 1.09: 0.79 | [6 empty] | 157
| 1 | 6 | 1.01: 0.8; 1.02: 0.67; 1.04: 0.8; 1.05: 0.78702; 1.06: 0.756; 1.07: < 2; 1.08: 0.84508; 1.09: 0.85
| Mean |  | 1.01: 0.81667; 1.02: 0.66667; 1.04: 0.75; 1.05: 0.78049; 1.06: 0.75667; 1.08: 0.82416; 1.09: 0.83333
| Median |  | 1.01: 0.8; 1.02: 0.665; 1.04: 0.75; 1.05: 0.78842; 1.06: 0.7565; 1.08: 0.82514; 1.09: 0.835
| Std Dev. |  | 1.01: 0.1169; 1.02: 0.021602; 1.04: 0.054772; 1.05: 0.023328; 1.06: 0.0026583; 1.08: 0.020275; 1.09: 0.03559
| Rel.Std.Dev. |  | 1.01: 14.31%; 1.02: 3.24%; 1.04: 7.303%; 1.05: 2.989%; 1.06: 0.3513%; 1.08: 2.46%; 1.09: 4.271%
| PDM3 |  | 1.01: 5.75%; 1.02: -13.67%; 1.04: -2.883%; 1.05: 1.065%; 1.06: -2.019%; 1.08: 6.72%; 1.09: 7.908%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 2.39; 1.04: NA; 1.05: 0.38; 1.06: 0.8; 1.07: 3.48; 1.08: 0.38; 1.09: 0.55
| Analytical results for Te in OREAS 625 (Indicative Value < 1 ppm) | [24 empty] | Indicative
Te | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.15 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 1
| 1 | 2 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 15
| 1 | 3 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 16
| 1 | 4 | 1.01: 6; 1.02: < 5; 1.04: < 1 | [6 empty] | < 1
| 1 | 5 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 21
| 1 | 6 | 1.01: < 6; 1.02: < 5; 1.04: < 1
| Mean |  | 1.01: 6
| Median |  | 1.01: 6
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0; 1.04: 1.35
| Analytical results for Th in OREAS 625 (Certified Value 13.7 ppm) | [24 empty] | Certified
Th | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.16; 1.08: 1.17; 1.09: 1.19; 1.10: 1.20; 1.11: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*OES/MS; 1.09: BF*MS; 1.10: PF*MS; 1.11: PF*MS | [6 empty] | 1
2
| 1 | 1 | 1.01: 12.3; 1.02: 12.2; 1.04: 15; 1.05: 14.6051; 1.06: 15.0504; 1.07: 14; 1.08: 14.1; 1.09: 12.366; 1.10: 14; 1.11: 13.8 | [6 empty] | 1
| 1 | 2 | 1.01: 12.4; 1.02: 12.8; 1.04: 14.7; 1.05: 14.2754; 1.06: 15.2089; 1.07: 13.9; 1.08: 14.1; 1.09: 12.242; 1.10: 13; 1.11: 13.8 | [6 empty] | 38
| 1 | 3 | 1.01: 12.3; 1.02: 12.2; 1.04: 14.6; 1.05: 14.2264; 1.06: 14.7645; 1.07: 14.5; 1.08: 13.9; 1.09: 12.256; 1.10: 14; 1.11: 14 | [6 empty] | 16
| 1 | 4 | 1.01: 13.5; 1.02: 12.3; 1.04: 14.6; 1.05: 14.208; 1.06: 14.9233; 1.07: 14.3; 1.08: 14.4; 1.09: 12.384; 1.10: 14; 1.11: 14.2 | [6 empty] | 13.7295
| 1 | 5 | 1.01: 12.9; 1.02: 12.6; 1.04: 14.7; 1.05: 13.0577; 1.06: 14.9241; 1.07: 14.4; 1.08: 14.2; 1.09: 12.355; 1.10: 13; 1.11: 13.5 | [6 empty] | 158
| 1 | 6 | 1.01: 13.2; 1.02: 12.3; 1.04: 14.3; 1.05: 14.1011; 1.06: 15.2429; 1.07: 13.9; 1.08: 14.3; 1.09: 12.354; 1.10: 14; 1.11: 13.8
| Mean |  | 1.01: 12.7667; 1.02: 12.4; 1.04: 14.65; 1.05: 14.079; 1.06: 15.019; 1.07: 14.1667; 1.08: 14.1667; 1.09: 12.3262; 1.10: 13.6667; 1.11: 13.85
| Median |  | 1.01: 12.65; 1.02: 12.3; 1.04: 14.65; 1.05: 14.2172; 1.06: 14.9873; 1.07: 14.15; 1.08: 14.15; 1.09: 12.3545; 1.10: 14; 1.11: 13.8
| Std Dev. |  | 1.01: 0.51251; 1.02: 0.24495; 1.04: 0.22583; 1.05: 0.52862; 1.06: 0.18444; 1.07: 0.26583; 1.08: 0.17512; 1.09: 0.060901; 1.10: 0.5164; 1.11: 0.23452
| Rel.Std.Dev. |  | 1.01: 4.014%; 1.02: 1.975%; 1.04: 1.542%; 1.05: 3.755%; 1.06: 1.228%; 1.07: 1.876%; 1.08: 1.236%; 1.09: 0.4941%; 1.10: 3.779%; 1.11: 1.693%
| PDM3 |  | 1.01: -7.013%; 1.02: -9.684%; 1.04: 6.704%; 1.05: 2.545%; 1.06: 9.392%; 1.07: 3.184%; 1.08: 3.184%; 1.09: -10.22%; 1.10: -0.4577%; 1.11: 0.8776%
| Z-Score (Absolute) |  | 1.01: 1.64; 1.02: 2.15; 1.04: 0.94; 1.05: 0.16; 1.06: 1.45; 1.07: 0.28; 1.08: 0.28; 1.09: 2.25; 1.10: 0.41; 1.11: 0.16
| Analytical results for Ti in OREAS 625 (Certified Value 0.129 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 3
3
| 1 | 1 | 0.14 | 0.13 | 0.13 | 0.126 | 0.126 | 0.126 | 0.126 | 0.126 | 0.13 | 0.12 | 0.12 | 0.13 | 0.13683 | 0.17467 | 0.12 | 0.129 | 0.1419 | 0.13 | 0.13 | 0.12 | 0.14 |  |  | 1
| 1 | 2 | 0.14 | 0.13 | 0.12 | 0.126 | 0.132 | 0.132 | 0.192 | 0.126 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13834 | 0.17995 | 0.12 | 0.129 | 0.139 | 0.13 | 0.13 | 0.12 | 0.14 |  |  | 39
| 1 | 3 | 0.14 | 0.13 | 0.12 | 0.126 | 0.126 | 0.132 | 0.204 | 0.126 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13522 | 0.13775 | 0.12 | 0.127 | 0.135 | 0.13 | 0.13 | 0.13 | 0.14 |  |  | 16
| 1 | 4 | 0.14 | 0.13 | 0.13 | 0.126 | 0.132 | 0.126 | 0.126 | 0.144 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13653 | 0.16351 | 0.12 | 0.126 | 0.1352 | 0.13 | 0.12 | 0.13 | 0.14 |  |  | 0.12943
| 1 | 5 | 0.14 | 0.13 | 0.12 | 0.126 | 0.126 | 0.132 | 0.126 | 0.132 | 0.13 | 0.12 | 0.12 | 0.13 | 0.13688 | 0.14013 | 0.12 | 0.127 | 0.1372 | 0.13 | 0.13 | 0.12 | 0.14 |  |  | 159
| 1 | 6 | 0.14 | 0.13 | 0.13 | 0.126 | 0.126 | 0.126 | 0.126 | 0.126 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13673 | 0.17307 | 0.12 | 0.129 | 0.1348 | 0.13 | 0.13 | 0.12 | 0.13
| Mean |  | 0.14 | 0.13 | 0.125 | 0.126 | 0.128 | 0.129 | 0.15 | 0.13 | 0.13 | 0.12667 | 0.12667 | 0.13 | 0.13675 | 0.16151 | 0.12 | 0.12783 | 0.13718 | 0.13 | 0.12833 | 0.12333 | 0.13833
| Median |  | 0.14 | 0.13 | 0.125 | 0.126 | 0.126 | 0.129 | 0.126 | 0.126 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13678 | 0.16829 | 0.12 | 0.128 | 0.1362 | 0.13 | 0.13 | 0.12 | 0.14
| Std Dev. |  | 0 | 0 | 0.0054772 | 0 | 0.0030984 | 0.0032863 | 0.037374 | 0.0072664 | 0 | 0.005164 | 0.005164 | 0 | 0.00099464 | 0.018289 | 0 | 0.0013292 | 0.0028259 | 0 | 0.0040825 | 0.005164 | 0.0040825
| Rel.Std.Dev. |  | 0% | 0% | 4.382% | 0% | 2.421% | 2.548% | 24.92% | 5.59% | 0% | 4.077% | 4.077% | 0% | 0.7273% | 11.32% | 0% | 1.04% | 2.06% | 0% | 3.181% | 4.187% | 2.951%
| PDM3 |  | 8.168% | 0.4422% | -3.421% | -2.648% | -1.103% | -0.3305% | 15.89% | 0.4422% | 0.4422% | -2.133% | -2.133% | 0.4422% | 5.66% | 24.79% | -7.284% | -1.232% | 5.992% | 0.4422% | -0.8455% | -4.709% | 6.881%
| Z-Score (Absolute) |  | 2.12 | 0.02 | 1.02 | 0.84 | 0.42 | 0.21 | 4.18 | 0 | 0.02 | 0.67 | 0.67 | 0.02 | 1.44 | 6.62 | 2.07 | 0.43 | 1.53 | 0.02 | 0.33 | 1.37 | 1.77
| Analytical results for Tl in OREAS 625 (Certified Value 9.99 ppm) | [24 empty] | Certified
Tl | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.16; 1.06: 1.17; 1.07: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 10; 1.02: 8.9; 1.03: 10.4; 1.04: 9.7314; 1.05: 10.7; 1.06: 11.1; 1.07: 9.3 |  |  | 1
| 1 | 2 | 1.01: 9.9; 1.02: 9.2; 1.03: 10.6; 1.04: 9.6794; 1.05: 10.5; 1.06: 10.8; 1.07: 9.3 |  |  | 0
| 1 | 3 | 1.01: 10.5; 1.02: 9; 1.03: 10.3; 1.04: 9.3645; 1.05: 10.7; 1.06: 11.1; 1.07: 9.1 |  |  | 16
| 1 | 4 | 1.01: 10; 1.02: 9; 1.03: 10.5; 1.04: 9.1769; 1.05: 10.6; 1.06: 10.8; 1.07: 9.7 |  |  | 9.9868
| 1 | 5 | 1.01: 9.9; 1.02: 9.3; 1.03: 10.3; 1.04: 9.3217; 1.05: 10.7; 1.06: 11.3; 1.07: 9.2 |  |  | 160
| 1 | 6 | 1.01: 9.9; 1.02: 9; 1.03: 10.4; 1.04: 9.8336; 1.05: 10.8; 1.06: 10.7; 1.07: 9.4
| Mean |  | 1.01: 10.0333; 1.02: 9.0667; 1.03: 10.4167; 1.04: 9.5179; 1.05: 10.6667; 1.06: 10.9667; 1.07: 9.3333
| Median |  | 1.01: 9.95; 1.02: 9; 1.03: 10.4; 1.04: 9.5219; 1.05: 10.7; 1.06: 10.95; 1.07: 9.3
| Std Dev. |  | 1.01: 0.23381; 1.02: 0.15055; 1.03: 0.1169; 1.04: 0.26444; 1.05: 0.10328; 1.06: 0.23381; 1.07: 0.20656
| Rel.Std.Dev. |  | 1.01: 2.33%; 1.02: 1.661%; 1.03: 1.122%; 1.04: 2.778%; 1.05: 0.9682%; 1.06: 2.132%; 1.07: 2.213%
| PDM3 |  | 1.01: 0.4655%; 1.02: -9.214%; 1.03: 4.304%; 1.04: -4.695%; 1.05: 6.807%; 1.06: 9.811%; 1.07: -6.544%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 1.03; 1.03: 0.41; 1.04: 0.55; 1.05: 0.67; 1.06: 0.99; 1.07: 0.75
| Analytical results for Tm in OREAS 625 (Certified Value 0.18 ppm) | [24 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.2; 1.02: 0.16; 1.03: 0.2; 1.04: 0.17464; 1.05: 0.164; 1.06: < 2; 1.07: 0.18 |  |  | 1
| 1 | 2 | 1.01: 0.2; 1.02: 0.17; 1.03: 0.2; 1.04: 0.17735; 1.05: 0.164; 1.06: < 2; 1.07: 0.19 |  |  | 9
| 1 | 3 | 1.01: 0.2; 1.02: 0.16; 1.03: 0.2; 1.04: 0.17001; 1.05: 0.167; 1.06: < 2; 1.07: 0.19 |  |  | 16
| 1 | 4 | 1.01: 0.2; 1.02: 0.16; 1.03: 0.2; 1.04: 0.19394; 1.05: 0.165; 1.06: < 2; 1.07: 0.19 |  |  | 0.18296
| 1 | 5 | 1.01: 0.2; 1.02: 0.16; 1.03: 0.2; 1.04: 0.21358; 1.05: 0.162; 1.06: < 2; 1.07: 0.19 |  |  | 161
| 1 | 6 | 1.01: 0.2; 1.02: 0.16; 1.03: 0.2; 1.04: 0.15903; 1.05: 0.166; 1.06: < 2; 1.07: 0.2
| Mean |  | 1.01: 0.2; 1.02: 0.16167; 1.03: 0.2; 1.04: 0.18142; 1.05: 0.16467; 1.07: 0.19
| Median |  | 1.01: 0.2; 1.02: 0.16; 1.03: 0.2; 1.04: 0.176; 1.05: 0.1645; 1.07: 0.19
| Std Dev. |  | 1.01: 3.0405e-17; 1.02: 0.0040825; 1.03: 3.0405e-17; 1.04: 0.019418; 1.05: 0.0017512; 1.07: 0.0063246
| Rel.Std.Dev. |  | 1.01: 1.52e-14%; 1.02: 2.525%; 1.03: 1.52e-14%; 1.04: 10.7%; 1.05: 1.063%; 1.07: 3.329%
| PDM3 |  | 1.01: 9.314%; 1.02: -11.64%; 1.03: 9.314%; 1.04: -0.839%; 1.05: -9.998%; 1.07: 3.848%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 1.91; 1.03: 0.67; 1.04: 0.58; 1.05: 1.71; 1.06: 54.62; 1.07: 0
| Analytical results for U in OREAS 625 (Certified Value 5.78 ppm) | [24 empty] | Certified
U | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.15; 1.06: 1.17; 1.07: 1.19; 1.08: 1.20; 1.09: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: BF*MS; 1.08: PF*MS; 1.09: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 5.3; 1.02: 5.1; 1.03: 6; 1.04: 7.1178; 1.05: 6.2224; 1.06: 6.3; 1.07: 5.219; 1.08: 8; 1.09: 5.97 |  |  | 1
| 1 | 2 | 1.01: 5.5; 1.02: 5.4; 1.03: 6; 1.04: 7.2836; 1.05: 5.9291; 1.06: 6.43; 1.07: 5.108; 1.08: 7; 1.09: 5.9 |  |  | 41
| 1 | 3 | 1.01: 5.2; 1.02: 5.1; 1.03: 5.7; 1.04: 5.7248; 1.05: 5.6687; 1.06: 6.33; 1.07: 5.214; 1.08: 7; 1.09: 5.98 |  |  | 16
| 1 | 4 | 1.01: 5.6; 1.02: 5.1; 1.03: 6.2; 1.04: 6.1776; 1.05: 5.6687; 1.06: 6.17; 1.07: 5.221; 1.08: 7; 1.09: 6.05 |  |  | 5.7822
| 1 | 5 | 1.01: 5.4; 1.02: 5.4; 1.03: 5.9; 1.04: 6.186; 1.05: 5.9468; 1.06: 6.23; 1.07: 5.226; 1.08: 7; 1.09: 5.74 |  |  | 162
| 1 | 6 | 1.01: 5.2; 1.02: 5.2; 1.03: 5.9; 1.04: 5.8312; 1.05: 6.1232; 1.06: 6.28; 1.07: 5.219; 1.08: 7; 1.09: 5.88
| Mean |  | 1.01: 5.3667; 1.02: 5.2167; 1.03: 5.95; 1.04: 6.3868; 1.05: 5.9265; 1.06: 6.29; 1.07: 5.2012; 1.08: 7.1667; 1.09: 5.92
| Median |  | 1.01: 5.35; 1.02: 5.15; 1.03: 5.95; 1.04: 6.1818; 1.05: 5.9379; 1.06: 6.29; 1.07: 5.219; 1.08: 7; 1.09: 5.935
| Std Dev. |  | 1.01: 0.1633; 1.02: 0.1472; 1.03: 0.16432; 1.04: 0.65873; 1.05: 0.22785; 1.06: 0.088769; 1.07: 0.045806; 1.08: 0.40825; 1.09: 0.10714
| Rel.Std.Dev. |  | 1.01: 3.043%; 1.02: 2.822%; 1.03: 2.762%; 1.04: 10.31%; 1.05: 3.845%; 1.06: 1.411%; 1.07: 0.8807%; 1.08: 5.696%; 1.09: 1.81%
| PDM3 |  | 1.01: -7.187%; 1.02: -9.781%; 1.03: 2.901%; 1.04: 10.46%; 1.05: 2.495%; 1.06: 8.782%; 1.07: -10.05%; 1.08: 23.94%; 1.09: 2.383%
| Z-Score (Absolute) |  | 1.01: 0.9; 1.02: 1.15; 1.03: 0.04; 1.04: 0.75; 1.05: 0.01; 1.06: 0.6; 1.07: 1.17; 1.08: NA; 1.09: 0.01
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 625 (Certified Value 11.5 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | [10 empty] | 1
1
| 1 | 1 | 12 | 9 | < 100 | 12 | < 20 | < 20 | < 20 | 17.1 | 10 | NA | 8 | < 10 | 10 | [10 empty] | 1
| 1 | 2 | 12 | 10 | < 100 | 12 | < 20 | < 20 | < 20 | 15.7 | 10 | NA | 9 | < 10 | 10 | [10 empty] | 23
| 1 | 3 | 12 | 15 | < 100 | 12 | < 20 | < 20 | < 20 | 14.9 | 10 | NA | 10 | < 10 | 10 | [10 empty] | 16
| 1 | 4 | 12 | 10 | < 100 | 12 | < 20 | < 20 | < 20 | 15.6 | 10 | NA | 11 | < 10 | 10 | [10 empty] | 11.5024
| 1 | 5 | 12 | 11 | < 100 | 12 | < 20 | < 20 | < 20 | 14.4 | 10 | NA | 8 | < 10 | 10 | [10 empty] | 163
| 1 | 6 | 12 | 15 | < 100 | 12 | < 20 | < 20 | < 20 | 15.4 | 10 | NA | 10 | < 10 | 10
| Mean |  | 1.01: 12; 1.02: 11.6667; 1.04: 12; 1.08: 15.5167; 1.09: 10; 1.11: 9.3333; 1.13: 10
| Median |  | 1.01: 12; 1.02: 10.5; 1.04: 12; 1.08: 15.5; 1.09: 10; 1.11: 9.5; 1.13: 10
| Std Dev. |  | 1.01: 0; 1.02: 2.6583; 1.04: 0; 1.08: 0.91524; 1.09: 0; 1.11: 1.2111; 1.13: 0
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 22.79%; 1.04: 0%; 1.08: 5.898%; 1.09: 0%; 1.11: 12.98%; 1.13: 0%
| PDM3 |  | 1.01: 4.326%; 1.02: 1.428%; 1.04: 4.326%; 1.08: 34.9%; 1.09: -13.06%; 1.11: -18.86%; 1.13: -13.06%
| Z-Score (Absolute) |  | 1.16 | 0.96 | 23.12 | 1.16 | 0 | 0 | 0 | 3.19 | 0 | NA | 0.39 | 2.89 | 0
| Analytical results for W in OREAS 625 (Certified Value 4.39 ppm) | [24 empty] | Certified
W | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.09; 1.05: 1.10; 1.06: 1.15; 1.07: 1.16; 1.08: 1.20; 1.09: 1.23; 1.10: 1.25; 1.11: 1.26 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*MS; 1.10: PF*OES; 1.11: PF*OES |  |  | 2
2
| 1 | 1 | 1.01: 346; 1.02: 5.3; 1.03: < 100; 1.04: < 5; 1.05: 4; 1.06: 4.5542; 1.07: < 5; 1.08: 4; 1.09: 4; 1.10: 66; 1.11: < 50 |  |  | 1
| 1 | 2 | 1.01: 346; 1.02: 4; 1.03: < 100; 1.04: < 5; 1.05: 5; 1.06: 4.288; 1.07: < 5; 1.08: 4; 1.09: 4; 1.10: 65; 1.11: < 50 |  |  | 43
| 1 | 3 | 1.01: 348; 1.02: 4.3; 1.03: < 100; 1.04: < 5; 1.05: 4; 1.06: 4.3163; 1.07: < 5; 1.08: 5; 1.09: 4; 1.10: 68; 1.11: < 50 |  |  | 16
| 1 | 4 | 1.01: 344; 1.02: 4; 1.03: < 100; 1.04: < 5; 1.05: 4; 1.06: 4.3497; 1.07: < 5; 1.08: 5; 1.09: 5; 1.10: 64; 1.11: < 50 |  |  | 4.3931
| 1 | 5 | 1.01: 346; 1.02: 4.4; 1.03: < 100; 1.04: < 5; 1.05: 4; 1.06: 4.3975; 1.07: < 5; 1.08: 6; 1.09: 4; 1.10: 64; 1.11: < 50 |  |  | 164
| 1 | 6 | 1.01: 349; 1.02: 4.4; 1.03: < 100; 1.04: < 5; 1.05: 4; 1.06: 4.5682; 1.07: < 5; 1.08: 6; 1.09: 4; 1.10: 68; 1.11: < 50
| Mean |  | 1.01: 346.5; 1.02: 4.4; 1.05: 4.1667; 1.06: 4.4123; 1.08: 5; 1.09: 4.1667; 1.10: 65.8333
| Median |  | 1.01: 346; 1.02: 4.35; 1.05: 4; 1.06: 4.3736; 1.08: 5; 1.09: 4; 1.10: 65.5
| Std Dev. |  | 1.01: 1.7607; 1.02: 0.47749; 1.05: 0.40825; 1.06: 0.12102; 1.08: 0.89443; 1.09: 0.40825; 1.10: 1.8348
| Rel.Std.Dev. |  | 1.01: 0.5081%; 1.02: 10.85%; 1.05: 9.798%; 1.06: 2.743%; 1.08: 17.89%; 1.09: 9.798%; 1.10: 2.787%
| PDM3 |  | 1.01: 7787%; 1.02: 0.1563%; 1.05: -5.155%; 1.06: 0.437%; 1.08: 13.81%; 1.09: -5.155%; 1.10: 1399%
| Z-Score (Absolute) |  | 1.01: 120.62; 1.02: 0; 1.03: 16.07; 1.04: 0.67; 1.05: 0.09; 1.06: 0; 1.07: 0.67; 1.08: 0.21; 1.09: 0.09; 1.10: 21.66; 1.11: 7.26
| Analytical results for Y in OREAS 625 (Certified Value 17.6 ppm) | [24 empty] | Certified
Y | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.17; 1.08: 1.19; 1.09: 1.20; 1.10: 1.23; 1.11: 1.25; 1.12: 1.26 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES/MS; 1.08: BF*MS; 1.09: PF*MS; 1.10: PF*MS; 1.11: PF*OES; 1.12: PF*OES |  |  | 1
2
| 1 | 1 | 1.01: 19; 1.02: 16.8; 1.03: 15.8; 1.04: 17.1; 1.05: 17.3288; 1.06: 18.9234; 1.07: 19.4; 1.08: 17.523; 1.09: 33; 1.10: 16.5; 1.11: 15; 1.12: 17 |  |  | 1
| 1 | 2 | 1.01: 19; 1.02: 17.6; 1.03: 16.8; 1.04: 16.9; 1.05: 17.9925; 1.06: 18.5899; 1.07: 18.9; 1.08: 17.55; 1.09: 31; 1.10: 17.4; 1.11: 15; 1.12: 17 |  |  | 44
| 1 | 3 | 1.01: 19; 1.02: 17.7; 1.03: 15.5; 1.04: 16.8; 1.05: 17.4263; 1.06: 18.0028; 1.07: 19.3; 1.08: 17.512; 1.09: 33; 1.10: 16.9; 1.11: 15; 1.12: 17 |  |  | 16
| 1 | 4 | 1.01: 19; 1.02: 17.8; 1.03: 16.2; 1.04: 16.8; 1.05: 17.4002; 1.06: 18.2877; 1.07: 18.8; 1.08: 17.551; 1.09: 31; 1.10: 17.2; 1.11: 15; 1.12: 17 |  |  | 17.5649
| 1 | 5 | 1.01: 19; 1.02: 17.4; 1.03: 16.3; 1.04: 17.2; 1.05: 17.7459; 1.06: 18.3099; 1.07: 19.8; 1.08: 17.544; 1.09: 22; 1.10: 16.6; 1.11: 15; 1.12: 17 |  |  | 165
| 1 | 6 | 1.01: 19; 1.02: 18.3; 1.03: 16.1; 1.04: 17.6; 1.05: 17.6862; 1.06: 18.7084; 1.07: 19.1; 1.08: 17.532; 1.09: 23; 1.10: 16.9; 1.11: 16; 1.12: 16
| Mean |  | 1.01: 19; 1.02: 17.6; 1.03: 16.1167; 1.04: 17.0667; 1.05: 17.5966; 1.06: 18.4704; 1.07: 19.2167; 1.08: 17.5353; 1.09: 28.8333; 1.10: 16.9167; 1.11: 15.1667; 1.12: 16.8333
| Median |  | 1.01: 19; 1.02: 17.65; 1.03: 16.15; 1.04: 17; 1.05: 17.5563; 1.06: 18.4499; 1.07: 19.2; 1.08: 17.538; 1.09: 31; 1.10: 16.9; 1.11: 15; 1.12: 17
| Std Dev. |  | 1.01: 0; 1.02: 0.49396; 1.03: 0.4446; 1.04: 0.30768; 1.05: 0.25547; 1.06: 0.33283; 1.07: 0.3656; 1.08: 0.015769; 1.09: 4.9967; 1.10: 0.34303; 1.11: 0.40825; 1.12: 0.40825
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 2.807%; 1.03: 2.759%; 1.04: 1.803%; 1.05: 1.452%; 1.06: 1.802%; 1.07: 1.903%; 1.08: 0.08993%; 1.09: 17.33%; 1.10: 2.028%; 1.11: 2.692%; 1.12: 2.425%
| PDM3 |  | 1.01: 8.17%; 1.02: 0.2%; 1.03: -8.245%; 1.04: -2.836%; 1.05: 0.1808%; 1.06: 5.155%; 1.07: 9.404%; 1.08: -0.1682%; 1.09: 64.15%; 1.10: -3.69%; 1.11: -13.65%; 1.12: -4.165%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.04; 1.03: 1.7; 1.04: 0.59; 1.05: 0.04; 1.06: 1.06; 1.07: 1.94; 1.08: 0.04; 1.09: NA; 1.10: 0.76; 1.11: NA; 1.12: NA
Results from laboratories 1.01, 1.20, 1.25 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 625 (Certified Value 1.05 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 1; 1.02: 0.9; 1.03: 1.1; 1.04: 1.093; 1.05: 1.041; 1.06: < 2; 1.07: 1.0929; 1.08: 1.1 |  |  | 1
| 1 | 2 | 1.01: 1; 1.02: 1; 1.03: 1; 1.04: 1.1353; 1.05: 1.023; 1.06: < 2; 1.07: 1.0885; 1.08: 1.1 |  |  | 45
| 1 | 3 | 1.01: 1; 1.02: 0.9; 1.03: 1.1; 1.04: 1.0059; 1.05: 1.032; 1.06: < 2; 1.07: 1.0357; 1.08: 1.1 |  |  | 16
| 1 | 4 | 1.01: 1.1; 1.02: 1; 1.03: 1; 1.04: 1.1031; 1.05: 1.034; 1.06: < 2; 1.07: 0.99795; 1.08: 1.2 |  |  | 1.0549
| 1 | 5 | 1.01: 1.1; 1.02: 1; 1.03: 1.2; 1.04: 1.0848; 1.05: 1.036; 1.06: < 2; 1.07: 1.0651; 1.08: 1 |  |  | 166
| 1 | 6 | 1.01: 1; 1.02: 1; 1.03: 1.3; 1.04: 1.1278; 1.05: 1.038; 1.06: < 2; 1.07: 1.0905; 1.08: 1.1
| Mean |  | 1.01: 1.0333; 1.02: 0.96667; 1.03: 1.1167; 1.04: 1.0917; 1.05: 1.034; 1.07: 1.0618; 1.08: 1.1
| Median |  | 1.01: 1; 1.02: 1; 1.03: 1.1; 1.04: 1.098; 1.05: 1.035; 1.07: 1.0768; 1.08: 1.1
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.05164; 1.03: 0.1169; 1.04: 0.046343; 1.05: 0.006229; 1.07: 0.038131; 1.08: 0.063246
| Rel.Std.Dev. |  | 1.01: 4.997%; 1.02: 5.342%; 1.03: 10.47%; 1.04: 4.245%; 1.05: 0.6024%; 1.07: 3.591%; 1.08: 5.75%
| PDM3 |  | 1.01: -2.048%; 1.02: -8.367%; 1.03: 5.851%; 1.04: 3.48%; 1.05: -1.985%; 1.07: 0.6478%; 1.08: 4.272%
| Z-Score (Absolute) |  | 1.01: 0.21; 1.02: 1.2; 1.03: 1.01; 1.04: 0.64; 1.05: 0.2; 1.06: 0.7; 1.07: 0.2; 1.08: 0.77
| Analytical results for Zn in OREAS 625 (Certified Value 3.2 wt.%) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 3.22 | 3.19 | 3.2284 | 3.2 | 3.24 | 3.27 | 3.15 | 3.21 | 3.39 | 3.1129 | 3 | 3.0009 | 2.9651 | > 1 | 3.14 | 3.3733 | 3.3102 | 3.12 | 3.2289 | 3.17 | 3.17 | 3.15 |  | 1
| 1 | 2 | 3.23 | 3.2 | 3.0729 | 3.19 | 3.29 | 3.01 | 3.13 | 3.16 | 3.35 | 3.1963 | 3 | 2.9307 | 3.3379 | > 1 | 3.12 | 3.36 | 3.3659 | 3.2 | 3.225 | 3.245 | 3.25 | 3.12 |  | 24
| 1 | 3 | 3.21 | 3.24 | 3.1063 | 3.21 | 3.24 | 3.05 | 3.09 | 3.2 | 3.22 | 3.1506 | 3.04 | 2.9433 | 3.2023 | > 1 | 3.18 | 3.3685 | 3.2975 | 3.14 | 3.2197 | 3.169 | 3.29 | 3.13 |  | 16
| 1 | 4 | 3.24 | 3.21 | 3.3109 | 3.2 | 3.29 | 3.21 | 3.24 | 3.21 | 3.37 | 3.187 | 3.04 | 2.9376 | 3.1072 | > 1 | 3.16 | 3.3712 | 3.353 | 3.1 | 3.2226 | 3.176 | 3.32 | 3.14 |  | 3.2007
| 1 | 5 | 3.23 | 3.19 | 3.1686 | 3.18 | 3.26 | 3.05 | 3.19 | 3.17 | 3.38 | 3.2243 | 3 | 2.9469 | 3.242 | > 1 | 3.17 | 3.3715 | 3.3066 | 3.11 | 3.2369 | 3.182 | 3.22 | 3.04 |  | 167
| 1 | 6 | 3.25 | 3.22 | 3.2356 | 3.22 | 3.27 | 3.21 | 3.14 | 3.16 | 3.29 | 3.2296 | 2.99 | 3.0049 | 2.9791 | > 1 | 3.17 | 3.3674 | 3.3232 | 3.09 | 3.237 | 3.165 | 3.29 | 3.1
| Mean |  | 3.23 | 3.2083 | 3.1871 | 3.2 | 3.265 | 3.1333 | 3.1567 | 3.185 | 3.3333 | 3.1834 | 3.0117 | 2.9607 | 3.1389 |  | 3.1567 | 3.3687 | 3.3261 | 3.1267 | 3.2283 | 3.1845 | 3.2567 | 3.1133
| Median |  | 3.23 | 3.205 | 3.1985 | 3.2 | 3.265 | 3.13 | 3.145 | 3.185 | 3.36 | 3.1916 | 3 | 2.9451 | 3.1548 |  | 3.165 | 3.3699 | 3.3167 | 3.115 | 3.2269 | 3.173 | 3.27 | 3.125
| Std Dev. |  | 0.014142 | 0.019408 | 0.088653 | 0.014142 | 0.022583 | 0.10912 | 0.052026 | 0.02429 | 0.065929 | 0.044804 | 0.022286 | 0.033156 | 0.14902 |  | 0.022509 | 0.0047467 | 0.027448 | 0.039833 | 0.0073121 | 0.030231 | 0.055015 | 0.039833
| Rel.Std.Dev. |  | 0.4378% | 0.6049% | 2.782% | 0.4419% | 0.6917% | 3.482% | 1.648% | 0.7626% | 1.978% | 1.407% | 0.74% | 1.12% | 4.747% |  | 0.7131% | 0.1409% | 0.8253% | 1.274% | 0.2265% | 0.9493% | 1.689% | 1.279%
| PDM3 |  | 0.9141% | 0.2371% | -0.4257% | -0.02321% | 2.008% | -2.106% | -1.377% | -0.4919% | 4.142% | -0.5412% | -5.907% | -7.499% | -1.932% |  | -1.377% | 5.246% | 3.915% | -2.314% | 0.8625% | -0.5075% | 1.747% | -2.731%
| Z-Score (Absolute) |  | 0.66 | 0.34 | 0.03 | 0.22 | 1.17 | 0.76 | 0.41 | 0 | 2.17 | 0.02 | 2.54 | 3.28 | 0.67 | NA | 0.41 | 2.69 | 2.06 | 0.85 | 0.63 | 0.01 | 1.05 | 1.05
| Analytical results for Zr in OREAS 625 (Indicative Value 248 ppm) | [24 empty] | Indicative
Zr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab |  | 1
| No | No | 1.01: 1.10; 1.02: 1.13; 1.03: 1.16; 1.04: 1.20; 1.05: 1.23 |  | ppm
|  |  | 1.01: PF*MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*MS; 1.05: PF*MS |  | 0
0
| 1 | 1 | 1.01: 260; 1.02: 253.3445; 1.03: 257.8; 1.04: 250; 1.05: 237 |  | 1
| 1 | 2 | 1.01: 256; 1.02: 248.3259; 1.03: 224.9; 1.04: 256; 1.05: 243 |  | 16
| 1 | 3 | 1.01: 250; 1.02: 252.7277; 1.03: 221.9; 1.04: 247; 1.05: 240 |  | 16
| 1 | 4 | 1.01: 249; 1.02: 249.5659; 1.03: 224.3; 1.04: 245; 1.05: 245 |  | 248.0719
| 1 | 5 | 1.01: 254; 1.02: 237.4442; 1.03: 222.5; 1.04: 246; 1.05: 238 |  | 22
| 1 | 6 | 1.01: 251; 1.02: 259.3175; 1.03: 221.6; 1.04: 243; 1.05: 243
| Mean |  | 1.01: 253.3333; 1.02: 250.121; 1.03: 228.8333; 1.04: 247.8333; 1.05: 241
| Median |  | 1.01: 252.5; 1.02: 251.1468; 1.03: 223.4; 1.04: 246.5; 1.05: 241.5
| Std Dev. |  | 1.01: 4.1793; 1.02: 7.2937; 1.03: 14.252; 1.04: 4.6224; 1.05: 3.1623
| Rel.Std.Dev. |  | 1.01: 1.65%; 1.02: 2.916%; 1.03: 6.228%; 1.04: 1.865%; 1.05: 1.312%
| PDM3 |  | 1.01: 2.121%; 1.02: 0.826%; 1.03: -7.755%; 1.04: -0.09617%; 1.05: -2.851%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.28; 1.03: 2.33; 1.04: 0; 1.05: 0.84
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C672: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C690: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C925: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C943: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C961: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1015: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1033: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 625 (Indicative Value 0.384 wt.%) | [28 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.12 | 1.20 | 3.01 | [24 empty] | wt.%
|  |  | IRC | IRC | IRC | [24 empty] | 3
3
| 1 | 1 | 0.38 | 0.4 | 0.37 | [24 empty] | 1
| 1 | 2 | 0.39 | 0.4 | 0.37 | [24 empty] | 18
| 1 | 3 | 0.41 | 0.39 | [25 empty] | 16
| 1 | 4 | 0.4 | 0.4 | [25 empty] | 0.38389
| 1 | 5 | 0.37 | 0.38 | [25 empty] | 24
| 1 | 6 | 0.38 | 0.39
| Mean |  | 0.38833 | 0.39333 | 0.37
| Median |  | 0.385 | 0.395 | 0.37
| Std Dev. |  | 0.01472 | 0.008165 | 0
| Rel.Std.Dev. |  | 3.79% | 2.076% | 0%
| PDM3 |  | 1.158% | 2.46% | -3.618%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.47
| Analytical results for S in OREAS 625 (Certified Value 3.91 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 3.01 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 3.81 | 4.05 | 3.96 | 3.82 | 4.02 | 3.93 | 3.79 | 3.8 | 4.24 | 4 | 4.01 | 3.84 | 3.8 | 3.97 | 3.7577 | 3.9 | 3.84 | 3.9 | 3.33 | 4.06 | 3.92 | 3.835 | 3.45 | 3.69 | 3.93 | 3.92 | 3.65 | 1
| 1 | 2 | 3.8 | 4.05 | 4.02 | 3.9 | 4.01 | 3.9 | 3.87 | 3.69 | 4.07 | 3.99 | 4.01 | 3.83 | 3.82 | 3.98 | 3.7288 | 3.88 | 3.79 | 3.9 | 3.35 | 4 | 3.89 | 3.802 | 3.4 | 3.62 | 3.95 | 3.96 | 3.94 | 26
| 1 | 3 | 3.79 | 4.01 | 4.05 | 3.89 | 4.03 | 4 | 3.89 | 4.04 | 4.27 | 3.99 | 4 | 3.84 | 3.8 | 3.87 | 3.7875 | 3.89 | 3.83 | 3.87 | 3.45 | 4 | 3.9 | 3.825 | 3.49 | 3.61 | 3.96 | 3.98 |  | 16
| 1 | 4 | 3.82 | 4 | 3.96 | 3.9 | 3.99 | 3.9 | 3.91 | 4.04 | 4.19 | 4.03 | 4.07 | 3.87 | 3.81 | 3.97 | 3.723 | 3.89 | 3.84 | 3.87 | 3.4 | 4.04 | 3.93 | 3.813 | 3.51 | 3.69 | 3.95 | 3.96 |  | 3.9116
| 1 | 5 | 3.8 | 4 | 3.97 | 3.93 | 4.1 | 3.89 | 3.81 | 4.06 | 4.2 | 3.99 | 4.01 | 3.85 | 3.79 | 4 | 3.8424 | 3.88 | 3.84 | 3.86 | 3.42 | 3.88 | 3.95 | 3.834 | 3.43 | 3.68 | 3.95 | 3.94 |  | 169
| 1 | 6 | 3.82 | 3.96 | 4.02 | 3.95 | 4.07 | 3.82 | 3.95 | 4.04 | 4.13 | 3.97 | 4.01 | 3.83 | 3.81 | 4.06 | 3.8121 | 3.88 | 3.86 | 3.88 | 3.55 | 3.96 | 3.91 | 3.851 | 3.52 | 3.67 | 3.94 | 3.97
| Mean |  | 3.8067 | 4.0117 | 3.9967 | 3.8983 | 4.0367 | 3.9067 | 3.87 | 3.945 | 4.1833 | 3.995 | 4.0183 | 3.8433 | 3.805 | 3.975 | 3.7752 | 3.8867 | 3.8333 | 3.88 | 3.4167 | 3.99 | 3.9167 | 3.8267 | 3.4667 | 3.66 | 3.9467 | 3.955 | 3.795
| Median |  | 3.805 | 4.005 | 3.995 | 3.9 | 4.025 | 3.9 | 3.88 | 4.04 | 4.195 | 3.99 | 4.01 | 3.84 | 3.805 | 3.975 | 3.7726 | 3.885 | 3.84 | 3.875 | 3.41 | 4 | 3.915 | 3.8295 | 3.47 | 3.675 | 3.95 | 3.96 | 3.795
| Std Dev. |  | 0.012111 | 0.034303 | 0.038297 | 0.04446 | 0.040825 | 0.058538 | 0.060663 | 0.15897 | 0.073121 | 0.019748 | 0.025626 | 0.015055 | 0.010488 | 0.061563 | 0.047349 | 0.008165 | 0.023381 | 0.016733 | 0.078909 | 0.064187 | 0.021602 | 0.017397 | 0.04761 | 0.035777 | 0.010328 | 0.021679 | 0.20506
| Rel.Std.Dev. |  | 0.3181% | 0.8551% | 0.9582% | 1.14% | 1.011% | 1.498% | 1.568% | 4.03% | 1.748% | 0.4943% | 0.6377% | 0.3917% | 0.2756% | 1.549% | 1.254% | 0.2101% | 0.6099% | 0.4313% | 2.31% | 1.609% | 0.5516% | 0.4546% | 1.373% | 0.9775% | 0.2617% | 0.5482% | 5.403%
| PDM3 |  | -2.683% | 2.558% | 2.174% | -0.3398% | 3.197% | -0.1267% | -1.064% | 0.8533% | 6.946% | 2.132% | 2.728% | -1.746% | -2.726% | 1.62% | -3.486% | -0.638% | -2.001% | -0.8084% | -12.65% | 2.004% | 0.1289% | -2.172% | -11.38% | -6.433% | 0.8959% | 1.109% | -2.981%
| Z-Score (Absolute) |  | 0.67 | 0.83 | 0.72 | 0 | 1.02 | 0.06 | 0.21 | 0.34 | 2.1 | 0.71 | 0.88 | 0.4 | 0.69 | 0.56 | 0.91 | 0.09 | 0.48 | 0.13 | 3.54 | 0.67 | 0.13 | 0.53 | 3.18 | 1.75 | 0.36 | 0.42 | 0.76
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 625 (Indicative Value 13.16 wt.%) |  |  |  | Indicative
Al2O3 | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 13.2462 | 13.12 | 1
| 1 | 2 | 13.063 | 13.2 | 20
| 1 | 3 | 13.2818 |  | 16
| 1 | 4 | 13.1677 |  | 13.1626
| 1 | 5 | 13.1514 |  | 26
| 1 | 6 | 13.081
| Mean |  | 13.1652 | 13.16
| Median |  | 13.1595 | 13.16
| Std Dev. |  | 0.087049 | 0.056569
| Rel.Std.Dev. |  | 0.6612% | 0.4299%
| PDM3 |  | 0.01969% | -0.01969%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for As in OREAS 625 (Indicative Value 3156 ppm) |  |  |  | Indicative
As | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 3110.094 |  | 1
| 1 | 2 | 3101.022 |  | 21
| 1 | 3 | 3179.729 |  | 16
| 1 | 4 | 3287.019 |  | 3156.1728
| 1 | 5 | 3126.346 |  | 27
| 1 | 6 | 3132.827
| Mean |  | 3156.1728
| Median |  | 3129.5865
| Std Dev. |  | 69.6781
| Rel.Std.Dev. |  | 2.208%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 625 (Indicative Value 3588 ppm) |  |  |  | Indicative
Ba | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 3538.899 |  | 1
| 1 | 2 | 3576.889 |  | 4
| 1 | 3 | 3577.22 |  | 16
| 1 | 4 | 3619.06 |  | 3587.7583
| 1 | 5 | 3609.535 |  | 28
| 1 | 6 | 3604.947
| Mean |  | 3587.7583
| Median |  | 3591.0835
| Std Dev. |  | 29.5496
| Rel.Std.Dev. |  | 0.8236%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 625 (Indicative Value 1.61 wt.%) |  |  |  | Indicative
CaO | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 1.6261 | 1.6 | 1
| 1 | 2 | 1.623 | 1.6 | 23
| 1 | 3 | 1.6362 |  | 16
| 1 | 4 | 1.6021 |  | 1.6087
| 1 | 5 | 1.5924 |  | 29
| 1 | 6 | 1.6246
| Mean |  | 1.6174 | 1.6
| Median |  | 1.6238 | 1.6
| Std Dev. |  | 0.016562 | 0
| Rel.Std.Dev. |  | 1.024% | 0%
| PDM3 |  | 0.5408% | -0.5408%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cr in OREAS 625 (Indicative Value 27.4 ppm) |  |  |  | Indicative
Cr | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 1
1
| 1 | 1 | 18.9484 |  | 1
| 1 | 2 | 54.4084 |  | 15
| 1 | 3 | 17.5561 |  | 16
| 1 | 4 | 40.7594 |  | 27.3813
| 1 | 5 | 5.2342 |  | 30
| Mean |  | 27.3813
| Median |  | 18.9484
| Std Dev. |  | 19.7963
| Rel.Std.Dev. |  | 72.3%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 625 (Indicative Value 0.169 wt.%) |  |  |  | Indicative
Cu | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 3
3
| 1 | 1 | 0.1717 |  | 1
| 1 | 2 | 0.182 |  | 25
| 1 | 3 | 0.1729 |  | 16
| 1 | 4 | 0.1638 |  | 0.16943
| 1 | 5 | 0.1656 |  | 31
| 1 | 6 | 0.1606
| Mean |  | 0.16943
| Median |  | 0.16865
| Std Dev. |  | 0.0077399
| Rel.Std.Dev. |  | 4.568%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 625 (Indicative Value 4.05 wt.%) |  |  |  | Indicative
Fe | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 4.0079 | 4.029 | 1
| 1 | 2 | 4.0544 | 4.043 | 26
| 1 | 3 | 4.0856 |  | 16
| 1 | 4 | 4.0276 |  | 4.0512
| 1 | 5 | 4.1702 |  | 32
| 1 | 6 | 4.0533
| Mean |  | 4.0665 | 4.036
| Median |  | 4.0538 | 4.036
| Std Dev. |  | 0.057247 | 0.0098995
| Rel.Std.Dev. |  | 1.408% | 0.2453%
| PDM3 |  | 0.3788% | -0.3741%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for K2O in OREAS 625 (Indicative Value 3.52 wt.%) |  |  |  | Indicative
K2O | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 3.5562 | 3.47 | 1
| 1 | 2 | 3.5007 | 3.5 | 27
| 1 | 3 | 3.579 |  | 16
| 1 | 4 | 3.5666 |  | 3.5187
| 1 | 5 | 3.5562 |  | 33
| 1 | 6 | 3.5556
| Mean |  | 3.5524 | 3.485
| Median |  | 3.5562 | 3.485
| Std Dev. |  | 0.026912 | 0.021213
| Rel.Std.Dev. |  | 0.7576% | 0.6087%
| PDM3 |  | 0.9575% | -0.9575%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 625 (Indicative Value 0.467 wt.%) |  |  |  | Indicative
MgO | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.4553 | 0.45 | 1
| 1 | 2 | 0.4551 | 0.49 | 20
| 1 | 3 | 0.4716 |  | 16
| 1 | 4 | 0.4665 |  | 0.46655
| 1 | 5 | 0.4651 |  | 26
| 1 | 6 | 0.465
| Mean |  | 0.4631 | 0.47
| Median |  | 0.46505 | 0.47
| Std Dev. |  | 0.0065766 | 0.028284
| Rel.Std.Dev. |  | 1.42% | 6.018%
| PDM3 |  | -0.7395% | 0.7395%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mn in OREAS 625 (Indicative Value 0.454 wt.%) |  |  |  | Indicative
Mn | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.4673 | 0.4469 | 1
| 1 | 2 | 0.4509 | 0.4507 | 21
| 1 | 3 | 0.4605 |  | 16
| 1 | 4 | 0.452 |  | 0.45362
| 1 | 5 | 0.4633 |  | 27
| 1 | 6 | 0.4566
| Mean |  | 0.45843 | 0.4488
| Median |  | 0.45855 | 0.4488
| Std Dev. |  | 0.0064509 | 0.002687
| Rel.Std.Dev. |  | 1.407% | 0.5987%
| PDM3 |  | 1.062% | -1.062%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Na2O in OREAS 625 (Indicative Value 2.69 wt.%) |  |  |  | Indicative
Na2O | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 2.713 |  | 1
| 1 | 2 | 2.6745 |  | 22
| 1 | 3 | 2.7234 |  | 16
| 1 | 4 | 2.7072 |  | 2.6942
| 1 | 5 | 2.6535 |  | 28
| 1 | 6 | 2.6933
| Mean |  | 2.6942
| Median |  | 2.7003
| Std Dev. |  | 0.026155
| Rel.Std.Dev. |  | 0.9708%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 625 (Indicative Value 10.6 ppm) |  |  |  | Indicative
Ni | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 1
1
| 1 | 1 | 6.9945 |  | 1
| 1 | 2 | 14.5163 |  | 14
| 1 | 3 | 10.2881 |  | 16
| 1 | 4 | 5.9685 |  | 10.5991
| 1 | 5 | 15.2279 |  | 29
| Mean |  | 10.5991
| Median |  | 10.2881
| Std Dev. |  | 4.222
| Rel.Std.Dev. |  | 39.83%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 625 (Indicative Value 0.029 wt.%) |  |  |  | Indicative
P | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.024 | 0.0314 | 1
| 1 | 2 | 0.0276 | 0.0319 | 24
| 1 | 3 | 0.0284 |  | 16
| 1 | 4 | 0.0251 |  | 0.02906
| 1 | 5 | 0.0265 |  | 30
| 1 | 6 | 0.0273
| Mean |  | 0.026483 | 0.03165
| Median |  | 0.0269 | 0.03165
| Std Dev. |  | 0.0016534 | 0.00035355
| Rel.Std.Dev. |  | 6.243% | 1.117%
| PDM3 |  | -8.868% | 8.911%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pb in OREAS 625 (Indicative Value 0.871 wt.%) |  |  |  | Indicative
Pb | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 3
3
| 1 | 1 | 0.8671 |  | 1
| 1 | 2 | 0.861 |  | 25
| 1 | 3 | 0.878 |  | 16
| 1 | 4 | 0.8733 |  | 0.8709
| 1 | 5 | 0.8719 |  | 31
| 1 | 6 | 0.8741
| Mean |  | 0.8709
| Median |  | 0.8726
| Std Dev. |  | 0.0059977
| Rel.Std.Dev. |  | 0.6887%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 625 (Indicative Value 3.33 wt.%) |  |  |  | Indicative
S | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 2.6283 | 3.9124 | 1
| 1 | 2 | 2.8424 | 3.9244 | 26
| 1 | 3 | 2.7173 |  | 16
| 1 | 4 | 2.7624 |  | 3.3264
| 1 | 5 | 2.7917 |  | 32
| 1 | 6 | 2.6645
| Mean |  | 2.7344 | 3.9184
| Median |  | 2.7399 | 3.9184
| Std Dev. |  | 0.080208 | 0.0084853
| Rel.Std.Dev. |  | 2.933% | 0.2165%
| PDM3 |  | -17.8% | 17.8%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for SiO2 in OREAS 625 (Indicative Value 62.85 wt.%) |  |  |  | Indicative
SiO2 | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 64.2169 | 61.7 | 1
| 1 | 2 | 63.4886 | 61.8 | 27
| 1 | 3 | 64.4199 |  | 16
| 1 | 4 | 63.8746 |  | 62.8471
| 1 | 5 | 63.5688 |  | 33
| 1 | 6 | 64.0959
| Mean |  | 63.9441 | 61.75
| Median |  | 63.9853 | 61.75
| Std Dev. |  | 0.36797 | 0.070711
| Rel.Std.Dev. |  | 0.5755% | 0.1145%
| PDM3 |  | 1.746% | -1.746%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 625 (Indicative Value 27.5 ppm) |  |  |  | Indicative
Sn | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 1
1
| 1 | 1 | 7.605 |  | 1
| 1 | 2 | 47.4069 |  | 13
| Mean |  | 27.5059 |  | 16
| Median |  | 27.5059 |  | 27.5059
| Std Dev. |  | 28.1442 |  | 26
| Rel.Std.Dev. |  | 102.3%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 625 (Indicative Value 99 ppm) |  |  |  | Indicative
Sr | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 129.9639 |  | 1
| 1 | 2 | 65.807 |  | 21
| 1 | 3 | 124.9393 |  | 16
| 1 | 4 | 75.792 |  | 99.1256
| Mean |  | 99.1256 |  | 27
| Median |  | 100.3657
| Std Dev. |  | 33.0249
| Rel.Std.Dev. |  | 33.32%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 625 (Indicative Value 0.211 wt.%) |  |  |  | Indicative
TiO2 | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.2084 | 0.21 | 1
| 1 | 2 | 0.1926 | 0.212 | 22
| 1 | 3 | 0.2129 |  | 16
| 1 | 4 | 0.2236 |  | 0.21066
| 1 | 5 | 0.2081 |  | 28
| 1 | 6 | 0.2163
| Mean |  | 0.21032 | 0.211
| Median |  | 0.21065 | 0.211
| Std Dev. |  | 0.010406 | 0.0014142
| Rel.Std.Dev. |  | 4.948% | 0.6702%
| PDM3 |  | -0.1622% | 0.1622%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V in OREAS 625 (Indicative Value 18 ppm) |  |  |  | Indicative
V | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | ppm
|  |  | BF*XRF | BF*XRF | 1
1
| 1 | 1 | 11.3923 | 22 | 1
| 1 | 2 | 5.4432 | 28 | 23
| 1 | 3 | 14.1986 |  | 16
| 1 | 4 | 8.2328 |  | 17.9829
| 1 | 5 | 14.5256 |  | 29
| Mean |  | 10.7585 | 25
| Median |  | 11.3923 | 25
| Std Dev. |  | 3.907 | 4.2426
| Rel.Std.Dev. |  | 36.32% | 16.97%
| PDM3 |  | -40.17% | 39.02%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 625 (Indicative Value 3.32 wt.%) |  |  |  | Indicative
Zn | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 3.3215 |  | 1
| 1 | 2 | 3.2151 |  | 24
| 1 | 3 | 3.3617 |  | 16
| 1 | 4 | 3.3475 |  | 3.3177
| 1 | 5 | 3.3222 |  | 30
| 1 | 6 | 3.338
| Mean |  | 3.3177
| Median |  | 3.3301
| Std Dev. |  | 0.052531
| Rel.Std.Dev. |  | 1.583%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 625 (Indicative Value 296 ppm) |  |  |  | Indicative
Zr | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 301.195 |  | 1
| 1 | 2 | 320.4198 |  | 16
| 1 | 3 | 345.7319 |  | 16
| 1 | 4 | 277.916 |  | 295.8355
| 1 | 5 | 256.2719 |  | 31
| 1 | 6 | 273.4783
| Mean |  | 295.8355
| Median |  | 289.5555
| Std Dev. |  | 33.1934
| Rel.Std.Dev. |  | 11.22%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C88: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C124: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C142: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C160: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C178: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C196: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C213: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C231: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C249: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C267: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C285: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C299: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C315: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C333: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C350: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C368: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C386: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 625 (Indicative Value 4.37 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV | 1
| No | No | Ankara | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 4.37 | 1
| 1 | 2 | 4.36 | 29
| 1 | 3 | 4.35 | 16
| 1 | 4 | 4.39 | 4.3667
| 1 | 5 | 4.37 | 35
| 1 | 6 | 4.36
| Mean |  | 4.3667
| Median |  | 4.365
| Std Dev. |  | 0.013663
| Rel.Std.Dev. |  | 0.3129%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 625 (Indicative Value 10.9 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.6 | 1
| 1 | 2 | 11.1 | 1
| Mean |  | 10.85 | 16
| Median |  | 10.85 | 10.85
| Std Dev. |  | 0.35355 | 37
| Rel.Std.Dev. |  | 3.259%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 625 (Indicative Value 142 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 143 | 1
| 1 | 2 | 140 | 21
| Mean |  | 141.5 | 16
| Median |  | 141.5 | 141.5
| Std Dev. |  | 2.1213 | 38
| Rel.Std.Dev. |  | 1.499%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 625 (Indicative Value 3450 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 3440 | 1
| 1 | 2 | 3460 | 4
| Mean |  | 3450 | 16
| Median |  | 3450 | 3450
| Std Dev. |  | 14.1421 | 39
| Rel.Std.Dev. |  | 0.4099%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 625 (Indicative Value 2.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 3 | 14
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0.56569 | 40
| Rel.Std.Dev. |  | 21.76%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 625 (Indicative Value 8.9 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9 | 1
| 1 | 2 | 8.8 | 35
| Mean |  | 8.9 | 16
| Median |  | 8.9 | 8.9
| Std Dev. |  | 0.14142 | 41
| Rel.Std.Dev. |  | 1.589%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 625 (Indicative Value 82 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 80.6 | 1
| 1 | 2 | 82.6 | 36
| Mean |  | 81.6 | 16
| Median |  | 81.6 | 81.6
| Std Dev. |  | 1.4142 | 42
| Rel.Std.Dev. |  | 1.733%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 625 (Indicative Value 5.05 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.2 | 1
| 1 | 2 | 4.9 | 37
| Mean |  | 5.05 | 16
| Median |  | 5.05 | 5.05
| Std Dev. |  | 0.21213 | 43
| Rel.Std.Dev. |  | 4.201%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 625 (Indicative Value 16 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14 | 1
| 1 | 2 | 18 | 15
| Mean |  | 16 | 16
| Median |  | 16 | 16
| Std Dev. |  | 2.8284 | 44
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 625 (Indicative Value 6.07 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.09 | 1
| 1 | 2 | 6.05 | 39
| Mean |  | 6.07 | 16
| Median |  | 6.07 | 6.07
| Std Dev. |  | 0.028284 | 45
| Rel.Std.Dev. |  | 0.466%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 625 (Indicative Value 0.207 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.185 | 1
| 1 | 2 | 0.228 | 25
| Mean |  | 0.2065 | 16
| Median |  | 0.2065 | 0.2065
| Std Dev. |  | 0.030406 | 46
| Rel.Std.Dev. |  | 14.72%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 625 (Indicative Value 3.83 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.66 | 1
| 1 | 2 | 3.99 | 7
| Mean |  | 3.825 | 16
| Median |  | 3.825 | 3.825
| Std Dev. |  | 0.23335 | 47
| Rel.Std.Dev. |  | 6.101%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 625 (Indicative Value 1.37 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.29 | 1
| 1 | 2 | 1.44 | 8
| Mean |  | 1.365 | 16
| Median |  | 1.365 | 1.365
| Std Dev. |  | 0.10607 | 48
| Rel.Std.Dev. |  | 7.77%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 625 (Indicative Value 1.43 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.34 | 1
| 1 | 2 | 1.51 | 9
| Mean |  | 1.425 | 16
| Median |  | 1.425 | 1.425
| Std Dev. |  | 0.12021 | 49
| Rel.Std.Dev. |  | 8.436%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 625 (Indicative Value 19.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.9 | 1
| 1 | 2 | 19.6 | 44
| Mean |  | 19.75 | 16
| Median |  | 19.75 | 19.75
| Std Dev. |  | 0.21213 | 50
| Rel.Std.Dev. |  | 1.074%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 625 (Indicative Value 5.36 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.37 | 1
| 1 | 2 | 5.35 | 10
| Mean |  | 5.36 | 16
| Median |  | 5.36 | 5.36
| Std Dev. |  | 0.014142 | 51
| Rel.Std.Dev. |  | 0.2638%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 625 (Indicative Value 1.08 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.15 | 1
| 1 | 2 | 1 | 46
| Mean |  | 1.075 | 16
| Median |  | 1.075 | 1.075
| Std Dev. |  | 0.10607 | 52
| Rel.Std.Dev. |  | 9.867%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 625 (Indicative Value 6.8 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.44 | 1
| 1 | 2 | 7.15 | 16
| Mean |  | 6.795 | 16
| Median |  | 6.795 | 6.795
| Std Dev. |  | 0.50205 | 37
| Rel.Std.Dev. |  | 7.388%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 625 (Indicative Value 0.6 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.59 | 7
| Mean |  | 0.595 | 16
| Median |  | 0.595 | 0.595
| Std Dev. |  | 0.0070711 | 38
| Rel.Std.Dev. |  | 1.188%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 625 (Indicative Value 1.95 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.9 | 1
| 1 | 2 | 2 | 33
| Mean |  | 1.95 | 16
| Median |  | 1.95 | 1.95
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 625 (Indicative Value 42.7 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 42.4 | 1
| 1 | 2 | 43 | 34
| Mean |  | 42.7 | 16
| Median |  | 42.7 | 42.7
| Std Dev. |  | 0.42426 | 40
| Rel.Std.Dev. |  | 0.9936%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 625 (Indicative Value 0.13 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.14 | 1
| 1 | 2 | 0.12 | 35
| Mean |  | 0.13 | 16
| Median |  | 0.13 | 0.13
| Std Dev. |  | 0.014142 | 41
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 625 (Indicative Value 0.454 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.437 | 1
| 1 | 2 | 0.471 | 21
| Mean |  | 0.454 | 16
| Median |  | 0.454 | 0.454
| Std Dev. |  | 0.024042 | 42
| Rel.Std.Dev. |  | 5.296%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 625 (Indicative Value 11.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.2 | 1
| 1 | 2 | 12 | 37
| Mean |  | 11.6 | 16
| Median |  | 11.6 | 11.6
| Std Dev. |  | 0.56569 | 43
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 625 (Indicative Value 15.6 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.3 | 1
| 1 | 2 | 15.9 | 38
| Mean |  | 15.6 | 16
| Median |  | 15.6 | 15.6
| Std Dev. |  | 0.42426 | 44
| Rel.Std.Dev. |  | 2.72%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 625 (Indicative Value 36.2 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.5 | 1
| 1 | 2 | 36.9 | 8
| Mean |  | 36.2 | 16
| Median |  | 36.2 | 36.2
| Std Dev. |  | 0.98995 | 45
| Rel.Std.Dev. |  | 2.735%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 625 (Indicative Value 8 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 12 | 14
| Mean |  | 8 | 16
| Median |  | 8 | 8
| Std Dev. |  | 5.6569 | 46
| Rel.Std.Dev. |  | 70.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 625 (Indicative Value 0.737 wt.%) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.717 | 1
| 1 | 2 | 0.756 | 25
| Mean |  | 0.7365 | 16
| Median |  | 0.7365 | 0.7365
| Std Dev. |  | 0.027577 | 47
| Rel.Std.Dev. |  | 3.744%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 625 (Indicative Value 10.2 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.1 | 1
| 1 | 2 | 10.3 | 10
| Mean |  | 10.2 | 16
| Median |  | 10.2 | 10.2
| Std Dev. |  | 0.14142 | 48
| Rel.Std.Dev. |  | 1.386%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 625 (Indicative Value 136 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 129 | 1
| 1 | 2 | 143 | 43
| Mean |  | 136 | 16
| Median |  | 136 | 136
| Std Dev. |  | 9.8995 | 49
| Rel.Std.Dev. |  | 7.279%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 625 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 15
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 50
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 625 (Indicative Value 36.4 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.2 | 1
| 1 | 2 | 37.5 | 45
| Mean |  | 36.35 | 16
| Median |  | 36.35 | 36.35
| Std Dev. |  | 1.6263 | 51
| Rel.Std.Dev. |  | 4.474%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 625 (Indicative Value 5.1 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.1 | 1
| 1 | 2 | 5.1 | 16
| Mean |  | 5.1 | 16
| Median |  | 5.1 | 5.1
| Std Dev. |  | 0 | 52
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 625 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 3
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 37
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 625 (Indicative Value 6.89 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.55 | 1
| 1 | 2 | 7.23 | 8
| Mean |  | 6.89 | 16
| Median |  | 6.89 | 6.89
| Std Dev. |  | 0.48083 | 38
| Rel.Std.Dev. |  | 6.979%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 625 (Indicative Value 3.4 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.2 | 1
| 1 | 2 | 3.6 | 13
| Mean |  | 3.4 | 16
| Median |  | 3.4 | 3.4
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 625 (Indicative Value 134 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 134 | 1
| 1 | 2 | 134 | 21
| Mean |  | 134 | 16
| Median |  | 134 | 134
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 625 (Indicative Value 1.18 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.19 | 1
| 1 | 2 | 1.16 | 14
| Mean |  | 1.175 | 16
| Median |  | 1.175 | 1.175
| Std Dev. |  | 0.021213 | 41
| Rel.Std.Dev. |  | 1.805%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 625 (Indicative Value 0.74 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.73 | 1
| 1 | 2 | 0.75 | 36
| Mean |  | 0.74 | 16
| Median |  | 0.74 | 0.74
| Std Dev. |  | 0.014142 | 42
| Rel.Std.Dev. |  | 1.911%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 625 (Indicative Value 0.15 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | 0.2 | 15
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 43
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 625 (Indicative Value 14.3 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.3 | 1
| 1 | 2 | 14.2 | 38
| Mean |  | 14.25 | 16
| Median |  | 14.25 | 14.25
| Std Dev. |  | 0.070711 | 44
| Rel.Std.Dev. |  | 0.4962%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 625 (Indicative Value 0.134 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.133 | 1
| 1 | 2 | 0.134 | 39
| Mean |  | 0.1335 | 16
| Median |  | 0.1335 | 0.1335
| Std Dev. |  | 0.00070711 | 45
| Rel.Std.Dev. |  | 0.5297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 625 (Indicative Value 0.16 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.16 | 1
| 1 | 2 | 0.16 | 9
| Mean |  | 0.16 | 16
| Median |  | 0.16 | 0.16
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 625 (Indicative Value 5.6 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.53 | 1
| 1 | 2 | 5.66 | 41
| Mean |  | 5.595 | 16
| Median |  | 5.595 | 5.595
| Std Dev. |  | 0.091924 | 47
| Rel.Std.Dev. |  | 1.643%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 625 (Indicative Value 11.1 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.4 | 1
| 1 | 2 | 10.8 | 23
| Mean |  | 11.1 | 16
| Median |  | 11.1 | 11.1
| Std Dev. |  | 0.42426 | 48
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 625 (Indicative Value 4.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.5 | 1
| 1 | 2 | 4.5 | 43
| Mean |  | 4.5 | 16
| Median |  | 4.5 | 4.5
| Std Dev. |  | 0 | 49
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 625 (Indicative Value 16.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16 | 1
| 1 | 2 | 17.1 | 44
| Mean |  | 16.55 | 16
| Median |  | 16.55 | 16.55
| Std Dev. |  | 0.77782 | 50
| Rel.Std.Dev. |  | 4.7%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 625 (Indicative Value 1.07 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.04 | 1
| 1 | 2 | 1.09 | 45
| Mean |  | 1.065 | 16
| Median |  | 1.065 | 1.065
| Std Dev. |  | 0.035355 | 51
| Rel.Std.Dev. |  | 3.32%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 625 (Indicative Value 241 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 239 | 1
| 1 | 2 | 242 | 16
| Mean |  | 240.5 | 16
| Median |  | 240.5 | 240.5
| Std Dev. |  | 2.1213 | 52
| Rel.Std.Dev. |  | 0.882%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.