## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 556b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 10.9015 | 10.7361 | 11.0668 | 10.786 | 11.0169
BaO, Barium oxide (ppm) | 242.1257 | 160.8065 | 323.4448 | 224.1963 | 260.055
CaO, Calcium oxide (wt.%) | 0.19527 | 0.18414 | 0.20639 | 0.18717 | 0.20336
Co, Cobalt (wt.%) | 1.9906 | 1.9399 | 2.0413 | 1.9663 | 2.0148
Cu, Copper (wt.%) | 8.5259 | 8.3357 | 8.7161 | 8.4264 | 8.6254
Fe2O3, Iron(III) oxide (wt.%) | 9.6806 | 9.5305 | 9.8308 | 9.5889 | 9.7724
K2O, Potassium oxide (wt.%) | 2.7415 | 2.7018 | 2.7813 | 2.7083 | 2.7748
MgO, Magnesium oxide (wt.%) | 4.1091 | 4.0541 | 4.1641 | 4.0635 | 4.1547
MnO, Manganese oxide (wt.%) | 0.15474 | 0.14788 | 0.16161 | 0.14953 | 0.15996
Ni, Nickel (ppm) | 343.9656 | 308.4062 | 379.5251 | 319.5767 | 368.3545
P2O5, Phosphorus(V) oxide (wt.%) | 0.079854 | 0.074347 | 0.085361 | 0.077078 | 0.082631
SiO2, Silicon dioxide (wt.%) | 42.7709 | 42.32 | 43.2218 | 42.3829 | 43.1589
SO3, Sulphur trioxide (wt.%) | 21.4269 | 20.0335 | 22.8203 | 21.1053 | 21.7486
Sr, Strontium (ppm) | 63.2432 | 26.9792 | 99.5072 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.70703 | 0.68662 | 0.72743 | 0.6936 | 0.72045
V2O5, Vanadium(V) oxide (ppm) | 640.9222 | 570.7999 | 711.0445 | 605.8375 | 676.0069
Zr, Zirconium (ppm) | 216.113 | 194.3087 | 237.9174 | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 14.409 | 14.2027 | 14.6154 | 14.3202 | 14.4978
Infrared Combustion
C, Carbon (wt.%) | 4.5805 | 4.483 | 4.678 | 4.5362 | 4.6248
S, Sulphur (wt.%) | 8.7844 | 8.6689 | 8.8998 | 8.7128 | 8.8559
Sulphuric Acid 5% Leach
Co, Cobalt (wt.%) | 0.68534 | 0.63853 | 0.73215 | 0.66427 | 0.7064
Cu, Copper (wt.%) | 1.318 | 1.2526 | 1.3834 | 1.2951 | 1.3409
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.7078 | 5.576 | 5.8397 | 5.5791 | 5.8366
As, Arsenic (ppm) | 101.3026 | 95.0989 | 107.5064 | 98.4413 | 104.1639
B, Boron (ppm) | 199.287 | 165.677 | 232.8971 | 185.8016 | 212.7725
Ba, Barium (ppm) | 195.6728 | 184.6941 | 206.6514 | 186.8425 | 204.503
Be, Beryllium (ppm) | 5.061 | 4.6211 | 5.5009 | IND | IND
Bi, Bismuth (ppm) | 3.0753 | 2.8675 | 3.2832 | 2.8122 | 3.3385
Ca, Calcium (wt.%) | 0.14866 | 0.12117 | 0.17614 | 0.13408 | 0.16323
Ce, Cerium (ppm) | 155.9579 | 143.5924 | 168.3234 | 142.7182 | 169.1976
Co, Cobalt (wt.%) | 1.9911 | 1.9125 | 2.0698 | 1.9431 | 2.0392
Cr, Chromium (ppm) | 112.6336 | 100.4478 | 124.8193 | 105.7775 | 119.4897
Cs, Caesium (ppm) | 1.8586 | 1.6601 | 2.057 | 1.6771 | 2.04
Cu, Copper (wt.%) | 8.6302 | 8.4133 | 8.8471 | 8.5011 | 8.7592
Dy, Dysprosium (ppm) | 6.278 | 5.7463 | 6.8097 | 5.9321 | 6.6238
Er, Erbium (ppm) | 3.4664 | 3.2186 | 3.7142 | 3.2998 | 3.6329
Eu, Europium (ppm) | 0.87762 | 0.70879 | 1.0464 | 0.82131 | 0.93394
Fe, Iron (wt.%) | 6.7301 | 6.5558 | 6.9044 | 6.6 | 6.8602
Ga, Gallium (ppm) | 28.6577 | 26.5258 | 30.7896 | 26.811 | 30.5045
Gd, Gadolinium (ppm) | 6.8993 | 6.475 | 7.3237 | 6.4855 | 7.3132
Ge, Germanium (ppm) | 2.8039 | 2.3861 | 3.2218 | 2.5094 | 3.0985
Ho, Holmium (ppm) | 1.2166 | 1.1082 | 1.3251 | 1.1499 | 1.2833
In, Indium (ppm) | 3.5174 | 3.2502 | 3.7846 | 3.3243 | 3.7106
K, Potassium (wt.%) | 2.2696 | 2.1989 | 2.3402 | 2.2019 | 2.3373
La, Lanthanum (ppm) | 75.2334 | 71.0157 | 79.451 | 72.4458 | 78.021
Li, Lithium (ppm) | 51.0031 | 45.437 | 56.5692 | 49.6914 | 52.3147
Lu, Lutetium (ppm) | 0.53412 | 0.41335 | 0.6549 | 0.49491 | 0.57333
Mg, Magnesium (wt.%) | 2.4886 | 2.4177 | 2.5596 | 2.4359 | 2.5414
Mn, Manganese (wt.%) | 0.12072 | 0.11359 | 0.12784 | 0.11528 | 0.12616
Mo, Molybdenum (ppm) | 18.7411 | 17.0281 | 20.454 | 17.3616 | 20.1206
Nb, Niobium (ppm) | 22.4636 | 19.7249 | 25.2023 | 20.7743 | 24.1529
Nd, Neodymium (ppm) | 61.8593 | 54.9995 | 68.7192 | 58.6863 | 65.0324
Ni, Nickel (ppm) | 342.6253 | 325.5059 | 359.7447 | 332.0093 | 353.2413
P, Phosphorus (wt.%) | 0.034683 | 0.024909 | 0.044458 | IND | IND
Pb, Lead (ppm) | < 100 | IND | IND | IND | IND
Pr, Praseodymium (ppm) | 18.2767 | 16.8423 | 19.7112 | 17.5227 | 19.0308
Rb, Rubidium (ppm) | 82.0004 | 77.5095 | 86.4913 | 79.1705 | 84.8303
S, Sulphur (wt.%) | 8.7525 | 8.4938 | 9.0113 | 8.5632 | 8.9419
Sb, Antimony (ppm) | 5.7737 | 5.0999 | 6.4474 | 5.0804 | 6.467
Sc, Scandium (ppm) | 9.5602 | 7.4608 | 11.6596 | IND | IND
Si, Silicon (wt.%) | 20.4762 | 19.9104 | 21.042 | 20.1352 | 20.8172
Sm, Samarium (ppm) | 8.8758 | 7.7645 | 9.9871 | 8.2936 | 9.4581
Sr, Strontium (ppm) | 73.3151 | 69.4895 | 77.1407 | 71.3163 | 75.3138
Ta, Tantalum (ppm) | 1.8409 | 1.5254 | 2.1564 | IND | IND
Tb, Terbium (ppm) | 1.0544 | 0.96266 | 1.1461 | 0.97875 | 1.13
Th, Thorium (ppm) | 15.5063 | 14.4816 | 16.531 | 14.9091 | 16.1036
Ti, Titanium (wt.%) | 0.41403 | 0.40265 | 0.4254 | 0.40173 | 0.42632
Tl, Thallium (ppm) | 1.06 | 0.941 | 1.1789 | IND | IND
Tm, Thulium (ppm) | 0.52179 | 0.47526 | 0.56832 | 0.48303 | 0.56055
U, Uranium (ppm) | 8.8503 | 8.0461 | 9.6544 | 8.4647 | 9.2358
V, Vanadium (ppm) | 352.7061 | 337.5868 | 367.8254 | 342.6013 | 362.8109
W, Tungsten (ppm) | 2.9639 | 2.1266 | 3.8011 | IND | IND
Y, Yttrium (ppm) | 32.611 | 30.7961 | 34.426 | 31.2048 | 34.0173
Yb, Ytterbium (ppm) | 3.1916 | 2.8627 | 3.5205 | 2.9492 | 3.4339
Zn, Zinc (ppm) | 41.353 | 35.8376 | 46.8683 | 38.4471 | 44.2588
4-Acid Digestion
Ag, Silver (ppm) | 0.14397 | 0.12893 | 0.15901 | 0.1365 | 0.15145
Al, Aluminium (wt.%) | 5.6681 | 5.5365 | 5.7997 | 5.5639 | 5.7724
As, Arsenic (ppm) | 98.2191 | 94.3231 | 102.1151 | 94.8578 | 101.5804
Ba, Barium (ppm) | 189.7248 | 182.6986 | 196.7509 | 183.1029 | 196.3466
Be, Beryllium (ppm) | 4.5183 | 4.3663 | 4.6703 | 4.3959 | 4.6407
Bi, Bismuth (ppm) | 2.894 | 2.7533 | 3.0346 | 2.8098 | 2.9781
Ca, Calcium (wt.%) | 0.14107 | 0.13168 | 0.15046 | 0.13839 | 0.14374
Cd, Cadmium (ppm) | 0.51006 | 0.46099 | 0.55913 | 0.48459 | 0.53554
Ce, Cerium (ppm) | 143.3697 | 137.1977 | 149.5417 | 138.9537 | 147.7857
Co, Cobalt (wt.%) | 1.9932 | 1.9478 | 2.0387 | 1.9584 | 2.0281
Cr, Chromium (ppm) | 88.9552 | 84.1467 | 93.7637 | 87.0221 | 90.8883
Cs, Caesium (ppm) | 1.7544 | 1.6632 | 1.8457 | 1.6849 | 1.8239
Cu, Copper (wt.%) | 8.6408 | 8.4778 | 8.8039 | 8.5277 | 8.754
Dy, Dysprosium (ppm) | 4.5357 | 4.1022 | 4.9692 | 4.3571 | 4.7143
Er, Erbium (ppm) | 2.6977 | 2.3156 | 3.0798 | 2.5879 | 2.8075
Eu, Europium (ppm) | 0.88976 | 0.81263 | 0.96688 | 0.8424 | 0.93711
Fe, Iron (wt.%) | 6.7076 | 6.5624 | 6.8529 | 6.603 | 6.8123
Ga, Gallium (ppm) | 28.1302 | 26.878 | 29.3825 | 27.3653 | 28.8952
Gd, Gadolinium (ppm) | 4.9221 | 3.9971 | 5.8471 | 4.7252 | 5.1191
Ge, Germanium (ppm) | 0.22821 | 0.18068 | 0.27574 | 0.21109 | 0.24533
Hf, Hafnium (ppm) | 4.9448 | 4.7235 | 5.1661 | 4.7807 | 5.1089
Ho, Holmium (ppm) | 0.88682 | 0.74397 | 1.0297 | 0.8298 | 0.94384
In, Indium (ppm) | 3.564 | 3.3814 | 3.7466 | 3.4583 | 3.6696
K, Potassium (wt.%) | 2.2457 | 2.2012 | 2.2902 | 2.1994 | 2.292
La, Lanthanum (ppm) | 69.7876 | 66.5504 | 73.0247 | 67.973 | 71.6022
Li, Lithium (ppm) | 51.6024 | 49.8747 | 53.3302 | 50.4364 | 52.7685
Lu, Lutetium (ppm) | 0.40632 | 0.35125 | 0.46139 | 0.38101 | 0.43164
Mg, Magnesium (wt.%) | 2.4254 | 2.353 | 2.4978 | 2.3789 | 2.4718
Mn, Manganese (wt.%) | 0.11719 | 0.11221 | 0.12218 | 0.11434 | 0.12005
Mo, Molybdenum (ppm) | 18.353 | 17.5662 | 19.1398 | 17.8306 | 18.8754
Na, Sodium (wt.%) | 0.049017 | 0.04361 | 0.054425 | 0.045688 | 0.052346
Nb, Niobium (ppm) | 14.0083 | 12.5077 | 15.5088 | 13.3997 | 14.6168
Nd, Neodymium (ppm) | 62.9858 | 59.275 | 66.6966 | 60.7913 | 65.1803
Ni, Nickel (ppm) | 336.7816 | 324.0201 | 349.5431 | 329.2839 | 344.2794
P, Phosphorus (wt.%) | 0.032675 | 0.031027 | 0.034322 | 0.03194 | 0.033409
Pb, Lead (ppm) | 17.2652 | 16.4818 | 18.0487 | 16.7997 | 17.7307
Pr, Praseodymium (ppm) | 16.5315 | 15.2363 | 17.8266 | 16.0057 | 17.0572
Rb, Rubidium (ppm) | 82.3175 | 78.711 | 85.9239 | 80.1734 | 84.4615
Re, Rhenium (ppm) | 0.032724 | 0.02876 | 0.036688 | 0.029785 | 0.035662
S, Sulphur (wt.%) | 8.5467 | 8.3585 | 8.7348 | 8.3778 | 8.7155
Sb, Antimony (ppm) | 5.2929 | 5.0059 | 5.5799 | 5.0866 | 5.4993
Sc, Scandium (ppm) | 11.03 | 10.4167 | 11.6433 | 10.632 | 11.4279
Se, Selenium (ppm) | 11.1614 | 10.1245 | 12.1984 | 10.5029 | 11.82
Sm, Samarium (ppm) | 8.5793 | 8.0085 | 9.1502 | 8.2925 | 8.8661
Sn, Tin (ppm) | 3.7046 | 3.4758 | 3.9335 | 3.5712 | 3.838
Sr, Strontium (ppm) | 67.5203 | 65.2566 | 69.784 | 65.7507 | 69.2899
Ta, Tantalum (ppm) | 1.0427 | 0.9337 | 1.1517 | 0.99112 | 1.0943
Tb, Terbium (ppm) | 0.77875 | 0.63997 | 0.91754 | 0.73178 | 0.82572
Te, Tellurium (ppm) | 0.089187 | 0.062799 | 0.11557 | IND | IND
Th, Thorium (ppm) | 15.4439 | 14.7908 | 16.097 | 14.9599 | 15.9278
Ti, Titanium (wt.%) | 0.23512 | 0.20847 | 0.26178 | 0.22785 | 0.2424
Tl, Thallium (ppm) | 1.0561 | 0.99567 | 1.1165 | 1.0228 | 1.0894
Tm, Thulium (ppm) | 0.39318 | 0.33837 | 0.44799 | 0.38103 | 0.40533
U, Uranium (ppm) | 8.7957 | 8.4148 | 9.1766 | 8.4654 | 9.126
V, Vanadium (ppm) | 330.3722 | 319.8529 | 340.8916 | 324.4092 | 336.3353
W, Tungsten (ppm) | 2.2981 | 2.1468 | 2.4494 | 2.16 | 2.4363
Y, Yttrium (ppm) | 23.4511 | 22.551 | 24.3512 | 22.7563 | 24.1458
Yb, Ytterbium (ppm) | 2.7218 | 2.4707 | 2.9729 | 2.6199 | 2.8237
Zn, Zinc (ppm) | 38.8027 | 37.1014 | 40.504 | 37.6809 | 39.9245
Zr, Zirconium (ppm) | 171.093 | 165.2496 | 176.9363 | 168.0528 | 174.1332
Aqua Regia Digestion
Ag, Silver (ppm) | 0.11978 | 0.10463 | 0.13493 | 0.1046 | 0.13496
Al, Aluminium (wt.%) | 0.86918 | 0.79148 | 0.94687 | 0.82678 | 0.91158
As, Arsenic (ppm) | 94.0686 | 90.3294 | 97.8079 | 92.1208 | 96.0165
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 52.5177 | 44.5221 | 60.5134 | 50.9243 | 54.1112
Be, Beryllium (ppm) | 2.2028 | 2.0497 | 2.356 | 2.1262 | 2.2794
Bi, Bismuth (ppm) | 2.8176 | 2.6681 | 2.967 | 2.7476 | 2.8876
Ca, Calcium (wt.%) | 0.13063 | 0.1249 | 0.13636 | 0.12639 | 0.13488
Cd, Cadmium (ppm) | 0.51043 | 0.46491 | 0.55594 | 0.48296 | 0.53789
Ce, Cerium (ppm) | 43.5722 | 40.2795 | 46.8648 | 41.8394 | 45.3049
Co, Cobalt (wt.%) | 1.9765 | 1.9254 | 2.0275 | 1.9444 | 2.0085
Cr, Chromium (ppm) | 35.0105 | 33.4785 | 36.5425 | 33.904 | 36.117
Cs, Caesium (ppm) | 0.2938 | 0.23911 | 0.34849 | 0.272 | 0.3156
Cu, Copper (wt.%) | 8.5536 | 8.3927 | 8.7144 | 8.4387 | 8.6684
Dy, Dysprosium (ppm) | 1.0157 | 0.84323 | 1.1882 | 0.95188 | 1.0795
Er, Erbium (ppm) | 0.41235 | 0.28755 | 0.53716 | 0.37706 | 0.44765
Eu, Europium (ppm) | 0.38311 | 0.34201 | 0.4242 | 0.36208 | 0.40413
Fe, Iron (wt.%) | 6.6451 | 6.4934 | 6.7969 | 6.5392 | 6.7511
Ga, Gallium (ppm) | 9.6132 | 8.7682 | 10.4582 | 9.2636 | 9.9628
Gd, Gadolinium (ppm) | 1.915 | 1.6397 | 2.1902 | 1.8234 | 2.0065
Ge, Germanium (ppm) | 0.12571 | 0.091861 | 0.15957 | IND | IND
Hf, Hafnium (ppm) | 0.36687 | 0.33921 | 0.39453 | 0.34736 | 0.38638
Hg, Mercury (ppm) | 0.10605 | 0.090425 | 0.12167 | 0.088825 | 0.12327
Ho, Holmium (ppm) | 0.16714 | 0.13151 | 0.20276 | IND | IND
In, Indium (ppm) | 2.9356 | 2.8141 | 3.0571 | 2.8659 | 3.0054
K, Potassium (wt.%) | 0.18728 | 0.1726 | 0.20196 | 0.17593 | 0.19863
La, Lanthanum (ppm) | 19.4745 | 18.0538 | 20.8953 | 18.8782 | 20.0709
Li, Lithium (ppm) | 31.2152 | 28.0637 | 34.3667 | 30.3049 | 32.1255
Mg, Magnesium (wt.%) | 1.8474 | 1.7417 | 1.9531 | 1.7983 | 1.8965
Mn, Manganese (wt.%) | 0.11668 | 0.11273 | 0.12063 | 0.1143 | 0.11906
Mo, Molybdenum (ppm) | 17.3801 | 16.5827 | 18.1775 | 16.9551 | 17.8051
Na, Sodium (wt.%) | 0.016202 | 0.013257 | 0.019148 | 0.014437 | 0.017968
Nb, Niobium (ppm) | 0.12569 | 0.099656 | 0.15173 | IND | IND
Nd, Neodymium (ppm) | 22.9668 | 18.0989 | 27.8347 | 21.7375 | 24.1961
Ni, Nickel (ppm) | 323.2012 | 309.5011 | 336.9012 | 316.347 | 330.0553
P, Phosphorus (wt.%) | 0.018787 | 0.017621 | 0.019953 | 0.01802 | 0.019554
Pb, Lead (ppm) | 14.8661 | 14.272 | 15.4602 | 14.4729 | 15.2594
Pr, Praseodymium (ppm) | 5.7531 | 4.5602 | 6.946 | 5.4631 | 6.0431
Rb, Rubidium (ppm) | 6.096 | 5.4309 | 6.7611 | 5.7742 | 6.4178
Re, Rhenium (ppm) | 0.029321 | 0.026951 | 0.031691 | 0.027088 | 0.031555
S, Sulphur (wt.%) | 8.3791 | 8.0829 | 8.6753 | 8.2418 | 8.5164
Sb, Antimony (ppm) | 3.3222 | 3.0468 | 3.5976 | 3.1762 | 3.4683
Sc, Scandium (ppm) | 2.762 | 2.4991 | 3.0249 | 2.6246 | 2.8994
Se, Selenium (ppm) | 10.8076 | 9.7281 | 11.8871 | 10.3907 | 11.2246
Sm, Samarium (ppm) | 3.5003 | 2.8952 | 4.1053 | 3.2892 | 3.7113
Sn, Tin (ppm) | 2.4509 | 2.345 | 2.5569 | 2.3618 | 2.5401
Sr, Strontium (ppm) | 15.4311 | 13.8202 | 17.0419 | 14.8099 | 16.0522
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.20255 | 0.15078 | 0.25431 | 0.18468 | 0.22041
Te, Tellurium (ppm) | 0.075122 | 0.057778 | 0.092466 | IND | IND
Th, Thorium (ppm) | 6.1639 | 5.6502 | 6.6776 | 5.903 | 6.4248
U, Uranium (ppm) | 3.5671 | 3.3743 | 3.76 | 3.4581 | 3.6762
V, Vanadium (ppm) | 60.3315 | 56.3647 | 64.2983 | 57.9864 | 62.6766
W, Tungsten (ppm) | 0.98283 | 0.90412 | 1.0615 | 0.9256 | 1.0401
Y, Yttrium (ppm) | 4.0253 | 3.744 | 4.3066 | 3.858 | 4.1925
Yb, Ytterbium (ppm) | 0.3827 | 0.31513 | 0.45027 | 0.357 | 0.4084
Zn, Zinc (ppm) | 37.1266 | 35.4744 | 38.7788 | 36.0515 | 38.2017
Zr, Zirconium (ppm) | 11.0788 | 10.2494 | 11.9081 | 10.6559 | 11.5017
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 556b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Ag | ppm | 0.021333 | HfO2 | ppm | 428.8131 | Sc | ppm | 11.8981
As | ppm | 85.5388 | Hg | ppm | < 100 | Se | ppm | 2.3333
Be | ppm | 3 | Ho | ppm | 1.2228 | Sm | ppm | 8.7876
Bi | ppm | < 100 | In | ppm | 3.8 | Sn | ppm | 105.9333
Cd | ppm | < 100 | La | ppm | 76.7525 | Ta | ppm | 1.25
Ce | ppm | 153.065 | Lu | ppm | 0.50808 | Tb | ppm | 1.0185
Cl | ppm | 249.1667 | Mo | ppm | < 50 | Te | ppm | 11.1333
Cr2O3 | ppm | 184.5868 | Na2O | wt.% | 0.10694 | Th | ppm | 15.6418
Cs | ppm | 1.9833 | Nb | ppm | 55.8883 | Tl | ppm | 0.84667
Dy | ppm | 5.9748 | Nd | ppm | 59.611 | Tm | ppm | 0.51608
Er | ppm | 3.4055 | Pb | ppm | 80.9596 | U | ppm | 8.9455
Eu | ppm | 0.83083 | Pr | ppm | 16.4624 | W | ppm | 6.26
Ga | ppm | 28.8167 | Rb | ppm | 98.7733 | Y | ppm | 33.0882
Gd | ppm | 6.4842 | Re | ppm | 22.5833 | Yb | ppm | 3.295
Ge | ppm | < 100 | Sb | ppm | < 50 | Zn | ppm | 44.3667
Peroxide Fusion ICP
Ag | ppm | 0.45333 | Na | wt.% | 0.045 | Te | ppm | < 2
Cd | ppm | < 10 | Re | ppm | 0.059167 | Zr | ppm | 209.2336
Hf | ppm | 6.4137 | Se | ppm | 16.7222
Hg | ppm | 0.083333 | Sn | ppm | 3.6359
4-Acid Digestion
Hg | ppm | 0.13 | Pt | ppb | 11.1667
Aqua Regia Digestion
Ir | ppm | < 0.003 | Pt | ppb | < 5 | Tm | ppm | 0.056284
Lu | ppm | 0.060117 | Ti | wt.% | 0.0026059
Pd | ppb | < 10 | Tl | ppm | 0.36686
3-Acid Digestion (no HF)
Ag | ppm | 0.39 | Gd | ppm | 5.4667 | S | wt.% | 7.8767
Al2O3 | wt.% | 10.5 | Hf | ppm | 5.5 | Sc | ppm | 11.4167
Ba | ppm | 180.1667 | Ho | ppm | 0.85667 | Sm | ppm | 8.1833
Be | ppm | 5.45 | K2O | wt.% | 2.6833 | Sn | ppm | 3.6167
Bi | ppm | 3.4667 | La | ppm | 77.3333 | Sr | ppm | 67.1667
CaO | wt.% | 0.17667 | Li | ppm | 52.6667 | Ta | ppm | 0.31
Cd | ppm | 0.65 | MgO | wt.% | 4.0283 | Tb | ppm | 0.675
Ce | ppm | 137 | MnO | wt.% | 0.14633 | Th | ppm | 17.25
Co | wt.% | 2.0817 | Mo | ppm | 20.8167 | TiO2 | wt.% | 0.66867
Cr | ppm | 97.6667 | Na2O | wt.% | 0.066667 | U | ppm | 9.4983
Cs | ppm | 1.7833 | Nb | ppm | 20.3167 | V | ppm | 332.3333
Cu | wt.% | 8.9367 | Nd | ppm | 64.4833 | W | ppm | 2.95
Dy | ppm | 4.6233 | Ni | ppm | 383 | Y | ppm | 25.8667
Er | ppm | 2.75 | P2O5 | wt.% | < 0.002 | Yb | ppm | 2.735
Eu | ppm | 0.93667 | Pb | ppm | 19.6833 | Zn | ppm | 75.8333
Fe2O3 | wt.% | 9.3383 | Pr | ppm | 17.1667 | Zr | ppm | 178.6667
Ga | ppm | 27.6167 | Rb | ppm | 81.8
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 556b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 10.9015 | 0.24315 | 10.4152 | 11.3878 | 10.172 | 11.6309 | 2.23% | 4.461% | 6.691% | 10.3564 | 11.4465
BaO, ppm | 242.1257 | 60.4157 | 121.2943 | 362.957 | 60.8786 | 423.3727 | 24.95% | 49.9% | 74.86% | 230.0194 | 254.232
CaO, wt.% | 0.19527 | 0.0080415 | 0.17918 | 0.21135 | 0.17114 | 0.21939 | 4.118% | 8.236% | 12.35% | 0.1855 | 0.20503
Co, wt.% | 1.9906 | 0.05579 | 1.879 | 2.1021 | 1.8232 | 2.1579 | 2.803% | 5.605% | 8.408% | 1.891 | 2.0901
Cu, wt.% | 8.5259 | 0.2568 | 8.0123 | 9.0395 | 7.7555 | 9.2963 | 3.012% | 6.024% | 9.036% | 8.0996 | 8.9522
Fe2O3, wt.% | 9.6806 | 0.21112 | 9.2584 | 10.1029 | 9.0473 | 10.314 | 2.181% | 4.362% | 6.542% | 9.1966 | 10.1647
K2O, wt.% | 2.7415 | 0.069862 | 2.6018 | 2.8813 | 2.532 | 2.9511 | 2.548% | 5.097% | 7.645% | 2.6045 | 2.8786
MgO, wt.% | 4.1091 | 0.080292 | 3.9485 | 4.2697 | 3.8683 | 4.35 | 1.954% | 3.908% | 5.862% | 3.9037 | 4.3146
MnO, wt.% | 0.15474 | 0.0059619 | 0.14282 | 0.16667 | 0.13686 | 0.17263 | 3.853% | 7.705% | 11.56% | 0.14701 | 0.16248
Ni, ppm | 343.9656 | 44.9831 | 253.9994 | 433.9318 | 209.0163 | 478.9149 | 13.08% | 26.16% | 39.23% | 326.7673 | 361.1639
P2O5, wt.% | 0.079854 | 0.0069467 | 0.065961 | 0.093747 | 0.059014 | 0.10069 | 8.699% | 17.4% | 26.1% | 0.075861 | 0.083847
SiO2, wt.% | 42.7709 | 0.70849 | 41.3539 | 44.1879 | 40.6454 | 44.8963 | 1.656% | 3.313% | 4.969% | 40.6323 | 44.9094
SO3, wt.% | 21.4269 | 1.4574 | 18.5121 | 24.3417 | 17.0547 | 25.7991 | 6.802% | 13.6% | 20.41% | 20.3556 | 22.4983
Sr, ppm | 63.2432 | 15.6746 | 31.8941 | 94.5923 | 16.2195 | 110.2669 | 24.78% | 49.57% | 74.35% | 60.081 | 66.4054
TiO2, wt.% | 0.70703 | 0.032371 | 0.64228 | 0.77177 | 0.60991 | 0.80414 | 4.578% | 9.157% | 13.74% | 0.67167 | 0.74238
V2O5, ppm | 640.9222 | 80.2714 | 480.3794 | 801.465 | 400.108 | 881.7364 | 12.52% | 25.05% | 37.57% | 608.8761 | 672.9683
Zr, ppm | 216.113 | 30.4217 | 155.2695 | 276.9565 | 124.8478 | 307.3783 | 14.08% | 28.15% | 42.23% | 205.3074 | 226.9187
Thermogravimetry
LOI1000, wt.% | 14.409 | 0.3317 | 13.7456 | 15.0724 | 13.4139 | 15.4041 | 2.302% | 4.604% | 6.906% | 13.6886 | 15.1295
Infrared Combustion
C, wt.% | 4.5805 | 0.19302 | 4.1944 | 4.9665 | 4.0014 | 5.1596 | 4.214% | 8.428% | 12.64% | 4.3515 | 4.8095
S, wt.% | 8.7844 | 0.13797 | 8.5084 | 9.0603 | 8.3705 | 9.1983 | 1.571% | 3.141% | 4.712% | 8.3452 | 9.2236
Sulphuric Acid 5% Leach
Co, wt.% | 0.68534 | 0.045222 | 0.5949 | 0.77578 | 0.54967 | 0.82101 | 6.599% | 13.2% | 19.8% | 0.65107 | 0.71961
Cu, wt.% | 1.318 | 0.093452 | 1.1311 | 1.5049 | 1.0376 | 1.5984 | 7.09% | 14.18% | 21.27% | 1.2521 | 1.3839
Peroxide Fusion ICP
Al, wt.% | 5.7078 | 0.16397 | 5.3799 | 6.0358 | 5.2159 | 6.1997 | 2.873% | 5.745% | 8.618% | 5.4225 | 5.9932
As, ppm | 101.3026 | 7.2135 | 86.8756 | 115.7297 | 79.6621 | 122.9432 | 7.121% | 14.24% | 21.36% | 96.2375 | 106.3678
B, ppm | 199.287 | 37.1604 | 124.9662 | 273.6079 | 87.8057 | 310.7683 | 18.65% | 37.29% | 55.94% | 189.3227 | 209.2514
Ba, ppm | 195.6728 | 8.8098 | 178.0532 | 213.2924 | 169.2434 | 222.1022 | 4.502% | 9.005% | 13.51% | 185.8891 | 205.4564
Be, ppm | 5.061 | 0.41275 | 4.2355 | 5.8865 | 3.8228 | 6.2993 | 8.156% | 16.31% | 24.47% | 4.808 | 5.3141
Bi, ppm | 3.0753 | 0.13075 | 2.8138 | 3.3368 | 2.6831 | 3.4676 | 4.251% | 8.503% | 12.75% | 2.9216 | 3.2291
Ca, wt.% | 0.14866 | 0.027938 | 0.09278 | 0.20453 | 0.064842 | 0.23247 | 18.79% | 37.59% | 56.38% | 0.14122 | 0.15609
Ce, ppm | 155.9579 | 5.7882 | 144.3815 | 167.5343 | 138.5932 | 173.3226 | 3.711% | 7.423% | 11.13% | 148.16 | 163.7558
Co, wt.% | 1.9911 | 0.078609 | 1.8339 | 2.1484 | 1.7553 | 2.227 | 3.948% | 7.896% | 11.84% | 1.8916 | 2.0907
Cr, ppm | 112.6336 | 17.8778 | 76.878 | 148.3892 | 59.0001 | 166.267 | 15.87% | 31.75% | 47.62% | 107.0019 | 118.2653
Cs, ppm | 1.8586 | 0.11761 | 1.6234 | 2.0938 | 1.5057 | 2.2114 | 6.328% | 12.66% | 18.98% | 1.7656 | 1.9515
Cu, wt.% | 8.6302 | 0.27055 | 8.0891 | 9.1713 | 7.8185 | 9.4418 | 3.135% | 6.27% | 9.405% | 8.1987 | 9.0617
Dy, ppm | 6.278 | 0.28988 | 5.6982 | 6.8578 | 5.4083 | 7.1476 | 4.617% | 9.235% | 13.85% | 5.9641 | 6.5919
Er, ppm | 3.4664 | 0.22555 | 3.0153 | 3.9175 | 2.7897 | 4.143 | 6.507% | 13.01% | 19.52% | 3.2931 | 3.6397
Eu, ppm | 0.87762 | 0.1294 | 0.61881 | 1.1364 | 0.48941 | 1.2658 | 14.74% | 29.49% | 44.23% | 0.83374 | 0.9215
Fe, wt.% | 6.7301 | 0.21063 | 6.3088 | 7.1514 | 6.0982 | 7.362 | 3.13% | 6.259% | 9.389% | 6.3936 | 7.0666
Ga, ppm | 28.6577 | 1.9003 | 24.8572 | 32.4582 | 22.957 | 34.3585 | 6.631% | 13.26% | 19.89% | 27.2248 | 30.0906
Gd, ppm | 6.8993 | 0.32697 | 6.2454 | 7.5533 | 5.9184 | 7.8803 | 4.739% | 9.478% | 14.22% | 6.5544 | 7.2443
Ge, ppm | 2.8039 | 0.36401 | 2.0759 | 3.5319 | 1.7119 | 3.896 | 12.98% | 25.96% | 38.95% | 2.6637 | 2.9441
Ho, ppm | 1.2166 | 0.092695 | 1.0312 | 1.402 | 0.93853 | 1.4947 | 7.619% | 15.24% | 22.86% | 1.1558 | 1.2774
In, ppm | 3.5174 | 0.21484 | 3.0877 | 3.9471 | 2.8729 | 4.1619 | 6.108% | 12.22% | 18.32% | 3.3415 | 3.6933
K, wt.% | 2.2696 | 0.084713 | 2.1001 | 2.439 | 2.0154 | 2.5237 | 3.733% | 7.465% | 11.2% | 2.1561 | 2.3831
La, ppm | 75.2334 | 3.89 | 67.4533 | 83.0134 | 63.5633 | 86.9035 | 5.171% | 10.34% | 15.51% | 71.4717 | 78.995
Li, ppm | 51.0031 | 5.1256 | 40.7519 | 61.2543 | 35.6263 | 66.3799 | 10.05% | 20.1% | 30.15% | 48.4529 | 53.5532
Lu, ppm | 0.53412 | 0.089766 | 0.35459 | 0.71366 | 0.26483 | 0.80342 | 16.81% | 33.61% | 50.42% | 0.50742 | 0.56083
Mg, wt.% | 2.4886 | 0.087284 | 2.3141 | 2.6632 | 2.2268 | 2.7505 | 3.507% | 7.015% | 10.52% | 2.3642 | 2.6131
Mn, wt.% | 0.12072 | 0.0073167 | 0.10609 | 0.13535 | 0.098769 | 0.14267 | 6.061% | 12.12% | 18.18% | 0.11468 | 0.12676
Mo, ppm | 18.7411 | 1.9746 | 14.7918 | 22.6903 | 12.8171 | 24.665 | 10.54% | 21.07% | 31.61% | 17.804 | 19.6781
Nb, ppm | 22.4636 | 1.7368 | 18.9899 | 25.9372 | 17.2531 | 27.674 | 7.732% | 15.46% | 23.2% | 21.3404 | 23.5867
Nd, ppm | 61.8593 | 3.3101 | 55.2392 | 68.4795 | 51.9291 | 71.7896 | 5.351% | 10.7% | 16.05% | 58.7664 | 64.9523
Ni, ppm | 342.6253 | 25.5389 | 291.5475 | 393.7031 | 266.0087 | 419.242 | 7.454% | 14.91% | 22.36% | 325.494 | 359.7566
P, wt.% | 0.034683 | 0.0065073 | 0.021669 | 0.047698 | 0.015161 | 0.054205 | 18.76% | 37.52% | 56.29% | 0.032949 | 0.036417
Pb, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Pr, ppm | 18.2767 | 0.83368 | 16.6094 | 19.9441 | 15.7757 | 20.7778 | 4.561% | 9.123% | 13.68% | 17.3629 | 19.1906
Rb, ppm | 82.0004 | 2.9194 | 76.1617 | 87.8391 | 73.2423 | 90.7585 | 3.56% | 7.12% | 10.68% | 77.9004 | 86.1004
S, wt.% | 8.7525 | 0.27655 | 8.1994 | 9.3057 | 7.9229 | 9.5822 | 3.16% | 6.319% | 9.479% | 8.3149 | 9.1902
Sb, ppm | 5.7737 | 0.44967 | 4.8743 | 6.673 | 4.4247 | 7.1227 | 7.788% | 15.58% | 23.36% | 5.485 | 6.0623
Sc, ppm | 9.5602 | 1.809 | 5.9422 | 13.1782 | 4.1332 | 14.9872 | 18.92% | 37.84% | 56.77% | 9.0822 | 10.0382
Si, wt.% | 20.4762 | 0.47864 | 19.5189 | 21.4335 | 19.0403 | 21.9121 | 2.338% | 4.675% | 7.013% | 19.4524 | 21.5
Sm, ppm | 8.8758 | 0.97446 | 6.9269 | 10.8248 | 5.9524 | 11.7992 | 10.98% | 21.96% | 32.94% | 8.432 | 9.3196
Sr, ppm | 73.3151 | 6.8649 | 59.5853 | 87.0449 | 52.7204 | 93.9097 | 9.364% | 18.73% | 28.09% | 69.6493 | 76.9808
Ta, ppm | 1.8409 | 0.26972 | 1.3014 | 2.3803 | 1.0317 | 2.65 | 14.65% | 29.3% | 43.96% | 1.7488 | 1.9329
Tb, ppm | 1.0544 | 0.061542 | 0.93131 | 1.1775 | 0.86977 | 1.239 | 5.837% | 11.67% | 17.51% | 1.0017 | 1.1071
Th, ppm | 15.5063 | 0.87279 | 13.7607 | 17.2519 | 12.8879 | 18.1247 | 5.629% | 11.26% | 16.89% | 14.731 | 16.2816
Ti, wt.% | 0.41403 | 0.014126 | 0.38577 | 0.44228 | 0.37165 | 0.4564 | 3.412% | 6.824% | 10.24% | 0.39332 | 0.43473
Tl, ppm | 1.06 | 0.095812 | 0.86834 | 1.2516 | 0.77253 | 1.3474 | 9.039% | 18.08% | 27.12% | 1.007 | 1.113
Tm, ppm | 0.52179 | 0.039512 | 0.44276 | 0.60081 | 0.40325 | 0.64033 | 7.572% | 15.14% | 22.72% | 0.4957 | 0.54788
U, ppm | 8.8503 | 0.8968 | 7.0567 | 10.6439 | 6.1599 | 11.5407 | 10.13% | 20.27% | 30.4% | 8.4077 | 9.2928
V, ppm | 352.7061 | 19.4963 | 313.7135 | 391.6987 | 294.2172 | 411.1949 | 5.528% | 11.06% | 16.58% | 335.0708 | 370.3414
W, ppm | 2.9639 | 0.47756 | 2.0088 | 3.919 | 1.5312 | 4.3966 | 16.11% | 32.23% | 48.34% | 2.8157 | 3.1121
Y, ppm | 32.611 | 2.5957 | 27.4197 | 37.8024 | 24.824 | 40.3981 | 7.96% | 15.92% | 23.88% | 30.9805 | 34.2416
Yb, ppm | 3.1916 | 0.271 | 2.6495 | 3.7336 | 2.3785 | 4.0046 | 8.491% | 16.98% | 25.47% | 3.032 | 3.3511
Zn, ppm | 41.353 | 3.5435 | 34.2659 | 48.44 | 30.7224 | 51.9836 | 8.569% | 17.14% | 25.71% | 39.2853 | 43.4206
4-Acid Digestion
Ag, ppm | 0.14397 | 0.012954 | 0.11806 | 0.16988 | 0.10511 | 0.18283 | 8.998% | 18% | 26.99% | 0.13677 | 0.15117
Al, wt.% | 5.6681 | 0.15242 | 5.3633 | 5.973 | 5.2109 | 6.1254 | 2.689% | 5.378% | 8.067% | 5.3847 | 5.9515
As, ppm | 98.2191 | 4.6485 | 88.9221 | 107.516 | 84.2737 | 112.1645 | 4.733% | 9.466% | 14.2% | 93.3081 | 103.13
Ba, ppm | 189.7248 | 9.9127 | 169.8994 | 209.5502 | 159.9867 | 219.4629 | 5.225% | 10.45% | 15.67% | 180.2385 | 199.211
Be, ppm | 4.5183 | 0.20982 | 4.0987 | 4.9379 | 3.8888 | 5.1478 | 4.644% | 9.288% | 13.93% | 4.2924 | 4.7442
Bi, ppm | 2.894 | 0.17753 | 2.5389 | 3.249 | 2.3614 | 3.4265 | 6.134% | 12.27% | 18.4% | 2.7493 | 3.0387
Ca, wt.% | 0.14107 | 0.0094556 | 0.12216 | 0.15998 | 0.1127 | 0.16944 | 6.703% | 13.41% | 20.11% | 0.13402 | 0.14812
Cd, ppm | 0.51006 | 0.058655 | 0.39275 | 0.62737 | 0.3341 | 0.68603 | 11.5% | 23% | 34.5% | 0.48456 | 0.53556
Ce, ppm | 143.3697 | 7.0009 | 129.3679 | 157.3715 | 122.367 | 164.3724 | 4.883% | 9.766% | 14.65% | 136.2012 | 150.5382
Co, wt.% | 1.9932 | 0.060505 | 1.8722 | 2.1143 | 1.8117 | 2.1748 | 3.035% | 6.071% | 9.106% | 1.8936 | 2.0929
Cr, ppm | 88.9552 | 7.8138 | 73.3275 | 104.5828 | 65.5137 | 112.3966 | 8.784% | 17.57% | 26.35% | 84.5074 | 93.4029
Cs, ppm | 1.7544 | 0.10111 | 1.5522 | 1.9566 | 1.4511 | 2.0578 | 5.763% | 11.53% | 17.29% | 1.6667 | 1.8421
Cu, wt.% | 8.6408 | 0.21857 | 8.2037 | 9.078 | 7.9851 | 9.2965 | 2.53% | 5.059% | 7.589% | 8.2088 | 9.0729
Dy, ppm | 4.5357 | 0.48418 | 3.5673 | 5.5041 | 3.0832 | 5.9883 | 10.67% | 21.35% | 32.02% | 4.3089 | 4.7625
Er, ppm | 2.6977 | 0.50748 | 1.6827 | 3.7127 | 1.1753 | 4.2201 | 18.81% | 37.62% | 56.43% | 2.5628 | 2.8326
Eu, ppm | 0.88976 | 0.066905 | 0.75595 | 1.0236 | 0.68904 | 1.0905 | 7.519% | 15.04% | 22.56% | 0.84527 | 0.93424
Fe, wt.% | 6.7076 | 0.24479 | 6.2181 | 7.1972 | 5.9733 | 7.442 | 3.649% | 7.299% | 10.95% | 6.3723 | 7.043
Ga, ppm | 28.1302 | 1.3039 | 25.5224 | 30.7381 | 24.2184 | 32.042 | 4.635% | 9.271% | 13.91% | 26.7237 | 29.5367
Gd, ppm | 4.9221 | 1.1397 | 2.6427 | 7.2016 | 1.5029 | 8.3413 | 23.16% | 46.31% | 69.47% | 4.676 | 5.1682
Ge, ppm | 0.22821 | 0.03067 | 0.16687 | 0.28955 | 0.1362 | 0.32022 | 13.44% | 26.88% | 40.32% | 0.2168 | 0.23962
Hf, ppm | 4.9448 | 0.21833 | 4.5081 | 5.3815 | 4.2898 | 5.5998 | 4.415% | 8.831% | 13.25% | 4.6976 | 5.192
Ho, ppm | 0.88682 | 0.1845 | 0.51782 | 1.2558 | 0.33332 | 1.4403 | 20.8% | 41.61% | 62.41% | 0.84248 | 0.93116
In, ppm | 3.564 | 0.18777 | 3.1884 | 3.9395 | 3.0007 | 4.1273 | 5.268% | 10.54% | 15.81% | 3.3858 | 3.7422
K, wt.% | 2.2457 | 0.082053 | 2.0816 | 2.4098 | 1.9995 | 2.4919 | 3.654% | 7.308% | 10.96% | 2.1334 | 2.358
La, ppm | 69.7876 | 4.3403 | 61.1071 | 78.4681 | 56.7668 | 82.8084 | 6.219% | 12.44% | 18.66% | 66.2982 | 73.277
Li, ppm | 51.6024 | 2.6754 | 46.2517 | 56.9532 | 43.5763 | 59.6285 | 5.185% | 10.37% | 15.55% | 49.0223 | 54.1826
Lu, ppm | 0.40632 | 0.047426 | 0.31147 | 0.50118 | 0.26405 | 0.5486 | 11.67% | 23.34% | 35.02% | 0.38601 | 0.42664
Mg, wt.% | 2.4254 | 0.10063 | 2.2241 | 2.6266 | 2.1235 | 2.7272 | 4.149% | 8.298% | 12.45% | 2.3041 | 2.5466
Mn, wt.% | 0.11719 | 0.0062713 | 0.10465 | 0.12974 | 0.09838 | 0.13601 | 5.351% | 10.7% | 16.05% | 0.11133 | 0.12305
Mo, ppm | 18.353 | 0.96526 | 16.4225 | 20.2835 | 15.4572 | 21.2488 | 5.259% | 10.52% | 15.78% | 17.4353 | 19.2706
Na, wt.% | 0.049017 | 0.0087808 | 0.031455 | 0.066579 | 0.022675 | 0.07536 | 17.91% | 35.83% | 53.74% | 0.046566 | 0.051468
Nb, ppm | 14.0083 | 2.3962 | 9.2158 | 18.8007 | 6.8196 | 21.1969 | 17.11% | 34.21% | 51.32% | 13.3078 | 14.7087
Nd, ppm | 62.9858 | 2.3165 | 58.3527 | 67.6189 | 56.0362 | 69.9355 | 3.678% | 7.356% | 11.03% | 59.8365 | 66.1351
Ni, ppm | 336.7816 | 18.63 | 299.5215 | 374.0417 | 280.8915 | 392.6718 | 5.532% | 11.06% | 16.6% | 319.9426 | 353.6207
P, wt.% | 0.032675 | 0.0023649 | 0.027945 | 0.037404 | 0.02558 | 0.039769 | 7.238% | 14.48% | 21.71% | 0.031041 | 0.034308
Pb, ppm | 17.2652 | 1.1536 | 14.958 | 19.5725 | 13.8044 | 20.7261 | 6.682% | 13.36% | 20.04% | 16.402 | 18.1285
Pr, ppm | 16.5315 | 1.6232 | 13.2851 | 19.7778 | 11.6619 | 21.401 | 9.819% | 19.64% | 29.46% | 15.7049 | 17.358
Rb, ppm | 82.3175 | 5.1535 | 72.0105 | 92.6245 | 66.857 | 97.778 | 6.261% | 12.52% | 18.78% | 78.2016 | 86.4333
Re, ppm | 0.032724 | 0.0033363 | 0.026051 | 0.039396 | 0.022715 | 0.042733 | 10.2% | 20.39% | 30.59% | 0.031088 | 0.03436
S, wt.% | 8.5467 | 0.33156 | 7.8836 | 9.2098 | 7.552 | 9.5414 | 3.879% | 7.759% | 11.64% | 8.1193 | 8.974
Sb, ppm | 5.2929 | 0.4733 | 4.3463 | 6.2395 | 3.873 | 6.7128 | 8.942% | 17.88% | 26.83% | 5.0283 | 5.5576
Sc, ppm | 11.03 | 0.80728 | 9.4154 | 12.6445 | 8.6081 | 13.4518 | 7.319% | 14.64% | 21.96% | 10.4785 | 11.5815
Se, ppm | 11.1614 | 1.2397 | 8.682 | 13.6409 | 7.4422 | 14.8807 | 11.11% | 22.21% | 33.32% | 10.6034 | 11.7195
Sm, ppm | 8.5793 | 0.57257 | 7.4342 | 9.7245 | 6.8616 | 10.297 | 6.674% | 13.35% | 20.02% | 8.1504 | 9.0083
Sn, ppm | 3.7046 | 0.30048 | 3.1037 | 4.3056 | 2.8032 | 4.6061 | 8.111% | 16.22% | 24.33% | 3.5194 | 3.8899
Sr, ppm | 67.5203 | 3.0461 | 61.4282 | 73.6124 | 58.3821 | 76.6585 | 4.511% | 9.023% | 13.53% | 64.1443 | 70.8963
Ta, ppm | 1.0427 | 0.1556 | 0.7315 | 1.3539 | 0.5759 | 1.5095 | 14.92% | 29.85% | 44.77% | 0.99057 | 1.0948
Tb, ppm | 0.77875 | 0.16626 | 0.44623 | 1.1113 | 0.27997 | 1.2775 | 21.35% | 42.7% | 64.05% | 0.73981 | 0.81769
Te, ppm | 0.089187 | 0.022804 | 0.043579 | 0.13479 | 0.020775 | 0.1576 | 25.57% | 51.14% | 76.71% | 0.084727 | 0.093646
Th, ppm | 15.4439 | 0.86869 | 13.7065 | 17.1813 | 12.8378 | 18.05 | 5.625% | 11.25% | 16.87% | 14.6717 | 16.2161
Ti, wt.% | 0.23512 | 0.051179 | 0.13277 | 0.33748 | 0.081587 | 0.38866 | 21.77% | 43.53% | 65.3% | 0.22337 | 0.24688
Tl, ppm | 1.0561 | 0.077169 | 0.90176 | 1.2104 | 0.8246 | 1.2876 | 7.307% | 14.61% | 21.92% | 1.0033 | 1.1089
Tm, ppm | 0.39318 | 0.043465 | 0.30625 | 0.48011 | 0.26278 | 0.52358 | 11.05% | 22.11% | 33.16% | 0.37352 | 0.41284
U, ppm | 8.7957 | 0.62463 | 7.5464 | 10.045 | 6.9218 | 10.6696 | 7.102% | 14.2% | 21.3% | 8.3559 | 9.2355
V, ppm | 330.3722 | 16.6725 | 297.0271 | 363.7173 | 280.3546 | 380.3899 | 5.047% | 10.09% | 15.14% | 313.8536 | 346.8908
W, ppm | 2.2981 | 0.18111 | 1.9359 | 2.6603 | 1.7548 | 2.8414 | 7.881% | 15.76% | 23.64% | 2.1832 | 2.413
Y, ppm | 23.4511 | 0.82176 | 21.8075 | 25.0946 | 20.9858 | 25.9164 | 3.504% | 7.008% | 10.51% | 22.2785 | 24.6236
Yb, ppm | 2.7218 | 0.34166 | 2.0385 | 3.4051 | 1.6968 | 3.7468 | 12.55% | 25.11% | 37.66% | 2.5857 | 2.8579
Zn, ppm | 38.8027 | 2.6133 | 33.5761 | 44.0293 | 30.9628 | 46.6426 | 6.735% | 13.47% | 20.2% | 36.8626 | 40.7428
Zr, ppm | 171.093 | 6.3044 | 158.4842 | 183.7018 | 152.1798 | 190.0062 | 3.685% | 7.37% | 11.05% | 162.5383 | 179.6476
Aqua Regia Digestion
Ag, ppm | 0.11978 | 0.016777 | 0.086227 | 0.15333 | 0.06945 | 0.17011 | 14.01% | 28.01% | 42.02% | 0.11379 | 0.12577
Al, wt.% | 0.86918 | 0.14448 | 0.58021 | 1.1581 | 0.43573 | 1.3026 | 16.62% | 33.25% | 49.87% | 0.82572 | 0.91264
As, ppm | 94.0686 | 6.3083 | 81.4521 | 106.6852 | 75.1438 | 112.9934 | 6.706% | 13.41% | 20.12% | 89.3652 | 98.7721
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 52.5177 | 13.0579 | 26.4019 | 78.6335 | 13.344 | 91.6914 | 24.86% | 49.73% | 74.59% | 49.8918 | 55.1436
Be, ppm | 2.2028 | 0.25165 | 1.6995 | 2.7061 | 1.4479 | 2.9578 | 11.42% | 22.85% | 34.27% | 2.0927 | 2.313
Bi, ppm | 2.8176 | 0.18182 | 2.4539 | 3.1812 | 2.2721 | 3.363 | 6.453% | 12.91% | 19.36% | 2.6767 | 2.9584
Ca, wt.% | 0.13063 | 0.0034286 | 0.12378 | 0.13749 | 0.12035 | 0.14092 | 2.625% | 5.249% | 7.874% | 0.1241 | 0.13717
Cd, ppm | 0.51043 | 0.028566 | 0.4533 | 0.56756 | 0.42473 | 0.59612 | 5.596% | 11.19% | 16.79% | 0.48491 | 0.53595
Ce, ppm | 43.5722 | 5.9993 | 31.5736 | 55.5707 | 25.5743 | 61.57 | 13.77% | 27.54% | 41.31% | 41.3935 | 45.7508
Co, wt.% | 1.9765 | 0.076468 | 1.8235 | 2.1294 | 1.7471 | 2.2059 | 3.869% | 7.738% | 11.61% | 1.8776 | 2.0753
Cr, ppm | 35.0105 | 2.2305 | 30.5495 | 39.4716 | 28.3189 | 41.7021 | 6.371% | 12.74% | 19.11% | 33.26 | 36.761
Cs, ppm | 0.2938 | 0.095934 | 0.10193 | 0.48566 | 0.0059952 | 0.5816 | 32.65% | 65.31% | 97.96% | 0.27911 | 0.30849
Cu, wt.% | 8.5536 | 0.1813 | 8.191 | 8.9162 | 8.0097 | 9.0975 | 2.12% | 4.239% | 6.359% | 8.1259 | 8.9812
Dy, ppm | 1.0157 | 0.10786 | 0.79999 | 1.2314 | 0.69214 | 1.3393 | 10.62% | 21.24% | 31.86% | 0.96492 | 1.0665
Er, ppm | 0.41235 | 0.090279 | 0.23179 | 0.59291 | 0.14152 | 0.68319 | 21.89% | 43.79% | 65.68% | 0.39174 | 0.43297
Eu, ppm | 0.38311 | 0.035097 | 0.31291 | 0.4533 | 0.27781 | 0.4884 | 9.161% | 18.32% | 27.48% | 0.36395 | 0.40226
Fe, wt.% | 6.6451 | 0.19686 | 6.2514 | 7.0389 | 6.0546 | 7.2357 | 2.962% | 5.925% | 8.887% | 6.3129 | 6.9774
Ga, ppm | 9.6132 | 1.6173 | 6.3787 | 12.8477 | 4.7614 | 14.465 | 16.82% | 33.65% | 50.47% | 9.1325 | 10.0938
Gd, ppm | 1.915 | 0.20263 | 1.5097 | 2.3202 | 1.3071 | 2.5228 | 10.58% | 21.16% | 31.74% | 1.8192 | 2.0107
Ge, ppm | 0.12571 | 0.029726 | 0.066263 | 0.18517 | 0.036537 | 0.21489 | 23.65% | 47.29% | 70.94% | 0.11943 | 0.132
Hf, ppm | 0.36687 | 0.038867 | 0.28914 | 0.44461 | 0.25027 | 0.48347 | 10.59% | 21.19% | 31.78% | 0.34853 | 0.38522
Hg, ppm | 0.10605 | 0.017363 | 0.071324 | 0.14077 | 0.053961 | 0.15814 | 16.37% | 32.74% | 49.12% | 0.10075 | 0.11135
Ho, ppm | 0.16714 | 0.030756 | 0.10563 | 0.22865 | 0.074873 | 0.25941 | 18.4% | 36.8% | 55.2% | 0.15878 | 0.1755
In, ppm | 2.9356 | 0.18377 | 2.5681 | 3.3032 | 2.3843 | 3.4869 | 6.26% | 12.52% | 18.78% | 2.7889 | 3.0824
K, wt.% | 0.18728 | 0.027662 | 0.13195 | 0.2426 | 0.10429 | 0.27026 | 14.77% | 29.54% | 44.31% | 0.17791 | 0.19664
La, ppm | 19.4745 | 2.9687 | 13.537 | 25.412 | 10.5683 | 28.3807 | 15.24% | 30.49% | 45.73% | 18.5008 | 20.4482
Li, ppm | 31.2152 | 6.2519 | 18.7114 | 43.719 | 12.4595 | 49.9709 | 20.03% | 40.06% | 60.09% | 29.6544 | 32.776
Mg, wt.% | 1.8474 | 0.21141 | 1.4246 | 2.2702 | 1.2132 | 2.4816 | 11.44% | 22.89% | 34.33% | 1.755 | 1.9398
Mn, wt.% | 0.11668 | 0.0051378 | 0.1064 | 0.12696 | 0.10127 | 0.13209 | 4.403% | 8.807% | 13.21% | 0.11085 | 0.12251
Mo, ppm | 17.3801 | 1.1468 | 15.0865 | 19.6737 | 13.9397 | 20.8205 | 6.598% | 13.2% | 19.79% | 16.5111 | 18.2491
Na, wt.% | 0.016202 | 0.0045072 | 0.007188 | 0.025217 | 0.0026808 | 0.029724 | 27.82% | 55.64% | 83.45% | 0.015392 | 0.017013
Nb, ppm | 0.12569 | 0.034254 | 0.057187 | 0.1942 | 0.022933 | 0.22846 | 27.25% | 54.5% | 81.75% | 0.11941 | 0.13198
Nd, ppm | 22.9668 | 3.4935 | 15.9798 | 29.9538 | 12.4863 | 33.4473 | 15.21% | 30.42% | 45.63% | 21.8185 | 24.1152
Ni, ppm | 323.2012 | 23.3109 | 276.5793 | 369.823 | 253.2684 | 393.1339 | 7.213% | 14.43% | 21.64% | 307.0411 | 339.3612
P, wt.% | 0.018787 | 0.0016877 | 0.015411 | 0.022162 | 0.013724 | 0.02385 | 8.983% | 17.97% | 26.95% | 0.017847 | 0.019726
Pb, ppm | 14.8661 | 0.82665 | 13.2128 | 16.5194 | 12.3862 | 17.3461 | 5.561% | 11.12% | 16.68% | 14.1228 | 15.6094
Pr, ppm | 5.7531 | 0.87234 | 4.0084 | 7.4978 | 3.1361 | 8.3701 | 15.16% | 30.33% | 45.49% | 5.4655 | 6.0408
Rb, ppm | 6.096 | 1.0865 | 3.923 | 8.269 | 2.8365 | 9.3555 | 17.82% | 35.65% | 53.47% | 5.7912 | 6.4008
Re, ppm | 0.029321 | 0.0023798 | 0.024562 | 0.034081 | 0.022182 | 0.036461 | 8.116% | 16.23% | 24.35% | 0.027855 | 0.030787
S, wt.% | 8.3791 | 0.52318 | 7.3327 | 9.4254 | 6.8096 | 9.9486 | 6.244% | 12.49% | 18.73% | 7.9601 | 8.798
Sb, ppm | 3.3222 | 0.31151 | 2.6992 | 3.9453 | 2.3877 | 4.2568 | 9.377% | 18.75% | 28.13% | 3.1561 | 3.4883
Sc, ppm | 2.762 | 0.30912 | 2.1437 | 3.3802 | 1.8346 | 3.6893 | 11.19% | 22.38% | 33.58% | 2.6239 | 2.9001
Se, ppm | 10.8076 | 1.803 | 7.2016 | 14.4137 | 5.3985 | 16.2168 | 16.68% | 33.37% | 50.05% | 10.2673 | 11.348
Sm, ppm | 3.5003 | 0.45273 | 2.5948 | 4.4057 | 2.1421 | 4.8584 | 12.93% | 25.87% | 38.8% | 3.3252 | 3.6753
Sn, ppm | 2.4509 | 0.10293 | 2.2451 | 2.6568 | 2.1422 | 2.7597 | 4.2% | 8.399% | 12.6% | 2.3284 | 2.5735
Sr, ppm | 15.4311 | 3.0329 | 9.3652 | 21.4969 | 6.3323 | 24.5299 | 19.65% | 39.31% | 58.96% | 14.6595 | 16.2026
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.20255 | 0.037549 | 0.12745 | 0.27764 | 0.089898 | 0.31519 | 18.54% | 37.08% | 55.62% | 0.19242 | 0.21267
Te, ppm | 0.075122 | 0.01469 | 0.045743 | 0.1045 | 0.031053 | 0.11919 | 19.55% | 39.11% | 58.66% | 0.071366 | 0.078878
Th, ppm | 6.1639 | 0.9393 | 4.2853 | 8.0425 | 3.346 | 8.9818 | 15.24% | 30.48% | 45.72% | 5.8557 | 6.4721
U, ppm | 3.5671 | 0.26471 | 3.0377 | 4.0966 | 2.773 | 4.3613 | 7.421% | 14.84% | 22.26% | 3.3888 | 3.7455
V, ppm | 60.3315 | 7.6553 | 45.0208 | 75.6422 | 37.3655 | 83.2975 | 12.69% | 25.38% | 38.07% | 57.3149 | 63.3481
W, ppm | 0.98283 | 0.09418 | 0.79447 | 1.1712 | 0.70029 | 1.2654 | 9.583% | 19.17% | 28.75% | 0.93369 | 1.032
Y, ppm | 4.0253 | 0.4586 | 3.1081 | 4.9425 | 2.6495 | 5.4011 | 11.39% | 22.79% | 34.18% | 3.824 | 4.2265
Yb, ppm | 0.3827 | 0.063601 | 0.2555 | 0.5099 | 0.1919 | 0.5735 | 16.62% | 33.24% | 49.86% | 0.36356 | 0.40183
Zn, ppm | 37.1266 | 2.3775 | 32.3717 | 41.8815 | 29.9942 | 44.2589 | 6.404% | 12.81% | 19.21% | 35.2703 | 38.9829
Zr, ppm | 11.0788 | 1.5816 | 7.9156 | 14.242 | 6.3339 | 15.8236 | 14.28% | 28.55% | 42.83% | 10.5248 | 11.6327
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 556b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AL-5S*AAS | 5% sulphuric acid laech with atomic absorption spectroscopy
AL-5S*OES | 5% sulphuric acid leach with inductively coupled plasma optical emission spectroscopy
AL-5S*OES/AAS | 5% sulphuric acid leach with AAS or ICP-OES finish
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PPP*XRF | pressed powder pellet with X-ray fluorescence spectroscopy
Text Values:
NA | Unable to report due to QC failure (Lab 19)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 556b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Labwest Minerals Analysis, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Fusion XRF
| Analytical results for Ag in OREAS 556b (Indicative Value 0.021 ppm) | [20 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | [17 empty] | 3
3
| 1 | 1 | < 100 | 0.026 | [17 empty] | 1
| 1 | 2 | < 100 | 0.007 | [17 empty] | 1
| 1 | 3 | < 100 | 0.025 | [17 empty] | 16
| 1 | 4 | < 100 | 0.04 | [17 empty] | 0.021333
| 1 | 5 | < 100 | 0.007 | [17 empty] | 7
| 1 | 6 | < 100 | 0.023
| Mean |  |  | 0.021333
| Median |  |  | 0.024
| Std Dev. |  |  | 0.012628
| Rel.Std.Dev. |  |  | 59.19%
| PDM3 |  |  | 1.554e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Al2O3 in OREAS 556b (Certified Value 10.9 wt.%) | [20 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 12.03 | 10.91 | 10.74 | 10.98 | 10.55 | 10.77 | 11.05 | 10.9 | 11 | 11.46 | 10.5907 | 8.4 | 10.91 | 11.27 | 10.6313 | 10.64 | 11.03 | 10.66 |  | 1
| 1 | 2 | 11.92 | 10.87 | 10.82 | 11.01 | 10.6 | 10.74 | 11.1 | 10.85 | 10.95 | 11.45 | 10.686 | 8.4 | 10.86 | 11.42 | 10.6565 | 11.03 | 11.11 | 10.79 |  | 29
| 1 | 3 | 12.09 | 10.84 | 10.83 | 10.85 | 10.63 | 10.71 | 11.1 | 10.75 | 11 | 11.39 | 10.6191 | 8.4 | 10.73 | 11.42 | 10.6636 | 11.86 | 11.03 | 10.77 |  | 16
| 1 | 4 | 11.94 | 10.86 | 10.82 | 11 | 10.66 | 10.69 | 11.05 | 10.8 | 11.05 | 11.39 | 10.7489 | 8.1 | 10.81 | 11.24 | 10.6303 | 11.22 | 10.95 | 10.84 |  | 10.9015
| 1 | 5 | 11.97 | 10.84 | 10.79 | 10.97 | 10.62 | 10.72 | 11 | 10.85 | 10.95 | 11.44 | 10.7276 | 8.4 | 10.92 | 11.37 | 10.5979 | 10.53 | 10.99 | 10.92 |  | 7
| 1 | 6 | 11.96 | 10.77 | 10.7 | 11.01 | 10.63 | 10.73 | 11.05 | 10.95 | 10.95 | 11.39 | 10.5689 | 8.4 | 10.86 | 11.54 | 10.5767 | 10.8 | 11 | 10.81
| Mean |  | 11.985 | 10.8483 | 10.7833 | 10.97 | 10.615 | 10.7267 | 11.0583 | 10.85 | 10.9833 | 11.42 | 10.6568 | 8.35 | 10.8483 | 11.3767 | 10.626 | 11.0133 | 11.0183 | 10.7983
| Median |  | 11.965 | 10.85 | 10.805 | 10.99 | 10.625 | 10.725 | 11.05 | 10.85 | 10.975 | 11.415 | 10.6525 | 8.4 | 10.86 | 11.395 | 10.6308 | 10.915 | 11.015 | 10.8
| Std Dev. |  | 0.063482 | 0.046224 | 0.052409 | 0.060992 | 0.037283 | 0.027325 | 0.037639 | 0.070711 | 0.040825 | 0.033466 | 0.074651 | 0.12247 | 0.070261 | 0.11003 | 0.033482 | 0.48546 | 0.053821 | 0.085654
| Rel.Std.Dev. |  | 0.5297% | 0.4261% | 0.486% | 0.556% | 0.3512% | 0.2547% | 0.3404% | 0.6517% | 0.3717% | 0.2931% | 0.7005% | 1.467% | 0.6477% | 0.9672% | 0.3151% | 4.408% | 0.4885% | 0.7932%
| PDM3 |  | 9.939% | -0.4874% | -1.084% | 0.6286% | -2.628% | -1.603% | 1.439% | -0.4721% | 0.7509% | 4.757% | -2.244% | -23.4% | -0.4874% | 4.359% | -2.527% | 1.026% | 1.072% | -0.9461%
| Z-Score (Absolute) |  | 4.6 | 0 | 0.27 | 0.49 | 0.95 | 0.5 | 0.85 | 0 | 0.54 | 2.31 | 0.78 | 10.11 | 0 | 2.13 | 0.9 | 0.66 | 0.68 | 0.21
| Analytical results for As in OREAS 556b (Indicative Value 86 ppm) | [20 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 08 | 12 | 13 | 17 | 24 | 25 | 28 | [11 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | 100 | 100 | < 100 | < 100 | 110 | 79.5274 | 84 | < 100 | [11 empty] | 1
| 1 | 2 | 100 | 100 | < 100 | < 100 | 110 | 79.5274 | 84 | < 100 | [11 empty] | 2
| 1 | 3 | < 100 | 100 | < 100 | < 100 | 120 | 79.5274 | 75 | < 100 | [11 empty] | 16
| 1 | 4 | 100 | 200 | 100 | < 100 | 130 | 79.5274 | 70 | < 100 | [11 empty] | 85.5388
| 1 | 5 | 100 | 200 | < 100 | < 100 | 130 | 79.5274 | 76 | < 100 | [11 empty] | 8
| 1 | 6 | 100 | 200 | < 100 | < 100 | 120 | 79.5274 | 80 | < 100
| Mean |  | 100 | 150 | 100 |  | 120 | 79.5274 | 78.1667
| Median |  | 100 | 150 | 100 |  | 120 | 79.5274 | 78
| Std Dev. |  | 0 | 54.7723 |  |  | 8.9443 | 0 | 5.5287
| Rel.Std.Dev. |  | 0% | 36.51% |  |  | 7.454% | 0% | 7.073%
| PDM3 |  | 16.91% | 75.36% | 16.91% |  | 40.29% | -7.028% | -8.618%
| Z-Score (Absolute) |  | 0.35 | 1.94 | 0.56 | 0.79 | 1.12 | 0.02 | 0.02 | 0.79
| Analytical results for BaO in OREAS 556b (Certified Value 242 ppm) | [20 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 28 |  |  |  |  | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | 300 | 260 | 300 | 200 | 200 | 410 | 276.48 | < 1000 | 227.8 | 300 | 244.8 | 193.2 | < 110 |  |  |  |  | 1
| 1 | 2 | < 100 | < 100 | 200 | 260 | 300 | 200 | 300 | 480 | 296.96 | < 1000 | 224.4 | 300 | 234.6 | 192 | 110 |  |  |  |  | 0
| 1 | 3 | < 100 | < 100 | 200 | 270 | 300 | 200 | 200 | 340 | 286.72 | < 1000 | 226.6 | 300 | 247.35 | 197.6 | 220 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | 300 | 250 | 300 | 300 | 200 | 460 | 286.72 | < 1000 | 233.3 | 300 | 295.8 | 194.3 | 110 |  |  |  |  | 242.1257
| 1 | 5 | < 100 | < 100 | 300 | 260 | 400 | 100 | 200 | 440 | 296.96 | < 1000 | 232.2 | 400 | 214.2 | 188.7 | < 110 |  |  |  |  | 8
| 1 | 6 | < 100 | < 100 | 200 | 270 | 300 | 200 | 300 | 400 | 276.48 | < 1000 | 223.3 | 300 | 247.35 | 224.4 | 110
| Mean |  |  |  | 250 | 261.6667 | 316.6667 | 200 | 233.3333 | 421.6667 | 286.72 |  | 227.9333 | 316.6667 | 247.35 | 198.3667 | 137.5
| Median |  |  |  | 250 | 260 | 300 | 200 | 200 | 425 | 286.72 |  | 227.2 | 300 | 246.075 | 193.75 | 110
| Std Dev. |  |  |  | 54.7723 | 7.5277 | 40.8248 | 63.2456 | 51.6398 | 49.9667 | 9.1589 |  | 4.0683 | 40.8248 | 26.8901 | 13.0808 | 55
| Rel.Std.Dev. |  |  |  | 21.91% | 2.877% | 12.89% | 31.62% | 22.13% | 11.85% | 3.194% |  | 1.785% | 12.89% | 10.87% | 6.594% | 40%
| PDM3 |  |  |  | 3.252% | 8.071% | 30.79% | -17.4% | -3.631% | 74.15% | 18.42% |  | -5.862% | 30.79% | 2.158% | -18.07% | -43.21%
| Z-Score (Absolute) |  | 2.72 | 2.72 | 0.04 | 0.2 | 0.95 | 0.65 | 0.19 | 2.4 | 0.54 | 3.48 | 0.27 | 0.95 | 0 | 0.67 | 1.87
| Analytical results for Be in OREAS 556b (Indicative Value 3 ppm) | [20 empty] | Indicative
Be | Round | Replicate | Lab | [18 empty] | 1
| No | No | 25 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 3 | [18 empty] | 1
| 1 | 2 | 3 | [18 empty] | 3
| 1 | 3 | 3 | [18 empty] | 16
| 1 | 4 | 3 | [18 empty] | 3
| 1 | 5 | 3 | [18 empty] | 9
| 1 | 6 | 3
| Mean |  | 3
| Median |  | 3
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 556b (Indicative Value < 100 ppm) | [20 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 08 | 12 | 13 | 25 | [14 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | 1.1 | [14 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | 1.1 | [14 empty] | 4
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | 1 | [14 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | 0.8 | [14 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | 1.1 | [14 empty] | 10
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | 1.3
| Mean | [5 empty] | 1.0667
| Median | [5 empty] | 1.1
| Std Dev. | [5 empty] | 0.1633
| Rel.Std.Dev. | [5 empty] | 15.31%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 556b (Certified Value 0.195 wt.%) | [20 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.23 | 0.15 | 0.2 | 0.2 | 0.19 | 0.198 | 0.2 | 0.19 | 0.2 | 0.09 | 0.19155 | 0.1 | 0.19 | 0.2 | 0.18382 | 0.15 | 0.19 | 0.19 |  | 1
| 1 | 2 | 0.23 | 0.15 | 0.2 | 0.2 | 0.2 | 0.196 | 0.2 | 0.19 | 0.2 | 0.09 | 0.19055 | 0.1 | 0.19 | 0.2 | 0.18685 | 0.16 | 0.2 | 0.21 |  | 33
| 1 | 3 | 0.21 | 0.15 | 0.2 | 0.21 | 0.19 | 0.199 | 0.21 | 0.2 | 0.2 | 0.09 | 0.18854 | 0.1 | 0.19 | 0.2 | 0.18887 | 0.18 | 0.2 | 0.19 |  | 16
| 1 | 4 | 0.21 | 0.15 | 0.2 | 0.2 | 0.23 | 0.194 | 0.21 | 0.2 | 0.2 | 0.08 | 0.19356 | 0.2 | 0.18 | 0.2 | 0.18685 | 0.16 | 0.2 | 0.2 |  | 0.19527
| 1 | 5 | 0.19 | 0.15 | 0.2 | 0.2 | 0.19 | 0.2 | 0.21 | 0.2 | 0.2 | 0.09 | 0.18854 | 0.1 | 0.18 | 0.2 | 0.18988 | 0.14 | 0.2 | 0.23 |  | 9
| 1 | 6 | 0.21 | 0.15 | 0.2 | 0.2 | 0.19 | 0.193 | 0.2 | 0.19 | 0.19 | 0.09 | 0.19206 | 0.1 | 0.18 | 0.2 | 0.18483 | 0.15 | 0.2 | 0.2
| Mean |  | 0.21333 | 0.15 | 0.2 | 0.20167 | 0.19833 | 0.19667 | 0.205 | 0.195 | 0.19833 | 0.088333 | 0.1908 | 0.11667 | 0.185 | 0.2 | 0.18685 | 0.15667 | 0.19833 | 0.20333
| Median |  | 0.21 | 0.15 | 0.2 | 0.2 | 0.19 | 0.197 | 0.205 | 0.195 | 0.2 | 0.09 | 0.19105 | 0.1 | 0.185 | 0.2 | 0.18685 | 0.155 | 0.2 | 0.2
| Std Dev. |  | 0.015055 | 0 | 3.0405e-17 | 0.0040825 | 0.016021 | 0.0028048 | 0.0054772 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0020016 | 0.040825 | 0.0054772 | 3.0405e-17 | 0.0023032 | 0.013663 | 0.0040825 | 0.015055
| Rel.Std.Dev. |  | 7.057% | 0% | 1.52e-14% | 2.024% | 8.078% | 1.426% | 2.672% | 2.809% | 2.058% | 4.622% | 1.049% | 34.99% | 2.961% | 1.52e-14% | 1.233% | 8.721% | 2.058% | 7.404%
| PDM3 |  | 9.253% | -23.18% | 2.425% | 3.278% | 1.571% | 0.7178% | 4.985% | -0.1358% | 1.571% | -54.76% | -2.286% | -40.25% | -5.257% | 2.425% | -4.31% | -19.77% | 1.571% | 4.132%
| Z-Score (Absolute) |  | 1.48 | 3.64 | NA | NA | 0.27 | 0.13 | 0.81 | 0 | 0.27 | 8.63 | 0.34 | NA | 0.81 | NA | 0.66 | 3.1 | NA | 0.67
Results from laboratories 4, 5, 22 and 27 were removed due to their 0.01 wt.% reading resolution. Results from laboratory 19 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 556b (Indicative Value < 100 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 13 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 0
0
| 1 | 1 | < 100 | < 100 | [17 empty] | 1
| 1 | 2 | < 100 | < 100 | [17 empty] | 5
| 1 | 3 | < 100 | < 100 | [17 empty] | 16
| 1 | 4 | < 100 | < 100 | [17 empty] | < 100
| 1 | 5 | < 100 | < 100 | [17 empty] | 11
| 1 | 6 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 556b (Indicative Value 153 ppm) | [20 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 12 | 13 | 25 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 160 | 80 | 143.85 | 146 | [15 empty] | 1
| 1 | 2 | 240 | 160 | 144.792 | 145 | [15 empty] | 6
| 1 | 3 | 160 | 80 | 144.274 | 139 | [15 empty] | 16
| 1 | 4 | 240 | 160 | 144.834 | 143 | [15 empty] | 153.065
| 1 | 5 | 240 | 160 | 144.563 | 140 | [15 empty] | 12
| 1 | 6 | 160 | 80 | 144.487 | 140
| Mean |  | 200 | 120 | 144.4667 | 142.1667
| Median |  | 200 | 120 | 144.525 | 141.5
| Std Dev. |  | 43.8178 | 43.8178 | 0.36548 | 2.9269
| Rel.Std.Dev. |  | 21.91% | 36.51% | 0.253% | 2.059%
| PDM3 |  | 30.66% | -21.6% | -5.617% | -7.12%
| Z-Score (Absolute) |  | 3.63 | 1.28 | 0.07 | 0.07
| Analytical results for Cl in OREAS 556b (Indicative Value 249 ppm) | [20 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 17 | 24 | 28 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 200 | 430 | 200 | [16 empty] | 1
| 1 | 2 | 240 | 420 | 100 | [16 empty] | 7
| 1 | 3 | 190 | 460 | 100 | [16 empty] | 16
| 1 | 4 | 210 | 420 | 100 | [16 empty] | 249.1667
| 1 | 5 | 210 | 450 | < 100 | [16 empty] | 13
| 1 | 6 | 210 | 395 | 100
| Mean |  | 210 | 429.1667 | 120
| Median |  | 210 | 425 | 100
| Std Dev. |  | 16.7332 | 23.3274 | 44.7214
| Rel.Std.Dev. |  | 7.968% | 5.436% | 37.27%
| PDM3 |  | -15.72% | 72.24% | -51.84%
| Z-Score (Absolute) |  | 0 | 1.45 | 0.67
| Analytical results for Co in OREAS 556b (Certified Value 1.99 wt.%) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 17 | 19 | 22 | 24 | 25 | 27 | 28 |  |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 2.47 | 2.06 | 2.002 | 1.9237 | 1.9867 | 1.94 | 1.96 | 2 | 1.98 | 2.0385 | 2.212 | 1.63 | 2.0292 | 1.916 | 1.87 | 1.975 | 2.08 |  |  | 1
| 1 | 2 | 2.42 | 2.07 | 1.992 | 1.9751 | 1.9803 | 1.96 | 1.94 | 1.97 | 1.97 | 2.021 | 2.243 | 1.61 | 2.034 | 1.9293 | 1.97 | 1.974 | 2.15 |  |  | 36
| 1 | 3 | 2.44 | 2.06 | 2.01 | 1.953 | 1.9699 | 1.99 | 1.97 | 1.97 | 1.96 | 2.0329 | 2.209 | 1.65 | 2.0169 | 1.9271 | 1.89 | 1.963 | 2.1 |  |  | 16
| 1 | 4 | 2.44 | 2.08 | 2.022 | 1.9751 | 1.9728 | 1.96 | 1.94 | 2.02 | 1.98 | 2.0343 | 2.196 | 1.56 | 2.0203 | 1.9211 | 1.88 | 1.974 | 2.1 |  |  | 1.9906
| 1 | 5 | 2.45 | 2.07 | 2.037 | 2.0044 | 1.9693 | 1.95 | 1.94 | 1.98 | 1.96 | 2.0354 | 2.194 | 1.64 | 2.0372 | 1.918 | 1.89 | 1.989 | 2.14 |  |  | 10
| 1 | 6 | 2.41 | 2.05 | 1.992 | 1.9604 | 1.9885 | 1.98 | 1.94 | 2 | 1.96 | 2.0321 | 2.222 | 1.65 | 2.0112 | 1.9618 | 2.22 | 2.012 | 2.11
| Mean |  | 2.4383 | 2.065 | 2.0092 | 1.9653 | 1.9779 | 1.9633 | 1.9483 | 1.99 | 1.9683 | 2.0324 | 2.2127 | 1.6233 | 2.0248 | 1.9289 | 1.9533 | 1.9812 | 2.1133
| Median |  | 2.44 | 2.065 | 2.006 | 1.9678 | 1.9766 | 1.96 | 1.94 | 1.99 | 1.965 | 2.0336 | 2.2105 | 1.635 | 2.0248 | 1.9241 | 1.89 | 1.9745 | 2.105
| Std Dev. |  | 0.02137 | 0.010488 | 0.017781 | 0.026932 | 0.0084797 | 0.018619 | 0.013292 | 0.02 | 0.0098319 | 0.0060019 | 0.018151 | 0.034448 | 0.010251 | 0.01691 | 0.13545 | 0.017221 | 0.026583
| Rel.Std.Dev. |  | 0.8764% | 0.5079% | 0.885% | 1.37% | 0.4287% | 0.9483% | 0.6822% | 1.005% | 0.4995% | 0.2953% | 0.8203% | 2.122% | 0.5063% | 0.8767% | 6.934% | 0.8692% | 1.258%
| PDM3 |  | 22.49% | 3.739% | 0.9345% | -1.27% | -0.6354% | -1.368% | -2.122% | -0.02836% | -1.117% | 2.1% | 11.16% | -18.45% | 1.72% | -3.099% | -1.87% | -0.4721% | 6.168%
| Z-Score (Absolute) |  | 9.39 | 1.72 | 0.58 | 0.33 | 0.07 | 0.37 | 0.67 | 0.18 | 0.26 | 1.05 | 4.75 | 7.35 | 0.9 | 1.07 | 0.57 | 0 | 2.71
| Analytical results for Cr2O3 in OREAS 556b (Indicative Value 185 ppm) | [20 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 25 | 27 | 28 |  |  |  | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | 200 | 100 | 200 | 150 | 290 | 100 | 150 | < 50 | 200 | < 1000 | 200 | 200 | 144 | 140 | 100 |  |  |  | 1
| 1 | 2 | < 100 | 200 | 100 | 100 | 150 | 290 | 100 | 150 | < 50 | 180 | < 1000 | 200 | 200 | 156 | 180 | 100 |  |  |  | 8
| 1 | 3 | < 100 | 200 | 200 | 100 | 150 | 150 | 100 | 290 | < 50 | 190 | < 1000 | 200 | 200 | 171 | 140 | < 100 |  |  |  | 16
| 1 | 4 | < 100 | 200 | 300 | 200 | 160 | 150 | 100 | 150 | < 50 | 180 | < 1000 | 200 | 200 | 132 | 140 | 100 |  |  |  | 184.5868
| 1 | 5 | 300 | 200 | 300 | 200 | 150 | 150 | 100 | 150 | < 50 | 200 | < 1000 | 200 | 200 | 132 | 150 | < 100 |  |  |  | 14
| 1 | 6 | 200 | 200 | 200 | 200 | 150 | 150 | 100 | 150 | < 50 | 195 | < 1000 | 200 | 200 | 144 | 150 | 100
| Mean |  | 250 | 200 | 200 | 166.6667 | 151.6667 | 196.6667 | 100 | 173.3333 |  | 190.8333 |  | 200 | 200 | 146.5 | 150 | 100
| Median |  | 250 | 200 | 200 | 200 | 150 | 150 | 100 | 150 |  | 192.5 |  | 200 | 200 | 144 | 145 | 100
| Std Dev. |  | 70.7107 | 0 | 89.4427 | 51.6398 | 4.0825 | 72.2957 | 0 | 57.1548 |  | 9.1742 |  | 0 | 0 | 14.99 | 15.4919 | 0
| Rel.Std.Dev. |  | 28.28% | 0% | 44.72% | 30.98% | 2.692% | 36.76% | 0% | 32.97% |  | 4.807% |  | 0% | 0% | 10.23% | 10.33% | 0%
| PDM3 |  | 35.44% | 8.35% | 8.35% | -9.708% | -17.83% | 6.544% | -45.82% | -6.097% |  | 3.384% |  | 8.35% | 8.35% | -20.63% | -18.74% | -45.82%
| Z-Score (Absolute) |  | 1.11 | 0.67 | 0.67 | 0.04 | 0.36 | 0.56 | 1.47 | 0.04 | 3.08 | 0.48 | 7.11 | 0.67 | 0.67 | 0.47 | 0.4 | 1.83
| Analytical results for Cs in OREAS 556b (Indicative Value 1.98 ppm) | [20 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 2
2
| 1 | 1 | < 100 | 1.9 | [17 empty] | 1
| 1 | 2 | < 100 | 1.9 | [17 empty] | 9
| 1 | 3 | < 100 | 1.9 | [17 empty] | 16
| 1 | 4 | < 100 | 2 | [17 empty] | 1.9833
| 1 | 5 | < 100 | 1.9 | [17 empty] | 15
| 1 | 6 | < 100 | 2.3
| Mean |  |  | 1.9833
| Median |  |  | 1.9
| Std Dev. |  |  | 0.16021
| Rel.Std.Dev. |  |  | 8.078%
| PDM3 |  |  | 1.554e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 556b (Certified Value 8.53 wt.%) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 17 | 19 | 22 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 8.94 | 8.47 | 8.816 | 8.2282 | 7.973 | 8.6454 | 8.7 | 8.43 | 8.54 | 8.39 | 8.3538 | > 5 | 7.005 | 8.843 | 8.4265 | 8.75 | > 8 | 8.87 |  | 1
| 1 | 2 | 8.81 | 8.4 | 8.794 | 8.308 | 7.909 | 8.332 | 8.78 | 8.1 | 8.44 | 8.43 | 8.3659 | > 5 | 7.061 | 8.404 | 8.514 | 8.65 | > 8 | 8.9 |  | 39
| 1 | 3 | 8.84 | 8.48 | 8.775 | 8.3879 | 7.877 | 8.726 | 8.77 | 8.22 | 8.39 | 8.42 | 8.3225 | > 5 | 7.075 | 8.54 | 8.5411 | 9.44 | > 8 | 8.99 |  | 16
| 1 | 4 | 8.91 | 8.51 | 8.796 | 8.2282 | 7.893 | 8.9144 | 8.77 | 8.44 | 8.52 | 8.54 | 8.3684 | > 5 | 6.777 | 8.388 | 8.4473 | 8.52 | > 8 | 8.82 |  | 8.5259
| 1 | 5 | 8.78 | 8.47 | 8.797 | 8.3879 | 7.949 | 8.7233 | 8.72 | 8.27 | 8.49 | 8.35 | 8.3315 | > 5 | 6.997 | 8.899 | 8.3768 | 8.5 | > 8 | 8.77 |  | 11
| 1 | 6 | 8.78 | 8.46 | 8.781 | 8.308 | 8.076 | 8.6804 | 8.77 | 8.37 | 8.6 | 8.41 | 8.3867 | > 5 | 7.041 | 8.468 | 8.5929 | 8.61 | > 8 | 8.76
| Mean |  | 8.8433 | 8.465 | 8.7932 | 8.308 | 7.9462 | 8.6702 | 8.7517 | 8.305 | 8.4967 | 8.4233 | 8.3548 |  | 6.9927 | 8.5903 | 8.4831 | 8.745 |  | 8.8517
| Median |  | 8.825 | 8.47 | 8.795 | 8.308 | 7.929 | 8.7019 | 8.77 | 8.32 | 8.505 | 8.415 | 8.3598 |  | 7.023 | 8.504 | 8.4806 | 8.63 |  | 8.845
| Std Dev. |  | 0.067725 | 0.036194 | 0.014303 | 0.07142 | 0.072907 | 0.19009 | 0.033116 | 0.13308 | 0.074476 | 0.06377 | 0.024137 |  | 0.10997 | 0.22462 | 0.080209 | 0.35241 |  | 0.087044
| Rel.Std.Dev. |  | 0.7658% | 0.4276% | 0.1627% | 0.8597% | 0.9175% | 2.192% | 0.3784% | 1.602% | 0.8765% | 0.7571% | 0.2889% |  | 1.573% | 2.615% | 0.9455% | 4.03% |  | 0.9834%
| PDM3 |  | 3.723% | -0.7141% | 3.135% | -2.555% | -6.799% | 1.693% | 2.648% | -2.591% | -0.3426% | -1.203% | -2.007% |  | -17.98% | 0.756% | -0.5018% | 2.57% |  | 3.821%
| Z-Score (Absolute) |  | 1.3 | 0.09 | 1.12 | 0.67 | 2 | 0.66 | 0.96 | 0.68 | 0.02 | 0.24 | 0.5 | NA | 5.51 | 0.37 | 0.02 | 0.94 | NA | 1.33
| Analytical results for Dy in OREAS 556b (Indicative Value 5.97 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 5.599 | 6.36 | [17 empty] | 1
| 1 | 2 | 5.711 | 6.34 | [17 empty] | 10
| 1 | 3 | 5.634 | 6.3 | [17 empty] | 16
| 1 | 4 | 5.687 | 6.29 | [17 empty] | 5.9748
| 1 | 5 | 5.701 | 6.03 | [17 empty] | 16
| 1 | 6 | 5.695 | 6.35
| Mean |  | 5.6712 | 6.2783
| Median |  | 5.691 | 6.32
| Std Dev. |  | 0.044463 | 0.12481
| Rel.Std.Dev. |  | 0.784% | 1.988%
| PDM3 |  | -5.081% | 5.081%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Er in OREAS 556b (Indicative Value 3.41 ppm) | [20 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 3.188 | 3.49 | [17 empty] | 1
| 1 | 2 | 3.238 | 3.72 | [17 empty] | 11
| 1 | 3 | 3.215 | 3.56 | [17 empty] | 16
| 1 | 4 | 3.224 | 3.68 | [17 empty] | 3.4055
| 1 | 5 | 3.237 | 3.45 | [17 empty] | 17
| 1 | 6 | 3.234 | 3.63
| Mean |  | 3.2227 | 3.5883
| Median |  | 3.229 | 3.595
| Std Dev. |  | 0.019138 | 0.10685
| Rel.Std.Dev. |  | 0.5939% | 2.978%
| PDM3 |  | -5.369% | 5.369%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 556b (Indicative Value 0.83 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 0.798 | 0.92 | [17 empty] | 1
| 1 | 2 | 0.81 | 0.88 | [17 empty] | 12
| 1 | 3 | 0.79 | 0.91 | [17 empty] | 16
| 1 | 4 | 0.801 | 0.73 | [17 empty] | 0.83083
| 1 | 5 | 0.811 | 0.88 | [17 empty] | 18
| 1 | 6 | 0.81 | 0.83
| Mean |  | 0.80333 | 0.85833
| Median |  | 0.8055 | 0.88
| Std Dev. |  | 0.0084774 | 0.070261
| Rel.Std.Dev. |  | 1.055% | 8.186%
| PDM3 |  | -3.31% | 3.31%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Fe2O3 in OREAS 556b (Certified Value 9.68 wt.%) | [20 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 7.24 | 10.63 | 9.69 | 9.75 | 9.55 | 9.53 | 9.836 | 9.83 | 9.722 | 11.395 | 10.222 | 9.4595 | 8.1 | 9.58 | 9.89 | 9.404 | 9.41 | 9.63 | 9.779 | 1
| 1 | 2 | 7.01 | 10.64 | 9.65 | 9.79 | 9.59 | 9.5 | 9.879 | 9.77 | 9.765 | 11.323 | 10.237 | 9.4935 | 8.1 | 9.66 | 9.89 | 9.362 | 9.7 | 9.75 | 9.751 | 43
| 1 | 3 | 7.18 | 10.66 | 9.64 | 9.76 | 9.57 | 9.47 | 9.865 | 9.75 | 9.722 | 11.466 | 10.094 | 9.4524 | 8.2 | 9.62 | 9.91 | 9.418 | 10.44 | 9.64 | 9.708 | 16
| 1 | 4 | 7.2 | 10.68 | 9.67 | 9.71 | 9.63 | 9.53 | 9.836 | 9.82 | 9.851 | 11.452 | 10.165 | 9.4775 | 7.8 | 9.67 | 9.88 | 9.499 | 9.55 | 9.58 | 9.622 | 9.6806
| 1 | 5 | 7.13 | 10.66 | 9.69 | 9.79 | 9.59 | 9.42 | 9.808 | 9.77 | 9.693 | 11.38 | 10.179 | 9.4584 | 8.1 | 9.56 | 9.84 | 9.283 | 9.08 | 9.67 | 9.851 | 12
| 1 | 6 | 7.21 | 10.68 | 9.65 | 9.73 | 9.58 | 9.48 | 9.865 | 9.9 | 9.751 | 11.423 | 10.108 | 9.4104 | 8.1 | 9.59 | 9.85 | 9.526 | 9.16 | 9.68 | 9.736
| Mean |  | 7.1617 | 10.6583 | 9.665 | 9.755 | 9.585 | 9.4883 | 9.8482 | 9.8067 | 9.7507 | 11.4065 | 10.1675 | 9.4586 | 8.0667 | 9.6133 | 9.8767 | 9.4153 | 9.5567 | 9.6583 | 9.7412
| Median |  | 7.19 | 10.66 | 9.66 | 9.755 | 9.585 | 9.49 | 9.8505 | 9.795 | 9.7365 | 11.409 | 10.172 | 9.4589 | 8.1 | 9.605 | 9.885 | 9.411 | 9.48 | 9.655 | 9.7435
| Std Dev. |  | 0.082805 | 0.020412 | 0.021679 | 0.032094 | 0.026646 | 0.041673 | 0.02618 | 0.055377 | 0.055211 | 0.052325 | 0.058106 | 0.028052 | 0.13663 | 0.044572 | 0.026583 | 0.089151 | 0.49131 | 0.057067 | 0.076051
| Rel.Std.Dev. |  | 1.156% | 0.1915% | 0.2243% | 0.329% | 0.278% | 0.4392% | 0.2658% | 0.5647% | 0.5662% | 0.4587% | 0.5715% | 0.2966% | 1.694% | 0.4636% | 0.2692% | 0.9469% | 5.141% | 0.5909% | 0.7807%
| PDM3 |  | -26.02% | 10.1% | -0.1616% | 0.7681% | -0.988% | -1.987% | 1.73% | 1.302% | 0.7233% | 17.83% | 5.029% | -2.294% | -16.67% | -0.6953% | 2.025% | -2.741% | -1.281% | -0.2305% | 0.6252%
| Z-Score (Absolute) |  | 9.55 | 3.79 | 0 | 0.34 | 0.31 | 0.67 | 0.7 | 0.54 | 0.33 | 6.65 | 1.92 | 0.79 | 6.1 | 0.2 | 0.81 | 0.95 | 0.41 | 0.03 | 0.29
| Analytical results for Ga in OREAS 556b (Indicative Value 28.8 ppm) | [20 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 1
1
| 1 | 1 | < 100 | 28.4 | [17 empty] | 1
| 1 | 2 | < 100 | 27.7 | [17 empty] | 13
| 1 | 3 | < 100 | 27.7 | [17 empty] | 16
| 1 | 4 | < 100 | 28.3 | [17 empty] | 28.8167
| 1 | 5 | < 100 | 27.9 | [17 empty] | 19
| 1 | 6 | < 100 | 32.9
| Mean |  |  | 28.8167
| Median |  |  | 28.1
| Std Dev. |  |  | 2.0223
| Rel.Std.Dev. |  |  | 7.018%
| PDM3 |  |  | -1.11e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Gd in OREAS 556b (Indicative Value 6.48 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 6.071 | 6.6 | [17 empty] | 1
| 1 | 2 | 6.16 | 7.16 | [17 empty] | 14
| 1 | 3 | 6.146 | 6.62 | [17 empty] | 16
| 1 | 4 | 6.124 | 6.96 | [17 empty] | 6.4842
| 1 | 5 | 6.132 | 6.75 | [17 empty] | 20
| 1 | 6 | 6.137 | 6.95
| Mean |  | 6.1283 | 6.84
| Median |  | 6.1345 | 6.85
| Std Dev. |  | 0.03069 | 0.22045
| Rel.Std.Dev. |  | 0.5008% | 3.223%
| PDM3 |  | -5.488% | 5.488%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ge in OREAS 556b (Indicative Value < 100 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | [18 empty] | 1
| No | No | 01 | [18 empty] | ppm
|  |  | PPP*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 15
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 21
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 556b (Indicative Value 429 ppm) | [20 empty] | Indicative
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 06 | 08 | 12 | 22 | 25 | [14 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | 800 | 7.78 | [14 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | 900 | 7.78 | [14 empty] | 1
| 1 | 3 | < 100 | < 100 | < 100 | 900 | 7.67 | [14 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | 800 | 7.43 | [14 empty] | 428.8131
| 1 | 5 | < 100 | < 100 | < 100 | 800 | 7.67 | [14 empty] | 7
| 1 | 6 | < 100 | < 100 | < 100 | 900 | 7.43
| Mean |  |  |  |  | 850 | 7.6267
| Median |  |  |  |  | 850 | 7.67
| Std Dev. |  |  |  |  | 54.7723 | 0.16008
| Rel.Std.Dev. |  |  |  |  | 6.444% | 2.099%
| PDM3 |  |  |  |  | 98.22% | -98.22%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Hg in OREAS 556b (Indicative Value < 100 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | [18 empty] | 1
| No | No | 01 | [18 empty] | ppm
|  |  | PPP*XRF | [18 empty] | 0
0
| 1 | 1 | < 100 | [18 empty] | 1
| 1 | 2 | < 100 | [18 empty] | 20
| 1 | 3 | < 100 | [18 empty] | 16
| 1 | 4 | < 100 | [18 empty] | < 100
| 1 | 5 | < 100 | [18 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 556b (Indicative Value 1.22 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 1.218 | 1.29 | [17 empty] | 1
| 1 | 2 | 1.226 | 1.23 | [17 empty] | 3
| 1 | 3 | 1.224 | 1.17 | [17 empty] | 16
| 1 | 4 | 1.226 | 1.18 | [17 empty] | 1.2228
| 1 | 5 | 1.223 | 1.2 | [17 empty] | 9
| 1 | 6 | 1.227 | 1.26
| Mean |  | 1.224 | 1.2217
| Median |  | 1.225 | 1.215
| Std Dev. |  | 0.0032863 | 0.047081
| Rel.Std.Dev. |  | 0.2685% | 3.854%
| PDM3 |  | 0.09541% | -0.09541%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 556b (Indicative Value 3.8 ppm) | [20 empty] | Indicative
In | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 2
2
| 1 | 1 | < 100 | 3.8 | [17 empty] | 1
| 1 | 2 | < 100 | 4 | [17 empty] | 4
| 1 | 3 | < 100 | 3.6 | [17 empty] | 16
| 1 | 4 | < 100 | 3.9 | [17 empty] | 3.8
| 1 | 5 | < 100 | 3.7 | [17 empty] | 10
| 1 | 6 | < 100 | 3.8
| Mean |  |  | 3.8
| Median |  |  | 3.8
| Std Dev. |  |  | 0.14142
| Rel.Std.Dev. |  |  | 3.722%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for K2O in OREAS 556b (Certified Value 2.74 wt.%) | [20 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 2.8 | 2.74 | 2.72 | 2.68 | 2.74 | 2.59 | 2.77 | 2.75 | 2.72 | 2.76 | 2.6675 | 2.2 | 2.54 | 2.64 | 2.684 | 2.83 | 2.739 | 2.83 |  | 1
| 1 | 2 | 2.91 | 2.74 | 2.71 | 2.7 | 2.76 | 2.52 | 2.77 | 2.72 | 2.73 | 2.78 | 2.6953 | 2.2 | 2.57 | 2.66 | 2.696 | 2.95 | 2.749 | 2.76 |  | 31
| 1 | 3 | 2.87 | 2.75 | 2.71 | 2.69 | 2.74 | 2.58 | 2.77 | 2.71 | 2.73 | 2.79 | 2.6744 | 2.2 | 2.52 | 2.65 | 2.687 | 3.04 | 2.744 | 2.9 |  | 16
| 1 | 4 | 2.97 | 2.78 | 2.72 | 2.66 | 2.75 | 2.52 | 2.79 | 2.75 | 2.73 | 2.76 | 2.6933 | 2.2 | 2.51 | 2.65 | 2.677 | 2.69 | 2.729 | 2.79 |  | 2.7415
| 1 | 5 | 2.87 | 2.76 | 2.72 | 2.69 | 2.77 | 2.51 | 2.81 | 2.75 | 2.71 | 2.77 | 2.6923 | 2.2 | 2.55 | 2.64 | 2.678 | 2.56 | 2.756 | 2.72 |  | 13
| 1 | 6 | 2.84 | 2.73 | 2.72 | 2.68 | 2.78 | 2.51 | 2.85 | 2.78 | 2.71 | 2.76 | 2.6596 | 2.2 | 2.54 | 2.64 | 2.687 | 2.69 | 2.746 | 2.87
| Mean |  | 2.8767 | 2.75 | 2.7167 | 2.6833 | 2.7567 | 2.5383 | 2.7933 | 2.7433 | 2.7217 | 2.77 | 2.6804 | 2.2 | 2.5383 | 2.6467 | 2.6848 | 2.7933 | 2.7438 | 2.8117
| Median |  | 2.87 | 2.745 | 2.72 | 2.685 | 2.755 | 2.52 | 2.78 | 2.75 | 2.725 | 2.765 | 2.6834 | 2.2 | 2.54 | 2.645 | 2.6855 | 2.76 | 2.745 | 2.81
| Std Dev. |  | 0.058538 | 0.017889 | 0.005164 | 0.013663 | 0.01633 | 0.03656 | 0.032042 | 0.025033 | 0.0098319 | 0.012649 | 0.015265 | 4.8648e-16 | 0.02137 | 0.008165 | 0.006969 | 0.1803 | 0.009196 | 0.067946
| Rel.Std.Dev. |  | 2.035% | 0.6505% | 0.1901% | 0.5092% | 0.5924% | 1.44% | 1.147% | 0.9125% | 0.3612% | 0.4566% | 0.5695% | 2.211e-14% | 0.8419% | 0.3085% | 0.2596% | 6.455% | 0.3352% | 2.417%
| PDM3 |  | 4.929% | 0.3087% | -0.9072% | -2.123% | 0.5519% | -7.412% | 1.889% | 0.06553% | -0.7248% | 1.038% | -2.231% | -19.75% | -7.412% | -3.46% | -2.068% | 1.889% | 0.08377% | 2.558%
| Z-Score (Absolute) |  | 1.92 | 0.23 | 0.21 | 0.65 | 0.32 | 2.59 | 0.81 | 0.14 | 0.14 | 0.5 | 0.69 | 7.09 | 2.59 | 1.14 | 0.63 | 0.81 | 0.15 | 1.05
| Analytical results for La in OREAS 556b (Indicative Value 77 ppm) | [20 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 12 | 13 | 25 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*MS | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 90 | 90 | 74.162 | 72.2 | [15 empty] | 1
| 1 | 2 | 170 | 90 | 74.438 | 71.6 | [15 empty] | 5
| 1 | 3 | 90 | 90 | 74.257 | 68.9 | [15 empty] | 16
| 1 | 4 | 170 | 90 | 74.368 | 71.2 | [15 empty] | 76.7525
| 1 | 5 | 170 | 90 | 74.415 | 69.9 | [15 empty] | 11
| 1 | 6 | 170 | < 90 | 74.309 | 70.2
| Mean |  | 143.3333 | 90 | 74.3248 | 70.6667
| Median |  | 170 | 90 | 74.3385 | 70.7
| Std Dev. |  | 41.3118 | 0 | 0.10408 | 1.2193
| Rel.Std.Dev. |  | 28.82% | 0% | 0.14% | 1.725%
| PDM3 |  | 86.75% | 17.26% | -3.163% | -7.929%
| Z-Score (Absolute) |  | 11.85 | 0.35 | 0.35 | 1
| Analytical results for Lu in OREAS 556b (Indicative Value 0.51 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 0.469 | 0.52 | [17 empty] | 1
| 1 | 2 | 0.469 | 0.52 | [17 empty] | 25
| 1 | 3 | 0.465 | 0.51 | [17 empty] | 16
| 1 | 4 | 0.467 | 0.6 | [17 empty] | 0.50808
| 1 | 5 | 0.469 | 0.56 | [17 empty] | 12
| 1 | 6 | 0.468 | 0.58
| Mean |  | 0.46783 | 0.54833
| Median |  | 0.4685 | 0.54
| Std Dev. |  | 0.0016021 | 0.037103
| Rel.Std.Dev. |  | 0.3424% | 6.767%
| PDM3 |  | -7.922% | 7.922%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 556b (Certified Value 4.11 wt.%) | [20 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 4.25 | 4.08 | 4.12 | 4.08 | 3.95 | 4.09 | 4.16 | 4.16 | 4.18 | 4.13 | 4.1204 | 3.3 | 4.27 | 4.01 | 3.984 | 4.02 | 4.23 | 4.23 |  | 1
| 1 | 2 | 4.44 | 4.07 | 4.1 | 4.12 | 3.96 | 4.06 | 4.18 | 4.14 | 4.17 | 4.14 | 4.1588 | 3.3 | 4.15 | 4.03 | 4.0221 | 4.21 | 4.2 | 4.2 |  | 34
| 1 | 3 | 4.39 | 4.07 | 4.11 | 4.07 | 3.98 | 4.06 | 4.25 | 4.1 | 4.16 | 4.14 | 4.1339 | 3.3 | 4.26 | 4.01 | 4.0057 | 4.49 | 4.2 | 4.2 |  | 16
| 1 | 4 | 4.26 | 4.06 | 4.1 | 4.1 | 3.99 | 4.05 | 4.16 | 4.13 | 4.21 | 4.12 | 4.1984 | 3.2 | 4.2 | 3.97 | 4.0108 | 4.16 | 4.19 | 4.17 |  | 4.1091
| 1 | 5 | 4.3 | 4.05 | 4.12 | 4.11 | 3.98 | 4.06 | 4.15 | 4.14 | 4.15 | 4.13 | 4.1765 | 3.3 | 4.1 | 3.98 | 4.0118 | 3.94 | 4.22 | 4.24 |  | 14
| 1 | 6 | 4.49 | 4.02 | 4.1 | 4.11 | 3.98 | 4.08 | 4.18 | 4.16 | 4.13 | 4.14 | 4.1188 | 3.3 | 4.14 | 4.01 | 3.9851 | 4.02 | 4.22 | 4.16
| Mean |  | 4.355 | 4.0583 | 4.1083 | 4.0983 | 3.9733 | 4.0667 | 4.18 | 4.1383 | 4.1667 | 4.1333 | 4.1511 | 3.2833 | 4.1867 | 4.0017 | 4.0033 | 4.14 | 4.21 | 4.2
| Median |  | 4.345 | 4.065 | 4.105 | 4.105 | 3.98 | 4.06 | 4.17 | 4.14 | 4.165 | 4.135 | 4.1464 | 3.3 | 4.175 | 4.01 | 4.0082 | 4.09 | 4.21 | 4.2
| Std Dev. |  | 0.09975 | 0.02137 | 0.0098319 | 0.019408 | 0.015055 | 0.015055 | 0.036332 | 0.022286 | 0.027325 | 0.008165 | 0.032324 | 0.040825 | 0.068605 | 0.022286 | 0.015452 | 0.19829 | 0.015492 | 0.031623
| Rel.Std.Dev. |  | 2.29% | 0.5266% | 0.2393% | 0.4736% | 0.3789% | 0.3702% | 0.8692% | 0.5385% | 0.6558% | 0.1975% | 0.7787% | 1.243% | 1.639% | 0.5569% | 0.386% | 4.79% | 0.368% | 0.7529%
| PDM3 |  | 5.984% | -1.236% | -0.01936% | -0.2627% | -3.305% | -1.033% | 1.725% | 0.7107% | 1.4% | 0.589% | 1.022% | -20.1% | 1.887% | -2.615% | -2.576% | 0.7513% | 2.455% | 2.211%
| Z-Score (Absolute) |  | 2.57 | 0.91 | 0.32 | 0.44 | 1.91 | 0.81 | 0.52 | 0.03 | 0.36 | 0.03 | 0.18 | 10 | 0.6 | 1.57 | 1.55 | 0.05 | 0.87 | 0.75
| Analytical results for MnO in OREAS 556b (Certified Value 0.155 wt.%) | [20 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.14 | 0.16 | 0.16 | 0.155 | 0.15 | 0.149 | 0.155 | 0.16 | 0.155 | 0.16528 | 0.15 | 0.15494 | 0.1 | 0.15 | 0.049427 | 0.15014 | 0.154 | 0.153 | 0.155 | 1
| 1 | 2 | 0.14 | 0.16 | 0.16 | 0.159 | 0.15 | 0.149 | 0.155 | 0.15 | 0.155 | 0.16528 | 0.14 | 0.15817 | 0.1 | 0.16 | 0.049427 | 0.14687 | 0.182 | 0.152 | 0.155 | 35
| 1 | 3 | 0.13 | 0.16 | 0.16 | 0.157 | 0.15 | 0.147 | 0.155 | 0.15 | 0.155 | 0.16269 | 0.14 | 0.15602 | 0.1 | 0.15 | 0.044933 | 0.14851 | 0.184 | 0.151 | 0.155 | 16
| 1 | 4 | 0.13 | 0.17 | 0.16 | 0.16 | 0.16 | 0.153 | 0.155 | 0.17 | 0.155 | 0.16528 | 0.14 | 0.15817 | 0.1 | 0.16 | 0.044933 | 0.14769 | 0.154 | 0.154 | 0.155 | 0.15474
| 1 | 5 | 0.13 | 0.16 | 0.16 | 0.158 | 0.15 | 0.148 | 0.155 | 0.16 | 0.155 | 0.16399 | 0.15 | 0.15387 | 0.1 | 0.15 | 0.044933 | 0.14606 | 0.149 | 0.156 | 0.155 | 15
| 1 | 6 | 0.14 | 0.16 | 0.15 | 0.155 | 0.14 | 0.151 | 0.155 | 0.16 | 0.155 | 0.16528 | 0.15 | 0.15763 | 0.1 | 0.15 | 0.049427 | 0.15259 | 0.159 | 0.157 | 0.155
| Mean |  | 0.135 | 0.16167 | 0.15833 | 0.15733 | 0.15 | 0.1495 | 0.155 | 0.15833 | 0.155 | 0.16463 | 0.145 | 0.15647 | 0.1 | 0.15333 | 0.04718 | 0.14864 | 0.16367 | 0.15383 | 0.155
| Median |  | 0.135 | 0.16 | 0.16 | 0.1575 | 0.15 | 0.149 | 0.155 | 0.16 | 0.155 | 0.16528 | 0.145 | 0.15683 | 0.1 | 0.15 | 0.04718 | 0.1481 | 0.1565 | 0.1535 | 0.155
| Std Dev. |  | 0.0054772 | 0.0040825 | 0.0040825 | 0.0020656 | 0.0063246 | 0.0021679 | 0 | 0.0075277 | 0 | 0.0010829 | 0.0054772 | 0.0018138 | 1.5202e-17 | 0.005164 | 0.0024611 | 0.0023882 | 0.015319 | 0.0023166 | 0
| Rel.Std.Dev. |  | 4.057% | 2.525% | 2.578% | 1.313% | 4.216% | 1.45% | 0% | 4.754% | 0% | 0.6578% | 3.777% | 1.159% | 1.52e-14% | 3.368% | 5.216% | 1.607% | 9.36% | 1.506% | 0%
| PDM3 |  | -12.76% | 4.473% | 2.319% | 1.673% | -3.066% | -3.389% | 0.1649% | 2.319% | 0.1649% | 6.39% | -6.297% | 1.114% | -35.38% | -0.9121% | -69.51% | -3.944% | 5.766% | -0.589% | 0.1649%
| Z-Score (Absolute) |  | 3.23 | 1.09 | 0.55 | 0.39 | 0.8 | 0.88 | 0 | 0.55 | 0 | 1.57 | 1.61 | 0.25 | NA | 0.26 | 17.44 | 1.02 | 1.41 | 0.18 | 0
Results from laboratory 19 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Mo in OREAS 556b (Indicative Value < 50 ppm) | [20 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 08 | 12 | 25 | [15 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | 16.8 | [15 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | 18.1 | [15 empty] | 7
| 1 | 3 | < 100 | < 50 | < 50 | 17.1 | [15 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | 17.7 | [15 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | 17.6 | [15 empty] | 13
| 1 | 6 | < 100 | < 50 | < 50 | 18.7
| Mean |  |  |  |  | 17.6667
| Median |  |  |  |  | 17.65
| Std Dev. |  |  |  |  | 0.68313
| Rel.Std.Dev. |  |  |  |  | 3.867%
PDM3
| Z-Score (Absolute) |  | 4.6 | 0 | 0 | 1.35
| Analytical results for Na2O in OREAS 556b (Indicative Value 0.107 wt.%) | [20 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 17 | 19 | 21 | 22 | 25 | 28 | [9 empty] | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 3
3
| 1 | 1 | < 0.01 | 0.22 | 0.06 | < 0.01 | 0.18 | < 0.1 | 0.06 | 0.08 | 0.44 | 0.05 | [9 empty] | 1
| 1 | 2 | < 0.01 | 0.19 | 0.04 | < 0.01 | 0.18 | < 0.1 | 0.06 | 0.08 | 0.59 | 0.06 | [9 empty] | 8
| 1 | 3 | < 0.01 | 0.21 | 0.07 | < 0.01 | 0.16 | < 0.1 | 0.06 | 0.09 | 0.5 | 0.04 | [9 empty] | 16
| 1 | 4 | < 0.01 | 0.2 | 0.07 | < 0.01 | 0.17 | 2.9 | 0.06 | 0.09 | 0.51 | 0.05 | [9 empty] | 0.10694
| 1 | 5 | < 0.01 | 0.19 | 0.06 | < 0.01 | 0.18 | < 0.1 | 0.06 | 0.09 | 0.51 | 0.07 | [9 empty] | 14
| 1 | 6 | < 0.01 | 0.19 | 0.09 | < 0.01 | 0.18 | < 0.1 | 0.06 | 0.09 | 0.76 | 0.06
| Mean |  |  | 0.2 | 0.065 |  | 0.175 | 2.9 | 0.06 | 0.086667 | 0.55167 | 0.055
| Median |  |  | 0.195 | 0.065 |  | 0.18 | 2.9 | 0.06 | 0.09 | 0.51 | 0.055
| Std Dev. |  |  | 0.012649 | 0.016432 |  | 0.0083666 |  | 0 | 0.005164 | 0.11268 | 0.010488
| Rel.Std.Dev. |  |  | 6.325% | 25.28% |  | 4.781% |  | 0% | 5.958% | 20.43% | 19.07%
| PDM3 |  |  | 87.01% | -39.22% |  | 63.64% | 2612% | -43.9% | -18.96% | 415.8% | -48.57%
| Z-Score (Absolute) |  | 0.67 | 1.18 | 0.1 | 0.67 | 0.94 | 4.28 | 0.15 | 0.1 | 4.53 | 0.2
| Analytical results for Nb in OREAS 556b (Indicative Value 56 ppm) | [20 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 08 | 12 | 21 | 25 | [15 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*OES | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 60 | 60 | 60 | 21.8 | [15 empty] | 1
| 1 | 2 | 60 | 70 | 59 | 22.2 | [15 empty] | 9
| 1 | 3 | 60 | 90 | 58 | 22.3 | [15 empty] | 16
| 1 | 4 | 60 | 100 | 61 | 22.8 | [15 empty] | 55.8883
| 1 | 5 | 70 | 80 | 60 | 22 | [15 empty] | 15
| 1 | 6 | 70 | 70 | 60 | 26.1
| Mean |  | 63.3333 | 78.3333 | 59.6667 | 22.8667
| Median |  | 60 | 75 | 60 | 22.25
| Std Dev. |  | 5.164 | 14.7196 | 1.0328 | 1.6195
| Rel.Std.Dev. |  | 8.154% | 18.79% | 1.731% | 7.082%
| PDM3 |  | 13.32% | 40.16% | 6.761% | -59.09%
| Z-Score (Absolute) |  | 0.13 | 1.22 | 0.13 | 2.79
| Analytical results for Nd in OREAS 556b (Indicative Value 60 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 0
0
| 1 | 1 | 61.021 | 59.5 | [17 empty] | 1
| 1 | 2 | 61.501 | 59.1 | [17 empty] | 10
| 1 | 3 | 61.254 | 56.3 | [17 empty] | 16
| 1 | 4 | 61.167 | 57.9 | [17 empty] | 59.611
| 1 | 5 | 61.341 | 56.9 | [17 empty] | 16
| 1 | 6 | 61.448 | 57.9
| Mean |  | 61.2887 | 57.9333
| Median |  | 61.2975 | 57.9
| Std Dev. |  | 0.17937 | 1.2291
| Rel.Std.Dev. |  | 0.2927% | 2.122%
| PDM3 |  | 2.814% | -2.814%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ni in OREAS 556b (Certified Value 344 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 08 | 12 | 17 | 22 | 24 | 25 | 27 | 28 | [7 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 200 | 300 | 283 | 240 | 400 | 360 | 240 | 310 | 333.9882 | 338 | 380 | 200 | [7 empty] | 1
| 1 | 2 | 300 | 300 | 283 | 240 | 400 | 330 | 310 | 310 | 333.9882 | 374 | 400 | 300 | [7 empty] | 40
| 1 | 3 | 200 | 300 | 346 | 240 | 400 | 350 | 210 | 310 | 333.9882 | 349 | 390 | 200 | [7 empty] | 16
| 1 | 4 | 200 | 300 | 361 | 310 | 380 | 350 | 230 | 310 | 340.668 | 331 | 390 | 200 | [7 empty] | 343.9656
| 1 | 5 | 300 | 300 | 369 | 240 | 380 | 360 | 410 | 310 | 333.9882 | 346 | 400 | 200 | [7 empty] | 16
| 1 | 6 | 300 | 300 | 338 | 240 | 390 | 360 | 350 | 310 | 347.3477 | 324 | 400 | 200
| Mean |  | 250 | 300 | 330 | 251.6667 | 391.6667 | 351.6667 | 291.6667 | 310 | 337.3281 | 343.6667 | 393.3333 | 216.6667
| Median |  | 250 | 300 | 342 | 240 | 395 | 355 | 275 | 310 | 333.9882 | 342 | 395 | 200
| Std Dev. |  | 54.7723 | 0 | 38 | 28.5774 | 9.8319 | 11.6905 | 78.5918 | 0 | 5.5887 | 17.5119 | 8.165 | 40.8248
| Rel.Std.Dev. |  | 21.91% | 0% | 11.52% | 11.36% | 2.51% | 3.324% | 26.95% | 0% | 1.657% | 5.096% | 2.076% | 18.84%
| PDM3 |  | -27.32% | -12.78% | -4.06% | -26.83% | 13.87% | 2.239% | -15.2% | -9.875% | -1.93% | -0.08691% | 14.35% | -37.01%
| Z-Score (Absolute) |  | NA | NA | 0.21 | 2.6 | 1.59 | 0.42 | 1.34 | 0.67 | 0 | 0.19 | 1.64 | NA
Results from laboratories 1, 3 and 28 were removed due to their 100 ppm reading resolution.
| Analytical results for P2O5 in OREAS 556b (Certified Value 0.08 wt.%) | [20 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.06 | 0.08 | 0.086 | 0.09 | 0.08 | 0.075 | 0.08 | 0.09 | 0.08 | 0.061 | 0.073375 | < 0.1 | 0.1 | 0.08 | 0.082 | 0.1 | 0.082 | 0.069 |  | 1
| 1 | 2 | 0.07 | 0.07 | 0.08 | 0.09 | 0.08 | 0.075 | 0.08 | 0.09 | 0.08 | 0.062 | 0.07441 | < 0.1 | 0.11 | 0.08 | 0.088 | 0.1 | 0.081 | 0.092 |  | 41
| 1 | 3 | 0.07 | 0.08 | 0.081 | 0.09 | 0.07 | 0.075 | 0.08 | 0.09 | 0.09 | 0.062 | 0.073375 | < 0.1 | 0.1 | 0.08 | 0.083 | 0.12 | 0.082 | 0.092 |  | 16
| 1 | 4 | 0.07 | 0.08 | 0.079 | 0.09 | 0.08 | 0.075 | 0.08 | 0.09 | 0.08 | 0.062 | 0.076481 | < 0.1 | 0.1 | 0.08 | 0.085 | 0.09 | 0.081 | 0.069 |  | 0.079854
| 1 | 5 | 0.07 | 0.08 | 0.082 | 0.09 | 0.08 | 0.077 | 0.08 | 0.09 | 0.08 | 0.064 | 0.07441 | < 0.1 | 0.1 | 0.08 | 0.082 | 0.13 | 0.082 | 0.069 |  | 17
| 1 | 6 | 0.06 | 0.08 | 0.079 | 0.09 | 0.08 | 0.075 | 0.08 | 0.09 | 0.08 | 0.063 | 0.073375 | < 0.1 | 0.1 | 0.08 | 0.084 | 0.1 | 0.083 | 0.069
| Mean |  | 0.066667 | 0.078333 | 0.081167 | 0.09 | 0.078333 | 0.075333 | 0.08 | 0.09 | 0.081667 | 0.062333 | 0.074238 |  | 0.10167 | 0.08 | 0.084 | 0.10667 | 0.081833 | 0.076667
| Median |  | 0.07 | 0.08 | 0.0805 | 0.09 | 0.08 | 0.075 | 0.08 | 0.09 | 0.08 | 0.062 | 0.073893 |  | 0.1 | 0.08 | 0.0835 | 0.1 | 0.082 | 0.069
| Std Dev. |  | 0.005164 | 0.0040825 | 0.0026394 | 1.5202e-17 | 0.0040825 | 0.0008165 | 0 | 1.5202e-17 | 0.0040825 | 0.0010328 | 0.0012101 |  | 0.0040825 | 0 | 0.0022804 | 0.015055 | 0.00075277 | 0.011877
| Rel.Std.Dev. |  | 7.746% | 5.212% | 3.252% | 1.689e-14% | 5.212% | 1.084% | 0% | 1.689e-14% | 4.999% | 1.657% | 1.63% |  | 4.016% | 0% | 2.715% | 14.11% | 0.9199% | 15.49%
| PDM3 |  | -16.51% | -1.904% | 1.644% | 12.71% | -1.904% | -5.661% | 0.1827% | 12.71% | 2.27% | -21.94% | -7.033% |  | 27.32% | 0.1827% | 5.192% | 33.58% | 2.479% | -3.992%
| Z-Score (Absolute) |  | 2.25 | 0.28 | 0.2 | 1.69 | 0.28 | 0.79 | 0 | 1.69 | 0.28 | 2.98 | 0.97 | 5.06 | NA | 0 | 0.67 | 4.5 | 0.31 | 0.61
Results from laboratory 21 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 556b (Indicative Value 81 ppm) | [20 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 12 | 13 | 17 | 22 | 25 | 28 | [8 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 90 | < 35 | 90 | 60 | < 100 | < 50 | 370 | 12.6 | < 100 | [8 empty] | 1
| 1 | 2 | < 100 | < 100 | 190 | < 35 | 90 | 60 | < 100 | < 50 | 370 | 16.4 | < 100 | [8 empty] | 11
| 1 | 3 | 200 | < 100 | 190 | < 35 | 90 | 70 | < 100 | < 50 | 370 | 18.1 | < 100 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | 190 | < 35 | 70 | 70 | < 100 | < 50 | 370 | 12.2 | < 100 | [8 empty] | 80.9596
| 1 | 5 | < 100 | < 100 | 190 | < 35 | 90 | 60 | < 100 | < 50 | 370 | 27.4 | < 100 | [8 empty] | 17
| 1 | 6 | < 100 | < 100 | 190 | < 35 | 90 | 60 | < 100 | < 50 | 370 | 40.5 | < 100
| Mean |  | 200 |  | 190 |  | 86.6667 | 63.3333 |  |  | 370 | 21.2
| Median |  | 200 |  | 190 |  | 90 | 60 |  |  | 370 | 17.25
| Std Dev. |  |  |  | 0 |  | 8.165 | 5.164 |  |  | 0 | 10.9411
| Rel.Std.Dev. |  |  |  | 0% |  | 9.421% | 8.154% |  |  | 0% | 51.61%
| PDM3 |  | 147% |  | 134.7% |  | 7.049% | -21.77% |  |  | 357% | -73.81%
| Z-Score (Absolute) |  | 0.67 | 0 | 3.03 | 0.88 | 0.99 | 0.36 | 0 | 0.67 | 8.67 | 0.78 | 0
| Analytical results for Pr in OREAS 556b (Indicative Value 16.5 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 1
1
| 1 | 1 | 16.495 | 16.6 | [17 empty] | 1
| 1 | 2 | 16.692 | 16.8 | [17 empty] | 12
| 1 | 3 | 16.527 | 16.2 | [17 empty] | 16
| 1 | 4 | 16.513 | 16.4 | [17 empty] | 16.4624
| 1 | 5 | 16.554 | 15.9 | [17 empty] | 18
| 1 | 6 | 16.568 | 16.3
| Mean |  | 16.5582 | 16.3667
| Median |  | 16.5405 | 16.35
| Std Dev. |  | 0.070749 | 0.31411
| Rel.Std.Dev. |  | 0.4273% | 1.919%
| PDM3 |  | 0.5816% | -0.5816%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 556b (Indicative Value 99 ppm) | [20 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 08 | 12 | 25 | [15 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 100 | 110 | < 50 | 86.2 | [15 empty] | 1
| 1 | 2 | 100 | 110 | < 50 | 86.4 | [15 empty] | 13
| 1 | 3 | 100 | 110 | < 50 | 84.3 | [15 empty] | 16
| 1 | 4 | 100 | 110 | < 50 | 88.4 | [15 empty] | 98.7733
| 1 | 5 | 100 | 110 | < 50 | 86.3 | [15 empty] | 19
| 1 | 6 | 100 | 110 | < 50 | 104
| Mean |  | 100 | 110 |  | 89.2667
| Median |  | 100 | 110 |  | 86.35
| Std Dev. |  | 0 | 0 |  | 7.3337
| Rel.Std.Dev. |  | 0% | 0% |  | 8.215%
| PDM3 |  | 1.242% | 11.37% |  | -9.625%
| Z-Score (Absolute) |  | 0.35 | 1 | 4.53 | 0.35
| Analytical results for Re in OREAS 556b (Indicative Value 22.6 ppm) | [20 empty] | Indicative
Re | Round | Replicate | Lab | [18 empty] | 1
| No | No | 25 | [18 empty] | ppm
|  |  | BF*XRF | [18 empty] | 1
1
| 1 | 1 | 19.9 | [18 empty] | 1
| 1 | 2 | 20.9 | [18 empty] | 14
| 1 | 3 | 17.5 | [18 empty] | 16
| 1 | 4 | 21.4 | [18 empty] | 22.5833
| 1 | 5 | 25.2 | [18 empty] | 20
| 1 | 6 | 30.6
| Mean |  | 22.5833
| Median |  | 21.15
| Std Dev. |  | 4.6559
| Rel.Std.Dev. |  | 20.62%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 556b (Indicative Value < 50 ppm) | [20 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 08 | 12 | 24 | 25 | [14 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 42 | 7 | [14 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 42 | 5 | [14 empty] | 15
| 1 | 3 | < 100 | < 50 | < 50 | < 42 | 5 | [14 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 42 | 6 | [14 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 42 | 5 | [14 empty] | 21
| 1 | 6 | < 100 | < 50 | < 50 | < 42 | 6
| Mean | [5 empty] | 5.6667
| Median | [5 empty] | 5.5
| Std Dev. | [5 empty] | 0.8165
| Rel.Std.Dev. | [5 empty] | 14.41%
PDM3
| Z-Score (Absolute) |  | 4.1 | 0 | 0 | 0.67 | 3.17
| Analytical results for Sc in OREAS 556b (Indicative Value 11.9 ppm) | [20 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 13 | 25 | [16 empty] | ppm
|  |  | PPP*XRF | BF*MS | BF*XRF | [16 empty] | 1
1
| 1 | 1 | < 100 | 12.908 | 11 | [16 empty] | 1
| 1 | 2 | < 100 | 12.679 | 11 | [16 empty] | 1
| 1 | 3 | < 100 | 12.845 | 11 | [16 empty] | 16
| 1 | 4 | < 100 | 12.745 | 11 | [16 empty] | 11.8981
| 1 | 5 | < 100 | 12.733 | 11 | [16 empty] | 7
| 1 | 6 | < 100 | 12.867 | 11
| Mean |  |  | 12.7962 | 11
| Median |  |  | 12.795 | 11
| Std Dev. |  |  | 0.089716 | 0
| Rel.Std.Dev. |  |  | 0.7011% | 0%
| PDM3 |  |  | 7.548% | -7.548%
| Z-Score (Absolute) |  | 13.97 | 0 | 0.67
| Analytical results for Se in OREAS 556b (Indicative Value 2.33 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 2
2
| 1 | 1 | < 100 | 1 | [17 empty] | 1
| 1 | 2 | < 100 | 3 | [17 empty] | 19
| 1 | 3 | < 100 | 4 | [17 empty] | 16
| 1 | 4 | < 100 | 2 | [17 empty] | 2.3333
| 1 | 5 | < 100 | 2 | [17 empty] | 8
| 1 | 6 | < 100 | 2
| Mean |  |  | 2.3333
| Median |  |  | 2
| Std Dev. |  |  | 1.0328
| Rel.Std.Dev. |  |  | 44.26%
| PDM3 |  |  | 1.554e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for SiO2 in OREAS 556b (Certified Value 42.77 wt.%) | [20 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 27 | 28 | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 36.99 | 42.44 | 43.22 | 42.97 | 42.19 | 42.26 | 43.3 | 43 | 42.9 | 90.253 | 42.89 | 42.0785 | 33 | 42.87 | 41.62 | 42.1948 | 44.29 | 43.41 | 43.05 | 1
| 1 | 2 | 37.24 | 42.78 | 42.96 | 43.18 | 42.53 | 42.07 | 43.5 | 42.6 | 43 | 89.461 | 43.15 | 42.3664 | 33 | 43.9 | 41.54 | 42.4086 | 44.6 | 43.42 | 43.12 | 0
| 1 | 3 | 37 | 42.69 | 42.96 | 43.12 | 42.7 | 42.07 | 43.4 | 42.3 | 42.9 | 89.739 | 42.27 | 42.1048 | 33.1 | 43.88 | 41.65 | 42.4017 | 47.81 | 43.54 | 42.95 | 16
| 1 | 4 | 36.91 | 42.88 | 42.94 | 42.91 | 42.9 | 41.93 | 43.3 | 42.7 | 43.2 | 89.761 | 42.67 | 42.5896 | 31.6 | 44.62 | 41.13 | 42.27 | 42.53 | 43.21 | 43.06 | 42.7709
| 1 | 5 | 37.3 | 42.75 | 42.83 | 43.22 | 42.65 | 42.04 | 43.3 | 42.8 | 42.6 | 89.59 | 42.72 | 42.3916 | 33.1 | 44.75 | 41.21 | 42.2304 | 40.81 | 43.41 | 43.05 | 18
| 1 | 6 | 37.08 | 42.05 | 42.8 | 42.97 | 42.48 | 42.2 | 43.5 | 43.2 | 42.8 | 89.568 | 42.61 | 41.8978 | 32.9 | 43.91 | 41.26 | 42.0087 | 42.68 | 43.37 | 42.94
| Mean |  | 37.0867 | 42.5983 | 42.9517 | 43.0617 | 42.575 | 42.095 | 43.3833 | 42.7667 | 42.9 | 89.7287 | 42.7183 | 42.2381 | 32.7833 | 43.9883 | 41.4017 | 42.2524 | 43.7867 | 43.3933 | 43.0283
| Median |  | 37.04 | 42.72 | 42.95 | 43.045 | 42.59 | 42.07 | 43.35 | 42.75 | 42.9 | 89.6645 | 42.695 | 42.2356 | 33 | 43.905 | 41.4 | 42.2502 | 43.485 | 43.41 | 43.05
| Std Dev. |  | 0.15306 | 0.30643 | 0.14838 | 0.12828 | 0.23923 | 0.11811 | 0.098319 | 0.31411 | 0.2 | 0.28027 | 0.29356 | 0.25403 | 0.58452 | 0.67164 | 0.22763 | 0.14847 | 2.3986 | 0.10671 | 0.06969
| Rel.Std.Dev. |  | 0.4127% | 0.7193% | 0.3455% | 0.2979% | 0.5619% | 0.2806% | 0.2266% | 0.7345% | 0.4662% | 0.3124% | 0.6872% | 0.6014% | 1.783% | 1.527% | 0.5498% | 0.3514% | 5.478% | 0.2459% | 0.162%
| PDM3 |  | -13.29% | -0.4034% | 0.4227% | 0.6799% | -0.458% | -1.58% | 1.432% | -0.009864% | 0.3019% | 109.8% | -0.1229% | -1.246% | -23.35% | 2.846% | -3.201% | -1.212% | 2.375% | 1.455% | 0.6019%
| Z-Score (Absolute) |  | 7.25 | 0.21 | 0.24 | 0.38 | 0.24 | 0.86 | 0.79 | 0 | 0.17 | 59.91 | 0.06 | 0.67 | 12.74 | 1.56 | 1.74 | 0.66 | 1.3 | 0.8 | 0.33
| Analytical results for Sm in OREAS 556b (Indicative Value 8.79 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 8.675 | 9.1 | [17 empty] | 1
| 1 | 2 | 8.767 | 9 | [17 empty] | 3
| 1 | 3 | 8.745 | 8.5 | [17 empty] | 16
| 1 | 4 | 8.749 | 8.6 | [17 empty] | 8.7876
| 1 | 5 | 8.559 | 9 | [17 empty] | 9
| 1 | 6 | 8.756 | 9
| Mean |  | 8.7085 | 8.8667
| Median |  | 8.747 | 9
| Std Dev. |  | 0.080154 | 0.25033
| Rel.Std.Dev. |  | 0.9204% | 2.823%
| PDM3 |  | -0.8999% | 0.8999%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 556b (Indicative Value 106 ppm) | [20 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 08 | 12 | 13 | 17 | 25 | 28 | [12 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [12 empty] | 0
0
| 1 | 1 | < 100 | 60 | 100 | < 100 | 400 | 3.6 | 200 | [12 empty] | 1
| 1 | 2 | < 100 | 170 | 140 | < 100 | 390 | 4.1 | 200 | [12 empty] | 20
| 1 | 3 | < 100 | 180 | 130 | < 100 | 350 | 3.3 | 100 | [12 empty] | 16
| 1 | 4 | < 100 | 110 | 130 | < 100 | 410 | 3.5 | 200 | [12 empty] | 105.9333
| 1 | 5 | < 100 | 80 | 110 | < 100 | 400 | 3.7 | 200 | [12 empty] | 10
| 1 | 6 | < 100 | 90 | 120 | < 100 | 400 | 4.2 | 200
| Mean |  |  | 115 | 121.6667 |  | 391.6667 | 3.7333 | 183.3333
| Median |  |  | 100 | 125 |  | 400 | 3.65 | 200
| Std Dev. |  |  | 49.295 | 14.7196 |  | 21.3698 | 0.35024 | 40.8248
| Rel.Std.Dev. |  |  | 42.87% | 12.1% |  | 5.456% | 9.381% | 22.27%
| PDM3 |  |  | 8.559% | 14.85% |  | 269.7% | -96.48% | 73.06%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.07 | 0.67 | 2.87 | 1.15 | 0.71
| Analytical results for SO3 in OREAS 556b (Certified Value 21.43 wt.%) | [20 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 08 | 10 | 12 | 17 | 18 | 19 | 22 | 24 | 28 | [6 empty] | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 2
3
| 1 | 1 | 11.85 | 0.09 | 22.82 | 21.87 | 21.875 | 22.4 | 22.225 | > 5 | 21.113 | 17.7 | 7.08 | 21.6977 | > 12.5 | [6 empty] | 1
| 1 | 2 | 12.03 | 0.11 | 22.34 | 21.67 | 22.05 | 22 | 22.55 | > 5 | 21.315 | 17.5 | 7.31 | 21.6607 | > 12.5 | [6 empty] | 0
| 1 | 3 | 11.83 | 0.1 | 22.77 | 21.71 | 21.925 | 22 | 22.325 | > 5 | 21.152 | 17.6 | 7.16 | 21.8717 | > 12.5 | [6 empty] | 16
| 1 | 4 | 11.78 | 0.1 | 22.4 | 21.8 | 22.2 | 22.3 | 22.55 | > 5 | 21.247 | 16.8 | 7.53 | 21.7247 | > 12.5 | [6 empty] | 21.4269
| 1 | 5 | 11.87 | 0.11 | 22.35 | 21.68 | 22.45 | 22.3 | 22.075 | > 5 | 21.161 | 17.5 | 7.53 | 21.3856 | > 12.5 | [6 empty] | 19
| 1 | 6 | 11.95 | 0.1 | 22.31 | 21.84 | 22.724 | 22.6 | 22.225 | > 5 | 21.034 | 17.4 | 7.16 | 21.5556 | > 12.5
| Mean |  | 11.885 | 0.10167 | 22.4983 | 21.7617 | 22.204 | 22.2667 | 22.325 |  | 21.1703 | 17.4167 | 7.295 | 21.6493
| Median |  | 11.86 | 0.1 | 22.375 | 21.755 | 22.125 | 22.3 | 22.275 |  | 21.1565 | 17.5 | 7.235 | 21.6792
| Std Dev. |  | 0.090277 | 0.0075277 | 0.23216 | 0.08612 | 0.32886 | 0.23381 | 0.1917 |  | 0.099099 | 0.31885 | 0.19665 | 0.16486
| Rel.Std.Dev. |  | 0.7596% | 7.404% | 1.032% | 0.3957% | 1.481% | 1.05% | 0.8587% |  | 0.4681% | 1.831% | 2.696% | 0.7615%
| PDM3 |  | -44.53% | -99.53% | 5% | 1.562% | 3.627% | 3.919% | 4.191% |  | -1.198% | -18.72% | -65.95% | 1.038%
| Z-Score (Absolute) |  | 9.75 | 21.51 | 0.85 | 0.11 | 0.55 | 0.62 | 0.67 | NA | 0.48 | 4.22 | 14.33 | 0 | NA
| Analytical results for Sr in OREAS 556b (Certified Value 63 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 06 | 07 | 08 | 10 | 12 | 21 | 22 | 25 | 28 | [9 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 80 | < 80 | 67 | < 100 | < 80 | 100 | 77 | < 3400 | 61 | < 100 | [9 empty] | 1
| 1 | 2 | < 80 | 80 | 66 | < 100 | < 80 | 100 | 77 | < 3400 | 64 | < 100 | [9 empty] | 32
| 1 | 3 | < 80 | < 80 | 65 | < 100 | < 80 | 100 | 73 | < 3400 | 66 | < 100 | [9 empty] | 16
| 1 | 4 | < 80 | < 80 | 67 | 100 | 80 | 100 | 77 | < 3400 | 60 | < 100 | [9 empty] | 63.2432
| 1 | 5 | < 80 | 80 | 65 | < 100 | 80 | 100 | 75 | < 3400 | 55 | < 100 | [9 empty] | 20
| 1 | 6 | < 80 | < 80 | 64 | < 100 | 80 | 100 | 71 | < 3400 | 58 | < 100
| Mean |  |  | 80 | 65.6667 | 100 | 80 | 100 | 75 |  | 60.6667
| Median |  |  | 80 | 65.5 | 100 | 80 | 100 | 76 |  | 60.5
| Std Dev. |  |  | 0 | 1.2111 |  | 0 | 0 | 2.5298 |  | 3.9833
| Rel.Std.Dev. |  |  | 0% | 1.844% |  | 0% | 0% | 3.373% |  | 6.566%
| PDM3 |  |  | 26.5% | 3.832% | 58.12% | 26.5% | 58.12% | 18.59% |  | -4.074%
| Z-Score (Absolute) |  | 1.5 | 0.43 | 0.28 | 0.28 | 0.1 | 2.89 | 0.99 | 124.04 | 0.1 | 0.92
| Analytical results for Ta in OREAS 556b (Indicative Value 1.25 ppm) | [20 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 2
2
| 1 | 1 | < 100 | 1.1 | [17 empty] | 1
| 1 | 2 | < 100 | 1.2 | [17 empty] | 5
| 1 | 3 | < 100 | 1.2 | [17 empty] | 16
| 1 | 4 | < 100 | 1.2 | [17 empty] | 1.25
| 1 | 5 | < 100 | 1.2 | [17 empty] | 11
| 1 | 6 | < 100 | 1.6
| Mean |  |  | 1.25
| Median |  |  | 1.2
| Std Dev. |  |  | 0.17607
| Rel.Std.Dev. |  |  | 14.09%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Tb in OREAS 556b (Indicative Value 1.02 ppm) | [20 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 1.038 | 1 | [17 empty] | 1
| 1 | 2 | 1.049 | 1.02 | [17 empty] | 6
| 1 | 3 | 1.038 | 0.98 | [17 empty] | 16
| 1 | 4 | 1.044 | 0.99 | [17 empty] | 1.0185
| 1 | 5 | 1.048 | 0.97 | [17 empty] | 12
| 1 | 6 | 1.045 | 1
| Mean |  | 1.0437 | 0.99333
| Median |  | 1.0445 | 0.995
| Std Dev. |  | 0.004761 | 0.017512
| Rel.Std.Dev. |  | 0.4562% | 1.763%
| PDM3 |  | 2.471% | -2.471%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 556b (Indicative Value 11.1 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 1
1
| 1 | 1 | < 100 | 8.5 | [17 empty] | 1
| 1 | 2 | < 100 | 10.2 | [17 empty] | 7
| 1 | 3 | < 100 | 8.9 | [17 empty] | 16
| 1 | 4 | < 100 | 15 | [17 empty] | 11.1333
| 1 | 5 | < 100 | 2.4 | [17 empty] | 13
| 1 | 6 | < 100 | 21.8
| Mean |  |  | 11.1333
| Median |  |  | 9.55
| Std Dev. |  |  | 6.5999
| Rel.Std.Dev. |  |  | 59.28%
| PDM3 |  |  | 2.887e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Th in OREAS 556b (Indicative Value 15.6 ppm) | [20 empty] | Indicative
Th | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 1
1
| 1 | 1 | 13.99 | 16.1 | [17 empty] | 1
| 1 | 2 | 14.28 | 16.7 | [17 empty] | 8
| 1 | 3 | 14.286 | 16.6 | [17 empty] | 16
| 1 | 4 | 14.284 | 16.7 | [17 empty] | 15.6418
| 1 | 5 | 14.283 | 16.5 | [17 empty] | 14
| 1 | 6 | 14.278 | 19.7
| Mean |  | 14.2335 | 17.05
| Median |  | 14.2815 | 16.65
| Std Dev. |  | 0.11932 | 1.3172
| Rel.Std.Dev. |  | 0.8383% | 7.725%
| PDM3 |  | -9.003% | 9.003%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for TiO2 in OREAS 556b (Certified Value 0.707 wt.%) | [20 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 17 | 18 | 19 | 21 | 22 | 25 | 27 | 28 |  |  | wt.%
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.48 | 0.72 | 0.7 | 0.71 | 0.72 | 0.66 | 0.71 | 0.73 | 0.72 | 0.7 | 0.65093 | 0.6 | 0.69 | 0.7 | 0.649 | 0.75 | 0.79 |  |  | 1
| 1 | 2 | 0.45 | 0.73 | 0.69 | 0.71 | 0.73 | 0.66 | 0.71 | 0.72 | 0.71 | 0.69 | 0.65374 | 0.6 | 0.7 | 0.7 | 0.674 | 0.74 | 0.75 |  |  | 36
| 1 | 3 | 0.55 | 0.72 | 0.69 | 0.71 | 0.72 | 0.66 | 0.71 | 0.73 | 0.72 | 0.71 | 0.64906 | 0.6 | 0.7 | 0.7 | 0.722 | 0.75 | 0.73 |  |  | 16
| 1 | 4 | 0.56 | 0.73 | 0.69 | 0.71 | 0.73 | 0.69 | 0.72 | 0.73 | 0.72 | 0.69 | 0.65936 | 0.6 | 0.7 | 0.7 | 0.666 | 0.75 | 0.76 |  |  | 0.70703
| 1 | 5 | 0.46 | 0.73 | 0.69 | 0.71 | 0.74 | 0.65 | 0.73 | 0.72 | 0.72 | 0.72 | 0.65561 | 0.6 | 0.7 | 0.7 | 0.63 | 0.74 | 0.78 |  |  | 21
| 1 | 6 | 0.48 | 0.72 | 0.69 | 0.71 | 0.73 | 0.65 | 0.73 | 0.74 | 0.72 | 0.7 | 0.65561 | 0.6 | 0.7 | 0.69 | 0.651 | 0.74 | 0.79
| Mean |  | 0.49667 | 0.725 | 0.69167 | 0.71 | 0.72833 | 0.66167 | 0.71833 | 0.72833 | 0.71833 | 0.70167 | 0.65405 | 0.6 | 0.69833 | 0.69833 | 0.66533 | 0.745 | 0.76667
| Median |  | 0.48 | 0.725 | 0.69 | 0.71 | 0.73 | 0.66 | 0.715 | 0.73 | 0.72 | 0.7 | 0.65468 | 0.6 | 0.7 | 0.7 | 0.6585 | 0.745 | 0.77
| Std Dev. |  | 0.046762 | 0.0054772 | 0.0040825 | 0 | 0.0075277 | 0.01472 | 0.0098319 | 0.0075277 | 0.0040825 | 0.01169 | 0.0036813 | 0 | 0.0040825 | 0.0040825 | 0.031646 | 0.0054772 | 0.024221
| Rel.Std.Dev. |  | 9.415% | 0.7555% | 0.5902% | 0% | 1.034% | 2.225% | 1.369% | 1.034% | 0.5683% | 1.666% | 0.5628% | 0% | 0.5846% | 0.5846% | 4.756% | 0.7352% | 3.159%
| PDM3 |  | -29.75% | 2.542% | -2.172% | 0.4207% | 3.014% | -6.415% | 1.599% | 3.014% | 1.599% | -0.758% | -7.493% | -15.14% | -1.229% | -1.229% | -5.897% | 5.371% | 8.435%
| Z-Score (Absolute) |  | 6.77 | 0.62 | 0.46 | 0.13 | 0.73 | 1.43 | 0.4 | 0.73 | 0.4 | 0.13 | 1.68 | NA | 0.24 | 0.24 | 1.31 | 1.27 | 1.97
Results from laboratory 19 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Tl in OREAS 556b (Indicative Value 0.85 ppm) | [20 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 01 | 25 | [17 empty] | ppm
|  |  | PPP*XRF | BF*XRF | [17 empty] | 2
2
| 1 | 1 | < 100 | 0.82 | [17 empty] | 1
| 1 | 2 | < 100 | 0.86 | [17 empty] | 26
| 1 | 3 | < 100 | 0.78 | [17 empty] | 16
| 1 | 4 | < 100 | 0.8 | [17 empty] | 0.84667
| 1 | 5 | < 100 | 0.83 | [17 empty] | 15
| 1 | 6 | < 100 | 0.99
| Mean |  |  | 0.84667
| Median |  |  | 0.825
| Std Dev. |  |  | 0.075277
| Rel.Std.Dev. |  |  | 8.891%
| PDM3 |  |  | -4.441e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Tm in OREAS 556b (Indicative Value 0.52 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 0.487 | 0.55 | [17 empty] | 1
| 1 | 2 | 0.493 | 0.55 | [17 empty] | 10
| 1 | 3 | 0.492 | 0.55 | [17 empty] | 16
| 1 | 4 | 0.488 | 0.57 | [17 empty] | 0.51608
| 1 | 5 | 0.496 | 0.49 | [17 empty] | 16
| 1 | 6 | 0.497 | 0.53
| Mean |  | 0.49217 | 0.54
| Median |  | 0.4925 | 0.55
| Std Dev. |  | 0.0040702 | 0.027568
| Rel.Std.Dev. |  | 0.827% | 5.105%
| PDM3 |  | -4.634% | 4.634%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 556b (Indicative Value 8.95 ppm) | [20 empty] | Indicative
U | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 7.88 | 9.36 | [17 empty] | 1
| 1 | 2 | 8.07 | 9.8 | [17 empty] | 11
| 1 | 3 | 7.968 | 9.56 | [17 empty] | 16
| 1 | 4 | 7.995 | 9.75 | [17 empty] | 8.9455
| 1 | 5 | 7.997 | 9.47 | [17 empty] | 17
| 1 | 6 | 7.996 | 11.5
| Mean |  | 7.9843 | 9.9067
| Median |  | 7.9955 | 9.655
| Std Dev. |  | 0.061484 | 0.79804
| Rel.Std.Dev. |  | 0.7701% | 8.056%
| PDM3 |  | -10.74% | 10.74%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V2O5 in OREAS 556b (Certified Value 641 ppm) | [20 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 12 | 17 | 19 | 22 | 25 | 28 | [7 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 710 | 700 | 680 | 600 | 560.6 | 540 | 710 | 500 | < 1000 | 400 | 651.6 | 661 | [7 empty] | 1
| 1 | 2 | < 180 | 700 | 700 | 600 | 557 | 540 | 710 | 500 | < 1000 | 400 | 706.9 | 678 | [7 empty] | 0
| 1 | 3 | < 180 | 700 | 690 | 600 | 548.1 | 710 | 710 | 470 | < 1000 | 400 | 774.8 | 661 | [7 empty] | 16
| 1 | 4 | 710 | 700 | 690 | 600 | 580.2 | 710 | 710 | 500 | < 1000 | 400 | 680.2 | 625 | [7 empty] | 640.9222
| 1 | 5 | < 180 | 700 | 700 | 600 | 530.2 | 710 | 710 | 510 | < 1000 | 400 | 667.7 | 643 | [7 empty] | 22
| 1 | 6 | < 180 | 700 | 680 | 600 | 546.3 | 540 | 710 | 490 | < 1000 | 400 | 689.1 | 643
| Mean |  | 710 | 700 | 690 | 600 | 553.7333 | 625 | 710 | 495 |  | 400 | 695.05 | 651.8333
| Median |  | 710 | 700 | 690 | 600 | 552.55 | 625 | 710 | 500 |  | 400 | 684.65 | 652
| Std Dev. |  | 0 | 0 | 8.9443 | 0 | 16.7279 | 93.1128 | 0 | 13.784 |  | 0 | 43.3413 | 18.5948
| Rel.Std.Dev. |  | 0% | 0% | 1.296% | 0% | 3.021% | 14.9% | 0% | 2.785% |  | 0% | 6.236% | 2.853%
| PDM3 |  | 10.78% | 9.218% | 7.657% | -6.385% | -13.6% | -2.484% | 10.78% | -22.77% |  | -37.59% | 8.445% | 1.702%
| Z-Score (Absolute) |  | 3.1 | NA | 0.62 | NA | 0.67 | 0 | 0.85 | 1.23 | 1.18 | NA | 0.67 | 0.25
Results from laboratories 3, 6 and 22 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 556b (Indicative Value 6.26 ppm) | [20 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 08 | 12 | 25 | [15 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 100 | 10 | < 10 | 2.4 | [15 empty] | 1
| 1 | 2 | < 100 | 10 | < 10 | 2.4 | [15 empty] | 12
| 1 | 3 | < 100 | 10 | < 10 | 6.8 | [15 empty] | 16
| 1 | 4 | < 100 | 10 | < 10 | 2.4 | [15 empty] | 6.26
| 1 | 5 | < 100 | 10 | < 10 | 2.4 | [15 empty] | 18
| 1 | 6 | < 100 | 10 | < 10 | 3
| Mean |  |  | 10 |  | 3.2333
| Median |  |  | 10 |  | 2.4
| Std Dev. |  |  | 0 |  | 1.7637
| Rel.Std.Dev. |  |  | 0% |  | 54.55%
| PDM3 |  |  | 59.74% |  | -48.35%
| Z-Score (Absolute) |  | 8.47 | 0.5 | 0.5 | 0.85
| Analytical results for Y in OREAS 556b (Indicative Value 33.1 ppm) | [20 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 08 | 12 | 13 | 21 | 25 | [14 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*MS | BF*OES | BF*XRF | [14 empty] | 1
1
| 1 | 1 | 39 | 39 | 30.486 | 35 | 33.2 | [14 empty] | 1
| 1 | 2 | 39 | < 39 | 30.621 | 33 | 34.1 | [14 empty] | 13
| 1 | 3 | 47 | < 39 | 30.256 | 33 | 34.1 | [14 empty] | 16
| 1 | 4 | 47 | < 39 | 30.315 | 33 | 34.1 | [14 empty] | 33.0882
| 1 | 5 | 47 | < 39 | 30.327 | 34 | 32.9 | [14 empty] | 19
| 1 | 6 | 47 | < 39 | 30.338 | 33 | 34.3
| Mean |  | 44.3333 | 39 | 30.3905 | 33.5 | 33.7833
| Median |  | 47 | 39 | 30.3325 | 33 | 34.1
| Std Dev. |  | 4.1312 |  | 0.13626 | 0.83666 | 0.58109
| Rel.Std.Dev. |  | 9.318% |  | 0.4484% | 2.497% | 1.72%
| PDM3 |  | 33.99% | 17.87% | -8.153% | 1.244% | 2.101%
| Z-Score (Absolute) |  | 2.41 | 2.28 | 0.67 | 0 | 0.06
| Analytical results for Yb in OREAS 556b (Indicative Value 3.3 ppm) | [20 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | [17 empty] | 1
| No | No | 13 | 25 | [17 empty] | ppm
|  |  | BF*MS | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 3.301 | 3.3 | [17 empty] | 1
| 1 | 2 | 3.307 | 3.4 | [17 empty] | 14
| 1 | 3 | 3.312 | 3.4 | [17 empty] | 16
| 1 | 4 | 3.311 | 3.1 | [17 empty] | 3.295
| 1 | 5 | 3.305 | 3.2 | [17 empty] | 20
| 1 | 6 | 3.304 | 3.3
| Mean |  | 3.3067 | 3.2833
| Median |  | 3.306 | 3.3
| Std Dev. |  | 0.0042269 | 0.1169
| Rel.Std.Dev. |  | 0.1278% | 3.561%
| PDM3 |  | 0.3541% | -0.3541%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 556b (Indicative Value 44.4 ppm) | [20 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 12 | 13 | 17 | 21 | 22 | 24 | 25 | 27 | 28 | [5 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | 2700 | < 80 | 67 | < 50 | 50 | < 100 | < 50 | 39 | 1210 | < 40 | 30 | 30 | 300 | [5 empty] | 1
| 1 | 2 | < 100 | 2700 | < 80 | 72 | < 50 | < 50 | < 100 | < 50 | 39 | 1210 | < 40 | 27 | 40 | 300 | [5 empty] | 15
| 1 | 3 | < 100 | 2800 | < 80 | 72 | < 50 | < 50 | < 100 | < 50 | 40 | 1210 | < 40 | 25 | 30 | 300 | [5 empty] | 16
| 1 | 4 | < 100 | 2800 | < 80 | 78 | < 50 | < 50 | < 100 | < 50 | 38 | 1120 | < 40 | 26 | 40 | 300 | [5 empty] | 44.3667
| 1 | 5 | < 100 | 2700 | < 80 | 68 | < 50 | < 50 | < 100 | < 50 | 39 | 1210 | < 40 | 28 | 40 | 300 | [5 empty] | 21
| 1 | 6 | < 100 | 2700 | < 80 | 68 | < 50 | < 50 | < 100 | < 50 | 38 | 1210 | < 40 | 27 | 30 | 300
| Mean |  |  | 2733.3333 |  | 70.8333 |  | 50 |  |  | 38.8333 | 1195 |  | 27.1667 | 35 | 300
| Median |  |  | 2700 |  | 70 |  | 50 |  |  | 39 | 1210 |  | 27 | 35 | 300
| Std Dev. |  |  | 51.6398 |  | 4.1191 |  |  |  |  | 0.75277 | 36.7423 |  | 1.7224 | 5.4772 | 0
| Rel.Std.Dev. |  |  | 1.889% |  | 5.815% |  |  |  |  | 1.938% | 3.075% |  | 6.34% | 15.65% | 0%
| PDM3 |  |  | 6061% |  | 59.65% |  | 12.7% |  |  | -12.47% | 2593% |  | -38.77% | -21.11% | 576.2%
| Z-Score (Absolute) |  | 0.53 | 135.39 | 0.03 | 1.57 | 0.73 | 0.52 | 0.53 | 0.73 | 0.03 | 57.9 | 0.98 | 0.62 | 0.23 | 13.09
| Analytical results for Zr in OREAS 556b (Certified Value 216 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 06 | 07 | 08 | 12 | 21 | 22 | 25 | 28 | [10 empty] | ppm
|  |  | PPP*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | BF*OES | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 200 | 220 | 211 | 300 | 200 | 216 | 70 | 215 | 200 | [10 empty] | 1
| 1 | 2 | 200 | 220 | 209 | 200 | 300 | 213 | < 70 | 219 | 200 | [10 empty] | 18
| 1 | 3 | 200 | 220 | 209 | 300 | 200 | 211 | < 70 | 206 | 200 | [10 empty] | 16
| 1 | 4 | 200 | 150 | 212 | 300 | 300 | 213 | < 70 | 212 | 200 | [10 empty] | 216.113
| 1 | 5 | 200 | 220 | 211 | 300 | 300 | 215 | 70 | 209 | 200 | [10 empty] | 23
| 1 | 6 | 200 | 220 | 211 | 300 | 300 | 214 | 70 | 248 | 200
| Mean |  | 200 | 208.3333 | 210.5 | 283.3333 | 266.6667 | 213.6667 | 70 | 218.1667 | 200
| Median |  | 200 | 220 | 211 | 300 | 300 | 213.5 | 70 | 213.5 | 200
| Std Dev. |  | 0 | 28.5774 | 1.2247 | 40.8248 | 51.6398 | 1.7512 | 0 | 15.3025 | 0
| Rel.Std.Dev. |  | 0% | 13.72% | 0.5818% | 14.41% | 19.36% | 0.8196% | 0% | 7.014% | 0%
| PDM3 |  | -7.456% | -3.6% | -2.597% | 31.1% | 23.39% | -1.132% | -67.61% | 0.9503% | -7.456%
| Z-Score (Absolute) |  | 0.67 | 0.05 | 0 | 4.68 | 3.61 | 0.2 | 9.95 | 0.49 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C813: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C831: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C867: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C885: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C903: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C921: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C939: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C957: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C976: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C994: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1012: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 556b (Certified Value 14.41 wt.%) | [25 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | wt.%
|  |  | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | BF*XRF | BF*XRF | BF*XRF | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 14.27 | 14.68 | 14.43 | 14.66 | 14.34 | 14.3 | 14.46 | 14.03 | 14.76 | 14.22 | 14.02 | 14.37 | 14.3 | 14.616 | 15.0935 | 14.001 | 14.0291 | 14.76 | 14.7 | 14.12 | 14.9569 | 14.7 | 14.1 | 14.1 | 1
| 1 | 2 | 14.18 | 14.42 | 14.42 | 14.72 | 14.37 | 14.3 | 14.54 | 13.93 | 14.84 | 14.21 | 14.06 | 14.42 | 14.4 | 14.6296 | 15.1066 | 13.9523 | 14.0525 | 14.86 | 14.8 | 14.08 | 14.6825 | 14.76 | 14.1 | 14.05 | 0
| 1 | 3 | 14.22 | 14.33 | 14.37 | 14.75 | 14.38 | 14.3 | 14.52 | 13.98 | 14.78 | 14.21 | 13.76 | 14.35 | 14.3 | 14.6087 | 15.06 | 13.9797 | 14.0576 | 14.52 | 14.9 | 14.02 | 14.9227 | 14.83 | 14.1 | 14.05 | 16
| 1 | 4 | 14.39 | 14.68 | 14.41 | 14.77 | 14.35 | 14.3 | 14.53 | 13.93 | 14.66 | 14.26 | 13.9 | 14.36 | 14.3 | 14.6235 | 15.0887 | 14.0742 | 14.4287 | 14.76 | 14.7 | 13.98 | 14.9227 | 15.07 | 14 | 14.09 | 14.409
| 1 | 5 | 14.33 | 14.43 | 14.38 | 14.76 | 14.39 | 14.3 | 14.53 | 13.98 | 14.72 | 14.14 | 13.56 | 14.41 | 14.3 | 14.5893 | 14.9594 | 13.8789 | 14.1008 | 14.79 | 14.6 | 14.12 | 14.6533 | 14.64 | 14 | 14.07 | 25
| 1 | 6 | 14.41 | 14.49 | 14.42 | 14.87 | 14.3 | 14.3 | 14.52 | 13.94 | 14.82 | 14.2 | 14.13 | 14.42 | 14.4 | 14.59 | 14.9564 | 13.9908 | 14.1919 | 14.89 | 14.8 | 13.98 | 14.992 | 14.79 | 14 | 14.09
| Mean |  | 14.3 | 14.505 | 14.405 | 14.755 | 14.355 | 14.3 | 14.5167 | 13.965 | 14.7633 | 14.2067 | 13.905 | 14.3883 | 14.3333 | 14.6095 | 15.0441 | 13.9795 | 14.1434 | 14.7633 | 14.75 | 14.05 | 14.855 | 14.7983 | 14.05 | 14.075
| Median |  | 14.3 | 14.46 | 14.415 | 14.755 | 14.36 | 14.3 | 14.525 | 13.96 | 14.77 | 14.21 | 13.96 | 14.39 | 14.3 | 14.6124 | 15.0744 | 13.9852 | 14.0792 | 14.775 | 14.75 | 14.05 | 14.9227 | 14.775 | 14.05 | 14.08
| Std Dev. |  | 0.09252 | 0.14488 | 0.02429 | 0.06892 | 0.032711 | 1.9459e-15 | 0.028752 | 0.03937 | 0.066232 | 0.038816 | 0.21352 | 0.031885 | 0.05164 | 0.016914 | 0.068487 | 0.063897 | 0.15117 | 0.13064 | 0.10488 | 0.065422 | 0.1475 | 0.14905 | 0.054772 | 0.021679
| Rel.Std.Dev. |  | 0.647% | 0.9988% | 0.1686% | 0.4671% | 0.2279% | 1.361e-14% | 0.1981% | 0.2819% | 0.4486% | 0.2732% | 1.536% | 0.2216% | 0.3603% | 0.1158% | 0.4552% | 0.4571% | 1.069% | 0.8849% | 0.7111% | 0.4656% | 0.9929% | 1.007% | 0.3898% | 0.154%
| PDM3 |  | -0.7566% | 0.6661% | -0.02792% | 2.401% | -0.3749% | -0.7566% | 0.7471% | -3.082% | 2.459% | -1.404% | -3.498% | -0.1436% | -0.5253% | 1.391% | 4.408% | -2.981% | -1.843% | 2.459% | 2.366% | -2.492% | 3.095% | 2.702% | -2.492% | -2.318%
| Z-Score (Absolute) |  | 0.16 | 0.29 | 0.07 | 0.84 | 0.04 | 0.16 | 0.32 | 0.89 | 0.85 | 0.36 | 1.02 | 0.04 | 0.08 | 0.52 | 1.47 | 0.86 | 0.5 | 0.85 | 0.83 | 0.7 | 1.05 | 0.93 | 0.7 | 0.65
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 556b (Certified Value 4.58 wt.%) | [27 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC |  | 2
3
| 1 | 1 | 4.3 | 4.66 | 4.35 | 4.47 | 4.84 | 4.67 | 4.4277 | 4.46 | 4.46 | 4.59 | 4.64 | 4.53 | 5.1 | 3.9675 | 4.89 | 4.5 | 4.57 | 4.51 | 4.5 | 4.72 | 4.74 | 4.94 | 4.78 | 4.25 | 4.54 |  | 1
| 1 | 2 | 4.36 | 4.6 | 4.38 | 4.57 | 4.78 | 4.7 | 4.3104 | 4.51 | 4.5 | 4.61 | 4.5 | 4.52 | 5.02 | 3.99 | 4.87 | 4.52 | 4.6 | 4.5 | 4.5 | 4.71 | 4.87 | 5.04 | 4.76 | 4.21 | 4.58 |  | 0
| 1 | 3 | 4.27 | 4.65 | 4.42 | 4.5 | 4.82 | 4.77 | 4.3762 | 4.48 | 4.39 | 4.56 | 4.58 | 4.53 | 4.96 | 3.9793 | 4.85 | 4.53 | 4.56 | 4.46 | 4.5 | 4.72 | 4.84 | 5.09 | 4.72 | 4.23 | 4.56 |  | 16
| 1 | 4 | 4.27 | 4.62 | 4.4 | 4.57 | 4.87 | 4.55 | 4.3279 | 4.46 | 4.34 | 4.58 | 4.54 | 4.52 | 5.04 | 3.9672 | 4.79 | 4.53 | 4.56 | 4.49 | 4.51 | 4.73 | 4.98 | 5.01 | 4.71 | 4.23 | 4.53 |  | 4.5805
| 1 | 5 | 4.34 | 4.68 | 4.36 | 4.57 | 4.84 | 4.92 | 4.324 | 4.46 | 4.46 | 4.54 | 4.6 | 4.51 | 4.99 | 3.9831 | 4.86 | 4.52 | 4.61 | 4.49 | 4.52 | 4.73 | 4.75 | 5.11 | 4.73 | 4.23 | 4.58 |  | 27
| 1 | 6 | 4.32 | 4.62 | 4.4 | 4.55 | 4.82 | 4.5 | 4.2875 | 4.48 | 4.43 | 4.59 | 4.61 | 4.47 | 5.08 | 3.9873 | 4.83 | 4.525 | 4.59 | 4.51 | 4.51 | 4.71 | 4.77 | 5.2 | 4.75 | 4.22 | 4.52
| Mean |  | 4.31 | 4.6383 | 4.385 | 4.5383 | 4.8283 | 4.685 | 4.3423 | 4.475 | 4.43 | 4.5783 | 4.5783 | 4.5133 | 5.0317 | 3.9791 | 4.8483 | 4.5208 | 4.5817 | 4.4933 | 4.5067 | 4.72 | 4.825 | 5.065 | 4.7417 | 4.2283 | 4.5517
| Median |  | 4.31 | 4.635 | 4.39 | 4.56 | 4.83 | 4.685 | 4.3259 | 4.47 | 4.445 | 4.585 | 4.59 | 4.52 | 5.03 | 3.9812 | 4.855 | 4.5225 | 4.58 | 4.495 | 4.505 | 4.72 | 4.805 | 5.065 | 4.74 | 4.23 | 4.55
| Std Dev. |  | 0.036878 | 0.029944 | 0.026646 | 0.043089 | 0.029944 | 0.15189 | 0.051001 | 0.019748 | 0.057271 | 0.024833 | 0.050761 | 0.022509 | 0.053072 | 0.009784 | 0.034881 | 0.011143 | 0.02137 | 0.018619 | 0.008165 | 0.0089443 | 0.091815 | 0.08961 | 0.026394 | 0.013292 | 0.025626
| Rel.Std.Dev. |  | 0.8556% | 0.6456% | 0.6077% | 0.9494% | 0.6202% | 3.242% | 1.175% | 0.4413% | 1.293% | 0.5424% | 1.109% | 0.4987% | 1.055% | 0.2459% | 0.7194% | 0.2465% | 0.4664% | 0.4144% | 0.1812% | 0.1895% | 1.903% | 1.769% | 0.5566% | 0.3143% | 0.563%
| PDM3 |  | -5.905% | 1.263% | -4.268% | -0.9205% | 5.411% | 2.281% | -5.201% | -2.303% | -3.286% | -0.04725% | -0.04725% | -1.466% | 9.85% | -13.13% | 5.847% | -1.303% | 0.02553% | -1.903% | -1.612% | 3.046% | 5.338% | 10.58% | 3.519% | -7.688% | -0.6294%
| Z-Score (Absolute) |  | 1.22 | 0.44 | 0.84 | 0.07 | 1.4 | 0.67 | 1.06 | 0.39 | 0.62 | 0.13 | 0.13 | 0.19 | 2.43 | 2.9 | 1.5 | 0.16 | 0.15 | 0.3 | 0.23 | 0.85 | 1.38 | 2.6 | 0.96 | 1.64 | 0
| Analytical results for S in OREAS 556b (Certified Value 8.78 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 9.24 | 8.88 | 8.63 | 8.77 | 8.9 | 9.45 | 9.0967 | 8.67 | 8.55 | 8.64 | 8.85 | 8.9 | 8.85 | 8.9 | 8.907 | 8.91 | 8.73 | 8.73 | 8.79 | 8.61 | 8.63 | 9.98 | 8.94 | 8.77 | 9.09 | 8.57 | 1
| 1 | 2 | 9.41 | 9 | 8.64 | 8.76 | 9.02 | 9.49 | 8.8385 | 8.75 | 8.56 | 8.74 | 8.63 | 8.7 | 8.86 | 8.9 | 8.8979 | 8.94 | 8.71 | 8.91 | 8.76 | 8.63 | 8.61 | 10.1 | 9.03 | 8.76 | 8.69 | 8.72 | 28
| 1 | 3 | 9.4 | 8.76 | 8.63 | 8.8 | 8.92 | 9.6 | 9.1299 | 8.54 | 8.6 | 8.74 | 8.66 | 8.74 | 8.82 | 8.9 | 8.929 | 8.91 | 8.73 | 8.66 | 8.71 | 8.62 | 8.64 | 10 | 9.03 | 8.76 | 8.83 | 8.71 | 16
| 1 | 4 | 9.27 | 8.8 | 8.69 | 8.77 | 8.91 | 9.35 | 8.7205 | 8.71 | 8.59 | 8.61 | 8.33 | 8.8 | 8.89 | 8.9 | 8.9297 | 8.92 | 8.75 | 8.61 | 8.75 | 8.7 | 8.68 | 10.1 | 8.93 | 8.75 | 9 | 8.64 | 8.7844
| 1 | 5 | 9.34 | 9.25 | 8.7 | 8.81 | 8.95 | 9.74 | 9.0204 | 8.59 | 8.63 | 8.68 | 8.65 | 8.85 | 8.87 | 8.9 | 8.8958 | 8.94 | 8.72 | 8.63 | 8.76 | 8.67 | 8.65 | 10 | 9.06 | 8.75 | 8.97 | 8.63 | 28
| 1 | 6 | 9.32 | 8.92 | 8.66 | 8.84 | 8.96 | 9.34 | 8.7315 | 8.58 | 8.65 | 8.75 | 8.81 | 8.78 | 8.88 | 8.8 | 8.9097 | 8.94 | 8.745 | 8.72 | 8.8 | 8.63 | 8.61 | 10 | 9.09 | 8.72 | 9 | 8.61
| Mean |  | 9.33 | 8.935 | 8.6583 | 8.7917 | 8.9433 | 9.495 | 8.9229 | 8.64 | 8.5967 | 8.6933 | 8.655 | 8.795 | 8.8617 | 8.8833 | 8.9115 | 8.9267 | 8.7308 | 8.71 | 8.7617 | 8.6433 | 8.6367 | 10.03 | 9.0133 | 8.7517 | 8.93 | 8.6467
| Median |  | 9.33 | 8.9 | 8.65 | 8.785 | 8.935 | 9.47 | 8.9294 | 8.63 | 8.595 | 8.71 | 8.655 | 8.79 | 8.865 | 8.9 | 8.9083 | 8.93 | 8.73 | 8.69 | 8.76 | 8.63 | 8.635 | 10 | 9.03 | 8.755 | 8.985 | 8.635
| Std Dev. |  | 0.068118 | 0.17638 | 0.030605 | 0.030605 | 0.044121 | 0.15372 | 0.18291 | 0.082462 | 0.038816 | 0.059217 | 0.18349 | 0.072595 | 0.024833 | 0.040825 | 0.014792 | 0.015055 | 0.014972 | 0.109 | 0.031885 | 0.034448 | 0.026583 | 0.054772 | 0.064704 | 0.017224 | 0.14464 | 0.058195
| Rel.Std.Dev. |  | 0.7301% | 1.974% | 0.3535% | 0.3481% | 0.4933% | 1.619% | 2.05% | 0.9544% | 0.4515% | 0.6812% | 2.12% | 0.8254% | 0.2802% | 0.4596% | 0.166% | 0.1687% | 0.1715% | 1.251% | 0.3639% | 0.3986% | 0.3078% | 0.5461% | 0.7179% | 0.1968% | 1.62% | 0.673%
| PDM3 |  | 6.211% | 1.715% | -1.435% | 0.08303% | 1.81% | 8.09% | 1.577% | -1.644% | -2.137% | -1.036% | -1.473% | 0.121% | 0.8799% | 1.127% | 1.447% | 1.62% | -0.6095% | -0.8467% | -0.2585% | -1.606% | -1.681% | 14.18% | 2.606% | -0.3723% | 1.658% | -1.568%
| Z-Score (Absolute) |  | 2.66 | 0.7 | 0.67 | 0.01 | 0.74 | 3.48 | 0.64 | 0.76 | 0.98 | 0.5 | 0.69 | 0.01 | 0.34 | 0.45 | 0.59 | 0.66 | 0.31 | 0.41 | 0.16 | 0.74 | 0.78 | 6.14 | 1.09 | 0.21 | 0.68 | 0.73
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach (05%)
| Analytical results for Co in OREAS 556b (Certified Value 0.685 wt.%) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 13 | 15 | 18 | 19 | 23 | 25 | [10 empty] | wt.%
|  |  | AL-5S*AAS | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | [10 empty] | 3
3
| 1 | 1 | 0.6955 | 0.6998 | 0.72 | 0.61 | 0.72167 | 0.74522 | 1.5365 | 0.66352 | 0.5873 | [10 empty] | 1
| 1 | 2 | 0.7045 | 0.7213 | 0.705 | 0.62 | 0.71266 | 0.73031 | 1.7188 | 0.65278 | 0.5916 | [10 empty] | 36
| 1 | 3 | 0.7065 | 0.6991 | 0.705 | 0.66 | 0.72639 | 0.75907 | 1.6846 | 0.66616 | 0.6126 | [10 empty] | 16
| 1 | 4 | 0.6926 | 0.6939 | 0.742 | 0.65 | 0.71896 | 0.75228 | 1.6448 | 0.65692 | 0.6218 | [10 empty] | 0.68534
| 1 | 5 | 0.6995 | 0.7078 | 0.702 | 0.64 | 0.71436 | 0.77066 | 1.729 | 0.66334 | 0.5968 | [10 empty] | 30
| 1 | 6 | 0.703 | 0.7232 | 0.708 | 0.66 | 0.72349 | 0.70678 | 1.771 | 0.66492 | 0.6684
| Mean |  | 0.70027 | 0.70752 | 0.71367 | 0.64 | 0.71959 | 0.74405 | 1.6808 | 0.66127 | 0.61308
| Median |  | 0.70125 | 0.7038 | 0.7065 | 0.645 | 0.72031 | 0.74875 | 1.7017 | 0.66343 | 0.6047
| Std Dev. |  | 0.0054099 | 0.012262 | 0.01524 | 0.020976 | 0.0053211 | 0.022718 | 0.082506 | 0.0052464 | 0.030092
| Rel.Std.Dev. |  | 0.7725% | 1.733% | 2.135% | 3.278% | 0.7395% | 3.053% | 4.909% | 0.7934% | 4.908%
| PDM3 |  | 2.178% | 3.236% | 4.133% | -6.616% | 4.997% | 8.567% | 145.2% | -3.511% | -10.54%
| Z-Score (Absolute) |  | 0.13 | 0 | 0.11 | 1.25 | 0.22 | 0.67 | 17.96 | 0.85 | 1.74
| Analytical results for Cu in OREAS 556b (Certified Value 1.32 wt.%) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 05 | 08 | 09 | 10 | 11 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 25 | 27 | 28 | wt.%
|  |  | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*OES | 2
3
| 1 | 1 | 1.329 | 1.7525 | 1.4183 | 1.37 | 1.225 | 1.34 | 1.315 | 1.245 | 1.63 | 1.2658 | 1.47 | 1.6301 | 1.1912 | 1.415 | 1.385 | 1.0272 | 1.062 | 1.4 | 1.34 | 1
| 1 | 2 | 1.338 | 1.7785 | 1.4163 | 1.38 | 1.21 | 1.335 | 1.27 | 1.33 | 1.62 | 1.2132 | 1.41 | 1.6541 | 1.2041 | 1.41 | 1.374 | 1.0364 | 1.053 | 1.34 | 1.346 | 39
| 1 | 3 | 1.316 | 1.7765 | 1.3997 | 1.37 | 1.24 | 1.36 | 1.3 | 1.27 | 1.61 | 1.2211 | 1.45 | 1.6368 | 1.2192 | 1.403 | 1.386 | 1.0106 | 1.074 | 1.324 | 1.329 | 16
| 1 | 4 | 1.34 | 1.7223 | 1.4121 | 1.36 | 1.225 | 1.34 | 1.295 | 1.29 | 1.65 | 1.293 | 1.45 | 1.6325 | 1.2284 | 1.4 | 1.372 | 1.0328 | 1.092 | 1.346 | 1.365 | 1.318
| 1 | 5 | 1.343 | 1.7221 | 1.4379 | 1.38 | 1.23 | 1.37 | 1.31 | 1.335 | 1.64 | 1.2369 | 1.43 | 1.6308 | 1.2241 | 1.404 | 1.387 | 1.0148 | 1.095 | 1.375 | 1.351 | 31
| 1 | 6 | 1.329 | 1.7555 | 1.4473 | 1.36 | 1.23 | 1.345 | 1.28 | 1.25 | 1.66 | 1.2744 | 1.45 | 1.6424 | 1.2455 | 1.419 | 1.376 | 1.0328 | 1.118 | 1.36 | 1.355
| Mean |  | 1.3325 | 1.7512 | 1.4219 | 1.37 | 1.2267 | 1.3483 | 1.295 | 1.2867 | 1.635 | 1.2507 | 1.4433 | 1.6378 | 1.2188 | 1.4085 | 1.38 | 1.0258 | 1.0823 | 1.3575 | 1.3477
| Median |  | 1.3335 | 1.754 | 1.4173 | 1.37 | 1.2275 | 1.3425 | 1.2975 | 1.28 | 1.635 | 1.2514 | 1.45 | 1.6347 | 1.2216 | 1.407 | 1.3805 | 1.03 | 1.083 | 1.353 | 1.3485
| Std Dev. |  | 0.0099348 | 0.02485 | 0.017518 | 0.0089443 | 0.0098319 | 0.013663 | 0.017321 | 0.038944 | 0.018708 | 0.03176 | 0.020656 | 0.0092089 | 0.019032 | 0.0074498 | 0.0067231 | 0.010612 | 0.023956 | 0.027127 | 0.012453
| Rel.Std.Dev. |  | 0.7456% | 1.419% | 1.232% | 0.6529% | 0.8015% | 1.013% | 1.337% | 3.027% | 1.144% | 2.539% | 1.431% | 0.5623% | 1.562% | 0.5289% | 0.4872% | 1.035% | 2.213% | 1.998% | 0.924%
| PDM3 |  | 1.101% | 32.87% | 7.886% | 3.946% | -6.929% | 2.302% | -1.745% | -2.377% | 24.05% | -5.103% | 9.51% | 24.26% | -7.53% | 6.867% | 4.705% | -22.17% | -17.88% | 2.997% | 2.251%
| Z-Score (Absolute) |  | 0.15 | 3.69 | 0.67 | 0.2 | 1.11 | 0 | 0.49 | 0.56 | 2.63 | 0.89 | 0.87 | 2.65 | 1.19 | 0.55 | 0.29 | 2.96 | 2.44 | 0.08 | 0.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 556b (Indicative Value 0.453 ppm) | [23 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [19 empty] | 1
| No | No | 01 | 25 | 26 | [19 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | [19 empty] | 3
3
| 1 | 1 | 0.67 | 0.2 | < 5 | [19 empty] | 1
| 1 | 2 | 0.79 | 0.2 | < 5 | [19 empty] | 1
| 1 | 3 | 0.65 | 0.2 | < 5 | [19 empty] | 16
| 1 | 4 | 0.74 | 0.3 | < 5 | [19 empty] | 0.45333
| 1 | 5 | 0.72 | 0.1 | < 5 | [19 empty] | 23
| 1 | 6 | 0.67 | 0.2 | < 5
| Mean |  | 0.70667 | 0.2
| Median |  | 0.695 | 0.2
| Std Dev. |  | 0.053166 | 0.063246
| Rel.Std.Dev. |  | 7.524% | 31.62%
| PDM3 |  | 55.88% | -55.88%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.39
| Analytical results for Al in OREAS 556b (Certified Value 5.71 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 5.27 | 5.77 | 5.61 | 5.93 | 5.68 | 5.91 | 5.61 | 5.531 | 5.663 | 5.636 | 5.663 | 5.7351 | 7.6437 | 6.11 | 5.6865 | 5.372 | 5.68 | 6.9 | 3.7 | 5.62 | 5.47 | 5.7 | 1
| 1 | 2 | 5.34 | 5.82 | 5.6 | 5.94 | 5.72 | 5.88 | 5.61 | 5.61 | 5.61 | 5.742 | 5.795 | 5.7127 | 7.6492 | 6.15 | 5.6753 | 5.4588 | 5.73 | 7.02 | 7.25 | 5.72 | 5.44 | 5.73 | 0
| 1 | 3 | 5.16 | 5.82 | 5.6 | 5.78 | 5.76 | 5.83 | 5.61 | 5.583 | 5.504 | 5.689 | 5.742 | 5.6924 | 7.6903 | 6.07 | 5.7305 | 5.3677 | 5.61 | 7.2 | 6.76 | 5.52 | 5.47 | 5.71 | 16
| 1 | 4 | 5.23 | 5.88 | 5.62 | 5.94 | 5.76 | 5.89 | 5.636 | 5.583 | 5.478 | 5.478 | 5.795 | 5.8167 | 7.6975 | 5.92 | 5.6803 | 5.5634 | 5.69 | 6.99 | 3.91 | 5.8 | 5.51 | 5.74 | 5.7078
| 1 | 5 | 5.21 | 5.86 | 5.65 | 5.88 | 5.98 | 5.87 | 5.61 | 5.61 | 5.636 | 5.636 | 5.769 | 5.7395 | 7.6918 | 6.1 | 5.6861 | 5.5287 | 5.68 | 6.73 | 4.9 | 5.76 | 5.51 | 5.79 | 33
| 1 | 6 | 5.09 | 5.86 | 5.61 | 5.84 | 5.96 | 5.81 | 5.61 | 5.583 | 5.372 | 5.531 | 5.769 | 5.8189 | 7.6256 | 6.11 | 5.7007 | 5.3683 | 5.75 | 6.87 | 4.19 | 5.88 | 5.56 | 5.64
| Mean |  | 5.2167 | 5.835 | 5.615 | 5.885 | 5.81 | 5.865 | 5.6143 | 5.5833 | 5.5438 | 5.6187 | 5.7555 | 5.7526 | 7.6664 | 6.0767 | 5.6932 | 5.4432 | 5.69 | 6.9517 | 5.1183 | 5.7167 | 5.4933 | 5.7183
| Median |  | 5.22 | 5.84 | 5.61 | 5.905 | 5.76 | 5.875 | 5.61 | 5.583 | 5.557 | 5.636 | 5.769 | 5.7373 | 7.6697 | 6.105 | 5.6863 | 5.4154 | 5.685 | 6.945 | 4.545 | 5.74 | 5.49 | 5.72
| Std Dev. |  | 0.086641 | 0.039875 | 0.018708 | 0.065038 | 0.12759 | 0.037815 | 0.010614 | 0.028849 | 0.11182 | 0.098215 | 0.049444 | 0.053308 | 0.030526 | 0.080911 | 0.020141 | 0.087614 | 0.04858 | 0.15892 | 1.5245 | 0.12925 | 0.042269 | 0.049565
| Rel.Std.Dev. |  | 1.661% | 0.6834% | 0.3332% | 1.105% | 2.196% | 0.6448% | 0.1891% | 0.5167% | 2.017% | 1.748% | 0.8591% | 0.9267% | 0.3982% | 1.332% | 0.3538% | 1.61% | 0.8538% | 2.286% | 29.78% | 2.261% | 0.7695% | 0.8668%
| PDM3 |  | -8.605% | 2.228% | -1.627% | 3.104% | 1.79% | 2.753% | -1.638% | -2.181% | -2.873% | -1.562% | 0.8349% | 0.7834% | 34.31% | 6.462% | -0.256% | -4.637% | -0.3126% | 21.79% | -10.33% | 0.1546% | -3.758% | 0.1838%
| Z-Score (Absolute) |  | 2.62 | 0.7 | 0.48 | 0.97 | 0.56 | 0.86 | 0.49 | 0.65 | 0.86 | 0.46 | 0.27 | 0.26 | 10.52 | 1.99 | 0.06 | 1.4 | 0.08 | 6.68 | 3.15 | 0.06 | 1.13 | 0.07
| Analytical results for As in OREAS 556b (Certified Value 101 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 21 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 0
0
| 1 | 1 | 82.9 | < 100 | 106 | 102 | 98 | 100 | 100 | 100 | 100 | 200 | 200 | 101.7469 | 110 | 90 | 71 | 94 | 99 | < 100 | 100 |  |  |  | 1
| 1 | 2 | 82.7 | 101 | 105 | 105 | 105 | 103 | 100 | 100 | 100 | 300 | 200 | 102.0496 | 110 | 80 | 74 | 106 | 96 | < 100 | 100 |  |  |  | 2
| 1 | 3 | 75.4 | < 100 | 106 | 108 | 101 | 99 | 100 | 200 | 200 | 200 | 200 | 100.4637 | 110 | 90 | 75 | 102 | 92 | < 100 | 100 |  |  |  | 16
| 1 | 4 | 79.6 | < 100 | 105 | 106 | 104 | 101 | 100 | 100 | 200 | 200 | 200 | 100.3205 | 120 | 90 | 72 | 108 | 95 | < 100 | < 100 |  |  |  | 101.3026
| 1 | 5 | 78.3 | < 100 | 106 | 106 | 102 | 100 | 100 | 100 | 100 | 200 | 200 | 101.6762 | 110 | 100 | 73 | 98 | 90 | 100 | < 100 |  |  |  | 34
| 1 | 6 | 81.6 | < 100 | 105 | 102 | 104 | 100 | 100 | 100 | 100 | 200 | 200 | 100.9018 | 120 | 90 | 71 | 96 | 90 | 100 | 100
| Mean |  | 80.0833 | 101 | 105.5 | 104.8333 | 102.3333 | 100.5 | 100 | 116.6667 | 133.3333 | 216.6667 | 200 | 101.1931 | 113.3333 | 90 | 72.6667 | 100.6667 | 93.6667 | 100 | 100
| Median |  | 80.6 | 101 | 105.5 | 105.5 | 103 | 100 | 100 | 100 | 100 | 200 | 200 | 101.289 | 110 | 90 | 72.5 | 100 | 93.5 | 100 | 100
| Std Dev. |  | 2.913 |  | 0.54772 | 2.4014 | 2.582 | 1.3784 | 0 | 40.8248 | 51.6398 | 40.8248 | 0 | 0.72826 | 5.164 | 6.3246 | 1.633 | 5.6095 | 3.6148 | 0 | 0
| Rel.Std.Dev. |  | 3.637% |  | 0.5192% | 2.291% | 2.523% | 1.372% | 0% | 34.99% | 38.73% | 18.84% | 0% | 0.7197% | 4.556% | 7.027% | 2.247% | 5.572% | 3.859% | 0% | 0%
| PDM3 |  | -20.95% | -0.2988% | 4.143% | 3.485% | 1.017% | -0.7923% | -1.286% | 15.17% | 31.62% | 113.9% | 97.43% | -0.1081% | 11.88% | -11.16% | -28.27% | -0.6278% | -7.538% | -1.286% | -1.286%
| Z-Score (Absolute) |  | 2.34 | 4.8 | 0.56 | 0.48 | 0.2 | 0.01 | NA | NA | NA | NA | NA | 0.07 | 1.45 | 1.21 | 3.18 | 0.01 | 0.79 | NA | NA
Results from laboratories 8, 9, 10, 11, 12, 27 and 28 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 556b (Certified Value 199 ppm) | [23 empty] | Certified
B | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 19; 06: 21; 08: 25; 09: 26 | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*OES; 05: PF*MS; 06: PF*OES; 08: PF*OES; 09: PF*OES | 0
0
| 1 | 1 | 01: 207; 02: 260; 03: 195; 04: 251.8911; 05: 217; 06: 180; 08: 74; 09: 178 | 1
| 1 | 2 | 01: 213; 02: 240; 03: 199; 04: 236.7207; 05: 221; 06: 160; 08: 193; 09: 174 | 0
| 1 | 3 | 01: 214; 02: 240; 03: 199; 04: 223.4028; 05: 219; 06: 170; 08: 51; 09: 177 | 16
| 1 | 4 | 01: 217; 02: 260; 03: 201; 04: 226.5043; 05: 222; 06: 160; 08: 187; 09: 180 | 199.287
| 1 | 5 | 01: 212; 02: 230; 03: 197; 04: 227.4793; 05: 216; 06: 180; 08: 106; 09: 181 | 35
| 1 | 6 | 01: 202; 02: 240; 03: 197; 04: 250.7792; 05: 215; 06: 170; 08: 125; 09: 184
| Mean |  | 01: 210.8333; 02: 245; 03: 198; 04: 236.1296; 05: 218.3333; 06: 170; 08: 122.6667; 09: 179
| Median |  | 01: 212.5; 02: 240; 03: 198; 04: 232.1; 05: 218; 06: 170; 08: 115.5; 09: 179
| Std Dev. |  | 01: 5.4191; 02: 12.2474; 03: 2.0976; 04: 12.5923; 05: 2.8048; 06: 8.9443; 08: 58.0919; 09: 3.4641
| Rel.Std.Dev. |  | 01: 2.57%; 02: 4.999%; 03: 1.059%; 04: 5.333%; 05: 1.285%; 06: 5.261%; 08: 47.36%; 09: 1.935%
| PDM3 |  | 01: 5.794%; 02: 22.94%; 03: -0.6458%; 04: 18.49%; 05: 9.557%; 06: -14.7%; 08: -38.45%; 09: -10.18%
| Z-Score (Absolute) |  | 01: 0.15; 02: 0.96; 03: 0.15; 04: 0.75; 05: 0.33; 06: 0.81; 08: 1.93; 09: 0.6
| Analytical results for Ba in OREAS 556b (Certified Value 196 ppm) | [23 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 201.1 | 222 | 179 | 198 | 194 | 233 | 192.2248 | 215 | 191.6906 | 191 | 190 | 204 | 131 | 188 | 191 | [7 empty] | 1
| 1 | 2 | 202.3 | 208 | 176 | 196 | 192 | 232 | 193.6352 | 210 | 190.812 | 209 | 170 | 204 | 248 | 192 | 192 | [7 empty] | 30
| 1 | 3 | 194 | 208 | 177 | 200 | 195 | 233 | 198.5256 | 217 | 197.3013 | 190 | 200 | 215 | 247 | 205 | 187 | [7 empty] | 16
| 1 | 4 | 203.9 | 208 | 176 | 201 | 200 | 238 | 202.2738 | 211 | 190.7879 | 194 | 180 | 185 | 132 | 195 | 192 | [7 empty] | 195.6728
| 1 | 5 | 197.3 | 210 | 179 | 197 | 199 | 235 | 197.0649 | 207 | 191.9398 | 240 | 190 | 185 | 173 | 201 | 196 | [7 empty] | 36
| 1 | 6 | 200.9 | 221 | 179 | 200 | 197 | 232 | 200.6733 | 209 | 196.2098 | 190 | 180 | 194 | 195 | 202 | 197
| Mean |  | 199.9167 | 212.8333 | 177.6667 | 198.6667 | 196.1667 | 233.8333 | 197.3996 | 211.5 | 193.1236 | 202.3333 | 185 | 197.8333 | 187.6667 | 197.1667 | 192.5
| Median |  | 201 | 209 | 178 | 199 | 196 | 233 | 197.7952 | 210.5 | 191.8152 | 192.5 | 185 | 199 | 184 | 198 | 192
| Std Dev. |  | 3.6268 | 6.7651 | 1.5055 | 1.9664 | 3.0605 | 2.3166 | 3.9193 | 3.7815 | 2.8716 | 19.8259 | 10.4881 | 11.9569 | 52.4201 | 6.5549 | 3.6194
| Rel.Std.Dev. |  | 1.814% | 3.179% | 0.8474% | 0.9898% | 1.56% | 0.9907% | 1.985% | 1.788% | 1.487% | 9.799% | 5.669% | 6.044% | 27.93% | 3.325% | 1.88%
| PDM3 |  | 2.169% | 8.77% | -9.202% | 1.53% | 0.2524% | 19.5% | 0.8825% | 8.089% | -1.303% | 3.404% | -5.454% | 1.104% | -4.092% | 0.7635% | -1.621%
| Z-Score (Absolute) |  | 0.35 | 2.12 | 2.72 | 0.17 | 0.17 | 5.01 | 0 | 1.94 | 0.59 | 0.68 | 1.71 | 0.06 | 1.34 | 0.03 | 0.67
| Analytical results for Be in OREAS 556b (Certified Value 5.06 ppm) | [23 empty] | Certified
Be | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 02; 03: 04; 04: 05; 05: 06; 06: 15; 08: 17; 09: 18; 10: 19; 11: 21; 12: 23; 15: 25; 16: 26 | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*OES/MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES; 12: PF*MS; 15: PF*OES; 16: PF*MS | 2
3
| 1 | 1 | 01: 4.8; 02: 12; 03: 5; 04: 5; 05: < 5; 06: 5.3372; 08: 5; 09: 5.2004; 10: NA; 11: 5; 12: 5; 15: 1; 16: 5 | 1
| 1 | 2 | 01: 5.2; 02: 13; 03: 5; 04: 6; 05: < 5; 06: 5.4736; 08: 5; 09: 5.1805; 10: NA; 11: 5; 12: 5; 15: 2; 16: 5 | 3
| 1 | 3 | 01: 4.7; 02: 11; 03: 5; 04: 6; 05: < 5; 06: 4.9973; 08: 5; 09: 5.2549; 10: NA; 11: 5; 12: 5; 15: 1; 16: 5 | 16
| 1 | 4 | 01: 4.9; 02: 11; 03: 5; 04: 6; 05: < 5; 06: 5.2574; 08: 5; 09: 5.1469; 10: NA; 11: 5; 12: 4; 15: 1; 16: 5 | 5.061
| 1 | 5 | 01: 4.7; 02: 11; 03: 5; 04: 5; 05: < 5; 06: 5.1003; 08: 5; 09: 5.3826; 10: NA; 11: 5; 12: 4; 15: 1; 16: 5 | 37
| 1 | 6 | 01: 4.8; 02: 11; 03: 5; 04: 6; 05: < 5; 06: 4.966; 08: 5; 09: 5.3176; 10: NA; 11: 6; 12: 4; 15: 1; 16: 5
| Mean |  | 01: 4.85; 02: 11.5; 03: 5; 04: 5.6667; 06: 5.1886; 08: 5; 09: 5.2471; 11: 5.1667; 12: 4.5; 15: 1.1667; 16: 5
| Median |  | 01: 4.8; 02: 11; 03: 5; 04: 6; 06: 5.1789; 08: 5; 09: 5.2276; 11: 5; 12: 4.5; 15: 1; 16: 5
| Std Dev. |  | 01: 0.18708; 02: 0.83666; 03: 0; 04: 0.5164; 06: 0.201; 08: 0; 09: 0.089504; 11: 0.40825; 12: 0.54772; 15: 0.40825; 16: 0
| Rel.Std.Dev. |  | 01: 3.857%; 02: 7.275%; 03: 0%; 04: 9.113%; 06: 3.874%; 08: 0%; 09: 1.706%; 11: 7.902%; 12: 12.17%; 15: 34.99%; 16: 0%
| PDM3 |  | 01: -4.169%; 02: 127.2%; 03: -1.206%; 04: 11.97%; 06: 2.522%; 08: -1.206%; 09: 3.677%; 11: 2.088%; 12: -11.08%; 15: -76.95%; 16: -1.206%
| Z-Score (Absolute) |  | 01: 0.46; 02: 20.12; 03: 0; 04: 2.06; 05: 7.74; 06: 0.58; 08: 0; 09: 0.76; 10: NA; 11: 0.52; 12: 1.55; 15: 11.86; 16: 0
| Analytical results for Bi in OREAS 556b (Certified Value 3.08 ppm) | [23 empty] | Certified
Bi | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 01; 02: 04; 03: 05; 04: 06; 05: 15; 06: 17; 08: 19; 09: 25; 10: 26 | ppm
|  |  | 01: PF*OES; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*OES; 10: PF*MS | 2
3
| 1 | 1 | 01: < 5; 02: 3; 03: 3; 04: 3; 05: 2.6177; 06: 3.33; 08: 3; 09: 2.9; 10: 3.1 | 1
| 1 | 2 | 01: < 5; 02: 3.1; 03: 3; 04: 2.9; 05: 2.3782; 06: 3.1; 08: 3; 09: 3.4; 10: 3 | 4
| 1 | 3 | 01: < 5; 02: 3.1; 03: 3; 04: 2.9; 05: 2.3532; 06: 3.23; 08: 3; 09: 3.1; 10: 3.2 | 16
| 1 | 4 | 01: < 5; 02: 3.1; 03: 3; 04: 3.2; 05: 2.5438; 06: 3.19; 08: 3; 09: 3.2; 10: 3.1 | 3.0753
| 1 | 5 | 01: < 5; 02: 2.9; 03: 3; 04: 3; 05: 2.4467; 06: 3.15; 08: 3; 09: 3.1; 10: 3 | 38
| 1 | 6 | 01: < 5; 02: 3; 03: 3; 04: 2.9; 05: 2.536; 06: 3.16; 08: 3; 09: 2.9; 10: 3
| Mean |  | 02: 3.0333; 03: 3; 04: 2.9833; 05: 2.4793; 06: 3.1933; 08: 3; 09: 3.1; 10: 3.0667
| Median |  | 02: 3.05; 03: 3; 04: 2.95; 05: 2.4913; 06: 3.175; 08: 3; 09: 3.1; 10: 3.05
| Std Dev. |  | 02: 0.08165; 03: 0; 04: 0.1169; 05: 0.10364; 06: 0.079666; 08: 0; 09: 0.18974; 10: 0.08165
| Rel.Std.Dev. |  | 02: 2.692%; 03: 0%; 04: 3.919%; 05: 4.18%; 06: 2.495%; 08: 0%; 09: 6.121%; 10: 2.662%
| PDM3 |  | 02: -1.366%; 03: -2.45%; 04: -2.992%; 05: -19.38%; 06: 3.837%; 08: -2.45%; 09: 0.8021%; 10: -0.2818%
| Z-Score (Absolute) |  | 01: 5.39; 02: 0; 03: NA; 04: 0.51; 05: 5.6; 06: 1.62; 08: NA; 09: 0.67; 10: 0.34
Results from laboratories 5 and 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 556b (Certified Value 0.149 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.15 | 0.18 | 0.12 | 0.1 | 0.16 | 0.08 | 0.15 | 0.143 | 0.15 | 0.164 | 0.157 | 0.12424 | 0.4946 | 0.17 | 0.14619 | 0.1636 | 0.1 | 0.2 | 0.14 | 0.3 | 0.19 | 0.1 | 1
| 1 | 2 | 0.15 | 0.15 | 0.12 | 0.1 | 0.18 | 0.07 | 0.1 | 0.229 | 0.157 | 0.186 | 0.25 | 0.1183 | 0.494 | 0.17 | 0.13915 | 0.2775 | 0.1 | 0.3 | 0.41 | 0.3 | 0.15 | 0.08 | 0
| 1 | 3 | 0.16 | 0.17 | 0.12 | 0.1 | 0.17 | 0.07 | 0.129 | 0.179 | 0.143 | 0.179 | 0.222 | 0.11951 | 0.4971 | 0.16 | 0.14566 | 0.161 | 0.1 | 0.4 | 0.37 | 0.3 | 0.24 | 0.1 | 16
| 1 | 4 | 0.15 | 0.19 | 0.12 | 0.1 | 0.19 | 0.08 | 0.093 | 0.157 | 0.15 | 0.15 | 0.157 | 0.13178 | 0.4972 | 0.16 | 0.14017 | 0.1673 | 0.1 | 0.2 | 0.14 | 0.3 | 0.11 | 0.09 | 0.14866
| 1 | 5 | 0.15 | 0.15 | 0.11 | 0.1 | 0.18 | 0.09 | 0.1 | 0.071 | 0.157 | 0.164 | 0.157 | 0.13333 | 0.491 | 0.17 | 0.14073 | 0.2808 | 0.1 | 0.2 | 0.28 | 0.3 | 0.14 | 0.08 | 39
| 1 | 6 | 0.15 | 0.14 | 0.12 | 0.1 | 0.17 | 0.07 | 0.093 | 0.157 | 0.15 | 0.157 | 0.157 | 0.12112 | 0.4947 | 0.18 | 0.14039 | 0.2748 | 0.1 | 0.3 | 0.29 | 0.3 | 0.12 | 0.09
| Mean |  | 0.15167 | 0.16333 | 0.11833 | 0.1 | 0.175 | 0.076667 | 0.11083 | 0.156 | 0.15117 | 0.16667 | 0.18333 | 0.12471 | 0.49477 | 0.16833 | 0.14205 | 0.22083 | 0.1 | 0.26667 | 0.27167 | 0.3 | 0.15833 | 0.09
| Median |  | 0.15 | 0.16 | 0.12 | 0.1 | 0.175 | 0.075 | 0.1 | 0.157 | 0.15 | 0.164 | 0.157 | 0.12268 | 0.49465 | 0.17 | 0.14056 | 0.22105 | 0.1 | 0.25 | 0.285 | 0.3 | 0.145 | 0.09
| Std Dev. |  | 0.0040825 | 0.019664 | 0.0040825 | 1.5202e-17 | 0.010488 | 0.008165 | 0.023387 | 0.051486 | 0.0052694 | 0.013501 | 0.041745 | 0.0064098 | 0.0022879 | 0.0075277 | 0.0030515 | 0.062355 | 1.5202e-17 | 0.08165 | 0.11303 | 0 | 0.048751 | 0.0089443
| Rel.Std.Dev. |  | 2.692% | 12.04% | 3.45% | 1.52e-14% | 5.993% | 10.65% | 21.1% | 33% | 3.486% | 8.1% | 22.77% | 5.14% | 0.4624% | 4.472% | 2.148% | 28.24% | 1.52e-14% | 30.62% | 41.61% | 0% | 30.79% | 9.938%
| PDM3 |  | 2.025% | 9.874% | -20.4% | -32.73% | 17.72% | -48.43% | -25.44% | 4.941% | 1.689% | 12.12% | 23.33% | -16.11% | 232.8% | 13.24% | -4.443% | 48.55% | -32.73% | 79.39% | 82.75% | 101.8% | 6.51% | -39.46%
| Z-Score (Absolute) |  | 0.17 | 0.19 | 1.19 | NA | 0.55 | 2.46 | 1.42 | 0.04 | 0.18 | 0.29 | 0.8 | 0.99 | 10.33 | 0.34 | 0.46 | 1.95 | NA | NA | 3.5 | NA | 0.04 | 2.06
Results from laboratory 21 were removed due to their 0.01 wt.% reading resolution. Results from laboratories 4, 23 and 26 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 556b (Indicative Value < 10 ppm) | [23 empty] | Indicative
Cd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 01; 02: 02; 03: 04; 04: 05; 05: 06; 06: 15; 08: 21; 09: 23; 10: 25; 11: 26 | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*OES; 09: PF*MS; 10: PF*OES; 11: PF*MS | 1
1
| 1 | 1 | 01: < 1; 02: < 5; 03: < 10; 04: < 2; 05: < 5; 06: < 10; 08: < 10; 09: 0.8; 10: 0.3; 11: < 10 | 1
| 1 | 2 | 01: < 1; 02: < 5; 03: < 10; 04: < 2; 05: < 5; 06: < 10; 08: 10; 09: 0.9; 10: 0.3; 11: < 10 | 5
| 1 | 3 | 01: < 1; 02: < 5; 03: < 10; 04: < 2; 05: < 5; 06: < 10; 08: 10; 09: 0.9; 10: 0.4; 11: < 10 | 16
| 1 | 4 | 01: < 1; 02: < 5; 03: < 10; 04: < 2; 05: < 5; 06: < 10; 08: 10; 09: 1.1; 10: 0.3; 11: < 10 | < 10
| 1 | 5 | 01: < 1; 02: < 5; 03: < 10; 04: < 2; 05: < 5; 06: < 10; 08: 10; 09: 0.9; 10: 0.4; 11: < 10 | 24
| 1 | 6 | 01: < 1; 02: < 5; 03: < 10; 04: < 2; 05: < 5; 06: < 10; 08: 10; 09: 0.8; 10: 0.4; 11: < 10
| Mean |  | 08: 10; 09: 0.9; 10: 0.35
| Median |  | 08: 10; 09: 0.9; 10: 0.35
| Std Dev. |  | 08: 0; 09: 0.10954; 10: 0.054772
| Rel.Std.Dev. |  | 08: 0%; 09: 12.17%; 10: 15.65%
PDM3
| Z-Score (Absolute) |  | 01: 0.65; 02: 0; 03: 0.81; 04: 0.49; 05: 0; 06: 0.81; 08: 2.17; 09: 0.52; 10: 0.7; 11: 0.81
| Analytical results for Ce in OREAS 556b (Certified Value 156 ppm) | [23 empty] | Certified
Ce | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 05; 03: 06; 04: 18; 05: 19; 06: 23; 08: 25; 09: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*MS; 08: PF*OES; 09: PF*OES | 0
0
| 1 | 1 | 01: 614; 02: 153; 03: 153; 04: 150.2095; 05: NA; 06: 127; 08: 157; 09: 156.2 | 1
| 1 | 2 | 01: 425; 02: 135; 03: 158; 04: 154.654; 05: NA; 06: 130; 08: 184; 09: 157.7 | 6
| 1 | 3 | 01: 469; 02: 155; 03: 156; 04: 159.2717; 05: NA; 06: 132; 08: 167; 09: 150.8 | 16
| 1 | 4 | 01: 390; 02: 155; 03: 154; 04: 155.6176; 05: NA; 06: 127; 08: 166; 09: 157.3 | 155.9579
| 1 | 5 | 01: 445; 02: 153; 03: 156; 04: 151.6683; 05: NA; 06: 136; 08: 166; 09: 160.1 | 40
| 1 | 6 | 01: 912; 02: 152; 03: 155; 04: 156.816; 05: NA; 06: 120; 08: 159; 09: 155.4
| Mean |  | 01: 542.5; 02: 150.5; 03: 155.3333; 04: 154.7062; 06: 128.6667; 08: 166.5; 09: 156.25
| Median |  | 01: 457; 02: 153; 03: 155.5; 04: 155.1358; 06: 128.5; 08: 166; 09: 156.75
| Std Dev. |  | 01: 196.7808; 02: 7.6877; 03: 1.7512; 04: 3.334; 06: 5.4283; 08: 9.5237; 09: 3.113
| Rel.Std.Dev. |  | 01: 36.27%; 02: 5.108%; 03: 1.127%; 04: 2.155%; 06: 4.219%; 08: 5.72%; 09: 1.992%
| PDM3 |  | 01: 247.9%; 02: -3.5%; 03: -0.4005%; 04: -0.8026%; 06: -17.5%; 08: 6.76%; 09: 0.1873%
| Z-Score (Absolute) |  | 01: 54.01; 02: 0.67; 03: 0; 04: 0.09; 05: NA; 06: 3.72; 08: 1.56; 09: 0.13
| Analytical results for Co in OREAS 556b (Certified Value 1.99 wt.%) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 2.2628 | 2.0231 | 1.9995 | 1.9846 | 1.96 | 1.93 | 1.97 | 1.96 | 1.87 | 1.98 | 2.06 | 2.0952 | 1.8098 | 2.23 | 1.8785 | 2.2784 | 1.983 | 2.01 | 2.11 | 1.938 | 2.035 |  | 1
| 1 | 2 | 2.2678 | 2.0777 | 1.9469 | 2.0004 | 2.01 | 2 | 2.02 | 2.04 | 1.885 | 2.06 | 2.06 | 2.0609 | 1.8878 | 2.18 | 1.9707 | 1.9827 | 1.853 | 2.16 | 2.0126 | 1.959 | 2.044 |  | 36
| 1 | 3 | 2.142 | 2.0216 | 1.9498 | 2.0936 | 1.97 | 2.05 | 2.13 | 1.965 | 1.78 | 2.08 | 2.04 | 2.0603 | 1.8784 | 2.22 | 1.8892 | 2.0342 | 1.984 | 2.1 | 2.0354 | 1.927 | 2.042 |  | 16
| 1 | 4 | 2.2268 | 2.0222 | 1.9467 | 2.0203 | 1.99 | 1.96 | 2.02 | 1.975 | 1.835 | 1.985 | 2.03 | 2.0811 | 1.8533 | 2.22 | 1.8995 | 2.0275 | 1.843 | 2.21 | 2.0728 | 1.914 | 2.014 |  | 1.9911
| 1 | 5 | 2.2173 | 2.0989 | 1.9881 | 2.0286 | 1.95 | 1.97 | 1.97 | 2.02 | 1.89 | 2.01 | 2.05 | 2.0754 | 1.8293 | 2.25 | 1.9764 | 1.9935 | 2.015 | 2.07 | 2.0783 | 1.929 | 2.036 |  | 41
| 1 | 6 | 2.2229 | 2.054 | 1.9851 | 2.0168 | 1.94 | 1.96 | 2.01 | 1.985 | 1.78 | 2.03 | 2.08 | 2.0443 | 1.8058 | 2.37 | 2.0035 | 1.9659 | 1.957 | 2.03 | 2.03 | 1.939 | 2.038
| Mean |  | 2.2233 | 2.0496 | 1.9693 | 2.024 | 1.97 | 1.9783 | 2.02 | 1.9908 | 1.84 | 2.0242 | 2.0533 | 2.0695 | 1.8441 | 2.245 | 1.9363 | 2.047 | 1.9392 | 2.0967 | 2.0565 | 1.9343 | 2.0348
| Median |  | 2.2249 | 2.0385 | 1.9674 | 2.0185 | 1.965 | 1.965 | 2.015 | 1.98 | 1.8525 | 2.02 | 2.055 | 2.0681 | 1.8413 | 2.225 | 1.9351 | 2.0105 | 1.97 | 2.085 | 2.0541 | 1.9335 | 2.037
| Std Dev. |  | 0.04514 | 0.033095 | 0.024116 | 0.037524 | 0.026077 | 0.041673 | 0.058652 | 0.032158 | 0.050299 | 0.040301 | 0.017512 | 0.01802 | 0.034737 | 0.065345 | 0.053341 | 0.11633 | 0.073038 | 0.076855 | 0.036545 | 0.015095 | 0.010778
| Rel.Std.Dev. |  | 2.03% | 1.615% | 1.225% | 1.854% | 1.324% | 2.106% | 2.904% | 1.615% | 2.734% | 1.991% | 0.8529% | 0.8707% | 1.884% | 2.911% | 2.755% | 5.683% | 3.766% | 3.666% | 1.777% | 0.7804% | 0.5297%
| PDM3 |  | 11.66% | 2.935% | -1.095% | 1.652% | -1.062% | -0.6435% | 1.449% | -0.01574% | -7.591% | 1.658% | 3.123% | 3.937% | -7.387% | 12.75% | -2.757% | 2.807% | -2.611% | 5.299% | 3.283% | -2.853% | 2.194%
| Z-Score (Absolute) |  | 2.94 | 0.38 | 0.81 | 0 | 0.8 | 0.67 | 0.06 | 0.49 | 2.71 | 0 | 0.43 | 0.67 | 2.65 | 3.26 | 1.29 | 0.34 | 1.25 | 1.07 | 0.48 | 1.32 | 0.16
| Analytical results for Cr in OREAS 556b (Certified Value 113 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | 100 | 102 | 118 | 130 | 91 | 140 | 70 | 140 | 140 | 70 | 77.948 | 100 | 106.5955 | 119 | 120 | 105 | 114 | 109 | 100 | < 100 |  |  | 1
| 1 | 2 | 100 | 102 | 119 | 130 | 116 | 140 | 70 | 140 | 140 | 70 | 70.5083 | 100 | 106.1353 | 118 | 100 | 114 | 215 | 126 | 100 | < 100 |  |  | 37
| 1 | 3 | 100 | 101 | 119 | 130 | 95 | 140 | 70 | 140 | 210 | 140 | 61.2093 | 100 | 108.3772 | 117 | 110 | 115 | 172 | 109 | 100 | < 100 |  |  | 16
| 1 | 4 | 100 | 101 | 119 | 140 | 90 | 140 | 70 | 140 | 210 | 140 | 73.9396 | 100 | 106.0514 | 124 | 110 | 90 | 129 | 108 | 100 | < 100 |  |  | 112.6336
| 1 | 5 | 100 | 103 | 116 | 130 | 92 | 140 | 70 | 140 | 140 | 70 | 64.5955 | 100 | 106.4676 | 119 | 100 | 95 | 143 | 114 | 100 | < 100 |  |  | 42
| 1 | 6 | 100 | 100 | 114 | 130 | 121 | 140 | 70 | 140 | 140 | 70 | 74.1717 | 100 | 109.6619 | 114 | 120 | 90 | 141 | 113 | 100 | < 100
| Mean |  | 100 | 101.5 | 117.5 | 131.6667 | 100.8333 | 140 | 70 | 140 | 163.3333 | 93.3333 | 70.3954 | 100 | 107.2148 | 118.5 | 110 | 101.5 | 152.3333 | 113.1667 | 100
| Median |  | 100 | 101.5 | 118.5 | 130 | 93.5 | 140 | 70 | 140 | 140 | 70 | 72.224 | 100 | 106.5315 | 118.5 | 110 | 100 | 142 | 111 | 100
| Std Dev. |  | 0 | 1.0488 | 2.0736 | 4.0825 | 13.8768 | 0 | 0 | 0 | 36.1478 | 36.1478 | 6.3545 | 0 | 1.4697 | 3.2711 | 8.9443 | 11.4673 | 36.1755 | 6.7355 | 0
| Rel.Std.Dev. |  | 0% | 1.033% | 1.765% | 3.101% | 13.76% | 0% | 0% | 0% | 22.13% | 38.73% | 9.027% | 0% | 1.371% | 2.76% | 8.131% | 11.3% | 23.75% | 5.952% | 0%
| PDM3 |  | -11.22% | -9.885% | 4.321% | 16.9% | -10.48% | 24.3% | -37.85% | 24.3% | 45.01% | -17.14% | -37.5% | -11.22% | -4.811% | 5.208% | -2.338% | -9.885% | 35.25% | 0.4733% | -11.22%
| Z-Score (Absolute) |  | NA | 0.35 | 0.44 | 1.14 | 0.38 | 1.4 | 1.99 | 1.4 | 2.52 | 0.86 | 1.89 | 0.43 | 0.07 | 0.49 | 0.07 | 0.35 | 2.16 | 0.23 | NA | 2.9
Results from laboratories 2 and 27 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 556b (Certified Value 1.86 ppm) | [23 empty] | Certified
Cs | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 17; 06: 18; 08: 19; 09: 23; 10: 25; 11: 26; 12: 28 |  | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*MS; 10: PF*OES; 11: PF*MS; 12: PF*OES |  | 2
3
| 1 | 1 | 01: 1.9; 02: 1.9; 03: 1.8; 04: 1.7064; 05: 2; 06: 2.0774; 08: < 2; 09: 1.8; 10: 2.1; 11: 1.9; 12: 1.8 |  | 1
| 1 | 2 | 01: 1.9; 02: 1.9; 03: 2; 04: 1.82; 05: 2; 06: 2.2229; 08: < 2; 09: 1.9; 10: 2.4; 11: 1.7; 12: 1.9 |  | 9
| 1 | 3 | 01: 1.9; 02: 2.1; 03: 1.7; 04: 1.6804; 05: 2; 06: 2.266; 08: < 2; 09: 1.8; 10: 2.5; 11: 1.8; 12: 1.8 |  | 16
| 1 | 4 | 01: 1.8; 02: 1.8; 03: 1.9; 04: 1.6435; 05: 2; 06: 2.1094; 08: < 2; 09: 1.8; 10: 2.4; 11: 2; 12: 1.8 |  | 1.8586
| 1 | 5 | 01: 2; 02: 2; 03: 1.8; 04: 1.7255; 05: 2.1; 06: 2.1326; 08: < 2; 09: 1.7; 10: 2.1; 11: 1.7; 12: 1.8 |  | 43
| 1 | 6 | 01: 1.8; 02: 1.9; 03: 1.9; 04: 1.8359; 05: 2.1; 06: 2.1994; 08: < 2; 09: 1.9; 10: 2.2; 11: 1.7; 12: 1.8
| Mean |  | 01: 1.8833; 02: 1.9333; 03: 1.85; 04: 1.7353; 05: 2.0333; 06: 2.168; 09: 1.8167; 10: 2.2833; 11: 1.8; 12: 1.8167
| Median |  | 01: 1.9; 02: 1.9; 03: 1.85; 04: 1.716; 05: 2; 06: 2.166; 09: 1.8; 10: 2.3; 11: 1.75; 12: 1.8
| Std Dev. |  | 01: 0.075277; 02: 0.10328; 03: 0.10488; 04: 0.077037; 05: 0.05164; 06: 0.072796; 09: 0.075277; 10: 0.17224; 11: 0.12649; 12: 0.040825
| Rel.Std.Dev. |  | 01: 3.997%; 02: 5.342%; 03: 5.669%; 04: 4.439%; 05: 2.54%; 06: 3.358%; 09: 4.144%; 10: 7.543%; 11: 7.027%; 12: 2.247%
| PDM3 |  | 01: 1.332%; 02: 4.022%; 03: -0.4615%; 04: -6.634%; 05: 9.403%; 06: 16.65%; 09: -2.255%; 10: 22.85%; 11: -3.152%; 12: -2.255%
| Z-Score (Absolute) |  | 01: 0.27; 02: 0.67; 03: 0; 04: 0.93; 05: 1.48; 06: 2.57; 08: 6.88; 09: 0.27; 10: 3.51; 11: 0.4; 12: 0.27
| Analytical results for Cu in OREAS 556b (Certified Value 8.63 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 21 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 8.7531 | 8.2885 | 8.7719 | 9.0954 | 8.05 | 8.29 | 8.68 | 8.5 | 8.71 | 8.51 | 8.63 | 8.3817 | 8.0973 | 9.07 | 8.1886 | 8.746 | 8.69 | 8.4614 | 8.155 | 8.681 |  |  | 1
| 1 | 2 | 8.8324 | 8.23 | 8.7934 | 9.0741 | 8.68 | 8.38 | 8.69 | 8.47 | 8.76 | 8.71 | 8.95 | 8.6275 | 8.0333 | 8.97 | 8.1875 | 8.767 | 9.24 | 8.607 | 8.161 | 8.603 |  |  | 39
| 1 | 3 | 8.4557 | 8.3429 | 8.7341 | 8.8097 | 8.77 | 8.32 | 8.62 | 8.56 | 8.52 | 8.47 | 8.88 | 8.3859 | 8.0373 | 9.1 | 7.9803 | 8.712 | 8.99 | 8.2643 | 8.163 | 8.755 |  |  | 16
| 1 | 4 | 8.8359 | 8.447 | 8.7458 | 9.0406 | 8.73 | 8.32 | 8.69 | 8.65 | 8.59 | 8.22 | 8.81 | 8.7648 | 7.9461 | 9.03 | 8.3971 | 8.747 | 9.2 | 8.6849 | 8.108 | 8.64 |  |  | 8.6302
| 1 | 5 | 8.6874 | 8.2975 | 8.725 | 9.0184 | 8.79 | 8.21 | 8.64 | 8.57 | 8.66 | 8.54 | 8.88 | 8.7191 | 7.9412 | 9.08 | 8.2558 | 8.839 | 8.77 | 8.6094 | 8.23 | 8.679 |  |  | 44
| 1 | 6 | 8.7371 | 8.3749 | 8.7772 | 8.9631 | 8.94 | 8.26 | 8.66 | 8.54 | 8.5 | 8.27 | 8.85 | 8.9051 | 8.0204 | 9.04 | 8.1421 | 8.712 | 8.65 | 8.8149 | 8.349 | 8.551
| Mean |  | 8.7169 | 8.3301 | 8.7579 | 9.0002 | 8.66 | 8.2967 | 8.6633 | 8.5483 | 8.6233 | 8.4533 | 8.8333 | 8.6307 | 8.0126 | 9.0483 | 8.1919 | 8.7538 | 8.9233 | 8.5737 | 8.1943 | 8.6515
| Median |  | 8.7451 | 8.3202 | 8.7589 | 9.0295 | 8.75 | 8.305 | 8.67 | 8.55 | 8.625 | 8.49 | 8.865 | 8.6733 | 8.0269 | 9.055 | 8.1881 | 8.7465 | 8.88 | 8.6082 | 8.162 | 8.6595
| Std Dev. |  | 0.14024 | 0.075697 | 0.026922 | 0.10407 | 0.31138 | 0.058195 | 0.028752 | 0.062423 | 0.10443 | 0.18162 | 0.10967 | 0.21119 | 0.059657 | 0.046224 | 0.13673 | 0.046978 | 0.25843 | 0.1904 | 0.085191 | 0.070602
| Rel.Std.Dev. |  | 1.609% | 0.9087% | 0.3074% | 1.156% | 3.596% | 0.7014% | 0.3319% | 0.7302% | 1.211% | 2.149% | 1.242% | 2.447% | 0.7445% | 0.5109% | 1.669% | 0.5367% | 2.896% | 2.221% | 1.04% | 0.8161%
| PDM3 |  | 1.005% | -3.477% | 1.48% | 4.288% | 0.3457% | -3.864% | 0.3844% | -0.9482% | -0.07912% | -2.049% | 2.354% | 0.006072% | -7.156% | 4.845% | -5.078% | 1.433% | 3.397% | -0.6548% | -5.05% | 0.2473%
| Z-Score (Absolute) |  | 0.34 | 1.38 | 0.52 | 1.59 | 0.08 | 1.53 | 0.1 | 0.41 | 0.08 | 0.83 | 0.85 | 0.05 | 2.78 | 1.8 | 1.99 | 0.5 | 1.25 | 0.3 | 1.98 | 0.05
| Analytical results for Dy in OREAS 556b (Certified Value 6.28 ppm) | [23 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab |  |  | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 |  |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES |  |  | 2
3
| 1 | 1 | 01: 5.3; 02: 6.3; 03: 5.7909; 04: 6; 05: 6.48; 06: 5.86; 08: 6.27 |  |  | 1
| 1 | 2 | 01: 5.1; 02: 6.36; 03: 5.9841; 04: 6; 05: 6.21; 06: 6.6; 08: 6.47 |  |  | 10
| 1 | 3 | 01: 5.5; 02: 6.26; 03: 6.2555; 04: 6; 05: 6.84; 06: 6.34; 08: 6.54 |  |  | 16
| 1 | 4 | 01: 5.7; 02: 6.57; 03: 6.0484; 04: 6; 05: 6.38; 06: 6.36; 08: 6.39 |  |  | 6.278
| 1 | 5 | 01: 5.7; 02: 6.44; 03: 5.8806; 04: 9; 05: 6.4; 06: 5.79; 08: 6.48 |  |  | 45
| 1 | 6 | 01: 5.5; 02: 6.42; 03: 6.1001; 04: 7; 05: 6.6; 06: 5.54; 08: 6.38
| Mean |  | 01: 5.4667; 02: 6.3917; 03: 6.0099; 04: 6.6667; 05: 6.485; 06: 6.0817; 08: 6.4217
| Median |  | 01: 5.5; 02: 6.39; 03: 6.0162; 04: 6; 05: 6.44; 06: 6.1; 08: 6.43
| Std Dev. |  | 01: 0.23381; 02: 0.11107; 03: 0.16449; 04: 1.2111; 05: 0.21594; 06: 0.41; 08: 0.095376
| Rel.Std.Dev. |  | 01: 4.277%; 02: 1.738%; 03: 2.737%; 04: 18.17%; 05: 3.33%; 06: 6.742%; 08: 1.485%
| PDM3 |  | 01: -12.92%; 02: 1.811%; 03: -4.27%; 04: 6.191%; 05: 3.297%; 06: -3.127%; 08: 2.289%
| Z-Score (Absolute) |  | 01: 2.52; 02: 0.51; 03: 0.74; 04: NA; 05: 0.81; 06: 0.51; 08: 0.61
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 556b (Certified Value 3.47 ppm) | [23 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab |  |  | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 |  |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES |  |  | 2
3
| 1 | 1 | 01: 3.4; 02: 3.53; 03: 3.3218; 04: 4; 05: 3.25; 06: 3.28; 08: 3.6 |  |  | 1
| 1 | 2 | 01: 3.2; 02: 3.64; 03: 3.2966; 04: 3; 05: 3.25; 06: 3.91; 08: 3.8 |  |  | 11
| 1 | 3 | 01: 3.1; 02: 3.55; 03: 3.3571; 04: 3; 05: 3.28; 06: 3.71; 08: 3.71 |  |  | 16
| 1 | 4 | 01: 3.3; 02: 3.59; 03: 3.2954; 04: 3; 05: 3.44; 06: 4.06; 08: 3.62 |  |  | 3.4664
| 1 | 5 | 01: 3.5; 02: 3.61; 03: 3.384; 04: 5; 05: 3.34; 06: 3.34; 08: 3.76 |  |  | 46
| 1 | 6 | 01: 3.4; 02: 3.63; 03: 3.3647; 04: 4; 05: 3.19; 06: 3.13; 08: 3.65
| Mean |  | 01: 3.3167; 02: 3.5917; 03: 3.3366; 04: 3.6667; 05: 3.2917; 06: 3.5717; 08: 3.69
| Median |  | 01: 3.35; 02: 3.6; 03: 3.3394; 04: 3.5; 05: 3.265; 06: 3.525; 08: 3.68
| Std Dev. |  | 01: 0.1472; 02: 0.044008; 03: 0.037332; 04: 0.8165; 05: 0.087502; 06: 0.37573; 08: 0.08
| Rel.Std.Dev. |  | 01: 4.438%; 02: 1.225%; 03: 1.119%; 04: 22.27%; 05: 2.658%; 06: 10.52%; 08: 2.168%
| PDM3 |  | 01: -4.319%; 02: 3.615%; 03: -3.744%; 04: 5.778%; 05: -5.04%; 06: 3.038%; 08: 6.451%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67; 03: 0.58; 04: NA; 05: 0.8; 06: 0.58; 08: 1.16
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 556b (Certified Value 0.88 ppm) | [23 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab |  |  | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 |  |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES |  |  | 2
2
| 1 | 1 | 01: 0.8; 02: 0.88; 03: 0.95234; 04: < 2; 05: 0.89; 06: 0.53; 08: 0.96 |  |  | 1
| 1 | 2 | 01: 0.8; 02: 0.89; 03: 0.99016; 04: < 2; 05: 0.85; 06: 0.85; 08: 1 |  |  | 12
| 1 | 3 | 01: 0.9; 02: 0.9; 03: 0.9595; 04: < 2; 05: 0.87; 06: 0.73; 08: 0.98 |  |  | 16
| 1 | 4 | 01: 0.9; 02: 0.97; 03: 0.96509; 04: < 2; 05: 0.91; 06: 0.7; 08: 1 |  |  | 0.87762
| 1 | 5 | 01: 0.9; 02: 0.92; 03: 0.98384; 04: < 2; 05: 0.87; 06: 0.59; 08: 0.97 |  |  | 47
| 1 | 6 | 01: 0.9; 02: 0.88; 03: 0.96368; 04: < 2; 05: 0.83; 06: 0.6; 08: 1.02
| Mean |  | 01: 0.86667; 02: 0.90667; 03: 0.9691; 05: 0.87; 06: 0.66667; 08: 0.98833
| Median |  | 01: 0.9; 02: 0.895; 03: 0.96439; 05: 0.87; 06: 0.65; 08: 0.99
| Std Dev. |  | 01: 0.05164; 02: 0.034448; 03: 0.014691; 05: 0.028284; 06: 0.11639; 08: 0.022286
| Rel.Std.Dev. |  | 01: 5.958%; 02: 3.799%; 03: 1.516%; 05: 3.251%; 06: 17.46%; 08: 2.255%
| PDM3 |  | 01: -1.248%; 02: 3.31%; 03: 10.42%; 05: -0.8683%; 06: -24.04%; 08: 12.62%
| Z-Score (Absolute) |  | 01: NA; 02: 0.37; 03: 0.37; 04: 0.74; 05: 0.81; 06: 3.25; 08: 0.6
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 556b (Certified Value 6.73 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 6.44 | 6.62 | 6.58 | 7.08 | 6.77 | 7.02 | 6.652 | 6.652 | 6.841 | 6.673 | 6.82 | 7.1322 | 6.7472 | 7.93 | 6.5514 | 6.5248 | 6.69 | 7.43 | 7.19 | 6.51 | 6.32 | 6.75 | 1
| 1 | 2 | 6.57 | 6.66 | 6.59 | 7.08 | 6.82 | 7.03 | 6.743 | 6.792 | 6.994 | 6.771 | 6.883 | 6.8405 | 6.7485 | 7.83 | 6.4391 | 6.6363 | 6.7 | 8.29 | 8.98 | 6.66 | 6.32 | 6.68 | 0
| 1 | 3 | 6.33 | 6.66 | 6.54 | 6.91 | 6.79 | 6.99 | 6.603 | 6.659 | 6.757 | 6.743 | 6.869 | 6.9115 | 6.7861 | 7.83 | 6.5522 | 6.5086 | 6.79 | 7.22 | 8.37 | 6.44 | 6.35 | 6.77 | 16
| 1 | 4 | 6.38 | 6.63 | 6.56 | 7.11 | 6.77 | 7.04 | 6.708 | 6.743 | 6.89 | 6.477 | 6.813 | 7.1922 | 6.7793 | 7.93 | 6.589 | 6.7446 | 6.68 | 7.45 | 7.72 | 6.75 | 6.31 | 6.87 | 6.7301
| 1 | 5 | 6.39 | 6.69 | 6.56 | 7.02 | 7.12 | 7.06 | 6.631 | 6.757 | 6.89 | 6.673 | 6.841 | 7.1131 | 6.7955 | 7.94 | 6.8192 | 6.6684 | 6.78 | 7.16 | 7.91 | 6.71 | 6.38 | 6.82 | 48
| 1 | 6 | 6.24 | 6.74 | 6.56 | 6.95 | 7.1 | 6.98 | 6.652 | 6.729 | 6.771 | 6.533 | 6.939 | 7.322 | 6.7106 | 7.75 | 6.6811 | 6.508 | 6.74 | 7.66 | 7.86 | 6.83 | 6.5 | 6.71
| Mean |  | 6.3917 | 6.6667 | 6.565 | 7.025 | 6.895 | 7.02 | 6.6648 | 6.722 | 6.8572 | 6.645 | 6.8608 | 7.0852 | 6.7612 | 7.8683 | 6.6053 | 6.5984 | 6.73 | 7.535 | 8.005 | 6.65 | 6.3633 | 6.7667
| Median |  | 6.385 | 6.66 | 6.56 | 7.05 | 6.805 | 7.025 | 6.652 | 6.736 | 6.8655 | 6.673 | 6.855 | 7.1227 | 6.7639 | 7.88 | 6.5706 | 6.5805 | 6.72 | 7.44 | 7.885 | 6.685 | 6.335 | 6.76
| Std Dev. |  | 0.11053 | 0.043665 | 0.017607 | 0.080187 | 0.16766 | 0.030332 | 0.051519 | 0.055649 | 0.087862 | 0.11646 | 0.04694 | 0.17921 | 0.031755 | 0.07705 | 0.13048 | 0.099362 | 0.047329 | 0.41079 | 0.60981 | 0.14819 | 0.07174 | 0.070048
| Rel.Std.Dev. |  | 1.729% | 0.655% | 0.2682% | 1.141% | 2.432% | 0.4321% | 0.773% | 0.8279% | 1.281% | 1.753% | 0.6842% | 2.529% | 0.4697% | 0.9792% | 1.975% | 1.506% | 0.7032% | 5.452% | 7.618% | 2.228% | 1.127% | 1.035%
| PDM3 |  | -5.029% | -0.9425% | -2.453% | 4.382% | 2.45% | 4.308% | -0.9698% | -0.1204% | 1.888% | -1.264% | 1.943% | 5.277% | 0.4621% | 16.91% | -1.854% | -1.956% | -0.001493% | 11.96% | 18.94% | -1.19% | -5.45% | 0.5433%
| Z-Score (Absolute) |  | 1.66 | 0.37 | 0.85 | 1.31 | 0.7 | 1.29 | 0.38 | 0.11 | 0.52 | 0.47 | 0.54 | 1.59 | 0.07 | 5.27 | 0.66 | 0.69 | 0.07 | 3.7 | 5.91 | 0.45 | 1.79 | 0.1
| Analytical results for Ga in OREAS 556b (Certified Value 28.7 ppm) | [23 empty] | Certified
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 04; 03: 05; 04: 06; 05: 15; 06: 17; 08: 18; 09: 19; 10: 23; 11: 25; 12: 26; 15: 28 | ppm
|  |  | 01: PF*OES; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*OES/MS; 08: PF*OES/MS; 09: PF*MS; 10: PF*MS; 11: PF*OES; 12: PF*MS; 15: PF*OES | 1
2
| 1 | 1 | 01: 32.7; 02: 29; 03: 29.1; 04: 28; 05: 30.4235; 06: 30.7; 08: 27.9512; 09: 25; 10: 29; 11: 25; 12: 30; 15: 22 | 1
| 1 | 2 | 01: 33.8; 02: 29; 03: 27.1; 04: 29; 05: 29.6763; 06: 30; 08: 29.1064; 09: 25; 10: 28; 11: 32; 12: 30; 15: 23 | 13
| 1 | 3 | 01: 32.2; 02: 31; 03: 29; 04: 27.5; 05: 28.935; 06: 30.4; 08: 29.4953; 09: 25; 10: 29; 11: 32; 12: 30; 15: 21 | 16
| 1 | 4 | 01: 34.3; 02: 28; 03: 27; 04: 28.2; 05: 30.9361; 06: 31.6; 08: 28.7787; 09: 27; 10: 28; 11: 28; 12: 29; 15: 20 | 28.6577
| 1 | 5 | 01: 33.4; 02: 29; 03: 26.6; 04: 28.2; 05: 28.9845; 06: 29.3; 08: 30.2227; 09: 25; 10: 30; 11: 26; 12: 30; 15: 20 | 49
| 1 | 6 | 01: 33.4; 02: 28; 03: 28.2; 04: 28.3; 05: 28.3612; 06: 32.8; 08: 29.592; 09: 26; 10: 28; 11: 24; 12: 31; 15: 22
| Mean |  | 01: 33.3; 02: 29; 03: 27.8333; 04: 28.2; 05: 29.5527; 06: 30.8; 08: 29.191; 09: 25.5; 10: 28.6667; 11: 27.8333; 12: 30; 15: 21.3333
| Median |  | 01: 33.4; 02: 29; 03: 27.65; 04: 28.2; 05: 29.3304; 06: 30.55; 08: 29.3009; 09: 25; 10: 28.5; 11: 27; 12: 30; 15: 21.5
| Std Dev. |  | 01: 0.75366; 02: 1.0954; 03: 1.082; 04: 0.4858; 05: 0.98096; 06: 1.241; 08: 0.77874; 09: 0.83666; 10: 0.8165; 11: 3.4881; 12: 0.63246; 15: 1.2111
| Rel.Std.Dev. |  | 01: 2.263%; 02: 3.777%; 03: 3.887%; 04: 1.723%; 05: 3.319%; 06: 4.029%; 08: 2.668%; 09: 3.281%; 10: 2.848%; 11: 12.53%; 12: 2.108%; 15: 5.677%
| PDM3 |  | 01: 16.2%; 02: 1.194%; 03: -2.877%; 04: -1.597%; 05: 3.123%; 06: 7.475%; 08: 1.861%; 09: -11.02%; 10: 0.03124%; 11: -2.877%; 12: 4.684%; 15: -25.56%
| Z-Score (Absolute) |  | 01: 3.01; 02: 0.11; 03: 0.67; 04: 0.43; 05: 0.49; 06: 1.33; 08: 0.24; 09: 2.25; 10: 0.11; 11: 0.67; 12: 0.79; 15: 5.06
| Analytical results for Gd in OREAS 556b (Certified Value 6.9 ppm) | [23 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES | 2
3
| 1 | 1 | 01: 6.6; 02: 6.3; 03: 7.0159; 04: 7; 05: 6.76; 06: 6.88; 08: 7 | 1
| 1 | 2 | 01: 6.6; 02: 6.63; 03: 7.1365; 04: 6; 05: 6.37; 06: 8.98; 08: 7 | 14
| 1 | 3 | 01: 6.5; 02: 6.7; 03: 7.0453; 04: 6; 05: 6.68; 06: 7.67; 08: 8 | 16
| 1 | 4 | 01: 6.9; 02: 6.75; 03: 7.2407; 04: 6; 05: 6.88; 06: 7.65; 08: 7 | 6.8993
| 1 | 5 | 01: 6.8; 02: 6.73; 03: 7.0806; 04: 9; 05: 6.87; 06: 7.17; 08: 8 | 50
| 1 | 6 | 01: 6.7; 02: 6.51; 03: 7.4295; 04: 7; 05: 6.84; 06: 6.86; 08: 9
| Mean |  | 01: 6.6833; 02: 6.6033; 03: 7.1581; 04: 6.8333; 05: 6.7333; 06: 7.535; 08: 7.6667
| Median |  | 01: 6.65; 02: 6.665; 03: 7.1085; 04: 6.5; 05: 6.8; 06: 7.41; 08: 7.5
| Std Dev. |  | 01: 0.1472; 02: 0.17224; 03: 0.15486; 04: 1.169; 05: 0.19346; 06: 0.79213; 08: 0.8165
| Rel.Std.Dev. |  | 01: 2.202%; 02: 2.608%; 03: 2.163%; 04: 17.11%; 05: 2.873%; 06: 10.51%; 08: 10.65%
| PDM3 |  | 01: -3.131%; 02: -4.29%; 03: 3.75%; 04: -0.9569%; 05: -2.406%; 06: 9.213%; 08: 11.12%
| Z-Score (Absolute) |  | 01: 0.26; 02: 0.67; 03: 2.2; 04: NA; 05: 0; 06: 4.16; 08: NA
Results from laboratories 19 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 556b (Certified Value 2.8 ppm) | [23 empty] | Certified
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 17; 06: 18; 08: 19; 09: 23; 10: 25; 11: 26 | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*MS; 10: PF*OES; 11: PF*MS | 2
2
| 1 | 1 | 01: 3; 02: 3.2; 03: 3; 04: 2.2757; 05: 2.7; 06: 2.2242; 08: NA; 09: 3; 10: < 5; 11: 3 | 1
| 1 | 2 | 01: 3; 02: 3.4; 03: 3; 04: 2.359; 05: 2.6; 06: 2.3606; 08: NA; 09: 3; 10: < 5; 11: 3 | 15
| 1 | 3 | 01: 3; 02: 3.3; 03: 3; 04: 2.7679; 05: 2.7; 06: 2.3217; 08: NA; 09: 3; 10: < 5; 11: 3 | 16
| 1 | 4 | 01: 2; 02: 3.4; 03: 3; 04: 2.6779; 05: 2.6; 06: 2.2877; 08: NA; 09: 2; 10: < 5; 11: 3 | 2.8039
| 1 | 5 | 01: 3; 02: 3.3; 03: 3; 04: 2.7976; 05: 2.9; 06: 2.4664; 08: NA; 09: 2; 10: < 5; 11: 3 | 51
| 1 | 6 | 01: 3; 02: 2.8; 03: 3; 04: 2.4725; 05: 2.7; 06: 2.4568; 08: NA; 09: 3; 10: < 5; 11: 3
| Mean |  | 01: 2.8333; 02: 3.2333; 03: 3; 04: 2.5584; 05: 2.7; 06: 2.3529; 09: 2.6667; 11: 3
| Median |  | 01: 3; 02: 3.3; 03: 3; 04: 2.5752; 05: 2.7; 06: 2.3411; 09: 3; 11: 3
| Std Dev. |  | 01: 0.40825; 02: 0.22509; 03: 0; 04: 0.2202; 05: 0.10954; 06: 0.095416; 09: 0.5164; 11: 0
| Rel.Std.Dev. |  | 01: 14.41%; 02: 6.962%; 03: 0%; 04: 8.607%; 05: 4.057%; 06: 4.055%; 09: 19.36%; 11: 0%
| PDM3 |  | 01: 1.049%; 02: 15.32%; 03: 6.993%; 04: -8.755%; 05: -3.706%; 06: -16.09%; 09: -4.895%; 11: 6.993%
| Z-Score (Absolute) |  | 01: 0.45; 02: 1.8; 03: 1.01; 04: 0.48; 05: 0; 06: 1.17; 08: NA; 09: 0.11; 10: 0.67; 11: 1.01
| Analytical results for Hf in OREAS 556b (Indicative Value 6.41 ppm) | [23 empty] | Indicative
Hf | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 25 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES | 2
2
| 1 | 1 | 01: < 10; 02: 6; 03: 6.0317; 04: 6; 05: 7 | 1
| 1 | 2 | 01: 10; 02: 6; 03: 6.0923; 04: 6; 05: 9 | 19
| 1 | 3 | 01: < 10; 02: 5; 03: 6.0675; 04: 6; 05: 8 | 16
| 1 | 4 | 01: < 10; 02: 5; 03: 6.0743; 04: 6; 05: 8 | 6.4137
| 1 | 5 | 01: 10; 02: 6; 03: 6.0776; 04: 6; 05: 8 | 25
| 1 | 6 | 01: < 10; 02: 5; 03: 6.318; 04: 6; 05: 7
| Mean |  | 01: 10; 02: 5.5; 03: 6.1102; 04: 6; 05: 7.8333
| Median |  | 01: 10; 02: 5.5; 03: 6.076; 04: 6; 05: 8
| Std Dev. |  | 01: 0; 02: 0.54772; 03: 0.10376; 04: 0; 05: 0.75277
| Rel.Std.Dev. |  | 01: 0%; 02: 9.959%; 03: 1.698%; 04: 0%; 05: 9.61%
| PDM3 |  | 01: 55.92%; 02: -14.25%; 03: -4.732%; 04: -6.451%; 05: 22.13%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.74; 03: 0; 04: 0.13; 05: 2.09
| Analytical results for Hg in OREAS 556b (Indicative Value 0.083 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 01; 02: 25 | ppm
|  |  | 01: PF*OES; 02: PF*OES | 3
3
| 1 | 1 | 01: < 5; 02: 0.1 | 1
| 1 | 2 | 01: < 5; 02: 0.1 | 20
| 1 | 3 | 01: < 5; 02: 0.1 | 16
| 1 | 4 | 01: < 5; 02: < 0.1 | 0.083333
| 1 | 5 | 01: < 5; 02: < 0.1 | 26
| 1 | 6 | 01: < 5; 02: 0.1
| Mean |  | 02: 0.1
| Median |  | 02: 0.1
| Std Dev. |  | 02: 0
Rel.Std.Dev. | bar: 02
| PDM3 |  | 02: 20%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Ho in OREAS 556b (Certified Value 1.22 ppm) | [23 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES | 2
3
| 1 | 1 | 01: 1.1; 02: 1.26; 03: 1.178; 04: < 2; 05: 1.1; 06: 1.25; 08: 1.26 | 1
| 1 | 2 | 01: 1.1; 02: 1.26; 03: 1.1625; 04: < 2; 05: 1.07; 06: 1.48; 08: 1.28 | 3
| 1 | 3 | 01: 1.2; 02: 1.19; 03: 1.2482; 04: < 2; 05: 1.24; 06: 1.31; 08: 1.29 | 16
| 1 | 4 | 01: 1.2; 02: 1.27; 03: 1.1825; 04: < 2; 05: 1.12; 06: 1.46; 08: 1.26 | 1.2166
| 1 | 5 | 01: 1.2; 02: 1.24; 03: 1.1914; 04: < 2; 05: 1.11; 06: 1.33; 08: 1.28 | 52
| 1 | 6 | 01: 1.1; 02: 1.22; 03: 1.2014; 04: < 2; 05: 1.12; 06: 1.21; 08: 1.26
| Mean |  | 01: 1.15; 02: 1.24; 03: 1.194; 05: 1.1267; 06: 1.34; 08: 1.2717
| Median |  | 01: 1.15; 02: 1.25; 03: 1.187; 05: 1.115; 06: 1.32; 08: 1.27
| Std Dev. |  | 01: 0.054772; 02: 0.030332; 03: 0.029618; 05: 0.058538; 06: 0.10954; 08: 0.013292
| Rel.Std.Dev. |  | 01: 4.763%; 02: 2.446%; 03: 2.481%; 05: 5.196%; 06: 8.175%; 08: 1.045%
| PDM3 |  | 01: -5.475%; 02: 1.922%; 03: -1.859%; 05: -7.393%; 06: 10.14%; 08: 4.525%
| Z-Score (Absolute) |  | 01: 0.44; 02: 0.46; 03: 0; 04: 1.94; 05: 0.67; 06: 1.46; 08: 0.78
| Analytical results for In in OREAS 556b (Certified Value 3.52 ppm) | [23 empty] | Certified
In | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 17; 06: 19; 08: 23; 09: 25; 10: 26 | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*MS; 09: PF*OES; 10: PF*MS | 2
3
| 1 | 1 | 01: 3.6; 02: 3.6; 03: 3.4; 04: 3.5814; 05: 3.93; 06: 3.6; 08: 3.1; 09: 4; 10: 3.4 | 1
| 1 | 2 | 01: 3.8; 02: 3.2; 03: 3.4; 04: 3.6449; 05: 3.79; 06: 3.5; 08: 3.2; 09: 4.5; 10: 3.5 | 4
| 1 | 3 | 01: 3.5; 02: 3.6; 03: 3.4; 04: 3.5787; 05: 3.72; 06: 3.4; 08: 3.3; 09: 4.2; 10: 3.6 | 16
| 1 | 4 | 01: 3.6; 02: 3.6; 03: 3.5; 04: 3.6797; 05: 3.86; 06: 3.6; 08: 3.2; 09: 4.1; 10: 3.4 | 3.5174
| 1 | 5 | 01: 3.7; 02: 3.2; 03: 3.4; 04: 3.8179; 05: 3.89; 06: 3.3; 08: 3.1; 09: 3.9; 10: 3.5 | 53
| 1 | 6 | 01: 3.5; 02: 3.5; 03: 3.5; 04: 3.8431; 05: 3.8; 06: 3.5; 08: 3.1; 09: 4; 10: 3.4
| Mean |  | 01: 3.6167; 02: 3.45; 03: 3.4333; 04: 3.6909; 05: 3.8317; 06: 3.4833; 08: 3.1667; 09: 4.1167; 10: 3.4667
| Median |  | 01: 3.6; 02: 3.55; 03: 3.4; 04: 3.6623; 05: 3.83; 06: 3.5; 08: 3.15; 09: 4.05; 10: 3.45
| Std Dev. |  | 01: 0.1169; 02: 0.19748; 03: 0.05164; 04: 0.11499; 05: 0.076267; 06: 0.1169; 08: 0.08165; 09: 0.2137; 10: 0.08165
| Rel.Std.Dev. |  | 01: 3.232%; 02: 5.724%; 03: 1.504%; 04: 3.115%; 05: 1.99%; 06: 3.356%; 08: 2.578%; 09: 5.191%; 10: 2.355%
| PDM3 |  | 01: 2.822%; 02: -1.916%; 03: -2.39%; 04: 4.934%; 05: 8.934%; 06: -0.9688%; 08: -9.972%; 09: 17.04%; 10: -1.443%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.17; 03: 0.25; 04: 1.05; 05: 1.76; 06: 0; 08: 1.6; 09: 3.2; 10: 0.08
| Analytical results for K in OREAS 556b (Certified Value 2.27 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 2.17 | 2.99 | 2.3 | 2.43 | 2.3 | 2.3 | 2.291 | 2.3 | 2.217 | 2.258 | 2.3 | 2.368 | 3.2577 | 2.74 | 2.1235 | 2.1 | 3.1 | 2.66 | 2.71 | 2.2 | 2.3 |  | 1
| 1 | 2 | 2.2 | 2.92 | 2.32 | 2.45 | 2.3 | 2.3 | 2.283 | 2.241 | 2.2 | 2.308 | 2.308 | 2.3212 | 3.2666 | 2.8 | 2.1725 | 2 | 3.1 | 3.83 | 2.74 | 2.22 | 2.3 |  | 0
| 1 | 3 | 2.13 | 2.99 | 2.3 | 2.51 | 2.3 | 2.29 | 2.283 | 2.308 | 2.167 | 2.291 | 2.308 | 2.3502 | 3.2456 | 2.74 | 2.1359 | 2.2 | 3.2 | 3.79 | 2.65 | 2.27 | 2.3 |  | 16
| 1 | 4 | 2.16 | 3.1 | 2.3 | 2.39 | 2.3 | 2.31 | 2.291 | 2.358 | 2.175 | 2.208 | 2.316 | 2.3488 | 3.2512 | 2.66 | 2.2367 | 2.1 | 2.8 | 2.97 | 2.79 | 2.29 | 2.3 |  | 2.2696
| 1 | 5 | 2.16 | 2.96 | 2.31 | 2.37 | 2.4 | 2.33 | 2.283 | 2.341 | 2.217 | 2.266 | 2.333 | 2.3316 | 3.2795 | 2.76 | 2.1911 | 2.1 | 3 | 3.42 | 2.76 | 2.23 | 2.3 |  | 54
| 1 | 6 | 2.11 | 2.95 | 2.33 | 2.34 | 2.4 | 2.29 | 2.291 | 2.3 | 2.1 | 2.25 | 2.316 | 2.3742 | 3.2531 | 2.76 | 2.1323 | 2.1 | 3.2 | 3.41 | 2.81 | 2.24 | 2.3
| Mean |  | 2.155 | 2.985 | 2.31 | 2.415 | 2.3333 | 2.3033 | 2.287 | 2.308 | 2.1793 | 2.2635 | 2.3135 | 2.349 | 3.259 | 2.7433 | 2.1653 | 2.1 | 3.0667 | 3.3467 | 2.7433 | 2.2417 | 2.3
| Median |  | 2.16 | 2.975 | 2.305 | 2.41 | 2.3 | 2.3 | 2.287 | 2.304 | 2.1875 | 2.262 | 2.312 | 2.3495 | 3.2554 | 2.75 | 2.1542 | 2.1 | 3.1 | 3.415 | 2.75 | 2.235 | 2.3
| Std Dev. |  | 0.031464 | 0.062209 | 0.012649 | 0.061237 | 0.05164 | 0.015055 | 0.0043818 | 0.040512 | 0.044112 | 0.034743 | 0.011274 | 0.020356 | 0.012287 | 0.046332 | 0.043626 | 0.063246 | 0.15055 | 0.4589 | 0.05785 | 0.033116 | 4.8648e-16
| Rel.Std.Dev. |  | 1.46% | 2.084% | 0.5476% | 2.536% | 2.213% | 0.6536% | 0.1916% | 1.755% | 2.024% | 1.535% | 0.4873% | 0.8666% | 0.377% | 1.689% | 2.015% | 3.012% | 4.909% | 13.71% | 2.109% | 1.477% | 2.115e-14%
| PDM3 |  | -5.048% | 31.52% | 1.781% | 6.408% | 2.809% | 1.488% | 0.7678% | 1.693% | -3.976% | -0.2676% | 1.935% | 3.5% | 43.59% | 20.87% | -4.593% | -7.472% | 35.12% | 47.46% | 20.87% | -1.23% | 1.341%
| Z-Score (Absolute) |  | 1 | 4.33 | 0 | 0.67 | 0.15 | 0.04 | 0.15 | 0.01 | 0.84 | 0.3 | 0.02 | 0.25 | 6.09 | 2.78 | 0.93 | 1.35 | 4.86 | 6.66 | 2.78 | 0.44 | 0.06
| Analytical results for La in OREAS 556b (Certified Value 75 ppm) | [23 empty] | Certified
La | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 05; 05: 06; 06: 18; 08: 19; 09: 21; 10: 23; 11: 25; 12: 28 |  | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*OES; 10: PF*MS; 11: PF*OES; 12: PF*OES |  | 0
1
| 1 | 1 | 01: 73.6; 02: 60; 03: 72; 04: 77.7; 05: 76.4; 06: 75.0036; 08: 74; 09: 80; 10: 69.8; 11: 69.2; 12: 79.46 |  | 1
| 1 | 2 | 01: 75; 02: 56; 03: 73; 04: 70.9; 05: 79.2; 06: 75.8158; 08: 73; 09: 70; 10: 67; 11: 81.9; 12: 80.14 |  | 5
| 1 | 3 | 01: 73.1; 02: 57; 03: 73; 04: 77.7; 05: 77.4; 06: 81.3055; 08: 70; 09: 80; 10: 68.9; 11: 77.6; 12: 76.89 |  | 16
| 1 | 4 | 01: 76.8; 02: 57; 03: 74; 04: 77.7; 05: 76.8; 06: 77.8564; 08: 74; 09: 70; 10: 69.6; 11: 79.3; 12: 79.64 |  | 75.2334
| 1 | 5 | 01: 74.4; 02: 56; 03: 74; 04: 76.9; 05: 77.4; 06: 73.9133; 08: 91; 09: 80; 10: 70.7; 11: 69.1; 12: 81.61 |  | 55
| 1 | 6 | 01: 74.3; 02: 64; 03: 74; 04: 76.7; 05: 77; 06: 78.997; 08: 82; 09: 80; 10: 68.1; 11: 68.9; 12: 78.83
| Mean |  | 01: 74.5333; 02: 58.3333; 03: 73.3333; 04: 76.2667; 05: 77.3667; 06: 77.1486; 08: 77.3333; 09: 76.6667; 10: 69.0167; 11: 74.3333; 12: 79.4283
| Median |  | 01: 74.35; 02: 57; 03: 73.5; 04: 77.3; 05: 77.2; 06: 76.8361; 08: 74; 09: 80; 10: 69.25; 11: 73.4; 12: 79.55
| Std Dev. |  | 01: 1.2925; 02: 3.1411; 03: 0.8165; 04: 2.6666; 05: 0.97502; 06: 2.7584; 08: 7.7889; 09: 5.164; 10: 1.3197; 11: 5.9305; 12: 1.5565
| Rel.Std.Dev. |  | 01: 1.734%; 02: 5.385%; 03: 1.113%; 04: 3.496%; 05: 1.26%; 06: 3.575%; 08: 10.07%; 09: 6.736%; 10: 1.912%; 11: 7.978%; 12: 1.96%
| PDM3 |  | 01: -0.9305%; 02: -22.46%; 03: -2.526%; 04: 1.373%; 05: 2.836%; 06: 2.546%; 08: 2.791%; 09: 1.905%; 10: -8.263%; 11: -1.196%; 12: 5.576%
| Z-Score (Absolute) |  | 01: 0.3; 02: 5.9; 03: 0.71; 04: 0.3; 05: 0.68; 06: 0.6; 08: 0.67; 09: NA; 10: 2.21; 11: 0.37; 12: 1.39
Results from laboratory 21 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 556b (Certified Value 51 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 15 | 17 | 19 | 21 | 25 | 26 | 27 | 28 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 0
0
| 1 | 1 | 58 | 46.5 | 48 | 56 | 55 | 52.2564 | 53.3 | 46 | 70 | 44.4 | < 100 | 50 | < 100 | [9 empty] | 1
| 1 | 2 | 58 | 46.7 | 49 | 55 | 55 | 48.9784 | 52.7 | 52 | 70 | 44.6 | < 100 | 50 | < 100 | [9 empty] | 33
| 1 | 3 | 69 | 46.2 | 49 | 56 | 55 | 48.5723 | 54.1 | 47 | 70 | 40.5 | < 100 | 50 | < 100 | [9 empty] | 16
| 1 | 4 | 63 | 46.5 | 51 | 56 | 55 | 52.6548 | 54.7 | 47 | 70 | 40.4 | < 100 | 50 | < 100 | [9 empty] | 51.0031
| 1 | 5 | 57 | 46.6 | 47 | 59 | 55 | 46.9991 | 53.8 | 50 | 60 | 42.6 | < 100 | 50 | < 100 | [9 empty] | 56
| 1 | 6 | 57 | 46.4 | 49 | 57 | 54 | 51.2048 | 54.2 | 48 | 70 | 39.7 | < 100 | 50 | < 100
| Mean |  | 01: 60.3333; 02: 46.4833; 03: 48.8333; 04: 56.5; 05: 54.8333; 06: 50.111; 08: 53.8; 09: 48.3333; 10: 68.3333; 11: 42.0333; 15: 50
| Median |  | 01: 58; 02: 46.5; 03: 49; 04: 56; 05: 55; 06: 50.0916; 08: 53.95; 09: 47.5; 10: 70; 11: 41.55; 15: 50
| Std Dev. |  | 01: 4.8028; 02: 0.17224; 03: 1.3292; 04: 1.3784; 05: 0.40825; 06: 2.2629; 08: 0.70993; 09: 2.2509; 10: 4.0825; 11: 2.1435; 15: 0
| Rel.Std.Dev. |  | 01: 7.96%; 02: 0.3705%; 03: 2.722%; 04: 2.44%; 05: 0.7445%; 06: 4.516%; 08: 1.32%; 09: 4.657%; 10: 5.974%; 11: 5.1%; 15: 0%
| PDM3 |  | 01: 18.29%; 02: -8.862%; 03: -4.254%; 04: 10.78%; 05: 7.51%; 06: -1.749%; 08: 5.484%; 09: -5.234%; 10: 33.98%; 11: -17.59%; 15: -1.967%
| Z-Score (Absolute) |  | 1.98 | 0.67 | 0.22 | 1.25 | 0.93 | 0.02 | 0.73 | 0.32 | NA | 1.53 | 0 | NA | 0
Results from laboratories 21 and 27 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 556b (Certified Value 0.53 ppm) | [23 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 06; 02: 18; 03: 19; 04: 23; 05: 25; 06: 28 |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*MS; 04: PF*MS; 05: PF*OES; 06: PF*OES |  | 2
2
| 1 | 1 | 01: 0.45; 02: 0.48283; 03: NA; 04: 0.68; 05: 0.54; 06: 0.51 |  | 1
| 1 | 2 | 01: 0.46; 02: 0.50295; 03: NA; 04: 0.71; 05: 0.6; 06: 0.53 |  | 25
| 1 | 3 | 01: 0.46; 02: 0.50711; 03: NA; 04: 0.72; 05: 0.48; 06: 0.53 |  | 16
| 1 | 4 | 01: 0.51; 02: 0.48559; 03: NA; 04: 0.68; 05: 0.56; 06: 0.52 |  | 0.53412
| 1 | 5 | 01: 0.45; 02: 0.46129; 03: NA; 04: 0.68; 05: 0.46; 06: 0.53 |  | 57
| 1 | 6 | 01: 0.43; 02: 0.50395; 03: NA; 04: 0.69; 05: 0.44; 06: 0.52
| Mean |  | 01: 0.46; 02: 0.49062; 04: 0.69333; 05: 0.51333; 06: 0.52333
| Median |  | 01: 0.455; 02: 0.49427; 04: 0.685; 05: 0.51; 06: 0.525
| Std Dev. |  | 01: 0.026833; 02: 0.017593; 04: 0.017512; 05: 0.062823; 06: 0.008165
| Rel.Std.Dev. |  | 01: 5.833%; 02: 3.586%; 04: 2.526%; 05: 12.24%; 06: 1.56%
| PDM3 |  | 01: -13.88%; 02: -8.145%; 04: 29.81%; 05: -3.892%; 06: -2.02%
| Z-Score (Absolute) |  | 01: 1.58; 02: 0.67; 03: NA; 04: 5.34; 05: 0; 06: 0.3
| Analytical results for Mg in OREAS 556b (Certified Value 2.49 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 2.49 | 2.33 | 2.42 | 2.59 | 2.46 | 2.59 | 2.388 | 2.442 | 2.37 | 2.442 | 2.503 | 2.578 | 4.4819 | 2.78 | 2.5121 | 2.5114 | 2.51 | 2.7 | 2.88 | 2.42 | 2.36 | 2.52 | 1
| 1 | 2 | 2.54 | 2.38 | 2.42 | 2.59 | 2.45 | 2.59 | 2.418 | 2.509 | 2.388 | 2.533 | 2.551 | 2.5628 | 4.4619 | 2.79 | 2.4952 | 2.4411 | 2.48 | 2.98 | 3.4 | 2.47 | 2.38 | 2.49 | 0
| 1 | 3 | 2.45 | 2.35 | 2.41 | 2.59 | 2.45 | 2.59 | 2.442 | 2.43 | 2.304 | 2.485 | 2.533 | 2.655 | 4.4343 | 2.74 | 2.4826 | 2.444 | 2.44 | 3.18 | 3.14 | 2.4 | 2.37 | 2.54 | 16
| 1 | 4 | 2.47 | 2.33 | 2.42 | 2.59 | 2.51 | 2.58 | 2.43 | 2.479 | 2.376 | 2.4 | 2.509 | 2.5884 | 4.4628 | 2.68 | 2.5101 | 2.5103 | 2.49 | 2.97 | 2.88 | 2.51 | 2.35 | 2.53 | 2.4886
| 1 | 5 | 2.47 | 2.4 | 2.4 | 2.56 | 2.58 | 2.57 | 2.418 | 2.479 | 2.364 | 2.467 | 2.539 | 2.571 | 4.4672 | 2.77 | 2.5041 | 2.5133 | 2.53 | 3.07 | 3.06 | 2.49 | 2.37 | 2.51 | 58
| 1 | 6 | 2.41 | 2.41 | 2.43 | 2.55 | 2.56 | 2.55 | 2.4 | 2.473 | 2.358 | 2.406 | 2.509 | 2.6427 | 4.4874 | 2.64 | 2.5258 | 2.4785 | 2.57 | 3.3 | 2.92 | 2.54 | 2.35 | 2.49
| Mean |  | 2.4717 | 2.3667 | 2.4167 | 2.5783 | 2.5017 | 2.5783 | 2.416 | 2.4687 | 2.36 | 2.4555 | 2.524 | 2.5996 | 4.4659 | 2.7333 | 2.505 | 2.4831 | 2.5033 | 3.0333 | 3.0467 | 2.4717 | 2.3633 | 2.5133
| Median |  | 2.47 | 2.365 | 2.42 | 2.59 | 2.485 | 2.585 | 2.418 | 2.476 | 2.367 | 2.4545 | 2.521 | 2.5832 | 4.465 | 2.755 | 2.5071 | 2.4944 | 2.5 | 3.025 | 2.99 | 2.48 | 2.365 | 2.515
| Std Dev. |  | 0.043089 | 0.035024 | 0.010328 | 0.018348 | 0.057764 | 0.016021 | 0.019596 | 0.028514 | 0.029312 | 0.050441 | 0.019627 | 0.0392 | 0.018677 | 0.060553 | 0.014873 | 0.033961 | 0.044572 | 0.20588 | 0.20265 | 0.053448 | 0.012111 | 0.020656
| Rel.Std.Dev. |  | 1.743% | 1.48% | 0.4274% | 0.7116% | 2.309% | 0.6214% | 0.8111% | 1.155% | 1.242% | 2.054% | 0.7776% | 1.508% | 0.4182% | 2.215% | 0.5937% | 1.368% | 1.781% | 6.787% | 6.652% | 2.162% | 0.5124% | 0.8219%
| PDM3 |  | -0.6824% | -4.902% | -2.892% | 3.604% | 0.5231% | 3.604% | -2.919% | -0.8029% | -5.169% | -1.332% | 1.42% | 4.46% | 79.45% | 9.832% | 0.6565% | -0.223% | 0.59% | 21.89% | 22.42% | -0.6824% | -5.035% | 0.9919%
| Z-Score (Absolute) |  | 0.27 | 1.21 | 0.76 | 0.67 | 0.01 | 0.67 | 0.77 | 0.3 | 1.27 | 0.42 | 0.19 | 0.86 | 17.46 | 2.05 | 0.02 | 0.17 | 0.01 | 4.72 | 4.84 | 0.27 | 1.24 | 0.1
| Analytical results for Mn in OREAS 556b (Certified Value 0.121 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.1121 | 0.12 | 0.1138 | 0.1259 | 0.123 | 0.131 | 0.116 | 0.116 | 0.116 | 0.116 | 0.124 | 0.12594 | 0.1287 | 0.124 | 0.12358 | 0.1662 | 0.11 | 0.1111 | 0.129 | 0.1316 | 0.12 | 0.12 | 1
| 1 | 2 | 0.1116 | 0.13 | 0.1129 | 0.1242 | 0.118 | 0.135 | 0.124 | 0.116 | 0.116 | 0.124 | 0.124 | 0.12228 | 0.1249 | 0.127 | 0.11627 | 0.122 | 0.114 | 0.1139 | 0.151 | 0.1337 | 0.12 | 0.11 | 0
| 1 | 3 | 0.1057 | 0.12 | 0.114 | 0.1326 | 0.127 | 0.143 | 0.147 | 0.116 | 0.108 | 0.132 | 0.124 | 0.12232 | 0.1213 | 0.122 | 0.11706 | 0.136 | 0.112 | 0.1394 | 0.145 | 0.1257 | 0.12 | 0.12 | 16
| 1 | 4 | 0.1068 | 0.12 | 0.1116 | 0.128 | 0.122 | 0.133 | 0.124 | 0.116 | 0.108 | 0.116 | 0.124 | 0.12414 | 0.1224 | 0.124 | 0.11403 | 0.1282 | 0.115 | 0.1234 | 0.138 | 0.1245 | 0.11 | 0.11 | 0.12072
| 1 | 5 | 0.1103 | 0.13 | 0.113 | 0.1237 | 0.118 | 0.133 | 0.116 | 0.116 | 0.108 | 0.124 | 0.124 | 0.12418 | 0.1296 | 0.125 | 0.11699 | 0.1219 | 0.11 | 0.1143 | 0.146 | 0.1312 | 0.12 | 0.12 | 59
| 1 | 6 | 0.1056 | 0.12 | 0.1125 | 0.1288 | 0.122 | 0.131 | 0.124 | 0.116 | 0.108 | 0.124 | 0.124 | 0.12528 | 0.1269 | 0.126 | 0.12349 | 0.1214 | 0.11 | 0.1493 | 0.138 | 0.1318 | 0.12 | 0.11
| Mean |  | 0.10868 | 0.12333 | 0.11297 | 0.1272 | 0.12167 | 0.13433 | 0.12517 | 0.116 | 0.11067 | 0.12267 | 0.124 | 0.12402 | 0.12563 | 0.12467 | 0.11857 | 0.13262 | 0.11183 | 0.12523 | 0.14117 | 0.12975 | 0.11833 | 0.115
| Median |  | 0.10855 | 0.12 | 0.11295 | 0.12695 | 0.122 | 0.133 | 0.124 | 0.116 | 0.108 | 0.124 | 0.124 | 0.12416 | 0.1259 | 0.1245 | 0.11702 | 0.1251 | 0.111 | 0.11885 | 0.1415 | 0.1314 | 0.12 | 0.115
| Std Dev. |  | 0.0029916 | 0.005164 | 0.00087788 | 0.0033257 | 0.0033862 | 0.0045019 | 0.011392 | 0 | 0.0041312 | 0.0060222 | 0 | 0.0014992 | 0.0033619 | 0.0017512 | 0.0040006 | 0.017391 | 0.0022286 | 0.015691 | 0.0077825 | 0.0037228 | 0.0040825 | 0.0054772
| Rel.Std.Dev. |  | 2.753% | 4.187% | 0.7771% | 2.615% | 2.783% | 3.351% | 9.101% | 0% | 3.733% | 4.909% | 0% | 1.209% | 2.676% | 1.405% | 3.374% | 13.11% | 1.993% | 12.53% | 5.513% | 2.869% | 3.45% | 4.763%
| PDM3 |  | -9.97% | 2.165% | -6.422% | 5.369% | 0.7849% | 11.28% | 3.684% | -3.909% | -8.327% | 1.613% | 2.718% | 2.736% | 4.071% | 3.27% | -1.781% | 9.856% | -7.361% | 3.739% | 16.94% | 7.481% | -1.976% | -4.738%
| Z-Score (Absolute) |  | 1.95 | 0.04 | 1.39 | 0.47 | 0.26 | 1.4 | 0.21 | 0.97 | 1.65 | 0.13 | 0.04 | 0.05 | 0.26 | 0.14 | 0.66 | 1.17 | 1.54 | 0.21 | 2.29 | 0.8 | 0.69 | 1.12
| Analytical results for Mo in OREAS 556b (Certified Value 18.7 ppm) | [23 empty] | Certified
Mo | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 05; 05: 06; 06: 15; 08: 17; 09: 18; 10: 19; 11: 21; 12: 23; 15: 25; 16: 26 | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*OES/MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES; 12: PF*MS; 15: PF*OES; 16: PF*OES | 1
1
| 1 | 1 | 01: 19.1; 02: 17.2; 03: 19; 04: 20; 05: 16; 06: 20.6364; 08: 20; 09: 17.591; 10: 19; 11: 20; 12: 18; 15: 14; 16: < 20 | 1
| 1 | 2 | 01: 18.2; 02: 17.1; 03: 19; 04: 18; 05: 16; 06: 20.6174; 08: 20; 09: 18.3241; 10: 28; 11: 20; 12: 17; 15: 18; 16: 23 | 7
| 1 | 3 | 01: 17.7; 02: 16.8; 03: 19; 04: 20; 05: 17; 06: 19.8595; 08: 21; 09: 18.349; 10: 19; 11: 10; 12: 18; 15: 17; 16: 24 | 16
| 1 | 4 | 01: 18.2; 02: 16.6; 03: 18; 04: 19; 05: 16; 06: 19.5149; 08: 20; 09: 17.5537; 10: 19; 11: 10; 12: 18; 15: 17; 16: 23 | 18.7411
| 1 | 5 | 01: 17.8; 02: 17.2; 03: 19; 04: 19; 05: 17; 06: 19.6404; 08: 21; 09: 17.8768; 10: 19; 11: 20; 12: 17; 15: 14; 16: < 20 | 60
| 1 | 6 | 01: 18; 02: 16.5; 03: 19; 04: 20; 05: 17; 06: 20.749; 08: 20; 09: 18.2446; 10: 24; 11: 20; 12: 17; 15: 15; 16: < 20
| Mean |  | 01: 18.1667; 02: 16.9; 03: 18.8333; 04: 19.3333; 05: 16.5; 06: 20.1696; 08: 20.3333; 09: 17.9899; 10: 21.3333; 11: 16.6667; 12: 17.5; 15: 15.8333; 16: 23.3333
| Median |  | 01: 18.1; 02: 16.95; 03: 19; 04: 19.5; 05: 16.5; 06: 20.2385; 08: 20; 09: 18.0607; 10: 19; 11: 20; 12: 17.5; 15: 16; 16: 23
| Std Dev. |  | 01: 0.50067; 02: 0.30984; 03: 0.40825; 04: 0.8165; 05: 0.54772; 06: 0.55841; 08: 0.5164; 09: 0.36543; 10: 3.8297; 11: 5.164; 12: 0.54772; 15: 1.7224; 16: 0.57735
| Rel.Std.Dev. |  | 01: 2.756%; 02: 1.833%; 03: 2.168%; 04: 4.223%; 05: 3.32%; 06: 2.769%; 08: 2.54%; 09: 2.031%; 10: 17.95%; 11: 30.98%; 12: 3.13%; 15: 10.88%; 16: 2.474%
| PDM3 |  | 01: -3.065%; 02: -9.824%; 03: 0.4923%; 04: 3.16%; 05: -11.96%; 06: 7.622%; 08: 8.496%; 09: -4.008%; 10: 13.83%; 11: -11.07%; 12: -6.622%; 15: -15.52%; 16: 24.5%
| Z-Score (Absolute) |  | 01: 0.04; 02: 0.6; 03: 0.38; 04: 0.63; 05: 0.8; 06: 1.06; 08: 1.14; 09: 0.04; 10: 1.65; 11: NA; 12: 0.29; 15: 1.14; 16: 0.71
Results from laboratory 21 were removed due to their 10 ppm reading resolution.
| Analytical results for Na in OREAS 556b (Indicative Value 0.045 wt.%) | [23 empty] | Indicative
Na | Round | Replicate | 01: Lab | 1
| No | No | 01: 01 | wt.%
|  |  | 01: PF*OES | 3
3
| 1 | 1 | 01: 0.04 | 1
| 1 | 2 | 01: 0.04 | 17
| 1 | 3 | 01: 0.06 | 16
| 1 | 4 | 01: 0.04 | 0.045
| 1 | 5 | 01: 0.04 | 23
| 1 | 6 | 01: 0.05
| Mean |  | 01: 0.045
| Median |  | 01: 0.04
| Std Dev. |  | 01: 0.0083666
| Rel.Std.Dev. |  | 01: 18.59%
| PDM3 |  | 01: 2.22e-14%
Z-Score (Absolute) | bar: 01
| Analytical results for Nb in OREAS 556b (Certified Value 22.5 ppm) | [23 empty] | Certified
Nb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 02; 02: 05; 03: 06; 04: 15; 05: 18; 06: 19; 08: 23; 09: 25; 10: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*MS; 09: PF*OES; 10: PF*OES | 1
2
| 1 | 1 | 01: 22; 02: 25; 03: 21; 04: 25.3904; 05: 22.1869; 06: 22; 08: 20; 09: 19.9; 10: < 20 | 1
| 1 | 2 | 01: 24; 02: 22.7; 03: 21; 04: 21.7617; 05: 23.495; 06: 22; 08: 22; 09: 21.8; 10: < 20 | 9
| 1 | 3 | 01: 25; 02: 25.2; 03: 20; 04: 21.8195; 05: 22.7252; 06: 21; 08: 20; 09: 23.1; 10: < 20 | 16
| 1 | 4 | 01: 23; 02: 24.4; 03: 19; 04: 24.6739; 05: 22.5601; 06: 22; 08: 20; 09: 17.5; 10: < 20 | 22.4636
| 1 | 5 | 01: 23; 02: 24.2; 03: 21; 04: 25.0387; 05: 22.2037; 06: 21; 08: 22; 09: 22.1; 10: < 20 | 61
| 1 | 6 | 01: 25; 02: 25.4; 03: 20; 04: 25.4311; 05: 22.9045; 06: 22; 08: 22; 09: 22.4; 10: < 20
| Mean |  | 01: 23.6667; 02: 24.4833; 03: 20.3333; 04: 24.0192; 05: 22.6792; 06: 21.6667; 08: 21; 09: 21.1333
| Median |  | 01: 23.5; 02: 24.7; 03: 20.5; 04: 24.8563; 05: 22.6427; 06: 22; 08: 21; 09: 21.95
| Std Dev. |  | 01: 1.2111; 02: 0.98877; 03: 0.8165; 04: 1.748; 05: 0.49011; 06: 0.5164; 08: 1.0954; 09: 2.0772
| Rel.Std.Dev. |  | 01: 5.117%; 02: 4.039%; 03: 4.016%; 04: 7.277%; 05: 2.161%; 06: 2.383%; 08: 5.216%; 09: 9.829%
| PDM3 |  | 01: 5.356%; 02: 8.991%; 03: -9.483%; 04: 6.925%; 05: 0.9602%; 06: -3.547%; 08: -6.515%; 09: -5.922%
| Z-Score (Absolute) |  | 01: 1.01; 02: 1.42; 03: 0.67; 04: 1.19; 05: 0.51; 06: 0; 08: 0.34; 09: 0.27; 10: 5.9
| Analytical results for Nd in OREAS 556b (Certified Value 62 ppm) | [23 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES | 0
1
| 1 | 1 | 01: 67.7; 02: 61; 03: 59.2156; 04: 62; 05: 59.3; 06: 67.8; 08: 65.5 | 1
| 1 | 2 | 01: 59.4; 02: 61; 03: 60.4004; 04: 61; 05: 53.2; 06: 78.9; 08: 67.5 | 10
| 1 | 3 | 01: 71.3; 02: 60; 03: 66.6071; 04: 58; 05: 64.1; 06: 74.4; 08: 66.4 | 16
| 1 | 4 | 01: 71.3; 02: 62; 03: 61.7182; 04: 62; 05: 58.7; 06: 74.8; 08: 66.3 | 61.8593
| 1 | 5 | 01: 69.2; 02: 62; 03: 56.8531; 04: 71; 05: 60.7; 06: 64.5; 08: 66.4 | 62
| 1 | 6 | 01: 69.6; 02: 61; 03: 62.9854; 04: 64; 05: 59.2; 06: 65; 08: 65.3
| Mean |  | 01: 68.0833; 02: 61.1667; 03: 61.2966; 04: 63; 05: 59.2; 06: 70.9; 08: 66.2333
| Median |  | 01: 69.4; 02: 61; 03: 61.0593; 04: 62; 05: 59.25; 06: 71.1; 08: 66.35
| Std Dev. |  | 01: 4.4674; 02: 0.75277; 03: 3.3499; 04: 4.3818; 05: 3.5361; 06: 5.9471; 08: 0.78401
| Rel.Std.Dev. |  | 01: 6.562%; 02: 1.231%; 03: 5.465%; 04: 6.955%; 05: 5.973%; 06: 8.388%; 08: 1.184%
| PDM3 |  | 01: 10.06%; 02: -1.12%; 03: -0.9096%; 04: 1.844%; 05: -4.299%; 06: 14.61%; 08: 7.071%
| Z-Score (Absolute) |  | 01: 1.06; 02: 0.38; 03: 0.36; 04: 0; 05: 0.79; 06: 1.65; 08: 0.67
| Analytical results for Ni in OREAS 556b (Certified Value 343 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 0
0
| 1 | 1 | 317.2 | 347 | 343 | 342 | 360 | 371 | 370 | 360 | 330 | 350 | 340 | 415.3905 | 418 | 380 | 287.6869 | NA | 280 | 331 | 327 | 349 | 360 | 410 | 1
| 1 | 2 | 315 | 351 | 345 | 348 | 360 | 399 | 330 | 380 | 330 | 420 | 330 | 405.0552 | 415 | 370 | 294.8085 | NA | 280 | 318 | 359 | 363 | 370 | 370 | 40
| 1 | 3 | 304 | 336 | 342 | 356 | 370 | 386 | 370 | 340 | 350 | 360 | 330 | 403.6994 | 417 | 360 | 294.0563 | NA | 300 | 324 | 345 | 327 | 360 | 370 | 16
| 1 | 4 | 314.3 | 355 | 343 | 337 | 370 | 372 | 330 | 320 | 330 | 340 | 330 | 406.5303 | 414 | 370 | 286.0396 | NA | 310 | 320 | 357 | 353 | 350 | 370 | 342.6253
| 1 | 5 | 315.6 | 356 | 342 | 334 | 360 | 366 | 320 | 340 | 340 | 350 | 340 | 406.127 | 410 | 370 | 294.7612 | NA | 290 | 311 | 331 | 338 | 370 | 380 | 63
| 1 | 6 | 304.7 | 344 | 343 | 351 | 360 | 383 | 350 | 340 | 330 | 350 | 340 | 415.04 | 414 | 380 | 295.1323 | NA | 320 | 328 | 322 | 367 | 360 | 380
| Mean |  | 311.8 | 348.1667 | 343 | 344.6667 | 363.3333 | 379.5 | 345 | 346.6667 | 335 | 361.6667 | 335 | 408.6404 | 414.6667 | 371.6667 | 292.0808 |  | 296.6667 | 322 | 340.1667 | 349.5 | 361.6667 | 380
| Median |  | 314.65 | 349 | 343 | 345 | 360 | 377.5 | 340 | 340 | 330 | 350 | 335 | 406.3287 | 414.5 | 370 | 294.4088 |  | 295 | 322 | 338 | 351 | 360 | 375
| Std Dev. |  | 5.8539 | 7.5211 | 1.0954 | 8.4774 | 5.164 | 12.2107 | 21.6795 | 20.6559 | 8.3666 | 29.2689 | 5.4772 | 5.1876 | 2.8048 | 7.5277 | 4.09 |  | 16.3299 | 7.2388 | 15.804 | 15.0964 | 7.5277 | 15.4919
| Rel.Std.Dev. |  | 1.877% | 2.16% | 0.3194% | 2.46% | 1.421% | 3.218% | 6.284% | 5.958% | 2.497% | 8.093% | 1.635% | 1.269% | 0.6764% | 2.025% | 1.4% |  | 5.504% | 2.248% | 4.646% | 4.319% | 2.081% | 4.077%
| PDM3 |  | -8.997% | 1.617% | 0.1094% | 0.5958% | 6.044% | 10.76% | 0.6931% | 1.18% | -2.226% | 5.557% | -2.226% | 19.27% | 21.03% | 8.476% | -14.75% |  | -13.41% | -6.02% | -0.7176% | 2.006% | 5.557% | 10.91%
| Z-Score (Absolute) |  | 1.57 | 0.07 | 0.16 | 0.09 | 0.75 | 1.48 | 0.07 | 0 | 0.52 | 0.67 | 0.52 | 2.79 | 3.06 | 1.12 | 2.45 | NA | 2.25 | 1.11 | 0.29 | 0.13 | 0.67 | 1.5
| Analytical results for P in OREAS 556b (Certified Value 0.035 wt.%) | [23 empty] | Certified
P | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 06; 05: 17; 06: 18; 08: 19; 09: 21; 10: 23; 11: 25; 12: 26 | wt.%
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*OES; 10: PF*OES; 11: PF*OES; 12: PF*OES | 3
3
| 1 | 1 | 01: 0.03; 02: 0.03; 03: 0.1244; 04: 0.26; 05: 0.04; 06: < 0.05; 08: 0.0332; 09: 0.04; 10: 0.17; 11: 0.04; 12: 0.03 | 1
| 1 | 2 | 01: 0.03; 02: 0.04; 03: 0.1238; 04: 0.27; 05: 0.04; 06: < 0.05; 08: 0.0328; 09: 0.03; 10: 0.19; 11: 0.05; 12: 0.03 | 0
| 1 | 3 | 01: 0.03; 02: 0.03; 03: 0.1238; 04: 0.26; 05: 0.04; 06: < 0.05; 08: 0.0319; 09: 0.03; 10: 0.17; 11: 0.05; 12: 0.04 | 16
| 1 | 4 | 01: 0.04; 02: 0.03; 03: 0.1212; 04: 0.27; 05: 0.04; 06: < 0.05; 08: 0.0326; 09: 0.02; 10: 0.16; 11: 0.04; 12: 0.03 | 0.034683
| 1 | 5 | 01: 0.03; 02: 0.03; 03: 0.1217; 04: 0.26; 05: 0.04; 06: < 0.05; 08: 0.0353; 09: 0.03; 10: 0.16; 11: 0.04; 12: 0.04 | 64
| 1 | 6 | 01: 0.03; 02: 0.03; 03: 0.1227; 04: 0.26; 05: 0.04; 06: < 0.05; 08: 0.0309; 09: 0.02; 10: 0.17; 11: 0.04; 12: 0.04
| Mean |  | 01: 0.031667; 02: 0.031667; 03: 0.12293; 04: 0.26333; 05: 0.04; 08: 0.032783; 09: 0.028333; 10: 0.17; 11: 0.043333; 12: 0.035
| Median |  | 01: 0.03; 02: 0.03; 03: 0.12325; 04: 0.26; 05: 0.04; 08: 0.0327; 09: 0.03; 10: 0.17; 11: 0.04; 12: 0.035
| Std Dev. |  | 01: 0.0040825; 02: 0.0040825; 03: 0.0012832; 04: 0.005164; 05: 0; 08: 0.0014743; 09: 0.0075277; 10: 0.010954; 11: 0.005164; 12: 0.0054772
| Rel.Std.Dev. |  | 01: 12.89%; 02: 12.89%; 03: 1.044%; 04: 1.961%; 05: 0%; 08: 4.497%; 09: 26.57%; 10: 6.444%; 11: 11.92%; 12: 15.65%
| PDM3 |  | 01: -8.698%; 02: -8.698%; 03: 254.4%; 04: 659.3%; 05: 15.33%; 08: -5.478%; 09: -18.31%; 10: 390.1%; 11: 24.94%; 12: 0.913%
| Z-Score (Absolute) |  | 01: 0.34; 02: 0.34; 03: 8.89; 04: 23.1; 05: 0.51; 06: 1.01; 08: 0.22; 09: 0.67; 10: 13.65; 11: 0.84; 12: 0
| Analytical results for Pb in OREAS 556b (Certified Value < 100 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 21 | 23 | 25 | 26 | 27 | 28 | [5 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 0
0
| 1 | 1 | 18 | < 100 | 21.4 | 17 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 23 | 20 | 31 | 13 | < 50 | < 300 | < 100 | [5 empty] | 1
| 1 | 2 | 20.8 | < 100 | 20.5 | 17 | < 100 | < 100 | < 100 | < 100 | 400 | < 20 | 25 | 20 | 31 | 29 | < 50 | < 300 | < 100 | [5 empty] | 11
| 1 | 3 | 16.9 | < 100 | 22.6 | 17 | < 100 | 200 | < 100 | < 100 | 500 | < 20 | 26 | 20 | 32 | 24 | < 50 | < 300 | < 100 | [5 empty] | 16
| 1 | 4 | 18.6 | < 100 | 23.5 | 18 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 27 | 30 | 34 | 15 | < 50 | < 300 | < 100 | [5 empty] | < 100
| 1 | 5 | 19.6 | < 100 | 24 | 18 | < 100 | 1400 | < 100 | < 100 | < 100 | < 20 | 25 | 20 | 31 | 21 | < 50 | < 300 | < 100 | [5 empty] | 65
| 1 | 6 | 17.2 | < 100 | 23.6 | 18 | < 100 | < 100 | < 100 | < 100 | < 100 | < 20 | 25 | 20 | 31 | 15 | < 50 | < 300 | < 100
| Mean |  | 01: 18.5167; 03: 22.6; 04: 17.5; 06: 800; 10: 450; 12: 25.1667; 15: 21.6667; 16: 31.6667; 17: 19.5
| Median |  | 01: 18.3; 03: 23.05; 04: 17.5; 06: 800; 10: 450; 12: 25; 15: 20; 16: 31; 17: 18
| Std Dev. |  | 01: 1.4838; 03: 1.3871; 04: 0.54772; 06: 848.5281; 10: 70.7107; 12: 1.3292; 15: 4.0825; 16: 1.2111; 17: 6.253
| Rel.Std.Dev. |  | 01: 8.013%; 03: 6.138%; 04: 3.13%; 06: 106.1%; 10: 15.71%; 12: 5.281%; 15: 18.84%; 16: 3.824%; 17: 32.07%
PDM3
| Z-Score (Absolute) |  | 0.48 | 0.67 | 0.33 | 0.52 | 0.67 | 9.87 | 0.67 | 0.67 | 5.58 | 0.8 | 0.24 | 0.37 | 0 | 0.45 | 0.25 | 4.35 | 0.67
| Analytical results for Pr in OREAS 556b (Certified Value 18.3 ppm) | [23 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES | 1
2
| 1 | 1 | 01: 18.6; 02: 15.8; 03: 16.8925; 04: 17; 05: 18.78; 06: 18.2; 08: 18.06 | 1
| 1 | 2 | 01: 16.1; 02: 15.8; 03: 17.2717; 04: 17; 05: 17.96; 06: 20.7; 08: 18.21 | 12
| 1 | 3 | 01: 18; 02: 15.8; 03: 18.8069; 04: 16; 05: 18.81; 06: 19.6; 08: 18.18 | 16
| 1 | 4 | 01: 17.9; 02: 16.2; 03: 17.6551; 04: 17; 05: 18.35; 06: 19.9; 08: 18.1 | 18.2767
| 1 | 5 | 01: 18.3; 02: 16.1; 03: 16.7201; 04: 19; 05: 19.32; 06: 18.2; 08: 18.27 | 66
| 1 | 6 | 01: 18.2; 02: 15.8; 03: 17.8758; 04: 18; 05: 17.8; 06: 17.8; 08: 17.64
| Mean |  | 01: 17.85; 02: 15.9167; 03: 17.537; 04: 17.3333; 05: 18.5033; 06: 19.0667; 08: 18.0767
| Median |  | 01: 18.1; 02: 15.8; 03: 17.4634; 04: 17; 05: 18.565; 06: 18.9; 08: 18.14
| Std Dev. |  | 01: 0.89163; 02: 0.18348; 03: 0.76085; 04: 1.0328; 05: 0.57462; 06: 1.1622; 08: 0.22686
| Rel.Std.Dev. |  | 01: 4.995%; 02: 1.153%; 03: 4.339%; 04: 5.958%; 05: 3.105%; 06: 6.095%; 08: 1.255%
| PDM3 |  | 01: -2.335%; 02: -12.91%; 03: -4.047%; 04: -5.162%; 05: 1.24%; 06: 4.322%; 08: -1.095%
| Z-Score (Absolute) |  | 01: 0.16; 02: 2.86; 03: 0.59; 04: NA; 05: 0.75; 06: 1.54; 08: 0.16
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 556b (Certified Value 82 ppm) | [23 empty] | Certified
Rb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 17; 06: 18; 08: 19; 09: 23; 10: 25; 11: 26; 12: 28 | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*MS; 10: PF*OES; 11: PF*MS; 12: PF*OES | 0
1
| 1 | 1 | 01: 83.1; 02: 82; 03: 79; 04: 83.977; 05: 92.1; 06: 83.4352; 08: 79; 09: 91; 10: 0.079; 11: 85.8; 12: 79 | 1
| 1 | 2 | 01: 82.5; 02: 79.9; 03: 79; 04: 77.5181; 05: 90.3; 06: 86.8381; 08: 81; 09: 92; 10: 0.089; 11: 86.3; 12: 84 | 13
| 1 | 3 | 01: 82.2; 02: 81.2; 03: 79; 04: 77.6688; 05: 89.8; 06: 84.289; 08: 77; 09: 94; 10: 0.085; 11: 85.8; 12: 80 | 16
| 1 | 4 | 01: 84.3; 02: 80.4; 03: 77; 04: 82.6131; 05: 93.2; 06: 83.4102; 08: 86; 09: 90; 10: 0.085; 11: 83.3; 12: 81 | 82.0004
| 1 | 5 | 01: 81.6; 02: 84.3; 03: 80; 04: 82.8691; 05: 93.1; 06: 83.9115; 08: 76; 09: 88; 10: 0.078; 11: 87.9; 12: 82 | 67
| 1 | 6 | 01: 82.3; 02: 84; 03: 79; 04: 78.9787; 05: 94.4; 06: 85.033; 08: 85; 09: 90; 10: 0.081; 11: 87.4; 12: 81
| Mean |  | 01: 82.6667; 02: 81.9667; 03: 78.8333; 04: 80.6041; 05: 92.15; 06: 84.4862; 08: 80.6667; 09: 90.8333; 10: 0.082833; 11: 86.0833; 12: 81.1667
| Median |  | 01: 82.4; 02: 81.6; 03: 79; 04: 80.7959; 05: 92.6; 06: 84.1002; 08: 80; 09: 90.5; 10: 0.083; 11: 86.05; 12: 81
| Std Dev. |  | 01: 0.93524; 02: 1.8381; 03: 0.98319; 04: 2.8748; 05: 1.7897; 06: 1.3009; 08: 4.1312; 09: 2.0412; 10: 0.0042151; 11: 1.6117; 12: 1.7224
| Rel.Std.Dev. |  | 01: 1.131%; 02: 2.243%; 03: 1.247%; 04: 3.567%; 05: 1.942%; 06: 1.54%; 08: 5.121%; 09: 2.247%; 10: 5.089%; 11: 1.872%; 12: 2.122%
| PDM3 |  | 01: 0.8125%; 02: -0.04114%; 03: -3.862%; 04: -1.703%; 05: 12.38%; 06: 3.031%; 08: -1.626%; 09: 10.77%; 10: -99.9%; 11: 4.979%; 12: -1.017%
| Z-Score (Absolute) |  | 01: 0.19; 02: 0; 03: 0.84; 04: 0.36; 05: 2.73; 06: 0.67; 08: 0.35; 09: 2.37; 10: 21.92; 11: 1.1; 12: 0.21
| Analytical results for Re in OREAS 556b (Indicative Value 0.059 ppm) | [23 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 04; 02: 15; 03: 17; 04: 18; 05: 19; 06: 23; 08: 25; 09: 26 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*MS; 08: PF*OES; 09: PF*MS | 3
3
| 1 | 1 | 01: < 0.1; 02: < 0.1; 03: < 0.05; 04: < 0.01; 05: 0.03; 06: 0.07; 08: 37.9; 09: < 0.1 | 1
| 1 | 2 | 01: < 0.1; 02: < 0.1; 03: < 0.05; 04: < 0.01; 05: 0.05; 06: 0.1; 08: 44.4; 09: < 0.1 | 14
| 1 | 3 | 01: < 0.1; 02: < 0.1; 03: < 0.05; 04: < 0.01; 05: 0.03; 06: 0.09; 08: 43.2; 09: < 0.1 | 16
| 1 | 4 | 01: < 0.1; 02: < 0.1; 03: < 0.05; 04: < 0.01; 05: 0.03; 06: 0.08; 08: 34; 09: < 0.1 | 0.059167
| 1 | 5 | 01: < 0.1; 02: < 0.1; 03: < 0.05; 04: < 0.01; 05: 0.04; 06: 0.07; 08: 36.2; 09: < 0.1 | 24
| 1 | 6 | 01: < 0.1; 02: < 0.1; 03: < 0.05; 04: < 0.01; 05: 0.04; 06: 0.08; 08: 34; 09: < 0.1
| Mean |  | 05: 0.036667; 06: 0.081667; 08: 38.2833
| Median |  | 05: 0.035; 06: 0.08; 08: 37.05
| Std Dev. |  | 05: 0.008165; 06: 0.01169; 08: 4.5336
| Rel.Std.Dev. |  | 05: 22.27%; 06: 14.31%; 08: 11.84%
| PDM3 |  | 05: -38.03%; 06: 38.03%; 08: 64604%
| Z-Score (Absolute) |  | 01: 0; 02: 0; 03: 0.88; 04: 1.58; 05: 0.47; 06: 1.11; 08: 1345.1; 09: 0
| Analytical results for S in OREAS 556b (Certified Value 8.75 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 5.92 | 9 | 8.79 | 8.61 | 8.4 | 9.02 | 8.53 | 8.55 | 8.64 | 8.75 | 8.89 | 8.9032 | 11.7528 | 9.39 | 8.1452 | 8.71 | 9.79 | 9.16 | 8.52 | 8.19 | 8.73 |  | 1
| 1 | 2 | 6.45 | 9 | 8.78 | 8.83 | 8.69 | 9.01 | 8.54 | 8.71 | 8.75 | 9.04 | 8.94 | 9.123 | 11.8838 | 9 | 8.2933 | 8.71 | > 10 | 10.4 | 8.6 | 8.27 | 8.54 |  | 28
| 1 | 3 | 6.01 | 9 | 8.8 | 8.97 | 8.69 | 8.86 | 8.48 | 8.64 | 8.6 | 8.86 | 8.93 | 8.7972 | 11.6684 | 9.17 | 8.1918 | 8.67 | > 10 | 9.78 | 8.29 | 8.08 | 8.59 |  | 16
| 1 | 4 | 7.49 | 9 | 8.79 | 8.82 | 8.62 | 9.03 | 8.78 | 8.69 | 8.64 | 8.58 | 8.88 | 8.8782 | 11.2177 | 9.31 | 8.4194 | 8.61 | > 10 | 9.63 | 8.78 | 8.3 | 8.41 |  | 8.7525
| 1 | 5 | 6.47 | 9 | 8.77 | 8.82 | 9.18 | 8.95 | 8.79 | 8.7 | 8.54 | 8.8 | 8.92 | 8.9126 | 11.8859 | 9.36 | 8.3198 | 8.65 | 9.03 | 9.53 | 8.56 | 8.38 | 8.65 |  | 68
| 1 | 6 | 6.56 | 9 | 8.8 | 8.5 | 8.89 | 8.78 | 8.51 | 8.67 | 8.63 | 8.53 | 8.84 | 9.1886 | 11.0255 | 9.05 | 8.1126 | 8.74 | 9.4 | 9.35 | 8.81 | 8.31 | 8.57
| Mean |  | 6.4833 | 9 | 8.7883 | 8.7583 | 8.745 | 8.9417 | 8.605 | 8.66 | 8.6333 | 8.76 | 8.9 | 8.9671 | 11.5724 | 9.2133 | 8.247 | 8.6817 | 9.4067 | 9.6417 | 8.5933 | 8.255 | 8.5817
| Median |  | 6.46 | 9 | 8.79 | 8.82 | 8.69 | 8.98 | 8.535 | 8.68 | 8.635 | 8.775 | 8.905 | 8.9079 | 11.7106 | 9.24 | 8.2425 | 8.69 | 9.4 | 9.58 | 8.58 | 8.285 | 8.58
| Std Dev. |  | 0.55892 | 0 | 0.01169 | 0.17105 | 0.26493 | 0.10147 | 0.14096 | 0.05933 | 0.068605 | 0.1873 | 0.037417 | 0.1531 | 0.36385 | 0.16501 | 0.11717 | 0.047504 | 0.38004 | 0.42986 | 0.18991 | 0.10559 | 0.10778
| Rel.Std.Dev. |  | 8.621% | 0% | 0.133% | 1.953% | 3.03% | 1.135% | 1.638% | 0.6851% | 0.7947% | 2.138% | 0.4204% | 1.707% | 3.144% | 1.791% | 1.421% | 0.5472% | 4.04% | 4.458% | 2.21% | 1.279% | 1.256%
| PDM3 |  | -25.93% | 2.827% | 0.4089% | 0.06613% | -0.0862% | 2.161% | -1.686% | -1.057% | -1.362% | 0.08518% | 1.685% | 2.452% | 32.22% | 5.265% | -5.776% | -0.8098% | 7.474% | 10.16% | -1.819% | -5.685% | -1.952%
| Z-Score (Absolute) |  | 8.68 | 0.92 | 0.11 | 0 | 0.05 | 0.7 | 0.59 | 0.38 | 0.48 | 0.01 | 0.54 | 0.8 | 10.74 | 1.74 | 1.95 | 0.29 | 2.47 | 3.37 | 0.63 | 1.92 | 0.67
| Analytical results for Sb in OREAS 556b (Certified Value 5.77 ppm) | [23 empty] | Certified
Sb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 01; 02: 04; 03: 05; 04: 06; 05: 15; 06: 21; 08: 23; 09: 25; 10: 26 |  | ppm
|  |  | 01: PF*OES; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*OES; 08: PF*MS; 09: PF*OES; 10: PF*MS |  | 2
3
| 1 | 1 | 01: 8.2; 02: 5.4; 03: 5; 04: 6; 05: 5.7734; 06: < 50; 08: 5; 09: 5; 10: 5.6 |  | 1
| 1 | 2 | 01: 6.4; 02: 5.6; 03: 6; 04: 6; 05: 5.4621; 06: < 50; 08: 4; 09: 7; 10: 5.4 |  | 15
| 1 | 3 | 01: 6.8; 02: 5.8; 03: 6; 04: 6; 05: 5.2903; 06: < 50; 08: 5; 09: 6; 10: 5.7 |  | 16
| 1 | 4 | 01: 6; 02: 6.4; 03: 6; 04: 5; 05: 5.7146; 06: < 50; 08: 4; 09: 5; 10: 5.3 |  | 5.7737
| 1 | 5 | 01: 6; 02: 5.4; 03: 6; 04: 6; 05: 5.5121; 06: < 50; 08: 5; 09: 6; 10: 5.6 |  | 69
| 1 | 6 | 01: 5.8; 02: 5.2; 03: 6; 04: 6; 05: 5.4412; 06: < 50; 08: 5; 09: 6; 10: 5.7
| Mean |  | 01: 6.5333; 02: 5.6333; 03: 5.8333; 04: 5.8333; 05: 5.5323; 08: 4.6667; 09: 5.8333; 10: 5.55
| Median |  | 01: 6.2; 02: 5.5; 03: 6; 04: 6; 05: 5.4871; 08: 5; 09: 6; 10: 5.6
| Std Dev. |  | 01: 0.89144; 02: 0.4274; 03: 0.40825; 04: 0.40825; 05: 0.18086; 08: 0.5164; 09: 0.75277; 10: 0.16432
| Rel.Std.Dev. |  | 01: 13.64%; 02: 7.587%; 03: 6.999%; 04: 6.999%; 05: 3.269%; 08: 11.07%; 09: 12.9%; 10: 2.961%
| PDM3 |  | 01: 13.16%; 02: -2.431%; 03: 1.033%; 04: 1.033%; 05: -4.18%; 08: -19.17%; 09: 1.033%; 10: -3.874%
| Z-Score (Absolute) |  | 01: 1.67; 02: 0.48; 03: 0; 04: 0; 05: 0.72; 06: 45.62; 08: 2.78; 09: 0; 10: 0.67
| Analytical results for Sc in OREAS 556b (Certified Value 9.56 ppm) | [23 empty] | Certified
Sc | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 01; 02: 04; 03: 06; 04: 15; 05: 17; 06: 18; 08: 19; 09: 21; 10: 25; 11: 26 |  | ppm
|  |  | 01: PF*OES; 02: PF*OES; 03: PF*OES/MS; 04: PF*OES; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*OES; 10: PF*OES; 11: PF*OES |  | 2
2
| 1 | 1 | 01: 11.2; 02: < 10; 03: < 10; 04: < 20; 05: 10; 06: 7.6643; 08: NA; 09: 11; 10: 12; 11: < 10 |  | 1
| 1 | 2 | 01: 11.2; 02: < 10; 03: < 10; 04: < 20; 05: 10; 06: 7.3702; 08: NA; 09: 11; 10: 17; 11: 10 |  | 1
| 1 | 3 | 01: 11; 02: < 10; 03: < 10; 04: < 20; 05: 9; 06: 7.8575; 08: NA; 09: 11; 10: 17; 11: < 10 |  | 16
| 1 | 4 | 01: 11.4; 02: < 10; 03: < 10; 04: < 20; 05: 9; 06: 7.7077; 08: NA; 09: 11; 10: 14; 11: 10 |  | 9.5602
| 1 | 5 | 01: 11.2; 02: < 10; 03: < 10; 04: < 20; 05: 10; 06: 7.7932; 08: NA; 09: 11; 10: 14; 11: 10 |  | 70
| 1 | 6 | 01: 10.8; 02: < 10; 03: < 10; 04: < 20; 05: 9; 06: 7.6129; 08: NA; 09: 11; 10: 13; 11: 11
| Mean |  | 01: 11.1333; 05: 9.5; 06: 7.6676; 09: 11; 10: 14.5; 11: 10.25
| Median |  | 01: 11.2; 05: 9.5; 06: 7.686; 09: 11; 10: 14; 11: 10
| Std Dev. |  | 01: 0.20656; 05: 0.54772; 06: 0.17023; 09: 0; 10: 2.0736; 11: 0.5
| Rel.Std.Dev. |  | 01: 1.855%; 05: 5.766%; 06: 2.22%; 09: 0%; 10: 14.3%; 11: 4.878%
| PDM3 |  | 01: 16.46%; 05: -0.6296%; 06: -19.8%; 09: 15.06%; 10: 51.67%; 11: 7.215%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.86; 03: 1.86; 04: 0.21; 05: 0; 06: 0.76; 08: NA; 09: 0.62; 10: 2.06; 11: 0.41
| Analytical results for Se in OREAS 556b (Indicative Value 16.7 ppm) | [23 empty] | Indicative
Se | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 25; 06: 26 |  | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES; 06: PF*MS |  | 1
1
| 1 | 1 | 01: < 20; 02: 20; 03: 12; 04: < 20; 05: 15; 06: < 20 |  | 1
| 1 | 2 | 01: < 20; 02: 17; 03: 12; 04: < 20; 05: 19; 06: < 20 |  | 19
| 1 | 3 | 01: < 20; 02: 21; 03: 13; 04: < 20; 05: 18; 06: < 20 |  | 16
| 1 | 4 | 01: < 20; 02: 23; 03: 12; 04: < 20; 05: 21; 06: < 20 |  | 16.7222
| 1 | 5 | 01: < 20; 02: 19; 03: 12; 04: < 20; 05: 17; 06: < 20 |  | 25
| 1 | 6 | 01: < 20; 02: 20; 03: 13; 04: < 20; 05: 17; 06: < 20
| Mean |  | 02: 20; 03: 12.3333; 05: 17.8333
| Median |  | 02: 20; 03: 12; 05: 17.5
| Std Dev. |  | 02: 2; 03: 0.5164; 05: 2.0412
| Rel.Std.Dev. |  | 02: 10%; 03: 4.187%; 05: 11.45%
| PDM3 |  | 02: 19.6%; 03: -26.25%; 05: 6.645%
| Z-Score (Absolute) |  | 01: 0.67; 02: 5.11; 03: 0.67; 04: 0.67; 05: 3.85; 06: 0.67
| Analytical results for Si in OREAS 556b (Certified Value 20.48 wt.%) | [23 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 18 | 19 | 21 | 25 | 26 | 28 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 2
3
| 1 | 1 | 22.54 | 20.64 | 20.7 | 20.7 | 20.5 | 20.01 | 20.62 | 20.29 | 20.38 | 20.38 | 20.7702 | 29.4033 | 20.3443 | 20.3341 | 20.5 | 1.12 | 20.7 | 20.3 |  |  |  |  | 1
| 1 | 2 | 22.93 | 20.53 | 20.8 | 20.3 | 20.5 | 20.01 | 20.9 | 20.1 | 20.62 | 20.71 | 20.7416 | 29.9582 | 20.5481 | 20.7036 | 19.6 | 19.92 | 21.1 | 19.9 |  |  |  |  | 0
| 1 | 3 | 23.82 | 20.43 | 20.2 | 20.8 | 20.3 | 19.91 | 20.9 | 19.45 | 20.38 | 20.52 | 20.5848 | 29.5571 | 20.0175 | 20.3133 | 20.6 | 17.35 | 20.4 | 20.1 |  |  |  |  | 16
| 1 | 4 | 22.38 | 20.29 | 20.8 | 20.8 | 20.2 | 20.01 | 21.08 | 19.4 | 19.82 | 20.52 | 21.2503 | 29.4425 | 21.0085 | 21.0593 | 19.5 | 0.65 | 21.3 | 20.2 |  |  |  |  | 20.4762
| 1 | 5 | 23.08 | 20.31 | 20.6 | 21.8 | 20.4 | 20.01 | 20.9 | 20.38 | 20.38 | 20.52 | 20.7182 | 29.3654 | 21.0313 | 20.5655 | 20.7 | 8.52 | 21.2 | 20.1 |  |  |  |  | 71
| 1 | 6 | 23.71 | 20.51 | 20.4 | 21.7 | 20.1 | 19.91 | 20.8 | 19.07 | 19.91 | 20.52 | 21.313 | 29.2934 | 20.4216 | 20.5182 | 20.3 | 7.29 | 21.5 | 19.9
| Mean |  | 23.0767 | 20.4517 | 20.5833 | 21.0167 | 20.3333 | 19.9767 | 20.8667 | 19.7817 | 20.2483 | 20.5283 | 20.8964 | 29.5033 | 20.5619 | 20.5823 | 20.2 | 9.1417 | 21.0333 | 20.0833
| Median |  | 23.005 | 20.47 | 20.65 | 20.8 | 20.35 | 20.01 | 20.9 | 19.775 | 20.38 | 20.52 | 20.7559 | 29.4229 | 20.4848 | 20.5419 | 20.4 | 7.905 | 21.15 | 20.1
| Std Dev. |  | 0.59136 | 0.13541 | 0.24014 | 0.59805 | 0.1633 | 0.05164 | 0.15108 | 0.54404 | 0.31244 | 0.10515 | 0.30579 | 0.23942 | 0.39582 | 0.2759 | 0.52154 | 8.0476 | 0.40825 | 0.16021
| Rel.Std.Dev. |  | 2.563% | 0.6621% | 1.167% | 2.846% | 0.8031% | 0.2585% | 0.724% | 2.75% | 1.543% | 0.5122% | 1.463% | 0.8115% | 1.925% | 1.34% | 2.582% | 88.03% | 1.941% | 0.7977%
| PDM3 |  | 12.7% | -0.1198% | 0.5232% | 2.639% | -0.6977% | -2.44% | 1.907% | -3.392% | -1.113% | 0.2546% | 2.052% | 44.09% | 0.4183% | 0.5183% | -1.349% | -55.35% | 2.721% | -1.919%
| Z-Score (Absolute) |  | 4.9 | 0.18 | 0.07 | 0.91 | 0.41 | 1.1 | 0.62 | 1.48 | 0.57 | 0.03 | 0.68 | 17.35 | 0.03 | 0.07 | 0.67 | 22.09 | 0.94 | 0.9
| Analytical results for Sm in OREAS 556b (Certified Value 8.88 ppm) | [23 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab |  | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES |  | 2
2
| 1 | 1 | 01: 8.9; 02: 8.7; 03: 8.5113; 04: 7; 05: 8; 06: 9.7; 08: 8.78 |  | 1
| 1 | 2 | 01: 7.1; 02: 8.9; 03: 8.834; 04: 8; 05: 8.1; 06: 11.6; 08: 9.13 |  | 3
| 1 | 3 | 01: 8.1; 02: 8.8; 03: 9.2998; 04: 7; 05: 8.3; 06: 11.5; 08: 9.02 |  | 16
| 1 | 4 | 01: 7.4; 02: 8.9; 03: 8.7073; 04: 7; 05: 7.9; 06: 10.9; 08: 8.86 |  | 8.8758
| 1 | 5 | 01: 8.5; 02: 9.1; 03: 9.1436; 04: 9; 05: 8.2; 06: 10.4; 08: 8.94 |  | 72
| 1 | 6 | 01: 8.5; 02: 8.5; 03: 9.3142; 04: 7; 05: 7.9; 06: 8.1; 08: 8.99
| Mean |  | 01: 8.0833; 02: 8.8167; 03: 8.9684; 04: 7.5; 05: 8.0667; 06: 10.3667; 08: 8.9533
| Median |  | 01: 8.3; 02: 8.85; 03: 8.9888; 04: 7; 05: 8.05; 06: 10.65; 08: 8.965
| Std Dev. |  | 01: 0.69976; 02: 0.20412; 03: 0.33323; 04: 0.83666; 05: 0.1633; 06: 1.3171; 08: 0.12323
| Rel.Std.Dev. |  | 01: 8.657%; 02: 2.315%; 03: 3.716%; 04: 11.16%; 05: 2.024%; 06: 12.7%; 08: 1.376%
| PDM3 |  | 01: -8.929%; 02: -0.6667%; 03: 1.042%; 04: -15.5%; 05: -9.117%; 06: 16.8%; 08: 0.8731%
| Z-Score (Absolute) |  | 01: 1.22; 02: 0.1; 03: 0.13; 04: NA; 05: 1.25; 06: 2.26; 08: 0.1
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 556b (Indicative Value 3.64 ppm) | [23 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 15 | 18 | 19 | 21 | 25 | 26 | 27 | 28 | [9 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | [9 empty] | 2
2
| 1 | 1 | < 5 | 22 | < 100 | 4.3 | 5 | < 100 | 2.8728 | < 5 | < 50 | 3 | < 100 | < 50 | < 100 | [9 empty] | 1
| 1 | 2 | < 5 | 24 | < 100 | 4.3 | 5 | < 100 | 3.0902 | < 5 | < 50 | 4 | < 100 | < 50 | < 100 | [9 empty] | 20
| 1 | 3 | < 5 | 26 | < 100 | 4.3 | 5 | < 100 | 2.871 | < 5 | < 50 | 4 | < 100 | < 50 | < 100 | [9 empty] | 16
| 1 | 4 | 5 | 22 | < 100 | 5 | 5 | < 100 | 3.0771 | < 5 | < 50 | 4 | < 100 | < 50 | < 100 | [9 empty] | 3.6359
| 1 | 5 | < 5 | 26 | < 100 | 4.1 | 4 | < 100 | 3.085 | < 5 | < 50 | 3 | < 100 | < 50 | < 100 | [9 empty] | 26
| 1 | 6 | < 5 | 24 | < 100 | 4.7 | 5 | < 100 | 3.0415 | < 5 | < 50 | 3 | < 100 | < 50 | < 100
| Mean |  | 01: 5; 02: 24; 04: 4.45; 05: 4.8333; 08: 3.0063; 11: 3.5
| Median |  | 01: 5; 02: 24; 04: 4.3; 05: 5; 08: 3.0593; 11: 3.5
| Std Dev. |  | 02: 1.7889; 04: 0.33317; 05: 0.40825; 08: 0.10544; 11: 0.54772
| Rel.Std.Dev. |  | 02: 7.454%; 04: 7.487%; 05: 8.447%; 08: 3.507%; 11: 15.65%
| PDM3 |  | 01: 37.52%; 02: 560.1%; 04: 22.39%; 05: 32.93%; 08: -17.32%; 11: -3.738%
| Z-Score (Absolute) |  | 0.68 | 0 | 0.84 | 0.63 | 0.62 | 0.84 | 0.67 | 0.69 | 0.03 | 0.66 | 0.84 | 0.03 | 0.84
| Analytical results for Sr in OREAS 556b (Certified Value 73 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | [6 empty] | 0
1
| 1 | 1 | 61.9 | 80 | 66 | 76 | 83 | 69 | 71.3185 | 104 | 70.8 | 68.6965 | NA | 70 | 80 | 63 | 70 | 73 | [6 empty] | 1
| 1 | 2 | 61.6 | 80 | 66.8 | 75 | 76 | 68 | 75.2576 | 101 | 71.7 | 73.9774 | NA | 70 | 90 | 90 | 71 | 75 | [6 empty] | 32
| 1 | 3 | 62.1 | 80 | 66.9 | 76 | 79 | 68 | 70.4965 | 105 | 71.3 | 71.3906 | NA | 80 | 100 | 81 | 70 | 73 | [6 empty] | 16
| 1 | 4 | 61.8 | 80 | 66.7 | 78 | 82 | 69 | 75.9535 | 103 | 72.1 | 75.426 | NA | 70 | 70 | < 10 | 69 | 74 | [6 empty] | 73.3151
| 1 | 5 | 61.3 | 80 | 66.6 | 77 | 82 | 70 | 73.3338 | 106 | 71.8 | 65.8822 | NA | 70 | 75 | 37 | 71 | 74 | [6 empty] | 73
| 1 | 6 | 60.4 | 80 | 65.6 | 76 | 81 | 68 | 77.0117 | 100 | 72.8 | 71.5236 | NA | 70 | 80 | 13 | 70 | 74
| Mean |  | 61.5167 | 80 | 66.4333 | 76.3333 | 80.5 | 68.6667 | 73.8953 | 103.1667 | 71.75 | 71.1494 |  | 71.6667 | 82.5 | 56.8 | 70.1667 | 73.8333
| Median |  | 61.7 | 80 | 66.65 | 76 | 81.5 | 68.5 | 74.2957 | 103.5 | 71.75 | 71.4571 |  | 70 | 80 | 63 | 70 | 74
| Std Dev. |  | 0.61128 | 0 | 0.5164 | 1.0328 | 2.5884 | 0.8165 | 2.6196 | 2.3166 | 0.68337 | 3.4712 |  | 4.0825 | 10.8397 | 31.7679 | 0.75277 | 0.75277
| Rel.Std.Dev. |  | 0.9937% | 0% | 0.7773% | 1.353% | 3.215% | 1.189% | 3.545% | 2.245% | 0.9524% | 4.879% |  | 5.696% | 13.14% | 55.93% | 1.073% | 1.02%
| PDM3 |  | -16.09% | 9.118% | -9.387% | 4.117% | 9.8% | -6.34% | 0.7914% | 40.72% | -2.135% | -2.954% |  | -2.248% | 12.53% | -22.53% | -4.294% | 0.7069%
| Z-Score (Absolute) |  | 1.51 | 1.21 | 0.78 | 0.67 | 1.29 | 0.45 | 0.32 | 4.62 | 0 | 0.09 | NA | 0.01 | 1.58 | 3.47 | 0.23 | 0.31
| Analytical results for Ta in OREAS 556b (Certified Value 1.84 ppm) | [23 empty] | Certified
Ta | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 15; 06: 18; 08: 19; 09: 23; 10: 25; 11: 26; 12: 28 |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*MS; 10: PF*OES; 11: PF*MS; 12: PF*OES |  | 2
2
| 1 | 1 | 01: 2; 02: 1.7; 03: 2.8; 04: 1.6; 05: 2.2891; 06: 1.6234; 08: NA; 09: 2.2; 10: 2; 11: 1.7; 12: < 5 |  | 1
| 1 | 2 | 01: 2; 02: 1.8; 03: 2.4; 04: 1.6; 05: 2.4415; 06: 1.5931; 08: NA; 09: 2.2; 10: 2; 11: 1.8; 12: < 5 |  | 5
| 1 | 3 | 01: 2; 02: 1.7; 03: 2.8; 04: 1.6; 05: 1.9288; 06: 1.5951; 08: NA; 09: 2.3; 10: 2; 11: 1.7; 12: < 5 |  | 16
| 1 | 4 | 01: 2; 02: 1.7; 03: 2.2; 04: 1.6; 05: 2.2152; 06: 1.6734; 08: NA; 09: 2.2; 10: 2; 11: 1.8; 12: < 5 |  | 1.8409
| 1 | 5 | 01: 2; 02: 1.7; 03: 3; 04: 1.6; 05: 1.8514; 06: 1.5908; 08: NA; 09: 2.1; 10: 2; 11: 1.8; 12: < 5 |  | 74
| 1 | 6 | 01: 2; 02: 1.8; 03: 2.6; 04: 1.6; 05: 2.1561; 06: 1.589; 08: NA; 09: 2.3; 10: 2; 11: 1.7; 12: < 5
| Mean |  | 01: 2; 02: 1.7333; 03: 2.6333; 04: 1.6; 05: 2.147; 06: 1.6108; 09: 2.2167; 10: 2; 11: 1.75
| Median |  | 01: 2; 02: 1.7; 03: 2.7; 04: 1.6; 05: 2.1856; 06: 1.5941; 09: 2.2; 10: 2; 11: 1.75
| Std Dev. |  | 01: 0; 02: 0.05164; 03: 0.29439; 04: 2.4324e-16; 05: 0.22211; 06: 0.0332; 09: 0.075277; 10: 0; 11: 0.054772
| Rel.Std.Dev. |  | 01: 0%; 02: 2.979%; 03: 11.18%; 04: 1.52e-14%; 05: 10.35%; 06: 2.061%; 09: 3.396%; 10: 0%; 11: 3.13%
| PDM3 |  | 01: 8.644%; 02: -5.842%; 03: 43.05%; 04: -13.09%; 05: 16.63%; 06: -12.5%; 09: 20.41%; 10: 8.644%; 11: -4.937%
| Z-Score (Absolute) |  | 01: NA; 02: 0.48; 03: 1.52; 04: 0.78; 05: 0.44; 06: 0.75; 08: NA; 09: 0.6; 10: NA; 11: 0.44; 12: 1.23
Results from laboratories 2 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 556b (Certified Value 1.05 ppm) | [23 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab |  | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES |  | 2
3
| 1 | 1 | 01: 1.1; 02: 0.97; 03: 1.0017; 04: NA; 05: 1.05; 06: 1.11; 08: 1.15 |  | 1
| 1 | 2 | 01: 1; 02: 1; 03: 1.0989; 04: NA; 05: 1; 06: 1.34; 08: 1.17 |  | 6
| 1 | 3 | 01: 1; 02: 1.02; 03: 1.0235; 04: NA; 05: 1.08; 06: 1.35; 08: 1.12 |  | 16
| 1 | 4 | 01: 1.1; 02: 1.07; 03: 1.0258; 04: NA; 05: 1.03; 06: 1.34; 08: 1.17 |  | 1.0544
| 1 | 5 | 01: 1; 02: 1.02; 03: 1.0287; 04: NA; 05: 1.03; 06: 1.14; 08: 1.18 |  | 75
| 1 | 6 | 01: 1; 02: 1.01; 03: 1.0635; 04: NA; 05: 1.04; 06: 1.08; 08: 1.15
| Mean |  | 01: 1.0333; 02: 1.015; 03: 1.0403; 05: 1.0383; 06: 1.2267; 08: 1.1567
| Median |  | 01: 1; 02: 1.015; 03: 1.0273; 05: 1.035; 06: 1.24; 08: 1.16
| Std Dev. |  | 01: 0.05164; 02: 0.032711; 03: 0.034905; 05: 0.026394; 06: 0.12925; 08: 0.021602
| Rel.Std.Dev. |  | 01: 4.997%; 02: 3.223%; 03: 3.355%; 05: 2.542%; 06: 10.54%; 08: 1.868%
| PDM3 |  | 01: -1.997%; 02: -3.736%; 03: -1.332%; 05: -1.523%; 06: 16.34%; 08: 9.7%
| Z-Score (Absolute) |  | 01: 0.27; 02: 1.08; 03: 0.04; 04: NA; 05: 0.04; 06: 8.32; 08: 5.21
| Analytical results for Te in OREAS 556b (Indicative Value < 2 ppm) | [23 empty] | Indicative
Te | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 25; 06: 26 |  | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES; 06: PF*MS |  | 2
2
| 1 | 1 | 01: < 2; 02: < 6; 03: < 5; 04: < 1; 05: 111; 06: < 2 |  | 1
| 1 | 2 | 01: < 2; 02: < 6; 03: < 5; 04: < 1; 05: 81; 06: < 2 |  | 7
| 1 | 3 | 01: < 2; 02: < 6; 03: < 5; 04: < 1; 05: 70; 06: < 2 |  | 16
| 1 | 4 | 01: < 2; 02: < 6; 03: < 5; 04: < 1; 05: 91; 06: < 2 |  | < 2
| 1 | 5 | 01: < 2; 02: < 6; 03: < 5; 04: < 1; 05: 73; 06: < 2 |  | 23
| 1 | 6 | 01: < 2; 02: < 6; 03: < 5; 04: < 1; 05: 125; 06: < 2
| Mean |  | 05: 91.8333
| Median |  | 05: 86
| Std Dev. |  | 05: 21.9856
| Rel.Std.Dev. |  | 05: 23.94%
PDM3
| Z-Score (Absolute) |  | 01: 0.51; 02: 0.84; 03: 0.51; 04: 0.84; 05: 60.74; 06: 0.51
| Analytical results for Th in OREAS 556b (Certified Value 15.5 ppm) | [23 empty] | Certified
Th | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 15: Lab |  | 1
| No | No | 01: 01; 02: 04; 03: 05; 04: 06; 05: 15; 06: 17; 08: 18; 09: 19; 10: 23; 11: 25; 12: 26; 15: 28 |  | ppm
|  |  | 01: PF*OES; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*OES/MS; 08: PF*OES/MS; 09: PF*MS; 10: PF*MS; 11: PF*OES; 12: PF*MS; 15: PF*OES |  | 1
2
| 1 | 1 | 01: 15.4; 02: 15.4; 03: 14.9; 04: 15.3; 05: 15.7746; 06: 17.1; 08: 14.4094; 09: NA; 10: 17.8; 11: 17.6; 12: 14.9; 15: 14 |  | 1
| 1 | 2 | 01: 15.5; 02: 14.7; 03: 13.5; 04: 15.5; 05: 14.879; 06: 16.4; 08: 14.7966; 09: NA; 10: 16.7; 11: 20.8; 12: 15; 15: 15 |  | 8
| 1 | 3 | 01: 15.1; 02: 15.5; 03: 15.1; 04: 15.2; 05: 14.9126; 06: 16.7; 08: 15.0778; 09: NA; 10: 17.1; 11: 18.2; 12: 15.2; 15: 15 |  | 16
| 1 | 4 | 01: 15.8; 02: 15.3; 03: 14.8; 04: 15.6; 05: 15.8898; 06: 16.9; 08: 14.6399; 09: NA; 10: 16.5; 11: 18.7; 12: 15; 15: 14 |  | 15.5063
| 1 | 5 | 01: 15.4; 02: 15.3; 03: 14.4; 04: 15.7; 05: 15.1884; 06: 26.9; 08: 14.3115; 09: NA; 10: 16.7; 11: 18.7; 12: 15.2; 15: 14 |  | 76
| 1 | 6 | 01: 15.6; 02: 15.3; 03: 14.4; 04: 15.5; 05: 15.4266; 06: 17.7; 08: 14.8751; 09: NA; 10: 16.3; 11: 17.9; 12: 14.8; 15: 15
| Mean |  | 01: 15.4667; 02: 15.25; 03: 14.5167; 04: 15.4667; 05: 15.3452; 06: 18.6167; 08: 14.6851; 10: 16.85; 11: 18.65; 12: 15.0167; 15: 14.5
| Median |  | 01: 15.45; 02: 15.3; 03: 14.6; 04: 15.5; 05: 15.3075; 06: 17; 08: 14.7183; 10: 16.7; 11: 18.45; 12: 15; 15: 14.5
| Std Dev. |  | 01: 0.23381; 02: 0.28107; 03: 0.57067; 04: 0.18619; 05: 0.42817; 06: 4.0814; 08: 0.28997; 10: 0.53572; 11: 1.1397; 12: 0.16021; 15: 0.54772
| Rel.Std.Dev. |  | 01: 1.512%; 02: 1.843%; 03: 3.931%; 04: 1.204%; 05: 2.79%; 06: 21.92%; 08: 1.975%; 10: 3.179%; 11: 6.111%; 12: 1.067%; 15: 3.777%
| PDM3 |  | 01: -0.2557%; 02: -1.653%; 03: -6.382%; 04: -0.2557%; 05: -1.039%; 06: 20.06%; 08: -5.296%; 10: 8.665%; 11: 20.27%; 12: -3.158%; 15: -6.49%
| Z-Score (Absolute) |  | 01: 0.07; 02: 0.19; 03: 1.08; 04: 0.07; 05: 0.07; 06: 3.9; 08: 0.88; 09: NA; 10: 1.75; 11: 3.94; 12: 0.47; 15: NA
Results from laboratory 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 556b (Certified Value 0.414 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.32 | 0.42 | 0.42 | 0.43 | 0.42 | 0.42 | 0.42 | 0.408 | 0.39 | 0.414 | 0.414 | 0.3921 | 0.4421 | 0.495 | 0.41819 | 0.4007 | 0.4 | 0.49 | 0.48 | 0.41 | 0.39 | 0.41 | 1
| 1 | 2 | 0.35 | 0.42 | 0.43 | 0.43 | 0.42 | 0.42 | 0.402 | 0.402 | 0.384 | 0.42 | 0.432 | 0.41983 | 0.446 | 0.497 | 0.4167 | 0.4053 | 0.39 | 0.48 | 0.63 | 0.42 | 0.39 | 0.41 | 25
| 1 | 3 | 0.33 | 0.42 | 0.42 | 0.42 | 0.41 | 0.41 | 0.402 | 0.408 | 0.372 | 0.414 | 0.426 | 0.39105 | 0.4408 | 0.49 | 0.41877 | 0.4024 | 0.4 | 0.47 | 0.58 | 0.4 | 0.39 | 0.44 | 16
| 1 | 4 | 0.36 | 0.42 | 0.42 | 0.43 | 0.41 | 0.42 | 0.39 | 0.42 | 0.372 | 0.402 | 0.426 | 0.4251 | 0.4412 | 0.478 | 0.41767 | 0.4129 | 0.39 | 0.46 | 0.48 | 0.43 | 0.39 | 0.41 | 0.41403
| 1 | 5 | 0.33 | 0.42 | 0.42 | 0.43 | 0.43 | 0.41 | 0.39 | 0.408 | 0.39 | 0.408 | 0.426 | 0.41951 | 0.4489 | 0.491 | 0.41758 | 0.4124 | 0.4 | 0.45 | 0.54 | 0.42 | 0.39 | 0.4 | 77
| 1 | 6 | 0.31 | 0.43 | 0.42 | 0.42 | 0.43 | 0.41 | 0.42 | 0.408 | 0.366 | 0.402 | 0.426 | 0.39758 | 0.4441 | 0.501 | 0.42013 | 0.4041 | 0.4 | 0.48 | 0.52 | 0.43 | 0.39 | 0.4
| Mean |  | 0.33333 | 0.42167 | 0.42167 | 0.42667 | 0.42 | 0.415 | 0.404 | 0.409 | 0.379 | 0.41 | 0.425 | 0.40753 | 0.44385 | 0.492 | 0.41817 | 0.4063 | 0.39667 | 0.47167 | 0.53833 | 0.41833 | 0.39 | 0.41167
| Median |  | 0.33 | 0.42 | 0.42 | 0.43 | 0.42 | 0.415 | 0.402 | 0.408 | 0.378 | 0.411 | 0.426 | 0.40855 | 0.4431 | 0.493 | 0.41793 | 0.4047 | 0.4 | 0.475 | 0.53 | 0.42 | 0.39 | 0.41
| Std Dev. |  | 0.018619 | 0.0040825 | 0.0040825 | 0.005164 | 0.0089443 | 0.0054772 | 0.013506 | 0.0058992 | 0.010334 | 0.0072664 | 0.0058992 | 0.015568 | 0.0031488 | 0.0079498 | 0.0011788 | 0.005161 | 0.005164 | 0.01472 | 0.058793 | 0.01169 | 6.0809e-17 | 0.01472
| Rel.Std.Dev. |  | 5.586% | 0.9682% | 0.9682% | 1.21% | 2.13% | 1.32% | 3.343% | 1.442% | 2.727% | 1.772% | 1.388% | 3.82% | 0.7094% | 1.616% | 0.2819% | 1.27% | 1.302% | 3.121% | 10.92% | 2.795% | 1.559e-14% | 3.576%
| PDM3 |  | -19.49% | 1.846% | 1.846% | 3.053% | 1.443% | 0.2354% | -2.421% | -1.214% | -8.46% | -0.9723% | 2.651% | -1.57% | 7.204% | 18.83% | 1.002% | -1.866% | -4.193% | 13.92% | 30.02% | 1.04% | -5.803% | -0.5697%
| Z-Score (Absolute) |  | 5.87 | 0.36 | 0.36 | 0.71 | 0.24 | 0.11 | 0.91 | 0.56 | 2.67 | 0.49 | 0.57 | 0.64 | 1.92 | 5.31 | 0.11 | 0.72 | 1.4 | 3.88 | 8.58 | 0.12 | 1.87 | 0.35
| Analytical results for Tl in OREAS 556b (Certified Value 1.06 ppm) | [23 empty] | Certified
Tl | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 01; 02: 02; 03: 04; 04: 05; 05: 06; 06: 15; 08: 17; 09: 23; 10: 25; 11: 26 | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: PF*MS | 2
3
| 1 | 1 | 01: < 5; 02: < 2; 03: 1.1; 04: 1; 05: 1.1; 06: 1.0249; 08: 1.16; 09: 0.9; 10: 1.1; 11: 1 | 1
| 1 | 2 | 01: < 5; 02: < 2; 03: 1.1; 04: 1.1; 05: 1; 06: 1.0099; 08: 1.11; 09: 0.9; 10: 1.3; 11: 1.1 | 26
| 1 | 3 | 01: < 5; 02: < 2; 03: 1.1; 04: 1; 05: 1.1; 06: 0.96018; 08: 1.14; 09: 0.9; 10: 1.2; 11: 1.1 | 16
| 1 | 4 | 01: < 5; 02: < 2; 03: 1.1; 04: 1; 05: 1.1; 06: 1.0148; 08: 1.14; 09: 1; 10: 1.2; 11: 1.1 | 1.06
| 1 | 5 | 01: < 5; 02: < 2; 03: 1.1; 04: 1; 05: 1; 06: 0.9493; 08: 1.14; 09: 0.9; 10: 1.2; 11: 1 | 78
| 1 | 6 | 01: < 5; 02: < 2; 03: 1.2; 04: 1.1; 05: 0.9; 06: 1.0195; 08: 1.11; 09: 0.9; 10: 1.2; 11: 1
| Mean |  | 03: 1.1167; 04: 1.0333; 05: 1.0333; 06: 0.99641; 08: 1.1333; 09: 0.91667; 10: 1.2; 11: 1.05
| Median |  | 03: 1.1; 04: 1; 05: 1.05; 06: 1.0123; 08: 1.14; 09: 0.9; 10: 1.2; 11: 1.05
| Std Dev. |  | 03: 0.040825; 04: 0.05164; 05: 0.08165; 06: 0.032842; 08: 0.019664; 09: 0.040825; 10: 0.063246; 11: 0.054772
| Rel.Std.Dev. |  | 03: 3.656%; 04: 4.997%; 05: 7.902%; 06: 3.296%; 08: 1.735%; 09: 4.454%; 10: 5.27%; 11: 5.216%
| PDM3 |  | 03: 5.349%; 04: -2.513%; 05: -2.513%; 06: -5.996%; 08: 6.921%; 09: -13.52%; 10: 13.21%; 11: -0.9404%
| Z-Score (Absolute) |  | 01: 16.35; 02: 0.47; 03: 0.84; 04: 0.09; 05: 0.09; 06: 0.51; 08: 1.03; 09: 1.4; 10: 1.78; 11: 0.09
| Analytical results for Tm in OREAS 556b (Certified Value 0.52 ppm) | [23 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES | 2
3
| 1 | 1 | 01: 0.5; 02: 0.51; 03: 0.54319; 04: NA; 05: 0.52; 06: 0.46; 08: 0.55 | 1
| 1 | 2 | 01: 0.5; 02: 0.53; 03: 0.53958; 04: NA; 05: 0.48; 06: 0.64; 08: 0.54 | 10
| 1 | 3 | 01: 0.5; 02: 0.5; 03: 0.53885; 04: NA; 05: 0.52; 06: 0.54; 08: 0.54 | 16
| 1 | 4 | 01: 0.5; 02: 0.57; 03: 0.51555; 04: NA; 05: 0.47; 06: 0.62; 08: 0.53 | 0.52179
| 1 | 5 | 01: 0.5; 02: 0.51; 03: 0.51351; 04: NA; 05: 0.46; 06: 0.52; 08: 0.54 | 79
| 1 | 6 | 01: 0.5; 02: 0.49; 03: 0.52496; 04: NA; 05: 0.49; 06: 0.48; 08: 0.53
| Mean |  | 01: 0.5; 02: 0.51833; 03: 0.52927; 05: 0.49; 06: 0.54333; 08: 0.53833
| Median |  | 01: 0.5; 02: 0.51; 03: 0.5319; 05: 0.485; 06: 0.53; 08: 0.54
| Std Dev. |  | 01: 0; 02: 0.028577; 03: 0.013014; 05: 0.025298; 06: 0.073121; 08: 0.0075277
| Rel.Std.Dev. |  | 01: 0%; 02: 5.513%; 03: 2.459%; 05: 5.163%; 06: 13.46%; 08: 1.398%
| PDM3 |  | 01: -4.176%; 02: -0.6621%; 03: 1.435%; 05: -6.092%; 06: 4.129%; 08: 3.171%
| Z-Score (Absolute) |  | 01: NA; 02: 0.67; 03: 0; 04: NA; 05: 2.42; 06: 0.87; 08: 0.56
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 556b (Certified Value 8.85 ppm) | [23 empty] | Certified
U | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 04; 02: 05; 03: 06; 04: 15; 05: 17; 06: 18; 08: 19; 09: 25; 10: 26; 11: 28 | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*OES/MS; 08: PF*MS; 09: PF*OES; 10: PF*MS; 11: PF*OES | 2
2
| 1 | 1 | 01: 8.7; 02: 8.7; 03: 8.7; 04: 10.0968; 05: 9.96; 06: 8.076; 08: 9; 09: 5.82; 10: 9.1; 11: 9 | 1
| 1 | 2 | 01: 8.5; 02: 8.3; 03: 8.6; 04: 9.4619; 05: 10.1; 06: 8.0909; 08: 9; 09: 9.45; 10: 9.2; 11: 9 | 11
| 1 | 3 | 01: 8.9; 02: 8.8; 03: 9; 04: 9.8501; 05: 9.92; 06: 8.2959; 08: 9; 09: 7.66; 10: 9.4; 11: 9 | 16
| 1 | 4 | 01: 8.9; 02: 8.8; 03: 8.8; 04: 10.116; 05: 9.88; 06: 7.9842; 08: 9; 09: 6.92; 10: 9.1; 11: 9 | 8.8503
| 1 | 5 | 01: 8.9; 02: 8.8; 03: 8.7; 04: 9.5379; 05: 10.2; 06: 8.6836; 08: 9; 09: 7.25; 10: 9; 11: 9 | 80
| 1 | 6 | 01: 8.8; 02: 8.8; 03: 8.6; 04: 9.7187; 05: 10.3; 06: 8.1854; 08: 9; 09: 5.83; 10: 8.8; 11: 9
| Mean |  | 01: 8.7833; 02: 8.7; 03: 8.7333; 04: 9.7969; 05: 10.06; 06: 8.2193; 08: 9; 09: 7.155; 10: 9.1; 11: 9
| Median |  | 01: 8.85; 02: 8.8; 03: 8.7; 04: 9.7844; 05: 10.03; 06: 8.1381; 08: 9; 09: 7.085; 10: 9.1; 11: 9
| Std Dev. |  | 01: 0.16021; 02: 0.2; 03: 0.15055; 04: 0.2757; 05: 0.16781; 06: 0.25096; 08: 0; 09: 1.3509; 10: 0.2; 11: 0
| Rel.Std.Dev. |  | 01: 1.824%; 02: 2.299%; 03: 1.724%; 04: 2.814%; 05: 1.668%; 06: 3.053%; 08: 0%; 09: 18.88%; 10: 2.198%; 11: 0%
| PDM3 |  | 01: -0.7562%; 02: -1.698%; 03: -1.321%; 04: 10.7%; 05: 13.67%; 06: -7.129%; 08: 1.692%; 09: -19.15%; 10: 2.822%; 11: 1.692%
| Z-Score (Absolute) |  | 01: 0.04; 02: 0.09; 03: 0.04; 04: 1.59; 05: 1.99; 06: 0.83; 08: NA; 09: 2.46; 10: 0.52; 11: NA
Results from laboratories 19 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 556b (Certified Value 353 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 15 | 16 | 17 | 19 | 21 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 320.5 | 320 | 335 | 363 | 377 | 368 | 373.706 | 363 | 350 | NA | 310 | 369 | 410 | 346 | 349 | [7 empty] | 1
| 1 | 2 | 320.5 | 330 | 338 | 369 | 370 | 368 | 362.21 | 365 | 360 | NA | 300 | 398 | 498 | 345 | 354 | [7 empty] | 38
| 1 | 3 | 312.8 | 330 | 336 | 366 | 378 | 365 | 359.9863 | 362 | 350 | NA | 310 | 405 | 461 | 349 | 341 | [7 empty] | 16
| 1 | 4 | 327.8 | 330 | 334 | 369 | 374 | 368 | 371.6755 | 365 | 350 | NA | 290 | 335 | 438 | 354 | 348 | [7 empty] | 352.7061
| 1 | 5 | 319 | 330 | 339 | 363 | 374 | 370 | 366.6708 | 360 | 350 | NA | 310 | 339 | 424 | 350 | 356 | [7 empty] | 81
| 1 | 6 | 308.5 | 330 | 335 | 351 | 380 | 363 | 373.4895 | 365 | 350 | NA | 310 | 328 | 426 | 356 | 343
| Mean |  | 318.1833 | 328.3333 | 336.1667 | 363.5 | 375.5 | 367 | 367.9564 | 363.3333 | 351.6667 |  | 305 | 362.3333 | 442.8333 | 350 | 348.5
| Median |  | 319.75 | 330 | 335.5 | 364.5 | 375.5 | 368 | 369.1732 | 364 | 350 |  | 310 | 354 | 432 | 349.5 | 348.5
| Std Dev. |  | 6.7324 | 4.0825 | 1.9408 | 6.6858 | 3.5637 | 2.5298 | 5.9279 | 2.0656 | 4.0825 |  | 8.3666 | 33.4883 | 31.9776 | 4.3359 | 5.8907
| Rel.Std.Dev. |  | 2.116% | 1.243% | 0.5773% | 1.839% | 0.9491% | 0.6893% | 1.611% | 0.5685% | 1.161% |  | 2.743% | 9.242% | 7.221% | 1.239% | 1.69%
| PDM3 |  | -9.788% | -6.91% | -4.689% | 3.06% | 6.463% | 4.053% | 4.324% | 3.013% | -0.2947% |  | -13.53% | 2.73% | 25.55% | -0.7672% | -1.193%
| Z-Score (Absolute) |  | 2.5 | 1.84 | 1.34 | 0.42 | 1.19 | 0.64 | 0.71 | 0.41 | 0.34 | NA | 3.35 | 0.34 | 5.52 | 0.45 | 0.55
| Analytical results for W in OREAS 556b (Certified Value 2.96 ppm) | [23 empty] | Certified
W | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 01; 02: 02; 03: 04; 04: 05; 05: 06; 06: 15; 08: 19; 09: 21; 10: 25; 11: 26; 12: 28 | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*MS; 08: PF*MS; 09: PF*OES; 10: PF*OES; 11: PF*MS; 12: PF*OES | 2
2
| 1 | 1 | 01: < 5; 02: < 10; 03: 3; 04: 3.9; 05: < 5; 06: 2.9072; 08: NA; 09: < 50; 10: < 1; 11: 3; 12: < 5 | 1
| 1 | 2 | 01: < 5; 02: < 10; 03: 3; 04: 3; 05: < 5; 06: 2.6269; 08: NA; 09: < 50; 10: 2; 11: 3; 12: < 5 | 12
| 1 | 3 | 01: < 5; 02: < 10; 03: 3; 04: 3.6; 05: < 5; 06: 2.4716; 08: NA; 09: < 50; 10: 2; 11: 3; 12: 7 | 16
| 1 | 4 | 01: < 5; 02: < 10; 03: 3; 04: 4.3; 05: < 5; 06: 2.4536; 08: NA; 09: < 50; 10: < 1; 11: 3; 12: < 5 | 2.9639
| 1 | 5 | 01: < 5; 02: < 10; 03: 3; 04: 4.2; 05: < 5; 06: 2.6949; 08: NA; 09: < 50; 10: 2; 11: 3; 12: < 5 | 82
| 1 | 6 | 01: < 5; 02: < 10; 03: 3; 04: 2.9; 05: < 5; 06: 2.8626; 08: NA; 09: < 50; 10: 3; 11: 3; 12: < 5
| Mean |  | 03: 3; 04: 3.65; 06: 2.6695; 10: 2.25; 11: 3; 12: 7
| Median |  | 03: 3; 04: 3.75; 06: 2.6609; 10: 2; 11: 3; 12: 7
| Std Dev. |  | 03: 0; 04: 0.59582; 06: 0.1908; 10: 0.5; 11: 0
| Rel.Std.Dev. |  | 03: 0%; 04: 16.32%; 06: 7.148%; 10: 22.22%; 11: 0%
| PDM3 |  | 03: 1.218%; 04: 23.15%; 06: -9.934%; 10: -24.09%; 11: 1.218%; 12: 136.2%
| Z-Score (Absolute) |  | 01: 0.67; 02: 2.7; 03: 0; 04: 0.88; 05: 0.67; 06: 0.45; 08: NA; 09: 29.67; 10: 1.8; 11: 0; 12: 0.34
| Analytical results for Y in OREAS 556b (Certified Value 32.6 ppm) | [23 empty] | Certified
Y | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 03; 02: 04; 03: 05; 04: 06; 05: 15; 06: 17; 08: 18; 09: 19; 10: 21; 11: 23; 12: 25; 15: 26; 16: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 05: PF*MS; 06: PF*OES/MS; 08: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: PF*MS; 12: PF*OES; 15: PF*MS; 16: PF*OES | 1
2
| 1 | 1 | 01: 32.5; 02: 32.6; 03: 33.3; 04: 31.2; 05: 29.3924; 06: 36.3; 08: 31.5164; 09: 34; 10: 30; 11: 29.5; 12: 32.6; 15: 33.1; 16: 34 | 1
| 1 | 2 | 01: 32.8; 02: 32.8; 03: 31.9; 04: 32.8; 05: 28.5357; 06: 35.4; 08: 32.2206; 09: 33; 10: 29; 11: 27.5; 12: 38.3; 15: 33.2; 16: 34 | 13
| 1 | 3 | 01: 32.3; 02: 32.9; 03: 34.5; 04: 31.7; 05: 28.7496; 06: 36.5; 08: 32.6061; 09: 33; 10: 31; 11: 28.7; 12: 36.8; 15: 33.9; 16: 33 | 16
| 1 | 4 | 01: 32.9; 02: 33.7; 03: 34.6; 04: 31.6; 05: 29.3909; 06: 36.8; 08: 31.9392; 09: 33; 10: 29; 11: 29; 12: 40.3; 15: 32.5; 16: 34 | 32.611
| 1 | 5 | 01: 32.5; 02: 33.3; 03: 33.6; 04: 31.5; 05: 28.8721; 06: 36; 08: 33.1137; 09: 49; 10: 31; 11: 28.8; 12: 32.2; 15: 33.8; 16: 35 | 83
| 1 | 6 | 01: 32.6; 02: 33.2; 03: 33.7; 04: 30.7; 05: 28.9465; 06: 35.6; 08: 33.1781; 09: 40; 10: 30; 11: 28.4; 12: 30.5; 15: 31.7; 16: 35
| Mean |  | 01: 32.6; 02: 33.0833; 03: 33.6; 04: 31.5833; 05: 28.9812; 06: 36.1; 08: 32.429; 09: 37; 10: 30; 11: 28.65; 12: 35.1167; 15: 33.0333; 16: 34.1667
| Median |  | 01: 32.55; 02: 33.05; 03: 33.65; 04: 31.55; 05: 28.9093; 06: 36.15; 08: 32.4134; 09: 33.5; 10: 30; 11: 28.75; 12: 34.7; 15: 33.15; 16: 34
| Std Dev. |  | 01: 0.21909; 02: 0.39707; 03: 0.9798; 04: 0.6969; 05: 0.34704; 06: 0.53666; 08: 0.65992; 09: 6.4807; 10: 0.89443; 11: 0.67157; 12: 3.8984; 15: 0.82865; 16: 0.75277
| Rel.Std.Dev. |  | 01: 0.6721%; 02: 1.2%; 03: 2.916%; 04: 2.207%; 05: 1.197%; 06: 1.487%; 08: 2.035%; 09: 17.52%; 10: 2.981%; 11: 2.344%; 12: 11.1%; 15: 2.509%; 16: 2.203%
| PDM3 |  | 01: -0.03386%; 02: 1.448%; 03: 3.033%; 04: -3.151%; 05: -11.13%; 06: 10.7%; 08: -0.5581%; 09: 13.46%; 10: -8.007%; 11: -12.15%; 12: 7.683%; 15: 1.295%; 16: 4.77%
| Z-Score (Absolute) |  | 01: 0.2; 02: 0.02; 03: 0.26; 04: 0.67; 05: 1.88; 06: 1.43; 08: 0.28; 09: 1.84; 10: 1.41; 11: 2.04; 12: 0.97; 15: 0; 16: 0.53
| Analytical results for Yb in OREAS 556b (Certified Value 3.19 ppm) | [23 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | 1
| No | No | 01: 05; 02: 06; 03: 18; 04: 19; 05: 23; 06: 25; 08: 28 | ppm
|  |  | 01: PF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS; 06: PF*OES; 08: PF*OES | 2
3
| 1 | 1 | 01: 3.1; 02: 3.4; 03: 3.3397; 04: 3; 05: 3; 06: 2.4; 08: 3.29 | 1
| 1 | 2 | 01: 2.8; 02: 3.5; 03: 3.1848; 04: 3; 05: 2.9; 06: 3.4; 08: 3.36 | 14
| 1 | 3 | 01: 3.4; 02: 3.3; 03: 3.286; 04: 3; 05: 3; 06: 2.9; 08: 3.37 | 16
| 1 | 4 | 01: 3.4; 02: 3.4; 03: 3.3391; 04: 3; 05: 3.1; 06: 3.5; 08: 3.39 | 3.1916
| 1 | 5 | 01: 3.3; 02: 3.4; 03: 3.2586; 04: 3; 05: 2.8; 06: 2.6; 08: 3.48 | 84
| 1 | 6 | 01: 3; 02: 3.3; 03: 3.4578; 04: 3; 05: 3.1; 06: 2.8; 08: 3.34
| Mean |  | 01: 3.1667; 02: 3.3833; 03: 3.311; 04: 3; 05: 2.9833; 06: 2.9333; 08: 3.3717
| Median |  | 01: 3.2; 02: 3.4; 03: 3.3126; 04: 3; 05: 3; 06: 2.85; 08: 3.365
| Std Dev. |  | 01: 0.24221; 02: 0.075277; 03: 0.092161; 04: 0; 05: 0.1169; 06: 0.43665; 08: 0.063061
| Rel.Std.Dev. |  | 01: 7.649%; 02: 2.225%; 03: 2.783%; 04: 0%; 05: 3.919%; 06: 14.89%; 08: 1.87%
| PDM3 |  | 01: -0.7798%; 02: 6.009%; 03: 3.742%; 04: -6.002%; 05: -6.524%; 06: -8.091%; 08: 5.643%
| Z-Score (Absolute) |  | 01: 0.35; 02: 0.7; 03: 0.35; 04: NA; 05: 1.24; 06: 1.49; 08: 0.65
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 556b (Certified Value 41.4 ppm) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 21 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 1
2
| 1 | 1 | 40.6 | 88 | 42.9 | 43.16 | 40 | 40 | 30 | 20 | 40 | 20 | 60 | 38.1314 | 60 | 40.6346 | NA | 40 | 35 | 26 | 39 | < 100 | < 100 |  | 1
| 1 | 2 | 46.4 | 88 | 42.3 | 46.27 | 40 | 39 | 20 | 50 | 40 | 20 | 190 | 39.0509 | 60 | 40.5423 | NA | 40 | 50 | 44 | 42 | < 100 | < 100 |  | 15
| 1 | 3 | 35.6 | 84 | 42.1 | 48.43 | 50 | 43 | 20 | 70 | 30 | 30 | 320 | 39.12 | 70 | 42.3687 | NA | 40 | 45 | 39 | 41 | < 100 | < 100 |  | 16
| 1 | 4 | 46.2 | 87 | 43.4 | 50.95 | 40 | 38 | < 20 | 30 | 30 | 20 | < 20 | 42.7839 | 60 | 42.0511 | NA | 40 | 45 | 41 | 46 | < 100 | < 100 |  | 41.353
| 1 | 5 | 39.2 | 84 | 42.1 | 54.68 | 60 | 41 | 20 | 30 | 40 | 20 | 20 | 41.3395 | 60 | 41.6403 | NA | 40 | 30 | 45 | 44 | < 100 | < 100 |  | 85
| 1 | 6 | 42.9 | 86 | 42.3 | 43.58 | 50 | 37 | < 20 | 20 | 30 | 30 | 60 | 43.6373 | 60 | 40.0425 | NA | 40 | 45 | 34 | 44 | < 100 | < 100
| Mean |  | 41.8167 | 86.1667 | 42.5167 | 47.845 | 46.6667 | 39.6667 | 22.5 | 36.6667 | 35 | 23.3333 | 130 | 40.6772 | 61.6667 | 41.2133 |  | 40 | 41.6667 | 38.1667 | 42.6667
| Median |  | 41.75 | 86.5 | 42.3 | 47.35 | 45 | 39.5 | 20 | 30 | 35 | 20 | 60 | 40.2297 | 60 | 41.1375 |  | 40 | 45 | 40 | 43
| Std Dev. |  | 4.204 | 1.8348 | 0.52313 | 4.4569 | 8.165 | 2.1602 | 5 | 19.6638 | 5.4772 | 5.164 | 124.0967 | 2.2446 | 4.0825 | 0.93542 |  | 0 | 7.5277 | 7.1391 | 2.5033
| Rel.Std.Dev. |  | 10.05% | 2.129% | 1.23% | 9.315% | 17.5% | 5.446% | 22.22% | 53.63% | 15.65% | 22.13% | 95.46% | 5.518% | 6.62% | 2.27% |  | 0% | 18.07% | 18.71% | 5.867%
| PDM3 |  | 1.121% | 108.4% | 2.814% | 15.7% | 12.85% | -4.078% | -45.59% | -11.33% | -15.36% | -43.58% | 214.4% | -1.634% | 49.12% | -0.3378% |  | -3.272% | 0.7586% | -7.705% | 3.177%
| Z-Score (Absolute) |  | 0.12 | 14.87 | 0.12 | 1.92 | NA | 0.85 | NA | NA | NA | NA | NA | 0.5 | NA | 0.32 | NA | NA | 0.17 | 1.35 | 0.17 | 2.65 | 2.65
Results from laboratories 5, 8, 9, 10, 11, 12, 17 and 21 were removed due to their 10 ppm reading resolution.
| Analytical results for Zr in OREAS 556b (Indicative Value 209 ppm) | [23 empty] | Indicative
Zr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 01: 01; 02: 06; 03: 18; 04: 19; 05: 25 |  | ppm
|  |  | 01: PF*OES; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES |  | 0
0
| 1 | 1 | 01: 168.3; 02: 202; 03: 217.0629; 04: 208; 05: 175 |  | 1
| 1 | 2 | 01: 167.9; 02: 208; 03: 219.9061; 04: 199; 05: 241 |  | 18
| 1 | 3 | 01: 162; 02: 206; 03: 223.1253; 04: 203; 05: 214 |  | 16
| 1 | 4 | 01: 171.8; 02: 197; 03: 207.3312; 04: 206; 05: 222 |  | 209.2336
| 1 | 5 | 01: 170.9; 02: 209; 03: 221.4965; 04: 195; 05: 216 |  | 24
| 1 | 6 | 01: 164.1; 02: 202; 03: 221.4144; 04: 208; 05: 187
| Mean |  | 01: 167.5; 02: 204; 03: 218.3894; 04: 203.1667; 05: 209.1667
| Median |  | 01: 168.1; 02: 204; 03: 220.6602; 04: 204.5; 05: 215
| Std Dev. |  | 01: 3.8121; 02: 4.5166; 03: 5.7893; 04: 5.2694; 05: 24.1115
| Rel.Std.Dev. |  | 01: 2.276%; 02: 2.214%; 03: 2.651%; 04: 2.594%; 05: 11.53%
| PDM3 |  | 01: -19.95%; 02: -2.501%; 03: 4.376%; 04: -2.9%; 05: -0.03199%
| Z-Score (Absolute) |  | 01: 4.76; 02: 0; 03: 1.88; 04: 0.11; 05: 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 556b (Certified Value 0.144 ppm) | [28 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 3.4 | 0.37 | < 0.5 | 0.13 | 0.5 | 0.16 | 0.14 | 0.05 | 0.11 | 0.15 | 0.14 | < 5 | 6.8 | 0.13041 | 0.1 | 0.23684 | < 0.3 | 0.2 | 0.2 | 0.3 | 0.15 | 0.14 | < 0.5 | 0.15 |  |  |  | 1
| 1 | 2 | 3.4 | 0.4 | < 0.5 | 0.14 | 0.7 | 0.15 | 0.15 | 0.06 | 0.06 | 0.15 | 0.14 | < 5 | 4.2 | 0.12459 | 0.1 | 0.23914 | < 0.3 | 0.2 | < 0.2 | 0.2 | 0.14 | 0.15 | < 0.5 | 0.14 |  |  |  | 1
| 1 | 3 | 3.54 | 0.37 | < 0.5 | 0.14 | < 0.3 | 0.16 | 0.15 | 0.07 | 0.08 | 0.16 | 0.13 | < 5 | 4.1 | 0.13166 | 0.1 | 0.23394 | < 0.3 | 0.2 | 0.2 | 0.3 | 0.15 | 0.14 | < 0.5 | 0.15 |  |  |  | 16
| 1 | 4 | 3.74 | 0.35 | < 0.5 | 0.14 | < 0.3 | 0.16 | 0.13 | 0.05 | 0.17 | 0.16 | 0.13 | < 5 | 7.5 | 0.12909 | 0.2 | 0.24052 | < 0.3 | 0.2 | < 0.2 | 0.2 | 0.15 | 0.14 | < 0.5 | 0.15 |  |  |  | 0.14397
| 1 | 5 | 3.45 | 0.39 | < 0.5 | 0.13 | 0.4 | 0.15 | 0.15 | 0.06 | 0.17 | 0.17 | 0.15 | < 5 | 3.4 | 0.12323 | 0.2 | 0.23942 | < 0.3 | 0.2 | 0.2 | 0.2 | 0.15 | 0.14 | < 0.5 | 0.15 |  |  |  | 87
| 1 | 6 | 3.44 | 0.33 | < 0.5 | 0.12 | 0.4 | 0.15 | 0.14 | 0.06 | 0.17 | 0.14 | 0.13 | < 5 | 5.2 | 0.12932 | 0.2 | 0.23327 | < 0.3 | 0.2 | < 0.2 | 0.2 | 0.14 | 0.14 | < 0.5 | 0.13
| Mean |  | 3.495 | 0.36833 |  | 0.13333 | 0.5 | 0.155 | 0.14333 | 0.058333 | 0.12667 | 0.155 | 0.13667 |  | 5.2 | 0.12805 | 0.15 | 0.23719 |  | 0.2 | 0.2 | 0.23333 | 0.14667 | 0.14167 |  | 0.145
| Median |  | 3.445 | 0.37 |  | 0.135 | 0.45 | 0.155 | 0.145 | 0.06 | 0.14 | 0.155 | 0.135 |  | 4.7 | 0.12921 | 0.15 | 0.23799 |  | 0.2 | 0.2 | 0.2 | 0.15 | 0.14 |  | 0.15
| Std Dev. |  | 0.1305 | 0.025626 |  | 0.008165 | 0.14142 | 0.0054772 | 0.008165 | 0.0075277 | 0.050067 | 0.010488 | 0.008165 |  | 1.631 | 0.0033634 | 0.054772 | 0.0030273 |  | 3.0405e-17 | 3.3993e-17 | 0.05164 | 0.005164 | 0.0040825 |  | 0.0083666
| Rel.Std.Dev. |  | 3.734% | 6.957% |  | 6.124% | 28.28% | 3.534% | 5.696% | 12.9% | 39.53% | 6.767% | 5.974% |  | 31.36% | 2.627% | 36.51% | 1.276% |  | 1.52e-14% | 1.7e-14% | 22.13% | 3.521% | 2.882% |  | 5.77%
| PDM3 |  | 2328% | 155.8% |  | -7.389% | 247.3% | 7.66% | -0.4434% | -59.48% | -12.02% | 7.66% | -5.074% |  | 3512% | -11.06% | 4.187% | 64.75% |  | 38.92% | 38.92% | 62.07% | 1.872% | -1.601% |  | 0.7142%
| Z-Score (Absolute) |  | 127.92 | 8.41 | 3.89 | 0.57 | NA | 0.25 | 0.19 | 3.44 | 0.83 | 0.25 | 0.45 | 89.89 | NA | 0.78 | NA | 3.4 | 0.06 | NA | NA | NA | 0.06 | 0.25 | 3.89 | 0.13
Results from laboratories 5, 14, 17, 20, 21 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 556b (Certified Value 5.67 wt.%) | [28 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.24 | 6.04 | 5.61 | 5.5575 | 5.77 | 5.03 | 5.49 | 5.49 | 5.71 | 5.63 | 5.88 | 5.7672 | 5.075 | 5.1673 | 5.6941 | 5.83 | > 5 | 5.4369 | 5.47 | 5.6 | 5.8 | 5.94 | 5.61 | 5.4728 | 5.72 | 5.6 |  | 1
| 1 | 2 | 5.21 | 5.82 | 5.58 | 5.6463 | 5.82 | 5.01 | 5.49 | 5.6 | 5.53 | 5.53 | 5.88 | 5.794 | 5.234 | 5.2131 | 5.6637 | 5.76 | > 5 | 5.5031 | 5.26 | 5.77 | 6.02 | 5.75 | 5.97 | 5.5434 | 5.71 | 5.68 |  | 0
| 1 | 3 | 5.13 | 6.27 | 5.65 | 5.5333 | 5.81 | 5 | 5.59 | 5.62 | 5.65 | 5.74 | 5.96 | 5.7842 | 5.261 | 5.5959 | 5.6368 | 5.82 | > 5 | 5.4428 | 5.52 | 5.6 | 5.84 | 5.59 | 5.69 | 5.5015 | 5.72 | 5.68 |  | 16
| 1 | 4 | 5.22 | 5.52 | 5.6 | 5.6016 | 5.83 | 5.08 | 5.6 | 5.64 | 5.73 | 5.64 | 5.87 | 5.7688 | 5.208 | 5.9332 | 5.6889 | 5.7 | > 5 | 5.4026 | 5.6 | 5.69 | 5.97 | 5.72 | 5.57 | 5.5398 | 5.67 | 5.64 |  | 5.6681
| 1 | 5 | 5.26 | 5.72 | 5.64 | 5.5286 | 5.79 | 5.15 | 5.74 | 5.61 | 5.76 | 5.72 | 5.75 | 5.7834 | 5.277 | 5.646 | 5.639 | 5.76 | > 5 | 5.349 | 5.44 | 5.72 | 5.91 | 5.74 | 5.5 | 5.5409 | 5.66 | 5.72 |  | 88
| 1 | 6 | 5.2 | 5.98 | 5.58 | 5.6403 | 5.7 | 5.1 | 5.71 | 5.54 | 5.78 | 5.95 | 5.74 | 5.7847 | 5.208 | 5.9554 | 5.6358 | 5.67 | > 5 | 5.4308 | 5.24 | 5.62 | 5.91 | 5.73 | 5.63 | 5.5068 | 5.69 | 5.67
| Mean |  | 5.21 | 5.8917 | 5.61 | 5.5846 | 5.7867 | 5.0617 | 5.6033 | 5.5833 | 5.6933 | 5.7017 | 5.8467 | 5.7804 | 5.2105 | 5.5852 | 5.6597 | 5.7567 |  | 5.4275 | 5.4217 | 5.6667 | 5.9083 | 5.745 | 5.6617 | 5.5175 | 5.695 | 5.665
| Median |  | 5.215 | 5.9 | 5.605 | 5.5796 | 5.8 | 5.055 | 5.595 | 5.605 | 5.72 | 5.68 | 5.875 | 5.7838 | 5.221 | 5.621 | 5.6513 | 5.76 |  | 5.4338 | 5.455 | 5.655 | 5.91 | 5.735 | 5.62 | 5.5233 | 5.7 | 5.675
| Std Dev. |  | 0.044721 | 0.26294 | 0.029665 | 0.052346 | 0.04761 | 0.058452 | 0.10577 | 0.056804 | 0.091797 | 0.14275 | 0.085245 | 0.010351 | 0.071969 | 0.33913 | 0.026744 | 0.063456 |  | 0.050658 | 0.14372 | 0.070899 | 0.08085 | 0.11221 | 0.16376 | 0.028581 | 0.025884 | 0.040866
| Rel.Std.Dev. |  | 0.8584% | 4.463% | 0.5288% | 0.9373% | 0.8227% | 1.155% | 1.888% | 1.017% | 1.612% | 2.504% | 1.458% | 0.1791% | 1.381% | 6.072% | 0.4725% | 1.102% |  | 0.9333% | 2.651% | 1.251% | 1.368% | 1.953% | 2.892% | 0.518% | 0.4545% | 0.7214%
| PDM3 |  | -8.083% | 3.944% | -1.026% | -1.474% | 2.091% | -10.7% | -1.143% | -1.496% | 0.4447% | 0.5917% | 3.15% | 1.98% | -8.074% | -1.464% | -0.1484% | 1.562% |  | -4.245% | -4.348% | -0.02579% | 4.238% | 1.356% | -0.114% | -2.657% | 0.4741% | -0.05519%
| Z-Score (Absolute) |  | 3.65 | 1.86 | 0.42 | 0.62 | 1.01 | 4.86 | 0.47 | 0.63 | 0.26 | 0.32 | 1.5 | 0.96 | 3.65 | 0.62 | 0.02 | 0.77 | NA | 1.89 | 1.94 | 0.04 | 2 | 0.67 | 0 | 1.17 | 0.27 | 0.03
| Analytical results for As in OREAS 556b (Certified Value 98 ppm) | [28 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 73.34 | 92.4 | 105.8 | 100.3 | 97.4 | 98.3 | 98.7 | 95.2 | 99.7 | 100 | 103 | 88.55 | 314 | 95.0252 | 106.19 | 103 | 100.2913 | 91.8 | 103 | 94 | 39.2 | 61 | 94.9 | 98 | 33 | 97 |  | 1
| 1 | 2 | 69.26 | 88.7 | 105.5 | 98.6 | 97.2 | 95.2 | 99.6 | 92.3 | 100.5 | 97.5 | 101.5 | 90.83 | 324 | 93.6477 | 102.89 | 103 | 100.215 | 92.9 | 101 | 91 | 39.6 | 59 | 91.8 | 98 | 31 | 96.2 |  | 2
| 1 | 3 | 66.51 | 99.6 | 103.3 | 98.9 | 96.6 | 101 | 104.5 | 100 | 99.7 | 100.5 | 101 | 90.12 | 296 | 94.4732 | 106.92 | 101 | 100.4798 | 98.8 | 103 | 97 | 41.4 | 59 | 92.7 | 99 | 34 | 96.1 |  | 16
| 1 | 4 | 71.18 | 86.3 | 104.3 | 94.8 | 97.6 | 94.8 | 104.5 | 98.5 | 102.5 | 100 | 104 | 90.24 | 404 | 96.5351 | 105 | 103 | 100.0914 | 91.5 | 105 | 95 | 38.8 | 58 | 96.7 | 98 | 35 | 97.6 |  | 98.2191
| 1 | 5 | 69.31 | 92.8 | 104.3 | 99.9 | 101 | 96.7 | 106 | 96.8 | 101 | 103 | 98.6 | 90.26 | 285 | 96.6647 | 103.9 | 104 | 99.0547 | 91.1 | 104 | 91 | 41.5 | 58 | 97.2 | 98 | 34 | 98.3 |  | 89
| 1 | 6 | 72.91 | 96 | 104.4 | 96.5 | 98.7 | 93.1 | 107.5 | 92.7 | 103 | 105 | 96.5 | 90.32 | 275 | 96.2209 | 100.55 | 102 | 98.4768 | 95 | 103 | 97 | 39.7 | 57 | 88.9 | 99 | 34 | 96.7
| Mean |  | 70.4183 | 92.6333 | 104.6 | 98.1667 | 98.0833 | 96.5167 | 103.4667 | 95.9167 | 101.0667 | 101 | 100.7667 | 90.0533 | 316.3333 | 95.4278 | 104.2417 | 102.6667 | 99.7682 | 93.5167 | 103.1667 | 94.1667 | 40.0333 | 58.6667 | 93.7 | 98.3333 | 33.5 | 96.9833
| Median |  | 70.245 | 92.6 | 104.35 | 98.75 | 97.5 | 95.95 | 104.5 | 96 | 100.75 | 100.25 | 101.25 | 90.25 | 305 | 95.623 | 104.45 | 103 | 100.1532 | 92.35 | 103 | 94.5 | 39.65 | 58.5 | 93.8 | 98 | 34 | 96.85
| Std Dev. |  | 2.5757 | 4.8069 | 0.91214 | 2.1163 | 1.5855 | 2.8167 | 3.5353 | 3.1006 | 1.4038 | 2.6268 | 2.7905 | 0.7767 | 46.5904 | 1.2351 | 2.3288 | 1.0328 | 0.80756 | 2.9444 | 1.3292 | 2.7142 | 1.1431 | 1.3663 | 3.173 | 0.5164 | 1.3784 | 0.84715
| Rel.Std.Dev. |  | 3.658% | 5.189% | 0.872% | 2.156% | 1.616% | 2.918% | 3.417% | 3.233% | 1.389% | 2.601% | 2.769% | 0.8625% | 14.73% | 1.294% | 2.234% | 1.006% | 0.8094% | 3.149% | 1.288% | 2.882% | 2.855% | 2.329% | 3.386% | 0.5252% | 4.115% | 0.8735%
| PDM3 |  | -28.3% | -5.687% | 6.497% | -0.05338% | -0.1382% | -1.733% | 5.343% | -2.344% | 2.899% | 2.831% | 2.594% | -8.314% | 222.1% | -2.842% | 6.132% | 4.528% | 1.577% | -4.788% | 5.037% | -4.126% | -59.24% | -40.27% | -4.601% | 0.1163% | -65.89% | -1.258%
| Z-Score (Absolute) |  | 4.66 | 0.84 | 1.21 | 0.11 | 0.09 | 0.17 | 1.02 | 0.28 | 0.61 | 0.6 | 0.56 | 1.29 | 37.59 | 0.36 | 1.15 | 0.88 | 0.38 | 0.69 | 0.97 | 0.58 | 9.88 | 6.68 | 0.66 | 0.14 | 11 | 0.09
| Analytical results for Ba in OREAS 556b (Certified Value 190 ppm) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 194.8 | 28 | 196 | 206.1 | 188 | 204 | 180.4448 | 190 | 190 | 100 | 170 | 200 | 172 | 183 | 189.7536 | 190.87 | 194 | 98.1191 | 188 | 194 | 394 | 158 | 196 | 198 | 87 |  |  | 1
| 1 | 2 | 192.4 | 25 | 191 | 206.6 | 181 | 200 | 175.5753 | 190 | 190 | 80 | 160 | 200 | 171 | 177 | 195.108 | 192.18 | 189 | 100.0029 | 190 | 201 | 336 | 157 | 196 | 197 | 72 |  |  | 30
| 1 | 3 | 182.7 | 18 | 193 | 204.4 | 181 | 205 | 180.5191 | 190 | 190 | 140 | 170 | 210 | 170 | 182 | 195.4212 | 191.52 | 192 | 102.5882 | 189 | 197 | 397 | 379 | 191 | 198 | 68 |  |  | 16
| 1 | 4 | 186.7 | 23 | 198 | 197.5 | 184 | 198 | 179.6858 | 190 | 190 | 130 | 170 | 210 | 173 | 177 | 193.0782 | 193.22 | 191 | 100.25 | 187 | 202 | 398 | 146 | 194 | 189 | 69 |  |  | 189.7248
| 1 | 5 | 193.3 | 27 | 193 | 198.6 | 187 | 207 | 168.7865 | 190 | 180 | 80 | 150 | 200 | 172 | 181 | 200.8657 | 188.1 | 189 | 102.5673 | 185 | 195 | 365 | 146 | 188 | 189 | 56 |  |  | 90
| 1 | 6 | 190.7 | 43 | 195 | 197.8 | 192 | 201 | 167.5671 | 190 | 190 | 100 | 190 | 200 | 173 | 185 | 199.7073 | 194.62 | 183 | 99.9853 | 187 | 201 | 387 | 149 | 193 | 185 | 51
| Mean |  | 190.1 | 27.3333 | 194.3333 | 201.8333 | 185.5 | 202.5 | 175.4298 | 190 | 188.3333 | 105 | 168.3333 | 203.3333 | 171.8333 | 180.8333 | 195.6557 | 191.7517 | 189.6667 | 100.5855 | 187.6667 | 198.3333 | 379.5 | 189.1667 | 193 | 192.6667 | 67.1667
| Median |  | 191.55 | 26 | 194 | 201.5 | 185.5 | 202.5 | 177.6305 | 190 | 190 | 100 | 170 | 200 | 172 | 181.5 | 195.2646 | 191.85 | 190 | 100.1265 | 187.5 | 199 | 390.5 | 153 | 193.5 | 193 | 68.5
| Std Dev. |  | 4.5673 | 8.4538 | 2.5033 | 4.3131 | 4.3243 | 3.3912 | 5.9185 | 0 | 4.0825 | 25.0998 | 13.2916 | 5.164 | 1.169 | 3.2506 | 4.1326 | 2.2241 | 3.7771 | 1.7224 | 1.7512 | 3.4448 | 24.5662 | 93.1481 | 3.0984 | 5.6804 | 12.703
| Rel.Std.Dev. |  | 2.403% | 30.93% | 1.288% | 2.137% | 2.331% | 1.675% | 3.374% | 0% | 2.168% | 23.9% | 7.896% | 2.54% | 0.6803% | 1.798% | 2.112% | 1.16% | 1.991% | 1.712% | 0.9331% | 1.737% | 6.473% | 49.24% | 1.605% | 2.948% | 18.91%
| PDM3 |  | 0.1978% | -85.59% | 2.429% | 6.382% | -2.227% | 6.734% | -7.535% | 0.1451% | -0.7334% | -44.66% | -11.27% | 7.173% | -9.43% | -4.686% | 3.126% | 1.068% | -0.03062% | -46.98% | -1.085% | 4.537% | 100% | -0.2942% | 1.726% | 1.551% | -64.6%
| Z-Score (Absolute) |  | 0.03 | 12.63 | 0.36 | 0.95 | 0.32 | 1 | 1.11 | 0.03 | 0.1 | 6.59 | 1.66 | 1.06 | 1.39 | 0.69 | 0.47 | 0.16 | 0 | 6.93 | 0.16 | 0.67 | 14.77 | 0.04 | 0.26 | 0.23 | 9.53
| Analytical results for Be in OREAS 556b (Certified Value 4.52 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 6.5 | 5 | < 5 | 4.64 | 5.5 | 4.65 | 4.8228 | 4.34 | 4.25 | 4.6 | 4.34 | 4.63 | < 5 | 4.3 | 4.7916 | 5.1039 | 4.5 | 4.5224 | 4.21 | 4.5 | 3.7 | 4 | 3.15 | 4.71 | 4.33 | < 5 | 4.76 | 1
| 1 | 2 | 6.81 | 4 | < 5 | 4.59 | 5.4 | 4.18 | 4.841 | 4.34 | 4.41 | 4.5 | 4.29 | 4.56 | < 5 | 4.3 | 4.6797 | 5.1936 | 4.4 | 4.5117 | 4.26 | 4.8 | 3.6 | 4 | 3.03 | 4.51 | 4.44 | < 5 | 4.73 | 3
| 1 | 3 | 6.41 | 5 | < 5 | 4.59 | 5.4 | 4.5 | 4.7369 | 4.51 | 4.37 | 4.62 | 4.41 | 4.61 | < 5 | 4.2 | 4.8389 | 5.0435 | 4.5 | 4.5249 | 4.23 | 5.2 | 3.9 | 4 | 3.13 | 4.8 | 4.35 | < 5 | 4.75 | 16
| 1 | 4 | 6.43 | 4 | < 5 | 4.61 | 5.3 | 4.46 | 4.8413 | 4.4 | 4.34 | 4.67 | 4.33 | 4.66 | < 5 | 4.4 | 4.7073 | 5.0061 | 4.4 | 4.5286 | 4.21 | 4.8 | 3.8 | 4 | 3.17 | 4.54 | 4.37 | < 5 | 4.76 | 4.5183
| 1 | 5 | 6.62 | 4 | < 5 | 4.65 | 5.4 | 4.44 | 4.8793 | 4.55 | 4.38 | 4.59 | 4.37 | 4.5 | < 5 | 4.1 | 4.7554 | 5.9129 | 4.5 | 4.4624 | 4.14 | 5.1 | 3.6 | 4 | 3.13 | 4.37 | 4.41 | < 5 | 4.81 | 91
| 1 | 6 | 6.53 | 5 | < 5 | 4.68 | 5.1 | 4.44 | 4.7624 | 4.46 | 4.32 | 4.66 | 4.62 | 4.45 | < 5 | 4.2 | 4.8025 | 5.832 | 4.5 | 4.3509 | 4.23 | 4.8 | 3.9 | 4 | 3.02 | 3.94 | 4.37 | < 5 | 4.78
| Mean |  | 6.55 | 4.5 |  | 4.6267 | 5.35 | 4.445 | 4.814 | 4.4333 | 4.345 | 4.6067 | 4.3933 | 4.5683 |  | 4.25 | 4.7626 | 5.3487 | 4.4667 | 4.4835 | 4.2133 | 4.8667 | 3.75 | 4 | 3.105 | 4.4783 | 4.3783 |  | 4.765
| Median |  | 6.515 | 4.5 |  | 4.625 | 5.4 | 4.45 | 4.8319 | 4.43 | 4.355 | 4.61 | 4.355 | 4.585 |  | 4.25 | 4.7735 | 5.1488 | 4.5 | 4.517 | 4.22 | 4.8 | 3.75 | 4 | 3.13 | 4.525 | 4.37 |  | 4.76
| Std Dev. |  | 0.14792 | 0.54772 |  | 0.036148 | 0.13784 | 0.15202 | 0.053715 | 0.088015 | 0.056125 | 0.06121 | 0.1181 | 0.08085 |  | 0.10488 | 0.060394 | 0.41143 | 0.05164 | 0.069419 | 0.040332 | 0.25033 | 0.13784 | 0 | 0.063797 | 0.30446 | 0.040208 |  | 0.027386
| Rel.Std.Dev. |  | 2.258% | 12.17% |  | 0.7813% | 2.576% | 3.42% | 1.116% | 1.985% | 1.292% | 1.329% | 2.688% | 1.77% |  | 2.468% | 1.268% | 7.692% | 1.156% | 1.548% | 0.9572% | 5.144% | 3.676% | 0% | 2.055% | 6.799% | 0.9183% |  | 0.5747%
| PDM3 |  | 44.97% | -0.4052% |  | 2.398% | 18.41% | -1.622% | 6.543% | -1.881% | -3.836% | 1.956% | -2.766% | 1.107% |  | -5.938% | 5.407% | 18.38% | -1.143% | -0.771% | -6.75% | 7.71% | -17% | -11.47% | -31.28% | -0.8847% | -3.098% |  | 5.46%
| Z-Score (Absolute) |  | 5.55 | NA | 5.23 | 0.43 | 2.35 | 0.06 | 0.92 | 0.09 | 0.32 | 0.37 | 0.2 | 0.27 | 5.23 | 0.58 | 0.79 | 2.35 | 0 | 0.04 | 0.67 | 1.06 | 1.91 | NA | 3.62 | 0.03 | 0.24 | 5.23 | 0.79
Results from laboratories 2 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 556b (Certified Value 2.89 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 5 | 2.71 | < 5 | 3.18 | 2.91 | 3.07 | 2.9417 | 2.69 | 2.54 | 2.92 | 2.78 | 3.18 | < 5 | 2.7037 | 2.97 | 2.8419 | 2.78 | 2.96 | < 5 | 2.8 | 3.44 | 3.15 | 3.02 | 3.2 | 2.85 |  |  | 1
| 1 | 2 | < 5 | 2.58 | < 5 | 3.13 | 2.99 | 3 | 2.9491 | 2.74 | 2.38 | 2.88 | 2.68 | 2.85 | < 5 | 2.7169 | 2.95 | 2.8194 | 2.9 | 2.85 | < 5 | 2.7 | 3.42 | 3.69 | 3 | 3.1 | 2.79 |  |  | 4
| 1 | 3 | < 5 | 2.69 | < 5 | 3.05 | 2.93 | 3.08 | 2.9684 | 2.76 | 2.61 | 3.05 | 2.77 | 3.06 | < 5 | 2.74 | 2.92 | 2.8011 | 3.65 | 2.89 | < 5 | 2.8 | 3.37 | 3.5 | 3.02 | 3.1 | 2.79 |  |  | 16
| 1 | 4 | < 5 | 2.56 | < 5 | 3.11 | 3.04 | 3.06 | 2.9593 | 2.69 | 2.57 | 3.01 | 2.72 | 3.06 | < 5 | 2.7438 | 3.02 | 2.8136 | 2.74 | 2.98 | < 5 | 2.8 | 3.51 | 3.43 | 3.06 | 3 | 2.85 |  |  | 2.894
| 1 | 5 | < 5 | 2.49 | < 5 | 3.03 | 3.01 | 2.99 | 2.9549 | 2.76 | 2.53 | 3.07 | 2.72 | 2.91 | < 5 | 2.7621 | 2.91 | 2.9152 | 2.72 | 2.95 | < 5 | 2.7 | 3.34 | 3.37 | 3.04 | 3 | 2.84 |  |  | 92
| 1 | 6 | < 5 | 2.53 | < 5 | 3.07 | 2.96 | 2.99 | 2.9614 | 2.77 | 2.51 | 2.92 | 2.85 | 2.97 | < 5 | 2.7142 | 2.9 | 2.7467 | 2.73 | 2.9 | < 5 | 2.9 | 3.38 | 3.2 | 3.08 | 3 | 2.93
| Mean |  |  | 2.5933 |  | 3.095 | 2.9733 | 3.0317 | 2.9558 | 2.735 | 2.5233 | 2.975 | 2.7533 | 3.005 |  | 2.7301 | 2.945 | 2.823 | 2.92 | 2.9217 |  | 2.7833 | 3.41 | 3.39 | 3.0367 | 3.0667 | 2.8417
| Median |  |  | 2.57 |  | 3.09 | 2.975 | 3.03 | 2.9571 | 2.75 | 2.535 | 2.965 | 2.745 | 3.015 |  | 2.7284 | 2.935 | 2.8165 | 2.76 | 2.925 |  | 2.8 | 3.4 | 3.4 | 3.03 | 3.05 | 2.845
| Std Dev. |  |  | 0.088242 |  | 0.055767 | 0.049261 | 0.042622 | 0.0094355 | 0.036194 | 0.078401 | 0.078677 | 0.059889 | 0.11912 |  | 0.022102 | 0.045056 | 0.055247 | 0.3637 | 0.049565 |  | 0.075277 | 0.060663 | 0.1989 | 0.029439 | 0.08165 | 0.051543
| Rel.Std.Dev. |  |  | 3.403% |  | 1.802% | 1.657% | 1.406% | 0.3192% | 1.323% | 3.107% | 2.645% | 2.175% | 3.964% |  | 0.8096% | 1.53% | 1.957% | 12.46% | 1.696% |  | 2.705% | 1.779% | 5.867% | 0.9695% | 2.662% | 1.814%
| PDM3 |  |  | -10.39% |  | 6.947% | 2.743% | 4.759% | 2.138% | -5.493% | -12.81% | 2.801% | -4.859% | 3.837% |  | -5.661% | 1.764% | -2.453% | 0.9% | 0.9576% |  | -3.822% | 17.83% | 17.14% | 4.931% | 5.968% | -1.807%
| Z-Score (Absolute) |  | 1.93 | 1.5 | 1.93 | 0.8 | 0.25 | 0.51 | 0.16 | 0.85 | 1.82 | 0.25 | 0.77 | 0.39 | 1.93 | 0.87 | 0.11 | 0.45 | 0 | 0.01 | 1.93 | 0.63 | 2.25 | 2.16 | 0.54 | 0.67 | 0.36
| Analytical results for Ca in OREAS 556b (Certified Value 0.141 wt.%) | [28 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.15 | 0.15 | 0.16 | 0.136 | 0.14 | 0.12 | 0.14 | 0.13 | 0.11 | 0.14 | 0.15 | 0.1406 | 0.107 | 0.12808 | 0.1424 | 0.147 | 0.145 | 0.1527 | 0.14 | 0.13 | 0.14 | 0.133 | 0.0902 | 0.1447 | 0.15 | 0.14 |  | 1
| 1 | 2 | 0.17 | 0.15 | 0.15 | 0.1371 | 0.15 | 0.13 | 0.14 | 0.13 | 0.1 | 0.14 | 0.15 | 0.1411 | 0.107 | 0.13058 | 0.14279 | 0.144 | 0.14348 | 0.1536 | 0.13 | 0.14 | 0.15 | 0.144 | 0.087 | 0.1434 | 0.14 | 0.14 |  | 0
| 1 | 3 | 0.18 | 0.16 | 0.14 | 0.1356 | 0.14 | 0.12 | 0.14 | 0.14 | 0.11 | 0.14 | 0.15 | 0.1408 | 0.114 | 0.13029 | 0.14611 | 0.149 | 0.14363 | 0.1529 | 0.13 | 0.13 | 0.15 | 0.12 | 0.0818 | 0.1427 | 0.15 | 0.14 |  | 16
| 1 | 4 | 0.15 | 0.14 | 0.15 | 0.1371 | 0.14 | 0.12 | 0.14 | 0.14 | 0.11 | 0.14 | 0.15 | 0.1406 | 0.107 | 0.13992 | 0.14054 | 0.143 | 0.14328 | 0.1528 | 0.13 | 0.14 | 0.16 | 0.116 | 0.0776 | 0.1413 | 0.14 | 0.14 |  | 0.14107
| 1 | 5 | 0.16 | 0.15 | 0.14 | 0.1348 | 0.15 | 0.13 | 0.14 | 0.14 | 0.11 | 0.14 | 0.14 | 0.1411 | 0.114 | 0.13087 | 0.14982 | 0.148 | 0.14119 | 0.1526 | 0.13 | 0.13 | 0.14 | 0.121 | 0.0741 | 0.1385 | 0.14 | 0.14 |  | 93
| 1 | 6 | 0.16 | 0.15 | 0.14 | 0.1357 | 0.15 | 0.12 | 0.14 | 0.13 | 0.11 | 0.15 | 0.14 | 0.141 | 0.107 | 0.14063 | 0.14043 | 0.144 | 0.14025 | 0.1535 | 0.12 | 0.14 | 0.16 | 0.122 | 0.0786 | 0.1406 | 0.14 | 0.14
| Mean |  | 0.16167 | 0.15 | 0.14667 | 0.13605 | 0.145 | 0.12333 | 0.14 | 0.135 | 0.10833 | 0.14167 | 0.14667 | 0.14087 | 0.10933 | 0.1334 | 0.14368 | 0.14583 | 0.14281 | 0.15302 | 0.13 | 0.135 | 0.15 | 0.126 | 0.08155 | 0.14187 | 0.14333 | 0.14
| Median |  | 0.16 | 0.15 | 0.145 | 0.13585 | 0.145 | 0.12 | 0.14 | 0.135 | 0.11 | 0.14 | 0.15 | 0.1409 | 0.107 | 0.13072 | 0.14259 | 0.1455 | 0.14338 | 0.15285 | 0.13 | 0.135 | 0.15 | 0.1215 | 0.0802 | 0.142 | 0.14 | 0.14
| Std Dev. |  | 0.01169 | 0.0063246 | 0.008165 | 0.00090499 | 0.0054772 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.0040825 | 0.005164 | 0.00023381 | 0.0036148 | 0.0054224 | 0.0036457 | 0.0024833 | 0.0017512 | 0.00042622 | 0.0063246 | 0.0054772 | 0.0089443 | 0.010488 | 0.0060728 | 0.0022061 | 0.005164 | 0
| Rel.Std.Dev. |  | 7.231% | 4.216% | 5.567% | 0.6652% | 3.777% | 4.187% | 0% | 4.057% | 3.768% | 2.882% | 3.521% | 0.166% | 3.306% | 4.065% | 2.537% | 1.703% | 1.226% | 0.2785% | 4.865% | 4.057% | 5.963% | 8.324% | 7.447% | 1.555% | 3.603% | 0%
| PDM3 |  | 14.6% | 6.331% | 3.968% | -3.558% | 2.787% | -12.57% | -0.7577% | -4.302% | -23.21% | 0.4238% | 3.968% | -0.1433% | -22.5% | -5.439% | 1.852% | 3.377% | 1.231% | 8.469% | -7.846% | -4.302% | 6.331% | -10.68% | -42.19% | 0.5656% | 1.605% | -0.7577%
| Z-Score (Absolute) |  | 2.36 | 1.01 | 0.62 | 0.6 | 0.43 | 2.07 | 0.15 | 0.72 | 3.81 | 0.05 | 0.62 | 0.05 | 3.66 | 0.91 | 0.28 | 0.53 | 0.18 | 1.36 | 1.3 | 0.72 | 1.01 | 1.76 | 6.9 | 0.07 | 0.24 | 0.15
| Analytical results for Cd in OREAS 556b (Certified Value 0.51 ppm) | [28 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | < 1 | 0.59 | < 0.5 | 0.49 | 0.4 | 0.58 | 1.1016 | 0.45 | 0.44 | 0.51 | 0.53 | 0.54 | < 5 | < 1 | 0.50971 | 0.47 | 0.60849 | 0.42 | 0.57 | < 1 | 0.7 | 0.34 | 0.59 | < 0.5 | 0.49 |  |  | 1
| 1 | 2 | < 1 | 0.61 | < 0.5 | 0.51 | < 0.3 | 0.56 | 1.165 | 0.44 | 0.42 | 0.55 | 0.51 | 0.53 | < 5 | < 1 | 0.5018 | 0.49 | 0.62204 | 0.43 | 0.56 | 1 | 0.6 | 0.37 | 0.5 | < 0.5 | 0.51 |  |  | 5
| 1 | 3 | < 1 | 0.59 | < 0.5 | 0.53 | < 0.3 | 0.59 | 1.1358 | 0.48 | 0.45 | 0.5 | 0.53 | 0.51 | < 5 | < 1 | 0.4741 | 0.47 | 0.60098 | 0.41 | 0.56 | 1 | 0.8 | 0.35 | 0.54 | 0.5 | 0.52 |  |  | 16
| 1 | 4 | < 1 | 0.59 | < 0.5 | 0.57 | < 0.3 | 0.51 | 1.0771 | 0.46 | 0.44 | 0.59 | 0.5 | 0.54 | < 5 | < 1 | 0.49625 | 0.49 | 0.61014 | 0.37 | 0.57 | < 1 | 0.6 | 0.37 | 0.57 | < 0.5 | 0.51 |  |  | 0.51006
| 1 | 5 | < 1 | 0.57 | < 0.5 | 0.55 | 0.4 | 0.53 | 1.1316 | 0.45 | 0.44 | 0.53 | 0.54 | 0.57 | < 5 | < 1 | 0.50816 | 0.51 | 0.60617 | 0.41 | 0.56 | 1 | 0.8 | 0.35 | 0.55 | < 0.5 | 0.51 |  |  | 94
| 1 | 6 | < 1 | 0.57 | < 0.5 | 0.48 | < 0.3 | 0.57 | 1.1343 | 0.48 | 0.42 | 0.54 | 0.57 | 0.49 | < 5 | < 1 | 0.47963 | 0.5 | 0.5988 | 0.4 | 0.57 | 1 | 0.8 | 0.36 | 0.51 | < 0.5 | 0.51
| Mean |  |  | 0.58667 |  | 0.52167 | 0.4 | 0.55667 | 1.1242 | 0.46 | 0.435 | 0.53667 | 0.53 | 0.53 |  |  | 0.49494 | 0.48833 | 0.60777 | 0.40667 | 0.565 | 1 | 0.71667 | 0.35667 | 0.54333 | 0.5 | 0.50833
| Median |  |  | 0.59 |  | 0.52 | 0.4 | 0.565 | 1.133 | 0.455 | 0.44 | 0.535 | 0.53 | 0.535 |  |  | 0.49902 | 0.49 | 0.60733 | 0.41 | 0.565 | 1 | 0.75 | 0.355 | 0.545 | 0.5 | 0.51
| Std Dev. |  |  | 0.015055 |  | 0.034881 | 0 | 0.030768 | 0.030623 | 0.016733 | 0.012247 | 0.032042 | 0.024495 | 0.027568 |  |  | 0.014901 | 0.016021 | 0.0082308 | 0.020656 | 0.0054772 | 0 | 0.098319 | 0.012111 | 0.034448 |  | 0.0098319
| Rel.Std.Dev. |  |  | 2.566% |  | 6.686% | 0% | 5.527% | 2.724% | 3.638% | 2.816% | 5.97% | 4.622% | 5.202% |  |  | 3.011% | 3.281% | 1.354% | 5.079% | 0.9694% | 0% | 13.72% | 3.395% | 6.34% |  | 1.934%
| PDM3 |  |  | 15.02% |  | 2.275% | -21.58% | 9.137% | 120.4% | -9.815% | -14.72% | 5.216% | 3.909% | 3.909% |  |  | -2.965% | -4.26% | 19.16% | -20.27% | 10.77% | 96.05% | 40.51% | -30.07% | 6.523% | -1.973% | -0.3388%
| Z-Score (Absolute) |  | 0.19 | 0.92 | 3.4 | 0.09 | 3.62 | 0.54 | 7.83 | 0.71 | 1.03 | 0.28 | 0.19 | 0.19 | 25.5 | 0.19 | 0.26 | 0.34 | 1.19 | 1.39 | 0.64 | 4.09 | NA | 2.03 | 0.36 | 2.87 | 0.09
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 556b (Certified Value 143 ppm) | [28 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [6 empty] | 0
0
| 1 | 1 | 147.08 | 150.13 | 153 | 143 | 147.2867 | 141 | 136.5 | 152 | 145.5 | 153.5 | 133 | 143 | 142.7031 | 134.8 | 147 | 70 | 98.45 | 156 | 147 | 148 | 137.37 | [6 empty] | 1
| 1 | 2 | 142.68 | 152.65 | 157 | 139 | 145.6212 | 138 | 127.5 | 155 | 137 | 142 | 131 | 140 | 142.711 | 134 | 144 | 70 | 97.03 | 137 | 144 | 149 | 136.91 | [6 empty] | 6
| 1 | 3 | 153.44 | 148.62 | 154 | 143 | 147.7134 | 150 | 138 | 172 | 139 | 144.5 | 132 | 140 | 141.7221 | 133.1 | 145 | 77 | 98.85 | 147 | 143 | 151 | 140.94 | [6 empty] | 16
| 1 | 4 | 136.64 | 146.8 | 155 | 141 | 147.3384 | 141.5 | 135.5 | 162 | 140 | 151.5 | 134 | 139 | 139.7559 | 132 | 148 | 71 | 102 | 143 | 142 | 147 | 133.29 | [6 empty] | 143.3697
| 1 | 5 | 140.28 | 147.52 | 162 | 136 | 146.2124 | 143 | 136 | 150 | 147 | 145 | 135 | 140 | 138.2308 | 130.7 | 147 | 69 | 96.55 | 141 | 139 | 145 | 136.97 | [6 empty] | 95
| 1 | 6 | 146.24 | 149.17 | 159 | 136 | 146.4704 | 147.5 | 133 | 148 | 149.5 | 151.5 | 131 | 139 | 139.0605 | 134.5 | 147 | 75 | 95.27 | 131 | 144 | 146 | 138.29
| Mean |  | 144.3933 | 149.1483 | 156.6667 | 139.6667 | 146.7738 | 143.5 | 134.4167 | 156.5 | 143 | 148 | 132.6667 | 140.1667 | 140.6972 | 133.1833 | 146.3333 | 72 | 98.025 | 142.5 | 143.1667 | 147.6667 | 137.295
| Median |  | 144.46 | 148.895 | 156 | 140 | 146.8785 | 142.25 | 135.75 | 153.5 | 142.75 | 148.25 | 132.5 | 140 | 140.739 | 133.55 | 147 | 70.5 | 97.74 | 142 | 143.5 | 147.5 | 137.17
| Std Dev. |  | 5.8726 | 2.0819 | 3.3862 | 3.2042 | 0.80005 | 4.4497 | 3.7605 | 9.0277 | 5.01 | 4.7329 | 1.633 | 1.472 | 1.9379 | 1.5867 | 1.5055 | 3.2249 | 2.3417 | 8.5732 | 2.6394 | 2.1602 | 2.4729
| Rel.Std.Dev. |  | 4.067% | 1.396% | 2.161% | 2.294% | 0.5451% | 3.101% | 2.798% | 5.769% | 3.503% | 3.198% | 1.231% | 1.05% | 1.377% | 1.191% | 1.029% | 4.479% | 2.389% | 6.016% | 1.844% | 1.463% | 1.801%
| PDM3 |  | 0.714% | 4.031% | 9.275% | -2.583% | 2.374% | 0.09088% | -6.245% | 9.158% | -0.2579% | 3.23% | -7.465% | -2.234% | -1.864% | -7.105% | 2.067% | -49.78% | -31.63% | -0.6066% | -0.1416% | 2.997% | -4.237%
| Z-Score (Absolute) |  | 0.2 | 0.89 | 1.97 | 0.48 | 0.55 | 0.07 | 1.24 | 1.95 | 0 | 0.72 | 1.49 | 0.41 | 0.33 | 1.42 | 0.48 | 10.26 | 6.5 | 0.07 | 0.02 | 0.67 | 0.82
| Analytical results for Co in OREAS 556b (Certified Value 1.99 wt.%) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.37 | 1.96 | 2.1088 | 1.9107 | 2.01 | 1.89 | 2.1359 | 1.94 | 1.945 | 1.93 | 2.05 | 2.03 | 2.0765 | 2.04 | 2.0105 | 1.9976 | 1.83 | 1.9663 | 1.9707 | 2.04 | 2 | 2.01 | 1.96 | 1.9628 | 2.0381 | 2.024 |  | 1
| 1 | 2 | 2.38 | 1.93 | 2.1035 | 1.9407 | 2.02 | 1.88 | 2.1109 | 1.95 | 1.98 | 1.96 | 2.06 | 1.965 | 2.0639 | 2.06 | 2.036 | 1.9413 | 1.8 | 1.9797 | 1.9761 | 2.02 | 1.96 | 2.01 | 2.03 | 1.9718 | 2.0136 | 2.058 |  | 36
| 1 | 3 | 2.4 | 1.94 | 2.1293 | 2.0632 | 2.02 | 1.89 | 2.03 | 1.975 | 1.955 | 1.945 | 2.02 | 1.975 | 2.0685 | 2.08 | 2.0545 | 1.9224 | 1.83 | 1.9939 | 1.9677 | 2.01 | 1.98 | 1.96 | 1.93 | 1.9884 | 1.99 | 2.081 |  | 16
| 1 | 4 | 2.36 | 1.96 | 2.1353 | 1.9109 | 2 | 1.93 | 1.9873 | 1.98 | 1.895 | 2 | 2.07 | 2 | 2.0667 | 2.04 | 2.0738 | 1.9299 | 1.81 | 1.96 | 1.9815 | 2.03 | 1.97 | 1.96 | 1.79 | 1.9142 | 2.0328 | 2.056 |  | 1.9932
| 1 | 5 | 2.33 | 1.94 | 2.1252 | 1.9493 | 2 | 1.93 | 2.0261 | 1.935 | 1.89 | 2.04 | 2.04 | 1.965 | 2.0724 | 1.99 | 2.002 | 1.906 | 1.82 | 1.961 | 1.9527 | 2.05 | 1.98 | 1.98 | 1.82 | 1.8911 | 1.9742 | 1.996 |  | 96
| 1 | 6 | 2.36 | 1.9 | 2.1364 | 1.8992 | 2.01 | 1.89 | 2.0564 | 1.94 | 1.95 | 2 | 2.06 | 1.96 | 2.0763 | 2.05 | 2.0618 | 1.9669 | 1.8 | 1.9781 | 1.9743 | 2.03 | 2.01 | 1.95 | 1.79 | 1.9649 | 1.9796 | 2.018
| Mean |  | 2.3667 | 1.9383 | 2.1231 | 1.9457 | 2.01 | 1.9017 | 2.0578 | 1.9533 | 1.9358 | 1.9792 | 2.05 | 1.9825 | 2.0707 | 2.0433 | 2.0398 | 1.944 | 1.815 | 1.9732 | 1.9705 | 2.03 | 1.9833 | 1.9783 | 1.8867 | 1.9489 | 2.0047 | 2.0388
| Median |  | 2.365 | 1.94 | 2.1273 | 1.9258 | 2.01 | 1.89 | 2.0432 | 1.945 | 1.9475 | 1.98 | 2.055 | 1.97 | 2.0705 | 2.045 | 2.0452 | 1.9356 | 1.815 | 1.9722 | 1.9725 | 2.03 | 1.98 | 1.97 | 1.875 | 1.9639 | 2.0018 | 2.04
| Std Dev. |  | 0.023381 | 0.022286 | 0.013836 | 0.060733 | 0.0089443 | 0.022286 | 0.056005 | 0.019408 | 0.035695 | 0.041282 | 0.017889 | 0.02734 | 0.0051955 | 0.030111 | 0.02884 | 0.033209 | 0.013784 | 0.013151 | 0.0099126 | 0.014142 | 0.018619 | 0.026394 | 0.10093 | 0.037627 | 0.027425 | 0.031397
| Rel.Std.Dev. |  | 0.9879% | 1.15% | 0.6517% | 3.121% | 0.445% | 1.172% | 2.722% | 0.9936% | 1.844% | 2.086% | 0.8726% | 1.379% | 0.2509% | 1.474% | 1.414% | 1.708% | 0.7595% | 0.6665% | 0.5031% | 0.6967% | 0.9388% | 1.334% | 5.35% | 1.931% | 1.368% | 1.54%
| PDM3 |  | 18.73% | -2.755% | 6.514% | -2.387% | 0.8405% | -4.595% | 3.237% | -2.002% | -2.88% | -0.7064% | 2.847% | -0.5391% | 3.887% | 2.513% | 2.334% | -2.47% | -8.943% | -1.008% | -1.141% | 1.844% | -0.4973% | -0.7482% | -5.347% | -2.227% | 0.5754% | 2.287%
| Z-Score (Absolute) |  | 5.95 | 0.66 | 2.19 | 0.54 | 0.45 | 1.22 | 1.19 | 0.42 | 0.69 | 0.03 | 1.07 | 0.03 | 1.39 | 0.96 | 0.91 | 0.57 | 2.56 | 0.12 | 0.16 | 0.76 | 0.04 | 0.04 | 1.45 | 0.49 | 0.37 | 0.89
| Analytical results for Cr in OREAS 556b (Certified Value 89 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 89.5 | 106 | 85.2 | 90 | 84 | 83.2 | 102.5084 | 91 | 89 | 97 | 91 | 98 | 77.89 | 105 | 89.682 | 85 | 90.9697 | 80.6 | 95 | 79 | 85 | 82 | 86 | 92 | 75 | 90 |  | 1
| 1 | 2 | 83.41 | 101 | 83.9 | 91 | 83 | 84.1 | 101.1067 | 90 | 91 | 94 | 90 | 97 | 78.14 | 99 | 88.3771 | 81 | 90.0631 | 80.8 | 93 | 74 | 86 | 94 | 86 | 91 | 74 | 90 |  | 37
| 1 | 3 | 85.81 | 108 | 85.8 | 89 | 72 | 84.3 | 100.3307 | 92 | 92 | 98 | 92 | 96 | 78.06 | 102 | 88.954 | 84 | 89.7927 | 80.5 | 97 | 78 | 87 | 82 | 85 | 89 | 75 | 92 |  | 16
| 1 | 4 | 92.32 | 98 | 85 | 90 | 85 | 85.4 | 99.7089 | 92 | 91 | 98 | 95 | 98 | 77.92 | 107 | 90.2617 | 83 | 89.5823 | 80 | 100 | 74 | 86 | 89 | 81 | 90 | 72 | 90 |  | 88.9552
| 1 | 5 | 91.74 | 98 | 85.4 | 89 | 86 | 86.1 | 100.849 | 94 | 91 | 98 | 88 | 90 | 77.94 | 106 | 90.9777 | 85 | 89.1485 | 78.9 | 97 | 79 | 86 | 88 | 79 | 87 | 70 | 91 |  | 97
| 1 | 6 | 82.29 | 107 | 84.4 | 90 | 86 | 84.9 | 100.2762 | 95 | 90 | 100 | 96 | 93 | 77.96 | 101 | 90.1162 | 88 | 89.4222 | 80.4 | 97 | 80 | 85 | 86 | 82 | 90 | 73 | 91
| Mean |  | 87.5117 | 103 | 84.95 | 89.8333 | 82.6667 | 84.6667 | 100.7966 | 92.3333 | 90.6667 | 97.5 | 92 | 95.3333 | 77.985 | 103.3333 | 89.7281 | 84.3333 | 89.8297 | 80.2 | 96.5 | 77.3333 | 85.8333 | 86.8333 | 83.1667 | 89.8333 | 73.1667 | 90.6667
| Median |  | 87.655 | 103.5 | 85.1 | 90 | 84.5 | 84.6 | 100.5898 | 92 | 91 | 98 | 91.5 | 96.5 | 77.95 | 103.5 | 89.8991 | 84.5 | 89.6875 | 80.45 | 97 | 78.5 | 86 | 87 | 83.5 | 90 | 73.5 | 90.5
| Std Dev. |  | 4.2881 | 4.5607 | 0.6921 | 0.75277 | 5.3541 | 1.025 | 0.96955 | 1.8619 | 1.0328 | 1.9748 | 3.0332 | 3.2042 | 0.095446 | 3.1411 | 0.94058 | 2.3381 | 0.63981 | 0.68993 | 2.3452 | 2.6583 | 0.75277 | 4.5789 | 2.9269 | 1.7224 | 1.9408 | 0.8165
| Rel.Std.Dev. |  | 4.9% | 4.428% | 0.8147% | 0.838% | 6.477% | 1.211% | 0.9619% | 2.016% | 1.139% | 2.025% | 3.297% | 3.361% | 0.1224% | 3.04% | 1.048% | 2.772% | 0.7122% | 0.8603% | 2.43% | 3.437% | 0.877% | 5.273% | 3.519% | 1.917% | 2.653% | 0.9005%
| PDM3 |  | -1.623% | 15.79% | -4.502% | 0.9872% | -7.069% | -4.821% | 13.31% | 3.798% | 1.924% | 9.606% | 3.423% | 7.17% | -12.33% | 16.16% | 0.8689% | -5.196% | 0.9832% | -9.842% | 8.482% | -13.06% | -3.509% | -2.385% | -6.507% | 0.9872% | -17.75% | 1.924%
| Z-Score (Absolute) |  | 0.29 | 1.69 | 0.62 | 0.01 | 0.91 | 0.65 | 1.41 | 0.33 | 0.11 | 0.99 | 0.28 | 0.71 | 1.51 | 1.73 | 0.01 | 0.7 | 0.01 | 1.22 | 0.86 | 1.59 | 0.5 | 0.38 | 0.84 | 0.01 | 2.12 | 0.11
| Analytical results for Cs in OREAS 556b (Certified Value 1.75 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.5 | 1.82 | 1.82 | 1.96 | 1.8144 | 1.63 | 1.56 | 1.72 | 1.73 | 1.91 | 1.4124 | 1.7905 | 1.8 | 1.7969 | 1.8 | 1.73 | 1.5 | 1.75 | 1 | 1.85 | 1.64 | [6 empty] | 1
| 1 | 2 | 1.5 | 1.83 | 1.88 | 1.85 | 1.7614 | 1.58 | 1.49 | 1.74 | 1.68 | 1.72 | 1.3738 | 1.7996 | 1.8 | 1.8254 | 1.9 | 1.64 | 1.6 | 1.71 | 1.2 | 1.84 | 1.64 | [6 empty] | 9
| 1 | 3 | 1.5 | 1.94 | 1.82 | 1.9 | 1.8061 | 1.72 | 1.59 | 1.78 | 1.72 | 1.82 | 1.7741 | 1.7198 | 1.8 | 1.8119 | 1.8 | 1.72 | 1.9 | 1.69 | 1 | 1.8 | 1.71 | [6 empty] | 16
| 1 | 4 | 1.4 | 1.76 | 1.78 | 1.89 | 1.8062 | 1.68 | 1.55 | 1.83 | 1.76 | 1.85 | 1.7502 | 1.7106 | 1.8 | 1.797 | 1.7 | 1.76 | 1.5 | 1.7 | 1.1 | 1.87 | 1.63 | [6 empty] | 1.7544
| 1 | 5 | 1.6 | 1.83 | 1.91 | 1.82 | 1.7732 | 1.73 | 1.56 | 1.78 | 1.78 | 1.74 | 1.7883 | 1.625 | 1.8 | 1.7916 | 1.8 | 1.72 | 1.8 | 1.65 | 1.1 | 1.87 | 1.65 | [6 empty] | 98
| 1 | 6 | 1.3 | 1.8 | 1.92 | 1.97 | 1.7956 | 1.69 | 1.53 | 1.74 | 1.82 | 1.79 | 1.8118 | 1.7431 | 1.8 | 1.7547 | 1.8 | 1.72 | 1.5 | 1.63 | 1.1 | 1.79 | 1.67
| Mean |  | 1.4667 | 1.83 | 1.855 | 1.8983 | 1.7928 | 1.6717 | 1.5467 | 1.765 | 1.7483 | 1.805 | 1.6518 | 1.7314 | 1.8 | 1.7962 | 1.8 | 1.715 | 1.6333 | 1.6883 | 1.0833 | 1.8367 | 1.6567
| Median |  | 1.5 | 1.825 | 1.85 | 1.895 | 1.8009 | 1.685 | 1.555 | 1.76 | 1.745 | 1.805 | 1.7622 | 1.7315 | 1.8 | 1.7969 | 1.8 | 1.72 | 1.55 | 1.695 | 1.1 | 1.845 | 1.645
| Std Dev. |  | 0.10328 | 0.06 | 0.05648 | 0.059133 | 0.020978 | 0.057067 | 0.033862 | 0.039875 | 0.04916 | 0.07064 | 0.20174 | 0.063501 | 0 | 0.023863 | 0.063246 | 0.039875 | 0.17512 | 0.043089 | 0.075277 | 0.034448 | 0.029439
| Rel.Std.Dev. |  | 7.042% | 3.279% | 3.045% | 3.115% | 1.17% | 3.414% | 2.189% | 2.259% | 2.812% | 3.914% | 12.21% | 3.668% | 0% | 1.328% | 3.514% | 2.325% | 10.72% | 2.552% | 6.949% | 1.876% | 1.777%
| PDM3 |  | -16.4% | 4.308% | 5.733% | 8.203% | 2.189% | -4.717% | -11.84% | 0.6031% | -0.3469% | 2.883% | -5.851% | -1.309% | 2.598% | 2.383% | 2.598% | -2.247% | -6.902% | -3.767% | -38.25% | 4.688% | -5.572%
| Z-Score (Absolute) |  | 2.48 | 0.72 | 0.94 | 1.32 | 0.39 | 0.67 | 1.77 | 0.15 | 0 | 0.5 | 0.85 | 0.15 | 0.45 | 0.42 | 0.45 | 0.29 | 1.01 | 0.53 | 5.85 | 0.78 | 0.81
| Analytical results for Cu in OREAS 556b (Certified Value 8.64 wt.%) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*AAS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 8.59 | 8.448 | 8.61 | 8.0709 | 8.64 | 7.85 | 9.1299 | 8.45 | 8.55 | 8.5 | 8.74 | 8.74 | 8.9408 | 8.8 | 8.6464 | 7.9937 | 9.15 | 8.7395 | 2.7372 | 8.77 | 8.809 | 8.75 | 9.425 | 8.6789 | 8.4916 | 8.781 |  | 1
| 1 | 2 | 8.54 | 8.383 | 8.525 | 8.2332 | 8.58 | 7.81 | 9.0984 | 8.58 | 8.67 | 8.6 | 8.73 | 8.57 | 8.9513 | 8.64 | 8.4857 | 7.9923 | 9.14 | 8.8276 | 2.0538 | 8.71 | 8.835 | 8.72 | 8.898 | 8.483 | 8.4119 | 8.692 |  | 39
| 1 | 3 | 8.46 | 8.473 | 8.595 | 8.5207 | 8.54 | 7.79 | 8.8321 | 8.47 | 8.48 | 8.68 | 8.73 | 8.57 | 8.9446 | 8.69 | 8.4696 | 7.9196 | 9.18 | 8.7882 | 3.1256 | 8.8 | 8.726 | 8.61 | 8.791 | 8.3998 | 8.3312 | 8.625 |  | 16
| 1 | 4 | 8.63 | 8.269 | 8.58 | 8.1849 | 8.51 | 7.93 | 8.6246 | 8.55 | 8.42 | 8.54 | 8.69 | 8.59 | 8.9506 | 8.6 | 8.7488 | 7.9265 | 9.02 | 8.8152 | 2.8153 | 8.78 | 8.723 | 8.66 | 8.882 | 8.3888 | 8.3588 | 8.756 |  | 8.6408
| 1 | 5 | 8.68 | 8.426 | 8.55 | 8.3567 | 8.53 | 8 | 8.8509 | 8.35 | 8.3 | 8.72 | 8.77 | 8.55 | 8.9504 | 8.62 | 8.464 | 7.9048 | 9.15 | 8.772 | 2.7168 | 8.73 | 8.622 | 8.63 | 9.007 | 8.1703 | 8.3846 | 8.568 |  | 99
| 1 | 6 | 8.58 | 8.146 | 8.605 | 8.1434 | 8.62 | 7.84 | 8.8433 | 8.49 | 8.6 | 8.6 | 8.78 | 8.6 | 8.9512 | 8.68 | 8.6625 | 7.9719 | 9.05 | 8.778 | 2.5621 | 8.81 | 8.806 | 8.59 | 8.391 | 8.3287 | 8.3798 | 8.746
| Mean |  | 8.58 | 8.3575 | 8.5775 | 8.2516 | 8.57 | 7.87 | 8.8965 | 8.4817 | 8.5033 | 8.6067 | 8.74 | 8.6033 | 8.9482 | 8.6717 | 8.5795 | 7.9515 | 9.115 | 8.7867 | 2.6685 | 8.7667 | 8.7535 | 8.66 | 8.899 | 8.4083 | 8.393 | 8.6947
| Median |  | 8.585 | 8.4045 | 8.5875 | 8.209 | 8.56 | 7.845 | 8.8471 | 8.48 | 8.515 | 8.6 | 8.735 | 8.58 | 8.9505 | 8.66 | 8.566 | 7.9492 | 9.145 | 8.7831 | 2.727 | 8.775 | 8.766 | 8.645 | 8.89 | 8.3943 | 8.3822 | 8.719
| Std Dev. |  | 0.075631 | 0.12605 | 0.033578 | 0.16285 | 0.052154 | 0.07975 | 0.18878 | 0.081097 | 0.13277 | 0.082624 | 0.032249 | 0.069186 | 0.0044026 | 0.071671 | 0.12186 | 0.039205 | 0.064109 | 0.031648 | 0.35404 | 0.039328 | 0.079319 | 0.063246 | 0.33417 | 0.16866 | 0.055347 | 0.083373
| Rel.Std.Dev. |  | 0.8815% | 1.508% | 0.3915% | 1.974% | 0.6086% | 1.013% | 2.122% | 0.9561% | 1.561% | 0.96% | 0.369% | 0.8042% | 0.0492% | 0.8265% | 1.42% | 0.4931% | 0.7033% | 0.3602% | 13.27% | 0.4486% | 0.9061% | 0.7303% | 3.755% | 2.006% | 0.6594% | 0.9589%
| PDM3 |  | -0.704% | -3.279% | -0.7329% | -4.504% | -0.8197% | -8.921% | 2.959% | -1.842% | -1.591% | -0.3954% | 1.148% | -0.4339% | 3.557% | 0.3569% | -0.7099% | -7.978% | 5.488% | 1.689% | -69.12% | 1.456% | 1.304% | 0.2219% | 2.988% | -2.692% | -2.868% | 0.6231%
| Z-Score (Absolute) |  | 0.05 | 0.94 | 0.06 | 1.36 | 0.09 | 2.89 | 1.22 | 0.44 | 0.35 | 0.06 | 0.59 | 0.05 | 1.43 | 0.32 | 0.05 | 2.56 | 2.1 | 0.78 | 23.72 | 0.7 | 0.65 | 0.27 | 1.23 | 0.73 | 0.8 | 0.41
| Analytical results for Dy in OREAS 556b (Certified Value 4.54 ppm) | [28 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 07; 05: 14; 06: 16; 07: 18; 08: 19; 09: 23; 10: 25; 11: 26 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS |  | 2
2
| 1 | 1 | 01: 3; 02: 4.47; 03: 5.7; 04: 6.1874; 05: 4.7; 06: 4.6064; 07: 4.6726; 08: 3.99; 09: 4.02; 10: 4.58; 11: 4.83 |  | 1
| 1 | 2 | 01: 3.1; 02: 4.54; 03: 5.5; 04: 6.1269; 05: 4.6; 06: 4.4989; 07: 4.6565; 08: 4.11; 09: 3.9; 10: 4.27; 11: 4.75 |  | 10
| 1 | 3 | 01: 3.2; 02: 4.45; 03: 5.4; 04: 6.1361; 05: 4.7; 06: 4.7007; 07: 4.6036; 08: 3.91; 09: 4.03; 10: 4.35; 11: 4.89 |  | 16
| 1 | 4 | 01: 3; 02: 4.43; 03: 5.5; 04: 6.2138; 05: 4.6; 06: 4.4629; 07: 4.5961; 08: 3.65; 09: 4.16; 10: 4.22; 11: 4.88 |  | 4.5357
| 1 | 5 | 01: 3; 02: 4.32; 03: 5.8; 04: 5.9869; 05: 4.5; 06: 4.3697; 07: 4.4507; 08: 3.82; 09: 4.1; 10: 4.17; 11: 4.82 |  | 100
| 1 | 6 | 01: 2.9; 02: 4.42; 03: 5.7; 04: 6.1723; 05: 4.6; 06: 4.7414; 07: 4.3885; 08: 3.92; 09: 4.01; 10: 4.06; 11: 4.81
| Mean |  | 01: 3.0333; 02: 4.4383; 03: 5.6; 04: 6.1372; 05: 4.6167; 06: 4.5634; 07: 4.5613; 08: 3.9; 09: 4.0367; 10: 4.275; 11: 4.83
| Median |  | 01: 3; 02: 4.44; 03: 5.6; 04: 6.1542; 05: 4.6; 06: 4.5526; 07: 4.5998; 08: 3.915; 09: 4.025; 10: 4.245; 11: 4.825
| Std Dev. |  | 01: 0.10328; 02: 0.071949; 03: 0.15492; 04: 0.080394; 05: 0.075277; 06: 0.14434; 07: 0.11538; 08: 0.15595; 09: 0.088242; 10: 0.1783; 11: 0.05099
| Rel.Std.Dev. |  | 01: 3.405%; 02: 1.621%; 03: 2.766%; 04: 1.31%; 05: 1.631%; 06: 3.163%; 07: 2.529%; 08: 3.999%; 09: 2.186%; 10: 4.171%; 11: 1.056%
| PDM3 |  | 01: -33.12%; 02: -2.147%; 03: 23.46%; 04: 35.31%; 05: 1.785%; 06: 0.6097%; 07: 0.5649%; 08: -14.02%; 09: -11%; 10: -5.748%; 11: 6.488%
| Z-Score (Absolute) |  | 01: 3.6; 02: 0.29; 03: 2.45; 04: 3.71; 05: 0.13; 06: 0; 07: 0; 08: 1.56; 09: 1.24; 10: 0.67; 11: 0.63
| Analytical results for Er in OREAS 556b (Certified Value 2.7 ppm) | [28 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 07; 05: 14; 06: 16; 07: 18; 08: 19; 09: 23; 10: 25; 11: 26 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS |  | 2
2
| 1 | 1 | 01: 1.9; 02: 2.75; 03: 3.2; 04: 3.4457; 05: < 0.2; 06: 3.1501; 07: 2.8039; 08: 2.09; 09: 2.33; 10: 2.7; 11: 2.94 |  | 1
| 1 | 2 | 01: 1.9; 02: 2.67; 03: 3.4; 04: 3.5106; 05: < 0.2; 06: 3.1489; 07: 2.7944; 08: 2.16; 09: 2.32; 10: 2.52; 11: 2.9 |  | 11
| 1 | 3 | 01: 1.9; 02: 2.71; 03: 3; 04: 3.4114; 05: < 0.2; 06: 3.2325; 07: 2.7804; 08: 2.04; 09: 2.39; 10: 2.7; 11: 2.86 |  | 16
| 1 | 4 | 01: 1.8; 02: 2.64; 03: 3.2; 04: 3.4879; 05: < 0.2; 06: 2.976; 07: 2.7932; 08: 1.93; 09: 2.37; 10: 2.56; 11: 2.94 |  | 2.6977
| 1 | 5 | 01: 1.9; 02: 2.61; 03: 3.5; 04: 3.4541; 05: < 0.2; 06: 3.9456; 07: 2.8853; 08: 1.99; 09: 2.24; 10: 2.58; 11: 2.91 |  | 101
| 1 | 6 | 01: 1.8; 02: 2.58; 03: 3.2; 04: 3.5377; 05: < 0.2; 06: 3.0933; 07: 2.7395; 08: 2.06; 09: 2.25; 10: 2.34; 11: 2.82
| Mean |  | 01: 1.8667; 02: 2.66; 03: 3.25; 04: 3.4746; 06: 3.2577; 07: 2.7994; 08: 2.045; 09: 2.3167; 10: 2.5667; 11: 2.895
| Median |  | 01: 1.9; 02: 2.655; 03: 3.2; 04: 3.471; 06: 3.1495; 07: 2.7938; 08: 2.05; 09: 2.325; 10: 2.57; 11: 2.905
| Std Dev. |  | 01: 0.05164; 02: 0.063246; 03: 0.17607; 04: 0.046261; 06: 0.34744; 07: 0.047775; 08: 0.079687; 09: 0.06121; 10: 0.13367; 11: 0.047223
| Rel.Std.Dev. |  | 01: 2.766%; 02: 2.378%; 03: 5.417%; 04: 1.331%; 06: 10.67%; 07: 1.707%; 08: 3.897%; 09: 2.642%; 10: 5.208%; 11: 1.631%
| PDM3 |  | 01: -30.81%; 02: -1.397%; 03: 20.47%; 04: 28.8%; 06: 20.76%; 07: 3.771%; 08: -24.19%; 09: -14.12%; 10: -4.857%; 11: 7.314%
| Z-Score (Absolute) |  | 01: 0.91; 02: 0; 03: 0.67; 04: 0.93; 05: 2.93; 06: 0.68; 07: 0.16; 08: 0.7; 09: 0.39; 10: 0.11; 11: 0.27
| Analytical results for Eu in OREAS 556b (Certified Value 0.89 ppm) | [28 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 07; 05: 14; 06: 18; 07: 19; 08: 23; 09: 25; 10: 26 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*MS |  | 2
3
| 1 | 1 | 01: 0.5; 02: 0.98; 03: 0.96; 04: 0.94095; 05: 1.5; 06: 0.88825; 07: 0.9; 08: 0.8; 09: 0.87; 10: 0.97 |  | 1
| 1 | 2 | 01: 0.5; 02: 0.92; 03: 0.94; 04: 0.92679; 05: 1.2; 06: 0.87916; 07: 0.94; 08: 0.78; 09: 0.81; 10: 0.96 |  | 12
| 1 | 3 | 01: 0.6; 02: 0.88; 03: 0.91; 04: 0.92741; 05: 1.4; 06: 0.87928; 07: 0.89; 08: 0.79; 09: 0.83; 10: 0.94 |  | 16
| 1 | 4 | 01: 0.5; 02: 0.9; 03: 0.95; 04: 0.90773; 05: 1.4; 06: 0.88841; 07: 0.83; 08: 0.84; 09: 0.81; 10: 0.91 |  | 0.88976
| 1 | 5 | 01: 0.6; 02: 0.93; 03: 0.99; 04: 0.94369; 05: 1.6; 06: 0.87305; 07: 0.87; 08: 0.77; 09: 0.8; 10: 0.95 |  | 102
| 1 | 6 | 01: 0.5; 02: 0.91; 03: 0.99; 04: 0.95853; 05: 1.2; 06: 0.85505; 07: 0.9; 08: 0.74; 09: 0.72; 10: 0.96
| Mean |  | 01: 0.53333; 02: 0.92; 03: 0.95667; 04: 0.93418; 05: 1.3833; 06: 0.8772; 07: 0.88833; 08: 0.78667; 09: 0.80667; 10: 0.94833
| Median |  | 01: 0.5; 02: 0.915; 03: 0.955; 04: 0.93418; 05: 1.4; 06: 0.87922; 07: 0.895; 08: 0.785; 09: 0.81; 10: 0.955
| Std Dev. |  | 01: 0.05164; 02: 0.034059; 03: 0.030768; 04: 0.017489; 05: 0.16021; 06: 0.012359; 07: 0.03656; 08: 0.033267; 09: 0.049261; 10: 0.02137
| Rel.Std.Dev. |  | 01: 9.682%; 02: 3.702%; 03: 3.216%; 04: 1.872%; 05: 11.58%; 06: 1.409%; 07: 4.116%; 08: 4.229%; 09: 6.107%; 10: 2.253%
| PDM3 |  | 01: -40.06%; 02: 3.399%; 03: 7.52%; 04: 4.993%; 05: 55.47%; 06: -1.411%; 07: -0.1599%; 08: -11.59%; 09: -9.338%; 10: 6.583%
| Z-Score (Absolute) |  | 01: NA; 02: 0.29; 03: 0.95; 04: 0.55; 05: NA; 06: 0.49; 07: 0.29; 08: 2.14; 09: 1.77; 10: 0.8
Results from laboratories 2 and 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 556b (Certified Value 6.71 wt.%) | [28 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 6.43 | 6.84 | 6.67 | 6.73 | 6.89 | 6.17 | 6.49 | 6.23 | 6.71 | 6.58 | 6.76 | 6.9431 | 6.61 | 6.6998 | 6.5426 | 7.05 | > 5 | 6.8757 | 6.64 | 6.73 | 6.62 | 6.87 | 5.93 | 7.14 | 6.95 | 6.76 |  | 1
| 1 | 2 | 6.39 | 6.56 | 6.66 | 6.8 | 7.13 | 6.2 | 6.48 | 6.39 | 6.49 | 6.47 | 6.76 | 7.001 | 6.498 | 6.6853 | 6.5218 | 7.08 | > 5 | 6.9286 | 6.44 | 6.7 | 6.74 | 6.69 | 6.65 | 7.03 | 6.92 | 6.8 |  | 0
| 1 | 3 | 6.32 | 6.98 | 6.66 | 6.7 | 7.05 | 6.11 | 6.63 | 6.39 | 6.66 | 6.7 | 6.86 | 6.9678 | 6.498 | 6.7621 | 6.5431 | 7.05 | > 5 | 6.9119 | 6.68 | 6.73 | 6.7 | 6.47 | 6.08 | 6.93 | 6.98 | 6.84 |  | 16
| 1 | 4 | 6.42 | 6.25 | 6.69 | 6.76 | 7.12 | 6.19 | 6.6 | 6.41 | 6.77 | 6.59 | 6.78 | 6.9235 | 6.505 | 7.2314 | 6.5489 | 6.89 | > 5 | 6.9047 | 6.65 | 6.69 | 6.75 | 6.63 | 5.94 | 6.98 | 6.93 | 6.81 |  | 6.7076
| 1 | 5 | 6.49 | 6.48 | 6.66 | 6.67 | 7.34 | 6.14 | 6.77 | 6.38 | 6.77 | 6.64 | 6.67 | 6.9317 | 6.491 | 6.7756 | 6.5819 | 7.05 | > 5 | 6.7573 | 6.56 | 6.74 | 6.68 | 6.71 | 5.87 | 6.81 | 6.87 | 6.9 |  | 103
| 1 | 6 | 6.41 | 6.72 | 6.67 | 6.8 | 7.29 | 6.05 | 6.7 | 6.32 | 6.78 | 6.96 | 6.63 | 6.9218 | 6.54 | 7.2745 | 6.5177 | 6.95 | > 5 | 6.9242 | 6.39 | 6.8 | 6.74 | 6.68 | 5.98 | 6.9 | 6.89 | 6.85
| Mean |  | 6.41 | 6.6383 | 6.6683 | 6.7433 | 7.1367 | 6.1433 | 6.6117 | 6.3533 | 6.6967 | 6.6567 | 6.7433 | 6.9482 | 6.5237 | 6.9048 | 6.5427 | 7.0117 |  | 6.8837 | 6.56 | 6.7317 | 6.705 | 6.675 | 6.075 | 6.965 | 6.9233 | 6.8267
| Median |  | 6.415 | 6.64 | 6.665 | 6.745 | 7.125 | 6.155 | 6.615 | 6.385 | 6.74 | 6.615 | 6.76 | 6.9374 | 6.5015 | 6.7688 | 6.5428 | 7.05 |  | 6.9083 | 6.6 | 6.73 | 6.72 | 6.685 | 5.96 | 6.955 | 6.925 | 6.825
| Std Dev. |  | 0.055498 | 0.26309 | 0.01169 | 0.053166 | 0.16342 | 0.056451 | 0.11444 | 0.067725 | 0.1113 | 0.16693 | 0.082138 | 0.03091 | 0.045724 | 0.27224 | 0.022965 | 0.074409 |  | 0.064705 | 0.12017 | 0.038687 | 0.049699 | 0.12927 | 0.29016 | 0.11362 | 0.039833 | 0.048028
| Rel.Std.Dev. |  | 0.8658% | 3.963% | 0.1753% | 0.7884% | 2.29% | 0.9189% | 1.731% | 1.066% | 1.662% | 2.508% | 1.218% | 0.4449% | 0.7009% | 3.943% | 0.351% | 1.061% |  | 0.94% | 1.832% | 0.5747% | 0.7412% | 1.937% | 4.776% | 1.631% | 0.5753% | 0.7035%
| PDM3 |  | -4.437% | -1.033% | -0.5861% | 0.532% | 6.396% | -8.413% | -1.431% | -5.282% | -0.1637% | -0.76% | 0.532% | 3.586% | -2.743% | 2.939% | -2.46% | 4.532% |  | 2.625% | -2.201% | 0.3581% | -0.03945% | -0.4867% | -9.432% | 3.837% | 3.216% | 1.774%
| Z-Score (Absolute) |  | 1.26 | 0.26 | 0.12 | 0.2 | 1.93 | 2.42 | 0.37 | 1.5 | 0 | 0.18 | 0.2 | 1.1 | 0.76 | 0.91 | 0.67 | 1.38 | NA | 0.82 | 0.6 | 0.15 | 0.04 | 0.09 | 2.72 | 1.17 | 0.99 | 0.57
| Analytical results for Ga in OREAS 556b (Certified Value 28.1 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 26.68 | 27.2 | 29.92 | 28.2 | 29.8 | 28.1256 | 28.1 | 25.6 | 27.4 | 29 | 28.4 | 23.75 | 13 | 27.5614 | 27.7 | 28.3535 | 28.25 | 30.4 | 27 | 29 | 21.3 | 28.21 | 25.31 |  |  |  |  | 1
| 1 | 2 | 29.28 | 28.2 | 29.3 | 29.3 | 28.5 | 27.3742 | 27.5 | 23.9 | 27.5 | 27.5 | 26.3 | 23.52 | 14 | 27.501 | 28.3 | 28.1155 | 28.52 | 28.69 | 26 | 29 | 23.2 | 28.8 | 25.19 |  |  |  |  | 13
| 1 | 3 | 28.15 | 28.7 | 28.95 | 28.4 | 29.7 | 27.0658 | 29.2 | 26.4 | 28.4 | 28.6 | 27.5 | 23.61 | 10 | 28.1445 | 29.6 | 27.8953 | 28.08 | 30.03 | 28 | 31 | 22.4 | 28.19 | 26.33 |  |  |  |  | 16
| 1 | 4 | 27.24 | 28.1 | 28.6 | 28.6 | 28.6 | 27.6544 | 29.2 | 25 | 29 | 28.4 | 27.5 | 23.63 | 12 | 28.0956 | 29.4 | 28.1737 | 27.88 | 30.98 | 26 | 29 | 22.2 | 28.34 | 26.5 |  |  |  |  | 28.1302
| 1 | 5 | 26.17 | 28.4 | 29.19 | 29.4 | 28.5 | 27.7765 | 29.5 | 25.7 | 27.5 | 29.6 | 26.6 | 23.67 | 13 | 28.0256 | 28.5 | 28.0067 | 27.63 | 30.29 | 27 | 31 | 21.6 | 28.74 | 26.77 |  |  |  |  | 104
| 1 | 6 | 25.85 | 28.1 | 28.32 | 28.8 | 29.2 | 27.0912 | 29.7 | 24.8 | 28.3 | 30 | 27.6 | 23.72 | 11 | 28.2747 | 28.6 | 27.5375 | 28.06 | 30.02 | 27 | 31 | 20.7 | 28.9 | 25.34
| Mean |  | 27.2283 | 28.1167 | 29.0467 | 28.7833 | 29.05 | 27.5146 | 28.8667 | 25.2333 | 28.0167 | 28.85 | 27.3167 | 23.65 | 12.1667 | 27.9338 | 28.6833 | 28.0137 | 28.07 | 30.0683 | 26.8333 | 30 | 21.9 | 28.53 | 25.9067
| Median |  | 26.96 | 28.15 | 29.07 | 28.7 | 28.9 | 27.5143 | 29.2 | 25.3 | 27.9 | 28.8 | 27.5 | 23.65 | 12.5 | 28.0606 | 28.55 | 28.0611 | 28.07 | 30.16 | 27 | 30 | 21.9 | 28.54 | 25.835
| Std Dev. |  | 1.2946 | 0.50365 | 0.56291 | 0.48339 | 0.60249 | 0.41517 | 0.86872 | 0.8641 | 0.64936 | 0.89387 | 0.75741 | 0.082704 | 1.472 | 0.32285 | 0.70828 | 0.28011 | 0.30489 | 0.76077 | 0.75277 | 1.0954 | 0.88544 | 0.31875 | 0.70247
| Rel.Std.Dev. |  | 4.754% | 1.791% | 1.938% | 1.679% | 2.074% | 1.509% | 3.009% | 3.424% | 2.318% | 3.098% | 2.773% | 0.3497% | 12.1% | 1.156% | 2.469% | 0.9999% | 1.086% | 2.53% | 2.805% | 3.651% | 4.043% | 1.117% | 2.712%
| PDM3 |  | -3.206% | -0.04819% | 3.258% | 2.322% | 3.27% | -2.188% | 2.618% | -10.3% | -0.4037% | 2.559% | -2.892% | -15.93% | -56.75% | -0.6983% | 1.966% | -0.4142% | -0.2141% | 6.89% | -4.61% | 6.647% | -22.15% | 1.421% | -7.905%
| Z-Score (Absolute) |  | 0.64 | 0.08 | 0.83 | 0.62 | 0.84 | 0.41 | 0.69 | 2.25 | 0 | 0.67 | 0.57 | 3.53 | 12.83 | 0.07 | 0.54 | 0 | 0.04 | 1.66 | 0.96 | 1.6 | 4.95 | 0.42 | 1.71
| Analytical results for Gd in OREAS 556b (Certified Value 4.92 ppm) | [28 empty] | Certified
Gd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  |  |  |  | 1
| No | No | 02: 02; 03: 04; 04: 05; 05: 07; 06: 14; 07: 16; 08: 18; 09: 19; 10: 23; 11: 25; 12: 26 |  |  |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*MS; 11: 4A*OES/MS; 12: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 02: 3.4; 03: 4.91; 04: 6.7; 05: 6.2683; 06: 3; 07: 11.3065; 08: 5.8836; 09: 3.05; 10: 4.45; 11: 5.42; 12: 5.02 |  |  |  |  | 1
| 1 | 2 | 02: 3.4; 03: 4.92; 04: 6.6; 05: 6.0868; 06: 3; 07: 11.4299; 08: 5.8086; 09: 3.17; 10: 4.37; 11: 4.93; 12: 5.06 |  |  |  |  | 14
| 1 | 3 | 02: 3.5; 03: 4.9; 04: 6; 05: 6.1142; 06: 3; 07: 11.7477; 08: 5.9067; 09: 2.97; 10: 4.44; 11: 5.08; 12: 5.03 |  |  |  |  | 16
| 1 | 4 | 02: 3.4; 03: 5.05; 04: 6.5; 05: 6.2759; 06: 3; 07: 11.9338; 08: 5.8752; 09: 3.13; 10: 4.64; 11: 4.92; 12: 5.21 |  |  |  |  | 4.9221
| 1 | 5 | 02: 3.4; 03: 4.84; 04: 6.8; 05: 6.1064; 06: 3; 07: 11.7565; 08: 5.7144; 09: 2.99; 10: 4.3; 11: 4.99; 12: 5.03 |  |  |  |  | 105
| 1 | 6 | 02: 3.3; 03: 4.75; 04: 6.5; 05: 6.3353; 06: 2; 07: 11.0687; 08: 5.7062; 09: 2.93; 10: 4.45; 11: 4.48; 12: 4.97
| Mean |  | 02: 3.4; 03: 4.895; 04: 6.5167; 05: 6.1978; 06: 2.8333; 07: 11.5405; 08: 5.8158; 09: 3.04; 10: 4.4417; 11: 4.97; 12: 5.0533
| Median |  | 02: 3.4; 03: 4.905; 04: 6.55; 05: 6.1912; 06: 3; 07: 11.5888; 08: 5.8419; 09: 3.02; 10: 4.445; 11: 4.96; 12: 5.03
| Std Dev. |  | 02: 0.063246; 03: 0.098944; 04: 0.27869; 05: 0.10738; 06: 0.40825; 07: 0.32675; 08: 0.087994; 09: 0.094446; 10: 0.11374; 11: 0.30305; 12: 0.082138
| Rel.Std.Dev. |  | 02: 1.86%; 03: 2.021%; 04: 4.277%; 05: 1.733%; 06: 14.41%; 07: 2.831%; 08: 1.513%; 09: 3.107%; 10: 2.561%; 11: 6.098%; 12: 1.625%
| PDM3 |  | 02: -30.92%; 03: -0.5507%; 04: 32.4%; 05: 25.92%; 06: -42.44%; 07: 134.5%; 08: 18.16%; 09: -38.24%; 10: -9.761%; 11: 0.9731%; 12: 2.666%
| Z-Score (Absolute) |  | 02: 1.09; 03: 0.08; 04: 1.02; 05: 0.8; 06: NA; 07: 4.42; 08: 0.54; 09: 1.34; 10: 0.39; 11: 0.03; 12: 0.03
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 556b (Certified Value 0.23 ppm) | [28 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [10 empty] | 2
2
| 1 | 1 | 0.2 | 2.6 | 0.1 | 0.5 | 0.22 | 0.22 | 0.23 | 0.2 | 0.27 | 1.317 | 2.3 | 0.23694 | NA | 8.3 | 1.85 | 1.8 | 0.1 | [10 empty] | 1
| 1 | 2 | 0.17 | 2.6 | < 0.1 | 0.49 | 0.21 | 0.2 | 0.29 | 0.21 | 0.24 | 1.2984 | 2.4 | 0.22183 | NA | 8.2 | 2.07 | 2 | 0.1 | [10 empty] | 15
| 1 | 3 | 0.15 | 2.4 | < 0.1 | 0.56 | 0.22 | 0.21 | 0.23 | 0.26 | 0.27 | 1.2896 | 2.3 | 0.24242 | NA | 8.3 | 1.96 | 1.8 | 0.12 | [10 empty] | 16
| 1 | 4 | 0.17 | 2.7 | < 0.1 | 0.52 | 0.22 | 0.21 | 0.35 | 0.2 | 0.28 | 1.3205 | 2.3 | 0.23526 | NA | 8.7 | 1.92 | 1.9 | 0.11 | [10 empty] | 0.22821
| 1 | 5 | 0.23 | 2.5 | < 0.1 | 0.49 | 0.23 | 0.21 | 0.28 | 0.24 | 0.23 | 1.3093 | 2.2 | 0.22684 | NA | 8.6 | 1.87 | 2 | 0.11 | [10 empty] | 106
| 1 | 6 | 0.26 | 2.7 | < 0.1 | 0.44 | 0.23 | 0.2 | 0.28 | 0.21 | 0.25 | 1.2998 | 2.2 | 0.22939 | NA | 8.2 | 1.83 | 1.9 | 0.1
| Mean |  | 0.19667 | 2.5833 | 0.1 | 0.5 | 0.22167 | 0.20833 | 0.27667 | 0.22 | 0.25667 | 1.3058 | 2.2833 | 0.23211 |  | 8.3833 | 1.9167 | 1.9 | 0.10667
| Median |  | 0.185 | 2.6 | 0.1 | 0.495 | 0.22 | 0.21 | 0.28 | 0.21 | 0.26 | 1.3045 | 2.3 | 0.23232 |  | 8.3 | 1.895 | 1.9 | 0.105
| Std Dev. |  | 0.041793 | 0.1169 |  | 0.039497 | 0.0075277 | 0.0075277 | 0.044572 | 0.024495 | 0.019664 | 0.011874 | 0.075277 | 0.0074895 |  | 0.2137 | 0.088919 | 0.089443 | 0.008165
| Rel.Std.Dev. |  | 21.25% | 4.525% |  | 7.899% | 3.396% | 3.613% | 16.11% | 11.13% | 7.661% | 0.9093% | 3.297% | 3.227% |  | 2.549% | 4.639% | 4.708% | 7.655%
| PDM3 |  | -13.82% | 1032% | -56.18% | 119.1% | -2.866% | -8.708% | 21.24% | -3.596% | 12.47% | 472.2% | 900.6% | 1.712% |  | 3574% | 739.9% | 732.6% | -53.26%
| Z-Score (Absolute) |  | 0.67 | NA | NA | 5.1 | 0.2 | 0.45 | 0.85 | 0.23 | 0.47 | 20.42 | NA | 0 | NA | NA | 32.05 | NA | 2.39
Results from laboratories 4, 5, 17, 22 and 26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 556b (Certified Value 4.94 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 5.23 | 5.1 | 4.8 | 5.15 | 6.0372 | 4.5 | 4.7 | 5.2 | 5.2 | 5.2 | 9 | 4.9252 | 4.8617 | 4.9 | 4.9165 | 4.97 | 4.6 | 4.6 | 4.32 | 5.61 | 4.93 | 5 | 4.7 |  |  |  |  | 1
| 1 | 2 | 5.25 | 5.03 | 5 | 5.1 | 6.037 | 4.4 | 4.4 | 5.1 | 5 | 4.8 | 7 | 4.813 | 4.0949 | 5 | 5.0189 | 5.16 | 4.7 | 4.7 | 4.26 | 5.19 | 4.92 | 5 | 4.8 |  |  |  |  | 19
| 1 | 3 | 5.27 | 5 | 5 | 5.21 | 6.0363 | 4.7 | 4.8 | 5.4 | 5.1 | 5 | 13 | 4.7902 | 4.8635 | 5 | 4.9514 | 4.9 | 4.88 | 4.9 | 4.26 | 5.47 | 4.87 | 5.1 | 5 |  |  |  |  | 16
| 1 | 4 | 5.22 | 5.04 | 5 | 5.13 | 6.1304 | 4.7 | 4.8 | 5 | 5.1 | 5 | 7 | 4.941 | 4.6075 | 5.1 | 4.934 | 4.51 | 4.95 | 4.9 | 4.32 | 5.35 | 5.09 | 4.9 | 5.1 |  |  |  |  | 4.9448
| 1 | 5 | 5.11 | 4.85 | 5.2 | 5.01 | 6.1136 | 4.8 | 4.7 | 4.9 | 5.3 | 4.6 | 7 | 4.8094 | 4.7903 | 5 | 4.9777 | 4.62 | 4.95 | 4.7 | 4.17 | 5.34 | 4.93 | 4.9 | 5.2 |  |  |  |  | 107
| 1 | 6 | 5.03 | 4.86 | 5.1 | 4.96 | 5.9837 | 4.7 | 4.6 | 4.8 | 5.4 | 4.8 | 13 | 4.9156 | 4.0216 | 5 | 4.9831 | 4.87 | 4.83 | 5 | 4.14 | 5.01 | 4.98 | 4.8 | 4.6
| Mean |  | 5.185 | 4.98 | 5.0167 | 5.0933 | 6.0564 | 4.6333 | 4.6667 | 5.0667 | 5.1833 | 4.9 | 9.3333 | 4.8657 | 4.5399 | 5 | 4.9636 | 4.8383 | 4.8183 | 4.8 | 4.245 | 5.3283 | 4.9533 | 4.95 | 4.9
| Median |  | 5.225 | 5.015 | 5 | 5.115 | 6.0371 | 4.7 | 4.7 | 5.05 | 5.15 | 4.9 | 8 | 4.8643 | 4.6989 | 5 | 4.9646 | 4.885 | 4.855 | 4.8 | 4.26 | 5.345 | 4.93 | 4.95 | 4.9
| Std Dev. |  | 0.094181 | 0.10218 | 0.13292 | 0.092664 | 0.055087 | 0.15055 | 0.15055 | 0.21602 | 0.1472 | 0.20976 | 2.9439 | 0.068306 | 0.38528 | 0.063246 | 0.037082 | 0.2371 | 0.14162 | 0.15492 | 0.075299 | 0.20999 | 0.075542 | 0.10488 | 0.23664
| Rel.Std.Dev. |  | 1.816% | 2.052% | 2.649% | 1.819% | 0.9096% | 3.249% | 3.226% | 4.264% | 2.84% | 4.281% | 31.54% | 1.404% | 8.487% | 1.265% | 0.7471% | 4.9% | 2.939% | 3.227% | 1.774% | 3.941% | 1.525% | 2.119% | 4.829%
| PDM3 |  | 4.858% | 0.7118% | 1.453% | 3.004% | 22.48% | -6.299% | -5.625% | 2.464% | 4.824% | -0.9061% | 88.75% | -1.599% | -8.188% | 1.116% | 0.3803% | -2.153% | -2.558% | -2.928% | -14.15% | 7.756% | 0.1725% | 0.1051% | -0.9061%
| Z-Score (Absolute) |  | 1.27 | 0.15 | 0.35 | 0.77 | 6 | 1.73 | 1.55 | 0.62 | 1.26 | 0.28 | NA | 0.47 | 2.24 | 0.26 | 0.06 | 0.62 | 0.72 | 0.82 | 3.84 | 2.05 | 0.01 | 0.01 | 0.28
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 556b (Indicative Value 0.13 ppm) | [28 empty] | Indicative
Hg | Round | Replicate | 02: Lab; 04: Lab; 05: Lab |  |  |  |  | 1
| No | No | 02: 18; 04: 19; 05: 25 |  |  |  |  | ppm
|  |  | 02: 4A*MS; 04: 4A*OES/MS; 05: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 02: < 1; 04: NA; 05: 0.1 |  |  |  |  | 1
| 1 | 2 | 02: < 1; 04: NA; 05: 0.14 |  |  |  |  | 20
| 1 | 3 | 02: < 1; 04: NA; 05: 0.15 |  |  |  |  | 16
| 1 | 4 | 02: < 1; 04: NA; 05: 0.13 |  |  |  |  | 0.13
| 1 | 5 | 02: < 1; 04: NA; 05: 0.12 |  |  |  |  | 28
| 1 | 6 | 02: < 1; 04: NA; 05: 0.14
| Mean |  | 05: 0.13
| Median |  | 05: 0.135
| Std Dev. |  | 05: 0.017889
| Rel.Std.Dev. |  | 05: 13.76%
PDM3 | bar: 05
| Z-Score (Absolute) |  | 02: 0.67; 04: NA; 05: 0.67
| Analytical results for Ho in OREAS 556b (Certified Value 0.89 ppm) | [28 empty] | Certified
Ho | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 14: Lab |  |  |  |  | 1
| No | No | 02: 02; 04: 04; 05: 05; 06: 07; 07: 14; 08: 16; 09: 18; 10: 19; 11: 23; 12: 25; 14: 26 |  |  |  |  | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*OES/MS; 14: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 02: 0.6; 04: 0.89; 05: 1; 06: 1.2282; 07: 1; 08: 0.95109; 09: 0.9246; 10: 0.78; 11: 0.66; 12: 0.91; 14: 0.96 |  |  |  |  | 1
| 1 | 2 | 02: 0.6; 04: 0.9; 05: 1.1; 06: 1.2343; 07: 1; 08: 0.9865; 09: 0.88501; 10: 0.81; 11: 0.68; 12: 0.82; 14: 0.97 |  |  |  |  | 3
| 1 | 3 | 02: 0.6; 04: 0.9; 05: 1; 06: 1.2023; 07: 1; 08: 0.95698; 09: 0.90178; 10: 0.77; 11: 0.66; 12: 0.92; 14: 0.96 |  |  |  |  | 16
| 1 | 4 | 02: 0.6; 04: 0.89; 05: 1.1; 06: 1.2417; 07: 1; 08: 1.9348; 09: 0.92152; 10: 0.73; 11: 0.66; 12: 0.84; 14: 0.99 |  |  |  |  | 0.88682
| 1 | 5 | 02: 0.6; 04: 0.89; 05: 1.1; 06: 1.1999; 07: 1; 08: 1.9142; 09: 0.88244; 10: 0.74; 11: 0.64; 12: 0.84; 14: 0.97 |  |  |  |  | 108
| 1 | 6 | 02: 0.5; 04: 0.89; 05: 1.1; 06: 1.2205; 07: 1; 08: 0.95429; 09: 0.86388; 10: 0.78; 11: 0.63; 12: 0.79; 14: 0.96
| Mean |  | 02: 0.58333; 04: 0.89333; 05: 1.0667; 06: 1.2211; 07: 1; 08: 1.283; 09: 0.89654; 10: 0.76833; 11: 0.655; 12: 0.85333; 14: 0.96833
| Median |  | 02: 0.6; 04: 0.89; 05: 1.1; 06: 1.2243; 07: 1; 08: 0.97174; 09: 0.8934; 10: 0.775; 11: 0.66; 12: 0.84; 14: 0.965
| Std Dev. |  | 02: 0.040825; 04: 0.005164; 05: 0.05164; 06: 0.017023; 07: 0; 08: 0.49713; 09: 0.02382; 10: 0.029269; 11: 0.017607; 12: 0.051251; 14: 0.01169
| Rel.Std.Dev. |  | 02: 6.999%; 04: 0.5781%; 05: 4.841%; 06: 1.394%; 07: 0%; 08: 38.75%; 09: 2.657%; 10: 3.809%; 11: 2.688%; 12: 6.006%; 14: 1.207%
| PDM3 |  | 02: -34.22%; 04: 0.7342%; 05: 20.28%; 06: 37.7%; 07: 12.76%; 08: 44.67%; 09: 1.096%; 10: -13.36%; 11: -26.14%; 12: -3.776%; 14: 9.191%
| Z-Score (Absolute) |  | 02: 1.41; 04: 0.01; 05: 0.78; 06: 1.47; 07: NA; 08: 1.75; 09: 0.01; 10: 0.57; 11: 1.08; 12: 0.19; 14: 0.33
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for In in OREAS 556b (Certified Value 3.56 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 19 | 20 | 22 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 3.49 | 3.57 | 3.7 | 3.7 | 3.84 | 3.5768 | 3.46 | 3.44 | 3.38 | 3.8 | 3.64 | 3.4854 | 3.76 | 3.468 | 3.8 | 3.21 | 2.824 | 3.48 | 3.77 | 3.289 | [7 empty] | 1
| 1 | 2 | 3.62 | 3.48 | 3.71 | 3.8 | 3.7 | 3.5031 | 3.42 | 3.16 | 3.38 | 3.66 | 3.53 | 3.4828 | 3.8 | 3.696 | 3.7 | 3.27 | 2.819 | 3.29 | 3.68 | 3.286 | [7 empty] | 4
| 1 | 3 | 3.56 | 3.55 | 3.68 | 3.7 | 3.82 | 3.4923 | 3.6 | 3.5 | 3.48 | 3.8 | 3.8 | 3.5088 | 3.75 | 3.454 | 3.8 | 3.61 | 2.753 | 3.34 | 3.65 | 3.299 | [7 empty] | 16
| 1 | 4 | 3.51 | 3.42 | 3.63 | 3.8 | 3.69 | 3.5417 | 3.53 | 3.39 | 3.51 | 3.69 | 3.65 | 3.492 | 3.8 | 3.283 | 4 | 3.31 | 2.762 | 3.39 | 3.66 | 3.342 | [7 empty] | 3.564
| 1 | 5 | 3.5 | 3.21 | 3.67 | 4 | 3.63 | 3.5911 | 3.67 | 3.35 | 3.46 | 3.81 | 3.67 | 3.5136 | 3.74 | 3.401 | 4 | 3.36 | 2.763 | 3.38 | 3.63 | 3.392 | [7 empty] | 109
| 1 | 6 | 3.26 | 3.3 | 3.56 | 3.8 | 3.69 | 3.5253 | 3.59 | 3.32 | 3.42 | 3.91 | 3.62 | 3.4893 | 3.7 | 3.541 | 3.8 | 3.48 | 2.73 | 3.16 | 3.6 | 3.212
| Mean |  | 3.49 | 3.4217 | 3.6583 | 3.8 | 3.7283 | 3.5384 | 3.545 | 3.36 | 3.4383 | 3.7783 | 3.6517 | 3.4953 | 3.7583 | 3.4738 | 3.85 | 3.3733 | 2.7752 | 3.34 | 3.665 | 3.3033
| Median |  | 3.505 | 3.45 | 3.675 | 3.8 | 3.695 | 3.5335 | 3.56 | 3.37 | 3.44 | 3.8 | 3.645 | 3.4907 | 3.755 | 3.461 | 3.8 | 3.335 | 2.7625 | 3.36 | 3.655 | 3.294
| Std Dev. |  | 0.12264 | 0.14247 | 0.055648 | 0.10954 | 0.082805 | 0.039492 | 0.093541 | 0.11713 | 0.053821 | 0.090646 | 0.087502 | 0.012803 | 0.038166 | 0.13858 | 0.12247 | 0.1476 | 0.037839 | 0.10826 | 0.058224 | 0.060371
| Rel.Std.Dev. |  | 3.514% | 4.164% | 1.521% | 2.883% | 2.221% | 1.116% | 2.639% | 3.486% | 1.565% | 2.399% | 2.396% | 0.3663% | 1.016% | 3.989% | 3.181% | 4.376% | 1.363% | 3.241% | 1.589% | 1.828%
| PDM3 |  | -2.075% | -3.993% | 2.648% | 6.623% | 4.612% | -0.7172% | -0.532% | -5.723% | -3.525% | 6.015% | 2.461% | -1.926% | 5.454% | -2.529% | 8.026% | -5.349% | -22.13% | -6.284% | 2.835% | -7.313%
| Z-Score (Absolute) |  | 0.12 | 0.44 | 0.65 | 1.31 | 0.98 | 0.1 | 0.13 | 0.73 | 0.36 | 1.21 | 0.62 | 0.1 | 1.12 | 0.2 | 1.54 | 0.66 | 3.43 | 0.82 | 0.68 | 0.99
| Analytical results for K in OREAS 556b (Certified Value 2.25 wt.%) | [28 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.2 | 2.4 | 2.32 | 2.2202 | 2.42 | 1.98 | 2.21 | 2.15 | 2.25 | 2.24 | 2.32 | 2.3763 | 1.951 | 2.127 | 2.2704 | 2.3 | 2.2497 | 2.2043 | 2.13 | 2.11 | 2.18 | 2.36 | 2.21 | 2.2778 | 2.33 | 2.27 |  | 1
| 1 | 2 | 2.18 | 1.99 | 2.34 | 2.2531 | 2.53 | 2.03 | 2.21 | 2.19 | 2.16 | 2.19 | 2.32 | 2.4003 | 1.943 | 2.0259 | 2.2971 | 2.28 | 2.2191 | 2.2353 | 2.04 | 2.16 | 2.19 | 2.35 | 2.37 | 2.3265 | 2.32 | 2.29 |  | 0
| 1 | 3 | 2.14 | 2.16 | 2.37 | 2.2145 | 2.51 | 1.97 | 2.25 | 2.19 | 2.26 | 2.29 | 2.34 | 2.3864 | 1.943 | 2.2039 | 2.2812 | 2.28 | 2.2544 | 2.2174 | 2.15 | 2.13 | 2.19 | 2.27 | 2.25 | 2.2879 | 2.34 | 2.29 |  | 16
| 1 | 4 | 2.18 | 2.08 | 2.31 | 2.2437 | 2.57 | 2.02 | 2.25 | 2.2 | 2.24 | 2.24 | 2.33 | 2.3866 | 1.918 | 2.3439 | 2.2181 | 2.21 | 2.2229 | 2.1905 | 2.18 | 2.16 | 2.21 | 2.25 | 2.24 | 2.2927 | 2.32 | 2.28 |  | 2.2457
| 1 | 5 | 2.19 | 2.03 | 2.34 | 2.2003 | 2.57 | 2.01 | 2.29 | 2.18 | 2.27 | 2.28 | 2.28 | 2.3967 | 1.926 | 2.203 | 2.2533 | 2.26 | 2.1927 | 2.1905 | 2.11 | 2.11 | 2.2 | 2.28 | 2.2 | 2.321 | 2.3 | 2.31 |  | 110
| 1 | 6 | 2.15 | 2.41 | 2.34 | 2.2543 | 2.35 | 1.98 | 2.24 | 2.16 | 2.28 | 2.36 | 2.27 | 2.3869 | 1.926 | 2.3613 | 2.2033 | 2.23 | 2.1711 | 2.1862 | 2.04 | 2.14 | 2.2 | 2.42 | 2.26 | 2.3019 | 2.31 | 2.29
| Mean |  | 2.1733 | 2.1783 | 2.3367 | 2.231 | 2.4917 | 1.9983 | 2.2417 | 2.1783 | 2.2433 | 2.2667 | 2.31 | 2.3889 | 1.9345 | 2.2108 | 2.2539 | 2.26 | 2.2183 | 2.204 | 2.1083 | 2.135 | 2.195 | 2.3217 | 2.255 | 2.3013 | 2.32 | 2.2883
| Median |  | 2.18 | 2.12 | 2.34 | 2.2319 | 2.52 | 1.995 | 2.245 | 2.185 | 2.255 | 2.26 | 2.32 | 2.3868 | 1.9345 | 2.2034 | 2.2618 | 2.27 | 2.221 | 2.1974 | 2.12 | 2.135 | 2.195 | 2.315 | 2.245 | 2.2973 | 2.32 | 2.29
| Std Dev. |  | 0.023381 | 0.18454 | 0.020656 | 0.022468 | 0.088638 | 0.024833 | 0.029944 | 0.019408 | 0.043205 | 0.05785 | 0.028284 | 0.0085451 | 0.012911 | 0.12783 | 0.036677 | 0.034059 | 0.032233 | 0.019187 | 0.057764 | 0.022583 | 0.010488 | 0.065549 | 0.06091 | 0.019127 | 0.014142 | 0.013292
| Rel.Std.Dev. |  | 1.076% | 8.472% | 0.884% | 1.007% | 3.557% | 1.243% | 1.336% | 0.891% | 1.926% | 2.552% | 1.224% | 0.3577% | 0.6674% | 5.782% | 1.627% | 1.507% | 1.453% | 0.8705% | 2.74% | 1.058% | 0.4778% | 2.823% | 2.701% | 0.8312% | 0.6096% | 0.5808%
| PDM3 |  | -3.223% | -3% | 4.05% | -0.6542% | 10.95% | -11.02% | -0.1799% | -3% | -0.1057% | 0.9333% | 2.863% | 6.375% | -13.86% | -1.553% | 0.3648% | 0.6364% | -1.219% | -1.856% | -6.117% | -4.93% | -2.258% | 3.382% | 0.4138% | 2.476% | 3.308% | 1.898%
| Z-Score (Absolute) |  | 0.76 | 0.7 | 1.03 | 0.13 | 2.73 | 2.68 | 0.01 | 0.7 | 0.01 | 0.27 | 0.74 | 1.61 | 3.38 | 0.35 | 0.13 | 0.19 | 0.27 | 0.42 | 1.47 | 1.18 | 0.52 | 0.87 | 0.14 | 0.64 | 0.85 | 0.5
| Analytical results for La in OREAS 556b (Certified Value 70 ppm) | [28 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  | 0
1
| 1 | 1 | 60.87 | 69.9 | 53.4 | 75.85 | 72.3 | 67.6 | 74.6467 | 66.8 | 69.2 | 68.9 | 68.3 | 72.7 | 55.28 | 70 | 67.0429 | 78.4 | 72.3 | 62.4 | 68 | 44.33 | 74.4 | 75 | 72.3 | 68.3 |  |  |  | 1
| 1 | 2 | 61.14 | 68.3 | 53.4 | 77.08 | 74 | 67.8 | 74.1018 | 63 | 62.5 | 70 | 65 | 68.8 | 53.86 | 71 | 65.6201 | 79.2 | 71.1 | 63.9 | 70 | 44.8 | 67.7 | 74 | 69.3 | 66.7 |  |  |  | 5
| 1 | 3 | 58.76 | 72.9 | 51.9 | 74.04 | 73.3 | 67.8 | 75.2425 | 68.3 | 67.5 | 70.9 | 66 | 71.5 | 54.48 | 71 | 66.7161 | 78.6 | 71.5 | 62.6 | 67 | 46.91 | 72.6 | 72 | 71.3 | 69.6 |  |  |  | 16
| 1 | 4 | 60.5 | 64.8 | 54.5 | 73.43 | 73.5 | 67 | 74.0125 | 67.3 | 65.3 | 70.1 | 66.3 | 73.9 | 54.83 | 70 | 65.6017 | 77.8 | 71.8 | 62.8 | 75 | 46.57 | 71.3 | 72 | 70.7 | 70.4 |  |  |  | 69.7876
| 1 | 5 | 62.06 | 66.4 | 51.9 | 73.43 | 77.2 | 65.2 | 74.0544 | 69.1 | 65.1 | 67.4 | 69.4 | 70.6 | 54.87 | 71 | 64.748 | 76 | 71.3 | 63.3 | 72 | 44.27 | 70.7 | 71 | 69.8 | 71.6 |  |  |  | 111
| 1 | 6 | 60.27 | 68.8 | 54.1 | 74.85 | 75 | 66.1 | 74.4516 | 68 | 64.6 | 70.9 | 69.2 | 73.7 | 54.92 | 70 | 66.316 | 79.2 | 71.3 | 62.4 | 74 | 43.76 | 66.2 | 73 | 68.3 | 72.2
| Mean |  | 60.6 | 68.5167 | 53.2 | 74.78 | 74.2167 | 66.9167 | 74.4182 | 67.0833 | 65.7 | 69.7 | 67.3667 | 71.8667 | 54.7067 | 70.5 | 66.0075 | 78.2 | 71.55 | 62.9 | 71 | 45.1067 | 70.4833 | 72.8333 | 70.2833 | 69.8
| Median |  | 60.685 | 68.55 | 53.4 | 74.445 | 73.75 | 67.3 | 74.2767 | 67.65 | 65.2 | 70.05 | 67.3 | 72.1 | 54.85 | 70.5 | 65.9681 | 78.5 | 71.4 | 62.7 | 71 | 44.565 | 71 | 72.5 | 70.25 | 70
| Std Dev. |  | 1.0955 | 2.8124 | 1.0918 | 1.4582 | 1.7081 | 1.0629 | 0.47536 | 2.1535 | 2.3435 | 1.3461 | 1.8425 | 1.9664 | 0.48669 | 0.54772 | 0.84504 | 1.2 | 0.43704 | 0.5933 | 3.2249 | 1.3118 | 3.0538 | 1.472 | 1.4428 | 2.062
| Rel.Std.Dev. |  | 1.808% | 4.105% | 2.052% | 1.95% | 2.302% | 1.588% | 0.6388% | 3.21% | 3.567% | 1.931% | 2.735% | 2.736% | 0.8896% | 0.7769% | 1.28% | 1.535% | 0.6108% | 0.9432% | 4.542% | 2.908% | 4.333% | 2.021% | 2.053% | 2.954%
| PDM3 |  | -13.17% | -1.821% | -23.77% | 7.154% | 6.347% | -4.114% | 6.635% | -3.875% | -5.857% | -0.1255% | -3.469% | 2.979% | -21.61% | 1.021% | -5.417% | 12.05% | 2.525% | -9.869% | 1.737% | -35.37% | 0.997% | 4.364% | 0.7104% | 0.01781%
| Z-Score (Absolute) |  | 2.09 | 0.28 | 3.77 | 1.15 | 1.02 | 0.65 | 1.06 | 0.61 | 0.92 | 0.01 | 0.54 | 0.48 | 3.43 | 0.17 | 0.85 | 1.93 | 0.41 | 1.56 | 0.28 | 5.62 | 0.17 | 0.7 | 0.12 | 0.01
| Analytical results for Li in OREAS 556b (Certified Value 52 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 53.11 | 58.2 | 56.5 | 53.8 | 49 | 54.8 | 50.4773 | 50.5 | 50 | 51.3 | 52.5 | 54.9 | 53.47 | 40 | 51.2136 | 46.0463 | 53.7 | 52.1919 | 48 | 53 | 52 | 51 | 43.6 | 53 | 48.9 | 41.7 |  | 1
| 1 | 2 | 53.46 | 55.7 | 53.8 | 52.6 | 50 | 54.4 | 49.1168 | 50.3 | 51 | 49.8 | 51.4 | 54.6 | 53.24 | 41 | 51.0922 | 46.0361 | 53.6 | 51.9824 | 48 | 50.4 | 51 | 51.8 | 49.4 | 53 | 46.9 | 41.5 |  | 33
| 1 | 3 | 52.54 | 59.5 | 53.4 | 52.1 | 50 | 56.7 | 49.1133 | 51.3 | 50.7 | 50.9 | 53.7 | 55.1 | 53.34 | 41 | 51.3253 | 46.6683 | 54.2 | 52.5649 | 48 | 52.8 | 50 | 53.5 | 44.8 | 52 | 48.3 | 43.4 |  | 16
| 1 | 4 | 53.43 | 52.7 | 54.6 | 50.2 | 50 | 54.3 | 50.0759 | 51.2 | 51.4 | 51.4 | 52.6 | 54.9 | 53.36 | 40 | 51.8593 | 46.5313 | 54.9 | 51.881 | 48 | 54.4 | 53 | 55.9 | 46.3 | 53 | 48 | 43.8 |  | 51.6024
| 1 | 5 | 53.99 | 53.9 | 53.3 | 49.7 | 50 | 53.1 | 49.2281 | 52.3 | 50.8 | 52 | 53.2 | 53.9 | 53.37 | 41 | 51.4551 | 46.7484 | 54.1 | 51.237 | 48 | 53.7 | 50 | 54.9 | 47.5 | 51 | 48.7 | 44.1 |  | 112
| 1 | 6 | 53.02 | 56.5 | 53.6 | 48.8 | 49 | 53.5 | 48.7943 | 52 | 50.3 | 52.2 | 55.2 | 53.3 | 53.43 | 40 | 51.8729 | 46.8909 | 54.8 | 51.0489 | 48 | 52.7 | 52 | 49.9 | 48.1 | 52 | 50.1 | 41.5
| Mean |  | 53.2583 | 56.0833 | 54.2 | 51.2 | 49.6667 | 54.4667 | 49.4676 | 51.2667 | 50.7 | 51.2667 | 53.1 | 54.45 | 53.3683 | 40.5 | 51.4697 | 46.4869 | 54.2167 | 51.8177 | 48 | 52.8333 | 51.3333 | 52.8333 | 46.6167 | 52.3333 | 48.4833 | 42.6667
| Median |  | 53.27 | 56.1 | 53.7 | 51.15 | 50 | 54.35 | 49.1724 | 51.25 | 50.75 | 51.35 | 52.9 | 54.75 | 53.365 | 40.5 | 51.3902 | 46.5998 | 54.15 | 51.9317 | 48 | 52.9 | 51.5 | 52.65 | 46.9 | 52.5 | 48.5 | 42.55
| Std Dev. |  | 0.48996 | 2.5569 | 1.2182 | 1.9256 | 0.5164 | 1.2596 | 0.65556 | 0.79162 | 0.498 | 0.86178 | 1.2869 | 0.70356 | 0.079352 | 0.54772 | 0.32968 | 0.36437 | 0.54191 | 0.5758 | 0 | 1.3545 | 1.2111 | 2.3304 | 2.1554 | 0.8165 | 1.0591 | 1.2275
| Rel.Std.Dev. |  | 0.92% | 4.559% | 2.248% | 3.761% | 1.04% | 2.313% | 1.325% | 1.544% | 0.9822% | 1.681% | 2.423% | 1.292% | 0.1487% | 1.352% | 0.6405% | 0.7838% | 0.9995% | 1.111% | 0% | 2.564% | 2.359% | 4.411% | 4.624% | 1.56% | 2.184% | 2.877%
| PDM3 |  | 3.209% | 8.683% | 5.034% | -0.7799% | -3.751% | 5.551% | -4.137% | -0.6507% | -1.749% | -0.6507% | 2.902% | 5.518% | 3.422% | -21.52% | -0.2572% | -9.913% | 5.066% | 0.4171% | -6.981% | 2.385% | -0.5215% | 2.385% | -9.662% | 1.416% | -6.044% | -17.32%
| Z-Score (Absolute) |  | 0.66 | 1.67 | 1 | 0.07 | 0.62 | 1.09 | 0.69 | 0.05 | 0.25 | 0.05 | 0.6 | 1.08 | 0.7 | 3.88 | 0.02 | 1.75 | 1 | 0.15 | 1.21 | 0.51 | 0.02 | 0.51 | 1.7 | 0.33 | 1.04 | 3.11
| Analytical results for Lu in OREAS 556b (Certified Value 0.41 ppm) | [28 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 07; 05: 18; 06: 19; 07: 23; 08: 25; 09: 26 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS |  | 2
2
| 1 | 1 | 01: 0.3; 02: 0.45; 03: 0.4; 04: 0.46761; 05: 0.40573; 06: NA; 07: 0.32; 08: 0.44; 09: 0.41 |  | 1
| 1 | 2 | 01: 0.4; 02: 0.44; 03: 0.4; 04: 0.48017; 05: 0.41636; 06: NA; 07: 0.32; 08: 0.39; 09: 0.42 |  | 25
| 1 | 3 | 01: 0.4; 02: 0.44; 03: 0.4; 04: 0.47233; 05: 0.40283; 06: NA; 07: 0.32; 08: 0.42; 09: 0.41 |  | 16
| 1 | 4 | 01: 0.3; 02: 0.42; 03: 0.4; 04: 0.47127; 05: 0.4117; 06: NA; 07: 0.33; 08: 0.41; 09: 0.42 |  | 0.40632
| 1 | 5 | 01: 0.3; 02: 0.42; 03: 0.4; 04: 0.45063; 05: 0.41778; 06: NA; 07: 0.31; 08: 0.39; 09: 0.41 |  | 113
| 1 | 6 | 01: 0.3; 02: 0.4; 03: 0.5; 04: 0.46635; 05: 0.39493; 06: NA; 07: 0.31; 08: 0.37; 09: 0.4
| Mean |  | 01: 0.33333; 02: 0.42833; 03: 0.41667; 04: 0.46806; 05: 0.40822; 07: 0.31833; 08: 0.40333; 09: 0.41167
| Median |  | 01: 0.3; 02: 0.43; 03: 0.4; 04: 0.46944; 05: 0.40872; 07: 0.32; 08: 0.4; 09: 0.41
| Std Dev. |  | 01: 0.05164; 02: 0.018348; 03: 0.040825; 04: 0.0098196; 05: 0.0087367; 07: 0.0075277; 08: 0.025033; 09: 0.0075277
| Rel.Std.Dev. |  | 01: 15.49%; 02: 4.284%; 03: 9.798%; 04: 2.098%; 05: 2.14%; 07: 2.365%; 08: 6.207%; 09: 1.829%
| PDM3 |  | 01: -17.96%; 02: 5.416%; 03: 2.545%; 04: 15.19%; 05: 0.4669%; 07: -21.66%; 08: -0.7362%; 09: 1.315%
| Z-Score (Absolute) |  | 01: NA; 02: 0.99; 03: NA; 04: 3.14; 05: 0.09; 06: NA; 07: 4.94; 08: 0.36; 09: 0.09
Results from laboratories 2 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 556b (Certified Value 2.43 wt.%) | [28 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.51 | 2.52 | 2.4 | 2.3704 | 2.43 | 2.31 | 2.34 | 2.33 | 2.32 | 2.42 | 2.52 | 2.5756 | 2.177 | 2.3273 | 2.6197 | 2.47 | 2.4277 | 2.4396 | 2.35 | 2.29 | 2.23 | 2.421 | 2.42 | 2.4997 | 2.49 | 2.5 |  | 1
| 1 | 2 | 2.49 | 2.41 | 2.4 | 2.4518 | 2.54 | 2.37 | 2.33 | 2.39 | 2.26 | 2.37 | 2.51 | 2.5813 | 2.165 | 2.33 | 2.6401 | 2.45 | 2.4049 | 2.4671 | 2.25 | 2.31 | 2.3 | 2.333 | 2.53 | 2.4842 | 2.47 | 2.51 |  | 0
| 1 | 3 | 2.46 | 2.61 | 2.37 | 2.3558 | 2.52 | 2.33 | 2.39 | 2.39 | 2.31 | 2.46 | 2.55 | 2.5764 | 2.171 | 2.4755 | 2.613 | 2.43 | 2.4308 | 2.4516 | 2.35 | 2.22 | 2.26 | 2.25 | 2.39 | 2.451 | 2.49 | 2.52 |  | 16
| 1 | 4 | 2.49 | 2.28 | 2.45 | 2.3915 | 2.59 | 2.37 | 2.39 | 2.4 | 2.33 | 2.42 | 2.52 | 2.5769 | 2.141 | 2.5986 | 2.6704 | 2.42 | 2.4298 | 2.4383 | 2.39 | 2.23 | 2.29 | 2.299 | 2.34 | 2.4471 | 2.48 | 2.5 |  | 2.4254
| 1 | 5 | 2.52 | 2.35 | 2.41 | 2.3608 | 2.58 | 2.39 | 2.44 | 2.39 | 2.33 | 2.45 | 2.48 | 2.5774 | 2.165 | 2.4726 | 2.6025 | 2.5 | 2.4047 | 2.4085 | 2.32 | 2.37 | 2.27 | 2.322 | 2.31 | 2.3853 | 2.46 | 2.53 |  | 114
| 1 | 6 | 2.48 | 2.47 | 2.4 | 2.4001 | 2.39 | 2.35 | 2.41 | 2.37 | 2.36 | 2.55 | 2.46 | 2.5779 | 2.165 | 2.6532 | 2.6404 | 2.44 | 2.8464 | 2.4469 | 2.24 | 2.23 | 2.28 | 2.304 | 2.37 | 2.4399 | 2.46 | 2.52
| Mean |  | 2.4917 | 2.44 | 2.405 | 2.3884 | 2.5083 | 2.3533 | 2.3833 | 2.3783 | 2.3183 | 2.445 | 2.5067 | 2.5776 | 2.164 | 2.4762 | 2.631 | 2.4517 | 2.4907 | 2.442 | 2.3167 | 2.275 | 2.2717 | 2.3215 | 2.3933 | 2.4512 | 2.475 | 2.5133
| Median |  | 2.49 | 2.44 | 2.4 | 2.3809 | 2.53 | 2.36 | 2.39 | 2.39 | 2.325 | 2.435 | 2.515 | 2.5772 | 2.165 | 2.4741 | 2.6299 | 2.445 | 2.4287 | 2.4432 | 2.335 | 2.26 | 2.275 | 2.313 | 2.38 | 2.4491 | 2.475 | 2.515
| Std Dev. |  | 0.02137 | 0.119 | 0.025884 | 0.035538 | 0.081343 | 0.029439 | 0.041793 | 0.025626 | 0.033116 | 0.060249 | 0.032042 | 0.0019874 | 0.012247 | 0.13405 | 0.024444 | 0.029269 | 0.17467 | 0.019431 | 0.059889 | 0.059245 | 0.024833 | 0.056493 | 0.077115 | 0.039821 | 0.013784 | 0.012111
| Rel.Std.Dev. |  | 0.8576% | 4.877% | 1.076% | 1.488% | 3.243% | 1.251% | 1.754% | 1.077% | 1.428% | 2.464% | 1.278% | 0.0771% | 0.566% | 5.413% | 0.9291% | 1.194% | 7.013% | 0.7957% | 2.585% | 2.604% | 1.093% | 2.433% | 3.222% | 1.625% | 0.5569% | 0.4819%
| PDM3 |  | 2.734% | 0.6034% | -0.8397% | -1.524% | 3.421% | -2.97% | -1.733% | -1.939% | -4.413% | 0.8095% | 3.352% | 6.276% | -10.78% | 2.096% | 8.479% | 1.084% | 2.694% | 0.6859% | -4.482% | -6.2% | -6.337% | -4.282% | -1.321% | 1.065% | 2.046% | 3.627%
| Z-Score (Absolute) |  | 0.57 | 0.01 | 0.4 | 0.59 | 0.75 | 0.98 | 0.65 | 0.7 | 1.37 | 0.04 | 0.74 | 1.53 | 3.1 | 0.39 | 2.13 | 0.12 | 0.56 | 0.01 | 1.39 | 1.86 | 1.9 | 1.34 | 0.53 | 0.11 | 0.38 | 0.81
| Analytical results for Mn in OREAS 556b (Certified Value 0.117 wt.%) | [28 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.1117 | 0.1239 | 0.1249 | 0.1159 | 0.12 | 0.114 | 0.109 | 0.112 | 0.1175 | 0.1135 | 0.1235 | 0.1182 | 0.101 | 0.11362 | 0.11523 | 0.119 | 0.11701 | 0.1193 | 0.1231 | 0.112 | 0.1232 | 0.1113 | 0.128 | 0.1237 | 0.1236 | 0.1202 |  | 1
| 1 | 2 | 0.1085 | 0.1259 | 0.1246 | 0.119 | 0.124 | 0.116 | 0.1115 | 0.1175 | 0.11 | 0.114 | 0.122 | 0.1153 | 0.108 | 0.11742 | 0.118 | 0.115 | 0.12069 | 0.1199 | 0.1217 | 0.1141 | 0.1272 | 0.1136 | 0.115 | 0.1274 | 0.1219 | 0.1138 |  | 0
| 1 | 3 | 0.1053 | 0.1296 | 0.1243 | 0.1163 | 0.118 | 0.114 | 0.1175 | 0.114 | 0.1115 | 0.1165 | 0.1255 | 0.1164 | 0.108 | 0.11681 | 0.11643 | 0.118 | 0.11585 | 0.1187 | 0.1213 | 0.114 | 0.1246 | 0.1089 | 0.108 | 0.1225 | 0.1217 | 0.1187 |  | 16
| 1 | 4 | 0.1071 | 0.1137 | 0.1275 | 0.118 | 0.129 | 0.116 | 0.1105 | 0.1125 | 0.116 | 0.1145 | 0.1205 | 0.1172 | 0.108 | 0.12713 | 0.11675 | 0.115 | 0.11803 | 0.1195 | 0.1208 | 0.1118 | 0.1318 | 0.1087 | 0.102 | 0.1293 | 0.1228 | 0.1141 |  | 0.11719
| 1 | 5 | 0.1116 | 0.1145 | 0.1248 | 0.1167 | 0.123 | 0.118 | 0.117 | 0.1145 | 0.1115 | 0.1215 | 0.124 | 0.1181 | 0.101 | 0.12269 | 0.11528 | 0.116 | 0.11883 | 0.119 | 0.1207 | 0.113 | 0.1239 | 0.1108 | 0.102 | 0.127 | 0.1204 | 0.1152 |  | 115
| 1 | 6 | 0.1081 | 0.1229 | 0.1271 | 0.1168 | 0.126 | 0.114 | 0.114 | 0.1185 | 0.1135 | 0.122 | 0.117 | 0.1178 | 0.108 | 0.12589 | 0.11017 | 0.112 | 0.11475 | 0.1187 | 0.1206 | 0.1121 | 0.1342 | 0.1108 | 0.103 | 0.1246 | 0.1225 | 0.114
| Mean |  | 0.10872 | 0.12175 | 0.12553 | 0.11712 | 0.12333 | 0.11533 | 0.11325 | 0.11483 | 0.11333 | 0.117 | 0.12208 | 0.11717 | 0.10567 | 0.12059 | 0.11531 | 0.11583 | 0.11753 | 0.11918 | 0.12137 | 0.11283 | 0.12748 | 0.11068 | 0.10967 | 0.12575 | 0.12215 | 0.116
| Median |  | 0.1083 | 0.1234 | 0.12485 | 0.11675 | 0.1235 | 0.115 | 0.11275 | 0.11425 | 0.1125 | 0.1155 | 0.12275 | 0.1175 | 0.108 | 0.12005 | 0.11585 | 0.1155 | 0.11752 | 0.11915 | 0.12105 | 0.11255 | 0.1259 | 0.1108 | 0.1055 | 0.1258 | 0.1222 | 0.11465
| Std Dev. |  | 0.002527 | 0.0063585 | 0.0013895 | 0.0011618 | 0.0039833 | 0.001633 | 0.0035036 | 0.0026394 | 0.0029098 | 0.003821 | 0.0030235 | 0.0011325 | 0.0036148 | 0.0054447 | 0.00272 | 0.0024833 | 0.0021316 | 0.00047504 | 0.00094587 | 0.0010289 | 0.004546 | 0.0017904 | 0.010289 | 0.0025681 | 0.0010932 | 0.0027575
| Rel.Std.Dev. |  | 2.324% | 5.223% | 1.107% | 0.992% | 3.23% | 1.416% | 3.094% | 2.299% | 2.567% | 3.266% | 2.477% | 0.9666% | 3.421% | 4.515% | 2.359% | 2.144% | 1.814% | 0.3986% | 0.7793% | 0.9119% | 3.566% | 1.618% | 9.382% | 2.042% | 0.8949% | 2.377%
| PDM3 |  | -7.233% | 3.888% | 7.116% | -0.0656% | 5.239% | -1.587% | -3.365% | -2.014% | -3.294% | -0.1652% | 4.172% | -0.02294% | -9.836% | 2.9% | -1.607% | -1.161% | 0.2842% | 1.698% | 3.561% | -3.721% | 8.78% | -5.555% | -6.423% | 7.301% | 4.229% | -1.018%
| Z-Score (Absolute) |  | 1.4 | 0.79 | 1.42 | 0.01 | 1.05 | 0.29 | 0.64 | 0.37 | 0.63 | 0.01 | 0.84 | 0.02 | 1.88 | 0.59 | 0.29 | 0.21 | 0.08 | 0.36 | 0.72 | 0.71 | 1.75 | 1.07 | 1.24 | 1.46 | 0.85 | 0.18
| Analytical results for Mo in OREAS 556b (Certified Value 18.4 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | < 1 | 16.56 | 19.3 | 19.5 | 19 | 18.9 | 19.0181 | 17.2 | 16.95 | 18.75 | 19.7 | 19.8 | 16.48 | 14 | 17.7888 | 18.8 | 17.8927 | 18.4 | 19.64 | 17 | 17.6 | 13.7 | 18.4 | 19.4 | 18.2 | 17.17 |  | 1
| 1 | 2 | < 1 | 17.35 | 19.3 | 19.4 | 20 | 18.3 | 18.4892 | 16.85 | 16.2 | 19.05 | 19.05 | 19.55 | 17.03 | 14 | 17.7034 | 18.5 | 18.0167 | 19.4 | 18.38 | 18 | 18.3 | 13.53 | 17.6 | 19.6 | 17.8 | 17.72 |  | 7
| 1 | 3 | < 1 | 16.86 | 18.6 | 19.1 | 18 | 18.9 | 18.616 | 18.05 | 17.75 | 19.15 | 19.85 | 19.45 | 16.84 | 16 | 17.3915 | 19 | 17.6907 | 18.5 | 18.88 | 16 | 19 | 13.56 | 17.7 | 19.3 | 17.6 | 18.31 |  | 16
| 1 | 4 | < 1 | 16.29 | 19.4 | 18.6 | 19 | 18.4 | 18.9288 | 17.45 | 17.05 | 19.6 | 19.85 | 19.8 | 16.88 | 15 | 17.7117 | 18.6 | 17.8282 | 17.4 | 19.17 | 17 | 18.7 | 13.91 | 17.2 | 19.7 | 18.1 | 18.4 |  | 18.353
| 1 | 5 | < 1 | 16.69 | 18.8 | 18.7 | 18 | 18.2 | 19.047 | 17.85 | 17.45 | 18.5 | 20.7 | 19.15 | 16.93 | 18 | 17.3925 | 18.5 | 17.9153 | 18.6 | 19.25 | 18 | 18.8 | 13.95 | 17.8 | 19.4 | 17.9 | 18.1 |  | 116
| 1 | 6 | < 1 | 15.29 | 19 | 18.2 | 19 | 18.4 | 19.0888 | 17.8 | 17 | 18.85 | 21 | 19.4 | 16.92 | 13 | 17.6936 | 18.6 | 17.5721 | 18.7 | 18.84 | 16 | 19 | 13.61 | 16.5 | 19.6 | 17.6 | 18.49
| Mean |  |  | 16.5067 | 19.0667 | 18.9167 | 18.8333 | 18.5167 | 18.8646 | 17.5333 | 17.0667 | 18.9833 | 20.025 | 19.525 | 16.8467 | 15 | 17.6136 | 18.6667 | 17.8193 | 18.5 | 19.0267 | 17 | 18.5667 | 13.71 | 17.5333 | 19.5 | 17.8667 | 18.0317
| Median |  |  | 16.625 | 19.15 | 18.9 | 19 | 18.4 | 18.9734 | 17.625 | 17.025 | 18.95 | 19.85 | 19.5 | 16.9 | 14.5 | 17.6985 | 18.6 | 17.8605 | 18.55 | 19.025 | 17 | 18.75 | 13.655 | 17.65 | 19.5 | 17.85 | 18.205
| Std Dev. |  |  | 0.69278 | 0.32042 | 0.50365 | 0.75277 | 0.30605 | 0.25056 | 0.4524 | 0.52599 | 0.37903 | 0.71046 | 0.2505 | 0.19054 | 1.7889 | 0.17492 | 0.19664 | 0.16207 | 0.64498 | 0.42922 | 0.89443 | 0.53914 | 0.18033 | 0.6377 | 0.15492 | 0.25033 | 0.50348
| Rel.Std.Dev. |  |  | 4.197% | 1.681% | 2.662% | 3.997% | 1.653% | 1.328% | 2.58% | 3.082% | 1.997% | 3.548% | 1.283% | 1.131% | 11.93% | 0.9931% | 1.053% | 0.9095% | 3.486% | 2.256% | 5.261% | 2.904% | 1.315% | 3.637% | 0.7945% | 1.401% | 2.792%
| PDM3 |  |  | -10.06% | 3.889% | 3.071% | 2.617% | 0.8919% | 2.788% | -4.466% | -7.009% | 3.435% | 9.11% | 6.386% | -8.207% | -18.27% | -4.029% | 1.709% | -2.908% | 0.8011% | 3.671% | -7.372% | 1.164% | -25.3% | -4.466% | 6.25% | -2.65% | -1.751%
| Z-Score (Absolute) |  | 17.83 | 1.97 | 0.56 | 0.41 | NA | 0.02 | 0.36 | 0.96 | 1.42 | 0.48 | 1.51 | 1.02 | 1.64 | NA | 0.88 | 0.17 | 0.67 | 0 | 0.52 | NA | 0.07 | 4.74 | 0.96 | 0.99 | 0.63 | 0.46
Results from laboratories 5, 14 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 556b (Certified Value 0.049 wt.%) | [28 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.04 | 0.072 | 0.04 | 0.0429 | 0.04 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.0515 | 0.037 | 0.054305 | 0.069784 | 0.046 | 0.047972 | NA | 0.07 | 0.45 | 0.11 | 0.0231 | 0.0446 | 0.04 | 0.05 |  |  | 1
| 1 | 2 | 0.06 | 0.066 | 0.04 | 0.0435 | 0.05 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.0512 | 0.037 | 0.062066 | 0.069357 | 0.043 | 0.047614 | NA | 0.06 | 0.48 | 0.1 | 0.0325 | 0.0445 | 0.04 | 0.05 |  |  | 17
| 1 | 3 | 0.06 | 0.07 | 0.04 | 0.0426 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.0513 | 0.037 | 0.05805 | 0.069729 | 0.044 | 0.047538 | NA | 0.06 | 0.38 | 0.12 | 0.0223 | 0.0431 | 0.04 | 0.05 |  |  | 16
| 1 | 4 | 0.05 | 0.062 | 0.04 | 0.0435 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.0514 | 0.037 | 0.060223 | 0.068995 | 0.044 | 0.047591 | NA | 0.06 | 0.4 | 0.11 | 0.021 | 0.0433 | 0.04 | 0.05 |  |  | 0.049017
| 1 | 5 | 0.04 | 0.549 | 0.04 | 0.0424 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.0515 | 0.037 | 0.05492 | 0.069778 | 0.045 | 0.047017 | NA | 0.06 | 0.37 | 0.1 | 0.0224 | 0.0423 | 0.04 | 0.05 |  |  | 117
| 1 | 6 | 0.05 | 0.067 | 0.04 | 0.0435 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.0515 | 0.037 | 0.053171 | 0.069606 | 0.044 | 0.047088 | NA | 0.05 | 0.39 | 0.11 | 0.0264 | 0.0438 | 0.05 | 0.05
| Mean |  | 0.05 | 0.14767 | 0.04 | 0.043067 | 0.048333 | 0.04 | 0.05 | 0.041667 | 0.048333 | 0.048333 | 0.05 | 0.0514 | 0.037 | 0.057123 | 0.069542 | 0.044333 | 0.04747 |  | 0.06 | 0.41167 | 0.10833 | 0.024617 | 0.0436 | 0.041667 | 0.05
| Median |  | 0.05 | 0.0685 | 0.04 | 0.0432 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05145 | 0.037 | 0.056485 | 0.069667 | 0.044 | 0.047564 |  | 0.06 | 0.395 | 0.11 | 0.02275 | 0.04355 | 0.04 | 0.05
| Std Dev. |  | 0.0089443 | 0.19664 | 0 | 0.00050067 | 0.0040825 | 0 | 7.6012e-18 | 0.0040825 | 0.0040825 | 0.0040825 | 7.6012e-18 | 0.00012649 | 0 | 0.0035586 | 0.00031201 | 0.0010328 | 0.0003587 |  | 0.0063246 | 0.043551 | 0.0075277 | 0.0042659 | 0.00088091 | 0.0040825 | 7.6012e-18
| Rel.Std.Dev. |  | 17.89% | 133.2% | 0% | 1.163% | 8.447% | 0% | 1.52e-14% | 9.798% | 8.447% | 8.447% | 1.52e-14% | 0.2461% | 0% | 6.23% | 0.4487% | 2.33% | 0.7556% |  | 10.54% | 10.58% | 6.949% | 17.33% | 2.02% | 9.798% | 1.52e-14%
| PDM3 |  | 2.005% | 201.3% | -18.4% | -12.14% | -1.395% | -18.4% | 2.005% | -15% | -1.395% | -1.395% | 2.005% | 4.861% | -24.52% | 16.54% | 41.87% | -9.555% | -3.157% |  | 22.41% | 739.8% | 121% | -49.78% | -11.05% | -15% | 2.005%
| Z-Score (Absolute) |  | 0.19 | 11.23 | 0.94 | 0.6 | 0 | 0.94 | 0.19 | 0.75 | 0 | 0 | 0.19 | 0.35 | 1.27 | 0.99 | 2.4 | 0.45 | 0.1 | NA | 1.32 | 41.06 | 6.78 | 2.68 | 0.53 | 0.75 | 0.19
| Analytical results for Nb in OREAS 556b (Certified Value 14 ppm) | [28 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 14.48 | 5.09 | 14.54 | 19.8 | 7.4 | 22.5826 | 15.9 | 13.4 | 15.4 | 14.2 | 17.8 | 63 | 16.1089 | 8.6 | 14.1063 | NA | 16.1 | 6 | 15.5 | 11.6 | 12.92 | 10 | 15 |  |  |  |  | 1
| 1 | 2 | 15.27 | 5.58 | 14 | 21.5 | 8.2 | 22.3635 | 15.1 | 13.1 | 15.6 | 13.2 | 17.4 | 60 | 15.223 | 8.9 | 14.696 | NA | 15.4 | 7 | 15.8 | 11.4 | 12.38 | 9.8 | 14.7 |  |  |  |  | 9
| 1 | 3 | 14.79 | 5.1 | 12.12 | 22.6 | 7.4 | 22.3907 | 16.3 | 14.2 | 16.4 | 14.1 | 16.9 | 61 | 16.07 | 8.6 | 14.4886 | NA | 16.1 | 6 | 16.3 | 13.1 | 12.73 | 9.7 | 15.1 |  |  |  |  | 16
| 1 | 4 | 14.03 | 5.08 | 13.42 | 23.3 | 7.2 | 22.713 | 15.6 | 13.6 | 16 | 14 | 17.8 | 64 | 15.658 | 8.5 | 14.3867 | NA | 16.5 | 7 | 16.3 | 11.7 | 12.3 | 9.8 | 15.4 |  |  |  |  | 14.0083
| 1 | 5 | 13.73 | 5.56 | 12.74 | 23.7 | 7.1 | 22.9143 | 16.5 | 14 | 15.5 | 13.7 | 17.6 | 61 | 15.2647 | 8.5 | 14.2533 | NA | 16.3 | 6 | 16.2 | 13.1 | 11.96 | 9.8 | 15.5 |  |  |  |  | 118
| 1 | 6 | 11.78 | 5.39 | 12.9 | 23 | 6.8 | 22.4136 | 16.4 | 12.8 | 15.8 | 14.9 | 17.6 | 60 | 15.8387 | 8.7 | 13.6883 | NA | 16.2 | 6 | 15.9 | 10.6 | 12.18 | 9.7 | 14.7
| Mean |  | 14.0133 | 5.3 | 13.2867 | 22.3167 | 7.35 | 22.5629 | 15.9667 | 13.5167 | 15.7833 | 14.0167 | 17.5167 | 61.5 | 15.6939 | 8.6333 | 14.2698 |  | 16.1 | 6.3333 | 16 | 11.9167 | 12.4117 | 9.8 | 15.0667
| Median |  | 14.255 | 5.245 | 13.16 | 22.8 | 7.3 | 22.4981 | 16.1 | 13.5 | 15.7 | 14.05 | 17.6 | 61 | 15.7483 | 8.6 | 14.32 |  | 16.15 | 6 | 16.05 | 11.65 | 12.34 | 9.8 | 15.05
| Std Dev. |  | 1.2222 | 0.23942 | 0.88396 | 1.4442 | 0.47223 | 0.21822 | 0.54283 | 0.53072 | 0.37103 | 0.56362 | 0.33714 | 1.6432 | 0.3852 | 0.15055 | 0.34882 |  | 0.37417 | 0.5164 | 0.32249 | 0.99482 | 0.35521 | 0.10954 | 0.33862
| Rel.Std.Dev. |  | 8.722% | 4.517% | 6.653% | 6.471% | 6.425% | 0.9672% | 3.4% | 3.926% | 2.351% | 4.021% | 1.925% | 2.672% | 2.454% | 1.744% | 2.444% |  | 2.324% | 8.154% | 2.016% | 8.348% | 2.862% | 1.118% | 2.248%
| PDM3 |  | 0.03626% | -62.17% | -5.151% | 59.31% | -47.53% | 61.07% | 13.98% | -3.509% | 12.67% | 0.06006% | 25.05% | 339% | 12.03% | -38.37% | 1.867% |  | 14.93% | -54.79% | 14.22% | -14.93% | -11.4% | -30.04% | 7.556%
| Z-Score (Absolute) |  | 0.05 | 3.24 | 0.31 | 3 | 2.49 | 3.09 | 0.67 | 0.23 | 0.6 | 0.05 | 1.24 | NA | 0.57 | 2.02 | 0.05 | NA | 0.72 | NA | 0.68 | 0.82 | 0.63 | 1.59 | 0.34
Results from laboratories 14 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 556b (Certified Value 63 ppm) | [28 empty] | Certified
Nd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  |  |  |  | 1
| No | No | 02: 02; 03: 04; 04: 05; 05: 07; 06: 14; 07: 16; 08: 18; 09: 19; 10: 23; 11: 25; 12: 26 |  |  |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*MS; 11: 4A*OES/MS; 12: 4A*MS |  |  |  |  | 0
1
| 1 | 1 | 02: 43.2; 03: 64.5; 04: 64; 05: 64.0262; 06: 66; 07: 62.062; 08: 60.7727; 09: 61.81; 10: 50.7; 11: 66.6; 12: 65.09 |  |  |  |  | 1
| 1 | 2 | 02: 46.3; 03: 64.18; 04: 64.2; 05: 62.7644; 06: 65; 07: 62.583; 08: 61.6581; 09: 64.03; 10: 50.4; 11: 61.3; 12: 64.47 |  |  |  |  | 10
| 1 | 3 | 02: 49.6; 03: 63.18; 04: 62.7; 05: 63.3667; 06: 64; 07: 60.3412; 08: 61.1987; 09: 60.09; 10: 52.4; 11: 64.9; 12: 64.47 |  |  |  |  | 16
| 1 | 4 | 02: 42.4; 03: 62.28; 04: 62.2; 05: 63.503; 06: 63; 07: 69.1314; 08: 61.7075; 09: 56.57; 10: 52.3; 11: 63.1; 12: 67.6 |  |  |  |  | 62.9858
| 1 | 5 | 02: 46.2; 03: 61.95; 04: 65.4; 05: 62.5906; 06: 64; 07: 67.5986; 08: 60.0553; 09: 59.48; 10: 50.2; 11: 62; 12: 64.52 |  |  |  |  | 119
| 1 | 6 | 02: 41.6; 03: 62.17; 04: 65.7; 05: 63.1444; 06: 64; 07: 60.5244; 08: 58.2862; 09: 61.6; 10: 49.6; 11: 58.2; 12: 64.6
| Mean |  | 02: 44.8833; 03: 63.0433; 04: 64.0333; 05: 63.2325; 06: 64.3333; 07: 63.7068; 08: 60.6131; 09: 60.5967; 10: 50.9333; 11: 62.6833; 12: 65.125
| Median |  | 02: 44.7; 03: 62.73; 04: 64.1; 05: 63.2555; 06: 64; 07: 62.3225; 08: 60.9857; 09: 60.845; 10: 50.55; 11: 62.55; 12: 64.56
| Std Dev. |  | 02: 3.0255; 03: 1.093; 04: 1.401; 05: 0.52157; 06: 1.0328; 07: 3.7415; 08: 1.2949; 09: 2.5292; 10: 1.1553; 11: 2.9294; 12: 1.235
| Rel.Std.Dev. |  | 02: 6.741%; 03: 1.734%; 04: 2.188%; 05: 0.8248%; 06: 1.605%; 07: 5.873%; 08: 2.136%; 09: 4.174%; 10: 2.268%; 11: 4.673%; 12: 1.896%
| PDM3 |  | 02: -28.74%; 03: 0.09131%; 04: 1.663%; 05: 0.3917%; 06: 2.139%; 07: 1.145%; 08: -3.767%; 09: -3.793%; 10: -19.14%; 11: -0.4802%; 12: 3.396%
| Z-Score (Absolute) |  | 02: 9.49; 03: 0; 04: 0.52; 05: 0.1; 06: 0.67; 07: 0.35; 08: 1.27; 09: 1.28; 10: 6.33; 11: 0.19; 12: 1.09
| Analytical results for Ni in OREAS 556b (Certified Value 337 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 315.8 | 332.5 | 341 | 345.6 | 333 | 325 | 354.9981 | 328 | 327 | 349 | 337 | 354 | 309 | 266 | 355.4623 | 347.97 | 355 | 276.2585 | 318.9 | 391 | 304 | 308.2 | 209 | 412 | 348 | 362.3 | 358.8 | 1
| 1 | 2 | 308.2 | 332 | 346 | 338.2 | 333 | 332 | 343.5958 | 327 | 341 | 331 | 337 | 351 | 305 | 283 | 352.0659 | 342.96 | 349 | 277.4365 | 321.8 | 366 | 300 | 306.4 | 206 | 335 | 358 | 354.2 | 341.1 | 40
| 1 | 3 | 293.8 | 348.6 | 342 | 346.9 | 343 | 351 | 339.0457 | 345 | 334 | 335 | 345 | 359 | 307 | 274 | 351.8813 | 344.32 | 348 | 271.8185 | 318.6 | 382 | 312 | 313.1 | 204 | 338 | 360 | 351.3 | 352.3 | 16
| 1 | 4 | 304.5 | 302.6 | 347 | 353.6 | 358 | 331 | 336.2039 | 327 | 334 | 350 | 342 | 350 | 308 | 284 | 347.7956 | 341.46 | 350 | 277.5663 | 320.8 | 386 | 310 | 309.6 | 203 | 326 | 338 | 349.8 | 341.9 | 336.7816
| 1 | 5 | 316.1 | 305.9 | 342 | 340.5 | 340 | 341 | 338.8194 | 340 | 332 | 341 | 344 | 356 | 309 | 273 | 348.7706 | 347.6 | 351 | 274.0866 | 317.8 | 376 | 305 | 307.9 | 205 | 331 | 339 | 348 | 342 | 120
| 1 | 6 | 309.2 | 329.9 | 342 | 335.8 | 334 | 322 | 349.7333 | 337 | 343 | 345 | 361 | 340 | 310 | 275 | 357.2141 | 345.02 | 353 | 266.5855 | 317.3 | 379 | 309 | 312.2 | 199 | 311 | 357 | 345.7 | 340.4
| Mean |  | 307.9333 | 325.25 | 343.3333 | 343.4333 | 340.1667 | 333.6667 | 343.7327 | 334 | 335.1667 | 341.8333 | 344.3333 | 351.6667 | 308 | 275.8333 | 352.1983 | 344.8883 | 351 | 273.9586 | 319.2 | 380 | 306.6667 | 309.5667 | 204.3333 | 342.1667 | 350 | 351.8833 | 346.0833
| Median |  | 308.7 | 330.95 | 342 | 343.05 | 337 | 331.5 | 341.3207 | 332.5 | 334 | 343 | 343 | 352.5 | 308.5 | 274.5 | 351.9736 | 344.67 | 350.5 | 275.1726 | 318.75 | 380.5 | 307 | 308.9 | 204.5 | 333 | 352.5 | 350.55 | 341.95
| Std Dev. |  | 8.2679 | 17.6205 | 2.5033 | 6.5473 | 9.6626 | 10.7269 | 7.2901 | 7.746 | 5.9133 | 7.6529 | 8.8468 | 6.5929 | 1.7889 | 6.7355 | 3.6623 | 2.5559 | 2.6077 | 4.2275 | 1.7516 | 8.6487 | 4.4572 | 2.6113 | 3.3267 | 35.5045 | 9.8184 | 5.8629 | 7.6376
| Rel.Std.Dev. |  | 2.685% | 5.418% | 0.7291% | 1.906% | 2.841% | 3.215% | 2.121% | 2.319% | 1.764% | 2.239% | 2.569% | 1.875% | 0.5808% | 2.442% | 1.04% | 0.7411% | 0.7429% | 1.543% | 0.5487% | 2.276% | 1.453% | 0.8435% | 1.628% | 10.38% | 2.805% | 1.666% | 2.207%
| PDM3 |  | -8.566% | -3.424% | 1.945% | 1.975% | 1.005% | -0.9249% | 2.064% | -0.8259% | -0.4795% | 1.5% | 2.242% | 4.42% | -8.546% | -18.1% | 4.578% | 2.407% | 4.222% | -18.65% | -5.22% | 12.83% | -8.942% | -8.081% | -39.33% | 1.599% | 3.925% | 4.484% | 2.762%
| Z-Score (Absolute) |  | 2.49 | 1.22 | 0.11 | 0.12 | 0.12 | 0.6 | 0.14 | 0.58 | 0.49 | 0 | 0.18 | 0.72 | 2.49 | 4.86 | 0.76 | 0.22 | 0.67 | 4.99 | 1.66 | 2.81 | 2.59 | 2.37 | 10.11 | 0.02 | 0.6 | 0.74 | 0.31
| Analytical results for P in OREAS 556b (Certified Value 0.033 wt.%) | [28 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.03 | 0.039 | 0.0921 | 0.0321 | 0.045 | 0.0605 | 0.032 | 0.031 | 0.032 | 0.033 | 0.035 | 0.0928 | 0.192 | 0.032732 | 0.035 | 0.030681 | 0.0291 | 0.03 | 0.03 | 0.032 | 0.0425 | 0.0333 | 0.03 | 0.0336 |  |  |  | 1
| 1 | 2 | 0.03 | 0.036 | 0.0927 | 0.0319 | 0.046 | 0.0627 | 0.032 | 0.033 | 0.031 | 0.032 | 0.035 | 0.0911 | 0.209 | 0.035108 | 0.033 | 0.0309 | 0.0299 | 0.03 | 0.03 | 0.033 | 0.0415 | 0.0329 | 0.03 | 0.0333 |  |  |  | 0
| 1 | 3 | 0.03 | 0.038 | 0.0919 | 0.0316 | 0.046 | 0.0609 | 0.032 | 0.032 | 0.032 | 0.033 | 0.036 | 0.092 | 0.209 | 0.03315 | 0.034 | 0.030347 | 0.0298 | 0.03 | 0.03 | 0.032 | 0.0407 | 0.0333 | 0.03 | 0.0339 |  |  |  | 16
| 1 | 4 | 0.03 | 0.035 | 0.0926 | 0.0319 | 0.046 | 0.062 | 0.032 | 0.032 | 0.033 | 0.032 | 0.035 | 0.0925 | 0.192 | 0.035324 | 0.034 | 0.030672 | 0.0292 | 0.03 | 0.03 | 0.033 | 0.0416 | 0.0326 | 0.03 | 0.0334 |  |  |  | 0.032675
| 1 | 5 | 0.03 | 0.036 | 0.0935 | 0.0316 | 0.046 | 0.0617 | 0.033 | 0.032 | 0.033 | 0.033 | 0.035 | 0.0926 | 0.209 | 0.032331 | 0.034 | 0.030191 | 0.0289 | 0.03 | 0.03 | 0.033 | 0.0409 | 0.0324 | 0.03 | 0.0337 |  |  |  | 121
| 1 | 6 | 0.03 | 0.037 | 0.0932 | 0.0322 | 0.045 | 0.0613 | 0.032 | 0.031 | 0.033 | 0.035 | 0.035 | 0.0928 | 0.192 | 0.035721 | 0.033 | 0.029352 | 0.0297 | 0.03 | 0.03 | 0.033 | 0.0352 | 0.0327 | 0.03 | 0.0336
| Mean |  | 0.03 | 0.036833 | 0.092667 | 0.031883 | 0.045667 | 0.061517 | 0.032167 | 0.031833 | 0.032333 | 0.033 | 0.035167 | 0.0923 | 0.2005 | 0.034061 | 0.033833 | 0.030357 | 0.029433 | 0.03 | 0.03 | 0.032667 | 0.0404 | 0.032867 | 0.03 | 0.033583
| Median |  | 0.03 | 0.0365 | 0.09265 | 0.0319 | 0.046 | 0.0615 | 0.032 | 0.032 | 0.0325 | 0.033 | 0.035 | 0.09255 | 0.2005 | 0.034129 | 0.034 | 0.030509 | 0.02945 | 0.03 | 0.03 | 0.033 | 0.0412 | 0.0328 | 0.03 | 0.0336
| Std Dev. |  | 0 | 0.001472 | 0.00061536 | 0.00024833 | 0.0005164 | 0.00079099 | 0.00040825 | 0.00075277 | 0.0008165 | 0.0010954 | 0.00040825 | 0.00065727 | 0.0093113 | 0.0014856 | 0.00075277 | 0.00055461 | 0.00041793 | 0 | 0 | 0.0005164 | 0.0026245 | 0.00037238 | 0 | 0.0002137
| Rel.Std.Dev. |  | 0% | 3.996% | 0.6641% | 0.7789% | 1.131% | 1.286% | 1.269% | 2.365% | 2.525% | 3.32% | 1.161% | 0.7121% | 4.644% | 4.362% | 2.225% | 1.827% | 1.42% | 0% | 0% | 1.581% | 6.496% | 1.133% | 0% | 0.6363%
| PDM3 |  | -8.186% | 12.73% | 183.6% | -2.422% | 39.76% | 88.27% | -1.555% | -2.575% | -1.045% | 0.9958% | 7.627% | 182.5% | 513.6% | 4.243% | 3.546% | -7.093% | -9.92% | -8.186% | -8.186% | -0.02441% | 23.64% | 0.5877% | -8.186% | 2.781%
| Z-Score (Absolute) |  | 0.72 | 0.95 | 14.62 | 0.26 | 3.12 | 7 | 0.19 | 0.27 | 0.15 | 0.02 | 0.55 | 14.53 | 41.07 | 0.28 | 0.22 | 0.63 | 0.86 | 0.72 | 0.72 | 0.07 | 1.83 | 0.02 | 0.72 | 0.16
| Analytical results for Pb in OREAS 556b (Certified Value 17.3 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 53.21 | 15.34 | 18.5 | 18.2 | 13 | 18.2 | 17.001 | 16.8 | 16.4 | 15 | 16.8 | 19 | 16.35 | 44 | 16.9982 | 18 | 18.2766 | 14 | 17.3 | 24 | 18.1 | 17.9 | 18.7 | 18 | 15.6 |  |  | 1
| 1 | 2 | 54.19 | 14.85 | 17.8 | 18.9 | 15 | 17.6 | 17.1881 | 16.9 | 15.4 | 15.1 | 16.1 | 17.7 | 17.68 | 36 | 16.9698 | 18 | 18.0932 | 13 | 16.8 | 26 | 17.6 | 21.1 | 18.7 | 18.1 | 15.8 |  |  | 11
| 1 | 3 | 52.49 | 15.24 | 17.7 | 18.4 | 16 | 18.1 | 17.2022 | 17.1 | 16.7 | 15.6 | 16.8 | 18.4 | 17.54 | 27 | 17.4107 | 18 | 18.1808 | 16 | 17.1 | 25 | 15.9 | 19.9 | 18.7 | 18.1 | 16.2 |  |  | 16
| 1 | 4 | 52.16 | 14.76 | 18 | 18 | 14 | 17.9 | 17.3531 | 16.8 | 16.7 | 15.5 | 16.4 | 18.2 | 17.52 | 18 | 17.8954 | 18 | 18.6052 | 14 | 17.8 | 24 | 15 | 19.4 | 18.6 | 17.2 | 16.5 |  |  | 17.2652
| 1 | 5 | 55.67 | 14.85 | 19.3 | 18 | 15 | 17.4 | 17.2638 | 17 | 16.3 | 15.9 | 17 | 17 | 17.66 | 57 | 17.6509 | 18 | 18.4117 | 15 | 16.9 | 26 | 17.5 | 19.1 | 18.7 | 17.3 | 16.4 |  |  | 122
| 1 | 6 | 51.48 | 14.44 | 19.2 | 18 | 13 | 17.5 | 17.2334 | 17 | 15.9 | 14.9 | 17.5 | 17.7 | 17.56 | 31 | 17.323 | 18 | 17.969 | 16 | 16.8 | 24 | 16.5 | 18 | 18.7 | 17.2 | 15.7
| Mean |  | 53.2 | 14.9133 | 18.4167 | 18.25 | 14.3333 | 17.7833 | 17.2069 | 16.9333 | 16.2333 | 15.3333 | 16.7667 | 18 | 17.385 | 35.5 | 17.3746 | 18 | 18.2561 | 14.6667 | 17.1167 | 24.8333 | 16.7667 | 19.2333 | 18.6833 | 17.65 | 16.0333
| Median |  | 52.85 | 14.85 | 18.25 | 18.1 | 14.5 | 17.75 | 17.2178 | 16.95 | 16.35 | 15.3 | 16.8 | 17.95 | 17.55 | 33.5 | 17.3669 | 18 | 18.2287 | 14.5 | 17 | 24.5 | 17 | 19.25 | 18.7 | 17.65 | 16
| Std Dev. |  | 1.5248 | 0.32995 | 0.70261 | 0.35637 | 1.2111 | 0.33116 | 0.11668 | 0.12111 | 0.50465 | 0.39328 | 0.48442 | 0.68993 | 0.51122 | 13.6638 | 0.3625 | 0 | 0.22855 | 1.2111 | 0.38687 | 0.98319 | 1.1793 | 1.2061 | 0.040825 | 0.45935 | 0.38297
| Rel.Std.Dev. |  | 2.866% | 2.212% | 3.815% | 1.953% | 8.449% | 1.862% | 0.6781% | 0.7152% | 3.109% | 2.565% | 2.889% | 3.833% | 2.941% | 38.49% | 2.086% | 0% | 1.252% | 8.257% | 2.26% | 3.959% | 7.033% | 6.271% | 0.2185% | 2.603% | 2.389%
| PDM3 |  | 208.1% | -13.62% | 6.669% | 5.704% | -16.98% | 3.001% | -0.3378% | -1.922% | -5.977% | -11.19% | -2.888% | 4.256% | 0.6936% | 105.6% | 0.6337% | 4.256% | 5.739% | -15.05% | -0.8606% | 43.83% | -2.888% | 11.4% | 8.214% | 2.228% | -7.135%
| Z-Score (Absolute) |  | 27.92 | 1.93 | 0.8 | 0.67 | NA | 0.31 | 0.14 | 0.35 | 0.9 | 1.6 | 0.48 | 0.48 | 0 | NA | 0.01 | NA | 0.68 | NA | 0.21 | 5.81 | 0.48 | 1.44 | 1.01 | 0.21 | 1.05
Results from laboratories 5, 14, 17 and 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 556b (Certified Value 16.5 ppm) | [28 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 07; 05: 14; 06: 16; 07: 18; 08: 19; 09: 23; 10: 25; 11: 26 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS |  |  | 1
2
| 1 | 1 | 01: 13.9; 02: 18.23; 03: 18.1; 04: 17.6835; 05: 11; 06: 16.6662; 07: 17.0295; 08: 14.57; 09: 14.11; 10: 18; 11: 18.51 |  |  | 1
| 1 | 2 | 01: 14.1; 02: 18.34; 03: 18.4; 04: 17.5391; 05: 13; 06: 16.9115; 07: 17.0822; 08: 15.01; 09: 14.14; 10: 16.6; 11: 18.2 |  |  | 12
| 1 | 3 | 01: 14.7; 02: 17.95; 03: 18.2; 04: 17.5882; 05: 12; 06: 16.7755; 07: 16.7232; 08: 14.74; 09: 14.58; 10: 18; 11: 17.64 |  |  | 16
| 1 | 4 | 01: 13.5; 02: 17.76; 03: 18; 04: 17.6184; 05: 12; 06: 16.3365; 07: 16.8648; 08: 14.51; 09: 14.61; 10: 17.3; 11: 18.17 |  |  | 16.5315
| 1 | 5 | 01: 13.9; 02: 17.58; 03: 18.9; 04: 17.3844; 05: 12; 06: 16.7121; 07: 16.3931; 08: 14.09; 09: 13.99; 10: 17.1; 11: 17.6 |  |  | 123
| 1 | 6 | 01: 13.5; 02: 17.82; 03: 18.6; 04: 17.4526; 05: 12; 06: 16.3177; 07: 15.6891; 08: 14.73; 09: 14.51; 10: 16.1; 11: 17.7
| Mean |  | 01: 13.9333; 02: 17.9467; 03: 18.3667; 04: 17.5444; 05: 12; 06: 16.6199; 07: 16.6303; 08: 14.6083; 09: 14.3233; 10: 17.1833; 11: 17.97
| Median |  | 01: 13.9; 02: 17.885; 03: 18.3; 04: 17.5637; 05: 12; 06: 16.6892; 07: 16.794; 08: 14.65; 09: 14.325; 10: 17.2; 11: 17.935
| Std Dev. |  | 01: 0.44572; 02: 0.28994; 03: 0.33862; 04: 0.11025; 05: 0.63246; 06: 0.24147; 07: 0.52331; 08: 0.30766; 09: 0.27318; 10: 0.75741; 11: 0.37502
| Rel.Std.Dev. |  | 01: 3.199%; 02: 1.616%; 03: 1.844%; 04: 0.6284%; 05: 5.27%; 06: 1.453%; 07: 3.147%; 08: 2.106%; 09: 1.907%; 10: 4.408%; 11: 2.087%
| PDM3 |  | 01: -15.72%; 02: 8.561%; 03: 11.1%; 04: 6.127%; 05: -27.41%; 06: 0.535%; 07: 0.5982%; 08: -11.63%; 09: -13.36%; 10: 3.943%; 11: 8.702%
| Z-Score (Absolute) |  | 01: 1.91; 02: 0.67; 03: 0.94; 04: 0.41; 05: NA; 06: 0.18; 07: 0.18; 08: 1.47; 09: 1.66; 10: 0.18; 11: 0.68
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Pt in OREAS 556b (Indicative Value 11.2 ppb) | [28 empty] | Indicative
Pt | Round | Replicate | 01: Lab |  |  | 1
| No | No | 01: 25 |  |  | ppb
|  |  | 01: 4A*OES/MS |  |  | 1
1
| 1 | 1 | 01: 8 |  |  | 1
| 1 | 2 | 01: 15 |  |  | 22
| 1 | 3 | 01: 12 |  |  | 16
| 1 | 4 | 01: 8 |  |  | 11.1667
| 1 | 5 | 01: 7 |  |  | 28
| 1 | 6 | 01: 17
| Mean |  | 01: 11.1667
| Median |  | 01: 10
| Std Dev. |  | 01: 4.1673
| Rel.Std.Dev. |  | 01: 37.32%
| PDM3 |  | 01: -2.998e-13%
Z-Score (Absolute) | bar: 01
| Analytical results for Rb in OREAS 556b (Certified Value 82 ppm) | [28 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 0
1
| 1 | 1 | 79 | 87.76 | 90.7 | 84.9 | 79.961 | 80.9 | 76.7 | 71.3 | 84.4 | 90.1 | 81.9392 | 90.9 | 81.0055 | 78.7 | 80.3 | 75.8 | 66.5 | 84.8 | 84.74 | 85.1 | 83.5 | [6 empty] | 1
| 1 | 2 | 69.4 | 87.16 | 91.2 | 82.3 | 79.0328 | 77 | 72.1 | 71.3 | 82.8 | 89.3 | 79.2893 | 91.5 | 82.8966 | 83 | 77.8 | 77.5 | 64.3 | 80.3 | 85.49 | 83.4 | 85.3 | [6 empty] | 13
| 1 | 3 | 73.2 | 86.57 | 81.1 | 84.2 | 80.809 | 83.9 | 78.2 | 71.8 | 86.1 | 88.4 | 81.524 | 90.6 | 83.052 | 79.2 | 80.7 | 75.8 | 67.9 | 80.6 | 83.48 | 84 | 88.9 | [6 empty] | 16
| 1 | 4 | 70.2 | 83.94 | 92.1 | 81.7 | 83.0919 | 82.7 | 77.3 | 76.2 | 84.2 | 87.7 | 79.1194 | 91.2 | 81.6508 | 75.2 | 81.4 | 81.1 | 69.1 | 81.1 | 84.31 | 83.3 | 88.5 | [6 empty] | 82.3175
| 1 | 5 | 74.7 | 85.23 | 95.4 | 80.8 | 81.8746 | 83.9 | 77.5 | 73.2 | 88.1 | 85.2 | 82.5893 | 91.2 | 82.8881 | 78.4 | 81.2 | 78.9 | 64.4 | 80.4 | 83.84 | 82.9 | 82.4 | [6 empty] | 124
| 1 | 6 | 74.2 | 83.55 | 89.3 | 81.9 | 77.7728 | 82.7 | 74.8 | 73.2 | 89 | 85.7 | 82.8337 | 91.2 | 79.6969 | 81 | 80.5 | 82.9 | 60.3 | 76 | 83.62 | 82.7 | 81.9
| Mean |  | 73.45 | 85.7017 | 89.9667 | 82.6333 | 80.4237 | 81.85 | 76.1 | 72.8333 | 85.7667 | 87.7333 | 81.2158 | 91.1 | 81.865 | 79.25 | 80.3167 | 78.6667 | 65.4167 | 80.5333 | 84.2467 | 83.5667 | 85.0833
| Median |  | 73.7 | 85.9 | 90.95 | 82.1 | 80.385 | 82.7 | 77 | 72.5 | 85.25 | 88.05 | 81.7316 | 91.2 | 82.2695 | 78.95 | 80.6 | 78.2 | 65.45 | 80.5 | 84.075 | 83.35 | 84.4
| Std Dev. |  | 3.4628 | 1.7363 | 4.7999 | 1.5795 | 1.9261 | 2.6182 | 2.2742 | 1.864 | 2.4138 | 1.9521 | 1.6265 | 0.30984 | 1.342 | 2.6289 | 1.3014 | 2.8877 | 3.1422 | 2.7983 | 0.7667 | 0.8756 | 3.0374
| Rel.Std.Dev. |  | 4.715% | 2.026% | 5.335% | 1.911% | 2.395% | 3.199% | 2.988% | 2.559% | 2.814% | 2.225% | 2.003% | 0.3401% | 1.639% | 3.317% | 1.62% | 3.671% | 4.803% | 3.475% | 0.9101% | 1.048% | 3.57%
| PDM3 |  | -10.77% | 4.111% | 9.292% | 0.3837% | -2.301% | -0.5679% | -7.553% | -11.52% | 4.19% | 6.579% | -1.338% | 10.67% | -0.5497% | -3.726% | -2.431% | -4.435% | -20.53% | -2.167% | 2.344% | 1.518% | 3.36%
| Z-Score (Absolute) |  | 1.78 | 0.82 | 1.72 | 0.17 | 0.3 | 0 | 1.22 | 1.91 | 0.83 | 1.25 | 0.13 | 1.96 | 0 | 0.55 | 0.32 | 0.67 | 3.48 | 0.28 | 0.51 | 0.36 | 0.68
| Analytical results for Re in OREAS 556b (Certified Value 0.033 ppm) | [28 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 28 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | 0.029 | 0.037 | 0.031 | 0.038 | 0.031 | 0.024 | 0.031 | 0.034 | 0.036 | < 0.05 | < 0.05 | 0.025631 | 0.034 | 0.033 | 0.031 | 0.03 | 0.0309 | 0.036 | 0.022 | [8 empty] | 1
| 1 | 2 | 0.03 | 0.04 | 0.032 | 0.034 | 0.035 | 0.023 | 0.028 | 0.036 | 0.033 | < 0.05 | < 0.05 | 0.026294 | 0.039 | 0.031 | 0.03 | 0.027 | 0.0375 | 0.037 | 0.021 | [8 empty] | 14
| 1 | 3 | 0.029 | 0.03 | 0.045 | 0.035 | 0.033 | 0.026 | 0.031 | 0.031 | 0.038 | < 0.05 | < 0.05 | 0.022831 | 0.033 | 0.035 | 0.031 | 0.031 | 0.0366 | 0.035 | 0.022 | [8 empty] | 16
| 1 | 4 | 0.028 | 0.035 | 0.032 | 0.033 | 0.031 | 0.024 | 0.035 | 0.031 | 0.036 | < 0.05 | < 0.05 | 0.025114 | 0.029 | 0.036 | 0.034 | 0.025 | 0.0362 | 0.032 | 0.024 | [8 empty] | 0.032724
| 1 | 5 | 0.03 | 0.038 | 0.032 | 0.029 | 0.033 | 0.023 | 0.028 | 0.028 | 0.035 | < 0.05 | < 0.05 | 0.024577 | 0.033 | 0.034 | 0.034 | 0.028 | 0.0351 | 0.032 | 0.023 | [8 empty] | 125
| 1 | 6 | 0.029 | 0.041 | 0.032 | 0.032 | 0.036 | 0.023 | 0.027 | 0.033 | 0.035 | < 0.05 | < 0.05 | 0.026533 | 0.034 | 0.032 | 0.036 | 0.025 | 0.0307 | 0.032 | 0.02
| Mean |  | 0.029167 | 0.036833 | 0.034 | 0.0335 | 0.033167 | 0.023833 | 0.03 | 0.032167 | 0.0355 |  |  | 0.025163 | 0.033667 | 0.0335 | 0.032667 | 0.027667 | 0.0345 | 0.034 | 0.022
| Median |  | 0.029 | 0.0375 | 0.032 | 0.0335 | 0.033 | 0.0235 | 0.0295 | 0.032 | 0.0355 |  |  | 0.025372 | 0.0335 | 0.0335 | 0.0325 | 0.0275 | 0.03565 | 0.0335 | 0.022
| Std Dev. |  | 0.00075277 | 0.0039707 | 0.0054037 | 0.0030166 | 0.0020412 | 0.001169 | 0.0029665 | 0.0027869 | 0.0016432 |  |  | 0.0013536 | 0.0032042 | 0.0018708 | 0.0023381 | 0.0025033 | 0.0029685 | 0.0022804 | 0.0014142
| Rel.Std.Dev. |  | 2.581% | 10.78% | 15.89% | 9.005% | 6.154% | 4.905% | 9.888% | 8.664% | 4.629% |  |  | 5.379% | 9.517% | 5.585% | 7.157% | 9.048% | 8.604% | 6.707% | 6.428%
| PDM3 |  | -10.87% | 12.56% | 3.9% | 2.372% | 1.353% | -27.17% | -8.324% | -1.703% | 8.484% |  |  | -23.1% | 2.881% | 2.372% | -0.1746% | -15.45% | 5.428% | 3.9% | -32.77%
| Z-Score (Absolute) |  | 0.89 | 1.05 | 0.34 | 0.21 | 0.13 | 2.23 | 0.67 | 0.13 | 0.72 | 1.94 | 1.94 | 1.9 | 0.25 | 0.21 | 0 | 1.26 | 0.46 | 0.34 | 2.7
| Analytical results for S in OREAS 556b (Certified Value 8.55 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.88 | 9.16 | 8.53 | 8.3822 | 9.13 | 6.18 | 8.23 | 8.06 | 8.9 | 8.19 | 8.82 | 8.91 | 8.3265 | 8.3728 | 8.69 | 8.8767 | 6.9826 | 8.21 | 8.66 | 8.14 | 8.581 | 8.57 | 8.29 | 8.28 | 8.55 |  |  | 1
| 1 | 2 | 6.51 | 8.72 | 8.48 | 8.4953 | 9.19 | 5.84 | 8.22 | 8.33 | 8.63 | 8.12 | 8.87 | 8.9 | 8.8742 | 8.3673 | 8.78 | 8.9737 | 7.0415 | 8.08 | 8.53 | 8.24 | 8.834 | 8.63 | 8.21 | 8.27 | 8.57 |  |  | 28
| 1 | 3 | 5.92 | 9.48 | 8.48 | 8.3173 | 9.21 | 6.07 | 8.35 | 8.24 | 8.79 | 8.29 | 8.99 | 8.91 | 8.5272 | 8.3408 | 8.67 | 8.9996 | 7.1135 | 8.08 | 8.5 | 8.16 | 8.958 | 8.2 | 8.27 | 8.29 | 8.63 |  |  | 16
| 1 | 4 | 7.23 | 8.35 | 8.48 | 8.4556 | 9.27 | 6.21 | 8.32 | 8.3 | 8.92 | 8.13 | 8.89 | 8.99 | 9.1355 | 8.323 | 8.55 | 8.946 | 6.9328 | 8.29 | 8.73 | 8.28 | 8.535 | 7.98 | 8.32 | 8.25 | 8.63 |  |  | 8.5467
| 1 | 5 | 6.62 | 8.97 | 8.42 | 8.308 | 9.15 | 6.23 | 8.58 | 8.3 | 8.94 | 8.4 | 8.71 | 8.95 | 8.5349 | 8.3998 | 8.72 | 8.9445 | 6.8362 | 8.15 | 8.52 | 8.21 | 8.387 | 7.85 | 8.38 | 8.17 | 8.67 |  |  | 126
| 1 | 6 | 6.71 | 9 | 8.39 | 8.5117 | 9 | 6.02 | 8.48 | 8.09 | 9.02 | 8.81 | 8.69 | 8.83 | 9.1502 | 8.3823 | 8.64 | 8.8542 | 7.0233 | 8.03 | 8.67 | 8.23 | 8.232 | 8.14 | 8.2 | 8.19 | 8.61
| Mean |  | 6.4783 | 8.9467 | 8.4633 | 8.4117 | 9.1583 | 6.0917 | 8.3633 | 8.22 | 8.8667 | 8.3233 | 8.8283 | 8.915 | 8.7581 | 8.3643 | 8.675 | 8.9324 | 6.9883 | 8.14 | 8.6017 | 8.21 | 8.5878 | 8.2283 | 8.2783 | 8.2417 | 8.61
| Median |  | 6.565 | 8.985 | 8.48 | 8.4189 | 9.17 | 6.125 | 8.335 | 8.27 | 8.91 | 8.24 | 8.845 | 8.91 | 8.7045 | 8.3701 | 8.68 | 8.9452 | 7.003 | 8.115 | 8.595 | 8.22 | 8.558 | 8.17 | 8.28 | 8.26 | 8.62
| Std Dev. |  | 0.51176 | 0.3849 | 0.050067 | 0.088827 | 0.091742 | 0.14825 | 0.14208 | 0.11645 | 0.13765 | 0.2609 | 0.11392 | 0.053572 | 0.34612 | 0.028012 | 0.077653 | 0.056161 | 0.095893 | 0.096747 | 0.096626 | 0.052154 | 0.27108 | 0.31352 | 0.067946 | 0.049967 | 0.043818
| Rel.Std.Dev. |  | 7.9% | 4.302% | 0.5916% | 1.056% | 1.002% | 2.434% | 1.699% | 1.417% | 1.552% | 3.135% | 1.29% | 0.6009% | 3.952% | 0.3349% | 0.8951% | 0.6287% | 1.372% | 1.189% | 1.123% | 0.6352% | 3.157% | 3.81% | 0.8208% | 0.6063% | 0.5089%
| PDM3 |  | -24.2% | 4.68% | -0.9752% | -1.58% | 7.157% | -28.72% | -2.145% | -3.822% | 3.744% | -2.613% | 3.295% | 4.309% | 2.473% | -2.134% | 1.501% | 4.514% | -18.23% | -4.758% | 0.6433% | -3.939% | 0.4815% | -3.725% | -3.14% | -3.569% | 0.7408%
| Z-Score (Absolute) |  | 6.46 | 1.79 | 0.17 | 0 | 2.5 | 7.76 | 0.16 | 0.64 | 1.52 | 0.3 | 1.39 | 1.68 | 1.16 | 0.16 | 0.88 | 1.74 | 4.76 | 0.91 | 0.64 | 0.67 | 0.59 | 0.61 | 0.45 | 0.57 | 0.66
| Analytical results for Sb in OREAS 556b (Certified Value 5.29 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 5 | 6.1 | < 5 | 5.58 | 4.3 | 4.67 | 5.3135 | 4.97 | 4.87 | 5.3 | 5.52 | 6.05 | < 5 | 5.5629 | 5.1 | 5.451 | NA | 5.41 | 9 | 6.1 | 4.67 | 4.3 | 5.82 | 4.6 | 5.2 |  |  | 1
| 1 | 2 | < 5 | 5.95 | < 5 | 5.51 | 5.6 | 5.14 | 5.1398 | 4.7 | 4.61 | 5.26 | 5.23 | 5.48 | < 5 | 5.7246 | 5 | 5.2947 | NA | 5.22 | 10 | 5.8 | 4.75 | 4.5 | 5.74 | 4.4 | 5.4 |  |  | 15
| 1 | 3 | < 5 | 6.18 | < 5 | 5.58 | 6.2 | 4.82 | 5.3299 | 5.02 | 5.01 | 5.42 | 5.5 | 5.52 | < 5 | 5.5086 | 5.3 | 5.3516 | NA | 5.39 | 11 | 6.2 | 4.67 | 4.4 | 5.54 | 4.5 | 5.7 |  |  | 16
| 1 | 4 | < 5 | 6.01 | < 5 | 5.25 | 6.1 | 4.71 | 5.4099 | 4.96 | 4.86 | 5.26 | 5.36 | 5.61 | < 5 | 5.6492 | 5.4 | 5.3195 | NA | 5.43 | 9 | 6 | 4.73 | 5 | 5.62 | 4.6 | 5.6 |  |  | 5.2929
| 1 | 5 | < 5 | 6.05 | < 5 | 5.39 | 5.9 | 4.46 | 5.266 | 5.19 | 4.87 | 5.17 | 5.54 | 5.8 | < 5 | 5.7361 | 5.4 | 5.1978 | NA | 5.43 | 11 | 5.9 | 4.72 | 4.5 | 5.68 | 4.4 | 5.7 |  |  | 127
| 1 | 6 | < 5 | 5.85 | < 5 | 5.33 | 5.7 | 4.76 | 5.2371 | 4.99 | 4.79 | 5.16 | 5.78 | 5.27 | < 5 | 5.4794 | 5.3 | 5.3097 | NA | 5.28 | 10 | 5.9 | 4.71 | 4.3 | 5.58 | 4.4 | 5.1
| Mean |  |  | 6.0233 |  | 5.44 | 5.6333 | 4.76 | 5.2827 | 4.9717 | 4.835 | 5.2617 | 5.4883 | 5.6217 |  | 5.6101 | 5.25 | 5.3207 |  | 5.36 | 10 | 5.9833 | 4.7083 | 4.5 | 5.6633 | 4.4833 | 5.45
| Median |  |  | 6.03 |  | 5.45 | 5.8 | 4.735 | 5.2897 | 4.98 | 4.865 | 5.26 | 5.51 | 5.565 |  | 5.6061 | 5.3 | 5.3146 |  | 5.4 | 10 | 5.95 | 4.715 | 4.45 | 5.65 | 4.45 | 5.5
| Std Dev. |  |  | 0.11553 |  | 0.1377 | 0.69186 | 0.22298 | 0.091798 | 0.15766 | 0.13142 | 0.095167 | 0.18552 | 0.27184 |  | 0.10967 | 0.16432 | 0.082264 |  | 0.088544 | 0.89443 | 0.1472 | 0.032506 | 0.26077 | 0.10463 | 0.098319 | 0.25884
| Rel.Std.Dev. |  |  | 1.918% |  | 2.531% | 12.28% | 4.684% | 1.738% | 3.171% | 2.718% | 1.809% | 3.38% | 4.836% |  | 1.955% | 3.13% | 1.546% |  | 1.652% | 8.944% | 2.46% | 0.6904% | 5.795% | 1.847% | 2.193% | 4.749%
| PDM3 |  |  | 13.8% |  | 2.779% | 6.431% | -10.07% | -0.1934% | -6.07% | -8.652% | -0.5906% | 3.692% | 6.211% |  | 5.993% | -0.811% | 0.5249% |  | 1.267% | 88.93% | 13.04% | -11.04% | -14.98% | 6.998% | -15.3% | 2.968%
| Z-Score (Absolute) |  | 5.35 | 1.42 | 5.35 | 0.3 | 0.67 | 1.01 | 0 | 0.6 | 0.86 | 0.04 | 0.4 | 0.65 | 5.35 | 0.63 | 0.06 | 0.07 | NA | 0.15 | NA | 1.35 | 1.1 | 1.51 | 0.73 | 1.54 | 0.32
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 556b (Certified Value 11 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 11.8 | 11.5 | 12 | 11 | 12 | 9.7923 | 10.8 | 10.5 | 10.4 | 11.2 | 11.7 | 10 | 12.3973 | 7.7807 | 11 | 11.722 | 10.39 | 11.3 | 10.7 | 13.5 | 8.1 | 11 | 11.2 | 11 | 9.5 |  |  | 1
| 1 | 2 | 11.2 | 10.9 | 11.9 | 11 | 11.5 | 9.4169 | 10.6 | 10 | 10.4 | 10.8 | 10.8 | 10 | 12.1485 | 7.6754 | 11 | 11.6502 | 10.59 | 10.4 | 10.2 | 14 | 7.8 | 11 | 11.5 | 11 | 9.4 |  |  | 1
| 1 | 3 | 12.2 | 10.8 | 11.9 | 11 | 12.1 | 9.2031 | 11.4 | 11.1 | 10.7 | 11.2 | 11.5 | 11 | 12.044 | 7.6545 | 11 | 11.7186 | 10.38 | 11 | 10.7 | 13.5 | 8.1 | 11 | 11.4 | 11 | 9.7 |  |  | 16
| 1 | 4 | 10.6 | 11.4 | 11.7 | 11 | 11.7 | 9.1296 | 11 | 10.6 | 12 | 11.1 | 11.2 | 10 | 12.3965 | 7.287 | 10 | 11.7465 | 10.33 | 11.3 | 10.1 | 14.2 | 8.1 | 11 | 11.1 | 11 | 9.8 |  |  | 11.03
| 1 | 5 | 11.1 | 11.1 | 11.9 | 11 | 11.8 | 8.4829 | 11.4 | 10.5 | 11.6 | 11.6 | 10.9 | 11 | 12.6839 | 7.8728 | 11 | 11.5554 | 10.29 | 11.1 | 10.6 | 13.8 | 7.9 | 10 | 11.1 | 11 | 10.1 |  |  | 128
| 1 | 6 | 11.6 | 11.1 | 11.5 | 11 | 11.7 | 8.7938 | 11.3 | 10.4 | 11.8 | 11.8 | 11.2 | 10 | 12.4648 | 7.3518 | 11 | 11.4456 | 10.42 | 10.9 | 10.1 | 14.2 | 8.2 | 11 | 11.4 | 11 | 9.6
| Mean |  | 11.4167 | 11.1333 | 11.8167 | 11 | 11.8 | 9.1364 | 11.0833 | 10.5167 | 11.15 | 11.2833 | 11.2167 | 10.3333 | 12.3558 | 7.6037 | 10.8333 | 11.6397 | 10.4 | 11 | 10.4 | 13.8667 | 8.0333 | 10.8333 | 11.2833 | 11 | 9.6833
| Median |  | 11.4 | 11.1 | 11.9 | 11 | 11.75 | 9.1663 | 11.15 | 10.5 | 11.15 | 11.2 | 11.2 | 10 | 12.3969 | 7.6649 | 11 | 11.6844 | 10.385 | 11.05 | 10.4 | 13.9 | 8.1 | 11 | 11.3 | 11 | 9.65
| Std Dev. |  | 0.56716 | 0.27325 | 0.18348 | 0 | 0.21909 | 0.45997 | 0.33714 | 0.35449 | 0.73144 | 0.36009 | 0.34303 | 0.5164 | 0.22937 | 0.23461 | 0.40825 | 0.1177 | 0.10392 | 0.33466 | 0.29665 | 0.32042 | 0.15055 | 0.40825 | 0.17224 | 0 | 0.24833
| Rel.Std.Dev. |  | 4.968% | 2.454% | 1.553% | 0% | 1.857% | 5.034% | 3.042% | 3.371% | 6.56% | 3.191% | 3.058% | 4.997% | 1.856% | 3.085% | 3.768% | 1.011% | 0.9993% | 3.042% | 2.852% | 2.311% | 1.874% | 3.768% | 1.527% | 0% | 2.564%
| PDM3 |  | 3.506% | 0.9372% | 7.132% | -0.2717% | 6.981% | -17.17% | 0.4839% | -4.654% | 1.088% | 2.297% | 1.693% | -6.316% | 12.02% | -31.06% | -1.783% | 5.528% | -5.711% | -0.2717% | -5.711% | 25.72% | -27.17% | -1.783% | 2.297% | -0.2717% | -12.21%
| Z-Score (Absolute) |  | 0.29 | 0.01 | 0.71 | NA | 0.69 | 2.11 | 0.06 | 0.66 | 0.01 | 0.15 | 0.08 | NA | 1.28 | 3.72 | NA | 0.52 | 0.78 | 0.15 | 0.78 | 2.86 | 3.27 | NA | 0.15 | NA | 1.53
Results from laboratories 5, 14, 17, 25 and 27 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 556b (Certified Value 11.2 ppm) | [28 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 25 | 26 | 27 | 28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 1
1
| 1 | 1 | 11.1 | 13.4 | 12.5 | 11.3 | 10.7 | 12 | 11 | 11 | 10 | 14 | < 10 | 11.5756 | 12 | 12.0538 | NA | 9.4 | 10.5 | 11.6 | 12 | 11 | [7 empty] | 1
| 1 | 2 | 11.7 | 13.2 | 12 | 11.8 | 10.1 | 12 | 11 | 12 | 10 | 12 | < 10 | 9.8819 | 12 | 12.0061 | NA | 9.3 | 9.19 | 12.1 | 10 | 11 | [7 empty] | 19
| 1 | 3 | 12.5 | 13.4 | 12.9 | 11.2 | 10.5 | 13 | 11 | 13 | 10 | 14 | < 10 | 11.0347 | 12 | 11.7452 | NA | 9.5 | 9.98 | 11.4 | 11 | 12 | [7 empty] | 16
| 1 | 4 | 11.2 | 13.2 | 12.4 | 11.7 | 9.8 | 13 | 11 | 16 | 11 | 12 | < 10 | 11.5364 | 12 | 11.6494 | NA | 9.2 | 9.33 | 11 | 13 | 12 | [7 empty] | 11.1614
| 1 | 5 | 11.1 | 13 | 12.1 | 11.8 | 9.9 | 13 | 11 | 15 | 14 | 14 | < 10 | 10.021 | 12 | 11.8919 | NA | 9.6 | 9.6 | 11.2 | 10 | 12 | [7 empty] | 129
| 1 | 6 | 10 | 13.4 | 12 | 11.5 | 10.3 | 13 | 11 | 14 | 13 | 14 | < 10 | 11.4302 | 12 | 11.3008 | NA | 8.9 | 8.86 | 11.2 | 12 | 11
| Mean |  | 11.2667 | 13.2667 | 12.3167 | 11.55 | 10.2167 | 12.6667 | 11 | 13.5 | 11.3333 | 13.3333 |  | 10.9133 | 12 | 11.7745 |  | 9.3167 | 9.5767 | 11.4167 | 11.3333 | 11.5
| Median |  | 11.15 | 13.3 | 12.25 | 11.6 | 10.2 | 13 | 11 | 13.5 | 10.5 | 14 |  | 11.2325 | 12 | 11.8185 |  | 9.35 | 9.465 | 11.3 | 11.5 | 11.5
| Std Dev. |  | 0.82138 | 0.1633 | 0.35449 | 0.25884 | 0.34881 | 0.5164 | 0 | 1.8708 | 1.7512 | 1.0328 |  | 0.77054 | 0 | 0.27795 |  | 0.24833 | 0.58987 | 0.392 | 1.2111 | 0.54772
| Rel.Std.Dev. |  | 7.29% | 1.231% | 2.878% | 2.241% | 3.414% | 4.077% | 0% | 13.86% | 15.45% | 7.746% |  | 7.061% | 0% | 2.361% |  | 2.665% | 6.159% | 3.434% | 10.69% | 4.763%
| PDM3 |  | 0.9427% | 18.86% | 10.35% | 3.481% | -8.465% | 13.49% | -1.446% | 20.95% | 1.54% | 19.46% |  | -2.223% | 7.513% | 5.493% |  | -16.53% | -14.2% | 2.287% | 1.54% | 3.033%
| Z-Score (Absolute) |  | 0 | 1.28 | 0.67 | 0.18 | 0.67 | NA | NA | NA | NA | NA | 4.02 | 0.23 | NA | 0.33 | NA | 1.25 | 1.09 | 0.1 | NA | NA
Results from laboratories 8, 9, 10, 11, 12, 17, 27 and 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 556b (Certified Value 8.58 ppm) | [28 empty] | Certified
Sm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab | [7 empty] | 1
| No | No | 02: 02; 03: 04; 04: 05; 05: 07; 06: 14; 08: 16; 09: 18; 10: 19; 11: 23; 12: 25; 13: 26 | [7 empty] | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*MS | [7 empty] | 2
3
| 1 | 1 | 02: 5.8; 03: 8.86; 04: 9.1; 05: 8.7688; 06: 7; 08: 8.0395; 09: 8.4724; 10: 8.08; 11: 7.56; 12: 9.41; 13: 9.24 | [7 empty] | 1
| 1 | 2 | 02: 5.8; 03: 8.98; 04: 8.8; 05: 8.7665; 06: 7; 08: 8.9323; 09: 8.647; 10: 8.35; 11: 7.47; 12: 8.74; 13: 8.81 | [7 empty] | 3
| 1 | 3 | 02: 5.9; 03: 8.74; 04: 8.8; 05: 8.5577; 06: 7; 08: 8.6024; 09: 8.5162; 10: 7.89; 11: 7.61; 12: 9.13; 13: 9.08 | [7 empty] | 16
| 1 | 4 | 02: 5.6; 03: 8.72; 04: 9.9; 05: 8.7171; 06: 6; 08: 8.8058; 09: 8.9158; 10: 7.49; 11: 7.63; 12: 8.76; 13: 9.21 | [7 empty] | 8.5793
| 1 | 5 | 02: 5.7; 03: 8.67; 04: 9.8; 05: 8.6795; 06: 7; 08: 8.3783; 09: 8.5085; 10: 7.65; 11: 7.59; 12: 8.7; 13: 9.22 | [7 empty] | 130
| 1 | 6 | 02: 5.6; 03: 8.8; 04: 9; 05: 8.7579; 06: 6; 08: 8.6707; 09: 8.4453; 10: 8.03; 11: 7.43; 12: 8.43; 13: 8.82
| Mean |  | 02: 5.7333; 03: 8.795; 04: 9.2333; 05: 8.7079; 06: 6.6667; 08: 8.5715; 09: 8.5842; 10: 7.915; 11: 7.5483; 12: 8.8617; 13: 9.0633
| Median |  | 02: 5.75; 03: 8.77; 04: 9.05; 05: 8.7375; 06: 7; 08: 8.6366; 09: 8.5123; 10: 7.96; 11: 7.575; 12: 8.75; 13: 9.145
| Std Dev. |  | 02: 0.12111; 03: 0.11203; 04: 0.49261; 05: 0.081346; 06: 0.5164; 08: 0.32133; 09: 0.17668; 10: 0.31021; 11: 0.080602; 12: 0.34937; 13: 0.20047
| Rel.Std.Dev. |  | 02: 2.112%; 03: 1.274%; 04: 5.335%; 05: 0.9342%; 06: 7.746%; 08: 3.749%; 09: 2.058%; 10: 3.919%; 11: 1.068%; 12: 3.942%; 13: 2.212%
| PDM3 |  | 02: -33.17%; 03: 2.514%; 04: 7.623%; 05: 1.499%; 06: -22.29%; 08: -0.09122%; 09: 0.05668%; 10: -7.743%; 11: -12.02%; 12: 3.291%; 13: 5.642%
| Z-Score (Absolute) |  | 02: 6.21; 03: 0.32; 04: 1.25; 05: 0.13; 06: NA; 08: 0.16; 09: 0.13; 10: 1.56; 11: 2.34; 12: 0.46; 13: 0.89
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 556b (Certified Value 3.7 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | < 5 | 4.2 | < 10 | 3.7 | 4 | 4.5 | 3.6825 | 3.2 | 3.3 | 3.5 | 3.6 | 3.8 | < 10 | < 1 | 3.8235 | 4 | 3.7093 | 3.57 | 3.9 | < 10 | 3.4 | 3.1 | 3.8 | 3.8 | 3.4 | 3.4 |  | 1
| 1 | 2 | < 5 | 4.3 | < 10 | 3.8 | 4 | 4.3 | 3.6543 | 3.1 | 3.1 | 3.6 | 3.6 | 3.7 | < 10 | < 1 | 3.6374 | 4.1 | 3.7321 | 3.67 | 3.8 | < 10 | 3.7 | 3.1 | 3.5 | 3.9 | 3.4 | 3.5 |  | 20
| 1 | 3 | < 5 | 4.4 | < 10 | 3.7 | 4 | 4.5 | 3.6588 | 3.3 | 3.4 | 3.6 | 3.8 | 3.8 | < 10 | < 1 | 3.6689 | 4 | 3.6856 | 3.57 | 4 | < 10 | 4 | 3 | 3.5 | 3.8 | 3.4 | 3.5 |  | 16
| 1 | 4 | < 5 | 4.1 | < 10 | 3.8 | 4 | 4.4 | 3.7191 | 3.2 | 3.3 | 3.6 | 3.6 | 3.7 | < 10 | < 1 | 3.9018 | 3.9 | 3.7916 | 3.47 | 4 | < 10 | 3.8 | 3.1 | 3.5 | 3.9 | 3.3 | 3.6 |  | 3.7046
| 1 | 5 | < 5 | 4.1 | < 10 | 3.9 | 4 | 4.3 | 3.7342 | 3.3 | 3.2 | 3.5 | 3.9 | 3.7 | < 10 | < 1 | 3.6359 | 4 | 3.7392 | 2.78 | 4 | < 10 | 3.7 | 3 | 3.7 | 3.9 | 3.4 | 3.6 |  | 131
| 1 | 6 | < 5 | 3.9 | < 10 | 3.8 | 4 | 4.2 | 3.7128 | 3.2 | 3.2 | 3.4 | 3.9 | 3.7 | < 10 | < 1 | 3.8799 | 3.9 | 3.6621 | 3.47 | 3.8 | < 10 | 3.7 | 2.9 | 3.3 | 3.9 | 3.4 | 3.2
| Mean |  |  | 4.1667 |  | 3.7833 | 4 | 4.3667 | 3.6936 | 3.2167 | 3.25 | 3.5333 | 3.7333 | 3.7333 |  |  | 3.7579 | 3.9833 | 3.72 | 3.4217 | 3.9167 |  | 3.7167 | 3.0333 | 3.55 | 3.8667 | 3.3833 | 3.4667
| Median |  |  | 4.15 |  | 3.8 | 4 | 4.35 | 3.6977 | 3.2 | 3.25 | 3.55 | 3.7 | 3.7 |  |  | 3.7462 | 4 | 3.7207 | 3.52 | 3.95 |  | 3.7 | 3.05 | 3.5 | 3.9 | 3.4 | 3.5
| Std Dev. |  |  | 0.17512 |  | 0.075277 | 0 | 0.12111 | 0.033288 | 0.075277 | 0.10488 | 0.08165 | 0.15055 | 0.05164 |  |  | 0.12429 | 0.075277 | 0.045378 | 0.32314 | 0.098319 |  | 0.19408 | 0.08165 | 0.17607 | 0.05164 | 0.040825 | 0.15055
| Rel.Std.Dev. |  |  | 4.203% |  | 1.99% | 0% | 2.773% | 0.9012% | 2.34% | 3.227% | 2.311% | 4.033% | 1.383% |  |  | 3.307% | 1.89% | 1.22% | 9.444% | 2.51% |  | 5.222% | 2.692% | 4.96% | 1.336% | 1.207% | 4.343%
| PDM3 |  |  | 12.47% |  | 2.124% | 7.973% | 17.87% | -0.2974% | -13.17% | -12.27% | -4.624% | 0.7745% | 0.7745% |  |  | 1.438% | 7.523% | 0.4138% | -7.638% | 5.723% |  | 0.3246% | -18.12% | -4.174% | 4.374% | -8.673% | -6.424%
| Z-Score (Absolute) |  | 3.12 | 1.14 | 3.28 | 0.16 | NA | 1.66 | 0.07 | 1.29 | 1.2 | 0.48 | 0.03 | 0.03 | 3.28 | 8.24 | 0.1 | 0.67 | 0 | 0.76 | 0.5 | 3.28 | 0.01 | 1.76 | 0.44 | 0.38 | 0.86 | 0.65
Results from laboratory 5 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 556b (Certified Value 68 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 64.49 | 71 | 71 | 73.82 | 68.7 | 66.5 | 66.1848 | 67.5 | 65.4 | 67.7 | 67.5 | 67.7 | 58.02 | 63 | 53.4122 | < 100 | 70.6 | 66.4822 | 75 | 65.6 | 64.4 | 70 | 50.9 | 54.7 | 69.61 | 66 | 67.3 | 1
| 1 | 2 | 66.34 | 69 | 73.4 | 72.34 | 70.1 | 64.8 | 63.995 | 65.3 | 63.7 | 67.1 | 65.3 | 67.1 | 57.12 | 63 | 55.4357 | < 100 | 69.8 | 66.7229 | 75 | 62.3 | 62.6 | 68 | 50.9 | 59.9 | 69.06 | 65 | 67.3 | 32
| 1 | 3 | 68.68 | 73 | 71.7 | 72.44 | 70.7 | 66.5 | 65.3564 | 69.4 | 66.9 | 68.3 | 67.8 | 67.6 | 57.86 | 64 | 67.7173 | < 100 | 69.1 | 66.8094 | 69 | 64.6 | 63.7 | 71 | 51.1 | 57.8 | 68.17 | 65 | 68.1 | 16
| 1 | 4 | 65.08 | 66 | 74.7 | 71.39 | 67.9 | 64.4 | 67.8814 | 68.4 | 65.6 | 69.3 | 68.1 | 66.8 | 57.74 | 61 | 67.8643 | < 100 | 70.9 | 66.856 | 76 | 66 | 64.4 | 72 | 50.4 | 55 | 69.43 | 64 | 66.9 | 67.5203
| 1 | 5 | 64.77 | 69 | 73.8 | 71.45 | 76 | 64 | 66.1885 | 70.6 | 65.8 | 68 | 69.3 | 67 | 57.49 | 62 | 68.5876 | < 100 | 69.5 | 66.2461 | 76 | 64.4 | 62.9 | 69 | 52.2 | 53.9 | 68.49 | 64 | 67.9 | 132
| 1 | 6 | 65.33 | 71 | 74 | 71.54 | 73.4 | 64.8 | 65.2052 | 69.3 | 64.4 | 67.7 | 71.4 | 66.2 | 57.73 | 63 | 68.6696 | < 100 | 68.5 | 64.6514 | 75 | 64.3 | 63.3 | 72 | 48.8 | 54.9 | 67.89 | 63 | 67
| Mean |  | 65.7817 | 69.8333 | 73.1 | 72.1633 | 71.1333 | 65.1667 | 65.8019 | 68.4167 | 65.3 | 68.0167 | 68.2333 | 67.0667 | 57.66 | 62.6667 | 63.6144 |  | 69.7333 | 66.2947 | 74.3333 | 64.5333 | 63.55 | 70.3333 | 50.7167 | 56.0333 | 68.775 | 64.5 | 67.4167
| Median |  | 65.205 | 70 | 73.6 | 71.94 | 70.4 | 64.8 | 65.7706 | 68.85 | 65.5 | 67.85 | 67.95 | 67.05 | 57.735 | 63 | 67.7908 |  | 69.65 | 66.6026 | 75 | 64.5 | 63.5 | 70.5 | 50.9 | 54.95 | 68.775 | 64.5 | 67.3
| Std Dev. |  | 1.5557 | 2.4014 | 1.4367 | 0.93249 | 3.0481 | 1.0746 | 1.2992 | 1.8476 | 1.1207 | 0.74409 | 2.0255 | 0.55015 | 0.31667 | 1.0328 | 7.1576 |  | 0.9048 | 0.83681 | 2.6583 | 1.2925 | 0.75565 | 1.633 | 1.1125 | 2.3132 | 0.69828 | 1.0488 | 0.48339
| Rel.Std.Dev. |  | 2.365% | 3.439% | 1.965% | 1.292% | 4.285% | 1.649% | 1.974% | 2.701% | 1.716% | 1.094% | 2.968% | 0.8203% | 0.5492% | 1.648% | 11.25% |  | 1.298% | 1.262% | 3.576% | 2.003% | 1.189% | 2.322% | 2.194% | 4.128% | 1.015% | 1.626% | 0.717%
| PDM3 |  | -2.575% | 3.426% | 8.264% | 6.877% | 5.351% | -3.486% | -2.545% | 1.328% | -3.288% | 0.7351% | 1.056% | -0.6718% | -14.6% | -7.188% | -5.785% |  | 3.278% | -1.815% | 10.09% | -4.424% | -5.88% | 4.166% | -24.89% | -17.01% | 1.858% | -4.473% | -0.1535%
| Z-Score (Absolute) |  | 0.13 | 0.89 | 1.71 | 1.48 | 1.22 | 0.28 | 0.12 | 0.53 | 0.25 | 0.43 | 0.49 | 0.19 | 2.17 | 0.91 | 0.67 | 4.1 | 0.87 | 0 | 2.02 | 0.44 | 0.69 | 1.02 | 3.92 | 2.58 | 0.62 | 0.45 | 0.28
| Analytical results for Ta in OREAS 556b (Certified Value 1.04 ppm) | [28 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.3 | < 5 | 1.03 | 1.4 | 0.49 | 1.6097 | 1.03 | 0.93 | 1.07 | 0.98 | 1.2 | < 1 | 1.0363 | 0.5 | 0.87906 | NA | 1.6 | 1.1 | 0.72 | 1 | 0.98 | 0.8 | 0.95 |  |  |  |  | 1
| 1 | 2 | 1.4 | < 5 | 1.04 | 1.6 | 0.56 | 1.5861 | 0.97 | 0.93 | 1.11 | 0.93 | 1.11 | < 1 | 1.0542 | 0.41 | 0.84191 | NA | 1.4 | 1.1 | 0.64 | 1.09 | 0.96 | 0.8 | 0.95 |  |  |  |  | 5
| 1 | 3 | 1.3 | < 5 | 0.85 | 1.6 | 0.52 | 1.5931 | 1.05 | 1.01 | 1.19 | 0.99 | 1.1 | 11 | 0.98527 | 0.48 | 0.88906 | NA | 1.4 | 1.2 | 0.67 | 1.3 | 0.91 | 0.8 | 0.97 |  |  |  |  | 16
| 1 | 4 | 1.3 | < 5 | 0.97 | 1.6 | 0.46 | 1.6225 | 1.04 | 0.96 | 1.1 | 0.97 | 1.18 | < 1 | 1.1063 | 0.53 | 0.88866 | NA | 1.4 | 1.1 | 0.63 | 1.06 | 0.89 | 0.8 | 1.03 |  |  |  |  | 1.0427
| 1 | 5 | 1.2 | < 5 | 0.9 | 1.7 | 0.49 | 1.6445 | 1.05 | 1 | 1.09 | 0.94 | 1.14 | < 1 | 1.0214 | 0.45 | 0.88325 | NA | 1.4 | 1.2 | 0.6 | 1.14 | 0.95 | 0.8 | 0.98 |  |  |  |  | 133
| 1 | 6 | 1.2 | < 5 | 0.94 | 1.6 | 0.44 | 1.6335 | 1.04 | 0.92 | 1.07 | 1.06 | 1.15 | 13 | 1.1035 | 0.5 | 0.83425 | NA | 1.3 | 1 | 0.64 | 0.93 | 0.96 | 0.7 | 0.95
| Mean |  | 1.2833 |  | 0.955 | 1.5833 | 0.49333 | 1.6149 | 1.03 | 0.95833 | 1.105 | 0.97833 | 1.1467 | 12 | 1.0512 | 0.47833 | 0.86936 |  | 1.4167 | 1.1167 | 0.65 | 1.0867 | 0.94167 | 0.78333 | 0.97167
| Median |  | 1.3 |  | 0.955 | 1.6 | 0.49 | 1.6161 | 1.04 | 0.945 | 1.095 | 0.975 | 1.145 | 12 | 1.0452 | 0.49 | 0.88115 |  | 1.4 | 1.1 | 0.64 | 1.075 | 0.955 | 0.8 | 0.96
| Std Dev. |  | 0.075277 |  | 0.073959 | 0.098319 | 0.04274 | 0.02285 | 0.030332 | 0.038687 | 0.044609 | 0.046224 | 0.038816 | 1.4142 | 0.047433 | 0.042622 | 0.02463 |  | 0.098319 | 0.075277 | 0.040988 | 0.12738 | 0.034303 | 0.040825 | 0.031252
| Rel.Std.Dev. |  | 5.866% |  | 7.744% | 6.21% | 8.663% | 1.415% | 2.945% | 4.037% | 4.037% | 4.725% | 3.385% | 11.79% | 4.512% | 8.911% | 2.833% |  | 6.94% | 6.741% | 6.306% | 11.72% | 3.643% | 5.212% | 3.216%
| PDM3 |  | 23.08% |  | -8.411% | 51.85% | -52.69% | 54.88% | -1.218% | -8.091% | 5.975% | -6.173% | 9.971% | 1051% | 0.8119% | -54.13% | -16.62% |  | 35.86% | 7.094% | -37.66% | 4.216% | -9.69% | -24.87% | -6.813%
| Z-Score (Absolute) |  | 1.18 | 7.1 | 0.42 | 2.64 | 2.66 | 2.79 | 0.05 | 0.4 | 0.31 | 0.3 | 0.52 | 16.01 | 0.05 | 2.73 | 0.83 | NA | 1.83 | 0.37 | 1.9 | 0.22 | 0.48 | 1.25 | 0.34
| Analytical results for Tb in OREAS 556b (Certified Value 0.78 ppm) | [28 empty] | Certified
Tb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 14: Lab |  |  |  |  | 1
| No | No | 02: 02; 03: 04; 04: 05; 05: 07; 06: 14; 07: 16; 08: 18; 09: 19; 10: 20; 11: 23; 12: 25; 14: 26 |  |  |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*OES/MS; 14: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 02: 0.5; 03: 0.73; 04: 1; 05: 1.033; 06: 2; 07: 1.1715; 08: 0.83433; 09: NA; 10: 0.8; 11: 0.56; 12: 0.81; 14: 0.81 |  |  |  |  | 1
| 1 | 2 | 02: 0.5; 03: 0.76; 04: 1; 05: 1.0175; 06: 1; 07: 1.2889; 08: 0.83955; 09: NA; 10: 0.8; 11: 0.55; 12: 0.73; 14: 0.76 |  |  |  |  | 6
| 1 | 3 | 02: 0.5; 03: 0.72; 04: 0.9; 05: 0.98457; 06: 2; 07: 1.1702; 08: 0.8217; 09: NA; 10: 0.9; 11: 0.56; 12: 0.76; 14: 0.79 |  |  |  |  | 16
| 1 | 4 | 02: 0.5; 03: 0.76; 04: 1; 05: 1.0248; 06: 1; 07: 1.1278; 08: 0.86088; 09: NA; 10: 0.9; 11: 0.57; 12: 0.73; 14: 0.78 |  |  |  |  | 0.77875
| 1 | 5 | 02: 0.5; 03: 0.74; 04: 1; 05: 1.0059; 06: 3; 07: 1.0873; 08: 0.81648; 09: NA; 10: 0.9; 11: 0.55; 12: 0.75; 14: 0.8 |  |  |  |  | 134
| 1 | 6 | 02: 0.5; 03: 0.73; 04: 1; 05: 1.0566; 06: 2; 07: 1.2027; 08: 0.79725; 09: NA; 10: 0.8; 11: 0.55; 12: 0.69; 14: 0.77
| Mean |  | 02: 0.5; 03: 0.74; 04: 0.98333; 05: 1.0204; 06: 1.8333; 07: 1.1747; 08: 0.82837; 10: 0.85; 11: 0.55667; 12: 0.745; 14: 0.785
| Median |  | 02: 0.5; 03: 0.735; 04: 1; 05: 1.0212; 06: 2; 07: 1.1709; 08: 0.82802; 10: 0.85; 11: 0.555; 12: 0.74; 14: 0.785
| Std Dev. |  | 02: 0; 03: 0.016733; 04: 0.040825; 05: 0.024461; 06: 0.75277; 07: 0.068832; 08: 0.021779; 10: 0.054772; 11: 0.008165; 12: 0.039875; 14: 0.018708
| Rel.Std.Dev. |  | 02: 0%; 03: 2.261%; 04: 4.152%; 05: 2.397%; 06: 41.06%; 07: 5.859%; 08: 2.629%; 10: 6.444%; 11: 1.467%; 12: 5.352%; 14: 2.383%
| PDM3 |  | 02: -35.79%; 03: -4.976%; 04: 26.27%; 05: 31.03%; 06: 135.4%; 07: 50.85%; 08: 6.371%; 10: 9.149%; 11: -28.52%; 12: -4.334%; 14: 0.8022%
| Z-Score (Absolute) |  | 02: 1.7; 03: 0.37; 04: 0.98; 05: 1.18; 06: NA; 07: 2.04; 08: 0.12; 09: NA; 10: 0.24; 11: 1.39; 12: 0.34; 14: 0.12
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Te in OREAS 556b (Certified Value 0.089 ppm) | [28 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 22 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.09 | < 5 | < 0.2 | 0.2 | 0.02 | 0.08864 | 0.08 | 0.08 | 0.1 | 0.09 | 0.13 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | 0.05 | < 0.2 | 0.11 | [7 empty] | 1
| 1 | 2 | 0.09 | < 5 | < 0.2 | < 0.1 | 0.11 | 0.08574 | 0.07 | 0.07 | 0.12 | 0.11 | 0.1 | < 10 | 0.17237 | < 0.1 | < 5 | NA | < 0.1 | 0.05 | < 0.2 | 0.12 | [7 empty] | 7
| 1 | 3 | 0.09 | < 5 | < 0.2 | 0.1 | 0.09 | 0.08806 | 0.09 | 0.08 | 0.11 | 0.07 | 0.11 | < 10 | < 0.1 | < 0.1 | < 5 | NA | 0.1 | 0.05 | < 0.2 | 0.1 | [7 empty] | 16
| 1 | 4 | 0.09 | < 5 | < 0.2 | < 0.1 | 0.08 | 0.09089 | 0.08 | 0.08 | 0.07 | 0.07 | 0.15 | < 10 | 0.17891 | < 0.1 | < 5 | NA | 0.1 | 0.06 | < 0.2 | 0.11 | [7 empty] | 0.089187
| 1 | 5 | 0.12 | < 5 | < 0.2 | < 0.1 | 0.17 | 0.08883 | 0.11 | 0.08 | 0.12 | 0.08 | 0.11 | < 10 | < 0.1 | < 0.1 | < 5 | NA | < 0.1 | 0.06 | < 0.2 | 0.09 | [7 empty] | 135
| 1 | 6 | 0.09 | < 5 | < 0.2 | < 0.1 | 0.11 | 0.08504 | 0.08 | 0.08 | 0.11 | 0.08 | 0.1 | < 10 | < 0.1 | < 0.1 | < 5 | NA | 0.2 | 0.05 | < 0.2 | 0.12
| Mean |  | 0.095 |  |  | 0.15 | 0.096667 | 0.087867 | 0.085 | 0.078333 | 0.105 | 0.083333 | 0.11667 |  | 0.17564 |  |  |  | 0.13333 | 0.053333 |  | 0.10833
| Median |  | 0.09 |  |  | 0.15 | 0.1 | 0.08835 | 0.08 | 0.08 | 0.11 | 0.08 | 0.11 |  | 0.17564 |  |  |  | 0.1 | 0.05 |  | 0.11
| Std Dev. |  | 0.012247 |  |  | 0.070711 | 0.048854 | 0.0021548 | 0.013784 | 0.0040825 | 0.018708 | 0.015055 | 0.019664 |  | 0.004624 |  |  |  | 0.057735 | 0.005164 |  | 0.01169
| Rel.Std.Dev. |  | 12.89% |  |  | 47.14% | 50.54% | 2.452% | 16.22% | 5.212% | 17.82% | 18.07% | 16.85% |  | 2.633% |  |  |  | 43.3% | 9.682% |  | 10.79%
| PDM3 |  | 6.518% |  |  | 68.19% | 8.387% | -1.48% | -4.694% | -12.17% | 17.73% | -6.563% | 30.81% |  | 96.93% |  |  |  | 49.5% | -40.2% |  | 21.47%
| Z-Score (Absolute) |  | 0.04 | 129.69 | 0.22 | 0.67 | 0.04 | 0.43 | 0.58 | 0.94 | 0.49 | 0.67 | 1.12 | 264.55 | 0.21 | 2.47 | 129.69 | NA | NA | 2.29 | 0.22 | 0.67
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 556b (Certified Value 15.4 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 14.1 | 15.38 | 15.6 | 15.7 | 15.6998 | 16.6 | 15.15 | 15.15 | 16.35 | 15.6 | 16 | 15.2375 | 16.1542 | 15.7 | 15.3272 | 15.52 | 14.27 | 15 | 13.93 | 15.3 | 16.36 | 16.1 | 13.9 |  |  |  |  | 1
| 1 | 2 | 13.4 | 15.33 | 16 | 15.5 | 15.5032 | 16.7 | 14.1 | 15.3 | 15.7 | 13.95 | 17 | 15.715 | 16.6402 | 15.5 | 15.4384 | 15.95 | 15.78 | 15.9 | 13.95 | 18.7 | 16.16 | 15.7 | 14.3 |  |  |  |  | 8
| 1 | 3 | 14.2 | 15.19 | 16 | 15.7 | 15.2047 | 16.35 | 15.4 | 15.3 | 16.35 | 14.8 | 16 | 15.3943 | 16.158 | 15.7 | 15.7968 | 14.87 | 14.55 | 15.8 | 14.6 | 17.8 | 16.27 | 16.2 | 14.9 |  |  |  |  | 16
| 1 | 4 | 13.6 | 14.97 | 16.4 | 15.6 | 15.6563 | 16.25 | 15.2 | 15.7 | 15.95 | 14.95 | 16 | 15.4707 | 16.5503 | 15.6 | 15.5164 | 13.98 | 15.09 | 16.5 | 14.84 | 17.1 | 16.4 | 15.5 | 15 |  |  |  |  | 15.4439
| 1 | 5 | 13.3 | 14.71 | 17 | 15.2 | 15.5978 | 16.85 | 14.85 | 15.4 | 16.5 | 14.7 | 16 | 15.7235 | 16.4144 | 15.4 | 15.4688 | 14.23 | 14.82 | 16.1 | 14.29 | 16.8 | 16.46 | 15.5 | 15.1 |  |  |  |  | 136
| 1 | 6 | 13.1 | 14.76 | 16.9 | 15.3 | 15.5195 | 16.65 | 14.7 | 15.05 | 16.85 | 14.45 | 17 | 15.8235 | 16.1122 | 15.4 | 14.7636 | 14.81 | 14.6 | 15.9 | 14.16 | 15.6 | 16.26 | 15.5 | 13.7
| Mean |  | 13.6167 | 15.0567 | 16.3167 | 15.5 | 15.5302 | 16.5667 | 14.9 | 15.3167 | 16.2833 | 14.7417 | 16.3333 | 15.5608 | 16.3382 | 15.55 | 15.3852 | 14.8933 | 14.8517 | 15.8667 | 14.295 | 16.8833 | 16.3183 | 15.75 | 14.4833
| Median |  | 13.5 | 15.08 | 16.2 | 15.55 | 15.5586 | 16.625 | 15 | 15.3 | 16.35 | 14.75 | 16 | 15.5929 | 16.2862 | 15.55 | 15.4536 | 14.84 | 14.71 | 15.9 | 14.225 | 16.95 | 16.315 | 15.6 | 14.6
| Std Dev. |  | 0.4446 | 0.28731 | 0.55287 | 0.20976 | 0.17668 | 0.22509 | 0.46583 | 0.22509 | 0.40702 | 0.54719 | 0.5164 | 0.22786 | 0.22777 | 0.13784 | 0.3424 | 0.74712 | 0.53147 | 0.49261 | 0.36347 | 1.2922 | 0.10889 | 0.32094 | 0.60139
| Rel.Std.Dev. |  | 3.265% | 1.908% | 3.388% | 1.353% | 1.138% | 1.359% | 3.126% | 1.47% | 2.5% | 3.712% | 3.162% | 1.464% | 1.394% | 0.8864% | 2.226% | 5.016% | 3.578% | 3.105% | 2.543% | 7.653% | 0.6673% | 2.038% | 4.152%
| PDM3 |  | -11.83% | -2.507% | 5.651% | 0.3634% | 0.559% | 7.27% | -3.522% | -0.8237% | 5.436% | -4.547% | 5.759% | 0.7569% | 5.791% | 0.6871% | -0.3798% | -3.565% | -3.835% | 2.738% | -7.439% | 9.321% | 5.662% | 1.982% | -6.22%
| Z-Score (Absolute) |  | 1.99 | 0.48 | 0.84 | 0.02 | 0.02 | 1.1 | 0.65 | 0.21 | 0.81 | 0.81 | NA | 0.05 | 0.86 | 0.04 | 0.14 | 0.65 | 0.7 | 0.37 | 1.28 | 1.44 | 0.84 | 0.25 | 1.08
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 556b (Certified Value 0.235 wt.%) | [28 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.32 | 0.262 | 0.1 | 0.2208 | 0.197 | 0.13 | 0.40883 | 0.254 | 0.226 | 0.257 | 0.233 | 0.283 | 0.372 | 0.3196 | 0.19253 | 0.126 | 0.24083 | NA | 0.27 | 0.11 | 0.266 | 0.092 | 0.24 | 0.1986 | 0.18 | 0.27 |  | 1
| 1 | 2 | 0.35 | 0.265 | 0.11 | 0.2212 | 0.145 | 0.13 | 0.40735 | 0.244 | 0.238 | 0.251 | 0.22 | 0.285 | 0.372 | 0.29803 | 0.19815 | 0.127 | 0.2359 | NA | 0.25 | 0.12 | 0.263 | 0.116 | 0.236 | 0.1987 | 0.18 | 0.27 |  | 25
| 1 | 3 | 0.33 | 0.268 | 0.1 | 0.2155 | 0.202 | 0.13 | 0.40033 | 0.254 | 0.238 | 0.271 | 0.235 | 0.283 | 0.378 | 0.31867 | 0.19749 | 0.139 | 0.23785 | NA | 0.27 | 0.12 | 0.274 | 0.092 | 0.27 | 0.195 | 0.181 | 0.27 |  | 16
| 1 | 4 | 0.35 | 0.236 | 0.1 | 0.2154 | 0.206 | 0.14 | 0.40779 | 0.255 | 0.243 | 0.255 | 0.234 | 0.289 | 0.372 | 0.2893 | 0.19545 | 0.137 | 0.23676 | NA | 0.27 | 0.11 | 0.26 | 0.106 | 0.235 | 0.205 | 0.181 | 0.28 |  | 0.23512
| 1 | 5 | 0.34 | 0.24 | 0.1 | 0.1982 | 0.202 | 0.14 | 0.41001 | 0.262 | 0.247 | 0.262 | 0.222 | 0.288 | 0.372 | 0.3196 | 0.19477 | 0.129 | 0.23208 | NA | 0.27 | 0.12 | 0.268 | 0.102 | 0.255 | 0.1944 | 0.177 | 0.27 |  | 137
| 1 | 6 | 0.32 | 0.238 | 0.1 | 0.2068 | 0.208 | 0.14 | 0.40957 | 0.257 | 0.224 | 0.267 | 0.246 | 0.289 | 0.372 | 0.29831 | 0.19793 | 0.131 | 0.23406 | NA | 0.27 | 0.11 | 0.254 | 0.09 | 0.228 | 0.2001 | 0.182 | 0.27
| Mean |  | 0.335 | 0.2515 | 0.10167 | 0.21298 | 0.19333 | 0.135 | 0.40731 | 0.25433 | 0.236 | 0.2605 | 0.23167 | 0.28617 | 0.373 | 0.30725 | 0.19605 | 0.1315 | 0.23625 |  | 0.26667 | 0.115 | 0.26417 | 0.099667 | 0.244 | 0.19863 | 0.18017 | 0.27167
| Median |  | 0.335 | 0.251 | 0.1 | 0.21545 | 0.202 | 0.135 | 0.40831 | 0.2545 | 0.238 | 0.2595 | 0.2335 | 0.2865 | 0.372 | 0.30849 | 0.19647 | 0.13 | 0.23633 |  | 0.27 | 0.115 | 0.2645 | 0.097 | 0.238 | 0.19865 | 0.1805 | 0.27
| Std Dev. |  | 0.013784 | 0.014963 | 0.0040825 | 0.008917 | 0.023981 | 0.0054772 | 0.0035712 | 0.0058878 | 0.0091869 | 0.0075829 | 0.0095219 | 0.0028577 | 0.0024495 | 0.013582 | 0.0022094 | 0.0053572 | 0.0030385 |  | 0.008165 | 0.0054772 | 0.0068823 | 0.010231 | 0.015582 | 0.0038422 | 0.0017224 | 0.0040825
| Rel.Std.Dev. |  | 4.115% | 5.95% | 4.016% | 4.187% | 12.4% | 4.057% | 0.8768% | 2.315% | 3.893% | 2.911% | 4.11% | 0.9986% | 0.6567% | 4.42% | 1.127% | 4.074% | 1.286% |  | 3.062% | 4.763% | 2.605% | 10.26% | 6.386% | 1.934% | 0.956% | 1.503%
| PDM3 |  | 42.48% | 6.964% | -56.76% | -9.417% | -17.77% | -42.58% | 73.23% | 8.169% | 0.3722% | 10.79% | -1.471% | 21.71% | 58.64% | 30.68% | -16.62% | -44.07% | 0.4767% |  | 13.41% | -51.09% | 12.35% | -57.61% | 3.775% | -15.52% | -23.37% | 15.54%
| Z-Score (Absolute) |  | 1.66 | 0.26 | 2.26 | 0.39 | 0.72 | 1.7 | 2.87 | 0.3 | 0 | 0.41 | 0.08 | 0.84 | 2.29 | 1.19 | 0.67 | 1.76 | 0 | NA | 0.51 | 2.03 | 0.47 | 2.29 | 0.13 | 0.63 | 0.94 | 0.59
| Analytical results for Tl in OREAS 556b (Certified Value 1.06 ppm) | [28 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1 | 1.16 | 1.04 | 1.11 | 1.0974 | 0.96 | 0.94 | 0.93 | 1.07 | 1.08 | 1.0637 | 1.16 | 1.0566 | 0.255 | 1.12 | 2 | 1 | 0.96 | 1.13 | 1.15 | 0.96 | [6 empty] | 1
| 1 | 2 | 0.99 | 1.15 | 1.07 | 1.08 | 1.0854 | 0.96 | 0.89 | 0.95 | 1.06 | 0.99 | 1.0443 | 1.15 | 1.0728 | 0.263 | 1.05 | < 2 | 1.1 | 0.95 | 1.33 | 1.13 | 0.96 | [6 empty] | 26
| 1 | 3 | 0.96 | 1.12 | 1.08 | 1.1 | 1.1039 | 1.01 | 0.96 | 0.99 | 1.13 | 1.06 | 1.0821 | 1.14 | 1.0457 | 0.237 | 1.08 | < 2 | 1 | 0.95 | 1.28 | 1.15 | 1 | [6 empty] | 16
| 1 | 4 | 0.95 | 1.13 | 1.09 | 1.1 | 1.0848 | 0.96 | 0.97 | 0.99 | 1.08 | 1.03 | 1.0694 | 1.15 | 1.0601 | 0.225 | 1.11 | 2 | 1.1 | 0.96 | 1.24 | 1.15 | 1.02 | [6 empty] | 1.0561
| 1 | 5 | 0.96 | 1.13 | 1.13 | 1.07 | 1.0967 | 1.01 | 0.93 | 0.98 | 1.09 | 0.99 | 1.1085 | 1.13 | 1.0763 | 0.236 | 1.08 | < 2 | 1 | 0.97 | 1.21 | 1.16 | 1.01 | [6 empty] | 138
| 1 | 6 | 0.91 | 1.11 | 1.08 | 1.08 | 1.0858 | 0.99 | 0.94 | 0.94 | 1.16 | 1.02 | 1.1243 | 1.11 | 1.03 | 0.239 | 1.04 | 2 | 1.1 | 0.94 | 1.13 | 1.14 | 0.98
| Mean |  | 0.96167 | 1.1333 | 1.0817 | 1.09 | 1.0923 | 0.98167 | 0.93833 | 0.96333 | 1.0983 | 1.0283 | 1.082 | 1.14 | 1.0569 | 0.2425 | 1.08 | 2 | 1.05 | 0.955 | 1.22 | 1.1467 | 0.98833
| Median |  | 0.96 | 1.13 | 1.08 | 1.09 | 1.0913 | 0.975 | 0.94 | 0.965 | 1.085 | 1.025 | 1.0757 | 1.145 | 1.0583 | 0.238 | 1.08 | 2 | 1.05 | 0.955 | 1.225 | 1.15 | 0.99
| Std Dev. |  | 0.031885 | 0.018619 | 0.029269 | 0.015492 | 0.0080682 | 0.024833 | 0.027869 | 0.026583 | 0.038687 | 0.03656 | 0.02968 | 0.017889 | 0.017267 | 0.01391 | 0.031623 | 0 | 0.054772 | 0.010488 | 0.080498 | 0.010328 | 0.025626
| Rel.Std.Dev. |  | 3.316% | 1.643% | 2.706% | 1.421% | 0.7386% | 2.53% | 2.97% | 2.76% | 3.522% | 3.555% | 2.743% | 1.569% | 1.634% | 5.736% | 2.928% | 0% | 5.216% | 1.098% | 6.598% | 0.9007% | 2.593%
| PDM3 |  | -8.942% | 7.313% | 2.421% | 3.21% | 3.429% | -7.048% | -11.15% | -8.784% | 3.999% | -2.629% | 2.457% | 7.944% | 0.07614% | -77.04% | 2.263% | 89.38% | -0.5778% | -9.573% | 15.52% | 8.575% | -6.417%
| Z-Score (Absolute) |  | 1.06 | 0.64 | 0.13 | 0.21 | 0.24 | 0.86 | 1.29 | 1.04 | 0.3 | 0.4 | 0.13 | 0.71 | 0.11 | 8.16 | 0.11 | NA | 0.18 | 1.12 | 1.5 | 0.77 | 0.79
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Tm in OREAS 556b (Certified Value 0.39 ppm) | [28 empty] | Certified
Tm | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | [6 empty] | 1
| No | No | 02: 02; 04: 04; 05: 05; 06: 07; 07: 16; 08: 18; 09: 19; 10: 23; 11: 25; 12: 26 | [6 empty] | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*MS; 11: 4A*OES/MS; 12: 4A*MS | [6 empty] | 2
2
| 1 | 1 | 02: 0.3; 04: 0.37; 05: 0.5; 06: 0.49938; 07: 0.42493; 08: 0.42105; 09: 0.32; 10: 0.31; 11: 0.44; 12: 0.43 | [6 empty] | 1
| 1 | 2 | 02: 0.3; 04: 0.4; 05: 0.5; 06: 0.50031; 07: 0.43976; 08: 0.41878; 09: 0.34; 10: 0.31; 11: 0.37; 12: 0.43 | [6 empty] | 10
| 1 | 3 | 02: 0.3; 04: 0.39; 05: 0.4; 06: 0.50753; 07: 0.41276; 08: 0.41287; 09: 0.31; 10: 0.3; 11: 0.41; 12: 0.43 | [6 empty] | 16
| 1 | 4 | 02: 0.2; 04: 0.37; 05: 0.5; 06: 0.50281; 07: 0.40911; 08: 0.42333; 09: 0.3; 10: 0.32; 11: 0.38; 12: 0.44 | [6 empty] | 0.39318
| 1 | 5 | 02: 0.2; 04: 0.39; 05: 0.5; 06: 0.50355; 07: 0.41714; 08: 0.4177; 09: 0.3; 10: 0.3; 11: 0.38; 12: 0.43 | [6 empty] | 139
| 1 | 6 | 02: 0.3; 04: 0.37; 05: 0.5; 06: 0.509; 07: 0.43831; 08: 0.39571; 09: 0.31; 10: 0.3; 11: 0.36; 12: 0.43
| Mean |  | 02: 0.26667; 04: 0.38167; 05: 0.48333; 06: 0.50376; 07: 0.42367; 08: 0.41491; 09: 0.31333; 10: 0.30667; 11: 0.39; 12: 0.43167
| Median |  | 02: 0.3; 04: 0.38; 05: 0.5; 06: 0.50318; 07: 0.42103; 08: 0.41824; 09: 0.31; 10: 0.305; 11: 0.38; 12: 0.43
| Std Dev. |  | 02: 0.05164; 04: 0.013292; 05: 0.040825; 06: 0.0038394; 07: 0.013026; 08: 0.010042; 09: 0.015055; 10: 0.008165; 11: 0.029665; 12: 0.0040825
| Rel.Std.Dev. |  | 02: 19.36%; 04: 3.483%; 05: 8.447%; 06: 0.7621%; 07: 3.075%; 08: 2.42%; 09: 4.805%; 10: 2.662%; 11: 7.606%; 12: 0.9457%
| PDM3 |  | 02: -32.18%; 04: -2.928%; 05: 22.93%; 06: 28.13%; 07: 7.754%; 08: 5.526%; 09: -20.31%; 10: -22%; 11: -0.8088%; 12: 9.789%
| Z-Score (Absolute) |  | 02: NA; 04: 0.56; 05: NA; 06: 2.71; 07: 0.57; 08: 0.33; 09: 2.38; 10: 2.56; 11: 0.33; 12: 0.78
Results from laboratories 2 and 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 556b (Certified Value 8.8 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 7.8 | 9.09 | 9.2 | 9.17 | 8.9053 | 9.6 | 7.9 | 8.4 | 9.1 | 9.5 | < 2 | 7.6588 | 9.4965 | 9.55 | 7.977 | 7.59 | 9.18 | 7.8 | 8.46 | 10.1 | 9.41 | 9.3 | 8 |  |  |  |  | 1
| 1 | 2 | 7.6 | 9.03 | 9.5 | 8.99 | 8.808 | 9.9 | 7.6 | 8.4 | 8.6 | 8.5 | 8 | 7.5538 | 9.5236 | 9.55 | 7.9515 | 7.89 | 8.84 | 7.8 | 8.54 | 11.9 | 9.21 | 9.1 | 8 |  |  |  |  | 11
| 1 | 3 | 7.9 | 9.11 | 9.3 | 9.09 | 8.7613 | 9.7 | 8.2 | 8.6 | 9 | 9 | 7 | 8.9049 | 9.2501 | 9.5 | 7.9585 | 7.4 | 9.07 | 7.9 | 8.63 | 11.4 | 9.34 | 9.3 | 8.4 |  |  |  |  | 16
| 1 | 4 | 7.7 | 8.77 | 9.5 | 9.21 | 8.9399 | 9.2 | 8.2 | 8.9 | 8.8 | 9 | 3 | 8.9968 | 9.1375 | 9.28 | 8.0426 | 6.84 | 9.58 | 8.1 | 8.91 | 10.8 | 9.37 | 9 | 8.5 |  |  |  |  | 8.7957
| 1 | 5 | 7.4 | 8.82 | 9.8 | 8.71 | 8.9208 | 9.8 | 8 | 8.8 | 9.3 | 8.4 | 2 | 9.1532 | 9.0422 | 9.36 | 8.0957 | 6.99 | 9.23 | 7.8 | 8.72 | 10.6 | 9.32 | 9.1 | 8.6 |  |  |  |  | 140
| 1 | 6 | 7.2 | 8.78 | 9.7 | 8.78 | 9.1166 | 9.7 | 7.9 | 8.8 | 9.2 | 8.7 | 8 | 9.0266 | 9.5686 | 9.46 | 7.7766 | 7.23 | 8.92 | 8.3 | 8.52 | 10.2 | 9.38 | 9.1 | 8.1
| Mean |  | 7.6 | 8.9333 | 9.5 | 8.9917 | 8.9086 | 9.65 | 7.9667 | 8.65 | 9 | 8.85 | 5.6 | 8.549 | 9.3364 | 9.45 | 7.967 | 7.3233 | 9.1367 | 7.95 | 8.63 | 10.8333 | 9.3383 | 9.15 | 8.2667
| Median |  | 7.65 | 8.925 | 9.5 | 9.04 | 8.913 | 9.7 | 7.95 | 8.7 | 9.05 | 8.85 | 7 | 8.9508 | 9.3733 | 9.48 | 7.9678 | 7.315 | 9.125 | 7.85 | 8.585 | 10.7 | 9.355 | 9.1 | 8.25
| Std Dev. |  | 0.26077 | 0.16008 | 0.22804 | 0.20653 | 0.12346 | 0.2429 | 0.22509 | 0.21679 | 0.26077 | 0.40373 | 2.881 | 0.7353 | 0.22281 | 0.10918 | 0.10858 | 0.38769 | 0.26326 | 0.20736 | 0.16468 | 0.70048 | 0.070261 | 0.12247 | 0.26583
| Rel.Std.Dev. |  | 3.431% | 1.792% | 2.4% | 2.297% | 1.386% | 2.517% | 2.825% | 2.506% | 2.897% | 4.562% | 51.45% | 8.601% | 2.386% | 1.155% | 1.363% | 5.294% | 2.881% | 2.608% | 1.908% | 6.466% | 0.7524% | 1.339% | 3.216%
| PDM3 |  | -13.59% | 1.565% | 8.007% | 2.228% | 1.284% | 9.712% | -9.426% | -1.657% | 2.323% | 0.6171% | -36.33% | -2.805% | 6.147% | 7.439% | -9.422% | -16.74% | 3.876% | -9.615% | -1.884% | 23.17% | 6.169% | 4.028% | -6.015%
| Z-Score (Absolute) |  | 2.14 | 0.02 | 0.94 | 0.11 | 0.02 | 1.18 | 1.55 | 0.44 | 0.13 | 0.11 | NA | 0.6 | 0.67 | 0.86 | 1.54 | 2.59 | 0.35 | 1.57 | 0.47 | 3.1 | 0.68 | 0.37 | 1.06
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 556b (Certified Value 330 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 321.5 | 351 | 351 | 322 | 350 | 318 | 347.1049 | 317 | 309 | 323 | 311 | 333 | 310 | 337 | 342.7395 | 346.79 | 344 | 332.1578 | 357 | 360 | 300 | 308 | 312 | 341 | 333 | 326 | 341 | 1
| 1 | 2 | 324.9 | 334 | 352 | 322 | 363 | 313 | 341.8965 | 317 | 315 | 318 | 306 | 332 | 308 | 330 | 340.5357 | 340.19 | 334 | 329.857 | 360 | 341 | 288 | 317 | 321 | 345 | 333 | 327 | 340 | 38
| 1 | 3 | 313.2 | 362 | 352 | 314 | 350 | 324 | 337.4312 | 322 | 321 | 323 | 320 | 337 | 312 | 341 | 340.3793 | 348.79 | 328 | 330.4206 | 358 | 357 | 297 | 325 | 298 | 336 | 344 | 325 | 341 | 16
| 1 | 4 | 319.7 | 319 | 347 | 324 | 356 | 312 | 339.175 | 322 | 319 | 329 | 313 | 337 | 311 | 334 | 344.0827 | 340.76 | 330 | 327.6035 | 356 | 370 | 303 | 324 | 325 | 324 | 339 | 326 | 340 | 330.3722
| 1 | 5 | 327.4 | 328 | 347 | 315 | 363 | 309 | 335.1468 | 329 | 315 | 328 | 318 | 334 | 310 | 340 | 342.5006 | 344.3 | 332 | 324.5787 | 352 | 362 | 289 | 317 | 287 | 319 | 332 | 326 | 344 | 141
| 1 | 6 | 320.5 | 346 | 348 | 322 | 357 | 312 | 344.0176 | 327 | 315 | 332 | 336 | 334 | 310 | 340 | 344.916 | 349.01 | 331 | 323.3188 | 359 | 356 | 299 | 319 | 309 | 328 | 331 | 325 | 341
| Mean |  | 321.2 | 340 | 349.5 | 319.8333 | 356.5 | 314.6667 | 340.7953 | 322.3333 | 315.6667 | 325.5 | 317.3333 | 334.5 | 310.1667 | 337 | 342.5256 | 344.9733 | 333.1667 | 327.9894 | 357 | 357.6667 | 296 | 318.3333 | 308.6667 | 332.1667 | 335.3333 | 325.8333 | 341.1667
| Median |  | 321 | 340 | 349.5 | 322 | 356.5 | 312.5 | 340.5358 | 322 | 315 | 325.5 | 315.5 | 334 | 310 | 338.5 | 342.6201 | 345.545 | 331.5 | 328.7302 | 357.5 | 358.5 | 298 | 318 | 310.5 | 332 | 333 | 326 | 341
| Std Dev. |  | 4.8777 | 15.8871 | 2.429 | 4.2151 | 5.8224 | 5.4283 | 4.4103 | 4.9666 | 4.1312 | 5.0892 | 10.4243 | 2.0736 | 1.3292 | 4.2895 | 1.8304 | 3.8794 | 5.6716 | 3.4748 | 2.8284 | 9.5638 | 6.1319 | 6.121 | 14.2361 | 10.1472 | 5.0859 | 0.75277 | 1.472
| Rel.Std.Dev. |  | 1.519% | 4.673% | 0.695% | 1.318% | 1.633% | 1.725% | 1.294% | 1.541% | 1.309% | 1.564% | 3.285% | 0.6199% | 0.4285% | 1.273% | 0.5344% | 1.125% | 1.702% | 1.059% | 0.7923% | 2.674% | 2.072% | 1.923% | 4.612% | 3.055% | 1.517% | 0.231% | 0.4314%
| PDM3 |  | -2.776% | 2.914% | 5.79% | -3.19% | 7.909% | -4.754% | 3.155% | -2.433% | -4.451% | -1.475% | -3.947% | 1.249% | -6.116% | 2.006% | 3.679% | 4.42% | 0.8458% | -0.7213% | 8.06% | 8.262% | -10.4% | -3.644% | -6.57% | 0.5432% | 1.502% | -1.374% | 3.267%
| Z-Score (Absolute) |  | 0.67 | 0.48 | 1.07 | 0.76 | 1.5 | 1.08 | 0.53 | 0.6 | 1.01 | 0.41 | 0.91 | 0.14 | 1.35 | 0.3 | 0.64 | 0.79 | 0.06 | 0.26 | 1.53 | 1.57 | 2.22 | 0.85 | 1.44 | 0 | 0.19 | 0.39 | 0.55
| Analytical results for W in OREAS 556b (Certified Value 2.3 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 5 | 2.2 | < 10 | 2.4 | 2.7 | 1.9 | 2.6787 | 2.3 | 2 | 2.4 | 2.3 | 2.5 | 16 | 2.479 | 1.8 | 2.2323 | NA | 2.4 | < 10 | 2.3 | 1.9 | 2.5 | 2.3 | 2.2 | 2.3 |  |  | 1
| 1 | 2 | < 5 | 2.3 | < 10 | 2.3 | 3.1 | 1.9 | 2.6982 | 2.2 | 2 | 2.4 | 2.1 | 2.3 | 17 | 2.2723 | 1.8 | 2.3852 | NA | 2.3 | < 10 | 2.4 | 2 | 3 | 2.2 | 2.1 | 2.3 |  |  | 12
| 1 | 3 | < 5 | 2.3 | < 10 | 2.2 | 2.7 | 1.9 | 2.6329 | 2.4 | 2.2 | 2.5 | 2.3 | 2.4 | < 1 | 2.2807 | 1.9 | 2.2689 | NA | 2.3 | < 10 | 2.6 | 2 | 3 | 2.2 | 2.2 | 2.5 |  |  | 16
| 1 | 4 | < 5 | 2.3 | < 10 | 2.2 | 2.7 | 1.8 | 2.6366 | 2.3 | 2 | 2.5 | 4.2 | 2.4 | 21 | 2.3751 | 1.9 | 2.3363 | NA | 2.5 | < 10 | 2.4 | 1.9 | 2.6 | 2.3 | 2.1 | 2.6 |  |  | 2.2981
| 1 | 5 | < 5 | 2.2 | < 10 | 2.1 | 2.9 | 1.8 | 2.6963 | 2.3 | 2.2 | 2.3 | 2.3 | 2.4 | 15 | 2.4721 | 1.9 | 2.2574 | NA | 2.4 | < 10 | 2.3 | 1.9 | 2.7 | 2.3 | 2.2 | 2.4 |  |  | 142
| 1 | 6 | < 5 | 2.2 | < 10 | 2.2 | 2.8 | 1.8 | 2.6942 | 2.3 | 2 | 2.4 | 2.4 | 2.3 | 23 | 2.3569 | 1.9 | 2.3852 | NA | 2.5 | < 10 | 2.3 | 1.9 | 2.6 | 2.2 | 2 | 2.2
| Mean |  |  | 2.25 |  | 2.2333 | 2.8167 | 1.85 | 2.6728 | 2.3 | 2.0667 | 2.4167 | 2.6 | 2.3833 | 18.4 | 2.3727 | 1.8667 | 2.3109 |  | 2.4 |  | 2.3833 | 1.9333 | 2.7333 | 2.25 | 2.1333 | 2.3833
| Median |  |  | 2.25 |  | 2.2 | 2.75 | 1.85 | 2.6865 | 2.3 | 2 | 2.4 | 2.3 | 2.4 | 17 | 2.366 | 1.9 | 2.3026 |  | 2.4 |  | 2.35 | 1.9 | 2.65 | 2.25 | 2.15 | 2.35
| Std Dev. |  |  | 0.054772 |  | 0.10328 | 0.16021 | 0.054772 | 0.030316 | 0.063246 | 0.10328 | 0.075277 | 0.78994 | 0.075277 | 3.4351 | 0.089407 | 0.05164 | 0.067033 |  | 0.089443 |  | 0.1169 | 0.05164 | 0.21602 | 0.054772 | 0.08165 | 0.1472
| Rel.Std.Dev. |  |  | 2.434% |  | 4.624% | 5.688% | 2.961% | 1.134% | 2.75% | 4.997% | 3.115% | 30.38% | 3.158% | 18.67% | 3.768% | 2.766% | 2.901% |  | 3.727% |  | 4.905% | 2.671% | 7.903% | 2.434% | 3.827% | 6.176%
| PDM3 |  |  | -2.093% |  | -2.819% | 22.56% | -19.5% | 16.31% | 0.08233% | -10.07% | 5.159% | 13.14% | 3.709% | 700.7% | 3.246% | -18.77% | 0.5558% |  | 4.434% |  | 3.709% | -15.87% | 18.94% | -2.093% | -7.17% | 3.709%
| Z-Score (Absolute) |  | 0.43 | 0.49 | 9.62 | 0.55 | 1.59 | 1.96 | 1.06 | 0.31 | 1.16 | 0.12 | 0.8 | 0 | 47.94 | 0.04 | 1.9 | 0.27 | NA | 0.06 | 9.62 | 0 | 1.66 | 1.29 | 0.49 | 0.92 | 0
| Analytical results for Y in OREAS 556b (Certified Value 23.5 ppm) | [28 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 23.2 | 19.2 | 23.74 | 29.4 | 24 | 32.8238 | 22.9 | 22.4 | 23 | 23.2 | 24.1 | 18.76 | 24 | 22.9456 | 24.5 | 23.4862 | 19.94 | 24.2 | 19.2 | 22 | 16.4 | 23.8 | 24.47 | 23.7 | 22.5 |  |  | 1
| 1 | 2 | 22.7 | 19.6 | 23.49 | 29.5 | 23.6 | 32.7724 | 22.3 | 21.2 | 23.1 | 22.5 | 24.3 | 18.45 | 23 | 22.8737 | 24.1 | 23.5522 | 19.81 | 23 | 20.2 | 22.7 | 16.1 | 21.6 | 24.98 | 23.7 | 22.1 |  |  | 13
| 1 | 3 | 24.2 | 19.4 | 23.92 | 27.5 | 24 | 32.8957 | 24.7 | 23.6 | 23.6 | 22.9 | 24 | 18.62 | 24 | 23.8549 | 24.1 | 23.6441 | 19.83 | 24.7 | 19.4 | 23.6 | 16.4 | 22.3 | 24.45 | 23.6 | 22.5 |  |  | 16
| 1 | 4 | 21.4 | 20.6 | 23.18 | 29.5 | 23.6 | 33.5584 | 23.1 | 22.9 | 23.8 | 22.8 | 23.9 | 18.72 | 24 | 23.77 | 24.7 | 23.6977 | 19.48 | 24.8 | 20.1 | 23.7 | 17.4 | 22 | 24.89 | 23.5 | 22 |  |  | 23.4511
| 1 | 5 | 22.2 | 19.1 | 23.68 | 30.6 | 23.2 | 32.2503 | 23.9 | 22.6 | 22.6 | 26.5 | 24.8 | 18.75 | 24 | 23.6307 | 24 | 23.0627 | 19.2 | 25.2 | 19.1 | 23.9 | 15.6 | 22.2 | 24.68 | 23.2 | 22.7 |  |  | 143
| 1 | 6 | 22.8 | 20.3 | 23.28 | 29.7 | 23.3 | 32.8564 | 23.3 | 22.4 | 28 | 24.6 | 23.6 | 18.77 | 24 | 23.6593 | 24.1 | 22.4352 | 19.68 | 24.4 | 20.3 | 23.5 | 15.3 | 20.6 | 23.97 | 22.5 | 22.6
| Mean |  | 22.75 | 19.7 | 23.5483 | 29.3667 | 23.6167 | 32.8595 | 23.3667 | 22.5167 | 24.0167 | 23.75 | 24.1167 | 18.6783 | 23.8333 | 23.4557 | 24.25 | 23.313 | 19.6567 | 24.3833 | 19.7167 | 23.2333 | 16.2 | 22.0833 | 24.5733 | 23.3667 | 22.4
| Median |  | 22.75 | 19.5 | 23.585 | 29.5 | 23.6 | 32.8401 | 23.2 | 22.5 | 23.35 | 23.05 | 24.05 | 18.735 | 24 | 23.645 | 24.1 | 23.5192 | 19.745 | 24.55 | 19.75 | 23.55 | 16.25 | 22.1 | 24.575 | 23.55 | 22.5
| Std Dev. |  | 0.94181 | 0.61319 | 0.284 | 1.0152 | 0.33714 | 0.41708 | 0.83586 | 0.78592 | 1.9984 | 1.5346 | 0.40702 | 0.12449 | 0.40825 | 0.43107 | 0.28107 | 0.48537 | 0.27354 | 0.76004 | 0.54191 | 0.73121 | 0.73485 | 1.0439 | 0.36522 | 0.46332 | 0.28284
| Rel.Std.Dev. |  | 4.14% | 3.113% | 1.206% | 3.457% | 1.428% | 1.269% | 3.577% | 3.49% | 8.321% | 6.461% | 1.688% | 0.6665% | 1.713% | 1.838% | 1.159% | 2.082% | 1.392% | 3.117% | 2.748% | 3.147% | 4.536% | 4.727% | 1.486% | 1.983% | 1.263%
| PDM3 |  | -2.99% | -16% | 0.4147% | 25.23% | 0.7061% | 40.12% | -0.3599% | -3.984% | 2.412% | 1.275% | 2.838% | -20.35% | 1.63% | 0.01977% | 3.407% | -0.5887% | -16.18% | 3.975% | -15.92% | -0.9285% | -30.92% | -5.832% | 4.786% | -0.3599% | -4.482%
| Z-Score (Absolute) |  | 0.49 | 2.91 | 0.14 | 4.76 | 0.2 | 7.53 | 0 | 0.67 | 0.52 | 0.3 | 0.59 | 3.72 | 0.37 | 0.07 | 0.7 | 0.04 | 2.94 | 0.81 | 2.9 | 0.11 | 5.69 | 1.02 | 0.96 | 0 | 0.77
| Analytical results for Yb in OREAS 556b (Certified Value 2.72 ppm) | [28 empty] | Certified
Yb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  | 1
| No | No | 02: 02; 03: 04; 04: 05; 05: 07; 06: 14; 07: 16; 08: 18; 09: 19; 10: 20; 11: 23; 12: 25; 13: 26 |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*MS |  |  | 2
2
| 1 | 1 | 02: 2.2; 03: 2.83; 04: 3; 05: 3.2231; 06: 3; 07: 2.9018; 08: 2.7365; 09: 2.53; 10: 2.8; 11: 2; 12: 2.71; 13: 2.95 |  |  | 1
| 1 | 2 | 02: 2.2; 03: 2.76; 04: 3.1; 05: 3.1711; 06: 3; 07: 2.0175; 08: 2.7043; 09: 2.54; 10: 2.6; 11: 2; 12: 2.61; 13: 2.9 |  |  | 14
| 1 | 3 | 02: 2.2; 03: 2.82; 04: 3; 05: 3.169; 06: 3; 07: 2.7924; 08: 2.7263; 09: 2.51; 10: 2.8; 11: 2; 12: 2.67; 13: 2.94 |  |  | 16
| 1 | 4 | 02: 2.3; 03: 2.82; 04: 3; 05: 3.1852; 06: 3; 07: 2.7777; 08: 2.782; 09: 2.49; 10: 2.8; 11: 2.1; 12: 2.65; 13: 3.03 |  |  | 2.7218
| 1 | 5 | 02: 2.1; 03: 2.71; 04: 3.2; 05: 3.1174; 06: 3; 07: 2.7819; 08: 2.7373; 09: 2.54; 10: 2.8; 11: 2; 12: 2.55; 13: 3.03 |  |  | 144
| 1 | 6 | 02: 2.1; 03: 2.7; 04: 3.1; 05: 3.1867; 06: 3.1; 07: 2.925; 08: 2.6057; 09: 2.63; 10: 2.8; 11: 2; 12: 2.44; 13: 2.95
| Mean |  | 02: 2.1833; 03: 2.7733; 04: 3.0667; 05: 3.1754; 06: 3.0167; 07: 2.6994; 08: 2.7153; 09: 2.54; 10: 2.7667; 11: 2.0167; 12: 2.605; 13: 2.9667
| Median |  | 02: 2.2; 03: 2.79; 04: 3.05; 05: 3.1781; 06: 3; 07: 2.7872; 08: 2.7314; 09: 2.535; 10: 2.8; 11: 2; 12: 2.63; 13: 2.95
| Std Dev. |  | 02: 0.075277; 03: 0.058538; 04: 0.08165; 05: 0.034392; 06: 0.040825; 07: 0.34012; 08: 0.059379; 09: 0.048166; 10: 0.08165; 11: 0.040825; 12: 0.097519; 13: 0.052409
| Rel.Std.Dev. |  | 02: 3.448%; 03: 2.111%; 04: 2.662%; 05: 1.083%; 06: 1.353%; 07: 12.6%; 08: 2.187%; 09: 1.896%; 10: 2.951%; 11: 2.024%; 12: 3.744%; 13: 1.767%
| PDM3 |  | 02: -19.78%; 03: 1.894%; 04: 12.67%; 05: 16.67%; 06: 10.83%; 07: -0.8227%; 08: -0.2371%; 09: -6.679%; 10: 1.649%; 11: -25.91%; 12: -4.291%; 13: 8.997%
| Z-Score (Absolute) |  | 02: 1.76; 03: 0.1; 04: 1.03; 05: 1.37; 06: 0.87; 07: 0.13; 08: 0.08; 09: 0.64; 10: 0.08; 11: 2.29; 12: 0.43; 13: 0.71
| Analytical results for Zn in OREAS 556b (Certified Value 38.8 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 40.18 | 46.7 | 43.1 | 40 | 39 | 37.1 | 39.426 | 36 | 37 | 42 | 36 | 40 | 35.39 | 218 | 38.3118 | 38 | 36.5464 | 41 | 43 | 41 | 38 | 27 | 42.8 | 35 | 41 | 35 |  | 1
| 1 | 2 | 40.35 | 46.1 | 42.4 | 40 | 40 | 35.7 | 39.491 | 37 | 37 | 38 | 36 | 40 | 35.49 | 220 | 39.0607 | 39 | 36.6065 | 41 | 40 | 41.7 | 37 | 28 | 36.4 | 36 | 41 | 35 |  | 15
| 1 | 3 | 39.31 | 50.8 | 41.9 | 42 | 40 | 36.4 | 38.1156 | 38 | 37 | 39 | 38 | 40 | 35.12 | 219 | 39.8565 | 40 | 36.0167 | 41 | 40.9 | 41.8 | 37 | 23 | 36.5 | 35 | 40 | 35 |  | 16
| 1 | 4 | 41.9 | 42.5 | 41.6 | 40 | 41 | 36.6 | 40.6041 | 37 | 37 | 39 | 37 | 40 | 35.21 | 210 | 38.0366 | 38 | 36.1868 | 41 | 41.2 | 41.5 | 35 | 25 | 38.4 | 35 | 41 | 35 |  | 38.8027
| 1 | 5 | 40.2 | 45.4 | 43 | 41 | 40 | 35.5 | 38.5173 | 37 | 36 | 38 | 43 | 40 | 35.16 | 217 | 38.3139 | 40 | 37.3095 | 40 | 41.5 | 41.1 | 36 | 26 | 37.5 | 37 | 39 | 36 |  | 145
| 1 | 6 | 39.73 | 47.1 | 42.9 | 40 | 40 | 35.7 | 41.4441 | 37 | 37 | 38 | 38 | 39 | 35.27 | 212 | 39.6124 | 39 | 36.2895 | 41 | 41 | 42 | 36 | 25 | 35.8 | 57 | 41 | 35
| Mean |  | 40.2783 | 46.4333 | 42.4833 | 40.5 | 40 | 36.1667 | 39.5997 | 37 | 36.8333 | 39 | 38 | 39.8333 | 35.2733 | 216 | 38.8653 | 39 | 36.4926 | 40.8333 | 41.2667 | 41.5167 | 36.5 | 25.6667 | 37.9 | 39.1667 | 40.5 | 35.1667
| Median |  | 40.19 | 46.4 | 42.65 | 40 | 40 | 36.05 | 39.4585 | 37 | 37 | 38.5 | 37.5 | 40 | 35.24 | 217.5 | 38.6873 | 39 | 36.418 | 41 | 41.1 | 41.6 | 36.5 | 25.5 | 37 | 35.5 | 41 | 35
| Std Dev. |  | 0.88194 | 2.692 | 0.62423 | 0.83666 | 0.63246 | 0.6314 | 1.2512 | 0.63246 | 0.40825 | 1.5492 | 2.6077 | 0.40825 | 0.14208 | 4.0497 | 0.75845 | 0.89443 | 0.45703 | 0.40825 | 0.98725 | 0.39707 | 1.0488 | 1.7512 | 2.5706 | 8.7731 | 0.83666 | 0.40825
| Rel.Std.Dev. |  | 2.19% | 5.797% | 1.469% | 2.066% | 1.581% | 1.746% | 3.16% | 1.709% | 1.108% | 3.972% | 6.862% | 1.025% | 0.4028% | 1.875% | 1.951% | 2.293% | 1.252% | 0.9998% | 2.392% | 0.9564% | 2.873% | 6.823% | 6.783% | 22.4% | 2.066% | 1.161%
| PDM3 |  | 3.803% | 19.67% | 9.485% | 4.374% | 3.086% | -6.793% | 2.054% | -4.646% | -5.075% | 0.5085% | -2.069% | 2.656% | -9.096% | 456.7% | 0.1614% | 0.5085% | -5.954% | 5.233% | 6.35% | 6.994% | -5.934% | -33.85% | -2.326% | 0.938% | 4.374% | -9.371%
| Z-Score (Absolute) |  | 0.42 | 2.59 | 1.2 | 0.5 | 0.32 | 1.03 | 0.18 | 0.73 | 0.79 | 0.03 | 0.38 | 0.26 | 1.34 | 62.24 | 0.08 | 0.03 | 0.91 | 0.62 | 0.77 | 0.86 | 0.91 | 4.72 | 0.42 | 0.03 | 0.5 | 1.38
| Analytical results for Zr in OREAS 556b (Certified Value 171 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 166.9 | 178 | 169 | 178.8 | 180 | 172 | 223.269 | 169 | 170 | 171.5 | 173.5 | 192 | 154 | 146 | 171.4722 | 161.374 | 175 | 169.4722 | 159.4 | 178 | 158.9 | 168.7 | 164 | 186 | 182.8 | 173.2 | 175.9 | 1
| 1 | 2 | 166.1 | 172.9 | 170 | 177.3 | 186 | 169 | 222.1377 | 167 | 166 | 170 | 166.5 | 188.5 | 150 | 148 | 169.8477 | 167.391 | 174 | 168.8478 | 159.5 | 174 | 157.9 | 168.1 | 171 | 176 | 175.9 | 168.2 | 175.1 | 18
| 1 | 3 | 163.2 | 183.6 | 169 | 174.2 | 184 | 173 | 223.359 | 175 | 175 | 174.5 | 174 | 185.5 | 153 | 148 | 170.9756 | 168.7587 | 177 | 168.3984 | 159.7 | 176 | 162.9 | 170.8 | 170 | 178 | 177.8 | 169.4 | 176.7 | 16
| 1 | 4 | 171.2 | 163.4 | 171 | 174.2 | 189 | 168 | 224.8436 | 173.5 | 173.5 | 176.5 | 170.5 | 189.5 | 152 | 146 | 171.7074 | 162.2672 | 174 | 170.4022 | 156.6 | 182 | 157.3 | 169.4 | 177 | 178 | 178.9 | 173 | 176.2 | 171.093
| 1 | 5 | 172.7 | 169.4 | 170 | 174.8 | 194 | 166 | 224.0417 | 178 | 174.5 | 179 | 176 | 186.5 | 151 | 148 | 169.9286 | 161.1239 | 177 | 167.5355 | 153.8 | 179 | 157 | 168.5 | 158 | 179 | 183.4 | 169.6 | 179.3 | 146
| 1 | 6 | 166 | 177.4 | 170 | 171.6 | 189 | 166 | 220.9936 | 176 | 167 | 173 | 182 | 183.5 | 155 | 145 | 171.1002 | 166.5811 | 176 | 162.471 | 159.9 | 177 | 156.6 | 169.1 | 161 | 163 | 177.8 | 170.1 | 178.3
| Mean |  | 167.6833 | 174.1167 | 169.8333 | 175.15 | 187 | 169 | 223.1074 | 173.0833 | 171 | 174.0833 | 173.75 | 187.5833 | 152.5 | 146.8333 | 170.8386 | 164.5826 | 175.5 | 167.8545 | 158.15 | 177.6667 | 158.4333 | 169.1 | 166.8333 | 176.6667 | 179.4333 | 170.5833 | 176.9167
| Median |  | 166.5 | 175.15 | 170 | 174.5 | 187.5 | 168.5 | 223.314 | 174.25 | 171.75 | 173.75 | 173.75 | 187.5 | 152.5 | 147 | 171.0379 | 164.4241 | 175.5 | 168.6231 | 159.45 | 177.5 | 157.6 | 168.9 | 167 | 178 | 178.35 | 169.85 | 176.45
| Std Dev. |  | 3.5662 | 7.1326 | 0.75277 | 2.5485 | 4.8166 | 2.9665 | 1.3701 | 4.2475 | 3.9115 | 3.3078 | 5.2225 | 3.04 | 1.8708 | 1.3292 | 0.78147 | 3.3746 | 1.3784 | 2.8098 | 2.4566 | 2.7325 | 2.3304 | 0.94868 | 7.0828 | 7.5277 | 3.0058 | 2.0478 | 1.5804
| Rel.Std.Dev. |  | 2.127% | 4.096% | 0.4432% | 1.455% | 2.576% | 1.755% | 0.6141% | 2.454% | 2.287% | 1.9% | 3.006% | 1.621% | 1.227% | 0.9052% | 0.4574% | 2.05% | 0.7854% | 1.674% | 1.553% | 1.538% | 1.471% | 0.561% | 4.245% | 4.261% | 1.675% | 1.2% | 0.8933%
| PDM3 |  | -1.993% | 1.767% | -0.7362% | 2.371% | 9.297% | -1.223% | 30.4% | 1.163% | -0.05435% | 1.748% | 1.553% | 9.638% | -10.87% | -14.18% | -0.1487% | -3.805% | 2.576% | -1.893% | -7.565% | 3.842% | -7.399% | -1.165% | -2.49% | 3.258% | 4.875% | -0.2979% | 3.404%
| Z-Score (Absolute) |  | 0.54 | 0.5 | 0.19 | 0.67 | 2.59 | 0.32 | 8.43 | 0.34 | 0 | 0.5 | 0.45 | 2.68 | 2.99 | 3.91 | 0.03 | 1.04 | 0.73 | 0.51 | 2.08 | 1.08 | 2.03 | 0.31 | 0.67 | 0.92 | 1.36 | 0.07 | 0.96
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C714: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C862: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C955: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C974: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C992: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1011: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1049: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1067: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1085: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 556b (Certified Value 0.12 ppm) | [28 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | 2.94 | 0.121 | < 0.5 | 0.13 | 0.118 | 0.14 | < 0.2 | 0.13 | 0.11 | 0.14 | 0.14 | 0.12 | < 5 | 0.18 | 0.094519 | 0.08 | 0.13006 | < 0.3 | 0.1 | 0.2 | 0.1 | 0.13 | 0.1 | < 0.5 | 0.13 |  |  | 1
| 1 | 2 | 2.99 | 0.11 | < 0.5 | 0.14 | 0.103 | 0.14 | < 0.2 | 0.12 | 0.12 | 0.15 | 0.14 | 0.12 | < 5 | 0.13 | 0.1088 | 0.1 | 0.12859 | < 0.3 | 0.1 | < 0.2 | < 0.1 | 0.12 | 0.11 | < 0.5 | 0.12 |  |  | 1
| 1 | 3 | 3.27 | 0.11 | < 0.5 | 0.13 | 0.106 | 0.14 | < 0.2 | 0.13 | 0.12 | 0.15 | 0.13 | 0.11 | < 5 | 0.12 | 0.097862 | 0.09 | 0.13171 | < 0.3 | 0.1 | < 0.2 | < 0.1 | 0.11 | 0.11 | < 0.5 | 0.12 |  |  | 16
| 1 | 4 | 3.24 | 0.099 | < 0.5 | 0.14 | 0.112 | 0.14 | < 0.2 | 0.13 | 0.12 | 0.13 | 0.14 | 0.12 | < 5 | 0.12 | 0.1094 | 0.08 | 0.13246 | < 0.3 | 0.1 | < 0.2 | 0.1 | 0.13 | 0.11 | < 0.5 | 0.12 |  |  | 0.11978
| 1 | 5 | 2.87 | 0.11 | < 0.5 | 0.14 | 0.096 | 0.14 | < 0.2 | 0.14 | 0.12 | 0.13 | 0.12 | 0.12 | < 5 | 0.11 | 0.095522 | 0.09 | 0.12757 | < 0.3 | 0.1 | 0.2 | 0.1 | 0.13 | 0.11 | < 0.5 | 0.12 |  |  | 148
| 1 | 6 | 3.02 | 0.098 | < 0.5 | 0.14 | 0.098 | 0.15 | < 0.2 | 0.12 | 0.12 | 0.21 | 0.12 | 0.13 | < 5 | 0.11 | 0.095661 | 0.07 | 0.1277 | < 0.3 | 0.1 | 0.2 | 0.1 | 0.12 | 0.09 | < 0.5 | 0.14
| Mean |  | 3.055 | 0.108 |  | 0.13667 | 0.1055 | 0.14167 |  | 0.12833 | 0.11833 | 0.15167 | 0.13167 | 0.12 |  | 0.12833 | 0.1003 | 0.085 | 0.12968 |  | 0.1 | 0.2 | 0.1 | 0.12333 | 0.105 |  | 0.125
| Median |  | 3.005 | 0.11 |  | 0.14 | 0.1045 | 0.14 |  | 0.13 | 0.12 | 0.145 | 0.135 | 0.12 |  | 0.12 | 0.096762 | 0.085 | 0.12932 |  | 0.1 | 0.2 | 0.1 | 0.125 | 0.11 |  | 0.12
| Std Dev. |  | 0.16331 | 0.0085088 |  | 0.005164 | 0.0083845 | 0.0040825 |  | 0.0075277 | 0.0040825 | 0.029944 | 0.0098319 | 0.0063246 |  | 0.026394 | 0.0069119 | 0.010488 | 0.0020766 |  | 1.5202e-17 | 3.3993e-17 | 0 | 0.008165 | 0.0083666 |  | 0.0083666
| Rel.Std.Dev. |  | 5.346% | 7.879% |  | 3.779% | 7.947% | 2.882% |  | 5.866% | 3.45% | 19.74% | 7.467% | 5.27% |  | 20.57% | 6.892% | 12.34% | 1.601% |  | 1.52e-14% | 1.7e-14% | 0% | 6.62% | 7.968% |  | 6.693%
| PDM3 |  | 2451% | -9.834% |  | 14.1% | -11.92% | 18.27% |  | 7.141% | -1.208% | 26.62% | 9.924% | 0.1839% |  | 7.141% | -16.27% | -29.04% | 8.266% |  | -16.51% | 66.97% | -16.51% | 2.967% | -12.34% |  | 4.358%
| Z-Score (Absolute) |  | 93.98 | 0.65 | 3.91 | 0.27 | 0.73 | 0.43 | 0.91 | 0 | 0.32 | 0.75 | 0.11 | 0.27 | 76.15 | 0 | 0.9 | 1.39 | 0.04 | 0.7 | NA | NA | NA | 0.16 | 0.75 | 3.91 | 0.11
Results from laboratories 20, 21 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 556b (Certified Value 0.869 wt.%) | [28 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.81 | 0.77 | 1.1 | 0.9708 | 1.24 | 0.5 | 0.82 | 0.93 | 0.74 | 0.79 | 0.82 | 0.8678 | 0.63 | 1.129 | 1.4 | 0.85532 | 0.59 | 0.62 | 0.91 | 0.954 | 1.36 | 1.3808 | 0.9 | 0.92 |  |  |  | 1
| 1 | 2 | 0.81 | 0.89 | 1.09 | 1.0016 | 1.01 | 0.5 | 0.82 | 0.96 | 0.78 | 0.77 | 0.84 | 0.8901 | 0.688 | 1.0799 | 1.41 | 0.85504 | 0.61 | 0.63 | 0.95 | 0.936 | 1.6 | 1.4218 | 0.87 | 0.88 |  |  |  | 0
| 1 | 3 | 0.9 | 0.81 | 1.06 | 0.9726 | 1.1 | 0.49 | 0.74 | 0.95 | 0.78 | 0.81 | 0.83 | 0.8824 | 0.662 | 1.099 | 1.4 | 0.87785 | 0.62 | 0.62 | 0.92 | 1.005 | 1.25 | 1.3242 | 0.89 | 0.85 |  |  |  | 16
| 1 | 4 | 0.87 | 0.74 | 1.04 | 1.0125 | 1.18 | 0.48 | 0.75 | 0.97 | 0.77 | 0.79 | 0.82 | 0.8826 | 0.656 | 1.071 | 1.45 | 0.86325 | 0.62 | 0.6 | 0.9 | 0.928 | 1.88 | 1.3034 | 0.93 | 0.84 |  |  |  | 0.86918
| 1 | 5 | 0.95 | 0.88 | 1.06 | 0.9591 | 1.14 | 0.51 | 0.74 | 0.97 | 0.79 | 0.79 | 0.84 | 0.8831 | 0.714 | 1.1107 | 1.36 | 0.86795 | 0.63 | 0.61 | 0.92 | 0.992 | 1.56 | 1.3555 | 0.91 | 0.83 |  |  |  | 149
| 1 | 6 | 0.83 | 0.83 | 1.06 | 0.9915 | 1.14 | 0.52 | 0.74 | 0.97 | 0.8 | 0.76 | 0.82 | 0.8836 | 0.656 | 1.1219 | 1.42 | 0.8532 | 0.64 | 0.62 | 0.9 | 0.998 | 1.67 | 1.3726 | 0.9 | 0.86
| Mean |  | 0.86167 | 0.82 | 1.0683 | 0.98468 | 1.135 | 0.5 | 0.76833 | 0.95833 | 0.77667 | 0.785 | 0.82833 | 0.8816 | 0.66767 | 1.1019 | 1.4067 | 0.8621 | 0.61833 | 0.61667 | 0.91667 | 0.96883 | 1.5533 | 1.3597 | 0.9 | 0.86333
| Median |  | 0.85 | 0.82 | 1.06 | 0.98205 | 1.14 | 0.5 | 0.745 | 0.965 | 0.78 | 0.79 | 0.825 | 0.88285 | 0.659 | 1.1048 | 1.405 | 0.85929 | 0.62 | 0.62 | 0.915 | 0.973 | 1.58 | 1.364 | 0.9 | 0.855
| Std Dev. |  | 0.056006 | 0.05933 | 0.022286 | 0.020492 | 0.077395 | 0.014142 | 0.040208 | 0.016021 | 0.020656 | 0.017607 | 0.0098319 | 0.0073564 | 0.029269 | 0.023043 | 0.029439 | 0.0095773 | 0.017224 | 0.010328 | 0.018619 | 0.033648 | 0.22447 | 0.042227 | 0.02 | 0.03266
| Rel.Std.Dev. |  | 6.5% | 7.235% | 2.086% | 2.081% | 6.819% | 2.828% | 5.233% | 1.672% | 2.66% | 2.243% | 1.187% | 0.8344% | 4.384% | 2.091% | 2.093% | 1.111% | 2.786% | 1.675% | 2.031% | 3.473% | 14.45% | 3.106% | 2.222% | 3.783%
| PDM3 |  | -0.864% | -5.658% | 22.91% | 13.29% | 30.58% | -42.47% | -11.6% | 10.26% | -10.64% | -9.685% | -4.699% | 1.429% | -23.18% | 26.78% | 61.84% | -0.814% | -28.86% | -29.05% | 5.464% | 11.47% | 78.71% | 56.44% | 3.546% | -0.6723%
| Z-Score (Absolute) |  | 0.07 | 0.35 | 1.32 | 0.75 | 1.77 | 2.51 | 0.7 | 0.58 | 0.64 | 0.59 | 0.3 | 0.06 | 1.38 | 1.54 | 3.59 | 0.07 | 1.71 | 1.72 | 0.3 | 0.65 | 4.58 | 3.28 | 0.19 | 0.06
| Analytical results for As in OREAS 556b (Certified Value 94 ppm) | [28 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | 1
| 1 | 1 | 73.6 | 82.4 | 93.9 | 99 | 99.6 | 102 | 94.1408 | 106.5 | 101 | 94.9 | 97.9 | 100 | 85.55 | 95.6 | 91.1559 | 89.49 | 96 | 94.3495 | 84.8 | 98.1 | 89 | 92.7 | 86 | 91.8 | 93 | 95 | 98.7 | 1
| 1 | 2 | 71 | 81.4 | 91.8 | 97 | 89 | 102 | 91.8679 | 107 | 104 | 100 | 97.7 | 103 | 84.86 | 109 | 90.9935 | 89.97 | 95 | 94.9668 | 83.7 | 100 | 88 | 91.9 | 82 | 87.3 | 93 | 94 | 96.2 | 2
| 1 | 3 | 71.6 | 84.4 | 91.1 | 100 | 93.2 | 105 | 93.7255 | 106 | 102 | 97.2 | 97.7 | 102.5 | 85.24 | 101 | 93.6059 | 89.05 | 97 | 97.9222 | 89 | 97.1 | 90 | 88.4 | 82 | 86.4 | 91 | 94 | 96.4 | 16
| 1 | 4 | 77 | 83.6 | 91.8 | 100 | 96.7 | 102 | 91.7444 | 102 | 102.5 | 99.6 | 97.6 | 103 | 85.26 | 100 | 92.3649 | 90.82 | 95 | 97.677 | 85.8 | 92.4 | 87 | 90.5 | 84 | 90.9 | 90 | 94 | 98.4 | 94.0686
| 1 | 5 | 74.4 | 83.3 | 91.7 | 100 | 90.7 | 102 | 91.3382 | 108.5 | 103 | 97.3 | 98.1 | 103.5 | 85.21 | 99 | 89.1044 | 90.24 | 97 | 98.028 | 86.9 | 92.5 | 88 | 87.8 | 82 | 86 | 94 | 95 | 96.7 | 150
| 1 | 6 | 71 | 96.1 | 93.2 | 98 | 88.8 | 104 | 92.502 | 107.5 | 103.5 | 100.5 | 96.7 | 103.5 | 85.16 | 97.7 | 89.1588 | 89.89 | 97 | 98.2398 | 87.7 | 94.5 | 91 | 87 | 83 | 87.3 | 94 | 96 | 100.3
| Mean |  | 73.1 | 85.2 | 92.25 | 99 | 93 | 102.8333 | 92.5531 | 106.25 | 102.6667 | 98.25 | 97.6167 | 102.5833 | 85.2133 | 100.3833 | 91.0639 | 89.91 | 96.1667 | 96.8639 | 86.3167 | 95.7667 | 88.8333 | 89.7167 | 83.1667 | 88.2833 | 92.5 | 94.6667 | 97.7833
| Median |  | 72.6 | 83.45 | 91.8 | 99.5 | 91.95 | 102 | 92.185 | 106.75 | 102.75 | 98.45 | 97.7 | 103 | 85.225 | 99.5 | 91.0747 | 89.93 | 96.5 | 97.7996 | 86.35 | 95.8 | 88.5 | 89.45 | 82.5 | 87.3 | 93 | 94.5 | 97.55
| Std Dev. |  | 2.3757 | 5.4387 | 1.0635 | 1.2649 | 4.3868 | 1.3292 | 1.1399 | 2.2528 | 1.0801 | 2.1529 | 0.48339 | 1.3197 | 0.22088 | 4.621 | 1.7684 | 0.60943 | 0.98319 | 1.7292 | 1.9426 | 3.1226 | 1.472 | 2.3267 | 1.6021 | 2.4457 | 1.6432 | 0.8165 | 1.6216
| Rel.Std.Dev. |  | 3.25% | 6.384% | 1.153% | 1.278% | 4.717% | 1.293% | 1.232% | 2.12% | 1.052% | 2.191% | 0.4952% | 1.286% | 0.2592% | 4.603% | 1.942% | 0.6778% | 1.022% | 1.785% | 2.251% | 3.261% | 1.657% | 2.593% | 1.926% | 2.77% | 1.776% | 0.8625% | 1.658%
| PDM3 |  | -22.29% | -9.428% | -1.933% | 5.242% | -1.136% | 9.317% | -1.611% | 12.95% | 9.14% | 4.445% | 3.772% | 9.052% | -9.414% | 6.713% | -3.194% | -4.421% | 2.23% | 2.971% | -8.241% | 1.805% | -5.565% | -4.626% | -11.59% | -6.15% | -1.668% | 0.6358% | 3.949%
| Z-Score (Absolute) |  | 2.84 | 1.12 | 0.11 | 0.86 | 0 | 1.41 | 0.06 | 1.89 | 1.38 | 0.75 | 0.66 | 1.37 | 1.11 | 1.06 | 0.28 | 0.44 | 0.45 | 0.55 | 0.96 | 0.4 | 0.6 | 0.47 | 1.41 | 0.67 | 0.07 | 0.24 | 0.68
| Analytical results for Au in OREAS 556b (Certified Value < 0.02 ppm) | [28 empty] | Certified
Au | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 07; 06: 08; 07: 09; 08: 10; 09: 11; 10: 12; 11: 15; 12: 25; 13: 26; 14: 28 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*MS; 09: AR*MS; 10: AR*MS; 11: AR*OES/MS; 12: AR*OES/MS; 13: AR*MS; 14: AR*OES/MS | 3
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g; 04: 0.5g; 05: 0.5g; 06: 0.5g; 07: 0.5g; 08: 0.5g; 09: 0.5g; 10: 0.5g; 11: 01g; 12: 0.1g; 13: 01g; 14: 0.5g | 3
| 1 | 1 | 01: < 0.0002; 02: < 0.001; 03: < 0.0005; 04: < 0.005; 05: < 0.0156; 06: < 0.02; 07: < 0.02; 08: < 0.02; 09: < 0.02; 10: < 0.02; 11: < 0.001; 12: 0.0013; 13: < 0.001; 14: 0.0006 | 1
| 1 | 2 | 01: < 0.0002; 02: < 0.001; 03: < 0.0005; 04: < 0.005; 05: 0.0143; 06: < 0.02; 07: < 0.02; 08: < 0.02; 09: < 0.02; 10: < 0.02; 11: < 0.001; 12: 0.0005; 13: < 0.001; 14: 0.0007 | 0
| 1 | 3 | 01: < 0.0002; 02: < 0.001; 03: < 0.0005; 04: < 0.005; 05: < 0.003; 06: < 0.02; 07: < 0.02; 08: < 0.02; 09: < 0.02; 10: < 0.02; 11: < 0.001; 12: 0.0013; 13: < 0.001; 14: 0.0007 | 16
| 1 | 4 | 01: < 0.0002; 02: < 0.001; 03: < 0.0005; 04: < 0.005; 05: < 0.003; 06: < 0.02; 07: < 0.02; 08: < 0.02; 09: < 0.02; 10: < 0.02; 11: < 0.001; 12: 0.001; 13: < 0.001; 14: 0.0006 | < 0.02
| 1 | 5 | 01: < 0.0002; 02: < 0.001; 03: < 0.0005; 04: < 0.005; 05: < 0.003; 06: < 0.02; 07: < 0.02; 08: < 0.02; 09: < 0.02; 10: < 0.02; 11: < 0.001; 12: 0.0013; 13: < 0.001; 14: < 0.0005 | 151
| 1 | 6 | 01: < 0.0002; 02: < 0.001; 03: < 0.0005; 04: < 0.005; 05: 0.00555; 06: < 0.02; 07: < 0.02; 08: < 0.02; 09: < 0.02; 10: < 0.02; 11: < 0.001; 12: 0.0005; 13: < 0.001; 14: < 0.0005
| Mean |  | 05: 0.009925; 12: 0.00098333; 14: 0.00065
| Median |  | 05: 0.009925; 12: 0.00115; 14: 0.00065
| Std Dev. |  | 05: 0.0061872; 12: 0.000392; 14: 5.7735e-05
| Rel.Std.Dev. |  | 05: 62.34%; 12: 39.86%; 14: 8.882%
PDM3
| Z-Score (Absolute) |  | 01: 0.71; 02: 0.53; 03: 0.64; 04: 0.33; 05: 1.56; 06: 3.55; 07: 3.55; 08: 3.55; 09: 3.55; 10: 3.55; 11: 0.53; 12: 0.33; 13: 0.53; 14: 0.53
| Analytical results for B in OREAS 556b (Certified Value < 10 ppm) | [28 empty] | Certified
B | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 08; 06: 09; 07: 10; 08: 11; 09: 12; 10: 15; 11: 18; 12: 21; 13: 25; 14: 26; 15: 28 | ppm
|  |  | 01: AR*MS; 02: AR*OES; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*MS; 09: AR*MS; 10: AR*OES/MS; 11: AR*MS; 12: AR*OES; 13: AR*OES/MS; 14: AR*OES; 15: AR*OES/MS | 1
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g; 04: 0.5g; 05: 0.5g; 06: 0.5g; 07: 0.5g; 08: 0.5g; 09: 0.5g; 10: 01g; 11: 0.2g; 12: 0.25g; 13: 0.1g; 14: 01g; 15: 0.5g | 1
| 1 | 1 | 01: < 20; 02: < 10; 03: 5; 04: 2.8; 05: < 10; 06: < 10; 07: < 10; 08: 10; 09: 10; 10: < 10; 11: 3.2875; 12: < 10; 13: < 10; 14: < 10; 15: < 10 | 1
| 1 | 2 | 01: < 20; 02: < 10; 03: 5; 04: 2.6; 05: < 10; 06: < 10; 07: < 10; 08: 10; 09: < 10; 10: < 10; 11: 3.0347; 12: < 10; 13: 13; 14: < 10; 15: < 10 | 0
| 1 | 3 | 01: < 20; 02: < 10; 03: 5; 04: 3; 05: < 10; 06: < 10; 07: < 10; 08: 10; 09: < 10; 10: < 10; 11: 2.9922; 12: 10; 13: < 10; 14: < 10; 15: < 10 | 16
| 1 | 4 | 01: < 20; 02: < 10; 03: 5; 04: 2.8; 05: < 10; 06: < 10; 07: < 10; 08: 10; 09: < 10; 10: < 10; 11: 3.3193; 12: 10; 13: 11; 14: < 10; 15: < 10 | < 10
| 1 | 5 | 01: < 20; 02: < 10; 03: 5; 04: 2.7; 05: < 10; 06: < 10; 07: < 10; 08: 10; 09: < 10; 10: < 10; 11: 3.4689; 12: < 10; 13: 13; 14: < 10; 15: < 10 | 152
| 1 | 6 | 01: < 20; 02: < 10; 03: 5; 04: 3.1; 05: < 10; 06: < 10; 07: < 10; 08: 10; 09: 10; 10: < 10; 11: 3.3218; 12: 10; 13: < 10; 14: < 10; 15: < 10
| Mean |  | 03: 5; 04: 2.8333; 08: 10; 09: 10; 11: 3.2374; 12: 10; 13: 12.3333
| Median |  | 03: 5; 04: 2.8; 08: 10; 09: 10; 11: 3.3034; 12: 10; 13: 13
| Std Dev. |  | 03: 0; 04: 0.18619; 08: 0; 09: 0; 11: 0.18502; 12: 0; 13: 1.1547
| Rel.Std.Dev. |  | 03: 0%; 04: 6.571%; 08: 0%; 09: 0%; 11: 5.715%; 12: 0%; 13: 9.362%
PDM3
Z-Score (Absolute) | bar: 01 to 15
| Analytical results for Ba in OREAS 556b (Certified Value 53 ppm) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 0
| 1 | 1 | 59.1 | 41.1 | 15.5 | 20 | 10.8 | 48 | 30 | 40 | 30 | 40 | 40 | 53.82 | < 1 | 53.1959 | 69 | 47.6434 | 51 | 25 | 59 | 70.7 | 67 | 55 | 33 |  |  |  |  | 1
| 1 | 2 | 59.9 | 44.3 | 15.5 | 16 | 10.8 | 48 | 40 | 40 | 40 | 40 | 40 | 52.51 | 2 | 52.6532 | 67 | 47.9527 | 50 | 24 | 58 | 73.4 | 68 | 54 | 35 |  |  |  |  | 30
| 1 | 3 | 80.1 | 41.7 | 15.8 | 17 | 9.6 | 50 | 40 | 40 | 30 | 40 | 50 | 53.42 | 3 | 53.7121 | 69 | 49.02 | 51 | 25 | 54 | 70.7 | 65 | 55 | 31 |  |  |  |  | 16
| 1 | 4 | 57.9 | 40.3 | 15.2 | 25 | 10.2 | 48 | 40 | 50 | 30 | 40 | 50 | 53.55 | < 1 | 53.4095 | 70 | 48.373 | 50 | 26 | 59 | 79.4 | 63 | 54 | 30 |  |  |  |  | 52.5177
| 1 | 5 | 61.5 | 44.1 | 15.7 | 17 | 9.3 | 48 | 40 | 40 | 40 | 40 | 50 | 53.61 | 1 | 52.4379 | 68 | 47.1531 | 50 | 27 | 59 | 74.2 | 66 | 55 | 30 |  |  |  |  | 153
| 1 | 6 | 59.7 | 44.5 | 15.4 | 18 | 10.8 | 49 | 40 | 40 | 30 | 40 | 40 | 53.77 | 2 | 53.5293 | 68 | 48.0095 | 51 | 28 | 59 | 74.8 | 68 | 56 | 29
| Mean |  | 63.0333 | 42.6667 | 15.5167 | 18.8333 | 10.25 | 48.5 | 38.3333 | 41.6667 | 33.3333 | 40 | 45 | 53.4467 | 2 | 53.1563 | 68.5 | 48.0253 | 50.5 | 25.8333 | 58 | 73.8667 | 66.1667 | 54.8333 | 31.3333
| Median |  | 59.8 | 42.9 | 15.5 | 17.5 | 10.5 | 48 | 40 | 40 | 30 | 40 | 45 | 53.58 | 2 | 53.3027 | 68.5 | 47.9811 | 50.5 | 25.5 | 59 | 73.8 | 66.5 | 55 | 30.5
| Std Dev. |  | 8.4424 | 1.8479 | 0.2137 | 3.3116 | 0.66858 | 0.83666 | 4.0825 | 4.0825 | 5.164 | 0 | 5.4772 | 0.48153 | 0.8165 | 0.50652 | 1.0488 | 0.63601 | 0.54772 | 1.472 | 2 | 3.221 | 1.9408 | 0.75277 | 2.2509
| Rel.Std.Dev. |  | 13.39% | 4.331% | 1.377% | 17.58% | 6.523% | 1.725% | 10.65% | 9.798% | 15.49% | 0% | 12.17% | 0.9009% | 40.82% | 0.9529% | 1.531% | 1.324% | 1.085% | 5.698% | 3.448% | 4.361% | 2.933% | 1.373% | 7.184%
| PDM3 |  | 20.02% | -18.76% | -70.45% | -64.14% | -80.48% | -7.65% | -27.01% | -20.66% | -36.53% | -23.84% | -14.31% | 1.769% | -96.19% | 1.216% | 30.43% | -8.554% | -3.842% | -50.81% | 10.44% | 40.65% | 25.99% | 4.409% | -40.34%
| Z-Score (Absolute) |  | 0.67 | 0.42 | 1.88 | 1.7 | 2.17 | 0.11 | NA | NA | NA | NA | NA | 0.16 | NA | 0.14 | 0.97 | 0.13 | 0 | 1.33 | 0.4 | 1.26 | 0.84 | 0.23 | 1.03
Results from laboratory 14 were removed due to their 1 ppm reading resolution. Results from laboratories 8, 9, 10, 11 and 12 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 556b (Certified Value 2.2 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [5 empty] | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [5 empty] | 2
| 1 | 1 | 2.6 | < 5 | 2.39 | 2.2 | 2.02 | 2.19 | 2.27 | 2.17 | 2.15 | 2.14 | < 5 | 1.78 | 2.0339 | 2.6 | 2.2222 | 2.8 | 2.1 | 3 | 2.27 | 1.62 | 2.48 | 2.54 | [5 empty] | 1
| 1 | 2 | 2.6 | < 5 | 2.41 | 1.9 | 2.11 | 2.21 | 2.36 | 2.23 | 2.12 | 2.19 | < 5 | 1.77 | 2.0822 | 2.6 | 2.2518 | 2.8 | 2.1 | 2 | 2.19 | 1.76 | 2.51 | 2.48 | [5 empty] | 3
| 1 | 3 | 2.7 | < 5 | 2.25 | 1.9 | 2.02 | 2.11 | 2.34 | 2.21 | 2.15 | 2.18 | < 5 | 1.78 | 2.0942 | 2.6 | 2.2977 | 2.7 | 2.1 | 2 | 2.1 | 1.55 | 2.52 | 2.47 | [5 empty] | 16
| 1 | 4 | 2.4 | < 5 | 2.36 | 2.3 | 2.02 | 2.13 | 2.34 | 2.22 | 2.14 | 2.17 | < 5 | 1.79 | 1.9965 | 2.7 | 2.2485 | 2.7 | 2 | 2 | 2.21 | 1.79 | 2.37 | 2.45 | [5 empty] | 2.2028
| 1 | 5 | 2.7 | < 5 | 2.27 | 2.1 | 2.12 | 2.12 | 2.33 | 2.26 | 2.14 | 2.22 | < 5 | 1.79 | 1.9901 | 2.6 | 2.2148 | 2.8 | 2 | 2 | 2.1 | 1.7 | 2.44 | 2.45 | [5 empty] | 154
| 1 | 6 | 2.6 | < 5 | 2.23 | 2.1 | 2 | 2.14 | 2.36 | 2.3 | 2.11 | 2.16 | < 5 | 1.74 | 2.0042 | 2.7 | 2.2688 | 2.7 | 2.1 | 2 | 2.21 | 1.71 | 2.47 | 2.5
| Mean |  | 2.6 |  | 2.3183 | 2.0833 | 2.0483 | 2.15 | 2.3333 | 2.2317 | 2.135 | 2.1767 |  | 1.775 | 2.0335 | 2.6333 | 2.2506 | 2.75 | 2.0667 | 2.1667 | 2.18 | 1.6883 | 2.465 | 2.4817
| Median |  | 2.6 |  | 2.315 | 2.1 | 2.02 | 2.135 | 2.34 | 2.225 | 2.14 | 2.175 |  | 1.78 | 2.019 | 2.6 | 2.2502 | 2.75 | 2.1 | 2 | 2.2 | 1.705 | 2.475 | 2.475
| Std Dev. |  | 0.10954 |  | 0.077567 | 0.16021 | 0.052313 | 0.040497 | 0.033267 | 0.04446 | 0.016432 | 0.027325 |  | 0.018708 | 0.045091 | 0.05164 | 0.03047 | 0.054772 | 0.05164 | 0.40825 | 0.067528 | 0.089312 | 0.054681 | 0.034303
| Rel.Std.Dev. |  | 4.213% |  | 3.346% | 7.69% | 2.554% | 1.884% | 1.426% | 1.992% | 0.7696% | 1.255% |  | 1.054% | 2.217% | 1.961% | 1.354% | 1.992% | 2.499% | 18.84% | 3.098% | 5.29% | 2.218% | 1.382%
| PDM3 |  | 18.03% |  | 5.244% | -5.424% | -7.013% | -2.398% | 5.925% | 1.309% | -3.079% | -1.187% |  | -19.42% | -7.687% | 19.54% | 2.17% | 24.84% | -6.181% | -1.641% | -1.036% | -23.36% | 11.9% | 12.66%
| Z-Score (Absolute) |  | 1.35 | 0.99 | 0.32 | 0.55 | 0.67 | 0.3 | 0.37 | 0 | 0.36 | 0.2 | 0.99 | 1.68 | 0.73 | 1.48 | 0.07 | 1.91 | 0.61 | NA | 0.19 | 2 | 0.86 | 0.92
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 556b (Certified Value 2.82 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | 3
| 1 | 1 | < 5 | 2.34 | < 2 | 3.06 | 2.68 | 3.05 | 2.9129 | 2.79 | 3.05 | 2.82 | 3 | 2.68 | < 5 | 2.41 | 2.7514 | 1.8076 | 2.89 | 2.766 | 2.91 | 2.87 | 5 | 2.5 | 3.17 | 2.8 | 2.9 | 2.9 | 2.91 | 1
| 1 | 2 | < 5 | 2.54 | < 2 | 2.99 | 2.53 | 2.95 | 2.9211 | 2.86 | 2.86 | 2.87 | 2.98 | 2.78 | < 5 | 2.71 | 2.7542 | 1.7681 | 2.9 | 2.6022 | 2.8 | 2.9 | 7 | 2.5 | 3.03 | 2.54 | 2.98 | 2.9 | 2.84 | 4
| 1 | 3 | < 5 | 2.47 | < 2 | 3.04 | 2.54 | 2.98 | 2.9748 | 2.62 | 3.11 | 2.79 | 2.98 | 2.76 | < 5 | 2.48 | 2.7906 | 1.8117 | 2.88 | 2.6307 | 2.78 | 2.84 | 7 | 2.4 | 2.99 | 2.71 | 3.03 | 2.9 | 2.83 | 16
| 1 | 4 | < 5 | 2.54 | < 2 | 3.11 | 2.64 | 2.98 | 2.9026 | 2.62 | 3.09 | 2.83 | 2.97 | 2.8 | < 5 | 2.52 | 2.7283 | 1.7531 | 2.91 | 3.0558 | 2.75 | 2.88 | < 5 | 2.5 | 3.03 | 2.7 | 2.88 | 2.9 | 2.82 | 2.8176
| 1 | 5 | < 5 | 2.7 | < 2 | 2.99 | 2.58 | 2.99 | 2.9776 | 2.73 | 3.04 | 2.81 | 2.97 | 2.81 | < 5 | 2.64 | 2.6979 | 1.7488 | 2.9 | 2.806 | 2.81 | 2.76 | 7 | 2.4 | 3 | 2.68 | 3.02 | 2.9 | 2.83 | 155
| 1 | 6 | < 5 | 1.8 | < 2 | 3.01 | 2.53 | 2.98 | 2.9491 | 2.73 | 3.23 | 2.91 | 2.95 | 2.74 | < 5 | 2.68 | 2.6759 | 1.7273 | 2.9 | 2.8272 | 2.93 | 2.88 | 5 | 2.5 | 3.01 | 2.54 | 3.03 | 2.9 | 2.92
| Mean |  |  | 2.3983 |  | 3.0333 | 2.5833 | 2.9883 | 2.9397 | 2.725 | 3.0633 | 2.8383 | 2.975 | 2.7617 |  | 2.5733 | 2.733 | 1.7694 | 2.8967 | 2.7813 | 2.83 | 2.855 | 6.2 | 2.4667 | 3.0383 | 2.6617 | 2.9733 | 2.9 | 2.8583
| Median |  |  | 2.505 |  | 3.025 | 2.56 | 2.98 | 2.9351 | 2.73 | 3.07 | 2.825 | 2.975 | 2.77 |  | 2.58 | 2.7399 | 1.7606 | 2.9 | 2.786 | 2.805 | 2.875 | 7 | 2.5 | 3.02 | 2.69 | 3 | 2.9 | 2.835
| Std Dev. |  |  | 0.31549 |  | 0.046762 | 0.063456 | 0.033116 | 0.032232 | 0.094393 | 0.12061 | 0.044008 | 0.016432 | 0.047504 |  | 0.12061 | 0.041504 | 0.033807 | 0.010328 | 0.16314 | 0.072938 | 0.050498 | 1.0954 | 0.05164 | 0.066458 | 0.10284 | 0.067429 | 0 | 0.04446
| Rel.Std.Dev. |  |  | 13.15% |  | 1.542% | 2.456% | 1.108% | 1.096% | 3.464% | 3.937% | 1.55% | 0.5523% | 1.72% |  | 4.687% | 1.519% | 1.911% | 0.3565% | 5.865% | 2.577% | 1.769% | 17.67% | 2.094% | 2.187% | 3.864% | 2.268% | 0% | 1.555%
| PDM3 |  |  | -14.88% |  | 7.658% | -8.313% | 6.061% | 4.334% | -3.285% | 8.723% | 0.7372% | 5.588% | -1.984% |  | -8.668% | -2.999% | -37.2% | 2.808% | -1.287% | 0.4414% | 1.329% | 120% | -12.45% | 7.836% | -5.533% | 5.529% | 2.926% | 1.447%
| Z-Score (Absolute) |  | 1.41 | 1.84 | 7.79 | 0.87 | 1.05 | 0.67 | 0.47 | 0.45 | 0.99 | 0.04 | 0.62 | 0.29 | 1.41 | 1.09 | 0.41 | 4.52 | 0.28 | 0.21 | 0 | 0.11 | 11.73 | 1.55 | 0.89 | 0.72 | 0.61 | 0.3 | 0.12
| Analytical results for Ca in OREAS 556b (Certified Value 0.131 wt.%) | [28 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | 0.17 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.1346 | 0.136 | 0.12715 | 0.13 | 0.13114 | 0.1454 | 0.13 | 0.13 | 0.14 | 0.135 | 0.122 | 0.14 | 0.13 | 0.13 |  |  | 1
| 1 | 2 | 0.16 | 0.12 | 0.13 | 0.13 | 0.11 | 0.13 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.1324 | 0.136 | 0.12694 | 0.13 | 0.12921 | 0.147 | 0.13 | 0.13 | 0.14 | 0.131 | 0.118 | 0.14 | 0.13 | 0.13 |  |  | 0
| 1 | 3 | 0.16 | 0.13 | 0.13 | 0.13 | 0.11 | 0.13 | 0.13 | 0.14 | 0.13 | 0.14 | 0.13 | 0.1329 | 0.129 | 0.12899 | 0.13 | 0.12937 | 0.147 | 0.13 | 0.13 | 0.14 | 0.137 | 0.124 | 0.14 | 0.13 | 0.13 |  |  | 16
| 1 | 4 | 0.15 | 0.13 | 0.12 | 0.13 | 0.12 | 0.12 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.1335 | 0.129 | 0.13252 | 0.13 | 0.13986 | 0.1469 | 0.12 | 0.13 | 0.14 | 0.133 | 0.117 | 0.13 | 0.13 | 0.13 |  |  | 0.13063
| 1 | 5 | 0.15 | 0.13 | 0.13 | 0.13 | 0.11 | 0.13 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.1338 | 0.143 | 0.12383 | 0.13 | 0.13554 | 0.1445 | 0.13 | 0.13 | 0.14 | 0.131 | 0.115 | 0.14 | 0.13 | 0.13 |  |  | 156
| 1 | 6 | 0.14 | 0.14 | 0.13 | 0.13 | 0.12 | 0.13 | 0.13 | 0.14 | 0.14 | 0.13 | 0.13 | 0.1349 | 0.136 | 0.13282 | 0.14 | 0.1324 | 0.1497 | 0.13 | 0.13 | 0.13 | 0.132 | 0.12 | 0.14 | 0.13 | 0.13
| Mean |  | 0.155 | 0.13 | 0.12833 | 0.13 | 0.11667 | 0.12833 | 0.13 | 0.14 | 0.13167 | 0.13167 | 0.13 | 0.13368 | 0.13483 | 0.12871 | 0.13167 | 0.13292 | 0.14675 | 0.12833 | 0.13 | 0.13833 | 0.13317 | 0.11933 | 0.13833 | 0.13 | 0.13
| Median |  | 0.155 | 0.13 | 0.13 | 0.13 | 0.115 | 0.13 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.13365 | 0.136 | 0.12807 | 0.13 | 0.13177 | 0.14695 | 0.13 | 0.13 | 0.14 | 0.1325 | 0.119 | 0.14 | 0.13 | 0.13
| Std Dev. |  | 0.010488 | 0.0063246 | 0.0040825 | 0 | 0.008165 | 0.0040825 | 0 | 0 | 0.0040825 | 0.0040825 | 0 | 0.00096212 | 0.0052694 | 0.003488 | 0.0040825 | 0.0041205 | 0.001774 | 0.0040825 | 0 | 0.0040825 | 0.0024014 | 0.0033267 | 0.0040825 | 0 | 0
| Rel.Std.Dev. |  | 6.767% | 4.865% | 3.181% | 0% | 6.999% | 3.181% | 0% | 0% | 3.101% | 3.101% | 0% | 0.7197% | 3.908% | 2.71% | 3.101% | 3.1% | 1.209% | 3.181% | 0% | 2.951% | 1.803% | 2.788% | 2.951% | 0% | 0%
| PDM3 |  | 18.65% | -0.4852% | -1.761% | -0.4852% | -10.69% | -1.761% | -0.4852% | 7.17% | 0.7906% | 0.7906% | -0.4852% | 2.334% | 3.215% | -1.474% | 0.7906% | 1.748% | 12.34% | -1.761% | -0.4852% | 5.894% | 1.939% | -8.651% | 5.894% | -0.4852% | -0.4852%
| Z-Score (Absolute) |  | 10.11 | 0 | 0.67 | 0 | 5.39 | 0.67 | 0 | 4.05 | 0.67 | 0.67 | 0 | 1.49 | 1.86 | 0.52 | 0.67 | 1.18 | 6.78 | 0.67 | 0 | 3.37 | 1.28 | 4.32 | 3.37 | 0 | 0
| Analytical results for Cd in OREAS 556b (Certified Value 0.51 ppm) | [28 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | < 1 | 0.47 | < 0.5 | 0.5 | 0.51 | 0.51 | 0.39519 | 0.5 | 0.48 | 0.51 | 0.5 | 0.48 | < 5 | 0.53 | 0.43376 | 0.45 | 0.51393 | 0.52 | 1 | 0.5 | 0.45 | 0.47 | 0.49 | 0.5 | 0.5 |  |  | 1
| 1 | 2 | < 1 | 0.5 | < 0.5 | 0.49 | 0.55 | 0.49 | 0.56999 | 0.54 | 0.51 | 0.5 | 0.49 | 0.49 | < 5 | 0.55 | 0.45612 | 0.45 | 0.53858 | 0.52 | 2 | 0.5 | 0.45 | 0.43 | 0.54 | < 0.5 | 0.53 |  |  | 5
| 1 | 3 | < 1 | 0.47 | < 0.5 | 0.48 | 0.54 | 0.53 | 0.51772 | 0.49 | 0.53 | 0.48 | 0.49 | 0.49 | < 5 | 0.53 | 0.45891 | 0.45 | 0.52389 | 0.51 | 1 | 0.5 | 0.5 | 0.47 | 0.56 | 0.5 | 0.54 |  |  | 16
| 1 | 4 | < 1 | 0.49 | < 0.5 | 0.52 | 0.46 | 0.51 | 0.52032 | 0.5 | 0.48 | 0.49 | 0.5 | 0.51 | < 5 | 0.54 | 0.43547 | 0.46 | 0.58014 | 0.52 | 1 | 0.4 | 0.47 | 0.44 | 0.5 | < 0.5 | 0.53 |  |  | 0.51043
| 1 | 5 | < 1 | 0.54 | < 0.5 | 0.56 | 0.5 | 0.5 | 0.51016 | 0.5 | 0.5 | 0.5 | 0.5 | 0.51 | < 5 | 0.56 | 0.44236 | 0.46 | 0.57792 | 0.5 | 1 | 0.4 | 0.45 | 0.44 | 0.5 | < 0.5 | 0.55 |  |  | 157
| 1 | 6 | < 1 | 0.49 | < 0.5 | 0.47 | 0.51 | 0.5 | 0.55617 | 0.52 | 0.51 | 0.5 | 0.51 | 0.5 | < 5 | 0.53 | 0.42922 | 0.43 | 0.52543 | 0.51 | 1 | 0.5 | 0.47 | 0.42 | 0.59 | < 0.5 | 0.53
| Mean |  |  | 0.49333 |  | 0.50333 | 0.51167 | 0.50667 | 0.51159 | 0.50833 | 0.50167 | 0.49667 | 0.49833 | 0.49667 |  | 0.54 | 0.44264 | 0.45 | 0.54331 | 0.51333 | 1.1667 | 0.46667 | 0.465 | 0.445 | 0.53 | 0.5 | 0.53
| Median |  |  | 0.49 |  | 0.495 | 0.51 | 0.505 | 0.51902 | 0.5 | 0.505 | 0.5 | 0.5 | 0.495 |  | 0.535 | 0.43891 | 0.45 | 0.53201 | 0.515 | 1 | 0.5 | 0.46 | 0.44 | 0.52 | 0.5 | 0.53
| Std Dev. |  |  | 0.02582 |  | 0.03266 | 0.031885 | 0.013663 | 0.061746 | 0.018348 | 0.019408 | 0.010328 | 0.0075277 | 0.012111 |  | 0.012649 | 0.012302 | 0.010954 | 0.028766 | 0.008165 | 0.40825 | 0.05164 | 0.019748 | 0.020736 | 0.04 | 0 | 0.016733
| Rel.Std.Dev. |  |  | 5.234% |  | 6.489% | 6.232% | 2.697% | 12.07% | 3.61% | 3.869% | 2.079% | 1.511% | 2.438% |  | 2.342% | 2.779% | 2.434% | 5.295% | 1.591% | 34.99% | 11.07% | 4.247% | 4.66% | 7.547% | 0% | 3.157%
| PDM3 |  |  | -3.349% |  | -1.39% | 0.2426% | -0.737% | 0.2279% | -0.4104% | -1.717% | -2.696% | -2.37% | -2.696% |  | 5.794% | -13.28% | -11.84% | 6.443% | 0.5691% | 128.6% | -8.574% | -8.9% | -12.82% | 3.834% | -2.043% | 3.834%
| Z-Score (Absolute) |  | 0.17 | 0.62 | 17.03 | 0.06 | 0.62 | 0.28 | 0.61 | 0.39 | 0.06 | 0.39 | 0.28 | 0.39 | 134.69 | 2.53 | 4.04 | 3.54 | 2.75 | 0.73 | NA | NA | 2.53 | 3.88 | 1.85 | NA | 1.85
Results from laboratories 22 and 27 were removed due to their 0.1 ppm reading resolution. Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 556b (Certified Value 43.6 ppm) | [28 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 25 | 26 | 28 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | [7 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [7 empty] | 1
| 1 | 1 | 47.669 | 54.8 | 37.1 | 41.8 | 47.9 | 38.1 | 39.7 | 41.4 | 41.72 | 36.8 | 51.1102 | 21.4382 | 54.5 | 41.5243 | 33.9 | 41 | 43.42 | 64.5 | 56 | 41.55 | [7 empty] | 1
| 1 | 2 | 47.736 | 47.6 | 36.4 | 42.3 | 49.5 | 41.3 | 38.8 | 42.5 | 41.86 | 40.1 | 48.7149 | 20.5104 | 53.9 | 41.2512 | 32.9 | 41 | 41.75 | 73.9 | 56 | 40.64 | [7 empty] | 6
| 1 | 3 | 48.121 | 53 | 37.2 | 39.2 | 48.9 | 41.4 | 40.6 | 42.3 | 41.89 | 37.4 | 47.3965 | 21.8818 | 54.7 | 40.736 | 33.9 | 40 | 41.23 | 58.5 | 52 | 40.04 | [7 empty] | 16
| 1 | 4 | 49.323 | 54.2 | 35.4 | 39.8 | 49.1 | 40.5 | 38.7 | 44.2 | 41.92 | 38.1 | 47.2301 | 20.6676 | 54.8 | 46.1432 | 33.6 | 40 | 43.19 | 82.5 | 51 | 39.66 | [7 empty] | 43.5722
| 1 | 5 | 48.626 | 50.8 | 35.8 | 39.1 | 47.4 | 41.9 | 40 | 44.1 | 41.71 | 39.2 | 48.4282 | 20.7027 | 54.8 | 41.4135 | 34.1 | 41 | 41.98 | 70.6 | 53 | 38.89 | [7 empty] | 158
| 1 | 6 | 48.657 | 49.9 | 36.4 | 39.6 | 50 | 41.5 | 38.3 | 43.2 | 41.81 | 38.2 | 51.2218 | 20.5238 | 54.8 | 39.7656 | 34.2 | 40 | 41.69 | 73.8 | 54 | 41.16
| Mean |  | 48.3553 | 51.7167 | 36.3833 | 40.3 | 48.8 | 40.7833 | 39.35 | 42.95 | 41.8183 | 38.3 | 49.017 | 20.9541 | 54.5833 | 41.8056 | 33.7667 | 40.5 | 42.21 | 70.6333 | 53.6667 | 40.3233
| Median |  | 48.3735 | 51.9 | 36.4 | 39.7 | 49 | 41.35 | 39.25 | 42.85 | 41.835 | 38.15 | 48.5716 | 20.6851 | 54.75 | 41.3323 | 33.9 | 40.5 | 41.865 | 72.2 | 53.5 | 40.34
| Std Dev. |  | 0.63406 | 2.7716 | 0.70545 | 1.3885 | 0.9798 | 1.392 | 0.88713 | 1.0932 | 0.087958 | 1.1967 | 1.7607 | 0.56962 | 0.35449 | 2.2208 | 0.47188 | 0.54772 | 0.88542 | 8.3229 | 2.0656 | 0.9879
| Rel.Std.Dev. |  | 1.311% | 5.359% | 1.939% | 3.445% | 2.008% | 3.413% | 2.254% | 2.545% | 0.2103% | 3.124% | 3.592% | 2.718% | 0.6495% | 5.312% | 1.397% | 1.352% | 2.098% | 11.78% | 3.849% | 2.45%
| PDM3 |  | 10.98% | 18.69% | -16.5% | -7.51% | 12% | -6.4% | -9.69% | -1.428% | -4.025% | -12.1% | 12.5% | -51.91% | 25.27% | -4.054% | -22.5% | -7.051% | -3.126% | 62.11% | 23.17% | -7.456%
| Z-Score (Absolute) |  | 0.99 | 1.49 | 0.82 | 0.23 | 1.05 | 0.16 | 0.37 | 0.17 | 0 | 0.53 | 1.09 | 3.15 | 1.93 | 0 | 1.21 | 0.2 | 0.06 | 4.35 | 1.79 | 0.22
| Analytical results for Co in OREAS 556b (Certified Value 1.98 wt.%) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | AR*AAS | AR*AAS | AR*AAS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*AAS | AR*AAS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.4g | 0.4g | 0.4g | 0.4g | 0.4g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.4g |  | 3
| 1 | 1 | 2.3 | 2.08 | 2.07 | 1.9295 | 2 | 2.01 | 1.8317 | 1.95 | 1.965 | 1.93 | 2.1 | 1.915 | 1.9755 | 1.81 | 2.1097 | > 0.5 | 2.0026 | 1.928 | 2.02 | 2.01 | > 1 | > 1 | 1.85 | > 0.5 | 1.94 | 1.99 |  | 1
| 1 | 2 | 2.28 | 2.05 | 2.08 | 1.9296 | 1.99 | 1.94 | 1.8028 | 1.96 | 1.955 | 1.915 | 2.1 | 1.975 | 1.9898 | 1.88 | 2.1323 | > 0.5 | 1.958 | 1.9778 | 2.03 | 1.99 | > 1 | > 1 | 1.77 | > 0.5 | 1.9234 | 1.96 |  | 36
| 1 | 3 | 2.34 | 2.08 | 2.06 | 1.9687 | 2 | 1.96 | 1.7232 | 1.985 | 1.93 | 1.935 | 2.06 | 1.965 | 1.9976 | 1.86 | 2.045 | > 0.5 | 1.9987 | 1.9581 | 2.01 | 2 | > 1 | > 1 | 1.81 | > 0.5 | 1.9282 | 2.07 |  | 16
| 1 | 4 | 2.23 | 2.09 | 2.11 | 1.9919 | 2.01 | 1.93 | 1.7592 | 1.955 | 1.93 | 1.94 | 2.1 | 1.95 | 1.9963 | 1.85 | 2.0396 | > 0.5 | 1.9804 | 1.9458 | 2.02 | 1.99 | > 1 | > 1 | 1.87 | > 0.5 | 1.9359 | 2.109 |  | 1.9765
| 1 | 5 | 2.43 | 2.04 | 2.1 | 1.9059 | 1.99 | 1.99 | 1.7872 | 1.96 | 1.875 | 1.93 | 2.1 | 1.9 | 1.9978 | 1.87 | 2.0447 | > 0.5 | 2.0287 | 1.9077 | 2 | 2.01 | > 1 | > 1 | 1.81 | > 0.5 | 1.9301 | 2.016 |  | 159
| 1 | 6 | 2.35 | 2.04 | 2.11 | 1.9345 | 1.99 | 1.95 | 1.7453 | 1.985 | 1.91 | 1.895 | 2.1 | 1.915 | 1.9975 | 1.81 | 2.1024 | > 0.5 | 1.9437 | 1.9539 | 2.02 | 1.98 | > 1 | > 1 | 1.78 | > 0.5 | 1.9074 | 1.99
| Mean |  | 2.3217 | 2.0633 | 2.0883 | 1.9433 | 1.9967 | 1.9633 | 1.7749 | 1.9658 | 1.9275 | 1.9242 | 2.0933 | 1.9367 | 1.9924 | 1.8467 | 2.0789 |  | 1.9853 | 1.9452 | 2.0167 | 1.9967 |  |  | 1.815 |  | 1.9275 | 2.0225
| Median |  | 2.32 | 2.065 | 2.09 | 1.9321 | 1.995 | 1.955 | 1.7732 | 1.96 | 1.93 | 1.93 | 2.1 | 1.9325 | 1.9969 | 1.855 | 2.0737 |  | 1.9896 | 1.9499 | 2.02 | 1.995 |  |  | 1.81 |  | 1.9292 | 2.003
| Std Dev. |  | 0.068532 | 0.022509 | 0.02137 | 0.031182 | 0.008165 | 0.030768 | 0.039867 | 0.015303 | 0.032365 | 0.016558 | 0.01633 | 0.030768 | 0.0088294 | 0.030111 | 0.040552 |  | 0.031186 | 0.024525 | 0.010328 | 0.012111 |  |  | 0.038859 |  | 0.011451 | 0.056213
| Rel.Std.Dev. |  | 2.952% | 1.091% | 1.023% | 1.605% | 0.4089% | 1.567% | 2.246% | 0.7784% | 1.679% | 0.8605% | 0.7801% | 1.589% | 0.4431% | 1.631% | 1.951% |  | 1.571% | 1.261% | 0.5121% | 0.6065% |  |  | 2.141% |  | 0.5941% | 2.779%
| PDM3 |  | 17.47% | 4.395% | 5.66% | -1.676% | 1.022% | -0.6648% | -10.2% | -0.5383% | -2.478% | -2.646% | 5.913% | -2.014% | 0.8067% | -6.568% | 5.184% |  | 0.4491% | -1.581% | 2.034% | 1.022% |  |  | -8.17% |  | -2.477% | 2.329%
| Z-Score (Absolute) |  | 4.91 | 1.25 | 1.6 | 0.46 | 0.3 | 0.17 | 2.85 | 0.14 | 0.68 | 0.73 | 1.67 | 0.55 | 0.24 | 1.83 | 1.47 | NA | 0.14 | 0.43 | 0.58 | 0.3 | NA | NA | 2.28 | NA | 0.68 | 0.67
| Analytical results for Cr in OREAS 556b (Certified Value 35 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 2
| 1 | 1 | 30.9 | 37.4 | 38 | 39 | 30.6 | 36 | 36 | 34 | 34 | 36 | 34.35 | 36 | 49.641 | 49 | 32.561 | 48 | 34 | 32 | 33 | 43 | 45 | 37.4 | 38 |  |  |  |  | 1
| 1 | 2 | 32 | 37.3 | 38 | 33 | 30.7 | 37 | 37 | 35 | 33 | 35 | 33.88 | 38 | 49.9798 | 48 | 32.6315 | 47 | 35 | 33 | 33 | 44 | 45 | 36.6 | 38 |  |  |  |  | 37
| 1 | 3 | 31.8 | 36 | 38 | 35 | 30.1 | 36 | 39 | 35 | 34 | 35 | 34.51 | 37 | 52.0638 | 47 | 33.2339 | 46 | 35 | 32 | 35 | 42 | 43 | 37.1 | 37 |  |  |  |  | 16
| 1 | 4 | 30.9 | 35.7 | 39 | 37 | 28.7 | 35 | 37 | 35 | 34 | 35 | 34.26 | 37 | 54.6428 | 49 | 35.6153 | 44 | 34 | 32 | 36 | 47 | 42 | 37.8 | 37 |  |  |  |  | 35.0105
| 1 | 5 | 33 | 36.1 | 38 | 35 | 30.2 | 35 | 37 | 36 | 34 | 36 | 34.06 | 41 | 51.4999 | 49 | 33.3899 | 44 | 35 | 31 | 34 | 44 | 44 | 37.4 | 37 |  |  |  |  | 160
| 1 | 6 | 34.4 | 36.7 | 39 | 35 | 30.7 | 35 | 38 | 36 | 33 | 35 | 34.08 | 36 | 53.4675 | 49 | 33.547 | 44 | 35 | 31 | 37 | 45 | 44 | 37.1 | 41
| Mean |  | 32.1667 | 36.5333 | 38.3333 | 35.6667 | 30.1667 | 35.6667 | 37.3333 | 35.1667 | 33.6667 | 35.3333 | 34.19 | 37.5 | 51.8825 | 48.5 | 33.4964 | 45.5 | 34.6667 | 31.8333 | 34.6667 | 44.1667 | 43.8333 | 37.2333 | 38
| Median |  | 31.9 | 36.4 | 38 | 35 | 30.4 | 35.5 | 37 | 35 | 34 | 35 | 34.17 | 37 | 51.7818 | 49 | 33.3119 | 45 | 35 | 32 | 34.5 | 44 | 44 | 37.25 | 37.5
| Std Dev. |  | 1.3456 | 0.71181 | 0.5164 | 2.0656 | 0.76333 | 0.8165 | 1.0328 | 0.75277 | 0.5164 | 0.5164 | 0.22716 | 1.8708 | 1.9471 | 0.83666 | 1.1131 | 1.7607 | 0.5164 | 0.75277 | 1.633 | 1.7224 | 1.169 | 0.40332 | 1.5492
| Rel.Std.Dev. |  | 4.183% | 1.948% | 1.347% | 5.791% | 2.53% | 2.289% | 2.766% | 2.141% | 1.534% | 1.462% | 0.6644% | 4.989% | 3.753% | 1.725% | 3.323% | 3.87% | 1.49% | 2.365% | 4.711% | 3.9% | 2.667% | 1.083% | 4.077%
| PDM3 |  | -8.123% | 4.35% | 9.491% | 1.874% | -13.84% | 1.874% | 6.635% | 0.446% | -3.838% | 0.922% | -2.344% | 7.111% | 48.19% | 38.53% | -4.325% | 29.96% | -0.9821% | -9.075% | -0.9821% | 26.15% | 25.2% | 6.349% | 8.539%
| Z-Score (Absolute) |  | 1.18 | 0.29 | 0.9 | 0 | 1.85 | 0 | 0.56 | 0.17 | 0.67 | 0.11 | 0.5 | 0.62 | 5.47 | 4.33 | 0.73 | 3.32 | 0.34 | 1.29 | 0.34 | 2.87 | 2.75 | 0.53 | 0.79
| Analytical results for Cs in OREAS 556b (Certified Value 0.29 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 25 | 26 | 28 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [8 empty] | 2
| 1 | 1 | 0.35 | 0.36 | 0.25 | 0.18 | 0.21 | 0.19 | 0.21 | 0.19 | 0.14 | 0.40639 | 0.30273 | 0.49 | 0.27904 | 0.32 | 0.2 | 0.33 | 0.5 | 0.52 | 0.3 | [8 empty] | 1
| 1 | 2 | 0.34 | 0.35 | 0.24 | 0.19 | 0.21 | 0.21 | 0.22 | 0.18 | 0.14 | 0.39247 | 0.313 | 0.48 | 0.26854 | 0.32 | 0.2 | 0.31 | 0.5 | 0.48 | 0.28 | [8 empty] | 9
| 1 | 3 | 0.36 | 0.35 | 0.26 | 0.2 | 0.21 | 0.21 | 0.24 | 0.19 | 0.14 | 0.40976 | 0.30187 | 0.47 | 0.26765 | 0.34 | 0.2 | 0.31 | 0.5 | 0.46 | 0.3 | [8 empty] | 16
| 1 | 4 | 0.39 | 0.37 | 0.24 | 0.19 | 0.22 | 0.2 | 0.22 | 0.19 | 0.13 | 0.4253 | 0.28785 | 0.48 | 0.28374 | 0.33 | 0.2 | 0.32 | 0.6 | 0.45 | 0.27 | [8 empty] | 0.2938
| 1 | 5 | 0.37 | 0.34 | 0.25 | 0.19 | 0.21 | 0.22 | 0.25 | 0.19 | 0.14 | 0.40431 | 0.3082 | 0.46 | 0.29409 | 0.35 | 0.2 | 0.32 | 0.5 | 0.45 | 0.26 | [8 empty] | 161
| 1 | 6 | 0.35 | 0.37 | 0.25 | 0.2 | 0.22 | 0.2 | 0.22 | 0.19 | 0.14 | 0.43055 | 0.30036 | 0.16 | 0.26541 | 0.33 | 0.2 | 0.3 | 0.5 | 0.48 | 0.27
| Mean |  | 0.36 | 0.35667 | 0.24833 | 0.19167 | 0.21333 | 0.205 | 0.22667 | 0.18833 | 0.13833 | 0.41146 | 0.30234 | 0.42333 | 0.27641 | 0.33167 | 0.2 | 0.315 | 0.51667 | 0.47333 | 0.28
| Median |  | 0.355 | 0.355 | 0.25 | 0.19 | 0.21 | 0.205 | 0.22 | 0.19 | 0.14 | 0.40807 | 0.3023 | 0.475 | 0.27379 | 0.33 | 0.2 | 0.315 | 0.5 | 0.47 | 0.275
| Std Dev. |  | 0.017889 | 0.012111 | 0.0075277 | 0.0075277 | 0.005164 | 0.010488 | 0.015055 | 0.0040825 | 0.0040825 | 0.014115 | 0.0085127 | 0.12941 | 0.011252 | 0.01169 | 3.0405e-17 | 0.010488 | 0.040825 | 0.026583 | 0.016733
| Rel.Std.Dev. |  | 4.969% | 3.395% | 3.031% | 3.928% | 2.421% | 5.116% | 6.642% | 2.168% | 2.951% | 3.43% | 2.816% | 30.57% | 4.071% | 3.525% | 1.52e-14% | 3.33% | 7.902% | 5.616% | 5.976%
| PDM3 |  | 22.53% | 21.4% | -15.47% | -34.76% | -27.39% | -30.22% | -22.85% | -35.9% | -52.92% | 40.05% | 2.906% | 44.09% | -5.917% | 12.89% | -31.93% | 7.217% | 75.86% | 61.11% | -4.696%
| Z-Score (Absolute) |  | 0.63 | 0.6 | 0.39 | 0.91 | 0.72 | 0.79 | 0.59 | 0.95 | 1.41 | 1.11 | 0.1 | 1.22 | 0.14 | 0.37 | NA | 0.22 | 2.07 | 1.68 | 0.1
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 556b (Certified Value 8.55 wt.%) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | AR*AAS | AR*AAS | AR*AAS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*AAS | AR*AAS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.4g |  |  | 3
| 1 | 1 | 8.53 | 8.453 | 8.21 | 8.4593 | 8.61 | 8.47 | 8.4785 | 8.46 | 8.64 | 8.54 | 8.8 | 8.46 | 8.7788 | 8.62 | 9.5795 | > 2 | 9.1356 | 8.3833 | 8.78 | 8.69 | > 1 | 8.46 | > 1 | 8.0644 | 8.348 |  |  | 1
| 1 | 2 | 8.54 | 8.491 | 8.26 | 8.5395 | 8.57 | 8.36 | 8.4251 | 8.46 | 8.65 | 8.5 | 8.71 | 8.43 | 8.7593 | 8.9 | 9.643 | > 2 | 9.0046 | 8.4294 | 8.81 | 8.72 | > 1 | 7.8 | > 1 | 8.0128 | 8.332 |  |  | 39
| 1 | 3 | 8.53 | 8.491 | 8.16 | 8.3299 | 8.55 | 8.4 | 8.4227 | 8.47 | 8.59 | 8.55 | 8.67 | 8.54 | 8.7769 | 8.93 | 9.8119 | > 2 | 9.2692 | 8.3824 | 8.65 | 8.78 | > 1 | 8.6 | > 1 | 8.0226 | 8.744 |  |  | 16
| 1 | 4 | 8.53 | 8.512 | 8.18 | 8.6188 | 8.56 | 8.2 | 8.5412 | 8.52 | 8.68 | 8.57 | 8.8 | 8.48 | 8.7752 | 8.91 | 8.5335 | > 2 | 9.0835 | 8.4802 | 8.71 | 8.71 | > 1 | 8.57 | > 1 | 8.0748 | 8.673 |  |  | 8.5536
| 1 | 5 | 8.54 | 8.872 | 8.16 | 8.3225 | 8.52 | 8.65 | 8.4055 | 8.48 | 8.45 | 8.7 | 8.88 | 8.48 | 8.7746 | 8.98 | 8.1953 | > 2 | 9.2614 | 8.3517 | 8.66 | 8.69 | > 1 | 8.2 | 8.6794 | 8.0305 | 8.557 |  |  | 162
| 1 | 6 | 8.54 | 8.41 | 8.18 | 8.5068 | 8.51 | 8.44 | 8.5208 | 8.49 | 8.51 | 8.4 | 8.6 | 8.4 | 8.7783 | 8.83 | 9.7613 | > 2 | 9.1827 | 8.4985 | 8.76 | 8.7 | > 1 | 8.13 | > 1 | 8.0365 | 8.438
| Mean |  | 8.535 | 8.5382 | 8.1917 | 8.4628 | 8.5533 | 8.42 | 8.4656 | 8.48 | 8.5867 | 8.5433 | 8.7433 | 8.465 | 8.7738 | 8.8617 | 9.2541 |  | 9.1562 | 8.4209 | 8.7283 | 8.715 |  | 8.2933 | 8.6794 | 8.0402 | 8.5153
| Median |  | 8.535 | 8.491 | 8.18 | 8.4831 | 8.555 | 8.42 | 8.4518 | 8.475 | 8.615 | 8.545 | 8.755 | 8.47 | 8.776 | 8.905 | 9.6112 |  | 9.1591 | 8.4063 | 8.735 | 8.705 |  | 8.33 | 8.6794 | 8.0335 | 8.4975
| Std Dev. |  | 0.0054772 | 0.16749 | 0.038166 | 0.11791 | 0.036148 | 0.14711 | 0.056584 | 0.022804 | 0.089592 | 0.097707 | 0.10211 | 0.048062 | 0.007317 | 0.12797 | 0.70224 |  | 0.10318 | 0.058814 | 0.065549 | 0.033912 |  | 0.30891 |  | 0.024291 | 0.17116
| Rel.Std.Dev. |  | 0.06417% | 1.962% | 0.4659% | 1.393% | 0.4226% | 1.747% | 0.6684% | 0.2689% | 1.043% | 1.144% | 1.168% | 0.5678% | 0.0834% | 1.444% | 7.588% |  | 1.127% | 0.6984% | 0.751% | 0.3891% |  | 3.725% |  | 0.3021% | 2.01%
| PDM3 |  | -0.2171% | -0.1801% | -4.231% | -1.061% | -0.002766% | -1.562% | -1.028% | -0.8601% | 0.3869% | -0.1197% | 2.219% | -1.035% | 2.575% | 3.602% | 8.19% |  | 7.045% | -1.551% | 2.043% | 1.887% |  | -3.042% | 1.471% | -6.001% | -0.447%
| Z-Score (Absolute) |  | 0.02 | 0 | 1.98 | 0.43 | 0.09 | 0.67 | 0.41 | 0.33 | 0.28 | 0.03 | 1.17 | 0.42 | 1.34 | 1.85 | 4.09 | NA | 3.53 | 0.67 | 1.09 | 1.01 | NA | 1.4 | 0.81 | 2.84 | 0.13
| Analytical results for Dy in OREAS 556b (Certified Value 1.02 ppm) | [28 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 26 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS |  |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 01g |  |  | 2
| 1 | 1 | 01: 0.993; 02: 1.1; 03: 0.9; 04: 0.65713; 05: 0.93539; 06: 1.21 |  |  | 1
| 1 | 2 | 01: 0.911; 02: 0.9; 03: 0.93; 04: 0.64228; 05: 0.88507; 06: 1.21 |  |  | 10
| 1 | 3 | 01: 1; 02: 1; 03: 0.93; 04: 0.68418; 05: 0.89678; 06: 1.16 |  |  | 16
| 1 | 4 | 01: 1.048; 02: 1.1; 03: 0.92; 04: 0.66749; 05: 0.93539; 06: 1.19 |  |  | 1.0157
| 1 | 5 | 01: 1.043; 02: 1.1; 03: 0.95; 04: 0.64247; 05: 0.8961; 06: 1.19 |  |  | 163
| 1 | 6 | 01: 1.003; 02: 1; 03: 0.95; 04: 0.6348; 05: 1.0045; 06: 1.18
| Mean |  | 01: 0.99967; 02: 1.0333; 03: 0.93; 04: 0.65472; 05: 0.92553; 06: 1.19
| Median |  | 01: 1.0015; 02: 1.05; 03: 0.93; 04: 0.6498; 05: 0.91609; 06: 1.19
| Std Dev. |  | 01: 0.049257; 02: 0.08165; 03: 0.018974; 04: 0.018646; 05: 0.044166; 06: 0.018974
| Rel.Std.Dev. |  | 01: 4.927%; 02: 7.902%; 03: 2.04%; 04: 2.848%; 05: 4.772%; 06: 1.594%
| PDM3 |  | 01: -1.579%; 02: 1.735%; 03: -8.438%; 04: -35.54%; 05: -8.878%; 06: 17.16%
| Z-Score (Absolute) |  | 01: 0.44; 02: 0.86; 03: 0.44; 04: 3.88; 05: 0.49; 06: 2.82
| Analytical results for Er in OREAS 556b (Certified Value 0.41 ppm) | [28 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 26 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS |  |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 01g |  |  | 2
| 1 | 1 | 01: 0.423; 02: 0.5; 03: 0.41; 04: 0.26079; 05: 0.40838; 06: 0.55 |  |  | 1
| 1 | 2 | 01: 0.416; 02: 0.4; 03: 0.41; 04: 0.25193; 05: 0.39899; 06: 0.55 |  |  | 11
| 1 | 3 | 01: 0.445; 02: 0.5; 03: 0.4; 04: 0.27538; 05: 0.40803; 06: 0.52 |  |  | 16
| 1 | 4 | 01: 0.494; 02: 0.5; 03: 0.42; 04: 0.26777; 05: 0.43071; 06: 0.53 |  |  | 0.41235
| 1 | 5 | 01: 0.437; 02: 0.5; 03: 0.43; 04: 0.27144; 05: 0.38634; 06: 0.52 |  |  | 164
| 1 | 6 | 01: 0.458; 02: 0.5; 03: 0.41; 04: 0.24786; 05: 0.40999; 06: 0.53
| Mean |  | 01: 0.4455; 02: 0.48333; 03: 0.41333; 04: 0.26253; 05: 0.40707; 06: 0.53333
| Median |  | 01: 0.441; 02: 0.5; 03: 0.41; 04: 0.26428; 05: 0.40821; 06: 0.53
| Std Dev. |  | 01: 0.028134; 02: 0.040825; 03: 0.010328; 04: 0.010976; 05: 0.014595; 06: 0.013663
| Rel.Std.Dev. |  | 01: 6.315%; 02: 8.447%; 03: 2.499%; 04: 4.181%; 05: 3.585%; 06: 2.562%
| PDM3 |  | 01: 8.038%; 02: 17.21%; 03: 0.2376%; 04: -36.33%; 05: -1.28%; 06: 29.34%
| Z-Score (Absolute) |  | 01: 0.67; 02: NA; 03: 0; 04: 3.16; 05: 0.13; 06: 2.52
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 556b (Certified Value 0.38 ppm) | [28 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 26 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS |  |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 01g |  |  | 3
| 1 | 1 | 01: 0.374; 02: 0.4; 03: 0.35; 04: 0.18947; 05: 0.35783; 06: 0.44 |  |  | 1
| 1 | 2 | 01: 0.332; 02: 0.3; 03: 0.36; 04: 0.1783; 05: 0.34936; 06: 0.44 |  |  | 12
| 1 | 3 | 01: 0.394; 02: 0.4; 03: 0.36; 04: 0.19397; 05: 0.35763; 06: 0.42 |  |  | 16
| 1 | 4 | 01: 0.408; 02: 0.4; 03: 0.36; 04: 0.17178; 05: 0.3855; 06: 0.42 |  |  | 0.38311
| 1 | 5 | 01: 0.421; 02: 0.4; 03: 0.36; 04: 0.1849; 05: 0.35729; 06: 0.42 |  |  | 165
| 1 | 6 | 01: 0.394; 02: 0.4; 03: 0.37; 04: 0.17733; 05: 0.35301; 06: 0.44
| Mean |  | 01: 0.38717; 02: 0.38333; 03: 0.36; 04: 0.18262; 05: 0.3601; 06: 0.43
| Median |  | 01: 0.394; 02: 0.4; 03: 0.36; 04: 0.1816; 05: 0.35746; 06: 0.43
| Std Dev. |  | 01: 0.031269; 02: 0.040825; 03: 0.0063246; 04: 0.0083056; 05: 0.012884; 06: 0.010954
| Rel.Std.Dev. |  | 01: 8.076%; 02: 10.65%; 03: 1.757%; 04: 4.548%; 05: 3.578%; 06: 2.548%
| PDM3 |  | 01: 1.06%; 02: 0.05958%; 03: -6.031%; 04: -52.33%; 05: -6.004%; 06: 12.24%
| Z-Score (Absolute) |  | 01: 0.77; 02: 0.58; 03: 0.58; 04: 9.39; 05: 0.58; 06: 2.89
| Analytical results for Fe in OREAS 556b (Certified Value 6.65 wt.%) | [28 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 3
| 1 | 1 | 6.82 | 5.79 | 6.42 | 6.86 | 5.94 | 6.61 | 6.81 | 6.65 | 6.19 | 6.74 | 6.64 | 6.6638 | 7.344 | 6.6406 | 7.12 | > 5 | 6.4501 | 6.49 | 6.69 | 6.11 | 7.15 | 6.5 | 6.57 | 6.61 | 6.91 |  |  | 1
| 1 | 2 | 6.7 | 5.68 | 6.42 | 6.84 | 5.09 | 6.44 | 6.85 | 6.82 | 6.33 | 6.79 | 6.81 | 6.6677 | 7.484 | 6.6412 | 6.99 | > 5 | 6.5466 | 6.51 | 6.76 | 6.21 | 7.04 | 6.31 | 6.57 | 6.6 | 6.82 |  |  | 0
| 1 | 3 | 6.7 | 5.97 | 6.39 | 6.79 | 5.45 | 6.43 | 6.59 | 6.78 | 6.23 | 6.71 | 6.78 | 6.6676 | 6.994 | 6.747 | 7.34 | > 5 | 6.5164 | 6.49 | 6.71 | 6.12 | 6.95 | 6.45 | 6.39 | 6.61 | 6.82 |  |  | 16
| 1 | 4 | 6.72 | 5.88 | 6.42 | 7.02 | 5.68 | 6.08 | 6.69 | 6.81 | 6.34 | 6.72 | 6.71 | 6.6672 | 7.064 | 6.3686 | 7.37 | > 5 | 6.5235 | 6.47 | 6.73 | 6 | 7.14 | 6.36 | 6.23 | 6.66 | 6.79 |  |  | 6.6451
| 1 | 5 | 6.84 | 5.93 | 6.31 | 6.73 | 5.53 | 6.51 | 6.7 | 6.75 | 6.3 | 6.78 | 6.83 | 6.6678 | 7.764 | 6.4609 | 7.13 | > 5 | 6.4465 | 6.53 | 6.65 | 6.08 | 6.94 | 6.28 | 6.55 | 6.6 | 6.83 |  |  | 166
| 1 | 6 | 6.77 | 6.08 | 6.39 | 6.94 | 5.39 | 6.61 | 6.73 | 6.84 | 6.49 | 6.68 | 6.75 | 6.667 | 7.274 | 6.8491 | 7.13 | > 5 | 6.5242 | 6.55 | 6.69 | 5.95 | 7.04 | 6.4 | 6.52 | 6.62 | 6.97
| Mean |  | 6.7583 | 5.8883 | 6.3917 | 6.8633 | 5.5133 | 6.4467 | 6.7283 | 6.775 | 6.3133 | 6.7367 | 6.7533 | 6.6669 | 7.3207 | 6.6179 | 7.18 |  | 6.5012 | 6.5067 | 6.705 | 6.0783 | 7.0433 | 6.3833 | 6.4717 | 6.6167 | 6.8567
| Median |  | 6.745 | 5.905 | 6.405 | 6.85 | 5.49 | 6.475 | 6.715 | 6.795 | 6.315 | 6.73 | 6.765 | 6.6674 | 7.309 | 6.6409 | 7.13 |  | 6.5199 | 6.5 | 6.7 | 6.095 | 7.04 | 6.38 | 6.535 | 6.61 | 6.825
| Std Dev. |  | 0.061455 | 0.1402 | 0.042622 | 0.10405 | 0.28584 | 0.19603 | 0.09261 | 0.06892 | 0.10443 | 0.042269 | 0.070048 | 0.0015255 | 0.28232 | 0.17776 | 0.14588 |  | 0.042242 | 0.029439 | 0.037815 | 0.092394 | 0.089592 | 0.083586 | 0.1363 | 0.022509 | 0.068605
| Rel.Std.Dev. |  | 0.9093% | 2.381% | 0.6668% | 1.516% | 5.185% | 3.041% | 1.376% | 1.017% | 1.654% | 0.6274% | 1.037% | 0.02288% | 3.857% | 2.686% | 2.032% |  | 0.6498% | 0.4524% | 0.564% | 1.52% | 1.272% | 1.309% | 2.106% | 0.3402% | 1.001%
| PDM3 |  | 1.703% | -11.39% | -3.814% | 3.284% | -17.03% | -2.987% | 1.252% | 1.954% | -4.993% | 1.377% | 1.628% | 0.3267% | 10.17% | -0.41% | 8.049% |  | -2.166% | -2.084% | 0.9008% | -8.53% | 5.992% | -3.94% | -2.611% | -0.4285% | 3.183%
| Z-Score (Absolute) |  | 0.43 | 2.78 | 0.92 | 0.81 | 4.16 | 0.72 | 0.32 | 0.49 | 1.21 | 0.35 | 0.41 | 0.09 | 2.5 | 0.09 | 1.98 | NA | 0.52 | 0.5 | 0.23 | 2.08 | 1.48 | 0.95 | 0.63 | 0.09 | 0.79
| Analytical results for Ga in OREAS 556b (Certified Value 9.61 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  | 2
| 1 | 1 | < 5 | 8.3 | 12.9 | 10.4 | 11.8 | 8.56 | 8.99 | 10.05 | 8.79 | 9.16 | 8.76 | 7.88 | 7.51 | 11.8021 | 6.7498 | 15.8 | 9.0277 | 8.31 | < 10 | 11 | 11.4 | 12 | 15.39 | 9 | 11.05 |  |  | 1
| 1 | 2 | < 5 | 9.7 | 12.9 | 10.7 | 9.95 | 8.47 | 9.14 | 10.55 | 9.16 | 9.15 | 8.93 | 8.04 | 7.74 | 11.1232 | 6.4007 | 15.8 | 9.1848 | 8.29 | < 10 | 11 | 10.8 | 13 | 15.78 | 9 | 10.64 |  |  | 13
| 1 | 3 | < 5 | 8.5 | 12.6 | 10.69 | 11 | 8.87 | 8.46 | 10.1 | 9.13 | 9.37 | 8.79 | 8.06 | 7.57 | 10.5785 | 6.5544 | 16.2 | 9.127 | 8.3 | < 10 | 11 | 10.8 | 11.4 | 15.4 | 9 | 10.45 |  |  | 16
| 1 | 4 | < 5 | 8.1 | 11.8 | 10.66 | 11.6 | 8.4 | 8.55 | 10.3 | 9.16 | 9.42 | 9.04 | 7.95 | 7.74 | 10.4217 | 6.3229 | 15.8 | 9.9167 | 8.22 | < 10 | 11 | 11.3 | 15.1 | 15.23 | 9 | 10.43 |  |  | 9.6132
| 1 | 5 | < 5 | 9.5 | 12.3 | 10.6 | 11.3 | 8.44 | 8.47 | 10.3 | 9.06 | 9.43 | 9.02 | 7.94 | 7.84 | 11.2152 | 6.4723 | 16.4 | 9.8782 | 8.38 | < 10 | 10 | 10.8 | 12.8 | 15.53 | 9 | 10.15 |  |  | 167
| 1 | 6 | < 5 | 7.1 | 12.3 | 10.41 | 11.1 | 8.61 | 8.8 | 10.4 | 9.34 | 9.22 | 8.79 | 7.96 | 7.66 | 12.1165 | 6.2321 | 15.7 | 9.1798 | 8.29 | < 10 | 11 | 11 | 13 | 15.78 | 9 | 10.69
| Mean |  |  | 8.5333 | 12.4667 | 10.5767 | 11.125 | 8.5583 | 8.735 | 10.2833 | 9.1067 | 9.2917 | 8.8883 | 7.9717 | 7.6767 | 11.2095 | 6.4554 | 15.95 | 9.3857 | 8.2983 |  | 10.8333 | 11.0167 | 12.8833 | 15.5183 | 9 | 10.5683
| Median |  |  | 8.4 | 12.45 | 10.63 | 11.2 | 8.515 | 8.675 | 10.3 | 9.145 | 9.295 | 8.86 | 7.955 | 7.7 | 11.1692 | 6.4365 | 15.8 | 9.1823 | 8.295 |  | 11 | 10.9 | 12.9 | 15.465 | 9 | 10.545
| Std Dev. |  |  | 0.95847 | 0.42269 | 0.1375 | 0.64942 | 0.17128 | 0.28752 | 0.18619 | 0.18063 | 0.12983 | 0.12481 | 0.067057 | 0.12176 | 0.66335 | 0.18286 | 0.28107 | 0.40058 | 0.051153 |  | 0.40825 | 0.27142 | 1.2592 | 0.22391 | 0 | 0.30334
| Rel.Std.Dev. |  |  | 11.23% | 3.391% | 1.3% | 5.838% | 2.001% | 3.292% | 1.811% | 1.983% | 1.397% | 1.404% | 0.8412% | 1.586% | 5.918% | 2.833% | 1.762% | 4.268% | 0.6164% |  | 3.768% | 2.464% | 9.774% | 1.443% | 0% | 2.87%
| PDM3 |  |  | -11.23% | 29.68% | 10.02% | 15.73% | -10.97% | -9.135% | 6.971% | -5.269% | -3.345% | -7.54% | -17.08% | -20.14% | 16.61% | -32.85% | 65.92% | -2.366% | -13.68% |  | 12.69% | 14.6% | 34.02% | 61.43% | -6.379% | 9.936%
| Z-Score (Absolute) |  | 3.47 | 0.39 | 1.62 | 0.66 | 0.94 | 0.37 | 0.28 | 0.51 | 0.09 | 0 | 0.21 | 0.67 | 0.83 | 0.98 | 1.45 | 3.4 | 0.05 | 0.51 | 2.19 | NA | 0.88 | 1.83 | 3.18 | NA | 0.65
Results from laboratories 22 and 27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 556b (Certified Value 1.91 ppm) | [28 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 26 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS |  |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 01g |  |  | 2
| 1 | 1 | 01: 1.9; 02: 2.1; 03: 1.56; 04: 1.8113; 05: 2.0764; 06: 2.11 |  |  | 1
| 1 | 2 | 01: 1.897; 02: 1.9; 03: 1.6; 04: 1.7186; 05: 1.9693; 06: 2.2 |  |  | 14
| 1 | 3 | 01: 1.908; 02: 2.1; 03: 1.56; 04: 1.8992; 05: 1.9745; 06: 2.12 |  |  | 16
| 1 | 4 | 01: 1.933; 02: 2.1; 03: 1.6; 04: 1.7321; 05: 2.2122; 06: 2.17 |  |  | 1.915
| 1 | 5 | 01: 1.837; 02: 2; 03: 1.63; 04: 1.7083; 05: 2.0427; 06: 2.09 |  |  | 168
| 1 | 6 | 01: 1.886; 02: 2; 03: 1.59; 04: 1.7576; 05: 1.9578; 06: 2.22
| Mean |  | 01: 1.8935; 02: 2.0333; 03: 1.59; 04: 1.7712; 05: 2.0388; 06: 2.1517
| Median |  | 01: 1.8985; 02: 2.05; 03: 1.595; 04: 1.7448; 05: 2.0086; 06: 2.145
| Std Dev. |  | 01: 0.031854; 02: 0.08165; 03: 0.026833; 04: 0.072699; 05: 0.09696; 06: 0.052694
| Rel.Std.Dev. |  | 01: 1.682%; 02: 4.016%; 03: 1.688%; 04: 4.105%; 05: 4.756%; 06: 2.449%
| PDM3 |  | 01: -1.121%; 02: 6.181%; 03: -16.97%; 04: -7.509%; 05: 6.467%; 06: 12.36%
| Z-Score (Absolute) |  | 01: 0.36; 02: 0.36; 03: 1.91; 04: 0.98; 05: 0.39; 06: 0.96
| Analytical results for Ge in OREAS 556b (Certified Value 0.13 ppm) | [28 empty] | Certified
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  | 1
| No | No | 01: 05; 02: 06; 03: 08; 04: 09; 05: 10; 06: 11; 08: 12; 09: 14; 10: 18; 11: 22; 12: 25; 13: 28 |  |  | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 08: AR*MS; 09: AR*MS; 10: AR*MS; 11: AR*OES/MS; 12: AR*OES/MS; 13: AR*OES/MS |  |  | 2
|  |  | 01: 0.5g; 02: 0.5g; 03: 0.5g; 04: 0.5g; 05: 0.5g; 06: 0.5g; 08: 0.5g; 09: 0.5g; 10: 0.2g; 11: 0.25g; 12: 0.1g; 13: 0.5g |  |  | 2
| 1 | 1 | 01: 0.1; 02: 0.17; 03: 0.13; 04: 0.13; 05: 0.12; 06: 0.16; 08: 0.1; 09: 0.03; 10: < 0.2; 11: 1.3; 12: 0.34; 13: 0.08 |  |  | 1
| 1 | 2 | 01: 0.1; 02: 0.14; 03: 0.15; 04: 0.13; 05: 0.12; 06: 0.14; 08: 0.1; 09: 0.02; 10: < 0.2; 11: 1.3; 12: 0.36; 13: 0.08 |  |  | 15
| 1 | 3 | 01: < 0.1; 02: 0.15; 03: 0.16; 04: 0.13; 05: 0.1; 06: 0.16; 08: 0.09; 09: < 0.01; 10: < 0.2; 11: 1.3; 12: 0.13; 13: 0.09 |  |  | 16
| 1 | 4 | 01: 0.1; 02: 0.15; 03: 0.14; 04: 0.13; 05: 0.11; 06: 0.17; 08: 0.09; 09: < 0.01; 10: < 0.2; 11: 1.2; 12: 0.37; 13: 0.08 |  |  | 0.12571
| 1 | 5 | 01: < 0.1; 02: 0.18; 03: 0.16; 04: 0.13; 05: 0.12; 06: 0.16; 08: 0.11; 09: 0.01; 10: < 0.2; 11: 1.2; 12: 0.34; 13: 0.08 |  |  | 169
| 1 | 6 | 01: 0.1; 02: 0.16; 03: 0.13; 04: 0.13; 05: 0.1; 06: 0.15; 08: 0.1; 09: 0.03; 10: < 0.2; 11: 1.3; 12: 0.34; 13: 0.07
| Mean |  | 01: 0.1; 02: 0.15833; 03: 0.145; 04: 0.13; 05: 0.11167; 06: 0.15667; 08: 0.098333; 09: 0.0225; 11: 1.2667; 12: 0.31333; 13: 0.08
| Median |  | 01: 0.1; 02: 0.155; 03: 0.145; 04: 0.13; 05: 0.115; 06: 0.16; 08: 0.1; 09: 0.025; 11: 1.3; 12: 0.34; 13: 0.08
| Std Dev. |  | 01: 0; 02: 0.01472; 03: 0.013784; 04: 0; 05: 0.0098319; 06: 0.010328; 08: 0.0075277; 09: 0.0095743; 11: 0.05164; 12: 0.090701; 13: 0.0063246
| Rel.Std.Dev. |  | 01: 0%; 02: 9.297%; 03: 9.506%; 04: 0%; 05: 8.805%; 06: 6.592%; 08: 7.655%; 09: 42.55%; 11: 4.077%; 12: 28.95%; 13: 7.906%
| PDM3 |  | 01: -20.45%; 02: 25.95%; 03: 15.34%; 04: 3.409%; 05: -11.17%; 06: 24.62%; 08: -21.78%; 09: -82.1%; 11: 907.6%; 12: 149.2%; 13: -36.36%
| Z-Score (Absolute) |  | 01: NA; 02: 0.84; 03: 0.54; 04: 0.21; 05: 0.21; 06: 0.81; 08: 0.51; 09: 2.34; 10: 0.47; 11: NA; 12: 4.33; 13: 0.92
Results from laboratories 5 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 556b (Certified Value 0.37 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 22 | 23 | 25 | 26 | 28 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.39 | 0.4 | 0.36 | 0.32 | 0.33 | 0.36 | 0.38 | 0.32 | 0.32 | 0.48193 | 0.37 | 0.33479 | 0.47 | 0.3 | 0.29 | 0.37 | 0.33 | 0.39 | [9 empty] | 1
| 1 | 2 | 0.42 | 0.4 | 0.37 | 0.34 | 0.32 | 0.37 | 0.35 | 0.32 | 0.33 | 0.40469 | 0.37 | 0.33143 | 0.43 | 0.4 | 0.33 | 0.39 | 0.35 | 0.39 | [9 empty] | 19
| 1 | 3 | 0.42 | 0.4 | 0.37 | 0.32 | 0.34 | 0.38 | 0.37 | 0.32 | 0.33 | 0.49268 | 0.36 | 0.34259 | 0.44 | 0.3 | 0.33 | 0.33 | 0.35 | 0.39 | [9 empty] | 16
| 1 | 4 | 0.42 | 0.4 | 0.36 | 0.35 | 0.33 | 0.36 | 0.38 | 0.33 | 0.34 | 0.48292 | 0.37 | 0.34482 | 0.43 | 0.4 | 0.36 | 0.43 | 0.34 | 0.4 | [9 empty] | 0.36687
| 1 | 5 | 0.4 | 0.4 | 0.37 | 0.33 | 0.32 | 0.36 | 0.36 | 0.32 | 0.32 | 0.40107 | 0.38 | 0.34567 | 0.42 | 0.4 | 0.36 | 0.4 | 0.35 | 0.39 | [9 empty] | 170
| 1 | 6 | 0.43 | 0.4 | 0.37 | 0.34 | 0.33 | 0.38 | 0.36 | 0.32 | 0.32 | 0.39214 | 0.38 | 0.34712 | 0.64 | 0.4 | 0.35 | 0.39 | 0.35 | 0.41
| Mean |  | 0.41333 | 0.4 | 0.36667 | 0.33333 | 0.32833 | 0.36833 | 0.36667 | 0.32167 | 0.32667 | 0.44257 | 0.37167 | 0.34107 | 0.47167 | 0.36667 | 0.33667 | 0.385 | 0.345 | 0.395
| Median |  | 0.42 | 0.4 | 0.37 | 0.335 | 0.33 | 0.365 | 0.365 | 0.32 | 0.325 | 0.44331 | 0.37 | 0.34371 | 0.435 | 0.4 | 0.34 | 0.39 | 0.35 | 0.39
| Std Dev. |  | 0.015055 | 6.0809e-17 | 0.005164 | 0.012111 | 0.0075277 | 0.0098319 | 0.012111 | 0.0040825 | 0.008165 | 0.047725 | 0.0075277 | 0.0064274 | 0.084242 | 0.05164 | 0.026583 | 0.033317 | 0.0083666 | 0.0083666
| Rel.Std.Dev. |  | 3.642% | 1.52e-14% | 1.408% | 3.633% | 2.293% | 2.669% | 3.303% | 1.269% | 2.499% | 10.78% | 2.025% | 1.884% | 17.86% | 14.08% | 7.896% | 8.654% | 2.425% | 2.118%
| PDM3 |  | 12.66% | 9.03% | -0.05599% | -9.142% | -10.5% | 0.3983% | -0.05599% | -12.32% | -10.96% | 20.63% | 1.307% | -7.033% | 28.56% | -0.05599% | -8.233% | 4.941% | -5.962% | 7.667%
| Z-Score (Absolute) |  | 1.08 | NA | 0 | 0.77 | 0.89 | 0.04 | 0 | 1.04 | 0.92 | 1.75 | 0.12 | 0.59 | 2.43 | NA | 0.69 | 0.42 | 0.5 | 0.66
Results from laboratories 5 and 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 556b (Certified Value 0.11 ppm) | [28 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 18 | 19 | 21 | 22 | 25 | 26 | 27 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | [6 empty] | 2
| 1 | 1 | < 5 | 0.03 | 0.1 | 0.11 | 0.13 | 0.10358 | 0.11 | 0.1 | 0.11 | 0.09 | 0.1 | 0.1 | 0.057212 | < 1 | 0.1 | 1 | 0.797 | 0.07 | 0.14 | 0.08 | 0.101 | [6 empty] | 1
| 1 | 2 | < 5 | 0.054 | 0.1 | 0.09 | 0.13 | 0.13125 | 0.1 | 0.1 | 0.14 | 0.1 | 0.11 | 0.1 | 0.081111 | < 1 | 0.1 | 1 | 0.798 | 0.08 | 0.15 | 0.09 | 0.093 | [6 empty] | 20
| 1 | 3 | < 5 | 0.036 | 0.1 | 0.1 | 0.12 | 0.10852 | 0.13 | 0.1 | 0.11 | 0.09 | 0.11 | 0.1 | 0.095277 | < 1 | 0.29 | 1 | 0.829 | 0.09 | 0.14 | 0.1 | 0.102 | [6 empty] | 16
| 1 | 4 | < 5 | 0.063 | 0.1 | 0.1 | 0.13 | 0.09962 | 0.1 | 0.11 | 0.11 | 0.1 | 0.1 | 0.09 | 0.11461 | < 1 | 0.14 | 1 | 0.81 | 0.08 | 0.13 | 0.12 | 0.101 | [6 empty] | 0.10605
| 1 | 5 | < 5 | 0.054 | < 0.1 | 0.09 | 0.12 | 0.10924 | 0.11 | 0.1 | 0.12 | 0.13 | 0.11 | 0.11 | 0.084333 | < 1 | 0.11 | 1 | 0.844 | 0.09 | 0.13 | 0.12 | 0.113 | [6 empty] | 171
| 1 | 6 | < 5 | 0.079 | 0.1 | 0.1 | 0.12 | 0.11098 | 0.14 | 0.11 | 0.12 | 0.1 | 0.1 | 0.1 | < 0.05 | < 1 | 0.12 | < 1 | 0.732 | 0.07 | 0.14 | 0.1 | 0.107
| Mean |  |  | 0.052667 | 0.1 | 0.098333 | 0.125 | 0.11053 | 0.115 | 0.10333 | 0.11833 | 0.10167 | 0.105 | 0.1 | 0.086509 |  | 0.14333 | 1 | 0.80167 | 0.08 | 0.13833 | 0.10167 | 0.10283
| Median |  |  | 0.054 | 0.1 | 0.1 | 0.125 | 0.10888 | 0.11 | 0.1 | 0.115 | 0.1 | 0.105 | 0.1 | 0.084333 |  | 0.115 | 1 | 0.804 | 0.08 | 0.14 | 0.1 | 0.1015
| Std Dev. |  |  | 0.017862 | 0 | 0.0075277 | 0.0054772 | 0.010978 | 0.016432 | 0.005164 | 0.01169 | 0.01472 | 0.0054772 | 0.0063246 | 0.020969 |  | 0.073394 | 0 | 0.038723 | 0.0089443 | 0.0075277 | 0.016021 | 0.0067057
| Rel.Std.Dev. |  |  | 33.92% | 0% | 7.655% | 4.382% | 9.932% | 14.29% | 4.997% | 9.879% | 14.48% | 5.216% | 6.325% | 24.24% |  | 51.21% | 0% | 4.83% | 11.18% | 5.442% | 15.76% | 6.521%
| PDM3 |  |  | -50.34% | -5.704% | -7.275% | 17.87% | 4.227% | 8.441% | -2.56% | 11.58% | -4.132% | -0.9888% | -5.704% | -18.43% |  | 35.16% | 843% | 655.9% | -24.56% | 30.44% | -4.132% | -3.032%
| Z-Score (Absolute) |  | 121.12 | 2.65 | NA | 0.34 | 1.01 | 0.28 | 0.51 | 0.08 | 0.67 | 0.17 | 0 | 0.25 | 1.45 | 19.98 | 1.94 | NA | 35.23 | 1.26 | 1.69 | 0.17 | 0.11
Results from laboratory 4 were removed due to their 0.1 ppm reading resolution. Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 556b (Certified Value 0.17 ppm) | [28 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | [6 empty] | 1
| No | No | 01: 04; 02: 05; 04: 14; 05: 18; 06: 26 | [6 empty] | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*MS | [6 empty] | 2
|  |  | 01: 01g; 02: 0.5g; 04: 0.5g; 05: 0.2g; 06: 01g | [6 empty] | 2
| 1 | 1 | 01: 0.172; 02: 0.2; 04: 0.13; 05: 0.15007; 06: 0.2 | [6 empty] | 1
| 1 | 2 | 01: 0.175; 02: 0.1; 04: 0.13; 05: 0.14762; 06: 0.2 | [6 empty] | 3
| 1 | 3 | 01: 0.168; 02: 0.2; 04: 0.13; 05: 0.14879; 06: 0.19 | [6 empty] | 16
| 1 | 4 | 01: 0.174; 02: 0.2; 04: 0.13; 05: 0.17327; 06: 0.2 | [6 empty] | 0.16714
| 1 | 5 | 01: 0.174; 02: 0.2; 04: 0.13; 05: 0.15598; 06: 0.21 | [6 empty] | 172
| 1 | 6 | 01: 0.169; 02: 0.2; 04: 0.13; 05: 0.14937; 06: 0.2
| Mean |  | 01: 0.172; 02: 0.18333; 04: 0.13; 05: 0.15418; 06: 0.2
| Median |  | 01: 0.173; 02: 0.2; 04: 0.13; 05: 0.14972; 06: 0.2
| Std Dev. |  | 01: 0.0028983; 02: 0.040825; 04: 0; 05: 0.0097972; 06: 0.0063246
| Rel.Std.Dev. |  | 01: 1.685%; 02: 22.27%; 04: 0%; 05: 6.354%; 06: 3.162%
| PDM3 |  | 01: 2.908%; 02: 9.689%; 04: -22.22%; 05: -7.752%; 06: 19.66%
| Z-Score (Absolute) |  | 01: 0; 02: 0.43; 04: 1.59; 05: 0.67; 06: 1.06
| Analytical results for In in OREAS 556b (Certified Value 2.94 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 20 | 22 | 23 | 25 | 26 | 28 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [9 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [9 empty] | 3
| 1 | 1 | < 5 | 3 | 2.7 | 2.78 | 2.8671 | 2.94 | 3.1 | 2.86 | 2.91 | 2.82 | 2.7569 | 3.27 | 3.1 | 2.79 | 2.612 | 3.14 | 3.07 | 2.889 | [9 empty] | 1
| 1 | 2 | < 5 | 2.98 | 2.39 | 2.77 | 2.847 | 3 | 3.19 | 2.99 | 2.94 | 2.89 | 2.7529 | 3.22 | 3.2 | 2.84 | 2.467 | 2.84 | 3.13 | 2.84 | [9 empty] | 4
| 1 | 3 | < 5 | 2.94 | 2.53 | 2.89 | 2.8217 | 2.86 | 3.27 | 2.98 | 2.9 | 2.86 | 2.8151 | 3.26 | 3.2 | 2.82 | 2.369 | 3.08 | 3.14 | 2.862 | [9 empty] | 16
| 1 | 4 | < 5 | 2.96 | 2.62 | 2.8 | 2.8306 | 2.91 | 3.18 | 2.94 | 2.95 | 2.96 | 2.7648 | 3.28 | 3 | 2.82 | 2.482 | 3.09 | 3.06 | 2.862 | [9 empty] | 2.9356
| 1 | 5 | < 5 | 3.02 | 2.46 | 2.79 | 2.8425 | 2.93 | 3.13 | 2.96 | 2.94 | 2.97 | 2.8253 | 3.28 | 3 | 2.71 | 2.334 | 3 | 3.18 | 2.813 | [9 empty] | 173
| 1 | 6 | < 5 | 2.95 | 2.43 | 2.82 | 2.8232 | 3.04 | 3.2 | 2.98 | 2.87 | 2.91 | 2.7928 | 3.2 | 3 | 2.79 | 2.389 | 2.93 | 3.2 | 2.953
| Mean |  |  | 2.975 | 2.5217 | 2.8083 | 2.8387 | 2.9467 | 3.1783 | 2.9517 | 2.9183 | 2.9017 | 2.7846 | 3.2517 | 3.0833 | 2.795 | 2.4422 | 3.0133 | 3.13 | 2.8698
| Median |  |  | 2.97 | 2.495 | 2.795 | 2.8366 | 2.935 | 3.185 | 2.97 | 2.925 | 2.9 | 2.7788 | 3.265 | 3.05 | 2.805 | 2.428 | 3.04 | 3.135 | 2.862
| Std Dev. |  |  | 0.030822 | 0.11923 | 0.043551 | 0.017223 | 0.064395 | 0.059133 | 0.048339 | 0.030605 | 0.057764 | 0.031019 | 0.033714 | 0.098319 | 0.045935 | 0.10092 | 0.11272 | 0.056569 | 0.048014
| Rel.Std.Dev. |  |  | 1.036% | 4.728% | 1.551% | 0.6067% | 2.185% | 1.86% | 1.638% | 1.049% | 1.991% | 1.114% | 1.037% | 3.189% | 1.643% | 4.133% | 3.741% | 1.807% | 1.673%
| PDM3 |  |  | 1.341% | -14.1% | -4.336% | -3.303% | 0.3758% | 8.267% | 0.5461% | -0.5894% | -1.157% | -5.143% | 10.77% | 5.031% | -4.791% | -16.81% | 2.647% | 6.621% | -2.241%
| Z-Score (Absolute) |  | 2.53 | 0.4 | 2.4 | 0.63 | 0.44 | 0.23 | 1.66 | 0.26 | 0.05 | 0.05 | 0.77 | 2.11 | 1.07 | 0.71 | 2.89 | 0.64 | 1.36 | 0.25
| Analytical results for Ir in OREAS 556b (Indicative Value < 0.003 ppm) | [28 empty] | Indicative
Ir | Round | Replicate | 01: Lab | [6 empty] | 1
| No | No | 01: 07 | [6 empty] | ppm
|  |  | 01: AR*MS | [6 empty] | 3
|  |  | 01: 0.5g | [6 empty] | 3
| 1 | 1 | 01: < 0.003 | [6 empty] | 1
| 1 | 2 | 01: < 0.003 | [6 empty] | 24
| 1 | 3 | 01: < 0.003 | [6 empty] | 16
| 1 | 4 | 01: < 0.003 | [6 empty] | < 0.003
| 1 | 5 | 01: < 0.003 | [6 empty] | 30
| 1 | 6 | 01: < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for K in OREAS 556b (Certified Value 0.187 wt.%) | [28 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.01 | 0.16 | 0.23 | 0.199 | 0.25 | 0.15 | 0.18 | 0.21 | 0.17 | 0.18 | 0.19 | 0.2115 | 0.1361 | 0.20162 | 0.308 | 0.18889 | 0.13 | 0.16 | 0.19 | 0.213 | 0.324 | 0.3354 | 0.21 | 0.2 |  |  |  | 1
| 1 | 2 | 0.01 | 0.2 | 0.23 | 0.2057 | 0.21 | 0.15 | 0.18 | 0.21 | 0.18 | 0.18 | 0.19 | 0.2016 | 0.1395 | 0.19074 | 0.312 | 0.19088 | 0.13 | 0.16 | 0.2 | 0.209 | 0.427 | 0.307 | 0.21 | 0.2 |  |  |  | 0
| 1 | 3 | 0.01 | 0.17 | 0.22 | 0.1983 | 0.23 | 0.15 | 0.17 | 0.21 | 0.17 | 0.19 | 0.19 | 0.2105 | 0.1337 | 0.19109 | 0.308 | 0.19767 | 0.13 | 0.16 | 0.18 | 0.209 | 0.311 | 0.2903 | 0.21 | 0.19 |  |  |  | 16
| 1 | 4 | 0.01 | 0.16 | 0.21 | 0.2068 | 0.25 | 0.14 | 0.17 | 0.21 | 0.17 | 0.18 | 0.19 | 0.2112 | 0.1361 | 0.18336 | 0.326 | 0.18846 | 0.13 | 0.16 | 0.19 | 0.221 | 0.523 | 0.3165 | 0.22 | 0.19 |  |  |  | 0.18728
| 1 | 5 | 0.01 | 0.19 | 0.21 | 0.1977 | 0.24 | 0.15 | 0.17 | 0.21 | 0.18 | 0.18 | 0.19 | 0.2088 | 0.1461 | 0.19241 | 0.32 | 0.1876 | 0.14 | 0.16 | 0.19 | 0.224 | 0.394 | 0.3214 | 0.21 | 0.19 |  |  |  | 174
| 1 | 6 | 0.01 | 0.18 | 0.22 | 0.2024 | 0.24 | 0.15 | 0.17 | 0.21 | 0.18 | 0.18 | 0.19 | 0.2093 | 0.1361 | 0.19957 | 0.301 | 0.19024 | 0.14 | 0.16 | 0.18 | 0.224 | 0.427 | 0.3036 | 0.21 | 0.19
| Mean |  | 0.01 | 0.17667 | 0.22 | 0.20165 | 0.23667 | 0.14833 | 0.17333 | 0.21 | 0.175 | 0.18167 | 0.19 | 0.20882 | 0.13793 | 0.19313 | 0.3125 | 0.19063 | 0.13333 | 0.16 | 0.18833 | 0.21667 | 0.401 | 0.31237 | 0.21167 | 0.19333
| Median |  | 0.01 | 0.175 | 0.22 | 0.2007 | 0.24 | 0.15 | 0.17 | 0.21 | 0.175 | 0.18 | 0.19 | 0.2099 | 0.1361 | 0.19175 | 0.31 | 0.18957 | 0.13 | 0.16 | 0.19 | 0.217 | 0.4105 | 0.31175 | 0.21 | 0.19
| Std Dev. |  | 0 | 0.01633 | 0.0089443 | 0.0039328 | 0.015055 | 0.0040825 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 3.0405e-17 | 0.0036886 | 0.004408 | 0.0066234 | 0.0090719 | 0.0036498 | 0.005164 | 0 | 0.0075277 | 0.007174 | 0.077864 | 0.015632 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 0% | 9.243% | 4.066% | 1.95% | 6.361% | 2.752% | 2.979% | 0% | 3.13% | 2.247% | 1.6e-14% | 1.766% | 3.196% | 3.43% | 2.903% | 1.915% | 3.873% | 0% | 3.997% | 3.311% | 19.42% | 5.004% | 1.929% | 2.671%
| PDM3 |  | -94.66% | -5.666% | 17.47% | 7.674% | 26.37% | -20.8% | -7.446% | 12.13% | -6.556% | -2.996% | 1.453% | 11.5% | -26.35% | 3.125% | 66.86% | 1.787% | -28.8% | -14.57% | 0.5633% | 15.69% | 114.1% | 66.79% | 13.02% | 3.233%
| Z-Score (Absolute) |  | 6.4 | 0.54 | 0.99 | 0.34 | 1.58 | 1.53 | 0.65 | 0.64 | 0.59 | 0.36 | 0.07 | 0.6 | 1.9 | 0.04 | 4.24 | 0.04 | 2.06 | 1.12 | 0.12 | 0.87 | 7.36 | 4.24 | 0.7 | 0.05
| Analytical results for La in OREAS 556b (Certified Value 19.5 ppm) | [28 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 1
| 1 | 1 | 23.9 | 16.1 | 20.9 | 20.685 | 24 | 16 | 18.8 | 21 | 17.3 | 18.6 | 18.9 | 16.04 | 15 | 20.707 | 9.0338 | 24.9 | 18.8789 | 14.2 | 18 | 20.1 | 28.4 | 24.42 | 19.6 | 20.3 |  |  |  | 1
| 1 | 2 | 23.9 | 18.2 | 20.8 | 20.541 | 20.7 | 16 | 19.4 | 21.7 | 18.4 | 18.6 | 19.1 | 16.02 | 15 | 20.9424 | 9.0244 | 24.7 | 18.7545 | 14.2 | 19 | 19.4 | 33.4 | 25.18 | 18.8 | 19.7 |  |  |  | 5
| 1 | 3 | 25.4 | 16.2 | 20.2 | 20.856 | 23.2 | 16.5 | 18 | 21.9 | 18.4 | 18.8 | 19 | 16.13 | 15 | 20.3907 | 9.0719 | 25 | 19.0572 | 14.3 | 19 | 19.3 | 26.7 | 24.26 | 19.1 | 19.4 |  |  |  | 16
| 1 | 4 | 20.4 | 15.3 | 20 | 21.44 | 23.9 | 15.7 | 17.7 | 22 | 18.2 | 18.8 | 19.3 | 16.08 | 14 | 20.4739 | 8.9672 | 25.3 | 19.9819 | 13.9 | 18 | 19.8 | 37.1 | 23.88 | 19 | 19.4 |  |  |  | 19.4745
| 1 | 5 | 26.7 | 17.2 | 20.2 | 21.238 | 22.5 | 15.6 | 17.7 | 21.6 | 18.3 | 18.5 | 19.7 | 16.11 | 16 | 20.9009 | 9.1306 | 24.8 | 18.9199 | 14.1 | 18 | 19.5 | 31.8 | 24.37 | 19.5 | 18.8 |  |  |  | 175
| 1 | 6 | 24.1 | 19.3 | 20.2 | 21.019 | 21.9 | 16.1 | 18 | 22.5 | 18.5 | 17.8 | 19.2 | 16.07 | 15 | 21.1992 | 9.0689 | 24.8 | 18.6993 | 14.3 | 18 | 19.3 | 32.9 | 24.79 | 19.5 | 20
| Mean |  | 24.0667 | 17.05 | 20.3833 | 20.9632 | 22.7 | 15.9833 | 18.2667 | 21.7833 | 18.1833 | 18.5167 | 19.2 | 16.075 | 15 | 20.769 | 9.0495 | 24.9167 | 19.0486 | 14.1667 | 18.3333 | 19.5667 | 31.7167 | 24.4833 | 19.25 | 19.6
| Median |  | 24 | 16.7 | 20.2 | 20.9375 | 22.85 | 16 | 18 | 21.8 | 18.35 | 18.6 | 19.15 | 16.075 | 15 | 20.8039 | 9.0514 | 24.85 | 18.8994 | 14.2 | 18 | 19.45 | 32.35 | 24.395 | 19.3 | 19.55
| Std Dev. |  | 2.1078 | 1.4896 | 0.37103 | 0.33856 | 1.2696 | 0.31885 | 0.68605 | 0.49565 | 0.4446 | 0.37103 | 0.28284 | 0.041352 | 0.63246 | 0.30547 | 0.054948 | 0.2137 | 0.47435 | 0.15055 | 0.5164 | 0.32042 | 3.7253 | 0.44947 | 0.32711 | 0.52536
| Rel.Std.Dev. |  | 8.758% | 8.737% | 1.82% | 1.615% | 5.593% | 1.995% | 3.756% | 2.275% | 2.445% | 2.004% | 1.473% | 0.2572% | 4.216% | 1.471% | 0.6072% | 0.8576% | 2.49% | 1.063% | 2.817% | 1.638% | 11.75% | 1.836% | 1.699% | 2.68%
| PDM3 |  | 23.58% | -12.45% | 4.667% | 7.644% | 16.56% | -17.93% | -6.202% | 11.86% | -6.63% | -4.918% | -1.41% | -17.46% | -22.98% | 6.647% | -53.53% | 27.95% | -2.187% | -27.26% | -5.86% | 0.4732% | 62.86% | 25.72% | -1.153% | 0.6444%
| Z-Score (Absolute) |  | 1.67 | 0.75 | 0.4 | 0.6 | 1.2 | 1.12 | 0.33 | 0.88 | 0.36 | 0.24 | 0.01 | 1.09 | 1.46 | 0.53 | 3.51 | 1.96 | 0.06 | 1.74 | 0.31 | 0.12 | 4.31 | 1.81 | 0.01 | 0.13
| Analytical results for Li in OREAS 556b (Certified Value 31.2 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [6 empty] | 1
| 1 | 1 | 39.5 | 31 | 38.5 | 19.8 | 28.7 | 31.4 | 25.7 | 26.8 | 28 | 40.22 | < 1 | 27.0085 | 38.5 | 28.0588 | 38.3 | 23 | 28 | 37 | 37.1 | 37.7 | 24.1 | [6 empty] | 1
| 1 | 2 | 39 | 31.2 | 33.7 | 18.7 | 28.7 | 32.5 | 26.9 | 26 | 28.8 | 40.17 | < 1 | 25.5107 | 38.5 | 27.394 | 37.2 | 24 | 29 | 37 | 38.6 | 39.7 | 23.1 | [6 empty] | 33
| 1 | 3 | 37.9 | 29.5 | 35.6 | 19.9 | 25.9 | 32.5 | 27.1 | 27.4 | 28.9 | 40.34 | < 1 | 26.1609 | 38.7 | 27.7207 | 37.5 | 24 | 28.3 | 37 | 34 | 38.3 | 24.1 | [6 empty] | 16
| 1 | 4 | 37.4 | 29.9 | 38 | 17.8 | 26.4 | 32.9 | 26.7 | 26.8 | 28.5 | 40.36 | < 1 | 25.9226 | 39.4 | 29.5235 | 35.8 | 23 | 28.5 | 38 | 39.5 | 37 | 23.2 | [6 empty] | 31.2152
| 1 | 5 | 37.6 | 29.7 | 35.9 | 17.9 | 26.1 | 32.8 | 27.3 | 26.8 | 29.1 | 40.18 | 1 | 26.8178 | 39.3 | 29.9228 | 35.2 | 23 | 26.5 | 37 | 36.9 | 38.4 | 23.8 | [6 empty] | 176
| 1 | 6 | 38.4 | 29.5 | 35.5 | 19.1 | 26.2 | 32.8 | 27.9 | 26 | 28.5 | 40.19 | < 1 | 27.1693 | 39.7 | 27.9552 | 36.3 | 24 | 27.3 | 38 | 38.8 | 38.7 | 22.8
| Mean |  | 38.3 | 30.1333 | 36.2 | 18.8667 | 27 | 32.4833 | 26.9333 | 26.6333 | 28.6333 | 40.2433 | 1 | 26.4317 | 39.0167 | 28.4292 | 36.7167 | 23.5 | 27.9333 | 37.3333 | 37.4833 | 38.3 | 23.5167
| Median |  | 38.15 | 29.8 | 35.75 | 18.9 | 26.3 | 32.65 | 27 | 26.8 | 28.65 | 40.205 | 1 | 26.4894 | 39 | 28.007 | 36.75 | 23.5 | 28.15 | 37 | 37.85 | 38.35 | 23.5
| Std Dev. |  | 0.82462 | 0.76594 | 1.7731 | 0.9048 | 1.3266 | 0.55648 | 0.73121 | 0.54283 | 0.38816 | 0.084538 |  | 0.66432 | 0.51543 | 1.0357 | 1.1548 | 0.54772 | 0.90037 | 0.5164 | 1.9833 | 0.91433 | 0.55648
| Rel.Std.Dev. |  | 2.153% | 2.542% | 4.898% | 4.796% | 4.914% | 1.713% | 2.715% | 2.038% | 1.356% | 0.2101% |  | 2.513% | 1.321% | 3.643% | 3.145% | 2.331% | 3.223% | 1.383% | 5.291% | 2.387% | 2.366%
| PDM3 |  | 22.7% | -3.466% | 15.97% | -39.56% | -13.5% | 4.063% | -13.72% | -14.68% | -8.271% | 28.92% | -96.8% | -15.32% | 24.99% | -8.925% | 17.62% | -24.72% | -10.51% | 19.6% | 20.08% | 22.7% | -24.66%
| Z-Score (Absolute) |  | 1.27 | 0.2 | 0.99 | 1.28 | 0.21 | 0.51 | 0.22 | 0.26 | 0 | 1.53 | 3.68 | 0.29 | 1.36 | 0.03 | 1.06 | 0.67 | 0.09 | 1.14 | 1.16 | 1.27 | 0.67
| Analytical results for Lu in OREAS 556b (Indicative Value 0.06 ppm) | [28 empty] | Indicative
Lu | Round | Replicate | 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab | [6 empty] | 1
| No | No | 03: 04; 04: 05; 05: 18; 06: 23; 08: 26 | [6 empty] | ppm
|  |  | 03: AR*MS; 04: AR*OES/MS; 05: AR*MS; 06: AR*MS; 08: AR*MS | [6 empty] | 3
|  |  | 03: 01g; 04: 0.5g; 05: 0.2g; 06: 0.25g; 08: 01g | [6 empty] | 3
| 1 | 1 | 03: 0.055; 04: < 0.1; 05: 0.052034; 06: 0.07; 08: 0.07 | [6 empty] | 1
| 1 | 2 | 03: 0.049; 04: < 0.1; 05: < 0.05; 06: 0.06; 08: 0.07 | [6 empty] | 25
| 1 | 3 | 03: 0.051; 04: < 0.1; 05: 0.051266; 06: 0.06; 08: 0.07 | [6 empty] | 16
| 1 | 4 | 03: 0.061; 04: < 0.1; 05: 0.060025; 06: 0.07; 08: 0.07 | [6 empty] | 0.060117
| 1 | 5 | 03: 0.054; 04: < 0.1; 05: 0.052952; 06: 0.06; 08: 0.07 | [6 empty] | 31
| 1 | 6 | 03: 0.051; 04: < 0.1; 05: 0.051891; 06: 0.06; 08: 0.07
| Mean |  | 03: 0.0535; 05: 0.053633; 06: 0.063333; 08: 0.07
| Median |  | 03: 0.0525; 05: 0.052034; 06: 0.06; 08: 0.07
| Std Dev. |  | 03: 0.0042778; 05: 0.0036234; 06: 0.005164; 08: 0
| Rel.Std.Dev. |  | 03: 7.996%; 05: 6.756%; 06: 8.154%; 08: 0%
| PDM3 |  | 03: -11.01%; 05: -10.78%; 06: 5.351%; 08: 16.44%
| Z-Score (Absolute) |  | 03: 0; 04: 0.51; 05: 0.67; 06: 1.43; 08: 2.4
| Analytical results for Mg in OREAS 556b (Certified Value 1.85 wt.%) | [28 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 2
| 1 | 1 | 1.92 | 1.65 | 2.17 | 1.97 | 1.82 | 1.43 | 1.77 | 1.87 | 1.66 | 1.72 | 1.73 | 1.8722 | 1.52 | 2.0133 | 2.36 | 1.8548 | 1.56 | 1.56 | 1.88 | 2.029 | 2.08 | 2.21 | 1.86 | 1.96 |  |  |  | 1
| 1 | 2 | 1.88 | 1.71 | 2.19 | 2.04 | 1.58 | 1.44 | 1.77 | 1.93 | 1.72 | 1.69 | 1.75 | 1.8822 | 1.616 | 1.9757 | 2.22 | 1.8255 | 1.56 | 1.58 | 1.85 | 1.996 | 2.06 | 2.23 | 1.83 | 1.9 |  |  |  | 0
| 1 | 3 | 1.96 | 1.71 | 2.15 | 1.98 | 1.66 | 1.42 | 1.66 | 1.92 | 1.7 | 1.74 | 1.76 | 1.8815 | 1.538 | 1.9891 | 2.32 | 1.8818 | 1.6 | 1.55 | 1.83 | 1.969 | 2.01 | 2.16 | 1.85 | 1.88 |  |  |  | 16
| 1 | 4 | 1.76 | 1.62 | 2.18 | 2.05 | 1.74 | 1.36 | 1.69 | 1.93 | 1.7 | 1.72 | 1.74 | 1.8806 | 1.556 | 1.944 | 2.36 | 1.8995 | 1.57 | 1.5 | 1.87 | 2.054 | 2.09 | 2.1 | 1.89 | 1.85 |  |  |  | 1.8474
| 1 | 5 | 2.02 | 1.79 | 2.12 | 1.96 | 1.69 | 1.45 | 1.66 | 1.92 | 1.73 | 1.73 | 1.78 | 1.8804 | 1.677 | 1.9445 | 2.27 | 1.8643 | 1.59 | 1.54 | 1.85 | 2.028 | 2.03 | 2.21 | 1.86 | 1.86 |  |  |  | 177
| 1 | 6 | 1.91 | 1.93 | 2.17 | 2.01 | 1.71 | 1.46 | 1.68 | 1.94 | 1.77 | 1.68 | 1.73 | 1.8789 | 1.55 | 1.9872 | 2.27 | 1.8725 | 1.62 | 1.56 | 1.85 | 2.045 | 2.06 | 2.21 | 1.86 | 1.89
| Mean |  | 1.9083 | 1.735 | 2.1633 | 2.0017 | 1.7 | 1.4267 | 1.705 | 1.9183 | 1.7133 | 1.7133 | 1.7483 | 1.8793 | 1.5762 | 1.9756 | 2.3 | 1.8664 | 1.5833 | 1.5483 | 1.855 | 2.0202 | 2.055 | 2.1867 | 1.8583 | 1.89
| Median |  | 1.915 | 1.71 | 2.17 | 1.995 | 1.7 | 1.435 | 1.685 | 1.925 | 1.71 | 1.72 | 1.745 | 1.8805 | 1.553 | 1.9814 | 2.295 | 1.8684 | 1.58 | 1.555 | 1.85 | 2.0285 | 2.06 | 2.21 | 1.86 | 1.885
| Std Dev. |  | 0.087274 | 0.11203 | 0.025033 | 0.037639 | 0.08025 | 0.03559 | 0.051672 | 0.024833 | 0.036697 | 0.023381 | 0.019408 | 0.0036529 | 0.059088 | 0.027222 | 0.056214 | 0.025219 | 0.024221 | 0.027142 | 0.017607 | 0.031946 | 0.030166 | 0.048442 | 0.019408 | 0.038987
| Rel.Std.Dev. |  | 4.573% | 6.457% | 1.157% | 1.88% | 4.721% | 2.495% | 3.031% | 1.294% | 2.142% | 1.365% | 1.11% | 0.1944% | 3.749% | 1.378% | 2.444% | 1.351% | 1.53% | 1.753% | 0.9492% | 1.581% | 1.468% | 2.215% | 1.044% | 2.063%
| PDM3 |  | 3.297% | -6.085% | 17.1% | 8.349% | -7.98% | -22.78% | -7.709% | 3.838% | -7.258% | -7.258% | -5.364% | 1.726% | -14.68% | 6.94% | 24.5% | 1.028% | -14.29% | -16.19% | 0.4103% | 9.351% | 11.24% | 18.36% | 0.5907% | 2.305%
| Z-Score (Absolute) |  | 0.21 | 0.58 | 1.36 | 0.63 | 0.73 | 1.97 | 0.71 | 0.25 | 0.67 | 0.67 | 0.52 | 0.08 | 1.3 | 0.51 | 1.98 | 0.02 | 1.26 | 1.42 | 0.03 | 0.71 | 0.87 | 1.47 | 0.02 | 0.12
| Analytical results for Mn in OREAS 556b (Certified Value 0.117 wt.%) | [28 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1172 | 0.1134 | 0.1124 | 0.1229 | 0.112 | 0.131 | 0.118 | 0.12 | 0.1195 | 0.1165 | 0.1165 | 0.1114 | 0.1193 | 0.12062 | 0.115 | 0.11086 | 0.1137 | 0.1169 | 0.0925 | 0.1442 | 0.113 | 0.1211 | 0.1214 | 0.1225 |  |  |  | 1
| 1 | 2 | 0.1138 | 0.1131 | 0.1134 | 0.1192 | 0.0961 | 0.124 | 0.118 | 0.126 | 0.1205 | 0.1145 | 0.1185 | 0.1047 | 0.1208 | 0.12132 | 0.116 | 0.11645 | 0.1134 | 0.1177 | 0.0924 | 0.1402 | 0.112 | 0.1201 | 0.1224 | 0.1208 |  |  |  | 0
| 1 | 3 | 0.1147 | 0.114 | 0.1111 | 0.1249 | 0.103 | 0.129 | 0.1185 | 0.124 | 0.118 | 0.116 | 0.117 | 0.1108 | 0.1138 | 0.12056 | 0.111 | 0.114 | 0.1103 | 0.1168 | 0.0855 | 0.1447 | 0.112 | 0.119 | 0.1193 | 0.1205 |  |  |  | 16
| 1 | 4 | 0.1165 | 0.114 | 0.1107 | 0.1262 | 0.107 | 0.127 | 0.115 | 0.1195 | 0.119 | 0.115 | 0.1175 | 0.1109 | 0.1193 | 0.12057 | 0.115 | 0.11695 | 0.1079 | 0.1166 | 0.0878 | 0.1412 | 0.115 | 0.1173 | 0.1219 | 0.1199 |  |  |  | 0.11668
| 1 | 5 | 0.1168 | 0.1138 | 0.1122 | 0.1217 | 0.101 | 0.124 | 0.115 | 0.1215 | 0.119 | 0.1155 | 0.1175 | 0.1101 | 0.127 | 0.11788 | 0.11 | 0.11702 | 0.1064 | 0.1138 | 0.0836 | 0.1337 | 0.11 | 0.1201 | 0.1209 | 0.1206 |  |  |  | 178
| 1 | 6 | 0.1165 | 0.1218 | 0.1119 | 0.1228 | 0.0973 | 0.124 | 0.118 | 0.128 | 0.1235 | 0.1145 | 0.1155 | 0.1112 | 0.1162 | 0.11803 | 0.112 | 0.11416 | 0.1082 | 0.1166 | 0.0851 | 0.1403 | 0.11 | 0.1214 | 0.1201 | 0.1215
| Mean |  | 0.11592 | 0.11502 | 0.11195 | 0.12295 | 0.10273 | 0.1265 | 0.11708 | 0.12317 | 0.11992 | 0.11533 | 0.11708 | 0.10985 | 0.1194 | 0.11983 | 0.11317 | 0.11491 | 0.10998 | 0.1164 | 0.087817 | 0.14072 | 0.112 | 0.11983 | 0.121 | 0.12097
| Median |  | 0.1165 | 0.1139 | 0.11205 | 0.12285 | 0.102 | 0.1255 | 0.118 | 0.12275 | 0.11925 | 0.11525 | 0.11725 | 0.11085 | 0.1193 | 0.12057 | 0.1135 | 0.11531 | 0.10925 | 0.1167 | 0.08665 | 0.14075 | 0.112 | 0.1201 | 0.12115 | 0.1207
| Std Dev. |  | 0.0013467 | 0.0033421 | 0.00096488 | 0.0024501 | 0.0060112 | 0.0030166 | 0.0016253 | 0.0034157 | 0.0019343 | 0.0008165 | 0.0010206 | 0.0025618 | 0.004502 | 0.001481 | 0.0024833 | 0.0023981 | 0.0030314 | 0.0013372 | 0.0038332 | 0.0039474 | 0.0018974 | 0.0015042 | 0.0011524 | 0.00091141
| Rel.Std.Dev. |  | 1.162% | 2.906% | 0.8619% | 1.993% | 5.851% | 2.385% | 1.388% | 2.773% | 1.613% | 0.7079% | 0.8717% | 2.332% | 3.771% | 1.236% | 2.194% | 2.087% | 2.756% | 1.149% | 4.365% | 2.805% | 1.694% | 1.255% | 0.9524% | 0.7534%
| PDM3 |  | -0.6541% | -1.425% | -4.054% | 5.374% | -11.95% | 8.416% | 0.3458% | 5.56% | 2.774% | -1.154% | 0.3458% | -5.853% | 2.331% | 2.701% | -3.011% | -1.52% | -5.739% | -0.2398% | -24.74% | 20.6% | -4.011% | 2.703% | 3.703% | 3.674%
| Z-Score (Absolute) |  | 0.14 | 0.3 | 0.83 | 1.07 | 2.42 | 1.69 | 0.06 | 1.11 | 0.55 | 0.24 | 0.06 | 1.19 | 0.46 | 0.53 | 0.62 | 0.32 | 1.17 | 0.06 | 5 | 4.15 | 0.82 | 0.53 | 0.74 | 0.73
| Analytical results for Mo in OREAS 556b (Certified Value 17.4 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 2
| 1 | 1 | < 1 | 15.09 | 16.9 | 17.7 | 16.2 | 18.3 | 17.9079 | 18.85 | 18 | 16.05 | 18.5 | 16.8 | 15.99 | 18 | 17.3492 | 17.7 | 17.7403 | 18.9 | 17.88 | 16 | 16.1 | 19.96 | 16.6 | 18.1 | 16.5 | 17.58 |  | 1
| 1 | 2 | < 1 | 16.32 | 16.8 | 17 | 14.1 | 18.4 | 17.5308 | 18.95 | 18.55 | 16.85 | 18.4 | 17.15 | 15.38 | 18 | 17.4923 | 17.4 | 18.0086 | 18.4 | 17.83 | 16 | 16.6 | 19.25 | 15.4 | 18.5 | 16.7 | 17.42 |  | 7
| 1 | 3 | < 1 | 15.4 | 16.4 | 17.2 | 14.9 | 19.1 | 17.8438 | 18.65 | 19.1 | 16.95 | 18.15 | 16.9 | 15.41 | 18.2 | 17.8787 | 17.6 | 18.4579 | 18.1 | 17.43 | 16 | 16.2 | 18.75 | 16.7 | 18.8 | 16.8 | 17.29 |  | 16
| 1 | 4 | < 1 | 15.3 | 16.7 | 17.5 | 15 | 18.7 | 17.6637 | 18.65 | 18.35 | 16.8 | 18.4 | 17.4 | 15.62 | 18 | 17.801 | 17.6 | 18.2106 | 17.4 | 16.91 | 16 | 16.3 | 19.6 | 16.9 | 18.2 | 16.9 | 17.46 |  | 17.3801
| 1 | 5 | < 1 | 16.1 | 16.6 | 17.6 | 14.6 | 18.4 | 17.2023 | 18.95 | 18.15 | 16.4 | 18.5 | 17.4 | 15.84 | 18.1 | 16.9492 | 17.6 | 17.8622 | 18.2 | 16.5 | 16 | 15.7 | 18.45 | 16.2 | 18.5 | 16.6 | 17.05 |  | 179
| 1 | 6 | < 1 | 15.17 | 16.6 | 17.4 | 14.5 | 18.6 | 17.5864 | 19.55 | 18.55 | 17.05 | 18.2 | 16.9 | 15.88 | 18.1 | 17.1721 | 17.6 | 18.2268 | 18.6 | 18.14 | 17 | 16.2 | 18.81 | 16.2 | 18.6 | 16.8 | 18.14
| Mean |  |  | 15.5633 | 16.6667 | 17.4 | 14.8833 | 18.5833 | 17.6225 | 18.9333 | 18.45 | 16.6833 | 18.3583 | 17.0917 | 15.6867 | 18.0667 | 17.4404 | 17.5833 | 18.0844 | 18.2667 | 17.4483 | 16.1667 | 16.1833 | 19.1367 | 16.3333 | 18.45 | 16.7167 | 17.49
| Median |  |  | 15.35 | 16.65 | 17.45 | 14.75 | 18.5 | 17.625 | 18.9 | 18.45 | 16.825 | 18.4 | 17.025 | 15.73 | 18.05 | 17.4208 | 17.6 | 18.1096 | 18.3 | 17.63 | 16 | 16.2 | 19.03 | 16.4 | 18.5 | 16.75 | 17.44
| Std Dev. |  |  | 0.51678 | 0.17512 | 0.26077 | 0.71949 | 0.29269 | 0.25234 | 0.33116 | 0.38601 | 0.38166 | 0.14972 | 0.26536 | 0.2561 | 0.08165 | 0.35957 | 0.098319 | 0.26429 | 0.51251 | 0.63244 | 0.40825 | 0.29269 | 0.57151 | 0.53541 | 0.25884 | 0.1472 | 0.36606
| Rel.Std.Dev. |  |  | 3.321% | 1.051% | 1.499% | 4.834% | 1.575% | 1.432% | 1.749% | 2.092% | 2.288% | 0.8156% | 1.553% | 1.633% | 0.4519% | 2.062% | 0.5592% | 1.461% | 2.806% | 3.625% | 2.525% | 1.809% | 2.986% | 3.278% | 1.403% | 0.8805% | 2.093%
| PDM3 |  |  | -10.45% | -4.105% | 0.1145% | -14.37% | 6.923% | 1.395% | 8.937% | 6.156% | -4.009% | 5.628% | -1.66% | -9.743% | 3.95% | 0.347% | 1.169% | 4.052% | 5.101% | 0.3926% | -6.982% | -6.886% | 10.11% | -6.023% | 6.156% | -3.817% | 0.6324%
| Z-Score (Absolute) |  | 13.97 | 1.55 | 0.64 | 0.04 | 2.11 | 0.94 | 0.14 | 1.22 | 0.83 | 0.63 | 0.75 | 0.29 | 1.45 | 0.51 | 0.01 | 0.11 | 0.52 | 0.67 | 0 | NA | 1.04 | 1.39 | 0.92 | 0.83 | 0.6 | 0.03
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 556b (Certified Value 0.016 wt.%) | [28 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | 1.54 | 0.014 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.03 | 0.0298 | 0.015 | 0.020512 | 0.02 | 0.01637 | 0.02 | 0.02 | 0.02 | 0.0149 | 0.02 | 0.01 | 0.02 |  |  |  |  | 1
| 1 | 2 | 1.2 | 0.014 | 0.02 | 0.01 | 0.017 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | 0.0292 | 0.015 | 0.020647 | 0.02 | 0.016736 | 0.02 | 0.02 | 0.02 | 0.0147 | 0.02 | 0.01 | 0.02 |  |  |  |  | 17
| 1 | 3 | 1.51 | 0.015 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.03 | 0.0296 | 0.015 | 0.022717 | 0.02 | 0.015839 | 0.02 | 0.02 | 0.02 | 0.014 | 0.02 | 0.01 | 0.02 |  |  |  |  | 16
| 1 | 4 | 1.41 | 0.014 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | 0.0297 | 0.015 | 0.018627 | 0.02 | 0.016105 | 0.02 | 0.02 | 0.02 | 0.0156 | 0.01 | 0.01 | 0.02 |  |  |  |  | 0.016202
| 1 | 5 | 1.39 | 0.014 | 0.02 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.03 | 0.0296 | 0.015 | 0.026881 | 0.02 | 0.015246 | 0.02 | 0.02 | 0.02 | 0.0147 | 0.02 | 0.01 | 0.02 |  |  |  |  | 180
| 1 | 6 | 1.38 | 0.013 | 0.02 | 0.01 | 0.019 | 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | 0.0296 | 0.015 | 0.026166 | 0.02 | 0.01622 | 0.02 | 0.02 | 0.02 | 0.0153 | 0.02 | 0.01 | 0.02
| Mean |  | 1.405 | 0.014 | 0.02 | 0.01 | 0.018333 | 0.01 | 0.01 | 0.01 | 0.015 | 0.02 | 0.03 | 0.029583 | 0.015 | 0.022592 | 0.02 | 0.016086 | 0.02 | 0.02 | 0.02 | 0.014867 | 0.018333 | 0.01 | 0.02
| Median |  | 1.4 | 0.014 | 0.02 | 0.01 | 0.018 | 0.01 | 0.01 | 0.01 | 0.015 | 0.02 | 0.03 | 0.0296 | 0.015 | 0.021682 | 0.02 | 0.016163 | 0.02 | 0.02 | 0.02 | 0.0148 | 0.02 | 0.01 | 0.02
| Std Dev. |  | 0.12012 | 0.00063246 | 0 | 0 | 0.0010328 | 0 | 0 | 0 | 0.0054772 | 0 | 0 | 0.00020412 | 0 | 0.003317 | 0 | 0.00050741 | 0 | 0 | 0 | 0.00055377 | 0.0040825 | 0 | 0
| Rel.Std.Dev. |  | 8.55% | 4.518% | 0% | 0% | 5.633% | 0% | 0% | 0% | 36.51% | 0% | 0% | 0.69% | 0% | 14.68% | 0% | 3.154% | 0% | 0% | 0% | 3.725% | 22.27% | 0% | 0%
| PDM3 |  | 8572% | -13.59% | 23.44% | -38.28% | 13.15% | -38.28% | -38.28% | -38.28% | -7.421% | 23.44% | 85.16% | 82.59% | -7.421% | 39.43% | 23.44% | -0.717% | 23.44% | 23.44% | 23.44% | -8.244% | 13.15% | -38.28% | 23.44%
| Z-Score (Absolute) |  | 269.72 | 0.84 | 0.32 | 1.62 | 0 | 1.62 | 1.62 | 1.62 | 0.65 | 0.32 | 2.27 | 2.19 | 0.68 | 0.83 | 0.32 | 0.44 | 0.32 | 0.32 | 0.32 | 0.67 | 0 | 1.62 | 0.32
| Analytical results for Nb in OREAS 556b (Certified Value 0.13 ppm) | [28 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 28 | [8 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [8 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [8 empty] | 2
| 1 | 1 | < 5 | 0.14 | 0.1 | 0.16 | 0.1 | 0.07 | 0.15 | 0.13 | 0.1 | 0.17 | < 0.2 | < 0.2 | 0.23325 | 1 | < 0.1 | 0.11 | 0.06 | 0.2 | 0.18 | [8 empty] | 1
| 1 | 2 | < 5 | 0.13 | 0.1 | 0.16 | 0.1 | 0.07 | 0.15 | 0.14 | 0.09 | 0.17 | < 0.2 | < 0.2 | 0.235 | 2 | < 0.1 | 0.1 | 0.1 | 0.12 | 0.17 | [8 empty] | 9
| 1 | 3 | < 5 | 0.13 | 0.1 | 0.17 | 0.13 | 0.08 | 0.14 | 0.12 | 0.09 | 0.15 | < 0.2 | < 0.2 | 0.2401 | 2 | < 0.1 | 0.11 | 0.09 | 0.12 | 0.18 | [8 empty] | 16
| 1 | 4 | < 5 | 0.14 | 0.1 | 0.16 | 0.13 | 0.07 | 0.15 | 0.13 | 0.1 | 0.17 | < 0.2 | < 0.2 | 0.25225 | 1 | < 0.1 | 0.1 | 0.08 | 0.13 | 0.16 | [8 empty] | 0.12569
| 1 | 5 | < 5 | 0.16 | 0.1 | 0.16 | 0.13 | 0.06 | 0.15 | 0.14 | 0.09 | 0.17 | < 0.2 | < 0.2 | 0.27582 | 2 | < 0.1 | 0.1 | 0.08 | 0.11 | 0.17 | [8 empty] | 181
| 1 | 6 | < 5 | 0.14 | 0.1 | 0.16 | 0.13 | 0.07 | 0.15 | 0.14 | 0.09 | 0.18 | < 0.2 | < 0.2 | 0.21508 | 1 | < 0.1 | 0.09 | 0.07 | 0.12 | 0.17
| Mean |  |  | 0.14 | 0.1 | 0.16167 | 0.12 | 0.07 | 0.14833 | 0.13333 | 0.093333 | 0.16833 |  |  | 0.24192 | 1.5 |  | 0.10167 | 0.08 | 0.13333 | 0.17167
| Median |  |  | 0.14 | 0.1 | 0.16 | 0.13 | 0.07 | 0.15 | 0.135 | 0.09 | 0.17 |  |  | 0.23755 | 1.5 |  | 0.1 | 0.08 | 0.12 | 0.17
| Std Dev. |  |  | 0.010954 | 1.5202e-17 | 0.0040825 | 0.015492 | 0.0063246 | 0.0040825 | 0.008165 | 0.005164 | 0.0098319 |  |  | 0.020507 | 0.54772 |  | 0.0075277 | 0.014142 | 0.033267 | 0.0075277
| Rel.Std.Dev. |  |  | 7.825% | 1.52e-14% | 2.525% | 12.91% | 9.035% | 2.752% | 6.124% | 5.533% | 5.841% |  |  | 8.477% | 36.51% |  | 7.404% | 17.68% | 24.95% | 4.385%
| PDM3 |  |  | 11.38% | -20.44% | 28.62% | -4.53% | -44.31% | 18.01% | 6.077% | -25.75% | 33.92% |  |  | 92.46% | 1093% |  | -19.12% | -36.35% | 6.077% | 36.57%
| Z-Score (Absolute) |  | 47.88 | 0.13 | NA | 0.57 | 0.27 | 1.28 | 0.3 | 0 | 0.81 | 0.71 | 0.67 | 0.67 | 2.2 | NA | 1.69 | 0.64 | 1.08 | 0 | 0.78
Results from laboratory 5 were removed due to their 0.1 ppm reading resolution. Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 556b (Certified Value 23 ppm) | [28 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 26 |  |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS |  |  |  |  | 1
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 01g |  |  |  |  | 1
| 1 | 1 | 01: 24.2; 02: 25.6; 03: 18.5; 04: 10.1829; 05: 20.033; 06: 27.81 |  |  |  |  | 1
| 1 | 2 | 01: 24.246; 02: 22.2; 03: 19.2; 04: 9.8233; 05: 19.3085; 06: 28.06 |  |  |  |  | 10
| 1 | 3 | 01: 23.819; 02: 24.2; 03: 18.6; 04: 10.0901; 05: 19.4721; 06: 27.28 |  |  |  |  | 16
| 1 | 4 | 01: 24.801; 02: 25.1; 03: 19.2; 04: 9.7269; 05: 20.2785; 06: 28.23 |  |  |  |  | 22.9668
| 1 | 5 | 01: 24.103; 02: 23.7; 03: 19.7; 04: 9.8017; 05: 18.6143; 06: 28.67 |  |  |  |  | 182
| 1 | 6 | 01: 24.034; 02: 23.2; 03: 19.1; 04: 9.7204; 05: 19.035; 06: 28.71
| Mean |  | 01: 24.2005; 02: 24; 03: 19.05; 04: 9.8909; 05: 19.4569; 06: 28.1267
| Median |  | 01: 24.1515; 02: 23.95; 03: 19.15; 04: 9.8125; 05: 19.3903; 06: 28.145
| Std Dev. |  | 01: 0.33029; 02: 1.2474; 03: 0.44159; 04: 0.19668; 05: 0.61921; 06: 0.5418
| Rel.Std.Dev. |  | 01: 1.365%; 02: 5.197%; 03: 2.318%; 04: 1.989%; 05: 3.182%; 06: 1.926%
| PDM3 |  | 01: 5.372%; 02: 4.499%; 03: -17.05%; 04: -56.93%; 05: -15.28%; 06: 22.47%
| Z-Score (Absolute) |  | 01: 0.65; 02: 0.59; 03: 0.7; 04: 3.1; 05: 0.59; 06: 1.68
| Analytical results for Ni in OREAS 556b (Certified Value 323 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 0
| 1 | 1 | 270.3 | 311.4 | 321 | 379.4 | 326 | 374 | 326.272 | 322 | 330 | 312 | 324 | 322 | 308 | 287 | 343.1051 | 296.39 | 352 | 319.3126 | 359 | 307 | 312.6 | 218 | 315 | 346 | 324 | 340.8 |  | 1
| 1 | 2 | 264.7 | 316.8 | 324 | 372.3 | 276 | 354 | 318.0602 | 321 | 342 | 318 | 315 | 332 | 304 | 306 | 335.5764 | 298.65 | 350 | 325.9667 | 358 | 313 | 310.8 | 210 | 296 | 346 | 328.5 | 339.2 |  | 40
| 1 | 3 | 278.5 | 320.6 | 314 | 378.7 | 297 | 371 | 317.9484 | 314 | 339 | 314 | 323 | 326 | 305 | 302 | 335.6113 | 299.72 | 348 | 328.9031 | 349 | 301 | 311 | 201 | 303 | 342 | 325 | 337.2 |  | 16
| 1 | 4 | 263.2 | 318 | 311 | 383.6 | 301 | 354 | 317.4041 | 306 | 332 | 315 | 318 | 328 | 305 | 297 | 346.2754 | 295.58 | 359 | 333.2133 | 332 | 306 | 310.4 | 231 | 314 | 336 | 326.9 | 333.6 |  | 323.2012
| 1 | 5 | 290.4 | 325.8 | 320 | 376.7 | 294 | 351 | 325.196 | 309 | 337 | 317 | 323 | 330 | 306 | 301 | 341.0489 | 295.29 | 355 | 337.7458 | 335 | 310 | 312.9 | 206 | 302 | 349 | 327.1 | 337.9 |  | 183
| 1 | 6 | 286.9 | 348.3 | 316 | 363.6 | 286 | 357 | 311.0387 | 314 | 350 | 328 | 316 | 323 | 307 | 297 | 346.6376 | 295.59 | 355 | 327.5181 | 340 | 306 | 311.9 | 207 | 296 | 349 | 323.4 | 341.3
| Mean |  | 275.6667 | 323.4833 | 317.6667 | 375.7167 | 296.6667 | 360.1667 | 319.3199 | 314.3333 | 338.3333 | 317.3333 | 319.8333 | 326.8333 | 305.8333 | 298.3333 | 341.3758 | 296.87 | 353.1667 | 328.7766 | 345.5 | 307.1667 | 311.6 | 212.1667 | 304.3333 | 344.6667 | 325.8167 | 338.3333
| Median |  | 274.4 | 319.3 | 318 | 377.7 | 295.5 | 355.5 | 318.0043 | 314 | 338 | 316 | 320.5 | 327 | 305.5 | 299 | 342.077 | 295.99 | 353.5 | 328.2106 | 344.5 | 306.5 | 311.45 | 208.5 | 302.5 | 346 | 325.95 | 338.55
| Std Dev. |  | 11.4523 | 13.0406 | 4.8442 | 6.9867 | 16.8958 | 9.786 | 5.6316 | 6.3456 | 7.2296 | 5.6451 | 3.9707 | 3.92 | 1.472 | 6.5013 | 4.9326 | 1.8612 | 3.9707 | 6.3118 | 11.606 | 4.0702 | 1.0218 | 10.7966 | 8.4063 | 4.9666 | 1.9914 | 2.8112
| Rel.Std.Dev. |  | 4.154% | 4.031% | 1.525% | 1.86% | 5.695% | 2.717% | 1.764% | 2.019% | 2.137% | 1.779% | 1.241% | 1.199% | 0.4813% | 2.179% | 1.445% | 0.6269% | 1.124% | 1.92% | 3.359% | 1.325% | 0.3279% | 5.089% | 2.762% | 1.441% | 0.6112% | 0.8309%
| PDM3 |  | -14.71% | 0.08731% | -1.712% | 16.25% | -8.21% | 11.44% | -1.201% | -2.744% | 4.682% | -1.816% | -1.042% | 1.124% | -5.374% | -7.694% | 5.623% | -8.147% | 9.271% | 1.725% | 6.899% | -4.961% | -3.589% | -34.35% | -5.838% | 6.642% | 0.8093% | 4.682%
| Z-Score (Absolute) |  | 1.74 | 0.15 | 0.08 | 2.23 | 0.91 | 1.61 | 0.01 | 0.21 | 0.74 | 0.09 | 0.01 | 0.29 | 0.55 | 0.84 | 0.86 | 0.9 | 1.33 | 0.36 | 1.03 | 0.49 | 0.32 | 4.26 | 0.6 | 1 | 0.25 | 0.74
| Analytical results for P in OREAS 556b (Certified Value 0.019 wt.%) | [28 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02 | 0.0164 | 0.0706 | 0.0185 | 0.045 | 0.051 | 0.019 | 0.018 | 0.018 | 0.017 | 0.018 | 0.0855 | 0.2583 | 0.019156 | 0.0233 | 0.018179 | 0.02 | 0.02 | 0.016 | 0.018 | 0.0212 | 0.0204 | 0.018 | 0.0179 |  |  |  | 1
| 1 | 2 | 0.02 | 0.0168 | 0.0696 | 0.0185 | 0.044 | 0.05 | 0.019 | 0.019 | 0.017 | 0.017 | 0.018 | 0.085 | 0.2479 | 0.017926 | 0.0232 | 0.017738 | 0.02 | 0.02 | 0.017 | 0.019 | 0.0197 | 0.0205 | 0.018 | 0.0177 |  |  |  | 0
| 1 | 3 | 0.02 | 0.0169 | 0.0681 | 0.0184 | 0.044 | 0.049 | 0.018 | 0.019 | 0.017 | 0.017 | 0.018 | 0.0857 | 0.2365 | 0.018002 | 0.0233 | 0.018647 | 0.02 | 0.02 | 0.016 | 0.02 | 0.0215 | 0.0201 | 0.018 | 0.0178 |  |  |  | 16
| 1 | 4 | 0.02 | 0.0165 | 0.0671 | 0.0189 | 0.045 | 0.046 | 0.018 | 0.019 | 0.017 | 0.017 | 0.018 | 0.0856 | 0.2378 | 0.017181 | 0.0225 | 0.019002 | 0.02 | 0.02 | 0.016 | 0.02 | 0.0251 | 0.0196 | 0.018 | 0.0177 |  |  |  | 0.018787
| 1 | 5 | 0.02 | 0.0175 | 0.0668 | 0.0183 | 0.045 | 0.048 | 0.018 | 0.019 | 0.018 | 0.017 | 0.018 | 0.0857 | 0.237 | 0.017716 | 0.023 | 0.019368 | 0.02 | 0.02 | 0.016 | 0.018 | 0.0209 | 0.0205 | 0.018 | 0.0173 |  |  |  | 184
| 1 | 6 | 0.02 | 0.019 | 0.0677 | 0.0186 | 0.043 | 0.048 | 0.018 | 0.02 | 0.019 | 0.017 | 0.018 | 0.0855 | 0.2509 | 0.019282 | 0.023 | 0.018721 | 0.02 | 0.02 | 0.014 | 0.02 | 0.0207 | 0.0203 | 0.017 | 0.018
| Mean |  | 0.02 | 0.017183 | 0.068317 | 0.018533 | 0.044333 | 0.048667 | 0.018333 | 0.019 | 0.017667 | 0.017 | 0.018 | 0.0855 | 0.24473 | 0.01821 | 0.02305 | 0.018609 | 0.02 | 0.02 | 0.015833 | 0.019167 | 0.021517 | 0.020233 | 0.017833 | 0.017733
| Median |  | 0.02 | 0.01685 | 0.0679 | 0.0185 | 0.0445 | 0.0485 | 0.018 | 0.019 | 0.0175 | 0.017 | 0.018 | 0.08555 | 0.24285 | 0.017964 | 0.0231 | 0.018684 | 0.02 | 0.02 | 0.016 | 0.0195 | 0.02105 | 0.02035 | 0.018 | 0.01775
| Std Dev. |  | 0 | 0.0009704 | 0.0014878 | 0.00020656 | 0.0008165 | 0.0017512 | 0.0005164 | 0.00063246 | 0.0008165 | 0 | 0 | 0.00026077 | 0.0090308 | 0.00083303 | 0.00030166 | 0.00058102 | 0 | 0 | 0.00098319 | 0.00098319 | 0.0018595 | 0.00034448 | 0.00040825 | 0.00024221
| Rel.Std.Dev. |  | 0% | 5.647% | 2.178% | 1.115% | 1.842% | 3.598% | 2.817% | 3.329% | 4.622% | 0% | 0% | 0.305% | 3.69% | 4.574% | 1.309% | 3.122% | 0% | 0% | 6.21% | 5.13% | 8.642% | 1.703% | 2.289% | 1.366%
| PDM3 |  | 6.458% | -8.535% | 263.6% | -1.349% | 136% | 159% | -2.414% | 1.135% | -5.962% | -9.511% | -4.188% | 355.1% | 1203% | -3.068% | 22.69% | -0.9457% | 6.458% | 6.458% | -15.72% | 2.022% | 14.53% | 7.7% | -5.075% | -5.608%
| Z-Score (Absolute) |  | 0.48 | 0.99 | 25.54 | 0.29 | 13.1 | 15.34 | 0.39 | 0.04 | 0.73 | 1.08 | 0.56 | 34.45 | 117.05 | 0.45 | 2.06 | 0.25 | 0.48 | 0.48 | 1.69 | 0.04 | 1.26 | 0.6 | 0.65 | 0.7
| Analytical results for Pb in OREAS 556b (Certified Value 14.9 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 25 | 26 | 27 | 28 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 2
| 1 | 1 | 44.7 | 11.05 | 15.3 | 14.8 | 14.7 | 16.4 | 14.1419 | 13.4 | 15 | 14 | 14.6 | 13.6 | 21.09 | 14.6 | 14.8846 | 15.1 | 16.6 | 16.344 | 14 | 15.2 | 21 | 13.8 | 18.7 | 15.2 | 14.7 | 15.4 |  | 1
| 1 | 2 | 45.7 | 12.03 | 15.9 | 14.7 | 15.5 | 15.8 | 14.237 | 13.6 | 15.1 | 14.4 | 14.2 | 13.8 | 20.53 | 15.4 | 15.1207 | 15.14 | 16.3 | 15.552 | 14 | 15.1 | 20 | 13.8 | 17.5 | 15.3 | 14.6 | 15 |  | 11
| 1 | 3 | 46.2 | 11.84 | 15.7 | 14.7 | 13.9 | 15.5 | 14.3042 | 13.3 | 15.4 | 14.2 | 14.5 | 13.8 | 20.76 | 15 | 15.111 | 15.76 | 16.2 | 16.2527 | 13 | 14.9 | 22 | 13.6 | 18.1 | 15.3 | 14.7 | 15 |  | 16
| 1 | 4 | 46.6 | 11.96 | 15.3 | 14.9 | 14.4 | 15.3 | 14.3078 | 13.2 | 14.8 | 14 | 14.4 | 14 | 20.84 | 15.4 | 15.0036 | 15.93 | 16.9 | 17.3071 | 14 | 14.6 | 23 | 13.8 | 19.8 | 14.8 | 14.4 | 14.9 |  | 14.8661
| 1 | 5 | 46.8 | 12.82 | 15.7 | 14.9 | 14 | 15.2 | 14.1968 | 13.5 | 14.6 | 14.2 | 14.2 | 14.2 | 20.85 | 15.1 | 15.8055 | 15 | 16.3 | 14.3172 | 13 | 14.4 | 22 | 13.5 | 18.4 | 15.3 | 14.7 | 15 |  | 185
| 1 | 6 | 48.6 | 9.65 | 15.1 | 14.8 | 13.6 | 15.9 | 14.4841 | 13.5 | 15 | 14.9 | 14 | 13.8 | 20.86 | 14.9 | 15.6531 | 15.68 | 16 | 16.1219 | 15 | 15.4 | 21 | 13.8 | 18.3 | 15.3 | 14.8 | 15.5
| Mean |  | 46.4333 | 11.5583 | 15.5 | 14.8 | 14.35 | 15.6833 | 14.2786 | 13.4167 | 14.9833 | 14.2833 | 14.3167 | 13.8667 | 20.8217 | 15.0667 | 15.2631 | 15.435 | 16.3833 | 15.9825 | 13.8333 | 14.9333 | 21.5 | 13.7167 | 18.4667 | 15.2 | 14.65 | 15.1333
| Median |  | 46.4 | 11.9 | 15.5 | 14.8 | 14.2 | 15.65 | 14.2706 | 13.45 | 15 | 14.2 | 14.3 | 13.8 | 20.845 | 15.05 | 15.1159 | 15.41 | 16.3 | 16.1873 | 14 | 15 | 21.5 | 13.8 | 18.35 | 15.3 | 14.7 | 15
| Std Dev. |  | 1.3003 | 1.0913 | 0.30984 | 0.089443 | 0.68337 | 0.4446 | 0.11904 | 0.1472 | 0.27142 | 0.33714 | 0.22286 | 0.20656 | 0.18082 | 0.30768 | 0.37426 | 0.39979 | 0.31885 | 0.9936 | 0.75277 | 0.37771 | 1.0488 | 0.13292 | 0.76594 | 0.2 | 0.13784 | 0.25033
| Rel.Std.Dev. |  | 2.8% | 9.442% | 1.999% | 0.6043% | 4.762% | 2.835% | 0.8337% | 1.097% | 1.811% | 2.36% | 1.557% | 1.49% | 0.8684% | 2.042% | 2.452% | 2.59% | 1.946% | 6.217% | 5.442% | 2.529% | 4.878% | 0.969% | 4.148% | 1.316% | 0.9409% | 1.654%
| PDM3 |  | 212.3% | -22.25% | 4.264% | -0.4448% | -3.472% | 5.497% | -3.952% | -9.75% | 0.7884% | -3.92% | -3.696% | -6.723% | 40.06% | 1.349% | 2.67% | 3.827% | 10.21% | 7.509% | -6.947% | 0.4521% | 44.62% | -7.732% | 24.22% | 2.246% | -1.454% | 1.797%
| Z-Score (Absolute) |  | 30.62 | 3.38 | 0.46 | 0.22 | 0.66 | 0.64 | 0.73 | 1.57 | 0.04 | 0.72 | 0.69 | 1.13 | 5.65 | 0.04 | 0.23 | 0.4 | 1.32 | 0.93 | NA | 0.09 | NA | 1.28 | 3.36 | 0.17 | 0.37 | 0.11
Results from laboratories 19 and 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 556b (Indicative Value < 10 ppb) | [28 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 01: 04; 02: 06; 03: 07; 04: 15; 05: 26 |  | ppb
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS |  | 1
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 01g; 05: 01g |  | 1
| 1 | 1 | 01: < 10; 02: 14; 03: < 20; 04: < 10; 05: < 10 |  | 1
| 1 | 2 | 01: < 10; 02: 26; 03: 21.01; 04: < 10; 05: < 10 |  | 26
| 1 | 3 | 01: < 10; 02: 17; 03: < 20; 04: < 10; 05: < 10 |  | 16
| 1 | 4 | 01: < 10; 02: 24; 03: < 20; 04: < 10; 05: < 10 |  | < 10
| 1 | 5 | 01: < 10; 02: 19; 03: 22.8; 04: < 10; 05: < 10 |  | 32
| 1 | 6 | 01: < 10; 02: 24; 03: < 20; 04: < 10; 05: < 10
| Mean |  | 02: 20.6667; 03: 21.905
| Median |  | 02: 21.5; 03: 21.905
| Std Dev. |  | 02: 4.7188; 03: 1.2657
| Rel.Std.Dev. |  | 02: 22.83%; 03: 5.778%
PDM3
Z-Score (Absolute) | bar: 01 to 05
| Analytical results for Pr in OREAS 556b (Certified Value 5.75 ppm) | [28 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 26 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 01g |  | 2
| 1 | 1 | 01: 6.12; 02: 6.3; 03: 4.57; 04: 2.6516; 05: 4.9914; 06: 6.96 |  | 1
| 1 | 2 | 01: 6.215; 02: 5.5; 03: 4.73; 04: 2.4978; 05: 4.8221; 06: 7.17 |  | 12
| 1 | 3 | 01: 6.149; 02: 6; 03: 4.65; 04: 2.6478; 05: 4.9473; 06: 6.98 |  | 16
| 1 | 4 | 01: 6.352; 02: 6.2; 03: 4.81; 04: 2.5407; 05: 5.3858; 06: 6.85 |  | 5.7531
| 1 | 5 | 01: 6.147; 02: 5.9; 03: 4.9; 04: 2.5337; 05: 4.9144; 06: 6.93 |  | 186
| 1 | 6 | 01: 6.236; 02: 5.8; 03: 4.76; 04: 2.5028; 05: 4.7287; 06: 7.08
| Mean |  | 01: 6.2032; 02: 5.95; 03: 4.7367; 04: 2.5624; 05: 4.965; 06: 6.995
| Median |  | 01: 6.182; 02: 5.95; 03: 4.745; 04: 2.5372; 05: 4.9309; 06: 6.97
| Std Dev. |  | 01: 0.085293; 02: 0.2881; 03: 0.11656; 04: 0.069689; 05: 0.22666; 06: 0.11362
| Rel.Std.Dev. |  | 01: 1.375%; 02: 4.842%; 03: 2.461%; 04: 2.72%; 05: 4.565%; 06: 1.624%
| PDM3 |  | 01: 7.823%; 02: 3.422%; 03: -17.67%; 04: -55.46%; 05: -13.7%; 06: 21.59%
| Z-Score (Absolute) |  | 01: 0.69; 02: 0.45; 03: 0.66; 04: 2.66; 05: 0.45; 06: 1.41
| Analytical results for Pt in OREAS 556b (Indicative Value < 5 ppb) | [28 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 04; 02: 06; 03: 07; 04: 15; 05: 25; 06: 26 |  | ppb
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 06: AR*MS |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 01g; 05: 0.1g; 06: 01g |  | 2
| 1 | 1 | 01: < 5; 02: < 10; 03: < 6.44; 04: < 5; 05: 3; 06: < 5 |  | 1
| 1 | 2 | 01: < 5; 02: < 10; 03: 15.75; 04: < 5; 05: 3; 06: < 5 |  | 22
| 1 | 3 | 01: < 5; 02: < 10; 03: < 5; 04: < 5; 05: 2; 06: < 5 |  | 16
| 1 | 4 | 01: < 5; 02: < 10; 03: < 6.83; 04: < 5; 05: 4; 06: < 5 |  | < 5
| 1 | 5 | 01: < 5; 02: < 10; 03: 12.74; 04: < 5; 05: 5; 06: < 5 |  | 30
| 1 | 6 | 01: < 5; 02: < 10; 03: 8.76; 04: < 5; 05: 5; 06: < 5
| Mean |  | 03: 12.4167; 05: 3.6667
| Median |  | 03: 12.74; 05: 3.5
| Std Dev. |  | 03: 3.5062; 05: 1.2111
| Rel.Std.Dev. |  | 03: 28.24%; 05: 33.03%
PDM3
| Z-Score (Absolute) |  | 01: 0.67; 02: 2.22; 03: 5.37; 04: 0.67; 05: 0.67; 06: 0.67
| Analytical results for Rb in OREAS 556b (Certified Value 6.1 ppm) | [28 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 22 | 23 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [10 empty] | 2
| 1 | 1 | 7.86 | 8.3 | 5.2 | 5.2 | 6.8 | 5.1 | 5.5 | 5.6 | 4.31 | 7.9484 | 11.9 | 5.8987 | 5.7 | 6.39 | 11.5 | 12.03 | 6 | [10 empty] | 1
| 1 | 2 | 7.68 | 7.2 | 5.1 | 5.4 | 7 | 5.4 | 5.4 | 5.6 | 4.4 | 7.2305 | 11.8 | 6.0992 | 5.8 | 6.01 | 14 | 12.61 | 6.6 | [10 empty] | 13
| 1 | 3 | 7.96 | 8.1 | 5.3 | 5.1 | 7.1 | 5.2 | 5.8 | 5.5 | 4.43 | 7.335 | 12 | 6.0165 | 5.6 | 6.1 | 10.8 | 11.64 | 6.1 | [10 empty] | 16
| 1 | 4 | 8.06 | 8.3 | 5.1 | 5.2 | 7.1 | 5 | 5.5 | 5.8 | 4.44 | 7.236 | 12 | 5.8752 | 5.5 | 6.18 | 17.8 | 11.7 | 6.1 | [10 empty] | 6.096
| 1 | 5 | 7.8 | 8 | 5 | 5.1 | 6.9 | 5.2 | 5.6 | 5.8 | 4.47 | 7.2266 | 12.2 | 6.788 | 5.5 | 6.35 | 13.3 | 11.88 | 6.1 | [10 empty] | 187
| 1 | 6 | 7.87 | 7.7 | 5.2 | 5.2 | 7.1 | 5.2 | 5.4 | 5.6 | 4.41 | 8.0479 | 12.1 | 5.9518 | 5.4 | 6.39 | 14.4 | 12.15 | 6.3
| Mean |  | 7.8717 | 7.9333 | 5.15 | 5.2 | 7 | 5.1833 | 5.5333 | 5.65 | 4.41 | 7.5041 | 12 | 6.1049 | 5.5833 | 6.2367 | 13.6333 | 12.0017 | 6.2
| Median |  | 7.865 | 8.05 | 5.15 | 5.2 | 7.05 | 5.2 | 5.5 | 5.6 | 4.42 | 7.2855 | 12 | 5.9841 | 5.55 | 6.265 | 13.65 | 11.955 | 6.1
| Std Dev. |  | 0.1306 | 0.42269 | 0.10488 | 0.10954 | 0.12649 | 0.13292 | 0.15055 | 0.12247 | 0.054772 | 0.38611 | 0.14142 | 0.34444 | 0.1472 | 0.16318 | 2.4809 | 0.35482 | 0.21909
| Rel.Std.Dev. |  | 1.659% | 5.328% | 2.037% | 2.107% | 1.807% | 2.564% | 2.721% | 2.168% | 1.242% | 5.145% | 1.179% | 5.642% | 2.636% | 2.616% | 18.2% | 2.956% | 3.534%
| PDM3 |  | 29.13% | 30.14% | -15.52% | -14.7% | 14.83% | -14.97% | -9.23% | -7.316% | -27.66% | 23.1% | 96.85% | 0.1459% | -8.41% | 2.308% | 123.6% | 96.88% | 1.706%
| Z-Score (Absolute) |  | 1.11 | 1.15 | 0.7 | 0.66 | 0.53 | 0.67 | 0.44 | 0.36 | 1.19 | 0.86 | 3.85 | 0.06 | 0.41 | 0.02 | 4.93 | 3.85 | 0
| Analytical results for Re in OREAS 556b (Certified Value 0.029 ppm) | [28 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 20 | 22 | 25 | 26 | 28 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [11 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.1g | 01g | 0.5g | [11 empty] | 3
| 1 | 1 | 0.031 | 0.028 | 0.031 | 0.029 | 0.03 | 0.029 | 0.03 | 0.029 | < 0.05 | < 0.05 | 0.026773 | 0.027 | 0.029 | 0.027 | 0.032 | 0.027 | [11 empty] | 1
| 1 | 2 | 0.03 | 0.025 | 0.031 | 0.029 | 0.029 | 0.03 | 0.031 | 0.028 | < 0.05 | < 0.05 | 0.024692 | 0.031 | 0.03 | 0.024 | 0.032 | 0.026 | [11 empty] | 14
| 1 | 3 | 0.024 | 0.026 | 0.033 | 0.029 | 0.03 | 0.026 | 0.028 | 0.028 | < 0.05 | < 0.05 | 0.028079 | 0.029 | 0.031 | 0.024 | 0.034 | 0.025 | [11 empty] | 16
| 1 | 4 | 0.032 | 0.027 | 0.033 | 0.027 | 0.031 | 0.028 | 0.032 | 0.029 | < 0.05 | < 0.05 | 0.030671 | 0.029 | 0.029 | 0.025 | 0.031 | 0.025 | [11 empty] | 0.029321
| 1 | 5 | 0.031 | 0.026 | 0.033 | 0.03 | 0.03 | 0.029 | 0.026 | 0.03 | < 0.05 | < 0.05 | 0.031443 | 0.025 | 0.03 | 0.024 | 0.032 | 0.025 | [11 empty] | 188
| 1 | 6 | 0.033 | 0.026 | 0.033 | 0.03 | 0.03 | 0.03 | 0.033 | 0.029 | < 0.05 | < 0.05 | 0.027007 | 0.03 | 0.031 | 0.025 | 0.033 | 0.026
| Mean |  | 01: 0.030167; 02: 0.026333; 03: 0.032333; 04: 0.029; 05: 0.03; 06: 0.028667; 07: 0.03; 08: 0.028833; 11: 0.028111; 12: 0.0285; 13: 0.03; 14: 0.024833; 15: 0.032333; 16: 0.025667
| Median |  | 01: 0.031; 02: 0.026; 03: 0.033; 04: 0.029; 05: 0.03; 06: 0.029; 07: 0.0305; 08: 0.029; 11: 0.027543; 12: 0.029; 13: 0.03; 14: 0.0245; 15: 0.032; 16: 0.0255
| Std Dev. |  | 01: 0.0031885; 02: 0.0010328; 03: 0.0010328; 04: 0.0010954; 05: 0.00063246; 06: 0.0015055; 07: 0.0026077; 08: 0.00075277; 11: 0.0025439; 12: 0.0021679; 13: 0.00089443; 14: 0.001169; 15: 0.0010328; 16: 0.0008165
| Rel.Std.Dev. |  | 01: 10.57%; 02: 3.922%; 03: 3.194%; 04: 3.777%; 05: 2.108%; 06: 5.252%; 07: 8.692%; 08: 2.611%; 11: 9.049%; 12: 7.607%; 13: 2.981%; 14: 4.708%; 15: 3.194%; 16: 3.181%
| PDM3 |  | 01: 2.883%; 02: -10.19%; 03: 10.27%; 04: -1.096%; 05: 2.315%; 06: -2.233%; 07: 2.315%; 08: -1.664%; 11: -4.128%; 12: -2.801%; 13: 2.315%; 14: -15.31%; 15: 10.27%; 16: -12.46%
| Z-Score (Absolute) |  | 0.72 | 1.22 | 1.81 | 0.13 | 0.63 | 0.04 | 0.63 | 0.04 | 1.9 | 1.9 | 0.32 | 0.13 | 0.63 | 1.98 | 1.81 | 1.56
| Analytical results for S in OREAS 556b (Certified Value 8.38 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/AAS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  | 3
| 1 | 1 | 6.26 | 7.59 | 7.48 | 7.8786 | 7.726 | 7.83 | 8.97 | 8.8 | 8.43 | 8.73 | 8.86 | 8.75 | 9.4 | 8.5128 | 7.9833 | > 5 | 8.6053 | 7.6503 | 8.32 | 8.63 | 8.2 | > 5 | 8.27 | > 5 | 8.5 | 9.01 |  | 1
| 1 | 2 | 5.96 | 7.59 | 7.49 | 7.9502 | 7.652 | 7.87 | 8.97 | 8.77 | 8.68 | 8.68 | 9.06 | 8.77 | 9.86 | 8.5745 | 7.9854 | > 5 | 8.3945 | 7.9429 | 8.43 | 8.63 | 8.25 | > 5 | 7.97 | > 5 | 8.45 | 8.87 |  | 28
| 1 | 3 | 5.78 | 7.9 | 7.46 | 8.1186 | 7.594 | 7.79 | 8.66 | 8.74 | 8.56 | 8.69 | 9.07 | 8.76 | 9.34 | 8.6414 | 7.9375 | > 5 | 8.8524 | 7.8572 | 8.34 | 8.83 | 8.14 | > 5 | 8.39 | > 5 | 8.47 | 8.9 |  | 16
| 1 | 4 | 6.43 | 7.76 | 7.32 | 7.9342 | 7.687 | 7.42 | 8.76 | 8.81 | 8.69 | 8.7 | 8.93 | 8.74 | 9.69 | 8.3947 | 7.908 | > 5 | 8.2958 | 7.8291 | 8.36 | 8.59 | 8.02 | > 5 | 7.93 | > 5 | 8.54 | 8.83 |  | 8.3791
| 1 | 5 | 6.02 | 7.88 | 7.32 | 7.9115 | 7.727 | 7.76 | 8.74 | 8.58 | 8.64 | 8.73 | 9.13 | 8.72 | 10.2 | 8.4217 | 7.9309 | > 5 | 8.4911 | 7.8725 | 8.47 | 8.55 | 7.92 | > 5 | 8 | > 5 | 8.51 | 8.88 |  | 189
| 1 | 6 | 6.17 | 8.56 | 7.59 | 7.9557 | 7.487 | 7.74 | 8.78 | 8.67 | 8.94 | 8.62 | 8.99 | 8.73 | 9.74 | 8.5096 | 7.9027 | > 5 | 8.8573 | 7.8491 | 8.51 | 8.68 | 7.95 | > 5 | 7.94 | > 5 | 8.5 | 9.04
| Mean |  | 6.1033 | 7.88 | 7.4433 | 7.9581 | 7.6455 | 7.735 | 8.8133 | 8.7283 | 8.6567 | 8.6917 | 9.0067 | 8.745 | 9.705 | 8.5091 | 7.9413 |  | 8.5827 | 7.8335 | 8.405 | 8.6517 | 8.08 |  | 8.0833 |  | 8.495 | 8.9217
| Median |  | 6.095 | 7.82 | 7.47 | 7.9422 | 7.6695 | 7.775 | 8.77 | 8.755 | 8.66 | 8.695 | 9.025 | 8.745 | 9.715 | 8.5112 | 7.9342 |  | 8.5482 | 7.8532 | 8.395 | 8.63 | 8.08 |  | 7.985 |  | 8.5 | 8.89
| Std Dev. |  | 0.2314 | 0.35928 | 0.10558 | 0.083536 | 0.09237 | 0.16134 | 0.12801 | 0.088412 | 0.16908 | 0.040702 | 0.099733 | 0.018708 | 0.3152 | 0.092225 | 0.035846 |  | 0.23442 | 0.097853 | 0.076616 | 0.097656 | 0.13638 |  | 0.19633 |  | 0.031464 | 0.083766
| Rel.Std.Dev. |  | 3.791% | 4.559% | 1.418% | 1.05% | 1.208% | 2.086% | 1.452% | 1.013% | 1.953% | 0.4683% | 1.107% | 0.2139% | 3.248% | 1.084% | 0.4514% |  | 2.731% | 1.249% | 0.9116% | 1.129% | 1.688% |  | 2.429% |  | 0.3704% | 0.9389%
| PDM3 |  | -27.16% | -5.956% | -11.17% | -5.024% | -8.755% | -7.687% | 5.182% | 4.168% | 3.313% | 3.73% | 7.49% | 4.367% | 15.82% | 1.552% | -5.224% |  | 2.43% | -6.511% | 0.3092% | 3.253% | -3.569% |  | -3.53% |  | 1.383% | 6.475%
| Z-Score (Absolute) |  | 3.89 | 1 | 1.71 | 0.87 | 1.38 | 1.23 | 0.52 | 0.38 | 0.26 | 0.32 | 0.83 | 0.41 | 1.97 | 0.02 | 0.9 | NA | 0.14 | 1.07 | 0.15 | 0.25 | 0.67 | NA | 0.67 | NA | 0 | 0.69
| Analytical results for Sb in OREAS 556b (Certified Value 3.32 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | < 5 | 3.52 | < 5 | 4.61 | 3.75 | 3.15 | 3.7015 | 2.68 | 2.79 | 3 | 3.39 | 2.52 | < 5 | 0.97 | 2.2794 | 3.22 | 3.2309 | NA | 7 | 3.37 | 2.11 | 3.56 | 3.7 | 3.73 |  |  |  | 1
| 1 | 2 | < 5 | 3.86 | < 5 | 4.45 | 3.38 | 3.09 | 3.6045 | 2.85 | 2.79 | 3.25 | 3.4 | 2.32 | < 5 | 1.09 | 2.4105 | 3.29 | 3.2412 | NA | 8 | 3.31 | 2.03 | 3.72 | 3.6 | 3.64 |  |  |  | 15
| 1 | 3 | < 5 | 3.58 | < 5 | 4.38 | 3.39 | 3.23 | 3.2736 | 2.58 | 2.83 | 3.11 | 3.41 | 2.36 | < 5 | 1.26 | 2.3553 | 3.35 | 3.235 | NA | 8 | 3.56 | 2.79 | 3.55 | 3.6 | 3.65 |  |  |  | 16
| 1 | 4 | < 5 | 3.5 | < 5 | 4.52 | 3.55 | 3.12 | 3.3233 | 2.71 | 2.88 | 3.19 | 3.4 | 2.49 | < 5 | 1.09 | 2.3445 | 3.2 | 3.7061 | NA | 8 | 3.43 | 2.27 | 3.42 | 3.6 | 3.53 |  |  |  | 3.3222
| 1 | 5 | < 5 | 3.8 | < 5 | 4.52 | 3.41 | 3.17 | 3.5619 | 2.71 | 2.67 | 3.2 | 3.53 | 2.42 | < 5 | 1.06 | 2.3179 | 3.13 | 3.5982 | NA | 8 | 3.26 | 2.17 | 3.31 | 3.5 | 3.68 |  |  |  | 190
| 1 | 6 | < 5 | 3.82 | < 5 | 4.46 | 3.47 | 3.21 | 3.3153 | 2.66 | 2.67 | 3.17 | 3.33 | 2.47 | < 5 | 1.04 | 2.4185 | 3.2 | 3.1301 | NA | 7 | 3.58 | 2.1 | 3.42 | 3.6 | 3.69
| Mean |  |  | 3.68 |  | 4.49 | 3.4917 | 3.1617 | 3.4634 | 2.6983 | 2.7717 | 3.1533 | 3.41 | 2.43 |  | 1.085 | 2.3543 | 3.2317 | 3.3569 |  | 7.6667 | 3.4183 | 2.245 | 3.4967 | 3.6 | 3.6533
| Median |  |  | 3.69 |  | 4.49 | 3.44 | 3.16 | 3.4426 | 2.695 | 2.79 | 3.18 | 3.4 | 2.445 |  | 1.075 | 2.3499 | 3.21 | 3.2381 |  | 8 | 3.4 | 2.14 | 3.485 | 3.6 | 3.665
| Std Dev. |  |  | 0.16395 |  | 0.078486 | 0.14148 | 0.053072 | 0.18105 | 0.088412 | 0.085421 | 0.087788 | 0.065422 | 0.077974 |  | 0.096488 | 0.053495 | 0.077309 | 0.23483 |  | 0.5164 | 0.13075 | 0.27884 | 0.14376 | 0.063246 | 0.068313
| Rel.Std.Dev. |  |  | 4.455% |  | 1.748% | 4.052% | 1.679% | 5.228% | 3.277% | 3.082% | 2.784% | 1.919% | 3.209% |  | 8.893% | 2.272% | 2.392% | 6.995% |  | 6.736% | 3.825% | 12.42% | 4.111% | 1.757% | 1.87%
| PDM3 |  |  | 10.77% |  | 35.15% | 5.1% | -4.833% | 4.248% | -18.78% | -16.57% | -5.084% | 2.642% | -26.86% |  | -67.34% | -29.13% | -2.726% | 1.044% |  | 130.8% | 2.893% | -32.43% | 5.25% | 8.361% | 9.966%
| Z-Score (Absolute) |  | 1.07 | 0.74 | 1.07 | 1.98 | 0.45 | 0.05 | 0.41 | 0.76 | 0.65 | 0.07 | 0.33 | 1.17 | 1.07 | 3.23 | 1.29 | 0.05 | 0.25 | NA | NA | 0.34 | 1.46 | 0.46 | 0.62 | 0.7
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 556b (Certified Value 2.76 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 2
| 1 | 1 | 3.8 | 2.5 | < 5 | 3 | 2.7 | 2.6 | 2.9 | 3 | 2.7 | 2.8 | 2.8 | 2.4 | 2.3609 | 3 | 3.0312 | 4.3 | 2.1 | 2.9 | 3.3 | 4 | 3.3 | 2.6 | 3 |  |  |  |  | 1
| 1 | 2 | 3.7 | 2.6 | < 5 | 2.7 | 2.3 | 2.7 | 3 | 3.1 | 2.7 | 2.8 | 2.9 | 2.6 | 2.208 | 3 | 2.9134 | 4.3 | 2.1 | 2.8 | 3 | 4 | 3.5 | 2.6 | 2.8 |  |  |  |  | 1
| 1 | 3 | 3.7 | 2.7 | < 5 | 2.9 | 2.4 | 2.7 | 2.9 | 3.2 | 2.7 | 2.8 | 2.9 | 2.6 | 2.3513 | 3 | 2.9679 | 4.1 | 2.2 | 2.7 | 2.9 | 4 | 3.3 | 2.6 | 2.8 |  |  |  |  | 16
| 1 | 4 | 3.6 | 2.6 | < 5 | 3 | 2.7 | 2.6 | 3 | 3 | 2.7 | 2.8 | 2.9 | 2.6 | 2.1608 | 3 | 3.3651 | 4 | 2.1 | 2.9 | 3 | 4 | 3.3 | 2.7 | 2.9 |  |  |  |  | 2.762
| 1 | 5 | 3.9 | 2.7 | < 5 | 2.8 | 2.3 | 2.6 | 3 | 3 | 2.8 | 2.7 | 2.9 | 2.6 | 2.2092 | 3 | 3.2181 | 4.1 | 2.2 | 2.5 | 2.9 | 4 | 3.5 | 2.6 | 2.8 |  |  |  |  | 191
| 1 | 6 | 3.8 | 1.5 | < 5 | 3.1 | 2.4 | 2.6 | 3.1 | 3 | 2.8 | 2.8 | 2.8 | 2.6 | 2.4427 | 3 | 3.0023 | 4.4 | 2.1 | 2.6 | 2.9 | 4 | 3.5 | 2.7 | 3
| Mean |  | 3.75 | 2.4333 |  | 2.9167 | 2.4667 | 2.6333 | 2.9833 | 3.05 | 2.7333 | 2.7833 | 2.8667 | 2.5667 | 2.2888 | 3 | 3.083 | 4.2 | 2.1333 | 2.7333 | 3 | 4 | 3.4 | 2.6333 | 2.8833
| Median |  | 3.75 | 2.6 |  | 2.95 | 2.4 | 2.6 | 3 | 3 | 2.7 | 2.8 | 2.9 | 2.6 | 2.2803 | 3 | 3.0168 | 4.2 | 2.1 | 2.75 | 2.95 | 4 | 3.4 | 2.6 | 2.85
| Std Dev. |  | 0.10488 | 0.46332 |  | 0.1472 | 0.18619 | 0.05164 | 0.075277 | 0.083666 | 0.05164 | 0.040825 | 0.05164 | 0.08165 | 0.11139 | 0 | 0.17266 | 0.15492 | 0.05164 | 0.1633 | 0.15492 | 0 | 0.10954 | 0.05164 | 0.098319
| Rel.Std.Dev. |  | 2.797% | 19.04% |  | 5.047% | 7.548% | 1.961% | 2.523% | 2.743% | 1.889% | 1.467% | 1.801% | 3.181% | 4.867% | 0% | 5.6% | 3.689% | 2.421% | 5.974% | 5.164% | 0% | 3.222% | 1.961% | 3.41%
| PDM3 |  | 35.77% | -11.9% |  | 5.602% | -10.69% | -4.657% | 8.015% | 10.43% | -1.036% | 0.7741% | 3.791% | -7.071% | -17.13% | 8.619% | 11.62% | 52.07% | -22.76% | -1.036% | 8.619% | 44.82% | 23.1% | -4.657% | 4.395%
| Z-Score (Absolute) |  | 3.01 | 1.09 | 0.88 | 0.41 | 0.99 | 0.47 | 0.62 | 0.83 | 0.16 | 0 | 0.26 | 0.67 | 1.54 | NA | 0.93 | 4.41 | 2.02 | 0.16 | 0.67 | NA | 1.92 | 0.47 | 0.31
Results from laboratories 17 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 556b (Certified Value 10.8 ppm) | [28 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 1
| 1 | 1 | 8.4 | 6.66 | 12 | 10.5 | 12.4 | 7.6332 | 13 | 11.8 | 10.8 | 12.7 | 12.3 | 13.35 | 0.4 | 10.5232 | 10 | 11.4221 | NA | 15 | 11 | 10.7 | 11 | 12 | 11.1 |  |  |  |  | 1
| 1 | 2 | 9 | 6.84 | 11 | 9.3 | 12.6 | 7.544 | 13.1 | 12.1 | 10.9 | 13.3 | 12.7 | 12.25 | 1 | 10.6607 | 11 | 11.5461 | NA | 13 | 12 | 10.1 | 12 | 12 | 10.6 |  |  |  |  | 19
| 1 | 3 | 8.5 | 6.91 | 11 | 10 | 12.9 | 7.405 | 11 | 12.1 | 11.2 | 12.9 | 12.7 | 13.25 | 1.7 | 10.3064 | 11 | 11.5998 | NA | 14 | 10.6 | 10.6 | 12 | 11 | 10.8 |  |  |  |  | 16
| 1 | 4 | 8.3 | 6.88 | 12 | 10.3 | 12.7 | 7.3824 | 10.6 | 12.2 | 10.8 | 12.9 | 12.8 | 12.33 | 1.8 | 10.7555 | 11 | 11.8982 | NA | 15 | 10.8 | 10.9 | 11 | 11 | 10.8 |  |  |  |  | 10.8076
| 1 | 5 | 8.9 | 6.26 | 12 | 9.7 | 13.2 | 7.685 | 11.6 | 11.9 | 10.6 | 12.9 | 13 | 13.05 | 1.1 | 10.4384 | 11 | 11.1055 | NA | 16 | 10.8 | 10.2 | 11 | 11 | 10.9 |  |  |  |  | 192
| 1 | 6 | 8.7 | 6.74 | 12 | 9.3 | 12.6 | 7.5606 | 11.6 | 11.9 | 11 | 12.9 | 12.7 | 13.16 | 1.3 | 10.496 | 11 | 11.8452 | NA | 13 | 10.5 | 10.4 | 12 | 11 | 11.1
| Mean |  | 8.6333 | 6.715 | 11.6667 | 9.85 | 12.7333 | 7.535 | 11.8167 | 12 | 10.8833 | 12.9333 | 12.7 | 12.8983 | 1.2167 | 10.53 | 10.8333 | 11.5695 |  | 14.3333 | 10.95 | 10.4833 | 11.5 | 11.3333 | 10.8833
| Median |  | 8.6 | 6.79 | 12 | 9.85 | 12.65 | 7.5523 | 11.6 | 12 | 10.85 | 12.9 | 12.7 | 13.105 | 1.2 | 10.5096 | 11 | 11.573 |  | 14.5 | 10.8 | 10.5 | 11.5 | 11 | 10.85
| Std Dev. |  | 0.28048 | 0.24147 | 0.5164 | 0.50498 | 0.28048 | 0.12088 | 1.0284 | 0.15492 | 0.20412 | 0.19664 | 0.22804 | 0.48218 | 0.51153 | 0.15969 | 0.40825 | 0.29073 |  | 1.2111 | 0.54314 | 0.30605 | 0.54772 | 0.5164 | 0.19408
| Rel.Std.Dev. |  | 3.249% | 3.596% | 4.426% | 5.127% | 2.203% | 1.604% | 8.703% | 1.291% | 1.876% | 1.52% | 1.796% | 3.738% | 42.04% | 1.517% | 3.768% | 2.513% |  | 8.449% | 4.96% | 2.919% | 4.763% | 4.556% | 1.783%
| PDM3 |  | -20.12% | -37.87% | 7.948% | -8.861% | 17.82% | -30.28% | 9.336% | 11.03% | 0.7004% | 19.67% | 17.51% | 19.34% | -88.74% | -2.569% | 0.2378% | 7.049% |  | 32.62% | 1.317% | -3.001% | 6.406% | 4.864% | 0.7004%
| Z-Score (Absolute) |  | 1.36 | 2.52 | NA | 0.62 | 1.12 | 2.02 | 0.56 | 0.67 | 0 | 1.24 | 1.1 | 1.22 | 5.84 | 0.21 | NA | 0.41 | NA | NA | 0.04 | 0.24 | NA | NA | 0
Results from laboratories 4, 17, 21, 26 and 27 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 556b (Certified Value 3.5 ppm) | [28 empty] | Certified
Sm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  |  |  |  | 1
| No | No | 02: 04; 03: 05; 04: 14; 05: 16; 06: 18; 07: 26 |  |  |  |  | ppm
|  |  | 02: AR*MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS |  |  |  |  | 2
|  |  | 02: 01g; 03: 0.5g; 04: 0.5g; 05: 0.25g; 06: 0.2g; 07: 01g |  |  |  |  | 2
| 1 | 1 | 02: 3.646; 03: 3.8; 04: 2.72; 05: 1.6437; 06: 3.3251; 07: 3.97 |  |  |  |  | 1
| 1 | 2 | 02: 3.724; 03: 3.5; 04: 2.82; 05: 1.5865; 06: 3.2097; 07: 4.06 |  |  |  |  | 3
| 1 | 3 | 02: 3.632; 03: 3.9; 04: 2.8; 05: 1.6617; 06: 3.1973; 07: 3.96 |  |  |  |  | 16
| 1 | 4 | 02: 3.725; 03: 4.1; 04: 2.87; 05: 1.5857; 06: 3.6372; 07: 3.91 |  |  |  |  | 3.5003
| 1 | 5 | 02: 3.557; 03: 4; 04: 2.89; 05: 1.5634; 06: 3.3061; 07: 3.95 |  |  |  |  | 193
| 1 | 6 | 02: 3.668; 03: 3.4; 04: 2.81; 05: 1.58; 06: 3.1665; 07: 4.15
| Mean |  | 02: 3.6587; 03: 3.7833; 04: 2.8183; 05: 1.6035; 06: 3.307; 07: 4
| Median |  | 02: 3.657; 03: 3.85; 04: 2.815; 05: 1.5861; 06: 3.2579; 07: 3.965
| Std Dev. |  | 02: 0.063188; 03: 0.27869; 04: 0.059805; 05: 0.039439; 06: 0.17355; 07: 0.088544
| Rel.Std.Dev. |  | 02: 1.727%; 03: 7.366%; 04: 2.122%; 05: 2.46%; 06: 5.248%; 07: 2.214%
| PDM3 |  | 02: 4.526%; 03: 8.087%; 04: -19.48%; 05: -54.19%; 06: -5.522%; 07: 14.28%
| Z-Score (Absolute) |  | 02: 0.29; 03: 0.5; 04: 1.1; 05: 3.1; 06: 0.29; 07: 0.85
| Analytical results for Sn in OREAS 556b (Certified Value 2.45 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 5 | < 10 | 2.58 | 2.27 | 2.5 | 2.4056 | 2.4 | 2.3 | 2.4 | 2.5 | 2.3 | < 10 | 1.9 | 2.0837 | 3.1 | 2.4035 | 2.5 | < 10 | 2.4 | 2.3 | 1.8 | 2.7 | 2.9 |  |  |  |  | 1
| 1 | 2 | < 5 | < 10 | 2.6 | 2.11 | 2.4 | 2.4609 | 2.4 | 2.3 | 2.6 | 2.5 | 2.3 | < 10 | 1.97 | 2.0618 | 3.3 | 2.4388 | 2.5 | < 10 | 2.5 | 2.4 | 1.8 | 2.72 | 2.7 |  |  |  |  | 20
| 1 | 3 | < 5 | < 10 | 2.55 | 2.11 | 2.5 | 2.4027 | 2.4 | 2.4 | 2.5 | 2.5 | 2.3 | < 10 | 1.99 | 2.1144 | 3.2 | 2.4842 | 2.4 | < 10 | 2.4 | 2.3 | 1.9 | 2.65 | 2.8 |  |  |  |  | 16
| 1 | 4 | < 5 | < 10 | 2.54 | 2.18 | 2.5 | 2.412 | 2.5 | 2.4 | 2.4 | 2.5 | 2.4 | < 10 | 1.92 | 2.1163 | 3.2 | 2.6438 | 2.3 | < 10 | 2.4 | 2.4 | 2 | 2.61 | 2.8 |  |  |  |  | 2.4509
| 1 | 5 | < 5 | < 10 | 2.53 | 2.11 | 2.4 | 2.4346 | 2.5 | 2.3 | 2.5 | 2.5 | 2.4 | < 10 | 1.99 | 2.1066 | 3 | 2.3831 | 2.3 | < 10 | 2.3 | 2.4 | 1.8 | 2.67 | 2.8 |  |  |  |  | 194
| 1 | 6 | < 5 | < 10 | 2.47 | 2.12 | 2.5 | 2.4577 | 2.5 | 2.3 | 2.6 | 2.5 | 2.4 | < 10 | 1.99 | 2.0569 | 3 | 2.4408 | 2.5 | < 10 | 2.4 | 2.4 | 1.8 | 2.7 | 2.8
| Mean |  |  |  | 2.545 | 2.15 | 2.4667 | 2.4289 | 2.45 | 2.3333 | 2.5 | 2.5 | 2.35 |  | 1.96 | 2.0899 | 3.1333 | 2.4657 | 2.4167 |  | 2.4 | 2.3667 | 1.85 | 2.675 | 2.8
| Median |  |  |  | 2.545 | 2.115 | 2.5 | 2.4233 | 2.45 | 2.3 | 2.5 | 2.5 | 2.35 |  | 1.98 | 2.0951 | 3.15 | 2.4398 | 2.45 |  | 2.4 | 2.4 | 1.8 | 2.685 | 2.8
| Std Dev. |  |  |  | 0.045056 | 0.064807 | 0.05164 | 0.026095 | 0.054772 | 0.05164 | 0.089443 | 0 | 0.054772 |  | 0.04 | 0.026432 | 0.12111 | 0.093894 | 0.098319 |  | 0.063246 | 0.05164 | 0.083666 | 0.040373 | 0.063246
| Rel.Std.Dev. |  |  |  | 1.77% | 3.014% | 2.094% | 1.074% | 2.236% | 2.213% | 3.578% | 0% | 2.331% |  | 2.041% | 1.265% | 3.865% | 3.808% | 4.068% |  | 2.635% | 2.182% | 4.522% | 1.509% | 2.259%
| PDM3 |  |  |  | 3.837% | -12.28% | 0.6413% | -0.8982% | -0.03873% | -4.799% | 2.001% | 2.001% | -4.119% |  | -20.03% | -14.73% | 27.84% | 0.6013% | -1.399% |  | -2.079% | -3.439% | -24.52% | 9.141% | 14.24%
| Z-Score (Absolute) |  | 0.2 | 14.77 | 0.46 | 1.84 | 0.01 | 0.21 | 0.09 | 0.77 | 0.2 | 0.2 | 0.67 | 14.77 | 2.95 | 2.19 | 3.89 | 0 | 0.29 | 14.77 | 0.38 | 0.58 | 3.59 | 1.22 | 1.95
| Analytical results for Sr in OREAS 556b (Certified Value 15.4 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 1
| 1 | 1 | 29 | 12.5 | 31.1 | 15.53 | 18.8 | 10.9 | 13.7 | 14.9 | 13.8 | 13.1 | 14.4 | 14.86 | 10 | 15.5589 | 21.7 | 16.6761 | 22.1 | 11.3 | 14 | 15.9 | 22 | 20.58 | 14 | 13.7 |  |  |  | 1
| 1 | 2 | 26.3 | 13.9 | 30.9 | 14.55 | 15.9 | 10.9 | 14.2 | 15.8 | 15.1 | 12.8 | 14.5 | 14.9 | 11 | 14.3098 | 21.3 | 16.7011 | 21.6 | 11.4 | 14 | 15.8 | 26.6 | 20.82 | 13 | 12.8 |  |  |  | 32
| 1 | 3 | 29.6 | 13.1 | 29.9 | 15.32 | 17.5 | 11.3 | 13.6 | 15.4 | 14.3 | 13.4 | 14.5 | 14.88 | 10 | 13.9175 | 21.5 | 17.2756 | 21.1 | 11.2 | 14 | 16.6 | 19.3 | 20.36 | 14 | 13.8 |  |  |  | 16
| 1 | 4 | 21.7 | 12.3 | 29.6 | 15.76 | 18.1 | 10.7 | 13.5 | 15.5 | 13.8 | 13.2 | 14.7 | 14.95 | 10 | 13.3241 | 22.3 | 16.1408 | 19.5 | 11.1 | 14 | 16.4 | 29.5 | 20.21 | 14 | 13.4 |  |  |  | 15.4311
| 1 | 5 | 28.4 | 13.7 | 29.8 | 15.56 | 17.1 | 10.5 | 13.6 | 15.4 | 14.4 | 13.2 | 15.1 | 14.96 | 11 | 14.4711 | 21.7 | 16.7143 | 20.3 | 11.1 | 14 | 16.5 | 24.7 | 20.19 | 14 | 13.4 |  |  |  | 195
| 1 | 6 | 25.6 | 14.8 | 30.4 | 15.22 | 16 | 11 | 13.7 | 15.5 | 14.2 | 13 | 14.4 | 14.87 | 10 | 16.039 | 21.4 | 16.6471 | 20.3 | 11.3 | 14 | 16.5 | 26.2 | 20.64 | 14 | 13.1
| Mean |  | 26.7667 | 13.3833 | 30.2833 | 15.3233 | 17.2333 | 10.8833 | 13.7167 | 15.4167 | 14.2667 | 13.1167 | 14.6 | 14.9033 | 10.3333 | 14.6034 | 21.65 | 16.6925 | 20.8167 | 11.2333 | 14 | 16.2833 | 24.7167 | 20.4667 | 13.8333 | 13.3667
| Median |  | 27.35 | 13.4 | 30.15 | 15.425 | 17.3 | 10.9 | 13.65 | 15.45 | 14.25 | 13.15 | 14.5 | 14.89 | 10 | 14.3905 | 21.6 | 16.6886 | 20.7 | 11.25 | 14 | 16.45 | 25.45 | 20.47 | 14 | 13.4
| Std Dev. |  | 2.9303 | 0.93897 | 0.61779 | 0.42382 | 1.1483 | 0.27142 | 0.24833 | 0.29269 | 0.48028 | 0.20412 | 0.26833 | 0.042269 | 0.5164 | 1.0183 | 0.35637 | 0.3598 | 0.96003 | 0.12111 | 0 | 0.34303 | 3.6152 | 0.25359 | 0.40825 | 0.37238
| Rel.Std.Dev. |  | 10.95% | 7.016% | 2.04% | 2.766% | 6.663% | 2.494% | 1.81% | 1.899% | 3.366% | 1.556% | 1.838% | 0.2836% | 4.997% | 6.973% | 1.646% | 2.155% | 4.612% | 1.078% | 0% | 2.107% | 14.63% | 1.239% | 2.951% | 2.786%
| PDM3 |  | 73.46% | -13.27% | 96.25% | -0.6982% | 11.68% | -29.47% | -11.11% | -0.09334% | -7.546% | -15% | -5.386% | -3.42% | -33.04% | -5.364% | 40.3% | 8.175% | 34.9% | -27.2% | -9.274% | 5.523% | 60.17% | 32.63% | -10.35% | -13.38%
| Z-Score (Absolute) |  | 3.98 | 0.67 | 5.2 | 0 | 0.66 | 1.54 | 0.56 | 0.03 | 0.37 | 0.77 | 0.25 | 0.15 | NA | 0.25 | 2.2 | 0.48 | 1.91 | 1.42 | NA | 0.33 | 3.26 | 1.79 | NA | 0.68
Results from laboratories 14, 22 and 27 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 556b (Certified Value < 0.01 ppm) | [28 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 22 | 23 | 25 | 26 | 28 | [10 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [10 empty] | 3
| 1 | 1 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.08 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 1
| 1 | 2 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.09 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 5
| 1 | 3 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 16
| 1 | 4 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | < 0.01
| 1 | 5 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.07 | < 0.01 | < 0.01 | < 0.01 | [10 empty] | 196
| 1 | 6 | < 5 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | 0.06 | < 0.01 | < 0.01 | < 0.01
| Mean |  | 15: 0.073333
| Median |  | 15: 0.07
| Std Dev. |  | 15: 0.010328
| Rel.Std.Dev. |  | 15: 14.08%
PDM3
Z-Score (Absolute) | bar: 01 to 20
| Analytical results for Tb in OREAS 556b (Certified Value 0.2 ppm) | [28 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 04; 02: 05; 03: 14; 04: 16; 05: 18; 06: 20; 08: 26 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 08: AR*MS |  |  |  | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 0.25g; 05: 0.2g; 06: 0.2g; 08: 01g |  |  |  | 2
| 1 | 1 | 01: 0.212; 02: 0.2; 03: 0.16; 04: 0.165; 05: 0.23732; 06: 0.3; 08: 0.25 |  |  |  | 1
| 1 | 2 | 01: 0.211; 02: 0.2; 03: 0.16; 04: 0.15241; 05: 0.2226; 06: 0.3; 08: 0.24 |  |  |  | 6
| 1 | 3 | 01: 0.211; 02: 0.2; 03: 0.16; 04: 0.16904; 05: 0.2314; 06: 0.3; 08: 0.25 |  |  |  | 16
| 1 | 4 | 01: 0.227; 02: 0.2; 03: 0.16; 04: 0.15455; 05: 0.27086; 06: 0.3; 08: 0.24 |  |  |  | 0.20255
| 1 | 5 | 01: 0.211; 02: 0.2; 03: 0.16; 04: 0.15757; 05: 0.23446; 06: 0.3; 08: 0.25 |  |  |  | 197
| 1 | 6 | 01: 0.217; 02: 0.2; 03: 0.16; 04: 0.15827; 05: 0.233; 06: 0.3; 08: 0.25
| Mean |  | 01: 0.21483; 02: 0.2; 03: 0.16; 04: 0.15947; 05: 0.23827; 06: 0.3; 08: 0.24667
| Median |  | 01: 0.2115; 02: 0.2; 03: 0.16; 04: 0.15792; 05: 0.23373; 06: 0.3; 08: 0.25
| Std Dev. |  | 01: 0.0064005; 02: 3.0405e-17; 03: 0; 04: 0.0063438; 05: 0.01672; 06: 0; 08: 0.005164
| Rel.Std.Dev. |  | 01: 2.979%; 02: 1.52e-14%; 03: 0%; 04: 3.978%; 05: 7.017%; 06: 0%; 08: 2.094%
| PDM3 |  | 01: 6.067%; 02: -1.257%; 03: -21.01%; 04: -21.27%; 05: 17.64%; 06: 48.11%; 08: 21.78%
| Z-Score (Absolute) |  | 01: 0; 02: NA; 03: 1.16; 04: 1.17; 05: 0.5; 06: NA; 08: 0.67
Results from laboratories 5 and 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 556b (Certified Value 0.075 ppm) | [28 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 22 | 25 | 26 | 28 | [5 empty] | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.07 | < 5 | < 0.1 | 0.08 | 0.08 | < 0.04 | 0.08 | 0.09 | 0.07 | 0.1 | 0.09 | < 10 | 0.06 | < 0.05 | 0.07 | < 5 | NA | < 10 | < 0.1 | 0.05 | 0.08 | 0.08 | [5 empty] | 1
| 1 | 2 | 0.07 | < 5 | < 0.1 | 0.05 | 0.08 | 0.10724 | 0.1 | 0.1 | 0.08 | 0.09 | 0.07 | < 10 | 0.06 | < 0.05 | 0.07 | < 5 | NA | < 10 | < 0.1 | 0.05 | 0.08 | 0.07 | [5 empty] | 7
| 1 | 3 | 0.08 | < 5 | < 0.1 | 0.08 | 0.09 | 0.04849 | 0.07 | 0.1 | 0.07 | 0.09 | 0.07 | < 10 | 0.06 | < 0.05 | 0.07 | < 5 | NA | < 10 | < 0.1 | 0.05 | 0.06 | 0.1 | [5 empty] | 16
| 1 | 4 | 0.06 | < 5 | < 0.1 | 0.07 | 0.07 | 0.05283 | 0.07 | 0.09 | 0.08 | 0.08 | 0.06 | < 10 | 0.07 | < 0.05 | 0.07 | < 5 | NA | < 10 | < 0.1 | 0.05 | 0.09 | 0.08 | [5 empty] | 0.075122
| 1 | 5 | 0.13 | < 5 | < 0.1 | 0.09 | 0.1 | 0.07259 | 0.07 | 0.08 | 0.08 | 0.09 | 0.06 | < 10 | 0.07 | < 0.05 | 0.06 | < 5 | NA | < 10 | < 0.1 | 0.04 | 0.08 | 0.07 | [5 empty] | 198
| 1 | 6 | 0.1 | < 5 | < 0.1 | 0.07 | 0.06 | 0.05476 | 0.07 | 0.09 | 0.07 | 0.07 | 0.07 | < 10 | 0.07 | < 0.05 | 0.06 | < 5 | NA | < 10 | 0.1 | 0.06 | 0.11 | 0.08
| Mean |  | 0.085 |  |  | 0.073333 | 0.08 | 0.067182 | 0.076667 | 0.091667 | 0.075 | 0.086667 | 0.07 |  | 0.065 |  | 0.066667 |  |  |  | 0.1 | 0.05 | 0.083333 | 0.08
| Median |  | 0.075 |  |  | 0.075 | 0.08 | 0.05476 | 0.07 | 0.09 | 0.075 | 0.09 | 0.07 |  | 0.065 |  | 0.07 |  |  |  | 0.1 | 0.05 | 0.08 | 0.08
| Std Dev. |  | 0.025884 |  |  | 0.013663 | 0.014142 | 0.024205 | 0.012111 | 0.0075277 | 0.0054772 | 0.010328 | 0.010954 |  | 0.0054772 |  | 0.005164 |  |  |  |  | 0.0063246 | 0.01633 | 0.010954
| Rel.Std.Dev. |  | 30.45% |  |  | 18.63% | 17.68% | 36.03% | 15.8% | 8.212% | 7.303% | 11.92% | 15.65% |  | 8.427% |  | 7.746% |  |  |  |  | 12.65% | 19.6% | 13.69%
| PDM3 |  | 13.15% |  |  | -2.381% | 6.493% | -10.57% | 2.056% | 22.02% | -0.1628% | 15.37% | -6.819% |  | -13.47% |  | -11.26% |  |  |  | 33.12% | -33.44% | 10.93% | 6.493%
| Z-Score (Absolute) |  | 0.48 | 140.06 | 1.54 | 0.19 | 0.19 | 1 | 0 | 0.87 | 0.1 | 0.58 | 0.39 | 284.56 | 0.67 | 2.99 | 0.58 | 140.06 | NA | 284.56 | 1.06 | 1.54 | 0.39 | 0.19
| Analytical results for Th in OREAS 556b (Certified Value 6.16 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | [5 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | [5 empty] | 2
| 1 | 1 | 12.8 | 4.8 | 6.44 | 7.6 | 5.1 | 5.5 | 6.9 | 5.8 | 5.4 | 5.6 | 5.28 | 6.5973 | 4.5495 | 8.13 | 6.1137 | < 20 | 5.9 | 7.02 | 10.2 | 7.82 | 6.1 | 6.6 | [5 empty] | 1
| 1 | 2 | 12.3 | 5.5 | 6.27 | 6.6 | 5.2 | 5.6 | 6.8 | 5.9 | 5.3 | 5.8 | 5.39 | 6.331 | 4.3753 | 8.02 | 5.8423 | < 20 | 6.3 | 6.86 | 10.5 | 7.95 | 6.1 | 6.5 | [5 empty] | 8
| 1 | 3 | 12.9 | 5 | 6.15 | 7.3 | 5.2 | 5.3 | 7.2 | 5.9 | 5.5 | 5.7 | 5.22 | 6.1502 | 4.5937 | 8.14 | 5.9242 | < 20 | 6.1 | 6.83 | 9.17 | 7.89 | 6.2 | 6.3 | [5 empty] | 16
| 1 | 4 | 10.9 | 5 | 6.5 | 7.5 | 5.1 | 5.3 | 7.1 | 5.6 | 5.4 | 5.9 | 5.52 | 6.1951 | 4.458 | 8.06 | 6.7535 | < 20 | 6.2 | 7.12 | 11.6 | 7.61 | 6.1 | 6.4 | [5 empty] | 6.1639
| 1 | 5 | 12.8 | 5.8 | 6.36 | 7.2 | 5.1 | 5.4 | 6.7 | 5.7 | 5.4 | 6 | 5.51 | 6.2621 | 4.4672 | 7.96 | 6.2681 | < 20 | 6.2 | 6.89 | 10.4 | 7.83 | 6.2 | 6.2 | [5 empty] | 199
| 1 | 6 | 12.1 | 4.3 | 6.36 | 7 | 5.2 | 5.4 | 7 | 6.2 | 5.3 | 5.7 | 5.4 | 6.6131 | 4.3298 | 8 | 6.1308 | < 20 | 6.3 | 6.98 | 10.5 | 7.92 | 6.3 | 6.5
| Mean |  | 12.3 | 5.0667 | 6.3467 | 7.2 | 5.15 | 5.4167 | 6.95 | 5.85 | 5.3833 | 5.7833 | 5.3867 | 6.3581 | 4.4622 | 8.0517 | 6.1721 |  | 6.1667 | 6.95 | 10.395 | 7.8367 | 6.1667 | 6.4167
| Median |  | 12.55 | 5 | 6.36 | 7.25 | 5.15 | 5.4 | 6.95 | 5.85 | 5.4 | 5.75 | 5.395 | 6.2966 | 4.4626 | 8.04 | 6.1223 |  | 6.2 | 6.935 | 10.45 | 7.86 | 6.15 | 6.45
| Std Dev. |  | 0.75631 | 0.52789 | 0.1242 | 0.36332 | 0.054772 | 0.1169 | 0.18708 | 0.20736 | 0.075277 | 0.1472 | 0.12028 | 0.20098 | 0.10004 | 0.072226 | 0.32329 |  | 0.15055 | 0.11027 | 0.77624 | 0.12193 | 0.08165 | 0.1472
| Rel.Std.Dev. |  | 6.149% | 10.42% | 1.957% | 5.046% | 1.064% | 2.158% | 2.692% | 3.545% | 1.398% | 2.545% | 2.233% | 3.161% | 2.242% | 0.897% | 5.238% |  | 2.441% | 1.587% | 7.467% | 1.556% | 1.324% | 2.294%
| PDM3 |  | 99.55% | -17.8% | 2.965% | 16.81% | -16.45% | -12.12% | 12.75% | -5.093% | -12.66% | -6.174% | -12.61% | 3.151% | -27.61% | 30.63% | 0.133% |  | 0.04485% | 12.75% | 68.64% | 27.14% | 0.04485% | 4.101%
| Z-Score (Absolute) |  | 4.75 | 0.94 | 0.07 | 0.74 | 0.87 | 0.66 | 0.54 | 0.32 | 0.69 | 0.37 | 0.69 | 0.08 | 1.41 | 1.41 | 0.07 | 2.94 | 0.07 | 0.54 | 3.25 | 1.24 | 0.07 | 0.12
| Analytical results for Ti in OREAS 556b (Indicative Value 0.003 wt.%) | [28 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 27 | 28 | [5 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | [5 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | [5 empty] | 3
| 1 | 1 | < 0.01 | 0.007 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0012341 | 0.0016 | 0.0013102 | < 0.01 | < 0.01 | < 0.005 | 0.0011 | 0.0016 | 0.001 | < 0.005 | [5 empty] | 1
| 1 | 2 | < 0.01 | 0.007 | < 0.01 | < 0.0005 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.00111 | 0.0016 | 0.0013005 | < 0.01 | < 0.01 | < 0.005 | 0.0023 | 0.0014 | 0.001 | < 0.005 | [5 empty] | 25
| 1 | 3 | < 0.01 | 0.007 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.00094811 | 0.0016 | 0.0013154 | < 0.01 | < 0.01 | < 0.005 | 0.0025 | 0.0014 | 0.001 | < 0.005 | [5 empty] | 16
| 1 | 4 | < 0.01 | 0.007 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.00093664 | 0.0016 | 0.001395 | < 0.01 | < 0.01 | < 0.005 | 0.0032 | 0.0013 | 0.001 | < 0.005 | [5 empty] | 0.0026059
| 1 | 5 | < 0.01 | 0.007 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0010059 | 0.0016 | 0.001329 | < 0.01 | < 0.01 | < 0.005 | 0.0019 | 0.0013 | 0.001 | < 0.005 | [5 empty] | 31
| 1 | 6 | < 0.01 | 0.001 | < 0.01 | 0.0006 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.0012 | 0.0011648 | 0.0015 | 0.0012831 | < 0.01 | < 0.01 | < 0.005 | 0.0021 | 0.0013 | 0.001 | < 0.005
| Mean |  | 02: 0.006; 04: 0.0006; 15: 0.0010666; 17: 0.0015833; 18: 0.0013222; 25: 0.0021833; 26: 0.0013833; 27: 0.001
| Median |  | 02: 0.007; 04: 0.0006; 15: 0.001058; 17: 0.0016; 18: 0.0013128; 25: 0.0022; 26: 0.00135; 27: 0.001
| Std Dev. |  | 02: 0.0024495; 04: 0; 15: 0.00012182; 17: 4.0825e-05; 18: 3.8827e-05; 25: 0.00069402; 26: 0.0001169; 27: 0
| Rel.Std.Dev. |  | 02: 40.82%; 04: 0%; 15: 11.42%; 17: 2.578%; 18: 2.936%; 25: 31.79%; 26: 8.451%; 27: 0%
| PDM3 |  | 02: 130.2%; 04: -76.98%; 15: -59.07%; 17: -39.24%; 18: -49.26%; 25: -16.22%; 26: -46.92%; 27: -61.63%
| Z-Score (Absolute) |  | 1.47 | 2.06 | 1.47 | 1.15 | 1.47 | 0 | 0 | 0 | 0 | 0 | 0 | 1.12 | 0.84 | 0.54 | 0.69 | 1.47 | 1.47 | 0 | 0.19 | 0.66 | 0.88 | 0
| Analytical results for Tl in OREAS 556b (Indicative Value 0.37 ppm) | [28 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 2
| 1 | 1 | 57.1 | 0.47 | < 10 | 0.76 | 0.07 | 0.53 | 0.32492 | 0.07 | 0.06 | 0.13 | 0.12 | 0.11 | 0.42848 | 0.52 | 0.26565 | 0.28 | 0.69 | 2 | < 0.05 | 0.45 | 0.85 | 0.77 | < 0.5 | 0.4 |  |  |  | 1
| 1 | 2 | 56.2 | 0.51 | < 10 | 0.73 | 0.08 | 0.54 | 0.34203 | 0.09 | 0.05 | 0.13 | 0.14 | 0.09 | 0.44764 | 0.52 | 0.23262 | 0.287 | 0.63 | 3 | 0.05 | 0.44 | 0.8 | 0.76 | < 0.5 | 0.4 |  |  |  | 26
| 1 | 3 | 63.1 | 0.52 | < 10 | 0.76 | 0.04 | 0.55 | 0.30689 | 0.1 | 0.06 | 0.11 | 0.14 | 0.1 | 0.52305 | 0.54 | 0.2587 | 0.457 | 0.64 | 2 | 0.05 | 0.42 | 0.68 | 0.71 | < 0.5 | 0.44 |  |  |  | 16
| 1 | 4 | 52.6 | 0.52 | < 10 | 0.79 | 0.05 | 0.56 | 0.32335 | 0.1 | 0.06 | 0.12 | 0.12 | 0.11 | 0.44269 | 0.52 | 0.28445 | 0.324 | 0.64 | 2 | 0.06 | 0.42 | 0.87 | 0.7 | < 0.5 | 0.43 |  |  |  | 0.36686
| 1 | 5 | 64.9 | 0.53 | < 10 | 0.78 | 0.04 | 0.56 | 0.33381 | 0.09 | 0.05 | 0.12 | 0.14 | 0.1 | 0.54206 | 0.52 | 0.28402 | 0.239 | 0.58 | 3 | 0.06 | 0.42 | 0.82 | 0.74 | < 0.5 | 0.42 |  |  |  | 32
| 1 | 6 | 60.8 | 0.53 | < 10 | 0.77 | 0.05 | 0.55 | 0.2947 | 0.09 | 0.06 | 0.11 | 0.14 | 0.1 | 0.40506 | 0.54 | 0.2673 | 0.276 | 0.66 | 3 | 0.06 | 0.41 | 0.8 | 0.77 | < 0.5 | 0.42
| Mean |  | 59.1167 | 0.51333 |  | 0.765 | 0.055 | 0.54833 | 0.32095 | 0.09 | 0.056667 | 0.12 | 0.13333 | 0.10167 | 0.46483 | 0.52667 | 0.26546 | 0.3105 | 0.64 | 2.5 | 0.056 | 0.42667 | 0.80333 | 0.74167 |  | 0.41833
| Median |  | 58.95 | 0.52 |  | 0.765 | 0.05 | 0.55 | 0.32414 | 0.09 | 0.06 | 0.12 | 0.14 | 0.1 | 0.44516 | 0.52 | 0.26648 | 0.2835 | 0.64 | 2.5 | 0.06 | 0.42 | 0.81 | 0.75 |  | 0.42
| Std Dev. |  | 4.6301 | 0.022509 |  | 0.020736 | 0.016432 | 0.01169 | 0.017424 | 0.010954 | 0.005164 | 0.0089443 | 0.010328 | 0.0075277 | 0.054833 | 0.010328 | 0.019144 | 0.07672 | 0.036332 | 0.54772 | 0.0054772 | 0.015055 | 0.066533 | 0.030605 |  | 0.016021
| Rel.Std.Dev. |  | 7.832% | 4.385% |  | 2.711% | 29.88% | 2.132% | 5.429% | 12.17% | 9.113% | 7.454% | 7.746% | 7.404% | 11.8% | 1.961% | 7.212% | 24.71% | 5.677% | 21.91% | 9.781% | 3.529% | 8.282% | 4.127% |  | 3.83%
| PDM3 |  | 16014% | 39.93% |  | 108.5% | -85.01% | 49.47% | -12.51% | -75.47% | -84.55% | -67.29% | -63.66% | -72.29% | 26.71% | 43.56% | -27.64% | -15.36% | 74.46% | 581.5% | -84.74% | 16.3% | 119% | 102.2% |  | 14.03%
| Z-Score (Absolute) |  | 133.78 | 0.21 | 10.43 | 0.78 | 0.84 | 0.29 | 0.23 | 0.76 | 0.83 | 0.69 | 0.66 | 0.73 | 0.1 | 0.24 | 0.36 | 0.26 | 0.5 | 4.74 | 0.85 | 0.01 | 0.87 | 0.73 | 0.39 | 0.01
| Analytical results for Tm in OREAS 556b (Indicative Value 0.056 ppm) | [28 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab |  |  |  | 1
| No | No | 01: 04; 02: 05; 03: 18; 04: 26 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS |  |  |  | 3
|  |  | 01: 01g; 02: 0.5g; 03: 0.2g; 04: 01g |  |  |  | 3
| 1 | 1 | 01: 0.057; 02: < 0.1; 03: 0.050011; 04: 0.07 |  |  |  | 1
| 1 | 2 | 01: 0.053; 02: < 0.1; 03: < 0.05; 04: 0.07 |  |  |  | 10
| 1 | 3 | 01: 0.058; 02: < 0.1; 03: < 0.05; 04: 0.07 |  |  |  | 16
| 1 | 4 | 01: 0.06; 02: < 0.1; 03: 0.053357; 04: 0.06 |  |  |  | 0.056284
| 1 | 5 | 01: 0.057; 02: < 0.1; 03: 0.055396; 04: 0.07 |  |  |  | 30
| 1 | 6 | 01: 0.059; 02: < 0.1; 03: 0.050348; 04: 0.07
| Mean |  | 01: 0.057333; 03: 0.052278; 04: 0.068333
| Median |  | 01: 0.0575; 03: 0.051852; 04: 0.07
| Std Dev. |  | 01: 0.0024221; 03: 0.0025658; 04: 0.0040825
| Rel.Std.Dev. |  | 01: 4.225%; 03: 4.908%; 04: 5.974%
| PDM3 |  | 01: 1.864%; 03: -7.117%; 04: 21.41%
| Z-Score (Absolute) |  | 01: 0.35; 02: 0.35; 03: 1; 04: 1.4
| Analytical results for U in OREAS 556b (Certified Value 3.57 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | 3.07 | 3.66 | 3.7 | 3.32 | 3.49 | 3.77 | 3.24 | 3.48 | 3.37 | 3.21 | 3.7664 | 2.3346 | 3.97 | 3.3908 | 3.37 | 3.67 | 10 | 3.3 | 3.84 | 4.3 | 3.93 | 3.6 | 3.91 |  |  |  |  | 1
| 1 | 2 | 3.44 | 3.57 | 3.3 | 3.33 | 3.51 | 3.72 | 3.4 | 3.41 | 3.45 | 3.34 | 3.6605 | 2.2765 | 3.96 | 3.1954 | 3.34 | 3.65 | 11 | 3.5 | 3.65 | 4.19 | 3.99 | 3.6 | 3.79 |  |  |  |  | 11
| 1 | 3 | 3.19 | 3.55 | 3.5 | 3.47 | 3.12 | 3.94 | 3.33 | 3.48 | 3.44 | 3.28 | 3.7212 | 2.4351 | 3.89 | 3.1996 | 3.38 | 3.59 | 11 | 3.4 | 3.67 | 4.04 | 3.97 | 3.6 | 3.85 |  |  |  |  | 16
| 1 | 4 | 3.22 | 3.72 | 3.8 | 3.34 | 3.11 | 3.75 | 3.34 | 3.48 | 3.48 | 3.35 | 3.692 | 2.5302 | 3.91 | 3.5096 | 3.17 | 3.56 | 11 | 3.5 | 3.82 | 4.64 | 3.87 | 3.5 | 3.81 |  |  |  |  | 3.5671
| 1 | 5 | 3.6 | 3.58 | 3.5 | 3.34 | 3.19 | 3.66 | 3.18 | 3.36 | 3.53 | 3.38 | 3.6074 | 2.2942 | 3.88 | 3.2814 | 3.19 | 3.5 | 10 | 3.5 | 3.64 | 4.25 | 3.97 | 3.6 | 3.76 |  |  |  |  | 200
| 1 | 6 | 2.32 | 3.6 | 3.5 | 3.33 | 3.26 | 3.85 | 3.42 | 3.31 | 3.34 | 3.3 | 3.7969 | 2.272 | 3.92 | 3.3866 | 3.29 | 3.63 | 11 | 3.5 | 3.64 | 4.15 | 3.97 | 3.6 | 3.92
| Mean |  | 3.14 | 3.6133 | 3.55 | 3.355 | 3.28 | 3.7817 | 3.3183 | 3.42 | 3.435 | 3.31 | 3.7074 | 2.3571 | 3.9217 | 3.3272 | 3.29 | 3.6 | 10.6667 | 3.45 | 3.71 | 4.2617 | 3.95 | 3.5833 | 3.84
| Median |  | 3.205 | 3.59 | 3.5 | 3.335 | 3.225 | 3.76 | 3.335 | 3.445 | 3.445 | 3.32 | 3.7066 | 2.3144 | 3.915 | 3.334 | 3.315 | 3.61 | 11 | 3.5 | 3.66 | 4.22 | 3.97 | 3.6 | 3.83
| Std Dev. |  | 0.44448 | 0.064395 | 0.17607 | 0.056833 | 0.17889 | 0.099482 | 0.09261 | 0.072938 | 0.070071 | 0.060663 | 0.06946 | 0.10417 | 0.03656 | 0.12377 | 0.090995 | 0.063246 | 0.5164 | 0.083666 | 0.093808 | 0.20566 | 0.043818 | 0.040825 | 0.065115
| Rel.Std.Dev. |  | 14.16% | 1.782% | 4.96% | 1.694% | 5.454% | 2.631% | 2.791% | 2.133% | 2.04% | 1.833% | 1.874% | 4.42% | 0.9323% | 3.72% | 2.766% | 1.757% | 4.841% | 2.425% | 2.529% | 4.826% | 1.109% | 1.139% | 1.696%
| PDM3 |  | -11.97% | 1.295% | -0.4806% | -5.947% | -8.05% | 6.014% | -6.975% | -4.125% | -3.704% | -7.209% | 3.933% | -33.92% | 9.939% | -6.725% | -7.769% | 0.9211% | 199% | -3.284% | 4.005% | 19.47% | 10.73% | 0.4539% | 7.649%
| Z-Score (Absolute) |  | 1.22 | 0.38 | 0.17 | 0.49 | 0.75 | 0.96 | 0.62 | 0.27 | 0.22 | 0.64 | 0.7 | 3.88 | 1.43 | 0.59 | 0.71 | 0.34 | NA | 0.17 | 0.71 | 2.58 | 1.53 | 0.28 | 1.15
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 556b (Certified Value 60 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  | 0
| 1 | 1 | 62.8 | 55 | 77.2 | 67 | 79 | 49.2 | 57 | 64 | 54 | 56 | 57 | 59.91 | 46.4 | 70.6938 | 100 | 56.3919 | 120 | 51 | 56 | 63 | 130 | 95 | 60 | 63 |  |  |  | 1
| 1 | 2 | 62.9 | 60 | 76.9 | 69 | 67 | 49.7 | 58 | 66 | 56 | 54 | 58 | 59.76 | 47.7 | 66.4637 | 96 | 56.609 | 117 | 51 | 57 | 63 | 142 | 98 | 59 | 61 |  |  |  | 38
| 1 | 3 | 67.4 | 57 | 76.5 | 67 | 73 | 51 | 53 | 65 | 55 | 57 | 58 | 59.82 | 47.8 | 66.0637 | 100 | 57.8102 | 114 | 52 | 57 | 66 | 125 | 93 | 60 | 59 |  |  |  | 16
| 1 | 4 | 56.6 | 54 | 76.2 | 69 | 76 | 48.8 | 54 | 66 | 56 | 56 | 57 | 59.87 | 49.3 | 65.3209 | 101 | 61.065 | 108 | 50 | 57 | 67 | 158 | 91 | 62 | 59 |  |  |  | 60.3315
| 1 | 5 | 71.7 | 60 | 76.7 | 66 | 73 | 48.2 | 54 | 66 | 56 | 56 | 59 | 59.83 | 47.8 | 66.4314 | 97 | 59.6431 | 108 | 51 | 55 | 66 | 142 | 95 | 61 | 58 |  |  |  | 201
| 1 | 6 | 66.4 | 66 | 76.4 | 68 | 71 | 49.6 | 54 | 66 | 57 | 54 | 57 | 59.77 | 47.3 | 71.39 | 97 | 59.4376 | 108 | 51 | 55 | 69 | 143 | 96 | 61 | 60
| Mean |  | 64.6333 | 58.6667 | 76.65 | 67.6667 | 73.1667 | 49.4167 | 55 | 65.5 | 55.6667 | 55.5 | 57.6667 | 59.8267 | 47.7167 | 67.7272 | 98.5 | 58.4928 | 112.5 | 51 | 56.1667 | 65.6667 | 140 | 94.6667 | 60.5 | 60
| Median |  | 64.65 | 58.5 | 76.6 | 67.5 | 73 | 49.4 | 54 | 66 | 56 | 56 | 57.5 | 59.825 | 47.75 | 66.4475 | 98.5 | 58.6239 | 111 | 51 | 56.5 | 66 | 142 | 95 | 60.5 | 59.5
| Std Dev. |  | 5.1306 | 4.3665 | 0.36194 | 1.2111 | 4.1191 | 0.95167 | 2 | 0.83666 | 1.0328 | 1.2247 | 0.8165 | 0.057504 | 0.9411 | 2.6096 | 2.0736 | 1.858 | 5.282 | 0.63246 | 0.98319 | 2.3381 | 11.5412 | 2.4221 | 1.0488 | 1.7889
| Rel.Std.Dev. |  | 7.938% | 7.443% | 0.4722% | 1.79% | 5.63% | 1.926% | 3.636% | 1.277% | 1.855% | 2.207% | 1.416% | 0.09612% | 1.972% | 3.853% | 2.105% | 3.176% | 4.695% | 1.24% | 1.75% | 3.561% | 8.244% | 2.559% | 1.734% | 2.981%
| PDM3 |  | 7.13% | -2.759% | 27.05% | 12.16% | 21.27% | -18.09% | -8.837% | 8.567% | -7.732% | -8.008% | -4.417% | -0.8368% | -20.91% | 12.26% | 63.26% | -3.048% | 86.47% | -15.47% | -6.903% | 8.843% | 132.1% | 56.91% | 0.2793% | -0.5495%
| Z-Score (Absolute) |  | 0.55 | 0.2 | 2.07 | 0.94 | 1.63 | 1.37 | 0.66 | 0.66 | 0.58 | 0.6 | 0.33 | 0.05 | 1.58 | 0.95 | 4.84 | 0.22 | 6.61 | 1.17 | 0.52 | 0.68 | 10.08 | 4.35 | 0.03 | 0.03
| Analytical results for W in OREAS 556b (Certified Value 0.98 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 5 | 0.8 | < 5 | 0.91 | 1 | 1.08 | 0.95 | 0.9 | 1.09 | 0.98 | 0.91 | 1.1 | 0.97245 | 0.9 | < 1 | 1.1 | < 10 | 0.9 | 1.23 | 0.93 | 0.93 | 1.1 | 1.05 |  |  |  |  | 1
| 1 | 2 | < 5 | 0.79 | < 5 | 0.8 | 0.8 | 1.1 | 0.97 | 0.89 | 1.16 | 1.02 | 0.91 | 1.1 | 0.95786 | 0.9 | < 1 | 1 | < 10 | 0.9 | 1.3 | 0.88 | 0.92 | 1 | 1.02 |  |  |  |  | 12
| 1 | 3 | < 5 | 0.79 | < 5 | 0.87 | 0.8 | 1.13 | 1.04 | 0.96 | 1.14 | 0.96 | 0.93 | 1 | 0.90818 | 0.8 | < 1 | 1.1 | < 10 | 0.9 | 1.17 | 0.93 | 0.9 | 1 | 1.02 |  |  |  |  | 16
| 1 | 4 | < 5 | 0.83 | < 5 | 0.94 | 0.8 | 1.13 | 1.05 | 0.97 | 1.16 | 0.99 | 0.94 | 1.1 | 0.8776 | 0.8 | 1.1448 | 1.3 | < 10 | 1 | 1.17 | 0.99 | 0.92 | 1 | 1.02 |  |  |  |  | 0.98283
| 1 | 5 | < 5 | 0.88 | < 5 | 0.91 | 0.9 | 1.18 | 1.06 | 0.89 | 1.06 | 0.98 | 0.92 | 1.1 | 0.91347 | 0.9 | < 1 | 1 | < 10 | 0.9 | 1.15 | 0.9 | 0.87 | 1 | 1.03 |  |  |  |  | 202
| 1 | 6 | < 5 | 1.23 | < 5 | 0.9 | 0.8 | 1.12 | 1.06 | 0.95 | 1.12 | 1 | 0.91 | 1.1 | 0.97813 | 0.9 | < 1 | 1.2 | < 10 | 0.9 | 1.08 | 0.91 | 0.93 | 1 | 1.08
| Mean |  |  | 0.88667 |  | 0.88833 | 0.85 | 1.1233 | 1.0217 | 0.92667 | 1.1217 | 0.98833 | 0.92 | 1.0833 | 0.93461 | 0.86667 | 1.1448 | 1.1167 |  | 0.91667 | 1.1833 | 0.92333 | 0.91167 | 1.0167 | 1.0367
| Median |  |  | 0.815 |  | 0.905 | 0.8 | 1.125 | 1.045 | 0.925 | 1.13 | 0.985 | 0.915 | 1.1 | 0.93567 | 0.9 | 1.1448 | 1.1 |  | 0.9 | 1.17 | 0.92 | 0.92 | 1 | 1.025
| Std Dev. |  |  | 0.17166 |  | 0.048751 | 0.083666 | 0.033862 | 0.048751 | 0.037238 | 0.040208 | 0.020412 | 0.012649 | 0.040825 | 0.04065 | 0.05164 |  | 0.1169 |  | 0.040825 | 0.074744 | 0.037771 | 0.023166 | 0.040825 | 0.024221
| Rel.Std.Dev. |  |  | 19.36% |  | 5.488% | 9.843% | 3.014% | 4.772% | 4.018% | 3.585% | 2.065% | 1.375% | 3.768% | 4.349% | 5.958% |  | 10.47% |  | 4.454% | 6.316% | 4.091% | 2.541% | 4.016% | 2.336%
| PDM3 |  |  | -9.784% |  | -9.615% | -13.52% | 14.3% | 3.952% | -5.714% | 14.13% | 0.5599% | -6.393% | 10.23% | -4.906% | -11.82% | 16.48% | 13.62% |  | -6.732% | 20.4% | -6.054% | -7.241% | 3.443% | 5.478%
| Z-Score (Absolute) |  | 10.19 | 0.69 | 10.19 | 0.67 | NA | 0.91 | 0.22 | 0.42 | 0.9 | 0 | 0.46 | NA | 0.36 | NA | 2.57 | NA | 27.05 | NA | 1.31 | 0.44 | 0.52 | NA | 0.33
| Analytical results for Y in OREAS 556b (Certified Value 4.03 ppm) | [28 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 23 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [6 empty] | 2
| 1 | 1 | < 5 | 4.36 | 4.33 | 3.55 | 4.04 | 4.41 | 3.57 | 3.66 | 3.75 | < 10 | 3.95 | 4.5178 | 5.28 | 3.8619 | 5.3 | 3.3 | 4 | 3.88 | 4.87 | 4.78 | 4.04 | [6 empty] | 1
| 1 | 2 | < 5 | 4.33 | 3.76 | 3.46 | 4.14 | 4.7 | 3.71 | 3.65 | 3.79 | < 10 | 3.96 | 4.3581 | 5.22 | 3.9625 | 5.2 | 3.4 | 4 | 3.7 | 5.26 | 4.86 | 3.91 | [6 empty] | 13
| 1 | 3 | < 5 | 4.34 | 4.05 | 3.61 | 3.56 | 4.63 | 3.62 | 3.62 | 3.75 | < 10 | 3.86 | 4.495 | 5.33 | 3.9908 | 4.9 | 3.4 | 3.9 | 3.74 | 4.53 | 4.76 | 3.9 | [6 empty] | 16
| 1 | 4 | < 5 | 4.43 | 4.24 | 3.47 | 3.59 | 4.57 | 3.67 | 3.71 | 3.85 | < 10 | 4.03 | 4.2966 | 5.41 | 4.2295 | 4.8 | 3.3 | 3.9 | 3.99 | 6.17 | 4.71 | 3.9 | [6 empty] | 4.0253
| 1 | 5 | < 5 | 4.38 | 3.87 | 3.39 | 3.6 | 4.47 | 3.63 | 3.64 | 3.86 | < 10 | 4.09 | 4.3202 | 5.21 | 3.9881 | 4.8 | 3.4 | 3.8 | 3.68 | 5.17 | 4.82 | 3.78 | [6 empty] | 203
| 1 | 6 | < 5 | 4.33 | 3.85 | 3.47 | 3.64 | 4.67 | 3.77 | 3.57 | 3.78 | < 10 | 4 | 4.514 | 5.15 | 3.9012 | 4.8 | 3.4 | 3.9 | 3.75 | 5.28 | 4.98 | 4.02
| Mean |  |  | 4.3617 | 4.0167 | 3.4917 | 3.7617 | 4.575 | 3.6617 | 3.6417 | 3.7967 |  | 3.9817 | 4.4169 | 5.2667 | 3.989 | 4.9667 | 3.3667 | 3.9167 | 3.79 | 5.2133 | 4.8183 | 3.925
| Median |  |  | 4.35 | 3.96 | 3.47 | 3.62 | 4.6 | 3.65 | 3.645 | 3.785 |  | 3.98 | 4.4266 | 5.25 | 3.9753 | 4.85 | 3.4 | 3.9 | 3.745 | 5.215 | 4.8 | 3.905
| Std Dev. |  |  | 0.038687 | 0.22993 | 0.07705 | 0.25756 | 0.11485 | 0.07111 | 0.046224 | 0.048028 |  | 0.078337 | 0.10297 | 0.093524 | 0.1284 | 0.22509 | 0.05164 | 0.075277 | 0.12033 | 0.54928 | 0.094322 | 0.094604
| Rel.Std.Dev. |  |  | 0.887% | 5.724% | 2.207% | 6.847% | 2.51% | 1.942% | 1.269% | 1.265% |  | 1.967% | 2.331% | 1.776% | 3.219% | 4.532% | 1.534% | 1.922% | 3.175% | 10.54% | 1.958% | 2.41%
| PDM3 |  |  | 8.357% | -0.2136% | -13.26% | -6.549% | 13.66% | -9.033% | -9.53% | -5.679% |  | -1.083% | 9.73% | 30.84% | -0.9008% | 23.39% | -16.36% | -2.698% | -5.845% | 29.52% | 19.7% | -2.491%
| Z-Score (Absolute) |  | 2.63 | 0.67 | 0.06 | 0.87 | 0.39 | 1.05 | 0.57 | 0.6 | 0.33 | 1.81 | 0 | 0.77 | 2.28 | 0.01 | 1.75 | 1.09 | 0.12 | 0.34 | 2.19 | 1.48 | 0.1
| Analytical results for Yb in OREAS 556b (Certified Value 0.38 ppm) | [28 empty] | Certified
Yb | Round | Replicate | 03: Lab; 04: Lab; 05: Lab; 06: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | [6 empty] | 1
| No | No | 03: 04; 04: 05; 05: 14; 06: 16; 08: 18; 09: 20; 10: 23; 11: 26 | [6 empty] | ppm
|  |  | 03: AR*MS; 04: AR*OES/MS; 05: AR*MS; 06: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS | [6 empty] | 2
|  |  | 03: 01g; 04: 0.5g; 05: 0.5g; 06: 0.25g; 08: 0.2g; 09: 0.2g; 10: 0.25g; 11: 01g | [6 empty] | 2
| 1 | 1 | 03: 0.366; 04: 0.4; 05: 0.28; 06: 0.19301; 08: 0.3389; 09: 0.6; 10: 0.4; 11: 0.38 | [6 empty] | 1
| 1 | 2 | 03: 0.34; 04: 0.4; 05: 0.28; 06: 0.18137; 08: 0.31681; 09: 0.5; 10: 0.4; 11: 0.39 | [6 empty] | 14
| 1 | 3 | 03: 0.378; 04: 0.4; 05: 0.28; 06: 0.19777; 08: 0.31648; 09: 0.5; 10: 0.4; 11: 0.41 | [6 empty] | 16
| 1 | 4 | 03: 0.387; 04: 0.4; 05: 0.28; 06: 0.19357; 08: 0.3857; 09: 0.5; 10: 0.4; 11: 0.4 | [6 empty] | 0.3827
| 1 | 5 | 03: 0.363; 04: 0.4; 05: 0.29; 06: 0.18544; 08: 0.3327; 09: 0.5; 10: 0.4; 11: 0.43 | [6 empty] | 204
| 1 | 6 | 03: 0.364; 04: 0.4; 05: 0.28; 06: 0.19188; 08: 0.32461; 09: 0.5; 10: 0.4; 11: 0.42
| Mean |  | 03: 0.36633; 04: 0.4; 05: 0.28167; 06: 0.19051; 08: 0.33587; 09: 0.51667; 10: 0.4; 11: 0.405
| Median |  | 03: 0.365; 04: 0.4; 05: 0.28; 06: 0.19245; 08: 0.32866; 09: 0.5; 10: 0.4; 11: 0.405
| Std Dev. |  | 03: 0.015958; 04: 6.0809e-17; 05: 0.0040825; 06: 0.0059914; 08: 0.025958; 09: 0.040825; 10: 6.0809e-17; 11: 0.018708
| Rel.Std.Dev. |  | 03: 4.356%; 04: 1.52e-14%; 05: 1.449%; 06: 3.145%; 08: 7.729%; 09: 7.902%; 10: 1.52e-14%; 11: 4.619%
| PDM3 |  | 03: -4.277%; 04: 4.521%; 05: -26.4%; 06: -50.22%; 08: -12.24%; 09: 35.01%; 10: 4.521%; 11: 5.827%
| Z-Score (Absolute) |  | 03: 0.33; 04: 0.33; 05: 1.98; 06: 3.76; 08: 0.92; 09: 2.6; 10: 0.33; 11: 0.43
| Analytical results for Zn in OREAS 556b (Certified Value 37.1 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | 27 | 28 | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.1g | 01g | 01g | 0.5g | 2
| 1 | 1 | 38.3 | 44.7 | 42.1 | 38 | 33.4 | 41.2 | 35.9394 | 34 | 36 | 35 | 37 | 35 | 36.16 | 41.6 | 38.2706 | 38.56 | 40 | 35.1414 | 37 | 41.3 | 36.2 | 39 | 31 | 31 | 37 | 38 | 34 | 1
| 1 | 2 | 38.1 | 48.3 | 42.4 | 37 | 29.3 | 37.9 | 36.9501 | 35 | 39 | 36 | 35 | 35 | 35.34 | 41.7 | 37.9666 | 38.26 | 38 | 36.5248 | 38 | 41.3 | 37.7 | 38 | 28 | 28.6 | 37 | 36 | 34 | 15
| 1 | 3 | 37.3 | 40 | 41.6 | 37 | 32.5 | 39.5 | 37.7661 | 36 | 38 | 35 | 34 | 36 | 36.11 | 41.9 | 37.9689 | 39.29 | 39 | 36.7267 | 37 | 41 | 35.8 | 37 | 27 | 30.9 | 37 | 37 | 34 | 16
| 1 | 4 | 36.8 | 43 | 41.3 | 38 | 33.8 | 38.5 | 36.3455 | 35 | 36 | 36 | 34 | 36 | 35.84 | 42.2 | 38.2704 | 38.93 | 40 | 40.4595 | 38 | 40.9 | 36.9 | 38 | 30 | 31.6 | 36 | 37 | 34 | 37.1266
| 1 | 5 | 35.1 | 44.4 | 41.7 | 38 | 30.1 | 37.3 | 35.2872 | 34 | 37 | 35 | 34 | 35 | 35.92 | 42.7 | 38.4228 | 38.31 | 38 | 36.4079 | 36 | 39.6 | 36.6 | 36 | 28 | 30.1 | 37 | 37 | 34 | 205
| 1 | 6 | 34.8 | 20.6 | 41.1 | 37 | 31.9 | 40.2 | 36.6458 | 35 | 38 | 36 | 34 | 34 | 35.99 | 41.4 | 38.4413 | 38.11 | 38 | 36.0599 | 36 | 41.3 | 37 | 36 | 26 | 28.7 | 38 | 38 | 35
| Mean |  | 36.7333 | 40.1667 | 41.7 | 37.5 | 31.8333 | 39.1 | 36.489 | 34.8333 | 37.3333 | 35.5 | 34.6667 | 35.1667 | 35.8933 | 41.9167 | 38.2234 | 38.5767 | 38.8333 | 36.8867 | 37 | 40.9 | 36.7 | 37.3333 | 28.3333 | 30.15 | 37 | 37.1667 | 34.1667
| Median |  | 37.05 | 43.7 | 41.65 | 37.5 | 32.2 | 39 | 36.4957 | 35 | 37.5 | 35.5 | 34 | 35 | 35.955 | 41.8 | 38.2705 | 38.435 | 38.5 | 36.4663 | 37 | 41.15 | 36.75 | 37.5 | 28 | 30.5 | 37 | 37 | 34
| Std Dev. |  | 1.4868 | 9.9552 | 0.4858 | 0.54772 | 1.7996 | 1.4711 | 0.85268 | 0.75277 | 1.2111 | 0.54772 | 1.2111 | 0.75277 | 0.29568 | 0.47081 | 0.21091 | 0.45236 | 0.98319 | 1.8374 | 0.89443 | 0.6603 | 0.66332 | 1.2111 | 1.8619 | 1.2566 | 0.63246 | 0.75277 | 0.40825
| Rel.Std.Dev. |  | 4.048% | 24.78% | 1.165% | 1.461% | 5.653% | 3.762% | 2.337% | 2.161% | 3.244% | 1.543% | 3.493% | 2.141% | 0.8238% | 1.123% | 0.5518% | 1.173% | 2.532% | 4.981% | 2.417% | 1.614% | 1.807% | 3.244% | 6.571% | 4.168% | 1.709% | 2.025% | 1.195%
| PDM3 |  | -1.059% | 8.188% | 12.32% | 1.006% | -14.26% | 5.315% | -1.717% | -6.177% | 0.5569% | -4.381% | -6.626% | -5.279% | -3.322% | 12.9% | 2.954% | 3.906% | 4.597% | -0.6461% | -0.3409% | 10.16% | -1.149% | 0.5569% | -23.68% | -18.79% | -0.3409% | 0.108% | -7.972%
| Z-Score (Absolute) |  | 0.11 | 1.35 | 2.01 | 0.21 | 2.21 | 0.9 | 0.22 | 0.93 | 0.14 | 0.64 | 1 | 0.78 | 0.47 | 2.1 | 0.52 | 0.67 | 0.78 | 0.05 | 0 | 1.67 | 0.13 | 0.14 | 3.71 | 2.93 | 0 | 0.07 | 1.21
| Analytical results for Zr in OREAS 556b (Certified Value 11.1 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 22 | 25 | 26 | 28 | [6 empty] | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 1
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 01g | 0.2g | 0.2g | 0.25g | 0.25g | 0.1g | 01g | 0.5g | [6 empty] | 1
| 1 | 1 | 12.6 | 13.1 | 12.3 | 13.1 | 9.1 | 9.6 | 10.5 | 9.3 | 10.2 | 9.6 | 9.78 | 9.91 | 13.6627 | 13.3 | 8.9378 | 14.9 | 4.4 | 10.1 | 9.3 | 11.5 | 10.8 | [6 empty] | 1
| 1 | 2 | 14 | 13 | 12.1 | 11.3 | 9 | 9.7 | 10.5 | 9.6 | 10.3 | 10.6 | 9.64 | 10.2 | 13.6857 | 12.7 | 9.2536 | 14.4 | 4.3 | 10 | 11 | 11.9 | 10.6 | [6 empty] | 18
| 1 | 3 | 14.7 | 12.7 | 12 | 12.5 | 9.3 | 9.6 | 10.9 | 10 | 10.4 | 9.5 | 9.84 | 10.3 | 12.8481 | 12.7 | 9.3585 | 14.1 | 4.5 | 10 | 8.4 | 11.5 | 10.5 | [6 empty] | 16
| 1 | 4 | 12.5 | 12.5 | 12.4 | 12.7 | 8.9 | 9.7 | 10.7 | 9.9 | 10.4 | 9.9 | 9.83 | 10.5 | 13.596 | 12.7 | 9.9722 | 13.3 | 4.3 | 9.9 | 12.1 | 11.3 | 10.4 | [6 empty] | 11.0788
| 1 | 5 | 15.1 | 12.7 | 12 | 11.9 | 8.9 | 9.7 | 10.5 | 9.6 | 10.4 | 9.9 | 9.77 | 10.3 | 12.7963 | 12.9 | 9.6334 | 13.1 | 4.4 | 9.5 | 10.6 | 11.4 | 10.2 | [6 empty] | 206
| 1 | 6 | 13.4 | 12.8 | 11.9 | 11.4 | 9.1 | 10 | 10.4 | 10.3 | 10.2 | 10 | 9.57 | 10.2 | 12.9862 | 12.9 | 9.2039 | 12.4 | 4.4 | 9.8 | 10.9 | 11.6 | 10.8
| Mean |  | 13.7167 | 12.8 | 12.1167 | 12.15 | 9.05 | 9.7167 | 10.5833 | 9.7833 | 10.3167 | 9.9167 | 9.7383 | 10.235 | 13.2625 | 12.8667 | 9.3932 | 13.7 | 4.3833 | 9.8833 | 10.3833 | 11.5333 | 10.55
| Median |  | 13.7 | 12.75 | 12.05 | 12.2 | 9.05 | 9.7 | 10.5 | 9.75 | 10.35 | 9.9 | 9.775 | 10.25 | 13.2911 | 12.8 | 9.3061 | 13.7 | 4.4 | 9.95 | 10.75 | 11.5 | 10.55
| Std Dev. |  | 1.0759 | 0.21909 | 0.19408 | 0.73144 | 0.15166 | 0.1472 | 0.18348 | 0.35449 | 0.098319 | 0.38687 | 0.10907 | 0.19326 | 0.42799 | 0.23381 | 0.36244 | 0.92736 | 0.075277 | 0.2137 | 1.3227 | 0.20656 | 0.23452
| Rel.Std.Dev. |  | 7.844% | 1.712% | 1.602% | 6.02% | 1.676% | 1.515% | 1.734% | 3.623% | 0.953% | 3.901% | 1.12% | 1.888% | 3.227% | 1.817% | 3.858% | 6.769% | 1.717% | 2.162% | 12.74% | 1.791% | 2.223%
| PDM3 |  | 23.81% | 15.54% | 9.368% | 9.669% | -18.31% | -12.29% | -4.472% | -11.69% | -6.879% | -10.49% | -12.1% | -7.616% | 19.71% | 16.14% | -15.21% | 23.66% | -60.43% | -10.79% | -6.277% | 4.103% | -4.773%
| Z-Score (Absolute) |  | 2.27 | 1.65 | 1.18 | 1.2 | 0.91 | 0.45 | 0.14 | 0.41 | 0.05 | 0.32 | 0.44 | 0.1 | 1.96 | 1.69 | 0.67 | 2.26 | 4.09 | 0.34 | 0 | 0.78 | 0.11
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C513: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1116: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1188: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1206: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 556b (Indicative Value 0.39 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.39 | 1
| 1 | 2 | 0.37 | 1
| 1 | 3 | 0.41 | 16
| 1 | 4 | 0.39 | 0.39
| 1 | 5 | 0.4 | 34
| 1 | 6 | 0.38
| Mean |  | 0.39
| Median |  | 0.39
| Std Dev. |  | 0.014142
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 556b (Indicative Value 10.5 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 10.6 | 1
| 1 | 2 | 10.6 | 29
| 1 | 3 | 10.6 | 16
| 1 | 4 | 10.5 | 10.5
| 1 | 5 | 10.4 | 35
| 1 | 6 | 10.3
| Mean |  | 10.5
| Median |  | 10.55
| Std Dev. |  | 0.12649
| Rel.Std.Dev. |  | 1.205%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 556b (Indicative Value 180 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 185 | 30
| 1 | 3 | 181 | 16
| 1 | 4 | 179 | 180.1667
| 1 | 5 | 178 | 36
| 1 | 6 | 178
| Mean |  | 180.1667
| Median |  | 179.5
| Std Dev. |  | 2.6394
| Rel.Std.Dev. |  | 1.465%
| PDM3 |  | -1.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 556b (Indicative Value 5.45 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 5.4 | 3
| 1 | 3 | 5.6 | 16
| 1 | 4 | 5.4 | 5.45
| 1 | 5 | 5.2 | 37
| 1 | 6 | 5.3
| Mean |  | 5.45
| Median |  | 5.4
| Std Dev. |  | 0.21679
| Rel.Std.Dev. |  | 3.978%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 556b (Indicative Value 3.47 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.5 | 4
| 1 | 3 | 3.5 | 16
| 1 | 4 | 3.5 | 3.4667
| 1 | 5 | 3.4 | 38
| 1 | 6 | 3.5
| Mean |  | 3.4667
| Median |  | 3.5
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 1.49%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 556b (Indicative Value 0.177 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.18 | 1
| 1 | 2 | 0.18 | 33
| 1 | 3 | 0.18 | 16
| 1 | 4 | 0.18 | 0.17667
| 1 | 5 | 0.17 | 39
| 1 | 6 | 0.17
| Mean |  | 0.17667
| Median |  | 0.18
| Std Dev. |  | 0.005164
| Rel.Std.Dev. |  | 2.923%
| PDM3 |  | -1.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 556b (Indicative Value 0.65 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.7 | 5
| 1 | 3 | 0.6 | 16
| 1 | 4 | 0.7 | 0.65
| 1 | 5 | 0.7 | 40
| 1 | 6 | 0.7
| Mean |  | 0.65
| Median |  | 0.7
| Std Dev. |  | 0.083666
| Rel.Std.Dev. |  | 12.87%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 556b (Indicative Value 137 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 139 | 1
| 1 | 2 | 137 | 6
| 1 | 3 | 140 | 16
| 1 | 4 | 138 | 137
| 1 | 5 | 133 | 41
| 1 | 6 | 135
| Mean |  | 137
| Median |  | 137.5
| Std Dev. |  | 2.6077
| Rel.Std.Dev. |  | 1.903%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 556b (Indicative Value 2.08 wt.%) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.04 | 1
| 1 | 2 | 2.07 | 36
| 1 | 3 | 2.1 | 16
| 1 | 4 | 2.14 | 2.0817
| 1 | 5 | 2.06 | 42
| 1 | 6 | 2.08
| Mean |  | 2.0817
| Median |  | 2.075
| Std Dev. |  | 0.034881
| Rel.Std.Dev. |  | 1.676%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 556b (Indicative Value 98 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 95 | 1
| 1 | 2 | 97 | 37
| 1 | 3 | 97 | 16
| 1 | 4 | 96 | 97.6667
| 1 | 5 | 107 | 43
| 1 | 6 | 94
| Mean |  | 97.6667
| Median |  | 96.5
| Std Dev. |  | 4.7188
| Rel.Std.Dev. |  | 4.831%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 556b (Indicative Value 1.78 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.7 | 1
| 1 | 2 | 1.8 | 9
| 1 | 3 | 1.8 | 16
| 1 | 4 | 1.8 | 1.7833
| 1 | 5 | 1.8 | 44
| 1 | 6 | 1.8
| Mean |  | 1.7833
| Median |  | 1.8
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 2.289%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 556b (Indicative Value 8.94 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*MS | 2
2
| 1 | 1 | 8.86 | 1
| 1 | 2 | 8.8 | 39
| 1 | 3 | 8.99 | 16
| 1 | 4 | 8.92 | 8.9367
| 1 | 5 | 8.93 | 45
| 1 | 6 | 9.12
| Mean |  | 8.9367
| Median |  | 8.925
| Std Dev. |  | 0.11075
| Rel.Std.Dev. |  | 1.239%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 556b (Indicative Value 4.62 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.63 | 1
| 1 | 2 | 4.46 | 10
| 1 | 3 | 4.66 | 16
| 1 | 4 | 4.64 | 4.6233
| 1 | 5 | 4.61 | 46
| 1 | 6 | 4.74
| Mean |  | 4.6233
| Median |  | 4.635
| Std Dev. |  | 0.091797
| Rel.Std.Dev. |  | 1.986%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 556b (Indicative Value 2.75 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.77 | 1
| 1 | 2 | 2.71 | 11
| 1 | 3 | 2.83 | 16
| 1 | 4 | 2.69 | 2.75
| 1 | 5 | 2.76 | 47
| 1 | 6 | 2.74
| Mean |  | 2.75
| Median |  | 2.75
| Std Dev. |  | 0.049396
| Rel.Std.Dev. |  | 1.796%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 556b (Indicative Value 0.94 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.93 | 1
| 1 | 2 | 0.92 | 12
| 1 | 3 | 0.97 | 16
| 1 | 4 | 0.9 | 0.93667
| 1 | 5 | 0.93 | 48
| 1 | 6 | 0.97
| Mean |  | 0.93667
| Median |  | 0.93
| Std Dev. |  | 0.028048
| Rel.Std.Dev. |  | 2.994%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 556b (Indicative Value 9.34 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 9.39 | 1
| 1 | 2 | 9.46 | 43
| 1 | 3 | 9.43 | 16
| 1 | 4 | 9.36 | 9.3383
| 1 | 5 | 9.27 | 49
| 1 | 6 | 9.12
| Mean |  | 9.3383
| Median |  | 9.375
| Std Dev. |  | 0.12545
| Rel.Std.Dev. |  | 1.343%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 556b (Indicative Value 27.6 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 27.3 | 1
| 1 | 2 | 27.1 | 13
| 1 | 3 | 28 | 16
| 1 | 4 | 27.5 | 27.6167
| 1 | 5 | 27.9 | 50
| 1 | 6 | 27.9
| Mean |  | 27.6167
| Median |  | 27.7
| Std Dev. |  | 0.37103
| Rel.Std.Dev. |  | 1.344%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 556b (Indicative Value 5.47 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.3 | 1
| 1 | 2 | 5.6 | 14
| 1 | 3 | 5.6 | 16
| 1 | 4 | 5.4 | 5.4667
| 1 | 5 | 5.6 | 34
| 1 | 6 | 5.3
| Mean |  | 5.4667
| Median |  | 5.5
| Std Dev. |  | 0.15055
| Rel.Std.Dev. |  | 2.754%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 556b (Indicative Value 5.5 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.4 | 1
| 1 | 2 | 5.4 | 19
| 1 | 3 | 5.6 | 16
| 1 | 4 | 5.6 | 5.5
| 1 | 5 | 5.6 | 35
| 1 | 6 | 5.4
| Mean |  | 5.5
| Median |  | 5.5
| Std Dev. |  | 0.10954
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 556b (Indicative Value 0.86 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.84 | 3
| 1 | 3 | 0.88 | 16
| 1 | 4 | 0.86 | 0.85667
| 1 | 5 | 0.87 | 36
| 1 | 6 | 0.85
| Mean |  | 0.85667
| Median |  | 0.855
| Std Dev. |  | 0.01633
| Rel.Std.Dev. |  | 1.906%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 556b (Indicative Value 2.68 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 2.7 | 1
| 1 | 2 | 2.71 | 31
| 1 | 3 | 2.73 | 16
| 1 | 4 | 2.69 | 2.6833
| 1 | 5 | 2.66 | 37
| 1 | 6 | 2.61
| Mean |  | 2.6833
| Median |  | 2.695
| Std Dev. |  | 0.04274
| Rel.Std.Dev. |  | 1.593%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 556b (Indicative Value 77 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 78 | 1
| 1 | 2 | 78 | 5
| 1 | 3 | 78 | 16
| 1 | 4 | 78 | 77.3333
| 1 | 5 | 76 | 38
| 1 | 6 | 76
| Mean |  | 77.3333
| Median |  | 78
| Std Dev. |  | 1.0328
| Rel.Std.Dev. |  | 1.336%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 556b (Indicative Value 53 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 53 | 1
| 1 | 2 | 53 | 33
| 1 | 3 | 54 | 16
| 1 | 4 | 53 | 52.6667
| 1 | 5 | 52 | 39
| 1 | 6 | 51
| Mean |  | 52.6667
| Median |  | 53
| Std Dev. |  | 1.0328
| Rel.Std.Dev. |  | 1.961%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 556b (Indicative Value 4.03 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 4.02 | 1
| 1 | 2 | 4.03 | 34
| 1 | 3 | 4.02 | 16
| 1 | 4 | 4 | 4.0283
| 1 | 5 | 3.96 | 40
| 1 | 6 | 4.14
| Mean |  | 4.0283
| Median |  | 4.02
| Std Dev. |  | 0.060139
| Rel.Std.Dev. |  | 1.493%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 556b (Indicative Value 0.146 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.149 | 1
| 1 | 2 | 0.149 | 35
| 1 | 3 | 0.144 | 16
| 1 | 4 | 0.15 | 0.14633
| 1 | 5 | 0.147 | 41
| 1 | 6 | 0.139
| Mean |  | 0.14633
| Median |  | 0.148
| Std Dev. |  | 0.0041793
| Rel.Std.Dev. |  | 2.856%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 556b (Indicative Value 20.8 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 20.4 | 1
| 1 | 2 | 21.1 | 7
| 1 | 3 | 21 | 16
| 1 | 4 | 20.5 | 20.8167
| 1 | 5 | 21 | 42
| 1 | 6 | 20.9
| Mean |  | 20.8167
| Median |  | 20.95
| Std Dev. |  | 0.29269
| Rel.Std.Dev. |  | 1.406%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 556b (Indicative Value 0.067 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.07 | 8
| 1 | 3 | 0.07 | 16
| 1 | 4 | 0.07 | 0.066667
| 1 | 5 | 0.06 | 43
| 1 | 6 | 0.06
| Mean |  | 0.066667
| Median |  | 0.07
| Std Dev. |  | 0.005164
| Rel.Std.Dev. |  | 7.746%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 556b (Indicative Value 20.3 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 20.3 | 1
| 1 | 2 | 19.8 | 9
| 1 | 3 | 20.5 | 16
| 1 | 4 | 20.6 | 20.3167
| 1 | 5 | 20.3 | 44
| 1 | 6 | 20.4
| Mean |  | 20.3167
| Median |  | 20.35
| Std Dev. |  | 0.27869
| Rel.Std.Dev. |  | 1.372%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 556b (Indicative Value 64 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 63.8 | 1
| 1 | 2 | 64 | 10
| 1 | 3 | 65.4 | 16
| 1 | 4 | 65 | 64.4833
| 1 | 5 | 65 | 45
| 1 | 6 | 63.7
| Mean |  | 64.4833
| Median |  | 64.5
| Std Dev. |  | 0.73326
| Rel.Std.Dev. |  | 1.137%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 556b (Indicative Value 383 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 365 | 1
| 1 | 2 | 382 | 40
| 1 | 3 | 383 | 16
| 1 | 4 | 395 | 383
| 1 | 5 | 378 | 46
| 1 | 6 | 395
| Mean |  | 383
| Median |  | 382.5
| Std Dev. |  | 11.296
| Rel.Std.Dev. |  | 2.949%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 556b (Indicative Value < 0.002 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | < 0.002 | 1
| 1 | 2 | < 0.002 | 41
| 1 | 3 | < 0.002 | 16
| 1 | 4 | < 0.002 | < 0.002
| 1 | 5 | < 0.002 | 47
| 1 | 6 | < 0.002
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 556b (Indicative Value 19.7 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 19.7 | 1
| 1 | 2 | 19.4 | 11
| 1 | 3 | 20 | 16
| 1 | 4 | 19.8 | 19.6833
| 1 | 5 | 19.4 | 48
| 1 | 6 | 19.8
| Mean |  | 19.6833
| Median |  | 19.75
| Std Dev. |  | 0.24014
| Rel.Std.Dev. |  | 1.22%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 556b (Indicative Value 17.2 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 17.2 | 1
| 1 | 2 | 16.9 | 12
| 1 | 3 | 17.2 | 16
| 1 | 4 | 17.1 | 17.1667
| 1 | 5 | 17.4 | 49
| 1 | 6 | 17.2
| Mean |  | 17.1667
| Median |  | 17.2
| Std Dev. |  | 0.1633
| Rel.Std.Dev. |  | 0.9513%
| PDM3 |  | -1.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 556b (Indicative Value 82 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 80.8 | 1
| 1 | 2 | 80.8 | 13
| 1 | 3 | 84 | 16
| 1 | 4 | 81.4 | 81.8
| 1 | 5 | 81.8 | 50
| 1 | 6 | 82
| Mean |  | 81.8
| Median |  | 81.6
| Std Dev. |  | 1.1866
| Rel.Std.Dev. |  | 1.451%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 556b (Indicative Value 7.88 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 7.84 | 1
| 1 | 2 | 7.94 | 28
| 1 | 3 | 8 | 16
| 1 | 4 | 8.04 | 7.8767
| 1 | 5 | 7.74 | 34
| 1 | 6 | 7.7
| Mean |  | 7.8767
| Median |  | 7.89
| Std Dev. |  | 0.13938
| Rel.Std.Dev. |  | 1.77%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 556b (Indicative Value 11.4 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 11 | 1
| 1 | 2 | 11.1 | 1
| 1 | 3 | 11.7 | 16
| 1 | 4 | 11.5 | 11.4167
| 1 | 5 | 11.6 | 35
| 1 | 6 | 11.6
| Mean |  | 11.4167
| Median |  | 11.55
| Std Dev. |  | 0.29269
| Rel.Std.Dev. |  | 2.564%
| PDM3 |  | -2.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 556b (Indicative Value 8.18 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 8 | 1
| 1 | 2 | 8.1 | 3
| 1 | 3 | 8.3 | 16
| 1 | 4 | 8.3 | 8.1833
| 1 | 5 | 8 | 36
| 1 | 6 | 8.4
| Mean |  | 8.1833
| Median |  | 8.2
| Std Dev. |  | 0.17224
| Rel.Std.Dev. |  | 2.105%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 556b (Indicative Value 3.62 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.66 | 1
| 1 | 2 | 3.59 | 20
| 1 | 3 | 3.58 | 16
| 1 | 4 | 3.67 | 3.6167
| 1 | 5 | 3.64 | 37
| 1 | 6 | 3.56
| Mean |  | 3.6167
| Median |  | 3.615
| Std Dev. |  | 0.045898
| Rel.Std.Dev. |  | 1.269%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 556b (Indicative Value 67 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 68 | 1
| 1 | 2 | 68 | 32
| 1 | 3 | 68 | 16
| 1 | 4 | 67 | 67.1667
| 1 | 5 | 66 | 38
| 1 | 6 | 66
| Mean |  | 67.1667
| Median |  | 67.5
| Std Dev. |  | 0.98319
| Rel.Std.Dev. |  | 1.464%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 556b (Indicative Value 0.31 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.31 | 1
| 1 | 2 | 0.31 | 5
| 1 | 3 | 0.3 | 16
| 1 | 4 | 0.32 | 0.31
| 1 | 5 | 0.31 | 39
| 1 | 6 | 0.31
| Mean |  | 0.31
| Median |  | 0.31
| Std Dev. |  | 0.0063246
| Rel.Std.Dev. |  | 2.04%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 556b (Indicative Value 0.68 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.69 | 1
| 1 | 2 | 0.66 | 6
| 1 | 3 | 0.7 | 16
| 1 | 4 | 0.66 | 0.675
| 1 | 5 | 0.67 | 40
| 1 | 6 | 0.67
| Mean |  | 0.675
| Median |  | 0.67
| Std Dev. |  | 0.016432
| Rel.Std.Dev. |  | 2.434%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 556b (Indicative Value 17.3 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 17.1 | 1
| 1 | 2 | 16.9 | 8
| 1 | 3 | 17.4 | 16
| 1 | 4 | 17.3 | 17.25
| 1 | 5 | 17.4 | 41
| 1 | 6 | 17.4
| Mean |  | 17.25
| Median |  | 17.35
| Std Dev. |  | 0.20736
| Rel.Std.Dev. |  | 1.202%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 556b (Indicative Value 0.669 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 14 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.674 | 1
| 1 | 2 | 0.68 | 36
| 1 | 3 | 0.674 | 16
| 1 | 4 | 0.67 | 0.66867
| 1 | 5 | 0.66 | 42
| 1 | 6 | 0.654
| Mean |  | 0.66867
| Median |  | 0.672
| Std Dev. |  | 0.0097707
| Rel.Std.Dev. |  | 1.461%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 556b (Indicative Value 9.5 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 9.55 | 1
| 1 | 2 | 9.4 | 11
| 1 | 3 | 9.7 | 16
| 1 | 4 | 9.54 | 9.4983
| 1 | 5 | 9.44 | 43
| 1 | 6 | 9.36
| Mean |  | 9.4983
| Median |  | 9.49
| Std Dev. |  | 0.12432
| Rel.Std.Dev. |  | 1.309%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 556b (Indicative Value 332 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 332 | 1
| 1 | 2 | 337 | 38
| 1 | 3 | 337 | 16
| 1 | 4 | 332 | 332.3333
| 1 | 5 | 327 | 44
| 1 | 6 | 329
| Mean |  | 332.3333
| Median |  | 332
| Std Dev. |  | 4.0825
| Rel.Std.Dev. |  | 1.228%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 556b (Indicative Value 2.95 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.9 | 1
| 1 | 2 | 2.9 | 12
| 1 | 3 | 3 | 16
| 1 | 4 | 2.9 | 2.95
| 1 | 5 | 3 | 45
| 1 | 6 | 3
| Mean |  | 2.95
| Median |  | 2.95
| Std Dev. |  | 0.054772
| Rel.Std.Dev. |  | 1.857%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 556b (Indicative Value 25.9 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 25.5 | 1
| 1 | 2 | 25.3 | 13
| 1 | 3 | 26.9 | 16
| 1 | 4 | 25.7 | 25.8667
| 1 | 5 | 25.9 | 46
| 1 | 6 | 25.9
| Mean |  | 25.8667
| Median |  | 25.8
| Std Dev. |  | 0.55737
| Rel.Std.Dev. |  | 2.155%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 556b (Indicative Value 2.74 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.74 | 1
| 1 | 2 | 2.64 | 14
| 1 | 3 | 2.75 | 16
| 1 | 4 | 2.7 | 2.735
| 1 | 5 | 2.79 | 47
| 1 | 6 | 2.79
| Mean |  | 2.735
| Median |  | 2.745
| Std Dev. |  | 0.057533
| Rel.Std.Dev. |  | 2.104%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 556b (Indicative Value 76 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 74 | 1
| 1 | 2 | 81 | 15
| 1 | 3 | 77 | 16
| 1 | 4 | 75 | 75.8333
| 1 | 5 | 73 | 48
| 1 | 6 | 75
| Mean |  | 75.8333
| Median |  | 75
| Std Dev. |  | 2.8577
| Rel.Std.Dev. |  | 3.768%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 556b (Indicative Value 179 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 14 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 179 | 1
| 1 | 2 | 182 | 18
| 1 | 3 | 183 | 16
| 1 | 4 | 179 | 178.6667
| 1 | 5 | 174 | 49
| 1 | 6 | 175
| Mean |  | 178.6667
| Median |  | 179
| Std Dev. |  | 3.6148
| Rel.Std.Dev. |  | 2.023%
| PDM3 |  | -1.887e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.