## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 285
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.6561 | 0.64808 | 0.66413 | 0.65191 | 0.66029
PhotonAssay
Au, Gold (ppm) | 0.66944 | 0.65922 | 0.67967 | 0.66819 | 0.6707
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.64066 | 0.62649 | 0.65484 | 0.63618 | 0.64514
Cyanide Leach
Au, Gold (ppm) | 0.63488 | 0.61526 | 0.6545 | 0.63331 | 0.63645
4-Acid Digestion
Ag, Silver (ppm) | 0.090636 | 0.068189 | 0.11308 | IND | IND
Al, Aluminium (wt.%) | 7.759 | 7.5063 | 8.0118 | 7.614 | 7.9041
As, Arsenic (ppm) | 288.5915 | 276.2403 | 300.9428 | 277.495 | 299.688
Ba, Barium (ppm) | 1031.6368 | 995.5639 | 1067.7096 | 1006.1133 | 1057.1603
Be, Beryllium (ppm) | 2.5058 | 2.3779 | 2.6338 | 2.4355 | 2.5762
Bi, Bismuth (ppm) | 0.586 | 0.45869 | 0.71331 | 0.48386 | 0.68815
Ca, Calcium (wt.%) | 1.939 | 1.8865 | 1.9916 | 1.8971 | 1.9809
Cd, Cadmium (ppm) | 0.20164 | 0.17344 | 0.22983 | 0.1769 | 0.22638
Ce, Cerium (ppm) | 78.0157 | 72.8684 | 83.1631 | 74.6246 | 81.4069
Co, Cobalt (ppm) | 14.1293 | 13.4421 | 14.8165 | 13.6432 | 14.6154
Cr, Chromium (ppm) | 48.8317 | 44.451 | 53.2123 | 46.4076 | 51.2557
Cs, Caesium (ppm) | 11.1126 | 10.6126 | 11.6126 | 10.8164 | 11.4087
Cu, Copper (ppm) | 30.8513 | 29.2473 | 32.4553 | 29.8445 | 31.8581
Dy, Dysprosium (ppm) | 3.8949 | 3.6218 | 4.1679 | 3.7646 | 4.0251
Er, Erbium (ppm) | 1.4221 | 1.2564 | 1.5879 | 1.2964 | 1.5479
Eu, Europium (ppm) | 1.458 | 1.344 | 1.5719 | 1.3852 | 1.5308
Fe, Iron (wt.%) | 3.3103 | 3.2238 | 3.3968 | 3.2352 | 3.3854
Ga, Gallium (ppm) | 20.971 | 20.0193 | 21.9227 | 20.2296 | 21.7123
Gd, Gadolinium (ppm) | 6.1133 | 5.6841 | 6.5425 | 5.6795 | 6.5471
Hf, Hafnium (ppm) | 1.8787 | 1.742 | 2.0154 | 1.7816 | 1.9758
Ho, Holmium (ppm) | 0.62462 | 0.58321 | 0.66603 | 0.59209 | 0.65715
In, Indium (ppm) | 0.070619 | 0.063633 | 0.077606 | 0.06456 | 0.076679
K, Potassium (wt.%) | 2.937 | 2.8591 | 3.0149 | 2.8778 | 2.9962
La, Lanthanum (ppm) | 37.8146 | 35.4482 | 40.1811 | 36.4543 | 39.175
Li, Lithium (ppm) | 53.3174 | 51.5208 | 55.1139 | 52.0294 | 54.6054
Lu, Lutetium (ppm) | 0.16918 | 0.14973 | 0.18864 | IND | IND
Mg, Magnesium (wt.%) | 0.92527 | 0.8944 | 0.95614 | 0.90654 | 0.944
Mn, Manganese (wt.%) | 0.038493 | 0.037274 | 0.039712 | 0.03739 | 0.039596
Mo, Molybdenum (ppm) | 2.7066 | 2.4236 | 2.9896 | 2.5419 | 2.8712
Na, Sodium (wt.%) | 2.0214 | 1.9694 | 2.0734 | 1.9847 | 2.0581
Nb, Niobium (ppm) | 12.7362 | 12.0979 | 13.3745 | 12.1374 | 13.335
Nd, Neodymium (ppm) | 35.0719 | 32.581 | 37.5628 | 32.9144 | 37.2294
Ni, Nickel (ppm) | 19.7083 | 18.8606 | 20.556 | 19.1672 | 20.2494
P, Phosphorus (wt.%) | 0.083839 | 0.080698 | 0.08698 | 0.081633 | 0.086045
Pb, Lead (ppm) | 23.1096 | 21.9516 | 24.2676 | 22.1811 | 24.0381
Pr, Praseodymium (ppm) | 9.1557 | 8.546 | 9.7654 | 8.7629 | 9.5485
Rb, Rubidium (ppm) | 168.2898 | 160.6157 | 175.9639 | 163.5027 | 173.077
S, Sulphur (wt.%) | 0.33866 | 0.32239 | 0.35492 | 0.32866 | 0.34865
Sb, Antimony (ppm) | 0.53169 | 0.48849 | 0.5749 | 0.48346 | 0.57993
Sc, Scandium (ppm) | 8.9903 | 8.5877 | 9.3929 | 8.7134 | 9.2672
Sm, Samarium (ppm) | 7.1915 | 6.3857 | 7.9972 | 6.8637 | 7.5192
Sn, Tin (ppm) | 4.658 | 4.4184 | 4.8976 | 4.4399 | 4.8762
Sr, Strontium (ppm) | 159.5549 | 154.184 | 164.9258 | 155.922 | 163.1878
Ta, Tantalum (ppm) | 1.0869 | 1.0166 | 1.1573 | 1.0541 | 1.1198
Tb, Terbium (ppm) | 0.78657 | 0.71923 | 0.8539 | 0.75459 | 0.81855
Te, Tellurium (ppm) | 0.06555 | 0.038731 | 0.092369 | IND | IND
Th, Thorium (ppm) | 15.3383 | 14.3844 | 16.2922 | 14.7416 | 15.935
Ti, Titanium (wt.%) | 0.3717 | 0.3605 | 0.3829 | 0.36322 | 0.38019
Tl, Thallium (ppm) | 0.93987 | 0.89569 | 0.98404 | 0.91301 | 0.96673
Tm, Thulium (ppm) | 0.18925 | 0.17339 | 0.20512 | IND | IND
U, Uranium (ppm) | 4.0423 | 3.4836 | 4.601 | 3.737 | 4.3477
V, Vanadium (ppm) | 63.4261 | 60.1388 | 66.7134 | 61.8872 | 64.9649
W, Tungsten (ppm) | 17.4569 | 16.3724 | 18.5415 | 16.4459 | 18.468
Y, Yttrium (ppm) | 15.936 | 15.0499 | 16.822 | 15.4021 | 16.4698
Yb, Ytterbium (ppm) | 1.1544 | 1.0287 | 1.2802 | 1.0871 | 1.2218
Zn, Zinc (ppm) | 81.0972 | 78.6994 | 83.4949 | 79.0924 | 83.1019
Zr, Zirconium (ppm) | 57.5206 | 53.8307 | 61.2105 | 55.004 | 60.0372
Aqua Regia Digestion
Ag, Silver (ppm) | 0.084067 | 0.071854 | 0.09628 | 0.078553 | 0.089581
Al, Aluminium (wt.%) | 2.1454 | 2.0535 | 2.2373 | 2.0895 | 2.2013
As, Arsenic (ppm) | 275.5946 | 266.9018 | 284.2874 | 269.9147 | 281.2745
Ba, Barium (ppm) | 496.0441 | 475.6339 | 516.4543 | 483.1616 | 508.9266
Be, Beryllium (ppm) | 1.6116 | 1.5067 | 1.7166 | 1.554 | 1.6693
Bi, Bismuth (ppm) | 0.63448 | 0.4951 | 0.77386 | 0.53924 | 0.72972
Ca, Calcium (wt.%) | 0.65917 | 0.63092 | 0.68742 | 0.64172 | 0.67661
Cd, Cadmium (ppm) | 0.084838 | 0.068385 | 0.10129 | 0.067698 | 0.10198
Ce, Cerium (ppm) | 39.288 | 34.8782 | 43.6978 | 37.6813 | 40.8947
Co, Cobalt (ppm) | 13.7157 | 13.1352 | 14.2961 | 13.2725 | 14.1589
Cr, Chromium (ppm) | 52.5403 | 50.4147 | 54.6659 | 51.054 | 54.0267
Cs, Caesium (ppm) | 9.1703 | 8.7776 | 9.5629 | 8.8562 | 9.4843
Cu, Copper (ppm) | 30.9125 | 29.4859 | 32.3391 | 29.9336 | 31.8914
Dy, Dysprosium (ppm) | 2.801 | 2.4431 | 3.1589 | 2.6963 | 2.9057
Er, Erbium (ppm) | 1.0014 | 0.84891 | 1.1538 | 0.96938 | 1.0333
Eu, Europium (ppm) | 0.38276 | 0.28403 | 0.48148 | 0.35315 | 0.41236
Fe, Iron (wt.%) | 3.1919 | 3.0921 | 3.2917 | 3.1287 | 3.2552
Ga, Gallium (ppm) | 9.778 | 9.2533 | 10.3028 | 9.3824 | 10.1737
Gd, Gadolinium (ppm) | 3.9858 | 2.8724 | 5.0992 | 3.7757 | 4.1959
Ge, Germanium (ppm) | 0.11833 | 0.089416 | 0.14725 | IND | IND
Hf, Hafnium (ppm) | 0.33142 | 0.29732 | 0.36552 | 0.31 | 0.35285
Ho, Holmium (ppm) | 0.40796 | 0.36101 | 0.4549 | IND | IND
In, Indium (ppm) | 0.06493 | 0.056128 | 0.073732 | 0.058337 | 0.071524
K, Potassium (wt.%) | 0.99992 | 0.95771 | 1.0421 | 0.97721 | 1.0226
La, Lanthanum (ppm) | 18.2427 | 16.0843 | 20.4011 | 17.3729 | 19.1125
Li, Lithium (ppm) | 43.5649 | 41.9861 | 45.1437 | 42.5302 | 44.5996
Lu, Lutetium (ppm) | 0.094899 | 0.08449 | 0.10531 | IND | IND
Mg, Magnesium (wt.%) | 0.87388 | 0.84643 | 0.90132 | 0.85406 | 0.89369
Mn, Manganese (wt.%) | 0.033739 | 0.032668 | 0.03481 | 0.033081 | 0.034397
Mo, Molybdenum (ppm) | 2.4542 | 2.2389 | 2.6695 | 2.3385 | 2.5699
Na, Sodium (wt.%) | 0.15554 | 0.1424 | 0.16868 | 0.14772 | 0.16335
Nb, Niobium (ppm) | 0.78801 | 0.596 | 0.98002 | 0.70347 | 0.87254
Ni, Nickel (ppm) | 19.2528 | 18.0564 | 20.4493 | 18.6379 | 19.8678
P, Phosphorus (wt.%) | 0.064556 | 0.062645 | 0.066466 | 0.063116 | 0.065996
Pb, Lead (ppm) | 4.8711 | 4.5287 | 5.2135 | 4.658 | 5.0842
Rb, Rubidium (ppm) | 105.4634 | 100.3227 | 110.6041 | 101.7861 | 109.1406
S, Sulphur (wt.%) | 0.33376 | 0.32293 | 0.3446 | 0.32641 | 0.34112
Sb, Antimony (ppm) | 0.27087 | 0.23682 | 0.30493 | 0.25214 | 0.28961
Sc, Scandium (ppm) | 8.1112 | 7.7098 | 8.5125 | 7.8571 | 8.3653
Se, Selenium (ppm) | 0.47633 | 0.24307 | 0.70959 | IND | IND
Sn, Tin (ppm) | 3.4693 | 3.2953 | 3.6434 | 3.3349 | 3.6038
Sr, Strontium (ppm) | 19.6321 | 18.329 | 20.9352 | 18.9398 | 20.3244
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.56752 | 0.50523 | 0.6298 | 0.53236 | 0.60268
Te, Tellurium (ppm) | 0.055541 | 0.033516 | 0.077566 | IND | IND
Th, Thorium (ppm) | 8.0075 | 7.0295 | 8.9855 | 7.7399 | 8.2751
Ti, Titanium (wt.%) | 0.27158 | 0.25744 | 0.28573 | 0.26354 | 0.27963
Tl, Thallium (ppm) | 0.66473 | 0.62824 | 0.70123 | 0.63979 | 0.68968
U, Uranium (ppm) | 3.4524 | 3.0426 | 3.8622 | 3.1824 | 3.7224
V, Vanadium (ppm) | 60.7756 | 58.8325 | 62.7188 | 59.3644 | 62.1868
W, Tungsten (ppm) | 11.2239 | 10.3774 | 12.0705 | 10.7146 | 11.7333
Y, Yttrium (ppm) | 10.895 | 10.3553 | 11.4348 | 10.5074 | 11.2827
Yb, Ytterbium (ppm) | 0.7142 | 0.68929 | 0.73911 | 0.67441 | 0.75398
Zn, Zinc (ppm) | 76.7123 | 74.2716 | 79.153 | 75.0051 | 78.4195
Zr, Zirconium (ppm) | 8.1134 | 7.6056 | 8.6211 | 7.7171 | 8.5096
Infrared Combustion
S, Sulphur (wt.%) | 0.33355 | 0.32034 | 0.34677 | 0.32253 | 0.34458
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 285
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 10.2835 | Hg | ppm | 0.13333 | Se | ppm | 0.95708
Ge | ppm | 0.27114 | Re | ppm | 0.0015238
Aqua Regia Digestion
B | ppm | 10.2263 | Pr | ppm | 5.4071 | Sm | ppm | 4.7353
Hg | ppm | 0.014367 | Pt | ppb | < 5 | Tm | ppm | 0.11563
Nd | ppm | 19.3271 | Re | ppm | 0.00117
Pd | ppb | 20.1667 | Si | wt.% | 0.13298
Infrared Combustion
C | wt.% | 0.18417
Borate Fusion XRF
Al2O3 | wt.% | 15.07 | Fe | wt.% | 3.365 | S | wt.% | 0.3405
As | ppm | 270 | K2O | wt.% | 3.505 | SiO2 | wt.% | 66.96
Ba | ppm | 1020 | MgO | wt.% | 1.605 | Sn | ppm | < 10
CaO | wt.% | 2.73 | MnO | wt.% | 0.052 | Sr | ppm | 170
Cl | ppm | 25 | Na2O | wt.% | 2.735 | TiO2 | wt.% | 0.634
Co | ppm | 20 | Ni | ppm | 25 | V | ppm | 75
Cr | ppm | 75 | P | wt.% | 0.086 | Zn | ppm | 80
Cu | ppm | 45 | Pb | ppm | 35 | Zr | ppm | 260
Thermogravimetry
LOI1000 | wt.% | 1.42
Laser Ablation ICP-MS
Ag | ppm | < 0.1 | Hf | ppm | 6.99 | Sm | ppm | 7.775
As | ppm | 263.5 | Ho | ppm | 1.15 | Sn | ppm | 4.9
Ba | ppm | 997 | In | ppm | 0.05 | Sr | ppm | 155
Be | ppm | 2.9 | La | ppm | 39.4 | Ta | ppm | 1.12
Bi | ppm | 0.58 | Lu | ppm | 0.4 | Tb | ppm | 1.1
Cd | ppm | 0.15 | Mn | wt.% | 0.0388 | Te | ppm | < 0.2
Ce | ppm | 78.1 | Mo | ppm | 2.7 | Th | ppm | 15.2
Co | ppm | 14.1 | Nb | ppm | 12.7 | Ti | wt.% | 0.373
Cr | ppm | 57.5 | Nd | ppm | 37 | Tl | ppm | 0.7
Cs | ppm | 10.5 | Ni | ppm | 20 | Tm | ppm | 0.435
Cu | ppm | 33 | Pb | ppm | 23 | U | ppm | 4.73
Dy | ppm | 5.905 | Pr | ppm | 9.73 | V | ppm | 66.3
Er | ppm | 3.16 | Rb | ppm | 160 | W | ppm | 19
Eu | ppm | 1.44 | Re | ppm | < 0.01 | Y | ppm | 30.3
Ga | ppm | 19.9 | Sb | ppm | 0.75 | Yb | ppm | 2.805
Gd | ppm | 7.04 | Sc | ppm | 8.95 | Zn | ppm | 75
Ge | ppm | 1.35 | Se | ppm | < 5 | Zr | ppm | 249
3-Acid Digestion (no HF)
Ag | ppm | 0.24167 | Gd | ppm | 5.1167 | S | wt.% | 0.376
Al2O3 | wt.% | 15.9167 | Hf | ppm | 2.5833 | Sc | ppm | 9.4667
Ba | ppm | 1063.3333 | Ho | ppm | 0.53833 | Sm | ppm | 7.2167
Be | ppm | 3.25 | K2O | wt.% | 3.705 | Sn | ppm | 4.61
Bi | ppm | 0.73333 | La | ppm | 39.6667 | Sr | ppm | 164
CaO | wt.% | 2.7617 | Li | ppm | 53 | Ta | ppm | 0.935
Cd | ppm | 0.26667 | MgO | wt.% | 1.6667 | Tb | ppm | 0.69167
Ce | ppm | 80 | MnO | wt.% | 0.050333 | Th | ppm | 17.1833
Co | ppm | 16.1667 | Mo | ppm | 3.275 | TiO2 | wt.% | 0.6625
Cr | ppm | 45.6667 | Na2O | wt.% | 2.81 | U | ppm | 3.706
Cs | ppm | 10.1167 | Nb | ppm | 14.8833 | V | ppm | 68.35
Cu | ppm | 32.7833 | Nd | ppm | 37.2167 | W | ppm | 20.8
Dy | ppm | 3.78 | Ni | ppm | 26.15 | Y | ppm | 15.7167
Er | ppm | 1.49 | P2O5 | wt.% | 0.20317 | Yb | ppm | 1.0383
Eu | ppm | 1.6667 | Pb | ppm | 24.9167 | Zn | ppm | 91.6667
Fe2O3 | wt.% | 4.9367 | Pr | ppm | 9.5167 | Zr | ppm | 58.3333
Ga | ppm | 20.9167 | Rb | ppm | 215.6667
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 285
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.6561 | 0.021777 | 0.61255 | 0.69966 | 0.59077 | 0.72143 | 3.319% | 6.638% | 9.957% | 0.6233 | 0.68891
PhotonAssay
Au, ppm | 0.66944 | 0.036382 | 0.59668 | 0.74221 | 0.5603 | 0.77859 | 5.435% | 10.87% | 16.3% | 0.63597 | 0.70292
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.64066 | 0.03245 | 0.57576 | 0.70556 | 0.54331 | 0.73801 | 5.065% | 10.13% | 15.2% | 0.60863 | 0.67269
Cyanide Leach
Au, ppm | 0.63488 | 0.04271 | 0.54946 | 0.7203 | 0.50675 | 0.76301 | 6.727% | 13.45% | 20.18% | 0.60313 | 0.66662
4-Acid Digestion
Ag, ppm | 0.090636 | 0.016117 | 0.058402 | 0.12287 | 0.042285 | 0.13899 | 17.78% | 35.56% | 53.35% | 0.086105 | 0.095168
Al, wt.% | 7.759 | 0.34331 | 7.0724 | 8.4457 | 6.7291 | 8.789 | 4.425% | 8.849% | 13.27% | 7.3711 | 8.147
As, ppm | 288.5915 | 10.9127 | 266.7661 | 310.417 | 255.8534 | 321.3297 | 3.781% | 7.563% | 11.34% | 274.162 | 303.0211
Ba, ppm | 1031.6368 | 46.4735 | 938.6898 | 1124.5838 | 892.2163 | 1171.0573 | 4.505% | 9.01% | 13.51% | 980.0549 | 1083.2186
Be, ppm | 2.5058 | 0.13988 | 2.2261 | 2.7856 | 2.0862 | 2.9255 | 5.582% | 11.16% | 16.75% | 2.3805 | 2.6311
Bi, ppm | 0.586 | 0.069039 | 0.44792 | 0.72408 | 0.37889 | 0.79312 | 11.78% | 23.56% | 35.34% | 0.5567 | 0.6153
Ca, wt.% | 1.939 | 0.083659 | 1.7717 | 2.1064 | 1.6881 | 2.19 | 4.314% | 8.629% | 12.94% | 1.8421 | 2.036
Cd, ppm | 0.20164 | 0.016826 | 0.16799 | 0.23529 | 0.15116 | 0.25212 | 8.345% | 16.69% | 25.03% | 0.19156 | 0.21172
Ce, ppm | 78.0157 | 4.3721 | 69.2715 | 86.76 | 64.8994 | 91.1321 | 5.604% | 11.21% | 16.81% | 74.115 | 81.9165
Co, ppm | 14.1293 | 0.73668 | 12.6559 | 15.6027 | 11.9193 | 16.3394 | 5.214% | 10.43% | 15.64% | 13.4228 | 14.8358
Cr, ppm | 48.8317 | 3.6547 | 41.5222 | 56.1411 | 37.8675 | 59.7958 | 7.484% | 14.97% | 22.45% | 46.3901 | 51.2733
Cs, ppm | 11.1126 | 0.44102 | 10.2305 | 11.9946 | 9.7895 | 12.4356 | 3.969% | 7.937% | 11.91% | 10.557 | 11.6682
Cu, ppm | 30.8513 | 1.5449 | 27.7616 | 33.941 | 26.2167 | 35.4859 | 5.007% | 10.01% | 15.02% | 29.3087 | 32.3939
Dy, ppm | 3.8949 | 0.20583 | 3.4832 | 4.3065 | 3.2774 | 4.5123 | 5.285% | 10.57% | 15.85% | 3.7001 | 4.0896
Er, ppm | 1.4221 | 0.076927 | 1.2683 | 1.576 | 1.1913 | 1.6529 | 5.409% | 10.82% | 16.23% | 1.351 | 1.4932
Eu, ppm | 1.458 | 0.057276 | 1.3434 | 1.5725 | 1.2862 | 1.6298 | 3.928% | 7.857% | 11.79% | 1.3851 | 1.5309
Fe, wt.% | 3.3103 | 0.11321 | 3.0839 | 3.5367 | 2.9707 | 3.6499 | 3.42% | 6.84% | 10.26% | 3.1448 | 3.4758
Ga, ppm | 20.971 | 1.3971 | 18.1767 | 23.7653 | 16.7795 | 25.1624 | 6.662% | 13.32% | 19.99% | 19.9224 | 22.0195
Gd, ppm | 6.1133 | 0.2379 | 5.6375 | 6.5891 | 5.3996 | 6.827 | 3.891% | 7.783% | 11.67% | 5.8077 | 6.419
Hf, ppm | 1.8787 | 0.13258 | 1.6136 | 2.1439 | 1.481 | 2.2765 | 7.057% | 14.11% | 21.17% | 1.7848 | 1.9726
Ho, ppm | 0.62462 | 0.039644 | 0.54533 | 0.70391 | 0.50569 | 0.74355 | 6.347% | 12.69% | 19.04% | 0.59339 | 0.65585
In, ppm | 0.070619 | 0.0048873 | 0.060845 | 0.080394 | 0.055958 | 0.085281 | 6.921% | 13.84% | 20.76% | 0.067088 | 0.07415
K, wt.% | 2.937 | 0.098929 | 2.7391 | 3.1348 | 2.6402 | 3.2338 | 3.368% | 6.737% | 10.11% | 2.7901 | 3.0838
La, ppm | 37.8146 | 2.0727 | 33.6693 | 41.96 | 31.5967 | 44.0326 | 5.481% | 10.96% | 16.44% | 35.9239 | 39.7054
Li, ppm | 53.3174 | 2.5096 | 48.2982 | 58.3366 | 45.7886 | 60.8461 | 4.707% | 9.414% | 14.12% | 50.6515 | 55.9833
Lu, ppm | 0.16918 | 0.013666 | 0.14185 | 0.19651 | 0.12818 | 0.21018 | 8.078% | 16.16% | 24.23% | 0.16072 | 0.17764
Mg, wt.% | 0.92527 | 0.040622 | 0.84402 | 1.0065 | 0.8034 | 1.0471 | 4.39% | 8.781% | 13.17% | 0.879 | 0.97153
Mn, wt.% | 0.038493 | 0.0014554 | 0.035582 | 0.041404 | 0.034127 | 0.042859 | 3.781% | 7.562% | 11.34% | 0.036568 | 0.040418
Mo, ppm | 2.7066 | 0.186 | 2.3346 | 3.0786 | 2.1486 | 3.2646 | 6.872% | 13.74% | 20.62% | 2.5712 | 2.8419
Na, wt.% | 2.0214 | 0.065298 | 1.8908 | 2.152 | 1.8255 | 2.2173 | 3.23% | 6.461% | 9.691% | 1.9203 | 2.1225
Nb, ppm | 12.7362 | 0.86565 | 11.0049 | 14.4675 | 10.1393 | 15.3331 | 6.797% | 13.59% | 20.39% | 12.0994 | 13.373
Nd, ppm | 35.0719 | 1.7815 | 31.5088 | 38.6349 | 29.7273 | 40.4164 | 5.08% | 10.16% | 15.24% | 33.3183 | 36.8255
Ni, ppm | 19.7083 | 0.82256 | 18.0632 | 21.3534 | 17.2406 | 22.176 | 4.174% | 8.347% | 12.52% | 18.7229 | 20.6937
P, wt.% | 0.083839 | 0.0035994 | 0.07664 | 0.091038 | 0.073041 | 0.094637 | 4.293% | 8.586% | 12.88% | 0.079647 | 0.088031
Pb, ppm | 23.1096 | 1.4973 | 20.1151 | 26.1041 | 18.6178 | 27.6014 | 6.479% | 12.96% | 19.44% | 21.9541 | 24.2651
Pr, ppm | 9.1557 | 0.52672 | 8.1022 | 10.2091 | 7.5755 | 10.7359 | 5.753% | 11.51% | 17.26% | 8.6979 | 9.6135
Rb, ppm | 168.2898 | 8.2239 | 151.842 | 184.7377 | 143.618 | 192.9616 | 4.887% | 9.774% | 14.66% | 159.8753 | 176.7043
S, wt.% | 0.33866 | 0.020173 | 0.29831 | 0.379 | 0.27814 | 0.39917 | 5.957% | 11.91% | 17.87% | 0.32172 | 0.35559
Sb, ppm | 0.53169 | 0.030187 | 0.47132 | 0.59207 | 0.44113 | 0.62225 | 5.677% | 11.35% | 17.03% | 0.50511 | 0.55828
Sc, ppm | 8.9903 | 0.41363 | 8.163 | 9.8176 | 7.7494 | 10.2312 | 4.601% | 9.202% | 13.8% | 8.5408 | 9.4398
Sm, ppm | 7.1915 | 0.4037 | 6.3841 | 7.9989 | 5.9804 | 8.4026 | 5.614% | 11.23% | 16.84% | 6.8319 | 7.551
Sn, ppm | 4.658 | 0.21326 | 4.2315 | 5.0845 | 4.0182 | 5.2978 | 4.578% | 9.157% | 13.73% | 4.4251 | 4.8909
Sr, ppm | 159.5549 | 6.3146 | 146.9256 | 172.1842 | 140.611 | 178.4988 | 3.958% | 7.915% | 11.87% | 151.5772 | 167.5327
Ta, ppm | 1.0869 | 0.095268 | 0.8964 | 1.2775 | 0.80113 | 1.3727 | 8.765% | 17.53% | 26.29% | 1.0326 | 1.1413
Tb, ppm | 0.78657 | 0.059383 | 0.6678 | 0.90533 | 0.60842 | 0.96472 | 7.55% | 15.1% | 22.65% | 0.74724 | 0.8259
Te, ppm | 0.06555 | 0.018862 | 0.027826 | 0.10327 | 0.0089634 | 0.12214 | 28.78% | 57.55% | 86.33% | 0.062272 | 0.068827
Th, ppm | 15.3383 | 0.67178 | 13.9947 | 16.6819 | 13.323 | 17.3536 | 4.38% | 8.76% | 13.14% | 14.5714 | 16.1052
Ti, wt.% | 0.3717 | 0.014924 | 0.34186 | 0.40155 | 0.32693 | 0.41647 | 4.015% | 8.03% | 12.04% | 0.35312 | 0.39029
Tl, ppm | 0.93987 | 0.045001 | 0.84987 | 1.0299 | 0.80486 | 1.0749 | 4.788% | 9.576% | 14.36% | 0.89287 | 0.98686
Tm, ppm | 0.18925 | 0.0098273 | 0.1696 | 0.20891 | 0.15977 | 0.21874 | 5.193% | 10.39% | 15.58% | 0.17979 | 0.19872
U, ppm | 4.0423 | 0.43915 | 3.164 | 4.9206 | 2.7249 | 5.3598 | 10.86% | 21.73% | 32.59% | 3.8402 | 4.2444
V, ppm | 63.4261 | 2.8775 | 57.6712 | 69.181 | 54.7937 | 72.0585 | 4.537% | 9.073% | 13.61% | 60.2548 | 66.5974
W, ppm | 17.4569 | 1.3091 | 14.8387 | 20.0752 | 13.5295 | 21.3843 | 7.499% | 15% | 22.5% | 16.5841 | 18.3298
Y, ppm | 15.936 | 1.1216 | 13.6928 | 18.1791 | 12.5713 | 19.3006 | 7.038% | 14.08% | 21.11% | 15.1392 | 16.7328
Yb, ppm | 1.1544 | 0.10524 | 0.94396 | 1.3649 | 0.83871 | 1.4702 | 9.116% | 18.23% | 27.35% | 1.0967 | 1.2122
Zn, ppm | 81.0972 | 2.9987 | 75.0998 | 87.0945 | 72.1012 | 90.0931 | 3.698% | 7.395% | 11.09% | 77.0423 | 85.152
Zr, ppm | 57.5206 | 5.3565 | 46.8077 | 68.2336 | 41.4512 | 73.59 | 9.312% | 18.62% | 27.94% | 54.6446 | 60.3967
Aqua Regia Digestion
Ag, ppm | 0.084067 | 0.0083997 | 0.067267 | 0.10087 | 0.058868 | 0.10927 | 9.992% | 19.98% | 29.98% | 0.079864 | 0.08827
Al, wt.% | 2.1454 | 0.16852 | 1.8083 | 2.4824 | 1.6398 | 2.6509 | 7.855% | 15.71% | 23.57% | 2.0381 | 2.2526
As, ppm | 275.5946 | 10.8914 | 253.8118 | 297.3774 | 242.9204 | 308.2688 | 3.952% | 7.904% | 11.86% | 261.8149 | 289.3743
Ba, ppm | 496.0441 | 33.961 | 428.122 | 563.9662 | 394.161 | 597.9272 | 6.846% | 13.69% | 20.54% | 471.2419 | 520.8463
Be, ppm | 1.6116 | 0.13418 | 1.3433 | 1.88 | 1.2091 | 2.0142 | 8.326% | 16.65% | 24.98% | 1.5311 | 1.6922
Bi, ppm | 0.63448 | 0.067239 | 0.5 | 0.76896 | 0.43276 | 0.8362 | 10.6% | 21.2% | 31.79% | 0.60275 | 0.6662
Ca, wt.% | 0.65917 | 0.054454 | 0.55026 | 0.76808 | 0.4958 | 0.82253 | 8.261% | 16.52% | 24.78% | 0.62621 | 0.69213
Cd, ppm | 0.084838 | 0.012118 | 0.060602 | 0.10907 | 0.048484 | 0.12119 | 14.28% | 28.57% | 42.85% | 0.080596 | 0.08908
Ce, ppm | 39.288 | 5.6016 | 28.0849 | 50.4912 | 22.4833 | 56.0927 | 14.26% | 28.52% | 42.77% | 37.3236 | 41.2524
Co, ppm | 13.7157 | 0.77494 | 12.1658 | 15.2655 | 11.3909 | 16.0405 | 5.65% | 11.3% | 16.95% | 13.0299 | 14.4014
Cr, ppm | 52.5403 | 2.3752 | 47.79 | 57.2907 | 45.4149 | 59.6658 | 4.521% | 9.041% | 13.56% | 49.9133 | 55.1674
Cs, ppm | 9.1703 | 0.39986 | 8.3705 | 9.97 | 7.9707 | 10.3698 | 4.36% | 8.721% | 13.08% | 8.7117 | 9.6288
Cu, ppm | 30.9125 | 1.8959 | 27.1207 | 34.7043 | 25.2248 | 36.6001 | 6.133% | 12.27% | 18.4% | 29.3669 | 32.4581
Dy, ppm | 2.801 | 0.24193 | 2.3171 | 3.2848 | 2.0752 | 3.5268 | 8.637% | 17.27% | 25.91% | 2.6609 | 2.941
Er, ppm | 1.0014 | 0.10426 | 0.79284 | 1.2099 | 0.68858 | 1.3141 | 10.41% | 20.82% | 31.24% | 0.95129 | 1.0514
Eu, ppm | 0.38276 | 0.068323 | 0.24611 | 0.5194 | 0.17779 | 0.58772 | 17.85% | 35.7% | 53.55% | 0.36362 | 0.40189
Fe, wt.% | 3.1919 | 0.14553 | 2.9008 | 3.483 | 2.7553 | 3.6285 | 4.559% | 9.119% | 13.68% | 3.0323 | 3.3515
Ga, ppm | 9.778 | 0.74609 | 8.2859 | 11.2702 | 7.5398 | 12.0163 | 7.63% | 15.26% | 22.89% | 9.2891 | 10.2669
Gd, ppm | 3.9858 | 0.80991 | 2.366 | 5.6056 | 1.5561 | 6.4155 | 20.32% | 40.64% | 60.96% | 3.7865 | 4.1851
Ge, ppm | 0.11833 | 0.021844 | 0.074646 | 0.16202 | 0.052803 | 0.18386 | 18.46% | 36.92% | 55.38% | 0.11242 | 0.12425
Hf, ppm | 0.33142 | 0.041875 | 0.24767 | 0.41517 | 0.2058 | 0.45705 | 12.63% | 25.27% | 37.9% | 0.31485 | 0.34799
Ho, ppm | 0.40796 | 0.036134 | 0.33569 | 0.48022 | 0.29955 | 0.51636 | 8.857% | 17.71% | 26.57% | 0.38756 | 0.42835
In, ppm | 0.06493 | 0.0051606 | 0.054609 | 0.075251 | 0.049448 | 0.080412 | 7.948% | 15.9% | 23.84% | 0.061684 | 0.068177
K, wt.% | 0.99992 | 0.075405 | 0.84911 | 1.1507 | 0.7737 | 1.2261 | 7.541% | 15.08% | 22.62% | 0.94992 | 1.0499
La, ppm | 18.2427 | 3.3753 | 11.4922 | 24.9932 | 8.1169 | 28.3685 | 18.5% | 37% | 55.51% | 17.3306 | 19.1548
Li, ppm | 43.5649 | 2.4133 | 38.7383 | 48.3915 | 36.3249 | 50.8049 | 5.54% | 11.08% | 16.62% | 41.3867 | 45.7431
Lu, ppm | 0.094899 | 0.0059376 | 0.083023 | 0.10677 | 0.077086 | 0.11271 | 6.257% | 12.51% | 18.77% | 0.090154 | 0.099643
Mg, wt.% | 0.87388 | 0.030233 | 0.81341 | 0.93434 | 0.78317 | 0.96458 | 3.46% | 6.919% | 10.38% | 0.83018 | 0.91757
Mn, wt.% | 0.033739 | 0.001272 | 0.031195 | 0.036283 | 0.029923 | 0.037555 | 3.77% | 7.54% | 11.31% | 0.032052 | 0.035426
Mo, ppm | 2.4542 | 0.31353 | 1.8272 | 3.0813 | 1.5136 | 3.3948 | 12.78% | 25.55% | 38.33% | 2.3315 | 2.5769
Na, wt.% | 0.15554 | 0.021256 | 0.11302 | 0.19805 | 0.091769 | 0.21931 | 13.67% | 27.33% | 41% | 0.14776 | 0.16331
Nb, ppm | 0.78801 | 0.2555 | 0.277 | 1.299 | 0.0215 | 1.5545 | 32.42% | 64.85% | 97.27% | 0.74861 | 0.82741
Ni, ppm | 19.2528 | 1.008 | 17.2368 | 21.2688 | 16.2288 | 22.2768 | 5.236% | 10.47% | 15.71% | 18.2902 | 20.2155
P, wt.% | 0.064556 | 0.0025655 | 0.059425 | 0.069687 | 0.056859 | 0.072252 | 3.974% | 7.948% | 11.92% | 0.061328 | 0.067784
Pb, ppm | 4.8711 | 0.51866 | 3.8338 | 5.9084 | 3.3151 | 6.4271 | 10.65% | 21.3% | 31.94% | 4.6276 | 5.1147
Rb, ppm | 105.4634 | 7.4605 | 90.5424 | 120.3843 | 83.082 | 127.8448 | 7.074% | 14.15% | 21.22% | 100.1902 | 110.7366
S, wt.% | 0.33376 | 0.014669 | 0.30443 | 0.3631 | 0.28976 | 0.37777 | 4.395% | 8.79% | 13.18% | 0.31707 | 0.35045
Sb, ppm | 0.27087 | 0.050829 | 0.16922 | 0.37253 | 0.11839 | 0.42336 | 18.76% | 37.53% | 56.29% | 0.25733 | 0.28442
Sc, ppm | 8.1112 | 0.25477 | 7.6016 | 8.6207 | 7.3469 | 8.8755 | 3.141% | 6.282% | 9.423% | 7.7056 | 8.5167
Se, ppm | 0.47633 | 0.13121 | 0.21392 | 0.73875 | 0.082708 | 0.86996 | 27.55% | 55.09% | 82.64% | 0.45252 | 0.50015
Sn, ppm | 3.4693 | 0.16147 | 3.1464 | 3.7923 | 2.9849 | 3.9538 | 4.654% | 9.309% | 13.96% | 3.2959 | 3.6428
Sr, ppm | 19.6321 | 2.0907 | 15.4507 | 23.8136 | 13.3599 | 25.9043 | 10.65% | 21.3% | 31.95% | 18.6505 | 20.6137
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.56752 | 0.049627 | 0.46827 | 0.66677 | 0.41864 | 0.7164 | 8.745% | 17.49% | 26.23% | 0.53914 | 0.59589
Te, ppm | 0.055541 | 0.012519 | 0.030502 | 0.08058 | 0.017983 | 0.093099 | 22.54% | 45.08% | 67.62% | 0.052764 | 0.058318
Th, ppm | 8.0075 | 1.4512 | 5.1051 | 10.9099 | 3.654 | 12.361 | 18.12% | 36.25% | 54.37% | 7.6071 | 8.4079
Ti, wt.% | 0.27158 | 0.022579 | 0.22643 | 0.31674 | 0.20385 | 0.33932 | 8.314% | 16.63% | 24.94% | 0.25801 | 0.28516
Tl, ppm | 0.66473 | 0.031962 | 0.60081 | 0.72866 | 0.56885 | 0.76062 | 4.808% | 9.617% | 14.42% | 0.6315 | 0.69797
U, ppm | 3.4524 | 0.34166 | 2.7691 | 4.1357 | 2.4274 | 4.4774 | 9.896% | 19.79% | 29.69% | 3.2798 | 3.625
V, ppm | 60.7756 | 2.4325 | 55.9106 | 65.6406 | 53.4781 | 68.0731 | 4.002% | 8.005% | 12.01% | 57.7368 | 63.8144
W, ppm | 11.2239 | 1.3207 | 8.5825 | 13.8653 | 7.2618 | 15.186 | 11.77% | 23.53% | 35.3% | 10.6627 | 11.7851
Y, ppm | 10.895 | 0.73818 | 9.4187 | 12.3714 | 8.6805 | 13.1096 | 6.775% | 13.55% | 20.33% | 10.3503 | 11.4398
Yb, ppm | 0.7142 | 0.023419 | 0.66736 | 0.76104 | 0.64394 | 0.78445 | 3.279% | 6.558% | 9.837% | 0.67849 | 0.74991
Zn, ppm | 76.7123 | 3.8784 | 68.9555 | 84.4691 | 65.0771 | 88.3475 | 5.056% | 10.11% | 15.17% | 72.8767 | 80.5479
Zr, ppm | 8.1134 | 0.77123 | 6.5709 | 9.6558 | 5.7997 | 10.427 | 9.506% | 19.01% | 28.52% | 7.7077 | 8.519
Infrared Combustion
S, wt.% | 0.33355 | 0.015602 | 0.30235 | 0.36476 | 0.28675 | 0.38036 | 4.677% | 9.355% | 14.03% | 0.31688 | 0.35023
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 285
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
I.S. | Insufficient Sample (Lab 1.32)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 285
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Mineral Solutions (BVMS), Al Wadi District, Jeddah, Saudi Arabia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Minerals Ltd, Bibiani, Western North Region, Ghana
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Ghana Ltd, Obuasi, Ashanti, Ghana
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 285 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 13 | 7 | 261064 | 0.086104 | 0.68719 | 0.6725
2 | 1 | ANSLu | 1 | 3 | 2 | 261065 | 0.085014 | 0.65953 | 0.67103
2 | 1 | ANSLu | 1 | 5 | 3 | 261066 | 0.082854 | 0.64527 | 0.67027
2 | 1 | ANSLu | 1 | 2 | 1 | 261067 | 0.08584 | 0.64857 | 0.67044
2 | 1 | ANSLu | 1 | 14 | 7 | 261068 | 0.08696 | 0.68631 | 0.67245
2 | 1 | ANSLu | 1 | 17 | 9 | 261069 | 0.084514 | 0.71207 | 0.67383
2 | 1 | ANSLu | 1 | 6 | 3 | 261070 | 0.083433 | 0.73357 | 0.67497
2 | 1 | ANSLu | 1 | 1 | 1 | 261071 | 0.084593 | 0.68244 | 0.67225
2 | 1 | ANSLu | 1 | 18 | 9 | 261072 | 0.084559 | 0.68384 | 0.67232
2 | 1 | ANSLu | 1 | 19 | 10 | 261073 | 0.085105 | 0.67222 | 0.6717
2 | 1 | ANSLu | 1 | 12 | 6 | 261074 | 0.084155 | 0.61329 | 0.66856
2 | 1 | ANSLu | 1 | 15 | 8 | 261075 | 0.087556 | 0.67162 | 0.67167
2 | 1 | ANSLu | 1 | 10 | 5 | 261076 | 0.085479 | 0.67782 | 0.672
2 | 1 | ANSLu | 1 | 7 | 4 | 261077 | 0.086981 | 0.68476 | 0.67237
2 | 1 | ANSLu | 1 | 11 | 6 | 261078 | 0.08709 | 0.64951 | 0.67049
2 | 1 | ANSLu | 1 | 8 | 4 | 261079 | 0.083441 | 0.66599 | 0.67137
2 | 1 | ANSLu | 1 | 16 | 8 | 261080 | 0.085463 | 0.6518 | 0.67061
2 | 1 | ANSLu | 1 | 20 | 10 | 261081 | 0.085349 | 0.66924 | 0.67154
2 | 1 | ANSLu | 1 | 4 | 2 | 261082 | 0.083095 | 0.66035 | 0.67107
2 | 1 | ANSLu | 1 | 9 | 5 | 261083 | 0.08502 | 0.67807 | 0.67201
Mean | [6 empty] | 0.08513 | 0.67167 | 0.67167
Median | [7 empty] | 0.67192 | 0.67169
SD | [7 empty] | 0.025413 | 0.0013537
RSD | [7 empty] | 3.783% | 0.2015%
RSD Upscale Factor | [7 empty] | 0.05327
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 285 (Certified Value 0.656 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 50g | 30g | 30g | 25g | 30g | 50g | 50g | 30g | 50g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 50g | 25g | 15g | 30g | 40g | 30g | 30g | 3
| 1 | 1 | 0.68244 | 0.74 | 0.6633 | 0.67735 | 0.686 | 0.65776 | 0.726 | 0.64 | 0.684 | 0.63 | 0.66 | 0.665 | 0.657 | 0.663 | 0.637 | 0.656 | 0.689 | 0.675 | 0.66825 | 0.637 | 0.651 | 0.663 | 0.55 | 0.682 | 0.583 | 0.65 | 0.635 | 0.658 | 0.643 | 0.7 | 0.64 | 0.64 | 1
| 1 | 2 | 0.64857 | 0.64 | 0.6633 | 0.67642 | 0.635 | 0.65754 | 0.726 | 0.64 | 0.668 | 0.63 | 0.65 | 0.68 | 0.663 | 0.659 | 0.634 | 0.638 | 0.67 | 0.688 | 0.67618 | 0.661 | 0.647 | 0.663 | 0.59 | 0.683 | 0.573 | 0.659 | 0.626 | 0.638 | 0.648 | 0.7 | 0.625 | 0.65 | #N/A
| 1 | 3 | 0.65953 | 0.66 | 0.6534 | 0.67555 | 0.658 | 0.61835 | 0.747 | 0.64 | 0.671 | 0.63 | 0.65 | 0.68 | 0.677 | 0.663 | 0.642 | 0.653 | 0.661 | 0.673 | 0.66825 | 0.633 | 0.659 | 0.663 | 0.61 | 0.696 | 0.638 | 0.642 | 0.647 | 0.639 | 0.612 | 0.72 | 0.604 | 0.65 | 16
| 1 | 4 | 0.66035 | 0.65 | 0.6435 | 0.66618 | 0.662 | 0.64775 | 0.737 | 0.62 | 0.662 | 0.59 | 0.66 | 0.674 | 0.665 | 0.674 | 0.637 | 0.655 | 0.663 | 0.668 | 0.6705 | 0.646 | 0.667 | 0.659 | 0.64 | 0.681 | 0.597 | 0.671 | 0.654 | 0.626 | 0.571 | 0.7 | 0.621 | 0.65 | 0.6561
| 1 | 5 | 0.64527 | 0.65 | 0.663 | 0.66932 | 0.686 | 0.64753 | 0.73 | 0.64 | 0.685 | 0.62 | 0.67 | 0.653 | 0.656 | 0.671 | 0.631 | 0.651 | 0.68 | 0.667 | 0.67419 | 0.666 | 0.672 | 0.663 | 0.62 | 0.7 | 0.609 | 0.647 | 0.645 | 0.638 | 0.606 | 0.71 | 0.627 | 0.66 | 7
| 1 | 6 | 0.73357 | 0.64 | 0.6629 | 0.67763 | 0.671 | 0.63207 | 0.739 | 0.64 | 0.675 | 0.63 | 0.66 | 0.702 | 0.664 | 0.667 | 0.621 | 0.65 | 0.671 | 0.663 | 0.67819 | 0.661 | 0.667 | 0.665 | 0.64 | 0.67 | 0.604 | 0.656 | 0.638 | 0.652 | 0.638 | 0.72 | 0.649 | 0.66
|  | 7 | 0.68476
|  | 8 | 0.66599
|  | 9 | 0.67807
|  | 10 | 0.67782
|  | 11 | 0.64951
|  | 12 | 0.61329
|  | 13 | 0.68719
|  | 14 | 0.68631
|  | 15 | 0.67162
|  | 16 | 0.6518
|  | 17 | 0.71207
|  | 18 | 0.68384
|  | 19 | 0.67222
|  | 20 | 0.66924
| Mean |  | 0.67167 | 0.66333 | 0.65823 | 0.67374 | 0.66633 | 0.6435 | 0.73417 | 0.63667 | 0.67417 | 0.62167 | 0.65833 | 0.67567 | 0.66367 | 0.66617 | 0.63367 | 0.6505 | 0.67233 | 0.67233 | 0.67259 | 0.65067 | 0.6605 | 0.66267 | 0.60833 | 0.68533 | 0.60067 | 0.65417 | 0.64083 | 0.64183 | 0.61967 | 0.70833 | 0.62767 | 0.65167
| Median |  | 0.67192 | 0.65 | 0.66295 | 0.67599 | 0.6665 | 0.64764 | 0.7335 | 0.64 | 0.673 | 0.63 | 0.66 | 0.677 | 0.6635 | 0.665 | 0.6355 | 0.652 | 0.6705 | 0.6705 | 0.67235 | 0.6535 | 0.663 | 0.663 | 0.615 | 0.6825 | 0.6005 | 0.653 | 0.6415 | 0.6385 | 0.625 | 0.705 | 0.626 | 0.65
| Std Dev. |  | 0.025413 | 0.038297 | 0.0082009 | 0.0048031 | 0.019315 | 0.015479 | 0.0083287 | 0.008165 | 0.0090646 | 0.016021 | 0.0075277 | 0.016501 | 0.0075277 | 0.0056006 | 0.0072019 | 0.0065345 | 0.010577 | 0.0088015 | 0.0042165 | 0.013924 | 0.0099146 | 0.0019664 | 0.034303 | 0.010948 | 0.022651 | 0.010265 | 0.0099079 | 0.011427 | 0.029317 | 0.0098319 | 0.015616 | 0.0075277
| Rel.Std.Dev. |  | 3.783% | 5.773% | 1.246% | 0.7129% | 2.899% | 2.405% | 1.134% | 1.282% | 1.345% | 2.577% | 1.143% | 2.442% | 1.134% | 0.8407% | 1.137% | 1.005% | 1.573% | 1.309% | 0.6269% | 2.14% | 1.501% | 0.2967% | 5.639% | 1.598% | 3.771% | 1.569% | 1.546% | 1.78% | 4.731% | 1.388% | 2.488% | 1.155%
| PDM3 |  | 2.373% | 1.102% | 0.3248% | 2.689% | 1.559% | -1.921% | 11.9% | -2.962% | 2.753% | -5.249% | 0.34% | 2.982% | 1.153% | 1.534% | -3.42% | -0.8539% | 2.474% | 2.474% | 2.514% | -0.8285% | 0.6703% | 1% | -7.281% | 4.455% | -8.449% | -0.295% | -2.327% | -2.175% | -5.553% | 7.961% | -4.334% | -0.6761%
| Z-Score (Absolute) |  | NA | 0.23 | 0 | 0.7 | 0.36 | 0.67 | 3.45 | 0.98 | 0.72 | 1.67 | 0 | 0.79 | 0.24 | 0.36 | 1.12 | 0.36 | 0.64 | 0.64 | 0.65 | 0.35 | 0.1 | 0.2 | 2.27 | 1.23 | 2.62 | 0.19 | 0.8 | 0.75 | 1.76 | 2.27 | 1.39 | 0.3
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 285 (Certified Value 0.669 ppm) | [17 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | 5.05 | 5.06 | 5.07 | 5.08 | 5.09 | 5.11 | 5.12 | 5.13 | 5.14 | 5.15 | 5.16 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 0.68244 | 0.699 | 0.661 | 0.747 | 0.674 | 0.662 | 0.685 | 0.69 | 0.713 | 0.68 | 0.73 | 0.65 | 0.53 | 0.65 | 0.66 | 0.719 | 1
| 1 | 2 | 0.64857 | 0.682 | 0.659 | 0.712 | 0.68 | 0.664 | 0.66 | 0.669 | 0.594 | 0.66 | 0.71 | 0.63 | 0.66 | 0.64 | 0.656 | 0.683 | #N/A
| 1 | 3 | 0.65953 | 0.776 | 0.723 | 0.66 | 0.671 | 0.641 | 0.651 | 0.631 | 0.664 | 0.69 | 0.66 | 0.7 | 0.69 | 0.64 | 0.687 | 0.715 | 16
| 1 | 4 | 0.66035 | 0.741 | 0.678 | 0.678 | 0.689 | 0.675 | 0.623 | 0.643 | 0.645 | 0.68 | 0.68 | 0.66 | 0.71 | 0.69 | 0.69 | 0.68 | 0.66944
| 1 | 5 | 0.64527 | 0.587 | 0.663 | 0.632 | 0.685 | 0.625 | 0.638 | 0.629 | 0.662 | 0.66 | 0.74 | 0.63 | 0.68 | 0.65 | 0.652 | 0.679 | 9
| 1 | 6 | 0.73357 | 0.579 | 0.715 | 0.648 | 0.669 | 0.678 | 0.642 | 0.629 | 0.632 | 0.62 | 0.67 | 0.57 | 0.68 | 0.7 | 0.649 | 0.733
|  | 7 | 0.68476
|  | 8 | 0.66599
|  | 9 | 0.67807
|  | 10 | 0.67782
|  | 11 | 0.64951
|  | 12 | 0.61329
|  | 13 | 0.68719
|  | 14 | 0.68631
|  | 15 | 0.67162
|  | 16 | 0.6518
|  | 17 | 0.71207
|  | 18 | 0.68384
|  | 19 | 0.67222
|  | 20 | 0.66924
| Mean |  | 0.67167 | 0.67733 | 0.68317 | 0.6795 | 0.678 | 0.6575 | 0.64983 | 0.6485 | 0.65167 | 0.665 | 0.69833 | 0.64 | 0.65833 | 0.66167 | 0.66567 | 0.7015
| Median |  | 0.67192 | 0.6905 | 0.6705 | 0.669 | 0.677 | 0.663 | 0.6465 | 0.637 | 0.6535 | 0.67 | 0.695 | 0.64 | 0.68 | 0.65 | 0.658 | 0.699
| Std Dev. |  | 0.025413 | 0.080137 | 0.028666 | 0.042997 | 0.008 | 0.020579 | 0.021274 | 0.025454 | 0.039439 | 0.0251 | 0.033116 | 0.042895 | 0.064936 | 0.026394 | 0.018096 | 0.023628
| Rel.Std.Dev. |  | 3.783% | 11.83% | 4.196% | 6.328% | 1.18% | 3.13% | 3.274% | 3.925% | 6.052% | 3.774% | 4.742% | 6.702% | 9.864% | 3.989% | 2.718% | 3.368%
| PDM3 |  | 0.333% | 1.178% | 2.05% | 1.502% | 1.278% | -1.784% | -2.929% | -3.129% | -2.656% | -0.6639% | 4.315% | -4.398% | -1.66% | -1.162% | -0.5643% | 4.788%
| Z-Score (Absolute) |  | NA | 0.62 | 0.92 | 0.73 | 0.66 | 0.38 | 0.77 | 0.83 | 0.67 | 0 | 1.69 | 1.26 | 0.34 | 0.17 | 0.03 | 1.85
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 285 (Certified Value 0.641 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.04 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.28 | 1.32 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | 3
|  |  | 0.085g | 50g | 25g | 50g | 30g | 50g | 30g | 10g | 15g | 25g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 25g | 15g | 25g | 10g | 50g | 30g | 3
| 1 | 1 | 0.68244 | 0.65 | 0.663 | 0.64 | 0.626 | 0.64 | 0.64 | 0.64 | 0.6347 | 0.671 | 0.608 | 0.66 | 0.707 | 0.664 | 0.66872 | 0.605 | 0.63 | 0.52 | 0.6494 | > 0.5 | 0.621 | 0.673 | 0.62 | 1
| 1 | 2 | 0.64857 | 0.61 | 0.667 | 0.6 | 0.641 | 0.64 | 0.64 | 0.65 | 0.6051 | 0.644 | 0.601 | 0.647 | 0.699 | 0.668 | 0.67926 | 0.588 | 0.63 | 0.54 | 0.6508 | > 0.5 | 0.597 | 0.6726 | 0.65 | #N/A
| 1 | 3 | 0.65953 | 0.59 | 0.65 | 0.64 | 0.62 | 0.63 | 0.64 | 0.68 | 0.571 | 0.647 | 0.616 | 0.657 | 0.701 | 0.677 | 0.68073 | 0.601 | 0.61 | 0.55 | 0.6697 | > 0.5 | 0.656 | 0.674 | 0.7 | 16
| 1 | 4 | 0.66035 | 0.6 | 0.659 | 0.62 | 0.644 | 0.6 | 0.64 | 0.65 | 0.5806 | 0.691 | 0.605 | 0.656 | 0.711 | 0.652 | 0.68348 | 0.598 | 0.63 | 0.53 | 0.6702 | > 0.5 | 0.61 | 0.666 | 0.55 | 0.64066
| 1 | 5 | 0.64527 | 0.62 | 0.652 | 0.62 | 0.639 | 0.6 | 0.65 | 0.65 | 0.581 | 0.673 | 0.612 | 0.661 | 0.699 | 0.69 | 0.6675 | 0.585 | 0.62 | 0.56 | 0.6195 | > 0.5 | 0.624 | 0.656 | 0.66 | 11
| 1 | 6 | 0.73357 | 0.61 | 0.651 | 0.6 | 0.587 | 0.6 | 0.65 | 0.77 | 0.5855 | 0.649 | 0.602 | 0.658 | 0.699 | 0.689 | 0.67309 | 0.59 | 0.62 | 0.52 | 0.6324 | > 0.5 | 0.604 | 0.673 | 0.65
|  | 7 | 0.68476
|  | 8 | 0.66599
|  | 9 | 0.67807
|  | 10 | 0.67782
|  | 11 | 0.64951
|  | 12 | 0.61329
|  | 13 | 0.68719
|  | 14 | 0.68631
|  | 15 | 0.67162
|  | 16 | 0.6518
|  | 17 | 0.71207
|  | 18 | 0.68384
|  | 19 | 0.67222
|  | 20 | 0.66924
| Mean |  | 0.67167 | 0.61333 | 0.657 | 0.62 | 0.62617 | 0.61833 | 0.64333 | 0.67333 | 0.59298 | 0.6625 | 0.60733 | 0.6565 | 0.70267 | 0.67333 | 0.67546 | 0.5945 | 0.62333 | 0.53667 | 0.64867 |  | 0.61867 | 0.6691 | 0.63833
| Median |  | 0.67192 | 0.61 | 0.6555 | 0.62 | 0.6325 | 0.615 | 0.64 | 0.65 | 0.58325 | 0.66 | 0.6065 | 0.6575 | 0.7 | 0.6725 | 0.67618 | 0.594 | 0.625 | 0.535 | 0.6501 |  | 0.6155 | 0.6728 | 0.65
| Std Dev. |  | 0.025413 | 0.020656 | 0.0070711 | 0.017889 | 0.02133 | 0.020412 | 0.005164 | 0.049261 | 0.023338 | 0.018759 | 0.0058538 | 0.00501 | 0.0051251 | 0.014882 | 0.0066457 | 0.0079687 | 0.008165 | 0.01633 | 0.020132 |  | 0.020916 | 0.0070413 | 0.050365
| Rel.Std.Dev. |  | 3.783% | 3.368% | 1.076% | 2.885% | 3.406% | 3.301% | 0.8027% | 7.316% | 3.936% | 2.832% | 0.9638% | 0.7631% | 0.7294% | 2.21% | 0.9839% | 1.34% | 1.31% | 3.043% | 3.104% |  | 3.381% | 1.052% | 7.89%
| PDM3 |  | 4.841% | -4.265% | 2.55% | -3.225% | -2.262% | -3.485% | 0.4172% | 5.1% | -7.442% | 3.409% | -5.202% | 2.472% | 9.678% | 5.1% | 5.432% | -7.205% | -2.705% | -16.23% | 1.25% |  | -3.433% | 4.439% | -0.3633%
| Z-Score (Absolute) |  | NA | 0.7 | 0.52 | 0.51 | 0.34 | 0.56 | 0.14 | 0.98 | 1.27 | 0.67 | 0.86 | 0.51 | 1.8 | 0.98 | 1.04 | 1.22 | 0.42 | 2.84 | 0.29 | NA | 0.55 | 0.86 | 0
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 285 (Certified Value 0.635 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.31 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 20g | 15g | 05g | 05g | 15g | 200g | 3
| 1 | 1 | 0.68244 | 0.61 | 0.64 | 0.63486 | 0.65 | 0.60697 | 0.65 | 0.66773 | 0.64 | 0.719 | 0.67 | 0.64 | 0.66 | 0.71 | 0.55 | 0.51 | 0.63268 | 0.57 | 0.65 | 0.6 | 0.6 | 0.492 | 1
| 1 | 2 | 0.64857 | 0.6 | 0.62 | 0.64211 | 0.65 | 0.62 | 0.64 | 0.66559 | 0.64 | 0.696 | 0.68 | 0.63 | 0.65 | 0.69 | 0.48 | 0.59 | 0.64671 | 0.59 | 0.64 | 0.6 | 0.57 | I.S. | #N/A
| 1 | 3 | 0.65953 | 0.6 | 0.63 | 0.64203 | 0.63 | 0.60396 | 0.67 | 0.67391 | 0.63 | 0.696 | 0.66 | 0.64 | 0.67 | 0.68 | 0.57 | 0.63 | 0.61638 | 0.59 | 0.64 | 0.6 | 0.59 | 0.438 | 16
| 1 | 4 | 0.66035 | 0.61 | 0.65 | 0.64286 | 0.65 | 0.60396 | 0.66 | 0.65238 | 0.64 | 0.71 | 0.67 | 0.64 | 0.66 | 0.7 | 0.57 | 0.58 | 0.64436 | 0.6 | 0.65 | 0.6 | 0.58 | 0.423 | 0.63488
| 1 | 5 | 0.64527 | 0.62 | 0.64 | 0.62511 | 0.63 | 0.59406 | 0.71 | 0.66104 | 0.64 | 0.718 | 0.68 | 0.64 | 0.68 | 0.72 | 0.5 | 0.67 | 0.62906 | 0.61 | 0.64 | 0.6 | 0.57 | 0.428 | 13
| 1 | 6 | 0.73357 | 0.57 | 0.6408 | 0.63283 | 0.64 | 0.60396 | 0.69 | 0.68411 | 0.63 | 0.684 | 0.67 | 0.64 | 0.66 | 0.7 | 0.56 | 0.54 | 0.62569 | 0.61 | 0.65 | 0.6 | 0.57 | 0.431
|  | 7 | 0.68476
|  | 8 | 0.66599
|  | 9 | 0.67807
|  | 10 | 0.67782
|  | 11 | 0.64951
|  | 12 | 0.61329
|  | 13 | 0.68719
|  | 14 | 0.68631
|  | 15 | 0.67162
|  | 16 | 0.6518
|  | 17 | 0.71207
|  | 18 | 0.68384
|  | 19 | 0.67222
|  | 20 | 0.66924
| Mean |  | 0.67167 | 0.60167 | 0.6368 | 0.63663 | 0.64167 | 0.60548 | 0.67 | 0.66746 | 0.63667 | 0.70383 | 0.67167 | 0.63833 | 0.66333 | 0.7 | 0.53833 | 0.58667 | 0.63248 | 0.595 | 0.645 | 0.6 | 0.58 | 0.4424
| Median |  | 0.67192 | 0.605 | 0.64 | 0.63844 | 0.645 | 0.60396 | 0.665 | 0.66666 | 0.64 | 0.703 | 0.67 | 0.64 | 0.66 | 0.7 | 0.555 | 0.585 | 0.63087 | 0.595 | 0.645 | 0.6 | 0.575 | 0.431
| Std Dev. |  | 0.025413 | 0.017224 | 0.010385 | 0.0070486 | 0.0098319 | 0.0083713 | 0.026077 | 0.010871 | 0.005164 | 0.014034 | 0.0075277 | 0.0040825 | 0.010328 | 0.014142 | 0.038687 | 0.058195 | 0.011493 | 0.015166 | 0.0054772 | 0 | 0.012649 | 0.028254
| Rel.Std.Dev. |  | 3.783% | 2.863% | 1.631% | 1.107% | 1.532% | 1.383% | 3.892% | 1.629% | 0.8111% | 1.994% | 1.121% | 0.6396% | 1.557% | 2.02% | 7.186% | 9.92% | 1.817% | 2.549% | 0.8492% | 0% | 2.181% | 6.387%
| PDM3 |  | 5.796% | -5.231% | 0.3027% | 0.2762% | 1.069% | -4.63% | 5.532% | 5.132% | 0.2817% | 10.86% | 5.795% | 0.5442% | 4.482% | 10.26% | -15.21% | -7.594% | -0.3779% | -6.281% | 1.594% | -5.494% | -8.644% | -30.32%
| Z-Score (Absolute) |  | NA | 0.73 | 0 | 0 | 0.1 | 0.65 | 0.7 | 0.64 | 0 | 1.4 | 0.73 | 0.03 | 0.56 | 1.32 | 2.06 | 1.05 | 0.09 | 0.87 | 0.17 | NA | 1.19 | 4.06
Results from laboratory 1.22 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 285 (Certified Value 0.091 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.16055 | 0.1 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.06 | 0.09 | 0.11 | 0.1 | 0.11 | 0.06 | < 0.1 | < 0.1 | 0.22 | < 0.1 | 0.1 | < 0.1 | 0.11 | 0.09 | 0.1 | 2.2 | 0.2 | 0.11 |  |  | 1
| 1 | 2 | 0.14439 | 0.09 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.04 | 0.09 | 0.11 | 0.11 | 0.1 | 0.04 | < 0.1 | < 0.1 | 0.21 | < 0.1 | 0.06 | < 0.1 | 0.11 | 0.08 | 0.09 | 2.4 | 0.224 | 0.1 |  |  | 22
| 1 | 3 | 0.17592 | 0.1 | 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.03 | 0.06 | 0.1 | 0.09 | 0.07 | 0.04 | < 0.1 | < 0.1 | 0.22 | < 0.1 | 0.07 | < 0.1 | 0.11 | 0.09 | 0.1 | 2.3 | 0.3 | 0.1 |  |  | 16
| 1 | 4 | 0.14903 | 0.1 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.06 | 0.06 | 0.09 | 0.09 | 0.09 | 0.06 | < 0.1 | < 0.1 | 0.21 | < 0.1 | 0.08 | < 0.1 | 0.11 | 0.1 | 0.09 | 2.2 | 0.22 | 0.09 |  |  | 0.090636
| 1 | 5 | 0.15308 | 0.09 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.06 | 0.07 | 0.1 | 0.1 | 0.09 | 0.05 | < 0.1 | < 0.1 | 0.2 | < 0.1 | 0.09 | < 0.1 | 0.1 | 0.09 | 0.08 | 2.2 | 0.55 | 0.1 |  |  | 15
| 1 | 6 | 0.14541 | 0.1 | < 0.3 | < 0.1 | < 0.5 | < 0.3 | 0.09 | 0.07 | 0.11 | 0.09 | 0.09 | 0.07 | < 0.1 | < 0.1 | 0.2 | < 0.1 | 0.12 | < 0.1 | 0.1 | 0.1 | 0.08 | 2.2 | 0.24 | 0.09
| Mean |  | 0.15473 | 0.096667 | 0.3 |  |  |  | 0.056667 | 0.073333 | 0.10333 | 0.096667 | 0.091667 | 0.053333 |  |  | 0.21 |  | 0.086667 |  | 0.10667 | 0.091667 | 0.09 | 2.25 | 0.289 | 0.098333
| Median |  | 0.15105 | 0.1 | 0.3 |  |  |  | 0.06 | 0.07 | 0.105 | 0.095 | 0.09 | 0.055 |  |  | 0.21 |  | 0.085 |  | 0.11 | 0.09 | 0.09 | 2.2 | 0.232 | 0.1
| Std Dev. |  | 0.011931 | 0.005164 |  |  |  |  | 0.020656 | 0.013663 | 0.008165 | 0.008165 | 0.013292 | 0.012111 |  |  | 0.0089443 |  | 0.021602 |  | 0.005164 | 0.0075277 | 0.0089443 | 0.083666 | 0.13233 | 0.0075277
| Rel.Std.Dev. |  | 7.711% | 5.342% |  |  |  |  | 36.45% | 18.63% | 7.902% | 8.447% | 14.5% | 22.71% |  |  | 4.259% |  | 24.93% |  | 4.841% | 8.212% | 9.938% | 3.718% | 45.79% | 7.655%
| PDM3 |  | 70.71% | 6.653% | 231% |  |  |  | -37.48% | -19.09% | 14.01% | 6.653% | 1.137% | -41.16% |  |  | 131.7% |  | -4.38% |  | 17.69% | 1.137% | -0.7021% | 2382% | 218.9% | 8.492%
| Z-Score (Absolute) |  | 1.16 | 0.09 | NA | 0.77 | 2.91 | 1.07 | 0.64 | 0.34 | 0.21 | 0.09 | 0 | 0.7 | 0.77 | 0.77 | 2.18 | 0.77 | 0.09 | 0.77 | 0.28 | 0 | 0.03 | NA | 3.63 | 0.12
Results from laboratories 1.09 and 1.29 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 285 (Certified Value 7.76 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 7.7234 | 7.4401 | 8.0313 | 7.58 | 7.5075 | 7.76 | 7.2223 | 7.87 | 7.53 | 7.77 | 7.66 | 7.32 | 7.69 | 7.8922 | 6.44 | 7.09 | 8.09 | 7.51 | 7.69 | 7.91 | 7.9413 | 7.8515 | 8.41 | 7.69 | 5.45 |  | 1
| 1 | 2 | 7.7458 | 7.5374 | 7.874 | 7.32 | 7.7833 | 7.69 | 7.2191 | 8.22 | 7.44 | 7.97 | 7.39 | 7.22 | 7.7 | 7.8544 | 6.29 | 7.1 | 8 | 7.81 | 7.7 | 8.03 | 8.0271 | 8.0103 | 8.36 | 8.15 | 5.34 |  | #N/A
| 1 | 3 | 7.7207 | 7.5256 | 8.0571 | 7.51 | 7.3188 | 7.92 | 7.4038 | 7.72 | 7.47 | 8.19 | 7.71 | 7.19 | 7.91 | 7.6901 | 6.25 | 7.16 | 8.04 | 7.63 | 7.14 | 7.95 | 8.1701 | 8.0729 | 8.26 | 8.57 | 5.41 |  | 16
| 1 | 4 | 7.6912 | 7.6456 | 8.1409 | 7.49 | 7.8553 | 7.84 | 7.3549 | 7.61 | 7.46 | 8.13 | 7.89 | 7.3 | 7.67 | 7.951 | 6.07 | 7.07 | 8.03 | 7.71 | 7.65 | 7.83 | 8.1978 | 8.1284 | 8.47 | 8.38 | 5.68 |  | 7.759
| 1 | 5 | 7.6912 | 7.522 | 7.8274 | 7.61 | 7.7714 | 7.67 | 7.3816 | 7.72 | 7.61 | 8.16 | 7.57 | 7.4 | 7.81 | 7.7616 | 6.11 | 7.1 | 8.05 | 7.56 | 7.59 | 7.95 | 8.2657 | 7.914 | 8.57 | 8.4 | 5.58 |  | 16
| 1 | 6 | 7.7704 | 7.6321 | 8.0805 | 7.29 | 7.5618 | 7.65 | 7.2971 | 7.82 | 7.48 | 8.08 | 7.65 | 7.49 | 7.75 | 7.6994 | 6.23 | 7.06 | 7.98 | 7.62 | 7.25 | 7.82 | 8.2716 | 7.9724 | 8.47 | 8.43 | 5.22
| Mean |  | 7.7238 | 7.5505 | 8.0019 | 7.4667 | 7.633 | 7.755 | 7.3131 | 7.8267 | 7.4983 | 8.05 | 7.645 | 7.32 | 7.755 | 7.8081 | 6.2317 | 7.0967 | 8.0317 | 7.64 | 7.5033 | 7.915 | 8.1456 | 7.9916 | 8.4233 | 8.27 | 5.4467
| Median |  | 7.722 | 7.5315 | 8.0442 | 7.5 | 7.6666 | 7.725 | 7.326 | 7.77 | 7.475 | 8.105 | 7.655 | 7.31 | 7.725 | 7.808 | 6.24 | 7.095 | 8.035 | 7.625 | 7.62 | 7.93 | 8.1839 | 7.9913 | 8.44 | 8.39 | 5.43
| Std Dev. |  | 0.030955 | 0.076802 | 0.12347 | 0.13307 | 0.20528 | 0.10672 | 0.080003 | 0.21276 | 0.062423 | 0.15735 | 0.16441 | 0.11189 | 0.091159 | 0.10726 | 0.13273 | 0.035024 | 0.038687 | 0.10733 | 0.24443 | 0.079937 | 0.13372 | 0.10168 | 0.10652 | 0.31477 | 0.16513
| Rel.Std.Dev. |  | 0.4008% | 1.017% | 1.543% | 1.782% | 2.689% | 1.376% | 1.094% | 2.718% | 0.8325% | 1.955% | 2.151% | 1.529% | 1.175% | 1.374% | 2.13% | 0.4935% | 0.4817% | 1.405% | 3.258% | 1.01% | 1.642% | 1.272% | 1.265% | 3.806% | 3.032%
| PDM3 |  | -0.4542% | -2.688% | 3.13% | -3.768% | -1.624% | -0.05213% | -5.747% | 0.8715% | -3.36% | 3.75% | -1.47% | -5.658% | -0.05213% | 0.6326% | -19.69% | -8.537% | 3.514% | -1.534% | -3.296% | 2.01% | 4.982% | 2.997% | 8.561% | 6.585% | -29.8%
| Z-Score (Absolute) |  | 0 | 0.45 | 0.73 | 0.67 | 0.24 | 0.08 | 1.08 | 0.27 | 0.59 | 0.86 | 0.21 | 1.06 | 0.08 | 0.22 | 3.91 | 1.64 | 0.81 | 0.22 | 0.58 | 0.5 | 1.11 | 0.7 | 1.84 | 1.43 | 5.97
| Analytical results for As in OREAS 285 (Certified Value 289 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 280.247 | 304.9 | 253 | 191.6688 | 294 | 311 | 285.4 | 283 | 281 | 268 | 280 | 287 | 271.2376 | 44 | 294 | 297 | 310 | 267 | 293 | 287 | 280.4 | 308.79 | 302 | 305 | 284 |  | 1
| 1 | 2 | 280.884 | 298.4 | 240 | 212.6676 | 289 | 310.4 | 298.8 | 279 | 287 | 263 | 279 | 287 | 289.635 | 49.2 | 283 | 289 | 301 | 278 | 292 | 287 | 288.5 | 294.18 | 311 | 303 | 290 |  | 12
| 1 | 3 | 284.2287 | 295.7 | 250 | 173.1731 | 278 | 320.3 | 274.2 | 273 | 298 | 270 | 278 | 293 | 292.6505 | 46.9 | 294 | 287 | 307 | 280 | 287 | 292 | 285.6 | 300.64 | 310 | 380 | 282 |  | 16
| 1 | 4 | 285.1856 | 298.4 | 249 | 185.3817 | 296 | 317.9 | 276.8 | 274 | 297 | 279 | 290 | 284 | 288.5022 | 44.3 | 292 | 292 | 298 | 271 | 285 | 296 | 290.8 | 297.17 | 321 | 282 | 288 |  | 288.5915
| 1 | 5 | 281.7715 | 299.4 | 256 | 273.7473 | 290 | 319 | 281.3 | 276 | 296 | 266 | 291 | 291 | 274.1232 | 48 | 282 | 291 | 305 | 268 | 299 | 289 | 281.1 | 303.98 | 301 | 292 | 271 |  | 17
| 1 | 6 | 288.5769 | 297.9 | 256 | 133.2334 | 287 | 312 | 282.3 | 276 | 295 | 270 | 297 | 287 | 287.3218 | 48.8 | 291 | 299 | 304 | 252 | 284 | 290 | 287.6 | 292.91 | 304 | 286 | 271
| Mean |  | 283.4823 | 299.1167 | 250.6667 | 194.9787 | 289 | 315.1 | 283.1333 | 276.8333 | 292.3333 | 269.3333 | 285.8333 | 288.1667 | 283.9117 | 46.8667 | 289.3333 | 292.5 | 304.1667 | 269.3333 | 290 | 290.1667 | 285.6667 | 299.6117 | 308.1667 | 308 | 281
| Median |  | 283.0001 | 298.4 | 251.5 | 188.5252 | 289.5 | 314.95 | 281.8 | 276 | 295.5 | 269 | 285 | 287 | 287.912 | 47.45 | 291.5 | 291.5 | 304.5 | 269.5 | 289.5 | 289.5 | 286.6 | 298.905 | 307 | 297.5 | 283
| Std Dev. |  | 3.1483 | 3.089 | 5.9889 | 46.7085 | 6.3246 | 4.4407 | 8.6551 | 3.656 | 6.802 | 5.4283 | 7.8846 | 3.2506 | 8.9249 | 2.2483 | 5.4283 | 4.6368 | 4.2622 | 9.9933 | 5.7271 | 3.4303 | 4.1645 | 6.0788 | 7.5211 | 36.4253 | 8.2462
| Rel.Std.Dev. |  | 1.111% | 1.033% | 2.389% | 23.96% | 2.188% | 1.409% | 3.057% | 1.321% | 2.327% | 2.015% | 2.758% | 1.128% | 3.144% | 4.797% | 1.876% | 1.585% | 1.401% | 3.71% | 1.975% | 1.182% | 1.458% | 2.029% | 2.441% | 11.83% | 2.935%
| PDM3 |  | -1.77% | 3.647% | -13.14% | -32.44% | 0.1415% | 9.185% | -1.891% | -4.074% | 1.297% | -6.673% | -0.9557% | -0.1472% | -1.622% | -83.76% | 0.257% | 1.354% | 5.397% | -6.673% | 0.488% | 0.5458% | -1.013% | 3.819% | 6.783% | 6.725% | -2.631%
| Z-Score (Absolute) |  | 0.44 | 1.03 | 3.53 | 8.77 | 0.08 | 2.53 | 0.47 | 1.07 | 0.39 | 1.77 | 0.22 | 0 | 0.4 | 22.7 | 0.11 | 0.41 | 1.51 | 1.77 | 0.17 | 0.19 | 0.24 | 1.08 | 1.88 | 1.87 | 0.67
| Analytical results for B in OREAS 285 (Indicative Value 10.3 ppm) | [27 empty] | Indicative
B | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 1.10 | 1.11 | [24 empty] | ppm
|  |  | 4A*OES | 4A*OES | [24 empty] | 1
1
| 1 | 1 | 0.5756 | 20 | [24 empty] | 1
| 1 | 2 | 1.3143 | 19 | [24 empty] | 1
| 1 | 3 | 1.6755 | 19 | [24 empty] | 16
| 1 | 4 | 3.0819 | 20 | [24 empty] | 10.2835
| 1 | 5 | 0.4942 | 19 | [24 empty] | 7
| 1 | 6 | 1.2603 | 18
| Mean |  | 1.4003 | 19.1667
| Median |  | 1.2873 | 19
| Std Dev. |  | 0.94145 | 0.75277
| Rel.Std.Dev. |  | 67.23% | 3.928%
| PDM3 |  | -86.38% | 86.38%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ba in OREAS 285 (Certified Value 1032 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 1017.4288 | 1046.9 | 1090 | 1029.2499 | 954 | 1152 | 1074 | 1030 | 1030 | 1010 | 1020 | 1070 | 1016.3798 | > 1000 | 986 | 965 | 1110 | 1052 | 1014 | 1030 | 1020.9 | 1050 | 910 | 1027 |  |  | 1
| 1 | 2 | 1020.6731 | 1045.2 | 1080 | 995.3163 | 952 | 1132 | 1124 | 1030 | 1150 | 990 | 1020 | 1070 | 1083.6296 | > 1000 | 934 | 955 | 1080 | 1079 | 1031 | 1033 | 1052.5 | 1070 | 957 | 1008 |  |  | 13
| 1 | 3 | 1027.0606 | 1025.8 | 1110 | 974.727 | 963 | 1188 | 1064 | 1000 | 1090 | 1020 | 1010 | 1100 | 1094.377 | > 1000 | 988 | 941 | 1100 | 1048 | 1018 | 1048 | 1052.8 | 1040 | 1100 | 1032 |  |  | 16
| 1 | 4 | 1063.5199 | 1039.4 | 1090 | 1238.7679 | 953 | 1183 | 1054 | 1020 | 1070 | 1040 | 1030 | 1070 | 1076.8389 | > 1000 | 975 | 952 | 1080 | 1065 | 1007 | 1057 | 1024.6 | 1060 | 980 | 1029 |  |  | 1031.6368
| 1 | 5 | 1018.1422 | 1050.7 | 1090 | 996.6721 | 946 | 1178 | 1067 | 1020 | 1140 | 1000 | 1050 | 1090 | 1021.2669 | > 1000 | 950 | 958 | 1060 | 1055 | 1018 | 1053 | 1026.2 | 1080 | 990 | 990 |  |  | 18
| 1 | 6 | 1036.7289 | 1039.4 | 1060 | 984.3115 | 955 | 1170 | 1081 | 1010 | 1070 | 1020 | 1070 | 1090 | 1065.39 | > 1000 | 968 | 958 | 1100 | 986 | 1003 | 1044 | 1031.6 | 1080 | 990 | 993
| Mean |  | 1030.5922 | 1041.2333 | 1086.6667 | 1036.5074 | 953.8333 | 1167.1667 | 1077.3333 | 1018.3333 | 1091.6667 | 1013.3333 | 1033.3333 | 1081.6667 | 1059.647 |  | 966.8333 | 954.8333 | 1088.3333 | 1047.5 | 1015.1667 | 1044.1667 | 1034.7667 | 1063.3333 | 987.8333 | 1013.1667
| Median |  | 1023.8669 | 1042.3 | 1090 | 995.9942 | 953.5 | 1174 | 1070.5 | 1020 | 1080 | 1015 | 1025 | 1080 | 1071.1144 |  | 971.5 | 956.5 | 1090 | 1053.5 | 1016 | 1046 | 1028.9 | 1065 | 985 | 1017.5
| Std Dev. |  | 17.6694 | 8.745 | 16.3299 | 100.7837 | 5.4924 | 21.3393 | 24.6225 | 11.6905 | 45.7894 | 17.5119 | 22.5093 | 13.2916 | 33.0302 |  | 21.1888 | 8.0353 | 18.3485 | 32.117 | 9.8268 | 10.7966 | 14.274 | 16.3299 | 62.707 | 18.798
| Rel.Std.Dev. |  | 1.714% | 0.8399% | 1.503% | 9.723% | 0.5758% | 1.828% | 2.286% | 1.148% | 4.194% | 1.728% | 2.178% | 1.229% | 3.117% |  | 2.192% | 0.8415% | 1.686% | 3.066% | 0.968% | 1.034% | 1.379% | 1.536% | 6.348% | 1.855%
| PDM3 |  | -0.1012% | 0.9302% | 5.334% | 0.4721% | -7.542% | 13.14% | 4.43% | -1.29% | 5.819% | -1.774% | 0.1645% | 4.85% | 2.715% |  | -6.282% | -7.445% | 5.496% | 1.538% | -1.597% | 1.215% | 0.3034% | 3.072% | -4.246% | -1.79%
| Z-Score (Absolute) |  | 0.17 | 0.14 | 1.45 | 0 | 2.39 | 3.77 | 1.18 | 0.53 | 1.59 | 0.67 | 0.09 | 1.3 | 0.67 | NA | 2.01 | 2.36 | 1.5 | 0.32 | 0.62 | 0.22 | 0.05 | 0.77 | 1.41 | 0.67
| Analytical results for Be in OREAS 285 (Certified Value 2.51 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.452 | 2.74 | 2.8 | 1.9058 | < 5 | 2.55 | 2 | 2.54 | 2.46 | 2.47 | 2.46 | 2.54 | 2.4666 | 2 | 2.6 | 2.2 | 2.41 | 3 | 2.72 | 2.48 | 2.54 | 3.1 | 3.28 | 2 |  |  | 1
| 1 | 2 | 2.4569 | 2.71 | 2.6 | 1.9347 | < 5 | 2.53 | 2 | 2.54 | 2.51 | 2.39 | 2.45 | 2.52 | 2.5163 | 2 | 2.3 | 2.2 | 2.33 | 3 | 2.74 | 2.61 | 2.62 | 3.2 | 2.56 | 2.2 |  |  | 25
| 1 | 3 | 2.4332 | 2.74 | 2.6 | 1.9361 | < 5 | 2.58 | 2 | 2.5 | 2.6 | 2.5 | 2.44 | 2.57 | 2.4993 | 2 | 2.5 | 2.4 | 2.51 | 3 | 2.7 | 2.6 | 2.7 | 3 | 3.57 | 2.2 |  |  | 16
| 1 | 4 | 2.4517 | 2.71 | 2.5 | 1.9646 | < 5 | 2.57 | 2 | 2.52 | 2.57 | 2.49 | 2.45 | 2.55 | 2.4484 | 2 | 2.4 | 2.3 | 2.24 | 3 | 2.67 | 2.56 | 2.59 | 3.1 | 3.21 | 2.3 |  |  | 2.5058
| 1 | 5 | 2.4367 | 2.72 | 2.7 | 1.8008 | < 5 | 2.58 | 2 | 2.59 | 2.56 | 2.44 | 2.51 | 2.56 | 2.4577 | 2 | 2.4 | 2.2 | 2.49 | 3 | 2.71 | 2.57 | 2.71 | 3.2 | 3.07 | 2.2 |  |  | 19
| 1 | 6 | 2.4614 | 2.74 | 2.5 | 1.9196 | < 5 | 2.57 | 2 | 2.49 | 2.58 | 2.46 | 2.6 | 2.53 | 2.4841 | 2 | 2.4 | 2.3 | 2.29 | 2 | 2.59 | 2.51 | 2.69 | 3.1 | 3.52 | 2.3
| Mean |  | 2.4486 | 2.7267 | 2.6167 | 1.9103 |  | 2.5633 | 2 | 2.53 | 2.5467 | 2.4583 | 2.485 | 2.545 | 2.4787 | 2 | 2.4333 | 2.2667 | 2.3783 | 2.8333 | 2.6883 | 2.555 | 2.6417 | 3.1167 | 3.2017 | 2.2
| Median |  | 2.4518 | 2.73 | 2.6 | 1.9272 |  | 2.57 | 2 | 2.53 | 2.565 | 2.465 | 2.455 | 2.545 | 2.4754 | 2 | 2.4 | 2.25 | 2.37 | 3 | 2.705 | 2.565 | 2.655 | 3.1 | 3.245 | 2.2
| Std Dev. |  | 0.011234 | 0.015055 | 0.1169 | 0.057108 |  | 0.019664 | 0 | 0.035777 | 0.052026 | 0.039707 | 0.061563 | 0.018708 | 0.025971 | 0 | 0.10328 | 0.08165 | 0.10962 | 0.40825 | 0.053448 | 0.050892 | 0.069113 | 0.075277 | 0.36668 | 0.10954
| Rel.Std.Dev. |  | 0.4588% | 0.5522% | 4.468% | 2.99% |  | 0.7671% | 0% | 1.414% | 2.043% | 1.615% | 2.477% | 0.7351% | 1.048% | 0% | 4.244% | 3.602% | 4.609% | 14.41% | 1.988% | 1.992% | 2.616% | 2.415% | 11.45% | 4.979%
| PDM3 |  | -2.282% | 8.813% | 4.423% | -23.77% |  | 2.295% | -20.19% | 0.9647% | 1.63% | -1.895% | -0.8311% | 1.563% | -1.081% | -20.19% | -2.893% | -9.544% | -5.088% | 13.07% | 7.283% | 1.962% | 5.421% | 24.38% | 27.77% | -12.2%
| Z-Score (Absolute) |  | 0.63 | 1.53 | 0.67 | 4.82 | 0.23 | 0.26 | NA | 0 | 0.13 | 0.56 | 0.35 | 0.12 | 0.4 | NA | 0.75 | 2.05 | 1.18 | NA | 1.23 | 0.19 | 0.87 | 4.56 | 5.23 | 2.57
Results from laboratories 1.13, 1.20 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 285 (Certified Value 0.59 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.69571 | 0.59 | 0.58 | < 0.1 | < 5 | 0.9 | 0.83 | 0.65 | 0.7 | 0.42 | 0.53 | 0.58 | 0.47911 | 0.5 | 0.62 | 0.6 | 1.28 | 0.5 | 0.56 | 0.67 | 0.64 | 0.69 | 0.61 |  |  |  | 1
| 1 | 2 | 0.60482 | 0.79 | 0.66 | < 0.1 | < 5 | 0.73 | 0.84 | 0.51 | 0.49 | 0.58 | 0.51 | 0.55 | 0.58582 | 0.6 | 0.62 | 0.6 | 1.4 | 0.6 | 0.68 | 0.68 | 0.51 | 0.525 | 0.51 |  |  |  | 26
| 1 | 3 | 0.67393 | 0.63 | 0.62 | < 0.1 | < 5 | 0.83 | 0.62 | 0.44 | 0.55 | 0.64 | 0.59 | 0.55 | 0.47836 | 0.5 | 0.57 | 0.58 | 0.62 | 0.7 | 0.65 | 0.62 | 0.69 | 0.61 | 0.45 |  |  |  | 16
| 1 | 4 | 0.6194 | 0.63 | 0.55 | < 0.1 | < 5 | 0.67 | 0.77 | 0.94 | 0.63 | 0.63 | 0.99 | 0.58 | 0.54855 | 0.4 | 0.55 | 0.59 | 0.61 | 0.6 | 0.56 | 0.63 | 0.52 | 0.59 | 0.49 |  |  |  | 0.586
| 1 | 5 | 0.63068 | 0.54 | 0.61 | < 0.1 | < 5 | 0.8 | 0.5 | 0.71 | 0.65 | 0.52 | 0.51 | 0.48 | 0.53892 | 0.4 | 0.58 | 0.59 | 0.59 | 0.5 | 0.56 | 0.54 | 0.49 | 0.7 | 0.94 |  |  |  | 20
| 1 | 6 | 0.65026 | 0.57 | 0.71 | < 0.1 | < 5 | 0.78 | 0.68 | 0.53 | 0.59 | 0.48 | 0.55 | 0.51 | 0.52157 | 0.5 | 0.63 | 0.6 | 0.87 | 0.6 | 0.6 | 0.69 | 0.64 | 0.65 | 0.75
| Mean |  | 0.6458 | 0.625 | 0.62167 |  |  | 0.785 | 0.70667 | 0.63 | 0.60167 | 0.545 | 0.61333 | 0.54167 | 0.52539 | 0.48333 | 0.595 | 0.59333 | 0.895 | 0.58333 | 0.60167 | 0.63833 | 0.58167 | 0.6275 | 0.625
| Median |  | 0.64047 | 0.61 | 0.615 |  |  | 0.79 | 0.725 | 0.59 | 0.61 | 0.55 | 0.54 | 0.55 | 0.53024 | 0.5 | 0.6 | 0.595 | 0.745 | 0.6 | 0.58 | 0.65 | 0.58 | 0.63 | 0.56
| Std Dev. |  | 0.03437 | 0.088034 | 0.057067 |  |  | 0.079687 | 0.13261 | 0.18078 | 0.074944 | 0.087121 | 0.18694 | 0.039707 | 0.041809 | 0.075277 | 0.032711 | 0.008165 | 0.36159 | 0.075277 | 0.052313 | 0.055648 | 0.084715 | 0.066163 | 0.18823
| Rel.Std.Dev. |  | 5.322% | 14.09% | 9.18% |  |  | 10.15% | 18.77% | 28.69% | 12.46% | 15.99% | 30.48% | 7.331% | 7.958% | 15.57% | 5.498% | 1.376% | 40.4% | 12.9% | 8.695% | 8.718% | 14.56% | 10.54% | 30.12%
| PDM3 |  | 10.2% | 6.655% | 6.086% |  |  | 33.96% | 20.59% | 7.508% | 2.673% | -6.997% | 4.664% | -7.566% | -10.34% | -17.52% | 1.536% | 1.251% | 52.73% | -0.4553% | 2.673% | 8.93% | -0.7397% | 7.082% | 6.655%
| Z-Score (Absolute) |  | 0.73 | 0.19 | 0.11 | 14.59 | 48.39 | 4.31 | 2.29 | 0.32 | 0.41 | 1.86 | 0.11 | 1.95 | 2.37 | NA | 0.58 | 0.62 | 7.13 | 0.88 | 0.41 | 0.54 | 0.92 | 0.26 | 0.19
Results from laboratory 1.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 285 (Certified Value 1.94 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.8447 | 1.8942 | 1.9988 | 1.88 | 1.992 | 1.93 | 2.0243 | 2.01 | 1.95 | 1.99 | 1.94 | 1.88 | 2.05 | 2.043 | 1.8 | 1.8 | 2 | 1.81 | 1.93 | 1.899 | 1.9206 | 1.9843 | 2.048 | 1.98 | 1.87 | 1.84 | 1
| 1 | 2 | 1.8762 | 1.88 | 1.9623 | 1.85 | 2.0467 | 1.91 | 2.0137 | 2.1 | 1.92 | 2.04 | 1.89 | 1.86 | 2.06 | 1.9606 | 1.8 | 1.81 | 2 | 1.84 | 1.92 | 1.937 | 1.9422 | 1.9922 | 1.997 | 1.973 | 1.86 | 1.85 | #N/A
| 1 | 3 | 1.8659 | 1.885 | 2.0097 | 1.95 | 1.9815 | 1.95 | 2.0646 | 1.91 | 1.93 | 2.11 | 1.95 | 1.84 | 2.11 | 1.9353 | 1.84 | 1.82 | 1.91 | 1.79 | 1.77 | 1.916 | 1.9384 | 2.033 | 1.859 | 1.937 | 1.87 | 1.85 | 16
| 1 | 4 | 1.8648 | 1.9401 | 2.0258 | 1.95 | 2.0701 | 1.92 | 2.0612 | 1.93 | 1.93 | 2.1 | 1.99 | 1.86 | 2.04 | 2.0499 | 1.83 | 1.8 | 2 | 1.82 | 1.91 | 1.884 | 1.9697 | 1.9699 | 1.904 | 1.98 | 1.98 | 1.88 | 1.939
| 1 | 5 | 1.8574 | 1.9146 | 1.9953 | 2.02 | 2.049 | 1.9 | 2.0724 | 1.95 | 1.95 | 2.09 | 1.92 | 1.9 | 2.06 | 1.9461 | 1.82 | 1.8 | 2.01 | 1.84 | 1.89 | 1.914 | 1.955 | 1.957 | 2.13 | 2.001 | 1.98 | 1.84 | 21
| 1 | 6 | 1.8805 | 1.9278 | 2.0086 | 2.03 | 2.0268 | 1.92 | 2.0581 | 2 | 1.92 | 2.09 | 1.93 | 2.01 | 2.04 | 1.9266 | 1.81 | 1.79 | 2 | 1.8 | 1.82 | 1.886 | 1.964 | 1.9995 | 1.889 | 1.973 | 1.91 | 1.79
| Mean |  | 1.8649 | 1.9069 | 2.0001 | 1.9467 | 2.0277 | 1.9217 | 2.049 | 1.9833 | 1.9333 | 2.07 | 1.9367 | 1.8917 | 2.06 | 1.9769 | 1.8167 | 1.8033 | 1.9867 | 1.8167 | 1.8733 | 1.906 | 1.9483 | 1.9893 | 1.9712 | 1.974 | 1.9117 | 1.8417
| Median |  | 1.8654 | 1.9044 | 2.0037 | 1.95 | 2.0368 | 1.92 | 2.0596 | 1.975 | 1.93 | 2.09 | 1.935 | 1.87 | 2.055 | 1.9534 | 1.815 | 1.8 | 2 | 1.815 | 1.9 | 1.9065 | 1.9486 | 1.9882 | 1.9505 | 1.9765 | 1.89 | 1.845
| Std Dev. |  | 0.012933 | 0.024354 | 0.021347 | 0.072296 | 0.034705 | 0.017224 | 0.023995 | 0.069186 | 0.013663 | 0.046043 | 0.033267 | 0.061455 | 0.026077 | 0.05507 | 0.01633 | 0.010328 | 0.037771 | 0.020656 | 0.064083 | 0.020288 | 0.018176 | 0.026332 | 0.10549 | 0.020842 | 0.055648 | 0.029269
| Rel.Std.Dev. |  | 0.6935% | 1.277% | 1.067% | 3.714% | 1.712% | 0.8963% | 1.171% | 3.488% | 0.7067% | 2.224% | 1.718% | 3.249% | 1.266% | 2.786% | 0.8989% | 0.5727% | 1.901% | 1.137% | 3.421% | 1.064% | 0.9329% | 1.324% | 5.352% | 1.056% | 2.911% | 1.589%
| PDM3 |  | -3.823% | -1.655% | 3.148% | 0.3931% | 4.571% | -0.8962% | 5.673% | 2.284% | -0.2945% | 6.754% | -0.1226% | -2.443% | 6.238% | 1.953% | -6.311% | -6.999% | 2.456% | -6.311% | -3.389% | -1.704% | 0.4782% | 2.593% | 1.657% | 1.803% | -1.412% | -5.022%
| Z-Score (Absolute) |  | 1.12 | 0.51 | 0.85 | 0.07 | 1.25 | 0.29 | 1.56 | 0.61 | 0.12 | 1.87 | 0.07 | 0.73 | 1.72 | 0.51 | 1.82 | 2.01 | 0.66 | 1.82 | 0.99 | 0.52 | 0.1 | 0.69 | 0.43 | 0.47 | 0.44 | 1.46
| Analytical results for Cd in OREAS 285 (Certified Value 0.2 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 0.5 | 0.22 | < 0.3 | 0.485 | 0.9 | 0.24 | 0.24 | 0.18 | 0.2 | 0.23 | 0.18 | 0.21 | 0.20015 | 0.4 | 0.2 | 0.15 | 0.19 | 0.2 | 0.23 | 0.22 | 0.2 | < 1 | 0.22 | 0.17 |  |  | 1
| 1 | 2 | < 0.5 | 0.2 | < 0.3 | 0.4094 | 0.9 | 0.24 | 0.32 | 0.18 | 0.2 | 0.19 | 0.18 | 0.22 | 0.2119 | 0.2 | 0.21 | 0.15 | 0.22 | 0.2 | 0.2 | 0.22 | 0.19 | < 1 | 0.19 | 0.21 |  |  | 27
| 1 | 3 | < 0.5 | 0.21 | < 0.3 | 0.4359 | 0.9 | 0.25 | 0.33 | 0.18 | 0.21 | 0.22 | 0.18 | 0.22 | 0.20796 | 0.2 | 0.22 | 0.11 | 0.22 | 0.2 | 0.21 | 0.22 | 0.21 | < 1 | 0.28 | 0.19 |  |  | 16
| 1 | 4 | < 0.5 | 0.2 | 0.5 | 0.7673 | 0.9 | 0.26 | 0.33 | 0.19 | 0.21 | 0.18 | 0.16 | 0.21 | 0.19132 | 0.2 | 0.2 | 0.15 | 0.2 | 0.2 | 0.21 | 0.19 | 0.2 | < 1 | 0.22 | 0.2 |  |  | 0.20164
| 1 | 5 | < 0.5 | 0.23 | < 0.3 | 0.5138 | 0.9 | 0.24 | 0.33 | 0.18 | 0.2 | 0.2 | 0.17 | 0.19 | 0.1798 | 0.2 | 0.2 | 0.15 | 0.23 | 0.2 | 0.2 | 0.24 | 0.2 | 1 | 0.23 | 0.16 |  |  | 22
| 1 | 6 | < 0.5 | 0.2 | < 0.3 | 0.685 | 0.9 | 0.25 | 0.33 | 0.21 | 0.19 | 0.18 | 0.21 | 0.2 | 0.19055 | 0.2 | 0.19 | 0.15 | 0.2 | 0.2 | 0.21 | 0.21 | 0.22 | 1 | 0.21 | 0.19
| Mean |  |  | 0.21 | 0.5 | 0.5494 | 0.9 | 0.24667 | 0.31333 | 0.18667 | 0.20167 | 0.2 | 0.18 | 0.20833 | 0.19695 | 0.23333 | 0.20333 | 0.14333 | 0.21 | 0.2 | 0.21 | 0.21667 | 0.20333 | 1 | 0.225 | 0.18667
| Median |  |  | 0.205 | 0.5 | 0.4994 | 0.9 | 0.245 | 0.33 | 0.18 | 0.2 | 0.195 | 0.18 | 0.21 | 0.19573 | 0.2 | 0.2 | 0.15 | 0.21 | 0.2 | 0.21 | 0.22 | 0.2 | 1 | 0.22 | 0.19
| Std Dev. |  |  | 0.012649 |  | 0.14406 | 0 | 0.008165 | 0.036148 | 0.012111 | 0.0075277 | 0.020976 | 0.016733 | 0.01169 | 0.012015 | 0.08165 | 0.010328 | 0.01633 | 0.015492 | 3.0405e-17 | 0.010954 | 0.01633 | 0.010328 | 0 | 0.030166 | 0.018619
| Rel.Std.Dev. |  |  | 6.023% |  | 26.22% | 0% | 3.31% | 11.54% | 6.488% | 3.733% | 10.49% | 9.296% | 5.611% | 6.101% | 34.99% | 5.079% | 11.39% | 7.377% | 1.52e-14% | 5.216% | 7.537% | 5.079% | 0% | 13.41% | 9.974%
| PDM3 |  |  | 4.146% | 148% | 172.5% | 346.3% | 22.33% | 55.39% | -7.425% | 0.0137% | -0.8129% | -10.73% | 3.32% | -2.327% | 15.72% | 0.8403% | -28.92% | 4.146% | -0.8129% | 4.146% | 7.453% | 0.8403% | 395.9% | 11.59% | -7.425%
| Z-Score (Absolute) |  | 1.96 | 0.04 | NA | 16.36 | 33.21 | 1.8 | 5.01 | 1.08 | 0.36 | 0.44 | 1.4 | 0.04 | 0.59 | NA | 0.28 | 3.16 | 0.04 | NA | 0.04 | 0.36 | 0.28 | NA | 0.76 | 1.08
Results from laboratories 1.09, 1.20 and 1.24 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 285 (Certified Value 78 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 73.9435 | 79.85 | 76.5 | 78.7447 | 49 | 71.19 | 76 | 72 | 78.4 | 75.6 | 77.6 | 87 | 71.85 | 93.4 | 76 | 81.07 | 81 | 79.71 | 77 | 56.2 | 68.09 | [5 empty] | 1
| 1 | 2 | 74.4942 | 80 | 76.2 | 81.0149 | 48.5 | 73.12 | 77.3 | 75.6 | 72 | 73.1 | 78.1 | 83 | 68.29 | 93.6 | 79 | 77.35 | 78 | 80.68 | 80 | 59.7 | 66.3 | [5 empty] | 28
| 1 | 3 | 75.2209 | 81.77 | 75.7 | 82.4012 | 50.4 | 70.69 | 77 | 80.7 | 80.3 | 74.1 | 81.7 | 99 | 70.05 | 93.1 | 79 | 77.28 | 82 | 80.97 | 76 | 60.8 | 65 | [5 empty] | 16
| 1 | 4 | 75.8397 | 78.96 | 77.4 | 84.6137 | 50.2 | 68.96 | 73.7 | 80.6 | 77.1 | 77.5 | 78.6 | 85 | 71.77 | 83.1 | 76 | 77.88 | 78 | 80.34 | 81 | 56.8 | 63.36 | [5 empty] | 78.0157
| 1 | 5 | 74.4221 | 77.3 | 77.6 | 88.9751 | 50.1 | 67.91 | 82.3 | 80.4 | 75.9 | 79.8 | 80.9 | 88 | 69.11 | 89.7 | 79 | 82.33 | 82 | 78.82 | 84 | 58.9 | 67.88 | [5 empty] | 23
| 1 | 6 | 76.5686 | 80.9 | 79.4 | 79.4674 | 49.5 | 69.38 | 75.3 | 83.9 | 77.3 | 80.3 | 78.4 | 89 | 69.86 | 86.3 | 74 | 76.95 | 79 | 81.86 | 82 | 53.8 | 64.71
| Mean |  | 75.0815 | 79.7967 | 77.1333 | 82.5362 | 49.6167 | 70.2083 | 76.9333 | 78.8667 | 76.8333 | 76.7333 | 79.2167 | 88.5 | 70.155 | 89.8667 | 77.1667 | 78.81 | 80 | 80.3967 | 80 | 57.7 | 65.89
| Median |  | 74.8575 | 79.925 | 76.95 | 81.7081 | 49.8 | 70.035 | 76.5 | 80.5 | 77.2 | 76.55 | 78.5 | 87.5 | 69.955 | 91.4 | 77.5 | 77.615 | 80 | 80.51 | 80.5 | 57.85 | 65.65
| Std Dev. |  | 0.98799 | 1.5549 | 1.3231 | 3.7921 | 0.75211 | 1.8547 | 2.9303 | 4.2866 | 2.7926 | 2.971 | 1.6678 | 5.5767 | 1.4247 | 4.3684 | 2.137 | 2.2933 | 1.8974 | 1.0499 | 3.0332 | 2.5815 | 1.8737
| Rel.Std.Dev. |  | 1.316% | 1.949% | 1.715% | 4.594% | 1.516% | 2.642% | 3.809% | 5.435% | 3.635% | 3.872% | 2.105% | 6.301% | 2.031% | 4.861% | 2.769% | 2.91% | 2.372% | 1.306% | 3.791% | 4.474% | 2.844%
| PDM3 |  | -3.761% | 2.283% | -1.131% | 5.794% | -36.4% | -10.01% | -1.387% | 1.091% | -1.516% | -1.644% | 1.539% | 13.44% | -10.08% | 15.19% | -1.088% | 1.018% | 2.543% | 3.052% | 2.543% | -26.04% | -15.54%
| Z-Score (Absolute) |  | 0.5 | 0.63 | 0.01 | 1.28 | 6.56 | 1.66 | 0.06 | 0.4 | 0.08 | 0.1 | 0.49 | 2.7 | 1.67 | 3.02 | 0 | 0.39 | 0.67 | 0.77 | 0.67 | 4.63 | 2.68
| Analytical results for Co in OREAS 285 (Certified Value 14.1 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 20.47 | 13.5811 | 14.2 | 15 | 14.5628 | 13 | 15 | 11.5 | 13.8 | 13 | 13.6 | 13.8 | 14.8 | 13.1098 | 13.7 | 13 | 14 | 15.2 | 14.6 | 14.7 | 14 | 14.1 | 14.84 | 15 | 13.5 | 13.7 | 1
| 1 | 2 | 20.94 | 13.7389 | 13.9 | 15 | 13.637 | 13 | 14.7 | 13.1 | 13.9 | 14.5 | 13.6 | 14.4 | 15.4 | 14.0478 | 13.3 | 13.4 | 14 | 14.6 | 14.6 | 14.5 | 14.2 | 14.4 | 14.3 | 15 | 14.4 | 13.9 | 16
| 1 | 3 | 20.92 | 13.736 | 13.8 | 15 | 14.7743 | 13 | 13.9 | 12.4 | 13.2 | 14.7 | 13.9 | 14 | 15.7 | 14.1936 | 13 | 13.7 | 13 | 15.2 | 14.6 | 14.3 | 14.5 | 14.7 | 14.59 | 14 | 14.6 | 14.1 | 16
| 1 | 4 | 20.85 | 13.9387 | 14 | 15 | 15.3491 | 13 | 13.7 | 13 | 13.5 | 14.4 | 13.6 | 14.3 | 15 | 14.2189 | 13.5 | 13.1 | 14 | 14.5 | 14.6 | 14.3 | 14.8 | 14.4 | 14.46 | 16 | 14.5 | 13.5 | 14.1293
| 1 | 5 | 20.9 | 13.759 | 13.8 | 14 | 16.1049 | 13 | 13.8 | 13.1 | 14 | 14.7 | 13.5 | 14.6 | 15.5 | 13.2667 | 13.5 | 12.5 | 14 | 14.6 | 14.3 | 14.7 | 14.7 | 14.4 | 14.74 | 15 | 15 | 13.5 | 24
| 1 | 6 | 21.5 | 14.0628 | 14 | 14 | 15.2368 | 13 | 13.4 | 12.4 | 13.2 | 14.2 | 13.2 | 14.8 | 15.6 | 13.7586 | 13.6 | 12.9 | 14 | 15.1 | 13.6 | 14 | 14.5 | 14.4 | 15.31 | 14 | 14.7 | 13.4
| Mean |  | 20.93 | 13.8027 | 13.95 | 14.6667 | 14.9442 | 13 | 14.0833 | 12.5833 | 13.6 | 14.25 | 13.5667 | 14.3167 | 15.3333 | 13.7659 | 13.4333 | 13.1 | 13.8333 | 14.8667 | 14.3833 | 14.4167 | 14.45 | 14.4 | 14.7067 | 14.8333 | 14.45 | 13.6833
| Median |  | 20.91 | 13.7489 | 13.95 | 15 | 15.0055 | 13 | 13.85 | 12.7 | 13.65 | 14.45 | 13.6 | 14.35 | 15.45 | 13.9032 | 13.5 | 13.05 | 14 | 14.85 | 14.6 | 14.4 | 14.5 | 14.4 | 14.665 | 15 | 14.55 | 13.6
| Std Dev. |  | 0.32985 | 0.17071 | 0.15166 | 0.5164 | 0.83425 | 0 | 0.62423 | 0.62423 | 0.35214 | 0.64109 | 0.22509 | 0.37103 | 0.3559 | 0.47904 | 0.25033 | 0.41473 | 0.40825 | 0.33267 | 0.40208 | 0.27142 | 0.30166 | 0.18974 | 0.35291 | 0.75277 | 0.50892 | 0.27142
| Rel.Std.Dev. |  | 1.576% | 1.237% | 1.087% | 3.521% | 5.582% | 0% | 4.432% | 4.961% | 2.589% | 4.499% | 1.659% | 2.592% | 2.321% | 3.48% | 1.864% | 3.166% | 2.951% | 2.238% | 2.795% | 1.883% | 2.088% | 1.318% | 2.4% | 5.075% | 3.522% | 1.984%
| PDM3 |  | 48.13% | -2.311% | -1.269% | 3.803% | 5.767% | -7.993% | -0.3254% | -10.94% | -3.746% | 0.8542% | -3.982% | 1.326% | 8.521% | -2.572% | -4.926% | -7.285% | -2.095% | 5.219% | 1.798% | 2.034% | 2.27% | 1.916% | 4.086% | 4.983% | 2.27% | -3.156%
| Z-Score (Absolute) |  | 8.98 | 0.65 | 0.45 | 0.52 | 0.89 | 1.73 | 0.27 | 2.3 | 0.92 | 0.05 | 0.97 | 0.05 | 1.42 | 0.7 | 1.15 | 1.6 | 0.61 | 0.79 | 0.14 | 0.18 | 0.23 | 0.16 | 0.57 | 0.74 | 0.23 | 0.81
| Analytical results for Cr in OREAS 285 (Certified Value 48.8 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 48.6998 | 43 | 50 | 57.2075 | 48 | 49.2 | 36 | 47 | 47 | 45 | 48 | 48 | 45.0334 | 51 | 55 | 54 | 45.3 | 46 | 34 | 44 | 52 | 51.43 | 47 | 50 | 27 |  | 1
| 1 | 2 | 48.3276 | 52 | 55 | 43.7847 | 51 | 48.9 | 38 | 48 | 49 | 46 | 46 | 48 | 47.8353 | 50 | 54 | 53 | 44 | 45 | 30 | 43 | 47 | 50.21 | 51 | 49 | 28 |  | 17
| 1 | 3 | 49.4326 | 53 | 46 | 61.1101 | 52 | 50.1 | 34 | 47 | 48 | 48 | 47 | 49 | 48.5351 | 45 | 55 | 53 | 41.7 | 45 | 30 | 52 | 48 | 54.07 | 47 | 50 | 28 |  | 16
| 1 | 4 | 47.4848 | 51 | 48 | 59.1853 | 49 | 50 | 33 | 48 | 49 | 47 | 47 | 50 | 41.0745 | 49 | 54 | 59 | 45.4 | 47 | 33 | 51 | 47 | 50.71 | 45 | 51 | 29 |  | 48.8317
| 1 | 5 | 47.0352 | 47 | 54 | 59.8896 | 48 | 50.3 | 35 | 49 | 48 | 45 | 48 | 50 | 42.6544 | 49 | 53 | 52 | 45.5 | 43 | 30 | 47 | 53 | 54.68 | 40 | 50 | 28 |  | 25
| 1 | 6 | 47.6056 | 47 | 56 | 44.083 | 50 | 50.3 | 36 | 47 | 48 | 45 | 50 | 50 | 43.1618 | 51 | 54 | 53 | 42.7 | 42 | 35 | 49 | 52 | 50.42 | 44 | 47 | 27
| Mean |  | 48.0976 | 48.8333 | 51.5 | 54.21 | 49.6667 | 49.8 | 35.3333 | 47.6667 | 48.1667 | 46 | 47.6667 | 49.1667 | 44.7158 | 49.1667 | 54.1667 | 54 | 44.1 | 44.6667 | 32 | 47.6667 | 49.8333 | 51.92 | 45.6667 | 49.5 | 27.8333
| Median |  | 47.9666 | 49 | 52 | 58.1964 | 49.5 | 50.05 | 35.5 | 47.5 | 48 | 45.5 | 47.5 | 49.5 | 44.0976 | 49.5 | 54 | 53 | 44.65 | 45 | 31.5 | 48 | 50 | 51.07 | 46 | 50 | 28
| Std Dev. |  | 0.88814 | 3.8166 | 4.0866 | 8.0604 | 1.633 | 0.6 | 1.7512 | 0.8165 | 0.75277 | 1.2649 | 1.3663 | 0.98319 | 2.978 | 2.2286 | 0.75277 | 2.5298 | 1.6012 | 1.8619 | 2.2804 | 3.6697 | 2.7869 | 1.9554 | 3.6697 | 1.3784 | 0.75277
| Rel.Std.Dev. |  | 1.847% | 7.816% | 7.935% | 14.87% | 3.288% | 1.205% | 4.956% | 1.713% | 1.563% | 2.75% | 2.866% | 2% | 6.66% | 4.533% | 1.39% | 4.685% | 3.631% | 4.168% | 7.126% | 7.699% | 5.592% | 3.766% | 8.036% | 2.785% | 2.705%
| PDM3 |  | -1.503% | 0.00341% | 5.464% | 11.01% | 1.71% | 1.983% | -27.64% | -2.386% | -1.362% | -5.799% | -2.386% | 0.686% | -8.429% | 0.686% | 10.93% | 10.58% | -9.69% | -8.529% | -34.47% | -2.386% | 2.051% | 6.324% | -6.481% | 1.369% | -43%
| Z-Score (Absolute) |  | 0.02 | 0.21 | 1.04 | 1.88 | 0.47 | 0.51 | 3.99 | 0.16 | 0 | 0.67 | 0.16 | 0.31 | 1.07 | 0.31 | 1.87 | 1.82 | 1.27 | 1.09 | 5.03 | 0.16 | 0.52 | 1.17 | 0.78 | 0.41 | 6.33
| Analytical results for Cs in OREAS 285 (Certified Value 11.1 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 10.7788 | 11.49 | 10.5 | 12.8 | 9 | 11.5 | 10.25 | 11.15 | 10.35 | 11.45 | 10.8894 | 11.2 | 11 | 10.9 | 11.9 | 11.43 | 11.26 | 11.04 | 10.6 | 11 | [6 empty] | 1
| 1 | 2 | 10.866 | 11.37 | 10.2 | 12.5 | 10.2 | 11.65 | 11.15 | 10.75 | 10.5 | 11.3 | 11.4304 | 11.4 | 10.5 | 10.4 | 11.6 | 11.63 | 11.47 | 11.28 | 11.3 | 11 | [6 empty] | 31
| 1 | 3 | 10.717 | 11.18 | 10.6 | 13 | 9.7 | 10.5 | 11.8 | 11.35 | 10.15 | 11.75 | 11.5537 | 11.6 | 10.6 | 10.4 | 12 | 11.17 | 11.39 | 11.23 | 11.5 | 10 | [6 empty] | 16
| 1 | 4 | 10.9562 | 11.4 | 10.5 | 13 | 10.3 | 11.05 | 11.5 | 11.3 | 10.4 | 11.45 | 11.4048 | 11.4 | 10.9 | 10.4 | 11.4 | 11.14 | 11.44 | 11.03 | 11 | 10 | [6 empty] | 11.1126
| 1 | 5 | 10.7434 | 11.21 | 10.6 | 12.9 | 10.3 | 11.1 | 11.8 | 10.8 | 10.45 | 11.55 | 10.8849 | 10.8 | 10.4 | 10.5 | 11.5 | 11.49 | 11.64 | 11.05 | 11.1 | 10 | [6 empty] | 26
| 1 | 6 | 11.0293 | 11.47 | 10.5 | 13 | 9.8 | 11.2 | 11.8 | 10.75 | 10.8 | 11.55 | 11.4302 | 11.6 | 10.9 | 10.6 | 11.8 | 11.01 | 11.63 | 11.13 | 11.1 | 10
| Mean |  | 10.8484 | 11.3533 | 10.4833 | 12.8667 | 9.8833 | 11.1667 | 11.3833 | 11.0167 | 10.4417 | 11.5083 | 11.2656 | 11.3333 | 10.7167 | 10.5333 | 11.7 | 11.3117 | 11.4717 | 11.1267 | 11.1 | 10.3333
| Median |  | 10.8224 | 11.385 | 10.5 | 12.95 | 10 | 11.15 | 11.65 | 10.975 | 10.425 | 11.5 | 11.4175 | 11.4 | 10.75 | 10.45 | 11.7 | 11.3 | 11.455 | 11.09 | 11.1 | 10
| Std Dev. |  | 0.12473 | 0.13064 | 0.1472 | 0.19664 | 0.50365 | 0.40208 | 0.6121 | 0.28225 | 0.21311 | 0.14972 | 0.2977 | 0.30111 | 0.24833 | 0.19664 | 0.23664 | 0.23987 | 0.14552 | 0.10671 | 0.30332 | 0.5164
| Rel.Std.Dev. |  | 1.15% | 1.151% | 1.404% | 1.528% | 5.096% | 3.601% | 5.377% | 2.562% | 2.041% | 1.301% | 2.643% | 2.657% | 2.317% | 1.867% | 2.023% | 2.121% | 1.269% | 0.959% | 2.733% | 4.997%
| PDM3 |  | -2.377% | 2.166% | -5.663% | 15.78% | -11.06% | 0.4866% | 2.436% | -0.8632% | -6.037% | 3.561% | 1.376% | 1.986% | -3.563% | -5.213% | 5.286% | 1.791% | 3.231% | 0.1267% | -0.1133% | -7.012%
| Z-Score (Absolute) |  | 0.7 | 0.41 | 1.51 | 3.76 | 2.84 | 0 | 0.48 | 0.33 | 1.6 | 0.76 | 0.22 | 0.37 | 0.99 | 1.4 | 1.18 | 0.32 | 0.67 | 0.09 | 0.15 | NA
Results from laboratory 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 285 (Certified Value 30.9 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 36 | 31.9472 | 31.3 | 31 | 32.5987 | 33 | 31.7 | 31.8 | 30 | 32.9 | 30.5 | 28.5 | 29.5 | 30.882 | 26 | 32.7 | 30 | 31.3 | 29.4 | 31.3 | 30 | 31.7 | 31.09 | 44 | 35.6 | 36.8 | 1
| 1 | 2 | 35.59 | 29.9881 | 31.1 | 31 | 34.5315 | 33 | 32.1 | 32.7 | 30 | 31 | 29 | 29.2 | 30.4 | 30.5416 | 25.5 | 33.6 | 29 | 30.3 | 30 | 32.9 | 30 | 32.1 | 30.98 | 40 | 39.1 | 28.2 | 18
| 1 | 3 | 36.13 | 30.76 | 31.6 | 31 | 33.332 | 34 | 32.6 | 31.3 | 28.5 | 33 | 30.1 | 29.1 | 30.6 | 28.3294 | 25.7 | 32.9 | 28 | 30.3 | 30.3 | 32.2 | 30 | 33.7 | 31.72 | 44 | 40.9 | 28 | 16
| 1 | 4 | 36.08 | 30.1552 | 31.2 | 31 | 33.1513 | 34 | 32.2 | 30 | 29.6 | 31.5 | 30.8 | 30.1 | 31.5 | 28.6916 | 27.6 | 32.9 | 30 | 31.6 | 30.1 | 30.5 | 30 | 30.8 | 30.48 | 41 | 39.8 | 28.9 | 30.8513
| 1 | 5 | 35.48 | 30.9258 | 31.5 | 30 | 32.9767 | 34 | 32.8 | 30.6 | 29.6 | 31.7 | 29.7 | 31.1 | 31 | 28.3545 | 26.2 | 33.6 | 29 | 30.2 | 28.8 | 30.6 | 30 | 30.4 | 30.24 | 39 | 38 | 29.6 | 27
| 1 | 6 | 35.1 | 31.5973 | 31.2 | 30 | 31.9932 | 34 | 32.3 | 31.2 | 28.9 | 31.8 | 28.7 | 30.9 | 30.4 | 28.8565 | 25.7 | 33.2 | 29 | 29.9 | 27.6 | 31 | 30 | 32.4 | 29.51 | 39 | 38.9 | 27.5
| Mean |  | 35.73 | 30.8956 | 31.3167 | 30.6667 | 33.0972 | 33.6667 | 32.2833 | 31.2667 | 29.4333 | 31.9833 | 29.8 | 29.8167 | 30.5667 | 29.2759 | 26.1167 | 33.15 | 29.1667 | 30.6 | 29.3667 | 31.4167 | 30 | 31.85 | 30.67 | 41.1667 | 38.7167 | 29.8333
| Median |  | 35.795 | 30.8429 | 31.25 | 31 | 33.064 | 34 | 32.25 | 31.25 | 29.6 | 31.75 | 29.9 | 29.65 | 30.5 | 28.7741 | 25.85 | 33.05 | 29 | 30.3 | 29.7 | 31.15 | 30 | 31.9 | 30.73 | 40.5 | 39 | 28.55
| Std Dev. |  | 0.40851 | 0.77319 | 0.19408 | 0.5164 | 0.84819 | 0.5164 | 0.38687 | 0.93737 | 0.60882 | 0.79854 | 0.82946 | 1.0515 | 0.67132 | 1.1352 | 0.7679 | 0.38341 | 0.75277 | 0.68118 | 1.025 | 0.94956 | 0 | 1.1845 | 0.76702 | 2.3166 | 1.8082 | 3.4909
| Rel.Std.Dev. |  | 1.143% | 2.503% | 0.6197% | 1.684% | 2.563% | 1.534% | 1.198% | 2.998% | 2.068% | 2.497% | 2.783% | 3.527% | 2.196% | 3.878% | 2.94% | 1.157% | 2.581% | 2.226% | 3.49% | 3.022% | 0% | 3.719% | 2.501% | 5.627% | 4.67% | 11.7%
| PDM3 |  | 15.81% | 0.1436% | 1.508% | -0.5985% | 7.28% | 9.126% | 4.642% | 1.346% | -4.596% | 3.669% | -3.408% | -3.354% | -0.9226% | -5.106% | -15.35% | 7.451% | -5.461% | -0.8146% | -4.812% | 1.833% | -2.759% | 3.237% | -0.5877% | 33.44% | 25.49% | -3.3%
| Z-Score (Absolute) |  | 2.94 | 0.07 | 0.32 | 0.07 | 1.38 | 1.72 | 0.89 | 0.29 | 0.8 | 0.71 | 0.58 | 0.57 | 0.13 | 0.9 | 2.77 | 1.41 | 0.96 | 0.11 | 0.84 | 0.38 | 0.47 | 0.63 | 0.07 | 6.18 | 4.72 | 0.56
| Analytical results for Dy in OREAS 285 (Certified Value 3.89 ppm) | [27 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 3.7616 | 4.02 | 4.2 | 3.5843 | 3.2 | 3.95 | 4.1 | [19 empty] | 1
| 1 | 2 | 3.7848 | 3.84 | 4.1 | 3.594 | 3.7 | 4.09 | 4.2 | [19 empty] | 33
| 1 | 3 | 3.7938 | 3.98 | 4.1 | 3.4436 | 3.8 | 4.02 | 4 | [19 empty] | 16
| 1 | 4 | 3.7851 | 3.92 | 4 | 3.6568 | 3.9 | 4.01 | 4.2 | [19 empty] | 3.8949
| 1 | 5 | 3.7883 | 3.81 | 4.1 | 3.864 | 3.7 | 4.11 | 4.3 | [19 empty] | 28
| 1 | 6 | 3.8346 | 3.81 | 3.9 | 3.6092 | 3.6 | 3.97 | 4.2
| Mean |  | 3.7914 | 3.8967 | 4.0667 | 3.6253 | 3.65 | 4.025 | 4.1667
| Median |  | 3.7867 | 3.88 | 4.1 | 3.6016 | 3.7 | 4.015 | 4.2
| Std Dev. |  | 0.023896 | 0.09048 | 0.10328 | 0.13704 | 0.2429 | 0.063797 | 0.10328
| Rel.Std.Dev. |  | 0.6303% | 2.322% | 2.54% | 3.78% | 6.655% | 1.585% | 2.479%
| PDM3 |  | -2.657% | 0.04663% | 4.411% | -6.92% | -6.287% | 3.342% | 6.979%
| Z-Score (Absolute) |  | 0.42 | 0 | 0.67 | 1.08 | 0.98 | 0.51 | 1.07
| Analytical results for Er in OREAS 285 (Certified Value 1.42 ppm) | [27 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 1.4087 | 1.44 | 1.6 | 1.3259 | 1.3 | 1.45 | 1.3 | [19 empty] | 1
| 1 | 2 | 1.4388 | 1.53 | 1.7 | 1.4496 | 1.5 | 1.46 | 1.2 | [19 empty] | 34
| 1 | 3 | 1.4818 | 1.47 | 1.8 | 1.2496 | 1.4 | 1.52 | 1.1 | [19 empty] | 16
| 1 | 4 | 1.4773 | 1.43 | 1.6 | 1.2801 | 1.4 | 1.52 | 1.2 | [19 empty] | 1.4221
| 1 | 5 | 1.4522 | 1.38 | 1.7 | 1.3889 | 1.3 | 1.47 | 1.2 | [19 empty] | 29
| 1 | 6 | 1.5119 | 1.39 | 1.7 | 1.329 | 1.4 | 1.51 | 1.3
| Mean |  | 1.4618 | 1.44 | 1.6833 | 1.3372 | 1.3833 | 1.4883 | 1.2167
| Median |  | 1.4648 | 1.435 | 1.7 | 1.3275 | 1.4 | 1.49 | 1.2
| Std Dev. |  | 0.036292 | 0.055136 | 0.075277 | 0.072713 | 0.075277 | 0.031885 | 0.075277
| Rel.Std.Dev. |  | 2.483% | 3.829% | 4.472% | 5.438% | 5.442% | 2.142% | 6.187%
| PDM3 |  | 2.789% | 1.257% | 18.37% | -5.973% | -2.728% | 4.655% | -14.45%
| Z-Score (Absolute) |  | 0.26 | 0 | 2.9 | 1.22 | 0.67 | 0.58 | 2.66
| Analytical results for Eu in OREAS 285 (Certified Value 1.46 ppm) | [27 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 1.4184 | 1.47 | 1.44 | 1.4674 | 1.3 | 1.47 | 1.6 | [19 empty] | 1
| 1 | 2 | 1.4812 | 1.45 | 1.33 | 1.464 | 1.5 | 1.52 | 1.6 | [19 empty] | 35
| 1 | 3 | 1.4402 | 1.44 | 1.4 | 1.4015 | 1.4 | 1.57 | 1.6 | [19 empty] | 16
| 1 | 4 | 1.4623 | 1.48 | 1.44 | 1.5057 | 1.4 | 1.54 | 1.7 | [19 empty] | 1.458
| 1 | 5 | 1.4254 | 1.41 | 1.43 | 1.6002 | 1.4 | 1.51 | 1.7 | [19 empty] | 30
| 1 | 6 | 1.4959 | 1.45 | 1.45 | 1.4931 | 1.4 | 1.53 | 1.6
| Mean |  | 1.4539 | 1.45 | 1.415 | 1.4887 | 1.4 | 1.5233 | 1.6333
| Median |  | 1.4512 | 1.45 | 1.435 | 1.4803 | 1.4 | 1.525 | 1.6
| Std Dev. |  | 0.031105 | 0.024495 | 0.045056 | 0.065434 | 0.063246 | 0.033267 | 0.05164
| Rel.Std.Dev. |  | 2.139% | 1.689% | 3.184% | 4.396% | 4.518% | 2.184% | 3.162%
| PDM3 |  | -0.2806% | -0.5472% | -2.948% | 2.104% | -3.977% | 4.483% | 12.03%
| Z-Score (Absolute) |  | 0 | 0.07 | 0.67 | 0.6 | 0.93 | 1.2 | 3.11
| Analytical results for Fe in OREAS 285 (Certified Value 3.31 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.2937 | 3.2311 | 3.39 | 3.29 | 3.196 | 3.4 | 3.2472 | 3.26 | 3.26 | 3.28 | 3.21 | 3.19 | 3.24 | 3.411 | 3.13 | 3.15 | 3.54 | 3.15 | 3.33 | 3.37 | 3.27 | 3.28 | 3.116 | 3.448 | 3.21 | 3.23 | 1
| 1 | 2 | 3.3138 | 3.2773 | 3.34 | 3.32 | 3.2937 | 3.38 | 3.2351 | 3.41 | 3.24 | 3.34 | 3.15 | 3.16 | 3.24 | 3.3818 | 3.12 | 3.14 | 3.6 | 3.2 | 3.39 | 3.42 | 3.27 | 3.36 | 2.863 | 3.448 | 3.4 | 3.22 | 11
| 1 | 3 | 3.3215 | 3.2605 | 3.39 | 3.25 | 3.2162 | 3.34 | 3.3198 | 3.18 | 3.16 | 3.44 | 3.23 | 3.16 | 3.32 | 3.3215 | 3.15 | 3.15 | 3.55 | 3.41 | 3.62 | 3.38 | 3.34 | 3.33 | 2.991 | 3.378 | 3.54 | 3.23 | 16
| 1 | 4 | 3.2781 | 3.2998 | 3.44 | 3.24 | 3.2986 | 3.39 | 3.3058 | 3.16 | 3.22 | 3.41 | 3.29 | 3.2 | 3.24 | 3.4407 | 3.19 | 3.12 | 3.74 | 3.64 | 3.4 | 3.34 | 3.34 | 3.28 | 2.938 | 3.476 | 3.46 | 3.27 | 3.3103
| 1 | 5 | 3.3002 | 3.2776 | 3.38 | 3.34 | 3.2887 | 3.39 | 3.3228 | 3.19 | 3.24 | 3.41 | 3.17 | 3.24 | 3.29 | 3.3344 | 3.13 | 3.12 | 3.49 | 3.33 | 3.34 | 3.35 | 3.34 | 3.32 | 2.994 | 3.525 | 3.47 | 3.18 | 31
| 1 | 6 | 3.314 | 3.36 | 3.4 | 3.38 | 3.2234 | 3.37 | 3.321 | 3.23 | 3.16 | 3.43 | 3.22 | 3.33 | 3.28 | 3.3143 | 3.21 | 3.11 | 3.57 | 3.36 | 3.16 | 3.32 | 3.33 | 3.3 | 3.004 | 3.441 | 3.53 | 3.14
| Mean |  | 3.3035 | 3.2844 | 3.39 | 3.3033 | 3.2528 | 3.3783 | 3.2919 | 3.2383 | 3.2133 | 3.385 | 3.2117 | 3.2133 | 3.2683 | 3.3673 | 3.155 | 3.1317 | 3.5817 | 3.3483 | 3.3733 | 3.3633 | 3.315 | 3.3117 | 2.9843 | 3.4527 | 3.435 | 3.2117
| Median |  | 3.307 | 3.2774 | 3.39 | 3.305 | 3.2561 | 3.385 | 3.3128 | 3.21 | 3.23 | 3.41 | 3.215 | 3.195 | 3.26 | 3.3581 | 3.14 | 3.13 | 3.56 | 3.345 | 3.365 | 3.36 | 3.335 | 3.31 | 2.9925 | 3.448 | 3.465 | 3.225
| Std Dev. |  | 0.016086 | 0.043477 | 0.032249 | 0.053914 | 0.045803 | 0.02137 | 0.039994 | 0.091524 | 0.043205 | 0.062209 | 0.04916 | 0.064395 | 0.033714 | 0.051968 | 0.036742 | 0.017224 | 0.085654 | 0.1736 | 0.14855 | 0.035024 | 0.035071 | 0.031252 | 0.083342 | 0.048024 | 0.12145 | 0.04535
| Rel.Std.Dev. |  | 0.4869% | 1.324% | 0.9513% | 1.632% | 1.408% | 0.6326% | 1.215% | 2.826% | 1.345% | 1.838% | 1.531% | 2.004% | 1.032% | 1.543% | 1.165% | 0.55% | 2.391% | 5.185% | 4.404% | 1.041% | 1.058% | 0.9437% | 2.793% | 1.391% | 3.536% | 1.412%
| PDM3 |  | -0.2048% | -0.7837% | 2.407% | -0.2109% | -1.738% | 2.055% | -0.5547% | -2.174% | -2.93% | 2.256% | -2.98% | -2.93% | -1.268% | 1.721% | -4.692% | -5.397% | 8.197% | 1.149% | 1.904% | 1.602% | 0.1416% | 0.04088% | -9.847% | 4.3% | 3.767% | -2.98%
| Z-Score (Absolute) |  | 0 | 0.18 | 0.81 | 0 | 0.47 | 0.7 | 0.11 | 0.61 | 0.84 | 0.76 | 0.85 | 0.84 | 0.33 | 0.59 | 1.38 | 1.6 | 2.59 | 0.42 | 0.65 | 0.56 | 0.11 | 0.08 | 2.97 | 1.39 | 1.23 | 0.85
| Analytical results for Ga in OREAS 285 (Certified Value 21 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 20.7581 | 21.36 | 22 | 21.2421 | 21 | 23.87 | 17.54 | 19.1 | 20.4 | 20.4 | 20.5 | 21.5 | 20.0506 | 24 | 19.3 | 20.7 | 23.6 | 21.45 | 20.95 | 20.71 | 19.2 | 20 | 60.6 | 19.6 |  |  | 1
| 1 | 2 | 20.5433 | 21.61 | 22.5 | 20.7306 | 21 | 23.52 | 19.4 | 19.35 | 19.7 | 20.5 | 20.8 | 21.9 | 21.6032 | 22 | 18.8 | 20.4 | 23 | 21.38 | 21.18 | 21.28 | 19.09 | 20 | 64.5 | 19.5 |  |  | 36
| 1 | 3 | 20.3053 | 21.14 | 22.6 | 21.3384 | 21 | 23.95 | 18.74 | 18.4 | 20.7 | 21.2 | 20.4 | 22.1 | 21.2093 | 23 | 19.7 | 20 | 23.4 | 21.02 | 21.72 | 21.15 | 19.63 | 19 | 72.6 | 19 |  |  | 16
| 1 | 4 | 20.6943 | 21.22 | 23.4 | 21.7673 | 22 | 23.75 | 19.34 | 18.75 | 20.1 | 20.4 | 20.9 | 22.2 | 21.3766 | 23 | 18.9 | 19.9 | 23.2 | 21.27 | 21.57 | 21.34 | 20.66 | 20 | 65.1 | 19.3 |  |  | 20.971
| 1 | 5 | 20.7252 | 21.33 | 23.7 | 22.8476 | 21 | 23.9 | 19.65 | 19.6 | 20.4 | 20 | 21 | 22.4 | 19.8821 | 23 | 18.5 | 20.7 | 22.7 | 21.66 | 21.33 | 20.88 | 19.33 | 20 | 65.7 | 19.1 |  |  | 32
| 1 | 6 | 21.0809 | 21.34 | 23.4 | 20.7387 | 21 | 23.81 | 19.05 | 18.6 | 20.3 | 19.95 | 21.2 | 21.9 | 21.6782 | 23 | 19.3 | 20.2 | 23.6 | 20.21 | 21.39 | 21.7 | 20.11 | 20 | 66.6 | 19.2
| Mean |  | 20.6845 | 21.3333 | 22.9333 | 21.4441 | 21.1667 | 23.8 | 18.9533 | 18.9667 | 20.2667 | 20.4083 | 20.8 | 22 | 20.9667 | 23 | 19.0833 | 20.3167 | 23.25 | 21.165 | 21.3567 | 21.1767 | 19.67 | 19.8333 | 65.85 | 19.2833
| Median |  | 20.7098 | 21.335 | 23 | 21.2903 | 21 | 23.84 | 19.195 | 18.925 | 20.35 | 20.4 | 20.85 | 22 | 21.2929 | 23 | 19.1 | 20.3 | 23.3 | 21.325 | 21.36 | 21.215 | 19.48 | 20 | 65.4 | 19.25
| Std Dev. |  | 0.25613 | 0.15971 | 0.66232 | 0.79157 | 0.40825 | 0.15388 | 0.75983 | 0.46224 | 0.33862 | 0.44991 | 0.30332 | 0.30984 | 0.79419 | 0.63246 | 0.43089 | 0.34303 | 0.35637 | 0.51314 | 0.2742 | 0.35149 | 0.60798 | 0.40825 | 3.9012 | 0.23166
| Rel.Std.Dev. |  | 1.238% | 0.7486% | 2.888% | 3.691% | 1.929% | 0.6466% | 4.009% | 2.437% | 1.671% | 2.205% | 1.458% | 1.408% | 3.788% | 2.75% | 2.258% | 1.688% | 1.533% | 2.424% | 1.284% | 1.66% | 3.091% | 2.058% | 5.924% | 1.201%
| PDM3 |  | -1.366% | 1.728% | 9.358% | 2.256% | 0.9332% | 13.49% | -9.621% | -9.558% | -3.358% | -2.683% | -0.8153% | 4.907% | -0.02053% | 9.675% | -9.001% | -3.12% | 10.87% | 0.9252% | 1.839% | 0.9809% | -6.204% | -5.425% | 214% | -8.047%
| Z-Score (Absolute) |  | 0.3 | 0.21 | 1.45 | 0.29 | 0.08 | 2.13 | 1.64 | 1.63 | 0.62 | 0.51 | 0.21 | 0.73 | 0.08 | 1.5 | 1.54 | 0.58 | 1.7 | 0.08 | 0.23 | 0.09 | 1.09 | 0.96 | 34.84 | 1.39
| Analytical results for Gd in OREAS 285 (Certified Value 6.11 ppm) | [27 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 6.1273 | 6 | 6.4 | 5.9106 | 4.9 | 6.46 | 6 | [19 empty] | 1
| 1 | 2 | 5.9606 | 5.98 | 6.1 | 5.9589 | 6 | 6.36 | 6 | [19 empty] | 37
| 1 | 3 | 6.0213 | 6.21 | 6.4 | 5.7966 | 5.9 | 6.33 | 5 | [19 empty] | 16
| 1 | 4 | 6.0868 | 5.97 | 6.3 | 6.2326 | 6 | 6.62 | 6 | [19 empty] | 6.1133
| 1 | 5 | 6.2387 | 5.83 | 6.4 | 6.9061 | 5.6 | 6.5 | 6 | [19 empty] | 33
| 1 | 6 | 6.3422 | 6.12 | 6 | 6.1371 | 5.7 | 6.24 | 6
| Mean |  | 6.1295 | 6.0183 | 6.2667 | 6.157 | 5.6833 | 6.4183 | 5.8333
| Median |  | 6.107 | 5.99 | 6.35 | 6.048 | 5.8 | 6.41 | 6
| Std Dev. |  | 0.14087 | 0.13167 | 0.17512 | 0.39928 | 0.41673 | 0.13571 | 0.40825
| Rel.Std.Dev. |  | 2.298% | 2.188% | 2.794% | 6.485% | 7.333% | 2.114% | 6.999%
| PDM3 |  | 0.2641% | -1.554% | 2.508% | 0.7141% | -7.034% | 4.989% | -4.58%
| Z-Score (Absolute) |  | 0.07 | 0.68 | 0.67 | 0.07 | 2.5 | 1.49 | NA
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 285 (Indicative Value 0.27 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | 0.20276 | 1.4 | < 0.1 | 0.22 | 0.1 | 0.11 | 0.16 | 0.12 | 0.1 | 0.42 | 0.367 | 3.6 | 0.4 | 1.4 | 0.28 | 1.2 | 1.5 | 0.56 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.2 | 1.6 | < 0.1 | 0.41 | 0.25 | 0.11 | 0.11 | 0.14 | 0.13 | 0.47 | 0.38844 | 3.6 | 0.4 | 1.4 | 0.24 | 1.4 | 1.5 | 0.56 | < 0.1 | [7 empty] | 2
| 1 | 3 | 0.21037 | 1.5 | 0.1 | 0.33 | 0.05 | 0.12 | 0.13 | 0.14 | 0.11 | 0.46 | 0.28942 | 3.4 | 0.3 | 1.4 | 0.17 | 1.4 | 1.4 | 0.58 | < 0.1 | [7 empty] | 16
| 1 | 4 | 0.20866 | 1.4 | 0.1 | 0.24 | 0.28 | 0.11 | 0.14 | 0.12 | 0.11 | 0.48 | 0.50416 | 3.6 | 0.4 | 1.4 | 0.31 | 1.6 | 1.5 | 0.57 | < 0.1 | [7 empty] | 0.27114
| 1 | 5 | 0.20746 | 1.4 | 0.2 | 0.31 | 0.16 | 0.1 | 0.12 | 0.14 | 0.13 | 0.47 | 0.61565 | 3.4 | 0.3 | 1.5 | 0.54 | 1.5 | 1.4 | 0.59 | < 0.1 | [7 empty] | 8
| 1 | 6 | 0.2244 | 1.4 | 0.2 | 0.31 | 0.13 | 0.11 | 0.12 | 0.13 | 0.13 | 0.47 | 0.6007 | 3.6 | 0.4 | 1.4 | 0.48 | 1.4 | 1.4 | 0.61 | < 0.1
| Mean |  | 0.21073 | 1.45 | 0.15 | 0.30333 | 0.16167 | 0.11 | 0.13 | 0.13167 | 0.11833 | 0.46167 | 0.46089 | 3.5333 | 0.36667 | 1.4167 | 0.33667 | 1.4167 | 1.45 | 0.57833
| Median |  | 0.20866 | 1.4 | 0.15 | 0.31 | 0.145 | 0.11 | 0.125 | 0.135 | 0.12 | 0.47 | 0.4463 | 3.6 | 0.4 | 1.4 | 0.295 | 1.4 | 1.45 | 0.575
| Std Dev. |  | 0.0081462 | 0.083666 | 0.057735 | 0.06802 | 0.088412 | 0.0063246 | 0.017889 | 0.0098319 | 0.013292 | 0.02137 | 0.13329 | 0.10328 | 0.05164 | 0.040825 | 0.14348 | 0.13292 | 0.054772 | 0.019408
| Rel.Std.Dev. |  | 3.866% | 5.77% | 38.49% | 22.42% | 54.69% | 5.75% | 13.76% | 7.467% | 11.23% | 4.629% | 28.92% | 2.923% | 14.08% | 2.882% | 42.62% | 9.382% | 3.777% | 3.356%
| PDM3 |  | -22.28% | 434.8% | -44.68% | 11.87% | -40.38% | -59.43% | -52.06% | -51.44% | -56.36% | 70.27% | 69.98% | 1203% | 35.23% | 422.5% | 24.17% | 422.5% | 434.8% | 113.3%
| Z-Score (Absolute) |  | 0.45 | 3.44 | 0.68 | 0.1 | 0.54 | 0.7 | 0.64 | 0.63 | 0.67 | 0.39 | 0.38 | 9.87 | 0.09 | 3.34 | 0 | 3.34 | 3.44 | 0.75 | 0.89
| Analytical results for Hf in OREAS 285 (Certified Value 1.88 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.8666 | 1.91 | 2.2 | 1.73 | 1.61 | 1.9 | 1.8 | 1.9 | 2 | 2 | 1.9388 | 2.1 | 2.44 | 1.8 | 1.8 | 1.7 | 1.79 | 1.81 | 1.88 | 3 | 1.8 | 1.81 |  |  |  |  | 1
| 1 | 2 | 1.8673 | 1.99 | 2.3 | 1.67 | 1.77 | 1.9 | 1.9 | 1.7 | 1.9 | 2 | 1.9619 | 2 | 2.55 | 1.7 | 2 | 1.7 | 1.82 | 1.81 | 1.96 | 4 | 1.9 | 1.87 |  |  |  |  | 22
| 1 | 3 | 1.8183 | 1.9 | 2 | 1.74 | 1.8 | 1.9 | 1.9 | 1.8 | 2 | 2.1 | 1.9425 | 2 | 2.45 | 1.7 | 1.9 | 1.6 | 1.89 | 1.92 | 1.97 | 4 | 2 | 1.78 |  |  |  |  | 16
| 1 | 4 | 1.833 | 1.96 | 2.1 | 1.67 | 1.9 | 1.8 | 1.9 | 1.8 | 2.1 | 2 | 1.9692 | 1.9 | 2.33 | 1.7 | 2 | 1.7 | 1.69 | 1.91 | 1.98 | 3 | 1.9 | 1.89 |  |  |  |  | 1.8787
| 1 | 5 | 1.8281 | 1.76 | 2.2 | 1.69 | 1.78 | 1.9 | 1.9 | 1.7 | 2.1 | 2.1 | 2.0185 | 1.9 | 2.38 | 1.7 | 1.9 | 1.6 | 1.82 | 1.84 | 1.96 | 4 | 1.9 | 1.64 |  |  |  |  | 34
| 1 | 6 | 1.9043 | 1.85 | 2.2 | 1.72 | 1.75 | 1.8 | 2 | 1.7 | 2 | 2.1 | 2.0149 | 2 | 2.4 | 1.8 | 2 | 1.6 | 1.84 | 1.86 | 2.16 | 4 | 1.9 | 1.81
| Mean |  | 1.8529 | 1.895 | 2.1667 | 1.7033 | 1.7683 | 1.8667 | 1.9 | 1.7667 | 2.0167 | 2.05 | 1.9743 | 1.9833 | 2.425 | 1.7333 | 1.9333 | 1.65 | 1.8083 | 1.8583 | 1.985 | 3.6667 | 1.9 | 1.8
| Median |  | 1.8498 | 1.905 | 2.2 | 1.705 | 1.775 | 1.9 | 1.9 | 1.75 | 2 | 2.05 | 1.9655 | 2 | 2.42 | 1.7 | 1.95 | 1.65 | 1.82 | 1.85 | 1.965 | 4 | 1.9 | 1.81
| Std Dev. |  | 0.032374 | 0.082158 | 0.10328 | 0.030768 | 0.093684 | 0.05164 | 0.063246 | 0.08165 | 0.075277 | 0.054772 | 0.034802 | 0.075277 | 0.075033 | 0.05164 | 0.08165 | 0.054772 | 0.066758 | 0.047924 | 0.092898 | 0.5164 | 0.063246 | 0.088544
| Rel.Std.Dev. |  | 1.747% | 4.336% | 4.767% | 1.806% | 5.298% | 2.766% | 3.329% | 4.622% | 3.733% | 2.672% | 1.763% | 3.795% | 3.094% | 2.979% | 4.223% | 3.32% | 3.692% | 2.579% | 4.68% | 14.08% | 3.329% | 4.919%
| PDM3 |  | -1.372% | 0.867% | 15.33% | -9.335% | -5.875% | -0.6411% | 1.133% | -5.964% | 7.343% | 9.117% | 5.087% | 5.569% | 29.08% | -7.738% | 2.907% | -12.17% | -3.746% | -1.085% | 5.658% | 95.17% | 1.133% | -4.19%
| Z-Score (Absolute) |  | 0.32 | 0 | 2.04 | 1.44 | 0.95 | 0.21 | 0.04 | 0.96 | 0.91 | 1.16 | 0.59 | 0.66 | 3.97 | 1.21 | 0.29 | 1.84 | 0.65 | 0.27 | 0.67 | NA | 0.04 | 0.71
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 285 (Indicative Value 0.13 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.12 | [24 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | [24 empty] | 2
2
| 1 | 1 | < 1 | 0.14 | [24 empty] | 1
| 1 | 2 | < 1 | 0.14 | [24 empty] | 5
| 1 | 3 | < 1 | 0.13 | [24 empty] | 16
| 1 | 4 | < 1 | 0.12 | [24 empty] | 0.13333
| 1 | 5 | < 1 | 0.13 | [24 empty] | 7
| 1 | 6 | < 1 | 0.14
| Mean |  |  | 0.13333
| Median |  |  | 0.135
| Std Dev. |  |  | 0.008165
| Rel.Std.Dev. |  |  | 6.124%
| PDM3 |  |  | 2.442e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ho in OREAS 285 (Certified Value 0.62 ppm) | [27 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 0.59673 | 0.6 | 0.6 | 0.5172 | 0.6 | 0.6 | < 1 | [19 empty] | 1
| 1 | 2 | 0.61652 | 0.6 | 0.6 | 0.5514 | 0.7 | 0.61 | < 1 | [19 empty] | 23
| 1 | 3 | 0.60473 | 0.61 | 0.7 | 0.4586 | 0.7 | 0.62 | < 1 | [19 empty] | 16
| 1 | 4 | 0.62283 | 0.65 | 0.7 | 0.4622 | 0.6 | 0.62 | < 1 | [19 empty] | 0.62462
| 1 | 5 | 0.60158 | 0.59 | 0.7 | 0.4915 | 0.6 | 0.61 | < 1 | [19 empty] | 35
| 1 | 6 | 0.62425 | 0.61 | 0.7 | 0.499 | 0.6 | 0.6 | < 1
| Mean |  | 0.61111 | 0.61 | 0.66667 | 0.49665 | 0.63333 | 0.61
| Median |  | 0.61063 | 0.605 | 0.7 | 0.49525 | 0.6 | 0.61
| Std Dev. |  | 0.011642 | 0.020976 | 0.05164 | 0.03489 | 0.05164 | 0.0089443
| Rel.Std.Dev. |  | 1.905% | 3.439% | 7.746% | 7.025% | 8.154% | 1.466%
| PDM3 |  | -2.164% | -2.341% | 6.731% | -20.49% | 1.395% | -2.341%
| Z-Score (Absolute) |  | 0.03 | 0 | 1.64 | 3.28 | 0.67 | 0 | 3.18
| Analytical results for In in OREAS 285 (Certified Value 0.071 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.07 | < 0.1 | < 2 | 0.099 | 0.06 | 0.073 | 0.073 | 0.065 | 0.068 | 0.081 | 0.079636 | 0.08 | 0.07 | 0.07 | 0.068 | 0.07 | 0.074 | 0.07 | 0.07 | 0.096 | 0.09 | [5 empty] | 1
| 1 | 2 | 0.08 | < 0.1 | < 2 | 0.097 | 0.06 | 0.07 | 0.068 | 0.062 | 0.071 | 0.073 | 0.067802 | 0.1 | 0.07 | 0.07 | 0.071 | 0.07 | 0.072 | 0.07 | 0.07 | 0.093 | 0.09 | [5 empty] | 24
| 1 | 3 | 0.07 | < 0.1 | < 2 | 0.103 | 0.08 | 0.07 | 0.077 | 0.066 | 0.064 | 0.08 | 0.063096 | 0.1 | 0.07 | 0.07 | 0.072 | 0.07 | 0.07 | 0.07 | 0.07 | 0.079 | 0.1 | [5 empty] | 16
| 1 | 4 | 0.07 | < 0.1 | < 2 | 0.095 | 0.07 | 0.072 | 0.075 | 0.069 | 0.066 | 0.078 | 0.066706 | 0.09 | 0.07 | 0.07 | 0.066 | 0.07 | 0.073 | 0.06 | 0.07 | 0.088 | 0.09 | [5 empty] | 0.070619
| 1 | 5 | 0.06 | < 0.1 | < 2 | 0.093 | 0.08 | 0.075 | 0.073 | 0.071 | 0.064 | 0.078 | 0.070006 | 0.09 | 0.07 | 0.07 | 0.071 | 0.07 | 0.073 | 0.07 | 0.08 | 0.094 | 0.09 | [5 empty] | 36
| 1 | 6 | 0.07 | < 0.1 | < 2 | 0.101 | 0.07 | 0.068 | 0.075 | 0.066 | 0.065 | 0.077 | 0.080904 | 0.11 | 0.08 | 0.07 | 0.071 | 0.06 | 0.069 | 0.07 | 0.07 | 0.109 | 0.08
| Mean |  | 0.07 |  |  | 0.098 | 0.07 | 0.071333 | 0.0735 | 0.0665 | 0.066333 | 0.077833 | 0.071358 | 0.095 | 0.071667 | 0.07 | 0.069833 | 0.068333 | 0.071833 | 0.068333 | 0.071667 | 0.093167 | 0.09
| Median |  | 0.07 |  |  | 0.098 | 0.07 | 0.071 | 0.074 | 0.066 | 0.0655 | 0.078 | 0.068904 | 0.095 | 0.07 | 0.07 | 0.071 | 0.07 | 0.0725 | 0.07 | 0.07 | 0.0935 | 0.09
| Std Dev. |  | 0.0063246 |  |  | 0.0037417 | 0.0089443 | 0.0025033 | 0.0030822 | 0.0031464 | 0.0027325 | 0.0027869 | 0.0072668 | 0.010488 | 0.0040825 | 0 | 0.0023166 | 0.0040825 | 0.0019408 | 0.0040825 | 0.0040825 | 0.0098675 | 0.0063246
| Rel.Std.Dev. |  | 9.035% |  |  | 3.818% | 12.78% | 3.509% | 4.193% | 4.731% | 4.119% | 3.581% | 10.18% | 11.04% | 5.696% | 0% | 3.317% | 5.974% | 2.702% | 5.974% | 5.696% | 10.59% | 7.027%
| PDM3 |  | -0.8772% |  |  | 38.77% | -0.8772% | 1.011% | 4.079% | -5.833% | -6.069% | 10.22% | 1.046% | 34.52% | 1.483% | -0.8772% | -1.113% | -3.237% | 1.719% | -3.237% | 1.483% | 31.93% | 27.44%
| Z-Score (Absolute) |  | 0.3 | 4.76 | 206.99 | 5.94 | 0.3 | 0.01 | 0.48 | 1.08 | 1.12 | 1.44 | 0 | 5.27 | 0.07 | 0.3 | 0.34 | 0.67 | 0.11 | 0.67 | 0.07 | 4.86 | 4.16
| Analytical results for K in OREAS 285 (Certified Value 2.94 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.739 | 2.8304 | 2.9998 | 3.14 | 2.9322 | 2.9 | 2.7904 | 2.48 | 2.95 | 2.9 | 2.88 | 2.82 | 3.03 | 2.887 | 2.92 | 2.99 | 2.87 | 2.71 | 2.91 | 3.03 | 2.8888 | 2.8892 |  | 3.072 | 1.95 | 2.95 | 1
| 1 | 2 | 2.8083 | 2.8558 | 2.9582 | 3.11 | 3.0219 | 2.87 | 2.7708 | 1.92 | 2.94 | 3 | 2.82 | 2.83 | 3.03 | 2.8873 | 2.9 | 3.01 | 3.09 | 2.77 | 2.88 | 3.07 | 2.889 | 2.9376 | 2.902 | 3.08 | 2.1 | 2.97 | #N/A
| 1 | 3 | 2.7977 | 2.8617 | 3.0035 | 3.19 | 2.9615 | 2.85 | 2.8363 | 2.14 | 2.88 | 3.07 | 2.92 | 2.81 | 3.1 | 2.8343 | 2.91 | 3.03 | 2.96 | 2.79 | 3 | 3.04 | 2.9339 | 2.9567 | 2.89 | 3.038 | 1.95 | 2.99 | 16
| 1 | 4 | 2.7472 | 2.9004 | 3.0325 | 3.1 | 3.0439 | 2.88 | 2.8207 | 1.87 | 2.92 | 3.03 | 2.95 | 2.87 | 3.01 | 2.8937 | 2.87 | 2.99 | 3.07 | 2.77 | 2.96 | 2.99 | 2.9617 | 2.9007 | 2.916 | 3.088 | 1.74 | 3.01 | 2.937
| 1 | 5 | 2.806 | 2.8656 | 2.9875 | 3.09 | 3.0304 | 2.9 | 2.8336 | 2.03 | 2.95 | 3.08 | 2.85 | 2.9 | 3.08 | 2.8517 | 2.84 | 2.98 | 3.07 | 2.76 | 2.93 | 3.04 | 2.9536 | 2.8862 | 3.074 | 3.088 | 1.84 | 2.87 | 37
| 1 | 6 | 2.8094 | 2.9113 | 3.0069 | 2.8 | 2.9528 | 2.91 | 2.8182 | 1.94 | 2.88 | 3.07 | 2.89 | 2.9 | 3.05 | 2.8343 | 2.89 | 2.97 | 3.05 | 2.68 | 2.89 | 2.99 | 2.9224 | 2.9219 | 2.899 | 3.08 | 1.93 | 2.89
| Mean |  | 2.7846 | 2.8709 | 2.9981 | 3.0717 | 2.9905 | 2.885 | 2.8117 | 2.0633 | 2.92 | 3.025 | 2.885 | 2.855 | 3.05 | 2.8647 | 2.8883 | 2.995 | 3.0183 | 2.7467 | 2.9283 | 3.0267 | 2.9249 | 2.9154 | 2.9362 | 3.0743 | 1.9183 | 2.9467
| Median |  | 2.8018 | 2.8637 | 3.0016 | 3.105 | 2.9917 | 2.89 | 2.8194 | 1.985 | 2.93 | 3.05 | 2.885 | 2.85 | 3.04 | 2.8694 | 2.895 | 2.99 | 3.06 | 2.765 | 2.92 | 3.035 | 2.9282 | 2.9113 | 2.902 | 3.08 | 1.94 | 2.96
| Std Dev. |  | 0.032499 | 0.029931 | 0.024484 | 0.1379 | 0.047097 | 0.022583 | 0.025824 | 0.22527 | 0.032863 | 0.068337 | 0.046797 | 0.040373 | 0.034059 | 0.027797 | 0.029269 | 0.021679 | 0.085888 | 0.042269 | 0.04535 | 0.031411 | 0.031168 | 0.02827 | 0.077596 | 0.018779 | 0.1209 | 0.055737
| Rel.Std.Dev. |  | 1.167% | 1.043% | 0.8166% | 4.489% | 1.575% | 0.7828% | 0.9185% | 10.92% | 1.125% | 2.259% | 1.622% | 1.414% | 1.117% | 0.9703% | 1.013% | 0.7239% | 2.846% | 1.539% | 1.549% | 1.038% | 1.066% | 0.9697% | 2.643% | 0.6108% | 6.302% | 1.892%
| PDM3 |  | -5.188% | -2.251% | 2.08% | 4.586% | 1.821% | -1.77% | -4.267% | -29.75% | -0.578% | 2.997% | -1.77% | -2.791% | 3.848% | -2.46% | -1.656% | 1.976% | 2.77% | -6.48% | -0.2943% | 3.054% | -0.4112% | -0.7352% | -0.0264% | 4.677% | -34.68% | 0.33%
| Z-Score (Absolute) |  | 1.32 | 0.49 | 0.73 | 1.43 | 0.65 | 0.36 | 1.06 | 8.24 | 0.02 | 0.98 | 0.36 | 0.65 | 1.22 | 0.55 | 0.33 | 0.7 | 0.92 | 1.69 | 0.06 | 1 | 0.02 | 0.07 | 0.13 | 1.46 | 9.63 | 0.23
| Analytical results for La in OREAS 285 (Certified Value 37.8 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 34.2615 | 38.84 | 37.7 | 36.589 | 36 | 28.2 | 38.4 | 35.5 | 35.7 | 39.6 | 36.7 | 37 | 45 | 34.9 | 45.6 | 38.5 | 39.63 | 40 | 38.96 | 39 | 25.9 | 30.9 |  |  |  |  | 1
| 1 | 2 | 34.8941 | 38.56 | 37.2 | 37.1831 | 35 | 27.7 | 39.4 | 36 | 36.2 | 36.4 | 35 | 36.9 | 43 | 33.3 | 46.1 | 39.4 | 38.09 | 38 | 39.53 | 39 | 27.4 | 29.7 |  |  |  |  | 25
| 1 | 3 | 34.7526 | 39.63 | 38.2 | 37.6957 | 36 | 28.8 | 38.6 | 36 | 38 | 40.7 | 34.2 | 38.5 | 42 | 34.7 | 44.8 | 39.1 | 37.95 | 40 | 39.27 | 38 | 28.1 | 28.7 |  |  |  |  | 16
| 1 | 4 | 35.016 | 38.21 | 38.8 | 38.8675 | 36 | 28.8 | 37.7 | 35.4 | 38.6 | 40.8 | 36.6 | 38.4 | 41 | 34.5 | 42.4 | 37.8 | 37.9 | 39 | 39.28 | 41 | 27.3 | 27.7 |  |  |  |  | 37.8146
| 1 | 5 | 34.111 | 38.67 | 37.7 | 40.7598 | 34 | 28.9 | 37.1 | 39.6 | 37.5 | 38.5 | 38.1 | 38.6 | 38 | 33 | 43.5 | 40.3 | 40 | 40 | 38.4 | 41 | 27.5 | 30.6 |  |  |  |  | 38
| 1 | 6 | 35.8656 | 38.79 | 38.5 | 36.3752 | 35 | 28.7 | 37.8 | 36 | 39.3 | 37.9 | 38.5 | 38.5 | 43 | 33.9 | 43.4 | 36.8 | 37.7 | 39 | 39.62 | 40 | 28.1 | 28
| Mean |  | 34.8168 | 38.7833 | 38.0167 | 37.9117 | 35.3333 | 28.5167 | 38.1667 | 36.4167 | 37.55 | 38.9833 | 36.5167 | 37.9833 | 42 | 34.05 | 44.3 | 38.65 | 38.545 | 39.3333 | 39.1767 | 39.6667 | 27.3833 | 29.2667
| Median |  | 34.8234 | 38.73 | 37.95 | 37.4394 | 35.5 | 28.75 | 38.1 | 36 | 37.75 | 39.05 | 36.65 | 38.45 | 42.5 | 34.2 | 44.15 | 38.8 | 38.02 | 39.5 | 39.275 | 39.5 | 27.45 | 29.2
| Std Dev. |  | 0.6255 | 0.47149 | 0.59133 | 1.6561 | 0.8165 | 0.47081 | 0.80664 | 1.5829 | 1.3867 | 1.7151 | 1.6822 | 0.80353 | 2.3664 | 0.7791 | 1.4311 | 1.2373 | 0.99853 | 0.8165 | 0.44518 | 1.2111 | 0.80602 | 1.3426
| Rel.Std.Dev. |  | 1.797% | 1.216% | 1.555% | 4.368% | 2.311% | 1.651% | 2.113% | 4.347% | 3.693% | 4.4% | 4.607% | 2.115% | 5.634% | 2.288% | 3.23% | 3.201% | 2.591% | 2.076% | 1.136% | 3.053% | 2.943% | 4.588%
| PDM3 |  | -7.928% | 2.562% | 0.5343% | 0.2567% | -6.562% | -24.59% | 0.9309% | -3.697% | -0.6998% | 3.091% | -3.432% | 0.4461% | 11.07% | -9.956% | 17.15% | 2.209% | 1.931% | 4.016% | 3.602% | 4.898% | -27.59% | -22.6%
| Z-Score (Absolute) |  | 1.52 | 0.38 | 0.01 | 0.04 | 1.28 | 4.54 | 0.08 | 0.76 | 0.22 | 0.47 | 0.71 | 0.01 | 1.92 | 1.89 | 3.02 | 0.31 | 0.26 | 0.64 | 0.56 | 0.8 | 5.08 | 4.18
| Analytical results for Li in OREAS 285 (Certified Value 53 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 51.5404 | 54.3 | 53 | 80.354 | 56 | 48 | 42 | 54.8 | 54.7 | 51.5 | 51.6 | 55.1 | 53.0656 | 56.9 | 54 | 48.9 | 49.8 | 53.5 | 53.7 | 53 | 53.3 | 45.19 | 52 | 50.4 | 67 |  | 1
| 1 | 2 | 51.7029 | 54.4 | 55 | 81.8544 | 57 | 47 | 49 | 54.4 | 56 | 50.7 | 51.7 | 54.2 | 54.0755 | 56.9 | 55 | 48.3 | 47.4 | 56 | 54.5 | 53 | 54.6 |  | 53 | 52 | 68 |  | 45
| 1 | 3 | 51.8435 | 54.5 | 54 | 81.6934 | 57 | 49 | 44.2 | 53.4 | 57.8 | 52.1 | 51.6 | 55.7 | 54.2162 | 58.3 | 55 | 47.6 | 49.6 | 56 | 53.8 | 54 | 53.5 | 44.13 | 52 | 56.5 | 67 |  | 16
| 1 | 4 | 53.4744 | 54.2 | 55 | 84.3462 | 57 | 48 | 47.2 | 54.1 | 57.6 | 52.7 | 52.6 | 54.3 | 53.534 | 54.9 | 54 | 48.4 | 48.2 | 55.1 | 53.2 | 54 | 52.9 | 44.76 | 54 | 54.5 | 67 |  | 53.3174
| 1 | 5 | 52.397 | 54 | 53 | 83.2058 | 54 | 49 | 45.6 | 54.9 | 57.5 | 51.4 | 53 | 55.4 | 53.2443 | 57.8 | 54 | 48.7 | 48.9 | 55.3 | 53.4 | 54 | 53.1 | 45.86 | 54 | 53.9 | 64 |  | 39
| 1 | 6 | 52.3636 | 54.1 | 53 | 80.8686 | 54 | 48 | 43.8 | 53.7 | 57.2 | 51.9 | 52.3 | 55.1 | 52.9338 | 57.6 | 54 | 49.3 | 50.5 | 51.2 | 52.7 | 54 | 53.5 | 45.39 | 53 | 55 | 66
| Mean |  | 52.2203 | 54.25 | 53.8333 | 82.0537 | 55.8333 | 48.1667 | 45.3 | 54.2167 | 56.8 | 51.7167 | 52.1333 | 54.9667 | 53.5116 | 57.0667 | 54.3333 | 48.5333 | 49.0667 | 54.5167 | 53.55 | 53.6667 | 53.4833 | 45.066 | 53 | 53.7167 | 66.5
| Median |  | 52.1035 | 54.25 | 53.5 | 81.7739 | 56.5 | 48 | 44.9 | 54.25 | 57.35 | 51.7 | 52 | 55.1 | 53.3892 | 57.25 | 54 | 48.55 | 49.25 | 55.2 | 53.55 | 54 | 53.4 | 45.19 | 53 | 54.2 | 67
| Std Dev. |  | 0.7066 | 0.18708 | 0.98319 | 1.4863 | 1.472 | 0.75277 | 2.5195 | 0.59805 | 1.2116 | 0.6824 | 0.59217 | 0.59889 | 0.53275 | 1.1911 | 0.5164 | 0.58195 | 1.1343 | 1.8649 | 0.6091 | 0.5164 | 0.5947 | 0.65584 | 0.89443 | 2.1904 | 1.3784
| Rel.Std.Dev. |  | 1.353% | 0.3449% | 1.826% | 1.811% | 2.636% | 1.563% | 5.562% | 1.103% | 2.133% | 1.319% | 1.136% | 1.09% | 0.9956% | 2.087% | 0.9504% | 1.199% | 2.312% | 3.421% | 1.137% | 0.9622% | 1.112% | 1.455% | 1.688% | 4.078% | 2.073%
| PDM3 |  | -2.058% | 1.749% | 0.9677% | 53.9% | 4.719% | -9.66% | -15.04% | 1.687% | 6.532% | -3.002% | -2.221% | 3.093% | 0.3642% | 7.032% | 1.905% | -8.973% | -7.972% | 2.249% | 0.4363% | 0.6551% | 0.3112% | -15.48% | -0.5953% | 0.7489% | 24.72%
| Z-Score (Absolute) |  | 0.67 | 0.27 | 0.08 | 13.23 | 1.01 | 2.56 | 3.9 | 0.26 | 1.46 | 0.91 | 0.71 | 0.61 | 0.07 | 1.59 | 0.31 | 2.39 | 2.14 | 0.4 | 0.05 | 0 | 0.09 | 4.01 | 0.31 | 0.02 | 5.98
| Analytical results for Lu in OREAS 285 (Certified Value 0.17 ppm) | [27 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.21 | 1.25 | 1.26 | 1.32 | [17 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | [17 empty] | 2
3
| 1 | 1 | 0.16465 | 0.17 | 0.2 | 0.1575 | 0.1 | 0.19 | 0.18 | 0.16 | 0.15 | [17 empty] | 1
| 1 | 2 | 0.16752 | 0.16 | 0.2 | 0.1306 | 0.2 | 0.2 | 0.18 | 0.16 | 0.15 | [17 empty] | 26
| 1 | 3 | 0.17116 | 0.17 | 0.2 | 0.1238 | 0.2 | 0.19 | 0.17 | 0.15 | 0.15 | [17 empty] | 16
| 1 | 4 | 0.17095 | 0.17 | 0.2 | 0.126 | 0.2 | 0.2 | 0.17 | 0.17 | 0.16 | [17 empty] | 0.16918
| 1 | 5 | 0.17453 | 0.17 | 0.2 | 0.1385 | 0.1 | 0.18 | 0.18 | 0.16 | 0.15 | [17 empty] | 40
| 1 | 6 | 0.17176 | 0.16 | 0.2 | 0.1319 | 0.2 | 0.19 | 0.17 | 0.16 | 0.15
| Mean |  | 0.17009 | 0.16667 | 0.2 | 0.13472 | 0.16667 | 0.19167 | 0.175 | 0.16 | 0.15167
| Median |  | 0.17106 | 0.17 | 0.2 | 0.13125 | 0.2 | 0.19 | 0.175 | 0.16 | 0.15
| Std Dev. |  | 0.0034802 | 0.005164 | 3.0405e-17 | 0.012275 | 0.05164 | 0.0075277 | 0.0054772 | 0.0063246 | 0.0040825
| Rel.Std.Dev. |  | 2.046% | 3.098% | 1.52e-14% | 9.112% | 30.98% | 3.928% | 3.13% | 3.953% | 2.692%
| PDM3 |  | 0.5387% | -1.487% | 18.22% | -20.37% | -1.487% | 13.29% | 3.439% | -5.427% | -10.35%
| Z-Score (Absolute) |  | 0.28 | 0 | NA | 2.59 | NA | 2.02 | 0.67 | 0.54 | 1.21
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 285 (Certified Value 0.925 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.97284 | 0.88192 | 0.9231 | 0.92 | 0.8572 | 0.89 | 0.9205 | 0.93 | 0.88 | 0.9 | 0.87 | 0.86 | 0.92 | 0.90848 | 0.93 | 0.82 | 0.981 | 0.92 | 0.91 | 0.931 | 0.9207 | 0.9151 | 0.983 | 1.001 | 0.92 | 0.94 | 1
| 1 | 2 | 0.98366 | 0.89908 | 0.903 | 0.91 | 0.8859 | 0.88 | 0.909 | 0.97 | 0.87 | 0.91 | 0.84 | 0.86 | 0.92 | 0.91689 | 0.91 | 0.82 | 0.985 | 0.94 | 0.91 | 0.946 | 0.9213 | 0.9419 |  | 1.007 | 0.985 | 0.96 | #N/A
| 1 | 3 | 0.97997 | 0.89669 | 0.9239 | 0.95 | 0.8674 | 0.91 | 0.9337 | 0.89 | 0.86 | 0.94 | 0.88 | 0.85 | 0.94 | 0.89758 | 0.91 | 0.83 | 0.997 | 0.93 | 0.84 | 0.935 | 0.9358 | 0.9495 | 0.912 | 0.983 | 1.02 | 0.95 | 16
| 1 | 4 | 0.97826 | 0.90955 | 0.9358 | 0.93 | 0.8934 | 0.92 | 0.9335 | 0.89 | 0.87 | 0.94 | 0.89 | 0.86 | 0.91 | 0.90988 | 0.88 | 0.82 | 0.99 | 0.95 | 0.91 | 0.92 | 0.9444 | 0.9548 | 0.957 | 1.013 | 0.99 | 0.94 | 0.92527
| 1 | 5 | 0.97326 | 0.89988 | 0.8955 | 0.96 | 0.8844 | 0.88 | 0.9352 | 0.9 | 0.88 | 0.94 | 0.86 | 0.88 | 0.93 | 0.90964 | 0.88 | 0.82 | 0.992 | 0.93 | 0.9 | 0.935 | 0.9418 | 0.9243 | 0.991 | 1.013 | 1 | 0.93 | 41
| 1 | 6 | 0.97666 | 0.91669 | 0.9284 | 0.97 | 0.8716 | 0.87 | 0.9284 | 0.91 | 0.86 | 0.94 | 0.87 | 0.93 | 0.92 | 0.89846 | 0.9 | 0.82 | 0.985 | 0.92 | 0.87 | 0.92 | 0.9326 | 0.9376 | 0.938 | 1.007 | 1.01 | 0.9
| Mean |  | 0.97744 | 0.90063 | 0.91828 | 0.94 | 0.87665 | 0.89167 | 0.92672 | 0.915 | 0.87 | 0.92833 | 0.86833 | 0.87333 | 0.92333 | 0.90682 | 0.90167 | 0.82167 | 0.98833 | 0.93167 | 0.89 | 0.93117 | 0.93277 | 0.9372 | 0.9562 | 1.004 | 0.9875 | 0.93667
| Median |  | 0.97746 | 0.89948 | 0.9235 | 0.94 | 0.878 | 0.885 | 0.93095 | 0.905 | 0.87 | 0.94 | 0.87 | 0.86 | 0.92 | 0.90906 | 0.905 | 0.82 | 0.9875 | 0.93 | 0.905 | 0.933 | 0.9342 | 0.93975 | 0.957 | 1.007 | 0.995 | 0.94
| Std Dev. |  | 0.0041252 | 0.011883 | 0.015599 | 0.023664 | 0.013532 | 0.019408 | 0.010219 | 0.030822 | 0.0089443 | 0.018348 | 0.017224 | 0.029439 | 0.010328 | 0.0074382 | 0.019408 | 0.0040825 | 0.005785 | 0.01169 | 0.028983 | 0.0099883 | 0.010032 | 0.015087 | 0.032461 | 0.011225 | 0.035461 | 0.020656
| Rel.Std.Dev. |  | 0.422% | 1.319% | 1.699% | 2.517% | 1.544% | 2.177% | 1.103% | 3.369% | 1.028% | 1.976% | 1.984% | 3.371% | 1.119% | 0.8203% | 2.152% | 0.4969% | 0.5853% | 1.255% | 3.256% | 1.073% | 1.076% | 1.61% | 3.395% | 1.118% | 3.591% | 2.205%
| PDM3 |  | 5.639% | -2.662% | -0.7547% | 1.592% | -5.254% | -3.631% | 0.1568% | -1.11% | -5.973% | 0.3315% | -6.153% | -5.613% | -0.2089% | -1.993% | -2.551% | -11.2% | 6.816% | 0.6918% | -3.811% | 0.6377% | 0.8107% | 1.29% | 3.343% | 8.509% | 6.726% | 1.232%
| Z-Score (Absolute) |  | 1.48 | 0.69 | 0.19 | 0.42 | 1.37 | 0.94 | 0.05 | 0.28 | 1.55 | 0.09 | 1.6 | 1.46 | 0.05 | 0.51 | 0.66 | 2.92 | 1.79 | 0.19 | 0.99 | 0.17 | 0.22 | 0.34 | 0.88 | 2.23 | 1.76 | 0.33
| Analytical results for Mn in OREAS 285 (Certified Value 0.038 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.037595 | 0.0391 | 0.0378 | 0.0538 | 0.042 | 0.036 | 0.0383 | 0.0373 | 0.0384 | 0.0369 | 0.0368 | 0.0392 | 0.039625 | 0.0377 | 0.0368 | 0.0388 | 0.0375 | 0.0386 | 0.0375 | 0.0393 | 0.0389 | 0.02446 | 0.039 | 0.0468 | 0.04 |  | 1
| 1 | 2 | 0.03825 | 0.0384 | 0.0391 | 0.0485 | 0.0416 | 0.0355 | 0.04 | 0.0371 | 0.0394 | 0.0363 | 0.0371 | 0.0392 | 0.038074 | 0.0384 | 0.0366 | 0.0382 | 0.0383 | 0.0393 | 0.0382 | 0.0396 | 0.0406 | 0.023216 | 0.039 | 0.0497 | 0.04 |  | 27
| 1 | 3 | 0.03787 | 0.0391 | 0.0377 | 0.0565 | 0.0427 | 0.0361 | 0.037 | 0.0367 | 0.0404 | 0.0375 | 0.0367 | 0.0397 | 0.039832 | 0.0374 | 0.0375 | 0.0369 | 0.0377 | 0.0422 | 0.0378 | 0.0403 | 0.0401 | 0.024959 | 0.039 | 0.0602 | 0.04 |  | 16
| 1 | 4 | 0.038251 | 0.0395 | 0.0384 | 0.0615 | 0.0423 | 0.0363 | 0.0371 | 0.037 | 0.0401 | 0.038 | 0.0372 | 0.039 | 0.03876 | 0.0378 | 0.0372 | 0.0386 | 0.0385 | 0.0396 | 0.0373 | 0.0406 | 0.0393 | 0.023655 | 0.039 | 0.0511 | 0.04 |  | 0.038493
| 1 | 5 | 0.037592 | 0.0389 | 0.0395 | 0.0546 | 0.0413 | 0.0363 | 0.0379 | 0.0376 | 0.0405 | 0.0369 | 0.0381 | 0.0398 | 0.038048 | 0.0369 | 0.0359 | 0.0384 | 0.0381 | 0.0384 | 0.0377 | 0.0402 | 0.0393 |  | 0.039 | 0.0509 | 0.04 |  | 42
| 1 | 6 | 0.038286 | 0.0392 | 0.0397 | 0.053 | 0.0417 | 0.036 | 0.0383 | 0.0368 | 0.0401 | 0.0368 | 0.0394 | 0.039 | 0.03913 | 0.0406 | 0.0368 | 0.0383 | 0.0377 | 0.0367 | 0.0374 | 0.04 | 0.0393 | 0.024017 | 0.039 | 0.0517 | 0.04
| Mean |  | 0.037974 | 0.039033 | 0.0387 | 0.05465 | 0.041933 | 0.036033 | 0.0381 | 0.037083 | 0.039817 | 0.037067 | 0.03755 | 0.039317 | 0.038911 | 0.038133 | 0.0368 | 0.0382 | 0.037967 | 0.039133 | 0.03765 | 0.04 | 0.039583 | 0.024061 | 0.039 | 0.051733 | 0.04
| Median |  | 0.03806 | 0.0391 | 0.03875 | 0.0542 | 0.04185 | 0.03605 | 0.0381 | 0.03705 | 0.0401 | 0.0369 | 0.03715 | 0.0392 | 0.038945 | 0.03775 | 0.0368 | 0.03835 | 0.0379 | 0.03895 | 0.0376 | 0.0401 | 0.0393 | 0.024017 | 0.039 | 0.051 | 0.04
| Std Dev. |  | 0.00033188 | 0.00036697 | 0.00086023 | 0.0042805 | 0.00050859 | 0.00029439 | 0.00109 | 0.00033116 | 0.00079352 | 0.00059554 | 0.001033 | 0.00034881 | 0.00075834 | 0.0013049 | 0.00054772 | 0.00067231 | 0.00039328 | 0.0018107 | 0.00032711 | 0.00047749 | 0.00063377 | 0.00067941 | 0 | 0.0045001 | 0
| Rel.Std.Dev. |  | 0.874% | 0.9401% | 2.223% | 7.833% | 1.213% | 0.817% | 2.861% | 0.893% | 1.993% | 1.607% | 2.751% | 0.8872% | 1.949% | 3.422% | 1.488% | 1.76% | 1.036% | 4.627% | 0.8688% | 1.194% | 1.601% | 2.824% | 0% | 8.699% | 0%
| PDM3 |  | -1.348% | 1.404% | 0.5379% | 41.97% | 8.938% | -6.39% | -1.021% | -3.662% | 3.439% | -3.705% | -2.45% | 2.14% | 1.087% | -0.9343% | -4.398% | -0.7611% | -1.367% | 1.664% | -2.19% | 3.915% | 2.833% | -37.49% | 1.317% | 34.4% | 3.915%
| Z-Score (Absolute) |  | 0.47 | 0.21 | 0 | 10.24 | 2.08 | 1.71 | 0.39 | 1.04 | 0.72 | 1.05 | 0.74 | 0.4 | 0.14 | 0.36 | 1.22 | 0.32 | 0.47 | 0.28 | 0.67 | 0.83 | 0.57 | 9.4 | 0.01 | 8.37 | 0.83
| Analytical results for Mo in OREAS 285 (Certified Value 2.71 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.5795 | 2.8 | 3 | 3.0823 | 3 | 3.1 | 2.16 | 2.73 | 2.76 | 2.75 | 2.61 | 2.7 | 2.4825 | 2.4 | 2.74 | 2.7 | 3.45 | 2.6 | 2.54 | 3 | 2.7 | 4 | 3.02 | 2.99 |  |  | 1
| 1 | 2 | 2.6938 | 2.7 | 2 | 3.0236 | 3 | 2.9 | 2.43 | 2.52 | 2.54 | 2.8 | 2.56 | 2.85 | 2.4305 | 2.1 | 2.76 | 2.7 | 2.95 | 2.7 | 2.57 | 3 | 2.7 | 4 | 3.27 | 2.5 |  |  | 28
| 1 | 3 | 2.5656 | 2.6 | 2 | 1.803 | 3 | 3.1 | 2.18 | 2.6 | 2.54 | 2.57 | 2.77 | 2.58 | 2.8161 | 2.1 | 2.79 | 2.7 | 2.88 | 2.4 | 2.66 | 3 | 2.7 | 4 | 3.32 | 2.4 |  |  | 16
| 1 | 4 | 2.6342 | 2.7 | 2 | 2.3506 | 3 | 2.9 | 2.36 | 2.88 | 2.79 | 2.77 | 2.53 | 2.54 | 2.7646 | 2.3 | 2.77 | 2.7 | 2.61 | 2.8 | 2.76 | 3 | 2.8 | 4 | 3.31 | 2.8 |  |  | 2.7066
| 1 | 5 | 2.657 | 2.7 | 2 | 2.6768 | 3 | 2.9 | 2.69 | 2.4 | 2.72 | 2.85 | 3.04 | 2.45 | 2.4394 | 2.1 | 2.83 | 2.8 | 2.99 | 2.5 | 2.95 | 3 | 2.7 | 5 | 3.57 | 2.42 |  |  | 43
| 1 | 6 | 2.6842 | 2.9 | 1 | 2.086 | 3 | 3 | 2.35 | 2.65 | 2.95 | 2.86 | 2.88 | 2.78 | 2.5455 | 2.1 | 2.7 | 2.8 | 2.76 | 2.2 | 2.72 | 3 | 2.7 | 5 | 3.96 | 2.58
| Mean |  | 2.6357 | 2.7333 | 2 | 2.5037 | 3 | 2.9833 | 2.3617 | 2.63 | 2.7167 | 2.7667 | 2.7317 | 2.65 | 2.5798 | 2.1833 | 2.765 | 2.7333 | 2.94 | 2.5333 | 2.7 | 3 | 2.7167 | 4.3333 | 3.4083 | 2.615
| Median |  | 2.6456 | 2.7 | 2 | 2.5137 | 3 | 2.95 | 2.355 | 2.625 | 2.74 | 2.785 | 2.69 | 2.64 | 2.514 | 2.1 | 2.765 | 2.7 | 2.915 | 2.55 | 2.69 | 3 | 2.7 | 4 | 3.315 | 2.54
| Std Dev. |  | 0.053381 | 0.10328 | 0.63246 | 0.51457 | 0 | 0.098319 | 0.19302 | 0.16661 | 0.15756 | 0.10558 | 0.20173 | 0.15258 | 0.16891 | 0.13292 | 0.044159 | 0.05164 | 0.28552 | 0.21602 | 0.14873 | 0 | 0.040825 | 0.5164 | 0.32183 | 0.2339
| Rel.Std.Dev. |  | 2.025% | 3.779% | 31.62% | 20.55% | 0% | 3.296% | 8.173% | 6.335% | 5.8% | 3.816% | 7.385% | 5.758% | 6.548% | 6.088% | 1.597% | 1.889% | 9.711% | 8.527% | 5.508% | 0% | 1.503% | 11.92% | 9.443% | 8.945%
| PDM3 |  | -2.618% | 0.9889% | -26.11% | -7.495% | 10.84% | 10.23% | -12.74% | -2.829% | 0.3731% | 2.22% | 0.9273% | -2.09% | -4.685% | -19.33% | 2.159% | 0.9889% | 8.625% | -6.401% | -0.2427% | 10.84% | 0.3731% | 60.1% | 25.93% | -3.383%
| Z-Score (Absolute) |  | 0.65 | 0.22 | NA | 1.83 | NA | 2.46 | 3.1 | 0.7 | 0.07 | 0.52 | 0.21 | 0.52 | 1.15 | 4.69 | 0.51 | 0.22 | 2.07 | 1.56 | 0.07 | NA | 0.07 | NA | 6.25 | 0.83
Results from laboratories 1.09, 1.11, 1.26 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 285 (Certified Value 2.02 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.9534 | 1.9417 | 2.0244 | 2.03 | 1.9504 | 1.97 | 2.1389 | 2.113 | 2.02 | 1.99 | 1.98 | 1.9 | 2.08 | 1.9266 | 1.94 | 2.07 | 1.74 | 1.92 | 2.061 | 2.032 | 2.0114 | 2.0031 |  | 2.077 | 1.91 | 2.32 | 1
| 1 | 2 | 1.9805 | 1.9628 | 1.9916 | 2.05 | 2.0039 | 1.95 | 2.1189 | 2.198 | 2.01 | 2.05 | 1.95 | 1.91 | 2.08 | 1.9786 | 1.95 | 2.06 | 1.88 | 1.93 | 2.012 | 2.062 | 2.0187 | 2.0536 | 1.645 | 2.077 | 2.04 | 2.34 | #N/A
| 1 | 3 | 1.9544 | 1.9481 | 2.0191 | 2.04 | 1.9943 | 1.94 | 2.1764 | 2.065 | 1.97 | 2.12 | 2 | 1.9 | 2.13 | 1.9437 | 1.94 | 2.08 | 1.7 | 1.95 | 2.07 | 2.043 | 2.0451 | 2.0822 | 1.604 | 2.062 | 2.12 | 2.34 | 16
| 1 | 4 | 1.957 | 1.9982 | 2.0368 | 2.04 | 2.0204 | 2.03 | 2.1723 | 2.038 | 1.99 | 2.09 | 2.02 | 1.92 | 2.07 | 1.9771 | 1.94 | 2.06 | 1.87 | 1.99 | 2.068 | 2.007 | 2.0711 | 2.0413 | 1.552 | 2.099 | 2.08 | 2.31 | 2.0214
| 1 | 5 | 1.9531 | 1.9439 | 2.0155 | 2.12 | 2.0312 | 1.96 | 2.1763 | 2.085 | 2.02 | 2.1 | 1.98 | 1.96 | 2.11 | 1.9509 | 1.92 | 2.08 | 1.71 | 1.94 | 2.026 | 2.037 | 2.0592 | 2.0158 | 1.8 | 2.092 | 2.08 | 2.25 | 44
| 1 | 6 | 1.9595 | 1.9714 | 2.0323 | 2.17 | 1.9785 | 1.94 | 2.1725 | 2.087 | 1.97 | 2.09 | 2 | 1.94 | 2.09 | 1.9481 | 1.94 | 2.06 | 1.72 | 1.92 | 1.986 | 2.004 | 2.0342 | 2.0501 | 1.608 | 2.099 | 2.1 | 2.23
| Mean |  | 1.9596 | 1.961 | 2.0199 | 2.075 | 1.9965 | 1.965 | 2.1592 | 2.0977 | 1.9967 | 2.0733 | 1.9883 | 1.9217 | 2.0933 | 1.9542 | 1.9383 | 2.0683 | 1.77 | 1.9417 | 2.0372 | 2.0308 | 2.04 | 2.041 | 1.6418 | 2.0843 | 2.055 | 2.2983
| Median |  | 1.9557 | 1.9554 | 2.0217 | 2.045 | 1.9991 | 1.955 | 2.1724 | 2.086 | 2 | 2.09 | 1.99 | 1.915 | 2.085 | 1.9495 | 1.94 | 2.065 | 1.73 | 1.935 | 2.0435 | 2.0345 | 2.0396 | 2.0457 | 1.608 | 2.0845 | 2.08 | 2.315
| Std Dev. |  | 0.010479 | 0.021567 | 0.016003 | 0.056833 | 0.029271 | 0.033912 | 0.024384 | 0.055142 | 0.023381 | 0.046762 | 0.024014 | 0.024014 | 0.022509 | 0.020199 | 0.0098319 | 0.0098319 | 0.082462 | 0.026394 | 0.034563 | 0.022122 | 0.023096 | 0.028322 | 0.094436 | 0.014801 | 0.075829 | 0.047081
| Rel.Std.Dev. |  | 0.5347% | 1.1% | 0.7923% | 2.739% | 1.466% | 1.726% | 1.129% | 2.629% | 1.171% | 2.255% | 1.208% | 1.25% | 1.075% | 1.034% | 0.5072% | 0.4754% | 4.659% | 1.359% | 1.697% | 1.089% | 1.132% | 1.388% | 5.752% | 0.7101% | 3.69% | 2.049%
| PDM3 |  | -3.054% | -2.987% | -0.07112% | 2.652% | -1.234% | -2.79% | 6.819% | 3.774% | -1.223% | 2.57% | -1.635% | -4.933% | 3.559% | -3.325% | -4.109% | 2.322% | -12.44% | -3.944% | 0.7806% | 0.4673% | 0.9183% | 0.9711% | -18.78% | 3.114% | 1.663% | 13.7%
| Z-Score (Absolute) |  | 0.74 | 0.73 | 0.06 | 0.56 | 0.33 | 0.68 | 1.51 | 0.82 | 0.32 | 0.54 | 0.42 | 1.17 | 0.77 | 0.8 | 0.98 | 0.48 | 2.88 | 0.94 | 0.13 | 0.06 | 0.16 | 0.18 | 4.33 | 0.67 | 0.33 | 3.08
| Analytical results for Nb in OREAS 285 (Certified Value 12.7 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 12.6214 | 13.13 | 5.9 | 15.8193 | 13 | 16.78 | 10.59 | 13.3 | 12.3 | 12.5 | 11.8 | 14.2 | 12.5774 | 13.6 | 13.1 | 11.4 | 13.6 | 11.8 | 13.1 | 12.04 | 12.86 | 10 | 12.8 | 12 |  |  | 1
| 1 | 2 | 12.943 | 13.14 | 5.6 | 10.2183 | 13 | 14.7 | 12.12 | 13.6 | 12.9 | 12.5 | 12.2 | 14 | 12.7351 | 13.9 | 12.6 | 11 | 13 | 12 | 12.9 | 12.33 | 13.03 | 10 | 14.1 | 12.4 |  |  | 29
| 1 | 3 | 12.4242 | 13.07 | 3.7 | 17.8138 | 13 | 14.56 | 11.58 | 13 | 13.6 | 12.9 | 11.8 | 14.4 | 12.4557 | 13.2 | 13.3 | 11.5 | 13.2 | 12.2 | 12.7 | 12.36 | 12.89 | 9 | 13.4 | 12.1 |  |  | 16
| 1 | 4 | 12.8951 | 13.19 | 3.9 | 10.5034 | 13 | 14.43 | 12.47 | 13.3 | 13.3 | 12.6 | 12.2 | 14 | 12.8961 | 14.1 | 13 | 11.2 | 13.2 | 12 | 12.7 | 12.57 | 12.81 | 10 | 13.2 | 12.1 |  |  | 12.7362
| 1 | 5 | 13.533 | 13.05 | 3.2 | 12.74 | 13 | 15.34 | 12.42 | 13.8 | 13.2 | 12.5 | 12.4 | 14.4 | 12.8759 | 13.8 | 12.6 | 11.2 | 13.1 | 11.9 | 13.3 | 12.49 | 13.06 | 10 | 13.4 | 11.4 |  |  | 45
| 1 | 6 | 13.0263 | 12.92 | 1.9 | 8.4897 | 13 | 15.73 | 11.66 | 12.9 | 13.5 | 12.2 | 12.4 | 14 | 12.4398 | 13.7 | 12.8 | 11 | 13.5 | 11.4 | 13 | 12.44 | 13.08 | 11 | 13.6 | 11.9
| Mean |  | 12.9072 | 13.0833 | 4.0333 | 12.5974 | 13 | 15.2567 | 11.8067 | 13.3167 | 13.1333 | 12.5333 | 12.1333 | 14.1667 | 12.6633 | 13.7167 | 12.9 | 11.2167 | 13.2667 | 11.8833 | 12.95 | 12.3717 | 12.955 | 10 | 13.4167 | 11.9833
| Median |  | 12.9191 | 13.1 | 3.8 | 11.6217 | 13 | 15.02 | 11.89 | 13.3 | 13.25 | 12.5 | 12.2 | 14.1 | 12.6563 | 13.75 | 12.9 | 11.2 | 13.2 | 11.95 | 12.95 | 12.4 | 12.96 | 10 | 13.4 | 12.05
| Std Dev. |  | 0.38006 | 0.094587 | 1.5042 | 3.5925 | 0 | 0.89779 | 0.7025 | 0.34303 | 0.4761 | 0.22509 | 0.27325 | 0.19664 | 0.20247 | 0.30605 | 0.28284 | 0.20412 | 0.23381 | 0.27142 | 0.23452 | 0.18433 | 0.11537 | 0.63246 | 0.43089 | 0.33116
| Rel.Std.Dev. |  | 2.945% | 0.723% | 37.29% | 28.52% | 0% | 5.885% | 5.95% | 2.576% | 3.625% | 1.796% | 2.252% | 1.388% | 1.599% | 2.231% | 2.193% | 1.82% | 1.762% | 2.284% | 1.811% | 1.49% | 0.8905% | 6.325% | 3.212% | 2.764%
| PDM3 |  | 1.342% | 2.726% | -68.33% | -1.09% | 2.071% | 19.79% | -7.298% | 4.558% | 3.118% | -1.593% | -4.734% | 11.23% | -0.5721% | 7.698% | 1.286% | -11.93% | 4.165% | -6.696% | 1.679% | -2.862% | 1.718% | -21.48% | 5.343% | -5.911%
| Z-Score (Absolute) |  | 0.01 | 0.26 | 12.92 | 0.45 | NA | 3.43 | 1.6 | 0.6 | 0.33 | 0.54 | 1.12 | 1.84 | 0.35 | 1.18 | 0.01 | 2.46 | 0.53 | 1.49 | 0.07 | 0.77 | 0.07 | NA | 0.75 | 1.34
Results from laboratories 1.11 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 285 (Certified Value 35.1 ppm) | [27 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 1
2
| 1 | 1 | 33.5817 | 36.13 | 35.1 | 32.7606 | 31.3 | 38.1 | 44 | [19 empty] | 1
| 1 | 2 | 33.6457 | 35.88 | 34 | 33.5045 | 35.7 | 37.32 | 44 | [19 empty] | 6
| 1 | 3 | 33.7336 | 37.25 | 33.8 | 32.4038 | 36.8 | 37.86 | 44 | [19 empty] | 16
| 1 | 4 | 34.5304 | 36.03 | 34.7 | 34.1625 | 36.1 | 37.13 | 47 | [19 empty] | 35.0719
| 1 | 5 | 33.5664 | 34.92 | 34.9 | 36.7371 | 33.6 | 39.11 | 47 | [19 empty] | 46
| 1 | 6 | 35.2897 | 35.99 | 36.3 | 33.6268 | 33.6 | 36.87 | 45
| Mean |  | 34.0579 | 36.0333 | 34.8 | 33.8659 | 34.5167 | 37.7317 | 45.1667
| Median |  | 33.6896 | 36.01 | 34.8 | 33.5657 | 34.65 | 37.59 | 44.5
| Std Dev. |  | 0.70486 | 0.74223 | 0.89443 | 1.5414 | 2.0566 | 0.81558 | 1.472
| Rel.Std.Dev. |  | 2.07% | 2.06% | 2.57% | 4.551% | 5.958% | 2.162% | 3.259%
| PDM3 |  | -2.891% | 2.741% | -0.7752% | -3.439% | -1.583% | 7.584% | 28.78%
| Z-Score (Absolute) |  | 0.54 | 0.89 | 0 | 0.67 | 0.2 | 2.12 | 7.48
| Analytical results for Ni in OREAS 285 (Certified Value 19.7 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 1
2
| 1 | 1 | 18.76 | 19.2506 | 20.3 | 21 | 23.5177 | 21 | 18.9 | 18.1 | 20 | 18.4 | 19.8 | 19.9 | 19 | 20.1563 | 16.4 | 18.6 | 20 | 19.3 | 20.1 | 18 | 20 | 19.9 | 18.35 | 25 | 19 | 20.7 | 1
| 1 | 2 | 19.2 | 19.3361 | 20 | 20 | 21.135 | 21 | 18.9 | 19 | 19.9 | 19.7 | 19.6 | 19.9 | 19.2 | 20.0718 | 16.3 | 18.7 | 20 | 19.5 | 21.7 | 19 | 20 | 20.7 | 17.76 | 23 | 18.1 | 20.7 | 16
| 1 | 3 | 18.53 | 19.2147 | 20.1 | 21 | 22.1467 | 20 | 19.4 | 17.5 | 19.5 | 20 | 20.7 | 19.8 | 19.3 | 19.8694 | 16.1 | 18.7 | 19 | 19 | 21 | 19 | 20 | 20.7 | 18.21 | 25 | 19.5 | 20.7 | 16
| 1 | 4 | 19.78 | 20.082 | 20 | 21 | 23.3576 | 21 | 19.4 | 17.5 | 19.4 | 20 | 20.1 | 20 | 19.4 | 20.2518 | 16.7 | 18.8 | 20 | 19.6 | 20.3 | 18 | 20 | 19.9 | 17.84 | 26 | 20.4 | 20.6 | 19.7083
| 1 | 5 | 19.37 | 20.5717 | 20.5 | 21 | 20.2703 | 22 | 19.5 | 17.7 | 20.2 | 19.8 | 19.6 | 20.1 | 19.6 | 19.8585 | 16.7 | 18.7 | 20 | 19.4 | 19.8 | 19 | 20 | 19.8 | 17.64 | 23 | 19.3 | 20.2 | 47
| 1 | 6 | 19.67 | 19.6636 | 20.4 | 20 | 22.4392 | 21 | 19.5 | 18 | 19.4 | 20.2 | 19.3 | 20.5 | 19.4 | 19.7403 | 16.9 | 18.7 | 20 | 19.3 | 18.6 | 18 | 20 | 19.9 | 17.71 | 25 | 19.7 | 19.7
| Mean |  | 19.2183 | 19.6864 | 20.2167 | 20.6667 | 22.1444 | 21 | 19.2667 | 17.9667 | 19.7333 | 19.6833 | 19.85 | 20.0333 | 19.3167 | 19.9913 | 16.5167 | 18.7 | 19.8333 | 19.35 | 20.25 | 18.5 | 20 | 20.15 | 17.9183 | 24.5 | 19.3333 | 20.4333
| Median |  | 19.285 | 19.4999 | 20.2 | 21 | 22.2929 | 21 | 19.4 | 17.85 | 19.7 | 19.9 | 19.7 | 19.95 | 19.35 | 19.9706 | 16.55 | 18.7 | 20 | 19.35 | 20.2 | 18.5 | 20 | 19.9 | 17.8 | 25 | 19.4 | 20.65
| Std Dev. |  | 0.49531 | 0.5433 | 0.2137 | 0.5164 | 1.263 | 0.63246 | 0.28752 | 0.56451 | 0.34448 | 0.65243 | 0.49295 | 0.25033 | 0.20412 | 0.19851 | 0.29944 | 0.063246 | 0.40825 | 0.20736 | 1.0597 | 0.54772 | 0 | 0.42778 | 0.29103 | 1.2247 | 0.76594 | 0.40825
| Rel.Std.Dev. |  | 2.577% | 2.76% | 1.057% | 2.499% | 5.703% | 3.012% | 1.492% | 3.142% | 1.746% | 3.315% | 2.483% | 1.25% | 1.057% | 0.993% | 1.813% | 0.3382% | 2.058% | 1.072% | 5.233% | 2.961% | 0% | 2.123% | 1.624% | 4.999% | 3.962% | 1.998%
| PDM3 |  | -2.486% | -0.1109% | 2.579% | 4.863% | 12.36% | 6.554% | -2.241% | -8.837% | 0.127% | -0.1267% | 0.7189% | 1.649% | -1.987% | 1.436% | -16.19% | -5.116% | 0.6344% | -1.818% | 2.749% | -6.131% | 1.48% | 2.241% | -9.082% | 24.31% | -1.903% | 3.679%
| Z-Score (Absolute) |  | 0.82 | 0.14 | 0.63 | 1.28 | 3.41 | 1.76 | 0.75 | 2.62 | 0.07 | 0.14 | 0.1 | 0.36 | 0.67 | 0.3 | 4.72 | 1.57 | 0.07 | 0.63 | 0.67 | 1.85 | 0.31 | 0.53 | 2.69 | 6.82 | 0.65 | 0.94
| Analytical results for P in OREAS 285 (Certified Value 0.084 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.084463 | 0.0867 | 0.077 | 0.0806 | 0.0811 | 0.0819 | 0.084 | 0.085 | 0.085 | 0.082 | 0.082 | 0.087 | 0.088479 | 0.083 | 0.0807 | 0.089 | 0.0785 | 0.085 | 0.083 | 0.0814 | 0.0805 | 0.087 | 0.069 | 0.0878 |  |  | 1
| 1 | 2 | 0.085523 | 0.0856 | 0.072 | 0.0839 | 0.081 | 0.0812 | 0.087 | 0.084 | 0.086 | 0.08 | 0.082 | 0.086 | 0.089254 | 0.082 | 0.0807 | 0.088 | 0.0807 | 0.087 | 0.085 | 0.082 | 0.0883 | 0.087 | 0.0767 | 0.0883 |  |  | 17
| 1 | 3 | 0.084932 | 0.0869 | 0.074 | 0.0814 | 0.08 | 0.0825 | 0.08 | 0.083 | 0.089 | 0.083 | 0.082 | 0.088 | 0.085738 | 0.082 | 0.0799 | 0.088 | 0.0767 | 0.086 | 0.084 | 0.0876 | 0.0858 | 0.087 | 0.072 | 0.0886 |  |  | 16
| 1 | 4 | 0.087696 | 0.088 | 0.078 | 0.0832 | 0.0807 | 0.0839 | 0.08 | 0.084 | 0.088 | 0.084 | 0.084 | 0.087 | 0.089307 | 0.083 | 0.0809 | 0.086 | 0.081 | 0.087 | 0.081 | 0.0832 | 0.086 | 0.087 | 0.078 | 0.0897 |  |  | 0.083839
| 1 | 5 | 0.085684 | 0.0874 | 0.079 | 0.0826 | 0.0815 | 0.083 | 0.081 | 0.085 | 0.088 | 0.082 | 0.085 | 0.087 | 0.087368 | 0.081 | 0.0799 | 0.088 | 0.08 | 0.085 | 0.084 | 0.0849 | 0.086 | 0.092 | 0.071 | 0.0846 |  |  | 48
| 1 | 6 | 0.087208 | 0.0869 | 0.083 | 0.0813 | 0.0802 | 0.0825 | 0.082 | 0.083 | 0.088 | 0.082 | 0.088 | 0.087 | 0.086563 | 0.083 | 0.0822 | 0.089 | 0.0782 | 0.08 | 0.082 | 0.0844 | 0.0881 | 0.087 | 0.08 | 0.0856
| Mean |  | 0.085918 | 0.086917 | 0.077167 | 0.082167 | 0.08075 | 0.0825 | 0.082333 | 0.084 | 0.087333 | 0.082167 | 0.083833 | 0.087 | 0.087785 | 0.082333 | 0.080717 | 0.088 | 0.079183 | 0.085 | 0.083167 | 0.083917 | 0.085783 | 0.087833 | 0.07445 | 0.087433
| Median |  | 0.085603 | 0.0869 | 0.0775 | 0.082 | 0.08085 | 0.0825 | 0.0815 | 0.084 | 0.088 | 0.082 | 0.083 | 0.087 | 0.087923 | 0.0825 | 0.0807 | 0.088 | 0.07925 | 0.0855 | 0.0835 | 0.0838 | 0.086 | 0.087 | 0.07435 | 0.08805
| Std Dev. |  | 0.0012748 | 0.00079854 | 0.0038687 | 0.0012691 | 0.00056833 | 0.00092304 | 0.0027325 | 0.00089443 | 0.0015055 | 0.0013292 | 0.0024014 | 0.00063246 | 0.0014706 | 0.0008165 | 0.00084479 | 0.0010954 | 0.0016654 | 0.0026077 | 0.001472 | 0.0022489 | 0.0028181 | 0.0020412 | 0.0043835 | 0.0019377
| Rel.Std.Dev. |  | 1.484% | 0.9187% | 5.013% | 1.545% | 0.7038% | 1.119% | 3.319% | 1.065% | 1.724% | 1.618% | 2.864% | 0.727% | 1.675% | 0.9917% | 1.047% | 1.245% | 2.103% | 3.068% | 1.77% | 2.68% | 3.285% | 2.324% | 5.888% | 2.216%
| PDM3 |  | 2.479% | 3.671% | -7.959% | -1.995% | -3.684% | -1.597% | -1.796% | 0.192% | 4.168% | -1.995% | -0.006793% | 3.77% | 4.706% | -1.796% | -3.724% | 4.963% | -5.553% | 1.385% | -0.802% | 0.0926% | 2.319% | 4.764% | -11.2% | 4.287%
| Z-Score (Absolute) |  | 0.54 | 0.81 | 1.78 | 0.45 | 0.83 | 0.36 | 0.41 | 0.03 | 0.92 | 0.45 | 0.01 | 0.83 | 1.04 | 0.41 | 0.84 | 1.09 | 1.24 | 0.3 | 0.19 | 0.01 | 0.51 | 1.09 | 2.5 | 0.94
| Analytical results for Pb in OREAS 285 (Certified Value 23.1 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 23.8681 | 25.2 | 21 | 13.4569 | 21 | 21 | 21.19 | 23.1 | 24.2 | 22.5 | 21.3 | 23.3 | 22.7038 | 21.7 | 23.1 | 24 | 26.2 | 25.2 | 25.5 | 21 | 24 | 33 | 17.6 | 23.1 |  |  | 1
| 1 | 2 | 23.7412 | 24.7 | 22 | 14.8957 | 21 | 22 | 23.89 | 22.9 | 23.8 | 22.8 | 20.9 | 23.4 | 24.4253 | 21.5 | 21.8 | 24 | 26.1 | 26.1 | 25.1 | 22 | 24.4 | 29 | 22.3 | 22.7 |  |  | 18
| 1 | 3 | 23.8854 | 25.1 | 22 | 16.9233 | 21 | 21 | 22.8 | 21.9 | 24.6 | 23.6 | 20.9 | 23.8 | 23.6958 | 22.2 | 23 | 23 | 25.8 | 25.3 | 25.2 | 22 | 24.2 | 32 | 23.3 | 22.4 |  |  | 16
| 1 | 4 | 22.9428 | 24.8 | 22 | 13.8554 | 22 | 21 | 24.84 | 22.5 | 24.2 | 23.6 | 21.2 | 23.2 | 24.5935 | 22.6 | 22.1 | 24 | 24.9 | 25 | 24.4 | 22 | 24.2 | 36 | 19.5 | 23.1 |  |  | 23.1096
| 1 | 5 | 23.2454 | 24.4 | 22 | 14.159 | 22 | 20 | 23.91 | 22.9 | 24.7 | 22.8 | 21.5 | 23.3 | 22.3624 | 22.1 | 21.3 | 24 | 25.5 | 25.2 | 25.3 | 22 | 24.2 | 29 | 19.3 | 22.4 |  |  | 49
| 1 | 6 | 23.9399 | 24.8 | 23 | 16.5242 | 21 | 21 | 22.98 | 22.2 | 24.3 | 22.8 | 22 | 23.7 | 23.534 | 22.7 | 23.3 | 25 | 25.8 | 24.4 | 24.4 | 22 | 24.1 | 28 | 20.7 | 22.3
| Mean |  | 23.6038 | 24.8333 | 22 | 14.9691 | 21.3333 | 21 | 23.2683 | 22.5833 | 24.3 | 23.0167 | 21.3 | 23.45 | 23.5525 | 22.1333 | 22.4333 | 24 | 25.7167 | 25.2 | 24.9833 | 21.8333 | 24.1833 | 31.1667 | 20.45 | 22.6667
| Median |  | 23.8046 | 24.8 | 22 | 14.5274 | 21 | 21 | 23.435 | 22.7 | 24.25 | 22.8 | 21.25 | 23.35 | 23.6149 | 22.15 | 22.55 | 24 | 25.8 | 25.2 | 25.15 | 22 | 24.2 | 30.5 | 20.1 | 22.55
| Std Dev. |  | 0.41144 | 0.28752 | 0.63246 | 1.444 | 0.5164 | 0.63246 | 1.2562 | 0.46655 | 0.32249 | 0.46655 | 0.41473 | 0.2429 | 0.89476 | 0.4761 | 0.81404 | 0.63246 | 0.47081 | 0.54772 | 0.47081 | 0.40825 | 0.13292 | 3.0605 | 2.0955 | 0.36148
| Rel.Std.Dev. |  | 1.743% | 1.158% | 2.875% | 9.647% | 2.421% | 3.012% | 5.399% | 2.066% | 1.327% | 2.027% | 1.947% | 1.036% | 3.799% | 2.151% | 3.629% | 2.635% | 1.831% | 2.174% | 1.885% | 1.87% | 0.5496% | 9.82% | 10.25% | 1.595%
| PDM3 |  | 2.138% | 7.459% | -4.801% | -35.23% | -7.686% | -9.129% | 0.6869% | -2.277% | 5.151% | -0.4022% | -7.831% | 1.473% | 1.916% | -4.225% | -2.926% | 3.853% | 11.28% | 9.046% | 8.108% | -5.523% | 4.646% | 34.86% | -11.51% | -1.917%
| Z-Score (Absolute) |  | 0.27 | 0.99 | 0.67 | 4.79 | 1.06 | 1.26 | 0.07 | 0.33 | 0.68 | 0.07 | 1.08 | 0.18 | 0.24 | 0.59 | 0.42 | 0.5 | 1.51 | 1.21 | 1.08 | 0.77 | 0.61 | 4.71 | 1.58 | 0.28
| Analytical results for Pr in OREAS 285 (Certified Value 9.16 ppm) | [27 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 8.7396 | 8.74 | 9.2 | 9.407 | 8.1 | 9.67 | 7 | [19 empty] | 1
| 1 | 2 | 8.9982 | 8.7 | 9 | 9.6558 | 9 | 9.59 | 7 | [19 empty] | 4
| 1 | 3 | 8.7097 | 8.94 | 9.2 | 9.6971 | 9.1 | 9.78 | 8 | [19 empty] | 16
| 1 | 4 | 8.8819 | 8.66 | 9.3 | 10.0217 | 8.9 | 9.71 | 7 | [19 empty] | 9.1557
| 1 | 5 | 8.8482 | 8.88 | 9.4 | 10.4981 | 8.3 | 9.96 | 8 | [19 empty] | 50
| 1 | 6 | 9.3809 | 8.71 | 9.3 | 9.3724 | 8.3 | 9.5 | 8
| Mean |  | 8.9264 | 8.7717 | 9.2333 | 9.7753 | 8.6167 | 9.7017 | 7.5
| Median |  | 8.865 | 8.725 | 9.25 | 9.6764 | 8.6 | 9.69 | 7.5
| Std Dev. |  | 0.24564 | 0.11179 | 0.13663 | 0.42462 | 0.43089 | 0.15943 | 0.54772
| Rel.Std.Dev. |  | 2.752% | 1.274% | 1.48% | 4.344% | 5.001% | 1.643% | 7.303%
| PDM3 |  | -2.504% | -4.194% | 0.8479% | 6.768% | -5.887% | 5.963% | -18.08%
| Z-Score (Absolute) |  | 0.27 | 0.54 | 0.27 | 1.22 | 0.81 | 1.09 | NA
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 285 (Certified Value 168 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 0
0
| 1 | 1 | 156.2393 | 177.68 | 154 | 180.5 | 110.7 | 161.5 | 143 | 165.5 | 154.5 | 164.5 | 162.6891 | 161.2 | 169.9 | 161 | 185 | 164.9 | 173 | 180.07 | 170.86 | 118 | 151.3 | [5 empty] | 1
| 1 | 2 | 158.5652 | 175.47 | 159 | 176 | 101.1 | 161.5 | 164.5 | 153 | 157.5 | 164 | 174.8535 | 166.3 | 165.2 | 160 | 183 | 167.7 | 173 | 180.91 | 175.69 | 129 | 147.7 | [5 empty] | 34
| 1 | 3 | 156.9932 | 174.1 | 162 | 184.7 | 120.1 | 159.5 | 173 | 167 | 151.5 | 166 | 175.0472 | 163.9 | 169.9 | 165 | 190 | 165.6 | 168 | 181.54 | 173.14 | 122 | 145.6 | [5 empty] | 16
| 1 | 4 | 161.2645 | 175.92 | 160 | 182.9 | 107.6 | 158 | 170 | 164.5 | 161 | 164.5 | 177.4702 | 165.3 | 169.3 | 163 | 182 | 167.6 | 170 | 184.88 | 175.85 | 109 | 144.8 | [5 empty] | 168.2898
| 1 | 5 | 157.0388 | 175.4 | 177 | 183.4 | 110.5 | 169 | 171.5 | 159.5 | 162.5 | 164.5 | 164.6593 | 151.6 | 162.3 | 166 | 181 | 164.4 | 175 | 185.81 | 172.58 | 117 | 145.8 | [5 empty] | 51
| 1 | 6 | 163.164 | 175.34 | 185 | 182.2 | 93 | 157 | 171 | 158.5 | 168.5 | 165.5 | 170.4386 | 170.2 | 168.7 | 162 | 188 | 155 | 166 | 183.69 | 173.65 | 117 | 141.6
| Mean |  | 158.8775 | 175.6517 | 166.1667 | 181.6167 | 107.1667 | 161.0833 | 165.5 | 161.3333 | 159.25 | 164.8333 | 170.8596 | 163.0833 | 167.55 | 162.8333 | 184.8333 | 164.2 | 170.8333 | 182.8167 | 173.6283 | 118.6667 | 146.1333
| Median |  | 157.802 | 175.435 | 161 | 182.55 | 109.05 | 160.5 | 170.5 | 162 | 159.25 | 164.5 | 172.6461 | 164.6 | 169 | 162.5 | 184 | 165.25 | 171.5 | 182.615 | 173.395 | 117.5 | 145.7
| Std Dev. |  | 2.759 | 1.1651 | 12.0568 | 3.0812 | 9.258 | 4.2827 | 11.4018 | 5.2978 | 6.0807 | 0.75277 | 6.0426 | 6.3559 | 3.115 | 2.3166 | 3.5449 | 4.7104 | 3.4303 | 2.3147 | 1.9072 | 6.5929 | 3.221
| Rel.Std.Dev. |  | 1.737% | 0.6633% | 7.256% | 1.697% | 8.639% | 2.659% | 6.889% | 3.284% | 3.818% | 0.4567% | 3.537% | 3.897% | 1.859% | 1.423% | 1.918% | 2.869% | 2.008% | 1.266% | 1.098% | 5.556% | 2.204%
| PDM3 |  | -5.593% | 4.374% | -1.262% | 7.919% | -36.32% | -4.282% | -1.658% | -4.134% | -5.372% | -2.054% | 1.527% | -3.094% | -0.4396% | -3.242% | 9.83% | -2.43% | 1.511% | 8.632% | 3.172% | -29.49% | -13.17%
| Z-Score (Absolute) |  | 0.67 | 1.22 | 0.15 | 1.9 | 6.53 | 0.42 | 0.08 | 0.4 | 0.63 | 0 | 0.68 | 0.2 | 0.31 | 0.23 | 2.26 | 0.07 | 0.68 | 2.04 | 1 | 5.23 | 2.12
| Analytical results for Re in OREAS 285 (Indicative Value 0.002 ppm) | [27 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.24 | 1.26 | 1.31 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [9 empty] | 3
3
| 1 | 1 | < 0.01 | < 0.002 | 0.008 | 0.012 | < 0.002 | 0.003 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | 0.002 | [9 empty] | 1
| 1 | 2 | < 0.01 | < 0.002 | 0.006 | 0.011 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | < 0.002 | [9 empty] | 2
| 1 | 3 | < 0.01 | < 0.002 | 0.008 | 0.011 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | 0.002 | [9 empty] | 16
| 1 | 4 | < 0.01 | < 0.002 | 0.007 | 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.002 | [9 empty] | 0.0015238
| 1 | 5 | < 0.01 | < 0.002 | 0.008 | 0.011 | < 0.002 | < 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | 0.002 | [9 empty] | 8
| 1 | 6 | < 0.01 | < 0.002 | 0.01 | 0.011 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.05 | < 0.005 | < 0.05 | < 0.002 | < 0.005 | < 0.002 | 0.002
| Mean |  |  |  | 0.0078333 | 0.011 |  | 0.003 | 0.002 | 0.002 | 0.002 | 0.002 | [5 empty] | 0.003 | 0.002
| Median |  |  |  | 0.008 | 0.011 |  | 0.003 | 0.002 | 0.002 | 0.002 | 0.002 | [5 empty] | 0.003 | 0.002
| Std Dev. |  |  |  | 0.0013292 | 0.00063246 |  |  | 0 |  | 0 | 0 | [6 empty] | 0
| Rel.Std.Dev. |  |  |  | 16.97% | 5.75% |  |  | 0% |  | 0% | 0% | [6 empty] | 0%
| PDM3 |  |  |  | 414.1% | 621.9% |  | 96.88% | 31.25% | 31.25% | 31.25% | 31.25% | [5 empty] | 96.88% | 31.25%
| Z-Score (Absolute) |  | 3.2 | 0.84 | 6.07 | 9.27 | 0.84 | 0.51 | 0.51 | 0.67 | 0.17 | 0.17 | 23.43 | 0.67 | 23.43 | 0.84 | 0.67 | 0.51 | 0
| Analytical results for S in OREAS 285 (Certified Value 0.339 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.31257 | 0.34123 | 0.3394 | 0.31 | 0.3823 | 0.4 | 0.3174 | 0.34 | 0.34 | 0.34 | 0.33 | 0.34 | 0.35 | 0.32836 | 0.31 | 0.34 | 0.357 | 0.34 | 0.3 | 0.35 | 0.34 | 0.37 | 0.386 | 0.32 | 0.35 |  | 1
| 1 | 2 | 0.31867 | 0.34485 | 0.335 | 0.31 | 0.3717 | 0.4 | 0.3134 | 0.36 | 0.34 | 0.35 | 0.33 | 0.34 | 0.35 | 0.33449 | 0.31 | 0.34 | 0.352 | 0.35 | 0.3 | 0.361 | 0.35 | 0.37 | 0.391 | 0.321 | 0.35 |  | 11
| 1 | 3 | 0.31001 | 0.34539 | 0.3413 | 0.29 | 0.3949 | 0.41 | 0.3219 | 0.33 | 0.33 | 0.36 | 0.34 | 0.34 | 0.36 | 0.32859 | 0.31 | 0.34 | 0.349 | 0.34 | 0.3 | 0.353 | 0.32 | 0.37 | 0.354 | 0.34 | 0.34 |  | 16
| 1 | 4 | 0.31691 | 0.35441 | 0.3459 | 0.3 | 0.4235 | 0.41 | 0.3228 | 0.33 | 0.34 | 0.36 | 0.34 | 0.35 | 0.36 | 0.33404 | 0.31 | 0.34 | 0.353 | 0.36 | 0.3 | 0.344 | 0.35 | 0.37 | 0.362 | 0.35 | 0.35 |  | 0.33866
| 1 | 5 | 0.31504 | 0.34703 | 0.3376 | 0.3 | 0.393 | 0.4 | 0.3215 | 0.33 | 0.34 | 0.36 | 0.33 | 0.35 | 0.36 | 0.32954 | 0.31 | 0.34 | 0.351 | 0.35 | 0.3 | 0.358 | 0.34 | 0.37 | 0.381 | 0.35 | 0.33 |  | 52
| 1 | 6 | 0.31629 | 0.35437 | 0.3423 | 0.32 | 0.3796 | 0.39 | 0.3204 | 0.33 | 0.33 | 0.36 | 0.34 | 0.36 | 0.36 | 0.32871 | 0.32 | 0.34 | 0.35 | 0.34 | 0.3 | 0.348 | 0.34 | 0.39 | 0.382 | 0.33 | 0.33
| Mean |  | 0.31492 | 0.34788 | 0.34025 | 0.305 | 0.39083 | 0.40167 | 0.31957 | 0.33667 | 0.33667 | 0.355 | 0.335 | 0.34667 | 0.35667 | 0.33062 | 0.31167 | 0.34 | 0.352 | 0.34667 | 0.3 | 0.35233 | 0.34 | 0.37333 | 0.376 | 0.33517 | 0.34167
| Median |  | 0.31566 | 0.34621 | 0.34035 | 0.305 | 0.38765 | 0.4 | 0.32095 | 0.33 | 0.34 | 0.36 | 0.335 | 0.345 | 0.36 | 0.32912 | 0.31 | 0.34 | 0.3515 | 0.345 | 0.3 | 0.3515 | 0.34 | 0.37 | 0.3815 | 0.335 | 0.345
| Std Dev. |  | 0.0031484 | 0.0053863 | 0.0038088 | 0.010488 | 0.018185 | 0.0075277 | 0.0035512 | 0.012111 | 0.005164 | 0.0083666 | 0.0054772 | 0.008165 | 0.005164 | 0.002854 | 0.0040825 | 0 | 0.0028284 | 0.008165 | 0 | 0.0063456 | 0.010954 | 0.008165 | 0.014601 | 0.013571 | 0.0098319
| Rel.Std.Dev. |  | 0.9998% | 1.548% | 1.119% | 3.439% | 4.653% | 1.874% | 1.111% | 3.597% | 1.534% | 2.357% | 1.635% | 2.355% | 1.448% | 0.8632% | 1.31% | 0% | 0.8035% | 2.355% | 0% | 1.801% | 3.222% | 2.187% | 3.883% | 4.049% | 2.878%
| PDM3 |  | -7.01% | 2.724% | 0.471% | -9.938% | 15.41% | 18.61% | -5.637% | -0.5871% | -0.5871% | 4.826% | -1.079% | 2.366% | 5.319% | -2.372% | -7.969% | 0.3972% | 3.941% | 2.366% | -11.41% | 4.039% | 0.3972% | 10.24% | 11.03% | -1.03% | 0.8893%
| Z-Score (Absolute) |  | 1.45 | 0.44 | 0 | 2.02 | 2.9 | 3.52 | 1.19 | 0.21 | 0.21 | 0.85 | 0.3 | 0.37 | 0.94 | 0.55 | 1.64 | 0.01 | 0.67 | 0.37 | 2.31 | 0.69 | 0.01 | 1.9 | 2.05 | 0.29 | 0.08
| Analytical results for Sb in OREAS 285 (Certified Value 0.53 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.89509 | 0.55 | 0.4 | < 5 | 0.7 | 0.42 | 0.5 | 0.47 | 0.52 | 0.5 | 0.56 | 0.52744 | 1.1 | 0.52 | 0.5 | 0.58 | 0.5 | 0.54 | 0.48 | 0.51 | 6.98 | 0.76 | 0.56 |  |  |  | 1
| 1 | 2 | 0.81628 | 0.55 | 0.5 | < 5 | 0.54 | 0.44 | 0.54 | 0.49 | 0.51 | 0.53 | 0.56 | 0.58808 | 0.8 | 0.55 | 0.5 | 0.51 | 0.5 | 0.54 | 0.53 | 0.54 | 5.12 | 0.79 | 0.58 |  |  |  | 36
| 1 | 3 | 0.81261 | 0.53 | 0.4 | < 5 | 0.54 | 0.42 | 0.49 | 0.52 | 0.52 | 0.52 | 0.56 | 0.53809 | 0.9 | 0.55 | 0.5 | 0.56 | 0.5 | 0.54 | 0.55 | 0.5 | 5.57 | 0.84 | 0.56 |  |  |  | 16
| 1 | 4 | 0.80656 | 0.51 | 0.4 | < 5 | 0.52 | 0.5 | 0.5 | 0.52 | 0.5 | 0.52 | 0.56 | 0.57823 | 0.9 | 0.53 | 0.5 | 0.57 | 0.5 | 0.55 | 0.55 | 0.56 | 5.59 | 0.72 | 0.55 |  |  |  | 0.53169
| 1 | 5 | 0.77555 | 0.51 | 0.3 | < 5 | 0.52 | 0.46 | 0.54 | 0.54 | 0.51 | 0.54 | 0.55 | 0.54982 | 0.7 | 0.52 | 0.5 | 0.58 | 0.5 | 0.55 | 0.53 | 0.53 | 6.14 | 0.72 | 0.53 |  |  |  | 53
| 1 | 6 | 0.843 | 0.55 | 0.3 | < 5 | 0.52 | 0.42 | 0.47 | 0.56 | 0.48 | 0.55 | 0.56 | 0.57488 | 0.9 | 0.54 | 0.5 | 0.6 | 0.5 | 0.52 | 0.52 | 0.56 |  | 1.15 | 0.58
| Mean |  | 0.82485 | 0.53333 | 0.38333 |  | 0.55667 | 0.44333 | 0.50667 | 0.51667 | 0.50667 | 0.52667 | 0.55833 | 0.55942 | 0.88333 | 0.535 | 0.5 | 0.56667 | 0.5 | 0.54 | 0.52667 | 0.53333 | 5.88 | 0.83 | 0.56
| Median |  | 0.81444 | 0.54 | 0.4 |  | 0.53 | 0.43 | 0.5 | 0.52 | 0.51 | 0.525 | 0.56 | 0.56235 | 0.9 | 0.535 | 0.5 | 0.575 | 0.5 | 0.54 | 0.53 | 0.535 | 5.59 | 0.775 | 0.56
| Std Dev. |  | 0.040622 | 0.019664 | 0.075277 |  | 0.070899 | 0.032042 | 0.028048 | 0.03266 | 0.015055 | 0.017512 | 0.0040825 | 0.024428 | 0.13292 | 0.013784 | 0 | 0.030768 | 0 | 0.010954 | 0.02582 | 0.025033 | 0.71334 | 0.16322 | 0.018974
| Rel.Std.Dev. |  | 4.925% | 3.687% | 19.64% |  | 12.74% | 7.227% | 5.536% | 6.321% | 2.971% | 3.325% | 0.7312% | 4.367% | 15.05% | 2.576% | 0% | 5.43% | 0% | 2.029% | 4.903% | 4.694% | 12.13% | 19.66% | 3.388%
| PDM3 |  | 55.14% | 0.3081% | -27.9% |  | 4.697% | -16.62% | -4.707% | -2.826% | -4.707% | -0.9457% | 5.01% | 5.215% | 66.14% | 0.6216% | -5.961% | 6.577% | -5.961% | 1.562% | -0.9457% | 0.3081% | 1006% | 56.1% | 5.324%
| Z-Score (Absolute) |  | 9.6 | 0.22 | NA | 66.08 | 0.56 | 3.26 | 1.12 | 0.79 | 1.12 | 0.45 | 0.62 | 0.65 | NA | 0.17 | NA | 0.9 | NA | 0 | 0.45 | 0.22 | 180.04 | 9.78 | 0.67
Results from laboratories 1.09, 1.20, 1.22 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 285 (Certified Value 8.99 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 8.5765 | 9.4 | 9 | 10.3274 | 9 | 8.36 | 8.8 | 8.9 | 8.9 | 8.7 | 8.6 | 9.5 | 8.7299 | 10.1 | 8.5 | 8 | 10.4 | 8 | 9.9 | 9 | 8.9 | 9 | 8.6 | 7.1 |  |  | 1
| 1 | 2 | 8.7296 | 9.4 | 9 | 10.1489 | 9 | 8.3 | 9.1 | 9 | 8.8 | 8.5 | 8.8 | 9.5 | 9.1898 | 9.7 | 8.6 | 8 | 9.9 | 9 | 9.8 | 9 | 8.7 | 9 | 8.95 | 7.2 |  |  | 37
| 1 | 3 | 8.6549 | 9.4 | 9 | 10.1667 | 9 | 8.48 | 8.5 | 8.6 | 9.1 | 8.8 | 8.6 | 9.7 | 9.3347 | 9.5 | 8.8 | 8 | 10.3 | 9 | 9.6 | 9.2 | 9 | 9 | 9.2 | 6.9 |  |  | 16
| 1 | 4 | 8.8149 | 9.2 | 9 | 10.5 | 9 | 8.53 | 8.4 | 8.8 | 9 | 8.5 | 8.8 | 9.5 | 9.199 | 9.3 | 8.7 | 8 | 10.3 | 9 | 9.7 | 9.2 | 9 | 9 | 9.5 | 7.2 |  |  | 8.9903
| 1 | 5 | 8.7389 | 9 | 9 | 10.9045 | 9 | 8.43 | 8.6 | 9.1 | 9 | 8.5 | 9 | 9.7 | 8.8294 | 8.7 | 8.7 | 8 | 9.7 | 8 | 10 | 9.2 | 8.4 | 9 | 9.1 | 7.2 |  |  | 54
| 1 | 6 | 8.8694 | 9.2 | 9 | 9.9075 | 9 | 8.37 | 8.7 | 8.6 | 9 | 8.3 | 9.2 | 9.6 | 8.8822 | 9.4 | 8.6 | 8 | 10.3 | 8 | 9.5 | 9.1 | 8.9 | 9 | 9.2 | 7
| Mean |  | 8.7307 | 9.2667 | 9 | 10.3258 | 9 | 8.4117 | 8.6833 | 8.8333 | 8.9667 | 8.55 | 8.8333 | 9.5833 | 9.0275 | 9.45 | 8.65 | 8 | 10.15 | 8.5 | 9.75 | 9.1167 | 8.8167 | 9 | 9.0917 | 7.1
| Median |  | 8.7342 | 9.3 | 9 | 10.2471 | 9 | 8.4 | 8.65 | 8.85 | 9 | 8.5 | 8.8 | 9.55 | 9.036 | 9.45 | 8.65 | 8 | 10.3 | 8.5 | 9.75 | 9.15 | 8.9 | 9 | 9.15 | 7.15
| Std Dev. |  | 0.10575 | 0.1633 | 0 | 0.34549 | 0 | 0.084715 | 0.24833 | 0.20656 | 0.10328 | 0.17607 | 0.23381 | 0.098319 | 0.24457 | 0.46368 | 0.10488 | 0 | 0.28107 | 0.54772 | 0.18708 | 0.098319 | 0.23166 | 0 | 0.30069 | 0.12649
| Rel.Std.Dev. |  | 1.211% | 1.762% | 0% | 3.346% | 0% | 1.007% | 2.86% | 2.338% | 1.152% | 2.059% | 2.647% | 1.026% | 2.709% | 4.907% | 1.212% | 0% | 2.769% | 6.444% | 1.919% | 1.078% | 2.628% | 0% | 3.307% | 1.782%
| PDM3 |  | -2.888% | 3.074% | 0.1079% | 14.86% | 0.1079% | -6.436% | -3.414% | -1.746% | -0.2629% | -4.898% | -1.746% | 6.596% | 0.4136% | 5.113% | -3.785% | -11.02% | 12.9% | -5.454% | 8.45% | 1.406% | -1.931% | 0.1079% | 1.127% | -21.03%
| Z-Score (Absolute) |  | 0.53 | 0.67 | NA | 3.05 | NA | 1.25 | 0.64 | 0.3 | 0 | 0.94 | 0.3 | 1.39 | 0.14 | 1.09 | 0.71 | NA | 2.66 | NA | 1.76 | 0.34 | 0.34 | NA | 0.28 | 4.2
Results from laboratories 1.09, 1.11, 1.22, 1.24 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 285 (Indicative Value 0.96 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | < 1.5 | 0.8 | 0.6 | < 0.1 | < 5 | 0.46 | 1.4 | 1 | < 1 | 1 | < 1 | 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 |  | 2 | < 2 |  |  |  |  | 1
| 1 | 2 | < 1.5 | 1.1 | 0.6 | < 0.1 | < 5 | 0.46 | 1 | 1 | 1 | 1 | < 1 | 1 | < 2 | < 2 | < 1 | 0.5 | < 1 | < 1 | < 0.5 |  | 3.4 | < 2 |  |  |  |  | 38
| 1 | 3 | < 1.5 | 0.9 | 0.4 | < 0.1 | < 5 | 0.46 | 0.9 | 1 | < 1 | 1 | < 1 | 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 |  | 3 | < 2 |  |  |  |  | 16
| 1 | 4 | < 1.5 | 0.8 | 0.4 | < 0.1 | < 5 | 0.45 | 1.5 | < 1 | < 1 | 1 | < 1 | < 1 | < 2 | < 2 | < 1 | 0.5 | < 1 | < 1 | < 0.5 | 0.7 | 3 | < 2 |  |  |  |  | 0.95708
| 1 | 5 | < 1.5 | 0.8 | 0.4 | < 0.1 | < 5 | 0.45 | 1.1 | 1 | < 1 | 1 | < 1 | 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 |  | 3 | < 2 |  |  |  |  | 7
| 1 | 6 | < 1.5 | 1.1 | 0.3 | < 0.1 | < 5 | 0.44 | 0.8 | 1 | < 1 | 1 | 1 | < 1 | < 2 | < 2 | < 1 | < 0.5 | < 1 | < 1 | < 0.5 |  | 3 | < 2
| Mean |  |  | 0.91667 | 0.45 |  |  | 0.45333 | 1.1167 | 1 | 1 | 1 | 1 | 1 |  |  |  | 0.5 |  |  |  | 0.7 | 2.9
| Median |  |  | 0.85 | 0.4 |  |  | 0.455 | 1.05 | 1 | 1 | 1 | 1 | 1 |  |  |  | 0.5 |  |  |  | 0.7 | 3
| Std Dev. |  |  | 0.1472 | 0.12247 |  |  | 0.008165 | 0.27869 | 0 |  | 0 |  | 0 |  |  |  | 0 |  |  |  |  | 0.46904
| Rel.Std.Dev. |  |  | 16.06% | 27.22% |  |  | 1.801% | 24.96% | 0% |  | 0% |  | 0% |  |  |  | 0% |  |  |  |  | 16.17%
| PDM3 |  |  | -4.223% | -52.98% |  |  | -52.63% | 16.67% | 4.484% | 4.484% | 4.484% | 4.484% | 4.484% |  |  |  | -47.76% |  |  |  | -26.86% | 203%
| Z-Score (Absolute) |  | 0.06 | 0.47 | 0.68 | 1.67 | 4.38 | 0.67 | 0.97 | 0.47 | 0.35 | 0.68 | 0.35 | 0.27 | 0.68 | 0.68 | 0.56 | 0.97 | 0.56 | 0.56 | 1.17 | 0.06 | 5.37 | 0.68
| Analytical results for Sm in OREAS 285 (Certified Value 7.19 ppm) | [27 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | 1.29 | [19 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 6.9953 | 7.43 | 6.7 | 6.6045 | 6.3 | 7.51 | 5 | [19 empty] | 1
| 1 | 2 | 7.1449 | 7.15 | 7.7 | 6.7555 | 6.9 | 7.52 | 5 | [19 empty] | 4
| 1 | 3 | 7.1375 | 7.43 | 8 | 6.6166 | 7 | 7.71 | 5 | [19 empty] | 16
| 1 | 4 | 7.0855 | 7.31 | 7.3 | 6.9964 | 7.3 | 7.54 | 5 | [19 empty] | 7.1915
| 1 | 5 | 7.0684 | 7.13 | 7.9 | 7.4929 | 6.8 | 7.8 | 6 | [19 empty] | 55
| 1 | 6 | 7.4273 | 7.41 | 6.8 | 6.8692 | 6.8 | 7.6 | 5
| Mean |  | 7.1432 | 7.31 | 7.4 | 6.8892 | 6.85 | 7.6133 | 5.1667
| Median |  | 7.1115 | 7.36 | 7.5 | 6.8124 | 6.85 | 7.57 | 5
| Std Dev. |  | 0.14936 | 0.13914 | 0.55857 | 0.33154 | 0.32711 | 0.11759 | 0.40825
| Rel.Std.Dev. |  | 2.091% | 1.903% | 7.548% | 4.812% | 4.775% | 1.544% | 7.902%
| PDM3 |  | -0.6719% | 1.648% | 2.9% | -4.203% | -4.748% | 5.866% | -28.16%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 0.46 | 0.89 | 0.99 | 1.02 | NA
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 285 (Certified Value 4.66 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 4.7289 | 5 | 3 | < 10 | 5.14 | 3.8 | 4.6 | 4.2 | 4.9 | 4.5 | 4.7 | 4.6117 | 4.7 | 4.6 | 4.8 | 4.9 | 4.5 | 4.7 | 4.9 | 4.9 | 16 | 3.3 | 4.8 |  |  |  | 1
| 1 | 2 | 4.743 | 4.8 | 2 | < 10 | 3.84 | 4.3 | 4.5 | 4.3 | 4.7 | 4.6 | 4.8 | 4.7247 | 4.8 | 4.7 | 4.7 | 4.6 | 4.5 | 4.7 | 5.1 | 4.8 | 14 | 3.5 | 4.7 |  |  |  | 13
| 1 | 3 | 4.5643 | 4.8 | 2 | < 10 | 4.94 | 4.1 | 4.2 | 4.5 | 4.8 | 4.5 | 4.9 | 4.5493 | 4.9 | 4.5 | 4.5 | 4.7 | 4.5 | 4.5 | 5.1 | 4.6 | 15 | 3.5 | 4.6 |  |  |  | 16
| 1 | 4 | 4.5089 | 4.7 | 3 | < 10 | 4.54 | 4.5 | 4.4 | 4.5 | 4.9 | 4.6 | 4.8 | 4.813 | 4.9 | 4.5 | 4.6 | 4.5 | 4.4 | 4.6 | 5.3 | 4.7 | 17 | 3.3 | 4.6 |  |  |  | 4.658
| 1 | 5 | 4.6658 | 4.7 | 3 | < 10 | 4.54 | 4.4 | 4.6 | 4.5 | 4.8 | 4.6 | 5 | 4.5416 | 4.8 | 4.4 | 4.6 | 4.6 | 4.4 | 4.8 | 5 | 4.7 | 12 | 3.4 | 4.6 |  |  |  | 56
| 1 | 6 | 4.7069 | 4.8 | 4 | < 10 | 4.04 | 4.4 | 4.4 | 4.7 | 4.7 | 4.6 | 4.7 | 4.7159 | 5 | 4.7 | 4.6 | 4.8 | 4.1 | 4.5 | 5.1 | 4.8 | 12 | 3.4 | 4.7
| Mean |  | 4.653 | 4.8 | 2.8333 |  | 4.5067 | 4.25 | 4.45 | 4.45 | 4.8 | 4.5667 | 4.8167 | 4.6593 | 4.85 | 4.5667 | 4.6333 | 4.6833 | 4.4 | 4.6333 | 5.0833 | 4.75 | 14.3333 | 3.4 | 4.6667
| Median |  | 4.6864 | 4.8 | 3 |  | 4.54 | 4.35 | 4.45 | 4.5 | 4.8 | 4.6 | 4.8 | 4.6638 | 4.85 | 4.55 | 4.6 | 4.65 | 4.45 | 4.65 | 5.1 | 4.75 | 14.5 | 3.4 | 4.65
| Std Dev. |  | 0.095442 | 0.10954 | 0.75277 |  | 0.50067 | 0.25884 | 0.15166 | 0.17607 | 0.089443 | 0.05164 | 0.1169 | 0.10894 | 0.10488 | 0.12111 | 0.10328 | 0.1472 | 0.15492 | 0.12111 | 0.13292 | 0.10488 | 2.0656 | 0.089443 | 0.08165
| Rel.Std.Dev. |  | 2.051% | 2.282% | 26.57% |  | 11.11% | 6.09% | 3.408% | 3.957% | 1.863% | 1.131% | 2.427% | 2.338% | 2.162% | 2.652% | 2.229% | 3.143% | 3.521% | 2.614% | 2.615% | 2.208% | 14.41% | 2.631% | 1.75%
| PDM3 |  | -0.1081% | 3.048% | -39.17% |  | -3.249% | -8.759% | -4.466% | -4.466% | 3.048% | -1.961% | 3.406% | 0.02841% | 4.122% | -1.961% | -0.5299% | 0.5435% | -5.539% | -0.5299% | 9.131% | 1.975% | 207.7% | -27.01% | 0.1857%
| Z-Score (Absolute) |  | 0.01 | 0.66 | NA | 1.58 | 0.69 | 1.87 | 0.95 | 0.95 | 0.66 | 0.41 | 0.74 | 0.01 | 0.89 | 0.41 | 0.1 | 0.12 | 1.18 | 0.1 | 1.96 | 0.43 | 44.49 | 5.78 | 0.05
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 285 (Certified Value 160 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
0
| 1 | 1 | 160.5518 | 154.8866 | 166.49 | 164 | 153.7164 | 162 | 173 | 164 | 159 | 156.5 | 158 | 150 | 168.5 | 158.1591 | 150 | 156.5 | 165 | 177 | 161 | 157 | 154 | 156.52 | 116.74 | 164 | 104 | 147.7 | 1
| 1 | 2 | 163.6682 | 155.1006 | 163.45 | 160 | 159.0729 | 160 | 172 | 172 | 158.5 | 162 | 154.5 | 151 | 167.5 | 164.1632 | 150 | 150.1 | 160 | 173 | 162 | 159 | 154 | 158.32 | 109.07 | 163 | 110 | 143.2 | 14
| 1 | 3 | 162.5608 | 156.8877 | 163.32 | 164 | 153.0778 | 157 | 176 | 161 | 153.5 | 168.5 | 160 | 148.5 | 170.5 | 163.8503 | 152 | 156.9 | 160 | 177 | 161 | 157 | 157 | 158.84 | 118.88 | 162 | 114 | 146.1 | 16
| 1 | 4 | 161.8748 | 156.5321 | 164.89 | 162 | 159.7271 | 167 | 176 | 159 | 155.5 | 168 | 160.5 | 153.5 | 164.5 | 164.5322 | 152 | 154.2 | 162 | 167 | 161 | 155 | 158 | 161.36 | 114.08 | 166 | 111 | 153.3 | 159.5549
| 1 | 5 | 161.9924 | 157.5589 | 164.57 | 164 | 158.4552 | 160 | 176 | 162 | 160.5 | 168 | 156 | 154.5 | 168.5 | 155.5392 | 150 | 150 | 162 | 170 | 157 | 157 | 158 | 156.48 |  | 164 | 111 | 143.6 | 57
| 1 | 6 | 163.0018 | 158.9126 | 165.2 | 171 | 155.9388 | 159 | 175 | 163 | 154.5 | 167 | 157.5 | 156 | 168.5 | 158.2476 | 153 | 152.9 | 164 | 174 | 148 | 155 | 155 | 159.75 | 107.7 | 165 | 112 | 144.3
| Mean |  | 162.275 | 156.6464 | 164.6533 | 164.1667 | 156.6647 | 160.8333 | 174.6667 | 163.5 | 156.9167 | 165 | 157.75 | 152.25 | 168 | 160.7486 | 151.1667 | 153.4333 | 162.1667 | 173 | 158.3333 | 156.6667 | 156 | 158.545 | 113.294 | 164 | 110.3333 | 146.3667
| Median |  | 162.2766 | 156.7099 | 164.73 | 164 | 157.197 | 160 | 175.5 | 162.5 | 157 | 167.5 | 157.75 | 152.25 | 168.5 | 161.0489 | 151 | 153.55 | 162 | 173.5 | 161 | 157 | 156 | 158.58 | 114.08 | 164 | 111 | 145.2
| Std Dev. |  | 1.0733 | 1.5183 | 1.1799 | 3.7103 | 2.845 | 3.4303 | 1.7512 | 4.5056 | 2.8003 | 4.8062 | 2.2967 | 2.8766 | 1.9748 | 3.8909 | 1.3292 | 3.0064 | 2.0412 | 3.9497 | 5.3541 | 1.5055 | 1.8974 | 1.891 | 4.8174 | 1.4142 | 3.3862 | 3.7914
| Rel.Std.Dev. |  | 0.6614% | 0.9692% | 0.7166% | 2.26% | 1.816% | 2.133% | 1.003% | 2.756% | 1.785% | 2.913% | 1.456% | 1.889% | 1.176% | 2.42% | 0.8793% | 1.959% | 1.259% | 2.283% | 3.382% | 0.961% | 1.216% | 1.193% | 4.252% | 0.8623% | 3.069% | 2.59%
| PDM3 |  | 1.705% | -1.823% | 3.195% | 2.89% | -1.811% | 0.8012% | 9.471% | 2.473% | -1.654% | 3.413% | -1.131% | -4.578% | 5.293% | 0.7481% | -5.257% | -3.837% | 1.637% | 8.427% | -0.7656% | -1.81% | -2.228% | -0.633% | -28.99% | 2.786% | -30.85% | -8.266%
| Z-Score (Absolute) |  | 0.51 | 0.24 | 0.83 | 0.77 | 0.24 | 0.32 | 2.17 | 0.68 | 0.2 | 0.88 | 0.09 | 0.83 | 1.28 | 0.31 | 0.97 | 0.67 | 0.5 | 1.95 | 0.01 | 0.24 | 0.33 | 0.01 | 6.05 | 0.74 | 6.44 | 1.62
| Analytical results for Ta in OREAS 285 (Certified Value 1.09 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.2016 | 1.12 | < 0.1 | 5 | 1.9 | 0.9 | 1.04 | 1.08 | 1.04 | 1.16 | 1.13 | 1.06 | 1.1 | 1.35 | 1.02 | 1.11 | 1 | 1.05 | 1.09 | 1.09 | < 1 | 0.95 | 1.08 |  |  |  | 1
| 1 | 2 | 1.2273 | 1.13 | 0.1 | 5 | 2.07 | 1 | 1.04 | 1.03 | 1.06 | 1.18 | 1.16 | 1.0681 | 1.3 | 1.26 | 1 | 1.1 | 1 | 1.03 | 1.08 | 1.1 | < 1 | 1.06 | 1.11 |  |  |  | 25
| 1 | 3 | 1.1915 | 1.13 | < 0.1 | 5 | 1.84 | 0.9 | 1.02 | 1.1 | 1 | 1.15 | 1.15 | 1.1113 | 1 | 1.25 | 1.01 | 1.06 | 1 | 1.05 | 1.09 | 1.07 | < 1 | 0.95 | 1.1 |  |  |  | 16
| 1 | 4 | 1.1933 | 1.14 | 0.1 | 5 | 1.82 | 1 | 1.02 | 1.09 | 1.03 | 1.13 | 1.13 | 1.1488 | 1.3 | 1.24 | 1.01 | 1.06 | 1 | 1.03 | 1.08 | 1.06 | < 1 | 0.86 | 1.05 |  |  |  | 1.0869
| 1 | 5 | 1.1576 | 1.1 | < 0.1 | 5 | 1.86 | 1 | 1.02 | 1.12 | 1.03 | 1.13 | 1.14 | 1.1445 | 1.3 | 1.32 | 1.01 | 1.1 | 1 | 1.04 | 1.11 | 1.11 | 1 | 0.79 | 0.98 |  |  |  | 58
| 1 | 6 | 1.229 | 1.12 | < 0.1 | 5 | 1.87 | 0.9 | 1.02 | 1.11 | 1.03 | 1.14 | 1.12 | 1.1137 | 1 | 1.28 | 1.01 | 1.1 | 1 | 1.01 | 1.12 | 1.12 | < 1 | 0.74 | 1.03
| Mean |  | 1.2001 | 1.1233 | 0.1 | 5 | 1.8933 | 0.95 | 1.0267 | 1.0883 | 1.0317 | 1.1483 | 1.1383 | 1.1078 | 1.1667 | 1.2833 | 1.01 | 1.0883 | 1 | 1.035 | 1.095 | 1.0917 | 1 | 0.89167 | 1.0583
| Median |  | 1.1975 | 1.125 | 0.1 | 5 | 1.865 | 0.95 | 1.02 | 1.095 | 1.03 | 1.145 | 1.135 | 1.1125 | 1.2 | 1.27 | 1.01 | 1.1 | 1 | 1.035 | 1.09 | 1.095 | 1 | 0.905 | 1.065
| Std Dev. |  | 0.026479 | 0.013663 | 0 | 0 | 0.090701 | 0.054772 | 0.010328 | 0.031885 | 0.019408 | 0.019408 | 0.01472 | 0.037235 | 0.15055 | 0.043205 | 0.0063246 | 0.022286 | 0 | 0.015166 | 0.016432 | 0.023166 |  | 0.11788 | 0.048751
| Rel.Std.Dev. |  | 2.206% | 1.216% | 0% | 0% | 4.791% | 5.766% | 1.006% | 2.93% | 1.881% | 1.69% | 1.293% | 3.361% | 12.9% | 3.367% | 0.6262% | 2.048% | 0% | 1.465% | 1.501% | 2.122% |  | 13.22% | 4.606%
| PDM3 |  | 10.41% | 3.349% | -90.8% | 360% | 74.19% | -12.6% | -5.545% | 0.1287% | -5.085% | 5.649% | 4.729% | 1.915% | 7.336% | 18.07% | -7.078% | 0.1287% | -7.998% | -4.778% | 0.7421% | 0.4354% | -7.998% | -17.96% | -2.631%
| Z-Score (Absolute) |  | 1.27 | 0.39 | 11.83 | NA | 9.29 | 1.62 | 0.73 | 0.02 | 0.67 | 0.67 | 0.56 | 0.21 | 0.89 | 2.23 | 0.92 | 0.02 | NA | 0.64 | 0.06 | 0.02 | NA | 2.29 | 0.37
Results from laboratory 1.24 were removed due to their 0.1 ppm reading resolution. Results from laboratories 1.11 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 285 (Certified Value 0.79 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.21 | 1.25 | 1.26 | 1.29 | 1.32 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [16 empty] | 2
3
| 1 | 1 | 0.80822 | 0.79 | 0.8 | 0.7052 | 0.7 | 0.83 | 0.81 | 0.87 | < 1 | 0.69 | [16 empty] | 1
| 1 | 2 | 0.84427 | 0.8 | 0.8 | 0.754 | 0.8 | 0.78 | 0.76 | 0.85 | < 1 | 0.68 | [16 empty] | 26
| 1 | 3 | 0.83766 | 0.83 | 0.8 | 0.697 | 0.8 | 0.8 | 0.76 | 0.86 | < 1 | 0.68 | [16 empty] | 16
| 1 | 4 | 0.83614 | 0.8 | 0.8 | 0.7425 | 0.8 | 0.79 | 0.78 | 0.87 | < 1 | 0.69 | [16 empty] | 0.78657
| 1 | 5 | 0.83472 | 0.77 | 0.9 | 0.8241 | 0.7 | 0.79 | 0.81 | 0.87 | < 1 | 0.7 | [16 empty] | 59
| 1 | 6 | 0.86619 | 0.77 | 0.8 | 0.7358 | 0.8 | 0.8 | 0.77 | 0.87 | < 1 | 0.68
| Mean |  | 0.83787 | 0.79333 | 0.81667 | 0.7431 | 0.76667 | 0.79833 | 0.78167 | 0.865 |  | 0.68667
| Median |  | 0.8369 | 0.795 | 0.8 | 0.73915 | 0.8 | 0.795 | 0.775 | 0.87 |  | 0.685
| Std Dev. |  | 0.018627 | 0.022509 | 0.040825 | 0.045376 | 0.05164 | 0.017224 | 0.023166 | 0.0083666 |  | 0.008165
| Rel.Std.Dev. |  | 2.223% | 2.837% | 4.999% | 6.106% | 6.736% | 2.157% | 2.964% | 0.9672% |  | 1.189%
| PDM3 |  | 6.522% | 0.8603% | 3.827% | -5.526% | -2.53% | 1.496% | -0.623% | 9.972% |  | -12.7%
| Z-Score (Absolute) |  | 0.72 | 0.08 | NA | 0.63 | NA | 0.15 | 0.08 | 1.1 | 4.09 | 1.43
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 285 (Certified Value 0.066 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | 0.1 | 6 | 0.06 | < 0.05 | 0.06 | 0.06 | 0.06 | 0.08 | < 0.05 | < 0.1 | < 0.1 | 0.07 | < 0.1 | 0.07 | < 0.5 | 0.06 | < 0.2 | 0.25 | 0.06 | [5 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | 0.1 | 6 | 0.08 | < 0.05 | 0.05 | 0.06 | 0.06 | 0.08 | 0.05 | < 0.1 | < 0.1 | 0.06 | < 0.1 | 0.08 | < 0.5 | 0.05 | < 0.2 | 0.054 | < 0.05 | [5 empty] | 27
| 1 | 3 | < 0.2 | < 0.2 | 0.1 | 6 | 0.1 | 0.09 | 0.06 | 0.07 | < 0.05 | 0.07 | < 0.05 | < 0.1 | 0.2 | 0.05 | < 0.1 | 0.11 | < 0.5 | 0.05 | < 0.2 | 0.11 | < 0.05 | [5 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 6 | 0.11 | 0.19 | 0.06 | 0.09 | < 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.1 | 0.06 | < 0.1 | 0.05 | < 0.5 | 0.06 | < 0.2 | 0.1 | 0.08 | [5 empty] | 0.06555
| 1 | 5 | < 0.2 | < 0.2 | 0.1 | 6 | 0.15 | 0.05 | 0.07 | 0.11 | 0.05 | 0.1 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.03 | < 0.5 | 0.05 | < 0.2 | 0.26 | 0.09 | [5 empty] | 60
| 1 | 6 | < 0.2 | < 0.2 | 0.2 | 6 | 0.03 | 0.14 | 0.06 | 0.06 | 0.08 | 0.06 | < 0.05 | < 0.1 | < 0.1 | 0.07 | < 0.1 | 0.03 | < 0.5 | 0.07 | < 0.2 | 0.05 | 0.07
| Mean |  |  |  | 0.12 | 6 | 0.088333 | 0.1175 | 0.06 | 0.075 | 0.0625 | 0.076667 | 0.05 |  | 0.2 | 0.062 |  | 0.061667 |  | 0.056667 |  | 0.13733 | 0.075
| Median |  |  |  | 0.1 | 6 | 0.09 | 0.115 | 0.06 | 0.065 | 0.06 | 0.075 | 0.05 |  | 0.2 | 0.06 |  | 0.06 |  | 0.055 |  | 0.105 | 0.075
| Std Dev. |  |  |  | 0.044721 | 0 | 0.041673 | 0.060759 | 0.0063246 | 0.020736 | 0.012583 | 0.013663 |  |  |  | 0.0083666 |  | 0.031252 |  | 0.008165 |  | 0.09429 | 0.01291
| Rel.Std.Dev. |  |  |  | 37.27% | 0% | 47.18% | 51.71% | 10.54% | 27.65% | 20.13% | 17.82% |  |  |  | 13.49% |  | 50.68% |  | 14.41% |  | 68.66% | 17.21%
| PDM3 |  |  |  | 83.07% | 9053% | 34.76% | 79.25% | -8.467% | 14.42% | -4.653% | 16.96% | -23.72% |  | 205.1% | -5.416% |  | -5.924% |  | -13.55% |  | 109.5% | 14.42%
| Z-Score (Absolute) |  | 1.16 | 1.16 | NA | NA | 0.74 | 0.67 | 0.31 | 0.25 | 0.67 | 0.31 | 1.44 | 0.67 | NA | 0.46 | 0.67 | 0.25 | 6.68 | 0.43 | 1.16 | 2.54 | 0.37
Results from laboratories 1.09 and 1.20 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 285 (Certified Value 15.3 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 14.9564 | 15.18 | 15.5 | 16.2866 | 10.55 | 13.3 | 14.65 | 14.3 | 15.9 | 14.8 | 15.1 | 15.4392 | 18.6 | 12.9 | 15.2 | 16.6 | 14.8 | 15.48 | 15.99 | 15.51 | 14 | 16.4 | 10.7 |  |  |  | 1
| 1 | 2 | 15.7356 | 15.28 | 15.3 | 15.9012 | 10.36 | 14.8 | 15.05 | 15.6 | 14.45 | 13.75 | 14.6 | 16.3443 | 17 | 12.7 | 14.8 | 16.4 | 15.1 | 14.89 | 15.36 | 15.47 | 13 | 17.3 | 10.6 |  |  |  | 28
| 1 | 3 | 14.4601 | 15.77 | 15.4 | 14.02 | 10.47 | 14.8 | 14.95 | 16.65 | 16.15 | 14.3 | 15.9 | 16.1057 | 16.1 | 12.9 | 14.9 | 16 | 14.8 | 15.3 | 15.85 | 15.74 | 14 | 17.9 | 10.1 |  |  |  | 16
| 1 | 4 | 15.9323 | 15.17 | 15.4 | 14.8355 | 10.53 | 16.8 | 14.55 | 16.2 | 16 | 15.1 | 15.45 | 16.1885 | 15 | 12.3 | 15.1 | 15.1 | 14.1 | 14.62 | 15.36 | 15.73 | 14 | 17.7 | 9.7 |  |  |  | 15.3383
| 1 | 5 | 15.0225 | 15.48 | 15.7 | 15.3692 | 10.62 | 15.2 | 15.2 | 15.75 | 15.65 | 15.15 | 15.4 | 15.6834 | 13.7 | 12.4 | 15.3 | 15.7 | 15 | 16.62 | 16.07 | 15.06 | 15 | 18 | 10.8 |  |  |  | 61
| 1 | 6 | 15.6732 | 15.43 | 15.8 | 14.1528 | 10.53 | 14.4 | 14.85 | 16.35 | 15.7 | 15.4 | 15.2 | 15.4382 | 16.1 | 12.7 | 15.3 | 15.2 | 14.3 | 14.87 | 15.47 | 15.77 | 14 | 18.3 | 9.8
| Mean |  | 15.2967 | 15.385 | 15.5167 | 15.0942 | 10.51 | 14.8833 | 14.875 | 15.8083 | 15.6417 | 14.75 | 15.275 | 15.8665 | 16.0833 | 12.65 | 15.1 | 15.8333 | 14.6833 | 15.2967 | 15.6833 | 15.5467 | 14 | 17.6 | 10.2833
| Median |  | 15.3479 | 15.355 | 15.45 | 15.1023 | 10.53 | 14.8 | 14.9 | 15.975 | 15.8 | 14.95 | 15.3 | 15.8945 | 16.1 | 12.7 | 15.15 | 15.85 | 14.8 | 15.095 | 15.66 | 15.62 | 14 | 17.8 | 10.35
| Std Dev. |  | 0.57084 | 0.2274 | 0.19408 | 0.92258 | 0.087864 | 1.1427 | 0.24444 | 0.83392 | 0.61271 | 0.61806 | 0.43099 | 0.39721 | 1.675 | 0.251 | 0.20976 | 0.61536 | 0.39707 | 0.71974 | 0.32432 | 0.27016 | 0.63246 | 0.67528 | 0.47924
| Rel.Std.Dev. |  | 3.732% | 1.478% | 1.251% | 6.112% | 0.836% | 7.677% | 1.643% | 5.275% | 3.917% | 4.19% | 2.822% | 2.503% | 10.41% | 1.984% | 1.389% | 3.886% | 2.704% | 4.705% | 2.068% | 1.738% | 4.518% | 3.837% | 4.66%
| PDM3 |  | -0.2712% | 0.3044% | 1.163% | -1.591% | -31.48% | -2.966% | -3.021% | 3.064% | 1.978% | -3.836% | -0.4127% | 3.444% | 4.857% | -17.53% | -1.554% | 3.227% | -4.27% | -0.2715% | 2.249% | 1.358% | -8.725% | 14.75% | -32.96%
| Z-Score (Absolute) |  | 0 | 0.14 | 0.36 | 0.33 | 7.73 | 0.67 | 0.68 | 0.83 | 0.56 | 0.88 | 0.04 | 0.92 | 1.27 | 4.27 | 0.32 | 0.87 | 0.99 | 0 | 0.62 | 0.4 | NA | 3.72 | 8.1
Results from laboratory 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 285 (Certified Value 0.372 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.36201 | 0.37068 | 0.3763 | 0.337 | 0.3305 | 0.41 | 0.3683 | 0.367 | 0.363 | 0.366 | 0.361 | 0.361 | 0.364 | 0.36895 | 0.385 | 0.36 | 0.395 | 0.34 | 0.394 | 0.377 | 0.3613 | 0.3777 | 0.36985 | 0.396 | 0.371 | 0.37 | 1
| 1 | 2 | 0.36858 | 0.36895 | 0.3736 | 0.337 | 0.3414 | 0.4 | 0.3678 | 0.38 | 0.359 | 0.377 | 0.357 | 0.36 | 0.365 | 0.37524 | 0.383 | 0.36 | 0.391 | 0.36 | 0.393 | 0.381 | 0.3617 | 0.3806 | 0.34637 | 0.396 | 0.394 | 0.37 | 29
| 1 | 3 | 0.36208 | 0.37238 | 0.3778 | 0.311 | 0.3379 | 0.42 | 0.3759 | 0.356 | 0.358 | 0.387 | 0.364 | 0.359 | 0.375 | 0.36682 | 0.385 | 0.37 | 0.378 | 0.36 | 0.397 | 0.378 | 0.3671 | 0.3778 | 0.35602 | 0.39 | 0.404 | 0.37 | 16
| 1 | 4 | 0.36184 | 0.38222 | 0.3843 | 0.323 | 0.3488 | 0.42 | 0.3725 | 0.354 | 0.361 | 0.386 | 0.368 | 0.36 | 0.36 | 0.37759 | 0.389 | 0.36 | 0.391 | 0.36 | 0.395 | 0.375 | 0.3752 | 0.3774 | 0.35458 | 0.402 | 0.398 | 0.38 | 0.3717
| 1 | 5 | 0.36468 | 0.37211 | 0.3749 | 0.303 | 0.3449 | 0.4 | 0.3751 | 0.358 | 0.365 | 0.384 | 0.357 | 0.367 | 0.367 | 0.37743 | 0.386 | 0.36 | 0.389 | 0.35 | 0.384 | 0.38 | 0.3747 | 0.3767 | 0.35482 | 0.396 | 0.402 | 0.35 | 62
| 1 | 6 | 0.3634 | 0.38228 | 0.3789 | 0.269 | 0.3375 | 0.4 | 0.3761 | 0.36 | 0.357 | 0.386 | 0.356 | 0.377 | 0.363 | 0.37103 | 0.388 | 0.36 | 0.386 | 0.36 | 0.363 | 0.369 | 0.3698 | 0.3799 | 0.36267 | 0.396 | 0.409 | 0.36
| Mean |  | 0.36377 | 0.37477 | 0.37763 | 0.31333 | 0.34017 | 0.40833 | 0.37262 | 0.3625 | 0.3605 | 0.381 | 0.3605 | 0.364 | 0.36567 | 0.37284 | 0.386 | 0.36167 | 0.38833 | 0.355 | 0.38767 | 0.37667 | 0.3683 | 0.37835 | 0.35738 | 0.396 | 0.39633 | 0.36667
| Median |  | 0.36274 | 0.37225 | 0.37705 | 0.317 | 0.33965 | 0.405 | 0.3738 | 0.359 | 0.36 | 0.385 | 0.359 | 0.3605 | 0.3645 | 0.37314 | 0.3855 | 0.36 | 0.39 | 0.36 | 0.3935 | 0.3775 | 0.36845 | 0.37775 | 0.35542 | 0.396 | 0.4 | 0.37
| Std Dev. |  | 0.0026001 | 0.0059218 | 0.003784 | 0.025657 | 0.0063899 | 0.0098319 | 0.0037664 | 0.0096695 | 0.0030822 | 0.0081976 | 0.0047645 | 0.0069857 | 0.0051251 | 0.0045597 | 0.0021909 | 0.0040825 | 0.0058538 | 0.0083666 | 0.012894 | 0.0043205 | 0.0060798 | 0.0015372 | 0.0080132 | 0.0037947 | 0.013426 | 0.010328
| Rel.Std.Dev. |  | 0.7148% | 1.58% | 1.002% | 8.188% | 1.878% | 2.408% | 1.011% | 2.667% | 0.855% | 2.152% | 1.322% | 1.919% | 1.402% | 1.223% | 0.5676% | 1.129% | 1.507% | 2.357% | 3.326% | 1.147% | 1.651% | 0.4063% | 2.242% | 0.9583% | 3.388% | 2.817%
| PDM3 |  | -2.135% | 0.8251% | 1.595% | -15.7% | -8.485% | 9.854% | 0.2456% | -2.476% | -3.014% | 2.501% | -3.014% | -2.073% | -1.624% | 0.3069% | 3.846% | -2.7% | 4.474% | -4.494% | 4.294% | 1.335% | -0.9158% | 1.788% | -3.852% | 6.536% | 6.626% | -1.355%
| Z-Score (Absolute) |  | 0.48 | 0.31 | 0.52 | 4.11 | 2.18 | 2.72 | 0.16 | 0.57 | 0.72 | 0.76 | 0.72 | 0.46 | 0.34 | 0.17 | 1.12 | 0.63 | 1.29 | 1.11 | 1.24 | 0.45 | 0.16 | 0.57 | 0.94 | 1.81 | 1.86 | 0.27
| Analytical results for Tl in OREAS 285 (Certified Value 0.94 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.92256 | 1.02 | 0.91 | 1.035 | 0.82 | 0.93 | 0.98 | 0.88 | 0.85 | 0.9 | 0.97501 | 0.9 | 0.94 | 0.92 | 1.01 | 1 | 0.98 | 0.96 | 0.96 | 0.92 | 0.92 | [5 empty] | 1
| 1 | 2 | 0.9512 | 0.96 | 0.89 | 1.036 | 0.92 | 0.93 | 0.97 | 0.9 | 0.8 | 0.9 | 1.003 | 0.7 | 0.91 | 0.89 | 1 | 1 | 0.96 | 0.96 | 1.04 | 0.97 | 0.93 | [5 empty] | 30
| 1 | 3 | 0.91846 | 0.97 | 0.91 | 1.04 | 0.9 | 0.88 | 0.98 | 0.97 | 0.85 | 0.93 | 1.0372 | 0.7 | 0.91 | 0.9 | 1.01 | 1 | 0.96 | 0.99 | 0.97 | 1.09 | 0.92 | [5 empty] | 16
| 1 | 4 | 0.89914 | 0.98 | 0.89 | 1.039 | 1.01 | 0.92 | 0.98 | 0.94 | 0.84 | 0.9 | 1.0196 | 0.9 | 0.9 | 0.92 | 0.95 | 1 | 0.93 | 0.97 | 0.97 | 0.98 | 0.91 | [5 empty] | 0.93987
| 1 | 5 | 0.93655 | 0.97 | 0.93 | 1.041 | 0.97 | 0.92 | 0.99 | 0.88 | 0.84 | 0.9 | 0.99016 | 0.6 | 0.89 | 0.92 | 0.98 | 1 | 0.97 | 0.96 | 0.96 | 0.99 | 0.9 | [5 empty] | 63
| 1 | 6 | 0.94439 | 0.98 | 0.95 | 1.053 | 0.97 | 0.89 | 0.96 | 0.94 | 0.86 | 0.92 | 1.0523 | 0.9 | 0.94 | 0.9 | 0.98 | 1 | 0.95 | 0.96 | 0.98 | 0.99 | 0.9
| Mean |  | 0.92872 | 0.98 | 0.91333 | 1.0407 | 0.93167 | 0.91167 | 0.97667 | 0.91833 | 0.84 | 0.90833 | 1.0129 | 0.78333 | 0.915 | 0.90833 | 0.98833 | 1 | 0.95833 | 0.96667 | 0.98 | 0.99 | 0.91333
| Median |  | 0.92955 | 0.975 | 0.91 | 1.0395 | 0.945 | 0.92 | 0.98 | 0.92 | 0.845 | 0.9 | 1.0113 | 0.8 | 0.91 | 0.91 | 0.99 | 1 | 0.96 | 0.96 | 0.97 | 0.985 | 0.915
| Std Dev. |  | 0.019135 | 0.020976 | 0.023381 | 0.0064704 | 0.067355 | 0.02137 | 0.010328 | 0.037103 | 0.020976 | 0.013292 | 0.029118 | 0.13292 | 0.020736 | 0.013292 | 0.023166 | 0 | 0.017224 | 0.012111 | 0.030332 | 0.055498 | 0.012111
| Rel.Std.Dev. |  | 2.06% | 2.14% | 2.56% | 0.6218% | 7.229% | 2.344% | 1.057% | 4.04% | 2.497% | 1.463% | 2.875% | 16.97% | 2.266% | 1.463% | 2.344% | 0% | 1.797% | 1.253% | 3.095% | 5.606% | 1.326%
| PDM3 |  | -1.186% | 4.27% | -2.823% | 10.72% | -0.8725% | -3% | 3.915% | -2.291% | -10.63% | -3.355% | 7.77% | -16.65% | -2.646% | -3.355% | 5.157% | 6.398% | 1.965% | 2.851% | 4.27% | 5.334% | -2.823%
| Z-Score (Absolute) |  | 0.07 | 1.22 | 0.46 | 2.76 | 0 | 0.51 | 1.14 | 0.34 | 2.32 | 0.59 | 2.05 | NA | 0.42 | 0.59 | 1.43 | NA | 0.67 | 0.89 | 1.22 | 1.48 | 0.46
Results from laboratories 1.20 and 1.24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 285 (Certified Value 0.19 ppm) | [27 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.26 | [20 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [20 empty] | 2
3
| 1 | 1 | 0.17767 | 0.18 | 0.2 | 0.1885 | 0.2 | 0.18 | [20 empty] | 1
| 1 | 2 | 0.18099 | 0.19 | 0.2 | 0.1801 | 0.2 | 0.19 | [20 empty] | 5
| 1 | 3 | 0.17463 | 0.2 | 0.2 | 0.153 | 0.2 | 0.19 | [20 empty] | 16
| 1 | 4 | 0.18059 | 0.18 | 0.2 | 0.1759 | 0.2 | 0.19 | [20 empty] | 0.18925
| 1 | 5 | 0.1833 | 0.2 | 0.2 | 0.1789 | 0.2 | 0.18 | [20 empty] | 64
| 1 | 6 | 0.18508 | 0.19 | 0.2 | 0.169 | 0.2 | 0.19
| Mean |  | 0.18038 | 0.19 | 0.2 | 0.17423 | 0.2 | 0.18667
| Median |  | 0.18079 | 0.19 | 0.2 | 0.1774 | 0.2 | 0.19
| Std Dev. |  | 0.0037794 | 0.0089443 | 3.0405e-17 | 0.012172 | 3.0405e-17 | 0.005164
| Rel.Std.Dev. |  | 2.095% | 4.708% | 1.52e-14% | 6.986% | 1.52e-14% | 2.766%
| PDM3 |  | -4.691% | 0.3942% | 5.678% | -7.937% | 5.678% | -1.367%
| Z-Score (Absolute) |  | 0.55 | 0.11 | 0.8 | 0.97 | 0.8 | 0.11
| Analytical results for U in OREAS 285 (Certified Value 4.04 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 3.8433 | 4.19 | 4.3 | 3.4886 | 4.42 | 3.2 | 4.8 | 3.7 | 3.2 | 4.1 | 4.1 | 4.1401 | 4.7 | 3.9 | 3.54 | 3.8 | 3.9 | 3.98 | 4.52 | 4.35 | < 2 | 4.49 | 3.7 |  |  |  | 1
| 1 | 2 | 3.7173 | 4.42 | 4.3 | 3.3621 | 3.88 | 4.3 | 4.7 | 3.9 | 3.3 | 3.3 | 4.4 | 4.9249 | 4.2 | 3.8 | 3.56 | 4.09 | 4.2 | 3.66 | 4.12 | 4.18 | < 2 | 4.22 | 3.1 |  |  |  | 32
| 1 | 3 | 3.7419 | 4.9 | 4.3 | 3.3096 | 4.41 | 4.2 | 4.8 | 4.2 | 3.9 | 3.8 | 3.7 | 4.6786 | 4.7 | 4 | 3.4 | 3.96 | 3.8 | 3.69 | 4.38 | 3.92 | 9 | 4.43 | 3.1 |  |  |  | 16
| 1 | 4 | 3.3967 | 4.74 | 4 | 3.9119 | 4.58 | 4.5 | 4.7 | 4.1 | 3.4 | 3.8 | 4.1 | 4.3838 | 4 | 3.7 | 3.43 | 4.11 | 3.9 | 4.07 | 4.21 | 4.53 | < 2 | 4.34 | 3.4 |  |  |  | 4.0423
| 1 | 5 | 4.0861 | 4.2 | 4.1 | 4.151 | 4.21 | 3.9 | 4.1 | 4.2 | 3.1 | 4.9 | 4.2 | 4.2892 | 4.3 | 3.6 | 3.6 | 3.82 | 3.8 | 4.33 | 4.31 | 4.55 | < 2 | 4.28 | 3.3 |  |  |  | 65
| 1 | 6 | 3.8624 | 4.01 | 4.2 | 3.7163 | 4.38 | 4.6 | 3.9 | 4.2 | 3.3 | 4.2 | 5 | 4.7227 | 4.5 | 4 | 3.57 | 4.15 | 3.8 | 3.83 | 4.32 | 4.33 | < 2 | 4.13 | 3.4
| Mean |  | 3.7746 | 4.41 | 4.2 | 3.6566 | 4.3133 | 4.1167 | 4.5 | 4.05 | 3.3667 | 4.0167 | 4.25 | 4.5232 | 4.4 | 3.8333 | 3.5167 | 3.9883 | 3.9 | 3.9267 | 4.31 | 4.31 | 9 | 4.315 | 3.3333
| Median |  | 3.7926 | 4.31 | 4.25 | 3.6025 | 4.395 | 4.25 | 4.7 | 4.15 | 3.3 | 3.95 | 4.15 | 4.5312 | 4.4 | 3.85 | 3.55 | 4.025 | 3.85 | 3.905 | 4.315 | 4.34 | 9 | 4.31 | 3.35
| Std Dev. |  | 0.22656 | 0.34687 | 0.12649 | 0.33122 | 0.24279 | 0.51153 | 0.39497 | 0.20736 | 0.28048 | 0.53448 | 0.43243 | 0.29878 | 0.28284 | 0.1633 | 0.08165 | 0.15224 | 0.15492 | 0.25398 | 0.13799 | 0.23537 |  | 0.13338 | 0.22509
| Rel.Std.Dev. |  | 6.002% | 7.866% | 3.012% | 9.058% | 5.629% | 12.43% | 8.777% | 5.12% | 8.331% | 13.31% | 10.17% | 6.605% | 6.428% | 4.26% | 2.322% | 3.817% | 3.972% | 6.468% | 3.202% | 5.461% |  | 3.091% | 6.753%
| PDM3 |  | -6.623% | 9.096% | 3.901% | -9.542% | 6.704% | 1.839% | 11.32% | 0.19% | -16.71% | -0.6346% | 5.138% | 11.9% | 8.848% | -5.17% | -13% | -1.336% | -3.521% | -2.861% | 6.622% | 6.622% | 122.6% | 6.746% | -17.54%
| Z-Score (Absolute) |  | 0.71 | 0.92 | 0.38 | 1.01 | 0.67 | 0.17 | 1.15 | 0 | 1.75 | 0.09 | 0.51 | 1.21 | 0.9 | 0.55 | 1.37 | 0.16 | 0.38 | 0.32 | 0.67 | 0.67 | 4.4 | 0.68 | 1.84
| Analytical results for V in OREAS 285 (Certified Value 63 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
1
| 1 | 1 | 66.1323 | 63.8096 | 65 | 61 | 56.4293 | 63 | 62 | 67 | 62 | 61 | 61 | 61 | 64 | 67.1752 | 67 | 59 | 63 | 65.7 | 62 | 64 | 65 | 60 | 65.3 | 67.8 | 59 | 65 | 1
| 1 | 2 | 67.0716 | 64.8288 | 65 | 63 | 58.7097 | 63 | 62 | 70 | 62 | 63 | 61 | 61 | 65 | 63.7981 | 66 | 59 | 62 | 60.9 | 66 | 64 | 66 | 63 | 62.99 | 67.9 | 62.9 | 59 | 16
| 1 | 3 | 66.6996 | 64.1894 | 61 | 62 | 56.8145 | 64 | 63 | 64 | 61 | 65 | 62 | 61 | 67 | 62.8685 | 67 | 59 | 62 | 63.1 | 64 | 64 | 67 | 63 | 64.76 | 64.7 | 66 | 61 | 16
| 1 | 4 | 67.053 | 65.5659 | 63 | 60 | 59.2823 | 64 | 63 | 64 | 62 | 64 | 63 | 61 | 65 | 68.3665 | 66 | 59 | 63 | 62.1 | 64 | 63 | 68 | 62 | 65.67 | 71.4 | 64 | 66 | 63.4261
| 1 | 5 | 66.193 | 64.1635 | 57 | 59 | 59.0532 | 64 | 63 | 66 | 63 | 65 | 61 | 62 | 65 | 60.7492 | 67 | 58 | 65 | 65.1 | 61 | 63 | 67 | 60 | 65.47 | 67 | 65 | 59 | 66
| 1 | 6 | 66.7084 | 66.2118 | 63 | 46 | 57.4481 | 63 | 63 | 66 | 61 | 65 | 61 | 64 | 65 | 62.5394 | 67 | 58 | 64 | 66.3 | 57 | 63 | 66 | 61 | 65.99 | 71.3 | 65 | 60
| Mean |  | 66.643 | 64.7948 | 62.3333 | 58.5 | 57.9562 | 63.5 | 62.6667 | 66.1667 | 61.8333 | 63.8333 | 61.5 | 61.6667 | 65.1667 | 64.2495 | 66.6667 | 58.6667 | 63.1667 | 63.8667 | 62.3333 | 63.5 | 66.5 | 61.5 | 65.03 | 68.35 | 63.65 | 61.6667
| Median |  | 66.704 | 64.5091 | 63 | 60.5 | 58.0789 | 63.5 | 63 | 66 | 62 | 64.5 | 61 | 61 | 65 | 63.3333 | 67 | 59 | 63 | 64.1 | 63 | 63.5 | 66.5 | 61.5 | 65.385 | 67.85 | 64.5 | 60.5
| Std Dev. |  | 0.40561 | 0.93158 | 3.0111 | 6.2849 | 1.2184 | 0.54772 | 0.5164 | 2.2286 | 0.75277 | 1.6021 | 0.83666 | 1.2111 | 0.98319 | 2.9256 | 0.5164 | 0.5164 | 1.169 | 2.1593 | 3.1411 | 0.54772 | 1.0488 | 1.3784 | 1.08 | 2.5944 | 2.509 | 3.0768
| Rel.Std.Dev. |  | 0.6086% | 1.438% | 4.831% | 10.74% | 2.102% | 0.8626% | 0.824% | 3.368% | 1.217% | 2.51% | 1.36% | 1.964% | 1.509% | 4.554% | 0.7746% | 0.8802% | 1.851% | 3.381% | 5.039% | 0.8626% | 1.577% | 2.241% | 1.661% | 3.796% | 3.942% | 4.989%
| PDM3 |  | 5.072% | 2.158% | -1.723% | -7.767% | -8.624% | 0.1165% | -1.197% | 4.321% | -2.511% | 0.6421% | -3.037% | -2.774% | 2.744% | 1.298% | 5.109% | -7.504% | -0.409% | 0.6946% | -1.723% | 0.1165% | 4.846% | -3.037% | 2.529% | 7.763% | 0.353% | -2.774%
| Z-Score (Absolute) |  | 1.27 | 0.52 | 0.47 | 2.02 | 2.24 | 0 | 0.34 | 1.08 | 0.67 | 0.13 | 0.81 | 0.74 | 0.67 | 0.3 | 1.28 | 1.96 | 0.13 | 0.15 | 0.47 | 0 | 1.21 | 0.81 | 0.62 | 1.96 | 0.06 | 0.74
| Analytical results for W in OREAS 285 (Certified Value 17.5 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 17.6309 | 17.2 | 11.5 | 14.214 | 24 | 19.75 | 15.2 | 16.6 | 18.1 | 18.2 | 17.9 | 17.8 | 18.0916 | 18.3 | 15.6 | 15.3 | 18.5 | 18 | 17.6 | 17.6 | 18.1 | 15 | 17.7 | 16.5 |  |  | 1
| 1 | 2 | 17.4727 | 18.1 | 11.4 | 16.3512 | 25 | 19.54 | 17.1 | 16.4 | 17.8 | 17 | 17.6 | 19.2 | 18.3867 | 15.6 | 15.2 | 15.5 | 19.2 | 18.8 | 16.9 | 18.1 | 19.2 | 11 | 18.3 | 17.7 |  |  | 34
| 1 | 3 | 18.4505 | 17.5 | 9.2 | 11.9006 | 23 | 18.94 | 15.9 | 17 | 18.2 | 18.2 | 18.5 | 18.8 | 17.4074 | 14.6 | 16.3 | 15.4 | 18.6 | 18.3 | 17.8 | 18.5 | 19.2 | 12 | 18 | 16.8 |  |  | 16
| 1 | 4 | 17.5831 | 18.1 | 8.9 | 14.5432 | 22 | 18.97 | 17.7 | 17.4 | 18.4 | 17.8 | 18.4 | 18.3 | 18.0312 | 15.3 | 15.9 | 15.6 | 18 | 17.6 | 17 | 18.8 | 17.9 | 9 | 17.6 | 17.3 |  |  | 17.4569
| 1 | 5 | 17.1026 | 17.8 | 5.2 | 16.8194 | 26 | 19.2 | 17.7 | 16.1 | 18.6 | 18.2 | 18.5 | 19 | 17.6486 | 14.8 | 15.6 | 15.2 | 19.7 | 17.3 | 17.6 | 17.7 | 18.7 | 13 | 18.4 | 15.4 |  |  | 67
| 1 | 6 | 17.9687 | 17.8 | 1.8 | 15.9255 | 24 | 19.57 | 16.6 | 16.8 | 19.2 | 17.4 | 18.7 | 17.8 | 17.3666 | 14.4 | 15.9 | 15.1 | 20.9 | 17.2 | 17.5 | 19 | 18.2 | 15 | 18.5 | 16.5
| Mean |  | 17.7014 | 17.75 | 8 | 14.959 | 24 | 19.3283 | 16.7 | 16.7167 | 18.3833 | 17.8 | 18.2667 | 18.4833 | 17.822 | 15.5 | 15.75 | 15.35 | 19.15 | 17.8667 | 17.4 | 18.2833 | 18.55 | 12.5 | 18.0833 | 16.7
| Median |  | 17.607 | 17.8 | 9.05 | 15.2343 | 24 | 19.37 | 16.85 | 16.7 | 18.3 | 18 | 18.45 | 18.55 | 17.8399 | 15.05 | 15.75 | 15.35 | 18.9 | 17.8 | 17.55 | 18.3 | 18.45 | 12.5 | 18.15 | 16.65
| Std Dev. |  | 0.46099 | 0.35071 | 3.8037 | 1.8119 | 1.4142 | 0.33961 | 1.006 | 0.45789 | 0.48339 | 0.50596 | 0.42269 | 0.608 | 0.41103 | 1.4422 | 0.37283 | 0.18708 | 1.0407 | 0.6186 | 0.36332 | 0.57764 | 0.56833 | 2.3452 | 0.37639 | 0.79246
| Rel.Std.Dev. |  | 2.604% | 1.976% | 47.55% | 12.11% | 5.893% | 1.757% | 6.024% | 2.739% | 2.63% | 2.842% | 2.314% | 3.289% | 2.306% | 9.305% | 2.367% | 1.219% | 5.434% | 3.462% | 2.088% | 3.159% | 3.064% | 18.76% | 2.081% | 4.745%
| PDM3 |  | 1.4% | 1.679% | -54.17% | -14.31% | 37.48% | 10.72% | -4.336% | -4.241% | 5.307% | 1.965% | 4.638% | 5.88% | 2.091% | -11.21% | -9.778% | -12.07% | 9.698% | 2.347% | -0.3262% | 4.734% | 6.261% | -28.4% | 3.588% | -4.336%
| Z-Score (Absolute) |  | 0.05 | 0.02 | 7.19 | 2.07 | 4.58 | 1.14 | 0.79 | 0.78 | 0.45 | 0.02 | 0.36 | 0.52 | 0.03 | 1.67 | 1.49 | 1.78 | 1.01 | 0.07 | 0.28 | 0.37 | 0.57 | 3.88 | 0.23 | 0.79
| Analytical results for Y in OREAS 285 (Certified Value 15.9 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 15.1643 | 16.06 | 16.7 | 15.8526 | 15 | 15.51 | 15.3 | 16.2 | 16.2 | 16.6 | 14.4 | 17.2 | 17.092 | 14.9 | 14.9 | 16.2 | 18 | 15 | 16.2 | 15.63 | 16.4 | 18 | 15.2 | 14 |  |  | 1
| 1 | 2 | 15.3287 | 15.87 | 17.1 | 15.7241 | 15 | 15.29 | 16 | 16.4 | 19.9 | 15.7 | 14.2 | 17.4 | 16.788 | 14.4 | 14.6 | 15.9 | 17.2 | 15.3 | 16 | 15.88 | 16.77 | 19 | 16.2 | 13.8 |  |  | 35
| 1 | 3 | 15.1498 | 16.3 | 17.6 | 15.2283 | 15 | 15.66 | 15.8 | 15.8 | 17.6 | 16.7 | 13.9 | 17.7 | 17.1588 | 14.2 | 15.3 | 15.7 | 18.3 | 15.6 | 15.6 | 16.01 | 16.04 | 18 | 17 | 13.6 |  |  | 16
| 1 | 4 | 15.6363 | 16.2 | 17.2 | 15.8522 | 16 | 16.25 | 14.6 | 15.8 | 17.5 | 16.6 | 14.7 | 17.6 | 17.241 | 13.9 | 14.9 | 16 | 18.1 | 14.9 | 15.9 | 16.33 | 15.67 | 18 | 16.7 | 13.9 |  |  | 15.936
| 1 | 5 | 15.5988 | 15.36 | 17.7 | 16.9925 | 15 | 15.79 | 14.8 | 17 | 18 | 16.1 | 14.7 | 17.6 | 16.9754 | 13 | 14.7 | 15.7 | 17.4 | 14.9 | 16.2 | 16.45 | 16.11 | 19 | 16.7 | 14.4 |  |  | 68
| 1 | 6 | 15.8351 | 16.13 | 17.5 | 15.5258 | 15 | 15.73 | 14.9 | 15.8 | 17.7 | 15.8 | 14.9 | 17.6 | 16.8366 | 14.3 | 14.9 | 16.3 | 17.3 | 14.4 | 15.3 | 15.85 | 15.82 | 18 | 17.2 | 13.4
| Mean |  | 15.4521 | 15.9867 | 17.3 | 15.8626 | 15.1667 | 15.705 | 15.2333 | 16.1667 | 17.8167 | 16.25 | 14.4667 | 17.5167 | 17.0153 | 14.1167 | 14.8833 | 15.9667 | 17.7167 | 15.0167 | 15.8667 | 16.025 | 16.135 | 18.3333 | 16.5 | 13.85
| Median |  | 15.4637 | 16.095 | 17.35 | 15.7881 | 15 | 15.695 | 15.1 | 16 | 17.65 | 16.35 | 14.55 | 17.6 | 17.0337 | 14.25 | 14.9 | 15.95 | 17.7 | 14.95 | 15.95 | 15.945 | 16.075 | 18 | 16.7 | 13.85
| Std Dev. |  | 0.27984 | 0.33939 | 0.37417 | 0.60197 | 0.40825 | 0.32161 | 0.56804 | 0.48028 | 1.1957 | 0.44159 | 0.37238 | 0.18348 | 0.18033 | 0.63692 | 0.24014 | 0.25033 | 0.47081 | 0.40702 | 0.3559 | 0.31034 | 0.39964 | 0.5164 | 0.72111 | 0.34496
| Rel.Std.Dev. |  | 1.811% | 2.123% | 2.163% | 3.795% | 2.692% | 2.048% | 3.729% | 2.971% | 6.711% | 2.717% | 2.574% | 1.047% | 1.06% | 4.512% | 1.613% | 1.568% | 2.657% | 2.71% | 2.243% | 1.937% | 2.477% | 2.817% | 4.37% | 2.491%
| PDM3 |  | -3.036% | 0.3182% | 8.56% | -0.4604% | -4.827% | -1.449% | -4.409% | 1.448% | 11.8% | 1.971% | -9.22% | 9.919% | 6.773% | -11.42% | -6.605% | 0.1927% | 11.17% | -5.769% | -0.4348% | 0.5587% | 1.249% | 15.04% | 3.539% | -13.09%
| Z-Score (Absolute) |  | 0.56 | 0.01 | 1.41 | 0.12 | NA | 0.29 | 0.79 | 0.2 | 1.96 | 0.29 | 1.61 | 1.64 | 1.1 | 1.98 | 1.16 | 0.01 | 1.85 | 1.02 | 0.12 | 0.05 | 0.17 | NA | 0.56 | 2.26
Results from laboratories 1.11 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 285 (Certified Value 1.15 ppm) | [27 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.13 | 1.21 | 1.25 | 1.26 | 1.29 | 1.32 | [16 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [16 empty] | 2
3
| 1 | 1 | 1.1741 | 1.14 | 1.4 | 1.132 | 0.9 | 1.2 | 1.1 | 1.21 | 1.4 | 1.1 | [16 empty] | 1
| 1 | 2 | 1.1669 | 1.15 | 1.3 | 1.0494 | 1.1 | 1.2 | 1.1 | 1.25 | 1.4 | 1.1 | [16 empty] | 36
| 1 | 3 | 1.2267 | 1.18 | 1.4 | 0.9444 | 1.1 | 1.2 | 1.1 | 1.21 | 1.4 | 1.1 | [16 empty] | 16
| 1 | 4 | 1.195 | 1.1 | 1.3 | 1.0369 | 1.1 | 1.1 | 1.1 | 1.65 | 1.4 | 1.1 | [16 empty] | 1.1544
| 1 | 5 | 1.2095 | 1.15 | 1.4 | 1.1609 | 1 | 1.1 | 1.1 | 1.22 | 1.4 | 1.1 | [16 empty] | 69
| 1 | 6 | 1.2027 | 1.14 | 1.4 | 1.0492 | 1 | 1.2 | 1.1 | 1.22 | 1.4 | 1.1
| Mean |  | 1.1958 | 1.1433 | 1.3667 | 1.0621 | 1.0333 | 1.1667 | 1.1 | 1.2933 | 1.4 | 1.1
| Median |  | 1.1988 | 1.145 | 1.4 | 1.0493 | 1.05 | 1.2 | 1.1 | 1.22 | 1.4 | 1.1
| Std Dev. |  | 0.022316 | 0.02582 | 0.05164 | 0.076764 | 0.08165 | 0.05164 | 2.4324e-16 | 0.17535 | 2.4324e-16 | 2.4324e-16
| Rel.Std.Dev. |  | 1.866% | 2.258% | 3.779% | 7.227% | 7.902% | 4.426% | 2.211e-14% | 13.56% | 1.737e-14% | 2.211e-14%
| PDM3 |  | 3.584% | -0.9619% | 18.38% | -7.996% | -10.49% | 1.059% | -4.716% | 12.03% | 21.27% | -4.716%
| Z-Score (Absolute) |  | 0.37 | 0.11 | 1.93 | 0.85 | 1.11 | 0.11 | 0.5 | 1.26 | 2.23 | 0.5
| Analytical results for Zn in OREAS 285 (Certified Value 81 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 83.4395 | 80.5011 | 87 | 81 | 76.7195 | 76 | 80 | 83.5 | 80 | 84 | 77 | 77 | 84 | 80.3923 | 78 | 81 | 80 | 86.2 | 80 | 78 | 81 | 80 | 94 | 99 | 84 |  | 1
| 1 | 2 | 82.3845 | 81.2141 | 85 | 83 | 80.6977 | 78 | 80 | 87.2 | 80 | 86 | 76 | 77 | 83 | 82.1555 | 77 | 79 | 79 | 83.2 | 83 | 79 | 81 | 81 | 96 | 106 | 85 |  | 17
| 1 | 3 | 82.4735 | 80.357 | 85 | 83 | 85.5505 | 76 | 82 | 80.1 | 78 | 90 | 77 | 76 | 85 | 82.1671 | 77 | 82 | 79 | 83.7 | 82 | 78 | 82 | 82 | 97 | 142 | 85 |  | 16
| 1 | 4 | 82.9665 | 83.2278 | 85 | 83 | 77.9682 | 76 | 80 | 80.6 | 80 | 89 | 78 | 78 | 83 | 80.0081 | 81 | 80 | 81 | 83.1 | 82 | 77 | 83 | 81 | 96 | 112 | 84 |  | 81.0972
| 1 | 5 | 82.227 | 81.6442 | 78 | 82 | 83.7408 | 75 | 81 | 81 | 80 | 88 | 77 | 79 | 84 | 80.9169 | 77 | 79 | 80 | 81.1 | 80 | 78 | 82 | 79 | 97 | 110 | 83 |  | 70
| 1 | 6 | 82.085 | 83.445 | 85 | 83 | 79.4614 | 78 | 79 | 82.4 | 79 | 88 | 77 | 83 | 83 | 81.9642 | 80 | 80 | 80 | 85.8 | 75 | 78 | 82 | 81 | 97 | 111 | 81
| Mean |  | 82.596 | 81.7315 | 84.1667 | 82.5 | 80.6897 | 76.5 | 80.3333 | 82.4667 | 79.5 | 87.5 | 77 | 78.3333 | 83.6667 | 81.2673 | 78.3333 | 80.1667 | 79.8333 | 83.85 | 80.3333 | 78 | 81.8333 | 80.6667 | 96.1667 | 113.3333 | 83.6667
| Median |  | 82.429 | 81.4291 | 85 | 83 | 80.0795 | 76 | 80 | 81.7 | 80 | 88 | 77 | 77.5 | 83.5 | 81.4405 | 77.5 | 80 | 80 | 83.45 | 81 | 78 | 82 | 81 | 96.5 | 110.5 | 84
| Std Dev. |  | 0.51103 | 1.3307 | 3.1252 | 0.83666 | 3.3945 | 1.2247 | 1.0328 | 2.6349 | 0.83666 | 2.1679 | 0.63246 | 2.5033 | 0.8165 | 0.95477 | 1.7512 | 1.169 | 0.75277 | 1.8918 | 2.8752 | 0.63246 | 0.75277 | 1.0328 | 1.169 | 14.8279 | 1.5055
| Rel.Std.Dev. |  | 0.6187% | 1.628% | 3.713% | 1.014% | 4.207% | 1.601% | 1.286% | 3.195% | 1.052% | 2.478% | 0.8214% | 3.196% | 0.9759% | 1.175% | 2.236% | 1.458% | 0.9429% | 2.256% | 3.579% | 0.8108% | 0.9199% | 1.28% | 1.216% | 13.08% | 1.799%
| PDM3 |  | 1.848% | 0.7822% | 3.785% | 1.73% | -0.5024% | -5.669% | -0.9419% | 1.689% | -1.969% | 7.895% | -5.052% | -3.408% | 3.168% | 0.2098% | -3.408% | -1.147% | -1.558% | 3.395% | -0.9419% | -3.819% | 0.9078% | -0.5308% | 18.58% | 39.75% | 3.168%
| Z-Score (Absolute) |  | 0.51 | 0.18 | 1.11 | 0.47 | 0.22 | 1.82 | 0.36 | 0.46 | 0.67 | 2.38 | 1.63 | 1.12 | 0.92 | 0 | 1.12 | 0.42 | 0.55 | 0.99 | 0.36 | 1.25 | 0.22 | 0.23 | 5.68 | 12.23 | 0.92
| Analytical results for Zr in OREAS 285 (Certified Value 58 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 60.4 | 56.1341 | 58.8 | 75 | 52.4331 | 49 | 50.1 | 42.8 | 65.5 | 57.1 | 61.8 | 58.3 | 62.4 | 59.3904 | 60.3 | 73.1 | 47.3 | 56 | 54 | 56.5 | 56.6 | 61.8 | 54 | 57.3 | 59.9 |  | 1
| 1 | 2 | 61.03 | 57.3094 | 68.4 | 79 | 52.0053 | 47 | 48.1 | 57.8 | 65.3 | 61.1 | 59.4 | 58.7 | 62.9 | 56.0371 | 60.4 | 70.8 | 47.4 | 58 | 53.6 | 55.8 | 57.7 | 57.4 | 56 | 61.3 | 59.5 |  | 18
| 1 | 3 | 61.25 | 56.1213 | 68.5 | 73 | 52.2758 | 50 | 51.2 | 60.2 | 63.5 | 62.1 | 64 | 55.3 | 65.4 | 56.5933 | 57.7 | 72.6 | 48.4 | 59.7 | 54.5 | 57 | 59.6 | 57.9 | 73 | 58.3 | 58.5 |  | 16
| 1 | 4 | 60.83 | 58.4049 | 72.2 | 73 | 68.3975 | 52 | 48.6 | 52.3 | 63.5 | 61.7 | 61.8 | 56.9 | 64.4 | 57.7086 | 58.5 | 71.7 | 46.8 | 60.9 | 54 | 54.8 | 63.1 | 59.7 | 57 | 57.5 | 59.7 |  | 57.5206
| 1 | 5 | 61.61 | 57.2266 | 56.2 | 75 | 55.1647 | 47 | 49.4 | 53.2 | 67.4 | 62.8 | 59 | 60.4 | 65.3 | 57.8853 | 59 | 74.8 | 47.2 | 58.3 | 53 | 55.5 | 60 | 62.2 | 56 | 60.6 | 55.6 |  | 71
| 1 | 6 | 61.92 | 58.317 | 65.3 | 65 | 54.3087 | 47 | 48.8 | 47.3 | 62.1 | 63 | 60.6 | 59.8 | 64.1 | 58.2926 | 60 | 73.3 | 46.2 | 59.9 | 50.6 | 55.6 | 62 | 65.8 | 54 | 57.8 | 57.8
| Mean |  | 61.1733 | 57.2522 | 64.9 | 73.3333 | 55.7642 | 48.6667 | 49.3667 | 52.2667 | 64.55 | 61.3 | 61.1 | 58.2333 | 64.0833 | 57.6512 | 59.3167 | 72.7167 | 47.2167 | 58.8 | 53.2833 | 55.8667 | 59.8333 | 60.8 | 58.3333 | 58.8 | 58.5
| Median |  | 61.14 | 57.268 | 66.85 | 74 | 53.3709 | 48 | 49.1 | 52.75 | 64.4 | 61.9 | 61.2 | 58.5 | 64.25 | 57.7969 | 59.5 | 72.85 | 47.25 | 59 | 53.8 | 55.7 | 59.8 | 60.75 | 56 | 58.05 | 59
| Std Dev. |  | 0.54599 | 0.99957 | 6.1897 | 4.6332 | 6.3162 | 2.0656 | 1.1325 | 6.4609 | 1.8865 | 2.1735 | 1.8407 | 1.8843 | 1.2287 | 1.2015 | 1.0944 | 1.3819 | 0.72778 | 1.7413 | 1.4063 | 0.77889 | 2.4679 | 3.135 | 7.2847 | 1.7135 | 1.631
| Rel.Std.Dev. |  | 0.8925% | 1.746% | 9.537% | 6.318% | 11.33% | 4.244% | 2.294% | 12.36% | 2.923% | 3.546% | 3.013% | 3.236% | 1.917% | 2.084% | 1.845% | 1.9% | 1.541% | 2.961% | 2.639% | 1.394% | 4.125% | 5.156% | 12.49% | 2.914% | 2.788%
| PDM3 |  | 6.35% | -0.4666% | 12.83% | 27.49% | -3.054% | -15.39% | -14.18% | -9.134% | 12.22% | 6.57% | 6.223% | 1.239% | 11.41% | 0.227% | 3.122% | 26.42% | -17.91% | 2.224% | -7.367% | -2.875% | 4.021% | 5.701% | 1.413% | 2.224% | 1.703%
| Z-Score (Absolute) |  | 0.64 | 0.42 | 1.65 | 3.92 | 0.82 | 2.73 | 2.54 | 1.76 | 1.55 | 0.67 | 0.62 | 0.15 | 1.43 | 0.31 | 0.14 | 3.75 | 3.12 | 0 | 1.49 | 0.79 | 0.28 | 0.54 | 0.13 | 0 | 0.08
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C934: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1100: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 285 (Certified Value 0.084 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | < 0.1 | 0.09 | 0.60269 | < 1 | 0.09 | < 0.5 | < 0.3 | 0.062 | 0.07 | 0.09 | 0.08 | 0.09 | 0.08 | 0.085954 | < 0.1 | 0.08 | 0.07 | 0.08 | 0.086 | 0.09 | 0.06 | 4.91 | < 2 | 0.23 | 0.09 |  |  |  |  | 1
| 1 | 2 | < 0.1 | 0.09 | 0.48672 | < 1 | 0.088 | < 0.5 | < 0.3 | 0.072 | 0.08 | 0.09 | 0.09 | 0.09 | 0.08 | 0.084591 | < 0.1 | 0.08 | 0.07 | 0.08 | 0.084 | 0.09 | 0.06 | 0.1 | < 2 | 0.232 | 0.08 |  |  |  |  | 22
| 1 | 3 | < 0.1 | 0.09 | 0.31434 | < 1 | 0.094 | < 0.5 | < 0.3 | 0.072 | 0.08 | 0.07 | 0.09 | 0.08 | 0.07 | 0.08526 | < 0.1 | 0.07 | 0.07 | 0.08 | 0.083 | 0.09 | 0.05 | 0.1 | < 2 | 0.224 | 0.07 |  |  |  |  | 16
| 1 | 4 | < 0.1 | 0.09 | 0.27708 | < 1 | 0.093 | < 0.5 | < 0.3 | 0.052 | 0.08 | 0.09 | 0.09 | 0.09 | 0.1 | 0.086325 | < 0.1 | 0.08 | 0.06 | 0.09 | 0.082 | 0.08 | 0.06 | 0.09 | < 2 | 0.23 | 0.09 |  |  |  |  | 0.084067
| 1 | 5 | < 0.1 | 0.09 | 0.1571 | < 1 | 0.097 | < 0.5 | < 0.3 | 0.062 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.086042 | < 0.1 | 0.08 | 0.07 | 0.1 | 0.077 | 0.08 | 0.06 | 0.09 | < 2 | 0.25 | 0.08 |  |  |  |  | 73
| 1 | 6 | < 0.1 | 0.09 | 0.35456 | < 1 | 0.095 | < 0.5 | < 0.3 | 0.062 | 0.07 | 0.09 | 0.1 | 0.08 | 0.08 | 0.084844 | < 0.1 | 0.07 | 0.07 | 0.09 | 0.08 | 0.08 | 0.06 | 0.09 | < 2 | 0.21 | 0.08
| Mean |  |  | 0.09 | 0.36541 |  | 0.092833 |  |  | 0.063667 | 0.076667 | 0.085 | 0.088333 | 0.086667 | 0.081667 | 0.085503 |  | 0.076667 | 0.068333 | 0.086667 | 0.082 | 0.085 | 0.058333 | 0.89667 |  | 0.22933 | 0.081667
| Median |  |  | 0.09 | 0.33445 |  | 0.0935 |  |  | 0.062 | 0.08 | 0.09 | 0.09 | 0.09 | 0.08 | 0.085607 |  | 0.08 | 0.07 | 0.085 | 0.0825 | 0.085 | 0.06 | 0.095 |  | 0.23 | 0.08
| Std Dev. |  |  | 1.5202e-17 | 0.15807 |  | 0.0033116 |  |  | 0.0075277 | 0.005164 | 0.0083666 | 0.0075277 | 0.005164 | 0.0098319 | 0.00070628 |  | 0.005164 | 0.0040825 | 0.008165 | 0.0031623 | 0.0054772 | 0.0040825 | 1.9661 |  | 0.012941 | 0.0075277
| Rel.Std.Dev. |  |  | 1.689e-14% | 43.26% |  | 3.567% |  |  | 11.82% | 6.736% | 9.843% | 8.522% | 5.958% | 12.04% | 0.826% |  | 6.736% | 5.974% | 9.421% | 3.856% | 6.444% | 6.999% | 219.3% |  | 5.643% | 9.218%
| PDM3 |  |  | 7.058% | 334.7% |  | 10.43% |  |  | -24.27% | -8.803% | 1.11% | 5.075% | 3.093% | -2.855% | 1.708% |  | -8.803% | -18.72% | 3.093% | -2.459% | 1.11% | -30.61% | 966.6% |  | 172.8% | -2.855%
| Z-Score (Absolute) |  | 2.71 | 0.34 | 21.36 | 31.63 | 0.56 | 12.55 | 4.92 | 1.67 | 0.67 | 0.04 | 0.22 | 0.09 | 0.29 | 0 | 2.71 | 0.67 | 1.31 | 0.09 | 0.27 | 0.04 | 2.07 | 61.9 | 69.79 | 10.98 | 0.29
| Analytical results for Al in OREAS 285 (Certified Value 2.15 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 2.0634 | 2.0402 | 2.2914 | 2.12 | 2.246 | 2.21 | 1.91 | 1.99 | 2.45 | 2.08 | 2.12 | 2.17 | 2.1221 | 2.08 | 1.88 | 2.37 | 1.79 | 2.12 | 2 | 1.9522 | 2.0464 | 2.245 | 2.212 | 2.448 | 2.12 | 2.29 |  |  |  | 1
| 1 | 2 | 2.0768 | 2.0356 | 2.2797 | 2.17 | 2.3187 | 2.22 | 2.07 | 1.98 | 2.24 | 2.07 | 2.12 | 2.15 | 2.1423 | 2.05 | 1.87 | 2.34 | 1.83 | 2.11 | 2.03 | 1.9838 | 2.0287 | 2.22 | 2.149 | 2.41 | 2.63 | 2.02 |  |  |  | #N/A
| 1 | 3 | 2.0831 | 2.0064 | 2.2857 | 2.1 | 2.2544 | 2.24 | 2.02 | 2.02 | 2.46 | 2.04 | 2.15 | 2.15 | 2.1325 | 2.02 | 2.02 | 2.37 | 1.76 | 2.12 | 1.99 | 2.0489 | 2.1452 | 2.173 | 2.149 | 2.442 | 2.52 | 2.17 |  |  |  | 16
| 1 | 4 | 2.0391 | 2.0922 | 2.2945 | 2.29 | 2.2666 | 2.21 | 1.87 | 2.05 | 2.49 | 2.06 | 2.15 | 2.12 | 2.1063 | 2.05 | 2.03 | 2.38 | 1.75 | 2.07 | 2 | 1.9485 | 2.0114 | 2.405 | 2.249 | 2.462 | 2.35 | 2.2 |  |  |  | 2.1454
| 1 | 5 | 2.0646 | 2.1517 | 2.3103 | 2.06 | 2.3265 | 2.22 | 1.92 | 1.99 | 2.47 | 2.03 | 2.18 | 2.13 | 2.153 | 2.03 | 1.99 | 2.36 | 1.8 | 2.05 | 1.99 | 1.9743 | 2.0752 | 2.394 | 2.159 | 2.509 | 2.35 | 2.24 |  |  |  | 74
| 1 | 6 | 2.0582 | 2.1394 | 2.2463 | 2.03 | 2.3455 | 2.26 | 2.11 | 2 | 2.43 | 2.05 | 2.18 | 2.12 | 2.0959 | 2.07 | 2.01 | 2.3 | 1.79 | 2.1 | 1.99 | 1.9728 | 2.1102 | 2.513 | 2.191 | 2.383 | 2.13 | 2.26
| Mean |  | 2.0642 | 2.0776 | 2.2846 | 2.1283 | 2.2929 | 2.2267 | 1.9833 | 2.005 | 2.4233 | 2.055 | 2.15 | 2.14 | 2.1253 | 2.05 | 1.9667 | 2.3533 | 1.7867 | 2.095 | 2 | 1.9801 | 2.0695 | 2.325 | 2.1848 | 2.4423 | 2.35 | 2.1967
| Median |  | 2.064 | 2.0662 | 2.2886 | 2.11 | 2.2927 | 2.22 | 1.97 | 1.995 | 2.455 | 2.055 | 2.15 | 2.14 | 2.1273 | 2.05 | 2 | 2.365 | 1.79 | 2.105 | 1.995 | 1.9735 | 2.0608 | 2.3195 | 2.175 | 2.445 | 2.35 | 2.22
| Std Dev. |  | 0.015367 | 0.059591 | 0.021444 | 0.092826 | 0.042273 | 0.019664 | 0.097091 | 0.025884 | 0.092014 | 0.018708 | 0.026833 | 0.02 | 0.021659 | 0.022804 | 0.072296 | 0.029439 | 0.028752 | 0.02881 | 0.015492 | 0.036362 | 0.050977 | 0.13193 | 0.040321 | 0.043427 | 0.20425 | 0.096471
| Rel.Std.Dev. |  | 0.7445% | 2.868% | 0.9386% | 4.361% | 1.844% | 0.8831% | 4.895% | 1.291% | 3.797% | 0.9104% | 1.248% | 0.9346% | 1.019% | 1.112% | 3.676% | 1.251% | 1.609% | 1.375% | 0.7746% | 1.836% | 2.463% | 5.675% | 1.845% | 1.778% | 8.692% | 4.392%
| PDM3 |  | -3.784% | -3.159% | 6.492% | -0.7938% | 6.879% | 3.79% | -7.553% | -6.543% | 12.96% | -4.212% | 0.2162% | -0.2499% | -0.933% | -4.445% | -8.329% | 9.694% | -16.72% | -2.347% | -6.776% | -7.704% | -3.535% | 8.373% | 1.84% | 13.84% | 9.539% | 2.391%
| Z-Score (Absolute) |  | 0.38 | 0.3 | 0.96 | 0.01 | 1.01 | 0.61 | 0.87 | 0.74 | 1.8 | 0.44 | 0.14 | 0.08 | 0.01 | 0.47 | 0.97 | 1.38 | 2.07 | 0.19 | 0.77 | 0.89 | 0.35 | 1.21 | 0.35 | 1.92 | 1.36 | 0.42
| Analytical results for As in OREAS 285 (Certified Value 276 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | 0
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0
| 1 | 1 | 259.47 | 282.6008 | 290 | 266.426 | 248.1022 | 306 | 268.254 | 274 | 267.1 | 268.6 | 278 | 315 | 281 | 287 | 279 | 275.4014 | 276.7 | 252 | 277 | 282 | 295.9 | 282 | 265 | 289 | 280.34 | 265 | 279.6 | 297 | 279 | 1
| 1 | 2 | 262.26 | 281.5739 | 292 | 272.857 | 255.5576 | 307 | 264.4039 | 275 | 269.8 | 288.1 | 279 | 310 | 278 | 288 | 277 | 265.1442 | 271 | 258 | 280 | 287 | 296.8 | 285 | 266 | 280 | 278.03 | 258 | 279.9 | 312 | 251 | 12
| 1 | 3 | 259.17 | 283.7905 | 294 | 273.15 | 258.9893 | 314 | 261.7878 | 276 | 283.1 | 285.7 | 285 | 315 | 272 | 293 | 278 | 276.5259 | 268.8 | 263 | 278 | 276 | 300 | 282 | 275 | 292 | 274.94 | 266 | 282.8 | 312 | 268 | 16
| 1 | 4 | 265.28 | 274.8806 | 293 | 263.73 | 256.815 | 319 | 274.9343 | 271 | 265.4 | 269.4 | 285 | 316 | 276 | 291 | 277 | 275.703 | 270.4 | 265 | 282 | 279 | 296.9 | 279 | 263 | 280 | 270.82 | 269 | 285.9 | 307 | 271 | 275.5946
| 1 | 5 | 261.88 | 282.6111 | 298 | 283.068 | 254.8517 | 309 | 265.1273 | 271 | 268.8 | 272 | 280 | 316 | 274 | 297 | 276 | 266.9828 | 275.8 | 267 | 273 | 275 | 289.9 | 309 | 268 | 282 | 270.21 | 262 | 287.5 | 309 | 273 | 75
| 1 | 6 | 258.93 | 282.4016 | 295 | 272.198 | 252.5442 | 306 | 272.4912 | 276 | 265.4 | 296.5 | 282 | 313 | 278 | 296 | 277 | 265.3353 | 269.5 | 262 | 285 | 276 | 297.9 | 289 | 268 | 283 | 270.59 | 260 | 274.2 | 303 | 277
| Mean |  | 261.165 | 281.3098 | 293.6667 | 271.9048 | 254.4767 | 310.1667 | 267.8331 | 273.8333 | 269.9333 | 280.05 | 281.5 | 314.1667 | 276.5 | 292 | 277.3333 | 270.8488 | 272.0333 | 261.1667 | 279.1667 | 279.1667 | 296.2333 | 287.6667 | 267.5 | 284.3333 | 274.155 | 263.3333 | 281.65 | 306.6667 | 269.8333
| Median |  | 260.675 | 282.5012 | 293.5 | 272.5275 | 255.2047 | 308 | 266.6907 | 274.5 | 267.95 | 278.85 | 281 | 315 | 277 | 292 | 277 | 271.1921 | 270.7 | 262.5 | 279 | 277.5 | 296.85 | 283.5 | 267 | 282.5 | 272.88 | 263.5 | 281.35 | 308 | 272
| Std Dev. |  | 2.4695 | 3.2283 | 2.7325 | 6.6934 | 3.7818 | 5.2694 | 5.0582 | 2.3166 | 6.6902 | 11.6332 | 3.0166 | 2.3166 | 3.2094 | 4.0988 | 1.0328 | 5.5571 | 3.3637 | 5.4191 | 4.1673 | 4.6224 | 3.4045 | 10.9848 | 4.1352 | 5.0067 | 4.3193 | 4.0825 | 4.8244 | 5.8195 | 10.0482
| Rel.Std.Dev. |  | 0.9456% | 1.148% | 0.9305% | 2.462% | 1.486% | 1.699% | 1.889% | 0.846% | 2.478% | 4.154% | 1.072% | 0.7374% | 1.161% | 1.404% | 0.3724% | 2.052% | 1.237% | 2.075% | 1.493% | 1.656% | 1.149% | 3.819% | 1.546% | 1.761% | 1.575% | 1.55% | 1.713% | 1.898% | 3.724%
| PDM3 |  | -5.236% | 2.074% | 6.557% | -1.339% | -7.663% | 12.54% | -2.816% | -0.6391% | -2.054% | 1.617% | 2.143% | 14% | 0.3285% | 5.953% | 0.6309% | -1.722% | -1.292% | -5.235% | 1.296% | 1.296% | 7.489% | 4.38% | -2.937% | 3.171% | -0.5224% | -4.449% | 2.197% | 11.27% | -2.09%
| Z-Score (Absolute) |  | 1.47 | 0.36 | 1.49 | 0.49 | 2.08 | 2.99 | 0.87 | 0.32 | 0.67 | 0.25 | 0.38 | 3.36 | 0.08 | 1.34 | 0 | 0.59 | 0.48 | 1.47 | 0.17 | 0.17 | 1.72 | 0.94 | 0.9 | 0.64 | 0.29 | 1.28 | 0.39 | 2.67 | 0.68
| Analytical results for B in OREAS 285 (Indicative Value 10.2 ppm) | [30 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.26 | 1.28 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | [11 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 15g | 01g | 0.5g | 0.25g | [11 empty] | 1
| 1 | 1 | 5.0686 | < 10 | 3 | 3.8751 | 17 | < 20 | < 10 | 10 | < 10 | 10 | < 10 | < 10 | < 10 | 19 | 3 | < 10 | 23.68 | < 10 | [11 empty] | 1
| 1 | 2 | 4.996 | < 10 | 3 | 3.939 | 15 | < 20 | < 10 | < 10 | < 10 | 10 | < 10 | < 10 | < 10 | 19 | 3 | < 10 |  | < 10 | [11 empty] | 1
| 1 | 3 | 4.5321 | < 10 | 3 | 3.8152 | 16 | < 20 | < 10 | 10 | < 10 | 10 | < 10 | < 10 | < 10 | 18 | 3 | < 10 | 20.69 | < 10 | [11 empty] | 16
| 1 | 4 | 4.701 | < 10 | 3 | 7.8871 | 14 | < 20 | < 10 | 10 | < 10 | 10 | < 10 | < 10 | < 10 | 18 | 3 | < 10 | 21.35 | < 10 | [11 empty] | 10.2263
| 1 | 5 | 5.325 | < 10 | 3 | 6.3715 | 16 | < 20 | < 10 | 10 | < 10 | 10 | < 10 | < 10 | < 10 | 17 | 3 | < 10 | 23.37 | < 10 | [11 empty] | 10
| 1 | 6 | 4.7868 | < 10 | 3 | 9.0207 | 14 | < 20 | < 10 | 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 17 | 3 | < 10 | 20.83 | < 10
| Mean |  | 4.9016 |  | 3 | 5.8181 | 15.3333 |  |  | 10 |  | 10 |  |  |  | 18 | 3 |  | 21.984
| Median |  | 4.8914 |  | 3 | 5.1553 | 15.5 |  |  | 10 |  | 10 |  |  |  | 18 | 3 |  | 21.35
| Std Dev. |  | 0.28497 |  | 0 | 2.2874 | 1.2111 |  |  | 0 |  | 0 |  |  |  | 0.89443 | 0 |  | 1.4323
| Rel.Std.Dev. |  | 5.814% |  | 0% | 39.32% | 7.898% |  |  | 0% |  | 0% |  |  |  | 4.969% | 0% |  | 6.515%
| PDM3 |  | -52.07% |  | -70.66% | -43.11% | 49.94% |  |  | -2.213% |  | -2.213% |  |  |  | 76.02% | -70.66% |  | 115%
| Z-Score (Absolute) |  | 0.14 | 0 | 2.94 | 1.2 | 15.21 | 7.36 | 0 | 6.13 | 0 | 6.13 | 0 | 0 | 0 | 19.13 | 2.94 | 0 | 24.99 | 0
| Analytical results for Ba in OREAS 285 (Certified Value 496 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  |  | 0
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 0
| 1 | 1 | 501.2745 | 499 | 599.9424 | 483.6185 | 173 | 490.2666 | 422 | 495.9 | 500 | 560 | 500 | 500 | 510 | 466.9255 | 563 | 433 | 522 | 505 | 467.3 | 486 | 459 | 535 | 225 | 514.62 | 458 | 495 |  |  |  | 1
| 1 | 2 | 504.3445 | 506 | 592.2681 | 499.0516 | 189 | 489.9808 | 438 | 529.3 | 500 | 510 | 500 | 500 | 500 | 457.1371 | 562 | 443 | 526 | 513 | 408.3 | 493 | 466 | 525 | 238 | 507.51 | 503 | 442 |  |  |  | 13
| 1 | 3 | 509.8091 | 501 | 592.1789 | 497.1182 | 178 | 492.6971 | 431 | 518.6 | 500 | 570 | 490 | 510 | 510 | 467.6231 | 554 | 459 | 522 | 491 | 480.8 | 487 | 481 | 541 | 237 | 512.37 | 497 | 471 |  |  |  | 16
| 1 | 4 | 501.0536 | 511 | 605.7669 | 485.053 | 184 | 497.6884 | 428 | 485.7 | 510 | 570 | 500 | 510 | 510 | 456.5953 | 575 | 459 | 528 | 505 | 430.8 | 490 | 458 | 527 | 224 | 514.44 | 470 | 479 |  |  |  | 496.0441
| 1 | 5 | 505.723 | 522 | 584.811 | 484.1127 | 190 | 506.7732 | 434 | 491.3 | 500 | 570 | 490 | 520 | 510 | 444.1564 | 578 | 463 | 511 | 482 | 428.6 | 489 | 466 | 525 | 227 | 524.34 | 480 | 483 |  |  |  | 76
| 1 | 6 | 503.4557 | 512 | 582.9099 | 495.435 | 191 | 505.2107 | 443 | 540.6 | 500 | 560 | 500 | 510 | 500 | 449.9437 | 556 | 455 | 514 | 481 | 439 | 490 | 462 | 525 | 227 | 511.56 | 460 | 492
| Mean |  | 504.2767 | 508.5 | 592.9795 | 490.7315 | 184.1667 | 497.1028 | 432.6667 | 510.2333 | 501.6667 | 556.6667 | 496.6667 | 508.3333 | 506.6667 | 457.0635 | 564.6667 | 452 | 520.5 | 496.1667 | 442.4667 | 489.1667 | 465.3333 | 529.6667 | 229.6667 | 514.14 | 478 | 477
| Median |  | 503.9001 | 508.5 | 592.2235 | 490.244 | 186.5 | 495.1927 | 432.5 | 507.25 | 500 | 565 | 500 | 510 | 510 | 456.8662 | 562.5 | 457 | 522 | 498 | 434.9 | 489.5 | 464 | 526 | 227 | 513.405 | 475 | 481
| Std Dev. |  | 3.249 | 8.4083 | 8.7323 | 7.1942 | 7.3052 | 7.4359 | 7.4207 | 22.4565 | 4.0825 | 23.3809 | 5.164 | 7.5277 | 5.164 | 9.2333 | 9.8319 | 11.5758 | 6.6858 | 13.3928 | 26.8101 | 2.4833 | 8.3825 | 6.7725 | 6.186 | 5.6233 | 18.8574 | 19.2354
| Rel.Std.Dev. |  | 0.6443% | 1.654% | 1.473% | 1.466% | 3.967% | 1.496% | 1.715% | 4.401% | 0.8138% | 4.2% | 1.04% | 1.481% | 1.019% | 2.02% | 1.741% | 2.561% | 1.284% | 2.699% | 6.059% | 0.5077% | 1.801% | 1.279% | 2.693% | 1.094% | 3.945% | 4.033%
| PDM3 |  | 1.66% | 2.511% | 19.54% | -1.071% | -62.87% | 0.2134% | -12.78% | 2.86% | 1.133% | 12.22% | 0.1255% | 2.477% | 2.141% | -7.858% | 13.83% | -8.879% | 4.93% | 0.02471% | -10.8% | -1.386% | -6.191% | 6.778% | -53.7% | 3.648% | -3.638% | -3.839%
| Z-Score (Absolute) |  | 0.26 | 0.4 | 3.34 | 0.21 | 10.88 | 0.01 | 2.23 | 0.46 | 0.17 | 2.08 | 0.01 | 0.4 | 0.34 | 1.39 | 2.36 | 1.56 | 0.82 | 0.02 | 1.89 | 0.27 | 1.1 | 1.14 | 9.3 | 0.6 | 0.66 | 0.69
| Analytical results for Be in OREAS 285 (Certified Value 1.61 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 1.6418 | 1.6 | 2.1564 | 1.9325 | 1.5 | 1.5409 | < 5 | 1.48 | 1.71 | 1.52 | 1.6 | 1.61 | 1.4453 | 2 | 1.6 | 1.7 | 1.43 | 1.7 | 1.66 | 1.22 | 1.48 | 1.6 | [7 empty] | 1
| 1 | 2 | 1.616 | 1.59 | 2.0521 | 1.9672 | 1.5 | 1.6268 | < 5 | 1.48 | 1.59 | 1.52 | 1.6 | 1.58 | 1.5221 | 1 | 1.6 | 1.7 | 1.49 | 1.7 | 1.58 | 1.22 | 1.16 | 1.4 | [7 empty] | 25
| 1 | 3 | 1.6986 | 1.58 | 2.0764 | 1.8979 | 1.6 | 1.6821 | < 5 | 1.51 | 1.7 | 1.5 | 1.62 | 1.58 | 1.5358 | 1 | 1.6 | 1.7 | 1.4 | 1.7 | 1.75 | 1.2 | 1.77 | 1.6 | [7 empty] | 16
| 1 | 4 | 1.6582 | 1.65 | 2.1362 | 1.9666 | 1.6 | 1.6014 | < 5 | 1.52 | 1.7 | 1.51 | 1.64 | 1.58 | 1.4938 | 1 | 1.6 | 1.7 | 1.35 | 1.7 | 1.66 | 1.25 | 1.24 | 1.7 | [7 empty] | 1.6116
| 1 | 5 | 1.6354 | 1.63 | 1.9894 | 1.9874 | 1.7 | 1.6145 | < 5 | 1.48 | 1.71 | 1.48 | 1.66 | 1.59 | 1.4459 | 1 | 1.7 | 1.7 | 1.38 | 1.8 | 1.7 | 1.22 | 1.47 | 1.6 | [7 empty] | 77
| 1 | 6 | 1.5678 | 1.59 | 1.9859 | 1.8996 | 1.5 | 1.6209 | < 5 | 1.49 | 1.69 | 1.51 | 1.65 | 1.57 | 1.4848 | 1 | 1.7 | 1.8 | 1.29 | 1.7 | 1.76 | 1.2 | 1.45 | 1.8
| Mean |  | 1.6363 | 1.6067 | 2.0661 | 1.9419 | 1.5667 | 1.6144 |  | 1.4933 | 1.6833 | 1.5067 | 1.6283 | 1.585 | 1.488 | 1.1667 | 1.6333 | 1.7167 | 1.39 | 1.7167 | 1.685 | 1.2183 | 1.4283 | 1.6167
| Median |  | 1.6386 | 1.595 | 2.0642 | 1.9495 | 1.55 | 1.6177 |  | 1.485 | 1.7 | 1.51 | 1.63 | 1.58 | 1.4893 | 1 | 1.6 | 1.7 | 1.39 | 1.7 | 1.68 | 1.22 | 1.46 | 1.6
| Std Dev. |  | 0.043596 | 0.027325 | 0.071668 | 0.037764 | 0.08165 | 0.045498 |  | 0.017512 | 0.046332 | 0.015055 | 0.025626 | 0.013784 | 0.037637 | 0.40825 | 0.05164 | 0.040825 | 0.068411 | 0.040825 | 0.066858 | 0.018348 | 0.21405 | 0.13292
| Rel.Std.Dev. |  | 2.664% | 1.701% | 3.469% | 1.945% | 5.212% | 2.818% |  | 1.173% | 2.752% | 0.9993% | 1.574% | 0.8697% | 2.529% | 34.99% | 3.162% | 2.378% | 4.922% | 2.378% | 3.968% | 1.506% | 14.99% | 8.222%
| PDM3 |  | 1.53% | -0.3088% | 28.2% | 20.49% | -2.791% | 0.1731% |  | -7.341% | 4.448% | -6.514% | 1.036% | -1.653% | -7.674% | -27.61% | 1.346% | 6.517% | -13.75% | 6.517% | 4.552% | -24.4% | -11.37% | 0.3117%
| Z-Score (Absolute) |  | 0.13 | 0.07 | 3.03 | 2.19 | 0.34 | 0.02 | 5.96 | 0.83 | 0.45 | 0.74 | 0.08 | 0.21 | 0.87 | NA | 0.11 | 0.67 | 1.53 | 0.67 | 0.46 | 2.69 | 1.27 | 0
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 285 (Certified Value 0.63 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.66478 | 0.65 | 0.64558 | 0.57 | < 2 | 0.58 | 0.58 | 0.58 | 0.63 | 0.6 | 0.57 | 0.55 | 0.68483 | 0.5 | 0.53 | 0.61 | 0.73 | 0.65 | 0.65 | < 2 | 0.6 | 0.66 | < 0.5 | 0.647 | 0.73 |  |  |  |  | 1
| 1 | 2 | 0.73649 | 0.58 | 0.62873 | 0.55 | < 2 | 0.8 | 0.74 | 0.53 | 0.78 | 0.67 | 0.59 | 0.77 | 0.64808 | 0.5 | 0.52 | 0.64 | 0.57 | 0.65 | 0.59 | < 2 | 0.6 | 0.64 | < 0.5 | 0.617 | 0.63 |  |  |  |  | 26
| 1 | 3 | 0.50728 | 0.75 | 0.71501 | 0.64 | < 2 | 0.82 | 0.57 | 0.48 | 0.76 | 0.73 | 0.62 | 0.55 | 0.61087 | 0.5 | 0.59 | 0.63 | 0.59 | 0.7 | 0.57 | < 2 | 0.66 | 0.52 | < 0.5 | 0.667 | 0.9 |  |  |  |  | 16
| 1 | 4 | 0.51213 | 0.67 | 0.70977 | 0.69 | < 2 | 0.5 | 0.74 | 0.54 | 0.52 | 0.62 | 0.95 | 0.7 | 0.67724 | 0.6 | 0.6 | 0.6 | 0.53 | 0.68 | 0.59 | < 2 | 0.69 | 0.61 | < 0.5 | 0.63 | 0.73 |  |  |  |  | 0.63448
| 1 | 5 | 0.60781 | 0.59 | 0.57067 | 0.8 | < 2 | 0.74 | 0.52 | 0.54 | 0.57 | 0.51 | 0.68 | 1.95 | 0.67207 | 0.6 | 0.52 | 0.6 | 0.54 | 0.59 | 0.68 | < 2 | 0.61 | 0.63 | < 0.5 | 0.71 | 0.9 |  |  |  |  | 78
| 1 | 6 | 0.81209 | 0.6 | 0.60214 | 0.73 | < 2 | 1.12 | 0.62 | 0.47 | 0.6 | 0.6 | 0.57 | 0.54 | 0.60341 | 0.7 | 0.52 | 0.6 | 0.6 | 0.64 | 0.64 | < 2 | 0.72 | 0.84 | < 0.5 | 0.62 | 0.59
| Mean |  | 0.6401 | 0.64 | 0.64532 | 0.66333 |  | 0.76 | 0.62833 | 0.52333 | 0.64333 | 0.62167 | 0.66333 | 0.84333 | 0.64942 | 0.56667 | 0.54667 | 0.61333 | 0.59333 | 0.65167 | 0.62 |  | 0.64667 | 0.65 |  | 0.6485 | 0.74667
| Median |  | 0.63629 | 0.625 | 0.63716 | 0.665 |  | 0.77 | 0.6 | 0.535 | 0.615 | 0.61 | 0.605 | 0.625 | 0.66008 | 0.55 | 0.525 | 0.605 | 0.58 | 0.65 | 0.615 |  | 0.635 | 0.635 |  | 0.6385 | 0.73
| Std Dev. |  | 0.1221 | 0.064498 | 0.057832 | 0.095847 |  | 0.21689 | 0.092177 | 0.041312 | 0.10482 | 0.07414 | 0.14638 | 0.55041 | 0.035059 | 0.08165 | 0.037771 | 0.017512 | 0.072296 | 0.037639 | 0.042895 |  | 0.051251 | 0.10507 |  | 0.035422 | 0.13095
| Rel.Std.Dev. |  | 19.08% | 10.08% | 8.962% | 14.45% |  | 28.54% | 14.67% | 7.894% | 16.29% | 11.93% | 22.07% | 65.27% | 5.399% | 14.41% | 6.909% | 2.855% | 12.18% | 5.776% | 6.919% |  | 7.925% | 16.16% |  | 5.462% | 17.54%
| PDM3 |  | 0.8852% | 0.8702% | 1.708% | 4.548% |  | 19.78% | -0.9686% | -17.52% | 1.396% | -2.019% | 4.548% | 32.92% | 2.354% | -10.69% | -13.84% | -3.333% | -6.485% | 2.709% | -2.282% |  | 1.921% | 2.446% |  | 2.21% | 17.68%
| Z-Score (Absolute) |  | 0.19 | 0.19 | 0.02 | 0.56 | 11.37 | 3.66 | 0.57 | 3.94 | 0.09 | 0.78 | 0.56 | 6.34 | 0.11 | NA | 3.19 | 1.05 | 1.69 | 0.18 | 0.83 | 11.37 | 0.02 | 0.13 | 12.72 | 0.08 | 3.23
Results from laboratory 1.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 285 (Certified Value 0.659 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.65542 | 0.62 | 0.71181 | 0.69783 | 0.79 | 0.8199 | 0.66 | 0.5186 | 0.65613 | 0.63 | 0.73 | 0.66 | 0.65 | 0.67 | 0.5746 | 0.67 | 0.59 | 0.7 | 0.6 | 0.65 | 0.61 | 0.61 | 0.62 | 0.693 | 0.722 | < 2 | 0.678 | 0.7 |  | 1
| 1 | 2 | 0.65548 | 0.61 | 0.70297 | 0.69652 | 0.78 | 0.8354 | 0.67 | 0.5307 | 0.70564 | 0.62 | 0.7 | 0.65 | 0.66 | 0.66 | 0.56013 | 0.66 | 0.6 | 0.7 | 0.61 | 0.64 | 0.62 | 0.62 | 0.62 | 0.67 | 0.693 | < 2 | 0.74 | 0.62 |  | #N/A
| 1 | 3 | 0.65515 | 0.61 | 0.70942 | 0.69457 | 0.74 | 0.7803 | 0.68 | 0.5383 | 0.70144 | 0.64 | 0.73 | 0.65 | 0.66 | 0.65 | 0.57288 | 0.66 | 0.64 | 0.7 | 0.6 | 0.65 | 0.61 | 0.64 | 0.66 | 0.691 | 0.693 | < 2 | 0.732 | 0.67 |  | 16
| 1 | 4 | 0.64818 | 0.63 | 0.70387 | 0.69641 | 0.82 | 0.821 | 0.66 | 0.5279 | 0.66551 | 0.64 | 0.73 | 0.65 | 0.67 | 0.66 | 0.55735 | 0.66 | 0.64 | 0.69 | 0.59 | 0.64 | 0.62 | 0.61 | 0.63 | 0.731 | 0.729 | < 2 | 0.71 | 0.68 |  | 0.65917
| 1 | 5 | 0.6405 | 0.64 | 0.71031 | 0.70618 | 0.79 | 0.84 | 0.67 | 0.5324 | 0.66691 | 0.62 | 0.73 | 0.64 | 0.67 | 0.65 | 0.55139 | 0.65 | 0.62 | 0.71 | 0.61 | 0.63 | 0.62 | 0.62 | 0.63 | 0.719 | 0.693 | < 2 | 0.73 | 0.69 |  | 79
| 1 | 6 | 0.6478 | 0.64 | 0.69847 | 0.67933 | 0.77 | 0.8449 | 0.68 | 0.5238 | 0.7232 | 0.63 | 0.73 | 0.65 | 0.67 | 0.65 | 0.54199 | 0.66 | 0.63 | 0.72 | 0.6 | 0.64 | 0.62 | 0.62 | 0.64 |  | 0.708 | < 2 | 0.69 | 0.7
| Mean |  | 0.65042 | 0.625 | 0.70614 | 0.69514 | 0.78167 | 0.82358 | 0.67 | 0.52862 | 0.68647 | 0.63 | 0.725 | 0.65 | 0.66333 | 0.65667 | 0.55972 | 0.66 | 0.62 | 0.70333 | 0.60167 | 0.64167 | 0.61667 | 0.62 | 0.63333 | 0.7008 | 0.70633 |  | 0.71333 | 0.67667
| Median |  | 0.65166 | 0.625 | 0.70665 | 0.69647 | 0.785 | 0.8282 | 0.67 | 0.5293 | 0.68418 | 0.63 | 0.73 | 0.65 | 0.665 | 0.655 | 0.55874 | 0.66 | 0.625 | 0.7 | 0.6 | 0.64 | 0.62 | 0.62 | 0.63 | 0.693 | 0.7005 |  | 0.72 | 0.685
| Std Dev. |  | 0.006055 | 0.013784 | 0.0051847 | 0.0087529 | 0.026394 | 0.023483 | 0.0089443 | 0.006877 | 0.02714 | 0.0089443 | 0.012247 | 0.0063246 | 0.008165 | 0.008165 | 0.012524 | 0.0063246 | 0.020976 | 0.010328 | 0.0075277 | 0.0075277 | 0.005164 | 0.010954 | 0.015055 | 0.024232 | 0.016096 |  | 0.025065 | 0.030111
| Rel.Std.Dev. |  | 0.9309% | 2.205% | 0.7342% | 1.259% | 3.377% | 2.851% | 1.335% | 1.301% | 3.954% | 1.42% | 1.689% | 0.973% | 1.231% | 1.243% | 2.237% | 0.9583% | 3.383% | 1.468% | 1.251% | 1.173% | 0.8374% | 1.767% | 2.377% | 3.458% | 2.279% |  | 3.514% | 4.45%
| PDM3 |  | -1.327% | -5.183% | 7.126% | 5.458% | 18.58% | 24.94% | 1.643% | -19.81% | 4.142% | -4.425% | 9.987% | -1.391% | 0.6321% | -0.3793% | -15.09% | 0.1264% | -5.942% | 6.7% | -8.723% | -2.655% | -6.448% | -5.942% | -3.919% | 6.316% | 7.155% |  | 8.217% | 2.655%
| Z-Score (Absolute) |  | 0.19 | 0.61 | 0.74 | 0.56 | 2 | 2.7 | 0.14 | 2.22 | 0.41 | 0.53 | 1.06 | 0.19 | 0.03 | 0.08 | 1.7 | 0.03 | 0.7 | 0.7 | 1 | 0.33 | 0.75 | 0.7 | 0.47 | 0.65 | 0.75 | 5.65 | 0.86 | 0.25
| Analytical results for Cd in OREAS 285 (Certified Value 0.085 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | < 0.2 | 0.1 | 0.082084 | < 1 | 0.08 | 0.8 | 0.06 | 0.08 | 0.09 | 0.11 | 0.11 | 0.07 | 0.076059 | < 0.1 | 0.09 | 0.07 | 0.09 | 0.08 | 0.08 | < 0.5 | 0.08 | 0.11 | 0.621 | 0.088 | 0.07 |  |  |  |  | 1
| 1 | 2 | < 0.2 | 0.09 | 0.090067 | < 1 | 0.09 | 0.7 | 0.08 | 0.09 | 0.08 | 0.07 | 0.08 | 0.08 | 0.086838 | 0.1 | 0.1 | 0.06 | 0.09 | 0.08 | 0.09 | < 0.5 | 0.09 | 0.1 | 0.621 | 0.076 | 0.07 |  |  |  |  | 27
| 1 | 3 | < 0.2 | 0.08 | 0.094553 | < 1 | 0.1 | 0.8 | 0.08 | 0.07 | 0.08 | 0.09 | 0.09 | 0.06 | 0.074276 | < 0.1 | 0.09 | 0.07 | 0.08 | 0.08 | 0.08 | < 0.5 | 0.09 | 0.11 | 0.621 | 0.078 | 0.07 |  |  |  |  | 16
| 1 | 4 | < 0.2 | 0.1 | 0.076848 | < 1 | 0.07 | 0.7 | 0.06 | 0.08 | 0.1 | 0.09 | 0.07 | 0.08 | 0.084158 | < 0.1 | 0.1 | 0.07 | 0.08 | 0.09 | 0.08 | < 0.5 | 0.09 | 0.11 | 0.594 | 0.08 | 0.09 |  |  |  |  | 0.084838
| 1 | 5 | < 0.2 | 0.09 | 0.092147 | < 1 | 0.12 | 0.7 | 0.08 | 0.07 | 0.09 | 0.09 | 0.09 | 0.07 | 0.067905 | < 0.1 | 0.09 | 0.08 | 0.1 | 0.08 | 0.09 | < 0.5 | 0.08 | 0.11 | 0.594 | 0.08 | 0.07 |  |  |  |  | 80
| 1 | 6 | < 0.2 | 0.09 | 0.086721 | < 1 | 0.08 | 0.7 | 0.02 | 0.07 | 0.1 | 0.09 | 0.08 | 0.09 | 0.074949 | < 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.08 | < 0.5 | 0.09 | 0.1 | 0.576 | 0.07 | 0.06
| Mean |  |  | 0.091667 | 0.08707 |  | 0.09 | 0.73333 | 0.063333 | 0.076667 | 0.09 | 0.09 | 0.086667 | 0.075 | 0.077364 | 0.1 | 0.093333 | 0.071667 | 0.086667 | 0.081667 | 0.083333 |  | 0.086667 | 0.10667 | 0.6045 | 0.078667 | 0.071667
| Median |  |  | 0.09 | 0.088394 |  | 0.085 | 0.7 | 0.07 | 0.075 | 0.09 | 0.09 | 0.085 | 0.075 | 0.075504 | 0.1 | 0.09 | 0.07 | 0.085 | 0.08 | 0.08 |  | 0.09 | 0.11 | 0.6075 | 0.079 | 0.07
| Std Dev. |  |  | 0.0075277 | 0.0066366 |  | 0.017889 | 0.05164 | 0.023381 | 0.008165 | 0.0089443 | 0.012649 | 0.013663 | 0.010488 | 0.0069637 |  | 0.005164 | 0.0075277 | 0.008165 | 0.0040825 | 0.005164 |  | 0.005164 | 0.005164 | 0.019233 | 0.0058878 | 0.0098319
| Rel.Std.Dev. |  |  | 8.212% | 7.622% |  | 19.88% | 7.042% | 36.92% | 10.65% | 9.938% | 14.05% | 15.76% | 13.98% | 9.001% |  | 5.533% | 10.5% | 9.421% | 4.999% | 6.197% |  | 5.958% | 4.841% | 3.182% | 7.485% | 13.72%
| PDM3 |  |  | 8.049% | 2.63% |  | 6.084% | 764.4% | -25.35% | -9.632% | 6.084% | 6.084% | 2.155% | -11.6% | -8.81% | 17.87% | 10.01% | -15.53% | 2.155% | -3.739% | -1.774% |  | 2.155% | 25.73% | 612.5% | -7.275% | -15.53%
| Z-Score (Absolute) |  | 1.04 | 0.39 | 0.03 | 32.22 | 0.26 | NA | 1.82 | 0.78 | 0.26 | 0.26 | 0 | 0.91 | 0.73 | 2.21 | 0.52 | 1.17 | 0 | 0.39 | 0.26 | 12.73 | 0 | 1.56 | 40.36 | 0.62 | 1.17
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 285 (Certified Value 39.3 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | [9 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [9 empty] | 1
| 1 | 1 | 43.0948 | 33.527 | 40.1 | 47.0936 | 32.2 | 41.4 | 38.4 | 34.7 | 33.2 | 41.1108 | 40 | 31.62 | 56.8 | 33.1 | 44.69 | 48.057 | 14.5 | 80.16 | 39.6 | 41.55 | [9 empty] | 1
| 1 | 2 | 43.0074 | 36.169 | 39.6 | 47.3484 | 32.2 | 37.6 | 38.6 | 37.5 | 31.6 | 43.5605 | 39 | 31.87 | 57 | 33.8 | 42.9 | 49.658 | 14.2 | 81.12 | 48.8 | 36.54 | [9 empty] | 28
| 1 | 3 | 42.3602 | 34.373 | 38.9 | 50.9369 | 31.8 | 40.6 | 39.9 | 36 | 30.2 | 44.9843 | 40 | 32.52 | 56 | 31.2 | 41.77 | 52.315 | 14.5 | 82.32 | 46.7 | 41.86 | [9 empty] | 16
| 1 | 4 | 42.0792 | 37.897 | 41.8 | 46.8413 | 32.7 | 40.7 | 38.8 | 36.3 | 32.9 | 40.9439 | 38 | 32.85 | 58.9 | 32.2 | 39.12 | 47.196 | 14.6 | 82.56 | 43.4 | 42.08 | [9 empty] | 39.288
| 1 | 5 | 42.9423 | 38.093 | 39.7 | 49.9184 | 31.5 | 37.4 | 37.7 | 37.7 | 30.5 | 37.6282 | 40 | 31.63 | 57.5 | 31 | 41.52 | 52.104 | 14.8 | 84.96 | 43.8 | 42.61 | [9 empty] | 81
| 1 | 6 | 42.9755 | 36.954 | 38.9 | 47.5829 | 31 | 42.2 | 39.2 | 36.4 | 34 | 39.0215 | 39 | 32.29 | 57.1 | 28.5 | 39.14 | 51.429 | 14.8 | 82.8 | 39.8 | 42.28
| Mean |  | 42.7432 | 36.1688 | 39.8333 | 48.2869 | 31.9 | 39.9833 | 38.7667 | 36.4333 | 32.0667 | 41.2082 | 39.3333 | 32.13 | 57.2167 | 31.6333 | 41.5233 | 50.1265 | 14.5667 | 82.32 | 43.6833 | 41.1533
| Median |  | 42.9589 | 36.5615 | 39.65 | 47.4656 | 32 | 40.65 | 38.7 | 36.35 | 32.25 | 41.0274 | 39.5 | 32.08 | 57.05 | 31.7 | 41.645 | 50.5435 | 14.55 | 82.44 | 43.6 | 41.97
| Std Dev. |  | 0.41822 | 1.8714 | 1.0727 | 1.7073 | 0.6 | 2.0084 | 0.74476 | 1.0912 | 1.541 | 2.7375 | 0.8165 | 0.50474 | 0.96212 | 1.8747 | 2.1652 | 2.1673 | 0.22509 | 1.6348 | 3.6641 | 2.2886
| Rel.Std.Dev. |  | 0.9784% | 5.174% | 2.693% | 3.536% | 1.881% | 5.023% | 1.921% | 2.995% | 4.806% | 6.643% | 2.076% | 1.571% | 1.682% | 5.926% | 5.214% | 4.324% | 1.545% | 1.986% | 8.388% | 5.561%
| PDM3 |  | 8.795% | -7.939% | 1.388% | 22.9% | -18.8% | 1.77% | -1.327% | -7.266% | -18.38% | 4.887% | 0.1153% | -18.22% | 45.63% | -19.48% | 5.69% | 27.59% | -62.92% | 109.5% | 11.19% | 4.748%
| Z-Score (Absolute) |  | 0.51 | 0.67 | 0.01 | 1.5 | 1.44 | 0.01 | 0.2 | 0.62 | 1.41 | 0.23 | 0.1 | 1.4 | 3.11 | 1.49 | 0.29 | 1.83 | 4.55 | 7.61 | 0.68 | 0.22
| Analytical results for Co in OREAS 285 (Certified Value 13.7 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 1
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 2
| 1 | 1 | 17.42 | 14.1697 | 14.1 | 19.4237 | 14.0994 | 14.2 | 13.0024 | 12 | 12.5 | 14 | 16 | 13.6 | 13.5 | 13 | 13.684 | 13 | 12.9 | 13.4 | 13.5 | 14.4 | 13.7 | 13 | 14.1 | 14.83 | 10.8 | 13.25 | 13.6 | 14 |  | 1
| 1 | 2 | 17.5 | 14.0569 | 14.4 | 18.8551 | 14.3357 | 14.7 | 11.4319 | 13 | 13 | 14.4 | 14.7 | 13.5 | 14.2 | 13 | 13.7369 | 12.7 | 13 | 14 | 13.6 | 13.7 | 13.9 | 13 | 13.7 | 14.79 | 10.5 | 13.22 | 14.3 | 12.2 |  | 16
| 1 | 3 | 17.23 | 14.254 | 14.2 | 18.3396 | 14.2867 | 15.1 | 11.6768 | 12 | 13.1 | 14.5 | 15.7 | 13.4 | 13.5 | 12.8 | 13.5065 | 12.5 | 13.6 | 13.4 | 13.1 | 13.7 | 13.9 | 13 | 14.2 | 14.08 | 10.6 | 13.17 | 14.4 | 13.5 |  | 16
| 1 | 4 | 17.59 | 14.006 | 14.3 | 18.8511 | 14.5156 | 15.8 | 12.6972 | 12 | 12.6 | 14.5 | 15.9 | 13.2 | 13.6 | 13.7 | 13.1985 | 12.6 | 13.6 | 13.8 | 13.3 | 14 | 13.6 | 13 | 13.6 | 12.45 | 11.1 | 13.22 | 14.2 | 14 |  | 13.7157
| 1 | 5 | 17.66 | 13.8691 | 14.4 | 18.8814 | 14.3615 | 14.8 | 12.8069 | 12 | 12.9 | 13.8 | 14.2 | 13.2 | 13.7 | 12.7 | 13.4239 | 12.9 | 13.4 | 13.7 | 13 | 14.1 | 15 | 13 | 13.8 | 12.95 | 10.5 | 13.36 | 14.2 | 14.1 |  | 82
| 1 | 6 | 17.33 | 14.1953 | 14.3 | 19.0391 | 14.0395 | 15.1 | 12.2398 | 12 | 12.7 | 13.4 | 15.9 | 13.8 | 14 | 13.3 | 13.0567 | 12.4 | 13.5 | 13.4 | 13.1 | 14.1 | 14.1 | 13 | 13.9 | 13.4 | 10.4 | 13.05 | 13.6 | 13.9
| Mean |  | 17.455 | 14.0919 | 14.2833 | 18.8983 | 14.2731 | 14.95 | 12.3092 | 12.1667 | 12.8 | 14.1 | 15.4 | 13.45 | 13.75 | 13.0833 | 13.4344 | 12.6833 | 13.3333 | 13.6167 | 13.2667 | 14 | 14.0333 | 13 | 13.8833 | 13.75 | 10.65 | 13.2117 | 14.05 | 13.6167
| Median |  | 17.46 | 14.1133 | 14.3 | 18.8682 | 14.3112 | 14.95 | 12.4685 | 12 | 12.8 | 14.2 | 15.8 | 13.45 | 13.65 | 13 | 13.4652 | 12.65 | 13.45 | 13.55 | 13.2 | 14.05 | 13.9 | 13 | 13.85 | 13.74 | 10.55 | 13.22 | 14.2 | 13.95
| Std Dev. |  | 0.16109 | 0.14229 | 0.1169 | 0.34997 | 0.17633 | 0.53198 | 0.64081 | 0.40825 | 0.23664 | 0.44721 | 0.75895 | 0.23452 | 0.2881 | 0.3656 | 0.26736 | 0.23166 | 0.30768 | 0.25626 | 0.24221 | 0.26833 | 0.50465 | 0 | 0.23166 | 0.98075 | 0.25884 | 0.10147 | 0.35637 | 0.72503
| Rel.Std.Dev. |  | 0.9229% | 1.01% | 0.8185% | 1.852% | 1.235% | 3.558% | 5.206% | 3.355% | 1.849% | 3.172% | 4.928% | 1.744% | 2.095% | 2.794% | 1.99% | 1.826% | 2.308% | 1.882% | 1.826% | 1.917% | 3.596% | 0% | 1.669% | 7.133% | 2.43% | 0.7681% | 2.536% | 5.325%
| PDM3 |  | 27.26% | 2.743% | 4.139% | 37.79% | 4.064% | 9% | -10.25% | -11.29% | -6.676% | 2.802% | 12.28% | -1.937% | 0.2504% | -4.61% | -2.051% | -7.527% | -2.788% | -0.7217% | -3.274% | 2.073% | 2.316% | -5.218% | 1.223% | 0.2504% | -22.35% | -3.675% | 2.438% | -0.7217%
| Z-Score (Absolute) |  | 5.55 | 0.51 | 0.8 | 7.71 | 0.78 | 1.8 | 2.16 | NA | 1.42 | 0.52 | 2.47 | 0.45 | 0 | 1 | 0.47 | 1.6 | 0.62 | 0.2 | 0.72 | 0.37 | 0.42 | NA | 0.2 | 0 | 4.65 | 0.81 | 0.45 | 0.2
Results from laboratories 1.11 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 285 (Certified Value 53 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  | 0
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 1
| 1 | 1 | 50.91 | 52.894 | 53 | 57.3582 | 51.6498 | 49 | 49.1678 | 47 | 47.7 | 51 | 55 | 55 | 53 | 53 | 51.8891 | 55 | 53 | 52 | 54.7 | 52.9 | 52 | 50 | 56 | 59.61 | 60 | 43.696 | 23.7 | 55 |  | 1
| 1 | 2 | 50.97 | 52.4201 | 52 | 56.6865 | 53.2398 | 50 | 48.0863 | 48 | 50.9 | 51 | 51 | 54 | 53 | 52 | 50.1679 | 55 | 51 | 54 | 56.4 | 51.2 | 53 | 50 | 55 | 60.67 | 59 | 45.896 | 25.4 | 49 |  | 17
| 1 | 3 | 51.4 | 52.9217 | 51 | 55.3108 | 52.7441 | 52 | 50.0407 | 47 | 50.3 | 52 | 55 | 54 | 53 | 52 | 52.382 | 54 | 55 | 53 | 54.9 | 52.6 | 53 | 52 | 58 | 55.99 | 59 | 43.584 | 24.9 | 52 |  | 16
| 1 | 4 | 51.84 | 53.1943 | 52 | 57.2948 | 53.1307 | 53 | 50.9062 | 47 | 47.7 | 52 | 55 | 54 | 53 | 52 | 49.4906 | 53 | 56 | 53 | 54 | 51.3 | 52 | 50 | 55 | 50.73 | 60 | 43.696 | 24 | 54 |  | 52.5403
| 1 | 5 | 52.09 | 52.5084 | 54 | 55.9498 | 52.003 | 52 | 51.5365 | 47 | 47.9 | 51 | 55 | 54 | 54 | 52 | 48.5392 | 56 | 55 | 52 | 54.9 | 50.8 | 53 | 50 | 56 | 51.64 | 58 | 46.072 | 24 | 53 |  | 83
| 1 | 6 | 50.59 | 53.4579 | 53 | 55.1438 | 52.3416 | 51 | 51.8231 | 48 | 52 | 51 | 55 | 54 | 54 | 51 | 49.8615 | 53 | 56 | 52 | 53.6 | 51.9 | 53 | 50 | 56 | 51.54 | 59 | 43.112 | 23 | 54
| Mean |  | 51.3 | 52.8994 | 52.5 | 56.2906 | 52.5182 | 51.1667 | 50.2601 | 47.3333 | 49.4167 | 51.3333 | 54.3333 | 54.1667 | 53.3333 | 52 | 50.3884 | 54.3333 | 54.3333 | 52.6667 | 54.75 | 51.7833 | 52.6667 | 50.3333 | 56 | 55.03 | 59.1667 | 44.3427 | 24.1667 | 52.8333
| Median |  | 51.185 | 52.9079 | 52.5 | 56.3181 | 52.5429 | 51.5 | 50.4734 | 47 | 49.1 | 51 | 55 | 54 | 53 | 52 | 50.0147 | 54.5 | 55 | 52.5 | 54.8 | 51.6 | 53 | 50 | 56 | 53.815 | 59 | 43.696 | 24 | 53.5
| Std Dev. |  | 0.58152 | 0.39545 | 1.0488 | 0.96914 | 0.63207 | 1.472 | 1.4461 | 0.5164 | 1.8894 | 0.5164 | 1.633 | 0.40825 | 0.5164 | 0.63246 | 1.4682 | 1.2111 | 1.9664 | 0.8165 | 0.96488 | 0.83287 | 0.5164 | 0.8165 | 1.0954 | 4.3788 | 0.75277 | 1.2907 | 0.85946 | 2.137
| Rel.Std.Dev. |  | 1.134% | 0.7476% | 1.998% | 1.722% | 1.204% | 2.877% | 2.877% | 1.091% | 3.823% | 1.006% | 3.006% | 0.7537% | 0.9682% | 1.216% | 2.914% | 2.229% | 3.619% | 1.55% | 1.762% | 1.608% | 0.9805% | 1.622% | 1.956% | 7.957% | 1.272% | 2.911% | 3.556% | 4.045%
| PDM3 |  | -2.361% | 0.6834% | -0.07679% | 7.138% | -0.04221% | -2.615% | -4.34% | -9.911% | -5.945% | -2.297% | 3.413% | 3.095% | 1.509% | -1.028% | -4.096% | 3.413% | 3.413% | 0.2404% | 4.206% | -1.441% | 0.2404% | -4.201% | 6.585% | 4.739% | 12.61% | -15.6% | -54% | 0.5576%
| Z-Score (Absolute) |  | 0.5 | 0.12 | 0.04 | 1.43 | 0.03 | 0.55 | 0.9 | 2.04 | 1.23 | 0.49 | 0.67 | 0.61 | 0.29 | 0.23 | 0.85 | 0.67 | 0.67 | 0.03 | 0.84 | 0.31 | 0.03 | 0.88 | 1.32 | 0.94 | 2.55 | 3.2 | 11.01 | 0.09
| Analytical results for Cs in OREAS 285 (Certified Value 9.17 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [11 empty] | 3
| 1 | 1 | 9.1039 | 9.07 | 8.45 | 9.18 | 10.3 | 9.34 | 9.51 | 8.77 | 9.1812 | 9.5 | 8.65 | 9.79 | 9.31 | 9.14 | 9.33 | 7.26 | 9.95 | 9.42 | [11 empty] | 1
| 1 | 2 | 9.1058 | 9.19 | 8.2 | 9.49 | 9.46 | 9.38 | 10.1 | 8.71 | 8.8971 | 9.4 | 8.98 | 9.83 | 9.49 | 9.12 | 9.22 | 7.32 | 10.3 | 8.3 | [11 empty] | 31
| 1 | 3 | 9.3697 | 9.1 | 8.38 | 9.25 | 10.55 | 9.36 | 9.52 | 8.29 | 9.204 | 9.3 | 9.02 | 9.95 | 9.08 | 9.11 | 9.54 | 7.44 | 10.4 | 9.08 | [11 empty] | 16
| 1 | 4 | 8.9315 | 9.14 | 8.65 | 9.31 | 10.4 | 9.33 | 9.58 | 9.14 | 8.9784 | 9.3 | 8.95 | 9.91 | 9.39 | 8.85 | 9.12 | 7.55 | 10.4 | 9.47 | [11 empty] | 9.1703
| 1 | 5 | 9.3528 | 9.34 | 8.35 | 8.89 | 9.81 | 9.21 | 9.65 | 8.18 | 8.7889 | 9.7 | 8.96 | 9.63 | 8.96 | 9.62 | 9.13 | 7.3 | 10.2 | 9.51 | [11 empty] | 84
| 1 | 6 | 9.3952 | 9.28 | 8.2 | 8.91 | 10.2 | 9.61 | 9.77 | 8.31 | 8.7086 | 9.4 | 8.85 | 9.67 | 8.88 | 9.22 | 9.36 | 7.32 | 9.88 | 9.63
| Mean |  | 9.2098 | 9.1867 | 8.3717 | 9.1717 | 10.12 | 9.3717 | 9.6883 | 8.5667 | 8.9597 | 9.4333 | 8.9017 | 9.7967 | 9.185 | 9.1767 | 9.2833 | 7.365 | 10.1883 | 9.235
| Median |  | 9.2293 | 9.165 | 8.365 | 9.215 | 10.25 | 9.35 | 9.615 | 8.51 | 8.9377 | 9.4 | 8.955 | 9.81 | 9.195 | 9.13 | 9.275 | 7.32 | 10.25 | 9.445
| Std Dev. |  | 0.18968 | 0.10539 | 0.16916 | 0.2343 | 0.40836 | 0.13106 | 0.22338 | 0.36947 | 0.20261 | 0.15055 | 0.13556 | 0.12754 | 0.24712 | 0.25065 | 0.16008 | 0.10877 | 0.22543 | 0.49383
| Rel.Std.Dev. |  | 2.06% | 1.147% | 2.021% | 2.555% | 4.035% | 1.398% | 2.306% | 4.313% | 2.261% | 1.596% | 1.523% | 1.302% | 2.691% | 2.731% | 1.724% | 1.477% | 2.213% | 5.347%
| PDM3 |  | 0.4314% | 0.179% | -8.708% | 0.01539% | 10.36% | 2.196% | 5.65% | -6.582% | -2.296% | 2.869% | -2.929% | 6.831% | 0.1608% | 0.06992% | 1.233% | -19.69% | 11.1% | 0.706%
| Z-Score (Absolute) |  | 0.03 | 0.03 | 2.35 | 0.08 | 2.62 | 0.49 | 1.4 | 1.8 | 0.68 | 0.67 | 0.84 | 1.7 | 0.04 | 0.06 | 0.24 | 5.22 | 2.82 | 0.1
| Analytical results for Cu in OREAS 285 (Certified Value 30.9 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 2
| 1 | 1 | 32.9693 | 32.9 | 31.9482 | 34.7268 | 33.3 | 49.6572 | 32 | 31 | 33.2 | 29.5 | 33.6 | 31.4 | 30.5 | 29 | 27.4958 | 23.3 | 30.8 | 28 | 30 | 30.35 | 30.7 | 30 | 32 | 29.25 | 29 | 30.08 | 35.6 | 33 |  | 1
| 1 | 2 | 31.099 | 32.7 | 31.5137 | 34.6256 | 34.1 | 29.2726 | 32 | 32 | 34.52 | 30.1 | 31.3 | 31.5 | 31.7 | 28.7 | 28.454 | 23.1 | 30.7 | 29 | 29.9 | 29.69 | 31.3 | 30 | 30.7 | 29.06 | 28.1 | 29.69 | 36 | 28.7 |  | 18
| 1 | 3 | 31.6436 | 32.3 | 31.4823 | 35.1239 | 34 | 29.6086 | 33 | 33 | 33.66 | 29.8 | 34.1 | 30.7 | 30 | 27.9 | 27.5038 | 22.4 | 31.7 | 29 | 29.9 | 29.91 | 30.6 | 31 | 32 | 28.13 | 28.5 | 29.07 | 36.2 | 30.7 |  | 16
| 1 | 4 | 31.6337 | 33.6 | 31.5343 | 35.6121 | 35.3 | 31.3743 | 32 | 32 | 31.84 | 31 | 34.6 | 30.5 | 30.5 | 29.9 | 27.1841 | 22.2 | 30 | 29 | 30.1 | 30.01 | 32.4 | 30 | 30.9 | 23.99 | 28.7 | 29.6 | 36.5 | 32 |  | 30.9125
| 1 | 5 | 30.8506 | 33.3 | 31.246 | 34.4535 | 33.2 | 30.1595 | 33 | 32 | 32.68 | 29 | 30.5 | 30.1 | 29.4 | 27.3 | 28.3351 | 22.8 | 31.5 | 29 | 29.8 | 30.59 | 31 | 30 | 31.7 | 24.22 | 28.4 | 29.99 | 36.4 | 31.3 |  | 85
| 1 | 6 | 30.9739 | 33.2 | 31.0908 | 34.076 | 33.1 | 30.3766 | 32 | 32 | 35.36 | 29.3 | 33 | 31.9 | 31.2 | 29.3 | 28.4119 | 22.5 | 30.7 | 29 | 30.2 | 30.48 | 31.8 | 29 | 31.5 | 26.12 | 28.2 | 29.85 | 35.4 | 32.2
| Mean |  | 31.5283 | 33 | 31.4692 | 34.7696 | 33.8333 | 33.4081 | 32.3333 | 32 | 33.5433 | 29.7833 | 32.85 | 31.0167 | 30.55 | 28.6833 | 27.8974 | 22.7167 | 30.9 | 28.8333 | 29.9833 | 30.1717 | 31.3 | 30 | 31.4667 | 26.795 | 28.4833 | 29.7133 | 36.0167 | 31.3167
| Median |  | 31.3663 | 33.05 | 31.498 | 34.6762 | 33.65 | 30.2681 | 32 | 32 | 33.43 | 29.65 | 33.3 | 31.05 | 30.5 | 28.85 | 27.9194 | 22.65 | 30.75 | 29 | 29.95 | 30.18 | 31.15 | 30 | 31.6 | 27.125 | 28.45 | 29.77 | 36.1 | 31.65
| Std Dev. |  | 0.78126 | 0.46476 | 0.29286 | 0.53646 | 0.83347 | 7.9931 | 0.5164 | 0.63246 | 1.2675 | 0.70828 | 1.6208 | 0.68823 | 0.82158 | 0.94745 | 0.56409 | 0.42622 | 0.61644 | 0.40825 | 0.1472 | 0.35454 | 0.69282 | 0.63246 | 0.55377 | 2.3619 | 0.33116 | 0.36247 | 0.44008 | 1.5039
| Rel.Std.Dev. |  | 2.478% | 1.408% | 0.9306% | 1.543% | 2.463% | 23.93% | 1.597% | 1.976% | 3.779% | 2.378% | 4.934% | 2.219% | 2.689% | 3.303% | 2.022% | 1.876% | 1.995% | 1.416% | 0.4909% | 1.175% | 2.213% | 2.108% | 1.76% | 8.815% | 1.163% | 1.22% | 1.222% | 4.802%
| PDM3 |  | 1.992% | 6.753% | 1.801% | 12.48% | 9.449% | 8.073% | 4.596% | 3.518% | 8.511% | -3.653% | 6.268% | 0.337% | -1.173% | -7.211% | -9.753% | -26.51% | -0.04039% | -6.726% | -3.006% | -2.397% | 1.254% | -2.952% | 1.793% | -13.32% | -7.858% | -3.879% | 16.51% | 1.308%
| Z-Score (Absolute) |  | 0.18 | 0.88 | 0.15 | 1.73 | 1.28 | 1.08 | 0.56 | 0.4 | 1.14 | 0.66 | 0.81 | 0.07 | 0.29 | 1.18 | 1.56 | 4.04 | 0.12 | 1.11 | 0.56 | 0.47 | 0.07 | 0.55 | 0.15 | 2.09 | 1.28 | 0.69 | 2.32 | 0.08
| Analytical results for Dy in OREAS 285 (Certified Value 2.8 ppm) | [30 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 3
| 1 | 1 | 2.8749 | 2.587 | 3 | 3.102 | 2.41 | [24 empty] | 1
| 1 | 2 | 2.8627 | 2.674 | 3 | 3.1484 | 2.33 | [24 empty] | 33
| 1 | 3 | 2.8629 | 2.579 | 2.9 | 3.1248 | 2.41 | [24 empty] | 16
| 1 | 4 | 2.8888 | 2.714 | 3 | 2.9997 | 2.44 | [24 empty] | 2.801
| 1 | 5 | 2.9353 | 2.731 | 2.9 | 3.137 | 2.39 | [24 empty] | 86
| 1 | 6 | 2.9422 | 2.661 | 2.9 | 3.0448 | 2.4
| Mean |  | 2.8945 | 2.6577 | 2.95 | 3.0928 | 2.3967
| Median |  | 2.8818 | 2.6675 | 2.95 | 3.1134 | 2.405
| Std Dev. |  | 0.035704 | 0.063263 | 0.054772 | 0.058524 | 0.036697
| Rel.Std.Dev. |  | 1.234% | 2.38% | 1.857% | 1.892% | 1.531%
| PDM3 |  | 3.338% | -5.117% | 5.32% | 10.42% | -14.43%
| Z-Score (Absolute) |  | 0 | 0.81 | 0.19 | 0.67 | 1.69
| Analytical results for Er in OREAS 285 (Certified Value 1 ppm) | [30 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 2
| 1 | 1 | 0.99929 | 0.937 | 1 | 1.2019 | 0.86 | [24 empty] | 1
| 1 | 2 | 1.0328 | 0.982 | 1 | 1.1755 | 0.84 | [24 empty] | 34
| 1 | 3 | 1.0385 | 0.934 | 1 | 1.1688 | 0.86 | [24 empty] | 16
| 1 | 4 | 1.0424 | 0.968 | 1 | 1.1258 | 0.88 | [24 empty] | 1.0014
| 1 | 5 | 1.0488 | 0.99 | 1 | 1.1524 | 0.85 | [24 empty] | 87
| 1 | 6 | 1.0404 | 0.987 | 0.9 | 1.1251 | 0.86
| Mean |  | 1.0337 | 0.96633 | 0.98333 | 1.1583 | 0.85833
| Median |  | 1.0394 | 0.975 | 1 | 1.1606 | 0.86
| Std Dev. |  | 0.017631 | 0.025065 | 0.040825 | 0.030001 | 0.013292
| Rel.Std.Dev. |  | 1.706% | 2.594% | 4.152% | 2.59% | 1.549%
| PDM3 |  | 3.227% | -3.498% | -1.8% | 15.67% | -14.28%
| Z-Score (Absolute) |  | 0.67 | 0.23 | 0 | 2.34 | 1.67
| Analytical results for Eu in OREAS 285 (Certified Value 0.38 ppm) | [30 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 2
| 1 | 1 | 0.3885 | 0.299 | 0.4 | 0.4995 | 0.31 | [24 empty] | 1
| 1 | 2 | 0.37667 | 0.321 | 0.4 | 0.5483 | 0.3 | [24 empty] | 35
| 1 | 3 | 0.39872 | 0.309 | 0.4 | 0.5009 | 0.3 | [24 empty] | 16
| 1 | 4 | 0.38185 | 0.334 | 0.4 | 0.4809 | 0.31 | [24 empty] | 0.38276
| 1 | 5 | 0.40404 | 0.325 | 0.4 | 0.4934 | 0.3 | [24 empty] | 88
| 1 | 6 | 0.39764 | 0.329 | 0.4 | 0.5155 | 0.31
| Mean |  | 0.39124 | 0.3195 | 0.4 | 0.50642 | 0.305
| Median |  | 0.39307 | 0.323 | 0.4 | 0.5002 | 0.305
| Std Dev. |  | 0.010666 | 0.013142 | 6.0809e-17 | 0.02339 | 0.0054772
| Rel.Std.Dev. |  | 2.726% | 4.113% | 1.52e-14% | 4.619% | 1.796%
| PDM3 |  | 2.216% | -16.53% | 4.505% | 32.31% | -20.31%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.08 | 1.08 | 0.81
| Analytical results for Fe in OREAS 285 (Certified Value 3.19 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 3.2211 | 3.21 | 3.4183 | 3.3182 | 3.42 | 3.0726 | 3.27 | 1.6833 | 2.9 | 3.09 | 3.46 | 3.22 | 3.2 | 3.16 | 3.1801 | 2.97 | 2.97 | 3.19 | 2.9 | 3.1 | 3.22 | 3.24 | 3.26 | 3.343 | 3.203 | 3.566 | 3.06 | 3.4 |  | 1
| 1 | 2 | 3.2124 | 3.16 | 3.4205 | 3.4382 | 3.51 | 3.0211 | 3.29 | 1.7774 | 3.11 | 3.1 | 3.32 | 3.19 | 3.2 | 3.14 | 3.2493 | 2.93 | 2.93 | 3.19 | 3 | 3.05 | 3.27 | 3.3 | 3.18 | 3.203 | 3.099 | 3.548 | 3.4 | 3.02 |  | 11
| 1 | 3 | 3.2501 | 3.15 | 3.3992 | 3.4185 | 3.6 | 3.0463 | 3.29 | 1.8486 | 3.07 | 3.15 | 3.47 | 3.15 | 3.23 | 3.16 | 3.305 | 2.88 | 3.03 | 3.22 | 2.95 | 3.06 | 3.23 | 3.41 | 3.34 | 3.182 | 3.127 | 3.551 | 3.34 | 3.23 |  | 16
| 1 | 4 | 3.2075 | 3.2 | 3.3337 | 3.3206 | 3.69 | 3.11 | 3.25 | 1.8541 | 2.89 | 3.15 | 3.51 | 3.18 | 3.24 | 3.15 | 3.2762 | 2.93 | 3.02 | 3.19 | 2.87 | 3 | 3.25 | 3.25 | 3.2 | 3.481 | 3.231 | 3.59 | 3.23 | 3.28 |  | 3.1919
| 1 | 5 | 3.2182 | 3.28 | 3.377 | 3.3144 | 3.48 | 3.1389 | 3.25 | 1.8289 | 2.92 | 3.1 | 3.48 | 3.15 | 3.26 | 3.15 | 3.145 | 2.91 | 3 | 3.16 | 3.07 | 2.94 | 3.25 | 3.3 | 3.22 | 3.429 | 3.064 | 3.641 | 3.23 | 3.33 |  | 89
| 1 | 6 | 3.2219 | 3.28 | 3.4151 | 3.3992 | 3.59 | 3.1247 | 3.31 | 1.7727 | 3.16 | 3.12 | 3.44 | 3.18 | 3.26 | 3.12 | 3.328 | 2.95 | 3.02 | 3.23 | 3.03 | 3.04 | 3.25 | 3.27 | 3.21 |  | 3.141 | 3.488 | 3.03 | 3.38
| Mean |  | 3.2219 | 3.2133 | 3.394 | 3.3682 | 3.5483 | 3.0856 | 3.2767 | 1.7942 | 3.0083 | 3.1183 | 3.4467 | 3.1783 | 3.2317 | 3.1467 | 3.2473 | 2.9283 | 2.995 | 3.1967 | 2.97 | 3.0317 | 3.245 | 3.295 | 3.235 | 3.3276 | 3.1442 | 3.564 | 3.215 | 3.2733
| Median |  | 3.2196 | 3.205 | 3.4072 | 3.3599 | 3.55 | 3.0913 | 3.28 | 1.8032 | 2.995 | 3.11 | 3.465 | 3.18 | 3.235 | 3.15 | 3.2628 | 2.93 | 3.01 | 3.19 | 2.975 | 3.045 | 3.25 | 3.285 | 3.215 | 3.343 | 3.134 | 3.5585 | 3.23 | 3.305
| Std Dev. |  | 0.014881 | 0.056451 | 0.033728 | 0.056663 | 0.09704 | 0.046559 | 0.024221 | 0.06449 | 0.1189 | 0.026394 | 0.066232 | 0.026394 | 0.027142 | 0.015055 | 0.071656 | 0.031252 | 0.038341 | 0.025033 | 0.077201 | 0.055287 | 0.017607 | 0.061563 | 0.057879 | 0.13302 | 0.062885 | 0.050616 | 0.14734 | 0.13909
| Rel.Std.Dev. |  | 0.4619% | 1.757% | 0.9938% | 1.682% | 2.735% | 1.509% | 0.7392% | 3.594% | 3.952% | 0.8464% | 1.922% | 0.8304% | 0.8399% | 0.4785% | 2.207% | 1.067% | 1.28% | 0.7831% | 2.599% | 1.824% | 0.5426% | 1.868% | 1.789% | 3.997% | 2% | 1.42% | 4.583% | 4.249%
| PDM3 |  | 0.9383% | 0.6711% | 6.33% | 5.522% | 11.17% | -3.331% | 2.655% | -43.79% | -5.751% | -2.305% | 7.981% | -0.4254% | 1.245% | -1.418% | 1.734% | -8.258% | -6.169% | 0.1489% | -6.952% | -5.02% | 1.663% | 3.23% | 1.35% | 4.251% | -1.496% | 11.66% | 0.7233% | 2.551%
| Z-Score (Absolute) |  | 0.03 | 0.04 | 1.34 | 1.14 | 2.52 | 1.01 | 0.44 | 10.87 | 1.6 | 0.76 | 1.74 | 0.31 | 0.1 | 0.55 | 0.22 | 2.21 | 1.71 | 0.17 | 1.9 | 1.43 | 0.2 | 0.58 | 0.13 | 0.83 | 0.57 | 2.64 | 0.03 | 0.42
| Analytical results for Ga in OREAS 285 (Certified Value 9.78 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 10.0511 | 9.66 | 10.9 | 8.5965 | 8 | 8.6 | 9.56 | 10.5 | 9.93 | 11.2 | 9.09 | 8.6814 | 11 | 10.1 | 9.6 | 9.97 | 8.9 | 9.7 | 10.5 | 6.08 | 46.2 | 9.3 | [7 empty] | 1
| 1 | 2 | 10.0802 | 9.91 | 11.2 | 9.3397 | 8 | 9.2 | 9.85 | 9.53 | 9.79 | 11.3 | 9.22 | 8.9499 | 11 | 10 | 9.6 | 10.1 | 8.9 | 9.8 | 10.48 | 5.94 | 49.8 | 8.5 | [7 empty] | 36
| 1 | 3 | 10.1503 | 9.65 | 11.3 | 9.3501 | 8 | 9.1 | 9.52 | 10.45 | 9.86 | 11.25 | 8.89 | 9.1701 | 11 | 10.5 | 9.9 | 9.65 | 8.9 | 9.7 | 10.81 | 6.11 | 48.9 | 9 | [7 empty] | 16
| 1 | 4 | 9.842 | 9.92 | 11.7 | 8.5963 | 8 | 8.7 | 9.73 | 10.55 | 9.87 | 11.15 | 9.66 | 9.1112 | 10 | 10.2 | 9.9 | 9.75 | 9 | 9.4 | 10.23 | 6.23 | 45.8 | 9.2 | [7 empty] | 9.778
| 1 | 5 | 10.3552 | 10.24 | 10.9 | 8.5516 | 8 | 9 | 9.41 | 9.33 | 9.68 | 11.35 | 8.96 | 8.8893 | 11 | 10.2 | 9.6 | 9.67 | 8.9 | 10.3 | 10.55 | 5.98 | 48.1 | 9.7 | [7 empty] | 90
| 1 | 6 | 9.8306 | 10.1 | 10.9 | 9.0414 | 8 | 8.8 | 9.45 | 10.6 | 10.15 | 11.4 | 9.37 | 8.8315 | 11 | 10.4 | 9.8 | 9.62 | 9 | 9.7 | 10.54 | 5.99 | 44.6 | 9.3
| Mean |  | 10.0516 | 9.9133 | 11.15 | 8.9126 | 8 | 8.9 | 9.5867 | 10.16 | 9.88 | 11.275 | 9.1983 | 8.9389 | 10.8333 | 10.2333 | 9.7333 | 9.7933 | 8.9333 | 9.7667 | 10.5183 | 6.055 | 47.2333 | 9.1667
| Median |  | 10.0656 | 9.915 | 11.05 | 8.819 | 8 | 8.9 | 9.54 | 10.475 | 9.865 | 11.275 | 9.155 | 8.9196 | 11 | 10.2 | 9.7 | 9.71 | 8.9 | 9.7 | 10.52 | 6.035 | 47.15 | 9.25
| Std Dev. |  | 0.19774 | 0.23458 | 0.32094 | 0.37966 | 0 | 0.23664 | 0.17025 | 0.57117 | 0.15748 | 0.093541 | 0.2852 | 0.18084 | 0.40825 | 0.18619 | 0.15055 | 0.19643 | 0.05164 | 0.29439 | 0.18519 | 0.1071 | 2.0086 | 0.39833
| Rel.Std.Dev. |  | 1.967% | 2.366% | 2.878% | 4.26% | 0% | 2.659% | 1.776% | 5.622% | 1.594% | 0.8296% | 3.101% | 2.023% | 3.768% | 1.819% | 1.547% | 2.006% | 0.5781% | 3.014% | 1.761% | 1.769% | 4.253% | 4.345%
| PDM3 |  | 2.797% | 1.384% | 14.03% | -8.851% | -18.18% | -8.98% | -1.957% | 3.906% | 1.043% | 15.31% | -5.929% | -8.582% | 10.79% | 4.656% | -0.4572% | 0.1564% | -8.639% | -0.1163% | 7.571% | -38.08% | 383.1% | -6.253%
| Z-Score (Absolute) |  | 0.31 | 0.15 | 1.55 | 0.98 | NA | 0.99 | 0.22 | 0.43 | 0.11 | 1.69 | 0.66 | 0.95 | NA | 0.51 | 0.05 | 0.02 | 0.96 | 0.02 | 0.83 | 4.2 | 42.27 | 0.69
Results from laboratories 1.11 and 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 285 (Certified Value 3.99 ppm) | [30 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 2
| 1 | 1 | 4.3137 | 3.518 | 4.2 | 4.8953 | 2.68 | [24 empty] | 1
| 1 | 2 | 4.344 | 3.675 | 4.2 | 5.0962 | 2.58 | [24 empty] | 37
| 1 | 3 | 4.582 | 3.602 | 4.2 | 5.2726 | 2.68 | [24 empty] | 16
| 1 | 4 | 4.2332 | 3.793 | 4.4 | 4.8987 | 2.78 | [24 empty] | 3.9858
| 1 | 5 | 4.3778 | 3.761 | 4.2 | 5.0168 | 2.65 | [24 empty] | 91
| 1 | 6 | 4.3443 | 3.784 | 4.1 | 4.9562 | 2.7
| Mean |  | 4.3658 | 3.6888 | 4.2167 | 5.0226 | 2.6783
| Median |  | 4.3442 | 3.718 | 4.2 | 4.9865 | 2.68
| Std Dev. |  | 0.11672 | 0.11138 | 0.098319 | 0.14416 | 0.065243
| Rel.Std.Dev. |  | 2.674% | 3.019% | 2.332% | 2.87% | 2.436%
| PDM3 |  | 9.535% | -7.451% | 5.792% | 26.01% | -32.8%
| Z-Score (Absolute) |  | 0.19 | 0.67 | 0 | 1.03 | 1.97
| Analytical results for Ge in OREAS 285 (Certified Value 0.12 ppm) | [30 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.03 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.29 | 1.31 | 1.32 | [15 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [15 empty] | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | [15 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.09 | 0.12 | 0.13 | 0.11 | 0.14 | 1.4 | 0.2 | 0.2 | < 0.1 | 0.04 | 0.1 | < 0.1 | [15 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.11 | 0.11 | 0.14 | 0.1 | 0.16 | 1.4 | 0.1 | 0.18 | 0.1 | 0.03 | 0.1 | < 0.1 | [15 empty] | 2
| 1 | 3 | < 0.2 | < 0.1 | 0.11 | 0.11 | 0.14 | 0.09 | 0.14 | 1.4 | 0.1 | 0.17 | 0.1 | 0.05 | 0.12 | < 0.1 | [15 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.1 | 0.11 | 0.13 | 0.1 | 0.17 | 1.3 | 0.1 | 0.19 | 0.1 | 0.05 | 0.1 | < 0.1 | [15 empty] | 0.11833
| 1 | 5 | < 0.2 | < 0.1 | 0.1 | 0.1 | 0.14 | 0.11 | 0.15 | 1.4 | 0.2 | 0.17 | 0.1 | 0.04 | 0.11 | < 0.1 | [15 empty] | 92
| 1 | 6 | < 0.2 | < 0.1 | 0.1 | 0.12 | 0.13 | 0.1 | 0.17 | 1.4 | 0.1 | 0.17 | 0.1 | 0.04 | 0.1 | < 0.1
| Mean |  |  |  | 0.10167 | 0.11167 | 0.135 | 0.10167 | 0.155 | 1.3833 | 0.13333 | 0.18 | 0.1 | 0.041667 | 0.105
| Median |  |  |  | 0.1 | 0.11 | 0.135 | 0.1 | 0.155 | 1.4 | 0.1 | 0.175 | 0.1 | 0.04 | 0.1
| Std Dev. |  |  |  | 0.0075277 | 0.0075277 | 0.0054772 | 0.0075277 | 0.013784 | 0.040825 | 0.05164 | 0.012649 | 0 | 0.0075277 | 0.0083666
| Rel.Std.Dev. |  |  |  | 7.404% | 6.741% | 4.057% | 7.404% | 8.893% | 2.951% | 38.73% | 7.027% | 0% | 18.07% | 7.968%
| PDM3 |  |  |  | -14.08% | -5.634% | 14.08% | -14.08% | 30.99% | 1069% | 12.68% | 52.11% | -15.49% | -64.79% | -11.27%
| Z-Score (Absolute) |  | 0.03 | 1.05 | 0 | 0.2 | 0.67 | 0 | 1.08 | NA | NA | 1.58 | NA | 1.21 | 0.07 | 1.05
Results from laboratories 1.20, 1.21 and 1.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 285 (Certified Value 0.33 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [11 empty] | 2
| 1 | 1 | 0.35434 | 0.27 | 0.3 | 0.29 | 0.32 | 0.33 | 0.31 | 0.27 | 0.31855 | 0.3 | 0.29 | 0.37 | 0.31 | 0.37 | 0.35 | 0.4 | 0.32 | 0.38 | [11 empty] | 1
| 1 | 2 | 0.3437 | 0.31 | 0.3 | 0.28 | 0.33 | 0.33 | 0.34 | 0.26 | 0.33041 | 0.3 | 0.29 | 0.37 | 0.31 | 0.4 | 0.36 | 0.4 | 0.42 | 0.34 | [11 empty] | 22
| 1 | 3 | 0.34947 | 0.29 | 0.3 | 0.28 | 0.29 | 0.32 | 0.33 | 0.26 | 0.31588 | 0.3 | 0.3 | 0.39 | 0.3 | 0.37 | 0.36 | 0.42 | 0.4 | 0.39 | [11 empty] | 16
| 1 | 4 | 0.34042 | 0.3 | 0.3 | 0.29 | 0.31 | 0.33 | 0.31 | 0.29 | 0.32458 | 0.3 | 0.29 | 0.37 | 0.3 | 0.35 | 0.35 | 0.43 | 0.35 | 0.35 | [11 empty] | 0.33142
| 1 | 5 | 0.3631 | 0.29 | 0.3 | 0.29 | 0.28 | 0.32 | 0.33 | 0.27 | 0.28998 | 0.3 | 0.28 | 0.37 | 0.28 | 0.38 | 0.38 | 0.41 | 0.37 | 0.35 | [11 empty] | 93
| 1 | 6 | 0.35744 | 0.29 | 0.3 | 0.29 | 0.32 | 0.34 | 0.32 | 0.27 | 0.28879 | 0.3 | 0.3 | 0.39 | 0.28 | 0.34 | 0.36 | 0.41 | 0.34 | 0.35
| Mean |  | 0.35141 | 0.29167 | 0.3 | 0.28667 | 0.30833 | 0.32833 | 0.32333 | 0.27 | 0.31137 | 0.3 | 0.29167 | 0.37667 | 0.29667 | 0.36833 | 0.36 | 0.41167 | 0.36667 | 0.36
| Median |  | 0.3519 | 0.29 | 0.3 | 0.29 | 0.315 | 0.33 | 0.325 | 0.27 | 0.31722 | 0.3 | 0.29 | 0.37 | 0.3 | 0.37 | 0.36 | 0.41 | 0.36 | 0.35
| Std Dev. |  | 0.0085525 | 0.013292 | 0 | 0.005164 | 0.019408 | 0.0075277 | 0.012111 | 0.010954 | 0.017756 | 0 | 0.0075277 | 0.010328 | 0.013663 | 0.02137 | 0.010954 | 0.01169 | 0.037771 | 0.02
| Rel.Std.Dev. |  | 2.434% | 4.557% | 0% | 1.801% | 6.294% | 2.293% | 3.746% | 4.057% | 5.703% | 0% | 2.581% | 2.742% | 4.605% | 5.802% | 3.043% | 2.84% | 10.3% | 5.556%
| PDM3 |  | 6.031% | -12% | -9.481% | -13.5% | -6.967% | -0.9324% | -2.441% | -18.53% | -6.052% | -9.481% | -12% | 13.65% | -10.49% | 11.14% | 8.622% | 24.21% | 10.63% | 8.622%
| Z-Score (Absolute) |  | 0.5 | 0.67 | NA | 0.77 | 0.35 | 0.05 | 0.05 | 1.1 | 0.29 | NA | 0.67 | 1 | 0.58 | 0.84 | 0.67 | 1.69 | 0.81 | 0.67
Results from laboratories 1.09 and 1.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 285 (Indicative Value 0.014 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.29 | 1.30 | 1.31 | 1.32 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [8 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 15g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | [8 empty] | 3
| 1 | 1 | < 1 | < 0.1 | < 0.01 | < 5 | 0.08 | 0.027 | 0.01 | 0.01 | 0.01 | 0.04 | 0.02 | < 0.05 | 0.795 | 0.02 | 0.14 | < 0.005 | < 0.01 | 0.22 | < 1 | < 0.01 | < 0.01 | [8 empty] | 1
| 1 | 2 | < 1 | < 0.1 | < 0.01 | < 5 | 0.08 | 0.027 | 0.01 | 0.02 | 0.01 | 0.02 | 0.01 | < 0.05 | 0.734 | < 0.01 | 0.14 | < 0.005 | < 0.01 | 0.24 | < 1 | < 0.01 | 0.02 | [8 empty] | 5
| 1 | 3 | < 1 | < 0.1 | < 0.01 | < 5 | 0.08 | 0.037 | 0.01 | 0.01 | 0.01 | 0.04 | 0.01 | < 0.05 | 0.727 | < 0.01 | 0.13 | < 0.005 | < 0.01 | 0.25 | < 1 | < 0.01 | < 0.01 | [8 empty] | 16
| 1 | 4 | < 1 | < 0.1 | 0.01 | < 5 | 0.08 | 0.014 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | < 0.05 | 0.722 | 0.01 | 0.12 | < 0.005 | < 0.01 | 0.23 | < 1 | < 0.01 | < 0.01 | [8 empty] | 0.014367
| 1 | 5 | < 1 | < 0.1 | < 0.01 | < 5 | 0.08 | 0.023 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | < 0.05 | 0.747 | < 0.01 | 0.12 | 0.005 | < 0.01 | 0.25 | < 1 | < 0.01 | < 0.01 | [8 empty] | 11
| 1 | 6 | < 1 | < 0.1 | < 0.01 | < 5 | 0.08 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.03 | 0.01 | < 0.05 | 0.755 | 0.01 | 0.12 | < 0.005 | < 0.01 | 0.23 | < 1 | < 0.01 | < 0.01
| Mean |  |  |  | 0.01 |  | 0.08 | 0.023 | 0.01 | 0.014 | 0.01 | 0.026667 | 0.011667 |  | 0.74667 | 0.013333 | 0.12833 | 0.005 |  | 0.23667 |  |  | 0.02
| Median |  |  |  | 0.01 |  | 0.08 | 0.025 | 0.01 | 0.01 | 0.01 | 0.025 | 0.01 |  | 0.7405 | 0.01 | 0.125 | 0.005 |  | 0.235 |  |  | 0.02
| Std Dev. | [5 empty] | 0 | 0.0097775 | 0 | 0.0054772 | 0 | 0.012111 | 0.0040825 |  | 0.026688 | 0.0057735 | 0.0098319 |  |  | 0.012111
| Rel.Std.Dev. | [5 empty] | 0% | 42.51% | 0% | 39.12% | 0% | 45.41% | 34.99% |  | 3.574% | 43.3% | 7.661% |  |  | 5.117%
| PDM3 |  |  |  | -30.39% |  | 456.8% | 60.09% | -30.39% | -2.552% | -30.39% | 85.61% | -18.79% |  | 5097% | -7.193% | 793.3% | -65.2% |  | 1547% |  |  | 39.21%
| Z-Score (Absolute) |  | 17.87 | 1.01 | 0.64 | 92.79 | 2.14 | 0 | 0.49 | 0.39 | 0.49 | 0.14 | 0.42 | 0.07 | 27.11 | 0.52 | 3.95 | 0.75 | 0.67 | 8 | 17.87 | 0.67 | 0.58
| Analytical results for Ho in OREAS 285 (Certified Value 0.41 ppm) | [30 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 3
| 1 | 1 | 0.44183 | 0.4 | 0.4 | 0.4291 | 0.35 | [24 empty] | 1
| 1 | 2 | 0.43364 | 0.414 | 0.4 | 0.4852 | 0.34 | [24 empty] | 23
| 1 | 3 | 0.44967 | 0.396 | 0.4 | 0.4313 | 0.35 | [24 empty] | 16
| 1 | 4 | 0.42712 | 0.411 | 0.5 | 0.4152 | 0.36 | [24 empty] | 0.40796
| 1 | 5 | 0.44324 | 0.415 | 0.4 | 0.4288 | 0.35 | [24 empty] | 94
| 1 | 6 | 0.44693 | 0.412 | 0.4 | 0.4108 | 0.36
| Mean |  | 0.4404 | 0.408 | 0.41667 | 0.4334 | 0.35167
| Median |  | 0.44253 | 0.4115 | 0.4 | 0.42895 | 0.35
| Std Dev. |  | 0.0084943 | 0.007975 | 0.040825 | 0.026718 | 0.0075277
| Rel.Std.Dev. |  | 1.929% | 1.955% | 9.798% | 6.165% | 2.141%
| PDM3 |  | 7.954% | 0.0109% | 2.135% | 6.237% | -13.8%
| Z-Score (Absolute) |  | 0.96 | 0.35 | 0 | 0.67 | 2.62
| Analytical results for In in OREAS 285 (Certified Value 0.065 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.04 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.31 | 1.32 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | [12 empty] | 3
| 1 | 1 | 0.07 | 0.07 | < 5 | 0.063 | 0.076 | 0.07 | 0.066 | 0.066 | 0.058098 | 0.04 | 0.06 | 0.07 | 0.063 | 0.07 | 0.07 | 0.08 | 0.07 | [12 empty] | 1
| 1 | 2 | 0.07 | 0.07 | < 5 | 0.064 | 0.069 | 0.07 | 0.063 | 0.057 | 0.06084 | 0.04 | 0.05 | 0.07 | 0.065 | 0.06 | 0.07 | 0.092 | 0.06 | [12 empty] | 24
| 1 | 3 | 0.06 | 0.07 | < 5 | 0.064 | 0.07 | 0.068 | 0.062 | 0.058 | 0.066229 | 0.04 | 0.06 | 0.06 | 0.066 | 0.07 | 0.07 | 0.083 | 0.06 | [12 empty] | 16
| 1 | 4 | 0.08 | 0.07 | < 5 | 0.064 | 0.072 | 0.066 | 0.058 | 0.068 | 0.067018 | 0.04 | 0.06 | 0.07 | 0.064 | 0.06 | 0.06 | 0.083 | 0.06 | [12 empty] | 0.06493
| 1 | 5 | 0.06 | 0.07 | < 5 | 0.061 | 0.064 | 0.066 | 0.064 | 0.058 | 0.067611 | 0.04 | 0.06 | 0.06 | 0.061 | 0.07 | 0.06 | 0.079 | 0.06 | [12 empty] | 95
| 1 | 6 | 0.07 | 0.07 | < 5 | 0.059 | 0.07 | 0.079 | 0.061 | 0.062 | 0.061334 | 0.04 | 0.06 | 0.06 | 0.067 | 0.07 | 0.06 | 0.072 | 0.07
| Mean |  | 0.068333 | 0.07 |  | 0.0625 | 0.070167 | 0.069833 | 0.062333 | 0.0615 | 0.063522 | 0.04 | 0.058333 | 0.065 | 0.064333 | 0.066667 | 0.065 | 0.0815 | 0.063333
| Median |  | 0.07 | 0.07 |  | 0.0635 | 0.07 | 0.069 | 0.0625 | 0.06 | 0.063781 | 0.04 | 0.06 | 0.065 | 0.0645 | 0.07 | 0.065 | 0.0815 | 0.06
| Std Dev. |  | 0.0075277 | 0 |  | 0.0020736 | 0.00392 | 0.0048339 | 0.0027325 | 0.0046368 | 0.0039415 | 0 | 0.0040825 | 0.0054772 | 0.0021602 | 0.005164 | 0.0054772 | 0.0065345 | 0.005164
| Rel.Std.Dev. |  | 11.02% | 0% |  | 3.318% | 5.587% | 6.922% | 4.384% | 7.54% | 6.205% | 0% | 6.999% | 8.427% | 3.358% | 7.746% | 8.427% | 8.018% | 8.154%
| PDM3 |  | 5.241% | 7.808% |  | -3.743% | 8.065% | 7.552% | -3.999% | -5.283% | -2.169% | -38.4% | -10.16% | 0.1076% | -0.9191% | 2.674% | 0.1076% | 25.52% | -2.459%
| Z-Score (Absolute) |  | 0.67 | 1.01 | 492.58 | 0.51 | 1.05 | 0.98 | 0.54 | 0.71 | 0.3 | 5.06 | 1.35 | 0 | 0.13 | 0.34 | 0 | 3.34 | 0.34
| Analytical results for K in OREAS 285 (Certified Value 1 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.98094 | 1.016 | 1.097 | 0.98563 | 0.94 | 0.962 | 0.89 | 0.99 | 1 | 1.17 | 1.02 | 1.02 | 1.04 | 0.92024 | 0.97 | 0.99 | 1.05 | 0.97 | 1.05 | 0.95 | 0.904 | 1.0227 | 1.262 | 1.1622 | 0.8072 | 1.25 | 1.07 |  |  | 1
| 1 | 2 | 0.98858 | 0.9991 | 1.1109 | 0.99409 | 0.93 | 0.929 | 0.9 | 1.06 | 1 | 1.08 | 1.01 | 1.01 | 1.03 | 0.96296 | 0.94 | 0.97 | 1.08 | 0.99 | 1.05 | 0.96 | 0.9168 | 1.0196 | 1.207 | 1.1373 | 0.8008 | 1.27 | 0.96 |  |  | #N/A
| 1 | 3 | 0.99033 | 0.9926 | 1.1024 | 0.98067 | 0.95 | 0.9255 | 0.9 | 1.04 | 1 | 1.17 | 1 | 1.02 | 1.03 | 0.9281 | 0.94 | 1.03 | 1.08 | 0.95 | 1.05 | 0.95 | 0.9472 | 1.068 | 1.188 | 1.129 | 0.82 | 1.25 | 1.03 |  |  | 16
| 1 | 4 | 0.9872 | 1.0153 | 1.0728 | 0.98603 | 0.97 | 0.9632 | 0.88 | 0.96 | 1.02 | 1.18 | 1.01 | 1.02 | 1.02 | 0.95507 | 0.95 | 1.03 | 1.07 | 0.95 | 1.03 | 0.96 | 0.9012 | 1.0178 | 1.321 | 1.1705 | 0.8184 | 1.27 | 1.04 |  |  | 0.99992
| 1 | 5 | 0.97663 | 1.0377 | 1.09 | 0.99647 | 0.97 | 0.9739 | 0.89 | 0.98 | 0.99 | 1.17 | 0.99 | 1.04 | 1.02 | 0.9655 | 0.94 | 1.01 | 1.06 | 0.99 | 1.02 | 0.96 | 0.9163 | 1.0264 | 1.294 | 1.1207 | 0.8408 | 1.24 | 1.06 |  |  | 96
| 1 | 6 | 0.98066 | 1.0388 | 1.1048 | 0.97808 | 0.9 | 0.9605 | 0.9 | 1.09 | 1.01 | 1.17 | 1 | 1.04 | 1.01 | 0.91218 | 0.95 | 1.03 | 1.06 | 0.99 | 1.03 | 0.96 | 0.9127 | 1.0491 | 1.397 | 1.1456 | 0.828 | 1.22 | 1.08
| Mean |  | 0.98406 | 1.0166 | 1.0963 | 0.98683 | 0.94333 | 0.95235 | 0.89333 | 1.02 | 1.0033 | 1.1567 | 1.005 | 1.025 | 1.025 | 0.94067 | 0.94833 | 1.01 | 1.0667 | 0.97333 | 1.0383 | 0.95667 | 0.91637 | 1.0339 | 1.2782 | 1.1442 | 0.8192 | 1.25 | 1.04
| Median |  | 0.98407 | 1.0156 | 1.0997 | 0.98583 | 0.945 | 0.96125 | 0.895 | 1.015 | 1 | 1.17 | 1.005 | 1.02 | 1.025 | 0.94158 | 0.945 | 1.02 | 1.065 | 0.98 | 1.04 | 0.96 | 0.9145 | 1.0246 | 1.278 | 1.1414 | 0.8192 | 1.25 | 1.05
| Std Dev. |  | 0.0054057 | 0.019091 | 0.01352 | 0.0072405 | 0.026583 | 0.020039 | 0.008165 | 0.05099 | 0.010328 | 0.037771 | 0.010488 | 0.012247 | 0.010488 | 0.02327 | 0.01169 | 0.025298 | 0.012111 | 0.019664 | 0.013292 | 0.005164 | 0.016415 | 0.020195 | 0.077023 | 0.019228 | 0.014338 | 0.018974 | 0.043359
| Rel.Std.Dev. |  | 0.5493% | 1.878% | 1.233% | 0.7337% | 2.818% | 2.104% | 0.914% | 4.999% | 1.029% | 3.266% | 1.044% | 1.195% | 1.023% | 2.474% | 1.233% | 2.505% | 1.135% | 2.02% | 1.28% | 0.5398% | 1.791% | 1.953% | 6.026% | 1.68% | 1.75% | 1.518% | 4.169%
| PDM3 |  | -1.586% | 1.667% | 9.642% | -1.309% | -5.659% | -4.757% | -10.66% | 2.009% | 0.3417% | 15.68% | 0.5084% | 2.509% | 2.509% | -5.925% | -5.159% | 1.008% | 6.676% | -2.659% | 3.842% | -4.325% | -8.356% | 3.402% | 27.83% | 14.43% | -18.07% | 25.01% | 4.009%
| Z-Score (Absolute) |  | 0.33 | 0.08 | 1.09 | 0.29 | 0.84 | 0.73 | 1.48 | 0.13 | 0.08 | 1.85 | 0.06 | 0.19 | 0.19 | 0.88 | 0.78 | 0 | 0.72 | 0.46 | 0.36 | 0.67 | 1.18 | 0.3 | 3.39 | 1.7 | 2.41 | 3.03 | 0.38
| Analytical results for La in OREAS 285 (Certified Value 18.2 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 1
| 1 | 1 | 19.3155 | 15.696 | 20.8 | 22.3044 | 17 | 17.4 | 15 | 21.8 | 17 | 15.8 | 16.4 | 19.7815 | 16 | 14.3 | 26.7 | 15.2 | 14.1 | 20.8 | 29 | 22.676 | 7 | 13.6448 | 22 | 19.2 | [5 empty] | 1
| 1 | 2 | 18.4669 | 16.836 | 20.5 | 22.1885 | 17 | 19.2 | 14.8 | 17.6 | 17.8 | 17 | 15.8 | 20.3087 | 16 | 14.4 | 26.9 | 15.2 | 12.9 | 20.1 | 30 | 23.421 | 7 | 13.9072 | 28 | 17 | [5 empty] | 25
| 1 | 3 | 18.7543 | 16.051 | 20.4 | 24.1144 | 17 | 19 | 15.1 | 20.6 | 17.4 | 16.2 | 14.8 | 22.7328 | 16 | 15 | 26.5 | 14.3 | 14 | 19.7 | 30 | 24.645 | 7 | 13.9456 | 26.4 | 19.4 | [5 empty] | 16
| 1 | 4 | 18.0316 | 17.68 | 22 | 22.304 | 17 | 16.7 | 15 | 20.9 | 17.2 | 16.4 | 16.4 | 22.8352 | 16 | 15.2 | 27.7 | 14.8 | 13.1 | 18.4 | 29 | 22.428 | 8 | 14.016 | 24.3 | 19.4 | [5 empty] | 18.2427
| 1 | 5 | 18.4068 | 17.896 | 20.5 | 23.3459 | 17 | 17.6 | 14.5 | 19.4 | 17 | 17 | 14.6 | 20.1931 | 16 | 15 | 27.4 | 13.8 | 12.7 | 19.3 | 30 | 24.508 | 8 | 14.432 | 24.5 | 19.7 | [5 empty] | 97
| 1 | 6 | 18.3362 | 17.52 | 20.1 | 22.3071 | 18 | 19.7 | 14.2 | 20.7 | 17.8 | 16.6 | 16.4 | 20.0465 | 16 | 15 | 27.5 | 12.9 | 14.4 | 18.6 | 29 | 24.225 | 8 | 14.1312 | 22.3 | 19.4
| Mean |  | 18.5519 | 16.9465 | 20.7167 | 22.7607 | 17.1667 | 18.2667 | 14.7667 | 20.1667 | 17.3667 | 16.5 | 15.7333 | 20.983 | 16 | 14.8167 | 27.1167 | 14.3667 | 13.5333 | 19.4833 | 29.5 | 23.6505 | 7.5 | 14.0128 | 24.5833 | 19.0167
| Median |  | 18.4369 | 17.178 | 20.5 | 22.3057 | 17 | 18.3 | 14.9 | 20.65 | 17.3 | 16.5 | 16.1 | 20.2509 | 16 | 15 | 27.15 | 14.55 | 13.55 | 19.5 | 29.5 | 23.823 | 7.5 | 13.9808 | 24.4 | 19.4
| Std Dev. |  | 0.44036 | 0.91064 | 0.66758 | 0.79056 | 0.40825 | 1.1928 | 0.35024 | 1.4733 | 0.36697 | 0.46904 | 0.83586 | 1.4065 | 0 | 0.37103 | 0.48339 | 0.90037 | 0.7174 | 0.91086 | 0.54772 | 0.95407 | 0.54772 | 0.26109 | 2.3216 | 1.0008
| Rel.Std.Dev. |  | 2.374% | 5.374% | 3.222% | 3.473% | 2.378% | 6.53% | 2.372% | 7.306% | 2.113% | 2.843% | 5.313% | 6.703% | 0% | 2.504% | 1.783% | 6.267% | 5.301% | 4.675% | 1.857% | 4.034% | 7.303% | 1.863% | 9.444% | 5.263%
| PDM3 |  | 1.695% | -7.105% | 13.56% | 24.77% | -5.898% | 0.1314% | -19.05% | 10.55% | -4.802% | -9.553% | -13.76% | 15.02% | -12.29% | -18.78% | 48.64% | -21.25% | -25.82% | 6.801% | 61.71% | 29.64% | -58.89% | -23.19% | 34.76% | 4.243%
| Z-Score (Absolute) |  | 0.04 | 0.38 | 0.59 | 1.12 | NA | 0.04 | 0.94 | 0.45 | 0.27 | 0.49 | 0.69 | 0.66 | NA | 0.92 | 2.24 | 1.04 | 1.25 | 0.28 | NA | 1.35 | NA | 1.13 | 1.59 | 0.16
Results from laboratories 1.11, 1.20, 1.26 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 285 (Certified Value 43.6 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 2
| 1 | 1 | 43.7313 | 41.9 | 55.0935 | 45.6399 | 45.2 | 41.8699 | 44 | 42.7 | 48 | 44.3 | 45 | 45.2 | 43.9482 | 40.9 | 42 | 41.5 | 38.8 | 44 | 43 | 43.8 | 39 | 102.96 | 44.2 | 44 | [5 empty] | 1
| 1 | 2 | 44.6017 | 41.5 | 56.2251 | 46.7639 | 45.6 | 46.0103 | 44 | 42.5 | 46.2 | 44 | 44.8 | 44.8 | 43.5304 | 39.7 | 41 | 42.7 | 38.6 | 44 | 44 | 42.5 | 38 | 102 | 46.7 | 39 | [5 empty] | 45
| 1 | 3 | 43.7736 | 41.1 | 56.3735 | 45.0622 | 45 | 48.4872 | 45 | 43.3 | 48 | 43.2 | 45.5 | 45 | 43.954 | 40.6 | 45 | 42.1 | 37.6 | 43 | 45 | 44.8 | 39 | 102.96 | 46.1 | 43 | [5 empty] | 16
| 1 | 4 | 43.7643 | 41.5 | 53.9002 | 46.9357 | 46.2 | 44.0294 | 44 | 43.6 | 48.3 | 43.8 | 45.6 | 45 | 43.2617 | 39.9 | 45 | 42.2 | 37.4 | 44 | 43 | 43 | 40 | 102.64 | 45.4 | 43 | [5 empty] | 43.5649
| 1 | 5 | 43.6022 | 42.4 | 55.6862 | 46.4008 | 46.1 | 46.0854 | 44 | 42.6 | 47.8 | 43 | 46 | 45.4 | 42.928 | 40.4 | 44 | 41.2 | 37.5 | 44 | 43 | 43.6 | 38 | 105.12 | 45.7 | 44 | [5 empty] | 98
| 1 | 6 | 43.5096 | 42.4 | 55.4662 | 45.7276 | 46.1 | 47.1598 | 45 | 42.9 | 47.8 | 43.6 | 46.1 | 44.4 | 42.7098 | 42.4 | 45 | 42.1 | 36.8 | 43 | 43 | 43.5 | 39 | 102.72 | 44.6 | 44
| Mean |  | 43.8304 | 41.8 | 55.4575 | 46.0883 | 45.7 | 45.607 | 44.3333 | 42.9333 | 47.6833 | 43.65 | 45.5 | 44.9667 | 43.3887 | 40.65 | 43.6667 | 41.9667 | 37.7833 | 43.6667 | 43.5 | 43.5333 | 38.8333 | 103.0667 | 45.45 | 42.8333
| Median |  | 43.7478 | 41.7 | 55.5762 | 46.0642 | 45.85 | 46.0478 | 44 | 42.8 | 47.9 | 43.7 | 45.55 | 45 | 43.396 | 40.5 | 44.5 | 42.1 | 37.55 | 44 | 43 | 43.55 | 39 | 102.84 | 45.55 | 43.5
| Std Dev. |  | 0.39174 | 0.52915 | 0.89868 | 0.72887 | 0.51381 | 2.3475 | 0.5164 | 0.43205 | 0.74944 | 0.48888 | 0.52154 | 0.34448 | 0.51805 | 0.96488 | 1.7512 | 0.53541 | 0.76529 | 0.5164 | 0.83666 | 0.77889 | 0.75277 | 1.0657 | 0.93113 | 1.9408
| Rel.Std.Dev. |  | 0.8938% | 1.266% | 1.62% | 1.581% | 1.124% | 5.147% | 1.165% | 1.006% | 1.572% | 1.12% | 1.146% | 0.7661% | 1.194% | 2.374% | 4.01% | 1.276% | 2.025% | 1.183% | 1.923% | 1.789% | 1.938% | 1.034% | 2.049% | 4.531%
| PDM3 |  | 0.6095% | -4.051% | 27.3% | 5.792% | 4.901% | 4.687% | 1.764% | -1.45% | 9.454% | 0.1953% | 4.442% | 3.218% | -0.4045% | -6.691% | 0.2336% | -3.669% | -13.27% | 0.2336% | -0.149% | -0.07246% | -10.86% | 136.6% | 4.327% | -1.679%
| Z-Score (Absolute) |  | 0.06 | 0.72 | 4.56 | 0.94 | 0.79 | 0.75 | 0.26 | 0.28 | 1.56 | 0.01 | 0.71 | 0.5 | 0.11 | 1.17 | 0 | 0.66 | 2.28 | 0 | 0.06 | 0.05 | 1.87 | 23 | 0.69 | 0.32
| Analytical results for Lu in OREAS 285 (Certified Value 0.095 ppm) | [30 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.21 | 1.25 | 1.32 | [22 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [22 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [22 empty] | 3
| 1 | 1 | 0.099536 | 0.088 | < 0.1 | 0.0931 | 0.1 | 0.1 | 0.09 | [22 empty] | 1
| 1 | 2 | 0.10231 | 0.089 | < 0.1 | 0.0925 | 0.1 | 0.09 | 0.07 | [22 empty] | 26
| 1 | 3 | 0.10651 | 0.087 | 0.1 | 0.0922 | 0.1 | 0.1 | 0.09 | [22 empty] | 16
| 1 | 4 | 0.10366 | 0.088 | < 0.1 | 0.0902 | 0.1 | 0.1 | 0.09 | [22 empty] | 0.094899
| 1 | 5 | 0.10287 | 0.091 | < 0.1 | 0.0943 | 0.1 | 0.1 | 0.1 | [22 empty] | 99
| 1 | 6 | 0.10231 | 0.09 | < 0.1 | 0.0869 | 0.1 | 0.1 | 0.09
| Mean |  | 0.10287 | 0.088833 | 0.1 | 0.091533 | 0.1 | 0.098333 | 0.088333
| Median |  | 0.10259 | 0.0885 | 0.1 | 0.09235 | 0.1 | 0.1 | 0.09
| Std Dev. |  | 0.0022646 | 0.001472 |  | 0.0026357 | 1.5202e-17 | 0.0040825 | 0.0098319
| Rel.Std.Dev. |  | 2.201% | 1.657% |  | 2.879% | 1.52e-14% | 4.152% | 11.13%
| PDM3 |  | 8.396% | -6.391% | 5.376% | -3.546% | 5.376% | 3.619% | -6.918%
| Z-Score (Absolute) |  | 2.39 | 0.57 | NA | 0 | NA | 1.43 | 0.67
Results from laboratory 1.21 were removed due to their 0.01 ppm reading resolution. Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 285 (Certified Value 0.874 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.89741 | 0.89 | 0.93271 | 0.87068 | 0.89 | 0.8399 | 0.84 | 0.83 | 0.83 | 0.96 | 0.87 | 0.87 | 0.89 | 0.87307 | 0.88 | 0.76 | 0.97 | 0.84 | 0.88 | 0.86 | 0.86 | 0.87 | 1.013 | 0.953 | 0.9539 | 0.886 | 0.95 |  |  | 1
| 1 | 2 | 0.89149 | 0.87 | 0.93962 | 0.89756 | 0.86 | 0.8229 | 0.85 | 0.89 | 0.84 | 0.93 | 0.86 | 0.86 | 0.88 | 0.86925 | 0.86 | 0.75 | 0.98 | 0.86 | 0.88 | 0.88 | 0.87 | 0.87 | 0.973 | 0.905 | 0.9397 | 0.843 | 0.84 |  |  | #N/A
| 1 | 3 | 0.89776 | 0.87 | 0.92904 | 0.89694 | 0.87 | 0.8314 | 0.85 | 0.88 | 0.85 | 0.97 | 0.85 | 0.88 | 0.88 | 0.88932 | 0.85 | 0.8 | 0.98 | 0.83 | 0.88 | 0.86 | 0.9 | 0.89 | 0.965 | 0.917 | 0.94 | 0.822 | 0.91 |  |  | 16
| 1 | 4 | 0.8875 | 0.89 | 0.90773 | 0.87433 | 0.91 | 0.8486 | 0.82 | 0.83 | 0.86 | 0.97 | 0.86 | 0.88 | 0.88 | 0.88075 | 0.86 | 0.8 | 0.98 | 0.83 | 0.87 | 0.87 | 0.86 | 0.85 | 1.053 | 0.953 | 0.943 | 0.81 | 0.92 |  |  | 0.87388
| 1 | 5 | 0.88297 | 0.91 | 0.92505 | 0.87277 | 0.86 | 0.8601 | 0.83 | 0.84 | 0.84 | 0.97 | 0.85 | 0.89 | 0.88 | 0.88071 | 0.85 | 0.79 | 0.95 | 0.87 | 0.86 | 0.87 | 0.87 | 0.87 | 1.035 | 0.905 | 0.9536 | 0.8 | 0.93 |  |  | 100
| 1 | 6 | 0.88358 | 0.91 | 0.93247 | 0.89036 | 0.87 | 0.8484 | 0.83 | 0.91 | 0.84 | 0.96 | 0.86 | 0.89 | 0.87 | 0.88793 | 0.86 | 0.79 | 0.95 | 0.87 | 0.86 | 0.87 | 0.86 | 0.88 | 1.104 | 0.923 | 0.9088 | 0.84 | 0.94
| Mean |  | 0.89012 | 0.89 | 0.92777 | 0.88377 | 0.87667 | 0.84188 | 0.83667 | 0.86333 | 0.84333 | 0.96 | 0.85833 | 0.87833 | 0.88 | 0.88017 | 0.86 | 0.78167 | 0.96833 | 0.85 | 0.87167 | 0.86833 | 0.87 | 0.87167 | 1.0238 | 0.926 | 0.93983 | 0.8335 | 0.915
| Median |  | 0.88949 | 0.89 | 0.93076 | 0.88235 | 0.87 | 0.84415 | 0.835 | 0.86 | 0.84 | 0.965 | 0.86 | 0.88 | 0.88 | 0.88073 | 0.86 | 0.79 | 0.975 | 0.85 | 0.875 | 0.87 | 0.865 | 0.87 | 1.024 | 0.92 | 0.9415 | 0.831 | 0.925
| Std Dev. |  | 0.0065446 | 0.017889 | 0.010932 | 0.012557 | 0.019664 | 0.013369 | 0.012111 | 0.034448 | 0.010328 | 0.015492 | 0.0075277 | 0.01169 | 0.0063246 | 0.0079277 | 0.010954 | 0.02137 | 0.01472 | 0.018974 | 0.0098319 | 0.0075277 | 0.015492 | 0.013292 | 0.052086 | 0.022045 | 0.016496 | 0.030645 | 0.03937
| Rel.Std.Dev. |  | 0.7352% | 2.01% | 1.178% | 1.421% | 2.243% | 1.588% | 1.447% | 3.99% | 1.225% | 1.614% | 0.877% | 1.331% | 0.7187% | 0.9007% | 1.274% | 2.734% | 1.52% | 2.232% | 1.128% | 0.8669% | 1.781% | 1.525% | 5.087% | 2.381% | 1.755% | 3.677% | 4.303%
| PDM3 |  | 1.859% | 1.845% | 6.167% | 1.133% | 0.3194% | -3.661% | -4.258% | -1.206% | -3.495% | 9.856% | -1.778% | 0.5102% | 0.7009% | 0.7207% | -1.588% | -10.55% | 10.81% | -2.732% | -0.2527% | -0.6342% | -0.4434% | -0.2527% | 17.16% | 5.965% | 7.548% | -4.62% | 4.706%
| Z-Score (Absolute) |  | 0.49 | 0.49 | 1.88 | 0.26 | 0 | 1.28 | 1.47 | 0.49 | 1.23 | 3.07 | 0.67 | 0.06 | 0.12 | 0.13 | 0.61 | 3.49 | 3.37 | 0.98 | 0.18 | 0.31 | 0.25 | 0.18 | 5.41 | 1.8 | 2.32 | 1.59 | 1.41
| Analytical results for Mn in OREAS 285 (Certified Value 0.034 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.034404 | 0.0326 | 0.035206 | 0.034347 | 0.0376 | 0.0322 | 0.0389 | 0.032 | 0.0317 | 0.0353 | 0.0336 | 0.0347 | 0.0339 | 0.033556 | 0.0311 | 0.0324 | 0.0341 | 0.0331 | 0.034 | 0.0337 | 0.033 | 0.0351 | 0.032 | 0.0356 | 0.04113 | 0.0366 | 0.0365 |  |  | 1
| 1 | 2 | 0.034494 | 0.0321 | 0.034731 | 0.034517 | 0.0386 | 0.0314 | 0.0393 | 0.0343 | 0.0317 | 0.0334 | 0.0334 | 0.0352 | 0.0332 | 0.03441 | 0.0313 | 0.0316 | 0.0344 | 0.0342 | 0.0336 | 0.0343 | 0.0334 | 0.0345 | 0.029 | 0.0341 | 0.04061 | 0.0402 | 0.0325 |  |  | 27
| 1 | 3 | 0.034996 | 0.0319 | 0.034604 | 0.034215 | 0.0385 | 0.0319 | 0.0393 | 0.034 | 0.0321 | 0.0351 | 0.0332 | 0.0348 | 0.0335 | 0.034641 | 0.0296 | 0.0326 | 0.0345 | 0.0335 | 0.0336 | 0.0336 | 0.0345 | 0.0362 | 0.027 | 0.0349 | 0.04085 | 0.0396 | 0.0349 |  |  | 16
| 1 | 4 | 0.033864 | 0.0326 | 0.03554 | 0.034258 | 0.04 | 0.0326 | 0.0387 | 0.032 | 0.0322 | 0.0352 | 0.0332 | 0.0356 | 0.0331 | 0.034927 | 0.0307 | 0.0325 | 0.034 | 0.0328 | 0.033 | 0.0339 | 0.0329 | 0.0345 |  | 0.0356 | 0.04083 | 0.0385 | 0.0353 |  |  | 0.033739
| 1 | 5 | 0.034031 | 0.0335 | 0.034311 | 0.034723 | 0.0384 | 0.0331 | 0.0389 | 0.0322 | 0.0317 | 0.0354 | 0.0329 | 0.0358 | 0.0334 | 0.033745 | 0.0314 | 0.0332 | 0.0342 | 0.0342 | 0.0324 | 0.0337 | 0.0335 | 0.0352 |  | 0.0341 | 0.04116 | 0.0385 | 0.0358 |  |  | 101
| 1 | 6 | 0.034746 | 0.0334 | 0.034049 | 0.033701 | 0.0374 | 0.0323 | 0.0394 | 0.035 | 0.0319 | 0.0353 | 0.0335 | 0.0357 | 0.0328 | 0.033981 | 0.0302 | 0.0334 | 0.0344 | 0.0332 | 0.0332 | 0.0336 | 0.0333 | 0.0352 | 0.03 | 0.0349 | 0.04057 | 0.0361 | 0.0362
| Mean |  | 0.034422 | 0.032683 | 0.03474 | 0.034293 | 0.038417 | 0.03225 | 0.039083 | 0.03325 | 0.031883 | 0.03495 | 0.0333 | 0.0353 | 0.033317 | 0.03421 | 0.030717 | 0.032617 | 0.034267 | 0.0335 | 0.0333 | 0.0338 | 0.033433 | 0.035117 | 0.0295 | 0.034867 | 0.040858 | 0.03825 | 0.0352
| Median |  | 0.034449 | 0.0326 | 0.034667 | 0.034303 | 0.03845 | 0.03225 | 0.0391 | 0.0331 | 0.0318 | 0.03525 | 0.0333 | 0.0354 | 0.0333 | 0.034195 | 0.0309 | 0.03255 | 0.0343 | 0.03335 | 0.0334 | 0.0337 | 0.03335 | 0.03515 | 0.0295 | 0.0349 | 0.04084 | 0.0385 | 0.03555
| Std Dev. |  | 0.0004251 | 0.00065549 | 0.00055443 | 0.00034521 | 0.00092177 | 0.00058224 | 0.00028577 | 0.0013383 | 0.00022286 | 0.00076616 | 0.00025298 | 0.00047329 | 0.00037639 | 0.00053593 | 0.00070261 | 0.00064005 | 0.00019664 | 0.00058652 | 0.00056214 | 0.00026833 | 0.00057155 | 0.00062423 | 0.0020817 | 0.00067132 | 0.00024919 | 0.0016183 | 0.001445
| Rel.Std.Dev. |  | 1.235% | 2.006% | 1.596% | 1.007% | 2.399% | 1.805% | 0.7312% | 4.025% | 0.699% | 2.192% | 0.7597% | 1.341% | 1.13% | 1.567% | 2.287% | 1.962% | 0.5738% | 1.751% | 1.688% | 0.7939% | 1.71% | 1.778% | 7.056% | 1.925% | 0.6099% | 4.231% | 4.105%
| PDM3 |  | 2.025% | -3.129% | 2.967% | 1.643% | 13.86% | -4.414% | 15.84% | -1.45% | -5.5% | 3.589% | -1.301% | 4.626% | -1.252% | 1.396% | -8.958% | -3.327% | 1.564% | -0.7086% | -1.301% | 0.1805% | -0.9062% | 4.083% | -12.56% | 3.342% | 21.1% | 13.37% | 4.33%
| Z-Score (Absolute) |  | 0.16 | 1.13 | 0.39 | 0.06 | 3.12 | 1.45 | 3.61 | 0.71 | 1.72 | 0.55 | 0.67 | 0.81 | 0.66 | 0 | 2.59 | 1.18 | 0.04 | 0.53 | 0.67 | 0.3 | 0.58 | 0.67 | 3.49 | 0.47 | 4.93 | 2.99 | 0.73
| Analytical results for Mo in OREAS 285 (Certified Value 2.45 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 2
| 1 | 1 | 2.785 | 2.5 | 2.327 | < 5 | 2.68 | 1.8055 | 3 | 1.7 | 2.17 | 1.92 | 2.18 | 2.51 | 2.24 | 2.3 | 2.4609 | 2.8 | 2.67 | 2.8 | 2.5 | 2.55 | 2.9 | 2 | 2.7 | 2.43 | 2.35 | 25.42 | 3.07 | 2.67 |  | 1
| 1 | 2 | 2.8866 | 2.4 | 2.7013 | < 5 | 2.71 | 2.6164 | 3 | 1.9 | 2.26 | 1.96 | 1.95 | 2.4 | 2.25 | 2.29 | 2.5656 | 2.7 | 2.71 | 2.8 | 2.35 | 2.56 | 2.57 | 3 | 2.4 | 2.51 | 2.28 | 25.11 | 3.21 | 2.2 |  | 28
| 1 | 3 | 2.7372 | 2.4 | 2.3877 | < 5 | 2.84 | 2.0261 | 3 | 2 | 2.37 | 1.95 | 2.07 | 2.51 | 2.26 | 2.07 | 2.5648 | 2.7 | 2.79 | 2.8 | 2.44 | 2.42 | 2.78 | 3 | 2.6 | 2.64 | 2.29 | 25.46 | 3.03 | 2.16 |  | 16
| 1 | 4 | 2.6896 | 2.4 | 2.4633 | < 5 | 3.04 | 1.6706 | 3 | 1.9 | 2.14 | 2.05 | 2.1 | 2.61 | 2.25 | 2.26 | 2.4795 | 2.6 | 2.92 | 2.7 | 2.44 | 2.6 | 2.57 | 3 | 2.6 |  | 2.46 | 25.17 | 3.88 | 2.55 |  | 2.4542
| 1 | 5 | 2.6557 | 2.7 | 2.4155 | < 5 | 2.66 | 2.1246 | 3 | 1.9 | 2.09 | 2.04 | 1.83 | 2.37 | 2.27 | 1.96 | 2.532 | 2.8 | 2.86 | 2.7 | 2.44 | 2.53 | 2.67 | 3 | 2.6 | 2.4 | 2.4 | 25.48 | 3.15 | 2.3 |  | 102
| 1 | 6 | 2.7912 | 2.6 | 2.4284 | < 5 | 2.73 | 2.8403 | 3 | 1.8 | 2.14 | 1.81 | 2.19 | 2.41 | 2.32 | 2.12 | 2.4946 | 2.8 | 2.68 | 2.7 | 2.59 | 2.52 | 2.77 | 2 | 2.6 |  | 2.25 | 24.46 | 3.05 | 2.37
| Mean |  | 2.7575 | 2.5 | 2.4539 |  | 2.7767 | 2.1806 | 3 | 1.8667 | 2.195 | 1.955 | 2.0533 | 2.4683 | 2.265 | 2.1667 | 2.5162 | 2.7333 | 2.7717 | 2.75 | 2.46 | 2.53 | 2.71 | 2.6667 | 2.5833 | 2.495 | 2.3383 | 25.1833 | 3.2317 | 2.375
| Median |  | 2.7611 | 2.45 | 2.4219 |  | 2.72 | 2.0754 | 3 | 1.9 | 2.155 | 1.955 | 2.085 | 2.46 | 2.255 | 2.19 | 2.5133 | 2.75 | 2.75 | 2.75 | 2.44 | 2.54 | 2.72 | 3 | 2.6 | 2.47 | 2.32 | 25.295 | 3.11 | 2.335
| Std Dev. |  | 0.082376 | 0.12649 | 0.12955 |  | 0.14348 | 0.4589 | 0 | 0.10328 | 0.10252 | 0.087807 | 0.13981 | 0.090866 | 0.02881 | 0.13852 | 0.044558 | 0.08165 | 0.10265 | 0.054772 | 0.07975 | 0.060663 | 0.13069 | 0.5164 | 0.098319 | 0.10724 | 0.080353 | 0.38713 | 0.32474 | 0.20007
| Rel.Std.Dev. |  | 2.987% | 5.06% | 5.279% |  | 5.167% | 21.04% | 0% | 5.533% | 4.671% | 4.491% | 6.809% | 3.681% | 1.272% | 6.393% | 1.771% | 2.987% | 3.703% | 1.992% | 3.242% | 2.398% | 4.823% | 19.36% | 3.806% | 4.298% | 3.436% | 1.537% | 10.05% | 8.424%
| PDM3 |  | 12.36% | 1.865% | -0.01507% |  | 13.14% | -11.15% | 22.24% | -23.94% | -10.56% | -20.34% | -16.33% | 0.5749% | -7.71% | -11.72% | 2.526% | 11.37% | 12.93% | 12.05% | 0.2353% | 3.088% | 10.42% | 8.656% | 5.261% | 1.661% | -4.722% | 926.1% | 31.68% | -3.228%
| Z-Score (Absolute) |  | 0.75 | 0.01 | 0.13 | 0.01 | 0.8 | 0.91 | NA | 1.82 | 0.87 | 1.56 | 1.28 | 0.08 | 0.67 | 0.95 | 0.05 | 0.68 | 0.79 | 0.73 | 0.11 | 0.09 | 0.61 | NA | 0.25 | 0.01 | 0.46 | 65.33 | 2.11 | 0.35
Results from laboratories 1.11 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 285 (Certified Value 0.156 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.15951 | 0.13 | 0.17231 | 0.18398 | 0.195 | 0.1955 | 0.16 | 0.138 | 0.14 | 0.31 | 0.17 | 0.15 | 0.18 | 0.15744 | 0.16 | 0.16 | 0.12 | 0.14 | 0.161 | 0.13 | 0.13 | 0.13 | 0.11529 | 0.156 | 0.1008 | 0.158 | 0.18 |  |  | 1
| 1 | 2 | 0.1576 | 0.13 | 0.17204 | 0.18165 | 0.192 | 0.2243 | 0.16 | 0.152 | 0.14 | 0.26 | 0.17 | 0.15 | 0.18 | 0.14984 | 0.16 | 0.17 | 0.13 | 0.14 | 0.16 | 0.13 | 0.13 | 0.13 | 0.12085 | 0.148 | 0.0902 | 0.181 | 0.16 |  |  | #N/A
| 1 | 3 | 0.15663 | 0.13 | 0.17682 | 0.18517 | 0.199 | 0.1933 | 0.16 | 0.148 | 0.14 | 0.31 | 0.16 | 0.16 | 0.17 | 0.14988 | 0.16 | 0.18 | 0.13 | 0.13 | 0.16 | 0.13 | 0.13 | 0.14 | 0.12907 | 0.148 | 0.0955 | 0.167 | 0.17 |  |  | 16
| 1 | 4 | 0.1542 | 0.14 | 0.17683 | 0.18649 | 0.209 | 0.1945 | 0.16 | 0.133 | 0.15 | 0.31 | 0.17 | 0.16 | 0.16 | 0.15468 | 0.16 | 0.18 | 0.13 | 0.13 | 0.156 | 0.13 | 0.13 | 0.13 | 0.11744 | 0.156 | 0.0919 | 0.16 | 0.17 |  |  | 0.15554
| 1 | 5 | 0.15345 | 0.14 | 0.17228 | 0.18815 | 0.193 | 0.2054 | 0.16 | 0.138 | 0.14 | 0.31 | 0.16 | 0.16 | 0.16 | 0.15716 | 0.15 | 0.17 | 0.13 | 0.14 | 0.154 | 0.13 | 0.13 | 0.14 | 0.12411 | 0.148 | 0.0967 | 0.16 | 0.18 |  |  | 103
| 1 | 6 | 0.1555 | 0.14 | 0.16614 | 0.18095 | 0.186 | 0.2191 | 0.17 | 0.155 | 0.14 | 0.31 | 0.16 | 0.16 | 0.16 | 0.15276 | 0.16 | 0.18 | 0.12 | 0.13 | 0.159 | 0.13 | 0.13 | 0.14 | 0.11859 | 0.156 | 0.0916 | 0.15 | 0.18
| Mean |  | 0.15615 | 0.135 | 0.17274 | 0.1844 | 0.19567 | 0.20535 | 0.16167 | 0.144 | 0.14167 | 0.30167 | 0.165 | 0.15667 | 0.16833 | 0.15363 | 0.15833 | 0.17333 | 0.12667 | 0.135 | 0.15833 | 0.13 | 0.13 | 0.135 | 0.12089 | 0.152 | 0.09445 | 0.16267 | 0.17333
| Median |  | 0.15606 | 0.135 | 0.17229 | 0.18457 | 0.194 | 0.20045 | 0.16 | 0.143 | 0.14 | 0.31 | 0.165 | 0.16 | 0.165 | 0.15372 | 0.16 | 0.175 | 0.13 | 0.135 | 0.1595 | 0.13 | 0.13 | 0.135 | 0.11972 | 0.152 | 0.0937 | 0.16 | 0.175
| Std Dev. |  | 0.0022416 | 0.0054772 | 0.0039478 | 0.0027822 | 0.0077889 | 0.013477 | 0.0040825 | 0.0088769 | 0.0040825 | 0.020412 | 0.0054772 | 0.005164 | 0.0098319 | 0.0033836 | 0.0040825 | 0.008165 | 0.005164 | 0.0054772 | 0.0027325 | 0 | 0 | 0.0054772 | 0.0050176 | 0.0043818 | 0.0039793 | 0.010501 | 0.008165
| Rel.Std.Dev. |  | 1.436% | 4.057% | 2.285% | 1.509% | 3.981% | 6.563% | 2.525% | 6.165% | 2.882% | 6.767% | 3.32% | 3.296% | 5.841% | 2.202% | 2.578% | 4.711% | 4.077% | 4.057% | 1.726% | 0% | 0% | 4.057% | 4.15% | 2.883% | 4.213% | 6.455% | 4.711%
| PDM3 |  | 0.3916% | -13.2% | 11.06% | 18.56% | 25.8% | 32.03% | 3.941% | -7.418% | -8.918% | 93.95% | 6.084% | 0.726% | 8.227% | -1.229% | 1.798% | 11.44% | -18.56% | -13.2% | 1.798% | -16.42% | -16.42% | -13.2% | -22.27% | -2.274% | -39.28% | 4.584% | 11.44%
| Z-Score (Absolute) |  | 0.02 | 0.88 | 0.65 | 1.12 | 1.58 | 1.97 | 0.2 | 0.51 | 0.61 | 5.87 | 0.34 | 0 | 0.47 | 0.12 | 0.07 | 0.67 | 1.21 | 0.88 | 0.07 | 1.08 | 1.08 | 0.88 | 1.45 | 0.19 | 2.52 | 0.24 | 0.67
| Analytical results for Nb in OREAS 285 (Certified Value 0.79 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [10 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [10 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [10 empty] | 2
| 1 | 1 | 1.0454 | 0.88 | 0.5 | < 5 | 0.72 | 0.9 | 0.63 | 0.94 | 0.93 | 1.1965 | 1.1 | 1.02 | 0.6 | 1.8 | 1.89 | 0.62 | 0.23 | 1.61 | 1.61 | [10 empty] | 1
| 1 | 2 | 1.0436 | 0.86 | 0.5 | < 5 | 0.68 | 0.77 | 0.64 | 1.08 | 0.77 | 1.2214 | 1.1 | 1.03 | 0.6 | 1.9 | 1.88 | 0.68 | 0.24 | 0.46 | 1.99 | [10 empty] | 29
| 1 | 3 | 1.0359 | 1 | 0.5 | < 5 | 0.63 | 0.84 | 0.65 | 0.8 | 0.82 | 1.2085 | 1 | 1.05 | 0.6 | 1.79 | 1.82 | 0.67 | 0.26 | 0.46 | 2.13 | [10 empty] | 16
| 1 | 4 | 1.012 | 0.85 | 0.5 | < 5 | 0.57 | 0.83 | 0.62 | 0.84 | 0.83 | 1.1337 | 1 | 1.13 | 0.6 | 1.78 | 1.66 | 0.6 | 0.26 | 0.59 | 2.03 | [10 empty] | 0.78801
| 1 | 5 | 1.0944 | 0.82 | 0.6 | < 5 | 0.58 | 0.81 | 0.71 | 0.99 | 0.76 | 1.1502 | 1 | 1.02 | 0.6 | 1.98 | 1.8 | 0.65 | 0.25 | 0.52 | 1.74 | [10 empty] | 104
| 1 | 6 | 1.0551 | 0.81 | 0.7 | < 5 | 0.68 | 0.92 | 0.68 | 0.83 | 0.84 | 1.1698 | 0.8 | 1.13 | 0.6 | 1.71 | 1.81 | 0.65 | 0.25 | 0.7 | 2.53
| Mean |  | 1.0477 | 0.87 | 0.55 |  | 0.64333 | 0.845 | 0.655 | 0.91333 | 0.825 | 1.18 | 1 | 1.0633 | 0.6 | 1.8267 | 1.81 | 0.645 | 0.24833 | 0.72333 | 2.005
| Median |  | 1.0445 | 0.855 | 0.5 |  | 0.655 | 0.835 | 0.645 | 0.89 | 0.825 | 1.1832 | 1 | 1.04 | 0.6 | 1.795 | 1.815 | 0.65 | 0.25 | 0.555 | 2.01
| Std Dev. |  | 0.027068 | 0.068702 | 0.083666 |  | 0.060222 | 0.056125 | 0.033912 | 0.10912 | 0.06091 | 0.034457 | 0.10954 | 0.052789 | 0 | 0.096678 | 0.082462 | 0.030166 | 0.01169 | 0.44374 | 0.32198
| Rel.Std.Dev. |  | 2.583% | 7.897% | 15.21% |  | 9.361% | 6.642% | 5.177% | 11.95% | 7.383% | 2.92% | 10.95% | 4.964% | 0% | 5.293% | 4.556% | 4.677% | 4.708% | 61.35% | 16.06%
| PDM3 |  | 32.96% | 10.41% | -30.2% |  | -18.36% | 7.233% | -16.88% | 15.9% | 4.694% | 49.75% | 26.9% | 34.94% | -23.86% | 131.8% | 129.7% | -18.15% | -68.49% | -8.207% | 154.4%
| Z-Score (Absolute) |  | 0.44 | 0.06 | NA | 4.49 | 0.69 | 0.13 | 0.66 | 0.06 | 0.19 | 0.8 | NA | 0.48 | NA | 2.61 | 2.56 | 0.69 | 1.8 | 0.47 | 3.11
Results from laboratories 1.09, 1.20 and 1.22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 285 (Indicative Value 19.3 ppm) | [30 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 1
| 1 | 1 | 22.2112 | 16.535 | 22.2 | 25.2508 | 8.94 | [24 empty] | 1
| 1 | 2 | 21.7261 | 17.831 | 22 | 25.9371 | 8.57 | [24 empty] | 6
| 1 | 3 | 22.7823 | 16.857 | 21.6 | 27.0048 | 9.02 | [24 empty] | 16
| 1 | 4 | 20.9589 | 18.478 | 23.5 | 24.8898 | 9.57 | [24 empty] | 19.3271
| 1 | 5 | 22.7489 | 18.646 | 21.9 | 26.2709 | 9.07 | [24 empty] | 12
| 1 | 6 | 22.3845 | 18.088 | 22 | 25.2016 | 9.2
| Mean |  | 22.1353 | 17.7392 | 22.2 | 25.7592 | 9.0617
| Median |  | 22.2978 | 17.9595 | 22 | 25.5939 | 9.045
| Std Dev. |  | 0.69486 | 0.86332 | 0.66633 | 0.79627 | 0.32738
| Rel.Std.Dev. |  | 3.139% | 4.867% | 3.002% | 3.091% | 3.613%
| PDM3 |  | 14.53% | -8.216% | 14.86% | 33.28% | -53.11%
| Z-Score (Absolute) |  | 0 | 0.82 | 0.01 | 0.67 | 2.43
| Analytical results for Ni in OREAS 285 (Certified Value 19.3 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 19.5736 | 20.5 | 21.3736 | 18.6681 | 18.9 | 18.9221 | 20 | 17.5 | 18.9 | 20.7 | 20.3 | 19 | 18.8 | 17.7334 | 17.4 | 17.2 | 20 | 19.2 | 19.1 | 18.7 | 18 | 20.5 | 18.52 | 15.3 | 19.28 | 18.8 | 19.5 |  |  | 1
| 1 | 2 | 19.7808 | 20.5 | 20.5266 | 19.4284 | 19.4 | 18.8483 | 20 | 18.6 | 19.5 | 18.3 | 20.4 | 19.8 | 18.8 | 18.5639 | 17.2 | 17.4 | 20 | 19.3 | 18.9 | 19.1 | 19 | 19.5 | 18.74 | 15 | 20.42 | 20.7 | 17.5 |  |  | 16
| 1 | 3 | 19.5307 | 20.4 | 19.6923 | 19.0975 | 20 | 20.8598 | 20 | 18.6 | 19.1 | 20.7 | 19.9 | 18.9 | 18.2 | 18.6865 | 17 | 18.8 | 20 | 18.5 | 18.5 | 19.3 | 20 | 20.9 | 18.27 | 15.4 | 19.16 | 20.4 | 19.1 |  |  | 16
| 1 | 4 | 19.3832 | 20.7 | 20.6344 | 19.3344 | 20.7 | 19.3355 | 20 | 17.4 | 19.5 | 21 | 20 | 19 | 20.2 | 18.2675 | 17 | 18 | 20 | 19.1 | 18.8 | 18.5 | 18 | 19.4 | 15.67 | 15.8 | 19.23 | 19.7 | 19.4 |  |  | 19.2528
| 1 | 5 | 19.6382 | 20.9 | 20.4945 | 19.1422 | 20 | 19.1863 | 20 | 17.6 | 18.5 | 18.8 | 19.4 | 19 | 18.2 | 17.7848 | 17.3 | 18.2 | 20 | 19.1 | 18.5 | 20.5 | 19 | 19.2 | 15.7 | 15.2 | 20.58 | 20.5 | 19.2 |  |  | 105
| 1 | 6 | 19.4159 | 20.7 | 20.5242 | 18.8715 | 19.8 | 21.339 | 20 | 19 | 18.5 | 21 | 20.5 | 19.6 | 19.2 | 17.8531 | 16.8 | 19 | 20 | 18.9 | 19.2 | 19.3 | 19 | 19.7 | 16.17 | 15.2 | 18.99 | 18.9 | 19.4
| Mean |  | 19.5537 | 20.6167 | 20.541 | 19.0904 | 19.8 | 19.7485 | 20 | 18.1167 | 19 | 20.0833 | 20.0833 | 19.2167 | 18.9 | 18.1482 | 17.1167 | 18.1 | 20 | 19.0167 | 18.8333 | 19.2333 | 18.8333 | 19.8667 | 17.1783 | 15.3167 | 19.61 | 19.8333 | 19.0167
| Median |  | 19.5521 | 20.6 | 20.5254 | 19.1199 | 19.9 | 19.2609 | 20 | 18.1 | 19 | 20.7 | 20.15 | 19 | 18.8 | 18.0603 | 17.1 | 18.1 | 20 | 19.1 | 18.85 | 19.2 | 19 | 19.6 | 17.22 | 15.25 | 19.255 | 20.05 | 19.3
| Std Dev. |  | 0.14677 | 0.18348 | 0.53383 | 0.2838 | 0.60992 | 1.0719 | 0 | 0.69402 | 0.45166 | 1.2057 | 0.40702 | 0.38166 | 0.74565 | 0.41667 | 0.22286 | 0.72388 | 0 | 0.28577 | 0.29439 | 0.70048 | 0.75277 | 0.67725 | 1.477 | 0.27142 | 0.69817 | 0.83347 | 0.75741
| Rel.Std.Dev. |  | 0.7506% | 0.89% | 2.599% | 1.487% | 3.08% | 5.428% | 0% | 3.831% | 2.377% | 6.003% | 2.027% | 1.986% | 3.945% | 2.296% | 1.302% | 3.999% | 0% | 1.503% | 1.563% | 3.642% | 3.997% | 3.409% | 8.598% | 1.772% | 3.56% | 4.202% | 3.983%
| PDM3 |  | 1.563% | 7.084% | 6.691% | -0.8438% | 2.842% | 2.575% | 3.881% | -5.901% | -1.313% | 4.314% | 4.314% | -0.1878% | -1.833% | -5.738% | -11.1% | -5.988% | 3.881% | -1.227% | -2.179% | -0.1012% | -2.179% | 3.188% | -10.77% | -20.44% | 1.855% | 3.015% | -1.227%
| Z-Score (Absolute) |  | 0.37 | 1.53 | 1.45 | 0.14 | 0.64 | 0.58 | 0.86 | 1.2 | 0.24 | 0.95 | 0.95 | 0 | 0.35 | 1.17 | 2.3 | 1.22 | 0.86 | 0.22 | 0.42 | 0.02 | 0.42 | 0.71 | 2.23 | 4.26 | 0.43 | 0.67 | 0.22
| Analytical results for P in OREAS 285 (Certified Value 0.065 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.065969 | 0.0642 | 0.068401 | 0.062059 | 0.064 | 0.0627 | 0.0578 | 0.063 | 0.067 | 0.063 | 0.065 | 0.064 | 0.064026 | 0.069 | 0.0633 | 0.0667 | 0.0679 | 0.064 | 0.06 | 0.0607 | 0.0666 | 0.0676 | 0.07281 | 0.0687 | 0.0701 |  |  |  |  | 1
| 1 | 2 | 0.065549 | 0.0625 | 0.066425 | 0.062588 | 0.063 | 0.0634 | 0.0619 | 0.063 | 0.062 | 0.062 | 0.065 | 0.064 | 0.064194 | 0.066 | 0.0623 | 0.0682 | 0.0703 | 0.062 | 0.06 | 0.0611 | 0.0655 | 0.0663 | 0.07117 | 0.0738 | 0.0627 |  |  |  |  | 17
| 1 | 3 | 0.065115 | 0.0621 | 0.064525 | 0.063132 | 0.063 | 0.0627 | 0.0618 | 0.064 | 0.067 | 0.062 | 0.065 | 0.064 | 0.065691 | 0.066 | 0.0647 | 0.0682 | 0.0699 | 0.064 | 0.06 | 0.0632 | 0.0705 | 0.0672 | 0.07109 | 0.0684 | 0.0676 |  |  |  |  | 16
| 1 | 4 | 0.06467 | 0.0633 | 0.066618 | 0.062926 | 0.067 | 0.0628 | 0.0585 | 0.064 | 0.067 | 0.063 | 0.065 | 0.064 | 0.064327 | 0.066 | 0.0646 | 0.0686 | 0.0676 | 0.063 | 0.06 | 0.0603 | 0.0653 | 0.0668 | 0.07138 | 0.068 | 0.0676 |  |  |  |  | 0.064556
| 1 | 5 | 0.064421 | 0.0648 | 0.067031 | 0.061905 | 0.064 | 0.0623 | 0.0584 | 0.063 | 0.066 | 0.062 | 0.066 | 0.063 | 0.065591 | 0.068 | 0.0649 | 0.0675 | 0.065 | 0.062 | 0.06 | 0.0619 | 0.0663 | 0.0668 | 0.07288 | 0.065 | 0.0687 |  |  |  |  | 106
| 1 | 6 | 0.064807 | 0.0645 | 0.06592 | 0.061602 | 0.065 | 0.0631 | 0.0637 | 0.063 | 0.067 | 0.061 | 0.066 | 0.064 | 0.064342 | 0.064 | 0.065 | 0.0667 | 0.0635 | 0.064 | 0.06 | 0.0614 | 0.0666 | 0.0676 | 0.07063 | 0.063 | 0.0701
| Mean |  | 0.065088 | 0.063567 | 0.066487 | 0.062368 | 0.064333 | 0.062833 | 0.06035 | 0.063333 | 0.066 | 0.062167 | 0.065333 | 0.063833 | 0.064695 | 0.0665 | 0.064133 | 0.06765 | 0.067367 | 0.063167 | 0.06 | 0.061433 | 0.0668 | 0.06705 | 0.07166 | 0.067817 | 0.0678
| Median |  | 0.064961 | 0.06375 | 0.066521 | 0.062323 | 0.064 | 0.06275 | 0.06015 | 0.063 | 0.067 | 0.062 | 0.065 | 0.064 | 0.064334 | 0.066 | 0.06465 | 0.06785 | 0.06775 | 0.0635 | 0.06 | 0.06125 | 0.06645 | 0.067 | 0.071275 | 0.0682 | 0.06815
| Std Dev. |  | 0.00058123 | 0.0011094 | 0.0012767 | 0.00060655 | 0.0015055 | 0.00037771 | 0.0024271 | 0.0005164 | 0.002 | 0.00075277 | 0.0005164 | 0.00040825 | 0.0007421 | 0.0017607 | 0.0010893 | 0.0008167 | 0.00268 | 0.00098319 | 0 | 0.001027 | 0.0018953 | 0.00051284 | 0.00095028 | 0.0036902 | 0.0027379
| Rel.Std.Dev. |  | 0.893% | 1.745% | 1.92% | 0.9725% | 2.34% | 0.6011% | 4.022% | 0.8154% | 3.03% | 1.211% | 0.7904% | 0.6396% | 1.147% | 2.648% | 1.699% | 1.207% | 3.978% | 1.557% | 0% | 1.672% | 2.837% | 0.7649% | 1.326% | 5.441% | 4.038%
| PDM3 |  | 0.825% | -1.532% | 2.991% | -3.388% | -0.3448% | -2.668% | -6.515% | -1.894% | 2.237% | -3.701% | 1.204% | -1.119% | 0.2156% | 3.011% | -0.6546% | 4.793% | 4.354% | -2.152% | -7.057% | -4.837% | 3.476% | 3.863% | 11% | 5.051% | 5.025%
| Z-Score (Absolute) |  | 0.14 | 0.41 | 0.65 | 0.84 | 0.13 | 0.67 | 1.57 | 0.49 | 0.47 | 0.92 | 0.23 | 0.31 | 0 | 0.65 | 0.2 | 1.07 | 0.97 | 0.55 | 1.7 | 1.18 | 0.76 | 0.86 | 2.52 | 1.13 | 1.12
| Analytical results for Pb in OREAS 285 (Certified Value 4.87 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 5.4612 | 5.1 | 6.016 | < 5 | 5.7 |  | 5 | 4 | 4.08 | 4.2 | 5.2 | 4.9 | 4.8 | 4.7 | 4.8463 | 5.9 | 4.3 | 5 | 5.1 | 4.85 | 4.8 | 5 | 4.6 | 4.32 | 52.432 | 8.6 | 4.3 |  |  | 1
| 1 | 2 | 6.1479 | 5.2 | 5.6673 | < 5 | 5.8 | 4.3867 | 5 | 5 | 4.4 | 4.3 | 4.7 | 4.8 | 5.1 | 4.5 | 4.7049 | 5.4 | 4.2 | 5.2 | 5.1 | 4.8 | 4.7 | 5 | 4.5 | 4.23 | 51.52 | 7.3 | 3.6 |  |  | 18
| 1 | 3 | 5.6121 | 5 | 5.668 | < 5 | 5.8 | 6.7678 | 5 | 9 | 4.49 | 4.2 | 5.1 | 5 | 4.8 | 4.5 | 4.7805 | 4.7 | 4.6 | 5.2 | 4.8 | 4.99 | 5 | 5 | 4.8 | 4.3 | 52.176 | 7.2 | 3.9 |  |  | 16
| 1 | 4 | 6.0269 | 5 | 5.6883 | < 5 | 5.8 |  | 5 | 5 | 4.48 | 4.4 | 5.3 | 4.9 | 4.8 | 4.9 | 4.7369 | 4.7 | 4.4 | 5.1 | 5 | 4.85 | 4.9 | 5 | 4.7 | 4.31 | 52.864 | 6.7 | 4 |  |  | 4.8711
| 1 | 5 | 5.5579 | 5.2 | 5.8751 | < 5 | 5.6 | 5.0739 | < 5 | 5 | 4.36 | 4.4 | 4.8 | 4.8 | 4.8 | 4.6 | 4.7256 | 4.9 | 4.5 | 5.1 | 7.9 | 4.47 | 5.1 | 5 | 4.6 | 4.17 | 53.952 | 12.9 | 4.1 |  |  | 107
| 1 | 6 | 5.895 | 5.1 | 5.8191 | < 5 | 5.6 | 8.4972 | 5 | 7 | 4.31 | 4.4 | 5.1 | 4.8 | 4.8 | 4.6 | 4.8679 | 4.7 | 4.4 | 5.1 | 6 | 4.74 | 4.9 | 5 | 4.6 | 4.23 | 52.4 | 7.1 | 4.2
| Mean |  | 5.7835 | 5.1 | 5.789 |  | 5.7167 | 6.1814 | 5 | 5.8333 | 4.3533 | 4.3167 | 5.0333 | 4.8667 | 4.85 | 4.6333 | 4.777 | 5.05 | 4.4 | 5.1167 | 5.65 | 4.7833 | 4.9 | 5 | 4.6333 | 4.26 | 52.5573 | 8.3 | 4.0167
| Median |  | 5.7535 | 5.1 | 5.7537 |  | 5.75 | 5.9208 | 5 | 5 | 4.38 | 4.35 | 5.1 | 4.85 | 4.8 | 4.6 | 4.7587 | 4.8 | 4.4 | 5.1 | 5.1 | 4.825 | 4.9 | 5 | 4.6 | 4.265 | 52.416 | 7.25 | 4.05
| Std Dev. |  | 0.27879 | 0.089443 | 0.14104 |  | 0.098319 | 1.8398 | 0 | 1.8348 | 0.15069 | 0.098319 | 0.23381 | 0.08165 | 0.12247 | 0.15055 | 0.067113 | 0.49699 | 0.14142 | 0.075277 | 1.1777 | 0.17432 | 0.14142 | 0 | 0.10328 | 0.05933 | 0.81254 | 2.3435 | 0.24833
| Rel.Std.Dev. |  | 4.82% | 1.754% | 2.436% |  | 1.72% | 29.76% | 0% | 31.45% | 3.461% | 2.278% | 4.645% | 1.678% | 2.525% | 3.249% | 1.405% | 9.841% | 3.214% | 1.471% | 20.84% | 3.644% | 2.886% | 0% | 2.229% | 1.393% | 1.546% | 28.23% | 6.182%
| PDM3 |  | 18.73% | 4.699% | 18.84% |  | 17.36% | 26.9% | 2.646% | 19.75% | -10.63% | -11.38% | 3.33% | -0.09117% | -0.4333% | -4.881% | -1.932% | 3.673% | -9.671% | 5.041% | 15.99% | -1.802% | 0.5931% | 2.646% | -4.881% | -12.55% | 979% | 70.39% | -17.54%
| Z-Score (Absolute) |  | 1.2 | 0.29 | 1.21 | 3.17 | 1.11 | 1.73 | NA | NA | 0.71 | 0.75 | 0.2 | 0.02 | 0.04 | 0.33 | 0.14 | 0.22 | 0.64 | 0.31 | 1.02 | 0.13 | 0.02 | NA | 0.33 | 0.83 | 63.45 | 4.55 | 1.15
Results from laboratories 1.11, 1.12 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 285 (Indicative Value 20.2 ppb) | [30 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.04 | 1.19 | 1.23 | [26 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | [26 empty] | 1
|  |  | 01g | 01g | 0.5g | [26 empty] | 1
| 1 | 1 | < 10 | < 10 | 22 | [26 empty] | 1
| 1 | 2 | < 10 | < 10 | 22 | [26 empty] | 7
| 1 | 3 | < 10 | < 10 | 20 | [26 empty] | 16
| 1 | 4 | < 10 | < 10 | 18 | [26 empty] | 20.1667
| 1 | 5 | < 10 | < 10 | 20 | [26 empty] | 13
| 1 | 6 | < 10 | < 10 | 19
| Mean |  |  |  | 20.1667
| Median |  |  |  | 20
| Std Dev. |  |  |  | 1.6021
| Rel.Std.Dev. |  |  |  | 7.944%
| PDM3 |  |  |  | -1.554e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Pr in OREAS 285 (Indicative Value 5.41 ppm) | [30 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 2
| 1 | 1 | 5.8883 | 4.105 | 5.5 | 6.0332 | 1.88 | [24 empty] | 1
| 1 | 2 | 5.565 | 4.415 | 5.4 | 6.0457 | 1.83 | [24 empty] | 4
| 1 | 3 | 5.6373 | 4.193 | 5.3 | 6.595 | 1.88 | [24 empty] | 16
| 1 | 4 | 5.6407 | 4.632 | 5.7 | 6.066 | 2 | [24 empty] | 5.4071
| 1 | 5 | 5.5753 | 4.694 | 5.4 | 6.3492 | 1.9 | [24 empty] | 10
| 1 | 6 | 5.6724 | 4.507 | 5.2 | 6.1275 | 1.92
| Mean |  | 5.6632 | 4.4243 | 5.4167 | 6.2028 | 1.9017
| Median |  | 5.639 | 4.461 | 5.4 | 6.0968 | 1.89
| Std Dev. |  | 0.11773 | 0.2359 | 0.17224 | 0.22498 | 0.056716
| Rel.Std.Dev. |  | 2.079% | 5.332% | 3.18% | 3.627% | 2.982%
| PDM3 |  | 4.735% | -18.18% | 0.1765% | 14.71% | -64.83%
| Z-Score (Absolute) |  | 0.21 | 0.85 | 0 | 0.67 | 3.02
| Analytical results for Pt in OREAS 285 (Indicative Value < 5 ppb) | [30 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.04 | 1.19 | 1.23 | 1.27 | [25 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [25 empty] | 2
|  |  | 01g | 01g | 0.5g | 01g | [25 empty] | 2
| 1 | 1 | < 5 | < 5 | < 10 | [26 empty] | 1
| 1 | 2 | < 5 | < 5 | < 10 | [26 empty] | 5
| 1 | 3 | < 5 | < 5 | < 10 | [26 empty] | 16
| 1 | 4 | < 5 | < 5 | < 10 | 8 | [25 empty] | < 5
| 1 | 5 | < 5 | < 5 | < 10 | [26 empty] | 11
| 1 | 6 | < 5 | < 5 | < 10
| Mean |  |  |  |  | 8
| Median |  |  |  |  | 8
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.67 | 2.29
| Analytical results for Rb in OREAS 285 (Certified Value 105 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 0
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [11 empty] | 0
| 1 | 1 | 103.621 | 106.46 | 88.8 | 100 | 113.5 | 108 | 113 | 99.8 | 95.7948 | 101 | 101.7 | 112 | 101 | 104 | 110.72 | 106 | 120 | 109.3 | [11 empty] | 1
| 1 | 2 | 105.867 | 108.68 | 90.1 | 105 | 108.5 | 105.5 | 118 | 97.4 | 95.2715 | 99.1 | 105.4 | 114 | 103 | 106 | 108.8 | 104 | 122 | 96.8 | [11 empty] | 34
| 1 | 3 | 103.0373 | 108.77 | 93.3 | 103 | 116 | 107.5 | 112 | 95.2 | 96.1264 | 98.2 | 104 | 113 | 97.8 | 104 | 112.73 | 106 | 121 | 106.5 | [11 empty] | 16
| 1 | 4 | 102.3665 | 109.83 | 94.5 | 102 | 115.5 | 105.5 | 112.5 | 102.5 | 92.3566 | 98.9 | 104.6 | 116 | 99.4 | 102 | 108.79 | 109 | 123 | 109.1 | [11 empty] | 105.4634
| 1 | 5 | 105.9228 | 111.29 | 93 | 99.6 | 105 | 103.5 | 110 | 92.6 | 94.1034 | 100.9 | 102.7 | 113 | 98.8 | 110 | 109.61 | 105 | 120 | 110.5 | [11 empty] | 108
| 1 | 6 | 102.3137 | 110.55 | 91.2 | 99.4 | 114.5 | 108 | 114.5 | 100.5 | 92.6945 | 98.3 | 103.9 | 114 | 99.3 | 105 | 111.12 | 104 | 118 | 110.7
| Mean |  | 103.8547 | 109.2633 | 91.8167 | 101.5 | 112.1667 | 106.3333 | 113.3333 | 98 | 94.3912 | 99.4 | 103.7167 | 113.6667 | 99.8833 | 105.1667 | 110.295 | 105.6667 | 120.6667 | 107.15
| Median |  | 103.3292 | 109.3 | 92.1 | 101 | 114 | 106.5 | 112.75 | 98.6 | 94.6875 | 99 | 103.95 | 113.5 | 99.35 | 104.5 | 110.165 | 105.5 | 120.5 | 109.2
| Std Dev. |  | 1.6516 | 1.7045 | 2.1554 | 2.237 | 4.4234 | 1.8074 | 2.7142 | 3.6633 | 1.6036 | 1.249 | 1.3288 | 1.3663 | 1.8465 | 2.7142 | 1.5336 | 1.8619 | 1.7512 | 5.2876
| Rel.Std.Dev. |  | 1.59% | 1.56% | 2.347% | 2.204% | 3.944% | 1.7% | 2.395% | 3.738% | 1.699% | 1.257% | 1.281% | 1.202% | 1.849% | 2.581% | 1.39% | 1.762% | 1.451% | 4.935%
| PDM3 |  | -1.525% | 3.603% | -12.94% | -3.758% | 6.356% | 0.8249% | 7.462% | -7.077% | -10.5% | -5.749% | -1.656% | 7.778% | -5.291% | -0.2813% | 4.581% | 0.1928% | 14.42% | 1.599%
| Z-Score (Absolute) |  | 0.2 | 0.5 | 1.76 | 0.51 | 0.87 | 0.12 | 1.03 | 0.96 | 1.43 | 0.78 | 0.22 | 1.07 | 0.72 | 0.03 | 0.63 | 0.03 | 1.98 | 0.22
| Analytical results for Re in OREAS 285 (Indicative Value 0.001 ppm) | [30 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.27 | 1.31 | 1.32 | [13 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [13 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 01g | 0.5g | 0.25g | [13 empty] | 3
| 1 | 1 | < 0.01 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.005 | < 0.002 | < 0.05 | 0.002 | 0.001 | 0.002 | < 0.1 | [13 empty] | 1
| 1 | 2 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.005 | 0.002 | < 0.05 | 0.002 |  | 0.001 | < 0.1 | [13 empty] | 2
| 1 | 3 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.005 | 0.002 | < 0.05 | 0.001 |  | 0.001 | < 0.1 | [13 empty] | 16
| 1 | 4 | < 0.01 | < 0.001 | 0.002 | 0.001 | 0.001 | 0.002 | 0.003 | 0.001 | < 0.05 | < 0.005 | 0.002 | < 0.05 | 0.001 | 0.001 | 0.001 | < 0.1 | [13 empty] | 0.00117
| 1 | 5 | < 0.01 | < 0.001 | 0.001 | 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.005 | < 0.002 | < 0.05 | 0.001 |  | 0.001 | < 0.1 | [13 empty] | 12
| 1 | 6 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | 0.002 | 0.001 | < 0.05 | < 0.005 | 0.002 | < 0.05 | 0.002 |  | 0.001 | < 0.1
| Mean |  |  | 0.0015 | 0.0011667 | 0.001 | 0.001 | 0.0013333 | 0.0015 | 0.001 |  |  | 0.002 |  | 0.0015 | 0.001 | 0.0011667
| Median |  |  | 0.0015 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 |  |  | 0.002 |  | 0.0015 | 0.001 | 0.001
| Std Dev. |  |  | 0.00070711 | 0.00040825 | 0 | 0 | 0.0005164 | 0.00083666 | 0 |  |  | 0 |  | 0.00054772 | 0 | 0.00040825
| Rel.Std.Dev. |  |  | 47.14% | 34.99% | 0% | 0% | 38.73% | 55.78% | 0% |  |  | 0% |  | 36.51% | 0% | 34.99%
| PDM3 |  |  | 28.21% | -0.2849% | -14.53% | -14.53% | 13.96% | 28.21% | -14.53% |  |  | 70.94% |  | 28.21% | -14.53% | -0.2849%
| Z-Score (Absolute) |  | 5.8 | 0.94 | 0.4 | 0.67 | 0.81 | 0.13 | 0.13 | 0.67 | 38.17 | 1.75 | 0.4 | 38.17 | 0.13 | 0.67 | 0.4 | 78.62
| Analytical results for S in OREAS 285 (Certified Value 0.334 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.25g | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.31477 | 0.3598 | 0.3352 | 0.32798 | 0.33484 | 0.322 | 0.2985 | 0.38 | 0.3015 | 0.32 | 0.33 | 0.38 | 0.34 | 0.35 | 0.38 | 0.32863 | 0.33 | 0.33 | 0.33 | 0.32 | 0.34 | 0.34 | 0.3197 | 0.31 | 0.348 | 0.344 | 0.36 |  |  | 1
| 1 | 2 | 0.31449 | 0.36044 | 0.3264 | 0.34668 | 0.33624 | 0.32 | 0.3054 | 0.38 | 0.3076 | 0.34 | 0.33 | 0.37 | 0.34 | 0.36 | 0.37 | 0.32712 | 0.33 | 0.33 | 0.34 | 0.33 | 0.32 | 0.34 | 0.3274 | 0.31 | 0.34 | 0.358 | 0.33 |  |  | 11
| 1 | 3 | 0.31181 | 0.35499 | 0.326 | 0.34729 | 0.33714 | 0.319 | 0.2855 | 0.39 | 0.3139 | 0.34 | 0.34 | 0.38 | 0.33 | 0.36 | 0.38 | 0.33152 | 0.32 | 0.34 | 0.34 | 0.32 | 0.32 | 0.33 | 0.3356 | 0.31 | 0.348 | 0.355 | 0.35 |  |  | 16
| 1 | 4 | 0.31733 | 0.35878 | 0.3315 | 0.33242 | 0.33294 | 0.334 | 0.2833 | 0.38 | 0.3099 | 0.32 | 0.34 | 0.38 | 0.34 | 0.36 | 0.38 | 0.32734 | 0.33 | 0.33 | 0.34 | 0.32 | 0.34 | 0.34 | 0.3179 | 0.3 | 0.345 | 0.35 | 0.35 |  |  | 0.33376
| 1 | 5 | 0.31278 | 0.35292 | 0.3397 | 0.34395 | 0.33397 | 0.319 | 0.2903 | 0.38 | 0.3049 | 0.32 | 0.33 | 0.38 | 0.34 | 0.36 | 0.37 | 0.33022 | 0.33 | 0.34 | 0.33 | 0.34 | 0.32 | 0.34 | 0.325 | 0.32 | 0.344 | 0.36 | 0.36 |  |  | 109
| 1 | 6 | 0.31124 | 0.35755 | 0.3394 | 0.34728 | 0.33454 | 0.323 | 0.3165 | 0.39 | 0.3048 | 0.35 | 0.33 | 0.38 | 0.34 | 0.36 | 0.38 | 0.32496 | 0.33 | 0.34 | 0.34 | 0.34 | 0.33 | 0.34 | 0.3223 | 0.31 | 0.343 | 0.35 | 0.36
| Mean |  | 0.31374 | 0.35741 | 0.33303 | 0.34093 | 0.33494 | 0.32283 | 0.29658 | 0.38333 | 0.3071 | 0.33167 | 0.33333 | 0.37833 | 0.33833 | 0.35833 | 0.37667 | 0.3283 | 0.32833 | 0.335 | 0.33667 | 0.32833 | 0.32833 | 0.33833 | 0.32465 | 0.31 | 0.34467 | 0.35283 | 0.35167
| Median |  | 0.31363 | 0.35817 | 0.33335 | 0.34531 | 0.33469 | 0.321 | 0.2944 | 0.38 | 0.30625 | 0.33 | 0.33 | 0.38 | 0.34 | 0.36 | 0.38 | 0.32799 | 0.33 | 0.335 | 0.34 | 0.325 | 0.325 | 0.34 | 0.32365 | 0.31 | 0.3445 | 0.3525 | 0.355
| Std Dev. |  | 0.0022542 | 0.0029289 | 0.0060909 | 0.008522 | 0.0015265 | 0.0057067 | 0.012782 | 0.005164 | 0.0043777 | 0.013292 | 0.005164 | 0.0040825 | 0.0040825 | 0.0040825 | 0.005164 | 0.0023485 | 0.0040825 | 0.0054772 | 0.005164 | 0.0098319 | 0.0098319 | 0.0040825 | 0.0063746 | 0.0063246 | 0.0030768 | 0.005947 | 0.01169
| Rel.Std.Dev. |  | 0.7185% | 0.8195% | 1.829% | 2.5% | 0.4557% | 1.768% | 4.31% | 1.347% | 1.425% | 4.008% | 1.549% | 1.079% | 1.207% | 1.139% | 1.371% | 0.7154% | 1.243% | 1.635% | 1.534% | 2.994% | 2.994% | 1.207% | 1.964% | 2.04% | 0.8927% | 1.685% | 3.324%
| PDM3 |  | -6% | 7.086% | -0.2185% | 2.148% | 0.3542% | -3.275% | -11.14% | 14.85% | -7.989% | -0.628% | -0.1287% | 13.35% | 1.369% | 7.362% | 12.85% | -1.637% | -1.627% | 0.3707% | 0.8701% | -1.627% | -1.627% | 1.369% | -2.73% | -7.12% | 3.267% | 5.714% | 5.364%
| Z-Score (Absolute) |  | 1.47 | 1.56 | 0.13 | 0.42 | 0 | 0.84 | 2.66 | 3.36 | 1.93 | 0.23 | 0.11 | 3.01 | 0.24 | 1.62 | 2.89 | 0.46 | 0.46 | 0 | 0.12 | 0.46 | 0.46 | 0.24 | 0.71 | 1.73 | 0.67 | 1.24 | 1.16
| Analytical results for Sb in OREAS 285 (Certified Value 0.27 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 0.33664 | 0.32 | 2.3273 | < 5 | 0.31 | < 5 | 0.31 | 0.32 | 0.24 | 0.31 | 0.31 | 0.31 | 0.28 | 0.28251 | 0.23 | 0.19 | 0.19 | 0.32 | 0.28 | 0.27 | < 2 | 0.21 | 0.32 | 0.531 | 0.36 | 0.18 |  |  |  | 1
| 1 | 2 | 0.31483 | 0.34 | 1.9596 | < 5 | 0.31 | < 5 | 0.31 | 0.3 | 0.25 | 0.3 | 0.3 | 0.31 | 0.27 | 0.28591 | 0.23 | 0.18 | 0.19 | 0.33 | 0.24 | 0.28 | < 2 | 0.21 | 0.3 | 0.45 | 0.25 | 0.13 |  |  |  | 36
| 1 | 3 | 0.31895 | 0.31 | 0.035485 | < 5 | 0.33 | < 5 | 0.32 | 0.32 | 0.25 | 0.28 | 0.31 | 0.28 | 0.29 | 0.2583 | 0.25 | 0.18 | 0.2 | 0.32 | 0.26 | 0.27 | < 2 | 0.21 | 0.32 | 0.405 | 0.28 | 0.16 |  |  |  | 16
| 1 | 4 | 0.3025 | 0.34 |  | < 5 | 0.31 | < 5 | 0.29 | 0.29 | 0.24 | 0.29 | 0.29 | 0.29 | 0.28 | 0.25778 | 0.22 | 0.19 | 0.18 | 0.33 | 0.26 | 0.27 | < 2 | 0.2 | 0.3 | 0.405 | 0.26 | 0.17 |  |  |  | 0.27087
| 1 | 5 | 0.33878 | 0.33 | 4.0204 | < 5 | 0.3 | < 5 | 0.3 | 0.33 | 0.25 | 0.25 | 0.34 | 0.3 | 0.26 | 0.28139 | 0.25 | 0.2 | 0.18 | 0.35 | 0.26 | 0.32 | < 2 | 0.21 | 0.31 | 0.693 | 0.25 | 0.15 |  |  |  | 110
| 1 | 6 | 0.32667 | 0.33 | 4.2797 | < 5 | 0.32 | < 5 | 0.3 | 0.44 | 0.24 | 0.29 | 0.32 | 0.33 | 0.29 | 0.29801 | 0.22 | 0.2 | 0.2 | 0.32 | 0.24 | 0.28 | < 2 | 0.2 | 0.3 | 0.477 | 0.22 | 0.2
| Mean |  | 0.32306 | 0.32833 | 2.5245 |  | 0.31333 |  | 0.305 | 0.33333 | 0.245 | 0.28667 | 0.31167 | 0.30333 | 0.27833 | 0.27731 | 0.23333 | 0.19 | 0.19 | 0.32833 | 0.25667 | 0.28167 |  | 0.20667 | 0.30833 | 0.4935 | 0.27 | 0.165
| Median |  | 0.32281 | 0.33 | 2.3273 |  | 0.31 |  | 0.305 | 0.32 | 0.245 | 0.29 | 0.31 | 0.305 | 0.28 | 0.28195 | 0.23 | 0.19 | 0.19 | 0.325 | 0.26 | 0.275 |  | 0.21 | 0.305 | 0.4635 | 0.255 | 0.165
| Std Dev. |  | 0.013797 | 0.01169 | 1.7228 |  | 0.010328 |  | 0.010488 | 0.054283 | 0.0054772 | 0.020656 | 0.017224 | 0.017512 | 0.01169 | 0.016056 | 0.013663 | 0.0089443 | 0.0089443 | 0.01169 | 0.015055 | 0.019408 |  | 0.005164 | 0.0098319 | 0.10866 | 0.048166 | 0.02429
| Rel.Std.Dev. |  | 4.271% | 3.561% | 68.24% |  | 3.296% |  | 3.439% | 16.28% | 2.236% | 7.206% | 5.526% | 5.773% | 4.2% | 5.79% | 5.855% | 4.708% | 4.708% | 3.561% | 5.866% | 6.89% |  | 2.499% | 3.189% | 22.02% | 17.84% | 14.72%
| PDM3 |  | 19.27% | 21.21% | 832% |  | 15.67% |  | 12.6% | 23.06% | -9.552% | 5.83% | 15.06% | 11.98% | 2.753% | 2.377% | -13.86% | -29.86% | -29.86% | 21.21% | -5.245% | 3.984% |  | -23.7% | 13.83% | 82.19% | -0.323% | -39.09%
| Z-Score (Absolute) |  | 0.4 | 0.51 | 47.28 | 46.76 | 0.2 | 46.76 | 0.02 | 0.62 | 1.26 | 0.37 | 0.16 | 0.02 | 0.55 | 0.57 | 1.51 | 2.43 | 2.43 | 0.51 | 1.01 | 0.48 | 14.82 | 2.08 | 0.09 | 4.03 | 0.73 | 2.96
| Analytical results for Sc in OREAS 285 (Certified Value 8.11 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 8.3893 | 8 | 7.8 | 6.8258 | 8 | 7.5 | 9.2 | 9.3 | 8.3 | 8.7 | 7.9 | 8.5068 | 8 | 7.9 | 8 | 8.2 | 7.9 | 8 | 8 | 8.3 | 5.1 | 8.28 | 9.6 | 8.1 | [5 empty] | 1
| 1 | 2 | 8.3724 | 7.9 | 7.9 | 7.2953 | 8 | 8.1 | 9.7 | 9.3 | 8.3 | 9.4 | 7.9 | 8.0767 | 7.8 | 7.7 | 8 | 8.5 | 8.3 | 8.1 | 8 | 8.2 | 5.1 | 8.15 | 11.2 | 7.1 | [5 empty] | 37
| 1 | 3 | 8.361 | 7.9 | 7.9 | 7.1712 | 8 | 8 | 9.8 | 8.9 | 8.2 | 8.8 | 7.7 | 8.4823 | 7.7 | 8.3 | 8 | 8.2 | 7.9 | 8.1 | 8 | 8.6 | 5.2 | 8.22 | 10.7 | 7.7 | [5 empty] | 16
| 1 | 4 | 8.318 | 8 | 8.4 | 6.7765 | 8 | 7.5 | 9.8 | 9.4 | 8.2 | 8.8 | 8.4 | 8.0623 | 7.7 | 8.3 | 8 | 8.1 | 8.3 | 7.9 | 8 | 8.3 | 5.4 | 8.25 | 10.5 | 8.2 | [5 empty] | 8.1112
| 1 | 5 | 8.4096 | 8.3 | 8.2 | 6.7101 | 8 | 7.6 | 9.3 | 8.6 | 8.2 | 8.8 | 7.7 | 8.1616 | 7.8 | 8.4 | 8 | 7.9 | 8 | 8.7 | 8 | 8.3 | 5.1 | 8.36 | 10.4 | 8.2 | [5 empty] | 111
| 1 | 6 | 8.3947 | 8.3 | 7.7 | 7.0438 | 8 | 8.2 | 9.3 | 9.2 | 8.4 | 9.3 | 8.2 | 8.1511 | 7.6 | 8.3 | 8 | 7.8 | 7.9 | 8.2 | 8 | 8.5 | 5.2 | 8.26 | 9.8 | 8.3
| Mean |  | 8.3742 | 8.0667 | 7.9833 | 6.9704 | 8 | 7.8167 | 9.5167 | 9.1167 | 8.2667 | 8.9667 | 7.9667 | 8.2401 | 7.7667 | 8.15 | 8 | 8.1167 | 8.05 | 8.1667 | 8 | 8.3667 | 5.1833 | 8.2533 | 10.3667 | 7.9333
| Median |  | 8.3808 | 8 | 7.9 | 6.9348 | 8 | 7.8 | 9.5 | 9.25 | 8.25 | 8.8 | 7.9 | 8.1563 | 7.75 | 8.3 | 8 | 8.15 | 7.95 | 8.1 | 8 | 8.3 | 5.15 | 8.255 | 10.45 | 8.15
| Std Dev. |  | 0.032349 | 0.18619 | 0.26394 | 0.2356 | 0 | 0.31885 | 0.27869 | 0.30605 | 0.08165 | 0.30111 | 0.28048 | 0.2011 | 0.13663 | 0.28107 | 0 | 0.24833 | 0.19748 | 0.28048 | 0 | 0.15055 | 0.1169 | 0.069186 | 0.58878 | 0.45898
| Rel.Std.Dev. |  | 0.3863% | 2.308% | 3.306% | 3.38% | 0% | 4.079% | 2.928% | 3.357% | 0.9877% | 3.358% | 3.521% | 2.441% | 1.759% | 3.449% | 0% | 3.059% | 2.453% | 3.434% | 0% | 1.799% | 2.255% | 0.8383% | 5.68% | 5.786%
| PDM3 |  | 3.242% | -0.5488% | -1.576% | -14.06% | -1.371% | -3.631% | 17.33% | 12.4% | 1.917% | 10.55% | -1.782% | 1.59% | -4.247% | 0.4786% | -1.371% | 0.06768% | -0.7542% | 0.6841% | -1.371% | 3.15% | -36.1% | 1.753% | 27.81% | -2.193%
| Z-Score (Absolute) |  | 0.7 | 0.26 | 0.52 | 3.67 | NA | 1.04 | 4.25 | 3.01 | 0.36 | 2.54 | 0.57 | 0.28 | 1.19 | 0 | NA | 0.1 | 0.31 | 0.05 | NA | 0.67 | 9.23 | 0.32 | 6.9 | 0.67
Results from laboratories 1.11, 1.22 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 285 (Certified Value 0.48 ppm) | [30 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.29 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | < 1.5 | < 1 | 0.79398 | < 5 | 0.7 | < 5 | 0.37 | 0.8 | 0.6 | 0.4 | 0.8 | 0.5 | 0.5 | < 1 | 3.3 | < 1 | < 1 | 0.4 | 0.4 | < 1 | < 0.1 | 0.9 | < 1 | [6 empty] | 1
| 1 | 2 | < 1.5 | < 1 | 0.89064 | < 5 | 0.7 | < 5 | 0.36 | 0.5 | 0.4 | 0.4 | 0.4 | 0.6 | 0.5 | < 1 | 3.4 | < 1 | < 1 | 0.4 | 0.4 | < 1 | < 0.1 | 0.7 | < 1 | [6 empty] | 38
| 1 | 3 | < 1.5 | < 1 | 0.81918 | < 5 | 0.8 | < 5 | 0.38 | 0.2 | 0.4 | 0.4 | 0.5 | 0.3 | 0.3 | < 1 | 3.2 | < 1 | < 1 | 0.4 | 0.4 | < 1 | < 0.1 | 0.8 | < 1 | [6 empty] | 16
| 1 | 4 | < 1.5 | < 1 | 0.87804 | < 5 | 0.6 | < 5 | 0.37 | 0.6 | 0.5 | 0.4 | 0.4 | 0.5 | 0.7 | < 1 | 2.8 | < 1 | < 1 | 0.4 | 0.5 | < 1 | < 0.1 | 1 | < 1 | [6 empty] | 0.47633
| 1 | 5 | < 1.5 | < 1 | 1.0499 | < 5 | 0.7 | < 5 | 0.38 | 0.7 | 0.4 | 0.6 | < 0.2 | 0.7 | 0.6 | < 1 | 3.1 | < 1 | < 1 | 0.4 | 0.5 | 1 | < 0.1 | 0.9 | < 1 | [6 empty] | 112
| 1 | 6 | < 1.5 | < 1 | 0.73713 | < 5 | 0.6 | < 5 | 0.37 | 0.5 | 0.5 | 0.5 | 0.4 | 0.3 | 0.4 | < 1 | 3.2 | < 1 | < 1 | 0.4 | 0.4 | < 1 | < 0.1 | 0.6 | < 1
| Mean |  |  |  | 0.86148 |  | 0.68333 |  | 0.37167 | 0.55 | 0.46667 | 0.45 | 0.5 | 0.48333 | 0.5 |  | 3.1667 |  |  | 0.4 | 0.43333 | 1 |  | 0.81667
| Median |  |  |  | 0.84861 |  | 0.7 |  | 0.37 | 0.55 | 0.45 | 0.4 | 0.4 | 0.5 | 0.5 |  | 3.2 |  |  | 0.4 | 0.4 | 1 |  | 0.85
| Std Dev. |  |  |  | 0.10811 |  | 0.075277 |  | 0.0075277 | 0.20736 | 0.08165 | 0.083666 | 0.17321 | 0.16021 | 0.14142 |  | 0.20656 |  |  | 6.0809e-17 | 0.05164 |  |  | 0.1472
| Rel.Std.Dev. |  |  |  | 12.55% |  | 11.02% |  | 2.025% | 37.7% | 17.5% | 18.59% | 34.64% | 33.15% | 28.28% |  | 6.523% |  |  | 1.52e-14% | 11.92% |  |  | 18.02%
| PDM3 |  |  |  | 80.86% |  | 43.46% |  | -21.97% | 15.47% | -2.029% | -5.528% | 4.969% | 1.47% | 4.969% |  | 564.8% |  |  | -16.03% | -9.027% | 109.9% |  | 71.45%
| Z-Score (Absolute) |  | 2.53 | 0 | 3.66 | 20.23 | 1.85 | 20.23 | 1.3 | 0.51 | 0.34 | 0.51 | 0.67 | 0.17 | 0 | 0 | 26.97 | 0 | 0 | 1.01 | 0.67 | 0.84 | 4.55 | 3.2 | 0
| Analytical results for Si in OREAS 285 (Indicative Value 0.133 wt.%) | [30 empty] | Indicative
Si | Round | Replicate | Lab | [28 empty] | 1
| No | No | 1.30 | [28 empty] | wt.%
|  |  | AR*OES | [28 empty] | 3
|  |  | 0.2g | [28 empty] | 3
| 1 | 1 | 0.1151 | [28 empty] | 1
| 1 | 2 | 0.1623 | [28 empty] | 7
| 1 | 3 | 0.1177 | [28 empty] | 16
| 1 | 4 | 0.1148 | [28 empty] | 0.13298
| 1 | 5 | 0.1438 | [28 empty] | 13
| 1 | 6 | 0.1442
| Mean |  | 0.13298
| Median |  | 0.13075
| Std Dev. |  | 0.019931
| Rel.Std.Dev. |  | 14.99%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 285 (Indicative Value 4.74 ppm) | [30 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.29 | [24 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [24 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | [24 empty] | 2
| 1 | 1 | 4.8902 | 3.783 | 4.6 | 5.6305 | 2.45 | [24 empty] | 1
| 1 | 2 | 4.8095 | 3.989 | 4.4 | 5.781 | 2.37 | [24 empty] | 4
| 1 | 3 | 4.9448 | 3.845 | 4.1 | 5.9813 | 2.46 | [24 empty] | 16
| 1 | 4 | 4.5587 | 4.076 | 4.8 | 5.5875 | 2.58 | [24 empty] | 4.7353
| 1 | 5 | 4.7986 | 4.239 | 4.3 | 5.7781 | 2.48 | [24 empty] | 10
| 1 | 6 | 4.8356 | 4.092 | 3.9 | 5.6302 | 2.46
| Mean |  | 4.8062 | 4.004 | 4.35 | 5.7314 | 2.4667
| Median |  | 4.8226 | 4.0325 | 4.35 | 5.7043 | 2.46
| Std Dev. |  | 0.13304 | 0.16882 | 0.32711 | 0.14712 | 0.067429
| Rel.Std.Dev. |  | 2.768% | 4.216% | 7.52% | 2.567% | 2.734%
| PDM3 |  | 1.498% | -15.44% | -8.137% | 21.04% | -47.91%
| Z-Score (Absolute) |  | 0.67 | 0.51 | 0 | 2.04 | 2.78
| Analytical results for Sn in OREAS 285 (Certified Value 3.47 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 3.5204 | 3.38 | 3.9951 | < 5 | 3.17 | < 10 | 3 | 3.6 | 3.6 | 3.5 | 3.2 | 3.3926 | 3.4 | 2.9 | 3.5 | 3.5 | 3.4 | 3.61 | 2.82 | < 0.1 | 2 | 3.9 | [7 empty] | 1
| 1 | 2 | 3.5446 | 3.58 | 3.4676 | < 5 | 3.2 | < 10 | 3.1 | 3.4 | 3.6 | 3.6 | 3.2 | 3.4962 | 3.4 | 2.7 | 3.5 | 3.5 | 3.4 | 3.62 | 2.77 | < 0.1 | 2.2 | 3.3 | [7 empty] | 13
| 1 | 3 | 3.6284 | 3.39 | 3.8438 | < 5 | 3.25 | < 10 | 3.1 | 3.6 | 3.5 | 3.4 | 3.1 | 3.4497 | 3.3 | 2.9 | 3.4 | 3.4 | 3.4 | 3.66 | 2.88 | < 0.1 | 2.2 | 3.7 | [7 empty] | 16
| 1 | 4 | 3.4089 | 3.51 | 3.796 | < 5 | 3.29 | < 10 | 3.1 | 3.6 | 3.4 | 3.4 | 3.3 | 3.5177 | 3.4 | 2.9 | 3.6 | 3.4 | 3.3 | 3.55 | 2.89 | < 0.1 | 2.1 | 3.8 | [7 empty] | 3.4693
| 1 | 5 | 3.6147 | 3.51 | 4.0163 | < 5 | 3.23 | < 10 | 3 | 3.2 | 3.5 | 3.4 | 3.1 | 3.396 | 3.4 | 3 | 3.3 | 3.5 | 3.6 | 3.47 | 2.83 | < 0.1 | 2.1 | 3.8 | [7 empty] | 113
| 1 | 6 | 3.5712 | 3.47 | 3.8818 | < 5 | 3.17 | < 10 | 3 | 3.6 | 3.6 | 3.5 | 3.2 | 3.4019 | 3.3 | 2.9 | 3.6 | 3.5 | 3.4 | 3.62 | 2.79 | < 0.1 | 1.9 | 3.8
| Mean |  | 3.5481 | 3.4733 | 3.8334 |  | 3.2183 |  | 3.05 | 3.5 | 3.5333 | 3.4667 | 3.1833 | 3.4423 | 3.3667 | 2.8833 | 3.4833 | 3.4667 | 3.4167 | 3.5883 | 2.83 |  | 2.0833 | 3.7167
| Median |  | 3.5579 | 3.49 | 3.8628 |  | 3.215 |  | 3.05 | 3.6 | 3.55 | 3.45 | 3.2 | 3.4258 | 3.4 | 2.9 | 3.5 | 3.5 | 3.4 | 3.615 | 2.825 |  | 2.1 | 3.8
| Std Dev. |  | 0.07947 | 0.077115 | 0.19861 |  | 0.047504 |  | 0.054772 | 0.16733 | 0.08165 | 0.08165 | 0.075277 | 0.054566 | 0.05164 | 0.098319 | 0.1169 | 0.05164 | 0.098319 | 0.067946 | 0.047749 |  | 0.1169 | 0.2137
| Rel.Std.Dev. |  | 2.24% | 2.22% | 5.181% |  | 1.476% |  | 1.796% | 4.781% | 2.311% | 2.355% | 2.365% | 1.585% | 1.534% | 3.41% | 3.356% | 1.49% | 2.878% | 1.894% | 1.687% |  | 5.611% | 5.75%
| PDM3 |  | 2.269% | 0.1151% | 10.49% |  | -7.235% |  | -12.09% | 0.8838% | 1.845% | -0.07702% | -8.244% | -0.778% | -2.959% | -16.89% | 0.4034% | -0.07702% | -1.518% | 3.43% | -18.43% |  | -39.95% | 7.129%
| Z-Score (Absolute) |  | 0.34 | 0.07 | 1.38 | 3.48 | 0.86 | 5.63 | 1.47 | 0.17 | 0.29 | 0.04 | 0.99 | 0.04 | 0.32 | 2.08 | 0.11 | 0.04 | 0.14 | 0.49 | 2.28 | 12.41 | 5 | 0.96
| Analytical results for Sr in OREAS 285 (Certified Value 19.6 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 1
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 21.1441 | 18.75 | 28.4669 | 23.3675 | 28.8 | 5.6369 | < 5 | 18.4 | 20.3 | 24.1 | 22.3 | 20.8 | 20.8 | 20.0627 | 21 | 16.9 | 19 | 18.3 | 19.7 | 18.2 | 17 | 16.96 | 12.11 | 16 | 29.65 | 16.3 | 21.9 |  |  | 1
| 1 | 2 | 21.9234 | 19.1 | 29.9012 | 22.9009 | 29.4 | 9.3383 | < 5 | 20.2 | 21.5 | 20.1 | 21.9 | 21.3 | 20.8 | 19.345 | 20 | 17.6 | 19.3 | 18.3 | 19.7 | 18.4 | 17 | 16.68 | 12.91 | 15 | 29.08 | 20.1 | 19.2 |  |  | 14
| 1 | 3 | 21.4745 | 18.93 | 30.5169 | 23.1956 | 29.6 | 4.8457 | < 5 | 19.5 | 20.8 | 23.8 | 22 | 20.8 | 20 | 19.1939 | 20 | 18.3 | 19 | 17.4 | 19.4 | 18.2 | 18 | 17.57 | 12.73 | 15 | 29.54 | 19.1 | 20.8 |  |  | 16
| 1 | 4 | 21.5228 | 19.98 | 27.6939 | 23.6007 | 30.3 | 5.0335 | < 5 | 17.8 | 20.9 | 23.7 | 21.5 | 20.9 | 22.3 | 20.1616 | 20 | 18.3 | 19.3 | 17.7 | 19.2 | 18.4 | 17 | 16.72 | 10.86 | 16 | 29.37 | 17.5 | 21 |  |  | 19.6321
| 1 | 5 | 22.2123 | 20.11 | 30.0347 | 23.4978 | 28.1 | 5.8886 | < 5 | 18.4 | 19.6 | 21.7 | 21.7 | 20.9 | 20.1 | 19.3534 | 20 | 17.9 | 19.2 | 17.4 | 19 | 18.2 | 17 | 17.21 | 11.06 | 15 | 29.96 | 17.9 | 21.4 |  |  | 114
| 1 | 6 | 21.2816 | 19.91 | 30.9934 | 22.7032 | 27.7 | 6.0292 | < 5 | 20.4 | 19.8 | 23.3 | 22.5 | 21.5 | 21.4 | 19.4653 | 20 | 18.1 | 19.6 | 16.9 | 19.7 | 18.3 | 17 | 17.36 | 11.35 | 16 | 29.1 | 16.1 | 21.6
| Mean |  | 21.5931 | 19.4633 | 29.6012 | 23.211 | 28.9833 | 6.1287 |  | 19.1167 | 20.4833 | 22.7833 | 21.9833 | 21.0333 | 20.9 | 19.597 | 20.1667 | 17.85 | 19.2333 | 17.6667 | 19.45 | 18.2833 | 17.1667 | 17.0833 | 11.8367 | 15.5 | 29.45 | 17.8333 | 20.9833
| Median |  | 21.4987 | 19.505 | 29.968 | 23.2815 | 29.1 | 5.7628 |  | 18.95 | 20.55 | 23.5 | 21.95 | 20.9 | 20.8 | 19.4094 | 20 | 18 | 19.25 | 17.55 | 19.55 | 18.25 | 17 | 17.085 | 11.73 | 15.5 | 29.455 | 17.7 | 21.2
| Std Dev. |  | 0.40255 | 0.60165 | 1.2632 | 0.35011 | 0.97451 | 1.6405 |  | 1.0704 | 0.71949 | 1.5651 | 0.37103 | 0.29439 | 0.8579 | 0.40947 | 0.40825 | 0.53572 | 0.22509 | 0.55377 | 0.30166 | 0.098319 | 0.40825 | 0.35758 | 0.87393 | 0.54772 | 0.33882 | 1.5629 | 0.96003
| Rel.Std.Dev. |  | 1.864% | 3.091% | 4.268% | 1.508% | 3.362% | 26.77% |  | 5.599% | 3.513% | 6.87% | 1.688% | 1.4% | 4.105% | 2.089% | 2.024% | 3.001% | 1.17% | 3.135% | 1.551% | 0.5378% | 2.378% | 2.093% | 7.383% | 3.534% | 1.15% | 8.764% | 4.575%
| PDM3 |  | 9.989% | -0.8597% | 50.78% | 18.23% | 47.63% | -68.78% |  | -2.626% | 4.336% | 16.05% | 11.98% | 7.137% | 6.458% | -0.1789% | 2.723% | -9.078% | -2.031% | -10.01% | -0.9276% | -6.87% | -12.56% | -12.98% | -39.71% | -21.05% | 50.01% | -9.162% | 6.883%
| Z-Score (Absolute) |  | 0.8 | 0 | 3.8 | 1.41 | 3.57 | 5 | 6.37 | 0.13 | 0.38 | 1.25 | 0.95 | 0.59 | 0.54 | 0.05 | 0.26 | 0.61 | 0.09 | 0.67 | 0.01 | 0.44 | 0.86 | 0.89 | 2.86 | 1.49 | 3.75 | 0.61 | 0.57
| Analytical results for Ta in OREAS 285 (Certified Value < 0.01 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.29 | 1.31 | 1.32 | [11 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [11 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | [11 empty] | 3
| 1 | 1 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 1
| 1 | 2 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 25
| 1 | 3 | < 0.01 | < 0.01 | < 0.05 | < 5 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 16
| 1 | 4 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.01 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | [11 empty] | 115
| 1 | 6 | < 0.01 | < 0.01 | < 0.05 | < 5 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05
| Mean | [5 empty] | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| Median | [5 empty] | 0.01 | 0.01 | 0.01 | 0.01 | 0.01
| Std Dev. | [5 empty] | 0 | 0 |  | 0 | 0
| Rel.Std.Dev. | [5 empty] | 0% | 0% |  | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 2.02 | 335.81 | 0.22 | 0.45 | 0.56 | 0 | 0 | 2.02 | 5.39 | 2.02 | 2.02 | 0.67 | 2.02 | 0.67 | 0.67 | 2.02
| Analytical results for Tb in OREAS 285 (Certified Value 0.57 ppm) | [30 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.21 | 1.25 | 1.29 | 1.32 | [21 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [21 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | [21 empty] | 3
| 1 | 1 | 0.61224 | 0.498 | 0.5 | 0.6477 | 0.52 | 0.58 | 0.37 | 0.53 | [21 empty] | 1
| 1 | 2 | 0.58112 | 0.531 | 0.5 | 0.6532 | 0.51 | 0.58 | 0.35 | 0.48 | [21 empty] | 26
| 1 | 3 | 0.60121 | 0.501 | 0.6 | 0.6662 | 0.54 | 0.58 | 0.36 | 0.52 | [21 empty] | 16
| 1 | 4 | 0.58169 | 0.522 | 0.6 | 0.6416 | 0.54 | 0.55 | 0.38 | 0.54 | [21 empty] | 0.56752
| 1 | 5 | 0.60706 | 0.53 | 0.5 | 0.6625 | 0.53 | 0.59 | 0.36 | 0.55 | [21 empty] | 116
| 1 | 6 | 0.60464 | 0.528 | 0.5 | 0.6455 | 0.52 | 0.56 | 0.37 | 0.54
| Mean |  | 0.59799 | 0.51833 | 0.53333 | 0.65278 | 0.52667 | 0.57333 | 0.365 | 0.52667
| Median |  | 0.60292 | 0.525 | 0.5 | 0.65045 | 0.525 | 0.58 | 0.365 | 0.535
| Std Dev. |  | 0.013341 | 0.014949 | 0.05164 | 0.0097831 | 0.012111 | 0.015055 | 0.010488 | 0.025033
| Rel.Std.Dev. |  | 2.231% | 2.884% | 9.682% | 1.499% | 2.299% | 2.626% | 2.873% | 4.753%
| PDM3 |  | 5.37% | -8.667% | -6.024% | 15.02% | -7.198% | 1.025% | -35.68% | -7.198%
| Z-Score (Absolute) |  | 1.03 | 0.12 | NA | 1.82 | 0 | 0.67 | 2.34 | 0
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 285 (Certified Value 0.056 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.29 | 1.30 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | [6 empty] | 3
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.2 | < 0.1 | 0.05 | < 5 | 0.04 | 0.05 | 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.05 | 0.1 | 0.06 | 0.08 | < 2 | 0.06 | < 0.2 | 0.032 | 0.05 | [6 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.03 | 5 | 0.06 | 0.07 | 0.05 | 0.06 | 0.03 | 0.08 | 0.03 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.06 | 0.07 | 0.06 | < 2 | 0.06 | < 0.2 | < 0.01 | 0.05 | [6 empty] | 27
| 1 | 3 | < 0.2 | < 0.1 | 0.06 | 5 | 0.05 | 0.04 | 0.06 | 0.04 | 0.04 | 0.04 | 0.06 | < 0.05 | < 0.1 | < 0.05 | 0.08 | 0.1 | 0.07 | 0.05 | < 2 | 0.05 | < 0.2 | 0.033 | < 0.05 | [6 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.05 | 5 | 0.06 | 0.05 | 0.07 | 0.04 | 0.03 | 0.06 | 0.03 | < 0.05 | < 0.1 | < 0.05 | 0.07 | 0.04 | 0.07 | 0.05 | < 2 | 0.07 | < 0.2 | < 0.01 | 0.07 | [6 empty] | 0.055541
| 1 | 5 | < 0.2 | < 0.1 | 0.06 | < 5 | 0.06 | 0.06 | 0.05 | 0.03 | 0.05 | 0.02 | 0.06 | 0.061604 | < 0.1 | < 0.05 | 0.06 | 0.05 | 0.06 | 0.06 | < 2 | 0.06 | < 0.2 | < 0.01 | 0.05 | [6 empty] | 117
| 1 | 6 | < 0.2 | < 0.1 | 0.07 | 5 | 0.04 | 0.13 | 0.05 | 0.04 | 0.03 | 0.05 | 0.07 | 0.057796 | < 0.1 | < 0.05 | < 0.05 | 0.09 | 0.06 | < 0.05 | < 2 | 0.05 | < 0.2 | < 0.01 | 0.13
| Mean |  |  |  | 0.053333 | 5 | 0.051667 | 0.066667 | 0.055 | 0.043333 | 0.038333 | 0.051667 | 0.05 | 0.0597 |  |  | 0.065 | 0.073333 | 0.065 | 0.06 |  | 0.058333 |  | 0.0325 | 0.07
| Median |  |  |  | 0.055 | 5 | 0.055 | 0.055 | 0.05 | 0.04 | 0.035 | 0.055 | 0.055 | 0.0597 |  |  | 0.065 | 0.075 | 0.065 | 0.06 |  | 0.06 |  | 0.0325 | 0.05
| Std Dev. |  |  |  | 0.013663 | 0 | 0.0098319 | 0.03266 | 0.0083666 | 0.010328 | 0.0098319 | 0.020412 | 0.016733 | 0.0026925 |  |  | 0.01291 | 0.026583 | 0.0054772 | 0.012247 |  | 0.0075277 |  | 0.00070711 | 0.034641
| Rel.Std.Dev. |  |  |  | 25.62% | 0% | 19.03% | 48.99% | 15.21% | 23.83% | 25.65% | 39.51% | 33.47% | 4.51% |  |  | 19.86% | 36.25% | 8.427% | 20.41% |  | 12.9% |  | 2.176% | 49.49%
| PDM3 |  |  |  | -3.975% | 8902% | -6.976% | 20.03% | -0.9741% | -21.98% | -30.98% | -6.976% | -9.976% | 7.488% |  |  | 17.03% | 32.03% | 17.03% | 8.028% |  | 5.028% |  | -41.48% | 26.03%
| Z-Score (Absolute) |  | 3.34 | 0.18 | 0.06 | NA | 0.06 | 1 | 0.18 | 0.64 | 1 | 0.06 | 0.18 | 1.12 | 0.18 | 1.93 | 0.06 | 1.47 | 0.88 | 0.12 | 66.67 | 0.41 | 3.34 | 2.7 | 0.7
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 285 (Certified Value 8.01 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.09 | 1.10 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 9.2187 | 6.69 | 7.8907 | 9.8 | 9.2239 | 7.3 | 6.6 | 9.7 | 7.5 | 7.1 | 6.9 | 7.7483 | 8.3 | 5.9 | 12 | 6.5 | 6.1 | 9.4 | 9.85 | 3.07 | 108.32 | 9.9 | 8.8 | [6 empty] | 1
| 1 | 2 | 8.6027 | 7.38 | 7.0458 | 9.7 | 8.5783 | 8 | 6.6 | 7.3 | 7.3 | 7.8 | 6.5 | 8.2536 | 8.1 | 6 | 11.9 | 6.5 | 5.6 | 9.3 | 10.13 | 2.89 | 106.88 | 12.8 | 7.8 | [6 empty] | 28
| 1 | 3 | 8.475 | 6.82 | 6.8894 | 9.4 | 10.0163 | 8 | 6.7 | 9.2 | 7.5 | 7.3 | 6.2 | 8.3911 | 8.2 | 6.3 | 11.5 | 6 | 6 | 9.1 | 10.71 | 3.04 | 107.6 | 12.2 | 8.9 | [6 empty] | 16
| 1 | 4 | 8.6375 | 7.57 | 8.1953 | 10 | 8.8161 | 7.1 | 6.6 | 9.4 | 7.2 | 7.5 | 6.8 | 7.5963 | 7.9 | 6.3 | 12 | 6.3 | 5.6 | 8.5 | 9.47 | 3.23 | 107.6 | 11.2 | 9 | [6 empty] | 8.0075
| 1 | 5 | 8.434 | 7.63 | 6.3999 | 9.3 | 9.9146 | 7.6 | 6.7 | 8.4 | 7.3 | 7.5 | 6.4 | 8.7097 | 8.2 | 6.2 | 11.8 | 6 | 5.2 | 9.1 | 10.39 | 3.08 | 108.56 | 11.3 | 9.1 | [6 empty] | 118
| 1 | 6 | 8.7715 | 7.34 | 6.2906 | 9.2 | 9.3952 | 7.8 | 6.7 | 9.5 | 7.3 | 7.3 | 6.5 | 8.519 | 8.2 | 6.2 | 11.6 | 5.7 | 6 | 8.4 | 10.3 | 3.09 | 105.84 | 10.1 | 9
| Mean |  | 8.6899 | 7.2383 | 7.1186 | 9.5667 | 9.3241 | 7.6333 | 6.65 | 8.9167 | 7.35 | 7.4167 | 6.55 | 8.203 | 8.15 | 6.15 | 11.8 | 6.1667 | 5.75 | 8.9667 | 10.1417 | 3.0667 | 107.4667 | 11.25 | 8.7667
| Median |  | 8.6201 | 7.36 | 6.9676 | 9.55 | 9.3096 | 7.7 | 6.65 | 9.3 | 7.3 | 7.4 | 6.5 | 8.3224 | 8.2 | 6.2 | 11.85 | 6.15 | 5.8 | 9.1 | 10.215 | 3.075 | 107.6 | 11.25 | 8.95
| Std Dev. |  | 0.28577 | 0.3923 | 0.77657 | 0.31411 | 0.57571 | 0.37238 | 0.054772 | 0.91086 | 0.12247 | 0.24014 | 0.25884 | 0.44033 | 0.13784 | 0.16432 | 0.20976 | 0.32042 | 0.34496 | 0.41793 | 0.43499 | 0.10893 | 0.99492 | 1.1362 | 0.48442
| Rel.Std.Dev. |  | 3.288% | 5.42% | 10.91% | 3.283% | 6.174% | 4.878% | 0.8236% | 10.22% | 1.666% | 3.238% | 3.952% | 5.368% | 1.691% | 2.672% | 1.778% | 5.196% | 5.999% | 4.661% | 4.289% | 3.552% | 0.9258% | 10.1% | 5.526%
| PDM3 |  | 8.522% | -9.605% | -11.1% | 19.47% | 16.44% | -4.673% | -16.95% | 11.35% | -8.211% | -7.378% | -18.2% | 2.442% | 1.78% | -23.2% | 47.36% | -22.99% | -28.19% | 11.98% | 26.65% | -61.7% | 1242% | 40.49% | 9.481%
| Z-Score (Absolute) |  | 0.31 | 0.52 | 0.59 | 0.81 | 0.67 | 0.3 | 0.86 | 0.44 | 0.46 | 0.42 | 0.92 | 0.03 | 0 | 1.15 | 2.1 | 1.14 | 1.38 | 0.47 | 1.14 | 2.92 | 57.04 | 1.78 | 0.35
| Analytical results for Ti in OREAS 285 (Certified Value 0.272 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES/MS |  |  | 3
|  |  | 0.2g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.2519 | 0.2552 | 0.30537 | 0.29152 | 0.22 | 0.2804 | 0.24 | 0.243 | 0.276 | 0.295 | 0.279 | 0.275 | 0.288 | 0.26762 | 0.238 | 0.27 | 0.31 | 0.26 | 0.264 | 0.26 | 0.2416 | 0.2766 | 0.355 | 0.2866 | 0.24784 | 0.327 | 0.3 |  |  | 1
| 1 | 2 | 0.25476 | 0.2564 | 0.30326 | 0.29999 | 0.22 | 0.2728 | 0.24 | 0.26 | 0.274 | 0.281 | 0.278 | 0.274 | 0.28 | 0.282 | 0.231 | 0.27 | 0.31 | 0.27 | 0.259 | 0.26 | 0.2444 | 0.2743 | 0.37 | 0.2788 | 0.24576 | 0.399 | 0.27 |  |  | 29
| 1 | 3 | 0.2584 | 0.2543 | 0.3054 | 0.29991 | 0.22 | 0.28 | 0.25 | 0.258 | 0.279 | 0.293 | 0.275 | 0.273 | 0.28 | 0.27401 | 0.23 | 0.29 | 0.312 | 0.27 | 0.249 | 0.26 | 0.2534 | 0.2877 | 0.355 | 0.2818 | 0.24552 | 0.386 | 0.29 |  |  | 16
| 1 | 4 | 0.25199 | 0.2598 | 0.3035 | 0.29219 | 0.24 | 0.2885 | 0.24 | 0.243 | 0.279 | 0.292 | 0.273 | 0.277 | 0.281 | 0.27886 | 0.234 | 0.29 | 0.315 | 0.26 | 0.258 | 0.26 | 0.2404 | 0.2715 | 0.302 | 0.2902 | 0.24696 | 0.363 | 0.29 |  |  | 0.27158
| 1 | 5 | 0.2511 | 0.2671 | 0.30194 | 0.29049 | 0.22 | 0.2901 | 0.24 | 0.245 | 0.278 | 0.295 | 0.272 | 0.282 | 0.278 | 0.26619 | 0.233 | 0.29 | 0.304 | 0.28 | 0.254 | 0.26 | 0.2474 | 0.2775 | 0.309 | 0.2782 | 0.24784 | 0.363 | 0.29 |  |  | 119
| 1 | 6 | 0.25029 | 0.2649 | 0.30152 | 0.29741 | 0.23 | 0.2842 | 0.25 | 0.264 | 0.276 | 0.294 | 0.274 | 0.281 | 0.279 | 0.27324 | 0.236 | 0.29 | 0.3 | 0.27 | 0.264 | 0.26 | 0.2436 | 0.2815 | 0.331 | 0.2836 | 0.24608 | 0.329 | 0.3
| Mean |  | 0.25307 | 0.25962 | 0.3035 | 0.29525 | 0.225 | 0.28267 | 0.24333 | 0.25217 | 0.277 | 0.29167 | 0.27517 | 0.277 | 0.281 | 0.27365 | 0.23367 | 0.28333 | 0.3085 | 0.26833 | 0.258 | 0.26 | 0.24513 | 0.27818 | 0.337 | 0.2832 | 0.24667 | 0.36117 | 0.29
| Median |  | 0.25195 | 0.2581 | 0.30338 | 0.2948 | 0.22 | 0.2823 | 0.24 | 0.2515 | 0.277 | 0.2935 | 0.2745 | 0.276 | 0.28 | 0.27363 | 0.2335 | 0.29 | 0.31 | 0.27 | 0.2585 | 0.26 | 0.244 | 0.27705 | 0.343 | 0.2827 | 0.24652 | 0.363 | 0.29
| Std Dev. |  | 0.0030144 | 0.0053304 | 0.0016459 | 0.0043547 | 0.0083666 | 0.006344 | 0.005164 | 0.0095376 | 0.002 | 0.0053541 | 0.0027869 | 0.0037417 | 0.0035777 | 0.0061532 | 0.0030111 | 0.010328 | 0.0055045 | 0.0075277 | 0.005831 | 0 | 0.0047171 | 0.0057322 | 0.027503 | 0.0046269 | 0.0010316 | 0.029178 | 0.010954
| Rel.Std.Dev. |  | 1.191% | 2.053% | 0.5423% | 1.475% | 3.718% | 2.244% | 2.122% | 3.782% | 0.722% | 1.836% | 1.013% | 1.351% | 1.273% | 2.249% | 1.289% | 3.645% | 1.784% | 2.805% | 2.26% | 0% | 1.924% | 2.061% | 8.161% | 1.634% | 0.4182% | 8.079% | 3.777%
| PDM3 |  | -6.816% | -4.407% | 11.75% | 8.715% | -17.15% | 4.081% | -10.4% | -7.15% | 1.994% | 7.394% | 1.319% | 1.994% | 3.467% | 0.7618% | -13.96% | 4.326% | 13.59% | -1.197% | -5.002% | -4.266% | -9.74% | 2.43% | 24.09% | 4.277% | -9.175% | 32.98% | 6.781%
| Z-Score (Absolute) |  | 0.93 | 0.67 | 1.03 | 0.71 | 2.02 | 0.22 | 1.31 | 0.96 | 0 | 0.57 | 0.07 | 0 | 0.16 | 0.13 | 1.68 | 0.25 | 1.22 | 0.34 | 0.74 | 0.66 | 1.24 | 0.05 | 2.33 | 0.24 | 1.18 | 3.26 | 0.5
| Analytical results for Tl in OREAS 285 (Certified Value 0.66 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.2g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 0.66076 | 0.68 | 0.70483 | 0.67 | < 10 | 0.407 | 0.58 | 0.59 | 0.67 | 0.67 | 0.66 | 0.56 | 0.62414 | 0.69 | 0.63 | 0.69 | 0.69 | 0.66 | 0.65 | < 5 | 0.72 | < 5 | 0.69 | 0.63 | [5 empty] | 1
| 1 | 2 | 0.66799 | 0.71 | 0.6798 | 0.65 | < 10 | 0.431 | 0.63 | 0.6 | 0.64 | 0.69 | 0.69 | 0.58 | 0.66018 | 0.69 | 0.64 | 0.69 | 0.7 | 0.64 | 0.66 | < 5 | 0.7 | < 5 | 0.73 | 0.57 | [5 empty] | 30
| 1 | 3 | 0.66073 | 0.69 | 0.68174 | 0.67 | < 10 | 0.471 | 0.62 | 0.6 | 0.68 | 0.7 | 0.65 | 0.55 | 0.66838 | 0.68 | 0.65 | 0.66 | 0.66 | 0.68 | 0.67 | < 5 | 0.71 | < 5 | 0.72 | 0.62 | [5 empty] | 16
| 1 | 4 | 0.65703 | 0.69 | 0.64605 | 0.69 | < 10 | 0.453 | 0.62 | 0.61 | 0.69 | 0.68 | 0.67 | 0.61 | 0.66438 | 0.67 | 0.66 | 0.7 | 0.68 | 0.66 | 0.65 | < 5 | 0.66 | < 5 | 0.72 | 0.62 | [5 empty] | 0.66473
| 1 | 5 | 0.66909 | 0.71 | 0.70651 | 0.65 | < 10 | 0.445 | 0.6 | 0.61 | 0.65 | 0.69 | 0.68 | 0.57 | 0.61686 | 0.71 | 0.66 | 0.68 | 0.68 | 0.6 | 0.71 | < 5 | 0.69 | < 5 | 0.7 | 0.63 | [5 empty] | 120
| 1 | 6 | 0.70136 | 0.7 | 0.67538 | 0.65 | < 10 | 0.428 | 0.6 | 0.61 | 0.66 | 0.66 | 0.67 | 0.56 | 0.6646 | 0.66 | 0.67 | 0.68 | 0.68 | 0.64 | 0.68 | < 5 | 0.7 | < 5 | 0.67 | 0.64
| Mean |  | 0.66949 | 0.69667 | 0.68238 | 0.66333 |  | 0.43917 | 0.60833 | 0.60333 | 0.665 | 0.68167 | 0.67 | 0.57167 | 0.64976 | 0.68333 | 0.65167 | 0.68333 | 0.68167 | 0.64667 | 0.67 |  | 0.69667 |  | 0.705 | 0.61833
| Median |  | 0.66438 | 0.695 | 0.68077 | 0.66 |  | 0.438 | 0.61 | 0.605 | 0.665 | 0.685 | 0.67 | 0.565 | 0.66228 | 0.685 | 0.655 | 0.685 | 0.68 | 0.65 | 0.665 |  | 0.7 |  | 0.71 | 0.625
| Std Dev. |  | 0.016289 | 0.012111 | 0.02219 | 0.01633 |  | 0.022203 | 0.018348 | 0.008165 | 0.018708 | 0.01472 | 0.014142 | 0.02137 | 0.022927 | 0.017512 | 0.01472 | 0.013663 | 0.013292 | 0.027325 | 0.022804 |  | 0.020656 |  | 0.022583 | 0.024833
| Rel.Std.Dev. |  | 2.433% | 1.738% | 3.252% | 2.462% |  | 5.056% | 3.016% | 1.353% | 2.813% | 2.159% | 2.111% | 3.738% | 3.529% | 2.563% | 2.259% | 1.999% | 1.95% | 4.226% | 3.404% |  | 2.965% |  | 3.203% | 4.016%
| PDM3 |  | 0.7161% | 4.804% | 2.655% | -0.2106% |  | -33.93% | -8.485% | -9.237% | 0.04016% | 2.547% | 0.7923% | -14% | -2.253% | 2.798% | -1.966% | 2.798% | 2.547% | -2.718% | 0.7923% |  | 4.804% |  | 6.058% | -6.98%
| Z-Score (Absolute) |  | 0.02 | 0.83 | 0.38 | 0.21 | 134.01 | 7.14 | 1.91 | 2.06 | 0.15 | 0.36 | 0 | 3.04 | 0.63 | 0.41 | 0.57 | 0.41 | 0.36 | 0.72 | 0 | 56.64 | 0.83 | 56.64 | 1.08 | 1.6
| Analytical results for Tm in OREAS 285 (Indicative Value 0.12 ppm) | [30 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [25 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | [25 empty] | 2
| 1 | 1 | 0.11076 | 0.116 | 0.1 | 0.1362 | [25 empty] | 1
| 1 | 2 | 0.11321 | 0.12 | 0.1 | 0.1333 | [25 empty] | 5
| 1 | 3 | 0.11874 | 0.116 | 0.1 | 0.1332 | [25 empty] | 16
| 1 | 4 | 0.11405 | 0.117 | 0.1 | 0.125 | [25 empty] | 0.11563
| 1 | 5 | 0.11823 | 0.113 | 0.1 | 0.1311 | [25 empty] | 11
| 1 | 6 | 0.11754 | 0.116 | 0.1 | 0.1258
| Mean |  | 0.11542 | 0.11633 | 0.1 | 0.13077
| Median |  | 0.1158 | 0.116 | 0.1 | 0.13215
| Std Dev. |  | 0.0032207 | 0.0022509 | 1.5202e-17 | 0.0044698
| Rel.Std.Dev. |  | 2.79% | 1.935% | 1.52e-14% | 3.418%
| PDM3 |  | -0.1801% | 0.6078% | -13.52% | 13.09%
| Z-Score (Absolute) |  | 0.04 | 0.04 | 1.4 | 1.31
| Analytical results for U in OREAS 285 (Certified Value 3.45 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 3.1602 | 3.56 | 3.4 | 3.3958 | 2.76 | 3.2 | 3.12 | 4.18 | 3.88 | 3.39 | 3.51 | 3.3508 | 3.6 | 3.55 | 3.69 | 3.58 | 3.5 | 3.6 | 3.64 | 3.18 | < 5 | 3.25 | 3.31 | [6 empty] | 1
| 1 | 2 | 3.3637 | 3.72 | 3.5 | 3.7351 | 3.01 | 3.49 | 2.8 | 3.34 | 3.46 | 3.42 | 2.87 | 3.4813 | 3.8 | 3.39 | 3.8 | 3.53 | 3.4 | 4.01 | 3.22 | 2.84 | < 5 | 4.13 | 3.4 | [6 empty] | 32
| 1 | 3 | 3.4945 | 3.63 | 4 | 3.5797 | 3.19 | 3.23 | 2.96 | 4.14 | 3.8 | 3.53 | 2.73 | 3.4012 | 3.3 | 3.3 | 3.75 | 3.32 | 3.5 | 3.39 | 3.79 | 2.93 | < 5 | 3.91 | 3.57 | [6 empty] | 16
| 1 | 4 | 3.101 | 3.52 | 3.7 | 2.9994 | 3.06 | 3.14 | 3.04 | 3.74 | 3.34 | 3.35 | 3.1 | 3.3053 | 4.2 | 3.51 | 3.85 | 3.29 | 3.4 | 3.43 | 3.57 | 3.05 | < 5 | 4.12 | 4.03 | [6 empty] | 3.4524
| 1 | 5 | 3.4079 | 3.57 | 3.5 | 3.2108 | 2.82 | 3.68 | 3.07 | 3.46 | 3.43 | 3.01 | 3.25 | 3.2228 | 3.9 | 3.47 | 3.81 | 3.16 | 3.2 | 3.94 | 3.74 | 3.22 | < 5 | 4.15 | 3.96 | [6 empty] | 121
| 1 | 6 | 3.642 | 3.44 | 3.3 | 3.3868 | 2.77 | 3.18 | 4.33 | 3.83 | 3.58 | 3.33 | 2.85 | 3.5024 | 2.9 | 3.45 | 3.91 | 3.18 | 3.2 | 3.79 | 3.89 | 3.01 | < 5 | 3.78 | 4.08
| Mean |  | 3.3616 | 3.5733 | 3.5667 | 3.3846 | 2.935 | 3.32 | 3.22 | 3.7817 | 3.5817 | 3.3383 | 3.0517 | 3.3773 | 3.6167 | 3.445 | 3.8017 | 3.3433 | 3.3667 | 3.6933 | 3.6417 | 3.0383 |  | 3.89 | 3.725
| Median |  | 3.3858 | 3.565 | 3.5 | 3.3913 | 2.915 | 3.215 | 3.055 | 3.785 | 3.52 | 3.37 | 2.985 | 3.376 | 3.7 | 3.46 | 3.805 | 3.305 | 3.4 | 3.695 | 3.69 | 3.03 |  | 4.015 | 3.765
| Std Dev. |  | 0.20342 | 0.095429 | 0.25033 | 0.26063 | 0.1774 | 0.21587 | 0.55516 | 0.34342 | 0.21582 | 0.17555 | 0.29288 | 0.10657 | 0.46224 | 0.089387 | 0.076529 | 0.1758 | 0.13663 | 0.26097 | 0.23507 | 0.14497 |  | 0.34647 | 0.33946
| Rel.Std.Dev. |  | 6.051% | 2.671% | 7.019% | 7.7% | 6.044% | 6.502% | 17.24% | 9.081% | 6.026% | 5.259% | 9.597% | 3.156% | 12.78% | 2.595% | 2.013% | 5.258% | 4.058% | 7.066% | 6.455% | 4.771% |  | 8.907% | 9.113%
| PDM3 |  | -2.632% | 3.502% | 3.309% | -1.964% | -14.99% | -3.835% | -6.732% | 9.537% | 3.744% | -3.304% | -11.61% | -2.175% | 4.758% | -0.2148% | 10.12% | -3.16% | -2.484% | 6.978% | 5.482% | -11.99% |  | 12.67% | 7.895%
| Z-Score (Absolute) |  | 0.08 | 0.65 | 0.62 | 0 | 1.54 | 0.22 | 0.56 | 1.36 | 0.67 | 0.16 | 1.14 | 0.02 | 0.79 | 0.21 | 1.43 | 0.14 | 0.06 | 1.06 | 0.88 | 1.18 | 3.03 | 1.73 | 1.16
| Analytical results for V in OREAS 285 (Certified Value 61 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  |  |  | 0
|  |  | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 60.1167 | 61 | 64.698 | 59 | 56.6959 | 61 | 59 | 59 | 62 | 61 | 63 | 60 | 57.2502 | 64 | 57 | 60 | 54.8 | 62 | 61 | 57 | 65 | 65.47 | 53.6 | 61.05 | 63.2 | 61 |  |  |  | 1
| 1 | 2 | 60.7073 | 60 | 65.1231 | 58 | 54.8361 | 61 | 63 | 59 | 60 | 60 | 62 | 59 | 57.1863 | 63 | 57 | 62 | 55.7 | 62 | 62 | 58 | 64 | 66.4 | 52.5 | 60.14 | 72.5 | 55 |  |  |  | 16
| 1 | 3 | 61.1608 | 60 | 65.0031 | 58 | 55.8605 | 61 | 62 | 60 | 62 | 60 | 63 | 60 | 56.7395 | 62 | 58 | 63 | 53.4 | 62 | 61 | 60 | 67 | 64.83 | 54.3 | 60.57 | 70.9 | 58 |  |  |  | 16
| 1 | 4 | 60.3014 | 61 | 65.0159 | 60 | 57.5405 | 61 | 59 | 60 | 62 | 60 | 63 | 60 | 57.6409 | 62 | 58 | 62 | 54.1 | 62 | 61 | 57 | 64 | 62.84 | 55.2 | 60.61 | 68 | 59 |  |  |  | 60.7756
| 1 | 5 | 59.917 | 63 | 63.1497 | 60 | 59.4116 | 61 | 59 | 59 | 62 | 59 | 64 | 60 | 58.4617 | 63 | 58 | 61 | 51.7 | 60 | 61 | 58 | 65 | 64.28 | 52.9 | 61.19 | 69 | 60 |  |  |  | 122
| 1 | 6 | 59.8763 | 62 | 64.1934 | 58 | 58.9843 | 62 | 64 | 59 | 62 | 60 | 64 | 59 | 58.5941 | 62 | 58 | 61 | 52 | 62 | 61 | 57 | 66 | 64.62 | 53.8 | 60.57 | 64 | 61
| Mean |  | 60.3466 | 61.1667 | 64.5305 | 58.8333 | 57.2215 | 61.1667 | 61 | 59.3333 | 61.6667 | 60 | 63.1667 | 59.6667 | 57.6455 | 62.6667 | 57.6667 | 61.5 | 53.6167 | 61.6667 | 61.1667 | 57.8333 | 65.1667 | 64.74 | 53.7167 | 60.6883 | 67.9333 | 59
| Median |  | 60.209 | 61 | 64.8506 | 58.5 | 57.1182 | 61 | 60.5 | 59 | 62 | 60 | 63 | 60 | 57.4456 | 62.5 | 58 | 61.5 | 53.75 | 62 | 61 | 57.5 | 65 | 64.725 | 53.7 | 60.59 | 68.5 | 59.5
| Std Dev. |  | 0.50072 | 1.169 | 0.75611 | 0.98319 | 1.779 | 0.40825 | 2.2804 | 0.5164 | 0.8165 | 0.63246 | 0.75277 | 0.5164 | 0.74217 | 0.8165 | 0.5164 | 1.0488 | 1.569 | 0.8165 | 0.40825 | 1.169 | 1.169 | 1.1944 | 0.9704 | 0.37876 | 3.706 | 2.2804
| Rel.Std.Dev. |  | 0.8297% | 1.911% | 1.172% | 1.671% | 3.109% | 0.6674% | 3.738% | 0.8703% | 1.324% | 1.054% | 1.192% | 0.8655% | 1.287% | 1.303% | 0.8955% | 1.705% | 2.926% | 1.324% | 0.6674% | 2.021% | 1.794% | 1.845% | 1.807% | 0.6241% | 5.455% | 3.865%
| PDM3 |  | -0.7059% | 0.6434% | 6.178% | -3.196% | -5.848% | 0.6434% | 0.3692% | -2.373% | 1.466% | -1.276% | 3.934% | -1.825% | -5.15% | 3.112% | -5.115% | 1.192% | -11.78% | 1.466% | 0.6434% | -4.841% | 7.225% | 6.523% | -11.61% | -0.1436% | 11.78% | -2.922%
| Z-Score (Absolute) |  | 0.18 | 0.12 | 1.36 | 0.74 | 1.33 | 0.12 | 0.06 | 0.56 | 0.3 | 0.31 | 0.85 | 0.43 | 1.18 | 0.67 | 1.17 | 0.24 | 2.66 | 0.3 | 0.12 | 1.11 | 1.59 | 1.43 | 2.62 | 0.06 | 2.61 | 0.68
| Analytical results for W in OREAS 285 (Certified Value 11.2 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [5 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [5 empty] | 1
| 1 | 1 | 11.4471 | 10.08 | 11.9 | 14.2994 | 14 | 8.86 | 11.2 | 13.35 | 11.6 | 11.35 | 11.35 | 12.9284 | 11.4 | 10.3 | 11.6 | 13.6 | 9.6 | 13.3 | 11 | 12.06 | 10.3 | 10.736 | 10.9 | 9.1 | [5 empty] | 1
| 1 | 2 | 11.3719 | 10.41 | 11.6 | 12.9414 | 14 | 8.95 | 10.8 | 12.1 | 11.75 | 11.3 | 11.1 | 12.241 | 10.8 | 10.4 | 11.7 | 13.1 | 8.8 | 13.1 | 10 | 11.75 | 9.9 | 10.44 | 10.7 | 7.1 | [5 empty] | 34
| 1 | 3 | 11.7474 | 10.37 | 11.8 | 11.147 | 13 | 8.89 | 10.8 | 12 | 11.7 | 10.85 | 11 | 12.4311 | 11.3 | 10.7 | 11.7 | 12.4 | 9.4 | 13.5 | 11 | 12.24 | 10.9 | 10.368 | 10.7 | 8.8 | [5 empty] | 16
| 1 | 4 | 11.7451 | 10.09 | 12.1 | 13.476 | 13 | 9.36 | 10.6 | 12.3 | 12 | 10.35 | 11.85 | 12.2773 | 10.7 | 10.9 | 12.2 | 12.7 | 8.7 | 12.7 | 10 | 11.53 | 10.7 | 10.592 | 10.9 | 8.5 | [5 empty] | 11.2239
| 1 | 5 | 11.2856 | 10.25 | 11.5 | 14.7388 | 12 | 9.54 | 10.65 | 11.75 | 11.7 | 10.8 | 11 | 11.9203 | 11.2 | 10.5 | 12.1 | 12.6 | 8.5 | 14.3 | 11 | 11.71 | 10.6 | 10.552 | 11 | 9.1 | [5 empty] | 123
| 1 | 6 | 11.4151 | 10.08 | 11.9 | 14.245 | 14 | 8.44 | 10.35 | 13.5 | 11.9 | 10.9 | 11.9 | 12.8687 | 11.5 | 10.8 | 12 | 12.2 | 9.1 | 13 | 11 | 11.93 | 10.2 | 10.064 | 10.9 | 8.7
| Mean |  | 11.502 | 10.2133 | 11.8 | 13.4746 | 13.3333 | 9.0067 | 10.7333 | 12.5 | 11.775 | 10.925 | 11.3667 | 12.4445 | 11.15 | 10.6 | 11.8833 | 12.7667 | 9.0167 | 13.3167 | 10.6667 | 11.87 | 10.4333 | 10.4587 | 10.85 | 8.55
| Median |  | 11.4311 | 10.17 | 11.85 | 13.8605 | 13.5 | 8.92 | 10.725 | 12.2 | 11.725 | 10.875 | 11.225 | 12.3542 | 11.25 | 10.6 | 11.85 | 12.65 | 8.95 | 13.2 | 11 | 11.84 | 10.45 | 10.496 | 10.9 | 8.75
| Std Dev. |  | 0.1968 | 0.15188 | 0.21909 | 1.3095 | 0.8165 | 0.39211 | 0.28225 | 0.73959 | 0.14748 | 0.36708 | 0.41433 | 0.38945 | 0.32711 | 0.23664 | 0.24833 | 0.50859 | 0.42622 | 0.55287 | 0.5164 | 0.25776 | 0.36697 | 0.2314 | 0.12247 | 0.74766
| Rel.Std.Dev. |  | 1.711% | 1.487% | 1.857% | 9.718% | 6.124% | 4.354% | 2.63% | 5.917% | 1.252% | 3.36% | 3.645% | 3.129% | 2.934% | 2.232% | 2.09% | 3.984% | 4.727% | 4.152% | 4.841% | 2.172% | 3.517% | 2.213% | 1.129% | 8.745%
| PDM3 |  | 2.478% | -9.004% | 5.133% | 20.05% | 18.79% | -19.75% | -4.371% | 11.37% | 4.91% | -2.663% | 1.272% | 10.87% | -0.6587% | -5.559% | 5.875% | 13.75% | -19.67% | 18.65% | -4.965% | 5.756% | -7.044% | -6.818% | -3.332% | -23.82%
| Z-Score (Absolute) |  | 0.23 | 0.97 | 0.5 | 2.05 | NA | 2.08 | 0.49 | 1.15 | 0.48 | 0.31 | 0.1 | 1.1 | 0.1 | 0.61 | 0.58 | 1.4 | 2.07 | 1.9 | NA | 0.57 | 0.76 | 0.74 | 0.38 | 2.51
Results from laboratories 1.11 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 285 (Certified Value 10.9 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [8 empty] | 1
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [8 empty] | 2
| 1 | 1 | 11.1745 | 10.63 | 10.5 | 10.7131 | 10 | 10.45 | 12.45 | 10.7 | 11.25 | 9.83 | 10.0329 | 10.7 | 9.51 | 12.6 | 10.2 | 10.89 | 11.36 | 6.78 | 14.121 | 11.5 | 11.13 | [8 empty] | 1
| 1 | 2 | 11.0167 | 10.89 | 10.7 | 10.6213 | 10 | 10.8 | 11.2 | 10.7 | 11.6 | 10.1 | 10.4654 | 10.4 | 9.92 | 12.5 | 10.3 | 10.91 | 11.41 | 6.5 | 13.887 | 12.4 | 9.55 | [8 empty] | 35
| 1 | 3 | 11.2796 | 10.62 | 10.9 | 11.1614 | 10 | 10.75 | 11.7 | 10.5 | 11.05 | 9.56 | 10.349 | 10.4 | 9.84 | 12.7 | 9.86 | 10.84 | 11.86 | 6.71 | 14.076 | 12.1 | 10.79 | [8 empty] | 16
| 1 | 4 | 10.9341 | 11.01 | 11.2 | 11.0337 | 10 | 10.75 | 11.85 | 10.5 | 11.25 | 10.4 | 10.35 | 10.4 | 9.98 | 12.8 | 10.1 | 10.5 | 11.11 | 6.92 | 14.103 | 11.8 | 11.01 | [8 empty] | 10.895
| 1 | 5 | 10.9386 | 11.09 | 10.6 | 11.2739 | 10 | 10.55 | 10.75 | 10.3 | 11.25 | 9.69 | 10.2246 | 10.6 | 9.72 | 12.6 | 9.98 | 11.17 | 11.66 | 6.78 | 14.373 | 12.1 | 11.27 | [8 empty] | 124
| 1 | 6 | 10.999 | 10.97 | 10.6 | 10.8464 | 11 | 10.2 | 11.8 | 10.8 | 11.4 | 10.2 | 10.0579 | 10.3 | 9.79 | 12.4 | 9.85 | 10.65 | 11.57 | 6.61 | 14.238 | 11.5 | 11.11
| Mean |  | 11.0571 | 10.8683 | 10.75 | 10.9416 | 10.1667 | 10.5833 | 11.625 | 10.5833 | 11.3 | 9.9633 | 10.2466 | 10.4667 | 9.7933 | 12.6 | 10.0483 | 10.8267 | 11.495 | 6.7167 | 14.133 | 11.9 | 10.81
| Median |  | 11.0079 | 10.93 | 10.65 | 10.94 | 10 | 10.65 | 11.75 | 10.6 | 11.25 | 9.965 | 10.2868 | 10.4 | 9.815 | 12.6 | 10.04 | 10.865 | 11.49 | 6.745 | 14.112 | 11.95 | 11.06
| Std Dev. |  | 0.13963 | 0.19924 | 0.25884 | 0.25733 | 0.40825 | 0.23166 | 0.58545 | 0.18348 | 0.18439 | 0.32266 | 0.17369 | 0.15055 | 0.16657 | 0.14142 | 0.18357 | 0.23106 | 0.26067 | 0.14679 | 0.16336 | 0.36332 | 0.6375
| Rel.Std.Dev. |  | 1.263% | 1.833% | 2.408% | 2.352% | 4.016% | 2.189% | 5.036% | 1.734% | 1.632% | 3.238% | 1.695% | 1.438% | 1.701% | 1.122% | 1.827% | 2.134% | 2.268% | 2.185% | 1.156% | 3.053% | 5.897%
| PDM3 |  | 1.487% | -0.2451% | -1.331% | 0.4277% | -6.685% | -2.861% | 6.7% | -2.861% | 3.717% | -8.552% | -5.952% | -3.932% | -10.11% | 15.65% | -7.771% | -0.6275% | 5.507% | -38.35% | 29.72% | 9.224% | -0.7805%
| Z-Score (Absolute) |  | 0.31 | 0.06 | 0.09 | 0.16 | NA | 0.3 | 1.03 | 0.3 | 0.62 | 1.09 | 0.73 | 0.45 | 1.31 | 2.28 | 0.99 | 0.01 | 0.87 | 5.25 | 4.24 | 1.38 | 0.01
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 285 (Certified Value 0.71 ppm) | [30 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.21 | 1.25 | 1.29 | 1.32 | [21 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [21 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | [21 empty] | 3
| 1 | 1 | 0.7438 | 0.74 | 0.7 | 0.859 | 0.7 | 0.7 | 0.61 | 0.7 | [21 empty] | 1
| 1 | 2 | 0.72448 | 0.717 | 0.7 | 0.8438 | 0.7 | 0.7 | 0.62 | 0.6 | [21 empty] | 36
| 1 | 3 | 0.73942 | 0.733 | 0.7 | 0.8505 | 0.7 | 0.7 | 0.64 | 0.7 | [21 empty] | 16
| 1 | 4 | 0.73871 | 0.718 | 0.7 | 0.8045 | 0.7 | 0.7 | 0.65 | 0.7 | [21 empty] | 0.7142
| 1 | 5 | 0.74945 | 0.718 | 0.7 | 0.807 | 0.7 | 0.8 | 0.62 | 0.7 | [21 empty] | 125
| 1 | 6 | 0.76023 | 0.729 | 0.7 | 0.8653 | 0.7 | 0.7 | 0.61 | 0.7
| Mean |  | 0.74268 | 0.72583 | 0.7 | 0.83835 | 0.7 | 0.71667 | 0.625 | 0.68333
| Median |  | 0.74161 | 0.7235 | 0.7 | 0.84715 | 0.7 | 0.7 | 0.62 | 0.7
| Std Dev. |  | 0.01194 | 0.0096212 | 1.2162e-16 | 0.026301 | 1.2162e-16 | 0.040825 | 0.016432 | 0.040825
| Rel.Std.Dev. |  | 1.608% | 1.326% | 1.737e-14% | 3.137% | 1.737e-14% | 5.696% | 2.629% | 5.974%
| PDM3 |  | 3.988% | 1.629% | -1.988% | 17.38% | -1.988% | 0.3458% | -12.49% | -4.321%
| Z-Score (Absolute) |  | 1.09 | 0.56 | 0.26 | 4.13 | 0.26 | 0.26 | 2.64 | 0.79
| Analytical results for Zn in OREAS 285 (Certified Value 77 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS |  | 0
|  |  | 0.25g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 15g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g |  | 1
| 1 | 1 | 70.01 | 79.7785 | 80 | 68.8092 | 110 | 79.3803 | 71 | 72 | 75.7 | 77 | 84 | 79 | 77 | 81 | 75.7308 | 78 | 79 | 81 | 78.4 | 73.7 | 76 | 68 | 79 | 68.24 | 73.9 | 93.84 | 54 | 75 |  | 1
| 1 | 2 | 71.34 | 79.8874 | 81 | 71.5635 | 108 | 74.3851 | 72 | 76 | 81.7 | 78 | 80 | 79 | 77 | 80 | 75.8606 | 77 | 79 | 81 | 79.9 | 76 | 77 | 69 | 78 |  | 72.7 | 91.84 | 55 | 67 |  | 17
| 1 | 3 | 70.19 | 80.0709 | 80 | 71.5766 | 111 | 73.9719 | 71 | 79 | 80.7 | 79 | 84 | 78 | 78 | 80 | 74.7495 | 75 | 81 | 80 | 77.1 | 73.9 | 76 | 71 | 80 | 66.85 | 74.3 | 95.44 | 55 | 72 |  | 16
| 1 | 4 | 70.65 | 79.1855 | 84 | 71.9907 | 114 | 78.4177 | 70 | 79 | 75.8 | 79 | 84 | 78 | 78 | 80 | 74.0582 | 76 | 81 | 80 | 78.2 | 75.2 | 76 | 67 | 77 | 56.68 | 75.1 | 95.92 | 54 | 73 |  | 76.7123
| 1 | 5 | 70.8 | 79.1318 | 82 | 70.2484 | 110 | 79.3039 | 71 | 78 | 77.2 | 78 | 84 | 78 | 79 | 80 | 73.9307 | 78 | 80 | 78 | 79.6 | 75.4 | 76 | 69 | 77 | 58.05 | 73.4 | 97.44 | 59 | 74 |  | 126
| 1 | 6 | 70.33 | 79.3197 | 82 | 69.719 | 109 | 78.0799 | 72 | 75 | 83.3 | 78 | 83 | 79 | 79 | 79 | 75.3996 | 76 | 82 | 81 | 77.4 | 75.3 | 76 | 69 | 79 | 61.06 | 73.2 | 92.88 | 51 | 75
| Mean |  | 70.5533 | 79.5623 | 81.5 | 70.6512 | 110.3333 | 77.2565 | 71.1667 | 76.5 | 79.0667 | 78.1667 | 83.1667 | 78.5 | 78 | 80 | 74.9549 | 76.6667 | 80.3333 | 80.1667 | 78.4333 | 74.9167 | 76.1667 | 68.8333 | 78.3333 | 62.176 | 73.7667 | 94.56 | 54.6667 | 72.6667
| Median |  | 70.49 | 79.5491 | 81.5 | 70.906 | 110 | 78.2488 | 71 | 77 | 78.95 | 78 | 84 | 78.5 | 78 | 80 | 75.0746 | 76.5 | 80.5 | 80.5 | 78.3 | 75.25 | 76 | 69 | 78.5 | 61.06 | 73.65 | 94.64 | 54.5 | 73.5
| Std Dev. |  | 0.48294 | 0.39931 | 1.5166 | 1.2576 | 2.0656 | 2.4398 | 0.75277 | 2.7386 | 3.2562 | 0.75277 | 1.6021 | 0.54772 | 0.89443 | 0.63246 | 0.8385 | 1.2111 | 1.2111 | 1.169 | 1.1325 | 0.91086 | 0.40825 | 1.3292 | 1.2111 | 5.1743 | 0.85713 | 2.0812 | 2.582 | 3.0111
| Rel.Std.Dev. |  | 0.6845% | 0.5019% | 1.861% | 1.78% | 1.872% | 3.158% | 1.058% | 3.58% | 4.118% | 0.963% | 1.926% | 0.6977% | 1.147% | 0.7906% | 1.119% | 1.58% | 1.508% | 1.458% | 1.444% | 1.216% | 0.536% | 1.931% | 1.546% | 8.322% | 1.162% | 2.201% | 4.723% | 4.144%
| PDM3 |  | -8.029% | 3.715% | 6.241% | -7.901% | 43.83% | 0.7094% | -7.229% | -0.2767% | 3.069% | 1.896% | 8.414% | 2.33% | 1.679% | 4.286% | -2.291% | -0.05947% | 4.72% | 4.503% | 2.244% | -2.341% | -0.7113% | -10.27% | 2.113% | -18.95% | -3.84% | 23.27% | -28.74% | -5.274%
| Z-Score (Absolute) |  | 1.77 | 0.48 | 0.97 | 1.75 | 8.2 | 0.09 | 1.62 | 0.28 | 0.36 | 0.14 | 1.39 | 0.22 | 0.09 | 0.59 | 0.67 | 0.24 | 0.68 | 0.64 | 0.2 | 0.68 | 0.37 | 2.21 | 0.18 | 3.87 | 0.97 | 4.25 | 5.76 | 1.24
| Analytical results for Zr in OREAS 285 (Certified Value 8.11 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.03 | 1.04 | 1.09 | 1.10 | 1.11 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | [8 empty] | 2
|  |  | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.2g | 0.5g | 0.25g | [8 empty] | 3
| 1 | 1 | 8.2057 | 7 | 8.5 | 3.3138 | 9 | 7 | 8.5 | 8.4 | 8 | 6.9 | 8.2629 | 8.1 | 7.5 | 8.9 | 7.2 | 8.3 | 9 | 8.38 | < 0.5 | 8.1 | 8.5 | [8 empty] | 1
| 1 | 2 | 7.9441 | 7.3 | 8.7 | 4.8611 | 9 | 7.2 | 8.2 | 8.4 | 8.4 | 6.8 | 7.8645 | 7.5 | 7.6 | 8.9 | 7.3 | 8.4 | 9 | 8.18 | < 0.5 | 10.3 | 7.4 | [8 empty] | 18
| 1 | 3 | 8.128 | 7.2 | 9.1 | 7.988 | 9 | 7.1 | 8.6 | 8.1 | 8.2 | 6.5 | 7.6605 | 7.7 | 7.7 | 9.1 | 6.9 | 8.1 | 9.2 | 8.53 | < 0.5 | 9.7 | 8.2 | [8 empty] | 16
| 1 | 4 | 7.7389 | 7.5 | 8.9 | 4.7071 | 9 | 7.2 | 8.5 | 8.2 | 8 | 7.2 | 7.6841 | 7.8 | 7.9 | 9.2 | 7 | 8.1 | 8.6 | 8.72 | < 0.5 | 9 | 8.2 | [8 empty] | 8.1134
| 1 | 5 | 7.9336 | 7.6 | 8.4 | 4.8747 | 9 | 6.9 | 7.8 | 8.1 | 8.3 | 6.8 | 7.0233 | 7.7 | 7.8 | 9.3 | 6.8 | 8 | 8.9 | 8.72 | < 0.5 | 9.1 | 8.3 | [8 empty] | 127
| 1 | 6 | 7.8035 | 7.4 | 8.6 | 4.6389 | 10 | 6.8 | 8.7 | 8.4 | 8.3 | 7 | 7.0834 | 7.5 | 7.9 | 9.1 | 6.7 | 8 | 9 | 8.44 | < 0.5 | 8.1 | 8.3
| Mean |  | 7.959 | 7.3333 | 8.7 | 5.0639 | 9.1667 | 7.0333 | 8.3833 | 8.2667 | 8.2 | 6.8667 | 7.5964 | 7.7167 | 7.7333 | 9.0833 | 6.9833 | 8.15 | 8.95 | 8.495 |  | 9.05 | 8.15
| Median |  | 7.9389 | 7.35 | 8.65 | 4.7841 | 9 | 7.05 | 8.5 | 8.3 | 8.25 | 6.85 | 7.6723 | 7.7 | 7.75 | 9.1 | 6.95 | 8.1 | 9 | 8.485 |  | 9.05 | 8.25
| Std Dev. |  | 0.18052 | 0.21602 | 0.26077 | 1.5491 | 0.40825 | 0.1633 | 0.33116 | 0.15055 | 0.16733 | 0.23381 | 0.4732 | 0.22286 | 0.1633 | 0.16021 | 0.23166 | 0.16432 | 0.19748 | 0.20878 |  | 0.87121 | 0.38341
| Rel.Std.Dev. |  | 2.268% | 2.946% | 2.997% | 30.59% | 4.454% | 2.322% | 3.95% | 1.821% | 2.041% | 3.405% | 6.229% | 2.888% | 2.112% | 1.764% | 3.317% | 2.016% | 2.207% | 2.458% |  | 9.627% | 4.704%
| PDM3 |  | -1.903% | -9.614% | 7.231% | -37.59% | 12.98% | -13.31% | 3.328% | 1.89% | 1.068% | -15.37% | -6.371% | -4.889% | -4.684% | 11.96% | -13.93% | 0.4517% | 10.31% | 4.704% |  | 11.54% | 0.4517%
| Z-Score (Absolute) |  | 0.23 | 0.99 | 0.67 | 3.76 | 1.24 | 1.36 | 0.28 | 0.14 | 0.06 | 1.56 | 0.67 | 0.53 | 0.51 | 1.14 | 1.42 | 0 | 0.97 | 0.42 | 9.62 | 1.1 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C930: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 285 (Indicative Value 0.184 wt.%) | [28 empty] | Indicative
C | Round | Replicate | Lab | BV Geo | [25 empty] | 1
| No | No | 1.12 | Perth | [25 empty] | wt.%
|  |  | IRC | IRC | [25 empty] | 3
3
| 1 | 1 | 0.19 | 0.18 | [25 empty] | 1
| 1 | 2 | 0.19 | 0.18 | [25 empty] | 9
| 1 | 3 | 0.19 | [26 empty] | 16
| 1 | 4 | 0.19 | [26 empty] | 0.18417
| 1 | 5 | 0.19 | [26 empty] | 15
| 1 | 6 | 0.18
| Mean |  | 0.18833 | 0.18
| Median |  | 0.19 | 0.18
| Std Dev. |  | 0.0040825 | 0
| Rel.Std.Dev. |  | 2.168% | 0%
| PDM3 |  | 2.262% | -2.262%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 285 (Certified Value 0.334 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.03 | 1.04 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.33346 | 0.34 | 0.364 | 0.33 | 0.35 | 0.35 | 0.35 | 0.33 | 0.3 | 0.34 | 0.3 | 0.33 | 0.35 | 0.31817 | 0.36 | 0.34 | 0.35 | 0.3 | 0.33 | 0.33 | 0.32 | 0.33 | 0.39 | 0.34 | 0.31 | 0.22 | 0.36 | 1
| 1 | 2 | 0.31301 | 0.34 | 0.387 | 0.33 | 0.34 | 0.36 | 0.35 | 0.33 | 0.32 | 0.34 | 0.31 | 0.33 | 0.35 | 0.34014 | 0.36 | 0.34 | 0.35 | 0.3 | 0.34 | 0.334 | 0.32 | 0.34 | 0.4 | 0.34 | 0.33 | 0.21 | 0.35 | 11
| 1 | 3 | 0.32842 | 0.34 | 0.361 | 0.33 | 0.35 | 0.34 | 0.34 | 0.33 | 0.33 | 0.34 | 0.31 | 0.33 | 0.34 | 0.34384 | 0.36 | 0.34 | 0.35 | 0.3 | 0.34 | 0.328 | 0.32 | 0.33 | 0.408 | 0.34 | 0.33 | 0.34 | 0.35 | 16
| 1 | 4 | 0.31961 | 0.33 | 0.38 | 0.33 | 0.35 | 0.35 | 0.34 | 0.33 | 0.32 | 0.34 | 0.3 | 0.32 | 0.34 | 0.34378 | 0.36 | 0.34 | 0.35 | 0.3 | 0.33 | 0.322 | 0.32 | 0.33 | 0.4 | 0.34 | 0.3 | 0.21 | 0.35 | 0.33355
| 1 | 5 | 0.32845 | 0.33 | 0.387 | 0.34 | 0.35 | 0.35 | 0.34 | 0.34 | 0.33 | 0.35 | 0.31 | 0.33 | 0.34 | 0.32611 | 0.36 | 0.34 | 0.36 | 0.3 | 0.34 | 0.329 | 0.32 | 0.33 | 0.4 | 0.34 | 0.275 | 0.22 | 0.35 | 129
| 1 | 6 | 0.33172 | 0.34 | 0.382 | 0.33 | 0.35 | 0.35 | 0.32 | 0.33 | 0.32 | 0.35 | 0.3 | 0.33 | 0.34 | 0.32309 | 0.36 | 0.34 | 0.36 | 0.3 | 0.32 | 0.325 | 0.32 | 0.33 | 0.4 | 0.34 | 0.3 | 0.22 | 0.33
| Mean |  | 0.32578 | 0.33667 | 0.37683 | 0.33167 | 0.34833 | 0.35 | 0.34 | 0.33167 | 0.32 | 0.34333 | 0.305 | 0.32833 | 0.34333 | 0.33252 | 0.36 | 0.34 | 0.35333 | 0.3 | 0.33333 | 0.328 | 0.32 | 0.33167 | 0.39967 | 0.34 | 0.3075 | 0.23667 | 0.34833
| Median |  | 0.32844 | 0.34 | 0.381 | 0.33 | 0.35 | 0.35 | 0.34 | 0.33 | 0.32 | 0.34 | 0.305 | 0.33 | 0.34 | 0.33312 | 0.36 | 0.34 | 0.35 | 0.3 | 0.335 | 0.3285 | 0.32 | 0.33 | 0.4 | 0.34 | 0.305 | 0.22 | 0.35
| Std Dev. |  | 0.0078676 | 0.005164 | 0.011479 | 0.0040825 | 0.0040825 | 0.0063246 | 0.010954 | 0.0040825 | 0.010954 | 0.005164 | 0.0054772 | 0.0040825 | 0.005164 | 0.011391 | 6.0809e-17 | 0 | 0.005164 | 0 | 0.008165 | 0.0041473 | 0 | 0.0040825 | 0.0057155 | 0 | 0.020917 | 0.050859 | 0.0098319
| Rel.Std.Dev. |  | 2.415% | 1.534% | 3.046% | 1.231% | 1.172% | 1.807% | 3.222% | 1.231% | 3.423% | 1.504% | 1.796% | 1.243% | 1.504% | 3.426% | 1.689e-14% | 0% | 1.462% | 0% | 2.449% | 1.264% | 0% | 1.231% | 1.43% | 0% | 6.802% | 21.49% | 2.823%
| PDM3 |  | -2.33% | 0.9331% | 12.98% | -0.5659% | 4.431% | 4.93% | 1.932% | -0.5659% | -4.064% | 2.932% | -8.561% | -1.565% | 2.932% | -0.31% | 7.928% | 1.932% | 5.93% | -10.06% | -0.06622% | -1.665% | -4.064% | -0.5659% | 19.82% | 1.932% | -7.811% | -29.05% | 4.431%
| Z-Score (Absolute) |  | 0.51 | 0.22 | 2.93 | 0.11 | 1.01 | 1.12 | 0.45 | 0.11 | 0.9 | 0.67 | 1.91 | 0.34 | 0.67 | 0.05 | 1.8 | 0.45 | 1.35 | 2.25 | 0 | 0.36 | 0.9 | 0.11 | 4.47 | 0.45 | 1.74 | 6.52 | 1.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 285 (Indicative Value 15.07 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 15.07 | 1
| 1 | 2 | 15.07 | 11
| Mean |  | 15.07 | 16
| Median |  | 15.07 | 15.07
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 285 (Indicative Value 270 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 270 | 1
| 1 | 2 | 270 | 12
| Mean |  | 270 | 16
| Median |  | 270 | 270
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 285 (Indicative Value 1020 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1010 | 1
| 1 | 2 | 1030 | 13
| Mean |  | 1020 | 16
| Median |  | 1020 | 1020
| Std Dev. |  | 14.1421 | 19
| Rel.Std.Dev. |  | 1.386%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 285 (Indicative Value 2.73 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.73 | 1
| 1 | 2 | 2.73 | 14
| Mean |  | 2.73 | 16
| Median |  | 2.73 | 2.73
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 285 (Indicative Value 25 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 20 | 15
| Mean |  | 25 | 16
| Median |  | 25 | 25
| Std Dev. |  | 7.0711 | 21
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 285 (Indicative Value 20 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 16
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 285 (Indicative Value 75 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 80 | 17
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 285 (Indicative Value 45 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 50 | 18
| Mean |  | 45 | 16
| Median |  | 45 | 45
| Std Dev. |  | 7.0711 | 24
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 285 (Indicative Value 3.37 wt.%) |  |  | Indicative
Fe | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.37 | 1
| 1 | 2 | 3.36 | 11
| Mean |  | 3.365 | 16
| Median |  | 3.365 | 3.365
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.2101%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 285 (Indicative Value 3.51 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.51 | 1
| 1 | 2 | 3.5 | 12
| Mean |  | 3.505 | 16
| Median |  | 3.505 | 3.505
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.2017%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 285 (Indicative Value 1.61 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.61 | 1
| 1 | 2 | 1.6 | 13
| Mean |  | 1.605 | 16
| Median |  | 1.605 | 1.605
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.4406%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 285 (Indicative Value 0.052 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.051 | 1
| 1 | 2 | 0.053 | 14
| Mean |  | 0.052 | 16
| Median |  | 0.052 | 0.052
| Std Dev. |  | 0.0014142 | 20
| Rel.Std.Dev. |  | 2.72%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 285 (Indicative Value 2.74 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.74 | 1
| 1 | 2 | 2.73 | 15
| Mean |  | 2.735 | 16
| Median |  | 2.735 | 2.735
| Std Dev. |  | 0.0070711 | 21
| Rel.Std.Dev. |  | 0.2585%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 285 (Indicative Value 25 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 30 | 16
| Mean |  | 25 | 16
| Median |  | 25 | 25
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 285 (Indicative Value 0.086 wt.%) |  |  | Indicative
P | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.086 | 1
| 1 | 2 | 0.086 | 17
| Mean |  | 0.086 | 16
| Median |  | 0.086 | 0.086
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 285 (Indicative Value 35 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 40 | 18
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 7.0711 | 24
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 285 (Indicative Value 0.341 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.341 | 11
| Mean |  | 0.3405 | 16
| Median |  | 0.3405 | 0.3405
| Std Dev. |  | 0.00070711 | 17
| Rel.Std.Dev. |  | 0.2077%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 285 (Indicative Value 66.96 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 66.97 | 1
| 1 | 2 | 66.95 | 12
| Mean |  | 66.96 | 16
| Median |  | 66.96 | 66.96
| Std Dev. |  | 0.014142 | 18
| Rel.Std.Dev. |  | 0.02112%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 285 (Indicative Value < 10 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 13
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 19
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 285 (Indicative Value 170 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 160 | 1
| 1 | 2 | 180 | 14
| Mean |  | 170 | 16
| Median |  | 170 | 170
| Std Dev. |  | 14.1421 | 20
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 285 (Indicative Value 0.634 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.633 | 1
| 1 | 2 | 0.635 | 15
| Mean |  | 0.634 | 16
| Median |  | 0.634 | 0.634
| Std Dev. |  | 0.0014142 | 21
| Rel.Std.Dev. |  | 0.2231%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 285 (Indicative Value 75 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 80 | 16
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 7.0711 | 22
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 285 (Indicative Value 80 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 80 | 17
| Mean |  | 80 | 16
| Median |  | 80 | 80
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 285 (Indicative Value 260 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 250 | 1
| 1 | 2 | 270 | 18
| Mean |  | 260 | 16
| Median |  | 260 | 260
| Std Dev. |  | 14.1421 | 24
| Rel.Std.Dev. |  | 5.439%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 285 (Indicative Value 1.42 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.43 | 1
| 1 | 2 | 1.41 | 20
| Mean |  | 1.42 | 16
| Median |  | 1.42 | 1.42
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 0.9959%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 285 (Indicative Value < 0.1 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 22
| Mean |  |  | 16
| Median |  |  | < 0.1
| Std Dev. |  |  | 28
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 285 (Indicative Value 264 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 263 | 1
| 1 | 2 | 264 | 12
| Mean |  | 263.5 | 16
| Median |  | 263.5 | 263.5
| Std Dev. |  | 0.70711 | 29
| Rel.Std.Dev. |  | 0.2684%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 285 (Indicative Value 997 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 995 | 1
| 1 | 2 | 999 | 13
| Mean |  | 997 | 16
| Median |  | 997 | 997
| Std Dev. |  | 2.8284 | 30
| Rel.Std.Dev. |  | 0.2837%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 285 (Indicative Value 2.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 2.8 | 25
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.14142 | 31
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 285 (Indicative Value 0.58 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.56 | 1
| 1 | 2 | 0.6 | 26
| Mean |  | 0.58 | 16
| Median |  | 0.58 | 0.58
| Std Dev. |  | 0.028284 | 32
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 285 (Indicative Value 0.15 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.1 | 27
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0.070711 | 33
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 285 (Indicative Value 78 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 77.5 | 1
| 1 | 2 | 78.7 | 28
| Mean |  | 78.1 | 16
| Median |  | 78.1 | 78.1
| Std Dev. |  | 0.84853 | 34
| Rel.Std.Dev. |  | 1.086%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 285 (Indicative Value 14.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.8 | 1
| 1 | 2 | 14.4 | 16
| Mean |  | 14.1 | 16
| Median |  | 14.1 | 14.1
| Std Dev. |  | 0.42426 | 35
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 285 (Indicative Value 58 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 59 | 1
| 1 | 2 | 56 | 17
| Mean |  | 57.5 | 16
| Median |  | 57.5 | 57.5
| Std Dev. |  | 2.1213 | 36
| Rel.Std.Dev. |  | 3.689%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 285 (Indicative Value 10.5 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.6 | 1
| 1 | 2 | 10.4 | 31
| Mean |  | 10.5 | 16
| Median |  | 10.5 | 10.5
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 285 (Indicative Value 33 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32 | 1
| 1 | 2 | 34 | 18
| Mean |  | 33 | 16
| Median |  | 33 | 33
| Std Dev. |  | 1.4142 | 38
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 285 (Indicative Value 5.91 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.88 | 1
| 1 | 2 | 5.93 | 33
| Mean |  | 5.905 | 16
| Median |  | 5.905 | 5.905
| Std Dev. |  | 0.035355 | 39
| Rel.Std.Dev. |  | 0.5987%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 285 (Indicative Value 3.16 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.22 | 1
| 1 | 2 | 3.1 | 34
| Mean |  | 3.16 | 16
| Median |  | 3.16 | 3.16
| Std Dev. |  | 0.084853 | 40
| Rel.Std.Dev. |  | 2.685%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 285 (Indicative Value 1.44 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.44 | 1
| 1 | 2 | 1.44 | 35
| Mean |  | 1.44 | 16
| Median |  | 1.44 | 1.44
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 285 (Indicative Value 19.9 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.5 | 1
| 1 | 2 | 20.3 | 36
| Mean |  | 19.9 | 16
| Median |  | 19.9 | 19.9
| Std Dev. |  | 0.56569 | 42
| Rel.Std.Dev. |  | 2.843%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 285 (Indicative Value 7.04 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.9 | 1
| 1 | 2 | 7.18 | 37
| Mean |  | 7.04 | 16
| Median |  | 7.04 | 7.04
| Std Dev. |  | 0.19799 | 43
| Rel.Std.Dev. |  | 2.812%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 285 (Indicative Value 1.35 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.35 | 2
| Mean |  | 1.35 | 16
| Median |  | 1.35 | 1.35
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 285 (Indicative Value 6.99 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.03 | 1
| 1 | 2 | 6.95 | 22
| Mean |  | 6.99 | 16
| Median |  | 6.99 | 6.99
| Std Dev. |  | 0.056569 | 28
| Rel.Std.Dev. |  | 0.8093%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 285 (Indicative Value 1.15 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.15 | 1
| 1 | 2 | 1.15 | 23
| Mean |  | 1.15 | 16
| Median |  | 1.15 | 1.15
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 285 (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 24
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 285 (Indicative Value 39.4 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39.1 | 1
| 1 | 2 | 39.7 | 25
| Mean |  | 39.4 | 16
| Median |  | 39.4 | 39.4
| Std Dev. |  | 0.42426 | 31
| Rel.Std.Dev. |  | 1.077%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 285 (Indicative Value 0.4 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 26
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 285 (Indicative Value 0.039 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0388 | 1
| 1 | 2 | 0.0388 | 27
| Mean |  | 0.0388 | 16
| Median |  | 0.0388 | 0.0388
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 285 (Indicative Value 2.7 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.8 | 28
| Mean |  | 2.7 | 16
| Median |  | 2.7 | 2.7
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 285 (Indicative Value 12.7 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 12.8 | 29
| Mean |  | 12.7 | 16
| Median |  | 12.7 | 12.7
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 1.114%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 285 (Indicative Value 37 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.9 | 1
| 1 | 2 | 37.1 | 6
| Mean |  | 37 | 16
| Median |  | 37 | 37
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 0.3822%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 285 (Indicative Value 20 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 16
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 285 (Indicative Value 23 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23 | 1
| 1 | 2 | 23 | 18
| Mean |  | 23 | 16
| Median |  | 23 | 23
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 285 (Indicative Value 9.73 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.63 | 1
| 1 | 2 | 9.83 | 4
| Mean |  | 9.73 | 16
| Median |  | 9.73 | 9.73
| Std Dev. |  | 0.14142 | 39
| Rel.Std.Dev. |  | 1.453%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 285 (Indicative Value 160 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 160 | 1
| 1 | 2 | 160 | 34
| Mean |  | 160 | 16
| Median |  | 160 | 160
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 285 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 2
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 285 (Indicative Value 0.75 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.8 | 36
| Mean |  | 0.75 | 16
| Median |  | 0.75 | 0.75
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 285 (Indicative Value 8.95 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.2 | 1
| 1 | 2 | 8.7 | 37
| Mean |  | 8.95 | 16
| Median |  | 8.95 | 8.95
| Std Dev. |  | 0.35355 | 43
| Rel.Std.Dev. |  | 3.95%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 285 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 38
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 285 (Indicative Value 7.78 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.88 | 1
| 1 | 2 | 7.67 | 4
| Mean |  | 7.775 | 16
| Median |  | 7.775 | 7.775
| Std Dev. |  | 0.14849 | 28
| Rel.Std.Dev. |  | 1.91%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 285 (Indicative Value 4.9 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 5 | 13
| Mean |  | 4.9 | 16
| Median |  | 4.9 | 4.9
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 285 (Indicative Value 155 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 155 | 1
| 1 | 2 | 155 | 14
| Mean |  | 155 | 16
| Median |  | 155 | 155
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 285 (Indicative Value 1.12 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.12 | 1
| 1 | 2 | 1.12 | 25
| Mean |  | 1.12 | 16
| Median |  | 1.12 | 1.12
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 285 (Indicative Value 1.1 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.11 | 1
| 1 | 2 | 1.09 | 26
| Mean |  | 1.1 | 16
| Median |  | 1.1 | 1.1
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 1.286%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 285 (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 27
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 285 (Indicative Value 15.2 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15 | 1
| 1 | 2 | 15.4 | 28
| Mean |  | 15.2 | 16
| Median |  | 15.2 | 15.2
| Std Dev. |  | 0.28284 | 34
| Rel.Std.Dev. |  | 1.861%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 285 (Indicative Value 0.373 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.373 | 1
| 1 | 2 | 0.373 | 29
| Mean |  | 0.373 | 16
| Median |  | 0.373 | 0.373
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 285 (Indicative Value 0.7 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.6 | 30
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 285 (Indicative Value 0.44 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.44 | 1
| 1 | 2 | 0.43 | 5
| Mean |  | 0.435 | 16
| Median |  | 0.435 | 0.435
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 1.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 285 (Indicative Value 4.73 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.87 | 1
| 1 | 2 | 4.59 | 32
| Mean |  | 4.73 | 16
| Median |  | 4.73 | 4.73
| Std Dev. |  | 0.19799 | 38
| Rel.Std.Dev. |  | 4.186%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 285 (Indicative Value 66 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 66 | 1
| 1 | 2 | 66.6 | 16
| Mean |  | 66.3 | 16
| Median |  | 66.3 | 66.3
| Std Dev. |  | 0.42426 | 39
| Rel.Std.Dev. |  | 0.6399%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 285 (Indicative Value 19 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.5 | 1
| 1 | 2 | 18.5 | 34
| Mean |  | 19 | 16
| Median |  | 19 | 19
| Std Dev. |  | 0.70711 | 40
| Rel.Std.Dev. |  | 3.722%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 285 (Indicative Value 30.3 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.3 | 1
| 1 | 2 | 30.3 | 35
| Mean |  | 30.3 | 16
| Median |  | 30.3 | 30.3
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 285 (Indicative Value 2.81 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.72 | 1
| 1 | 2 | 2.89 | 36
| Mean |  | 2.805 | 16
| Median |  | 2.805 | 2.805
| Std Dev. |  | 0.12021 | 42
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 285 (Indicative Value 75 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75 | 1
| 1 | 2 | 75 | 17
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 285 (Indicative Value 249 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 247 | 1
| 1 | 2 | 251 | 18
| Mean |  | 249 | 16
| Median |  | 249 | 249
| Std Dev. |  | 2.8284 | 44
| Rel.Std.Dev. |  | 1.136%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 285 (Indicative Value 0.242 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.26 | 22
| 1 | 3 | 0.24 | 16
| 1 | 4 | 0.24 | 0.24167
| 1 | 5 | 0.28 | 46
| 1 | 6 | 0.21
| Mean |  | 0.24167
| Median |  | 0.24
| Std Dev. |  | 0.025626
| Rel.Std.Dev. |  | 10.6%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 285 (Indicative Value 15.92 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 15.9 | 1
| 1 | 2 | 15.8 | 11
| 1 | 3 | 15.6 | 16
| 1 | 4 | 16 | 15.9167
| 1 | 5 | 16.2 | 47
| 1 | 6 | 16
| Mean |  | 15.9167
| Median |  | 15.95
| Std Dev. |  | 0.20412
| Rel.Std.Dev. |  | 1.282%
| PDM3 |  | -2.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 285 (Indicative Value 1063 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 1050 | 1
| 1 | 2 | 1070 | 13
| 1 | 3 | 1040 | 16
| 1 | 4 | 1060 | 1063.3333
| 1 | 5 | 1080 | 48
| 1 | 6 | 1080
| Mean |  | 1063.3333
| Median |  | 1065
| Std Dev. |  | 16.3299
| Rel.Std.Dev. |  | 1.536%
| PDM3 |  | 2.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 285 (Indicative Value 3.25 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 3.4 | 25
| 1 | 3 | 3.2 | 16
| 1 | 4 | 3.4 | 3.25
| 1 | 5 | 3.4 | 49
| 1 | 6 | 3.1
| Mean |  | 3.25
| Median |  | 3.3
| Std Dev. |  | 0.17607
| Rel.Std.Dev. |  | 5.417%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 285 (Indicative Value 0.73 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.6 | 26
| 1 | 3 | 0.8 | 16
| 1 | 4 | 0.8 | 0.73333
| 1 | 5 | 0.6 | 50
| 1 | 6 | 0.8
| Mean |  | 0.73333
| Median |  | 0.8
| Std Dev. |  | 0.10328
| Rel.Std.Dev. |  | 14.08%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 285 (Indicative Value 2.76 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 2.77 | 1
| 1 | 2 | 2.76 | 14
| 1 | 3 | 2.71 | 16
| 1 | 4 | 2.77 | 2.7617
| 1 | 5 | 2.8 | 51
| 1 | 6 | 2.76
| Mean |  | 2.7617
| Median |  | 2.765
| Std Dev. |  | 0.029269
| Rel.Std.Dev. |  | 1.06%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 285 (Indicative Value 0.27 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.3 | 27
| 1 | 3 | 0.3 | 16
| 1 | 4 | 0.3 | 0.26667
| 1 | 5 | 0.3 | 52
| 1 | 6 | 0.2
| Mean |  | 0.26667
| Median |  | 0.3
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 19.36%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 285 (Indicative Value 80 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 77 | 1
| 1 | 2 | 80 | 28
| 1 | 3 | 76 | 16
| 1 | 4 | 81 | 80
| 1 | 5 | 84 | 53
| 1 | 6 | 82
| Mean |  | 80
| Median |  | 80.5
| Std Dev. |  | 3.0332
| Rel.Std.Dev. |  | 3.791%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 285 (Indicative Value 16.2 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 14.5 | 1
| 1 | 2 | 16.6 | 16
| 1 | 3 | 16.6 | 16
| 1 | 4 | 17.4 | 16.1667
| 1 | 5 | 17.2 | 54
| 1 | 6 | 14.7
| Mean |  | 16.1667
| Median |  | 16.6
| Std Dev. |  | 1.2565
| Rel.Std.Dev. |  | 7.772%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 285 (Indicative Value 45.7 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 47 | 1
| 1 | 2 | 51 | 17
| 1 | 3 | 47 | 16
| 1 | 4 | 45 | 45.6667
| 1 | 5 | 40 | 55
| 1 | 6 | 44
| Mean |  | 45.6667
| Median |  | 46
| Std Dev. |  | 3.6697
| Rel.Std.Dev. |  | 8.036%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 285 (Indicative Value 10.1 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 9.1 | 1
| 1 | 2 | 10.6 | 31
| 1 | 3 | 10.3 | 16
| 1 | 4 | 10.8 | 10.1167
| 1 | 5 | 10.7 | 56
| 1 | 6 | 9.2
| Mean |  | 10.1167
| Median |  | 10.45
| Std Dev. |  | 0.7679
| Rel.Std.Dev. |  | 7.59%
| PDM3 |  | -3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 285 (Indicative Value 32.8 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 31.2 | 1
| 1 | 2 | 32.3 | 18
| 1 | 3 | 35.9 | 16
| 1 | 4 | 34.8 | 32.7833
| 1 | 5 | 33.5 | 57
| 1 | 6 | 29
| Mean |  | 32.7833
| Median |  | 32.9
| Std Dev. |  | 2.5039
| Rel.Std.Dev. |  | 7.638%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 285 (Indicative Value 3.78 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.6 | 1
| 1 | 2 | 3.85 | 33
| 1 | 3 | 3.79 | 16
| 1 | 4 | 3.97 | 3.78
| 1 | 5 | 3.87 | 58
| 1 | 6 | 3.6
| Mean |  | 3.78
| Median |  | 3.82
| Std Dev. |  | 0.151
| Rel.Std.Dev. |  | 3.995%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 285 (Indicative Value 1.49 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.41 | 1
| 1 | 2 | 1.52 | 34
| 1 | 3 | 1.44 | 16
| 1 | 4 | 1.56 | 1.49
| 1 | 5 | 1.57 | 59
| 1 | 6 | 1.44
| Mean |  | 1.49
| Median |  | 1.48
| Std Dev. |  | 0.068702
| Rel.Std.Dev. |  | 4.611%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 285 (Indicative Value 1.67 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.58 | 1
| 1 | 2 | 1.71 | 35
| 1 | 3 | 1.66 | 16
| 1 | 4 | 1.72 | 1.6667
| 1 | 5 | 1.72 | 60
| 1 | 6 | 1.61
| Mean |  | 1.6667
| Median |  | 1.685
| Std Dev. |  | 0.060553
| Rel.Std.Dev. |  | 3.633%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 285 (Indicative Value 4.94 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 4.93 | 1
| 1 | 2 | 4.93 | 55
| 1 | 3 | 4.83 | 16
| 1 | 4 | 4.97 | 4.9367
| 1 | 5 | 5.04 | 61
| 1 | 6 | 4.92
| Mean |  | 4.9367
| Median |  | 4.93
| Std Dev. |  | 0.068605
| Rel.Std.Dev. |  | 1.39%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 285 (Indicative Value 20.9 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 18.6 | 1
| 1 | 2 | 21.8 | 36
| 1 | 3 | 21.4 | 16
| 1 | 4 | 22.6 | 20.9167
| 1 | 5 | 22.1 | 62
| 1 | 6 | 19
| Mean |  | 20.9167
| Median |  | 21.6
| Std Dev. |  | 1.6905
| Rel.Std.Dev. |  | 8.082%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 285 (Indicative Value 5.12 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 5.1 | 37
| 1 | 3 | 5 | 16
| 1 | 4 | 5.4 | 5.1167
| 1 | 5 | 5.5 | 46
| 1 | 6 | 4.9
| Mean |  | 5.1167
| Median |  | 5.05
| Std Dev. |  | 0.27869
| Rel.Std.Dev. |  | 5.447%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 285 (Indicative Value 2.58 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.3 | 1
| 1 | 2 | 3.1 | 22
| 1 | 3 | 2.8 | 16
| 1 | 4 | 2.5 | 2.5833
| 1 | 5 | 2.5 | 47
| 1 | 6 | 2.3
| Mean |  | 2.5833
| Median |  | 2.5
| Std Dev. |  | 0.31252
| Rel.Std.Dev. |  | 12.1%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 285 (Indicative Value 0.54 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.52 | 1
| 1 | 2 | 0.55 | 23
| 1 | 3 | 0.53 | 16
| 1 | 4 | 0.56 | 0.53833
| 1 | 5 | 0.56 | 48
| 1 | 6 | 0.51
| Mean |  | 0.53833
| Median |  | 0.54
| Std Dev. |  | 0.02137
| Rel.Std.Dev. |  | 3.97%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 285 (Indicative Value 3.71 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.7 | 1
| 1 | 2 | 3.71 | 12
| 1 | 3 | 3.66 | 16
| 1 | 4 | 3.72 | 3.705
| 1 | 5 | 3.72 | 49
| 1 | 6 | 3.72
| Mean |  | 3.705
| Median |  | 3.715
| Std Dev. |  | 0.023452
| Rel.Std.Dev. |  | 0.633%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 285 (Indicative Value 39.7 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 39 | 1
| 1 | 2 | 39 | 25
| 1 | 3 | 38 | 16
| 1 | 4 | 41 | 39.6667
| 1 | 5 | 41 | 50
| 1 | 6 | 40
| Mean |  | 39.6667
| Median |  | 39.5
| Std Dev. |  | 1.2111
| Rel.Std.Dev. |  | 3.053%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 285 (Indicative Value 53 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 52 | 1
| 1 | 2 | 53 | 45
| 1 | 3 | 52 | 16
| 1 | 4 | 54 | 53
| 1 | 5 | 54 | 51
| 1 | 6 | 53
| Mean |  | 53
| Median |  | 53
| Std Dev. |  | 0.89443
| Rel.Std.Dev. |  | 1.688%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 285 (Indicative Value 1.67 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.66 | 1
| 1 | 2 | 1.67 | 13
| 1 | 3 | 1.64 | 16
| 1 | 4 | 1.68 | 1.6667
| 1 | 5 | 1.68 | 52
| 1 | 6 | 1.67
| Mean |  | 1.6667
| Median |  | 1.67
| Std Dev. |  | 0.015055
| Rel.Std.Dev. |  | 0.9033%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 285 (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 14
| 1 | 3 | 0.05 | 16
| 1 | 4 | 0.051 | 0.050333
| 1 | 5 | 0.051 | 53
| 1 | 6 | 0.05
| Mean |  | 0.050333
| Median |  | 0.05
| Std Dev. |  | 0.0005164
| Rel.Std.Dev. |  | 1.026%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 285 (Indicative Value 3.28 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3.25 | 28
| 1 | 3 | 3.96 | 16
| 1 | 4 | 3.36 | 3.275
| 1 | 5 | 3.43 | 54
| 1 | 6 | 2.85
| Mean |  | 3.275
| Median |  | 3.305
| Std Dev. |  | 0.42599
| Rel.Std.Dev. |  | 13.01%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 285 (Indicative Value 2.81 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.8 | 15
| 1 | 3 | 2.78 | 16
| 1 | 4 | 2.83 | 2.81
| 1 | 5 | 2.82 | 55
| 1 | 6 | 2.83
| Mean |  | 2.81
| Median |  | 2.81
| Std Dev. |  | 0.02
| Rel.Std.Dev. |  | 0.7117%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 285 (Indicative Value 14.9 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13 | 1
| 1 | 2 | 15.8 | 29
| 1 | 3 | 15.4 | 16
| 1 | 4 | 16 | 14.8833
| 1 | 5 | 15.8 | 56
| 1 | 6 | 13.3
| Mean |  | 14.8833
| Median |  | 15.6
| Std Dev. |  | 1.36
| Rel.Std.Dev. |  | 9.138%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 285 (Indicative Value 37.2 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 34.3 | 1
| 1 | 2 | 37.8 | 6
| 1 | 3 | 36.9 | 16
| 1 | 4 | 39.4 | 37.2167
| 1 | 5 | 39.9 | 57
| 1 | 6 | 35
| Mean |  | 37.2167
| Median |  | 37.35
| Std Dev. |  | 2.2728
| Rel.Std.Dev. |  | 6.107%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 285 (Indicative Value 26.2 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 24 | 1
| 1 | 2 | 26.4 | 16
| 1 | 3 | 28.2 | 16
| 1 | 4 | 27.6 | 26.15
| 1 | 5 | 27.4 | 58
| 1 | 6 | 23.3
| Mean |  | 26.15
| Median |  | 26.9
| Std Dev. |  | 2.0335
| Rel.Std.Dev. |  | 7.776%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 285 (Indicative Value 0.203 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.203 | 1
| 1 | 2 | 0.202 | 53
| 1 | 3 | 0.196 | 16
| 1 | 4 | 0.205 | 0.20317
| 1 | 5 | 0.208 | 59
| 1 | 6 | 0.205
| Mean |  | 0.20317
| Median |  | 0.204
| Std Dev. |  | 0.0040702
| Rel.Std.Dev. |  | 2.003%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 285 (Indicative Value 24.9 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 23.5 | 1
| 1 | 2 | 25.2 | 18
| 1 | 3 | 25.7 | 16
| 1 | 4 | 25.9 | 24.9167
| 1 | 5 | 25.6 | 60
| 1 | 6 | 23.6
| Mean |  | 24.9167
| Median |  | 25.4
| Std Dev. |  | 1.0834
| Rel.Std.Dev. |  | 4.348%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 285 (Indicative Value 9.52 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 8.6 | 1
| 1 | 2 | 9.6 | 4
| 1 | 3 | 9.5 | 16
| 1 | 4 | 10.2 | 9.5167
| 1 | 5 | 10.3 | 61
| 1 | 6 | 8.9
| Mean |  | 9.5167
| Median |  | 9.55
| Std Dev. |  | 0.67946
| Rel.Std.Dev. |  | 7.14%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 285 (Indicative Value 216 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 197 | 1
| 1 | 2 | 216 | 34
| 1 | 3 | 219 | 16
| 1 | 4 | 233 | 215.6667
| 1 | 5 | 234 | 62
| 1 | 6 | 195
| Mean |  | 215.6667
| Median |  | 217.5
| Std Dev. |  | 16.8721
| Rel.Std.Dev. |  | 7.823%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 285 (Indicative Value 0.376 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.386 | 1
| 1 | 2 | 0.391 | 11
| 1 | 3 | 0.354 | 16
| 1 | 4 | 0.362 | 0.376
| 1 | 5 | 0.381 | 46
| 1 | 6 | 0.382
| Mean |  | 0.376
| Median |  | 0.3815
| Std Dev. |  | 0.014601
| Rel.Std.Dev. |  | 3.883%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 285 (Indicative Value 9.47 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 8.5 | 1
| 1 | 2 | 9.7 | 37
| 1 | 3 | 9.6 | 16
| 1 | 4 | 10.2 | 9.4667
| 1 | 5 | 10 | 47
| 1 | 6 | 8.8
| Mean |  | 9.4667
| Median |  | 9.65
| Std Dev. |  | 0.67429
| Rel.Std.Dev. |  | 7.123%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 285 (Indicative Value 7.22 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 7.2 | 4
| 1 | 3 | 7.2 | 16
| 1 | 4 | 7.7 | 7.2167
| 1 | 5 | 7.8 | 48
| 1 | 6 | 6.8
| Mean |  | 7.2167
| Median |  | 7.2
| Std Dev. |  | 0.47504
| Rel.Std.Dev. |  | 6.583%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 285 (Indicative Value 4.61 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.13 | 1
| 1 | 2 | 4.81 | 13
| 1 | 3 | 4.78 | 16
| 1 | 4 | 4.93 | 4.61
| 1 | 5 | 4.87 | 49
| 1 | 6 | 4.14
| Mean |  | 4.61
| Median |  | 4.795
| Std Dev. |  | 0.37154
| Rel.Std.Dev. |  | 8.059%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 285 (Indicative Value 164 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 163 | 14
| 1 | 3 | 162 | 16
| 1 | 4 | 166 | 164
| 1 | 5 | 164 | 50
| 1 | 6 | 165
| Mean |  | 164
| Median |  | 164
| Std Dev. |  | 1.4142
| Rel.Std.Dev. |  | 0.8623%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 285 (Indicative Value 0.94 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.93 | 1
| 1 | 2 | 0.93 | 25
| 1 | 3 | 0.92 | 16
| 1 | 4 | 0.96 | 0.935
| 1 | 5 | 0.95 | 51
| 1 | 6 | 0.92
| Mean |  | 0.935
| Median |  | 0.93
| Std Dev. |  | 0.016432
| Rel.Std.Dev. |  | 1.757%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 285 (Indicative Value 0.69 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.66 | 1
| 1 | 2 | 0.7 | 26
| 1 | 3 | 0.68 | 16
| 1 | 4 | 0.72 | 0.69167
| 1 | 5 | 0.73 | 52
| 1 | 6 | 0.66
| Mean |  | 0.69167
| Median |  | 0.69
| Std Dev. |  | 0.029944
| Rel.Std.Dev. |  | 4.329%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 285 (Indicative Value 17.2 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 15.4 | 1
| 1 | 2 | 17.6 | 28
| 1 | 3 | 17.6 | 16
| 1 | 4 | 18.6 | 17.1833
| 1 | 5 | 18.6 | 53
| 1 | 6 | 15.3
| Mean |  | 17.1833
| Median |  | 17.6
| Std Dev. |  | 1.4892
| Rel.Std.Dev. |  | 8.666%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 285 (Indicative Value 0.663 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 1.29 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.661 | 1
| 1 | 2 | 0.662 | 15
| 1 | 3 | 0.652 | 16
| 1 | 4 | 0.669 | 0.6625
| 1 | 5 | 0.666 | 54
| 1 | 6 | 0.665
| Mean |  | 0.6625
| Median |  | 0.6635
| Std Dev. |  | 0.0058907
| Rel.Std.Dev. |  | 0.8892%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 285 (Indicative Value 3.71 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.74 | 1
| 1 | 2 | 3.64 | 32
| 1 | 3 | 11.2 | 16
| 1 | 4 | 3.58 | 3.706
| 1 | 5 | 4.26 | 55
| 1 | 6 | 3.31
| Mean |  | 4.955
| Median |  | 3.69
| Std Dev. |  | 3.0752
| Rel.Std.Dev. |  | 62.06%
| PDM3 |  | 33.7%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 285 (Indicative Value 68 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 67.8 | 1
| 1 | 2 | 67.9 | 16
| 1 | 3 | 64.7 | 16
| 1 | 4 | 71.4 | 68.35
| 1 | 5 | 67 | 56
| 1 | 6 | 71.3
| Mean |  | 68.35
| Median |  | 67.85
| Std Dev. |  | 2.5944
| Rel.Std.Dev. |  | 3.796%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 285 (Indicative Value 20.8 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 21.3 | 1
| 1 | 2 | 19.7 | 34
| 1 | 3 | 20.3 | 16
| 1 | 4 | 21.1 | 20.8
| 1 | 5 | 21.3 | 57
| 1 | 6 | 21.1
| Mean |  | 20.8
| Median |  | 21.1
| Std Dev. |  | 0.65422
| Rel.Std.Dev. |  | 3.145%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 285 (Indicative Value 15.7 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13.9 | 1
| 1 | 2 | 16.8 | 35
| 1 | 3 | 16.3 | 16
| 1 | 4 | 16.7 | 15.7167
| 1 | 5 | 16.6 | 58
| 1 | 6 | 14
| Mean |  | 15.7167
| Median |  | 16.45
| Std Dev. |  | 1.379
| Rel.Std.Dev. |  | 8.774%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 285 (Indicative Value 1.04 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.99 | 1
| 1 | 2 | 1.08 | 36
| 1 | 3 | 1.06 | 16
| 1 | 4 | 1.05 | 1.0383
| 1 | 5 | 1.06 | 59
| 1 | 6 | 0.99
| Mean |  | 1.0383
| Median |  | 1.055
| Std Dev. |  | 0.038687
| Rel.Std.Dev. |  | 3.726%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 285 (Indicative Value 92 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 81 | 1
| 1 | 2 | 92 | 17
| 1 | 3 | 92 | 16
| 1 | 4 | 106 | 91.6667
| 1 | 5 | 96 | 60
| 1 | 6 | 83
| Mean |  | 91.6667
| Median |  | 92
| Std Dev. |  | 9.0921
| Rel.Std.Dev. |  | 9.919%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 285 (Indicative Value 58 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 1.29 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 54 | 1
| 1 | 2 | 56 | 18
| 1 | 3 | 73 | 16
| 1 | 4 | 57 | 58.3333
| 1 | 5 | 56 | 61
| 1 | 6 | 54
| Mean |  | 58.3333
| Median |  | 56
| Std Dev. |  | 7.2847
| Rel.Std.Dev. |  | 12.49%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.