## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 255c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 4.1722 | 4.1353 | 4.2092 | 4.1595 | 4.185
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 4.1303 | 4.1017 | 4.1589 | 4.1165 | 4.1441
Cyanide Leach
Au, Gold (ppm) | 3.9653 | 3.8253 | 4.1052 | 3.9606 | 3.9699
PhotonAssay
Au, Gold (ppm) | 4.2308 | 4.1784 | 4.2832 | 4.227 | 4.2346
4-Acid Digestion
Ag, Silver (ppm) | 1.107 | 1.0102 | 1.2038 | 1.0516 | 1.1624
Al, Aluminium (wt.%) | 6.3074 | 6.1855 | 6.4292 | 6.203 | 6.4118
As, Arsenic (ppm) | 173.5351 | 168.9257 | 178.1445 | 168.9662 | 178.104
Ba, Barium (ppm) | 444.7703 | 434.4882 | 455.0524 | 437.3691 | 452.1715
Be, Beryllium (ppm) | 2.1034 | 1.9672 | 2.2396 | 2.0384 | 2.1685
Bi, Bismuth (ppm) | 3.1584 | 2.9453 | 3.3716 | 2.9766 | 3.3402
Ca, Calcium (wt.%) | 1.2642 | 1.2189 | 1.3095 | 1.2388 | 1.2896
Cd, Cadmium (ppm) | 0.094642 | 0.076214 | 0.11307 | 0.085988 | 0.1033
Ce, Cerium (ppm) | 73.3982 | 70.9914 | 75.8051 | 71.8139 | 74.9826
Co, Cobalt (ppm) | 20.9693 | 20.0651 | 21.8735 | 20.5143 | 21.4243
Cr, Chromium (ppm) | 168.6707 | 160.4596 | 176.8817 | 164.4904 | 172.8509
Cs, Caesium (ppm) | 3.8137 | 3.6518 | 3.9757 | 3.7077 | 3.9198
Cu, Copper (ppm) | 58.5061 | 56.015 | 60.9972 | 56.9755 | 60.0367
Dy, Dysprosium (ppm) | 3.584 | 3.3497 | 3.8183 | 3.4731 | 3.6949
Er, Erbium (ppm) | 1.6869 | 1.5661 | 1.8076 | 1.6263 | 1.7474
Eu, Europium (ppm) | 1.4858 | 1.3338 | 1.6379 | 1.4301 | 1.5415
Fe, Iron (wt.%) | 4.4003 | 4.2936 | 4.5071 | 4.3297 | 4.4709
Ga, Gallium (ppm) | 17.6579 | 16.6766 | 18.6391 | 17.2536 | 18.0621
Gd, Gadolinium (ppm) | 5.0515 | 4.7307 | 5.3723 | 4.8935 | 5.2096
Hf, Hafnium (ppm) | 4.3858 | 4.1123 | 4.6593 | 4.2044 | 4.5672
Ho, Holmium (ppm) | 0.64202 | 0.58361 | 0.70043 | 0.60789 | 0.67615
In, Indium (ppm) | 0.065344 | 0.05669 | 0.073998 | 0.061421 | 0.069267
K, Potassium (wt.%) | 1.5804 | 1.5338 | 1.627 | 1.5541 | 1.6068
La, Lanthanum (ppm) | 39.606 | 38.0354 | 41.1766 | 38.8681 | 40.3439
Li, Lithium (ppm) | 26.5387 | 25.6019 | 27.4754 | 25.8333 | 27.244
Lu, Lutetium (ppm) | 0.20483 | 0.17565 | 0.23402 | 0.17433 | 0.23534
Mg, Magnesium (wt.%) | 1.6157 | 1.5643 | 1.6671 | 1.5785 | 1.6529
Mn, Manganese (wt.%) | 0.04255 | 0.041445 | 0.043656 | 0.041887 | 0.043214
Mo, Molybdenum (ppm) | 4.8536 | 4.5771 | 5.1301 | 4.6774 | 5.0299
Na, Sodium (wt.%) | 0.81126 | 0.78577 | 0.83675 | 0.79381 | 0.8287
Nb, Niobium (ppm) | 26.3482 | 25.1045 | 27.5918 | 25.4241 | 27.2722
Nd, Neodymium (ppm) | 32.4699 | 31.0037 | 33.9361 | 31.7802 | 33.1595
Ni, Nickel (ppm) | 95.2758 | 92.197 | 98.3546 | 93.5162 | 97.0354
P, Phosphorus (wt.%) | 0.077298 | 0.074407 | 0.080189 | 0.075423 | 0.079173
Pb, Lead (ppm) | 15.1754 | 14.5317 | 15.8191 | 14.6697 | 15.681
Pr, Praseodymium (ppm) | 8.6683 | 8.1359 | 9.2008 | 8.4244 | 8.9123
Rb, Rubidium (ppm) | 78.7047 | 73.9929 | 83.4164 | 76.8269 | 80.5825
Re, Rhenium (ppm) | 0.0014367 | 0.00056403 | 0.0023093 | IND | IND
S, Sulphur (wt.%) | 0.038779 | 0.035757 | 0.041801 | 0.037185 | 0.040373
Sb, Antimony (ppm) | 8.0054 | 7.6292 | 8.3816 | 7.7066 | 8.3042
Sc, Scandium (ppm) | 14.6472 | 13.9498 | 15.3446 | 14.2731 | 15.0213
Sm, Samarium (ppm) | 6.1999 | 6.0112 | 6.3886 | 5.9977 | 6.4021
Sn, Tin (ppm) | 5.6548 | 5.4432 | 5.8664 | 5.4162 | 5.8934
Sr, Strontium (ppm) | 240.9621 | 230.0102 | 251.9139 | 235.3041 | 246.62
Ta, Tantalum (ppm) | 1.755 | 1.6111 | 1.8989 | 1.6688 | 1.8412
Tb, Terbium (ppm) | 0.68982 | 0.63634 | 0.7433 | 0.66995 | 0.70969
Te, Tellurium (ppm) | 0.17905 | 0.12795 | 0.23014 | IND | IND
Th, Thorium (ppm) | 9.8962 | 9.3836 | 10.4087 | 9.6511 | 10.1413
Ti, Titanium (wt.%) | 0.51492 | 0.49745 | 0.53239 | 0.50472 | 0.52513
Tl, Thallium (ppm) | 0.45698 | 0.43123 | 0.48273 | 0.43652 | 0.47744
Tm, Thulium (ppm) | 0.2127 | 0.19227 | 0.23313 | 0.19764 | 0.22776
U, Uranium (ppm) | 1.8168 | 1.7082 | 1.9254 | 1.7697 | 1.8638
V, Vanadium (ppm) | 98.0578 | 95.4075 | 100.708 | 96.2284 | 99.8872
W, Tungsten (ppm) | 33.2382 | 30.5136 | 35.9627 | 31.8863 | 34.5901
Y, Yttrium (ppm) | 16.1353 | 15.371 | 16.8997 | 15.6904 | 16.5803
Yb, Ytterbium (ppm) | 1.4168 | 1.2988 | 1.5349 | 1.3498 | 1.4838
Zn, Zinc (ppm) | 83.7642 | 80.952 | 86.5764 | 82.2181 | 85.3103
Zr, Zirconium (ppm) | 180.828 | 173.7758 | 187.8803 | 175.2463 | 186.4098
Aqua Regia Digestion
Ag, Silver (ppm) | 1.0415 | 0.98701 | 1.0959 | 1.009 | 1.0739
Al, Aluminium (wt.%) | 1.8413 | 1.7072 | 1.9754 | 1.7919 | 1.8907
As, Arsenic (ppm) | 153.4735 | 147.2424 | 159.7046 | 148.9096 | 158.0374
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 94.2241 | 89.7515 | 98.6967 | 91.6298 | 96.8183
Be, Beryllium (ppm) | 1.0252 | 0.97982 | 1.0705 | 0.99209 | 1.0582
Bi, Bismuth (ppm) | 2.665 | 2.4915 | 2.8386 | 2.5882 | 2.7419
Ca, Calcium (wt.%) | 0.47371 | 0.4444 | 0.50303 | 0.45932 | 0.48811
Cd, Cadmium (ppm) | 0.081041 | 0.063583 | 0.098499 | IND | IND
Ce, Cerium (ppm) | 43.3601 | 41.4069 | 45.3133 | 42.4054 | 44.3148
Co, Cobalt (ppm) | 18.0346 | 17.4004 | 18.6689 | 17.497 | 18.5723
Cr, Chromium (ppm) | 107.5263 | 100.7433 | 114.3092 | 104.1469 | 110.9056
Cs, Caesium (ppm) | 0.99336 | 0.89956 | 1.0872 | 0.96029 | 1.0264
Cu, Copper (ppm) | 55.0235 | 52.7479 | 57.2992 | 53.5118 | 56.5352
Fe, Iron (wt.%) | 3.4883 | 3.3383 | 3.6383 | 3.3995 | 3.5771
Ga, Gallium (ppm) | 6.1299 | 5.6753 | 6.5845 | 5.9123 | 6.3474
Ge, Germanium (ppm) | 0.11583 | 0.079917 | 0.15175 | IND | IND
In, Indium (ppm) | 0.0225 | 0.018635 | 0.026365 | 0.020093 | 0.024907
K, Potassium (wt.%) | 0.28601 | 0.26435 | 0.30767 | 0.27631 | 0.29571
La, Lanthanum (ppm) | 23.7222 | 22.5392 | 24.9051 | 22.9611 | 24.4833
Li, Lithium (ppm) | 10.4311 | 9.953 | 10.9092 | 10.1102 | 10.7521
Mg, Magnesium (wt.%) | 1.1627 | 1.1199 | 1.2055 | 1.1297 | 1.1957
Mn, Manganese (wt.%) | 0.031772 | 0.030576 | 0.032967 | 0.031153 | 0.03239
Mo, Molybdenum (ppm) | 3.9427 | 3.6741 | 4.2114 | 3.8272 | 4.0583
Na, Sodium (wt.%) | 0.26762 | 0.25171 | 0.28354 | 0.25776 | 0.27749
Ni, Nickel (ppm) | 83.9296 | 80.3968 | 87.4625 | 81.9606 | 85.8987
P, Phosphorus (wt.%) | 0.048376 | 0.046528 | 0.050223 | 0.047378 | 0.049373
Pb, Lead (ppm) | 11.4625 | 10.7934 | 12.1315 | 11.0594 | 11.8655
Rb, Rubidium (ppm) | 14.5777 | 13.6639 | 15.4915 | 14.059 | 15.0964
Re, Rhenium (ppm) | 0.0008125 | 0.00045206 | 0.0011729 | IND | IND
S, Sulphur (wt.%) | 0.03421 | 0.031848 | 0.036572 | 0.032499 | 0.035921
Sb, Antimony (ppm) | 5.2302 | 4.8459 | 5.6146 | 4.9627 | 5.4978
Sc, Scandium (ppm) | 5.1955 | 4.8785 | 5.5125 | 5.0132 | 5.3778
Sn, Tin (ppm) | 1.3707 | 1.2218 | 1.5196 | 1.2816 | 1.4598
Sr, Strontium (ppm) | 52.4683 | 50.581 | 54.3556 | 50.7667 | 54.1698
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.45396 | 0.40463 | 0.5033 | IND | IND
Te, Tellurium (ppm) | 0.11747 | 0.085021 | 0.14993 | IND | IND
Th, Thorium (ppm) | 6.6701 | 6.3101 | 7.0301 | 6.4692 | 6.8709
Ti, Titanium (wt.%) | 0.11412 | 0.10475 | 0.12349 | 0.10991 | 0.11833
Tl, Thallium (ppm) | 0.12453 | 0.1095 | 0.13956 | IND | IND
U, Uranium (ppm) | 1.0067 | 0.9551 | 1.0584 | 0.97915 | 1.0343
V, Vanadium (ppm) | 44.8323 | 42.5008 | 47.1637 | 43.4426 | 46.2219
W, Tungsten (ppm) | 10.1128 | 8.6676 | 11.5579 | 9.5666 | 10.6589
Y, Yttrium (ppm) | 9.0862 | 8.696 | 9.4764 | 8.851 | 9.3214
Yb, Ytterbium (ppm) | 0.68075 | 0.56506 | 0.79645 | IND | IND
Zn, Zinc (ppm) | 58.7129 | 55.6158 | 61.81 | 57.1806 | 60.2452
Zr, Zirconium (ppm) | 37.3243 | 32.9622 | 41.6865 | 35.3034 | 39.3453
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 255c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
Ge | ppm | 0.15656 | Hg | ppm | 0.23617 | Se | ppm | 0.93303
Aqua Regia Digestion
Dy | ppm | 2.28 | Ho | ppm | 0.36195 | Pt | ppb | < 5
Er | ppm | 0.87966 | Lu | ppm | 0.097283 | Se | ppm | 0.43651
Eu | ppm | 0.95656 | Nb | ppm | 0.49362 | Sm | ppm | 4.0668
Gd | ppm | 3.395 | Nd | ppm | 22.8353 | Tm | ppm | 0.10957
Hf | ppm | 0.62079 | Pd | ppb | < 10
Hg | ppm | 0.088875 | Pr | ppm | 5.8877
Borate Fusion XRF
Al2O3 | wt.% | 12.305 | MgO | wt.% | 2.785 | SiO2 | wt.% | 69.025
CaO | wt.% | 1.7 | MnO | wt.% | 0.056 | SO3 | wt.% | 0.092
Fe2O3 | wt.% | 6.385 | Na2O | wt.% | 1.125 | TiO2 | wt.% | 0.932
K2O | wt.% | 1.92 | P2O5 | wt.% | 0.1785
Thermogravimetry
LOI1000 | wt.% | 3.265
Infrared Combustion
C | wt.% | 0.105 | S | wt.% | 0.02
Laser Ablation ICP-MS
Ag | ppm | 1.15 | Hf | ppm | 6.425 | Sm | ppm | 6.41
As | ppm | 167 | Ho | ppm | 0.87 | Sn | ppm | 7.7
Ba | ppm | 442.5 | In | ppm | 0.05 | Sr | ppm | 237
Be | ppm | 2.2 | La | ppm | 39.95 | Ta | ppm | 1.98
Bi | ppm | 3.34 | Lu | ppm | 0.295 | Tb | ppm | 0.82
Cd | ppm | 0.1 | Mn | wt.% | 0.0441 | Te | ppm | < 0.2
Ce | ppm | 71.5 | Mo | ppm | 4.8 | Th | ppm | 10.2
Co | ppm | 21.65 | Nb | ppm | 28.55 | Ti | wt.% | 0.571
Cr | ppm | 195 | Nd | ppm | 33.65 | Tl | ppm | 0.4
Cs | ppm | 3.76 | Ni | ppm | 103 | Tm | ppm | 0.32
Cu | ppm | 63 | Pb | ppm | 16 | U | ppm | 1.97
Dy | ppm | 4.47 | Pr | ppm | 9.15 | V | ppm | 104
Er | ppm | 2.23 | Rb | ppm | 77.9 | W | ppm | 35.25
Eu | ppm | 1.52 | Re | ppm | < 0.01 | Y | ppm | 21.85
Ga | ppm | 17.5 | Sb | ppm | 8.55 | Yb | ppm | 2.095
Gd | ppm | 5.285 | Sc | ppm | 14.35 | Zn | ppm | 85
Ge | ppm | 1.325 | Se | ppm | < 5 | Zr | ppm | 246.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 255c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 4.1722 | 0.084924 | 4.0024 | 4.3421 | 3.9175 | 4.427 | 2.035% | 4.071% | 6.106% | 3.9636 | 4.3809
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 4.1303 | 0.083916 | 3.9625 | 4.2981 | 3.8786 | 4.382 | 2.032% | 4.063% | 6.095% | 3.9238 | 4.3368
Cyanide Leach
Au, ppm | 3.9653 | 0.27217 | 3.4209 | 4.5096 | 3.1487 | 4.7818 | 6.864% | 13.73% | 20.59% | 3.767 | 4.1635
PhotonAssay
Au, ppm | 4.2308 | 0.12275 | 3.9853 | 4.4763 | 3.8625 | 4.599 | 2.901% | 5.803% | 8.704% | 4.0192 | 4.4423
4-Acid Digestion
Ag, ppm | 1.107 | 0.098305 | 0.91036 | 1.3036 | 0.81205 | 1.4019 | 8.881% | 17.76% | 26.64% | 1.0516 | 1.1623
Al, wt.% | 6.3074 | 0.10017 | 6.107 | 6.5077 | 6.0069 | 6.6079 | 1.588% | 3.176% | 4.764% | 5.992 | 6.6227
As, ppm | 173.5351 | 3.3588 | 166.8175 | 180.2528 | 163.4586 | 183.6116 | 1.936% | 3.871% | 5.807% | 164.8584 | 182.2119
Ba, ppm | 444.7703 | 10.9114 | 422.9475 | 466.5931 | 412.0361 | 477.5045 | 2.453% | 4.907% | 7.36% | 422.5318 | 467.0088
Be, ppm | 2.1034 | 0.11628 | 1.8709 | 2.336 | 1.7546 | 2.4522 | 5.528% | 11.06% | 16.58% | 1.9982 | 2.2086
Bi, ppm | 3.1584 | 0.26591 | 2.6266 | 3.6903 | 2.3607 | 3.9562 | 8.419% | 16.84% | 25.26% | 3.0005 | 3.3163
Ca, wt.% | 1.2642 | 0.033889 | 1.1964 | 1.332 | 1.1625 | 1.3658 | 2.681% | 5.361% | 8.042% | 1.201 | 1.3274
Cd, ppm | 0.094642 | 0.015756 | 0.06313 | 0.12615 | 0.047374 | 0.14191 | 16.65% | 33.3% | 49.94% | 0.08991 | 0.099374
Ce, ppm | 73.3982 | 1.8706 | 69.657 | 77.1394 | 67.7864 | 79.01 | 2.549% | 5.097% | 7.646% | 69.7283 | 77.0681
Co, ppm | 20.9693 | 0.76864 | 19.432 | 22.5066 | 18.6634 | 23.2752 | 3.666% | 7.331% | 11% | 19.9208 | 22.0178
Cr, ppm | 168.6707 | 11.4508 | 145.769 | 191.5723 | 134.3182 | 203.0232 | 6.789% | 13.58% | 20.37% | 160.2371 | 177.1042
Cs, ppm | 3.8137 | 0.14545 | 3.5228 | 4.1046 | 3.3774 | 4.2501 | 3.814% | 7.628% | 11.44% | 3.623 | 4.0044
Cu, ppm | 58.5061 | 3.4313 | 51.6434 | 65.3688 | 48.2121 | 68.8001 | 5.865% | 11.73% | 17.59% | 55.5808 | 61.4314
Dy, ppm | 3.584 | 0.17161 | 3.2408 | 3.9272 | 3.0692 | 4.0988 | 4.788% | 9.576% | 14.36% | 3.4048 | 3.7632
Er, ppm | 1.6869 | 0.066675 | 1.5535 | 1.8202 | 1.4868 | 1.8869 | 3.953% | 7.905% | 11.86% | 1.6025 | 1.7712
Eu, ppm | 1.4858 | 0.088606 | 1.3086 | 1.6631 | 1.22 | 1.7517 | 5.963% | 11.93% | 17.89% | 1.4115 | 1.5601
Fe, wt.% | 4.4003 | 0.085588 | 4.2291 | 4.5715 | 4.1436 | 4.6571 | 1.945% | 3.89% | 5.835% | 4.1803 | 4.6203
Ga, ppm | 17.6579 | 1.3321 | 14.9937 | 20.3221 | 13.6615 | 21.6542 | 7.544% | 15.09% | 22.63% | 16.775 | 18.5408
Gd, ppm | 5.0515 | 0.20618 | 4.6392 | 5.4639 | 4.433 | 5.6701 | 4.082% | 8.163% | 12.24% | 4.799 | 5.3041
Hf, ppm | 4.3858 | 0.29798 | 3.7899 | 4.9818 | 3.4919 | 5.2797 | 6.794% | 13.59% | 20.38% | 4.1665 | 4.6051
Ho, ppm | 0.64202 | 0.035108 | 0.5718 | 0.71224 | 0.5367 | 0.74734 | 5.468% | 10.94% | 16.41% | 0.60992 | 0.67412
In, ppm | 0.065344 | 0.0069032 | 0.051538 | 0.07915 | 0.044634 | 0.086054 | 10.56% | 21.13% | 31.69% | 0.062077 | 0.068611
K, wt.% | 1.5804 | 0.029802 | 1.5208 | 1.64 | 1.491 | 1.6698 | 1.886% | 3.771% | 5.657% | 1.5014 | 1.6594
La, ppm | 39.606 | 1.358 | 36.8899 | 42.3221 | 35.5319 | 43.6802 | 3.429% | 6.858% | 10.29% | 37.6257 | 41.5863
Li, ppm | 26.5387 | 1.1812 | 24.1762 | 28.9012 | 22.9949 | 30.0824 | 4.451% | 8.902% | 13.35% | 25.2117 | 27.8656
Lu, ppm | 0.20483 | 0.020178 | 0.16448 | 0.24519 | 0.1443 | 0.26537 | 9.851% | 19.7% | 29.55% | 0.19459 | 0.21508
Mg, wt.% | 1.6157 | 0.031788 | 1.5521 | 1.6793 | 1.5203 | 1.7111 | 1.967% | 3.935% | 5.902% | 1.5349 | 1.6965
Mn, wt.% | 0.04255 | 0.0012162 | 0.040118 | 0.044983 | 0.038902 | 0.046199 | 2.858% | 5.717% | 8.575% | 0.040423 | 0.044678
Mo, ppm | 4.8536 | 0.2877 | 4.2782 | 5.429 | 3.9905 | 5.7167 | 5.928% | 11.86% | 17.78% | 4.6109 | 5.0963
Na, wt.% | 0.81126 | 0.020692 | 0.76987 | 0.85264 | 0.74918 | 0.87333 | 2.551% | 5.101% | 7.652% | 0.77069 | 0.85182
Nb, ppm | 26.3482 | 1.3028 | 23.7427 | 28.9537 | 22.4399 | 30.2564 | 4.944% | 9.889% | 14.83% | 25.0308 | 27.6656
Nd, ppm | 32.4699 | 0.74325 | 30.9834 | 33.9564 | 30.2401 | 34.6996 | 2.289% | 4.578% | 6.867% | 30.8464 | 34.0934
Ni, ppm | 95.2758 | 3.3912 | 88.4933 | 102.0583 | 85.1021 | 105.4495 | 3.559% | 7.119% | 10.68% | 90.512 | 100.0396
P, wt.% | 0.077298 | 0.0026071 | 0.072084 | 0.082512 | 0.069477 | 0.085119 | 3.373% | 6.746% | 10.12% | 0.073433 | 0.081163
Pb, ppm | 15.1754 | 0.77178 | 13.6318 | 16.7189 | 12.86 | 17.4907 | 5.086% | 10.17% | 15.26% | 14.4166 | 15.9341
Pr, ppm | 8.6683 | 0.33547 | 7.9974 | 9.3393 | 7.6619 | 9.6748 | 3.87% | 7.74% | 11.61% | 8.2349 | 9.1018
Rb, ppm | 78.7047 | 5.2384 | 68.228 | 89.1814 | 62.9896 | 94.4198 | 6.656% | 13.31% | 19.97% | 74.7695 | 82.6399
Re, ppm | 0.0014367 | 0.00068741 | 6.1849e-05 | 0.0028115 | 0 | 0.0034989 | 47.85% | 95.69% | 143.5% | 0.0013648 | 0.0015085
S, wt.% | 0.038779 | 0.0049627 | 0.028854 | 0.048705 | 0.023891 | 0.053667 | 12.8% | 25.59% | 38.39% | 0.03684 | 0.040718
Sb, ppm | 8.0054 | 0.50888 | 6.9876 | 9.0231 | 6.4787 | 9.532 | 6.357% | 12.71% | 19.07% | 7.6051 | 8.4056
Sc, ppm | 14.6472 | 0.368 | 13.9112 | 15.3832 | 13.5432 | 15.7512 | 2.512% | 5.025% | 7.537% | 13.9148 | 15.3796
Sm, ppm | 6.1999 | 0.13887 | 5.9222 | 6.4776 | 5.7833 | 6.6165 | 2.24% | 4.48% | 6.72% | 5.8899 | 6.5099
Sn, ppm | 5.6548 | 0.27484 | 5.1051 | 6.2045 | 4.8303 | 6.4793 | 4.86% | 9.721% | 14.58% | 5.3721 | 5.9375
Sr, ppm | 240.9621 | 10.9553 | 219.0515 | 262.8726 | 208.0963 | 273.8279 | 4.546% | 9.093% | 13.64% | 228.914 | 253.0102
Ta, ppm | 1.755 | 0.15519 | 1.4446 | 2.0654 | 1.2894 | 2.2206 | 8.843% | 17.69% | 26.53% | 1.6673 | 1.8428
Tb, ppm | 0.68982 | 0.030914 | 0.62799 | 0.75165 | 0.59708 | 0.78256 | 4.481% | 8.963% | 13.44% | 0.65533 | 0.72431
Te, ppm | 0.17905 | 0.026868 | 0.12531 | 0.23278 | 0.098442 | 0.25965 | 15.01% | 30.01% | 45.02% | 0.17009 | 0.188
Th, ppm | 9.8962 | 0.68809 | 8.52 | 11.2724 | 7.8319 | 11.9604 | 6.953% | 13.91% | 20.86% | 9.4014 | 10.391
Ti, wt.% | 0.51492 | 0.015111 | 0.4847 | 0.54514 | 0.46959 | 0.56026 | 2.935% | 5.869% | 8.804% | 0.48918 | 0.54067
Tl, ppm | 0.45698 | 0.021905 | 0.41317 | 0.50079 | 0.39127 | 0.5227 | 4.793% | 9.587% | 14.38% | 0.43413 | 0.47983
Tm, ppm | 0.2127 | 0.012219 | 0.18826 | 0.23714 | 0.17604 | 0.24936 | 5.745% | 11.49% | 17.23% | 0.20207 | 0.22334
U, ppm | 1.8168 | 0.09763 | 1.6215 | 2.012 | 1.5239 | 2.1097 | 5.374% | 10.75% | 16.12% | 1.7259 | 1.9076
V, ppm | 98.0578 | 3.0175 | 92.0228 | 104.0927 | 89.0053 | 107.1102 | 3.077% | 6.155% | 9.232% | 93.1549 | 102.9607
W, ppm | 33.2382 | 3.0864 | 27.0654 | 39.4109 | 23.9791 | 42.4973 | 9.286% | 18.57% | 27.86% | 31.5763 | 34.9001
Y, ppm | 16.1353 | 0.84327 | 14.4488 | 17.8219 | 13.6055 | 18.6652 | 5.226% | 10.45% | 15.68% | 15.3286 | 16.9421
Yb, ppm | 1.4168 | 0.098093 | 1.2206 | 1.613 | 1.1226 | 1.7111 | 6.923% | 13.85% | 20.77% | 1.346 | 1.4877
Zn, ppm | 83.7642 | 2.67 | 78.4242 | 89.1042 | 75.7542 | 91.7742 | 3.188% | 6.375% | 9.563% | 79.576 | 87.9524
Zr, ppm | 180.828 | 8.451 | 163.9261 | 197.73 | 155.4751 | 206.181 | 4.673% | 9.347% | 14.02% | 171.7866 | 189.8694
Aqua Regia Digestion
Ag, ppm | 1.0415 | 0.037009 | 0.96745 | 1.1155 | 0.93044 | 1.1525 | 3.554% | 7.107% | 10.66% | 0.9894 | 1.0935
Al, wt.% | 1.8413 | 0.19332 | 1.4547 | 2.2279 | 1.2614 | 2.4213 | 10.5% | 21% | 31.5% | 1.7492 | 1.9334
As, ppm | 153.4735 | 6.1316 | 141.2104 | 165.7366 | 135.0788 | 171.8682 | 3.995% | 7.99% | 11.99% | 145.7998 | 161.1472
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 94.2241 | 4.295 | 85.6341 | 102.814 | 81.3392 | 107.109 | 4.558% | 9.116% | 13.67% | 89.5128 | 98.9353
Be, ppm | 1.0252 | 0.043031 | 0.9391 | 1.1112 | 0.89607 | 1.1543 | 4.197% | 8.395% | 12.59% | 0.9739 | 1.0764
Bi, ppm | 2.665 | 0.16487 | 2.3353 | 2.9948 | 2.1704 | 3.1596 | 6.187% | 12.37% | 18.56% | 2.5318 | 2.7983
Ca, wt.% | 0.47371 | 0.050005 | 0.37371 | 0.57372 | 0.3237 | 0.62373 | 10.56% | 21.11% | 31.67% | 0.45003 | 0.4974
Cd, ppm | 0.081041 | 0.011767 | 0.057508 | 0.10457 | 0.045741 | 0.11634 | 14.52% | 29.04% | 43.56% | 0.076989 | 0.085093
Ce, ppm | 43.3601 | 2.0877 | 39.1847 | 47.5354 | 37.0971 | 49.6231 | 4.815% | 9.629% | 14.44% | 41.1921 | 45.5281
Co, ppm | 18.0346 | 0.76457 | 16.5055 | 19.5638 | 15.7409 | 20.3284 | 4.239% | 8.479% | 12.72% | 17.1329 | 18.9364
Cr, ppm | 107.5263 | 9.7643 | 87.9977 | 127.0548 | 78.2334 | 136.8191 | 9.081% | 18.16% | 27.24% | 102.1499 | 112.9026
Cs, ppm | 0.99336 | 0.12956 | 0.73424 | 1.2525 | 0.60468 | 1.382 | 13.04% | 26.09% | 39.13% | 0.9437 | 1.043
Cu, ppm | 55.0235 | 3.4914 | 48.0407 | 62.0063 | 44.5493 | 65.4978 | 6.345% | 12.69% | 19.04% | 52.2724 | 57.7747
Fe, wt.% | 3.4883 | 0.24745 | 2.9934 | 3.9832 | 2.746 | 4.2307 | 7.094% | 14.19% | 21.28% | 3.3139 | 3.6627
Ga, ppm | 6.1299 | 0.6269 | 4.8761 | 7.3837 | 4.2492 | 8.0106 | 10.23% | 20.45% | 30.68% | 5.8234 | 6.4364
Ge, ppm | 0.11583 | 0.028018 | 0.059798 | 0.17187 | 0.03178 | 0.19989 | 24.19% | 48.38% | 72.56% | 0.11004 | 0.12162
In, ppm | 0.0225 | 0.0019446 | 0.018611 | 0.026389 | 0.016666 | 0.028334 | 8.643% | 17.29% | 25.93% | 0.021375 | 0.023625
K, wt.% | 0.28601 | 0.035707 | 0.2146 | 0.35742 | 0.17889 | 0.39313 | 12.48% | 24.97% | 37.45% | 0.27171 | 0.30031
La, ppm | 23.7222 | 1.0206 | 21.681 | 25.7634 | 20.6604 | 26.784 | 4.302% | 8.605% | 12.91% | 22.5361 | 24.9083
Li, ppm | 10.4311 | 0.52618 | 9.3788 | 11.4835 | 8.8526 | 12.0097 | 5.044% | 10.09% | 15.13% | 9.9096 | 10.9527
Mg, wt.% | 1.1627 | 0.044577 | 1.0735 | 1.2519 | 1.029 | 1.2964 | 3.834% | 7.668% | 11.5% | 1.1046 | 1.2208
Mn, wt.% | 0.031772 | 0.0018311 | 0.028109 | 0.035434 | 0.026278 | 0.037265 | 5.763% | 11.53% | 17.29% | 0.030183 | 0.03336
Mo, ppm | 3.9427 | 0.38603 | 3.1707 | 4.7148 | 2.7847 | 5.1008 | 9.791% | 19.58% | 29.37% | 3.7456 | 4.1399
Na, wt.% | 0.26762 | 0.019775 | 0.22807 | 0.30717 | 0.2083 | 0.32695 | 7.389% | 14.78% | 22.17% | 0.25424 | 0.281
Ni, ppm | 83.9296 | 3.7347 | 76.4603 | 91.399 | 72.7257 | 95.1336 | 4.45% | 8.899% | 13.35% | 79.7332 | 88.1261
P, wt.% | 0.048376 | 0.0025945 | 0.043187 | 0.053565 | 0.040592 | 0.056159 | 5.363% | 10.73% | 16.09% | 0.045957 | 0.050795
Pb, ppm | 11.4625 | 0.93898 | 9.5845 | 13.3404 | 8.6455 | 14.2794 | 8.192% | 16.38% | 24.58% | 10.8893 | 12.0356
Rb, ppm | 14.5777 | 1.0305 | 12.5168 | 16.6387 | 11.4863 | 17.6691 | 7.069% | 14.14% | 21.21% | 13.8488 | 15.3066
Re, ppm | 0.0008125 | 0.00024401 | 0.00032448 | 0.0013005 | 8.0469e-05 | 0.0015445 | 30.03% | 60.06% | 90.1% | 0.00077187 | 0.00085313
S, wt.% | 0.03421 | 0.0042534 | 0.025703 | 0.042717 | 0.02145 | 0.04697 | 12.43% | 24.87% | 37.3% | 0.032499 | 0.03592
Sb, ppm | 5.2302 | 0.56403 | 4.1022 | 6.3583 | 3.5382 | 6.9223 | 10.78% | 21.57% | 32.35% | 4.9687 | 5.4918
Sc, ppm | 5.1955 | 0.40915 | 4.3772 | 6.0138 | 3.9681 | 6.423 | 7.875% | 15.75% | 23.63% | 4.9357 | 5.4553
Sn, ppm | 1.3707 | 0.19716 | 0.97637 | 1.765 | 0.77921 | 1.9622 | 14.38% | 28.77% | 43.15% | 1.3021 | 1.4392
Sr, ppm | 52.4683 | 2.0211 | 48.4261 | 56.5105 | 46.405 | 58.5316 | 3.852% | 7.704% | 11.56% | 49.8449 | 55.0917
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.45396 | 0.036941 | 0.38008 | 0.52785 | 0.34314 | 0.56479 | 8.137% | 16.27% | 24.41% | 0.43127 | 0.47666
Te, ppm | 0.11747 | 0.023863 | 0.069747 | 0.1652 | 0.045884 | 0.18906 | 20.31% | 40.63% | 60.94% | 0.1116 | 0.12335
Th, ppm | 6.6701 | 0.33602 | 5.998 | 7.3421 | 5.662 | 7.6782 | 5.038% | 10.08% | 15.11% | 6.3366 | 7.0036
Ti, wt.% | 0.11412 | 0.01413 | 0.08586 | 0.14238 | 0.071731 | 0.15651 | 12.38% | 24.76% | 37.14% | 0.10841 | 0.11983
Tl, ppm | 0.12453 | 0.016857 | 0.090814 | 0.15824 | 0.073957 | 0.1751 | 13.54% | 27.07% | 40.61% | 0.1183 | 0.13075
U, ppm | 1.0067 | 0.065325 | 0.87609 | 1.1374 | 0.81076 | 1.2027 | 6.489% | 12.98% | 19.47% | 0.9564 | 1.0571
V, ppm | 44.8323 | 3.1267 | 38.5788 | 51.0857 | 35.452 | 54.2125 | 6.974% | 13.95% | 20.92% | 42.5906 | 47.0739
W, ppm | 10.1128 | 1.9723 | 6.1681 | 14.0574 | 4.1958 | 16.0297 | 19.5% | 39.01% | 58.51% | 9.6071 | 10.6184
Y, ppm | 9.0862 | 0.39576 | 8.2947 | 9.8777 | 7.8989 | 10.2735 | 4.356% | 8.711% | 13.07% | 8.6319 | 9.5405
Yb, ppm | 0.68075 | 0.048877 | 0.583 | 0.77851 | 0.53412 | 0.82739 | 7.18% | 14.36% | 21.54% | 0.64672 | 0.71479
Zn, ppm | 58.7129 | 5.0495 | 48.6138 | 68.812 | 43.5643 | 73.8615 | 8.6% | 17.2% | 25.8% | 55.7773 | 61.6485
Zr, ppm | 37.3243 | 6.4753 | 24.3736 | 50.275 | 17.8983 | 56.7503 | 17.35% | 34.7% | 52.05% | 35.4581 | 39.1905
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 255c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 255c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 255c (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 2 | 1 | 245460 | 0.085372 | 4.4817 | 4.3554
2 | 1 | ANSLu | 1 | 10 | 5 | 245461 | 0.083347 | 4.333 | 4.3475
2 | 1 | ANSLu | 1 | 11 | 6 | 245462 | 0.084299 | 4.4307 | 4.3527
2 | 1 | ANSLu | 1 | 6 | 3 | 245463 | 0.087069 | 4.3919 | 4.3507
2 | 1 | ANSLu | 1 | 8 | 4 | 245464 | 0.083906 | 4.3275 | 4.3472
2 | 1 | ANSLu | 1 | 20 | 10 | 245465 | 0.084654 | 4.4151 | 4.3519
2 | 1 | ANSLu | 1 | 13 | 7 | 245466 | 0.083404 | 4.233 | 4.3422
2 | 1 | ANSLu | 1 | 16 | 8 | 245467 | 0.084405 | 4.2389 | 4.3425
2 | 1 | ANSLu | 1 | 1 | 1 | 245468 | 0.086671 | 4.397 | 4.3509
2 | 1 | ANSLu | 1 | 17 | 9 | 245469 | 0.086031 | 4.4305 | 4.3527
2 | 1 | ANSLu | 1 | 19 | 10 | 245470 | 0.085689 | 4.2594 | 4.3436
2 | 1 | ANSLu | 1 | 14 | 7 | 245471 | 0.087901 | 4.3278 | 4.3472
2 | 1 | ANSLu | 1 | 12 | 6 | 245472 | 0.08345 | 4.3321 | 4.3475
2 | 1 | ANSLu | 1 | 3 | 2 | 245473 | 0.085091 | 4.2139 | 4.3412
2 | 1 | ANSLu | 1 | 18 | 9 | 245474 | 0.08695 | 4.4011 | 4.3511
2 | 1 | ANSLu | 1 | 4 | 2 | 245475 | 0.086194 | 4.2774 | 4.3445
2 | 1 | ANSLu | 1 | 15 | 8 | 245476 | 0.086297 | 4.4524 | 4.3539
2 | 1 | ANSLu | 1 | 7 | 4 | 245477 | 0.087878 | 4.3184 | 4.3467
2 | 1 | ANSLu | 1 | 5 | 3 | 245478 | 0.084224 | 4.3994 | 4.3511
2 | 1 | ANSLu | 1 | 9 | 5 | 245479 | 0.08716 | 4.3055 | 4.346
Mean | [6 empty] | 0.0855 | 4.3483 | 4.3483
Median | [7 empty] | 4.3325 | 4.3475
SD | [7 empty] | 0.078365 | 0.0041835
RSD | [7 empty] | 1.802% | 0.09621%
RSD Upscale Factor | [7 empty] | 0.053385
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 255c (Certified Value 4.17 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*AAS | 2
|  |  | 0.085g | 25g | 50g | 30g | 50g | 30g | 50g | 50g | 50g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 4.397 | 4.2 | 3.88 | 4.0973 | 4.34 | 4.2 | 3.97 | 4.13 | 4.358 | 4.09 | 4.11 | 4.311 | 4.09 | 4.308 | 4.275 | 4.19 | 4.1429 | 4.233 | 4.15 | 4.23 | 4.16 | 4.18 | 1
| 1 | 2 | 4.4817 | 4.265 | 3.886 | 4.2314 | 4.1 | 4.22 | 4.24 | 4.1 | 4.469 | 4.08 | 4.1 | 4.356 | 4.08 | 4.421 | 4.275 | 4.19 | 4.1686 | 4.237 | 4.21 | 4.34 | 4.12 | 4.13 | 0
| 1 | 3 | 4.2139 | 4.215 | 3.885 | 4.0321 | 4.78 | 4.15 | 4.23 | 4.19 | 4.379 | 4.09 | 4.11 | 4.33 | 4.17 | 4.388 | 4.273 | 4.23 | 4.1736 | 4.18 | 4.17 | 4.22 | 4.21 | 4.14 | 16
| 1 | 4 | 4.2774 | 4.252 | 3.898 | 4.2074 | 4.68 | 4.27 | 4.05 | 4.04 | 4.237 | 3.98 | 4.12 | 4.325 | 4.2 | 4.45 | 4.273 | 4.12 | 4.1678 | 4.217 | 4.17 | 4.17 | 4.12 | 4.16 | 4.1722
| 1 | 5 | 4.3994 | 4.231 | 3.918 | 4.2208 | 4.31 | 4.2 | 3.97 | 4.16 | 4.434 | 4.04 | 4.14 | 4.3247 | 4.05 | 4.342 | 4.2715 | 4.22 | 4.1463 | 4.166 | 4.13 | 4.21 | 4.2 | 4.07 | 7
| 1 | 6 | 4.3919 | 4.159 | 3.828 | 4.1058 | 4.16 | 4.11 | 3.97 | 4.15 | 4.388 | 4.08 | 4.1 | 4.336 | 4.02 | 4.471 | 4.276 | 4.31 | 4.1653 | 4.238 | 4.16 | 4.16 | 4.21 | 4.21
|  | 7 | 2.00: 4.3184
|  | 8 | 2.00: 4.3275
|  | 9 | 2.00: 4.3055
|  | 10 | 2.00: 4.333
|  | 11 | 2.00: 4.4307
|  | 12 | 2.00: 4.3321
|  | 13 | 2.00: 4.233
|  | 14 | 2.00: 4.3278
|  | 15 | 2.00: 4.4524
|  | 16 | 2.00: 4.2389
|  | 17 | 2.00: 4.4305
|  | 18 | 2.00: 4.4011
|  | 19 | 2.00: 4.2594
|  | 20 | 2.00: 4.4151
| Mean |  | 4.3483 | 4.2203 | 3.8825 | 4.1491 | 4.395 | 4.1917 | 4.0717 | 4.1283 | 4.3775 | 4.06 | 4.1133 | 4.3304 | 4.1017 | 4.3967 | 4.2739 | 4.21 | 4.1607 | 4.2118 | 4.165 | 4.2217 | 4.17 | 4.1483
| Median |  | 4.3325 | 4.223 | 3.8855 | 4.1566 | 4.325 | 4.2 | 4.01 | 4.14 | 4.3835 | 4.08 | 4.11 | 4.3275 | 4.085 | 4.4045 | 4.274 | 4.205 | 4.1665 | 4.225 | 4.165 | 4.215 | 4.18 | 4.15
| Std Dev. |  | 0.078365 | 0.03825 | 0.029985 | 0.081931 | 0.27639 | 0.055648 | 0.13029 | 0.052694 | 0.079746 | 0.043359 | 0.015055 | 0.015002 | 0.06969 | 0.063067 | 0.0016857 | 0.06229 | 0.012848 | 0.031327 | 0.026646 | 0.064317 | 0.042895 | 0.047924
| Rel.Std.Dev. |  | 1.802% | 0.9063% | 0.7723% | 1.975% | 6.289% | 1.328% | 3.2% | 1.276% | 1.822% | 1.068% | 0.366% | 0.3464% | 1.699% | 1.434% | 0.03944% | 1.48% | 0.3088% | 0.7438% | 0.6398% | 1.523% | 1.029% | 1.155%
| PDM3 |  | 4.221% | 1.153% | -6.944% | -0.5542% | 5.339% | 0.4657% | -2.41% | -1.052% | 4.92% | -2.69% | -1.412% | 3.792% | -1.691% | 5.379% | 2.437% | 0.9051% | -0.2755% | 0.949% | -0.1735% | 1.185% | -0.05365% | -0.573%
| Z-Score (Absolute) |  | NA | 0.66 | 3.75 | 0.27 | 2.94 | 0.28 | 1.28 | 0.54 | 2.71 | 1.44 | 0.74 | 2.09 | 0.89 | 2.96 | 1.36 | 0.52 | 0.12 | 0.55 | 0.07 | 0.67 | 0 | 0.28
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 255c (Certified Value 4.13 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | 2
|  |  | 0.085g | 25g | 30g | 50g | 30g | 50g | 30g | 30g | 25g | 15g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 3
| 1 | 1 | 4.397 | 4.26 | 3.9473 | 4.18 | 4.03 | 3.93 | 3.76 | 4.1128 | 4.25 | 4.1303 | 4.168 | 4.16 | 4.0381 | 4.06 | 4.06 | 4.19 | 3.688 | 1
| 1 | 2 | 4.4817 | 4.22 | 3.9534 | 3.94 | 4 | 4 | 3.8 | 4.1248 | 4.15 | 4.1406 | 4.165 | 4.11 | 4.0463 | 4.11 | 4.22 | 4.21 | 3.785 | 0
| 1 | 3 | 4.2139 | 4.19 | 3.9524 | 4.38 | 4.15 | 3.9 | 3.63 | 4.102 | 4.12 | 4.1612 | 4.136 | 4.16 | 4.1651 | 4.1 | 4.11 | 4.27 | 3.891 | 16
| 1 | 4 | 4.2774 | 4.19 | 3.9655 | 4.42 | 4 | 3.77 | 3.91 | 4.174 | 4.1 | 4.0582 | 4.167 | 4.07 | 4.0675 | 4.1 | 4.02 | 4.16 | 3.708 | 4.1303
| 1 | 5 | 4.3994 | 4.15 | 3.9858 | 4.04 | 4.11 | 4.07 | 3.69 | 4.132 | 4.27 | 4.0994 | 4.165 | 4.18 | 4.1508 | 4.25 | 4.16 | 3.99 | 3.9 | 9
| 1 | 6 | 4.3919 | 4.15 | 3.8946 | 4.02 | 4.1 | 3.9 | 3.99 | 4.18 | 4.29 | 3.7492 | 4.169 | 3.85 | 3.9528 | 3.95 | 4.2 | 4.08 | 3.891
|  | 7 | 2.00: 4.3184
|  | 8 | 2.00: 4.3275
|  | 9 | 2.00: 4.3055
|  | 10 | 2.00: 4.333
|  | 11 | 2.00: 4.4307
|  | 12 | 2.00: 4.3321
|  | 13 | 2.00: 4.233
|  | 14 | 2.00: 4.3278
|  | 15 | 2.00: 4.4524
|  | 16 | 2.00: 4.2389
|  | 17 | 2.00: 4.4305
|  | 18 | 2.00: 4.4011
|  | 19 | 2.00: 4.2594
|  | 20 | 2.00: 4.4151
| Mean |  | 4.3483 | 4.1933 | 3.9498 | 4.1633 | 4.065 | 3.9283 | 3.7967 | 4.1376 | 4.1967 | 4.0565 | 4.1617 | 4.0883 | 4.0701 | 4.095 | 4.1283 | 4.15 | 3.8105
| Median |  | 4.3325 | 4.19 | 3.9529 | 4.11 | 4.065 | 3.915 | 3.78 | 4.1284 | 4.2 | 4.1148 | 4.166 | 4.135 | 4.0569 | 4.1 | 4.135 | 4.175 | 3.838
| Std Dev. |  | 0.078365 | 0.042269 | 0.030367 | 0.19937 | 0.063482 | 0.10187 | 0.13471 | 0.032244 | 0.082865 | 0.15476 | 0.012675 | 0.12352 | 0.078603 | 0.096488 | 0.079099 | 0.1002 | 0.097091
| Rel.Std.Dev. |  | 1.802% | 1.008% | 0.7688% | 4.789% | 1.562% | 2.593% | 3.548% | 0.7793% | 1.975% | 3.815% | 0.3046% | 3.021% | 1.931% | 2.356% | 1.916% | 2.414% | 2.548%
| PDM3 |  | 5.279% | 1.526% | -4.369% | 0.7997% | -1.581% | -4.89% | -8.078% | 0.1768% | 1.607% | -1.787% | 0.7594% | -1.016% | -1.458% | -0.8547% | -0.04768% | 0.4769% | -7.743%
| Z-Score (Absolute) |  | NA | 1.07 | 1.49 | 0.75 | 0.28 | 1.72 | 3.1 | 0.48 | 1.1 | 0.37 | 0.74 | 0.04 | 0.23 | 0.04 | 0.39 | 0.61 | 2.95
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 255c (Certified Value 3.97 ppm) | [18 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 05g | 50g | 50g | 200g | 50g | 30g | 30g | 30g | 30g | 20g | 05g | 3
| 1 | 1 | 4.397 | 4.13 | 3.803 | 3.7766 | 3.34 | 3.99 | 3.8 | 4.0591 | 3.8 | 3.98 | 4.245 | 3.343 | 4.1 | 4.32 | 4.34 | 3.694 | 4.164 | 1
| 1 | 2 | 4.4817 | 4.22 | 3.8088 | 3.8143 | 3.62 | 3.97 | 3.8 | 3.8878 | 3.75 | 3.99 | 4.205 | 3.31 | 4.09 | 4.32 | 4.32 | 3.8791 | 4.16 | 0
| 1 | 3 | 4.2139 | 4.3 | 3.8078 | 3.8301 | 4.14 | 4.01 | 3.7 | 4.2519 | 3.74 | 4.03 | 4.275 | 3.41 | 4.1 | 4.322 | 4.35 | 3.8521 | 4.152 | 16
| 1 | 4 | 4.2774 | 4.31 | 3.8205 | 3.8496 | 4.01 | 3.98 | 3.9 | 4.1822 | 3.71 | 4.02 | 4.252 | 3.375 | 4.18 | 4.296 | 4.35 | 3.6402 | 4.124 | 3.9653
| 1 | 5 | 4.3994 | 4.27 | 3.84 | 3.7887 | 3.92 | 3.95 | 3.8 | 4.0711 | 3.82 | 4 | 4.236 | 3.392 | 4.1 | 4.323 | 4.33 | 3.738 | 4.175 | 11
| 1 | 6 | 4.3919 | 4.28 | 3.7522 | 3.7809 | 3.56 | 4.04 | 3.8 | 3.81 | 3.69 | 3.99 | 4.2295 | 3.374 | 4.16 | 4.321 | 4.06 | 3.5844 | 4.178
|  | 7 | 2.00: 4.3184
|  | 8 | 2.00: 4.3275
|  | 9 | 2.00: 4.3055
|  | 10 | 2.00: 4.333
|  | 11 | 2.00: 4.4307
|  | 12 | 2.00: 4.3321
|  | 13 | 2.00: 4.233
|  | 14 | 2.00: 4.3278
|  | 15 | 2.00: 4.4524
|  | 16 | 2.00: 4.2389
|  | 17 | 2.00: 4.4305
|  | 18 | 2.00: 4.4011
|  | 19 | 2.00: 4.2594
|  | 20 | 2.00: 4.4151
| Mean |  | 4.3483 | 4.2517 | 3.8054 | 3.8067 | 3.765 | 3.99 | 3.8 | 4.0437 | 3.7517 | 4.0017 | 4.2404 | 3.3673 | 4.1217 | 4.317 | 4.2917 | 3.7313 | 4.1588
| Median |  | 4.3325 | 4.275 | 3.8083 | 3.8015 | 3.77 | 3.985 | 3.8 | 4.0651 | 3.745 | 3.995 | 4.2405 | 3.3745 | 4.1 | 4.3205 | 4.335 | 3.716 | 4.162
| Std Dev. |  | 0.078365 | 0.067355 | 0.029257 | 0.029466 | 0.30606 | 0.031623 | 0.063246 | 0.16881 | 0.050365 | 0.019408 | 0.023427 | 0.035809 | 0.038166 | 0.010354 | 0.11409 | 0.11642 | 0.01958
| Rel.Std.Dev. |  | 1.802% | 1.584% | 0.7688% | 0.7741% | 8.129% | 0.7926% | 1.664% | 4.175% | 1.342% | 0.485% | 0.5525% | 1.063% | 0.926% | 0.2398% | 2.658% | 3.12% | 0.4708%
| PDM3 |  | 9.661% | 7.223% | -4.031% | -3.999% | -5.05% | 0.6241% | -4.168% | 1.978% | -5.386% | 0.9183% | 6.939% | -15.08% | 3.945% | 8.871% | 8.232% | -5.9% | 4.882%
| Z-Score (Absolute) |  | NA | 0.81 | 0.6 | 0.6 | 0.73 | 0.02 | 0.62 | 0.15 | 0.77 | 0.02 | 0.77 | 1.99 | 0.4 | 1.02 | 0.94 | 0.84 | 0.52
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 255c (Certified Value 4.23 ppm) | [13 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 4.397 | 4.392 | 4.46 | 4.18 | 4.22 | 4.34 | 4.11 | 4.18 | 4.301 | 4.318 | 4.203 | 4.313 | 1
| 1 | 2 | 4.4817 | 4.366 | 4.447 | 4.16 | 4.11 | 4.12 | 4.21 | 3.97 | 4.27 | 4.327 | 4.221 | 4.275 | 0
| 1 | 3 | 4.2139 | 4.366 | 4.206 | 4.37 | 4.16 | 4.47 | 4.14 | 4.04 | 4.269 | 4.214 | 4.208 | 4.27 | 16
| 1 | 4 | 4.2774 | 4.438 | 4.302 | 4.19 | 4.19 | 4.25 | 4.24 | 4.09 | 4.253 | 3.953 | 4.22 | 4.118 | 4.2308
| 1 | 5 | 4.3994 | 4.284 | 4.165 | 4.32 | 3.97 | 4.37 | 4.32 | 4.33 | 4.41 | 4.313 | 4.202 | 4.091 | 13
| 1 | 6 | 4.3919 | 4.336 | 4.143 | 4.24 | 4.12 | 4.26 | 4.16 | 4 | 4.163 | 4.382 | 3.981 | 4.221
|  | 7 | 2.00: 4.3184
|  | 8 | 2.00: 4.3275
|  | 9 | 2.00: 4.3055
|  | 10 | 2.00: 4.333
|  | 11 | 2.00: 4.4307
|  | 12 | 2.00: 4.3321
|  | 13 | 2.00: 4.233
|  | 14 | 2.00: 4.3278
|  | 15 | 2.00: 4.4524
|  | 16 | 2.00: 4.2389
|  | 17 | 2.00: 4.4305
|  | 18 | 2.00: 4.4011
|  | 19 | 2.00: 4.2594
|  | 20 | 2.00: 4.4151
| Mean |  | 4.3483 | 4.3637 | 4.2872 | 4.2433 | 4.1283 | 4.3017 | 4.1967 | 4.1017 | 4.2777 | 4.2512 | 4.1725 | 4.2147
| Median |  | 4.3325 | 4.366 | 4.254 | 4.215 | 4.14 | 4.3 | 4.185 | 4.065 | 4.2695 | 4.3155 | 4.2055 | 4.2455
| Std Dev. |  | 0.078365 | 0.051868 | 0.13995 | 0.084538 | 0.087958 | 0.1199 | 0.076594 | 0.13408 | 0.079934 | 0.15585 | 0.094172 | 0.090608
| Rel.Std.Dev. |  | 1.802% | 1.189% | 3.264% | 1.992% | 2.131% | 2.787% | 1.825% | 3.269% | 1.869% | 3.666% | 2.257% | 2.15%
| PDM3 |  | 2.779% | 3.141% | 1.333% | 0.2969% | -2.421% | 1.676% | -0.8061% | -3.052% | 1.108% | 0.482% | -1.377% | -0.3807%
| Z-Score (Absolute) |  | NA | 1.74 | 0.63 | 0 | 1.66 | 0.84 | 0.67 | 2.05 | 0.5 | 0.11 | 1.02 | 0.41
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 255c (Certified Value 1.11 ppm) | [19 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.07 | 1.0716 | 0.5571 | 1.2 | 1.3 | 1 | 1.355 | 1.1 | 1.1 | 1.1366 | 1.1 | 1.0141 | 1.4 | 1.04 | 1.05 | 1.14 |  |  | 1
| 1 | 2 | 1.1 | 0.99213 | 0.5016 | 1.2 | 1.5 | 1 | 1.396 | 1.14 | 1.1 | 1.1349 | 1.06 | 0.96849 | 1.4 | 1.06 | 1.03 | 1.2 |  |  | 19
| 1 | 3 | 1.15 | 0.9881 | 0.5839 | 1.1 | 1.3 | 1 | 1.3999 | 1.15 | 1.2 | 1.1065 | 1.1 | 1.0299 | 1.3 | 1.02 | 1.06 | 1.27 |  |  | 16
| 1 | 4 | 1.09 | 1.006 | 0.5784 | 1.1 | 1.1 | 1 | 1.3542 | 1.13 | 1.1 | 1.138 | 1.07 | 1.0024 | 1.3 | 1.08 | 1 | 1.08 |  |  | 1.107
| 1 | 5 | 1.13 | 0.99454 | 0.6292 | 1.2 | 1.1 | 1 | 1.4652 | 1.06 | 1.2 | 1.132 | 1.1 | 0.97915 | 1.3 | 1 | 1.04 | 1.16 |  |  | 15
| 1 | 6 | 1.08 | 1.0666 | 0.8181 | 1.2 | 1.1 | 0.9 | 1.387 | 1.15 | 1.1 | 1.1359 | 1.12 | 1.0182 | 1.3 | 1.06 | 1.04 | 1.06
| Mean |  | 1.1033 | 1.0198 | 0.61138 | 1.1667 | 1.2333 | 0.98333 | 1.3929 | 1.1217 | 1.1333 | 1.1307 | 1.0917 | 1.002 | 1.3333 | 1.0433 | 1.0367 | 1.1517
| Median |  | 1.095 | 1.0003 | 0.58115 | 1.2 | 1.2 | 1 | 1.3915 | 1.135 | 1.1 | 1.1354 | 1.1 | 1.0082 | 1.3 | 1.05 | 1.04 | 1.15
| Std Dev. |  | 0.030768 | 0.038664 | 0.10945 | 0.05164 | 0.1633 | 0.040825 | 0.040632 | 0.035449 | 0.05164 | 0.012012 | 0.022286 | 0.023789 | 0.05164 | 0.029439 | 0.020656 | 0.077567
| Rel.Std.Dev. |  | 2.789% | 3.791% | 17.9% | 4.426% | 13.24% | 4.152% | 2.917% | 3.16% | 4.556% | 1.062% | 2.041% | 2.374% | 3.873% | 2.822% | 1.993% | 6.735%
| PDM3 |  | -0.3282% | -7.871% | -44.77% | 5.393% | 11.42% | -11.17% | 25.83% | 1.328% | 2.382% | 2.14% | -1.382% | -9.48% | 20.45% | -5.748% | -6.351% | 4.038%
| Z-Score (Absolute) |  | 0.09 | 0.86 | 4.66 | 0.5 | 1.12 | 1.2 | 2.61 | 0.09 | 0.19 | 0.17 | 0.19 | 1.03 | 2.05 | 0.64 | 0.71 | 0.36
| Analytical results for Al in OREAS 255c (Certified Value 6.31 wt.%) | [19 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 6.445 | 6.2653 | 6.5302 | 5.6645 | 6.2 | 6.61 | 6.33 | 6.35 | 6.25 | 5.96 | 5.8988 | 6.35 | 6.2853 | 6.45 | 6.2 | 6.23 | 6.21 | 5.85 | 1
| 1 | 2 | 6.3436 | 6.2666 | 6.4258 | 5.8372 | 6.29 | 6.65 | 6.32 | 6.32 | 6.42 | 6.14 | 5.8818 | 6.4 | 6.3155 | 6.42 | 6.03 | 6.17 | 6.22 | 6.13 | 0
| 1 | 3 | 6.4271 | 6.2418 | 6.6224 | 5.7774 | 6.33 | 6.7 | 6.22 | 6.37 | 6.43 | 6.48 | 5.8578 | 6.34 | 6.347 | 6.3 | 6.37 | 6.21 | 6.21 | 5.79 | 16
| 1 | 4 | 6.5285 | 6.2911 | 6.5364 | 5.7793 | 6.33 | 6.79 | 6.3 | 6.29 | 6.21 | 6.56 | 5.903 | 6.17 | 6.1881 | 6.36 | 6.4 | 6.2 | 6.15 | 5.86 | 6.3074
| 1 | 5 | 6.4675 | 6.2163 | 6.56 | 5.7577 | 6.32 | 6.68 | 6.32 | 6.32 | 6.42 | 6.18 | 5.9132 | 6.32 | 6.2621 | 6.42 | 6.23 | 6.26 | 6.23 | 5.61 | 16
| 1 | 6 | 6.4192 | 6.2304 | 6.6019 | 5.8281 | 6.27 | 6.66 | 6.34 | 6.35 | 6.36 | 6.51 | 5.8726 | 6.32 | 6.2201 | 6.38 | 6.29 | 6.33 | 6.14 | 5.82
| Mean |  | 6.4385 | 6.2519 | 6.5461 | 5.774 | 6.29 | 6.6817 | 6.305 | 6.3333 | 6.3483 | 6.305 | 5.8879 | 6.3167 | 6.2697 | 6.3883 | 6.2533 | 6.2333 | 6.1933 | 5.8433
| Median |  | 6.436 | 6.2536 | 6.5482 | 5.7784 | 6.305 | 6.67 | 6.32 | 6.335 | 6.39 | 6.33 | 5.8903 | 6.33 | 6.2737 | 6.4 | 6.26 | 6.22 | 6.21 | 5.835
| Std Dev. |  | 0.060814 | 0.027428 | 0.069179 | 0.061985 | 0.0502 | 0.061128 | 0.043704 | 0.028752 | 0.095795 | 0.24477 | 0.020805 | 0.077632 | 0.059138 | 0.053821 | 0.13397 | 0.056095 | 0.038297 | 0.16753
| Rel.Std.Dev. |  | 0.9445% | 0.4387% | 1.057% | 1.074% | 0.7981% | 0.9149% | 0.6932% | 0.454% | 1.509% | 3.882% | 0.3534% | 1.229% | 0.9432% | 0.8425% | 2.142% | 0.8999% | 0.6184% | 2.867%
| PDM3 |  | 2.079% | -0.879% | 3.785% | -8.456% | -0.2753% | 5.934% | -0.03752% | 0.4117% | 0.6495% | -0.03752% | -6.651% | 0.1474% | -0.5973% | 1.284% | -0.8567% | -1.174% | -1.808% | -7.357%
| Z-Score (Absolute) |  | 1.65 | 0.53 | 2.92 | 6.14 | 0.09 | 4.51 | 0.09 | 0.42 | 0.6 | 0.09 | 4.8 | 0.22 | 0.33 | 1.07 | 0.52 | 0.75 | 1.22 | 5.33
| Analytical results for As in OREAS 255c (Certified Value 174 ppm) | [19 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 173.5 | 179.0725 | 273.3533 | 177 | 138 | 174 | 172.9367 | 173 | 167 | 108.4699 | 171 | 170.459 | 168 | 173 | 175.5 | 178.5 | 174 |  | 1
| 1 | 2 | 173.1 | 174.2678 | 239.5508 | 171 | 147 | 173 | 171.02 | 175.5 | 168 | 108.205 | 176 | 161.6458 | 167 | 173.5 | 174 | 173.5 | 176 |  | 20
| 1 | 3 | 177.2 | 179.9004 | 237.1345 | 173 | 143 | 173 | 173.8633 | 177.5 | 171 | 109.6874 | 174.5 | 167.302 | 166 | 178.5 | 176 | 177 | 170 |  | 16
| 1 | 4 | 175.8 | 179.2793 | 243.5995 | 177 | 142 | 170 | 169.89 | 172 | 170 | 108.6196 | 169.5 | 167.371 | 168 | 177.5 | 172 | 174.5 | 175 |  | 173.5351
| 1 | 5 | 174.8 | 176.8312 | 247.4004 | 175 | 140 | 171 | 170.11 | 177.5 | 168 | 108.9359 | 175 | 168.3704 | 168 | 175.5 | 175.5 | 172.5 | 175 |  | 17
| 1 | 6 | 175.9 | 177.5244 | 286.1231 | 172 | 149 | 173 | 176.2433 | 176.5 | 168 | 108.7372 | 174.5 | 162.5288 | 169 | 179 | 176.5 | 171.5 | 171
| Mean |  | 175.05 | 177.8126 | 254.5269 | 174.1667 | 143.1667 | 172.3333 | 172.3439 | 175.3333 | 168.6667 | 108.7758 | 173.4167 | 166.2795 | 167.6667 | 176.1667 | 174.9167 | 174.5833 | 173.5
| Median |  | 175.3 | 178.2984 | 245.5 | 174 | 142.5 | 173 | 171.9783 | 176 | 168 | 108.6784 | 174.5 | 167.3365 | 168 | 176.5 | 175.5 | 174 | 174.5
| Std Dev. |  | 1.5604 | 2.0837 | 20.2471 | 2.5626 | 4.1673 | 1.5055 | 2.4751 | 2.3381 | 1.5055 | 0.51028 | 2.5577 | 3.453 | 1.0328 | 2.5626 | 1.6558 | 2.691 | 2.429
| Rel.Std.Dev. |  | 0.8914% | 1.172% | 7.955% | 1.471% | 2.911% | 0.8736% | 1.436% | 1.334% | 0.8926% | 0.4691% | 1.475% | 2.077% | 0.616% | 1.455% | 0.9466% | 1.541% | 1.4%
| PDM3 |  | 0.873% | 2.465% | 46.67% | 0.3639% | -17.5% | -0.6925% | -0.6864% | 1.036% | -2.805% | -37.32% | -0.06826% | -4.181% | -3.382% | 1.516% | 0.7961% | 0.604% | -0.02023%
| Z-Score (Absolute) |  | 0.57 | 1.59 | 29.8 | 0.25 | 11.16 | 0.43 | 0.43 | 0.67 | 1.78 | 23.81 | 0.03 | 2.66 | 2.15 | 0.98 | 0.52 | 0.4 | 0
| Analytical results for Ba in OREAS 255c (Certified Value 445 ppm) | [19 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 460.5 | 443.366 | 446.3161 | 397.8778 | 427 | 390 | 448 | 456 | 440 | 273 | 426.8264 | 440 | 459.6232 | 447 | 440 | 460 | 440 | 390 | 1
| 1 | 2 | 460.3 | 444.455 | 437.4816 | 407.7944 | 435 | 405 | 447 | 457 | 450 | 280 | 427.7554 | 450 | 451.8606 | 442 | 430 | 450 | 430 | 412 | 21
| 1 | 3 | 460.7 | 443.564 | 443.933 | 406.7147 | 437 | 411 | 445 | 450 | 450 | 289 | 432.9929 | 450 | 452.3818 | 443 | 460 | 460 | 440 | 413 | 16
| 1 | 4 | 460.2 | 450.945 | 447.0257 | 405.4694 | 430 | 409 | 443 | 447 | 440 | 284 | 426.1566 | 430 | 421.9634 | 441 | 460 | 460 | 430 | 406 | 444.7703
| 1 | 5 | 459.8 | 442.541 | 448.6241 | 403.741 | 434 | 404 | 444 | 456 | 450 | 278 | 426.1169 | 440 | 440.9982 | 445 | 450 | 460 | 440 | 403 | 18
| 1 | 6 | 457.3 | 445.423 | 449.7187 | 406.5213 | 431 | 419 | 449 | 456 | 450 | 283 | 424.1898 | 440 | 423.6471 | 444 | 450 | 470 | 430 | 401
| Mean |  | 459.8 | 445.049 | 445.5165 | 404.6864 | 432.3333 | 406.3333 | 446 | 453.6667 | 446.6667 | 281.1667 | 427.3397 | 441.6667 | 441.7457 | 443.6667 | 448.3333 | 460 | 435 | 404.1667
| Median |  | 460.25 | 444.0095 | 446.6709 | 405.9953 | 432.5 | 407 | 446 | 456 | 450 | 281.5 | 426.4915 | 440 | 446.4294 | 443.5 | 450 | 460 | 435 | 404.5
| Std Dev. |  | 1.2617 | 3.0523 | 4.4114 | 3.6057 | 3.6697 | 9.6264 | 2.3664 | 4.1312 | 5.164 | 5.4924 | 3.0071 | 7.5277 | 15.8398 | 2.1602 | 11.6905 | 6.3246 | 5.4772 | 8.4242
| Rel.Std.Dev. |  | 0.2744% | 0.6858% | 0.9902% | 0.891% | 0.8488% | 2.369% | 0.5306% | 0.9106% | 1.156% | 1.953% | 0.7037% | 1.704% | 3.586% | 0.4869% | 2.608% | 1.375% | 1.259% | 2.084%
| PDM3 |  | 3.379% | 0.06266% | 0.1678% | -9.012% | -2.796% | -8.642% | 0.2765% | 2% | 0.4264% | -36.78% | -3.919% | -0.6978% | -0.68% | -0.2481% | 0.8011% | 3.424% | -2.197% | -9.129%
| Z-Score (Absolute) |  | 1.28 | 0.17 | 0.21 | 2.84 | 0.77 | 2.71 | 0.25 | 0.82 | 0.3 | 12.05 | 1.15 | 0.08 | 0.07 | 0.07 | 0.42 | 1.29 | 0.57 | 2.87
| Analytical results for Be in OREAS 255c (Certified Value 2.1 ppm) | [19 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 2.08 | 2.2308 | 1.4973 | 1.8 | 2.2 | 2.12 | < 5 | 2.12 | 2 | 1.7384 | 2.14 | 2.3888 | 2.2 | 2.11 | 2.13 | 2.03 | 2.49 |  | 1
| 1 | 2 | 2 | 2.0929 | 1.5902 | 1.8 | 2.1 | 2.08 | < 5 | 2.17 | 2 | 1.7232 | 2.16 | 2.2254 | 2 | 2.05 | 2.11 | 2.02 | 2.36 |  | 22
| 1 | 3 | 1.97 | 2.2213 | 1.4713 | 1.8 | 2.4 | 2.15 | < 5 | 2.17 | 2 | 1.7268 | 2.14 | 2.2734 | 2.1 | 2.15 | 2.13 | 2.03 | 2.45 |  | 16
| 1 | 4 | 2.12 | 2.1983 | 1.416 | 2 | 2.3 | 2.1 | < 5 | 2.08 | 2 | 1.7456 | 2.1 | 1.8762 | 2 | 2.17 | 2.13 | 2 | 2.37 |  | 2.1034
| 1 | 5 | 2.04 | 2.1346 | 1.5314 | 2 | 2.2 | 2.12 | < 5 | 2.19 | 2 | 1.7433 | 2.14 | 1.9036 | 2 | 2.12 | 2.16 | 2.02 | 2.41 |  | 19
| 1 | 6 | 2.03 | 2.2071 | 1.5527 | 2 | 2.3 | 2.1 | < 5 | 2.16 | 2 | 1.736 | 2.16 | 2.203 | 2 | 2.13 | 2.16 | 2 | 2.53
| Mean |  | 2.04 | 2.1808 | 1.5098 | 1.9 | 2.25 | 2.1117 |  | 2.1483 | 2 | 1.7355 | 2.14 | 2.1451 | 2.05 | 2.1217 | 2.1367 | 2.0167 | 2.435
| Median |  | 2.035 | 2.2027 | 1.5144 | 1.9 | 2.25 | 2.11 |  | 2.165 | 2 | 1.7372 | 2.14 | 2.2142 | 2 | 2.125 | 2.13 | 2.02 | 2.43
| Std Dev. |  | 0.054037 | 0.054775 | 0.061981 | 0.10954 | 0.10488 | 0.024014 |  | 0.040702 | 0 | 0.00894 | 0.021909 | 0.20798 | 0.083666 | 0.041191 | 0.019664 | 0.013663 | 0.067454
| Rel.Std.Dev. |  | 2.649% | 2.512% | 4.105% | 5.766% | 4.661% | 1.137% |  | 1.895% | 0% | 0.5151% | 1.024% | 9.696% | 4.081% | 1.941% | 0.9203% | 0.6775% | 2.77%
| PDM3 |  | -3.015% | 3.681% | -28.22% | -9.67% | 6.969% | 0.3926% |  | 2.136% | -4.916% | -17.49% | 1.74% | 1.981% | -2.539% | 0.868% | 1.581% | -4.124% | 15.76%
| Z-Score (Absolute) |  | 0.71 | 0.41 | 4.96 | 1.84 | 0.97 | 0.14 | 2.97 | 0.15 | NA | 3.15 | 0.09 | 0.13 | 0.63 | 0.06 | 0.06 | 0.9 | 2.45
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 255c (Certified Value 3.16 ppm) | [19 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 2
3
| 1 | 1 | 3.02 | 2.7436 | 3.2948 | 4.1863 | 3.01 | 3.32 | 2.88 | < 5 | 3.31 | 3.9 | 3.308 | 3.04 | 3.4365 | 3.26 | 2.96 | 3.22 | 3.25 | 3.08 | 1
| 1 | 2 | 3.33 | 2.7106 | 3.2571 | 2.7342 | 3.2 | 3.36 | 2.94 | < 5 | 3.31 | 3.5 | 3.3316 | 2.95 | 3.3089 | 3.18 | 3.04 | 3.22 | 3.01 | 3.09 | 23
| 1 | 3 | 3.17 | 2.6559 | 3.3267 | 3.0018 | 3.26 | 3.38 | 2.95 | < 5 | 3.74 | 3 | 3.3584 | 2.98 | 3.0977 | 3.13 | 3.07 | 2.91 | 3.12 | 3.01 | 16
| 1 | 4 | 3.17 | 2.7567 | 3.3301 | 3.5998 | 3.23 | 3.63 | 2.84 | < 5 | 4.02 | 3 | 3.3232 | 3.2 | 2.5978 | 3.13 | 2.94 | 2.95 | 3.42 | 3.02 | 3.1584
| 1 | 5 | 3.16 | 2.7008 | 3.1222 | 3.5437 | 3.44 | 3.52 | 2.9 | < 5 | 3.46 | 3.2 | 3.3572 | 3.07 | 2.5773 | 3.33 | 2.86 | 3.1 | 3.11 | 3.01 | 20
| 1 | 6 | 3.25 | 2.7621 | 3.7856 | 2.1991 | 3.2 | 3.49 | 2.86 | < 5 | 3.62 | 3.2 | 4.0223 | 3.09 | 2.9953 | 3.67 | 2.93 | 3.02 | 3.19 | 3.02
| Mean |  | 3.1833 | 2.7216 | 3.3528 | 3.2108 | 3.2233 | 3.45 | 2.895 |  | 3.5767 | 3.3 | 3.4501 | 3.055 | 3.0023 | 3.2833 | 2.9667 | 3.07 | 3.1833 | 3.0383
| Median |  | 3.17 | 2.7271 | 3.3108 | 3.2727 | 3.215 | 3.435 | 2.89 |  | 3.54 | 3.2 | 3.3444 | 3.055 | 3.0465 | 3.22 | 2.95 | 3.06 | 3.155 | 3.02
| Std Dev. |  | 0.10347 | 0.04055 | 0.22549 | 0.70825 | 0.13779 | 0.1173 | 0.043704 |  | 0.27602 | 0.34641 | 0.28099 | 0.088713 | 0.35662 | 0.20491 | 0.076855 | 0.13297 | 0.14137 | 0.03656
| Rel.Std.Dev. |  | 3.25% | 1.49% | 6.725% | 22.06% | 4.275% | 3.4% | 1.51% |  | 7.717% | 10.5% | 8.145% | 2.904% | 11.88% | 6.241% | 2.591% | 4.331% | 4.441% | 1.203%
| PDM3 |  | 0.7885% | -13.83% | 6.153% | 1.659% | 2.055% | 9.232% | -8.34% |  | 13.24% | 4.482% | 9.235% | -3.275% | -4.944% | 3.955% | -6.071% | -2.8% | 0.7885% | -3.802%
| Z-Score (Absolute) |  | 0 | 1.98 | 0.73 | 0.12 | 0.17 | 1.14 | 1.24 | 2.93 | 1.69 | 0.5 | 1.14 | 0.55 | 0.78 | 0.43 | 0.93 | 0.49 | 0 | 0.62
| Analytical results for Ca in OREAS 255c (Certified Value 1.26 wt.%) | [19 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.2727 | 1.2329 | 1.2291 | 1.9058 | 1.22 | 1.2 | 1.22 | 1.27 | 1.24 | 1.26 | 1.3074 | 1.33 | 1.2761 | 1.25 | 1.24 | 1.28 | 1.24 | 1.096 | 1
| 1 | 2 | 1.2381 | 1.2444 | 1.2208 | 1.9495 | 1.25 | 1.23 | 1.22 | 1.24 | 1.3 | 1.3 | 1.2669 | 1.31 | 1.2817 | 1.25 | 1.22 | 1.29 | 1.24 | 1.169 | 0
| 1 | 3 | 1.2608 | 1.257 | 1.2636 | 1.9404 | 1.24 | 1.28 | 1.22 | 1.26 | 1.28 | 1.35 | 1.2749 | 1.3 | 1.2829 | 1.24 | 1.3 | 1.3 | 1.25 | 1.173 | 16
| 1 | 4 | 1.2791 | 1.2632 | 1.2544 | 1.9273 | 1.24 | 1.29 | 1.22 | 1.23 | 1.24 | 1.35 | 1.2967 | 1.27 | 1.2693 | 1.2 | 1.3 | 1.3 | 1.23 | 1.185 | 1.2642
| 1 | 5 | 1.2809 | 1.257 | 1.2484 | 1.9624 | 1.24 | 1.23 | 1.23 | 1.26 | 1.31 | 1.31 | 1.2941 | 1.31 | 1.2693 | 1.23 | 1.26 | 1.32 | 1.26 | 1.19 | 21
| 1 | 6 | 1.2374 | 1.284 | 1.2891 | 1.9428 | 1.24 | 1.26 | 1.22 | 1.25 | 1.28 | 1.36 | 1.2856 | 1.32 | 1.263 | 1.22 | 1.26 | 1.3 | 1.22 | 1.167
| Mean |  | 1.2615 | 1.2564 | 1.2509 | 1.938 | 1.2383 | 1.2483 | 1.2217 | 1.2517 | 1.275 | 1.3217 | 1.2876 | 1.3067 | 1.2737 | 1.2317 | 1.2633 | 1.2983 | 1.24 | 1.1633
| Median |  | 1.2668 | 1.257 | 1.2514 | 1.9416 | 1.24 | 1.245 | 1.22 | 1.255 | 1.28 | 1.33 | 1.2899 | 1.31 | 1.2727 | 1.235 | 1.26 | 1.3 | 1.24 | 1.171
| Std Dev. |  | 0.0197 | 0.017363 | 0.024573 | 0.019533 | 0.0098319 | 0.034303 | 0.0040825 | 0.01472 | 0.029496 | 0.038687 | 0.014904 | 0.020656 | 0.0078497 | 0.019408 | 0.032042 | 0.013292 | 0.014142 | 0.034215
| Rel.Std.Dev. |  | 1.562% | 1.382% | 1.964% | 1.008% | 0.794% | 2.748% | 0.3342% | 1.176% | 2.313% | 2.927% | 1.158% | 1.581% | 0.6163% | 1.576% | 2.536% | 1.024% | 1.14% | 2.941%
| PDM3 |  | -0.2117% | -0.6136% | -1.049% | 53.3% | -2.044% | -1.253% | -3.363% | -0.9896% | 0.8562% | 4.548% | 1.855% | 3.361% | 0.7534% | -2.572% | -0.06669% | 2.702% | -1.912% | -7.977%
| Z-Score (Absolute) |  | 0.09 | 0.09 | 0.27 | 23.13 | 0.7 | 0.36 | 1.27 | 0.25 | 0.55 | 2.14 | 0.98 | 1.63 | 0.5 | 0.93 | 0.15 | 1.34 | 0.65 | 3.26
| Analytical results for Cd in OREAS 255c (Certified Value 0.095 ppm) | [19 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.1 | < 0.5 | 0.5838 | 0.11 | < 0.3 | 0.28 | 1.3091 | 0.08 | 0.2 | 0.11095 | 0.1 | 0.073371 | 0.31 | 0.09 | 0.08 | 0.08 |  |  | 1
| 1 | 2 | 0.11 | < 0.5 | 0.4894 | 0.12 | < 0.3 | 0.29 | 1.1461 | 0.09 | 0.3 | 0.1175 | 0.08 | 0.076637 | 0.29 | 0.08 | 0.09 | 0.09 |  |  | 24
| 1 | 3 | 0.14 | < 0.5 | 0.6031 | 0.11 | < 0.3 | 0.29 | 1.2094 | 0.09 | 0.3 | 0.1114 | 0.09 | 0.10799 | 0.32 | 0.08 | 0.08 | 0.1 |  |  | 16
| 1 | 4 | 0.11 | < 0.5 | 0.7161 | 0.12 | < 0.3 | 0.28 | 1.0943 | 0.11 | 0.3 | 0.11205 | 0.08 | 0.099857 | 0.3 | 0.08 | 0.09 | 0.1 |  |  | 0.094642
| 1 | 5 | 0.08 | < 0.5 | 0.4673 | 0.11 | < 0.3 | 0.28 | 1.28 | 0.08 | 0.2 | 0.1148 | 0.09 | 0.07126 | 0.3 | 0.09 | 0.09 | 0.08 |  |  | 22
| 1 | 6 | 0.1 | < 0.5 | 0.5806 | 0.1 | < 0.3 | 0.28 | 1.2435 | 0.15 | 0.2 | 0.1148 | 0.09 | 0.060037 | 0.31 | 0.08 | 0.09 | 0.09
| Mean |  | 0.10667 |  | 0.57338 | 0.11167 |  | 0.28333 | 1.2137 | 0.1 | 0.25 | 0.11358 | 0.088333 | 0.081525 | 0.305 | 0.083333 | 0.086667 | 0.09
| Median |  | 0.105 |  | 0.5822 | 0.11 |  | 0.28 | 1.2265 | 0.09 | 0.25 | 0.11342 | 0.09 | 0.075004 | 0.305 | 0.08 | 0.09 | 0.09
| Std Dev. |  | 0.019664 |  | 0.089128 | 0.0075277 |  | 0.005164 | 0.081528 | 0.026833 | 0.054772 | 0.0025438 | 0.0075277 | 0.018406 | 0.010488 | 0.005164 | 0.005164 | 0.0089443
| Rel.Std.Dev. |  | 18.43% |  | 15.54% | 6.741% |  | 1.823% | 6.717% | 26.83% | 21.91% | 2.24% | 8.522% | 22.58% | 3.439% | 6.197% | 5.958% | 9.938%
| PDM3 |  | 12.71% |  | 505.8% | 17.99% |  | 199.4% | 1182% | 5.662% | 164.2% | 20.01% | -6.665% | -13.86% | 222.3% | -11.95% | -8.426% | -4.904%
| Z-Score (Absolute) |  | 0.12 | 3.29 | 10.99 | 0 | 0.91 | 4.09 | 26.23 | 0.28 | NA | 0.05 | 0.56 | 0.72 | 4.6 | 0.67 | 0.59 | 0.52
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 255c (Certified Value 73 ppm) | [19 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
1
| 1 | 1 | 69.81 | 72.2586 | 73.1176 | 68.1 | 61 | 74.2 | 72 | 55.9271 | 72.9 | 74.9 | 74.2 | 77.1 | 73.7 | 80.3 |  |  |  |  | 1
| 1 | 2 | 72.15 | 74.7645 | 70.2737 | 70.9 | 62 | 73.7 | 68 | 56.5614 | 72.3 | 74.4 | 72.7 | 73.8 | 73.4 | 80.21 |  |  |  |  | 25
| 1 | 3 | 73.69 | 73.8081 | 71.462 | 71.5 | 61 | 76.3 | 74 | 56.7686 | 75.6 | 74 | 75.5 | 75.7 | 74 | 80.91 |  |  |  |  | 16
| 1 | 4 | 69.46 | 73.9375 | 72.9498 | 72.1 | 61 | 73.7 | 73 | 56.3188 | 72.3 | 74.6 | 73.8 | 73.6 | 76.1 | 81.81 |  |  |  |  | 73.3982
| 1 | 5 | 72.23 | 73.877 | 73.1046 | 69.7 | 63 | 75.1 | 77 | 55.9839 | 72.2 | 74.7 | 73.9 | 75.2 | 74.8 | 80.77 |  |  |  |  | 23
| 1 | 6 | 72.72 | 73.2878 | 73.3238 | 71.8 | 63 | 75 | 77 | 55.9231 | 73.5 | 74.4 | 74.6 | 74.4 | 73.1 | 79.83
| Mean |  | 71.6767 | 73.6556 | 72.3719 | 70.6833 | 61.8333 | 74.6667 | 73.5 | 56.2471 | 73.1333 | 74.5 | 74.1167 | 74.9667 | 74.1833 | 80.6383
| Median |  | 72.19 | 73.8426 | 73.0272 | 71.2 | 61.5 | 74.6 | 73.5 | 56.1514 | 72.6 | 74.5 | 74.05 | 74.8 | 73.85 | 80.535
| Std Dev. |  | 1.6775 | 0.83285 | 1.2299 | 1.5237 | 0.98319 | 1.0053 | 3.3912 | 0.36126 | 1.3064 | 0.30984 | 0.92826 | 1.3186 | 1.1053 | 0.69502
| Rel.Std.Dev. |  | 2.34% | 1.131% | 1.699% | 2.156% | 1.59% | 1.346% | 4.614% | 0.6423% | 1.786% | 0.4159% | 1.252% | 1.759% | 1.49% | 0.8619%
| PDM3 |  | -2.346% | 0.3507% | -1.398% | -3.699% | -15.76% | 1.728% | 0.1387% | -23.37% | -0.3609% | 1.501% | 0.9788% | 2.137% | 1.07% | 9.864%
| Z-Score (Absolute) |  | 1.12 | 0.05 | 0.71 | 1.7 | 6.9 | 0.64 | 0.05 | 10.19 | 0.26 | 0.54 | 0.32 | 0.82 | 0.36 | 4.15
| Analytical results for Co in OREAS 255c (Certified Value 21 ppm) | [19 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 21.2 | 20.934 | 21.3169 | 20.297 | 20 | 21 | 21.3 | 23.7592 | 20.2 | 19.6 | 18.4176 | 19.5 | 20.5397 | 21.19 | 20.9 | 22.8 | 20.6 | 23.1 | 1
| 1 | 2 | 21.5 | 21.267 | 20.6144 | 20.9936 | 21 | 21 | 21.2 | 23.2932 | 20.5 | 19.8 | 18.1727 | 19.4 | 21.0684 | 21.13 | 21.7 | 22.2 | 20.3 | 23.7 | 26
| 1 | 3 | 21.4 | 21.042 | 21.4536 | 21.5559 | 21 | 21 | 21.3 | 23.6114 | 21.1 | 20.2 | 18.1398 | 20.3 | 21.5007 | 21.14 | 22.2 | 22.6 | 21 | 22.1 | 16
| 1 | 4 | 21.4 | 21.537 | 21.4835 | 22.0815 | 20 | 21 | 21.6 | 22.9567 | 20 | 20 | 18.5131 | 19 | 19.7887 | 20.98 | 22.4 | 22.2 | 21 | 22.6 | 20.9693
| 1 | 5 | 21.4 | 21.447 | 21.1165 | 21.0561 | 20 | 20 | 21.6 | 23.4462 | 21.2 | 19.8 | 18.4358 | 19.8 | 21.7318 | 21.14 | 21.6 | 22 | 20.8 | 23.5 | 24
| 1 | 6 | 21.1 | 21.699 | 21.5586 | 20.9066 | 20 | 21 | 21.4 | 23.7591 | 20.9 | 19.8 | 18.4279 | 20 | 20.1871 | 21.21 | 21.1 | 22.2 | 20.4 | 23
| Mean |  | 21.3333 | 21.321 | 21.2572 | 21.1485 | 20.3333 | 20.8333 | 21.4 | 23.471 | 20.65 | 19.8667 | 18.3511 | 19.6667 | 20.8027 | 21.1317 | 21.65 | 22.3333 | 20.6833 | 23
| Median |  | 21.4 | 21.357 | 21.3852 | 21.0249 | 20 | 21 | 21.35 | 23.5288 | 20.7 | 19.8 | 18.4228 | 19.65 | 20.804 | 21.14 | 21.65 | 22.2 | 20.7 | 23.05
| Std Dev. |  | 0.15055 | 0.29528 | 0.35133 | 0.6087 | 0.5164 | 0.40825 | 0.16733 | 0.31065 | 0.49295 | 0.20656 | 0.15508 | 0.46332 | 0.76122 | 0.08085 | 0.58907 | 0.30111 | 0.29944 | 0.58652
| Rel.Std.Dev. |  | 0.7057% | 1.385% | 1.653% | 2.878% | 2.54% | 1.96% | 0.7819% | 1.324% | 2.387% | 1.04% | 0.8451% | 2.356% | 3.659% | 0.3826% | 2.721% | 1.348% | 1.448% | 2.55%
| PDM3 |  | 1.736% | 1.677% | 1.373% | 0.8543% | -3.033% | -0.6485% | 2.054% | 11.93% | -1.523% | -5.258% | -12.49% | -6.212% | -0.7944% | 0.7743% | 3.246% | 6.505% | -1.364% | 9.684%
| Z-Score (Absolute) |  | 0.28 | 0.26 | 0.17 | 0.01 | 1.15 | 0.44 | 0.37 | 3.32 | 0.7 | 1.81 | 3.97 | 2.1 | 0.48 | 0.01 | 0.73 | 1.7 | 0.65 | 2.65
| Analytical results for Cr in OREAS 255c (Certified Value 169 ppm) | [19 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 193 | 168.3158 | 174.5542 | 164 | 137 | 183 | 161.837 | 166 | 140 | 152.5473 | 177 | 169.5886 | 167 | 170 | 170 | 160 | 160 |  | 1
| 1 | 2 | 191 | 162.5188 | 181.4278 | 172 | 167 | 183 | 161.0233 | 170 | 154 | 152.0487 | 178 | 157.675 | 165 | 171 | 166 | 156 | 167 |  | 27
| 1 | 3 | 189 | 166.1998 | 195.0158 | 164 | 143 | 183 | 160.5567 | 173 | 156 | 151.4467 | 176 | 169.6752 | 167 | 176 | 170 | 163 | 162 |  | 16
| 1 | 4 | 196 | 172.52 | 168.6178 | 168 | 144 | 185 | 157.258 | 167 | 159 | 152.8471 | 173 | 156.8181 | 168 | 170 | 167 | 158 | 168 |  | 168.6707
| 1 | 5 | 195 | 164.5421 | 188.4294 | 165 | 135 | 183 | 160.622 | 171 | 147 | 152.9438 | 180 | 153.4293 | 169 | 172 | 171 | 161 | 173 |  | 25
| 1 | 6 | 194 | 167.8957 | 176.2417 | 175 | 140 | 184 | 162.7301 | 170 | 149 | 151.9475 | 176 | 159.1115 | 167 | 173 | 169 | 161 | 176
| Mean |  | 193 | 166.9987 | 180.7144 | 168 | 144.3333 | 183.5 | 160.6712 | 169.5 | 150.8333 | 152.2968 | 176.6667 | 161.0496 | 167.1667 | 172 | 168.8333 | 159.8333 | 167.6667
| Median |  | 193.5 | 167.0477 | 178.8348 | 166.5 | 141.5 | 183 | 160.8226 | 170 | 151.5 | 152.298 | 176.5 | 158.3933 | 167 | 171.5 | 169.5 | 160.5 | 167.5
| Std Dev. |  | 2.6077 | 3.4574 | 9.6834 | 4.6043 | 11.6218 | 0.83666 | 1.8648 | 2.5884 | 6.9113 | 0.58145 | 2.3381 | 6.9056 | 1.3292 | 2.2804 | 1.9408 | 2.4833 | 6.1536
| Rel.Std.Dev. |  | 1.351% | 2.07% | 5.358% | 2.741% | 8.052% | 0.4559% | 1.161% | 1.527% | 4.582% | 0.3818% | 1.323% | 4.288% | 0.7951% | 1.326% | 1.15% | 1.554% | 3.67%
| PDM3 |  | 14.42% | -0.9913% | 7.14% | -0.3976% | -14.43% | 8.792% | -4.743% | 0.4917% | -10.58% | -9.708% | 4.741% | -4.518% | -0.8917% | 1.974% | 0.09644% | -5.239% | -0.5952%
| Z-Score (Absolute) |  | 2.44 | 0.06 | 1.26 | 0.03 | 2.25 | 1.53 | 0.67 | 0.18 | 1.62 | 1.48 | 0.87 | 0.64 | 0.05 | 0.42 | 0.11 | 0.76 | 0
| Analytical results for Cs in OREAS 255c (Certified Value 3.81 ppm) | [19 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 3.83 | 4.0873 | 3.7 | 3.53 | 3.87 | 3.87 | 3.4813 | 3.67 | 3.8749 | 3.65 | 3.57 | 3.94 | 3.85 | 4.1 |  |  |  |  | 1
| 1 | 2 | 3.85 | 3.8168 | 3.8 | 3.69 | 3.78 | 3.88 | 3.6271 | 3.82 | 3.8207 | 3.61 | 3.71 | 3.93 | 3.89 | 4.11 |  |  |  |  | 28
| 1 | 3 | 4.01 | 3.8473 | 3.7 | 3.76 | 3.74 | 3.98 | 3.6117 | 3.89 | 3.7506 | 3.63 | 3.69 | 3.96 | 3.96 | 4.01 |  |  |  |  | 16
| 1 | 4 | 3.92 | 3.8853 | 3.7 | 3.64 | 3.79 | 3.77 | 3.6638 | 3.72 | 3.2354 | 3.63 | 3.71 | 4.11 | 3.99 | 4.02 |  |  |  |  | 3.8137
| 1 | 5 | 3.92 | 4.0826 | 3.8 | 3.66 | 3.78 | 3.99 | 3.7276 | 3.76 | 3.1483 | 3.65 | 3.67 | 3.94 | 3.89 | 4.16 |  |  |  |  | 26
| 1 | 6 | 3.95 | 3.8137 | 3.7 | 3.81 | 3.86 | 3.96 | 3.6252 | 3.86 | 3.715 | 3.65 | 3.69 | 3.87 | 3.81 | 4.04
| Mean |  | 3.9133 | 3.9222 | 3.7333 | 3.6817 | 3.8033 | 3.9083 | 3.6228 | 3.7867 | 3.5908 | 3.6367 | 3.6733 | 3.9583 | 3.8983 | 4.0733
| Median |  | 3.92 | 3.8663 | 3.7 | 3.675 | 3.785 | 3.92 | 3.6261 | 3.79 | 3.7328 | 3.64 | 3.69 | 3.94 | 3.89 | 4.07
| Std Dev. |  | 0.065929 | 0.12869 | 0.05164 | 0.09786 | 0.050859 | 0.084715 | 0.081035 | 0.084774 | 0.31517 | 0.01633 | 0.052789 | 0.080353 | 0.067057 | 0.059217
| Rel.Std.Dev. |  | 1.685% | 3.281% | 1.383% | 2.658% | 1.337% | 2.168% | 2.237% | 2.239% | 8.777% | 0.449% | 1.437% | 2.03% | 1.72% | 1.454%
| PDM3 |  | 2.612% | 2.843% | -2.108% | -3.463% | -0.2727% | 2.481% | -5.007% | -0.7097% | -5.845% | -4.643% | -3.681% | 3.792% | 2.218% | 6.807%
| Z-Score (Absolute) |  | 0.66 | 0.71 | 0.35 | 0.64 | 0.05 | 0.64 | 0.97 | 0.05 | 1.15 | 0.89 | 0.68 | 0.92 | 0.58 | 1.56
| Analytical results for Cu in OREAS 255c (Certified Value 59 ppm) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 60.2 | 57.3141 | 52.473 | 57 | 60 | 62 | 60.544 | 56.9 | 61.7 | 44.7935 | 54.6 | 51.9344 | 57.9 | 60.4 | 62.1 | 61.1 | 51.4 |  | 1
| 1 | 2 | 59.1 | 56.146 | 53.3998 | 59 | 61 | 63 | 61.985 | 57.7 | 63.5 | 44.9142 | 54.3 | 52.2522 | 58.2 | 61.7 | 58.6 | 59.4 | 52.2 |  | 29
| 1 | 3 | 61.1 | 58.0924 | 53.7199 | 58 | 63 | 62 | 61.9905 | 58.3 | 64.7 | 46.136 | 56.5 | 53.4657 | 59.5 | 61.6 | 59.9 | 60.7 | 55.4 |  | 16
| 1 | 4 | 60.1 | 58.1071 | 53.8791 | 58 | 61 | 63 | 59.9775 | 56.3 | 63.8 | 45.6092 | 53.3 | 50.9239 | 58.7 | 64.5 | 59.3 | 61.3 | 55.6 |  | 58.5061
| 1 | 5 | 60.2 | 57.3644 | 53.4905 | 59 | 60 | 63 | 60.577 | 59 | 63 | 46.1878 | 55.3 | 52.8008 | 57.8 | 61.5 | 58.8 | 61.1 | 55.3 |  | 27
| 1 | 6 | 59.6 | 58.1621 | 53.7185 | 58 | 63 | 62 | 62.88 | 58.3 | 63.7 | 45.8587 | 55.7 | 51.029 | 57.7 | 61.8 | 59.7 | 59.4 | 54.8
| Mean |  | 60.05 | 57.531 | 53.4468 | 58.1667 | 61.3333 | 62.5 | 61.3257 | 57.75 | 63.4 | 45.5832 | 54.95 | 52.0677 | 58.3 | 61.9167 | 59.7333 | 60.5 | 54.1167
| Median |  | 60.15 | 57.7284 | 53.6045 | 58 | 61 | 62.5 | 61.281 | 58 | 63.6 | 45.7339 | 54.95 | 52.0933 | 58.05 | 61.65 | 59.5 | 60.9 | 55.05
| Std Dev. |  | 0.67157 | 0.77954 | 0.50748 | 0.75277 | 1.3663 | 0.54772 | 1.1209 | 0.9995 | 1 | 0.60315 | 1.1274 | 0.99305 | 0.68993 | 1.3644 | 1.2628 | 0.87407 | 1.8313
| Rel.Std.Dev. |  | 1.118% | 1.355% | 0.9495% | 1.294% | 2.228% | 0.8764% | 1.828% | 1.731% | 1.577% | 1.323% | 2.052% | 1.907% | 1.183% | 2.204% | 2.114% | 1.445% | 3.384%
| PDM3 |  | 2.639% | -1.667% | -8.647% | -0.5802% | 4.832% | 6.826% | 4.819% | -1.292% | 8.365% | -22.09% | -6.078% | -11% | -0.3523% | 5.829% | 2.098% | 3.408% | -7.503%
| Z-Score (Absolute) |  | 0.39 | 0.17 | 1.08 | 0.03 | 0.67 | 0.93 | 0.67 | 0.12 | 1.13 | 2.83 | 0.74 | 1.39 | 0 | 0.8 | 0.32 | 0.49 | 0.93
| Analytical results for Dy in OREAS 255c (Certified Value 3.58 ppm) | [19 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 3.48; 1.03: 3.5197; 1.06: 4.0494; 1.07: 3.6; 1.09: 3.49; 1.10: 2.6092; 1.11: 3.38 |  | 1
| 1 | 2 | 1.01: 3.53; 1.03: 3.6061; 1.06: 3.7358; 1.07: 3.7; 1.09: 3.53; 1.10: 2.5609; 1.11: 3.29 |  | 3
| 1 | 3 | 1.01: 3.58; 1.03: 3.4858; 1.06: 3.8807; 1.07: 3.7; 1.09: 3.5; 1.10: 2.5932; 1.11: 3.31 |  | 16
| 1 | 4 | 1.01: 3.62; 1.03: 3.6193; 1.06: 3.8599; 1.07: 3.8; 1.09: 3.49; 1.10: 2.6278; 1.11: 3.28 |  | 3.584
| 1 | 5 | 1.01: 3.6; 1.03: 3.4096; 1.06: 4.2097; 1.07: 3.7; 1.09: 3.51; 1.10: 2.6222; 1.11: 3.43 |  | 28
| 1 | 6 | 1.01: 3.57; 1.03: 3.5379; 1.06: 3.7873; 1.07: 3.7; 1.09: 3.45; 1.10: 2.5898; 1.11: 3.43
| Mean |  | 1.01: 3.5633; 1.03: 3.5297; 1.06: 3.9205; 1.07: 3.7; 1.09: 3.495; 1.10: 2.6005; 1.11: 3.3533
| Median |  | 1.01: 3.575; 1.03: 3.5288; 1.06: 3.8703; 1.07: 3.7; 1.09: 3.495; 1.10: 2.6012; 1.11: 3.345
| Std Dev. |  | 1.01: 0.050859; 1.03: 0.077941; 1.06: 0.17741; 1.07: 0.063246; 1.09: 0.026646; 1.10: 0.024631; 1.11: 0.068896
| Rel.Std.Dev. |  | 1.01: 1.427%; 1.03: 2.208%; 1.06: 4.525%; 1.07: 1.709%; 1.09: 0.7624%; 1.10: 0.9472%; 1.11: 2.055%
| PDM3 |  | 1.01: -0.5767%; 1.03: -1.515%; 1.06: 9.389%; 1.07: 3.237%; 1.09: -2.483%; 1.10: -27.44%; 1.11: -6.436%
| Z-Score (Absolute) |  | 1.01: 0.13; 1.03: 0; 1.06: 1.55; 1.07: 0.67; 1.09: 0.14; 1.10: 3.68; 1.11: 0.7
| Analytical results for Er in OREAS 255c (Certified Value 1.69 ppm) | [19 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 1.61; 1.03: 1.6849; 1.06: 1.6746; 1.07: 1.7; 1.09: 1.62; 1.10: 2.918; 1.11: 1.75 |  | 1
| 1 | 2 | 1.01: 1.69; 1.03: 1.7245; 1.06: 1.6625; 1.07: 1.7; 1.09: 1.57; 1.10: 2.8981; 1.11: 1.75 |  | 4
| 1 | 3 | 1.01: 1.69; 1.03: 1.7117; 1.06: 1.7051; 1.07: 1.8; 1.09: 1.62; 1.10: 2.8632; 1.11: 1.68 |  | 16
| 1 | 4 | 1.01: 1.68; 1.03: 1.6752; 1.06: 1.5961; 1.07: 1.8; 1.09: 1.59; 1.10: 2.947; 1.11: 1.65 |  | 1.6869
| 1 | 5 | 1.01: 1.64; 1.03: 1.6763; 1.06: 1.749; 1.07: 1.8; 1.09: 1.62; 1.10: 2.9214; 1.11: 1.72 |  | 29
| 1 | 6 | 1.01: 1.68; 1.03: 1.7224; 1.06: 1.5483; 1.07: 1.8; 1.09: 1.6; 1.10: 2.9096; 1.11: 1.77
| Mean |  | 1.01: 1.665; 1.03: 1.6992; 1.06: 1.6559; 1.07: 1.7667; 1.09: 1.6033; 1.10: 2.9095; 1.11: 1.72
| Median |  | 1.01: 1.68; 1.03: 1.6983; 1.06: 1.6686; 1.07: 1.8; 1.09: 1.61; 1.10: 2.9138; 1.11: 1.735
| Std Dev. |  | 1.01: 0.032711; 1.03: 0.02297; 1.06: 0.072974; 1.07: 0.05164; 1.09: 0.020656; 1.10: 0.027886; 1.11: 0.046476
| Rel.Std.Dev. |  | 1.01: 1.965%; 1.03: 1.352%; 1.06: 4.407%; 1.07: 2.923%; 1.09: 1.288%; 1.10: 0.9584%; 1.11: 2.702%
| PDM3 |  | 1.01: -1.295%; 1.03: 0.7306%; 1.06: -1.833%; 1.07: 4.732%; 1.09: -4.951%; 1.10: 72.48%; 1.11: 1.965%
| Z-Score (Absolute) |  | 1.01: 0.53; 1.03: 0; 1.06: 0.67; 1.07: 1.05; 1.09: 1.49; 1.10: 18.88; 1.11: 0.32
| Analytical results for Eu in OREAS 255c (Certified Value 1.49 ppm) | [19 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 1.56; 1.03: 0.43455; 1.06: 1.4964; 1.07: 1.24; 1.09: 1.53; 1.10: 1.2868; 1.11: 1.54 |  | 1
| 1 | 2 | 1.01: 1.58; 1.03: 0.43529; 1.06: 1.4719; 1.07: 1.33; 1.09: 1.5; 1.10: 1.2568; 1.11: 1.6 |  | 5
| 1 | 3 | 1.01: 1.54; 1.03: 0.36483; 1.06: 1.4823; 1.07: 1.37; 1.09: 1.5; 1.10: 1.2693; 1.11: 1.48 |  | 16
| 1 | 4 | 1.01: 1.5; 1.03: 0.53602; 1.06: 1.5173; 1.07: 1.41; 1.09: 1.51; 1.10: 1.2749; 1.11: 1.47 |  | 1.4858
| 1 | 5 | 1.01: 1.58; 1.03: 0.44414; 1.06: 1.5557; 1.07: 1.31; 1.09: 1.53; 1.10: 1.2868; 1.11: 1.57 |  | 30
| 1 | 6 | 1.01: 1.58; 1.03: 0.48987; 1.06: 1.4616; 1.07: 1.34; 1.09: 1.52; 1.10: 1.2674; 1.11: 1.5
| Mean |  | 1.01: 1.5567; 1.03: 0.45078; 1.06: 1.4975; 1.07: 1.3333; 1.09: 1.515; 1.10: 1.2737; 1.11: 1.5267
| Median |  | 1.01: 1.57; 1.03: 0.43972; 1.06: 1.4893; 1.07: 1.335; 1.09: 1.515; 1.10: 1.2721; 1.11: 1.52
| Std Dev. |  | 1.01: 0.032042; 1.03: 0.057842; 1.06: 0.034512; 1.07: 0.057504; 1.09: 0.013784; 1.10: 0.011732; 1.11: 0.052026
| Rel.Std.Dev. |  | 1.01: 2.058%; 1.03: 12.83%; 1.06: 2.305%; 1.07: 4.313%; 1.09: 0.9098%; 1.10: 0.9211%; 1.11: 3.408%
| PDM3 |  | 1.01: 4.767%; 1.03: -69.66%; 1.06: 0.7872%; 1.07: -10.26%; 1.09: 1.962%; 1.10: -14.28%; 1.11: 2.748%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 11.94; 1.06: 0; 1.07: 1.87; 1.09: 0.2; 1.10: 2.55; 1.11: 0.33
| Analytical results for Fe in OREAS 255c (Certified Value 4.4 wt.%) | [19 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 4.44 | 4.527 | 4.3899 | 3.3689 | 4.31 | 4 | 4.49 | 4.35 | 4.34 | 4 | 4.3772 | 4.44 | 4.4342 | 4.37 | 4.3 | 4.44 | 4.28 | 3.8 | 1
| 1 | 2 | 4.35 | 4.5714 | 4.4037 | 3.4356 | 4.35 | 4.07 | 4.46 | 4.41 | 4.48 | 4.08 | 4.34 | 4.45 | 4.4715 | 4.34 | 4.21 | 4.39 | 4.26 | 4.02 | 0
| 1 | 3 | 4.41 | 4.5519 | 4.4376 | 3.4549 | 4.43 | 4.22 | 4.47 | 4.39 | 4.48 | 4.18 | 4.2947 | 4.42 | 4.5025 | 4.29 | 4.44 | 4.38 | 4.29 | 4.07 | 16
| 1 | 4 | 4.47 | 4.541 | 4.4507 | 3.4035 | 4.42 | 4.18 | 4.45 | 4.34 | 4.32 | 4.15 | 4.3597 | 4.3 | 4.2998 | 4.25 | 4.47 | 4.43 | 4.26 | 4.04 | 4.4003
| 1 | 5 | 4.45 | 4.5874 | 4.4379 | 3.4181 | 4.42 | 4.14 | 4.48 | 4.38 | 4.52 | 4.08 | 4.3549 | 4.4 | 4.5339 | 4.33 | 4.3 | 4.48 | 4.32 | 4.08 | 31
| 1 | 6 | 4.38 | 4.6005 | 4.4453 | 3.4471 | 4.37 | 4.21 | 4.45 | 4.36 | 4.49 | 4.19 | 4.3425 | 4.41 | 4.3122 | 4.29 | 4.36 | 4.48 | 4.22 | 4.03
| Mean |  | 4.4167 | 4.5632 | 4.4275 | 3.4213 | 4.3833 | 4.1367 | 4.4667 | 4.3717 | 4.4383 | 4.1133 | 4.3448 | 4.4033 | 4.4257 | 4.3117 | 4.3467 | 4.4333 | 4.2717 | 4.0067
| Median |  | 4.425 | 4.5617 | 4.4378 | 3.4268 | 4.395 | 4.16 | 4.465 | 4.37 | 4.48 | 4.115 | 4.3487 | 4.415 | 4.4529 | 4.31 | 4.33 | 4.435 | 4.27 | 4.035
| Std Dev. |  | 0.045461 | 0.028189 | 0.02467 | 0.031831 | 0.048028 | 0.08641 | 0.01633 | 0.026394 | 0.085421 | 0.073121 | 0.027971 | 0.053914 | 0.098484 | 0.043089 | 0.097091 | 0.04274 | 0.033714 | 0.10386
| Rel.Std.Dev. |  | 1.029% | 0.6177% | 0.5572% | 0.9304% | 1.096% | 2.089% | 0.3656% | 0.6038% | 1.925% | 1.778% | 0.6438% | 1.224% | 2.225% | 0.9994% | 2.234% | 0.964% | 0.7893% | 2.592%
| PDM3 |  | 0.3714% | 3.701% | 0.6182% | -22.25% | -0.3861% | -5.992% | 1.508% | -0.6512% | 0.8638% | -6.522% | -1.262% | 0.0684% | 0.5761% | -2.015% | -1.219% | 0.7502% | -2.924% | -8.946%
| Z-Score (Absolute) |  | 0.45 | 2.15 | 0.58 | 11.05 | 0.07 | 2.78 | 1.03 | 0.07 | 0.7 | 3.05 | 0.38 | 0.3 | 0.56 | 0.76 | 0.36 | 0.65 | 1.22 | 4.29
| Analytical results for Ga in OREAS 255c (Certified Value 17.7 ppm) | [19 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 18.66 | 16.32 | 17.6532 | 18.8 | 10.5 | 19.67 | 19.6405 | 17.65 | 14.7353 | 16.2 | 18.0247 | 18.09 | 16 | 17.85 | 17.8 | 19.13 |  |  | 1
| 1 | 2 | 18.45 | 15.34 | 16.9929 | 18.7 | 9.9 | 19.31 | 18.942 | 17.6 | 14.8894 | 16.25 | 18.0411 | 17.93 | 16.5 | 17.35 | 17.25 | 19.76 |  |  | 33
| 1 | 3 | 18.8 | 16.45 | 17.3882 | 19 | 10.8 | 19.23 | 19.8319 | 18.1 | 14.5375 | 17 | 17.8231 | 17.85 | 16.55 | 17.45 | 17.6 | 19.22 |  |  | 16
| 1 | 4 | 18.69 | 16.68 | 17.0764 | 19.2 | 10.6 | 19.43 | 19.0042 | 17.2 | 14.7557 | 16 | 17.6116 | 18.02 | 16.25 | 17.15 | 17.8 | 18.3 |  |  | 17.6579
| 1 | 5 | 18.63 | 15.56 | 17.0486 | 19 | 10.1 | 19.41 | 19.4617 | 18 | 14.75 | 16.55 | 17.143 | 18.03 | 16.3 | 17.1 | 17.3 | 19.12 |  |  | 32
| 1 | 6 | 18.65 | 16.74 | 17.4173 | 18.8 | 10.7 | 19.53 | 19.9042 | 18.2 | 15.0858 | 16.75 | 17.0109 | 18.14 | 16.15 | 17.7 | 17.4 | 18.8
| Mean |  | 18.6467 | 16.1817 | 17.2628 | 18.9167 | 10.4333 | 19.43 | 19.4641 | 17.7917 | 14.7923 | 16.4583 | 17.6091 | 18.01 | 16.2917 | 17.4333 | 17.525 | 19.055
| Median |  | 18.655 | 16.385 | 17.2323 | 18.9 | 10.55 | 19.42 | 19.5511 | 17.825 | 14.7528 | 16.4 | 17.7174 | 18.025 | 16.275 | 17.4 | 17.5 | 19.125
| Std Dev. |  | 0.11361 | 0.59095 | 0.26287 | 0.18348 | 0.3559 | 0.15646 | 0.4109 | 0.37738 | 0.18271 | 0.37605 | 0.44276 | 0.10564 | 0.20837 | 0.29777 | 0.24444 | 0.48356
| Rel.Std.Dev. |  | 0.6093% | 3.652% | 1.523% | 0.97% | 3.411% | 0.8053% | 2.111% | 2.121% | 1.235% | 2.285% | 2.514% | 0.5866% | 1.279% | 1.708% | 1.395% | 2.538%
| PDM3 |  | 5.6% | -8.36% | -2.238% | 7.129% | -40.91% | 10.04% | 10.23% | 0.7577% | -16.23% | -6.793% | -0.2765% | 1.994% | -7.737% | -1.272% | -0.7525% | 7.912%
| Z-Score (Absolute) |  | 0.61 | 0.78 | 0.17 | 0.76 | 4.04 | 1.05 | 1.07 | 0.13 | 1.57 | 0.63 | 0.02 | 0.25 | 0.72 | 0.08 | 0.02 | 0.84
| Analytical results for Gd in OREAS 255c (Certified Value 5.05 ppm) | [19 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 4.9; 1.02: 5.7953; 1.03: 5.5783; 1.07: 4.5; 1.09: 5.07; 1.10: 2.4151; 1.11: 4.91 |  |  | 1
| 1 | 2 | 1.01: 5.1; 1.02: 6.0828; 1.03: 5.0661; 1.07: 4.9; 1.09: 5.05; 1.10: 2.2894; 1.11: 5.04 |  |  | 6
| 1 | 3 | 1.01: 5.04; 1.02: 5.8194; 1.03: 5.2571; 1.07: 5; 1.09: 5.02; 1.10: 2.5773; 1.11: 4.76 |  |  | 16
| 1 | 4 | 1.01: 5.2; 1.02: 5.8804; 1.03: 5.3885; 1.07: 5; 1.09: 5.05; 1.10: 2.4466; 1.11: 4.72 |  |  | 5.0515
| 1 | 5 | 1.01: 5.09; 1.02: 5.7394; 1.03: 5.6302; 1.07: 4.8; 1.09: 5.08; 1.10: 2.5596; 1.11: 4.92 |  |  | 33
| 1 | 6 | 1.01: 5.04; 1.02: 5.9953; 1.03: 5.1357; 1.07: 4.9; 1.09: 5.06; 1.10: 2.2858; 1.11: 4.92
| Mean |  | 1.01: 5.0617; 1.02: 5.8854; 1.03: 5.3426; 1.07: 4.85; 1.09: 5.055; 1.10: 2.429; 1.11: 4.8783
| Median |  | 1.01: 5.065; 1.02: 5.8499; 1.03: 5.3228; 1.07: 4.9; 1.09: 5.055; 1.10: 2.4308; 1.11: 4.915
| Std Dev. |  | 1.01: 0.098472; 1.02: 0.13027; 1.03: 0.23105; 1.07: 0.18708; 1.09: 0.020736; 1.10: 0.12613; 1.11: 0.11805
| Rel.Std.Dev. |  | 1.01: 1.945%; 1.02: 2.213%; 1.03: 4.325%; 1.07: 3.857%; 1.09: 0.4102%; 1.10: 5.193%; 1.11: 2.42%
| PDM3 |  | 1.01: 0.2007%; 1.02: 16.51%; 1.03: 5.763%; 1.07: -3.989%; 1.09: 0.06872%; 1.10: -51.92%; 1.11: -3.429%
| Z-Score (Absolute) |  | 1.01: 0.02; 1.02: 2.73; 1.03: 0.95; 1.07: 0.67; 1.09: 0; 1.10: 8.64; 1.11: 0.58
| Analytical results for Ge in OREAS 255c (Indicative Value 0.16 ppm) | [19 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.5 | < 0.2 | 1.7 | < 0.1 | 0.1 | 0.24 | < 0.05 | 0.2202 | 0.11 | 0.1981 | 0.14 | 0.23 | 0.1 | [5 empty] | 1
| 1 | 2 | 1.4 | < 0.2 | 1.7 | < 0.1 | 0.11 | 0.22 | 0.07 | 0.2319 | 0.1 | 0.19037 | 0.14 | 0.23 | 0.11 | [5 empty] | 1
| 1 | 3 | 1.5 | < 0.2 | 1.7 | < 0.1 | 0.11 | 0.25 | 0.24 | 0.1238 | 0.11 | 0.19529 | 0.13 | 0.23 | 0.12 | [5 empty] | 16
| 1 | 4 | 1.6 | < 0.2 | 1.8 | < 0.1 | 0.11 | 0.21 | 0.14 | 0.1403 | 0.09 | 0.14696 | 0.14 | 0.23 | 0.11 | [5 empty] | 0.15656
| 1 | 5 | 1.5 | < 0.2 | 1.8 | < 0.1 | 0.11 | 0.29 | 0.08 | 0.0862 | 0.11 | 0.15215 | 0.15 | 0.23 | 0.13 | [5 empty] | 7
| 1 | 6 | 1.4 | < 0.2 | 1.7 | < 0.1 | 0.11 | 0.21 | < 0.05 | 0.1703 | 0.11 | 0.20379 | 0.14 | 0.22 | 0.12
| Mean |  | 1.4833 |  | 1.7333 |  | 0.10833 | 0.23667 | 0.1325 | 0.16212 | 0.105 | 0.18111 | 0.14 | 0.22833 | 0.115
| Median |  | 1.5 |  | 1.7 |  | 0.11 | 0.23 | 0.11 | 0.1553 | 0.11 | 0.19283 | 0.14 | 0.23 | 0.115
| Std Dev. |  | 0.075277 |  | 0.05164 |  | 0.0040825 | 0.030768 | 0.078049 | 0.0566 | 0.0083666 | 0.024878 | 0.0063246 | 0.0040825 | 0.010488
| Rel.Std.Dev. |  | 5.075% |  | 2.979% |  | 3.768% | 13% | 58.91% | 34.91% | 7.968% | 13.74% | 4.518% | 1.788% | 9.12%
| PDM3 |  | 847.4% |  | 1007% |  | -30.8% | 51.16% | -15.37% | 3.548% | -32.93% | 15.68% | -10.58% | 45.84% | -26.55%
| Z-Score (Absolute) |  | 22.03 | 0.66 | 26.14 | 1.48 | 0.52 | 1.59 | 0.71 | 0.36 | 0.57 | 0.67 | 0 | 1.45 | 0.41
| Analytical results for Hf in OREAS 255c (Certified Value 4.39 ppm) | [19 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 4.54 | 4.7949 | 4.5011 | 4.1 | 1.1 | 4 | 4.5 | 4.4 | 4.5305 | 4 | 4.516 | 4.13 | 4.3 | 4.7 | 4.5 | 4.22 |  |  | 1
| 1 | 2 | 4.67 | 4.8316 | 4.1923 | 4.3 | 4.1 | 4 | 4.6 | 4.1 | 4.6323 | 4.1 | 4.2837 | 4.08 | 4.3 | 4.5 | 4.2 | 4.25 |  |  | 7
| 1 | 3 | 4.67 | 5.2602 | 4.3526 | 4.3 | 2.3 | 4 | 4.7 | 4.5 | 4.7433 | 4.2 | 4.2887 | 4.17 | 4.5 | 4.6 | 4.4 | 4.03 |  |  | 16
| 1 | 4 | 4.67 | 4.9767 | 4.3261 | 4.4 | 0.2 | 3.9 | 4.6 | 4.3 | 4.6943 | 4 | 3.4498 | 4.12 | 4.4 | 4.4 | 4.5 | 4.09 |  |  | 4.3858
| 1 | 5 | 4.7 | 5.0911 | 4.7389 | 4.4 | 0.5 | 3.9 | 4.6 | 4.4 | 4.836 | 4.2 | 3.3371 | 4.09 | 4.3 | 4.6 | 4.3 | 4.22 |  |  | 34
| 1 | 6 | 4.66 | 4.9206 | 4.2349 | 4.2 | 0.4 | 3.9 | 4.6 | 4.6 | 4.7282 | 4.2 | 4.1595 | 4.09 | 4.4 | 4.5 | 4.5 | 4.2
| Mean |  | 4.6517 | 4.9792 | 4.391 | 4.2833 | 1.4333 | 3.95 | 4.6 | 4.3833 | 4.6941 | 4.1167 | 4.0058 | 4.1133 | 4.3667 | 4.55 | 4.4 | 4.1683
| Median |  | 4.67 | 4.9486 | 4.3394 | 4.3 | 0.8 | 3.95 | 4.6 | 4.4 | 4.7112 | 4.15 | 4.2216 | 4.105 | 4.35 | 4.55 | 4.45 | 4.21
| Std Dev. |  | 0.056362 | 0.1736 | 0.20133 | 0.1169 | 1.5122 | 0.054772 | 0.063246 | 0.17224 | 0.10421 | 0.098319 | 0.48939 | 0.033862 | 0.08165 | 0.10488 | 0.12649 | 0.087502
| Rel.Std.Dev. |  | 1.212% | 3.487% | 4.585% | 2.729% | 105.5% | 1.387% | 1.375% | 3.929% | 2.22% | 2.388% | 12.22% | 0.8232% | 1.87% | 2.305% | 2.875% | 2.099%
| PDM3 |  | 6.062% | 13.53% | 0.1175% | -2.337% | -67.32% | -9.937% | 4.884% | -0.05643% | 7.029% | -6.137% | -8.664% | -6.213% | -0.4364% | 3.744% | 0.3236% | -4.959%
| Z-Score (Absolute) |  | 0.77 | 1.69 | 0.04 | 0.26 | 8.21 | 1.19 | 0.63 | 0.02 | 0.89 | 0.72 | 1.03 | 0.73 | 0.02 | 0.49 | 0.07 | 0.58
| Analytical results for Hg in OREAS 255c (Indicative Value 0.24 ppm) | [19 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.02: 1.10; 1.03: 1.14 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*OES/MS; 1.03: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.01: < 1; 1.02: < 2; 1.03: 0.2352 |  |  | 1
| 1 | 2 | 1.01: < 1; 1.02: < 2; 1.03: 0.2348 |  |  | 8
| 1 | 3 | 1.01: < 1; 1.02: < 2; 1.03: 0.2379 |  |  | 16
| 1 | 4 | 1.01: < 1; 1.02: < 2; 1.03: 0.2373 |  |  | 0.23617
| 1 | 5 | 1.01: < 1; 1.02: < 2; 1.03: 0.2342 |  |  | 7
| 1 | 6 | 1.01: < 1; 1.02: < 2; 1.03: 0.2376
| Mean |  | 1.03: 0.23617
| Median |  | 1.03: 0.23625
| Std Dev. |  | 1.03: 0.0016133
| Rel.Std.Dev. |  | 1.03: 0.6831%
| PDM3 |  | 1.03: -1.443e-13%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 1.28; 1.03: 0.67
| Analytical results for Ho in OREAS 255c (Certified Value 0.64 ppm) | [19 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.12: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.12: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.62; 1.02: 0.85406; 1.03: 0.73768; 1.07: 0.6; 1.09: 0.63; 1.10: 0.6001; 1.12: 0.57 |  |  | 1
| 1 | 2 | 1.01: 0.63; 1.02: 0.59921; 1.03: 0.65111; 1.07: 0.7; 1.09: 0.62; 1.10: 0.6147; 1.12: 0.6 |  |  | 3
| 1 | 3 | 1.01: 0.64; 1.02: 0.6317; 1.03: 0.66218; 1.07: 0.7; 1.09: 0.61; 1.10: 0.6352; 1.12: 0.55 |  |  | 16
| 1 | 4 | 1.01: 0.64; 1.02: 0.7404; 1.03: 0.6763; 1.07: 0.7; 1.09: 0.61; 1.10: 0.633; 1.12: 0.55 |  |  | 0.64202
| 1 | 5 | 1.01: 0.64; 1.02: 0.63291; 1.03: 0.70009; 1.07: 0.6; 1.09: 0.62; 1.10: 0.635; 1.12: 0.6 |  |  | 35
| 1 | 6 | 1.01: 0.66; 1.02: 0.63547; 1.03: 0.6691; 1.07: 0.7; 1.09: 0.62; 1.10: 0.6185; 1.12: 0.58
| Mean |  | 1.01: 0.63833; 1.02: 0.68229; 1.03: 0.68274; 1.07: 0.66667; 1.09: 0.61833; 1.10: 0.62275; 1.12: 0.575
| Median |  | 1.01: 0.64; 1.02: 0.63419; 1.03: 0.6727; 1.07: 0.7; 1.09: 0.62; 1.10: 0.62575; 1.12: 0.575
| Std Dev. |  | 1.01: 0.013292; 1.02: 0.096927; 1.03: 0.031528; 1.07: 0.05164; 1.09: 0.0075277; 1.10: 0.014185; 1.12: 0.022583
| Rel.Std.Dev. |  | 1.01: 2.082%; 1.02: 14.21%; 1.03: 4.618%; 1.07: 7.746%; 1.09: 1.217%; 1.10: 2.278%; 1.12: 3.928%
| PDM3 |  | 1.01: -0.5742%; 1.02: 6.273%; 1.03: 6.343%; 1.07: 3.839%; 1.09: -3.689%; 1.10: -3.001%; 1.12: -10.44%
| Z-Score (Absolute) |  | 1.01: 0.16; 1.02: 1.09; 1.03: 1.1; 1.07: NA; 1.09: 0.26; 1.10: 0.16; 1.12: 1.17
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 255c (Certified Value 0.065 ppm) | [19 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.07 | 0.06 | < 0.1 | < 2 | 0.065 | 0.05 | 0.078 | 0.07 | 0.065942 | < 0.1 | 0.06 | 0.064 | 0.06 | 0.069 |  |  |  |  | 1
| 1 | 2 | 0.06 | 0.06 | < 0.1 | < 2 | 0.066 | 0.06 | 0.0756 | 0.067 | 0.078525 | < 0.1 | 0.065 | 0.062 | 0.065 | 0.072 |  |  |  |  | 21
| 1 | 3 | 0.07 | 0.06 | < 0.1 | < 2 | 0.056 | 0.06 | 0.0769 | 0.073 | 0.069888 | < 0.1 | 0.062 | 0.066 | 0.057 | 0.067 |  |  |  |  | 16
| 1 | 4 | 0.07 | 0.05 | < 0.1 | < 2 | 0.057 | 0.06 | 0.0805 | 0.065 | < 0.05 | < 0.1 | 0.065 | 0.063 | 0.061 | 0.059 |  |  |  |  | 0.065344
| 1 | 5 | 0.06 | 0.06 | < 0.1 | < 2 | 0.057 | 0.07 | 0.079 | 0.069 | < 0.05 | < 0.1 | 0.066 | 0.063 | 0.066 | 0.072 |  |  |  |  | 36
| 1 | 6 | 0.07 | 0.06 | < 0.1 | < 2 | 0.061 | 0.06 | 0.0759 | 0.072 | 0.053781 | < 0.1 | 0.067 | 0.063 | 0.061 | 0.073
| Mean |  | 0.066667 | 0.058333 |  |  | 0.060333 | 0.06 | 0.07765 | 0.069333 | 0.067034 |  | 0.064167 | 0.0635 | 0.061667 | 0.068667
| Median |  | 0.07 | 0.06 |  |  | 0.059 | 0.06 | 0.07745 | 0.0695 | 0.067915 |  | 0.065 | 0.063 | 0.061 | 0.0705
| Std Dev. |  | 0.005164 | 0.0040825 |  |  | 0.0043665 | 0.0063246 | 0.0018918 | 0.0030111 | 0.01028 |  | 0.0026394 | 0.0013784 | 0.0033267 | 0.0052409
| Rel.Std.Dev. |  | 7.746% | 6.999% |  |  | 7.237% | 10.54% | 2.436% | 4.343% | 15.33% |  | 4.113% | 2.171% | 5.395% | 7.632%
| PDM3 |  | 2.024% | -10.73% |  |  | -7.668% | -8.178% | 18.83% | 6.105% | 2.586% |  | -1.802% | -2.822% | -5.628% | 5.085%
| Z-Score (Absolute) |  | 0.53 | 0.55 | 1.64 | 122.34 | 0.29 | 0.34 | 1.97 | 0.88 | 1.25 | 1.64 | 0.21 | 0.12 | 0.12 | 0.79
| Analytical results for K in OREAS 255c (Certified Value 1.58 wt.%) | [19 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.5925 | 1.6751 | 1.6096 | 1.4618 | 1.52 | 1.43 | 1.57 | 1.64 | 1.54 | 1.54 | 1.5688 | 1.59 | 1.5768 | 1.59 | 1.71 | 1.59 | 1.55 | 1.51 | 1
| 1 | 2 | 1.5572 | 1.687 | 1.592 | 1.5389 | 1.56 | 1.34 | 1.58 | 1.59 | 1.59 | 1.64 | 1.564 | 1.61 | 1.5934 | 1.58 | 1.7 | 1.57 | 1.54 | 1.58 | 0
| 1 | 3 | 1.5835 | 1.6835 | 1.6496 | 1.4804 | 1.57 | 1.4 | 1.59 | 1.63 | 1.57 | 1.64 | 1.5643 | 1.61 | 1.6336 | 1.56 | 1.88 | 1.59 | 1.55 | 1.56 | 16
| 1 | 4 | 1.5965 | 1.7028 | 1.6157 | 1.4928 | 1.57 | 1.4 | 1.58 | 1.63 | 1.52 | 1.57 | 1.5766 | 1.55 | 1.5562 | 1.56 | 1.72 | 1.59 | 1.54 | 1.58 | 1.5804
| 1 | 5 | 1.5878 | 1.6811 | 1.6198 | 1.463 | 1.58 | 1.41 | 1.57 | 1.6 | 1.58 | 1.57 | 1.5801 | 1.59 | 1.6217 | 1.58 | 1.72 | 1.6 | 1.55 | 1.6 | 37
| 1 | 6 | 1.5728 | 1.6986 | 1.639 | 1.5041 | 1.55 | 1.49 | 1.57 | 1.61 | 1.57 | 1.58 | 1.5703 | 1.59 | 1.564 | 1.57 | 1.72 | 1.62 | 1.52 | 1.53
| Mean |  | 1.5817 | 1.688 | 1.6209 | 1.4902 | 1.5583 | 1.4117 | 1.5767 | 1.6167 | 1.5617 | 1.59 | 1.5707 | 1.59 | 1.591 | 1.5733 | 1.7417 | 1.5933 | 1.5417 | 1.56
| Median |  | 1.5857 | 1.6852 | 1.6178 | 1.4866 | 1.565 | 1.405 | 1.575 | 1.62 | 1.57 | 1.575 | 1.5696 | 1.59 | 1.5851 | 1.575 | 1.72 | 1.59 | 1.545 | 1.57
| Std Dev. |  | 0.01453 | 0.010662 | 0.020686 | 0.029043 | 0.02137 | 0.048751 | 0.008165 | 0.019664 | 0.026394 | 0.040988 | 0.0065231 | 0.021909 | 0.031333 | 0.012111 | 0.06824 | 0.01633 | 0.01169 | 0.034059
| Rel.Std.Dev. |  | 0.9186% | 0.6316% | 1.276% | 1.949% | 1.371% | 3.453% | 0.5179% | 1.216% | 1.69% | 2.578% | 0.4153% | 1.378% | 1.969% | 0.7697% | 3.918% | 1.025% | 0.7583% | 2.183%
| PDM3 |  | 0.08152% | 6.808% | 2.564% | -5.711% | -1.398% | -10.68% | -0.238% | 2.293% | -1.187% | 0.6056% | -0.6148% | 0.6056% | 0.6662% | -0.4489% | 10.2% | 0.8166% | -2.453% | -1.293%
| Z-Score (Absolute) |  | 0.09 | 4 | 1.53 | 3.27 | 0.77 | 6.15 | 0.09 | 1.38 | 0.64 | 0.4 | 0.31 | 0.4 | 0.43 | 0.22 | 5.97 | 0.52 | 1.38 | 0.7
| Analytical results for La in OREAS 255c (Certified Value 39.6 ppm) | [19 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 38.63 | 39.0694 | 41.3368 | 35 | 35.8 | 40.4293 | 39.2 | 37.3 | 31.3399 | 40.9 | 40 | 38.6 | 41.4 | 40.2 | 39 |  |  |  | 1
| 1 | 2 | 39.95 | 39.8478 | 40.4702 | 36.5 | 35.5 | 38.8865 | 39.4 | 38.1 | 31.1498 | 40.4 | 39.9 | 40.7 | 39.4 | 40 | 38.8 |  |  |  | 22
| 1 | 3 | 39.04 | 38.3062 | 41.2457 | 36.9 | 35.2 | 39.9293 | 40.8 | 40.9 | 31.4154 | 41.4 | 39.7 | 40.2 | 40.6 | 41.3 | 36.98 |  |  |  | 16
| 1 | 4 | 37.93 | 39.5644 | 41.8777 | 36.8 | 35.2 | 38.1999 | 39.2 | 40.6 | 31.5097 | 40.6 | 39.8 | 40.1 | 40 | 41.6 | 37.36 |  |  |  | 39.606
| 1 | 5 | 39.4 | 38.8972 | 41.2165 | 36.5 | 35.7 | 39.1353 | 40.5 | 39.7 | 31.7746 | 40.7 | 40 | 39.1 | 41 | 40.9 | 37.08 |  |  |  | 38
| 1 | 6 | 40.06 | 38.5613 | 42.4388 | 37 | 35.7 | 40.7277 | 40.3 | 40.4 | 31.4712 | 40.7 | 40.2 | 39.3 | 40.8 | 39.9 | 38.66
| Mean |  | 39.1683 | 39.0411 | 41.4309 | 36.45 | 35.5167 | 39.5513 | 39.9 | 39.5 | 31.4434 | 40.7833 | 39.9333 | 39.6667 | 40.5333 | 40.65 | 37.98
| Median |  | 39.22 | 38.9833 | 41.2913 | 36.65 | 35.6 | 39.5323 | 39.85 | 40.05 | 31.4433 | 40.7 | 39.95 | 39.7 | 40.7 | 40.55 | 38.01
| Std Dev. |  | 0.81234 | 0.58589 | 0.66746 | 0.73959 | 0.26394 | 0.97354 | 0.71554 | 1.4711 | 0.2062 | 0.34303 | 0.17512 | 0.79162 | 0.72296 | 0.71764 | 0.93484
| Rel.Std.Dev. |  | 2.074% | 1.501% | 1.611% | 2.029% | 0.7432% | 2.461% | 1.793% | 3.724% | 0.6558% | 0.8411% | 0.4385% | 1.996% | 1.784% | 1.765% | 2.461%
| PDM3 |  | -1.105% | -1.426% | 4.608% | -7.969% | -10.33% | -0.1381% | 0.7422% | -0.2677% | -20.61% | 2.973% | 0.8264% | 0.1531% | 2.341% | 2.636% | -4.106%
| Z-Score (Absolute) |  | 0.26 | 0.35 | 1.29 | 2.13 | 2.77 | 0 | 0.24 | 0.04 | 5.57 | 0.85 | 0.26 | 0.08 | 0.67 | 0.75 | 1.08
| Analytical results for Li in OREAS 255c (Certified Value 26.5 ppm) | [19 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 26.1 | 27.0309 | 26.1607 | 42.3058 | 26.7 | 27 | 28.9 | 26.9586 | 25.7 | 26.5 | 24.414 | 27.1 | 28.4175 | 24.8 | 27.3 | 27.8 | 26.9 | 27 | 1
| 1 | 2 | 25.5 | 27.9278 | 25.6729 | 43.6525 | 26.3 | 27 | 27.8 | 25.8467 | 26.4 | 25 | 24.492 | 27.5 | 23.5877 | 24.7 | 26.6 | 27.6 | 26.8 | 29 | 0
| 1 | 3 | 25.9 | 27.0201 | 26.9657 | 43.5963 | 25.9 | 27 | 28.3 | 27.1529 | 26.6 | 26.7 | 24.258 | 27.5 | 28.3505 | 24.3 | 28.1 | 27.8 | 26.9 | 28 | 16
| 1 | 4 | 26.3 | 27.4485 | 26.0451 | 43.5546 | 26.6 | 27 | 28 | 26.0591 | 25.6 | 25.6 | 24.648 | 26.8 | 23.1471 | 24.9 | 28.1 | 28.3 | 26.7 | 25 | 26.5387
| 1 | 5 | 25.5 | 27.2374 | 26.4992 | 44.1498 | 26 | 27 | 28.1 | 26.1979 | 26.8 | 26.8 | 24.57 | 27.4 | 26.481 | 25.3 | 27.5 | 28.1 | 27.3 | 25 | 39
| 1 | 6 | 25.8 | 27.0274 | 26.2254 | 43.8466 | 26.7 | 27 | 27.9 | 26.4912 | 26.1 | 25.1 | 24.492 | 27.2 | 27.0995 | 25.1 | 27.9 | 28.1 | 26.3 | 26
| Mean |  | 25.85 | 27.282 | 26.2615 | 43.5176 | 26.3667 | 27 | 28.1667 | 26.4511 | 26.2 | 25.95 | 24.479 | 27.25 | 26.1805 | 24.85 | 27.5833 | 27.95 | 26.8167 | 26.6667
| Median |  | 25.85 | 27.1341 | 26.193 | 43.6244 | 26.45 | 27 | 28.05 | 26.3445 | 26.25 | 26.05 | 24.492 | 27.3 | 26.7902 | 24.85 | 27.7 | 27.95 | 26.85 | 26.5
| Std Dev. |  | 0.32094 | 0.35867 | 0.43754 | 0.63279 | 0.3559 | 0 | 0.39833 | 0.51667 | 0.4858 | 0.8167 | 0.13435 | 0.27386 | 2.3053 | 0.34496 | 0.58109 | 0.25884 | 0.32506 | 1.633
| Rel.Std.Dev. |  | 1.242% | 1.315% | 1.666% | 1.454% | 1.35% | 0% | 1.414% | 1.953% | 1.854% | 3.147% | 0.5488% | 1.005% | 8.805% | 1.388% | 2.107% | 0.9261% | 1.212% | 6.124%
| PDM3 |  | -2.595% | 2.801% | -1.044% | 63.98% | -0.6481% | 1.738% | 6.134% | -0.3301% | -1.276% | -2.218% | -7.761% | 2.68% | -1.35% | -6.363% | 3.936% | 5.318% | 1.048% | 0.4823%
| Z-Score (Absolute) |  | 0.74 | 0.75 | 0.31 | 17.59 | 0.2 | 0.46 | 1.67 | 0.11 | 0.37 | 0.63 | 2.16 | 0.72 | 0.39 | 1.77 | 1.06 | 1.44 | 0.27 | 0.11
| Analytical results for Lu in OREAS 255c (Certified Value 0.2 ppm) | [19 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.06: 1.09; 1.07: 1.10; 1.09: 1.14; 1.10: 1.22 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS | 2
3
| 1 | 1 | 1.01: 0.2; 1.02: 0.2007; 1.03: 0.22127; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0685; 1.10: 0.17 | 1
| 1 | 2 | 1.01: 0.23; 1.02: 0.19082; 1.03: 0.20717; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0865; 1.10: 0.18 | 4
| 1 | 3 | 1.01: 0.26; 1.02: 0.1932; 1.03: 0.2104; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0672; 1.10: 0.17 | 16
| 1 | 4 | 1.01: 0.21; 1.02: 0.21023; 1.03: 0.22063; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0616; 1.10: 0.17 | 0.20483
| 1 | 5 | 1.01: 0.22; 1.02: 0.20392; 1.03: 0.22217; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0978; 1.10: 0.18 | 40
| 1 | 6 | 1.01: 0.22; 1.02: 0.21601; 1.03: 0.20853; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0885; 1.10: 0.17
| Mean |  | 1.01: 0.22333; 1.02: 0.20248; 1.03: 0.21503; 1.06: 0.2; 1.07: 0.21; 1.09: 0.07835; 1.10: 0.17333
| Median |  | 1.01: 0.22; 1.02: 0.20231; 1.03: 0.21551; 1.06: 0.2; 1.07: 0.21; 1.09: 0.0775; 1.10: 0.17
| Std Dev. |  | 1.01: 0.020656; 1.02: 0.0097009; 1.03: 0.0070238; 1.06: 3.0405e-17; 1.07: 0; 1.09: 0.014489; 1.10: 0.005164
| Rel.Std.Dev. |  | 1.01: 9.249%; 1.02: 4.791%; 1.03: 3.266%; 1.06: 1.52e-14%; 1.07: 0%; 1.09: 18.49%; 1.10: 2.979%
| PDM3 |  | 1.01: 9.031%; 1.02: -1.15%; 1.03: 4.976%; 1.06: -2.36%; 1.07: 2.522%; 1.09: -61.75%; 1.10: -15.38%
| Z-Score (Absolute) |  | 1.01: 0.89; 1.02: 0.2; 1.03: 0.46; 1.06: NA; 1.07: 0.2; 1.09: 6.66; 1.10: 1.71
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 255c (Certified Value 1.62 wt.%) | [19 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.624 | 1.7745 | 1.621 | 1.731 | 1.64 | 1.57 | 1.56 | 1.58 | 1.58 | 1.57 | 1.6026 | 1.64 | 1.6421 | 1.66 | 1.6 | 1.62 | 1.55 | 1.461 | 1
| 1 | 2 | 1.5983 | 1.7482 | 1.5977 | 1.7736 | 1.62 | 1.59 | 1.57 | 1.6 | 1.63 | 1.62 | 1.5771 | 1.64 | 1.6501 | 1.66 | 1.56 | 1.61 | 1.53 | 1.546 | 0
| 1 | 3 | 1.6125 | 1.7096 | 1.6617 | 1.7505 | 1.6 | 1.65 | 1.58 | 1.62 | 1.63 | 1.67 | 1.5812 | 1.64 | 1.6715 | 1.64 | 1.66 | 1.62 | 1.55 | 1.576 | 16
| 1 | 4 | 1.6433 | 1.7825 | 1.6385 | 1.7472 | 1.64 | 1.64 | 1.57 | 1.57 | 1.57 | 1.64 | 1.6039 | 1.58 | 1.5836 | 1.61 | 1.66 | 1.61 | 1.54 | 1.593 | 1.6157
| 1 | 5 | 1.6358 | 1.7225 | 1.6321 | 1.7636 | 1.63 | 1.62 | 1.58 | 1.62 | 1.66 | 1.6 | 1.6019 | 1.64 | 1.67 | 1.64 | 1.6 | 1.64 | 1.56 | 1.586 | 41
| 1 | 6 | 1.6021 | 1.8059 | 1.6482 | 1.7748 | 1.64 | 1.65 | 1.55 | 1.62 | 1.64 | 1.66 | 1.5909 | 1.64 | 1.5991 | 1.63 | 1.62 | 1.64 | 1.52 | 1.553
| Mean |  | 1.6193 | 1.7572 | 1.6332 | 1.7568 | 1.6283 | 1.62 | 1.5683 | 1.6017 | 1.6183 | 1.6267 | 1.5929 | 1.63 | 1.6361 | 1.64 | 1.6167 | 1.6233 | 1.5417 | 1.5525
| Median |  | 1.6182 | 1.7614 | 1.6353 | 1.757 | 1.635 | 1.63 | 1.57 | 1.61 | 1.63 | 1.63 | 1.5964 | 1.64 | 1.6461 | 1.64 | 1.61 | 1.62 | 1.545 | 1.5645
| Std Dev. |  | 0.018185 | 0.037039 | 0.022252 | 0.017015 | 0.016021 | 0.033466 | 0.01169 | 0.022286 | 0.035449 | 0.037771 | 0.011738 | 0.024495 | 0.036767 | 0.018974 | 0.038816 | 0.013663 | 0.01472 | 0.048434
| Rel.Std.Dev. |  | 1.123% | 2.108% | 1.363% | 0.9685% | 0.9839% | 2.066% | 0.7454% | 1.391% | 2.19% | 2.322% | 0.7369% | 1.503% | 2.247% | 1.157% | 2.401% | 0.8416% | 0.9548% | 3.12%
| PDM3 |  | 0.2242% | 8.758% | 1.082% | 8.731% | 0.7812% | 0.2654% | -2.932% | -0.8693% | 0.1623% | 0.678% | -1.41% | 0.8843% | 1.261% | 1.503% | 0.05912% | 0.4717% | -4.583% | -3.912%
| Z-Score (Absolute) |  | 0.12 | 7.05 | 0.6 | 7.02 | 0.35 | 0.09 | 2.77 | 1.04 | 0.17 | 0.26 | 1.49 | 0.43 | 0.75 | 0.95 | 0.26 | 0.09 | 4.16 | 3.6
| Analytical results for Mn in OREAS 255c (Certified Value 0.043 wt.%) | [19 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0436 | 0.040819 | 0.042905 | < 0.1 | 0.0421 | 0.0391 | 0.0431 | 0.0444 | 0.042 | 0.0404 | 0.037866 | 0.0433 | 0.045015 | 0.0432 | 0.0424 | 0.0436 | 0.0418 | 0.0381 | 1
| 1 | 2 | 0.0424 | 0.041373 | 0.04178 | < 0.1 | 0.0418 | 0.0424 | 0.0429 | 0.0435 | 0.043 | 0.0413 | 0.037112 | 0.0432 | 0.044727 | 0.0429 | 0.0412 | 0.0436 | 0.0417 | 0.0397 | 24
| 1 | 3 | 0.0429 | 0.041334 | 0.043362 | < 0.1 | 0.0422 | 0.0422 | 0.0429 | 0.0444 | 0.0425 | 0.0425 | 0.037437 | 0.0432 | 0.043498 | 0.0425 | 0.0441 | 0.0438 | 0.0417 | 0.04 | 16
| 1 | 4 | 0.0438 | 0.040988 | 0.043377 | < 0.1 | 0.0416 | 0.042 | 0.043 | 0.0437 | 0.0412 | 0.0419 | 0.037868 | 0.0419 | 0.042108 | 0.0426 | 0.0443 | 0.0442 | 0.0414 | 0.0397 | 0.04255
| 1 | 5 | 0.0438 | 0.041423 | 0.043247 | < 0.1 | 0.042 | 0.0407 | 0.0431 | 0.0438 | 0.0435 | 0.0413 | 0.037759 | 0.0432 | 0.043729 | 0.0429 | 0.0426 | 0.0444 | 0.0419 | 0.0403 | 42
| 1 | 6 | 0.0425 | 0.041148 | 0.043675 | < 0.1 | 0.0418 | 0.0418 | 0.0428 | 0.0445 | 0.0428 | 0.0425 | 0.037545 | 0.0438 | 0.042208 | 0.0432 | 0.0432 | 0.0446 | 0.0408 | 0.0399
| Mean |  | 0.043167 | 0.041181 | 0.043058 |  | 0.041917 | 0.041367 | 0.042967 | 0.04405 | 0.0425 | 0.04165 | 0.037598 | 0.0431 | 0.043547 | 0.042883 | 0.042967 | 0.044033 | 0.04155 | 0.039617
| Median |  | 0.04325 | 0.041241 | 0.043304 |  | 0.0419 | 0.0419 | 0.04295 | 0.0441 | 0.04265 | 0.0416 | 0.037652 | 0.0432 | 0.043613 | 0.0429 | 0.0429 | 0.044 | 0.0417 | 0.0398
| Std Dev. |  | 0.00064704 | 0.00024016 | 0.00067332 |  | 0.00022286 | 0.0012596 | 0.00012111 | 0.00043243 | 0.00080994 | 0.00081425 | 0.00029504 | 0.00063246 | 0.0012206 | 0.00029269 | 0.001157 | 0.0004274 | 0.00040373 | 0.00077567
| Rel.Std.Dev. |  | 1.499% | 0.5832% | 1.564% |  | 0.5317% | 3.045% | 0.2819% | 0.9817% | 1.906% | 1.955% | 0.7847% | 1.467% | 2.803% | 0.6825% | 2.693% | 0.9706% | 0.9717% | 1.958%
| PDM3 |  | 1.449% | -3.219% | 1.192% |  | -1.489% | -2.782% | 0.9786% | 3.525% | -0.1182% | -2.116% | -11.64% | 1.292% | 2.343% | 0.7827% | 0.9786% | 3.485% | -2.351% | -6.894%
| Z-Score (Absolute) |  | 0.15 | 1.11 | 0.08 | 4.51 | 0.64 | 0.99 | 0.03 | 0.72 | 0.27 | 0.81 | 3.39 | 0.11 | 0.4 | 0.03 | 0.03 | 0.71 | 0.88 | 2.11
| Analytical results for Mo in OREAS 255c (Certified Value 4.85 ppm) | [19 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 4.9 | 4.7315 | 6.0674 | 5 | 2 | 5.39 | 4.8818 | 4.67 | 4.7 | 5.1509 | 4.46 | 4.8513 | 4.59 | 4.37 | 5.14 | 5.14 | 4.89 |  | 1
| 1 | 2 | 4.7 | 4.6625 | 6.7862 | 5.1 | 2 | 5.32 | 4.6778 | 4.9 | 4.5 | 5.1542 | 4.4 | 4.6038 | 4.59 | 4.59 | 5.01 | 5.02 | 4.81 |  | 25
| 1 | 3 | 5 | 4.983 | 6.2713 | 5 | 2 | 5.34 | 5.1705 | 5.06 | 4.9 | 5.1542 | 4.72 | 4.454 | 4.66 | 4.67 | 5.07 | 5.07 | 4.89 |  | 16
| 1 | 4 | 4.8 | 5.1177 | 6.0395 | 5 | < 1 | 5.41 | 4.6882 | 4.8 | 4.7 | 5.1309 | 4.39 | 3.761 | 4.41 | 4.61 | 5.01 | 5.2 | 4.83 |  | 4.8536
| 1 | 5 | 4.9 | 4.9073 | 6.2624 | 4.9 | 2 | 5.37 | 4.8672 | 4.97 | 4.8 | 5.2814 | 4.61 | 3.9967 | 4.45 | 4.44 | 5.03 | 5.09 | 5.02 |  | 43
| 1 | 6 | 4.7 | 4.9666 | 5.6967 | 5.1 | 4 | 5.36 | 4.7622 | 5.09 | 4.5 | 5.2932 | 4.6 | 4.3316 | 4.43 | 4.54 | 5.02 | 4.98 | 4.93
| Mean |  | 4.8333 | 4.8948 | 6.1872 | 5.0167 | 2.4 | 5.365 | 4.8413 | 4.915 | 4.6833 | 5.1941 | 4.53 | 4.3331 | 4.5217 | 4.5367 | 5.0467 | 5.0833 | 4.895
| Median |  | 4.85 | 4.937 | 6.1649 | 5 | 2 | 5.365 | 4.8147 | 4.935 | 4.7 | 5.1542 | 4.53 | 4.3928 | 4.52 | 4.565 | 5.025 | 5.08 | 4.89
| Std Dev. |  | 0.12111 | 0.16935 | 0.36008 | 0.075277 | 0.89443 | 0.032711 | 0.18282 | 0.16009 | 0.16021 | 0.072783 | 0.13327 | 0.3992 | 0.10439 | 0.11237 | 0.050859 | 0.079666 | 0.075299
| Rel.Std.Dev. |  | 2.506% | 3.46% | 5.82% | 1.501% | 37.27% | 0.6097% | 3.776% | 3.257% | 3.421% | 1.401% | 2.942% | 9.213% | 2.309% | 2.477% | 1.008% | 1.567% | 1.538%
| PDM3 |  | -0.418% | 0.8479% | 27.48% | 3.359% | -50.55% | 10.54% | -0.2544% | 1.265% | -3.509% | 7.016% | -6.668% | -10.72% | -6.839% | -6.53% | 3.977% | 4.733% | 0.8525%
| Z-Score (Absolute) |  | 0.21 | 0 | 4.36 | 0.41 | NA | 1.59 | 0.18 | 0.07 | 0.71 | 1.01 | 1.23 | 1.89 | 1.26 | 1.21 | 0.51 | 0.64 | 0
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 255c (Certified Value 0.811 wt.%) | [19 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.8115 | 0.73306 | 0.80914 | 2.3668 | 0.8 | 0.8 | 0.8 | 0.86 | 0.78 | 0.788 | 0.7903 | 0.82 | 0.8221 | 0.83 | 0.8 | 0.82 | 0.79 | 0.781 | 1
| 1 | 2 | 0.7951 | 0.7219 | 0.80939 | 2.3887 | 0.81 | 0.82 | 0.8 | 0.83 | 0.82 | 0.826 | 0.7637 | 0.81 | 0.8356 | 0.84 | 0.79 | 0.82 | 0.79 | 0.787 | 0
| 1 | 3 | 0.8091 | 0.723 | 0.8326 | 2.3467 | 0.808 | 0.83 | 0.79 | 0.85 | 0.81 | 0.842 | 0.7691 | 0.81 | 0.83657 | 0.83 | 0.84 | 0.82 | 0.8 | 0.802 | 16
| 1 | 4 | 0.8174 | 0.73305 | 0.82015 | 2.34 | 0.808 | 0.83 | 0.8 | 0.85 | 0.79 | 0.829 | 0.7846 | 0.79 | 0.81791 | 0.8 | 0.84 | 0.83 | 0.79 | 0.764 | 0.81126
| 1 | 5 | 0.8256 | 0.73152 | 0.82623 | 2.4284 | 0.808 | 0.8 | 0.8 | 0.85 | 0.83 | 0.81 | 0.7836 | 0.81 | 0.82708 | 0.82 | 0.8 | 0.84 | 0.81 | 0.793 | 44
| 1 | 6 | 0.797 | 0.73635 | 0.81888 | 2.4318 | 0.81 | 0.81 | 0.79 | 0.84 | 0.82 | 0.837 | 0.7775 | 0.82 | 0.82757 | 0.82 | 0.81 | 0.83 | 0.77 | 0.78
| Mean |  | 0.80928 | 0.72981 | 0.8194 | 2.3837 | 0.80733 | 0.815 | 0.79667 | 0.84667 | 0.80833 | 0.822 | 0.77813 | 0.81 | 0.82781 | 0.82333 | 0.81333 | 0.82667 | 0.79167 | 0.7845
| Median |  | 0.8103 | 0.73228 | 0.81951 | 2.3778 | 0.808 | 0.815 | 0.8 | 0.85 | 0.815 | 0.8275 | 0.78055 | 0.81 | 0.82733 | 0.825 | 0.805 | 0.825 | 0.79 | 0.784
| Std Dev. |  | 0.01174 | 0.0059286 | 0.0092428 | 0.039725 | 0.0037238 | 0.013784 | 0.005164 | 0.010328 | 0.019408 | 0.01995 | 0.0101 | 0.010954 | 0.007329 | 0.013663 | 0.021602 | 0.008165 | 0.013292 | 0.012942
| Rel.Std.Dev. |  | 1.451% | 0.8123% | 1.128% | 1.667% | 0.4612% | 1.691% | 0.6482% | 1.22% | 2.401% | 2.427% | 1.298% | 1.352% | 0.8854% | 1.659% | 2.656% | 0.9877% | 1.679% | 1.65%
| PDM3 |  | -0.2433% | -10.04% | 1.003% | 193.8% | -0.4837% | 0.4613% | -1.799% | 4.365% | -0.3604% | 1.324% | -4.083% | -0.155% | 2.04% | 1.489% | 0.2559% | 1.899% | -2.415% | -3.298%
| Z-Score (Absolute) |  | 0.12 | 4.14 | 0.39 | 79.5 | 0.22 | 0.17 | 0.76 | 1.77 | 0.17 | 0.52 | 1.7 | 0.08 | 0.82 | 0.59 | 0.08 | 0.76 | 1.01 | 1.37
| Analytical results for Nb in OREAS 255c (Certified Value 26.3 ppm) | [19 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 26.53 | 33.835 | 78.9353 | 25.2 | 0.6 | 24 | 26.3032 | 25.6 | 27.8 | 36.4844 | 25.7 | 30.9937 | 26.2 | 27.3 | 28.3 | 27 | 25.7 |  | 1
| 1 | 2 | 26.52 | 31.8782 | 54.5263 | 26 | 0.7 | 25 | 23.8989 | 25.9 | 26.5 | 36.3335 | 25.7 | 26.844 | 26.3 | 27.8 | 27.9 | 27.3 | 26.6 |  | 5
| 1 | 3 | 26.91 | 32.5555 | 53.3703 | 25.9 | 0.3 | 24 | 25.3111 | 26.1 | 27.9 | 37.7437 | 26.6 | 28.2296 | 26.3 | 28.7 | 28.2 | 27.1 | 25 |  | 16
| 1 | 4 | 27.02 | 32.3036 | 51.9179 | 26.9 | 0.8 | 24 | 24.077 | 25.4 | 27.3 | 37.0346 | 25.1 | 23.4443 | 26.6 | 28.4 | 28 | 26.5 | 25.2 |  | 26.3482
| 1 | 5 | 26.92 | 35.0655 | 54.0732 | 25.9 | 0.5 | 24 | 25.1022 | 26.5 | 28.5 | 38.4707 | 25.9 | 23.3027 | 26.4 | 27.3 | 27.8 | 27.4 | 26.4 |  | 45
| 1 | 6 | 26.78 | 31.355 | 59.7314 | 25.9 | 0.8 | 24 | 25.8349 | 26.7 | 27.2 | 37.6837 | 26.6 | 28.4877 | 26.4 | 27.3 | 27.8 | 27.4 | 26.3
| Mean |  | 26.78 | 32.8321 | 58.7591 | 25.9667 | 0.61667 | 24.1667 | 25.0879 | 26.0333 | 27.5333 | 37.2918 | 25.9333 | 26.8837 | 26.3667 | 27.8 | 28 | 27.1167 | 25.8667
| Median |  | 26.845 | 32.4295 | 54.2998 | 25.9 | 0.65 | 24 | 25.2066 | 26 | 27.55 | 37.3591 | 25.8 | 27.5368 | 26.35 | 27.55 | 27.95 | 27.2 | 26
| Std Dev. |  | 0.21175 | 1.3741 | 10.2346 | 0.54283 | 0.19408 | 0.40825 | 0.95084 | 0.50465 | 0.68896 | 0.82268 | 0.58195 | 3.0309 | 0.13663 | 0.61968 | 0.20976 | 0.34303 | 0.66833
| Rel.Std.Dev. |  | 0.7907% | 4.185% | 17.42% | 2.09% | 31.47% | 1.689% | 3.79% | 1.938% | 2.502% | 2.206% | 2.244% | 11.27% | 0.5182% | 2.229% | 0.7491% | 1.265% | 2.584%
| PDM3 |  | 1.639% | 24.61% | 123% | -1.448% | -97.66% | -8.28% | -4.783% | -1.195% | 4.498% | 41.53% | -1.574% | 2.032% | 0.0702% | 5.51% | 6.269% | 2.917% | -1.827%
| Z-Score (Absolute) |  | 0.04 | 4.34 | 23.1 | 0.63 | NA | 1.93 | 1.26 | 0.58 | 0.51 | 7.57 | 0.65 | 0.04 | 0.34 | 0.7 | 0.85 | 0.21 | 0.7
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 255c (Certified Value 32.5 ppm) | [19 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.09; 1.07: 1.10; 1.09: 1.14; 1.10: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS |  | 1
2
| 1 | 1 | 1.01: 32.02; 1.03: 32.6709; 1.06: 30.7; 1.07: 32; 1.09: 25.1927; 1.10: 32 |  | 1
| 1 | 2 | 1.01: 33.37; 1.03: 33.1923; 1.06: 33.1; 1.07: 31.7; 1.09: 24.9653; 1.10: 32.2 |  | 6
| 1 | 3 | 1.01: 33.19; 1.03: 33.2218; 1.06: 33; 1.07: 31.5; 1.09: 25.337; 1.10: 31.1 |  | 16
| 1 | 4 | 1.01: 32.11; 1.03: 33.3208; 1.06: 33.4; 1.07: 31.7; 1.09: 25.3829; 1.10: 31.2 |  | 32.4699
| 1 | 5 | 1.01: 33.03; 1.03: 32.6532; 1.06: 32; 1.07: 31.9; 1.09: 25.4067; 1.10: 32.5 |  | 46
| 1 | 6 | 1.01: 33.15; 1.03: 33.687; 1.06: 32.9; 1.07: 31.9; 1.09: 25.248; 1.10: 31.5
| Mean |  | 1.01: 32.8117; 1.03: 33.1243; 1.06: 32.5167; 1.07: 31.7833; 1.09: 25.2554; 1.10: 31.75
| Median |  | 1.01: 33.09; 1.03: 33.2071; 1.06: 32.95; 1.07: 31.8; 1.09: 25.2925; 1.10: 31.75
| Std Dev. |  | 1.01: 0.58925; 1.03: 0.39922; 1.06: 1.0068; 1.07: 0.18348; 1.09: 0.16368; 1.10: 0.56833
| Rel.Std.Dev. |  | 1.01: 1.796%; 1.03: 1.205%; 1.06: 3.096%; 1.07: 0.5773%; 1.09: 0.6481%; 1.10: 1.79%
| PDM3 |  | 1.01: 1.053%; 1.03: 2.016%; 1.06: 0.1441%; 1.07: -2.114%; 1.09: -22.22%; 1.10: -2.217%
| Z-Score (Absolute) |  | 1.01: 0.84; 1.03: 1.24; 1.06: 0.47; 1.07: 0.47; 1.09: 8.76; 1.10: 0.51
| Analytical results for Ni in OREAS 255c (Certified Value 95 ppm) | [19 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
1
| 1 | 1 | 96.1 | 96.1093 | 79.6783 | 91 | 98 | 95.1 | 91.8933 | 90.3 | 88.5 | 77.7481 | 91.3 | 93.7842 | 96.1 | 94.1 | 99.1 | 96.7 | 103 |  | 1
| 1 | 2 | 96.4 | 92.976 | 86.272 | 92 | 99 | 95.6 | 91.815 | 93.1 | 90.4 | 76.152 | 90.4 | 94.8659 | 97.1 | 94.8 | 97.5 | 96 | 103 |  | 28
| 1 | 3 | 97.2 | 97.4787 | 80.562 | 91 | 100 | 94.6 | 91.8 | 95.6 | 91.6 | 76.6611 | 93.3 | 96.4446 | 95.8 | 96.2 | 98 | 97.8 | 100 |  | 16
| 1 | 4 | 95.9 | 97.7696 | 79.4435 | 92 | 101 | 95.7 | 88.9267 | 91.5 | 90.1 | 77.7037 | 88.5 | 90.5169 | 95.4 | 96.2 | 97.5 | 97.5 | 101 |  | 95.2758
| 1 | 5 | 98.5 | 96.3145 | 80.3027 | 91 | 96 | 96 | 90.5567 | 97.4 | 89.9 | 77.6115 | 93.1 | 95.0482 | 96.8 | 93.6 | 97.8 | 97.9 | 104 |  | 47
| 1 | 6 | 96.6 | 97.384 | 82.2006 | 92 | 101 | 95.4 | 92.9933 | 95.4 | 90.9 | 77.2234 | 92.5 | 91.7249 | 96.6 | 95.9 | 98.4 | 94.8 | 100
| Mean |  | 96.7833 | 96.3387 | 81.4099 | 91.5 | 99.1667 | 95.4 | 91.3308 | 93.8833 | 90.2333 | 77.1833 | 91.5167 | 93.7308 | 96.3 | 95.1333 | 98.05 | 96.7833 | 101.8333
| Median |  | 96.5 | 96.8492 | 80.4323 | 91.5 | 99.5 | 95.5 | 91.8075 | 94.25 | 90.25 | 77.4175 | 91.9 | 94.3251 | 96.35 | 95.35 | 97.9 | 97.1 | 102
| Std Dev. |  | 0.95376 | 1.7777 | 2.572 | 0.54772 | 1.9408 | 0.49396 | 1.4081 | 2.711 | 1.0463 | 0.64979 | 1.8465 | 2.2246 | 0.64498 | 1.1308 | 0.61563 | 1.2123 | 1.7224
| Rel.Std.Dev. |  | 0.9855% | 1.845% | 3.159% | 0.5986% | 1.957% | 0.5178% | 1.542% | 2.888% | 1.16% | 0.8419% | 2.018% | 2.373% | 0.6698% | 1.189% | 0.6279% | 1.253% | 1.691%
| PDM3 |  | 1.582% | 1.116% | -14.55% | -3.963% | 4.084% | 0.1304% | -4.141% | -1.462% | -5.292% | -18.99% | -3.946% | -1.622% | 1.075% | -0.1495% | 2.912% | 1.582% | 6.883%
| Z-Score (Absolute) |  | 0.38 | 0.28 | 3.17 | 0.84 | 0.93 | 0.06 | 0.88 | 0.29 | 1.13 | 4.15 | 0.84 | 0.32 | 0.27 | 0 | 0.67 | 0.38 | 1.55
| Analytical results for P in OREAS 255c (Certified Value 0.077 wt.%) | [19 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0768 | 0.07668 | < 0.1 | 0.079 | 0.067 | 0.0797 | 0.0777 | 0.076 | 0.074 | 0.072371 | 0.079 | 0.080412 | 0.08 | 0.077 | 0.078 | 0.074 | 0.068 |  | 1
| 1 | 2 | 0.0755 | 0.075642 | < 0.1 | 0.081 | 0.073 | 0.0791 | 0.0766 | 0.079 | 0.077 | 0.070272 | 0.078 | 0.081204 | 0.08 | 0.076 | 0.078 | 0.075 | 0.072 |  | 0
| 1 | 3 | 0.0764 | 0.079749 | < 0.1 | 0.08 | 0.068 | 0.0791 | 0.0758 | 0.079 | 0.079 | 0.071072 | 0.078 | 0.081191 | 0.08 | 0.08 | 0.079 | 0.074 | 0.076 |  | 16
| 1 | 4 | 0.0758 | 0.0773 | < 0.1 | 0.079 | 0.065 | 0.0791 | 0.0758 | 0.075 | 0.078 | 0.07227 | 0.077 | 0.076724 | 0.08 | 0.081 | 0.08 | 0.074 | 0.075 |  | 0.077298
| 1 | 5 | 0.077 | 0.077194 | < 0.1 | 0.08 | 0.065 | 0.0794 | 0.0748 | 0.08 | 0.076 | 0.07217 | 0.078 | 0.080911 | 0.08 | 0.077 | 0.08 | 0.075 | 0.076 |  | 48
| 1 | 6 | 0.0744 | 0.079098 | < 0.1 | 0.08 | 0.062 | 0.0789 | 0.0761 | 0.079 | 0.079 | 0.071371 | 0.079 | 0.077404 | 0.08 | 0.078 | 0.079 | 0.073 | 0.075
| Mean |  | 0.075983 | 0.077611 |  | 0.079833 | 0.066667 | 0.079217 | 0.076133 | 0.078 | 0.077167 | 0.071588 | 0.078167 | 0.079641 | 0.08 | 0.078167 | 0.079 | 0.074167 | 0.073667
| Median |  | 0.0761 | 0.077247 |  | 0.08 | 0.066 | 0.0791 | 0.07595 | 0.079 | 0.0775 | 0.071771 | 0.078 | 0.080661 | 0.08 | 0.0775 | 0.079 | 0.074 | 0.075
| Std Dev. |  | 0.00096419 | 0.0015362 |  | 0.00075277 | 0.0037238 | 0.00028577 | 0.00096678 | 0.002 | 0.0019408 | 0.00083223 | 0.00075277 | 0.0020281 | 0 | 0.0019408 | 0.00089443 | 0.00075277 | 0.0031411
| Rel.Std.Dev. |  | 1.269% | 1.979% |  | 0.9429% | 5.586% | 0.3607% | 1.27% | 2.564% | 2.515% | 1.163% | 0.963% | 2.547% | 0% | 2.483% | 1.132% | 1.015% | 4.264%
| PDM3 |  | -1.701% | 0.4042% |  | 3.28% | -13.75% | 2.482% | -1.507% | 0.9079% | -0.1701% | -7.388% | 1.124% | 3.031% | 3.495% | 1.124% | 2.202% | -4.051% | -4.698%
| Z-Score (Absolute) |  | 0.67 | 0 | 11.44 | 0.92 | 4.53 | 0.67 | 0.61 | 0.16 | 0.18 | 2.5 | 0.23 | 0.84 | 0.99 | 0.23 | 0.58 | 1.43 | 1.63
| Analytical results for Pb in OREAS 255c (Certified Value 15.2 ppm) | [19 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 15.4 | 16.4052 | 9.9706 | 16 | 15 | 12.8 | 14.0151 | 15 | 15 | 11.8482 | 15.1 | 16.1039 | 15.1 | 14 | 15.5 | 15.6 | 14.9 |  | 1
| 1 | 2 | 15.8 | 16.0105 | 9.7322 | 16 | 14 | 11.9 | 13.3031 | 15 | 14.8 | 11.874 | 15 | 15.1749 | 15.3 | 14.2 | 15.4 | 15.1 | 13.9 |  | 29
| 1 | 3 | 15.7 | 15.6388 | 9.852 | 16 | 15 | 11.8 | 13.8928 | 15.4 | 15.1 | 11.6877 | 15.5 | 15.3486 | 15.5 | 14.5 | 15.3 | 15.4 | 14.8 |  | 16
| 1 | 4 | 15.6 | 16.3442 | 12.0542 | 17 | 15 | 12.9 | 14.1184 | 15 | 15 | 11.4549 | 15 | 12.0787 | 15.3 | 14.1 | 14.9 | 15.7 | 14.1 |  | 15.1754
| 1 | 5 | 15.6 | 15.9127 | 9.2364 | 17 | 15 | 13.1 | 13.4781 | 15.4 | 15 | 11.5527 | 15.2 | 11.7067 | 15.4 | 14.1 | 15.4 | 15.2 | 16 |  | 49
| 1 | 6 | 15.7 | 15.8973 | 7.2084 | 16 | 17 | 12.5 | 13.7502 | 15 | 15.5 | 11.9651 | 15.5 | 14.8292 | 15.5 | 14.3 | 15.3 | 15.5 | 13.7
| Mean |  | 15.6333 | 16.0348 | 9.6756 | 16.3333 | 15.1667 | 12.5 | 13.7596 | 15.1333 | 15.0667 | 11.7304 | 15.2167 | 14.207 | 15.35 | 14.2 | 15.3 | 15.4167 | 14.5667
| Median |  | 15.65 | 15.9616 | 9.7921 | 16 | 15 | 12.65 | 13.8215 | 15 | 15 | 11.7679 | 15.15 | 15.002 | 15.35 | 14.15 | 15.35 | 15.45 | 14.45
| Std Dev. |  | 0.13663 | 0.29125 | 1.553 | 0.5164 | 0.98319 | 0.54037 | 0.31604 | 0.20656 | 0.23381 | 0.19945 | 0.23166 | 1.8443 | 0.15166 | 0.17889 | 0.20976 | 0.23166 | 0.85245
| Rel.Std.Dev. |  | 0.8739% | 1.816% | 16.05% | 3.162% | 6.483% | 4.323% | 2.297% | 1.365% | 1.552% | 1.7% | 1.522% | 12.98% | 0.988% | 1.26% | 1.371% | 1.503% | 5.852%
| PDM3 |  | 3.018% | 5.663% | -36.24% | 7.631% | -0.05737% | -17.63% | -9.329% | -0.277% | -0.7163% | -22.7% | 0.2721% | -6.381% | 1.151% | -6.427% | 0.8212% | 1.59% | -4.011%
| Z-Score (Absolute) |  | 0.59 | 1.07 | 6.49 | 1.43 | 0.04 | 3.13 | 1.63 | 0 | 0.08 | 4.05 | 0.1 | 1.1 | 0.26 | 1.11 | 0.2 | 0.34 | 0.67
| Analytical results for Pr in OREAS 255c (Certified Value 8.67 ppm) | [19 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 8.55; 1.03: 8.9904; 1.06: 9.4373; 1.07: 8.2; 1.09: 8.31; 1.10: 4.4607; 1.11: 8.69 |  | 1
| 1 | 2 | 1.01: 8.98; 1.03: 9.022; 1.06: 8.677; 1.07: 8.5; 1.09: 8.16; 1.10: 4.4334; 1.11: 8.88 |  | 8
| 1 | 3 | 1.01: 8.71; 1.03: 8.6297; 1.06: 8.7174; 1.07: 8.7; 1.09: 8.21; 1.10: 4.5248; 1.11: 8.46 |  | 16
| 1 | 4 | 1.01: 8.61; 1.03: 9.0271; 1.06: 9.2321; 1.07: 8.8; 1.09: 8.13; 1.10: 4.5017; 1.11: 8.44 |  | 8.6683
| 1 | 5 | 1.01: 9.1; 1.03: 8.6627; 1.06: 9.3491; 1.07: 8.4; 1.09: 8.24; 1.10: 4.695; 1.11: 8.34 |  | 50
| 1 | 6 | 1.01: 8.84; 1.03: 8.8454; 1.06: 8.7604; 1.07: 8.6; 1.09: 8.26; 1.10: 4.4711; 1.11: 8.6
| Mean |  | 1.01: 8.7983; 1.03: 8.8629; 1.06: 9.0289; 1.07: 8.5333; 1.09: 8.2183; 1.10: 4.5145; 1.11: 8.5683
| Median |  | 1.01: 8.775; 1.03: 8.9179; 1.06: 8.9962; 1.07: 8.55; 1.09: 8.225; 1.10: 4.4864; 1.11: 8.53
| Std Dev. |  | 1.01: 0.21479; 1.03: 0.18074; 1.06: 0.34744; 1.07: 0.21602; 1.09: 0.066156; 1.10: 0.094004; 1.11: 0.19661
| Rel.Std.Dev. |  | 1.01: 2.441%; 1.03: 2.039%; 1.06: 3.848%; 1.07: 2.532%; 1.09: 0.805%; 1.10: 2.082%; 1.11: 2.295%
| PDM3 |  | 1.01: 1.5%; 1.03: 2.244%; 1.06: 4.159%; 1.07: -1.558%; 1.09: -5.191%; 1.10: -47.92%; 1.11: -1.154%
| Z-Score (Absolute) |  | 1.01: 0.53; 1.03: 0.67; 1.06: 1.05; 1.07: 0.08; 1.09: 0.8; 1.10: 9.28; 1.11: 0
| Analytical results for Rb in OREAS 255c (Certified Value 79 ppm) | [19 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 80.46 | 85.3212 | 78.7 | 64.7 | 83.1 | 80.5 | 74.4 | 53.9531 | 73.9 | 80.9398 | 74.4 | 80.1 | 85.7 | 82.5 | 79.3 |  |  |  | 1
| 1 | 2 | 81.37 | 83.3783 | 79.4 | 63.8 | 81.8 | 80.7 | 70.8 | 54.6761 | 76.2 | 79.3189 | 74 | 81.8 | 81.7 | 80.6 | 76.7 |  |  |  | 31
| 1 | 3 | 81.96 | 84.9445 | 79.1 | 66.9 | 81.5 | 83.7 | 75.8 | 54.6629 | 76.8 | 80.0269 | 73.5 | 88.9 | 83.6 | 83.5 | 76.6 |  |  |  | 16
| 1 | 4 | 81.7 | 84.8424 | 80.4 | 66.1 | 81.7 | 78.7 | 73 | 52.5512 | 73 | 73.2096 | 74.5 | 83.9 | 82.3 | 82.1 | 76.6 |  |  |  | 78.7047
| 1 | 5 | 81.67 | 84.6589 | 79.7 | 65.1 | 81.3 | 82.9 | 77.8 | 53.9207 | 75.3 | 72.9521 | 74.3 | 82.3 | 83 | 82 | 79 |  |  |  | 51
| 1 | 6 | 81.65 | 85.524 | 80.3 | 69.1 | 81.6 | 82.6 | 69.5 | 52.5577 | 77.3 | 76.9673 | 74.7 | 83.6 | 84 | 79.7 | 78.2
| Mean |  | 81.4683 | 84.7782 | 79.6 | 65.95 | 81.8333 | 81.5167 | 73.55 | 53.7203 | 75.4167 | 77.2358 | 74.2333 | 83.4333 | 83.3833 | 81.7333 | 77.7333
| Median |  | 81.66 | 84.8935 | 79.55 | 65.6 | 81.65 | 81.65 | 73.7 | 53.9369 | 75.75 | 78.1431 | 74.35 | 82.95 | 83.3 | 82.05 | 77.45
| Std Dev. |  | 0.5283 | 0.75588 | 0.66933 | 1.8844 | 0.64395 | 1.8702 | 3.1021 | 0.96072 | 1.6869 | 3.478 | 0.4274 | 3.0064 | 1.4106 | 1.3663 | 1.258
| Rel.Std.Dev. |  | 0.6485% | 0.8916% | 0.8409% | 2.857% | 0.7869% | 2.294% | 4.218% | 1.788% | 2.237% | 4.503% | 0.5757% | 3.603% | 1.692% | 1.672% | 1.618%
| PDM3 |  | 3.511% | 7.717% | 1.138% | -16.21% | 3.975% | 3.573% | -6.549% | -31.74% | -4.178% | -1.866% | -5.681% | 6.008% | 5.945% | 3.848% | -1.234%
| Z-Score (Absolute) |  | 0.33 | 0.92 | 0 | 2.43 | 0.4 | 0.34 | 1.08 | 4.61 | 0.75 | 0.42 | 0.96 | 0.68 | 0.67 | 0.38 | 0.33
| Analytical results for Re in OREAS 255c (Certified Value 0.001 ppm) | [19 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.0148 | 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.003 | [5 empty] | 1
| 1 | 2 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | 0.0174 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.002 | [5 empty] | 32
| 1 | 3 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | 0.0163 | 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.002 | [5 empty] | 16
| 1 | 4 | 0.002 | < 0.05 | < 0.001 | < 0.002 | < 0.002 | < 0.005 | 0.0134 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | < 0.002 | [5 empty] | 0.0014367
| 1 | 5 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | 0.017 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.006 | [5 empty] | 52
| 1 | 6 | < 0.002 | < 0.05 | < 0.001 | < 0.002 | 0.002 | < 0.005 | 0.0165 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.003
| Mean |  | 1.01: 0.002; 1.09: 0.002; 1.11: 0.0159; 1.12: 0.002; 1.15: 0.002; 1.17: 0.0032
| Median |  | 1.01: 0.002; 1.09: 0.002; 1.11: 0.0164; 1.12: 0.002; 1.15: 0.002; 1.17: 0.003
| Std Dev. |  | 1.09: 0; 1.11: 0.0015126; 1.12: 0; 1.15: 0; 1.17: 0.0016432
| Rel.Std.Dev. |  | 1.09: 0%; 1.11: 9.513%; 1.12: 0%; 1.15: 0%; 1.17: 51.35%
| PDM3 |  | 1.01: 39.21%; 1.09: 39.21%; 1.11: 1007%; 1.12: 39.21%; 1.15: 39.21%; 1.17: 122.7%
| Z-Score (Absolute) |  | 0.51 | 23.6 | 1.18 | 0.67 | 0 | 0.84 | 14.4 | 0.34 | 23.6 | 0.67 | 0.34 | 0.67 | 1.18
| Analytical results for S in OREAS 255c (Certified Value 0.039 wt.%) | [19 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0359 | 0.043718 | 0.45295 | 0.033 | 0.04 | 0.04 | 0.04 | 0.05 | < 0.1 | 0.032014 | 0.04 | 0.035956 | 0.04 | 0.02 | 0.04 | 0.04 | 0.032 |  | 1
| 1 | 2 | 0.034 | 0.041771 | 0.44581 | 0.033 | 0.04 | 0.04 | 0.04 | 0.05 | < 0.1 | 0.03114 | 0.04 | 0.036579 | 0.04 | 0.02 | 0.04 | 0.04 | 0.034 |  | 17
| 1 | 3 | 0.0343 | 0.042849 | 0.44471 | 0.034 | 0.04 | 0.04 | 0.04 | 0.05 | < 0.1 | 0.032206 | 0.04 | 0.036426 | 0.04 | 0.03 | 0.04 | 0.04 | 0.035 |  | 16
| 1 | 4 | 0.035 | 0.044532 | 0.45127 | 0.034 | 0.04 | 0.04 | 0.04 | 0.05 | < 0.1 | 0.032161 | 0.04 | 0.035637 | 0.037 | 0.03 | 0.05 | 0.04 | 0.035 |  | 0.038779
| 1 | 5 | 0.0343 | 0.041281 | 0.46761 | 0.034 | 0.03 | 0.04 | 0.04 | 0.05 | < 0.1 | 0.032005 | 0.04 | 0.036868 | 0.037 | 0.02 | 0.05 | 0.04 | 0.035 |  | 53
| 1 | 6 | 0.0344 | 0.043738 | 0.46801 | 0.034 | 0.03 | 0.04 | 0.04 | 0.05 | < 0.1 | 0.032079 | 0.04 | 0.035291 | 0.04 | 0.02 | 0.04 | 0.04 | 0.035
| Mean |  | 0.03465 | 0.042981 | 0.45506 | 0.033667 | 0.036667 | 0.04 | 0.04 | 0.05 |  | 0.031934 | 0.04 | 0.036126 | 0.039 | 0.023333 | 0.043333 | 0.04 | 0.034333
| Median |  | 0.03435 | 0.043283 | 0.45211 | 0.034 | 0.04 | 0.04 | 0.04 | 0.05 |  | 0.032047 | 0.04 | 0.036191 | 0.04 | 0.02 | 0.04 | 0.04 | 0.035
| Std Dev. |  | 0.00069498 | 0.0012566 | 0.010359 | 0.0005164 | 0.005164 | 0 | 0 | 7.6012e-18 |  | 0.000397 | 0 | 0.00060176 | 0.0015492 | 0.005164 | 0.005164 | 0 | 0.0012111
| Rel.Std.Dev. |  | 2.006% | 2.923% | 2.276% | 1.534% | 14.08% | 0% | 0% | 1.52e-14% |  | 1.243% | 0% | 1.666% | 3.972% | 22.13% | 11.92% | 0% | 3.527%
| PDM3 |  | -10.65% | 10.84% | 1073% | -13.18% | -5.448% | 3.148% | 3.148% | 28.94% |  | -17.65% | 3.148% | -6.842% | 0.5695% | -39.83% | 11.74% | 3.148% | -11.46%
| Z-Score (Absolute) |  | 0.93 | 0.52 | 72.24 | 1.1 | 0.58 | 0 | 0 | 1.74 | 1.74 | 1.4 | 0 | 0.67 | 0.17 | 2.9 | 0.58 | 0 | 0.99
| Analytical results for Sb in OREAS 255c (Certified Value 8.01 ppm) | [19 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 8.14 | 8.1084 | 10.0823 | 7.7 | 0.2 | 8.01 | 8.2067 | 10.4 | 8 | 8.5198 | 7.43 | 8.2366 | 7.61 | 7.42 | 8.97 | 7.78 | 8.51 |  | 1
| 1 | 2 | 8.38 | 7.612 | 5.5143 | 7.8 | 0.2 | 7.91 | 8.2833 | 8.18 | 7.6 | 8.7147 | 7.59 | 8.2336 | 7.51 | 7.59 | 8.46 | 7.8 | 8.12 |  | 33
| 1 | 3 | 8.41 | 7.882 | 7.3591 | 7.8 | 0.1 | 7.87 | 8.04 | 8.52 | 8.2 | 8.7344 | 7.77 | 6.5494 | 7.66 | 7.54 | 8.57 | 8.27 | 8.04 |  | 16
| 1 | 4 | 8.39 | 7.7223 | 8.842 | 7.7 | 0.2 | 7.87 | 8.13 | 8.26 | 7.8 | 8.8138 | 7.46 | 6.8619 | 7.84 | 7.58 | 8.43 | 7.78 | 8.1 |  | 8.0054
| 1 | 5 | 8.33 | 7.7933 | 9.9421 | 7.9 | 0.2 | 7.87 | 8.3333 | 8.51 | 8.2 | 9.1882 | 7.55 | 7.7853 | 7.49 | 7.57 | 8.26 | 7.63 | 8.17 |  | 54
| 1 | 6 | 8.4 | 7.7491 | 6.8007 | 7.8 | 0.3 | 7.95 | 8.8567 | 13.5 | 7.4 | 8.8153 | 7.84 | 7.7387 | 7.59 | 7.6 | 8.58 | 7.76 | 8.24
| Mean |  | 8.3417 | 7.8112 | 8.0901 | 7.7833 | 0.2 | 7.9133 | 8.3083 | 9.5617 | 7.8667 | 8.7977 | 7.6067 | 7.5676 | 7.6167 | 7.55 | 8.545 | 7.8367 | 8.1967
| Median |  | 8.385 | 7.7712 | 8.1006 | 7.8 | 0.2 | 7.89 | 8.245 | 8.515 | 7.9 | 8.7741 | 7.57 | 7.762 | 7.6 | 7.575 | 8.515 | 7.78 | 8.145
| Std Dev. |  | 0.10265 | 0.17036 | 1.8328 | 0.075277 | 0.063246 | 0.057155 | 0.28847 | 2.098 | 0.3266 | 0.21963 | 0.16573 | 0.70746 | 0.12644 | 0.066933 | 0.23839 | 0.22097 | 0.16765
| Rel.Std.Dev. |  | 1.231% | 2.181% | 22.65% | 0.9672% | 31.62% | 0.7223% | 3.472% | 21.94% | 4.152% | 2.496% | 2.179% | 9.349% | 1.66% | 0.8865% | 2.79% | 2.82% | 2.045%
| PDM3 |  | 4.201% | -2.425% | 1.058% | -2.774% | -97.5% | -1.15% | 3.784% | 19.44% | -1.733% | 9.897% | -4.98% | -5.469% | -4.856% | -5.688% | 6.741% | -2.107% | 2.39%
| Z-Score (Absolute) |  | 1.07 | 0.13 | 0.5 | 0.19 | 17.28 | 0.11 | 1 | 3.82 | 0 | 2.1 | 0.59 | 0.67 | 0.56 | 0.71 | 1.53 | 0.07 | 0.74
| Analytical results for Sc in OREAS 255c (Certified Value 14.6 ppm) | [19 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 14.2 | 14.6797 | 13 | 15 | 14.9 | 14.3634 | 14.3 | 14 | 13.466 | 12.3 | 15.4904 | 13.7 | 14.7 | 15.2 | 14.5 | 16.9 |  |  | 1
| 1 | 2 | 14.8 | 14.3363 | 13 | 15 | 14.4 | 13.9934 | 14.4 | 14 | 13.3564 | 12.3 | 14.5541 | 13.8 | 15.1 | 14.6 | 14.2 | 16.3 |  |  | 34
| 1 | 3 | 15.3 | 14.8426 | 13 | 15 | 14.1 | 14.3079 | 15.2 | 15 | 13.2863 | 13 | 14.9744 | 13.5 | 15.2 | 14.8 | 14.6 | 15.8 |  |  | 16
| 1 | 4 | 14.6 | 14.5073 | 13 | 16 | 14.6 | 13.8195 | 14.4 | 15 | 13.4457 | 12.1 | 14.258 | 13.8 | 15.1 | 14.6 | 14.6 | 16.2 |  |  | 14.6472
| 1 | 5 | 15 | 14.6738 | 13 | 15 | 14.6 | 14.0156 | 15 | 14 | 13.3557 | 12.6 | 15.0557 | 13.8 | 14.9 | 15 | 14.4 | 16.8 |  |  | 55
| 1 | 6 | 15 | 14.7888 | 13 | 15 | 14.8 | 14.5262 | 14.7 | 15 | 13.3561 | 12.6 | 13.9613 | 13.8 | 14.8 | 14.8 | 14.4 | 16.6
| Mean |  | 14.8167 | 14.6381 | 13 | 15.1667 | 14.5667 | 14.171 | 14.6667 | 14.5 | 13.3777 | 12.4833 | 14.7156 | 13.7333 | 14.9667 | 14.8333 | 14.45 | 16.4333
| Median |  | 14.9 | 14.6768 | 13 | 15 | 14.6 | 14.1617 | 14.55 | 14.5 | 13.3562 | 12.45 | 14.7643 | 13.8 | 15 | 14.8 | 14.45 | 16.45
| Std Dev. |  | 0.38166 | 0.18747 | 0 | 0.40825 | 0.28752 | 0.26876 | 0.36697 | 0.54772 | 0.066629 | 0.31885 | 0.56348 | 0.12111 | 0.19664 | 0.23381 | 0.15166 | 0.41312
| Rel.Std.Dev. |  | 2.576% | 1.281% | 0% | 2.692% | 1.974% | 1.897% | 2.502% | 3.777% | 0.4981% | 2.554% | 3.829% | 0.8818% | 1.314% | 1.576% | 1.05% | 2.514%
| PDM3 |  | 1.157% | -0.06219% | -11.25% | 3.547% | -0.5498% | -3.251% | 0.133% | -1.005% | -8.667% | -14.77% | 0.4673% | -6.239% | 2.181% | 1.271% | -1.346% | 12.19%
| Z-Score (Absolute) |  | 0.62 | 0 | NA | NA | 0.25 | 1.61 | 0.1 | NA | 4.35 | 7.44 | 0.27 | 3.12 | 1.13 | 0.67 | 0.65 | 6.2
Results from laboratories 1.07, 1.09 and 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 255c (Indicative Value 0.93 ppm) | [19 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.8 | < 1.5 | 13.5834 | < 1 | 0.7 | < 1 | < 5 | 2 | < 1 | 0.2983 | < 1 | < 2 | 1.6 | < 1 | < 1 | 1 |  |  | 1
| 1 | 2 | 0.9 | < 1.5 | 16.997 | < 1 | 0.7 | < 1 | < 5 | 3 | 1 | 0.2929 | < 1 | < 2 | 1.5 | 1 | < 1 | < 1 |  |  | 35
| 1 | 3 | 1.2 | < 1.5 | 13.6609 | < 1 | 0.8 | < 1 | < 5 | 2 | 1 | 0.3024 | 1 | < 2 | 1.5 | < 1 | < 1 | < 1 |  |  | 16
| 1 | 4 | 0.7 | < 1.5 | 11.2864 | < 1 | 0.8 | < 1 | < 5 | 2 | < 1 | 0.297 | < 1 | < 2 | 1.4 | 1 | < 1 | 1 |  |  | 0.93303
| 1 | 5 | 0.9 | < 1.5 | 13.3249 | < 1 | 0.5 | < 1 | < 5 | 3 | < 1 | 0.3184 | < 1 | < 2 | 1.6 | 1 | < 1 | < 1 |  |  | 7
| 1 | 6 | 1 | < 1.5 | 16.837 | < 1 | 0.7 | < 1 | < 5 | 3 | < 1 | 0.2973 | < 1 | < 2 | 1.7 | < 1 | < 1 | < 1
| Mean |  | 0.91667 |  | 14.2816 |  | 0.7 |  |  | 2.5 | 1 | 0.30105 | 1 |  | 1.55 | 1 |  | 1
| Median |  | 0.9 |  | 13.6221 |  | 0.7 |  |  | 2.5 | 1 | 0.2978 | 1 |  | 1.55 | 1 |  | 1
| Std Dev. |  | 1.01: 0.17224; 1.06: 2.2209; 1.09: 0.10954; 1.12: 0.54772; 1.13: 0; 1.14: 0.0090255; 1.17: 0.10488; 1.19: 0; 1.21: 0
| Rel.Std.Dev. |  | 1.01: 18.79%; 1.06: 15.55%; 1.09: 15.65%; 1.12: 21.91%; 1.13: 0%; 1.14: 2.998%; 1.17: 6.767%; 1.19: 0%; 1.21: 0%
| PDM3 |  | -1.754% |  | 1431% |  | -24.98% |  |  | 167.9% | 7.178% | -67.73% | 7.178% |  | 66.13% | 7.178% |  | 7.178%
| Z-Score (Absolute) |  | 0.57 | 0.07 | 40.63 | 0.67 | 0.07 | 0.67 | 5.32 | 5.32 | 0.17 | 1.27 | 0.42 | 0.82 | 2.47 | 0.07 | 0.67 | 0.17
| Analytical results for Sm in OREAS 255c (Certified Value 6.2 ppm) | [19 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 6.16; 1.03: 6.1324; 1.06: 6.9489; 1.07: 5.5; 1.09: 6.13; 1.10: 6.1867; 1.11: 6.18 |  |  | 1
| 1 | 2 | 1.01: 6.38; 1.03: 6.1835; 1.06: 6.5502; 1.07: 6.2; 1.09: 6.07; 1.10: 6.1452; 1.11: 6.05 |  |  | 5
| 1 | 3 | 1.01: 6.39; 1.03: 6.0354; 1.06: 6.5596; 1.07: 6; 1.09: 6.02; 1.10: 6.3583; 1.11: 6.11 |  |  | 16
| 1 | 4 | 1.01: 6.23; 1.03: 6.1801; 1.06: 6.8888; 1.07: 6.6; 1.09: 6.06; 1.10: 6.2298; 1.11: 6.01 |  |  | 6.1999
| 1 | 5 | 1.01: 6.51; 1.03: 6.2211; 1.06: 7.0886; 1.07: 6.2; 1.09: 6.15; 1.10: 6.2799; 1.11: 6.11 |  |  | 56
| 1 | 6 | 1.01: 6.29; 1.03: 6.2239; 1.06: 6.6249; 1.07: 6.4; 1.09: 6.17; 1.10: 6.2502; 1.11: 6.07
| Mean |  | 1.01: 6.3267; 1.03: 6.1627; 1.06: 6.7768; 1.07: 6.15; 1.09: 6.1; 1.10: 6.2417; 1.11: 6.0883
| Median |  | 1.01: 6.335; 1.03: 6.1818; 1.06: 6.7568; 1.07: 6.2; 1.09: 6.1; 1.10: 6.24; 1.11: 6.09
| Std Dev. |  | 1.01: 0.12565; 1.03: 0.070743; 1.06: 0.22846; 1.07: 0.37815; 1.09: 0.058652; 1.10: 0.07428; 1.11: 0.058793
| Rel.Std.Dev. |  | 1.01: 1.986%; 1.03: 1.148%; 1.06: 3.371%; 1.07: 6.149%; 1.09: 0.9615%; 1.10: 1.19%; 1.11: 0.9657%
| PDM3 |  | 1.01: 2.045%; 1.03: -0.5996%; 1.06: 9.305%; 1.07: -0.8049%; 1.09: -1.611%; 1.10: 0.6739%; 1.11: -1.8%
| Z-Score (Absolute) |  | 1.01: 1.49; 1.03: 0; 1.06: 5.57; 1.07: 0.12; 1.09: 0.57; 1.10: 0.72; 1.11: 0.67
| Analytical results for Sn in OREAS 255c (Certified Value 5.65 ppm) | [19 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 5.8 | 5.7998 | 42.8841 | 5.9 | 2 | 5.4 | < 10 | 6 | 5.8 | 5.2641 | 5.3 | 4.7911 | 5.8 | 5.2 | 6.1 | 5.9 | 5.7 |  | 1
| 1 | 2 | 6.1 | 5.4595 | 42.5422 | 5.8 | 3 | 5.4 | < 10 | 5.9 | 5.9 | 5.3967 | 5.4 | 5.8241 | 5.7 | 5.3 | 5.6 | 5.6 | 5.8 |  | 20
| 1 | 3 | 6.2 | 5.5786 | 40.743 | 5.9 | 2 | 5.4 | < 10 | 6 | 5.8 | 5.3611 | 5.5 | 5.6216 | 5.7 | 5.4 | 5.8 | 5.7 | 5.6 |  | 16
| 1 | 4 | 6 | 5.5812 | 41.5773 | 6 | 2 | 5.2 | < 10 | 5.8 | 5.7 | 5.3783 | 5.3 | 4.7654 | 5.8 | 5.3 | 5.6 | 5.8 | 5.6 |  | 5.6548
| 1 | 5 | 6.3 | 5.4151 | 39.5383 | 5.8 | 2 | 5.4 | < 10 | 6 | 5.9 | 5.497 | 5.4 | 5.3342 | 5.8 | 5.4 | 6 | 5.6 | 5.7 |  | 57
| 1 | 6 | 6.2 | 5.773 | 44.2124 | 5.9 | 2 | 5.3 | < 10 | 5.9 | 5.4 | 5.3389 | 5.5 | 5.6296 | 5.7 | 5.4 | 5.8 | 5.6 | 5.6
| Mean |  | 6.1 | 5.6012 | 41.9162 | 5.8833 | 2.1667 | 5.35 |  | 5.9333 | 5.75 | 5.3727 | 5.4 | 5.3277 | 5.75 | 5.3333 | 5.8167 | 5.7 | 5.6667
| Median |  | 6.15 | 5.5799 | 42.0598 | 5.9 | 2 | 5.4 |  | 5.95 | 5.8 | 5.3697 | 5.4 | 5.4779 | 5.75 | 5.35 | 5.8 | 5.65 | 5.65
| Std Dev. |  | 0.17889 | 0.15785 | 1.6581 | 0.075277 | 0.40825 | 0.083666 |  | 0.08165 | 0.18708 | 0.076327 | 0.089443 | 0.45347 | 0.054772 | 0.08165 | 0.20412 | 0.12649 | 0.08165
| Rel.Std.Dev. |  | 2.933% | 2.818% | 3.956% | 1.28% | 18.84% | 1.564% |  | 1.376% | 3.254% | 1.421% | 1.656% | 8.512% | 0.9526% | 1.531% | 3.509% | 2.219% | 1.441%
| PDM3 |  | 7.873% | -0.9483% | 641.2% | 4.041% | -61.68% | -5.39% |  | 4.925% | 1.683% | -4.989% | -4.506% | -5.785% | 1.683% | -5.685% | 2.862% | 0.7992% | 0.2097%
| Z-Score (Absolute) |  | 1.05 | 0.21 | 91.62 | 0.51 | NA | 0.84 | 1.73 | 0.63 | 0.17 | 0.79 | 0.72 | 0.9 | 0.17 | 0.89 | 0.34 | 0.04 | 0.04
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 255c (Certified Value 241 ppm) | [19 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 233.94 | 258.9516 | 239.2881 | 358.1455 | 231 | 224 | 238 | 261 | 235 | 238 | 216.2658 | 245 | 241.1134 | 244 | 238 | 251 | 239 | 224 | 1
| 1 | 2 | 233.75 | 261.6516 | 235.013 | 367.6344 | 231 | 225 | 238 | 255 | 245 | 243 | 217.0131 | 242 | 242.815 | 246 | 234 | 247 | 239 | 213 | 21
| 1 | 3 | 235.43 | 257.3208 | 241.3392 | 360.6974 | 232 | 240 | 237 | 260 | 246 | 256 | 230.2949 | 241 | 241.3345 | 243 | 249 | 249 | 241 | 225 | 16
| 1 | 4 | 233.35 | 268.3044 | 242.8246 | 361.4293 | 234 | 229 | 239 | 253 | 238 | 255 | 219.8573 | 236 | 238.6654 | 241 | 250 | 254 | 238 | 231 | 240.9621
| 1 | 5 | 237.35 | 265.464 | 242.5436 | 371.4202 | 235 | 221 | 240 | 260 | 250 | 246 | 232.0545 | 243 | 237.6933 | 243 | 238 | 256 | 244 | 236 | 58
| 1 | 6 | 231.39 | 262.116 | 247.6649 | 366.8779 | 234 | 234 | 236 | 257 | 249 | 256 | 228.3784 | 243 | 236.9534 | 242 | 240 | 254 | 232 | 236
| Mean |  | 234.2017 | 262.3014 | 241.4456 | 364.3675 | 232.8333 | 228.8333 | 238 | 257.6667 | 243.8333 | 249 | 223.9773 | 241.6667 | 239.7625 | 243.1667 | 241.5 | 251.8333 | 238.8333 | 227.5
| Median |  | 233.845 | 261.8838 | 241.9414 | 364.1536 | 233 | 227 | 238 | 258.5 | 245.5 | 250.5 | 224.1179 | 242.5 | 239.8894 | 243 | 239 | 252.5 | 239 | 228
| Std Dev. |  | 2.0161 | 4.0626 | 4.1912 | 5.0504 | 1.7224 | 7.0828 | 1.4142 | 3.2042 | 6.047 | 7.746 | 7.0635 | 3.0768 | 2.3234 | 1.7224 | 6.5038 | 3.4303 | 3.9707 | 8.7807
| Rel.Std.Dev. |  | 0.8609% | 1.549% | 1.736% | 1.386% | 0.7398% | 3.095% | 0.5942% | 1.244% | 2.48% | 3.111% | 3.154% | 1.273% | 0.969% | 0.7083% | 2.693% | 1.362% | 1.663% | 3.86%
| PDM3 |  | -2.806% | 8.856% | 0.2007% | 51.21% | -3.373% | -5.033% | -1.229% | 6.932% | 1.192% | 3.336% | -7.049% | 0.2924% | -0.4978% | 0.9149% | 0.2232% | 4.512% | -0.8834% | -5.587%
| Z-Score (Absolute) |  | 0.66 | 1.9 | 0 | 11.2 | 0.79 | 1.15 | 0.32 | 1.48 | 0.22 | 0.69 | 1.59 | 0.02 | 0.16 | 0.15 | 0 | 0.94 | 0.24 | 1.27
| Analytical results for Ta in OREAS 255c (Certified Value 1.76 ppm) | [19 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.82 | 2.1242 | 1.57 | < 0.1 | < 0.05 | 7.9067 | 1.81 | 1.7 | 2.265 | 1.56 | 1.8087 | 1.6 | 1.73 | 1.93 | 1.8 | 1.61 |  |  | 1
| 1 | 2 | 1.91 | 2.0291 | 1.61 | < 0.1 | < 0.05 | 7.4433 | 1.84 | 1.7 | 2.2928 | 1.55 | 1.7969 | 1.6 | 1.75 | 1.97 | 1.7 | 1.57 |  |  | 22
| 1 | 3 | 1.92 | 1.9829 | 1.56 | < 0.1 | < 0.05 | 8.1867 | 1.87 | 1.8 | 2.3485 | 1.59 | 1.6952 | 1.6 | 1.78 | 1.95 | 1.77 | 1.56 |  |  | 16
| 1 | 4 | 1.93 | 1.9899 | 1.71 | < 0.1 | < 0.05 | 7.3433 | 1.77 | 1.7 | 2.3803 | 1.56 | 1.6036 | 1.6 | 1.74 | 1.99 | 1.84 | 1.56 |  |  | 1.755
| 1 | 5 | 1.92 | 1.9504 | 1.62 | < 0.1 | < 0.05 | 7.7 | 1.85 | 1.7 | 2.4774 | 1.6 | 1.6036 | 1.6 | 1.75 | 1.92 | 1.68 | 1.67 |  |  | 59
| 1 | 6 | 1.9 | 2.1939 | 1.56 | < 0.1 | < 0.05 | 6.9233 | 1.86 | 1.8 | 2.3439 | 1.56 | 1.7057 | 1.6 | 1.74 | 1.91 | 1.74 | 1.67
| Mean |  | 1.9 | 2.0451 | 1.605 |  |  | 7.5839 | 1.8333 | 1.7333 | 2.3513 | 1.57 | 1.7023 | 1.6 | 1.7483 | 1.945 | 1.755 | 1.6067
| Median |  | 1.915 | 2.0095 | 1.59 |  |  | 7.5717 | 1.845 | 1.7 | 2.3462 | 1.56 | 1.7005 | 1.6 | 1.745 | 1.94 | 1.755 | 1.59
| Std Dev. |  | 0.040497 | 0.094391 | 0.057533 |  |  | 0.44609 | 0.037238 | 0.05164 | 0.074397 | 0.02 | 0.089251 | 2.4324e-16 | 0.017224 | 0.030822 | 0.060581 | 0.052409
| Rel.Std.Dev. |  | 2.131% | 4.616% | 3.585% |  |  | 5.882% | 2.031% | 2.979% | 3.164% | 1.274% | 5.243% | 1.52e-14% | 0.9852% | 1.585% | 3.452% | 3.262%
| PDM3 |  | 8.262% | 16.53% | -8.547% |  |  | 332.1% | 4.463% | -1.235% | 33.98% | -10.54% | -3.004% | -8.832% | -0.3799% | 10.83% | -3.614e-05% | -8.452%
| Z-Score (Absolute) |  | 0.72 | 1.37 | 0.61 | 7.6 | 7.71 | 26.27 | 0.42 | 0.03 | 2.74 | 0.77 | 0.17 | 0.63 | 0.03 | 0.92 | 0.06 | 0.6
| Analytical results for Tb in OREAS 255c (Certified Value 0.69 ppm) | [19 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.07: 1.03; 1.09: 1.09; 1.10: 1.10; 1.11: 1.14; 1.12: 1.17; 1.13: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.07: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*OES/MS; 1.13: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.68; 1.02: 0.74106; 1.07: 0.68485; 1.09: 0.6; 1.10: 0.68; 1.11: 0.36419; 1.12: 0.7; 1.13: 0.67 |  |  | 1
| 1 | 2 | 1.01: 0.72; 1.02: 0.74033; 1.07: 0.65482; 1.09: 0.6; 1.10: 0.68; 1.11: 0.36966; 1.12: 0.7; 1.13: 0.68 |  |  | 23
| 1 | 3 | 1.01: 0.7; 1.02: 0.72238; 1.07: 0.65697; 1.09: 0.7; 1.10: 0.67; 1.11: 0.32457; 1.12: 0.8; 1.13: 0.65 |  |  | 16
| 1 | 4 | 1.01: 0.7; 1.02: 0.74638; 1.07: 0.67293; 1.09: 0.7; 1.10: 0.68; 1.11: 0.36023; 1.12: 0.7; 1.13: 0.64 |  |  | 0.68982
| 1 | 5 | 1.01: 0.71; 1.02: 0.71233; 1.07: 0.69814; 1.09: 0.6; 1.10: 0.68; 1.11: 0.37169; 1.12: 0.7; 1.13: 0.67 |  |  | 60
| 1 | 6 | 1.01: 0.7; 1.02: 0.76217; 1.07: 0.6522; 1.09: 0.7; 1.10: 0.68; 1.11: 0.36492; 1.12: 0.7; 1.13: 0.66
| Mean |  | 1.01: 0.70167; 1.02: 0.73744; 1.07: 0.66999; 1.09: 0.65; 1.10: 0.67833; 1.11: 0.35921; 1.12: 0.71667; 1.13: 0.66167
| Median |  | 1.01: 0.7; 1.02: 0.7407; 1.07: 0.66495; 1.09: 0.65; 1.10: 0.68; 1.11: 0.36456; 1.12: 0.7; 1.13: 0.665
| Std Dev. |  | 1.01: 0.013292; 1.02: 0.01772; 1.07: 0.018646; 1.09: 0.054772; 1.10: 0.0040825; 1.11: 0.017456; 1.12: 0.040825; 1.13: 0.01472
| Rel.Std.Dev. |  | 1.01: 1.894%; 1.02: 2.403%; 1.07: 2.783%; 1.09: 8.427%; 1.10: 0.6018%; 1.11: 4.86%; 1.12: 5.696%; 1.13: 2.225%
| PDM3 |  | 1.01: 1.717%; 1.02: 6.904%; 1.07: -2.875%; 1.09: -5.772%; 1.10: -1.665%; 1.11: -47.93%; 1.12: 3.892%; 1.13: -4.081%
| Z-Score (Absolute) |  | 1.01: 0.93; 1.02: 2.13; 1.07: 0.14; 1.09: NA; 1.10: 0.14; 1.11: 10.62; 1.12: NA; 1.13: 0.42
Results from laboratories 1.09 and 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 255c (Certified Value 0.18 ppm) | [19 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.2 | 0.22405 | < 0.1 | < 0.1 | 0.17 | 7.248 | 0.16 | < 0.5 | 0.2217 | 0.18 | 0.30959 | 0.16 | 0.17 | 0.17 | 0.11 | 0.22 |  |  | 1
| 1 | 2 | < 0.2 | 0.21767 | < 0.1 | < 0.1 | 0.18 | 6.8507 | 0.14 | < 0.5 | 0.1301 | 0.17 | 0.24263 | 0.12 | 0.15 | 0.18 | 0.05 | 0.2 |  |  | 24
| 1 | 3 | < 0.2 | 0.23032 | 0.1 | < 0.1 | 0.18 | 7.7573 | 0.22 | < 0.5 | 0.1638 | 0.16 | 0.17849 | 0.19 | 0.17 | 0.16 | 0.15 | 0.19 |  |  | 16
| 1 | 4 | < 0.2 | 0.17368 | 0.1 | < 0.1 | 0.16 | 7.712 | 0.23 | < 0.5 | 0.1976 | 0.13 | 0.22121 | 0.15 | 0.18 | 0.17 | 0.1 | 0.18 |  |  | 0.17905
| 1 | 5 | < 0.2 | 0.18708 | 0.1 | < 0.1 | 0.17 | 7.56 | 0.18 | < 0.5 | 0.1879 | 0.18 | 0.15699 | 0.2 | 0.17 | 0.16 | 0.08 | 0.17 |  |  | 61
| 1 | 6 | < 0.2 | 0.23152 | < 0.1 | < 0.1 | 0.17 | 7.8187 | 0.23 | < 0.5 | 0.1831 | 0.15 | 0.32198 | 0.17 | 0.16 | 0.16 | 0.14 | 0.21
| Mean |  | 0.2 | 0.21072 | 0.1 |  | 0.17167 | 7.4911 | 0.19333 |  | 0.1807 | 0.16167 | 0.23848 | 0.165 | 0.16667 | 0.16667 | 0.105 | 0.195
| Median |  | 0.2 | 0.22086 | 0.1 |  | 0.17 | 7.636 | 0.2 |  | 0.1855 | 0.165 | 0.23192 | 0.165 | 0.17 | 0.165 | 0.105 | 0.195
| Std Dev. |  |  | 0.024388 | 1.6997e-17 |  | 0.0075277 | 0.37443 | 0.038816 |  | 0.03121 | 0.019408 | 0.067205 | 0.02881 | 0.010328 | 0.008165 | 0.037283 | 0.018708
| Rel.Std.Dev. |  |  | 11.57% | 1.7e-14% |  | 4.385% | 4.998% | 20.08% |  | 17.27% | 12% | 28.18% | 17.46% | 6.197% | 4.899% | 35.51% | 9.594%
| PDM3 |  | 11.7% | 17.69% | -44.15% |  | -4.122% | 4084% | 7.979% |  | 0.9234% | -9.707% | 33.2% | -7.845% | -6.914% | -6.914% | -41.36% | 8.91%
| Z-Score (Absolute) |  | NA | 1.3 | NA | 4.73 | 0.17 | 274.28 | 0.64 | 2.77 | 0.17 | 0.54 | 2.34 | 0.42 | 0.36 | 0.36 | 2.67 | 0.71
Results from laboratories 1.01 and 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 255c (Certified Value 9.9 ppm) | [19 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 10.04 | 9.4496 | 11.1858 | 9.78 | 10.1 | 8.87 | 10.05 | 10.4 | 12.5261 | 8.82 | 10.1508 | 9.18 | 10.1 | 10.5 | 9.82 | 9.5 |  |  | 1
| 1 | 2 | 10.29 | 9.5748 | 10.659 | 9.55 | 10.7 | 8.77 | 10 | 9.9 | 12.8951 | 8.73 | 10.1946 | 9.12 | 9.94 | 10.05 | 9.82 | 9.46 |  |  | 25
| 1 | 3 | 10.38 | 9.4425 | 10.7991 | 9.88 | 10.9 | 8.7 | 10.45 | 10.9 | 13.2658 | 9.28 | 10.3024 | 9.2 | 10.55 | 10.35 | 9.71 | 9.19 |  |  | 16
| 1 | 4 | 10.17 | 9.4856 | 11.1223 | 9.71 | 10.8 | 8.74 | 10.05 | 10.6 | 13.0343 | 8.81 | 9.278 | 9.73 | 10.35 | 9.9 | 10.1 | 9.11 |  |  | 9.8962
| 1 | 5 | 10.24 | 9.3456 | 10.9417 | 9.79 | 10.8 | 8.7 | 10.25 | 11.2 | 13.4362 | 9.12 | 9.2534 | 9.18 | 9.75 | 10.2 | 9.93 | 9.38 |  |  | 62
| 1 | 6 | 10.55 | 9.453 | 11.4644 | 9.82 | 11 | 8.88 | 10.05 | 11 | 13.1147 | 9.02 | 9.7298 | 9.19 | 10.55 | 10.2 | 9.85 | 9.31
| Mean |  | 10.2783 | 9.4585 | 11.0287 | 9.755 | 10.7167 | 8.7767 | 10.1417 | 10.6667 | 13.0454 | 8.9633 | 9.8182 | 9.2667 | 10.2067 | 10.2 | 9.8717 | 9.325
| Median |  | 10.265 | 9.4513 | 11.032 | 9.785 | 10.8 | 8.755 | 10.05 | 10.75 | 13.0745 | 8.92 | 9.9403 | 9.185 | 10.225 | 10.2 | 9.835 | 9.345
| Std Dev. |  | 0.17566 | 0.073986 | 0.2899 | 0.11467 | 0.31885 | 0.080664 | 0.1744 | 0.47188 | 0.31575 | 0.21285 | 0.4702 | 0.22871 | 0.33086 | 0.21213 | 0.13227 | 0.15268
| Rel.Std.Dev. |  | 1.709% | 0.7822% | 2.629% | 1.176% | 2.975% | 0.9191% | 1.72% | 4.424% | 2.42% | 2.375% | 4.789% | 2.468% | 3.242% | 2.08% | 1.34% | 1.637%
| PDM3 |  | 3.862% | -4.423% | 11.44% | -1.427% | 8.291% | -11.31% | 2.481% | 7.786% | 31.82% | -9.426% | -0.7885% | -6.361% | 3.137% | 3.07% | -0.2477% | -5.772%
| Z-Score (Absolute) |  | 0.3 | 0.61 | 1.14 | 0.28 | 0.79 | 1.37 | 0.15 | 0.74 | 3.39 | 1.16 | 0.21 | 0.83 | 0.22 | 0.22 | 0.15 | 0.76
| Analytical results for Ti in OREAS 255c (Certified Value 0.515 wt.%) | [19 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.5142 | 0.5347 | 0.51511 | < 0.1 | 0.499 | 0.351 | 0.511 | 0.5 | 0.517 | 0.504 | 0.36618 | 0.514 | 0.54949 | 0.52 | 0.514 | 0.511 | 0.509 | 0.411 | 1
| 1 | 2 | 0.5152 | 0.53347 | 0.50026 | 0.4442 | 0.509 | 0.377 | 0.509 | 0.48 | 0.531 | 0.512 | 0.36066 | 0.515 | 0.55591 | 0.52 | 0.506 | 0.52 | 0.505 | 0.435 | 26
| 1 | 3 | 0.5279 | 0.53778 | 0.50583 | < 0.1 | 0.491 | 0.315 | 0.51 | 0.49 | 0.523 | 0.513 | 0.36312 | 0.507 | 0.5335 | 0.51 | 0.533 | 0.52 | 0.504 | 0.44 | 16
| 1 | 4 | 0.536 | 0.54432 | 0.51721 | < 0.1 | 0.509 | 0.378 | 0.503 | 0.47 | 0.506 | 0.506 | 0.36923 | 0.5 | 0.52351 | 0.51 | 0.53 | 0.515 | 0.502 | 0.434 | 0.51492
| 1 | 5 | 0.5324 | 0.54551 | 0.49908 | 0.44733 | 0.497 | 0.385 | 0.509 | 0.49 | 0.535 | 0.505 | 0.36962 | 0.51 | 0.56859 | 0.52 | 0.524 | 0.521 | 0.511 | 0.432 | 63
| 1 | 6 | 0.5249 | 0.53636 | 0.52136 | < 0.1 | 0.501 | 0.392 | 0.506 | 0.49 | 0.521 | 0.519 | 0.36626 | 0.514 | 0.52992 | 0.51 | 0.515 | 0.52 | 0.505 | 0.433
| Mean |  | 0.5251 | 0.53869 | 0.50981 | 0.44576 | 0.501 | 0.36633 | 0.508 | 0.48667 | 0.52217 | 0.50983 | 0.36584 | 0.51 | 0.54349 | 0.515 | 0.52033 | 0.51783 | 0.506 | 0.43083
| Median |  | 0.5264 | 0.53707 | 0.51047 | 0.44576 | 0.5 | 0.3775 | 0.509 | 0.49 | 0.522 | 0.509 | 0.36622 | 0.512 | 0.54149 | 0.515 | 0.5195 | 0.52 | 0.505 | 0.4335
| Std Dev. |  | 0.0089084 | 0.0050518 | 0.009366 | 0.0022123 | 0.0070427 | 0.028731 | 0.0029665 | 0.010328 | 0.010323 | 0.0058452 | 0.0034729 | 0.0057619 | 0.01734 | 0.0054772 | 0.010405 | 0.0039707 | 0.0033466 | 0.010108
| Rel.Std.Dev. |  | 1.697% | 0.9378% | 1.837% | 0.4963% | 1.406% | 7.843% | 0.584% | 2.122% | 1.977% | 1.146% | 0.9493% | 1.13% | 3.19% | 1.064% | 2% | 0.7668% | 0.6614% | 2.346%
| PDM3 |  | 1.977% | 4.616% | -0.9931% | -13.43% | -2.704% | -28.86% | -1.344% | -5.487% | 1.407% | -0.9881% | -28.95% | -0.9558% | 5.548% | 0.01527% | 1.051% | 0.5655% | -1.733% | -16.33%
| Z-Score (Absolute) |  | 0.9 | 1.7 | 0 | 19.35 | 0.52 | 8.47 | 0.11 | 1.37 | 0.73 | 0 | 8.49 | 0.01 | 1.99 | 0.31 | 0.62 | 0.47 | 0.23 | 4.66
| Analytical results for Tl in OREAS 255c (Certified Value 0.46 ppm) | [19 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.46 | 0.14343 | 0.49588 | 0.47 | 0.46 | < 0.02 | 0.41 | < 0.5 | < 0.5727 | 0.44 | 0.46204 | 0.46 | 0.43 | 0.45 | 0.44 | 0.46 |  |  | 1
| 1 | 2 | 0.48 | 0.17249 | 0.48036 | 0.47 | 0.48 | < 0.02 | 0.44 | < 0.5 | < 0.5746 | 0.43 | 0.46762 | 0.46 | 0.43 | 0.44 | 0.37 | 0.47 |  |  | 27
| 1 | 3 | 0.47 | 0.17219 | 0.51038 | 0.49 | 0.47 | < 0.02 | 0.45 | < 0.5 | < 0.5761 | 0.44 | 0.46657 | 0.47 | 0.45 | 0.45 | 0.47 | 0.44 |  |  | 16
| 1 | 4 | 0.48 | 0.16454 | 0.48462 | 0.49 | 0.48 | < 0.02 | 0.42 | < 0.5 | < 0.5813 | 0.42 | 0.42348 | 0.46 | 0.45 | 0.43 | 0.45 | 0.43 |  |  | 0.45698
| 1 | 5 | 0.49 | 0.16943 | 0.48611 | 0.48 | 0.48 | < 0.02 | 0.44 | < 0.5 | < 0.5656 | 0.45 | 0.41639 | 0.46 | 0.43 | 0.45 | 0.45 | 0.46 |  |  | 64
| 1 | 6 | 0.46 | 0.16881 | 0.49168 | 0.48 | 0.48 | < 0.02 | 0.43 | < 0.5 | < 0.5743 | 0.45 | 0.44952 | 0.47 | 0.44 | 0.44 | 0.43 | 0.46
| Mean |  | 0.47333 | 0.16515 | 0.4915 | 0.48 | 0.475 |  | 0.43167 |  |  | 0.43833 | 0.4476 | 0.46333 | 0.43833 | 0.44333 | 0.435 | 0.45333
| Median |  | 0.475 | 0.16912 | 0.4889 | 0.48 | 0.48 |  | 0.435 |  |  | 0.44 | 0.45578 | 0.46 | 0.435 | 0.445 | 0.445 | 0.46
| Std Dev. |  | 0.012111 | 0.01102 | 0.010736 | 0.0089443 | 0.0083666 |  | 0.01472 |  |  | 0.01169 | 0.022487 | 0.005164 | 0.0098319 | 0.008165 | 0.034496 | 0.015055
| Rel.Std.Dev. |  | 2.559% | 6.673% | 2.184% | 1.863% | 1.761% |  | 3.41% |  |  | 2.667% | 5.024% | 1.115% | 2.243% | 1.842% | 7.93% | 3.321%
| PDM3 |  | 3.578% | -63.86% | 7.555% | 5.037% | 3.943% |  | -5.54% |  |  | -4.081% | -2.052% | 1.39% | -4.081% | -2.987% | -4.81% | -0.7983%
| Z-Score (Absolute) |  | 0.8 | 6.76 | 1.24 | 0.96 | 0.84 | 10.56 | 0.22 | 4.68 | 3.77 | 0.06 | 0.17 | 0.55 | 0.06 | 0.06 | 0.14 | 0.31
| Analytical results for Tm in OREAS 255c (Certified Value 0.21 ppm) | [19 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.12: 1.22 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.12: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.21; 1.02: 0.21836; 1.03: 0.2109; 1.07: 0.2; 1.09: 0.21; 1.10: 0.2113; 1.12: 0.2 |  |  | 1
| 1 | 2 | 1.01: 0.22; 1.02: 0.21819; 1.03: 0.20183; 1.07: 0.2; 1.09: 0.22; 1.10: 0.2105; 1.12: 0.2 |  |  | 6
| 1 | 3 | 1.01: 0.22; 1.02: 0.23304; 1.03: 0.20389; 1.07: 0.2; 1.09: 0.21; 1.10: 0.22; 1.12: 0.19 |  |  | 16
| 1 | 4 | 1.01: 0.24; 1.02: 0.21939; 1.03: 0.20579; 1.07: 0.2; 1.09: 0.21; 1.10: 0.2112; 1.12: 0.19 |  |  | 0.2127
| 1 | 5 | 1.01: 0.22; 1.02: 0.22853; 1.03: 0.22137; 1.07: 0.2; 1.09: 0.21; 1.10: 0.2163; 1.12: 0.2 |  |  | 65
| 1 | 6 | 1.01: 0.23; 1.02: 0.23456; 1.03: 0.19711; 1.07: 0.2; 1.09: 0.21; 1.10: 0.215; 1.12: 0.19
| Mean |  | 1.01: 0.22333; 1.02: 0.22535; 1.03: 0.20682; 1.07: 0.2; 1.09: 0.21167; 1.10: 0.21405; 1.12: 0.195
| Median |  | 1.01: 0.22; 1.02: 0.22396; 1.03: 0.20484; 1.07: 0.2; 1.09: 0.21; 1.10: 0.21315; 1.12: 0.195
| Std Dev. |  | 1.01: 0.010328; 1.02: 0.0076131; 1.03: 0.0084514; 1.07: 3.0405e-17; 1.09: 0.0040825; 1.10: 0.0037324; 1.12: 0.0054772
| Rel.Std.Dev. |  | 1.01: 4.624%; 1.02: 3.378%; 1.03: 4.086%; 1.07: 1.52e-14%; 1.09: 1.929%; 1.10: 1.744%; 1.12: 2.809%
| PDM3 |  | 1.01: 4.998%; 1.02: 5.944%; 1.03: -2.767%; 1.07: -5.972%; 1.09: -0.4868%; 1.10: 0.6337%; 1.12: -8.322%
| Z-Score (Absolute) |  | 1.01: 0.86; 1.02: 1.02; 1.03: 0.49; 1.07: NA; 1.09: 0.1; 1.10: 0.1; 1.12: 1.46
Results from laboratory 1.09 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 255c (Certified Value 1.82 ppm) | [19 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.88 | 1.7415 | 1.87 | 1.8 | 1.9 | 1.8 | 1.9 | 1.8 | 1.8991 | 1.7 | 1.8717 | 1.83 | 1.7 | 2 | 1.8 | 1.69 |  |  | 1
| 1 | 2 | 1.87 | 1.7409 | 1.7971 | 1.86 | 2 | 1.8 | 1.8 | 1.7 | 1.9741 | 1.7 | 1.8963 | 1.82 | 1.7 | 1.8 | 1.8 | 1.77 |  |  | 29
| 1 | 3 | 1.88 | 1.7025 | 1.775 | 1.81 | 2 | 1.7 | 1.9 | 1.8 | 2.0137 | 1.8 | 1.8485 | 1.83 | 1.7 | 1.9 | 1.8 | 1.63 |  |  | 16
| 1 | 4 | 1.85 | 1.7016 | 1.8174 | 1.82 | 2 | 1.7 | 1.8 | 1.7 | 1.9682 | 1.7 | 1.6777 | 1.82 | 1.7 | 1.9 | 1.9 | 1.76 |  |  | 1.8168
| 1 | 5 | 1.87 | 1.6614 | 1.7867 | 1.86 | 2 | 1.8 | 1.8 | 1.8 | 2.0419 | 1.7 | 1.7093 | 1.82 | 1.7 | 1.9 | 1.8 | 1.71 |  |  | 66
| 1 | 6 | 1.86 | 1.7363 | 1.8732 | 1.81 | 2.1 | 1.8 | 1.8 | 1.8 | 2.0015 | 1.8 | 1.8254 | 1.85 | 1.7 | 1.8 | 1.9 | 1.68
| Mean |  | 1.8683 | 1.714 | 1.8199 | 1.8267 | 2 | 1.7667 | 1.8333 | 1.7667 | 1.9831 | 1.7333 | 1.8048 | 1.8283 | 1.7 | 1.8833 | 1.8333 | 1.7067
| Median |  | 1.87 | 1.7194 | 1.8072 | 1.815 | 2 | 1.8 | 1.8 | 1.8 | 1.9878 | 1.7 | 1.8369 | 1.825 | 1.7 | 1.9 | 1.8 | 1.7
| Std Dev. |  | 0.01169 | 0.031712 | 0.042422 | 0.026583 | 0.063246 | 0.05164 | 0.05164 | 0.05164 | 0.049176 | 0.05164 | 0.089934 | 0.01169 | 0 | 0.075277 | 0.05164 | 0.052409
| Rel.Std.Dev. |  | 0.6257% | 1.85% | 2.331% | 1.455% | 3.162% | 2.923% | 2.817% | 2.923% | 2.48% | 2.979% | 4.983% | 0.6394% | 0% | 3.997% | 2.817% | 3.071%
| PDM3 |  | 2.838% | -5.656% | 0.1713% | 0.5442% | 10.08% | -2.758% | 0.9112% | -2.758% | 9.154% | -4.593% | -0.6593% | 0.6359% | -6.428% | 3.663% | 0.9112% | -6.061%
| Z-Score (Absolute) |  | 0.54 | 1.3 | 0.04 | 0.04 | 2.1 | 0.67 | 0.12 | 0.67 | 1.9 | 1.07 | 0.22 | 0.06 | 1.47 | 0.72 | 0.12 | 1.39
| Analytical results for V in OREAS 255c (Certified Value 98 ppm) | [19 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 99 | 92.0024 | 98.5432 | 79.8197 | 94 | 72 | 96 | 101.1867 | 99 | 101 | 86.2525 | 100 | 102.3201 | 96 | 98 | 99 | 98 | 89 | 1
| 1 | 2 | 98 | 91.63 | 97.5801 | 81.782 | 95 | 76 | 95 | 98.8133 | 100 | 103 | 86.2605 | 102 | 102.8457 | 97 | 93 | 99 | 99 | 94 | 30
| 1 | 3 | 98 | 92.0612 | 99.4592 | 81.8931 | 95 | 72 | 95 | 101.0533 | 99 | 104 | 85.274 | 100 | 100.6188 | 95 | 102 | 101 | 99 | 97 | 16
| 1 | 4 | 97 | 93.8742 | 98.6582 | 81.5804 | 96 | 78 | 95 | 99.495 | 98 | 103 | 86.8539 | 97 | 99.8849 | 96 | 100 | 99 | 98 | 95 | 98.0578
| 1 | 5 | 98 | 89.817 | 97.3356 | 81.4883 | 95 | 76 | 96 | 99.05 | 99 | 102 | 86.1689 | 100 | 103.8374 | 96 | 98 | 100 | 98 | 95 | 67
| 1 | 6 | 98 | 92.855 | 99.2345 | 81.976 | 96 | 81 | 96 | 102.4667 | 101 | 104 | 85.9771 | 100 | 100.3763 | 95 | 100 | 101 | 98 | 95
| Mean |  | 98 | 92.04 | 98.4685 | 81.4232 | 95.1667 | 75.8333 | 95.5 | 100.3442 | 99.3333 | 102.8333 | 86.1312 | 99.8333 | 101.6472 | 95.8333 | 98.5 | 99.8333 | 98.3333 | 94.1667
| Median |  | 98 | 92.0318 | 98.6007 | 81.6812 | 95 | 76 | 95.5 | 100.2742 | 99 | 103 | 86.2107 | 100 | 101.4695 | 96 | 99 | 99.5 | 98 | 95
| Std Dev. |  | 0.63246 | 1.3521 | 0.85819 | 0.80688 | 0.75277 | 3.4881 | 0.54772 | 1.4462 | 1.0328 | 1.169 | 0.51258 | 1.6021 | 1.5789 | 0.75277 | 3.0822 | 0.98319 | 0.5164 | 2.7142
| Rel.Std.Dev. |  | 0.6454% | 1.469% | 0.8715% | 0.991% | 0.791% | 4.6% | 0.5735% | 1.441% | 1.04% | 1.137% | 0.5951% | 1.605% | 1.553% | 0.7855% | 3.129% | 0.9848% | 0.5252% | 2.882%
| PDM3 |  | -0.05891% | -6.137% | 0.4188% | -16.96% | -2.948% | -22.66% | -2.608% | 2.332% | 1.301% | 4.87% | -12.16% | 1.811% | 3.661% | -2.268% | 0.451% | 1.811% | 0.281% | -3.968%
| Z-Score (Absolute) |  | 0.04 | 1.65 | 0.08 | 4.52 | 0.81 | 6.02 | 0.72 | 0.59 | 0.31 | 1.26 | 3.25 | 0.45 | 0.94 | 0.63 | 0.09 | 0.45 | 0.04 | 1.08
| Analytical results for W in OREAS 255c (Certified Value 33.2 ppm) | [19 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 1
2
| 1 | 1 | 34.4 | 33.6407 | 30.9361 | 34.1651 | 30.5 | 0.8 | 31 | 36.4088 | 34.8 | 32.5 | 35.5289 | 30.8 | 32.2574 | 26.2 | 33.1 | 36.1 | 36.9 | 32.9 | 1
| 1 | 2 | 36.7 | 32.919 | 31.9447 | 35.0365 | 30.1 | 1 | 30 | 34.8733 | 35.5 | 30.7 | 37.9465 | 31.9 | 32.1767 | 25 | 33.8 | 41 | 33 | 31.9 | 31
| 1 | 3 | 37.2 | 32.9108 | 31.0594 | 39.43 | 30.9 | 0.4 | 30 | 37.2388 | 33.1 | 31.2 | 38.2282 | 33.2 | 32.9116 | 26.6 | 33.2 | 35.4 | 34.6 | 31.5 | 16
| 1 | 4 | 34.8 | 33.4453 | 33.1885 | 36.7071 | 31.1 | 0.8 | 29 | 35.8216 | 33.3 | 30.6 | 37.787 | 31.4 | 31.5636 | 27.1 | 34 | 34.1 | 34.2 | 31.3 | 33.2382
| 1 | 5 | 33.8 | 33.0896 | 31.0018 | 39.0104 | 31.4 | 0.7 | 30 | 37.6859 | 32.9 | 31.8 | 38.935 | 31.9 | 32.8458 | 26.2 | 35 | 34.5 | 32 | 31.6 | 68
| 1 | 6 | 35 | 32.9764 | 32.919 | 37.7361 | 30.1 | 1 | 31 | 40.9716 | 34.5 | 30.9 | 38.4975 | 33.6 | 32.0094 | 25.6 | 35.9 | 36.6 | 35.7 | 30.8
| Mean |  | 35.3167 | 33.1636 | 31.8416 | 37.0142 | 30.6833 | 0.78333 | 30.1667 | 37.1667 | 34.0167 | 31.2833 | 37.8205 | 32.1333 | 32.2941 | 26.1167 | 34.1667 | 36.2833 | 34.4 | 31.6667
| Median |  | 34.9 | 33.033 | 31.5021 | 37.2216 | 30.7 | 0.8 | 30 | 36.8238 | 33.9 | 31.05 | 38.0874 | 31.9 | 32.217 | 26.2 | 33.9 | 35.75 | 34.4 | 31.55
| Std Dev. |  | 1.3393 | 0.30699 | 1.0122 | 2.1197 | 0.53821 | 0.22286 | 0.75277 | 2.1164 | 1.0629 | 0.73598 | 1.1946 | 1.069 | 0.51294 | 0.73869 | 1.0893 | 2.4943 | 1.7743 | 0.70616
| Rel.Std.Dev. |  | 3.792% | 0.9257% | 3.179% | 5.727% | 1.754% | 28.45% | 2.495% | 5.694% | 3.125% | 2.353% | 3.159% | 3.327% | 1.588% | 2.828% | 3.188% | 6.875% | 5.158% | 2.23%
| PDM3 |  | 6.253% | -0.2243% | -4.202% | 11.36% | -7.686% | -97.64% | -9.241% | 11.82% | 2.342% | -5.881% | 13.79% | -3.324% | -2.84% | -21.43% | 2.793% | 9.162% | 3.495% | -4.728%
| Z-Score (Absolute) |  | 0.94 | 0.16 | 0.32 | 1.55 | 0.74 | 11.59 | 0.93 | 1.61 | 0.47 | 0.52 | 1.85 | 0.22 | 0.16 | 2.4 | 0.52 | 1.29 | 0.61 | 0.39
| Analytical results for Y in OREAS 255c (Certified Value 16.1 ppm) | [19 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 16.33 | 17.0663 | 15.5996 | 16.2 | 15.8 | 15.5 | 15.35 | 16.6 | 14.8 | 13.5491 | 16.8 | 15.8507 | 16.6 | 16.1 | 17 | 15.8 | 14.7 |  | 1
| 1 | 2 | 16.43 | 17.0535 | 15.7789 | 16.8 | 15.5 | 15.2 | 14.9067 | 17.1 | 14 | 13.2492 | 16.4 | 16.1537 | 16.6 | 16.9 | 16.6 | 15.2 | 15.2 |  | 32
| 1 | 3 | 16.93 | 17.4604 | 15.7059 | 16.3 | 16.7 | 15.3 | 15.29 | 18.4 | 15.6 | 13.269 | 17 | 16.1224 | 16.5 | 16.7 | 17.2 | 16.7 | 14.1 |  | 16
| 1 | 4 | 16.86 | 17.5449 | 15.6974 | 16.8 | 16.1 | 15.2 | 14.7867 | 16.4 | 15.2 | 13.389 | 15.8 | 15.8166 | 16.3 | 16.3 | 16.6 | 17 | 14.4 |  | 16.1353
| 1 | 5 | 16.95 | 17.3393 | 15.7718 | 16.6 | 15.6 | 15.2 | 15.24 | 17.2 | 15.9 | 13.3989 | 16.5 | 15.9742 | 16.6 | 16.4 | 16.8 | 15.6 | 14.8 |  | 69
| 1 | 6 | 16.76 | 17.411 | 16.0762 | 16.7 | 16.6 | 15.2 | 15.4467 | 17.5 | 15 | 13.5091 | 16.8 | 16.1899 | 16.6 | 16.3 | 17 | 15.4 | 14.8
| Mean |  | 16.71 | 17.3126 | 15.7716 | 16.5667 | 16.05 | 15.2667 | 15.17 | 17.2 | 15.0833 | 13.394 | 16.55 | 16.0179 | 16.5333 | 16.45 | 16.8667 | 15.95 | 14.6667
| Median |  | 16.81 | 17.3751 | 15.7388 | 16.65 | 15.95 | 15.2 | 15.265 | 17.15 | 15.1 | 13.394 | 16.65 | 16.0483 | 16.6 | 16.35 | 16.9 | 15.7 | 14.75
| Std Dev. |  | 0.26601 | 0.20689 | 0.16263 | 0.2582 | 0.50892 | 0.12111 | 0.2625 | 0.71274 | 0.66458 | 0.1216 | 0.42778 | 0.16082 | 0.12111 | 0.29496 | 0.24221 | 0.73144 | 0.37771
| Rel.Std.Dev. |  | 1.592% | 1.195% | 1.031% | 1.559% | 3.171% | 0.7933% | 1.73% | 4.144% | 4.406% | 0.9079% | 2.585% | 1.004% | 0.7325% | 1.793% | 1.436% | 4.586% | 2.575%
| PDM3 |  | 3.561% | 7.296% | -2.254% | 2.673% | -0.5289% | -5.384% | -5.983% | 6.598% | -6.52% | -16.99% | 2.57% | -0.7278% | 2.467% | 1.95% | 4.532% | -1.149% | -9.102%
| Z-Score (Absolute) |  | 0.67 | 1.29 | 0.28 | 0.53 | 0 | 0.8 | 0.9 | 1.17 | 0.99 | 2.71 | 0.51 | 0.03 | 0.49 | 0.41 | 0.83 | 0.1 | 1.41
| Analytical results for Yb in OREAS 255c (Certified Value 1.42 ppm) | [19 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.06: 1.03; 1.07: 1.09; 1.09: 1.10; 1.10: 1.14; 1.11: 1.17; 1.12: 1.22 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*OES/MS; 1.12: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 1.51; 1.03: 1.3207; 1.06: 1.5179; 1.07: 1.5; 1.09: 1.42; 1.10: 1.3583; 1.11: 1.4; 1.12: 1.3 |  | 1
| 1 | 2 | 1.01: 1.55; 1.03: 1.4312; 1.06: 1.3478; 1.07: 1.5; 1.09: 1.41; 1.10: 1.3184; 1.11: 1.4; 1.12: 1.3 |  | 33
| 1 | 3 | 1.01: 1.49; 1.03: 1.4476; 1.06: 1.4025; 1.07: 1.6; 1.09: 1.41; 1.10: 1.3381; 1.11: 1.4; 1.12: 1.2 |  | 16
| 1 | 4 | 1.01: 1.52; 1.03: 1.4227; 1.06: 1.4065; 1.07: 1.6; 1.09: 1.38; 1.10: 1.3383; 1.11: 1.4; 1.12: 1.2 |  | 1.4168
| 1 | 5 | 1.01: 1.5; 1.03: 1.4421; 1.06: 1.4651; 1.07: 1.6; 1.09: 1.4; 1.10: 1.3381; 1.11: 1.4; 1.12: 1.3 |  | 70
| 1 | 6 | 1.01: 1.51; 1.03: 1.4989; 1.06: 1.3276; 1.07: 1.6; 1.09: 1.37; 1.10: 1.2883; 1.11: 1.4; 1.12: 1.3
| Mean |  | 1.01: 1.5133; 1.03: 1.4272; 1.06: 1.4112; 1.07: 1.5667; 1.09: 1.3983; 1.10: 1.3299; 1.11: 1.4; 1.12: 1.2667
| Median |  | 1.01: 1.51; 1.03: 1.4367; 1.06: 1.4045; 1.07: 1.6; 1.09: 1.405; 1.10: 1.3381; 1.11: 1.4; 1.12: 1.3
| Std Dev. |  | 1.01: 0.020656; 1.03: 0.058565; 1.06: 0.071261; 1.07: 0.05164; 1.09: 0.019408; 1.10: 0.023976; 1.11: 2.4324e-16; 1.12: 0.05164
| Rel.Std.Dev. |  | 1.01: 1.365%; 1.03: 4.103%; 1.06: 5.05%; 1.07: 3.296%; 1.09: 1.388%; 1.10: 1.803%; 1.11: 1.737e-14%; 1.12: 4.077%
| PDM3 |  | 1.01: 6.811%; 1.03: 0.732%; 1.06: -0.3955%; 1.07: 10.58%; 1.09: -1.306%; 1.10: -6.134%; 1.11: -1.188%; 1.12: -10.6%
| Z-Score (Absolute) |  | 1.01: 1.49; 1.03: 0.3; 1.06: 0.08; 1.07: 2.23; 1.09: 0.1; 1.10: 1.05; 1.11: 0.08; 1.12: 1.93
| Analytical results for Zn in OREAS 255c (Certified Value 84 ppm) | [19 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 84 | 86.618 | 83.3439 | 67.1365 | 83 | 86 | 83 | 80.7806 | 82 | 78 | 73.8875 | 86 | 86.2371 | 82.9 | 83 | 87 | 82 | 71 | 1
| 1 | 2 | 84 | 88.466 | 81.4948 | 69.3489 | 85 | 85 | 81 | 80.2465 | 85 | 80 | 73.0984 | 86 | 84.7154 | 83.1 | 81 | 88 | 82 | 75 | 34
| 1 | 3 | 86 | 86.893 | 84.4809 | 68.4134 | 83 | 85 | 81 | 82.6178 | 84 | 81 | 73.7153 | 85 | 84.0284 | 87 | 86 | 88 | 82 | 75 | 16
| 1 | 4 | 85 | 89.401 | 83.0405 | 70.6785 | 85 | 86 | 81 | 78.3607 | 81 | 80 | 74.7404 | 83 | 81.3077 | 83.4 | 86 | 88 | 82 | 79 | 83.7642
| 1 | 5 | 85 | 89.456 | 83.0452 | 68.4164 | 83 | 85 | 80 | 77.6128 | 85 | 79 | 74.5438 | 85 | 83.9012 | 83.5 | 83 | 89 | 83 | 78 | 71
| 1 | 6 | 84 | 89.192 | 84.7647 | 68.1711 | 85 | 85 | 83 | 80.3408 | 84 | 80 | 73.9469 | 86 | 81.874 | 83.4 | 84 | 89 | 81 | 75
| Mean |  | 84.6667 | 88.3377 | 83.3617 | 68.6941 | 84 | 85.3333 | 81.5 | 79.9932 | 83.5 | 79.6667 | 73.9887 | 85.1667 | 83.6773 | 83.8833 | 83.8333 | 88.1667 | 82 | 75.5
| Median |  | 84.5 | 88.829 | 83.1946 | 68.4149 | 84 | 85 | 81 | 80.2937 | 84 | 80 | 73.9172 | 85.5 | 83.9648 | 83.4 | 83.5 | 88 | 82 | 75
| Std Dev. |  | 0.8165 | 1.2784 | 1.1756 | 1.2018 | 1.0954 | 0.5164 | 1.2247 | 1.7902 | 1.6432 | 1.0328 | 0.59206 | 1.169 | 1.8259 | 1.5433 | 1.9408 | 0.75277 | 0.63246 | 2.8107
| Rel.Std.Dev. |  | 0.9644% | 1.447% | 1.41% | 1.75% | 1.304% | 0.6052% | 1.503% | 2.238% | 1.968% | 1.296% | 0.8002% | 1.373% | 2.182% | 1.84% | 2.315% | 0.8538% | 0.7713% | 3.723%
| PDM3 |  | 1.077% | 5.46% | -0.4806% | -17.99% | 0.2815% | 1.873% | -2.703% | -4.502% | -0.3154% | -4.892% | -11.67% | 1.674% | -0.1038% | 0.1422% | 0.08252% | 5.256% | -2.106% | -9.866%
| Z-Score (Absolute) |  | 0.44 | 1.92 | 0.09 | 6.03 | 0.17 | 0.71 | 0.85 | 1.45 | 0.04 | 1.59 | 3.88 | 0.64 | 0.04 | 0.12 | 0.1 | 1.85 | 0.64 | 3.27
| Analytical results for Zr in OREAS 255c (Certified Value 181 ppm) | [19 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 179.9 | 171.342 | 185.581 | 192.1246 | 178 | 75 | 176 | 167 | 188 | 178.9 | 148.368 | 177.5 | 185.7927 | 185 | 187 | 187.5 | 184 | 163 | 1
| 1 | 2 | 191.3 | 174.1608 | 180.0891 | 195.5962 | 182 | 159 | 175 | 162 | 186 | 182.6 | 145.5771 | 178 | 182.9318 | 185 | 186.5 | 183 | 179 | 162 | 35
| 1 | 3 | 187.3 | 171.2664 | 184.1756 | 186.8852 | 183 | 113 | 175 | 166 | 188.5 | 186.1 | 146.518 | 183.5 | 175.6754 | 185 | 193 | 187.5 | 184.5 | 165 | 16
| 1 | 4 | 190.2 | 177.39 | 185.307 | 186.1096 | 185 | 20 | 174 | 162 | 180.5 | 182.8 | 147.5837 | 174.5 | 175.7906 | 186 | 193 | 185.5 | 184.5 | 163 | 180.828
| 1 | 5 | 188.6 | 172.854 | 181.646 | 186.4894 | 181 | 40 | 175 | 163 | 190.5 | 182.1 | 147.8874 | 181 | 182.6397 | 188 | 186.5 | 187 | 183.5 | 161 | 72
| 1 | 6 | 188.1 | 172.8216 | 187.0564 | 193.8621 | 179 | 39 | 175 | 168 | 190.5 | 186.6 | 146.6845 | 181 | 173.2048 | 187 | 188 | 186.5 | 181 | 162
| Mean |  | 187.5667 | 173.3058 | 183.9759 | 190.1779 | 181.3333 | 74.3333 | 175 | 164.6667 | 187.3333 | 183.1833 | 147.1031 | 179.25 | 179.3392 | 186 | 189 | 186.1667 | 182.75 | 162.6667
| Median |  | 188.35 | 172.8378 | 184.7413 | 189.5049 | 181.5 | 57.5 | 175 | 164.5 | 188.25 | 182.7 | 147.1341 | 179.5 | 179.2151 | 185.5 | 187.5 | 186.75 | 183.75 | 162.5
| Std Dev. |  | 4.0267 | 2.2745 | 2.6235 | 4.1885 | 2.582 | 52.9969 | 0.63246 | 2.6583 | 3.7506 | 2.8351 | 1.0288 | 3.2055 | 5.0812 | 1.2649 | 3.1464 | 1.7224 | 2.2528 | 1.3663
| Rel.Std.Dev. |  | 2.147% | 1.312% | 1.426% | 2.202% | 1.424% | 71.3% | 0.3614% | 1.614% | 2.002% | 1.548% | 0.6994% | 1.788% | 2.833% | 0.6801% | 1.665% | 0.9252% | 1.233% | 0.8399%
| PDM3 |  | 3.727% | -4.16% | 1.741% | 5.171% | 0.2794% | -58.89% | -3.223% | -8.937% | 3.598% | 1.303% | -18.65% | -0.8727% | -0.8234% | 2.86% | 4.519% | 2.952% | 1.063% | -10.04%
| Z-Score (Absolute) |  | 0.69 | 1.09 | 0.24 | 1.01 | 0.09 | 13.43 | 0.88 | 2.17 | 0.66 | 0.14 | 4.36 | 0.35 | 0.34 | 0.49 | 0.87 | 0.51 | 0.09 | 2.42
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C707: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C725: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C743: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C761: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 255c (Certified Value 1.04 ppm) | [20 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.07 | 1.0344 | < 10 | 1.01 | 1.06 | 1.02 | 1.0239 | 1.03 | 1 | 1.162 | 1 | 0.92129 | 1.1 | 0.99 | 1.07 | 1.06 |  |  |  | 1
| 1 | 2 | 1.03 | 1.0581 | < 10 | 1.03 | 1.13 | 1.03 | 0.9893 | 1.02 | 1 | 1.174 | 1.01 | 0.89683 | 1.1 | 1.09 | 1.08 | 1.09 |  |  |  | 19
| 1 | 3 | 1.05 | 1.0421 | < 10 | 0.98 | 1.04 | 1.03 | 0.985 | 1.02 | 1 | 1.194 | 1.03 | 0.96257 | 1.1 | 1.03 | 1.07 | 1.06 |  |  |  | 16
| 1 | 4 | 1.07 | 1.0766 | < 10 | 1 | 1.09 | 1.02 | 1.0198 | 1.05 | 1 | 1.177 | 1.02 | 1.0075 | 1.1 | 1.03 | 1.08 | 1.1 |  |  |  | 1.0415
| 1 | 5 | 1.05 | 1.0997 | < 10 | 1 | 0.994 | 1.03 | 0.99607 | 1.05 | 1.1 | 1.209 | 1 | 0.989 | 1.1 | 1.01 | 1.08 | 1.04 |  |  |  | 74
| 1 | 6 | 1.07 | 1.0166 | < 10 | 1.01 | 1.03 | 1.04 | 0.9692 | 1.06 | 1 | 1.17 | 1.06 | 0.91569 | 1.1 | 1.01 | 1.06 | 1.04
| Mean |  | 1.0567 | 1.0546 |  | 1.005 | 1.0573 | 1.0283 | 0.99722 | 1.0383 | 1.0167 | 1.181 | 1.02 | 0.94881 | 1.1 | 1.0267 | 1.0733 | 1.065
| Median |  | 1.06 | 1.0501 |  | 1.005 | 1.05 | 1.03 | 0.99268 | 1.04 | 1 | 1.1755 | 1.015 | 0.94193 | 1.1 | 1.02 | 1.075 | 1.06
| Std Dev. |  | 0.01633 | 0.03013 |  | 0.016432 | 0.047777 | 0.0075277 | 0.021081 | 0.017224 | 0.040825 | 0.017321 | 0.022804 | 0.044289 | 2.4324e-16 | 0.034448 | 0.008165 | 0.0251
| Rel.Std.Dev. |  | 1.545% | 2.857% |  | 1.635% | 4.519% | 0.732% | 2.114% | 1.659% | 4.016% | 1.467% | 2.236% | 4.668% | 2.211e-14% | 3.355% | 0.7607% | 2.357%
| PDM3 |  | 1.459% | 1.259% |  | -3.502% | 1.523% | -1.262% | -4.249% | -0.3013% | -2.382% | 13.4% | -2.062% | -8.897% | 5.62% | -1.422% | 3.059% | 2.259%
| Z-Score (Absolute) |  | 0.26 | 0.21 | 99.97 | 1.05 | 0.28 | 0.46 | 1.25 | 0.21 | 0.75 | 3.4 | 0.67 | 2.47 | 1.35 | 0.5 | 0.68 | 0.47
| Analytical results for Al in OREAS 255c (Certified Value 1.84 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 1.9302 | 2.1805 | 1.8003 | 1.54 | 2.05 | 1.96 | 1.61 | 1.94 | 1.74 | 1.78 | 1.65 | 2.0959 | 2.13 | 1.76 | 1.75 | 1.72 | 1.637 |  |  | 1
| 1 | 2 | 1.9018 | 2.2002 | 1.8367 | 1.62 | 2.09 | 2.04 | 1.6 | 1.92 | 1.77 | 1.59 | 1.68 | 1.9421 | 2.15 | 1.84 | 1.78 | 1.82 | 1.635 |  |  | 0
| 1 | 3 | 1.8719 | 2.1726 | 1.8168 | 1.4 | 2.15 | 1.86 | 1.65 | 1.94 | 1.68 | 1.75 | 1.7 | 1.9896 | 2.13 | 1.85 | 1.75 | 1.74 | 1.65 |  |  | 16
| 1 | 4 | 1.8603 | 2.2059 | 1.8061 | 1.32 | 2.11 | 2.01 | 1.63 | 1.95 | 1.75 | 1.71 | 1.63 | 2.0607 | 2.12 | 1.8 | 1.85 | 1.78 | 1.656 |  |  | 1.8413
| 1 | 5 | 1.9583 | 2.2123 | 1.826 | 1.68 | 2.17 | 1.82 | 1.6 | 1.94 | 1.75 | 1.71 | 1.66 | 1.9902 | 2.11 | 1.77 | 1.83 | 1.79 | 1.614 |  |  | 75
| 1 | 6 | 1.9009 | 2.1906 | 1.8151 | 1.53 | 2.12 | 1.88 | 1.59 | 1.91 | 1.76 | 1.92 | 1.69 | 1.9014 | 2.11 | 1.75 | 1.8 | 1.74 | 1.661
| Mean |  | 1.9039 | 2.1937 | 1.8168 | 1.515 | 2.115 | 1.9283 | 1.6133 | 1.9333 | 1.7417 | 1.7433 | 1.6683 | 1.9966 | 2.125 | 1.795 | 1.7933 | 1.765 | 1.6422
| Median |  | 1.9014 | 2.1954 | 1.816 | 1.535 | 2.115 | 1.92 | 1.605 | 1.94 | 1.75 | 1.73 | 1.67 | 1.9899 | 2.125 | 1.785 | 1.79 | 1.76 | 1.6435
| Std Dev. |  | 0.036303 | 0.015251 | 0.013199 | 0.13442 | 0.042778 | 0.088185 | 0.022509 | 0.015055 | 0.031885 | 0.10801 | 0.026394 | 0.072219 | 0.015166 | 0.042308 | 0.041312 | 0.037815 | 0.017198
| Rel.Std.Dev. |  | 1.907% | 0.6952% | 0.7265% | 8.873% | 2.023% | 4.573% | 1.395% | 0.7787% | 1.831% | 6.196% | 1.582% | 3.617% | 0.7137% | 2.357% | 2.304% | 2.143% | 1.047%
| PDM3 |  | 3.399% | 19.14% | -1.329% | -17.72% | 14.86% | 4.726% | -12.38% | 4.998% | -5.411% | -5.321% | -9.394% | 8.436% | 15.41% | -2.515% | -2.605% | -4.144% | -10.82%
| Z-Score (Absolute) |  | 0.55 | 2.02 | 0.11 | 1.42 | 1.62 | 0.67 | 0.92 | 0.7 | 0.27 | 0.26 | 0.64 | 1.02 | 1.67 | 0 | 0.01 | 0.15 | 0.77
| Analytical results for As in OREAS 255c (Certified Value 153 ppm) | [20 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 0
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 0
| 1 | 1 | 160 | 158.5359 | 111.816 | 143 | 150 | 154 | 156 | 151 | 158 | 144.4 | 161.137 | 155.5 | 144.6901 | 157 | 156 | 156 | 150 | 151 |  | 1
| 1 | 2 | 157 | 165.251 | 115.465 | 152 | 150 | 158 | 150 | 152 | 158.5 | 136.5 | 166.645 | 160 | 146.3285 | 160 | 162 | 158 | 156 | 143 |  | 20
| 1 | 3 | 158 | 155.8413 | < 100 | 139 | 149 | 151 | 153 | 152 | 155 | 158.8 | 156.832 | 162 | 147.4235 | 155 | 159.5 | 159 | 150.5 | 144 |  | 16
| 1 | 4 | 154 | 159.837 | 134.007 | 133 | 151 | 152 | 151 | 153 | 157.5 | 137.2 | 161.761 | 159 | 145.5273 | 155 | 157.5 | 160 | 156 | 149 |  | 153.4735
| 1 | 5 | 156 | 156.0957 | 132.197 | 150 | 152 | 145 | 148 | 152 | 155 | 140.7 | 165.731 | 158 | 147.7875 | 154 | 156 | 156.5 | 150 | 149 |  | 76
| 1 | 6 | 158 | 161.3142 | 103.6 | 146 | 148 | 149 | 147 | 152 | 158.5 | 149.2 | 160.697 | 159 | 145.5621 | 160 | 155 | 160 | 154.5 | 150
| Mean |  | 157.1667 | 159.4792 | 119.417 | 143.8333 | 150 | 151.5 | 150.8333 | 152 | 157.0833 | 144.4667 | 162.1338 | 158.9167 | 146.2199 | 156.8333 | 157.6667 | 158.25 | 152.8333 | 147.6667
| Median |  | 157.5 | 159.1865 | 115.465 | 144.5 | 150 | 151.5 | 150.5 | 152 | 157.75 | 142.55 | 161.449 | 159 | 145.9453 | 156 | 156.75 | 158.5 | 152.5 | 149
| Std Dev. |  | 2.0412 | 3.5322 | 13.2266 | 7.0828 | 1.4142 | 4.4159 | 3.3116 | 0.63246 | 1.6558 | 8.4654 | 3.5944 | 2.1545 | 1.1975 | 2.6394 | 2.6394 | 1.7248 | 2.9777 | 3.3267
| Rel.Std.Dev. |  | 1.299% | 2.215% | 11.08% | 4.924% | 0.9428% | 2.915% | 2.196% | 0.4161% | 1.054% | 5.86% | 2.217% | 1.356% | 0.819% | 1.683% | 1.674% | 1.09% | 1.948% | 2.253%
| PDM3 |  | 2.406% | 3.913% | -22.19% | -6.281% | -2.263% | -1.286% | -1.72% | -0.9601% | 2.352% | -5.869% | 5.643% | 3.547% | -4.726% | 2.189% | 2.732% | 3.112% | -0.4171% | -3.784%
| Z-Score (Absolute) |  | 0.64 | 0.95 | 6.01 | 1.16 | 0.33 | 0.12 | 0.21 | 0.06 | 0.63 | 1.07 | 1.31 | 0.88 | 0.84 | 0.6 | 0.71 | 0.79 | 0.06 | 0.64
| Analytical results for B in OREAS 255c (Certified Value < 10 ppm) | [20 empty] | Certified
B | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.06; 1.05: 1.09; 1.06: 1.12; 1.07: 1.13; 1.09: 1.15; 1.10: 1.16; 1.11: 1.19; 1.12: 1.20; 1.13: 1.21 |  | ppm
|  |  | 1.01: AR*OES; 1.03: AR*OES; 1.05: AR*OES/MS; 1.06: AR*MS; 1.07: AR*MS; 1.09: AR*MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS |  | 1
|  |  | 1.01: 25g; 1.03: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.09: 0.5g; 1.10: 01g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g |  | 1
| 1 | 1 | 1.01: < 10; 1.03: 4.3; 1.05: 6; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 1
| 1 | 2 | 1.01: < 10; 1.03: 4.2; 1.05: 6; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 0
| 1 | 3 | 1.01: < 10; 1.03: 12.4; 1.05: 5; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 16
| 1 | 4 | 1.01: < 10; 1.03: 7.5; 1.05: 6; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | < 10
| 1 | 5 | 1.01: < 10; 1.03: 5.2; 1.05: 5; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10 |  | 77
| 1 | 6 | 1.01: < 10; 1.03: 15.8; 1.05: 5; 1.06: < 10; 1.07: < 20; 1.09: < 10; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 10
| Mean |  | 1.03: 8.2333; 1.05: 5.5
| Median |  | 1.03: 6.35; 1.05: 5.5
| Std Dev. |  | 1.03: 4.8186; 1.05: 0.54772
| Rel.Std.Dev. |  | 1.03: 58.53%; 1.05: 9.959%
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.13
| Analytical results for Ba in OREAS 255c (Certified Value 94 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 98 | 109.9764 | 95.2297 | 96.15 | 101 | 93 | 93.9 | 95 | 93.5633 | 90 | 87 | 90 | 99.8713 | 90 | 90 | 90 | 89 |  |  | 1
| 1 | 2 | 95 | 111.7044 | 96.7495 | 98.468 | 103 | 94 | 94.5 | 94 | 92.5133 | 90 | 79 | 90 | 95.5355 | 100 | 90 | 90 | 87 |  |  | 21
| 1 | 3 | 96 | 109.7604 | 94.7025 | 96.483 | 95 | 94 | 92.2 | 91 | 93.2767 | 90 | 88 | 90 | 96.0216 | 100 | 90 | 90 | 87 |  |  | 16
| 1 | 4 | 94 | 112.2336 | 95.0246 | 97.739 | 93 | 96 | 93.2 | 97 | 94.5133 | 90 | 89 | 90 | 96.444 | 100 | 100 | 90 | 88 |  |  | 94.2241
| 1 | 5 | 97 | 111.78 | 96.3178 | 101.256 | 106 | 93 | 83.1 | 94 | 93.05 | 90 | 83 | 90 | 96.0562 | 90 | 100 | 90 | 86 |  |  | 78
| 1 | 6 | 94 | 110.2356 | 95.6399 | 101.223 | 102 | 94 | 89.8 | 95 | 92.0367 | 90 | 88 | 90 | 99.6024 | 90 | 90 | 90 | 92
| Mean |  | 95.6667 | 110.9484 | 95.6107 | 98.5532 | 100 | 94 | 91.1167 | 94.3333 | 93.1589 | 90 | 85.6667 | 90 | 97.2552 | 95 | 93.3333 | 90 | 88.1667
| Median |  | 95.5 | 110.97 | 95.4348 | 98.1035 | 101.5 | 94 | 92.7 | 94.5 | 93.1633 | 90 | 87.5 | 90 | 96.2501 | 95 | 90 | 90 | 87.5
| Std Dev. |  | 1.633 | 1.0751 | 0.78881 | 2.2434 | 4.98 | 1.0954 | 4.2593 | 1.9664 | 0.86023 | 0 | 3.8816 | 0 | 1.9457 | 5.4772 | 5.164 | 0 | 2.137
| Rel.Std.Dev. |  | 1.707% | 0.969% | 0.825% | 2.276% | 4.98% | 1.165% | 4.675% | 2.085% | 0.9234% | 0% | 4.531% | 0% | 2.001% | 5.766% | 5.533% | 0% | 2.424%
| PDM3 |  | 1.531% | 17.75% | 1.472% | 4.594% | 6.13% | -0.2378% | -3.298% | 0.116% | -1.13% | -4.483% | -9.082% | -4.483% | 3.217% | 0.8235% | -0.9453% | -4.483% | -6.429%
| Z-Score (Absolute) |  | 0.16 | 3.67 | 0.15 | 0.82 | 1.16 | 0.22 | 0.89 | 0.15 | 0.42 | NA | 2.14 | NA | 0.53 | NA | NA | NA | 1.57
Results from laboratories 1.12, 1.15, 1.19, 1.20 and 1.21 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 255c (Certified Value 1.03 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 1.06 | 1.0414 | < 10 | 1.2 | 0.9 | 1 | 0.92 | < 5 | 0.99 | 0.97 | 1.0666 | 1 | 1.02 | 1.01 | 0.96 | 1.03 |  |  |  | 1
| 1 | 2 | 1.07 | 1.09 | < 10 | 1.2 | 1 | 1.1 | 0.93 | < 5 | 1.01 | 0.97 | 1.0491 | 1.1 | 1.05 | 1.03 | 1.01 | 1.06 |  |  |  | 22
| 1 | 3 | 1.02 | 1.0467 | < 10 | 1.2 | 1 | 1 | 0.94 | < 5 | 0.97 | 0.97 | 1.0347 | 1 | 1.06 | 1.03 | 0.96 | 1.03 |  |  |  | 16
| 1 | 4 | 1.09 | 1.0595 | < 10 | 1.2 | 1 | 1.1 | 0.93 | < 5 | 1 | 0.95 | 1.0238 | 1 | 1.03 | 1.06 | 1 | 1.05 |  |  |  | 1.0252
| 1 | 5 | 1.05 | 1.0653 | < 10 | 1.3 | 1 | 1 | 0.93 | < 5 | 1 | 0.96 | 1.078 | 1.1 | 1.01 | 1.04 | 0.99 | 1.05 |  |  |  | 79
| 1 | 6 | 1.05 | 1.0899 | < 10 | 1.2 | 1 | 1.1 | 0.91 | < 5 | 1.01 | 0.97 | 1.0364 | 1.2 | 1 | 1.04 | 0.97 | 1.09
| Mean |  | 1.0567 | 1.0655 |  | 1.2167 | 0.98333 | 1.05 | 0.92667 |  | 0.99667 | 0.965 | 1.0481 | 1.0667 | 1.0283 | 1.035 | 0.98167 | 1.0517
| Median |  | 1.055 | 1.0624 |  | 1.2 | 1 | 1.05 | 0.93 |  | 1 | 0.97 | 1.0428 | 1.05 | 1.025 | 1.035 | 0.98 | 1.05
| Std Dev. |  | 0.023381 | 0.020807 |  | 0.040825 | 0.040825 | 0.054772 | 0.010328 |  | 0.015055 | 0.0083666 | 0.020716 | 0.08165 | 0.023166 | 0.016432 | 0.02137 | 0.022286
| Rel.Std.Dev. |  | 2.213% | 1.953% |  | 3.355% | 4.152% | 5.216% | 1.115% |  | 1.511% | 0.867% | 1.977% | 7.655% | 2.253% | 1.588% | 2.177% | 2.119%
| PDM3 |  | 3.073% | 3.932% |  | 18.68% | -4.08% | 2.423% | -9.608% |  | -2.779% | -5.868% | 2.239% | 4.049% | 0.3096% | 0.9599% | -4.243% | 2.586%
| Z-Score (Absolute) |  | 0.14 | 0.3 | 72.88 | 3.09 | 1.21 | 0.02 | 2.26 | 26.76 | 0.97 | 1.55 | 0.02 | 0.32 | 0.38 | 0.26 | 1.24 | 0.05
| Analytical results for Bi in OREAS 255c (Certified Value 2.67 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2.85 | 2.617 | < 10 | < 0.3 | 2.92 | 2.87 | 2.59 | < 2 | 2.68 | 2.5 | 2.78 | 2.43 | 2.7229 | 3.28 | 2.83 | 2.71 | 2.41 | 2.65 |  | 1
| 1 | 2 | 2.87 | 2.5098 | < 10 | 0.8 | 2.94 | 2.98 | 2.53 | < 2 | 2.56 | 2.3 | 2.699 | 2.41 | 2.7796 | 3.5 | 2.86 | 2.71 | 2.44 | 2.58 |  | 23
| 1 | 3 | 2.76 | 2.619 | < 10 | < 0.3 | 2.91 | 2.79 | 2.54 | < 2 | 2.46 | 2.5 | 2.731 | 2.42 | 2.7529 | 3.27 | 2.97 | 2.77 | 2.35 | 2.85 |  | 16
| 1 | 4 | 2.81 | 2.6102 | < 10 | < 0.3 | 2.98 | 2.79 | 2.61 | < 2 | 2.48 | 2.6 | 2.743 | 2.24 | 2.7676 | 3.13 | 2.94 | 2.8 | 2.53 | 2.61 |  | 2.665
| 1 | 5 | 2.79 | 2.7067 | < 10 | < 0.3 | 2.85 | 2.57 | 2.6 | < 2 | 2.55 | 2.7 | 2.739 | 2.43 | 2.7349 | 3.21 | 2.88 | 2.78 | 2.42 | 2.6 |  | 80
| 1 | 6 | 2.71 | 2.5749 | < 10 | < 0.3 | 2.93 | 2.7 | 2.52 | < 2 | 2.51 | 2.6 | 2.763 | 2.43 | 2.7909 | 3.36 | 2.66 | 2.65 | 2.48 | 2.59
| Mean |  | 2.7983 | 2.6063 |  | 0.8 | 2.9217 | 2.7833 | 2.565 |  | 2.54 | 2.5333 | 2.7425 | 2.3933 | 2.7581 | 3.2917 | 2.8567 | 2.7367 | 2.4383 | 2.6467
| Median |  | 2.8 | 2.6136 |  | 0.8 | 2.925 | 2.79 | 2.565 |  | 2.53 | 2.55 | 2.741 | 2.425 | 2.7602 | 3.275 | 2.87 | 2.74 | 2.43 | 2.605
| Std Dev. |  | 0.058793 | 0.064295 |  |  | 0.042622 | 0.14052 | 0.03937 |  | 0.07874 | 0.13663 | 0.027797 | 0.075542 | 0.026195 | 0.12766 | 0.1093 | 0.056451 | 0.061779 | 0.1025
| Rel.Std.Dev. |  | 2.101% | 2.467% |  |  | 1.459% | 5.049% | 1.535% |  | 3.1% | 5.393% | 1.014% | 3.156% | 0.9498% | 3.878% | 3.826% | 2.063% | 2.534% | 3.873%
| PDM3 |  | 5.003% | -2.204% |  | -69.98% | 9.63% | 4.44% | -3.753% |  | -4.691% | -4.941% | 2.907% | -10.19% | 3.493% | 23.51% | 7.191% | 2.689% | -8.506% | -0.6885%
| Z-Score (Absolute) |  | 0.46 | 0.37 | 10.04 | 10.59 | 1 | 0.4 | 0.55 | 7.36 | 0.66 | 0.69 | 0.22 | 1.3 | 0.29 | 2.61 | 0.72 | 0.2 | 1.1 | 0.2
| Analytical results for Ca in OREAS 255c (Certified Value 0.474 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 0.5 | 0.54908 | 0.47894 | 0.40068 | 1.09 | 0.5 | 0.59 | 0.42 | 0.49 | 0.47 | 0.46 | 0.4034 | 0.44 | 0.50875 | 0.54 | 0.48 | 0.46 | 0.44 | 0.404 | 1
| 1 | 2 | 0.49 | 0.56595 | 0.48836 | 0.41264 | 1.12 | 0.5 | 0.59 | 0.41 | 0.49 | 0.48 | 0.42 | 0.4055 | 0.44 | 0.50549 | 0.54 | 0.5 | 0.47 | 0.47 | 0.399 | 0
| 1 | 3 | 0.48 | 0.56693 | 0.47185 | 0.40241 | 1.07 | 0.5 | 0.56 | 0.42 | 0.49 | 0.46 | 0.4 | 0.4262 | 0.44 | 0.48475 | 0.53 | 0.5 | 0.47 | 0.45 | 0.399 | 16
| 1 | 4 | 0.48 | 0.56611 | 0.47857 | 0.40576 | 1.04 | 0.5 | 0.59 | 0.42 | 0.5 | 0.47 | 0.44 | 0.4056 | 0.43 | 0.49855 | 0.52 | 0.49 | 0.49 | 0.46 | 0.4 | 0.47371
| 1 | 5 | 0.51 | 0.56189 | 0.48116 | 0.42633 | 1.14 | 0.51 | 0.53 | 0.42 | 0.49 | 0.47 | 0.44 | 0.4125 | 0.44 | 0.49039 | 0.53 | 0.48 | 0.48 | 0.46 | 0.394 | 81
| 1 | 6 | 0.49 | 0.56193 | 0.48878 | 0.40719 | 1.13 | 0.5 | 0.57 | 0.41 | 0.48 | 0.48 | 0.49 | 0.4132 | 0.44 | 0.48781 | 0.53 | 0.47 | 0.48 | 0.45 | 0.399
| Mean |  | 0.49167 | 0.56198 | 0.48128 | 0.40917 | 1.0983 | 0.50167 | 0.57167 | 0.41667 | 0.49 | 0.47167 | 0.44167 | 0.41107 | 0.43833 | 0.49596 | 0.53167 | 0.48667 | 0.475 | 0.455 | 0.39917
| Median |  | 0.49 | 0.56394 | 0.48005 | 0.40647 | 1.105 | 0.5 | 0.58 | 0.42 | 0.49 | 0.47 | 0.44 | 0.40905 | 0.44 | 0.49447 | 0.53 | 0.485 | 0.475 | 0.455 | 0.399
| Std Dev. |  | 0.01169 | 0.0066897 | 0.0064516 | 0.0093778 | 0.038687 | 0.0040825 | 0.024014 | 0.005164 | 0.0063246 | 0.0075277 | 0.031252 | 0.0084273 | 0.0040825 | 0.0098399 | 0.0075277 | 0.012111 | 0.010488 | 0.010488 | 0.0031885
| Rel.Std.Dev. |  | 2.378% | 1.19% | 1.341% | 2.292% | 3.522% | 0.8138% | 4.201% | 1.239% | 1.291% | 1.596% | 7.076% | 2.05% | 0.9314% | 1.984% | 1.416% | 2.488% | 2.208% | 2.305% | 0.7988%
| PDM3 |  | 3.79% | 18.63% | 1.596% | -13.63% | 131.9% | 5.901% | 20.68% | -12.04% | 3.438% | -0.4322% | -6.765% | -13.22% | -7.469% | 4.695% | 12.23% | 2.734% | 0.2715% | -3.95% | -15.74%
| Z-Score (Absolute) |  | 0.18 | 1.37 | 0 | 1.23 | 10.5 | 0.35 | 1.54 | 1.1 | 0.15 | 0.16 | 0.67 | 1.2 | 0.73 | 0.25 | 0.86 | 0.09 | 0.11 | 0.45 | 1.4
| Analytical results for Cd in OREAS 255c (Certified Value 0.081 ppm) | [20 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1 | < 0.2 | < 10 | < 0.3 | 0.08 | 0.13 | 0.14 | 0.97667 | 0.07 | 0.1 | 0.08 | 0.065112 | 0.1 | 0.07 | 0.07 | 0.07 |  |  |  | 1
| 1 | 2 | 0.08 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.07 | 0.14 | 1.029 | 0.08 | 0.1 | 0.07 | 0.07365 | 0.1 | 0.07 | 0.08 | 0.07 |  |  |  | 24
| 1 | 3 | 0.08 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.07 | 0.14 | 0.95233 | 0.07 | 0.1 | 0.07 | 0.087447 | 0.1 | 0.07 | 0.08 | 0.07 |  |  |  | 16
| 1 | 4 | 0.09 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.1 | 0.14 | 1.0067 | 0.08 | 0.1 | 0.08 | 0.076596 | 0.1 | 0.07 | 0.07 | 0.07 |  |  |  | 0.081041
| 1 | 5 | 0.07 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.08 | 0.14 | 0.977 | 0.08 | 0.1 | 0.08 | 0.08747 | 0.11 | 0.07 | 0.1 | 0.07 |  |  |  | 82
| 1 | 6 | 0.08 | < 0.2 | < 10 | < 0.3 | 0.09 | 0.1 | 0.14 | 1.0857 | 0.08 | 0.1 | 0.08 | 0.078176 | 0.11 | 0.07 | 0.08 | 0.07
| Mean |  | 0.083333 |  |  |  | 0.088333 | 0.091667 | 0.14 | 1.0046 | 0.076667 | 0.1 | 0.076667 | 0.078075 | 0.10333 | 0.07 | 0.08 | 0.07
| Median |  | 0.08 |  |  |  | 0.09 | 0.09 | 0.14 | 0.99183 | 0.08 | 0.1 | 0.08 | 0.077386 | 0.1 | 0.07 | 0.08 | 0.07
| Std Dev. |  | 0.010328 |  |  |  | 0.0040825 | 0.023166 | 0 | 0.047844 | 0.005164 | 1.5202e-17 | 0.005164 | 0.0085549 | 0.005164 | 0 | 0.010954 | 0
| Rel.Std.Dev. |  | 12.39% |  |  |  | 4.622% | 25.27% | 0% | 4.763% | 6.736% | 1.52e-14% | 6.736% | 10.96% | 4.997% | 0% | 13.69% | 0%
| PDM3 |  | 2.829% |  |  |  | 8.999% | 13.11% | 72.75% | 1140% | -5.397% | 23.39% | -5.397% | -3.66% | 27.51% | -13.62% | -1.284% | -13.62%
| Z-Score (Absolute) |  | 0.29 | 0.67 | 283.88 | 3.56 | 0 | 0.19 | 2.99 | 52.96 | 0.67 | NA | 0.67 | 0.59 | 0.87 | 1.06 | 0.48 | 1.06
| Analytical results for Ce in OREAS 255c (Certified Value 43.4 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 45.785 | 41.6608 | 44.4 | 46.3 | 41.96 | 41.9 | 40.8 | 45.8463 | 41.3 | 42.8 | 47.2 | 41.2 | 41.38 | [6 empty] | 1
| 1 | 2 | 44.488 | 41.8873 | 45.3 | 46.6 | 41.85 | 43.5 | 42.5 | 46.1157 | 42 | 46.1 | 47.3 | 41.7 | 40.93 | [6 empty] | 25
| 1 | 3 | 44.096 | 41.6132 | 45.4 | 45.6 | 40.86 | 40.9 | 43 | 46.119 | 41.5 | 43.7 | 44.7 | 40.4 | 42.65 | [6 empty] | 16
| 1 | 4 | 43.324 | 40.823 | 45.4 | 45.6 | 42.53 | 41.9 | 41.5 | 45.9918 | 39.6 | 44.4 | 46.2 | 42.6 | 41.37 | [6 empty] | 43.3601
| 1 | 5 | 44.612 | 41.8441 | 47.4 | 42.7 | 41.76 | 42.2 | 42.1 | 46.2426 | 40.2 | 42.8 | 47 | 42.1 | 41.19 | [6 empty] | 83
| 1 | 6 | 44.006 | 40.9999 | 45.1 | 45.1 | 41.93 | 43.3 | 42.4 | 45.9918 | 41.3 | 43.1 | 46.7 | 42.1 | 40.29
| Mean |  | 44.3852 | 41.4714 | 45.5 | 45.3167 | 41.815 | 42.2833 | 42.05 | 46.0512 | 40.9833 | 43.8167 | 46.5167 | 41.6833 | 41.3017
| Median |  | 44.292 | 41.637 | 45.35 | 45.6 | 41.89 | 42.05 | 42.25 | 46.0537 | 41.3 | 43.4 | 46.85 | 41.9 | 41.28
| Std Dev. |  | 0.82159 | 0.44956 | 1.004 | 1.3906 | 0.54069 | 0.97245 | 0.78677 | 0.13744 | 0.89759 | 1.2766 | 0.97451 | 0.78337 | 0.77497
| Rel.Std.Dev. |  | 1.851% | 1.084% | 2.207% | 3.069% | 1.293% | 2.3% | 1.871% | 0.2985% | 2.19% | 2.913% | 2.095% | 1.879% | 1.876%
| PDM3 |  | 2.364% | -4.356% | 4.935% | 4.512% | -3.563% | -2.483% | -3.021% | 6.206% | -5.481% | 1.053% | 7.28% | -3.867% | -4.747%
| Z-Score (Absolute) |  | 1.09 | 0.42 | 1.67 | 1.57 | 0.24 | 0 | 0.12 | 1.95 | 0.67 | 0.8 | 2.2 | 0.31 | 0.51
| Analytical results for Co in OREAS 255c (Certified Value 18 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 1
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 18.2 | 18.8784 | 18.3934 | 20.685 | 44 | 18.5 | 19.2 | 17.4 | 18.2267 | 17.2 | 17 | 16.9 | 18.0705 | 19.27 | 17.6 | 17.6 | 17.1 | 17.4 |  | 1
| 1 | 2 | 17.9 | 19.7532 | 18.7165 | 26.159 | 21 | 18.2 | 19.6 | 17.4 | 18.1867 | 17.4 | 16 | 18.1 | 18.3127 | 19.76 | 18.9 | 17.6 | 17.3 | 17.2 |  | 26
| 1 | 3 | 17.8 | 19.278 | 18.101 | 20.752 | 16 | 18.3 | 19 | 17.6 | 18.2633 | 16.8 | 16.7 | 18.7 | 18.2414 | 19.22 | 17.9 | 17.4 | 16.7 | 17.9 |  | 16
| 1 | 4 | 17.6 | 19.494 | 18.2478 | 22.344 | 46 | 18 | 19.2 | 17.7 | 18.3767 | 17.2 | 17.5 | 18.2 | 18.1783 | 18.66 | 17.9 | 18.2 | 17.6 | 17.5 |  | 18.0346
| 1 | 5 | 18.1 | 19.062 | 18.2212 | 22.75 | 57 | 18.7 | 17.5 | 17.6 | 18.2067 | 17.2 | 17 | 18.2 | 18.1376 | 19.16 | 17.4 | 18.2 | 17.2 | 17.7 |  | 84
| 1 | 6 | 18 | 19.4508 | 18.5072 | 23.005 | 33 | 18.1 | 18.3 | 17.5 | 17.9533 | 18 | 17.5 | 17.9 | 18.2989 | 19.72 | 17.4 | 17.5 | 17.1 | 17.3
| Mean |  | 17.9333 | 19.3194 | 18.3645 | 22.6158 | 36.1667 | 18.3 | 18.8 | 17.5333 | 18.2022 | 17.3 | 16.95 | 18 | 18.2066 | 19.2983 | 17.85 | 17.75 | 17.1667 | 17.5
| Median |  | 17.95 | 19.3644 | 18.3206 | 22.547 | 38.5 | 18.25 | 19.1 | 17.55 | 18.2167 | 17.2 | 17 | 18.15 | 18.2099 | 19.245 | 17.75 | 17.6 | 17.15 | 17.45
| Std Dev. |  | 0.21602 | 0.31572 | 0.22301 | 2 | 15.7406 | 0.26077 | 0.76681 | 0.12111 | 0.13926 | 0.39497 | 0.56125 | 0.6 | 0.094951 | 0.40607 | 0.56125 | 0.35637 | 0.29439 | 0.26077
| Rel.Std.Dev. |  | 1.205% | 1.634% | 1.214% | 8.843% | 43.52% | 1.425% | 4.079% | 0.6907% | 0.7651% | 2.283% | 3.311% | 3.333% | 0.5215% | 2.104% | 3.144% | 2.008% | 1.715% | 1.49%
| PDM3 |  | -0.5618% | 7.124% | 1.829% | 25.4% | 100.5% | 1.471% | 4.244% | -2.78% | 0.9292% | -4.074% | -6.014% | -0.1921% | 0.9533% | 7.007% | -1.024% | -1.578% | -4.813% | -2.965%
| Z-Score (Absolute) |  | 0.19 | 1.41 | 0.3 | 5.21 | 20.84 | 0.23 | 0.81 | 0.66 | 0.12 | 0.92 | 1.33 | 0.12 | 0.12 | 1.38 | 0.29 | 0.41 | 1.08 | 0.69
| Analytical results for Cr in OREAS 255c (Certified Value 108 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 0
| 1 | 1 | 118 | 116.6076 | 109.2775 | 81.178 | 85 | 122 | 108 | 101 | 110.0462 | 106 | 103 | 100 | 112.8893 | 124 | 106 | 106 | 101 | 102 |  | 1
| 1 | 2 | 116 | 117.5904 | 111.603 | 85.372 | 88 | 124 | 111 | 91 | 109.3531 | 109 | 95 | 99 | 111.8624 | 129 | 110 | 107 | 106 | 103 |  | 27
| 1 | 3 | 114 | 116.64 | 106.5752 | 83.15 | 76 | 121 | 107 | 94 | 110.6825 | 104 | 96 | 100 | 114.3213 | 122 | 110 | 106 | 98 | 102 |  | 16
| 1 | 4 | 114 | 118.2708 | 108.2465 | 84.397 | 73 | 122 | 110 | 100 | 111.796 | 107 | 99 | 97 | 113.4225 | 125 | 108 | 112 | 104 | 103 |  | 107.5263
| 1 | 5 | 119 | 117.4932 | 106.2798 | 91.335 | 91 | 123 | 101 | 99 | 110.8072 | 107 | 100 | 99 | 113.6403 | 122 | 106 | 109 | 104 | 99 |  | 85
| 1 | 6 | 116 | 117.1584 | 108.9356 | 88.794 | 84 | 125 | 105 | 97 | 108.6273 | 108 | 109 | 100 | 111.326 | 126 | 104 | 109 | 102 | 99
| Mean |  | 116.1667 | 117.2934 | 108.4863 | 85.7043 | 82.8333 | 122.8333 | 107 | 97 | 110.2187 | 106.8333 | 100.3333 | 99.1667 | 112.9103 | 124.6667 | 107.3333 | 108.1667 | 102.5 | 101.3333
| Median |  | 116 | 117.3258 | 108.5911 | 84.8845 | 84.5 | 122.5 | 107.5 | 98 | 110.3644 | 107 | 99.5 | 99.5 | 113.1559 | 124.5 | 107 | 108 | 103 | 102
| Std Dev. |  | 2.0412 | 0.63228 | 1.9555 | 3.7439 | 6.969 | 1.472 | 3.6332 | 3.8471 | 1.1279 | 1.7224 | 5.1251 | 1.169 | 1.1309 | 2.6583 | 2.4221 | 2.3166 | 2.8107 | 1.8619
| Rel.Std.Dev. |  | 1.757% | 0.5391% | 1.803% | 4.368% | 8.413% | 1.198% | 3.395% | 3.966% | 1.023% | 1.612% | 5.108% | 1.179% | 1.002% | 2.132% | 2.257% | 2.142% | 2.742% | 1.837%
| PDM3 |  | 8.036% | 9.083% | 0.8928% | -20.29% | -22.96% | 14.24% | -0.4894% | -9.789% | 2.504% | -0.6444% | -6.689% | -7.774% | 5.007% | 15.94% | -0.1794% | 0.5956% | -4.674% | -5.759%
| Z-Score (Absolute) |  | 0.96 | 1.08 | 0.14 | 2.29 | 2.59 | 1.67 | 0.02 | 1.08 | 0.32 | 0.04 | 0.73 | 0.85 | 0.61 | 1.86 | 0.02 | 0.11 | 0.5 | 0.62
| Analytical results for Cs in OREAS 255c (Certified Value 0.99 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 1.15 | 1.0429 | 1.29 | 0.94 | 1.07 | 0.88 | 0.81 | 1.1036 | 1.45 | 0.88 | 1 | 0.84 | 0.94 | [6 empty] | 1
| 1 | 2 | 1.11 | 1.1347 | 1.31 | 1.07 | 1.05 | 0.93 | 0.85 | 1.072 | 1.47 | 0.95 | 0.98 | 0.86 | 0.93 | [6 empty] | 28
| 1 | 3 | 1.13 | 1.0042 | 1.34 | 0.96 | 1.04 | 0.85 | 0.86 | 1.0862 | 1.43 | 0.89 | 0.96 | 0.84 | 0.95 | [6 empty] | 16
| 1 | 4 | 1.09 | 1.0634 | 1.27 | 0.95 | 1.07 | 0.89 | 0.83 | 1.0988 | 1.43 | 0.9 | 0.96 | 0.87 | 0.94 | [6 empty] | 0.99336
| 1 | 5 | 1.14 | 1.0065 | 1.23 | 0.9 | 1.06 | 0.9 | 0.84 | 1.0715 | 1.41 | 0.88 | 0.95 | 0.86 | 0.94 | [6 empty] | 86
| 1 | 6 | 1.11 | 1.0453 | 1.29 | 0.96 | 1.06 | 0.93 | 0.84 | 1.065 | 1.49 | 0.88 | 0.94 | 0.86 | 0.92
| Mean |  | 1.1217 | 1.0495 | 1.2883 | 0.96333 | 1.0583 | 0.89667 | 0.83833 | 1.0829 | 1.4467 | 0.89667 | 0.965 | 0.855 | 0.93667
| Median |  | 1.12 | 1.0441 | 1.29 | 0.955 | 1.06 | 0.895 | 0.84 | 1.0791 | 1.44 | 0.885 | 0.96 | 0.86 | 0.94
| Std Dev. |  | 0.022286 | 0.047775 | 0.037103 | 0.056804 | 0.01169 | 0.030768 | 0.017224 | 0.015856 | 0.029439 | 0.027325 | 0.021679 | 0.012247 | 0.010328
| Rel.Std.Dev. |  | 1.987% | 4.552% | 2.88% | 5.897% | 1.105% | 3.431% | 2.055% | 1.464% | 2.035% | 3.047% | 2.247% | 1.432% | 1.103%
| PDM3 |  | 12.92% | 5.653% | 29.69% | -3.023% | 6.54% | -9.734% | -15.61% | 9.009% | 45.63% | -9.734% | -2.855% | -13.93% | -5.708%
| Z-Score (Absolute) |  | 1.13 | 0.61 | 2.34 | 0.01 | 0.67 | 0.49 | 0.92 | 0.85 | 3.48 | 0.49 | 0 | 0.79 | 0.2
| Analytical results for Cu in OREAS 255c (Certified Value 55 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 55.3 | 55.3932 | 54.534 | 48.57 | 58 | 58 | 57 | 56.8 | 59.1667 | 49.9 | 57.5 | 49.9628 | 55 | 51.3777 | 59.1 | 54.8 | 57.4 | 53 | 49.5 | 1
| 1 | 2 | 54.4 | 57.0024 | 56.4901 | 48.599 | 56 | 58 | 58.2 | 58.7 | 59.7167 | 50.2 | 53.7 | 50.4409 | 56.9 | 51.1144 | 59.9 | 57.6 | 57.8 | 55.3 | 48.8 | 29
| 1 | 3 | 54.5 | 55.2204 | 54.6577 | 47.813 | 53 | 59 | 58.9 | 57.9 | 59.7667 | 48.7 | 61.4 | 52.0531 | 56.4 | 52.7606 | 58.6 | 55.9 | 57.6 | 53.3 | 48.4 | 16
| 1 | 4 | 53.7 | 56.5812 | 55.0968 | 49.458 | 52 | 58 | 58.1 | 58 | 59.9667 | 49.7 | 54.9 | 50.6329 | 55.1 | 53.5601 | 57.8 | 55.5 | 58.9 | 55.9 | 49.1 | 55.0235
| 1 | 5 | 54.7 | 55.0692 | 55.0167 | 50.486 | 57 | 59 | 57.4 | 57.8 | 58.92 | 49.6 | 55.4 | 50.8538 | 55.8 | 52.7689 | 58.7 | 54.3 | 58.3 | 53.4 | 48.1 | 87
| 1 | 6 | 54.6 | 55.4796 | 55.4098 | 50.546 | 60 | 59 | 58.5 | 58.1 | 59.12 | 51.7 | 60.2 | 50.9367 | 57.1 | 50.501 | 60.1 | 54.4 | 57.5 | 53.3 | 51.3
| Mean |  | 54.5333 | 55.791 | 55.2008 | 49.2453 | 56 | 58.5 | 58.0167 | 57.8833 | 59.4428 | 49.9667 | 57.1833 | 50.8134 | 56.05 | 52.0138 | 59.0333 | 55.4167 | 57.9167 | 54.0333 | 49.2
| Median |  | 54.55 | 55.4364 | 55.0567 | 49.0285 | 56.5 | 58.5 | 58.15 | 57.95 | 59.4417 | 49.8 | 56.45 | 50.7434 | 56.1 | 52.0692 | 58.9 | 55.15 | 57.7 | 53.35 | 48.95
| Std Dev. |  | 0.5164 | 0.79921 | 0.70572 | 1.1137 | 3.0332 | 0.54772 | 0.70261 | 0.61779 | 0.42617 | 0.98725 | 3.0825 | 0.6998 | 0.8961 | 1.185 | 0.86178 | 1.2384 | 0.57764 | 1.2356 | 1.1419
| Rel.Std.Dev. |  | 0.9469% | 1.433% | 1.278% | 2.262% | 5.416% | 0.9363% | 1.211% | 1.067% | 0.7169% | 1.976% | 5.391% | 1.377% | 1.599% | 2.278% | 1.46% | 2.235% | 0.9974% | 2.287% | 2.321%
| PDM3 |  | -0.8909% | 1.395% | 0.3222% | -10.5% | 1.775% | 6.318% | 5.44% | 5.197% | 8.032% | -9.19% | 3.925% | -7.652% | 1.866% | -5.47% | 7.287% | 0.7145% | 5.258% | -1.8% | -10.58%
| Z-Score (Absolute) |  | 0.4 | 0 | 0.19 | 2.08 | 0.07 | 0.86 | 0.71 | 0.66 | 1.16 | 1.85 | 0.44 | 1.58 | 0.08 | 1.2 | 1.03 | 0.12 | 0.67 | 0.56 | 2.09
| Analytical results for Dy in OREAS 255c (Indicative Value 2.28 ppm) | [20 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 2.358; 1.02: 2.0997; 1.03: 2.3 | 1
| 1 | 2 | 1.01: 2.255; 1.02: 2.2197; 1.03: 2.4 | 3
| 1 | 3 | 1.01: 2.246; 1.02: 2.226; 1.03: 2.2 | 16
| 1 | 4 | 1.01: 2.251; 1.02: 2.1725; 1.03: 2.3 | 2.28
| 1 | 5 | 1.01: 2.276; 1.02: 2.1533; 1.03: 2.2 | 9
| 1 | 6 | 1.01: 2.274; 1.02: 2.1318; 1.03: 2.3
| Mean |  | 1.01: 2.2767; 1.02: 2.1672; 1.03: 2.2833
| Median |  | 1.01: 2.2645; 1.02: 2.1629; 1.03: 2.3
| Std Dev. |  | 1.01: 0.041693; 1.02: 0.049468; 1.03: 0.075277
| Rel.Std.Dev. |  | 1.01: 1.831%; 1.02: 2.283%; 1.03: 3.297%
| PDM3 |  | 1.01: -0.1462%; 1.02: -4.948%; 1.03: 0.1462%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 11.07; 1.03: 0.67
| Analytical results for Er in OREAS 255c (Indicative Value 0.88 ppm) | [20 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 0.936; 1.02: 0.82464; 1.03: 0.9 | 1
| 1 | 2 | 1.01: 0.927; 1.02: 0.86762; 1.03: 0.9 | 4
| 1 | 3 | 1.01: 0.922; 1.02: 0.90323; 1.03: 0.8 | 16
| 1 | 4 | 1.01: 0.913; 1.02: 0.8866; 1.03: 0.9 | 0.87966
| 1 | 5 | 1.01: 0.959; 1.02: 0.83304; 1.03: 0.8 | 10
| 1 | 6 | 1.01: 0.942; 1.02: 0.81979; 1.03: 0.8
| Mean |  | 1.01: 0.93317; 1.02: 0.85582; 1.03: 0.85
| Median |  | 1.01: 0.9315; 1.02: 0.85033; 1.03: 0.85
| Std Dev. |  | 1.01: 0.016266; 1.02: 0.034995; 1.03: 0.054772
| Rel.Std.Dev. |  | 1.01: 1.743%; 1.02: 4.089%; 1.03: 6.444%
| PDM3 |  | 1.01: 6.082%; 1.02: -2.71%; 1.03: -3.372%
| Z-Score (Absolute) |  | 1.01: 8.96; 1.02: 0; 1.03: 0.67
| Analytical results for Eu in OREAS 255c (Indicative Value 0.96 ppm) | [20 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 1.06; 1.02: 0.86944; 1.03: 1 | 1
| 1 | 2 | 1.01: 1.048; 1.02: 0.89472; 1.03: 0.9 | 5
| 1 | 3 | 1.01: 1.045; 1.02: 0.9329; 1.03: 1 | 16
| 1 | 4 | 1.01: 1.012; 1.02: 0.86658; 1.03: 0.9 | 0.95656
| 1 | 5 | 1.01: 1.085; 1.02: 0.91613; 1.03: 0.9 | 11
| 1 | 6 | 1.01: 1.017; 1.02: 0.87134; 1.03: 0.9
| Mean |  | 1.01: 1.0445; 1.02: 0.89185; 1.03: 0.93333
| Median |  | 1.01: 1.0465; 1.02: 0.88303; 1.03: 0.9
| Std Dev. |  | 1.01: 0.027223; 1.02: 0.027728; 1.03: 0.05164
| Rel.Std.Dev. |  | 1.01: 2.606%; 1.02: 3.109%; 1.03: 5.533%
| PDM3 |  | 1.01: 9.193%; 1.02: -6.765%; 1.03: -2.428%
| Z-Score (Absolute) |  | 1.01: 1.81; 1.02: 0.67; 1.03: 0
| Analytical results for Fe in OREAS 255c (Certified Value 3.49 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 3
| 1 | 1 | 3.63 | 3.8069 | 3.4769 | 3.084 | 3.52 | 3.73 | 3.87 | 3.09 | 3.52 | 3.44 | 3.43 | 3.2119 | 3.27 | 3.7808 | 3.85 | 3.45 | 3.33 | 3.35 | 3.23 | 1
| 1 | 2 | 3.59 | 3.8824 | 3.605 | 3.2005 | 3.69 | 3.79 | 3.93 | 3.02 | 3.5 | 3.5 | 3.19 | 3.2626 | 3.3 | 3.6897 | 3.83 | 3.59 | 3.36 | 3.51 | 3.19 | 0
| 1 | 3 | 3.54 | 3.8088 | 3.5157 | 3.1323 | 3.36 | 3.77 | 3.97 | 3.02 | 3.52 | 3.36 | 3.22 | 3.3676 | 3.35 | 3.6668 | 3.79 | 3.6 | 3.38 | 3.35 | 3.22 | 16
| 1 | 4 | 3.53 | 3.8866 | 3.4796 | 3.1445 | 3.21 | 3.82 | 3.94 | 3.04 | 3.55 | 3.45 | 3.27 | 3.255 | 3.25 | 3.8438 | 3.77 | 3.52 | 3.47 | 3.42 | 3.25 | 3.4883
| 1 | 5 | 3.67 | 3.8248 | 3.512 | 3.303 | 3.75 | 3.84 | 3.86 | 3.03 | 3.53 | 3.45 | 3.33 | 3.3048 | 3.29 | 3.8038 | 3.78 | 3.48 | 3.41 | 3.4 | 3.18 | 88
| 1 | 6 | 3.61 | 3.807 | 3.4997 | 3.155 | 3.56 | 3.76 | 3.89 | 3.02 | 3.48 | 3.47 | 3.58 | 3.2861 | 3.32 | 3.6454 | 3.8 | 3.44 | 3.43 | 3.37 | 3.05
| Mean |  | 3.595 | 3.8361 | 3.5148 | 3.1699 | 3.515 | 3.785 | 3.91 | 3.0367 | 3.5167 | 3.445 | 3.3367 | 3.2813 | 3.2967 | 3.7384 | 3.8033 | 3.5133 | 3.3967 | 3.4 | 3.1867
| Median |  | 3.6 | 3.8168 | 3.5058 | 3.1497 | 3.54 | 3.78 | 3.91 | 3.025 | 3.52 | 3.45 | 3.3 | 3.2744 | 3.295 | 3.7352 | 3.795 | 3.5 | 3.395 | 3.385 | 3.205
| Std Dev. |  | 0.053572 | 0.038112 | 0.047027 | 0.075237 | 0.20246 | 0.040373 | 0.043359 | 0.027325 | 0.024221 | 0.046797 | 0.14665 | 0.052684 | 0.03559 | 0.081659 | 0.030768 | 0.069186 | 0.050465 | 0.060663 | 0.07174
| Rel.Std.Dev. |  | 1.49% | 0.9935% | 1.338% | 2.374% | 5.76% | 1.067% | 1.109% | 0.8998% | 0.6888% | 1.358% | 4.395% | 1.606% | 1.08% | 2.184% | 0.809% | 1.969% | 1.486% | 1.784% | 2.251%
| PDM3 |  | 3.058% | 9.97% | 0.7598% | -9.129% | 0.7651% | 8.505% | 12.09% | -12.95% | 0.8129% | -1.242% | -4.347% | -5.933% | -5.494% | 7.169% | 9.031% | 0.7173% | -2.627% | -2.532% | -8.647%
| Z-Score (Absolute) |  | 0.25 | 1 | 0 | 1.07 | 0.01 | 0.85 | 1.23 | 1.48 | 0.01 | 0.21 | 0.55 | 0.72 | 0.67 | 0.7 | 0.9 | 0 | 0.36 | 0.35 | 1.02
| Analytical results for Ga in OREAS 255c (Certified Value 6.13 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 6.28 | 6.3054 | 6.8 | 6.73 | 6.27 | 6.4612 | 5.87 | 5 | 5.15 | 6.6076 | 7.11 | 5.71 | 6.49 | 5.22 | [5 empty] | 1
| 1 | 2 | 6.03 | 6.1782 | 7.1 | 7.19 | 6.17 | 6.444 | 5.96 | 5 | 5.36 | 6.2353 | 7.31 | 6.33 | 6.58 | 5.42 | [5 empty] | 33
| 1 | 3 | 6.23 | 6.0326 | 7 | 6.69 | 6.23 | 6.4723 | 5.63 | 5 | 5.49 | 6.6505 | 6.88 | 5.79 | 6.55 | 5.2 | [5 empty] | 16
| 1 | 4 | 6.1 | 5.6901 | 6.9 | 7.1 | 6.37 | 6.5676 | 5.8 | 5 | 5.18 | 6.347 | 6.8 | 5.83 | 6.42 | 5.43 | [5 empty] | 6.1299
| 1 | 5 | 6.32 | 6.01 | 6.8 | 6.24 | 6.26 | 6.5551 | 5.79 | 5 | 5.32 | 6.3734 | 6.91 | 5.62 | 6.6 | 5.36 | [5 empty] | 89
| 1 | 6 | 6.07 | 5.8534 | 7 | 6.54 | 6.23 | 6.43 | 6.08 | 5 | 5.52 | 6.2116 | 7.38 | 5.63 | 6.43 | 5.3
| Mean |  | 6.1717 | 6.0116 | 6.9333 | 6.7483 | 6.255 | 6.4884 | 5.855 | 5 | 5.3367 | 6.4042 | 7.065 | 5.8183 | 6.5117 | 5.3217
| Median |  | 6.165 | 6.0213 | 6.95 | 6.71 | 6.245 | 6.4668 | 5.835 | 5 | 5.34 | 6.3602 | 7.01 | 5.75 | 6.52 | 5.33
| Std Dev. |  | 0.12057 | 0.22044 | 0.12111 | 0.35335 | 0.066257 | 0.058484 | 0.15476 | 0 | 0.15319 | 0.1854 | 0.24073 | 0.2643 | 0.07679 | 0.098472
| Rel.Std.Dev. |  | 1.954% | 3.667% | 1.747% | 5.236% | 1.059% | 0.9014% | 2.643% | 0% | 2.87% | 2.895% | 3.407% | 4.543% | 1.179% | 1.85%
| PDM3 |  | 0.6814% | -1.93% | 13.11% | 10.09% | 2.041% | 5.848% | -4.485% | -18.43% | -12.94% | 4.475% | 15.25% | -5.083% | 6.228% | -13.19%
| Z-Score (Absolute) |  | 0.07 | 0.36 | 1.29 | 0.96 | 0.07 | 0.49 | 0.64 | 2.17 | 1.57 | 0.34 | 1.52 | 0.71 | 0.53 | 1.6
| Analytical results for Gd in OREAS 255c (Indicative Value 3.39 ppm) | [20 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 3.482; 1.02: 3.3148; 1.03: 3.3 | 1
| 1 | 2 | 1.01: 3.304; 1.02: 3.5521; 1.03: 3.4 | 6
| 1 | 3 | 1.01: 3.339; 1.02: 3.501; 1.03: 3.6 | 16
| 1 | 4 | 1.01: 3.318; 1.02: 3.4749; 1.03: 3.3 | 3.395
| 1 | 5 | 1.01: 3.386; 1.02: 3.4698; 1.03: 3.1 | 12
| 1 | 6 | 1.01: 3.328; 1.02: 3.4398; 1.03: 3.5
| Mean |  | 1.01: 3.3595; 1.02: 3.4587; 1.03: 3.3667
| Median |  | 1.01: 3.3335; 1.02: 3.4724; 1.03: 3.35
| Std Dev. |  | 1.01: 0.066217; 1.02: 0.079953; 1.03: 0.17512
| Rel.Std.Dev. |  | 1.01: 1.971%; 1.02: 2.312%; 1.03: 5.202%
| PDM3 |  | 1.01: -1.045%; 1.02: 1.878%; 1.03: -0.8336%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 8.66; 1.03: 0
| Analytical results for Ge in OREAS 255c (Certified Value 0.12 ppm) | [20 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | 1
| No | No | 1.01: 1.03; 1.02: 1.09; 1.03: 1.10; 1.05: 1.12; 1.06: 1.15; 1.07: 1.19; 1.09: 1.20; 1.10: 1.21; 1.11: 1.22 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*OES/MS; 1.03: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*MS; 1.09: AR*MS; 1.10: AR*MS; 1.11: AR*MS | 2
|  |  | 1.01: 0.2g; 1.02: 0.5g; 1.03: 0.2g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.25g | 2
| 1 | 1 | 1.01: < 0.2; 1.02: < 0.1; 1.03: 0.14; 1.05: 0.13; 1.06: 0.1; 1.07: 0.09; 1.09: 0.15; 1.10: 0.08; 1.11: 0.2 | 1
| 1 | 2 | 1.01: < 0.2; 1.02: < 0.1; 1.03: 0.15; 1.05: 0.13; 1.06: 0.1; 1.07: 0.09; 1.09: 0.14; 1.10: 0.09; 1.11: 0.1 | 1
| 1 | 3 | 1.01: < 0.2; 1.02: < 0.1; 1.03: 0.16; 1.05: 0.13; 1.06: 0.1; 1.07: 0.08; 1.09: 0.12; 1.10: 0.08; 1.11: 0.2 | 16
| 1 | 4 | 1.01: < 0.2; 1.02: < 0.1; 1.03: 0.15; 1.05: 0.14; 1.06: 0.1; 1.07: 0.09; 1.09: 0.15; 1.10: 0.08; 1.11: 0.1 | 0.11583
| 1 | 5 | 1.01: < 0.2; 1.02: < 0.1; 1.03: 0.15; 1.05: 0.14; 1.06: 0.09; 1.07: 0.09; 1.09: 0.14; 1.10: 0.08; 1.11: 0.1 | 90
| 1 | 6 | 1.01: < 0.2; 1.02: < 0.1; 1.03: 0.16; 1.05: 0.15; 1.06: 0.1; 1.07: 0.09; 1.09: 0.12; 1.10: 0.09; 1.11: 0.1
| Mean |  | 1.03: 0.15167; 1.05: 0.13667; 1.06: 0.098333; 1.07: 0.088333; 1.09: 0.13667; 1.10: 0.083333; 1.11: 0.13333
| Median |  | 1.03: 0.15; 1.05: 0.135; 1.06: 0.1; 1.07: 0.09; 1.09: 0.14; 1.10: 0.08; 1.11: 0.1
| Std Dev. |  | 1.03: 0.0075277; 1.05: 0.008165; 1.06: 0.0040825; 1.07: 0.0040825; 1.09: 0.013663; 1.10: 0.005164; 1.11: 0.05164
| Rel.Std.Dev. |  | 1.03: 4.963%; 1.05: 5.974%; 1.06: 4.152%; 1.07: 4.622%; 1.09: 9.997%; 1.10: 6.197%; 1.11: 38.73%
| PDM3 |  | 1.03: 30.94%; 1.05: 17.99%; 1.06: -15.11%; 1.07: -23.74%; 1.09: 17.99%; 1.10: -28.06%; 1.11: 15.11%
| Z-Score (Absolute) |  | 1.01: 0.02; 1.02: 1.24; 1.03: 1.33; 1.05: 0.95; 1.06: 0.02; 1.07: 0.27; 1.09: 0.95; 1.10: 0.4; 1.11: NA
Results from laboratory 1.22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 255c (Indicative Value 0.62 ppm) | [20 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 0.11 | 0.72354 | 1.01 | < 0.1 | 0.52 | 0.54 | 0.8 | 0.89479 | 0.8 | 0.74 | 0.64 | 0.45 | 0.97 | [6 empty] | 1
| 1 | 2 | 0.09 | 0.80252 | 1.01 | < 0.1 | 0.54 | 0.63 | 0.74 | 0.90125 | 0.87 | 0.63 | 0.54 | 0.36 | 0.98 | [6 empty] | 7
| 1 | 3 | 0.07 | 0.80894 | 1 | < 0.1 | 0.51 | 0.69 | 0.63 | 0.9086 | 0.82 | 0.52 | 0.54 | 0.35 | 0.93 | [6 empty] | 16
| 1 | 4 | 0.12 | 0.77785 | 1.02 | < 0.1 | 0.53 | 0.69 | 0.75 | 0.87146 | 0.81 | 0.59 | 0.6 | 0.45 | 0.81 | [6 empty] | 0.62079
| 1 | 5 | 0.11 | 0.77634 | 0.97 | 0.1 | 0.49 | 0.72 | 0.65 | 0.89067 | 0.81 | 0.7 | 0.51 | 0.3 | 0.97 | [6 empty] | 13
| 1 | 6 | 0.1 | 0.7591 | 0.99 | < 0.1 | 0.5 | 0.6 | 0.58 | 0.86635 | 0.86 | 0.61 | 0.57 | 0.53 | 0.87
| Mean |  | 0.1 | 0.77472 | 1 | 0.1 | 0.515 | 0.645 | 0.69167 | 0.88885 | 0.82833 | 0.63167 | 0.56667 | 0.40667 | 0.92167
| Median |  | 0.105 | 0.77709 | 1.005 | 0.1 | 0.515 | 0.66 | 0.695 | 0.89273 | 0.815 | 0.62 | 0.555 | 0.405 | 0.95
| Std Dev. |  | 0.017889 | 0.031045 | 0.017889 |  | 0.018708 | 0.06775 | 0.084242 | 0.016681 | 0.029269 | 0.078846 | 0.047188 | 0.084538 | 0.06824
| Rel.Std.Dev. |  | 17.89% | 4.007% | 1.789% |  | 3.633% | 10.5% | 12.18% | 1.877% | 3.533% | 12.48% | 8.327% | 20.79% | 7.404%
| PDM3 |  | -83.89% | 24.8% | 61.09% | -83.89% | -17.04% | 3.9% | 11.42% | 43.18% | 33.43% | 1.753% | -8.718% | -34.49% | 48.47%
| Z-Score (Absolute) |  | 2 | 0.48 | 1.31 | 2.16 | 0.48 | 0 | 0.17 | 0.9 | 0.67 | 0.05 | 0.29 | 0.88 | 1.02
| Analytical results for Hg in OREAS 255c (Indicative Value 0.089 ppm) | [20 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [5 empty] | 3
|  |  | 25g | 0.2g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 0.1 | < 1 | 0.09 | < 2 | 0.1 | 0.13 | 0.179 | 0.08 | < 0.05 | 1 | 0.08 | 0.04 | 0.09 | 0.41 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 1 | 0.1 | < 2 | 0.11 | 0.1 | 0.194 | 0.07 | < 0.05 | 1 | 0.09 | 0.09 | 0.08 | 0.41 | [5 empty] | 8
| 1 | 3 | < 0.1 | < 1 | 0.06 | < 2 | 0.1 | 0.28 | 0.188 | 0.08 | < 0.05 | 1 | 0.09 | 0.1 | 0.09 | 0.42 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 1 | 0.1 | < 2 | 0.11 | 0.1 | 0.182 | 0.08 | < 0.05 | 1 | 0.07 | 0.05 | 0.08 | 0.4 | [5 empty] | 0.088875
| 1 | 5 | < 0.1 | < 1 | 0.06 | < 2 | 0.09 | 0.09 | 0.186 | 0.09 | < 0.05 | 1 | 0.09 | 0.07 | 0.07 | 0.39 | [5 empty] | 14
| 1 | 6 | < 0.1 | < 1 | 0.12 | < 2 | 0.1 | 0.11 | 0.186 | 0.08 | < 0.05 | 1 | 0.09 | 0.05 | 0.09 | 0.39
| Mean |  | 1.01: 0.1; 1.03: 0.088333; 1.06: 0.10167; 1.07: 0.135; 1.09: 0.18583; 1.10: 0.08; 1.12: 1; 1.13: 0.085; 1.14: 0.066667; 1.15: 0.083333; 1.16: 0.40333
| Median |  | 1.01: 0.1; 1.03: 0.095; 1.06: 0.1; 1.07: 0.105; 1.09: 0.186; 1.10: 0.08; 1.12: 1; 1.13: 0.09; 1.14: 0.06; 1.15: 0.085; 1.16: 0.405
| Std Dev. |  | 1.03: 0.024014; 1.06: 0.0075277; 1.07: 0.072319; 1.09: 0.0051543; 1.10: 0.0063246; 1.12: 0; 1.13: 0.0083666; 1.14: 0.024221; 1.15: 0.008165; 1.16: 0.012111
| Rel.Std.Dev. |  | 1.03: 27.19%; 1.06: 7.404%; 1.07: 53.57%; 1.09: 2.774%; 1.10: 7.906%; 1.12: 0%; 1.13: 9.843%; 1.14: 36.33%; 1.15: 9.798%; 1.16: 3.003%
| PDM3 |  | 1.01: 12.52%; 1.03: -0.6095%; 1.06: 14.39%; 1.07: 51.9%; 1.09: 109.1%; 1.10: -9.986%; 1.12: 1025%; 1.13: -4.36%; 1.14: -24.99%; 1.15: -6.235%; 1.16: 353.8%
| Z-Score (Absolute) |  | 0.67 | 9.25 | 0 | 20.49 | 0.3 | 1.05 | 2.19 | 0.19 | 1.42 | NA | 0.07 | 0.49 | 0.11 | 7.08
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 255c (Indicative Value 0.36 ppm) | [20 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS | 2
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g | 2
| 1 | 1 | 1.01: 0.395; 1.02: 0.33562; 1.03: 0.4 | 1
| 1 | 2 | 1.01: 0.377; 1.02: 0.35346; 1.03: 0.4 | 3
| 1 | 3 | 1.01: 0.368; 1.02: 0.35616; 1.03: 0.4 | 16
| 1 | 4 | 1.01: 0.364; 1.02: 0.33384; 1.03: 0.4 | 0.36195
| 1 | 5 | 1.01: 0.37; 1.02: 0.35336; 1.03: 0.3 | 9
| 1 | 6 | 1.01: 0.357; 1.02: 0.35172; 1.03: 0.3
| Mean |  | 1.01: 0.37183; 1.02: 0.34736; 1.03: 0.36667
| Median |  | 1.01: 0.369; 1.02: 0.35254; 1.03: 0.4
| Std Dev. |  | 1.01: 0.013136; 1.02: 0.0099028; 1.03: 0.05164
| Rel.Std.Dev. |  | 1.01: 3.533%; 1.02: 2.851%; 1.03: 14.08%
| PDM3 |  | 1.01: 2.73%; 1.02: -4.032%; 1.03: 1.302%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 2.52; 1.03: 0
| Analytical results for In in OREAS 255c (Certified Value 0.023 ppm) | [20 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.02 | < 0.05 | 0.03 | < 5 | 0.021 | 0.021 | < 0.05 | 0.1 | 0.022 | 0.023 | 0.022 | 0.02 | [7 empty] | 1
| 1 | 2 | 0.02 | < 0.05 | 0.03 | < 5 | 0.023 | 0.02 | < 0.05 | 0.1 | 0.024 | 0.026 | 0.022 | 0.019 | [7 empty] | 21
| 1 | 3 | 0.02 | < 0.05 | 0.03 | < 5 | 0.022 | 0.023 | < 0.05 | 0.1 | 0.023 | 0.025 | 0.023 | 0.019 | [7 empty] | 16
| 1 | 4 | 0.03 | < 0.05 | 0.03 | < 5 | 0.023 | 0.026 | < 0.05 | 0.1 | 0.023 | 0.026 | 0.023 | 0.02 | [7 empty] | 0.0225
| 1 | 5 | 0.02 | < 0.05 | 0.02 | < 5 | 0.021 | 0.026 | < 0.05 | 0.1 | 0.022 | 0.027 | 0.023 | 0.02 | [7 empty] | 91
| 1 | 6 | 0.02 | < 0.05 | 0.03 | < 5 | 0.023 | 0.021 | < 0.05 | 0.1 | 0.022 | 0.022 | 0.022 | 0.021
| Mean |  | 1.01: 0.021667; 1.03: 0.028333; 1.06: 0.022167; 1.07: 0.022833; 1.10: 0.1; 1.11: 0.022667; 1.12: 0.024833; 1.13: 0.0225; 1.14: 0.019833
| Median |  | 1.01: 0.02; 1.03: 0.03; 1.06: 0.0225; 1.07: 0.022; 1.10: 0.1; 1.11: 0.0225; 1.12: 0.0255; 1.13: 0.0225; 1.14: 0.02
| Std Dev. |  | 1.01: 0.0040825; 1.03: 0.0040825; 1.06: 0.00098319; 1.07: 0.0026394; 1.10: 1.5202e-17; 1.11: 0.0008165; 1.12: 0.0019408; 1.13: 0.00054772; 1.14: 0.00075277
| Rel.Std.Dev. |  | 1.01: 18.84%; 1.03: 14.41%; 1.06: 4.435%; 1.07: 11.56%; 1.10: 1.52e-14%; 1.11: 3.602%; 1.12: 7.815%; 1.13: 2.434%; 1.14: 3.795%
| PDM3 |  | 1.01: -3.704%; 1.03: 25.93%; 1.06: -1.481%; 1.07: 1.481%; 1.10: 344.4%; 1.11: 0.7407%; 1.12: 10.37%; 1.13: -3.331e-14%; 1.14: -11.85%
| Z-Score (Absolute) |  | 0.39 | 0.73 | 1.85 | 835.19 | 0.22 | 0 | 0.73 | NA | 0.06 | 0.67 | 0.11 | 1.01
Results from laboratory 1.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 255c (Certified Value 0.286 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.2937 | 0.28688 | 0.23018 | 0.35 | 0.324 | 0.29 | 0.23 | 0.3 | 0.29 | 0.27 | 0.26 | 0.31905 | 0.36 | 0.28 | 0.27 | 0.26 | 0.236 |  |  | 1
| 1 | 2 | 0.2868 | 0.29515 | 0.23978 | 0.36 | 0.329 | 0.29 | 0.23 | 0.3 | 0.29 | 0.24 | 0.27 | 0.30284 | 0.36 | 0.3 | 0.27 | 0.28 | 0.236 |  |  | 0
| 1 | 3 | 0.2824 | 0.28808 | 0.23404 | 0.33 | 0.331 | 0.29 | 0.23 | 0.3 | 0.28 | 0.26 | 0.27 | 0.30564 | 0.35 | 0.3 | 0.28 | 0.26 | 0.239 |  |  | 16
| 1 | 4 | 0.2829 | 0.28477 | 0.23299 | 0.32 | 0.331 | 0.3 | 0.23 | 0.31 | 0.29 | 0.26 | 0.26 | 0.3202 | 0.35 | 0.29 | 0.28 | 0.27 | 0.24 |  |  | 0.28601
| 1 | 5 | 0.2963 | 0.28754 | 0.2496 | 0.37 | 0.327 | 0.29 | 0.22 | 0.3 | 0.29 | 0.26 | 0.26 | 0.31077 | 0.35 | 0.29 | 0.28 | 0.27 | 0.234 |  |  | 92
| 1 | 6 | 0.2867 | 0.29043 | 0.23765 | 0.35 | 0.329 | 0.29 | 0.22 | 0.3 | 0.29 | 0.29 | 0.27 | 0.29566 | 0.36 | 0.28 | 0.28 | 0.27 | 0.257
| Mean |  | 0.28813 | 0.28881 | 0.23737 | 0.34667 | 0.3285 | 0.29167 | 0.22667 | 0.30167 | 0.28833 | 0.26333 | 0.265 | 0.30903 | 0.355 | 0.29 | 0.27667 | 0.26833 | 0.24033
| Median |  | 0.28675 | 0.28781 | 0.23585 | 0.35 | 0.329 | 0.29 | 0.23 | 0.3 | 0.29 | 0.26 | 0.265 | 0.30821 | 0.355 | 0.29 | 0.28 | 0.27 | 0.2375
| Std Dev. |  | 0.0056881 | 0.0036038 | 0.0068908 | 0.018619 | 0.0026646 | 0.0040825 | 0.005164 | 0.0040825 | 0.0040825 | 0.01633 | 0.0054772 | 0.0095575 | 0.0054772 | 0.0089443 | 0.005164 | 0.0075277 | 0.0084538
| Rel.Std.Dev. |  | 1.974% | 1.248% | 2.903% | 5.371% | 0.8111% | 1.4% | 2.278% | 1.353% | 1.416% | 6.201% | 2.067% | 3.093% | 1.543% | 3.084% | 1.866% | 2.805% | 3.518%
| PDM3 |  | 0.7423% | 0.9785% | -17.01% | 21.21% | 14.86% | 1.978% | -20.75% | 5.474% | 0.8122% | -7.929% | -7.346% | 8.047% | 24.12% | 1.395% | -3.267% | -6.181% | -15.97%
| Z-Score (Absolute) |  | 0.01 | 0.02 | 1.66 | 1.9 | 1.31 | 0.11 | 2.01 | 0.43 | 0 | 0.81 | 0.76 | 0.67 | 2.17 | 0.05 | 0.38 | 0.65 | 1.56
| Analytical results for La in OREAS 255c (Certified Value 23.7 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 23.799 | 22.7005 | 24.3 | 25.4 | 22.8 | 23.2633 | 23.6 | 24 | 23.2 | 25.3216 | 19.5 | 23.6 | 26.3 | 23.1 | 22.5 |  |  |  |  | 1
| 1 | 2 | 22.887 | 22.9229 | 24.9 | 26.3 | 22.6 | 22.9533 | 24.2 | 21 | 24.4 | 24.2885 | 19.8 | 25.5 | 26.3 | 23.8 | 22.4 |  |  |  |  | 22
| 1 | 3 | 22.78 | 23.005 | 24.9 | 25.6 | 23.5 | 23.1967 | 23.3 | 23 | 25.1 | 25.1959 | 19 | 24.2 | 25.1 | 22.8 | 23.1 |  |  |  |  | 16
| 1 | 4 | 22.918 | 22.4813 | 25.1 | 25.4 | 23.3 | 23.4767 | 23.6 | 22 | 23.4 | 24.9345 | 18.7 | 24.5 | 25.9 | 23.9 | 22.5 |  |  |  |  | 23.7222
| 1 | 5 | 23.679 | 23.3398 | 25.9 | 23.8 | 22.9 | 23.2067 | 23.9 | 23 | 23.4 | 24.9031 | 18.9 | 23.6 | 26.3 | 23.4 | 22.3 |  |  |  |  | 93
| 1 | 6 | 22.952 | 22.8129 | 24.7 | 24.8 | 22.9 | 22.9733 | 24.8 | 25 | 24.5 | 24.2381 | 19.4 | 23.7 | 26.2 | 23.6 | 23.1
| Mean |  | 23.1692 | 22.8771 | 24.9667 | 25.2167 | 23 | 23.1783 | 23.9 | 23 | 24 | 24.8136 | 19.2167 | 24.1833 | 26.0167 | 23.4333 | 22.65
| Median |  | 22.935 | 22.8679 | 24.9 | 25.4 | 22.9 | 23.2017 | 23.75 | 23 | 23.9 | 24.9188 | 19.2 | 23.95 | 26.25 | 23.5 | 22.5
| Std Dev. |  | 0.44675 | 0.29138 | 0.53166 | 0.84479 | 0.33466 | 0.19494 | 0.53666 | 1.4142 | 0.77201 | 0.4547 | 0.41673 | 0.7414 | 0.47504 | 0.42269 | 0.35637
| Rel.Std.Dev. |  | 1.928% | 1.274% | 2.129% | 3.35% | 1.455% | 0.8411% | 2.245% | 6.149% | 3.217% | 1.832% | 2.169% | 3.066% | 1.826% | 1.804% | 1.573%
| PDM3 |  | -2.331% | -3.563% | 5.246% | 6.3% | -3.044% | -2.293% | 0.7496% | -3.044% | 1.171% | 4.601% | -18.99% | 1.944% | 9.672% | -1.218% | -4.52%
| Z-Score (Absolute) |  | 0.31 | 0.66 | 1.82 | 2.12 | 0.52 | 0.3 | 0.56 | 0.52 | 0.67 | 1.64 | 5.02 | 0.89 | 3.07 | 0 | 0.93
| Analytical results for Li in OREAS 255c (Certified Value 10.4 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 10.8 | 10.7371 | 9 | 10.9 | 11.1 | 9.3 | 10.7977 | 10 | 9.8 | 8.9463 | 12.2 | 10.6 | 10.2 | 10 | 12 |  |  |  |  | 1
| 1 | 2 | 10.5 | 10.8547 | 10 | 10.9 | 11.4 | 9.3 | 10.6657 | 10.1 | 10 | 8.2546 | 13 | 11 | 10.3 | 10.7 | 12 |  |  |  |  | 0
| 1 | 3 | 10.6 | 10.7663 | 9 | 10.9 | 10.9 | 9.4 | 10.8862 | 9.8 | 10 | 8.9463 | 12.1 | 11 | 10.3 | 10.1 | 12 |  |  |  |  | 16
| 1 | 4 | 10.4 | 10.7505 | 9 | 11.1 | 11 | 9.5 | 11.0963 | 10 | 9.6 | 7.8253 | 11.6 | 10.8 | 10.5 | 10.3 | 12 |  |  |  |  | 10.4311
| 1 | 5 | 10.5 | 10.643 | 11 | 11.2 | 10.2 | 9.4 | 10.6043 | 10 | 9.8 | 8.4188 | 12.4 | 10.6 | 10.3 | 10.3 | 12 |  |  |  |  | 94
| 1 | 6 | 10.6 | 10.9144 | 10 | 11.4 | 10.9 | 9.4 | 10.6393 | 10.2 | 9.9 | 8.0919 | 12.8 | 10.6 | 10.4 | 10.3 | 12
| Mean |  | 10.5667 | 10.7777 | 9.6667 | 11.0667 | 10.9167 | 9.3833 | 10.7816 | 10.0167 | 9.85 | 8.4138 | 12.35 | 10.7667 | 10.3333 | 10.2833 | 12
| Median |  | 10.55 | 10.7584 | 9.5 | 11 | 10.95 | 9.4 | 10.7317 | 10 | 9.85 | 8.3367 | 12.3 | 10.7 | 10.3 | 10.3 | 12
| Std Dev. |  | 0.13663 | 0.095163 | 0.8165 | 0.20656 | 0.39707 | 0.075277 | 0.18727 | 0.13292 | 0.15166 | 0.45658 | 0.50498 | 0.19664 | 0.10328 | 0.24014 | 0
| Rel.Std.Dev. |  | 1.293% | 0.883% | 8.447% | 1.866% | 3.637% | 0.8022% | 1.737% | 1.327% | 1.54% | 5.427% | 4.089% | 1.826% | 0.9995% | 2.335% | 0%
| PDM3 |  | 1.299% | 3.322% | -7.329% | 6.093% | 4.655% | -10.04% | 3.359% | -3.973% | -5.571% | -19.34% | 18.4% | 3.217% | -0.9377% | -1.417% | 15.04%
| Z-Score (Absolute) |  | 0 | 0.41 | NA | 0.96 | 0.67 | 2.28 | 0.41 | 1.06 | 1.38 | 4.15 | 3.44 | 0.39 | 0.45 | 0.55 | NA
Results from laboratories 1.06 and 1.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 255c (Indicative Value 0.097 ppm) | [20 empty] | Indicative
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.05: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.09; 1.05: 1.22 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*OES/MS; 1.05: AR*MS | 3
|  |  | 1.01: 25g; 1.02: 0.2g; 1.03: 0.5g; 1.05: 0.25g | 3
| 1 | 1 | 1.01: 0.101; 1.02: 0.089262; 1.03: 0.1; 1.05: 0.1 | 1
| 1 | 2 | 1.01: 0.09; 1.02: 0.10043; 1.03: < 0.1; 1.05: 0.1 | 4
| 1 | 3 | 1.01: 0.095; 1.02: 0.095721; 1.03: < 0.1; 1.05: 0.1 | 16
| 1 | 4 | 1.01: 0.095; 1.02: 0.096705; 1.03: < 0.1; 1.05: 0.1 | 0.097283
| 1 | 5 | 1.01: 0.097; 1.02: 0.10273; 1.03: < 0.1; 1.05: 0.1 | 10
| 1 | 6 | 1.01: 0.092; 1.02: 0.096252; 1.03: < 0.1; 1.05: 0.1
| Mean |  | 1.01: 0.095; 1.02: 0.096849; 1.03: 0.1; 1.05: 0.1
| Median |  | 1.01: 0.095; 1.02: 0.096479; 1.03: 0.1; 1.05: 0.1
| Std Dev. |  | 1.01: 0.0038471; 1.02: 0.004617; 1.05: 1.5202e-17
| Rel.Std.Dev. |  | 1.01: 4.05%; 1.02: 4.767%; 1.05: 1.52e-14%
| PDM3 |  | 1.01: -2.347%; 1.02: -0.446%; 1.03: 2.793%; 1.05: 2.793%
| Z-Score (Absolute) |  | 1.01: 0.25; 1.02: 0.25; 1.03: 10.14; 1.05: 1.1
| Analytical results for Mg in OREAS 255c (Certified Value 1.16 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 2
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 1.19 | 1.3373 | 1.1919 | 0.99954 | 1.39 | 1.2 | 1.37 | 1 | 1.17 | 1.14 | 1.19 | 1.09 | 1.1753 | 1.22 | 1.15 | 1.13 | 1.09 | 1.115 |  | 1
| 1 | 2 | 1.17 | 1.3726 | 1.2207 | 1.0351 | 1.43 | 1.21 | 1.36 | 1.01 | 1.17 | 1.16 | 1.11 | 1.09 | 1.1828 | 1.22 | 1.2 | 1.14 | 1.14 | 1.109 |  | 0
| 1 | 3 | 1.15 | 1.3578 | 1.1666 | 1.0121 | 1.32 | 1.22 | 1.35 | 1 | 1.17 | 1.12 | 1.26 | 1.1 | 1.2172 | 1.2 | 1.21 | 1.15 | 1.09 | 1.117 |  | 16
| 1 | 4 | 1.15 | 1.381 | 1.1947 | 1.0224 | 1.28 | 1.21 | 1.38 | 0.98 | 1.2 | 1.14 | 1.14 | 1.07 | 1.2294 | 1.19 | 1.17 | 1.19 | 1.12 | 1.12 |  | 1.1627
| 1 | 5 | 1.19 | 1.3652 | 1.1665 | 1.0719 | 1.45 | 1.21 | 1.23 | 0.98 | 1.15 | 1.14 | 1.15 | 1.09 | 1.2173 | 1.2 | 1.15 | 1.16 | 1.11 | 1.101 |  | 95
| 1 | 6 | 1.18 | 1.343 | 1.2027 | 1.026 | 1.41 | 1.21 | 1.27 | 0.97 | 1.18 | 1.15 | 1.26 | 1.1 | 1.2013 | 1.21 | 1.14 | 1.16 | 1.09 | 1.102
| Mean |  | 1.1717 | 1.3595 | 1.1905 | 1.0278 | 1.38 | 1.21 | 1.3267 | 0.99 | 1.1733 | 1.1417 | 1.185 | 1.09 | 1.2039 | 1.2067 | 1.17 | 1.155 | 1.1067 | 1.1107
| Median |  | 1.175 | 1.3615 | 1.1933 | 1.0242 | 1.4 | 1.21 | 1.355 | 0.99 | 1.17 | 1.14 | 1.17 | 1.09 | 1.2092 | 1.205 | 1.16 | 1.155 | 1.1 | 1.112
| Std Dev. |  | 0.018348 | 0.01695 | 0.021105 | 0.0248 | 0.066332 | 0.0063246 | 0.061536 | 0.015492 | 0.01633 | 0.013292 | 0.063482 | 0.010954 | 0.021334 | 0.012111 | 0.028983 | 0.020736 | 0.020656 | 0.0079666
| Rel.Std.Dev. |  | 1.566% | 1.247% | 1.773% | 2.413% | 4.807% | 0.5227% | 4.638% | 1.565% | 1.392% | 1.164% | 5.357% | 1.005% | 1.772% | 1.004% | 2.477% | 1.795% | 1.866% | 0.7173%
| PDM3 |  | 0.7713% | 16.92% | 2.392% | -11.6% | 18.69% | 4.068% | 14.1% | -14.85% | 0.9147% | -1.809% | 1.918% | -6.253% | 3.544% | 3.782% | 0.628% | -0.6621% | -4.819% | -4.475%
| Z-Score (Absolute) |  | 0.02 | 3.52 | 0.34 | 2.72 | 3.9 | 0.71 | 2.9 | 3.43 | 0.02 | 0.58 | 0.24 | 1.55 | 0.59 | 0.64 | 0.05 | 0.33 | 1.24 | 1.16
| Analytical results for Mn in OREAS 255c (Certified Value 0.032 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.0324 | 0.032432 | 0.032035 | 0.027588 | 0.02 | 0.0331 | 0.0333 | 0.0282 | 0.0344 | 0.0319 | 0.0313 | 0.0297 | 0.032869 | 0.0342 | 0.0314 | 0.0306 | 0.0296 | 0.029 |  | 1
| 1 | 2 | 0.032 | 0.033354 | 0.032886 | 0.028037 | 0.03 | 0.0332 | 0.0344 | 0.0284 | 0.0343 | 0.0324 | 0.0291 | 0.0295 | 0.032205 | 0.0358 | 0.0328 | 0.031 | 0.0306 | 0.029 |  | 24
| 1 | 3 | 0.0316 | 0.032945 | 0.031955 | 0.027269 | 0.02 | 0.0334 | 0.0334 | 0.0285 | 0.0345 | 0.0308 | 0.0321 | 0.03 | 0.032843 | 0.0346 | 0.033 | 0.0309 | 0.0296 | 0.0293 |  | 16
| 1 | 4 | 0.0315 | 0.033032 | 0.032021 | 0.027678 | 0.02 | 0.033 | 0.0341 | 0.0287 | 0.0347 | 0.0324 | 0.03 | 0.0292 | 0.032377 | 0.0331 | 0.0322 | 0.0321 | 0.0305 | 0.0293 |  | 0.031772
| 1 | 5 | 0.0328 | 0.032732 | 0.032142 | 0.028891 | 0.03 | 0.033 | 0.0313 | 0.0282 | 0.0346 | 0.0321 | 0.0304 | 0.0296 | 0.033106 | 0.0341 | 0.0315 | 0.0312 | 0.0307 | 0.0289 |  | 96
| 1 | 6 | 0.0321 | 0.032675 | 0.032319 | 0.027569 | 0.02 | 0.033 | 0.0322 | 0.0281 | 0.0341 | 0.0324 | 0.033 | 0.0297 | 0.03264 | 0.0348 | 0.0313 | 0.0311 | 0.0298 | 0.0311
| Mean |  | 0.032067 | 0.032862 | 0.032226 | 0.027839 | 0.023333 | 0.033117 | 0.033117 | 0.02835 | 0.034433 | 0.032 | 0.030983 | 0.029617 | 0.032673 | 0.034433 | 0.032033 | 0.03115 | 0.030133 | 0.029433
| Median |  | 0.03205 | 0.032838 | 0.032088 | 0.027633 | 0.02 | 0.03305 | 0.03335 | 0.0283 | 0.03445 | 0.03225 | 0.03085 | 0.02965 | 0.032742 | 0.0344 | 0.03185 | 0.03105 | 0.03015 | 0.02915
| Std Dev. |  | 0.00048854 | 0.00032059 | 0.00034759 | 0.00057134 | 0.005164 | 0.00016021 | 0.001172 | 0.00022583 | 0.00021602 | 0.0006229 | 0.0014331 | 0.00026394 | 0.00033536 | 0.00089144 | 0.00074476 | 0.00050892 | 0.00052026 | 0.00083347
| Rel.Std.Dev. |  | 1.523% | 0.9756% | 1.079% | 2.052% | 22.13% | 0.4838% | 3.539% | 0.7966% | 0.6274% | 1.947% | 4.625% | 0.8912% | 1.026% | 2.589% | 2.325% | 1.634% | 1.727% | 2.832%
| PDM3 |  | 0.9289% | 3.431% | 1.431% | -12.38% | -26.56% | 4.234% | 4.234% | -10.77% | 8.378% | 0.719% | -2.481% | -6.782% | 2.839% | 8.378% | 0.824% | -1.956% | -5.156% | -7.359%
| Z-Score (Absolute) |  | 0.03 | 0.52 | 0.13 | 2.56 | 5.32 | 0.67 | 0.67 | 2.25 | 1.48 | 0.01 | 0.63 | 1.47 | 0.4 | 1.48 | 0.01 | 0.53 | 1.15 | 1.58
| Analytical results for Mo in OREAS 255c (Certified Value 3.94 ppm) | [20 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 4.4 | 3.9975 | < 10 | 4.3 | 3.89 | 3.3 | 3.8033 | 3.46 | 4 | 4.724 | 3.65 | 3.6647 | 4.09 | 3.66 | 4.28 | 3.64 | 4.12 |  |  | 1
| 1 | 2 | 4.4 | 4.0903 | < 10 | 4.3 | 3.92 | 3.35 | 3.7767 | 3.53 | 3.8 | 4.763 | 3.79 | 3.4282 | 4.28 | 3.85 | 4.24 | 3.62 | 4.22 |  |  | 25
| 1 | 3 | 4.3 | 3.987 | < 10 | 4.3 | 3.77 | 3.37 | 3.8467 | 3.42 | 3.3 | 4.958 | 3.85 | 3.495 | 4.08 | 3.69 | 4.38 | 3.58 | 4.04 |  |  | 16
| 1 | 4 | 4.3 | 3.965 | < 10 | 4.4 | 3.92 | 3.38 | 3.85 | 3.46 | 4.2 | 4.742 | 3.8 | 3.4706 | 4.01 | 3.71 | 4.31 | 3.7 | 4.16 |  |  | 3.9427
| 1 | 5 | 4.3 | 4 | < 10 | 4.4 | 3.56 | 3.36 | 3.83 | 3.56 | 3.9 | 4.922 | 3.75 | 3.6923 | 4.07 | 3.61 | 4.4 | 3.57 | 4.13 |  |  | 97
| 1 | 6 | 4.3 | 4.3868 | < 10 | 4.3 | 3.77 | 3.35 | 3.7867 | 3.7 | 4 | 4.852 | 3.96 | 3.4563 | 4.17 | 3.57 | 4.05 | 3.54 | 4.07
| Mean |  | 4.3333 | 4.0711 |  | 4.3333 | 3.805 | 3.3517 | 3.8156 | 3.5217 | 3.8667 | 4.8268 | 3.8 | 3.5345 | 4.1167 | 3.6817 | 4.2767 | 3.6083 | 4.1233
| Median |  | 4.3 | 3.9987 |  | 4.3 | 3.83 | 3.355 | 3.8167 | 3.495 | 3.95 | 4.8075 | 3.795 | 3.4828 | 4.085 | 3.675 | 4.295 | 3.6 | 4.125
| Std Dev. |  | 0.05164 | 0.16053 |  | 0.05164 | 0.13867 | 0.027869 | 0.031175 | 0.10128 | 0.30768 | 0.098733 | 0.10315 | 0.11394 | 0.095009 | 0.097245 | 0.12628 | 0.057417 | 0.064083
| Rel.Std.Dev. |  | 1.192% | 3.943% |  | 1.192% | 3.644% | 0.8315% | 0.817% | 2.876% | 7.957% | 2.045% | 2.714% | 3.224% | 2.308% | 2.641% | 2.953% | 1.591% | 1.554%
| PDM3 |  | 9.907% | 3.256% |  | 9.907% | -3.494% | -14.99% | -3.226% | -10.68% | -1.929% | 22.42% | -3.62% | -10.35% | 4.411% | -6.622% | 8.469% | -8.482% | 4.58%
| Z-Score (Absolute) |  | 1.22 | 0.53 | 2.96 | 1.22 | 0.16 | 1.34 | 0.13 | 0.9 | 0 | 2.51 | 0.17 | 0.87 | 0.65 | 0.48 | 1.07 | 0.67 | 0.67
| Analytical results for Na in OREAS 255c (Certified Value 0.268 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.29 | 0.26977 | 0.26569 | 1.18 | 0.27 | 0.283 | 0.22 | 0.27 | 0.27 | 0.248 | 0.25 | 0.27165 | 0.3 | 0.25 | 0.27 | 0.26 | 0.201 |  |  | 1
| 1 | 2 | 0.29 | 0.28034 | 0.27498 | 1.23 | 0.27 | 0.306 | 0.23 | 0.27 | 0.27 | 0.228 | 0.25 | 0.25289 | 0.31 | 0.26 | 0.28 | 0.27 | 0.199 |  |  | 0
| 1 | 3 | 0.28 | 0.27814 | 0.26694 | 1.18 | 0.28 | 0.265 | 0.23 | 0.27 | 0.26 | 0.271 | 0.26 | 0.24908 | 0.3 | 0.26 | 0.28 | 0.26 | 0.204 |  |  | 16
| 1 | 4 | 0.28 | 0.27408 | 0.26981 | 1.15 | 0.27 | 0.308 | 0.23 | 0.27 | 0.27 | 0.24 | 0.25 | 0.25252 | 0.29 | 0.26 | 0.29 | 0.27 | 0.206 |  |  | 0.26762
| 1 | 5 | 0.3 | 0.28012 | 0.28399 | 1.17 | 0.27 | 0.249 | 0.22 | 0.27 | 0.27 | 0.242 | 0.25 | 0.25965 | 0.3 | 0.25 | 0.28 | 0.27 | 0.201 |  |  | 98
| 1 | 6 | 0.29 | 0.27987 | 0.26927 | 1.17 | 0.27 | 0.282 | 0.22 | 0.26 | 0.28 | 0.266 | 0.26 | 0.27499 | 0.3 | 0.24 | 0.28 | 0.26 | 0.229
| Mean |  | 0.28833 | 0.27705 | 0.27178 | 1.18 | 0.27167 | 0.28217 | 0.225 | 0.26833 | 0.27 | 0.24917 | 0.25333 | 0.26013 | 0.3 | 0.25333 | 0.28 | 0.265 | 0.20667
| Median |  | 0.29 | 0.27901 | 0.26954 | 1.175 | 0.27 | 0.2825 | 0.225 | 0.27 | 0.27 | 0.245 | 0.25 | 0.25627 | 0.3 | 0.255 | 0.28 | 0.265 | 0.2025
| Std Dev. |  | 0.0075277 | 0.0042717 | 0.0067813 | 0.026833 | 0.0040825 | 0.022938 | 0.0054772 | 0.0040825 | 0.0063246 | 0.0164 | 0.005164 | 0.010827 | 0.0063246 | 0.008165 | 0.0063246 | 0.0054772 | 0.011219
| Rel.Std.Dev. |  | 2.611% | 1.542% | 2.495% | 2.274% | 1.503% | 8.129% | 2.434% | 1.521% | 2.342% | 6.582% | 2.038% | 4.162% | 2.108% | 3.223% | 2.259% | 2.067% | 5.429%
| PDM3 |  | 7.739% | 3.525% | 1.554% | 340.9% | 1.512% | 5.435% | -15.93% | 0.2662% | 0.8889% | -6.896% | -5.339% | -2.799% | 12.1% | -5.339% | 4.626% | -0.9794% | -22.78%
| Z-Score (Absolute) |  | 1.02 | 0.39 | 0.1 | 50.43 | 0.09 | 0.67 | 2.49 | 0.09 | 0 | 1.15 | 0.92 | 0.55 | 1.66 | 0.92 | 0.55 | 0.28 | 3.51
| Analytical results for Nb in OREAS 255c (Indicative Value 0.49 ppm) | [20 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 0.52 | 0.39693 | 25 | 0.6 | 0.3 | < 1 | < 5 | 0.47 | 0.39 | 0.86189 | 0.7 | 0.49 | 0.73 | 0.34 | 0.25 |  |  |  |  | 1
| 1 | 2 | 0.51 | 0.42279 | 25 | 0.6 | 0.3 | < 1 | < 5 | 0.54 | 0.39 | 0.76053 | 0.7 | 0.51 | 0.62 | 0.32 | 0.26 |  |  |  |  | 5
| 1 | 3 | 0.54 | 0.42109 | 15 | 0.6 | 0.3 | < 1 | < 5 | 0.47 | 0.38 | 0.83969 | 0.7 | 0.47 | 0.63 | 0.31 | 0.27 |  |  |  |  | 16
| 1 | 4 | 0.47 | 0.40828 | 17 | 0.6 | 0.3 | < 1 | < 5 | 0.53 | 0.41 | 0.87951 | 0.6 | 0.47 | 0.61 | 0.36 | 0.25 |  |  |  |  | 0.49362
| 1 | 5 | 0.51 | 0.40605 | 27 | 0.6 | 0.3 | < 1 | < 5 | 0.54 | 0.38 | 0.79543 | 0.7 | 0.46 | 0.69 | 0.32 | 0.25 |  |  |  |  | 11
| 1 | 6 | 0.46 | 0.40676 | 25 | 0.6 | 0.3 | < 1 | < 5 | 0.51 | 0.36 | 0.76189 | 0.7 | 0.43 | 0.65 | 0.36 | 0.25
| Mean |  | 0.50167 | 0.41032 | 22.3333 | 0.6 | 0.3 |  |  | 0.51 | 0.385 | 0.81649 | 0.68333 | 0.47167 | 0.655 | 0.335 | 0.255
| Median |  | 0.51 | 0.40752 | 25 | 0.6 | 0.3 |  |  | 0.52 | 0.385 | 0.81756 | 0.7 | 0.47 | 0.64 | 0.33 | 0.25
| Std Dev. |  | 0.030605 | 0.0098578 | 5.0067 | 0 | 0 |  |  | 0.032863 | 0.016432 | 0.051238 | 0.040825 | 0.027142 | 0.046368 | 0.021679 | 0.0083666
| Rel.Std.Dev. |  | 6.101% | 2.402% | 22.42% | 0% | 0% |  |  | 6.444% | 4.268% | 6.275% | 5.974% | 5.754% | 7.079% | 6.471% | 3.281%
| PDM3 |  | 1.63% | -16.88% | 4424% | 21.55% | -39.22% |  |  | 3.318% | -22.01% | 65.41% | 38.43% | -4.448% | 32.69% | -32.13% | -48.34%
| Z-Score (Absolute) |  | 0 | 0.4 | 96.01 | 0.43 | 0.89 | 0.01 | 8.79 | 0.04 | 0.51 | 1.38 | 0.8 | 0.13 | 0.67 | 0.73 | 1.08
| Analytical results for Nd in OREAS 255c (Indicative Value 22.8 ppm) | [20 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.09 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*OES/MS |  | 1
|  |  | 1.01: 25g; 1.02: 0.5g |  | 1
| 1 | 1 | 1.01: 23.474; 1.02: 23.1 |  | 1
| 1 | 2 | 1.01: 23.476; 1.02: 23.4 |  | 6
| 1 | 3 | 1.01: 23.015; 1.02: 22.2 |  | 16
| 1 | 4 | 1.01: 22.68; 1.02: 22.7 |  | 22.8353
| 1 | 5 | 1.01: 23.395; 1.02: 20.8 |  | 12
| 1 | 6 | 1.01: 23.084; 1.02: 22.7
| Mean |  | 1.01: 23.1873; 1.02: 22.4833
| Median |  | 1.01: 23.2395; 1.02: 22.7
| Std Dev. |  | 1.01: 0.31824; 1.02: 0.9196
| Rel.Std.Dev. |  | 1.01: 1.372%; 1.02: 4.09%
| PDM3 |  | 1.01: 1.541%; 1.02: -1.541%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ni in OREAS 255c (Certified Value 84 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 1
| 1 | 1 | 85 | 85.1364 | 81.3954 | 75.705 | 70 | 81 | 89.6 | 81.7 | 80.4827 | 82.9 | 80.3 | 83.8 | 86.435 | 89.3 | 83.4 | 84.5 | 82.8 | 61.7 |  | 1
| 1 | 2 | 84 | 87.7392 | 83.7296 | 77.211 | 72 | 81 | 91.9 | 82.1 | 80.2153 | 85.2 | 75.1 | 85.9 | 87.1399 | 90.6 | 87.8 | 85.4 | 84.6 | 61.4 |  | 28
| 1 | 3 | 82.8 | 86.0436 | 82.2262 | 75.417 | 66 | 80 | 92.7 | 83.6 | 80.7733 | 81.7 | 83.4 | 88.8 | 87.0876 | 89.2 | 85.9 | 83.8 | 81.1 | 61.9 |  | 16
| 1 | 4 | 82.5 | 88.182 | 81.2188 | 76.034 | 69 | 80 | 91.7 | 83 | 80.984 | 83.3 | 80.3 | 86.5 | 85.9814 | 90.5 | 84.8 | 88.3 | 85.1 | 62.4 |  | 83.9296
| 1 | 5 | 85.6 | 86.292 | 82.6056 | 79.173 | 74 | 81 | 89.1 | 82.7 | 80.61 | 83.7 | 77.1 | 84.5 | 87.4353 | 88.8 | 82.7 | 86.5 | 82.7 | 62.6 |  | 99
| 1 | 6 | 83.7 | 86.4972 | 83.3107 | 79.71 | 72 | 81 | 87.9 | 81.7 | 79.4023 | 86.4 | 80.4 | 87.1 | 86.1725 | 88.1 | 82.4 | 85.3 | 81.6 | 63.5
| Mean |  | 83.9333 | 86.6484 | 82.4144 | 77.2083 | 70.5 | 80.6667 | 90.4833 | 82.4667 | 80.4113 | 83.8667 | 79.4333 | 86.1 | 86.7086 | 89.4167 | 84.5 | 85.6333 | 82.9833 | 62.25
| Median |  | 83.85 | 86.3946 | 82.4159 | 76.6225 | 71 | 81 | 90.65 | 82.4 | 80.5463 | 83.5 | 80.3 | 86.2 | 86.7613 | 89.25 | 84.1 | 85.35 | 82.75 | 62.15
| Std Dev. |  | 1.2094 | 1.1264 | 1.0073 | 1.8421 | 2.8107 | 0.5164 | 1.8851 | 0.76594 | 0.55858 | 1.6836 | 2.9118 | 1.8078 | 0.59141 | 0.97451 | 2.0881 | 1.5921 | 1.5917 | 0.75565
| Rel.Std.Dev. |  | 1.441% | 1.3% | 1.222% | 2.386% | 3.987% | 0.6402% | 2.083% | 0.9288% | 0.6947% | 2.008% | 3.666% | 2.1% | 0.6821% | 1.09% | 2.471% | 1.859% | 1.918% | 1.214%
| PDM3 |  | 0.004393% | 3.239% | -1.805% | -8.008% | -16% | -3.888% | 7.809% | -1.743% | -4.192% | -0.07504% | -5.357% | 2.586% | 3.311% | 6.538% | 0.6796% | 2.03% | -1.128% | -25.83%
| Z-Score (Absolute) |  | 0.12 | 0.75 | 0.24 | 1.45 | 3.02 | 0.64 | 1.65 | 0.22 | 0.7 | 0.1 | 0.93 | 0.63 | 0.77 | 1.4 | 0.25 | 0.52 | 0.1 | 4.95
| Analytical results for P in OREAS 255c (Certified Value 0.048 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0481 | 0.050131 | 0.045697 | 0.05 | 0.0516 | 0.05 | 0.0442 | 0.0452 | 0.048 | 0.047 | 0.047 | 0.053651 | 0.05 | 0.049 | 0.048 | 0.046 | 0.041 |  |  | 1
| 1 | 2 | 0.0469 | 0.051398 | 0.04512 | 0.06 | 0.0518 | 0.05 | 0.0443 | 0.0452 | 0.05 | 0.044 | 0.046 | 0.053784 | 0.05 | 0.049 | 0.048 | 0.048 | 0.042 |  |  | 0
| 1 | 3 | 0.0467 | 0.04898 | 0.043135 | 0.05 | 0.0514 | 0.051 | 0.044 | 0.0453 | 0.048 | 0.049 | 0.047 | 0.052925 | 0.05 | 0.049 | 0.049 | 0.046 | 0.042 |  |  | 16
| 1 | 4 | 0.0463 | 0.05036 | 0.045769 | 0.05 | 0.052 | 0.051 | 0.0441 | 0.0457 | 0.049 | 0.046 | 0.046 | 0.053026 | 0.05 | 0.05 | 0.05 | 0.048 | 0.043 |  |  | 0.048376
| 1 | 5 | 0.0484 | 0.049316 | 0.04469 | 0.06 | 0.0521 | 0.05 | 0.044 | 0.0454 | 0.049 | 0.046 | 0.046 | 0.052412 | 0.05 | 0.049 | 0.049 | 0.047 | 0.041 |  |  | 100
| 1 | 6 | 0.047 | 0.050998 | 0.044296 | 0.05 | 0.0515 | 0.049 | 0.0444 | 0.0448 | 0.049 | 0.051 | 0.047 | 0.052981 | 0.05 | 0.048 | 0.049 | 0.047 | 0.043
| Mean |  | 0.047233 | 0.050197 | 0.044785 | 0.053333 | 0.051733 | 0.050167 | 0.044167 | 0.045267 | 0.048833 | 0.047167 | 0.0465 | 0.05313 | 0.05 | 0.049 | 0.048833 | 0.047 | 0.042
| Median |  | 0.04695 | 0.050245 | 0.044905 | 0.05 | 0.0517 | 0.05 | 0.04415 | 0.04525 | 0.049 | 0.0465 | 0.0465 | 0.053003 | 0.05 | 0.049 | 0.049 | 0.047 | 0.042
| Std Dev. |  | 0.00082865 | 0.00093524 | 0.00098848 | 0.005164 | 0.00028048 | 0.00075277 | 0.0001633 | 0.00029439 | 0.00075277 | 0.0024833 | 0.00054772 | 0.00050789 | 7.6012e-18 | 0.00063246 | 0.00075277 | 0.00089443 | 0.00089443
| Rel.Std.Dev. |  | 1.754% | 1.863% | 2.207% | 9.682% | 0.5422% | 1.501% | 0.3697% | 0.6504% | 1.542% | 5.265% | 1.178% | 0.9559% | 1.52e-14% | 1.291% | 1.542% | 1.903% | 2.13%
| PDM3 |  | -2.361% | 3.765% | -7.423% | 10.25% | 6.941% | 3.702% | -8.701% | -6.427% | 0.946% | -2.499% | -3.877% | 9.828% | 3.358% | 1.29% | 0.946% | -2.844% | -13.18%
| Z-Score (Absolute) |  | 0.59 | 0.5 | 1.49 | 1.66 | 1.07 | 0.49 | 1.72 | 1.31 | 0 | 0.61 | 0.86 | 1.58 | 0.43 | 0.06 | 0 | 0.67 | 2.51
| Analytical results for Pb in OREAS 255c (Certified Value 11.5 ppm) | [20 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 12 | 11.6481 | < 100 | 19 | 11.5 | 11.5 | 10.8 | 11.5062 | 10.4 | 11 | 7.2199 | 9.6 | 12.7059 | 13.2 | 11.1 | 12.4 | 11.4 |  |  | 1
| 1 | 2 | 11.8 | 11.8311 | < 100 | 17 | 11.3 | 11.9 | 11 | 11.2982 | 10.6 | 9.9 | 7.2955 | 9.9 | 12.8682 | 13.2 | 12 | 12.2 | 11.7 |  |  | 29
| 1 | 3 | 11.6 | 11.5424 | < 100 | 13 | 12 | 11.8 | 11.2 | 11.7778 | 10.1 | 9.3 | 7.4414 | 9.9 | 12.7715 | 13.1 | 11.3 | 12.1 | 11.8 |  |  | 16
| 1 | 4 | 11.5 | 11.865 | < 100 | 19 | 11.6 | 12 | 11.1 | 11.4925 | 10.2 | 11 | 7.1001 | 9.6 | 12.8031 | 13.1 | 11.4 | 12.1 | 11.4 |  |  | 11.4625
| 1 | 5 | 11.7 | 11.9649 | < 100 | 13 | 11.5 | 11.2 | 10.3 | 11.3025 | 10.2 | 10.5 | 7.1754 | 9.9 | 13.144 | 12.8 | 11.4 | 12.4 | 11.4 |  |  | 101
| 1 | 6 | 11.6 | 11.7105 | < 100 | 17 | 11.5 | 11.1 | 10.8 | 11.3622 | 10.4 | 10.8 | 7.1822 | 10.2 | 13.1921 | 13.5 | 11.2 | 11.7 | 11.6
| Mean |  | 11.7 | 11.7603 |  | 16.3333 | 11.5667 | 11.5833 | 10.8667 | 11.4566 | 10.3167 | 10.4167 | 7.2358 | 9.85 | 12.9141 | 13.15 | 11.4 | 12.15 | 11.55
| Median |  | 11.65 | 11.7708 |  | 17 | 11.5 | 11.65 | 10.9 | 11.4274 | 10.3 | 10.65 | 7.201 | 9.9 | 12.8356 | 13.15 | 11.35 | 12.15 | 11.5
| Std Dev. |  | 0.17889 | 0.15518 |  | 2.7325 | 0.23381 | 0.37639 | 0.32042 | 0.18145 | 0.18348 | 0.68532 | 0.11921 | 0.22583 | 0.20407 | 0.22583 | 0.31623 | 0.25884 | 0.17607
| Rel.Std.Dev. |  | 1.529% | 1.32% |  | 16.73% | 2.021% | 3.249% | 2.949% | 1.584% | 1.779% | 6.579% | 1.647% | 2.293% | 1.58% | 1.717% | 2.774% | 2.13% | 1.524%
| PDM3 |  | 2.072% | 2.599% |  | 42.49% | 0.9092% | 1.055% | -5.198% | -0.05123% | -9.996% | -9.124% | -36.87% | -14.07% | 12.66% | 14.72% | -0.5449% | 5.998% | 0.7638%
| Z-Score (Absolute) |  | 0.13 | 0.19 | 37.02 | 4.59 | 0 | 0.02 | 0.67 | 0.11 | 1.2 | 1.11 | 4.17 | 1.65 | 1.3 | 1.53 | 0.16 | 0.56 | 0.02
| Analytical results for Pd in OREAS 255c (Indicative Value < 10 ppb) | [20 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.16 |  |  | ppb
|  |  | 1.01: AR*MS; 1.03: AR*OES/MS |  |  | 1
|  |  | 1.01: 25g; 1.03: 01g |  |  | 1
| 1 | 1 | 1.01: < 10; 1.03: < 10 |  |  | 1
| 1 | 2 | 1.01: < 10; 1.03: < 10 |  |  | 7
| 1 | 3 | 1.01: < 10; 1.03: < 10 |  |  | 16
| 1 | 4 | 1.01: < 10; 1.03: < 10 |  |  | < 10
| 1 | 5 | 1.01: < 10; 1.03: < 10 |  |  | 13
| 1 | 6 | 1.01: < 10; 1.03: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Pr in OREAS 255c (Indicative Value 5.89 ppm) | [20 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.05: 1.09 |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.05: AR*OES/MS |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.05: 0.5g |  |  | 2
| 1 | 1 | 1.01: 6.239; 1.03: 5.3925; 1.05: 6.3 |  |  | 1
| 1 | 2 | 1.01: 6.024; 1.03: 5.7284; 1.05: 6.3 |  |  | 8
| 1 | 3 | 1.01: 5.882; 1.03: 5.8109; 1.05: 6.1 |  |  | 16
| 1 | 4 | 1.01: 5.79; 1.03: 5.5966; 1.05: 6.2 |  |  | 5.8877
| 1 | 5 | 1.01: 5.928; 1.03: 5.6934; 1.05: 5.6 |  |  | 14
| 1 | 6 | 1.01: 5.924; 1.03: 5.37; 1.05: 6.1
| Mean |  | 1.01: 5.9645; 1.03: 5.5986; 1.05: 6.1
| Median |  | 1.01: 5.926; 1.03: 5.645; 1.05: 6.15
| Std Dev. |  | 1.01: 0.15435; 1.03: 0.18199; 1.05: 0.26077
| Rel.Std.Dev. |  | 1.01: 2.588%; 1.03: 3.251%; 1.05: 4.275%
| PDM3 |  | 1.01: 1.304%; 1.03: -4.91%; 1.05: 3.606%
| Z-Score (Absolute) |  | 1.01: 0; 1.03: 1.82; 1.05: 0.67
| Analytical results for Pt in OREAS 255c (Indicative Value < 5 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.16 |  |  | ppb
|  |  | 1.01: AR*MS; 1.03: AR*OES/MS |  |  | 2
|  |  | 1.01: 25g; 1.03: 01g |  |  | 2
| 1 | 1 | 1.01: < 5; 1.03: < 5 |  |  | 1
| 1 | 2 | 1.01: < 5; 1.03: < 5 |  |  | 3
| 1 | 3 | 1.01: < 5; 1.03: < 5 |  |  | 16
| 1 | 4 | 1.01: < 5; 1.03: < 5 |  |  | < 5
| 1 | 5 | 1.01: < 5; 1.03: < 5 |  |  | 9
| 1 | 6 | 1.01: < 5; 1.03: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Rb in OREAS 255c (Certified Value 14.6 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 16.32 | 15.5055 | 13.9 | 15.6 | 14.1 | 13 | 13.9 | 15.3027 | 19.8 | 14.4 | 15.9 | 13.7 | 12.67 | [6 empty] | 1
| 1 | 2 | 16.16 | 15.8247 | 13.7 | 15.9 | 14.1 | 13.2 | 14.8 | 14.8191 | 20.6 | 15.6 | 15.7 | 14.3 | 12.45 | [6 empty] | 31
| 1 | 3 | 16.37 | 14.4539 | 14.4 | 15.4 | 14.4 | 12.6 | 14.9 | 15.3078 | 20 | 14.9 | 15.6 | 13.8 | 12.93 | [6 empty] | 16
| 1 | 4 | 15.57 | 15.4878 | 13.9 | 15.6 | 14.3 | 13 | 14.2 | 15.4263 | 19.4 | 14.9 | 15.6 | 14.5 | 12.86 | [6 empty] | 14.5777
| 1 | 5 | 15.96 | 14.4862 | 14.6 | 14.2 | 14.2 | 13.1 | 14.6 | 15.1524 | 19.6 | 14.4 | 15.9 | 14.1 | 12.47 | [6 empty] | 102
| 1 | 6 | 15.71 | 15.3771 | 14.1 | 15.1 | 14.2 | 13.8 | 14.4 | 14.7317 | 20.2 | 14.4 | 15.3 | 14.3 | 12.43
| Mean |  | 16.015 | 15.1892 | 14.1 | 15.3 | 14.2167 | 13.1167 | 14.4667 | 15.1233 | 19.9333 | 14.7667 | 15.6667 | 14.1167 | 12.635
| Median |  | 16.06 | 15.4325 | 14 | 15.5 | 14.2 | 13.05 | 14.5 | 15.2275 | 19.9 | 14.65 | 15.65 | 14.2 | 12.57
| Std Dev. |  | 0.32685 | 0.57677 | 0.34059 | 0.6 | 0.1169 | 0.392 | 0.37771 | 0.28451 | 0.43205 | 0.4761 | 0.22509 | 0.31252 | 0.22016
| Rel.Std.Dev. |  | 2.041% | 3.797% | 2.416% | 3.922% | 0.8223% | 2.989% | 2.611% | 1.881% | 2.167% | 3.224% | 1.437% | 2.214% | 1.742%
| PDM3 |  | 9.86% | 4.195% | -3.277% | 4.955% | -2.477% | -10.02% | -0.7617% | 3.743% | 36.74% | 1.296% | 7.47% | -3.163% | -13.33%
| Z-Score (Absolute) |  | 1.3 | 0.44 | 0.69 | 0.55 | 0.57 | 1.71 | 0.31 | 0.37 | 5.36 | 0 | 0.93 | 0.67 | 2.21
| Analytical results for Re in OREAS 255c (Certified Value 0.001 ppm) | [20 empty] | Certified
Re | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.07; 1.05: 1.09; 1.06: 1.10; 1.07: 1.12; 1.09: 1.15; 1.10: 1.16; 1.11: 1.19; 1.12: 1.20; 1.13: 1.21 |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*OES/MS; 1.05: AR*OES/MS; 1.06: AR*MS; 1.07: AR*MS; 1.09: AR*MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS |  |  | 3
|  |  | 1.01: 25g; 1.03: 01g; 1.05: 0.5g; 1.06: 0.2g; 1.07: 0.5g; 1.09: 0.5g; 1.10: 01g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g |  |  | 3
| 1 | 1 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.002; 1.06: < 0.002; 1.07: 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: 0.001; 1.12: < 0.001; 1.13: < 0.001 |  |  | 1
| 1 | 2 | 1.01: < 0.001; 1.03: < 0.05; 1.05: < 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: 0.001; 1.12: < 0.001; 1.13: < 0.001 |  |  | 32
| 1 | 3 | 1.01: < 0.001; 1.03: < 0.05; 1.05: < 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: 0.001; 1.12: < 0.001; 1.13: 0.001 |  |  | 16
| 1 | 4 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: 0.001; 1.12: < 0.001; 1.13: 0.001 |  |  | 0.0008125
| 1 | 5 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.002; 1.06: < 0.002; 1.07: < 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: 0.001; 1.12: < 0.001; 1.13: 0.001 |  |  | 103
| 1 | 6 | 1.01: < 0.001; 1.03: < 0.05; 1.05: 0.001; 1.06: < 0.002; 1.07: 0.001; 1.09: 0.001; 1.10: < 0.05; 1.11: 0.001; 1.12: < 0.001; 1.13: 0.001
| Mean |  | 1.05: 0.0015; 1.07: 0.001; 1.09: 0.001; 1.11: 0.001; 1.13: 0.001
| Median |  | 1.05: 0.0015; 1.07: 0.001; 1.09: 0.001; 1.11: 0.001; 1.13: 0.001
| Std Dev. |  | 1.05: 0.00057735; 1.07: 0; 1.09: 0; 1.11: 0; 1.13: 0
| Rel.Std.Dev. |  | 1.05: 38.49%; 1.07: 0%; 1.09: 0%; 1.11: 0%; 1.13: 0%
| PDM3 |  | 1.05: 84.62%; 1.07: 23.08%; 1.09: 23.08%; 1.11: 23.08%; 1.13: 23.08%
| Z-Score (Absolute) |  | 1.01: 2.02; 1.03: 97.1; 1.05: 0.67; 1.06: 0; 1.07: 0.34; 1.09: 0; 1.10: 97.1; 1.11: 0; 1.12: 2.02; 1.13: 0.67
| Analytical results for S in OREAS 255c (Certified Value 0.034 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.035 | 0.034876 |  | 0.57 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.04 | < 0.05 | 0.032225 | 0.04 | 0.035517 | 0.04 | 0.03 | 0.03 | 0.04 | 0.03 |  | 1
| 1 | 2 | 0.0344 | 0.03558 |  | 0.59 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.032724 | 0.04 | 0.036587 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 |  | 17
| 1 | 3 | 0.0341 | 0.034214 |  | 0.56 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.033221 | 0.04 | 0.037698 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 |  | 16
| 1 | 4 | 0.0344 | 0.034886 |  | 0.56 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.032724 | 0.04 | 0.037201 | 0.03 | 0.03 | 0.04 | 0.04 | 0.03 |  | 0.03421
| 1 | 5 | 0.0353 | 0.034781 |  | 0.6 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.032623 | 0.04 | 0.037361 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 |  | 104
| 1 | 6 | 0.0348 | 0.035171 |  | 0.61 | 0.03 | < 0.001 | 0.03 | 0.03 | 0.03 | < 0.05 | 0.033024 | 0.04 | 0.03522 | 0.04 | 0.03 | 0.04 | 0.04 | 0.04
| Mean |  | 0.034667 | 0.034918 |  | 0.58167 | 0.03 |  | 0.03 | 0.03 | 0.031667 |  | 0.032757 | 0.04 | 0.036597 | 0.038333 | 0.03 | 0.038333 | 0.04 | 0.031667
| Median |  | 0.0346 | 0.034881 |  | 0.58 | 0.03 |  | 0.03 | 0.03 | 0.03 |  | 0.032724 | 0.04 | 0.036894 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03
| Std Dev. |  | 0.00044572 | 0.00045164 |  | 0.02137 | 0 |  | 0 | 0 | 0.0040825 |  | 0.00034344 | 0 | 0.001022 | 0.0040825 | 0 | 0.0040825 | 0 | 0.0040825
| Rel.Std.Dev. |  | 1.286% | 1.293% |  | 3.674% | 0% |  | 0% | 0% | 12.89% |  | 1.048% | 0% | 2.793% | 10.65% | 0% | 10.65% | 0% | 12.89%
| PDM3 |  | 1.335% | 2.07% |  | 1600% | -12.31% |  | -12.31% | -12.31% | -7.434% |  | -4.248% | 16.93% | 6.979% | 12.05% | -12.31% | 12.05% | 16.93% | -7.434%
| Z-Score (Absolute) |  | 0.47 | 0.53 | NA | 134.26 | 0.67 | 7.89 | 0.67 | 0.67 | 0.27 | 1.9 | 0 | 1.77 | 0.94 | 1.36 | 0.67 | 1.36 | 1.77 | 0.27
| Analytical results for Sb in OREAS 255c (Certified Value 5.23 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  | 2
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  | 2
| 1 | 1 | 7.34 | 4.4309 | < 10 | 5.4 | 5.77 | 5.4 | 4.81 | 5.17 | 5.07 | 5.6 | 5.979 | 4.18 | 5.9181 | 7.15 | 4.93 | 6.04 | 4.74 | 5.08 |  | 1
| 1 | 2 | 7.18 | 4.7768 | < 10 | 5.5 | 5.71 | 5.18 | 4.62 | 5.52 | 5.2 | 5.3 | 6.318 | 4.56 | 5.8197 | 7.23 | 5.07 | 5.76 | 4.7 | 4.95 |  | 33
| 1 | 3 | 7.35 | 4.6835 | < 10 | 3.4 | 5.54 | 5.06 | 4.58 | 5.3767 | 4.96 | 6.1 | 6.412 | 4.58 | 5.9767 | 7.01 | 4.67 | 5.8 | 4.57 | 5.03 |  | 16
| 1 | 4 | 7.08 | 4.6109 | < 10 | 3.9 | 5.87 | 5.19 | 4.67 | 5.4267 | 5.13 | 5.7 | 6.162 | 4.52 | 5.7359 | 6.96 | 4.81 | 5.9 | 4.82 | 5.19 |  | 5.2302
| 1 | 5 | 7.06 | 4.6284 | < 10 | 5.1 | 5.61 | 5.1 | 4.56 | 5.2433 | 5.16 | 5.4 | 6.285 | 4.34 | 5.7186 | 7.08 | 4.94 | 5.75 | 4.48 | 5.18 |  | 105
| 1 | 6 | 7.24 | 4.5779 | < 10 | 6.4 | 5.68 | 5.09 | 4.68 | 5.3 | 5.14 | 5.5 | 6.208 | 4.39 | 5.8345 | 7.24 | 4.93 | 5.49 | 4.77 | 5.32
| Mean |  | 7.2083 | 4.6181 |  | 4.95 | 5.6967 | 5.17 | 4.6533 | 5.3394 | 5.11 | 5.6 | 6.2273 | 4.4283 | 5.8339 | 7.1117 | 4.8917 | 5.79 | 4.68 | 5.125
| Median |  | 7.21 | 4.6196 |  | 5.25 | 5.695 | 5.14 | 4.645 | 5.3383 | 5.135 | 5.55 | 6.2465 | 4.455 | 5.8271 | 7.115 | 4.93 | 5.78 | 4.72 | 5.13
| Std Dev. |  | 0.12465 | 0.11514 |  | 1.1077 | 0.11656 | 0.12394 | 0.090259 | 0.12736 | 0.084853 | 0.28284 | 0.14964 | 0.15471 | 0.10058 | 0.11514 | 0.1363 | 0.18287 | 0.12946 | 0.13187
| Rel.Std.Dev. |  | 1.729% | 2.493% |  | 22.38% | 2.046% | 2.397% | 1.94% | 2.385% | 1.661% | 5.051% | 2.403% | 3.494% | 1.724% | 1.619% | 2.786% | 3.158% | 2.766% | 2.573%
| PDM3 |  | 37.82% | -11.7% |  | -5.358% | 8.918% | -1.152% | -11.03% | 2.088% | -2.299% | 7.069% | 19.06% | -15.33% | 11.54% | 35.97% | -6.474% | 10.7% | -10.52% | -2.012%
| Z-Score (Absolute) |  | 2.89 | 0.74 | 0.21 | 0.28 | 0.77 | 0.03 | 0.69 | 0.27 | 0.05 | 0.63 | 1.51 | 1.01 | 0.96 | 2.75 | 0.36 | 0.9 | 0.66 | 0.03
| Analytical results for Sc in OREAS 255c (Certified Value 5.2 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 5.5 | 5.0468 | 5 | 5.2 | 4.7 | 5.1227 | 5.1 | 4.8 | 4.5 | 4.8962 | 6 | 5.5 | 5.4 | 5.1 | 5.7 |  |  |  |  | 1
| 1 | 2 | 5.4 | 5.2031 | 5 | 5.6 | 4.7 | 5.0853 | 5.2 | 4.3 | 4.8 | 4.8148 | 6.2 | 5.9 | 5.5 | 5.2 | 5.6 |  |  |  |  | 34
| 1 | 3 | 5.3 | 5.0987 | 5 | 5.3 | 4.7 | 5.1161 | 5 | 5.1 | 4.9 | 4.8657 | 6 | 5.6 | 5.3 | 5 | 5.6 |  |  |  |  | 16
| 1 | 4 | 5.3 | 5.0797 | 5 | 5.5 | 4.7 | 5.1655 | 5 | 4.8 | 4.8 | 4.6547 | 5.9 | 5.6 | 5.4 | 5.2 | 5.8 |  |  |  |  | 5.1955
| 1 | 5 | 5.5 | 5.0722 | 5 | 4.8 | 4.7 | 5.0791 | 5 | 4.5 | 4.7 | 4.69 | 5.9 | 5.4 | 5.5 | 5.1 | 5.8 |  |  |  |  | 106
| 1 | 6 | 5.4 | 5.0999 | 5 | 5.4 | 4.7 | 5.0667 | 5.3 | 4.8 | 4.9 | 4.5658 | 6.2 | 5.4 | 5.3 | 5.1 | 5.6
| Mean |  | 5.4 | 5.1001 | 5 | 5.3 | 4.7 | 5.1059 | 5.1 | 4.7167 | 4.7667 | 4.7479 | 6.0333 | 5.5667 | 5.4 | 5.1167 | 5.6833
| Median |  | 5.4 | 5.0892 | 5 | 5.35 | 4.7 | 5.1007 | 5.05 | 4.8 | 4.8 | 4.7524 | 6 | 5.55 | 5.4 | 5.1 | 5.65
| Std Dev. |  | 0.089443 | 0.054101 | 0 | 0.28284 | 0 | 0.036377 | 0.12649 | 0.27869 | 0.15055 | 0.13079 | 0.13663 | 0.18619 | 0.089443 | 0.075277 | 0.098319
| Rel.Std.Dev. |  | 1.656% | 1.061% | 0% | 5.337% | 0% | 0.7124% | 2.48% | 5.909% | 3.158% | 2.755% | 2.265% | 3.345% | 1.656% | 1.471% | 1.73%
| PDM3 |  | 3.936% | -1.837% | -3.763% | 2.011% | -9.537% | -1.725% | -1.838% | -9.217% | -8.254% | -8.616% | 16.13% | 7.144% | 3.936% | -1.518% | 9.389%
| Z-Score (Absolute) |  | 0.61 | 0.02 | NA | 0.4 | 0.88 | 0.01 | 0.02 | 0.84 | 0.73 | 0.77 | 1.96 | 0.97 | 0.61 | 0.01 | 1.22
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 255c (Indicative Value 0.44 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | < 1.5 | < 1 | 0.7 | 0.5 | < 5 | 0.5 | < 0.5 | 0.272 | 0.4 | < 1 | 1.1 | 0.3 | < 0.2 | 0.3 |  |  |  |  | 1
| 1 | 2 | < 1 | < 1.5 | < 1 | 0.6 | 0.5 | < 5 | 0.4 | 0.6 | 0.294 | 0.6 | < 1 | 1.5 | 0.3 | < 0.2 | 0.3 |  |  |  |  | 35
| 1 | 3 | < 1 | < 1.5 | < 1 | 0.7 | 0.5 | < 5 | 0.5 | 0.7 | 0.274 | 0.3 | < 1 | 1.5 | 0.2 | < 0.2 | 0.2 |  |  |  |  | 16
| 1 | 4 | < 1 | < 1.5 | < 1 | 0.6 | 0.5 | < 5 | 0.6 | 0.5 | 0.276 | 0.3 | < 1 | 1.5 | 0.2 | < 0.2 | 0.4 |  |  |  |  | 0.43651
| 1 | 5 | < 1 | < 1.5 | < 1 | 0.6 | 0.6 | < 5 | 0.6 | 0.6 | 0.281 | 0.5 | < 1 | 1.5 | 0.3 | < 0.2 | 0.2 |  |  |  |  | 10
| 1 | 6 | < 1 | < 1.5 | < 1 | 0.6 | 0.5 | < 5 | 0.5 | < 0.5 | 0.274 | 0.4 | < 1 | 1.6 | 0.3 | < 0.2 | 0.2
| Mean |  |  |  |  | 0.63333 | 0.51667 |  | 0.51667 | 0.6 | 0.2785 | 0.41667 |  | 1.45 | 0.26667 |  | 0.26667
| Median |  |  |  |  | 0.6 | 0.5 |  | 0.5 | 0.6 | 0.275 | 0.4 |  | 1.5 | 0.3 |  | 0.25
| Std Dev. |  |  |  |  | 0.05164 | 0.040825 |  | 0.075277 | 0.08165 | 0.0081915 | 0.1169 |  | 0.17607 | 0.05164 |  | 0.08165
| Rel.Std.Dev. |  |  |  |  | 8.154% | 7.902% |  | 14.57% | 13.61% | 2.941% | 28.06% |  | 12.14% | 19.36% |  | 30.62%
| PDM3 |  |  |  |  | 45.09% | 18.36% |  | 18.36% | 37.45% | -36.2% | -4.546% |  | 232.2% | -38.91% |  | -38.91%
| Z-Score (Absolute) |  | 0 | 1.26 | 0 | 0.67 | 0.08 | 10.11 | 0.08 | 0.08 | 1.12 | 0.42 | 0 | 4.8 | 1.18 | 2.02 | 1.18
| Analytical results for Sm in OREAS 255c (Indicative Value 4.07 ppm) | [20 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g |  |  |  |  | 2
| 1 | 1 | 1.01: 4.181; 1.03: 3.7705; 1.07: 4.3 |  |  |  |  | 1
| 1 | 2 | 1.01: 4.097; 1.03: 4.0962; 1.07: 3.7 |  |  |  |  | 5
| 1 | 3 | 1.01: 4.016; 1.03: 4.1029; 1.07: 4.4 |  |  |  |  | 16
| 1 | 4 | 1.01: 3.989; 1.03: 3.9899; 1.07: 4.1 |  |  |  |  | 4.0668
| 1 | 5 | 1.01: 4.112; 1.03: 3.9961; 1.07: 4.5 |  |  |  |  | 11
| 1 | 6 | 1.01: 4.042; 1.03: 3.9105; 1.07: 3.9
| Mean |  | 1.01: 4.0728; 1.03: 3.9777; 1.07: 4.15
| Median |  | 1.01: 4.0695; 1.03: 3.993; 1.07: 4.2
| Std Dev. |  | 1.01: 0.070749; 1.03: 0.1246; 1.07: 0.30822
| Rel.Std.Dev. |  | 1.01: 1.737%; 1.03: 3.132%; 1.07: 7.427%
| PDM3 |  | 1.01: 0.1474%; 1.03: -2.192%; 1.07: 2.045%
| Z-Score (Absolute) |  | 1.01: 0; 1.03: 0.83; 1.07: 0.67
| Analytical results for Sn in OREAS 255c (Certified Value 1.37 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 1.46 | 1.3691 | 3 | 1.6 | 1.63 | 1.1 | < 10 | 1.3 | 1.1 | 1.4483 | 1.8 | 1.2 | 1.5 | 1.2 | 1.4 |  |  |  |  | 1
| 1 | 2 | 1.38 | 1.39 | 2 | 1.6 | 2.42 | 1.2 | < 10 | 1.3 | 1.1 | 1.2073 | 1.8 | 1.3 | 1.5 | 1.2 | 1.3 |  |  |  |  | 20
| 1 | 3 | 1.39 | 1.3872 | 2 | 1.7 | 1.61 | 1.2 | < 10 | 1.2 | 1.2 | 1.3865 | 1.7 | 1.2 | 1.4 | 1.2 | 1.4 |  |  |  |  | 16
| 1 | 4 | 1.38 | 1.3199 | < 1 | 1.7 | 1.64 | 1.1 | < 10 | 1.3 | 1.2 | 1.1615 | 1.7 | 1.2 | 1.4 | 1.2 | 1.3 |  |  |  |  | 1.3707
| 1 | 5 | 1.42 | 1.3821 | < 1 | 1.7 | 1.5 | 1.1 | < 10 | 1.3 | 1.2 | 1.2217 | 1.7 | 1.2 | 1.5 | 1.2 | 1.3 |  |  |  |  | 107
| 1 | 6 | 1.39 | 1.3508 | 1 | 1.7 | 1.57 | 1.1 | < 10 | 1.3 | 1.2 | 1.0968 | 1.8 | 1.2 | 1.4 | 1.2 | 1.4
| Mean |  | 1.4033 | 1.3665 | 2 | 1.6667 | 1.7283 | 1.1333 |  | 1.2833 | 1.1667 | 1.2537 | 1.75 | 1.2167 | 1.45 | 1.2 | 1.35
| Median |  | 1.39 | 1.3756 | 2 | 1.7 | 1.62 | 1.1 |  | 1.3 | 1.2 | 1.2145 | 1.75 | 1.2 | 1.45 | 1.2 | 1.35
| Std Dev. |  | 0.031411 | 0.027001 | 0.8165 | 0.05164 | 0.34266 | 0.05164 |  | 0.040825 | 0.05164 | 0.13552 | 0.054772 | 0.040825 | 0.054772 | 0 | 0.054772
| Rel.Std.Dev. |  | 2.238% | 1.976% | 40.82% | 3.098% | 19.83% | 4.556% |  | 3.181% | 4.426% | 10.81% | 3.13% | 3.355% | 3.777% | 0% | 4.057%
| PDM3 |  | 2.382% | -0.3028% | 45.91% | 21.59% | 26.09% | -17.32% |  | -6.373% | -14.88% | -8.536% | 27.67% | -11.24% | 5.787% | -12.45% | -1.509%
| Z-Score (Absolute) |  | 0.2 | 0.04 | NA | 1.39 | 1.66 | 1.01 | 16.38 | 0.34 | 0.86 | 0.47 | 1.76 | 0.64 | 0.41 | 0.71 | 0.04
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 255c (Certified Value 52 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  | 0
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 55.34 | 51.6677 | 214 | 51.1 | 66.5 | 54 | 54.9233 | 52.7 | 48.9 | 50.5 | 50.7817 | 53.5 | 49.8 | 55.7 | 50.3 | 39.9 |  |  |  | 1
| 1 | 2 | 54.02 | 52.866 | 219 | 52.5 | 67.3 | 53 | 54.3533 | 53.1 | 44.4 | 51.9 | 49.9129 | 55 | 53.7 | 56 | 52.1 | 39.9 |  |  |  | 21
| 1 | 3 | 53.96 | 53.6669 | 209 | 52.1 | 63.4 | 53 | 54.9767 | 51.7 | 50.7 | 54.9 | 50.8048 | 54.1 | 52.1 | 54.5 | 49.7 | 40.2 |  |  |  | 16
| 1 | 4 | 53.12 | 52.3951 | 202 | 51.9 | 66.6 | 53 | 55.7933 | 52.6 | 46.8 | 51.4 | 50.7976 | 53 | 51.9 | 55.9 | 52.1 | 40.5 |  |  |  | 52.4683
| 1 | 5 | 55.02 | 53.3572 | 223 | 50.5 | 60.7 | 53 | 54.3633 | 52.3 | 47.3 | 50.3 | 50.9542 | 53.3 | 49.7 | 56.8 | 51 | 39.5 |  |  |  | 108
| 1 | 6 | 53.61 | 51.6328 | 217 | 50.3 | 61.3 | 52 | 53.9133 | 55.6 | 52.2 | 54.6 | 50.0366 | 54.3 | 50.1 | 54.5 | 51.2 | 39.8
| Mean |  | 54.1783 | 52.5976 | 214 | 51.4 | 64.3 | 53 | 54.7206 | 53 | 48.3833 | 52.2667 | 50.548 | 53.8667 | 51.2167 | 55.5667 | 51.0667 | 39.9667
| Median |  | 53.99 | 52.6305 | 215.5 | 51.5 | 64.95 | 53 | 54.6433 | 52.65 | 48.1 | 51.65 | 50.7897 | 53.8 | 51 | 55.8 | 51.1 | 39.9
| Std Dev. |  | 0.84542 | 0.85194 | 7.5366 | 0.90111 | 2.8948 | 0.63246 | 0.65903 | 1.3565 | 2.8195 | 2.0126 | 0.45004 | 0.73937 | 1.6105 | 0.90701 | 0.96056 | 0.34448
| Rel.Std.Dev. |  | 1.56% | 1.62% | 3.522% | 1.753% | 4.502% | 1.193% | 1.204% | 2.559% | 5.827% | 3.851% | 0.8903% | 1.373% | 3.144% | 1.632% | 1.881% | 0.8619%
| PDM3 |  | 3.259% | 0.2465% | 307.9% | -2.036% | 22.55% | 1.013% | 4.293% | 1.013% | -7.786% | -0.3843% | -3.66% | 2.665% | -2.385% | 5.905% | -2.671% | -23.83%
| Z-Score (Absolute) |  | 0.56 | 0.08 | 65.59 | 0.57 | 4.68 | 0.08 | 0.78 | 0.08 | 1.8 | 0.22 | 0.92 | 0.43 | 0.64 | 1.13 | 0.7 | 5.22
| Analytical results for Ta in OREAS 255c (Certified Value < 0.01 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 22
| 1 | 3 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01 | [7 empty] | 109
| 1 | 6 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 5 | < 0.01 | < 0.01 | < 0.05 | 0.3 | < 0.01 | 0.01 | < 0.01
| Mean |  | 1.11: 0.01; 1.12: 0.01; 1.14: 0.3; 1.16: 0.01
| Median |  | 1.11: 0.01; 1.12: 0.01; 1.14: 0.3; 1.16: 0.01
| Std Dev. |  | 1.11: 0; 1.14: 0; 1.16: 0
Rel.Std.Dev. | bar: 1.11; 1.14; 1.16
PDM3
| Z-Score (Absolute) |  | 0.67 | 3.37 | 3.37 | 0.67 | 504.05 | 0.34 | 0.51 | 3.37 | 59 | 0.67 | 0.34 | 0.67
| Analytical results for Tb in OREAS 255c (Certified Value 0.45 ppm) | [20 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09; 1.09: 1.17; 1.10: 1.22 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g; 1.09: 0.2g; 1.10: 0.25g |  |  |  |  | 3
| 1 | 1 | 1.01: 0.45; 1.03: 0.42108; 1.07: 0.5; 1.09: 0.5; 1.10: 0.48 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.434; 1.03: 0.44151; 1.07: 0.4; 1.09: 0.5; 1.10: 0.47 |  |  |  |  | 23
| 1 | 3 | 1.01: 0.42; 1.03: 0.44441; 1.07: 0.4; 1.09: 0.5; 1.10: 0.48 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.432; 1.03: 0.43008; 1.07: 0.4; 1.09: 0.5; 1.10: 0.49 |  |  |  |  | 0.45396
| 1 | 5 | 1.01: 0.437; 1.03: 0.45616; 1.07: 0.4; 1.09: 0.5; 1.10: 0.49 |  |  |  |  | 110
| 1 | 6 | 1.01: 0.435; 1.03: 0.4277; 1.07: 0.4; 1.09: 0.5; 1.10: 0.48
| Mean |  | 1.01: 0.43467; 1.03: 0.43682; 1.07: 0.41667; 1.09: 0.5; 1.10: 0.48167
| Median |  | 1.01: 0.4345; 1.03: 0.43579; 1.07: 0.4; 1.09: 0.5; 1.10: 0.48
| Std Dev. |  | 1.01: 0.0096264; 1.03: 0.012884; 1.07: 0.040825; 1.09: 0; 1.10: 0.0075277
| Rel.Std.Dev. |  | 1.01: 2.215%; 1.03: 2.949%; 1.07: 9.798%; 1.09: 0%; 1.10: 1.563%
| PDM3 |  | 1.01: -4.251%; 1.03: -3.776%; 1.07: -8.216%; 1.09: 10.14%; 1.10: 6.102%
| Z-Score (Absolute) |  | 1.01: 0.07; 1.03: 0; 1.07: 0.67; 1.09: 2.11; 1.10: 1.5
| Analytical results for Te in OREAS 255c (Certified Value 0.12 ppm) | [20 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.1 | 0.12869 | 0.1 | 0.13 | 0.15 | < 5 | 0.13 | < 0.2 | 0.13 | 0.09 | < 0.05 | 0.12 | 0.1 | 0.11 | 0.1 |  |  |  |  | 1
| 1 | 2 | 0.1 | 0.12845 | 0.08 | 0.13 | 0.14 | < 5 | 0.13 | < 0.2 | 0.12 | 0.11 | < 0.05 | 0.17 | 0.12 | 0.12 | 0.1 |  |  |  |  | 24
| 1 | 3 | 0.1 | 0.15193 | 0.07 | 0.13 | 0.15 | < 5 | 0.11 | < 0.2 | 0.113 | 0.11 | < 0.05 | 0.15 | 0.1 | 0.13 | 0.11 |  |  |  |  | 16
| 1 | 4 | 0.2 | 0.10774 | 0.1 | 0.09 | 0.15 | < 5 | 0.13 | < 0.2 | 0.138 | 0.11 | < 0.05 | 0.11 | 0.1 | 0.09 | 0.09 |  |  |  |  | 0.11747
| 1 | 5 | 0.2 | 0.11979 | 0.06 | 0.14 | 0.15 | < 5 | 0.14 | < 0.2 | 0.092 | 0.1 | < 0.05 | 0.14 | 0.1 | 0.07 | 0.1 |  |  |  |  | 111
| 1 | 6 | 0.1 | 0.13264 | 0.07 | 0.15 | 0.14 | < 5 | 0.12 | < 0.2 | 0.115 | 0.11 | < 0.05 | 0.17 | 0.1 | 0.11 | 0.1
| Mean |  | 0.13333 | 0.12821 | 0.08 | 0.12833 | 0.14667 |  | 0.12667 |  | 0.118 | 0.105 |  | 0.14333 | 0.10333 | 0.105 | 0.1
| Median |  | 0.1 | 0.12857 | 0.075 | 0.13 | 0.15 |  | 0.13 |  | 0.1175 | 0.11 |  | 0.145 | 0.1 | 0.11 | 0.1
| Std Dev. |  | 0.05164 | 0.014646 | 0.016733 | 0.020412 | 0.005164 |  | 0.010328 |  | 0.015862 | 0.0083666 |  | 0.025033 | 0.008165 | 0.021679 | 0.0063246
| Rel.Std.Dev. |  | 38.73% | 11.42% | 20.92% | 15.91% | 3.521% |  | 8.154% |  | 13.44% | 7.968% |  | 17.47% | 7.902% | 20.65% | 6.325%
| PDM3 |  | 13.5% | 9.137% | -31.9% | 9.245% | 24.85% |  | 7.826% |  | 0.4483% | -10.62% |  | 22.01% | -12.04% | -10.62% | -14.87%
| Z-Score (Absolute) |  | NA | 0.71 | 1.33 | 0.71 | 1.49 | 101.06 | 0.64 | 0.49 | 0.28 | 0.28 | 3.66 | 1.35 | 0.35 | 0.28 | 0.49
Results from laboratory 1.01 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 255c (Certified Value 6.67 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 7.23 | 6.4488 | < 100 | 6.42 | 7.3 | 5.1 | 6.5 | 6.9 | 5.5 | 7.4083 | 6.4 | 6.3 | 7.4 | 6.5 | 6.62 |  |  |  |  | 1
| 1 | 2 | 7.01 | 6.6375 | < 100 | 6.59 | 7.4 | 5.1 | 6.4 | 6.2 | 5.6 | 7.6384 | 6.39 | 6.8 | 7.4 | 6.7 | 6.41 |  |  |  |  | 25
| 1 | 3 | 6.93 | 6.7388 | < 100 | 6.55 | 7.1 | 5.1 | 6.2 | 6.7 | 5.6 | 7.9273 | 6.37 | 6.3 | 7.3 | 6.4 | 6.4 |  |  |  |  | 16
| 1 | 4 | 6.87 | 6.7088 | < 100 | 6.69 | 7.1 | 5.2 | 6.2 | 6.8 | 5.4 | 7.6922 | 6.22 | 6.4 | 7.2 | 6.8 | 6.55 |  |  |  |  | 6.6701
| 1 | 5 | 6.9 | 6.7626 | < 100 | 6.64 | 6.9 | 5.2 | 6.2 | 6.5 | 5.6 | 7.7654 | 6.32 | 6.5 | 7.3 | 6.6 | 6.46 |  |  |  |  | 112
| 1 | 6 | 6.94 | 6.3294 | < 100 | 6.59 | 7.1 | 5.2 | 6.4 | 6.7 | 5.8 | 7.3175 | 6.61 | 6.4 | 7.2 | 6.6 | 6.51
| Mean |  | 6.98 | 6.6043 |  | 6.58 | 7.15 | 5.15 | 6.3167 | 6.6333 | 5.5833 | 7.6248 | 6.385 | 6.45 | 7.3 | 6.6 | 6.4917
| Median |  | 6.935 | 6.6732 |  | 6.59 | 7.1 | 5.15 | 6.3 | 6.7 | 5.6 | 7.6653 | 6.38 | 6.4 | 7.3 | 6.6 | 6.485
| Std Dev. |  | 0.13115 | 0.17604 |  | 0.092087 | 0.17607 | 0.054772 | 0.13292 | 0.25033 | 0.13292 | 0.22686 | 0.12849 | 0.18708 | 0.089443 | 0.14142 | 0.085186
| Rel.Std.Dev. |  | 1.879% | 2.666% |  | 1.399% | 2.462% | 1.064% | 2.104% | 3.774% | 2.381% | 2.975% | 2.012% | 2.901% | 1.225% | 2.143% | 1.312%
| PDM3 |  | 4.646% | -0.9861% |  | -1.351% | 7.195% | -22.79% | -5.299% | -0.5511% | -16.29% | 14.31% | -4.274% | -3.3% | 9.444% | -1.051% | -2.675%
| Z-Score (Absolute) |  | 0.9 | 0.01 | 103.29 | 0.05 | 1.31 | 3.45 | 0.67 | 0.08 | 2.42 | 2.44 | 0.51 | 0.36 | 1.67 | 0 | 0.26
| Analytical results for Ti in OREAS 255c (Certified Value 0.114 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  | 3
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.1256 | 0.11764 | 0.087133 | 0.13 | 0.14 | 0.15 | 0.1 | 0.11 | 0.115 | 0.114 | 0.113 | 0.18618 | 0.15 | 0.124 | 0.118 | 0.113 | 0.089 |  |  | 1
| 1 | 2 | 0.1205 | 0.11789 | 0.093256 | 0.13 | 0.143 | 0.15 | 0.102 | 0.11 | 0.119 | 0.103 | 0.112 | 0.18139 | 0.15 | 0.13 | 0.12 | 0.118 | 0.09 |  |  | 26
| 1 | 3 | 0.1188 | 0.11889 | 0.089195 | 0.11 | 0.144 | 0.15 | 0.102 | 0.11 | 0.111 | 0.105 | 0.111 | 0.17947 | 0.15 | 0.13 | 0.117 | 0.113 | 0.091 |  |  | 16
| 1 | 4 | 0.1187 | 0.11581 | 0.087326 | 0.1 | 0.141 | 0.16 | 0.103 | 0.11 | 0.118 | 0.11 | 0.109 | 0.18915 | 0.15 | 0.127 | 0.124 | 0.119 | 0.091 |  |  | 0.11412
| 1 | 5 | 0.1267 | 0.11933 | 0.096908 | 0.14 | 0.145 | 0.15 | 0.1 | 0.11 | 0.117 | 0.111 | 0.111 | 0.17863 | 0.15 | 0.124 | 0.122 | 0.118 | 0.089 |  |  | 113
| 1 | 6 | 0.1215 | 0.11949 | 0.089363 | 0.12 | 0.142 | 0.15 | 0.099 | 0.11 | 0.12 | 0.123 | 0.114 | 0.17367 | 0.15 | 0.122 | 0.121 | 0.115 | 0.09
| Mean |  | 0.12197 | 0.11818 | 0.09053 | 0.12167 | 0.1425 | 0.15167 | 0.101 | 0.11 | 0.11667 | 0.111 | 0.11167 | 0.18141 | 0.15 | 0.12617 | 0.12033 | 0.116 | 0.09
| Median |  | 0.121 | 0.11839 | 0.089279 | 0.125 | 0.1425 | 0.15 | 0.101 | 0.11 | 0.1175 | 0.1105 | 0.1115 | 0.18043 | 0.15 | 0.1255 | 0.1205 | 0.1165 | 0.09
| Std Dev. |  | 0.0034256 | 0.0013804 | 0.0038224 | 0.01472 | 0.0018708 | 0.0040825 | 0.0015492 | 0 | 0.003266 | 0.0071274 | 0.0017512 | 0.0055511 | 0 | 0.0033714 | 0.002582 | 0.0026833 | 0.00089443
| Rel.Std.Dev. |  | 2.809% | 1.168% | 4.222% | 12.1% | 1.313% | 2.692% | 1.534% | 0% | 2.799% | 6.421% | 1.568% | 3.06% | 0% | 2.672% | 2.146% | 2.313% | 0.9938%
| PDM3 |  | 6.876% | 3.554% | -20.67% | 6.613% | 24.87% | 32.9% | -11.5% | -3.61% | 2.232% | -2.733% | -2.149% | 58.97% | 31.44% | 10.56% | 5.445% | 1.648% | -21.14%
| Z-Score (Absolute) |  | 0.32 | 0 | 2.33 | 0.29 | 2.05 | 2.83 | 1.45 | 0.69 | 0.13 | 0.61 | 0.55 | 5.34 | 2.69 | 0.67 | 0.18 | 0.18 | 2.38
| Analytical results for Tl in OREAS 255c (Certified Value 0.12 ppm) | [20 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.12 | 0.12129 | 0.14 | 0.14 | 0.07 | < 10 | 0.11 | < 0.1 | 0.116 | 0.1 | 0.13122 | 0.16 | 0.11 | 0.13 | 0.11 |  |  |  |  | 1
| 1 | 2 | 0.13 | 0.12553 | 0.14 | 0.15 | 0.08 | < 10 | 0.11 | < 0.1 | 0.117 | 0.11 | 0.12327 | 0.18 | 0.11 | 0.13 | 0.11 |  |  |  |  | 27
| 1 | 3 | 0.12 | 0.12263 | 0.14 | 0.14 | 0.08 | < 10 | 0.11 | < 0.1 | 0.121 | 0.1 | 0.13493 | 0.17 | 0.11 | 0.13 | 0.1 |  |  |  |  | 16
| 1 | 4 | 0.12 | 0.11722 | 0.13 | 0.15 | 0.07 | < 10 | 0.11 | < 0.1 | 0.117 | 0.1 | 0.12665 | 0.16 | 0.11 | 0.12 | 0.11 |  |  |  |  | 0.12453
| 1 | 5 | 0.13 | 0.12771 | 0.13 | 0.14 | 0.07 | < 10 | 0.11 | < 0.1 | 0.118 | 0.1 | 0.12191 | 0.17 | 0.11 | 0.13 | 0.11 |  |  |  |  | 114
| 1 | 6 | 0.12 | 0.12462 | 0.14 | 0.14 | 0.06 | < 10 | 0.12 | < 0.1 | 0.118 | 0.11 | 0.11596 | 0.17 | 0.11 | 0.11 | 0.11
| Mean |  | 0.12333 | 0.12317 | 0.13667 | 0.14333 | 0.071667 |  | 0.11167 |  | 0.11783 | 0.10333 | 0.12566 | 0.16833 | 0.11 | 0.125 | 0.10833
| Median |  | 0.12 | 0.12363 | 0.14 | 0.14 | 0.07 |  | 0.11 |  | 0.1175 | 0.1 | 0.12496 | 0.17 | 0.11 | 0.13 | 0.11
| Std Dev. |  | 0.005164 | 0.0036767 | 0.005164 | 0.005164 | 0.0075277 |  | 0.0040825 |  | 0.0017224 | 0.005164 | 0.0068067 | 0.0075277 | 0 | 0.0083666 | 0.0040825
| Rel.Std.Dev. |  | 4.187% | 2.985% | 3.779% | 3.603% | 10.5% |  | 3.656% |  | 1.462% | 4.997% | 5.417% | 4.472% | 0% | 6.693% | 3.768%
| PDM3 |  | -0.9587% | -1.092% | 9.748% | 15.1% | -42.45% |  | -10.33% |  | -5.375% | -17.02% | 0.9089% | 35.18% | -11.67% | 0.3797% | -13%
| Z-Score (Absolute) |  | 0.01 | 0 | 0.67 | 1.01 | 2.57 | 243.6 | 0.57 | 3.65 | 0.27 | 0.99 | 0.12 | 2.26 | 0.66 | 0.09 | 0.74
| Analytical results for Tm in OREAS 255c (Indicative Value 0.11 ppm) | [20 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g |  |  |  |  | 2
| 1 | 1 | 1.01: 0.128; 1.03: 0.10104; 1.07: 0.1 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.122; 1.03: 0.10569; 1.07: 0.1 |  |  |  |  | 6
| 1 | 3 | 1.01: 0.127; 1.03: 0.11553; 1.07: 0.1 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.116; 1.03: 0.10617; 1.07: 0.1 |  |  |  |  | 0.10957
| 1 | 5 | 1.01: 0.123; 1.03: 0.10628; 1.07: 0.1 |  |  |  |  | 12
| 1 | 6 | 1.01: 0.117; 1.03: 0.10449; 1.07: 0.1
| Mean |  | 1.01: 0.12217; 1.03: 0.10653; 1.07: 0.1
| Median |  | 1.01: 0.1225; 1.03: 0.10593; 1.07: 0.1
| Std Dev. |  | 1.01: 0.0049565; 1.03: 0.0048228; 1.07: 1.5202e-17
| Rel.Std.Dev. |  | 1.01: 4.057%; 1.03: 4.527%; 1.07: 1.52e-14%
| PDM3 |  | 1.01: 11.5%; 1.03: -2.769%; 1.07: -8.731%
| Z-Score (Absolute) |  | 1.01: 1.61; 1.03: 0; 1.07: 0.67
| Analytical results for U in OREAS 255c (Certified Value 1.01 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 1.06 | 1.0349 | 0.97 | 1 | 0.92 | 0.98 | 1 | 1.02 | 0.82 | 1.0414 | 1.15 | 0.89 | 1.17 | 0.95 | 1.02 |  |  |  |  | 1
| 1 | 2 | 1.02 | 1.003 | 0.97 | 1.1 | 0.94 | 0.99 | 0.9 | 1.043 | 0.85 | 1.0373 | 1.15 | 0.97 | 1.05 | 0.94 | 1.04 |  |  |  |  | 29
| 1 | 3 | 1 | 1.0116 | 0.96 | 1 | 0.93 | 0.95 | 1 | 1.065 | 0.85 | 1.0283 | 1.14 | 0.91 | 1.07 | 0.89 | 1.04 |  |  |  |  | 16
| 1 | 4 | 1 | 0.98182 | 1 | 1 | 0.94 | 0.99 | 1 | 1.041 | 0.82 | 1.046 | 1.14 | 0.91 | 1.07 | 1.02 | 1.01 |  |  |  |  | 1.0067
| 1 | 5 | 1 | 1.0291 | 1.01 | 0.9 | 0.93 | 0.97 | 1 | 1.042 | 0.85 | 1.0515 | 1.14 | 0.91 | 1.04 | 0.92 | 1.04 |  |  |  |  | 115
| 1 | 6 | 0.99 | 0.98596 | 1 | 1 | 0.94 | 1 | 1 | 1.045 | 0.85 | 1.0341 | 1.18 | 0.91 | 1.05 | 0.94 | 1.06
| Mean |  | 1.0117 | 1.0077 | 0.985 | 1 | 0.93333 | 0.98 | 0.98333 | 1.0427 | 0.84 | 1.0398 | 1.15 | 0.91667 | 1.075 | 0.94333 | 1.035
| Median |  | 1 | 1.0073 | 0.985 | 1 | 0.935 | 0.985 | 1 | 1.0425 | 0.85 | 1.0393 | 1.145 | 0.91 | 1.06 | 0.94 | 1.04
| Std Dev. |  | 0.025626 | 0.021819 | 0.020736 | 0.063246 | 0.008165 | 0.017889 | 0.040825 | 0.014292 | 0.015492 | 0.0083421 | 0.015492 | 0.027325 | 0.048062 | 0.043205 | 0.017607
| Rel.Std.Dev. |  | 2.533% | 2.165% | 2.105% | 6.325% | 0.8748% | 1.825% | 4.152% | 1.371% | 1.844% | 0.8023% | 1.347% | 2.981% | 4.471% | 4.58% | 1.701%
| PDM3 |  | 0.4897% | 0.1001% | -2.159% | -0.6692% | -7.291% | -2.656% | -2.325% | 3.569% | -16.56% | 3.281% | 14.23% | -8.947% | 6.781% | -6.298% | 2.807%
| Z-Score (Absolute) |  | 0.08 | 0 | 0.44 | NA | 1.44 | 0.54 | NA | 0.67 | 3.24 | 0.62 | 2.75 | 1.76 | 1.3 | 1.24 | 0.53
Results from laboratories 1.09 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 255c (Certified Value 44.8 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 25g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 48 | 47.6338 | 40.924 | 42 | 49 | 41 | 41 | 47.7193 | 45 | 47 | 43 | 47.0728 | 48 | 45 | 44 | 43 |  |  |  | 1
| 1 | 2 | 47 | 48.6852 | 41.518 | 43 | 50 | 41 | 40 | 47.4714 | 46 | 43 | 43 | 44.8029 | 49 | 47 | 45 | 45 |  |  |  | 30
| 1 | 3 | 46 | 45.2162 | 40.725 | 39 | 50 | 40 | 39 | 47.8107 | 44 | 46 | 43 | 46.4811 | 48 | 47 | 45 | 43 |  |  |  | 16
| 1 | 4 | 46 | 46.9448 | 40.363 | 38 | 50 | 41 | 40 | 48.098 | 46 | 45 | 42 | 46.8976 | 47 | 46 | 46 | 44 |  |  |  | 44.8323
| 1 | 5 | 48 | 45.6533 | 42.821 | 45 | 51 | 38 | 40 | 47.5358 | 45 | 46 | 43 | 46.293 | 48 | 45 | 46 | 44 |  |  |  | 116
| 1 | 6 | 47 | 46.9966 | 41.154 | 42 | 50 | 40 | 39 | 47.4463 | 46 | 50 | 43 | 44.5165 | 49 | 44 | 46 | 44
| Mean |  | 47 | 46.855 | 41.2508 | 41.5 | 50 | 40.1667 | 39.8333 | 47.6803 | 45.3333 | 46.1667 | 42.8333 | 46.0106 | 48.1667 | 45.6667 | 45.3333 | 43.8333
| Median |  | 47 | 46.9707 | 41.039 | 42 | 50 | 40.5 | 40 | 47.6275 | 45.5 | 46 | 43 | 46.3871 | 48 | 45.5 | 45.5 | 44
| Std Dev. |  | 0.89443 | 1.2739 | 0.86229 | 2.5884 | 0.63246 | 1.169 | 0.75277 | 0.2499 | 0.8165 | 2.3166 | 0.40825 | 1.0869 | 0.75277 | 1.2111 | 0.8165 | 0.75277
| Rel.Std.Dev. |  | 1.903% | 2.719% | 2.09% | 6.237% | 1.265% | 2.91% | 1.89% | 0.5241% | 1.801% | 5.018% | 0.9531% | 2.362% | 1.563% | 2.652% | 1.801% | 1.717%
| PDM3 |  | 4.835% | 4.512% | -7.988% | -7.433% | 11.53% | -10.41% | -11.15% | 6.353% | 1.118% | 2.976% | -4.459% | 2.628% | 7.438% | 1.861% | 1.118% | -2.228%
| Z-Score (Absolute) |  | 0.53 | 0.48 | 1.49 | 1.4 | 1.58 | 1.87 | 1.99 | 0.76 | 0.06 | 0.23 | 0.93 | 0.18 | 0.93 | 0.06 | 0.06 | 0.58
| Analytical results for W in OREAS 255c (Certified Value 10.1 ppm) | [20 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 10.97 | 11.1458 | 16 | 7.2 | 7.5 | < 20 | 17.84 | 11.2 | 9.3 | 11.35 | 10.9238 | 10.9 | 12.45 | 13 | 9.45 | 7.09 |  |  |  | 1
| 1 | 2 | 10.62 | 12.2103 | 15 | 7 | 6.8 | < 20 | 16.0467 | 11.15 | 9.1 | 11.95 | 10.5519 | 10.7 | 13 | 12.7 | 9.9 | 6.79 |  |  |  | 31
| 1 | 3 | 10.7 | 11.4759 | 15 | 7 | 6.7 | < 20 | 16.9167 | 10.85 | 9.3 | 12.1 | 11.4246 | 11.1 | 11.25 | 12.8 | 9.63 | 7.1 |  |  |  | 16
| 1 | 4 | 10.2 | 11.3486 | 16 | 7.5 | 6.9 | < 20 | 17.79 | 11.3 | 8.9 | 11.8 | 11.4728 | 11.1 | 11.85 | 13.05 | 9.79 | 6.97 |  |  |  | 10.1128
| 1 | 5 | 10.15 | 11.8906 | 15 | 7 | 7.6 | < 20 | 15.9967 | 11.05 | 8.8 | 10.95 | 10.3038 | 11.8 | 12.65 | 12.2 | 8.9 | 6.83 |  |  |  | 117
| 1 | 6 | 10.62 | 15.3592 | 17 | 7.1 | 6.7 | < 20 | 16.7233 | 10.65 | 8.5 | 11.85 | 10.4022 | 11.3 | 12.3 | 12.45 | 10.25 | 6.81
| Mean |  | 10.5433 | 12.2384 | 15.6667 | 7.1333 | 7.0333 |  | 16.8856 | 11.0333 | 8.9833 | 11.6667 | 10.8465 | 11.15 | 12.25 | 12.7 | 9.6533 | 6.9317
| Median |  | 10.62 | 11.6832 | 15.5 | 7.05 | 6.85 |  | 16.82 | 11.1 | 9 | 11.825 | 10.7378 | 11.1 | 12.375 | 12.75 | 9.71 | 6.9
| Std Dev. |  | 0.31335 | 1.5766 | 0.8165 | 0.19664 | 0.40825 |  | 0.80624 | 0.24221 | 0.31252 | 0.43205 | 0.51209 | 0.37815 | 0.62048 | 0.32711 | 0.45715 | 0.14148
| Rel.Std.Dev. |  | 2.972% | 12.88% | 5.212% | 2.757% | 5.804% |  | 4.775% | 2.195% | 3.479% | 3.703% | 4.721% | 3.392% | 5.065% | 2.576% | 4.736% | 2.041%
| PDM3 |  | 4.258% | 21.02% | 54.92% | -29.46% | -30.45% |  | 66.97% | 9.103% | -11.17% | 15.37% | 7.256% | 10.26% | 21.13% | 25.58% | -4.543% | -31.46%
| Z-Score (Absolute) |  | 0.15 | 0.67 | NA | 1.8 | 1.85 | 0.41 | 2.93 | 0.09 | 0.9 | 0.4 | 0 | 0.15 | 0.68 | 0.9 | 0.58 | 1.9
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 255c (Certified Value 9.09 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 9.47 | 9.1056 | 10 | 8.74 | 8.89 | 8.23 | 9.2 | 8.83 | 8.7 | 9.2367 | 9.5 | 8.82 | 9.87 | 8.94 | 8.9 |  |  |  |  | 1
| 1 | 2 | 9.43 | 9.3536 | 10 | 8.9 | 8.98 | 8.26 | 9.15 | 8.89 | 9.06 | 9.242 | 9.8 | 9.56 | 10 | 9.33 | 8.81 |  |  |  |  | 32
| 1 | 3 | 9.31 | 9.1929 | 9 | 8.95 | 8.82 | 8.32 | 9.1767 | 8.61 | 9.32 | 9.371 | 9.4 | 9.19 | 9.64 | 8.95 | 9.04 |  |  |  |  | 16
| 1 | 4 | 9.13 | 9.136 | 9 | 8.9 | 8.84 | 8.48 | 9.28 | 8.71 | 9.07 | 9.3694 | 9.3 | 9.13 | 9.81 | 9.17 | 8.97 |  |  |  |  | 9.0862
| 1 | 5 | 9.45 | 9.2575 | 10 | 8.75 | 8.31 | 8.2 | 9.1967 | 8.83 | 8.92 | 9.3673 | 9.5 | 8.91 | 10 | 8.92 | 8.93 |  |  |  |  | 118
| 1 | 6 | 9.21 | 9.1946 | 10 | 8.99 | 8.41 | 8.26 | 9.06 | 9.12 | 9.15 | 9.2394 | 9.7 | 8.92 | 9.86 | 9.02 | 8.78
| Mean |  | 9.3333 | 9.2067 | 9.6667 | 8.8717 | 8.7083 | 8.2917 | 9.1772 | 8.8317 | 9.0367 | 9.3043 | 9.5333 | 9.0883 | 9.8633 | 9.055 | 8.905
| Median |  | 9.37 | 9.1938 | 10 | 8.9 | 8.83 | 8.26 | 9.1867 | 8.83 | 9.065 | 9.3046 | 9.5 | 9.025 | 9.865 | 8.985 | 8.915
| Std Dev. |  | 0.14052 | 0.089138 | 0.5164 | 0.10381 | 0.27723 | 0.10048 | 0.072031 | 0.17348 | 0.21059 | 0.071173 | 0.18619 | 0.27081 | 0.13426 | 0.16282 | 0.097724
| Rel.Std.Dev. |  | 1.506% | 0.9682% | 5.342% | 1.17% | 3.184% | 1.212% | 0.7849% | 1.964% | 2.33% | 0.765% | 1.953% | 2.98% | 1.361% | 1.798% | 1.097%
| PDM3 |  | 2.72% | 1.327% | 6.389% | -2.361% | -4.159% | -8.744% | 1.002% | -2.801% | -0.545% | 2.4% | 4.921% | 0.02366% | 8.553% | -0.3432% | -1.994%
| Z-Score (Absolute) |  | 0.82 | 0.42 | NA | 0.62 | 1.13 | 2.43 | 0.33 | 0.75 | 0.11 | 0.73 | 1.44 | 0.05 | 2.47 | 0.05 | 0.52
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 255c (Certified Value 0.68 ppm) | [20 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.07: Lab; 1.09: Lab; 1.10: Lab |  |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.03; 1.07: 1.09; 1.09: 1.17; 1.10: 1.22 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.03: AR*MS; 1.07: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*MS |  |  |  |  | 2
|  |  | 1.01: 25g; 1.03: 0.2g; 1.07: 0.5g; 1.09: 0.2g; 1.10: 0.25g |  |  |  |  | 3
| 1 | 1 | 1.01: 0.7; 1.03: 0.62138; 1.07: 0.7; 1.09: 0.6; 1.10: 0.7 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.688; 1.03: 0.64296; 1.07: 0.7; 1.09: 0.8; 1.10: 0.7 |  |  |  |  | 33
| 1 | 3 | 1.01: 0.678; 1.03: 0.63419; 1.07: 0.7; 1.09: 0.7; 1.10: 0.7 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.691; 1.03: 0.62561; 1.07: 0.7; 1.09: 0.7; 1.10: 0.7 |  |  |  |  | 0.68075
| 1 | 5 | 1.01: 0.693; 1.03: 0.78094; 1.07: 0.6; 1.09: 0.7; 1.10: 0.7 |  |  |  |  | 119
| 1 | 6 | 1.01: 0.694; 1.03: 0.62472; 1.07: 0.6; 1.09: 0.7; 1.10: 0.8
| Mean |  | 1.01: 0.69067; 1.03: 0.65497; 1.07: 0.66667; 1.09: 0.7; 1.10: 0.71667
| Median |  | 1.01: 0.692; 1.03: 0.6299; 1.07: 0.7; 1.09: 0.7; 1.10: 0.7
| Std Dev. |  | 1.01: 0.0073666; 1.03: 0.06221; 1.07: 0.05164; 1.09: 0.063246; 1.10: 0.040825
| Rel.Std.Dev. |  | 1.01: 1.067%; 1.03: 9.498%; 1.07: 7.746%; 1.09: 9.035%; 1.10: 5.696%
| PDM3 |  | 1.01: 1.456%; 1.03: -3.788%; 1.07: -2.069%; 1.09: 2.827%; 1.10: 5.275%
| Z-Score (Absolute) |  | 1.01: 0; 1.03: 1; 1.07: 0.67; 1.09: 0.26; 1.10: 0.73
| Analytical results for Zn in OREAS 255c (Certified Value 59 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | 0
|  |  | 25g | 0.5g | 0.2g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 1
| 1 | 1 | 61 | 66.9276 | 60.8929 | 49.919 | 58 | 62 | 65.2 | 52 | 54.3733 | 60 | 60 | 55.954 | 57 | 62.598 | 66 | 60 | 58 | 56 | 51 | 1
| 1 | 2 | 60 | 67.5216 | 62.3378 | 51.351 | 54 | 62 | 65.8 | 53 | 53.8667 | 62 | 55 | 55.527 | 57 | 64.2274 | 68.2 | 62 | 59 | 58 | 50 | 34
| 1 | 3 | 61 | 66.42 | 59.979 | 50.597 | 49 | 63 | 65.3 | 53 | 54.65 | 59 | 60 | 55.219 | 57 | 63.2924 | 66.2 | 63 | 59 | 56 | 50 | 16
| 1 | 4 | 59 | 67.9428 | 61.1547 | 50.419 | 48 | 62 | 66.3 | 51 | 55.17 | 61 | 58 | 56.021 | 56 | 60.588 | 64.6 | 61 | 60 | 60 | 51 | 58.7129
| 1 | 5 | 61 | 67.284 | 60.0961 | 52.286 | 55 | 62 | 64.2 | 52 | 53.75 | 61 | 58 | 56.072 | 57 | 63.0974 | 65.7 | 60 | 59 | 57 | 49 | 120
| 1 | 6 | 60 | 67.8456 | 61.5518 | 50.966 | 54 | 62 | 64.1 | 53 | 53.5133 | 61 | 63 | 55.613 | 57 | 61.3466 | 67.3 | 59 | 60 | 56 | 54
| Mean |  | 60.3333 | 67.3236 | 61.002 | 50.923 | 53 | 62.1667 | 65.15 | 52.3333 | 54.2206 | 60.6667 | 59 | 55.7343 | 56.8333 | 62.5249 | 66.3333 | 60.8333 | 59.1667 | 57.1667 | 50.8333
| Median |  | 60.5 | 67.4028 | 61.0238 | 50.7815 | 54 | 62 | 65.25 | 52.5 | 54.12 | 61 | 59 | 55.7835 | 57 | 62.8477 | 66.1 | 60.5 | 59 | 56.5 | 50.5
| Std Dev. |  | 0.8165 | 0.57792 | 0.89318 | 0.82597 | 3.7947 | 0.40825 | 0.86891 | 0.8165 | 0.62546 | 1.0328 | 2.6833 | 0.33697 | 0.40825 | 1.3387 | 1.2612 | 1.472 | 0.75277 | 1.6021 | 1.7224
| Rel.Std.Dev. |  | 1.353% | 0.8584% | 1.464% | 1.622% | 7.16% | 0.6567% | 1.334% | 1.56% | 1.154% | 1.702% | 4.548% | 0.6046% | 0.7183% | 2.141% | 1.901% | 2.42% | 1.272% | 2.802% | 3.388%
| PDM3 |  | 2.76% | 14.67% | 3.899% | -13.27% | -9.73% | 5.882% | 10.96% | -10.87% | -7.651% | 3.328% | 0.489% | -5.073% | -3.201% | 6.493% | 12.98% | 3.612% | 0.7729% | -2.634% | -13.42%
| Z-Score (Absolute) |  | 0.23 | 1.64 | 0.37 | 1.66 | 1.24 | 0.6 | 1.2 | 1.37 | 0.99 | 0.3 | 0.03 | 0.69 | 0.47 | 0.67 | 1.44 | 0.33 | 0 | 0.4 | 1.67
| Analytical results for Zr in OREAS 255c (Certified Value 37.3 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 25g | 0.2g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 10.7 | 41.3899 | 93 | 36.1 | 8.2 | 29 | 35.8286 | 32.6 | 39.7 | 49.534 | 43.9 | 39.6 | 40.5 | 28.8 | 42.6 |  |  |  |  | 1
| 1 | 2 | 10 | 41.4983 | 105 | 36.8 | 6.4 | 29 | 35.3794 | 35.9 | 39.8 | 48.3695 | 44 | 36.5 | 35.6 | 24.1 | 42.6 |  |  |  |  | 35
| 1 | 3 | 8.9 | 41.0516 | 75 | 34.7 | 6 | 28 | 35.9188 | 38.4 | 36.3 | 50.265 | 44.2 | 30.4 | 36.7 | 22.9 | 42.9 |  |  |  |  | 16
| 1 | 4 | 10.8 | 40.2808 | 73 | 35.9 | 7.2 | 28 | 36.5674 | 39.5 | 40.3 | 47.6054 | 42.4 | 34.1 | 37.1 | 28.3 | 43.3 |  |  |  |  | 37.3243
| 1 | 5 | 11.5 | 42.1768 | 82 | 33.6 | 11.4 | 27 | 36.2976 | 40.6 | 36.3 | 49.4104 | 42.7 | 37.6 | 34.7 | 20 | 42.1 |  |  |  |  | 121
| 1 | 6 | 10 | 41.7999 | 86 | 36.2 | 7.1 | 27 | 35.7714 | 35.7 | 31.1 | 48.2452 | 43.3 | 34 | 36.7 | 32 | 43.2
| Mean |  | 10.3167 | 41.3662 | 85.6667 | 35.55 | 7.7167 | 28 | 35.9605 | 37.1167 | 37.25 | 48.9049 | 43.4167 | 35.3667 | 36.8833 | 26.0167 | 42.7833
| Median |  | 10.35 | 41.4441 | 84 | 36 | 7.15 | 28 | 35.8737 | 37.15 | 38 | 48.89 | 43.6 | 35.3 | 36.7 | 26.2 | 42.75
| Std Dev. |  | 0.89312 | 0.65424 | 11.961 | 1.1777 | 1.9559 | 0.89443 | 0.41824 | 2.9444 | 3.5053 | 0.99114 | 0.7414 | 3.2352 | 1.9803 | 4.435 | 0.4446
| Rel.Std.Dev. |  | 8.657% | 1.582% | 13.96% | 3.313% | 25.35% | 3.194% | 1.163% | 7.933% | 9.41% | 2.027% | 1.708% | 9.148% | 5.369% | 17.05% | 1.039%
| PDM3 |  | -72.36% | 10.83% | 129.5% | -4.754% | -79.33% | -24.98% | -3.654% | -0.5563% | -0.1991% | 31.03% | 16.32% | -5.245% | -1.181% | -30.3% | 14.63%
| Z-Score (Absolute) |  | 3.04 | 0.51 | 5.58 | 0.15 | 3.33 | 1.02 | 0.11 | 0.03 | 0.04 | 1.37 | 0.75 | 0.17 | 0 | 1.24 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C580: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C598: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C634: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C652: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C670: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C688: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C706: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C851: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C869: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C887: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C906: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C942: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C979: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C997: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1015: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1033: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1051: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 255c (Indicative Value 12.31 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.31 | 1
| 1 | 2 | 12.3 | 10
| Mean |  | 12.305 | 16
| Median |  | 12.305 | 12.305
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.05746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 255c (Indicative Value 1.7 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.69 | 1
| 1 | 2 | 1.71 | 11
| Mean |  | 1.7 | 16
| Median |  | 1.7 | 1.7
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.8319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 255c (Indicative Value 6.39 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.37 | 1
| 1 | 2 | 6.4 | 12
| Mean |  | 6.385 | 16
| Median |  | 6.385 | 6.385
| Std Dev. |  | 0.021213 | 18
| Rel.Std.Dev. |  | 0.3322%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 255c (Indicative Value 1.92 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.92 | 1
| 1 | 2 | 1.92 | 13
| Mean |  | 1.92 | 16
| Median |  | 1.92 | 1.92
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 255c (Indicative Value 2.79 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.78 | 1
| 1 | 2 | 2.79 | 10
| Mean |  | 2.785 | 16
| Median |  | 2.785 | 2.785
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.2539%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 255c (Indicative Value 0.056 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.055 | 1
| 1 | 2 | 0.057 | 11
| Mean |  | 0.056 | 16
| Median |  | 0.056 | 0.056
| Std Dev. |  | 0.0014142 | 17
| Rel.Std.Dev. |  | 2.525%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 255c (Indicative Value 1.13 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.12 | 1
| 1 | 2 | 1.13 | 12
| Mean |  | 1.125 | 16
| Median |  | 1.125 | 1.125
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.6285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 255c (Indicative Value 0.179 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.178 | 1
| 1 | 2 | 0.179 | 13
| Mean |  | 0.1785 | 16
| Median |  | 0.1785 | 0.1785
| Std Dev. |  | 0.00070711 | 19
| Rel.Std.Dev. |  | 0.3961%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 255c (Indicative Value 69.03 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69.06 | 1
| 1 | 2 | 68.99 | 10
| Mean |  | 69.025 | 16
| Median |  | 69.025 | 69.025
| Std Dev. |  | 0.049497 | 16
| Rel.Std.Dev. |  | 0.07171%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 255c (Indicative Value 0.092 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.091 | 1
| 1 | 2 | 0.093 | 11
| Mean |  | 0.092 | 16
| Median |  | 0.092 | 0.092
| Std Dev. |  | 0.0014142 | 17
| Rel.Std.Dev. |  | 1.537%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 255c (Indicative Value 0.932 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.93 | 1
| 1 | 2 | 0.934 | 12
| Mean |  | 0.932 | 16
| Median |  | 0.932 | 0.932
| Std Dev. |  | 0.0028284 | 18
| Rel.Std.Dev. |  | 0.3035%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 255c (Indicative Value 3.27 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.28 | 1
| 1 | 2 | 3.25 | 15
| Mean |  | 3.265 | 16
| Median |  | 3.265 | 3.265
| Std Dev. |  | 0.021213 | 21
| Rel.Std.Dev. |  | 0.6497%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 255c (Indicative Value 0.105 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.1 | 17
| Mean |  | 0.105 | 16
| Median |  | 0.105 | 0.105
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 255c (Indicative Value 0.02 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.02 | 17
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.02
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 255c (Indicative Value 1.15 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.1 | 1
| 1 | 2 | 1.2 | 19
| Mean |  | 1.15 | 16
| Median |  | 1.15 | 1.15
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 255c (Indicative Value 167 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 167 | 1
| 1 | 2 | 167 | 20
| Mean |  | 167 | 16
| Median |  | 167 | 167
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 255c (Indicative Value 443 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 447 | 1
| 1 | 2 | 438 | 21
| Mean |  | 442.5 | 16
| Median |  | 442.5 | 442.5
| Std Dev. |  | 6.364 | 27
| Rel.Std.Dev. |  | 1.438%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 255c (Indicative Value 2.2 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2 | 22
| Mean |  | 2.2 | 16
| Median |  | 2.2 | 2.2
| Std Dev. |  | 0.28284 | 28
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 255c (Indicative Value 3.34 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.36 | 1
| 1 | 2 | 3.32 | 23
| Mean |  | 3.34 | 16
| Median |  | 3.34 | 3.34
| Std Dev. |  | 0.028284 | 29
| Rel.Std.Dev. |  | 0.8468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 255c (Indicative Value 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.1 | 24
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.1
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 255c (Indicative Value 72 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72.1 | 1
| 1 | 2 | 70.9 | 25
| Mean |  | 71.5 | 16
| Median |  | 71.5 | 71.5
| Std Dev. |  | 0.84853 | 31
| Rel.Std.Dev. |  | 1.187%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 255c (Indicative Value 21.7 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.6 | 1
| 1 | 2 | 21.7 | 26
| Mean |  | 21.65 | 16
| Median |  | 21.65 | 21.65
| Std Dev. |  | 0.070711 | 32
| Rel.Std.Dev. |  | 0.3266%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 255c (Indicative Value 195 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 193 | 1
| 1 | 2 | 197 | 27
| Mean |  | 195 | 16
| Median |  | 195 | 195
| Std Dev. |  | 2.8284 | 33
| Rel.Std.Dev. |  | 1.45%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 255c (Indicative Value 3.76 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.79 | 1
| 1 | 2 | 3.73 | 28
| Mean |  | 3.76 | 16
| Median |  | 3.76 | 3.76
| Std Dev. |  | 0.042426 | 34
| Rel.Std.Dev. |  | 1.128%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 255c (Indicative Value 63 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 66 | 1
| 1 | 2 | 60 | 29
| Mean |  | 63 | 16
| Median |  | 63 | 63
| Std Dev. |  | 4.2426 | 35
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 255c (Indicative Value 4.47 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.5 | 1
| 1 | 2 | 4.44 | 3
| Mean |  | 4.47 | 16
| Median |  | 4.47 | 4.47
| Std Dev. |  | 0.042426 | 36
| Rel.Std.Dev. |  | 0.9491%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 255c (Indicative Value 2.23 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.25 | 1
| 1 | 2 | 2.21 | 4
| Mean |  | 2.23 | 16
| Median |  | 2.23 | 2.23
| Std Dev. |  | 0.028284 | 37
| Rel.Std.Dev. |  | 1.268%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 255c (Indicative Value 1.52 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.55 | 1
| 1 | 2 | 1.49 | 5
| Mean |  | 1.52 | 16
| Median |  | 1.52 | 1.52
| Std Dev. |  | 0.042426 | 38
| Rel.Std.Dev. |  | 2.791%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 255c (Indicative Value 17.5 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.6 | 1
| 1 | 2 | 17.4 | 33
| Mean |  | 17.5 | 16
| Median |  | 17.5 | 17.5
| Std Dev. |  | 0.14142 | 39
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 255c (Indicative Value 5.29 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.3 | 1
| 1 | 2 | 5.27 | 6
| Mean |  | 5.285 | 16
| Median |  | 5.285 | 5.285
| Std Dev. |  | 0.021213 | 40
| Rel.Std.Dev. |  | 0.4014%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 255c (Indicative Value 1.33 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.3 | 1
| Mean |  | 1.325 | 16
| Median |  | 1.325 | 1.325
| Std Dev. |  | 0.035355 | 41
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 255c (Indicative Value 6.43 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.53 | 1
| 1 | 2 | 6.32 | 7
| Mean |  | 6.425 | 16
| Median |  | 6.425 | 6.425
| Std Dev. |  | 0.14849 | 25
| Rel.Std.Dev. |  | 2.311%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 255c (Indicative Value 0.87 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.87 | 1
| 1 | 2 | 0.87 | 3
| Mean |  | 0.87 | 16
| Median |  | 0.87 | 0.87
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 255c (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 21
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 255c (Indicative Value 40 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.2 | 1
| 1 | 2 | 39.7 | 22
| Mean |  | 39.95 | 16
| Median |  | 39.95 | 39.95
| Std Dev. |  | 0.35355 | 28
| Rel.Std.Dev. |  | 0.885%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 255c (Indicative Value 0.3 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.3 | 4
| Mean |  | 0.295 | 16
| Median |  | 0.295 | 0.295
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 255c (Indicative Value 0.044 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0442 | 1
| 1 | 2 | 0.044 | 24
| Mean |  | 0.0441 | 16
| Median |  | 0.0441 | 0.0441
| Std Dev. |  | 0.00014142 | 30
| Rel.Std.Dev. |  | 0.3207%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 255c (Indicative Value 4.8 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 4.8 | 25
| Mean |  | 4.8 | 16
| Median |  | 4.8 | 4.8
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 255c (Indicative Value 28.6 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28.5 | 1
| 1 | 2 | 28.6 | 5
| Mean |  | 28.55 | 16
| Median |  | 28.55 | 28.55
| Std Dev. |  | 0.070711 | 32
| Rel.Std.Dev. |  | 0.2477%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 255c (Indicative Value 33.7 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.5 | 1
| 1 | 2 | 33.8 | 6
| Mean |  | 33.65 | 16
| Median |  | 33.65 | 33.65
| Std Dev. |  | 0.21213 | 33
| Rel.Std.Dev. |  | 0.6304%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 255c (Indicative Value 103 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 106 | 1
| 1 | 2 | 100 | 28
| Mean |  | 103 | 16
| Median |  | 103 | 103
| Std Dev. |  | 4.2426 | 34
| Rel.Std.Dev. |  | 4.119%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 255c (Indicative Value 16 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16 | 1
| 1 | 2 | 16 | 29
| Mean |  | 16 | 16
| Median |  | 16 | 16
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 255c (Indicative Value 9.15 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.28 | 1
| 1 | 2 | 9.02 | 8
| Mean |  | 9.15 | 16
| Median |  | 9.15 | 9.15
| Std Dev. |  | 0.18385 | 36
| Rel.Std.Dev. |  | 2.009%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 255c (Indicative Value 78 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 78 | 1
| 1 | 2 | 77.8 | 31
| Mean |  | 77.9 | 16
| Median |  | 77.9 | 77.9
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 0.1815%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 255c (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 32
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 38
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 255c (Indicative Value 8.55 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.6 | 1
| 1 | 2 | 8.5 | 33
| Mean |  | 8.55 | 16
| Median |  | 8.55 | 8.55
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.827%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 255c (Indicative Value 14.4 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.2 | 1
| 1 | 2 | 14.5 | 34
| Mean |  | 14.35 | 16
| Median |  | 14.35 | 14.35
| Std Dev. |  | 0.21213 | 40
| Rel.Std.Dev. |  | 1.478%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 255c (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 35
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 41
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 255c (Indicative Value 6.41 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.46 | 1
| 1 | 2 | 6.36 | 5
| Mean |  | 6.41 | 16
| Median |  | 6.41 | 6.41
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 1.103%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 255c (Indicative Value 7.7 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.2 | 1
| 1 | 2 | 8.2 | 20
| Mean |  | 7.7 | 16
| Median |  | 7.7 | 7.7
| Std Dev. |  | 0.70711 | 26
| Rel.Std.Dev. |  | 9.183%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 255c (Indicative Value 237 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 236 | 1
| 1 | 2 | 238 | 21
| Mean |  | 237 | 16
| Median |  | 237 | 237
| Std Dev. |  | 1.4142 | 27
| Rel.Std.Dev. |  | 0.5967%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 255c (Indicative Value 1.98 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.97 | 1
| 1 | 2 | 1.99 | 22
| Mean |  | 1.98 | 16
| Median |  | 1.98 | 1.98
| Std Dev. |  | 0.014142 | 28
| Rel.Std.Dev. |  | 0.7142%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 255c (Indicative Value 0.82 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.82 | 1
| 1 | 2 | 0.82 | 23
| Mean |  | 0.82 | 16
| Median |  | 0.82 | 0.82
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 255c (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 24
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 255c (Indicative Value 10.2 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.2 | 1
| 1 | 2 | 10.2 | 25
| Mean |  | 10.2 | 16
| Median |  | 10.2 | 10.2
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 255c (Indicative Value 0.571 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.572 | 1
| 1 | 2 | 0.57 | 26
| Mean |  | 0.571 | 16
| Median |  | 0.571 | 0.571
| Std Dev. |  | 0.0014142 | 32
| Rel.Std.Dev. |  | 0.2477%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 255c (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 27
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 255c (Indicative Value 0.32 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.32 | 1
| 1 | 2 | 0.32 | 6
| Mean |  | 0.32 | 16
| Median |  | 0.32 | 0.32
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 255c (Indicative Value 1.97 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.98 | 1
| 1 | 2 | 1.96 | 29
| Mean |  | 1.97 | 16
| Median |  | 1.97 | 1.97
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 0.7179%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 255c (Indicative Value 104 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 104 | 1
| 1 | 2 | 104 | 30
| Mean |  | 104 | 16
| Median |  | 104 | 104
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 255c (Indicative Value 35.3 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36 | 1
| 1 | 2 | 34.5 | 31
| Mean |  | 35.25 | 16
| Median |  | 35.25 | 35.25
| Std Dev. |  | 1.0607 | 37
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 255c (Indicative Value 21.9 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.8 | 1
| 1 | 2 | 21.9 | 32
| Mean |  | 21.85 | 16
| Median |  | 21.85 | 21.85
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.3236%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 255c (Indicative Value 2.1 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.13 | 1
| 1 | 2 | 2.06 | 33
| Mean |  | 2.095 | 16
| Median |  | 2.095 | 2.095
| Std Dev. |  | 0.049497 | 39
| Rel.Std.Dev. |  | 2.363%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 255c (Indicative Value 85 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 85 | 1
| 1 | 2 | 85 | 34
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 255c (Indicative Value 247 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 248 | 1
| 1 | 2 | 245 | 35
| Mean |  | 246.5 | 16
| Median |  | 246.5 | 246.5
| Std Dev. |  | 2.1213 | 41
| Rel.Std.Dev. |  | 0.8606%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.