## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 606b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.38554 | 0.37991 | 0.39116 | 0.38365 | 0.38742
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.37491 | 0.3617 | 0.38813 | 0.3729 | 0.37692
Infrared Combustion
S, Sulphur (wt.%) | 0.45984 | 0.44674 | 0.47294 | 0.44585 | 0.47383
4-Acid Digestion
Ag, Silver (ppm) | 1.5274 | 1.4387 | 1.6161 | 1.4639 | 1.5909
Al, Aluminium (wt.%) | 7.5356 | 7.2862 | 7.7849 | 7.3866 | 7.6846
As, Arsenic (ppm) | 89.3478 | 85.2426 | 93.4531 | 86.2284 | 92.4672
Ba, Barium (ppm) | 2891.6874 | 2796.7412 | 2986.6335 | 2820.7938 | 2962.581
Be, Beryllium (ppm) | 2.9785 | 2.8102 | 3.1468 | 2.8817 | 3.0752
Bi, Bismuth (ppm) | 4.8283 | 4.5998 | 5.0569 | 4.7175 | 4.9391
Ca, Calcium (wt.%) | 1.1 | 1.0674 | 1.1327 | 1.0744 | 1.1257
Cd, Cadmium (ppm) | 1.2405 | 1.1277 | 1.3532 | 1.1636 | 1.3174
Ce, Cerium (ppm) | 88.8532 | 84.7141 | 92.9924 | 86.7423 | 90.9642
Co, Cobalt (ppm) | 2.2525 | 2.1258 | 2.3791 | 2.1306 | 2.3743
Cr, Chromium (ppm) | 14.7162 | 12.9443 | 16.4882 | 13.2344 | 16.1981
Cs, Caesium (ppm) | 6.112 | 5.7875 | 6.4365 | 5.9779 | 6.2461
Cu, Copper (wt.%) | 0.033391 | 0.03235 | 0.034433 | 0.032658 | 0.034125
Dy, Dysprosium (ppm) | 3.5543 | 3.38 | 3.7285 | 3.3758 | 3.7327
Er, Erbium (ppm) | 1.007 | 0.934 | 1.08 | 0.9632 | 1.0508
Eu, Europium (ppm) | 1.3977 | 1.2633 | 1.5321 | 1.356 | 1.4394
Fe, Iron (wt.%) | 2.3276 | 2.2592 | 2.396 | 2.2682 | 2.387
Ga, Gallium (ppm) | 23.044 | 22.0416 | 24.0464 | 22.4863 | 23.6018
Gd, Gadolinium (ppm) | 5.6028 | 5.087 | 6.1187 | 5.4284 | 5.7772
Ge, Germanium (ppm) | 0.19297 | 0.12681 | 0.25913 | 0.16848 | 0.21746
Hf, Hafnium (ppm) | 6.4822 | 6.1653 | 6.7992 | 6.2516 | 6.7129
Ho, Holmium (ppm) | 0.48276 | 0.4367 | 0.52883 | 0.45462 | 0.51091
In, Indium (ppm) | 0.14224 | 0.13016 | 0.15432 | 0.13259 | 0.15189
K, Potassium (wt.%) | 3.0169 | 2.9207 | 3.1131 | 2.9446 | 3.0893
La, Lanthanum (ppm) | 44.9941 | 42.7599 | 47.2283 | 43.6058 | 46.3824
Li, Lithium (ppm) | 32.1337 | 30.8423 | 33.4252 | 31.3262 | 32.9412
Lu, Lutetium (ppm) | 0.074113 | 0.057508 | 0.090719 | IND | IND
Mg, Magnesium (wt.%) | 0.12667 | 0.11999 | 0.13334 | 0.12429 | 0.12904
Mn, Manganese (wt.%) | 0.028045 | 0.027165 | 0.028924 | 0.027515 | 0.028574
Mo, Molybdenum (ppm) | 4.3841 | 4.1197 | 4.6485 | 4.2097 | 4.5585
Na, Sodium (wt.%) | 2.6734 | 2.5883 | 2.7585 | 2.6109 | 2.7359
Nb, Niobium (ppm) | 18.4985 | 17.7371 | 19.26 | 17.9514 | 19.0456
Nd, Neodymium (ppm) | 37.4253 | 34.9975 | 39.853 | 36.3448 | 38.5058
Ni, Nickel (ppm) | 3.561 | 3.1307 | 3.9913 | 3.2886 | 3.8334
P, Phosphorus (wt.%) | 0.029991 | 0.028863 | 0.031119 | 0.029227 | 0.030754
Pb, Lead (ppm) | 157.0775 | 151.3194 | 162.8355 | 152.9966 | 161.1583
Pr, Praseodymium (ppm) | 10.0749 | 9.4068 | 10.743 | 9.7708 | 10.3789
Rb, Rubidium (ppm) | 139.0542 | 132.8853 | 145.2232 | 135.7017 | 142.4067
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.4683 | 0.45011 | 0.48649 | 0.45736 | 0.47924
Sb, Antimony (ppm) | 6.3769 | 6.0435 | 6.7103 | 6.0691 | 6.6847
Sc, Scandium (ppm) | 4.2592 | 3.9106 | 4.6078 | 4.1287 | 4.3896
Se, Selenium (ppm) | 1.7977 | 1.1128 | 2.4827 | 1.6704 | 1.925
Sm, Samarium (ppm) | 7.2675 | 6.8859 | 7.6492 | 7.0753 | 7.4597
Sn, Tin (ppm) | 3.7162 | 3.5376 | 3.8947 | 3.5938 | 3.8385
Sr, Strontium (ppm) | 199.5368 | 192.2829 | 206.7906 | 194.9335 | 204.1401
Ta, Tantalum (ppm) | 1.3643 | 1.2876 | 1.441 | 1.3187 | 1.4099
Tb, Terbium (ppm) | 0.73014 | 0.66365 | 0.79663 | 0.70941 | 0.75088
Te, Tellurium (ppm) | 0.83512 | 0.73456 | 0.93568 | 0.73247 | 0.93777
Th, Thorium (ppm) | 14.794 | 14.0997 | 15.4882 | 14.3812 | 15.2068
Ti, Titanium (wt.%) | 0.11614 | 0.11264 | 0.11964 | 0.11365 | 0.11863
Tl, Thallium (ppm) | 0.88616 | 0.84791 | 0.9244 | 0.86336 | 0.90895
Tm, Thulium (ppm) | 0.10423 | 0.095365 | 0.11309 | IND | IND
U, Uranium (ppm) | 5.5657 | 5.2769 | 5.8545 | 5.3921 | 5.7394
V, Vanadium (ppm) | 4.9654 | 4.481 | 5.4497 | 4.6455 | 5.2853
W, Tungsten (ppm) | 1.8512 | 1.7143 | 1.9881 | 1.7265 | 1.9759
Y, Yttrium (ppm) | 14.85 | 14.2047 | 15.4952 | 14.4468 | 15.2531
Yb, Ytterbium (ppm) | 0.60919 | 0.54022 | 0.67817 | 0.58278 | 0.63561
Zn, Zinc (ppm) | 315.9317 | 304.668 | 327.1954 | 308.2647 | 323.5988
Zr, Zirconium (ppm) | 242.9697 | 231.7677 | 254.1717 | 236.7893 | 249.1501
Aqua Regia Digestion
Ag, Silver (ppm) | 1.4853 | 1.403 | 1.5677 | 1.2488 | 1.7219
Al, Aluminium (wt.%) | 0.80363 | 0.77176 | 0.83549 | 0.78065 | 0.8266
As, Arsenic (ppm) | 85.8337 | 82.6112 | 89.0561 | 84.0868 | 87.5806
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 211.7273 | 198.3746 | 225.0799 | 204.1197 | 219.3348
Be, Beryllium (ppm) | 0.51276 | 0.47304 | 0.55248 | 0.48666 | 0.53885
Bi, Bismuth (ppm) | 4.9574 | 4.7263 | 5.1885 | 4.727 | 5.1877
Ca, Calcium (wt.%) | 0.63219 | 0.61433 | 0.65005 | 0.61883 | 0.64554
Cd, Cadmium (ppm) | 1.1946 | 1.106 | 1.2832 | 1.1312 | 1.2581
Ce, Cerium (ppm) | 56.153 | 53.1717 | 59.1342 | 54.7924 | 57.5136
Co, Cobalt (ppm) | 2.0231 | 1.9022 | 2.1441 | 1.9221 | 2.1242
Cr, Chromium (ppm) | 16.1418 | 15.0667 | 17.2169 | 15.1878 | 17.0957
Cs, Caesium (ppm) | 1.4778 | 1.3617 | 1.5939 | 1.4239 | 1.5316
Cu, Copper (wt.%) | 0.033727 | 0.032801 | 0.034652 | 0.033099 | 0.034354
Dy, Dysprosium (ppm) | 2.084 | 1.8794 | 2.2885 | 1.9856 | 2.1823
Er, Erbium (ppm) | 0.51123 | 0.44384 | 0.57862 | 0.4692 | 0.55326
Eu, Europium (ppm) | 0.69609 | 0.60292 | 0.78926 | 0.66719 | 0.725
Fe, Iron (wt.%) | 1.9992 | 1.9393 | 2.059 | 1.9665 | 2.0319
Ga, Gallium (ppm) | 5.8994 | 5.5883 | 6.2104 | 5.7005 | 6.0982
Gd, Gadolinium (ppm) | 3.8181 | 3.5522 | 4.0841 | 3.6633 | 3.973
Ge, Germanium (ppm) | 0.12292 | 0.085304 | 0.16053 | IND | IND
Hf, Hafnium (ppm) | 1.2159 | 1.1388 | 1.2929 | 1.1572 | 1.2745
Hg, Mercury (ppm) | 0.048397 | 0.032629 | 0.064166 | IND | IND
Ho, Holmium (ppm) | 0.24504 | 0.21929 | 0.27079 | 0.22741 | 0.26267
In, Indium (ppm) | 0.10472 | 0.096046 | 0.11339 | 0.096717 | 0.11272
K, Potassium (wt.%) | 0.23532 | 0.22034 | 0.2503 | 0.2244 | 0.24624
La, Lanthanum (ppm) | 27.9003 | 26.434 | 29.3666 | 27.2402 | 28.5603
Li, Lithium (ppm) | 19.8592 | 19.011 | 20.7074 | 19.2966 | 20.4218
Lu, Lutetium (ppm) | 0.027669 | 0.01904 | 0.036299 | IND | IND
Mg, Magnesium (wt.%) | 0.088435 | 0.086597 | 0.090273 | 0.08661 | 0.09026
Mn, Manganese (wt.%) | 0.025177 | 0.024328 | 0.026026 | 0.024759 | 0.025595
Mo, Molybdenum (ppm) | 3.9978 | 3.6998 | 4.2957 | 3.8167 | 4.1788
Na, Sodium (wt.%) | 0.075404 | 0.068855 | 0.081952 | 0.069893 | 0.080914
Nb, Niobium (ppm) | 0.90081 | 0.74085 | 1.0608 | 0.80523 | 0.99639
Nd, Neodymium (ppm) | 25.5332 | 22.8297 | 28.2367 | 24.8222 | 26.2443
Ni, Nickel (ppm) | 3.4101 | 3.0977 | 3.7224 | 3.1232 | 3.6969
P, Phosphorus (wt.%) | 0.021608 | 0.020699 | 0.022517 | 0.021138 | 0.022078
Pb, Lead (ppm) | 127.5224 | 123.3107 | 131.7341 | 124.8979 | 130.1469
Pr, Praseodymium (ppm) | 6.4858 | 5.8341 | 7.1374 | 6.1147 | 6.8569
Rb, Rubidium (ppm) | 13.3302 | 12.4667 | 14.1938 | 12.8107 | 13.8497
Re, Rhenium (ppm) | < 0.05 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.25123 | 0.24097 | 0.2615 | 0.24439 | 0.25808
Sb, Antimony (ppm) | 4.8405 | 4.4961 | 5.1849 | 4.6106 | 5.0704
Sc, Scandium (ppm) | 1.5165 | 1.3145 | 1.7184 | 1.4194 | 1.6135
Se, Selenium (ppm) | 1.2844 | 1.0633 | 1.5054 | IND | IND
Sm, Samarium (ppm) | 4.6177 | 4.0684 | 5.167 | 4.364 | 4.8714
Sn, Tin (ppm) | 1.147 | 1.0261 | 1.268 | 1.0948 | 1.1993
Sr, Strontium (ppm) | 24.8338 | 23.7639 | 25.9037 | 24.1277 | 25.5399
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.46139 | 0.40611 | 0.51668 | 0.43523 | 0.48756
Te, Tellurium (ppm) | 0.79864 | 0.71029 | 0.88699 | 0.74738 | 0.84989
Th, Thorium (ppm) | 10.8095 | 10.2682 | 11.3508 | 10.4671 | 11.1518
Ti, Titanium (wt.%) | 0.021017 | 0.019198 | 0.022836 | 0.019983 | 0.022051
Tl, Thallium (ppm) | 0.25128 | 0.23482 | 0.26774 | 0.23911 | 0.26345
U, Uranium (ppm) | 2.76 | 2.5593 | 2.9607 | 2.6698 | 2.8502
V, Vanadium (ppm) | 1.9682 | 1.8754 | 2.0609 | 1.8633 | 2.073
W, Tungsten (ppm) | 0.52141 | 0.47434 | 0.56848 | 0.48492 | 0.55789
Y, Yttrium (ppm) | 7.7506 | 7.3455 | 8.1556 | 7.5003 | 8.0008
Yb, Ytterbium (ppm) | 0.22761 | 0.18875 | 0.26647 | 0.20966 | 0.24555
Zn, Zinc (ppm) | 282.1638 | 273.5967 | 290.7309 | 275.7383 | 288.5892
Zr, Zirconium (ppm) | 42.9689 | 39.7609 | 46.1769 | 41.506 | 44.4318
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 606b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Infrared Combustion
C | wt.% | 0.16944 | C-(Inorganic) | wt.% | 0.12667 | C-(Organic) | wt.% | 0.043333
4-Acid Digestion
B | ppm | 14.1667 | Hg | ppm | 0.023333
Aqua Regia Digestion
Pd | ppb | 18.3333 | Pt | ppb | < 5 | Tm | ppm | 0.047479
Alkaline Leach
S-(Sulphate) | wt.% | 0.24491 | S-(Sulphide) | wt.% | 0.20838
Acid Leach
S-(Sulphate) | wt.% | 0.15833
Borate Fusion XRF
Al2O3 | wt.% | 14.855 | Fe2O3 | wt.% | 3.4098 | S | wt.% | 0.473
As | ppm | 80 | K2O | wt.% | 3.69 | SiO2 | wt.% | 69.235
BaO | ppm | 3438.82 | MgO | wt.% | 0.24 | Sn | ppm | 55
CaO | wt.% | 1.535 | MnO | wt.% | 0.034 | Sr | ppm | 200
Cl | ppm | 20 | Na2O | wt.% | 3.665 | TiO2 | wt.% | 0.203
Co | ppm | < 10 | Ni | ppm | < 10 | V2O5 | ppm | 17.852
Cr2O3 | ppm | 36.5375 | P2O5 | wt.% | 0.066456 | Zn | ppm | 320
Cu | wt.% | 0.036 | Pb | ppm | 180 | Zr | ppm | 270
Thermogravimetry
LOI1000 | wt.% | 1.975
Laser Ablation ICP-MS
Ag | ppm | 2.15 | Hf | ppm | 7.67 | Sm | ppm | 7.33
As | ppm | 101.4 | Ho | ppm | 0.515 | Sn | ppm | 4.5
Ba | ppm | 2890 | In | ppm | 0.125 | Sr | ppm | 194.5
Be | ppm | 3.3 | La | ppm | 44.9 | Ta | ppm | 1.395
Bi | ppm | 6.15 | Lu | ppm | 0.07 | Tb | ppm | 0.75
Cd | ppm | 1.6 | Mn | wt.% | 0.02905 | Te | ppm | 1
Ce | ppm | 86.75 | Mo | ppm | 4.3 | Th | ppm | 14.3
Co | ppm | 2.8 | Nb | ppm | 18.25 | Ti | wt.% | 0.119
Cr | ppm | 18 | Nd | ppm | 38.35 | Tl | ppm | 1.2
Cs | ppm | 6.155 | Ni | ppm | 5 | Tm | ppm | 0.115
Cu | wt.% | 0.035 | Pb | ppm | 189 | U | ppm | 5.755
Dy | ppm | 3.48 | Pr | ppm | 10.3 | V | ppm | 6.15
Er | ppm | 0.985 | Rb | ppm | 136.5 | W | ppm | 1.5
Eu | ppm | 1.25 | Re | ppm | 0.0075 | Y | ppm | 14.45
Ga | ppm | 23.3 | Sb | ppm | 7.7 | Yb | ppm | 0.575
Gd | ppm | 5.605 | Sc | ppm | 3.75 | Zn | ppm | 372.5
Ge | ppm | 1.525 | Se | ppm | < 5 | Zr | ppm | 275
3-Acid Digestion (no HF)
Ag | ppm | 1.785 | Gd | ppm | 5.6 | S | wt.% | 0.46833
Al2O3 | wt.% | 14.2833 | Hf | ppm | 6.3833 | Sc | ppm | 4.05
Ba | ppm | 2796.6667 | Ho | ppm | 0.42667 | Sm | ppm | 6.9667
Be | ppm | 3 | K2O | wt.% | 3.5583 | Sn | ppm | 3.7167
Bi | ppm | 4.9 | La | ppm | 45.3333 | Sr | ppm | 203
CaO | wt.% | 1.545 | Li | ppm | 31.5 | Ta | ppm | 1.3117
Cd | ppm | 1.25 | MgO | wt.% | 0.213 | Tb | ppm | 0.67667
Ce | ppm | 82.5 | MnO | wt.% | 0.034833 | Th | ppm | 13.1667
Co | ppm | 2.1433 | Mo | ppm | 4.3983 | TiO2 | wt.% | 0.18817
Cr | ppm | 14.6667 | Na2O | wt.% | 3.5667 | U | ppm | 5.5367
Cs | ppm | 6 | Nb | ppm | 17.6167 | V | ppm | 4.9833
Cu | wt.% | 0.032867 | Nd | ppm | 36.6167 | W | ppm | 2.0167
Dy | ppm | 3.4283 | Ni | ppm | 4.1167 | Y | ppm | 13.9333
Er | ppm | 0.97333 | P2O5 | wt.% | 0.0705 | Yb | ppm | 0.56167
Eu | ppm | 1.825 | Pb | ppm | 155.8333 | Zn | ppm | 323.8333
Fe2O3 | wt.% | 3.155 | Pr | ppm | 10.0167 | Zr | ppm | 287.1667
Ga | ppm | 21.4333 | Rb | ppm | 134
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 606b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.38554 | 0.015005 | 0.35553 | 0.41555 | 0.34052 | 0.43055 | 3.892% | 7.784% | 11.68% | 0.36626 | 0.40481
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.37491 | 0.032323 | 0.31027 | 0.43956 | 0.27794 | 0.47188 | 8.622% | 17.24% | 25.86% | 0.35617 | 0.39366
Infrared Combustion
S, wt.% | 0.45984 | 0.017593 | 0.42465 | 0.49502 | 0.40706 | 0.51262 | 3.826% | 7.652% | 11.48% | 0.43685 | 0.48283
4-Acid Digestion
Ag, ppm | 1.5274 | 0.088053 | 1.3513 | 1.7035 | 1.2632 | 1.7916 | 5.765% | 11.53% | 17.29% | 1.451 | 1.6038
Al, wt.% | 7.5356 | 0.30663 | 6.9223 | 8.1488 | 6.6157 | 8.4555 | 4.069% | 8.138% | 12.21% | 7.1588 | 7.9124
As, ppm | 89.3478 | 5.0909 | 79.166 | 99.5296 | 74.0752 | 104.6204 | 5.698% | 11.4% | 17.09% | 84.8804 | 93.8152
Ba, ppm | 2891.6874 | 127.5147 | 2636.658 | 3146.7167 | 2509.1433 | 3274.2314 | 4.41% | 8.819% | 13.23% | 2747.103 | 3036.2717
Be, ppm | 2.9785 | 0.15587 | 2.6667 | 3.2902 | 2.5109 | 3.4461 | 5.233% | 10.47% | 15.7% | 2.8296 | 3.1274
Bi, ppm | 4.8283 | 0.24306 | 4.3422 | 5.3145 | 4.0991 | 5.5575 | 5.034% | 10.07% | 15.1% | 4.5869 | 5.0697
Ca, wt.% | 1.1 | 0.030776 | 1.0385 | 1.1616 | 1.0077 | 1.1924 | 2.798% | 5.595% | 8.393% | 1.045 | 1.155
Cd, ppm | 1.2405 | 0.09758 | 1.0453 | 1.4356 | 0.94773 | 1.5332 | 7.866% | 15.73% | 23.6% | 1.1784 | 1.3025
Ce, ppm | 88.8532 | 5.3134 | 78.2265 | 99.48 | 72.9132 | 104.7933 | 5.98% | 11.96% | 17.94% | 84.4106 | 93.2959
Co, ppm | 2.2525 | 0.17141 | 1.9096 | 2.5953 | 1.7382 | 2.7667 | 7.61% | 15.22% | 22.83% | 2.1398 | 2.3651
Cr, ppm | 14.7162 | 1.6572 | 11.4018 | 18.0307 | 9.7446 | 19.6879 | 11.26% | 22.52% | 33.78% | 13.9804 | 15.452
Cs, ppm | 6.112 | 0.40638 | 5.2993 | 6.9248 | 4.8929 | 7.3312 | 6.649% | 13.3% | 19.95% | 5.8064 | 6.4176
Cu, wt.% | 0.033391 | 0.0014235 | 0.030544 | 0.036238 | 0.029121 | 0.037662 | 4.263% | 8.526% | 12.79% | 0.031722 | 0.035061
Dy, ppm | 3.5543 | 0.11761 | 3.319 | 3.7895 | 3.2014 | 3.9071 | 3.309% | 6.618% | 9.927% | 3.3766 | 3.732
Er, ppm | 1.007 | 0.069532 | 0.86792 | 1.146 | 0.79839 | 1.2156 | 6.905% | 13.81% | 20.72% | 0.95663 | 1.0573
Eu, ppm | 1.3977 | 0.13655 | 1.1246 | 1.6708 | 0.98805 | 1.8074 | 9.77% | 19.54% | 29.31% | 1.3278 | 1.4676
Fe, wt.% | 2.3276 | 0.086388 | 2.1548 | 2.5004 | 2.0684 | 2.5868 | 3.711% | 7.423% | 11.13% | 2.2112 | 2.444
Ga, ppm | 23.044 | 1.4059 | 20.2323 | 25.8557 | 18.8265 | 27.2616 | 6.101% | 12.2% | 18.3% | 21.8918 | 24.1962
Gd, ppm | 5.6028 | 0.40467 | 4.7935 | 6.4122 | 4.3888 | 6.8168 | 7.223% | 14.45% | 21.67% | 5.3227 | 5.883
Ge, ppm | 0.19297 | 0.079256 | 0.034457 | 0.35148 | 0 | 0.43074 | 41.07% | 82.14% | 123.2% | 0.18332 | 0.20262
Hf, ppm | 6.4822 | 0.28413 | 5.914 | 7.0505 | 5.6298 | 7.3346 | 4.383% | 8.767% | 13.15% | 6.1581 | 6.8064
Ho, ppm | 0.48276 | 0.042909 | 0.39695 | 0.56858 | 0.35404 | 0.61149 | 8.888% | 17.78% | 26.66% | 0.45863 | 0.5069
In, ppm | 0.14224 | 0.010893 | 0.12045 | 0.16403 | 0.10956 | 0.17492 | 7.658% | 15.32% | 22.97% | 0.13513 | 0.14935
K, wt.% | 3.0169 | 0.15448 | 2.708 | 3.3259 | 2.5535 | 3.4804 | 5.121% | 10.24% | 15.36% | 2.8661 | 3.1678
La, ppm | 44.9941 | 2.9516 | 39.0908 | 50.8974 | 36.1392 | 53.849 | 6.56% | 13.12% | 19.68% | 42.7444 | 47.2438
Li, ppm | 32.1337 | 1.4483 | 29.2371 | 35.0303 | 27.7888 | 36.4786 | 4.507% | 9.014% | 13.52% | 30.527 | 33.7404
Lu, ppm | 0.074113 | 0.015043 | 0.044027 | 0.1042 | 0.028984 | 0.11924 | 20.3% | 40.59% | 60.89% | 0.070408 | 0.077819
Mg, wt.% | 0.12667 | 0.0063751 | 0.11392 | 0.13942 | 0.10754 | 0.14579 | 5.033% | 10.07% | 15.1% | 0.12033 | 0.133
Mn, wt.% | 0.028045 | 0.0012531 | 0.025538 | 0.030551 | 0.024285 | 0.031804 | 4.468% | 8.936% | 13.4% | 0.026642 | 0.029447
Mo, ppm | 4.3841 | 0.18201 | 4.0201 | 4.7481 | 3.8381 | 4.9301 | 4.152% | 8.303% | 12.45% | 4.1649 | 4.6033
Na, wt.% | 2.6734 | 0.12509 | 2.4232 | 2.9235 | 2.2981 | 3.0486 | 4.679% | 9.358% | 14.04% | 2.5397 | 2.807
Nb, ppm | 18.4985 | 0.98707 | 16.5244 | 20.4727 | 15.5373 | 21.4597 | 5.336% | 10.67% | 16.01% | 17.5736 | 19.4234
Nd, ppm | 37.4253 | 3.2249 | 30.9755 | 43.8751 | 27.7505 | 47.1001 | 8.617% | 17.23% | 25.85% | 35.554 | 39.2966
Ni, ppm | 3.561 | 0.28876 | 2.9835 | 4.1386 | 2.6948 | 4.4273 | 8.109% | 16.22% | 24.33% | 3.383 | 3.7391
P, wt.% | 0.029991 | 0.00098869 | 0.028013 | 0.031968 | 0.027025 | 0.032957 | 3.297% | 6.593% | 9.89% | 0.028491 | 0.03149
Pb, ppm | 157.0775 | 7.4547 | 142.168 | 171.9869 | 134.7133 | 179.4416 | 4.746% | 9.492% | 14.24% | 149.2236 | 164.9313
Pr, ppm | 10.0749 | 0.63738 | 8.8001 | 11.3496 | 8.1627 | 11.987 | 6.326% | 12.65% | 18.98% | 9.5711 | 10.5786
Rb, ppm | 139.0542 | 9.774 | 119.5062 | 158.6023 | 109.7321 | 168.3763 | 7.029% | 14.06% | 21.09% | 132.1015 | 146.0069
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.4683 | 0.020157 | 0.42799 | 0.50862 | 0.40783 | 0.52877 | 4.304% | 8.609% | 12.91% | 0.44489 | 0.49172
Sb, ppm | 6.3769 | 0.39797 | 5.5809 | 7.1728 | 5.183 | 7.5708 | 6.241% | 12.48% | 18.72% | 6.058 | 6.6957
Sc, ppm | 4.2592 | 0.43315 | 3.3929 | 5.1255 | 2.9597 | 5.5586 | 10.17% | 20.34% | 30.51% | 4.0462 | 4.4721
Se, ppm | 1.7977 | 0.67874 | 0.44022 | 3.1552 | 0 | 3.8339 | 37.76% | 75.51% | 113.3% | 1.7078 | 1.8876
Sm, ppm | 7.2675 | 0.43353 | 6.4005 | 8.1346 | 5.9669 | 8.5681 | 5.965% | 11.93% | 17.9% | 6.9042 | 7.6309
Sn, ppm | 3.7162 | 0.16258 | 3.391 | 4.0413 | 3.2284 | 4.2039 | 4.375% | 8.75% | 13.13% | 3.5303 | 3.902
Sr, ppm | 199.5368 | 9.9715 | 179.5938 | 219.4798 | 169.6223 | 229.4513 | 4.997% | 9.995% | 14.99% | 189.56 | 209.5136
Ta, ppm | 1.3643 | 0.087031 | 1.1902 | 1.5384 | 1.1032 | 1.6254 | 6.379% | 12.76% | 19.14% | 1.2961 | 1.4325
Tb, ppm | 0.73014 | 0.049617 | 0.63091 | 0.82938 | 0.58129 | 0.87899 | 6.795% | 13.59% | 20.39% | 0.69364 | 0.76665
Te, ppm | 0.83512 | 0.07464 | 0.68584 | 0.9844 | 0.6112 | 1.059 | 8.938% | 17.88% | 26.81% | 0.79336 | 0.87687
Th, ppm | 14.794 | 0.71413 | 13.3657 | 16.2222 | 12.6516 | 16.9364 | 4.827% | 9.654% | 14.48% | 14.0543 | 15.5337
Ti, wt.% | 0.11614 | 0.0041079 | 0.10792 | 0.12436 | 0.10382 | 0.12846 | 3.537% | 7.074% | 10.61% | 0.11033 | 0.12195
Tl, ppm | 0.88616 | 0.043397 | 0.79936 | 0.97295 | 0.75597 | 1.0163 | 4.897% | 9.794% | 14.69% | 0.84185 | 0.93046
Tm, ppm | 0.10423 | 0.0071235 | 0.089978 | 0.11847 | 0.082855 | 0.1256 | 6.835% | 13.67% | 20.5% | 0.099014 | 0.10944
U, ppm | 5.5657 | 0.29122 | 4.9833 | 6.1481 | 4.6921 | 6.4394 | 5.232% | 10.46% | 15.7% | 5.2874 | 5.844
V, ppm | 4.9654 | 0.60933 | 3.7467 | 6.1841 | 3.1374 | 6.7934 | 12.27% | 24.54% | 36.81% | 4.7171 | 5.2137
W, ppm | 1.8512 | 0.14898 | 1.5532 | 2.1492 | 1.4043 | 2.2981 | 8.048% | 16.1% | 24.14% | 1.7586 | 1.9438
Y, ppm | 14.85 | 0.70752 | 13.4349 | 16.265 | 12.7274 | 16.9725 | 4.764% | 9.529% | 14.29% | 14.1075 | 15.5925
Yb, ppm | 0.60919 | 0.063722 | 0.48175 | 0.73664 | 0.41803 | 0.80036 | 10.46% | 20.92% | 31.38% | 0.57873 | 0.63965
Zn, ppm | 315.9317 | 15.5168 | 284.8981 | 346.9653 | 269.3813 | 362.4821 | 4.911% | 9.823% | 14.73% | 300.1351 | 331.7283
Zr, ppm | 242.9697 | 12.2937 | 218.3823 | 267.5571 | 206.0886 | 279.8508 | 5.06% | 10.12% | 15.18% | 230.8212 | 255.1182
Aqua Regia Digestion
Ag, ppm | 1.4853 | 0.11287 | 1.2596 | 1.7111 | 1.1467 | 1.824 | 7.599% | 15.2% | 22.8% | 1.4111 | 1.5596
Al, wt.% | 0.80363 | 0.049122 | 0.70538 | 0.90187 | 0.65626 | 0.95099 | 6.113% | 12.23% | 18.34% | 0.76345 | 0.84381
As, ppm | 85.8337 | 4.5478 | 76.7381 | 94.9292 | 72.1904 | 99.4769 | 5.298% | 10.6% | 15.9% | 81.542 | 90.1253
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 211.7273 | 19.6054 | 172.5165 | 250.938 | 152.9111 | 270.5434 | 9.26% | 18.52% | 27.78% | 201.1409 | 222.3136
Be, ppm | 0.51276 | 0.04165 | 0.42946 | 0.59606 | 0.38781 | 0.63771 | 8.123% | 16.25% | 24.37% | 0.48712 | 0.5384
Bi, ppm | 4.9574 | 0.32602 | 4.3053 | 5.6094 | 3.9793 | 5.9354 | 6.576% | 13.15% | 19.73% | 4.7095 | 5.2052
Ca, wt.% | 0.63219 | 0.026573 | 0.57904 | 0.68533 | 0.55247 | 0.71191 | 4.203% | 8.407% | 12.61% | 0.60058 | 0.6638
Cd, ppm | 1.1946 | 0.068929 | 1.0568 | 1.3325 | 0.98784 | 1.4014 | 5.77% | 11.54% | 17.31% | 1.1349 | 1.2544
Ce, ppm | 56.153 | 3.9669 | 48.2192 | 64.0867 | 44.2523 | 68.0536 | 7.064% | 14.13% | 21.19% | 53.3453 | 58.9606
Co, ppm | 2.0231 | 0.061817 | 1.8995 | 2.1468 | 1.8377 | 2.2086 | 3.055% | 6.111% | 9.166% | 1.922 | 2.1243
Cr, ppm | 16.1418 | 0.83438 | 14.473 | 17.8105 | 13.6386 | 18.6449 | 5.169% | 10.34% | 15.51% | 15.3347 | 16.9489
Cs, ppm | 1.4778 | 0.15365 | 1.1705 | 1.7851 | 1.0168 | 1.9387 | 10.4% | 20.8% | 31.19% | 1.4039 | 1.5517
Cu, wt.% | 0.033727 | 0.0012017 | 0.031323 | 0.03613 | 0.030121 | 0.037332 | 3.563% | 7.126% | 10.69% | 0.03204 | 0.035413
Dy, ppm | 2.084 | 0.16436 | 1.7552 | 2.4127 | 1.5909 | 2.577 | 7.887% | 15.77% | 23.66% | 1.9798 | 2.1881
Er, ppm | 0.51123 | 0.065095 | 0.38104 | 0.64142 | 0.31594 | 0.70652 | 12.73% | 25.47% | 38.2% | 0.48567 | 0.53679
Eu, ppm | 0.69609 | 0.060202 | 0.57569 | 0.81649 | 0.51549 | 0.8767 | 8.649% | 17.3% | 25.95% | 0.66129 | 0.7309
Fe, wt.% | 1.9992 | 0.065705 | 1.8677 | 2.1306 | 1.802 | 2.1963 | 3.287% | 6.573% | 9.86% | 1.8992 | 2.0991
Ga, ppm | 5.8994 | 0.4317 | 5.036 | 6.7628 | 4.6043 | 7.1945 | 7.318% | 14.64% | 21.95% | 5.6044 | 6.1943
Gd, ppm | 3.8181 | 0.23566 | 3.3468 | 4.2895 | 3.1111 | 4.5251 | 6.172% | 12.34% | 18.52% | 3.6272 | 4.009
Ge, ppm | 0.12292 | 0.037923 | 0.047072 | 0.19876 | 0.009149 | 0.23668 | 30.85% | 61.7% | 92.56% | 0.11677 | 0.12906
Hf, ppm | 1.2159 | 0.068084 | 1.0797 | 1.352 | 1.0116 | 1.4201 | 5.6% | 11.2% | 16.8% | 1.1551 | 1.2766
Hg, ppm | 0.048397 | 0.0089868 | 0.030424 | 0.066371 | 0.021437 | 0.075358 | 18.57% | 37.14% | 55.71% | 0.045978 | 0.050817
Ho, ppm | 0.24504 | 0.017378 | 0.21029 | 0.2798 | 0.19291 | 0.29718 | 7.092% | 14.18% | 21.28% | 0.23279 | 0.25729
In, ppm | 0.10472 | 0.0061235 | 0.09247 | 0.11696 | 0.086346 | 0.12309 | 5.848% | 11.7% | 17.54% | 0.099481 | 0.10995
K, wt.% | 0.23532 | 0.017989 | 0.19934 | 0.2713 | 0.18135 | 0.28929 | 7.645% | 15.29% | 22.93% | 0.22355 | 0.24709
La, ppm | 27.9003 | 2.0748 | 23.7506 | 32.0499 | 21.6758 | 34.1248 | 7.437% | 14.87% | 22.31% | 26.5053 | 29.2953
Li, ppm | 19.8592 | 1.2582 | 17.3428 | 22.3755 | 16.0847 | 23.6337 | 6.335% | 12.67% | 19.01% | 18.8662 | 20.8521
Lu, ppm | 0.027669 | 0.0049958 | 0.017678 | 0.037661 | 0.012682 | 0.042657 | 18.06% | 36.11% | 54.17% | 0.026286 | 0.029053
Mg, wt.% | 0.088435 | 0.0034005 | 0.081634 | 0.095236 | 0.078233 | 0.098636 | 3.845% | 7.69% | 11.54% | 0.084013 | 0.092857
Mn, wt.% | 0.025177 | 0.00093697 | 0.023303 | 0.027051 | 0.022366 | 0.027988 | 3.721% | 7.443% | 11.16% | 0.023918 | 0.026436
Mo, ppm | 3.9978 | 0.30181 | 3.3942 | 4.6014 | 3.0923 | 4.9032 | 7.549% | 15.1% | 22.65% | 3.7979 | 4.1977
Na, wt.% | 0.075404 | 0.0060526 | 0.063298 | 0.087509 | 0.057246 | 0.093562 | 8.027% | 16.05% | 24.08% | 0.071634 | 0.079174
Nb, ppm | 0.90081 | 0.19387 | 0.51306 | 1.2886 | 0.31919 | 1.4824 | 21.52% | 43.04% | 64.57% | 0.85577 | 0.94585
Nd, ppm | 25.5332 | 2.2598 | 21.0136 | 30.0529 | 18.7537 | 32.3127 | 8.851% | 17.7% | 26.55% | 24.2566 | 26.8099
Ni, ppm | 3.4101 | 0.24384 | 2.9224 | 3.8977 | 2.6786 | 4.1416 | 7.15% | 14.3% | 21.45% | 3.2396 | 3.5806
P, wt.% | 0.021608 | 0.00082856 | 0.019951 | 0.023265 | 0.019122 | 0.024094 | 3.835% | 7.669% | 11.5% | 0.020528 | 0.022688
Pb, ppm | 127.5224 | 4.0239 | 119.4747 | 135.5702 | 115.4508 | 139.594 | 3.155% | 6.311% | 9.466% | 121.1463 | 133.8985
Pr, ppm | 6.4858 | 0.2903 | 5.9052 | 7.0664 | 5.6149 | 7.3567 | 4.476% | 8.952% | 13.43% | 6.1615 | 6.8101
Rb, ppm | 13.3302 | 1.2611 | 10.808 | 15.8524 | 9.547 | 17.1135 | 9.46% | 18.92% | 28.38% | 12.6637 | 13.9967
Re, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.25123 | 0.014851 | 0.22153 | 0.28093 | 0.20668 | 0.29578 | 5.911% | 11.82% | 17.73% | 0.23867 | 0.26379
Sb, ppm | 4.8405 | 0.48785 | 3.8648 | 5.8162 | 3.377 | 6.304 | 10.08% | 20.16% | 30.24% | 4.5985 | 5.0825
Sc, ppm | 1.5165 | 0.17435 | 1.1678 | 1.8652 | 0.99342 | 2.0395 | 11.5% | 22.99% | 34.49% | 1.4406 | 1.5923
Se, ppm | 1.2844 | 0.16051 | 0.96337 | 1.6054 | 0.80286 | 1.7659 | 12.5% | 24.99% | 37.49% | 1.2202 | 1.3486
Sm, ppm | 4.6177 | 0.62563 | 3.3664 | 5.869 | 2.7408 | 6.4946 | 13.55% | 27.1% | 40.65% | 4.3868 | 4.8486
Sn, ppm | 1.147 | 0.12704 | 0.89296 | 1.4011 | 0.76593 | 1.5282 | 11.08% | 22.15% | 33.23% | 1.0897 | 1.2044
Sr, ppm | 24.8338 | 1.3085 | 22.2167 | 27.4509 | 20.9082 | 28.7594 | 5.269% | 10.54% | 15.81% | 23.5921 | 26.0755
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.46139 | 0.02882 | 0.40376 | 0.51903 | 0.37494 | 0.54785 | 6.246% | 12.49% | 18.74% | 0.43832 | 0.48446
Te, ppm | 0.79864 | 0.066739 | 0.66516 | 0.93211 | 0.59842 | 0.99885 | 8.357% | 16.71% | 25.07% | 0.7587 | 0.83857
Th, ppm | 10.8095 | 0.88165 | 9.0462 | 12.5728 | 8.1645 | 13.4544 | 8.156% | 16.31% | 24.47% | 10.269 | 11.3499
Ti, wt.% | 0.021017 | 0.0021847 | 0.016648 | 0.025386 | 0.014463 | 0.027571 | 10.39% | 20.79% | 31.18% | 0.019966 | 0.022068
Tl, ppm | 0.25128 | 0.018691 | 0.2139 | 0.28866 | 0.19521 | 0.30736 | 7.438% | 14.88% | 22.32% | 0.23872 | 0.26385
U, ppm | 2.76 | 0.20287 | 2.3543 | 3.1657 | 2.1514 | 3.3686 | 7.35% | 14.7% | 22.05% | 2.622 | 2.898
V, ppm | 1.9682 | 0.12966 | 1.7088 | 2.2275 | 1.5792 | 2.3571 | 6.588% | 13.18% | 19.76% | 1.8697 | 2.0666
W, ppm | 0.52141 | 0.05595 | 0.40951 | 0.63331 | 0.35356 | 0.68926 | 10.73% | 21.46% | 32.19% | 0.49534 | 0.54748
Y, ppm | 7.7506 | 0.54614 | 6.6583 | 8.8428 | 6.1121 | 9.389 | 7.046% | 14.09% | 21.14% | 7.363 | 8.1381
Yb, ppm | 0.22761 | 0.025438 | 0.17673 | 0.27848 | 0.15129 | 0.30392 | 11.18% | 22.35% | 33.53% | 0.21623 | 0.23899
Zn, ppm | 282.1638 | 13.3814 | 255.4009 | 308.9267 | 242.0195 | 322.3081 | 4.742% | 9.485% | 14.23% | 268.0556 | 296.272
Zr, ppm | 42.9689 | 5.4783 | 32.0124 | 53.9254 | 26.5341 | 59.4036 | 12.75% | 25.5% | 38.25% | 40.8204 | 45.1173
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 606b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
HCl*IRC | hydrochloric acid digestion with infrared combustion
HCl*OES | hydrochloric acid digest with inductively coupled plasma optical emission spectroscopy
I-550*IRC | ignition @ 550 degrees Celsius with infrared combustion
INAA | instrumental neutron activation analysis
IRC | infrared combustion
IRC-x*Calc | calculation by difference from IRC total value
LOI*TGA | loss on ignition by thermogravimetric analysis
Na2CO3*GRAV | sodium carbonate solution with Gravimetric finish
Na2CO3*IRC | sodium carbonate leach with infrared combustion
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 606b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek Genalysis, Perth, WA, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Intertek Testing Services, Townsville, QLD, Australia
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Mineral Services, Townsville, QLD, Australia
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 606b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
1 | 1 | ANSLu | 1 | 2 | 1 | 261680 | 0.087241 | 0.41182 | 0.40804
1 | 1 | ANSLu | 1 | 15 | 4 | 261681 | 0.086873 | 0.41304 | 0.40811
1 | 1 | ANSLu | 1 | 18 | 6 | 261682 | 0.082736 | 0.39686 | 0.40724
1 | 1 | ANSLu | 1 | 20 | 8 | 261683 | 0.085819 | 0.39812 | 0.40731
1 | 1 | ANSLu | 1 | 12 | 11 | 261684 | 0.086591 | 0.40344 | 0.4076
1 | 1 | ANSLu | 1 | 13 | 2 | 261685 | 0.087109 | 0.4028 | 0.40756
1 | 1 | ANSLu | 1 | 9 | 9 | 261686 | 0.083717 | 0.39367 | 0.40707
1 | 1 | ANSLu | 1 | 10 | 9 | 261687 | 0.086041 | 0.43105 | 0.40907
1 | 1 | ANSLu | 1 | 4 | 3 | 261688 | 0.08411 | 0.39677 | 0.40724
1 | 1 | ANSLu | 1 | 17 | 6 | 261689 | 0.084055 | 0.42785 | 0.4089
1 | 1 | ANSLu | 1 | 8 | 7 | 261690 | 0.084194 | 0.40132 | 0.40748
1 | 1 | ANSLu | 1 | 7 | 7 | 261691 | 0.086091 | 0.40036 | 0.40743
1 | 1 | ANSLu | 1 | 1 | 1 | 261692 | 0.088275 | 0.41101 | 0.408
1 | 1 | ANSLu | 1 | 16 | 4 | 261693 | 0.083968 | 0.4102 | 0.40796
1 | 1 | ANSLu | 1 | 11 | 11 | 261694 | 0.085748 | 0.41294 | 0.4081
1 | 1 | ANSLu | 1 | 5 | 5 | 261695 | 0.083647 | 0.41481 | 0.4082
1 | 1 | ANSLu | 1 | 14 | 2 | 261696 | 0.08285 | 0.40542 | 0.4077
1 | 1 | ANSLu | 1 | 19 | 8 | 261697 | 0.083697 | 0.39482 | 0.40714
1 | 1 | ANSLu | 1 | 3 | 3 | 261698 | 0.083429 | 0.39764 | 0.40729
1 | 1 | ANSLu | 1 | 6 | 5 | 261699 | 0.087323 | 0.43264 | 0.40915
Mean | [6 empty] | 0.085176 | 0.40783 | 0.40783
Median | [7 empty] | 0.40443 | 0.40765
SD | [7 empty] | 0.011806 | 0.00062909
RSD | [7 empty] | 2.895% | 0.1543%
RSD Upscale Factor | [7 empty] | 0.053284
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 606b (Certified Value 0.386 ppm) | [30 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | 3
|  |  | 0.085g | 40g | 50g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 50g | 50g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 30g | 30g | 15g | 50g | 3
| 1 | 1 | 0.41101 | 0.41 | 0.4 | 0.371 | 0.401 | 0.395 | 0.39077 | 0.402 | 0.38 | 0.34564 | 0.388 | 0.389 | 0.382 | 0.348 | 0.397 | 0.391 | 0.397 | 0.384 | 0.36 | 0.4 | 0.395 | 0.391 | 0.38 | 0.35 | 0.36 | 0.369 | 0.40671 | 0.351 | 0.38 | 1
| 1 | 2 | 0.41182 | 0.4 | 0.39 | 0.383 | 0.393 | 0.399 | 0.3873 | 0.406 | 0.367 | 0.35661 | 0.385 | 0.375 | 0.376 | 0.355 | 0.391 | 0.386 | 0.389 | 0.384 | 0.392 | 0.41 | 0.396 | 0.394 | 0.41 | 0.37 | 0.35 | 0.377 | 0.38882 | 0.36 | 0.38 | 0
| 1 | 3 | 0.39764 | 0.41 | 0.4 | 0.376 | 0.388 | 0.384 | 0.39146 | 0.399 | 0.353 | 0.35952 | 0.387 | 0.393 | 0.381 | 0.371 | 0.385 | 0.382 | 0.401 | 0.392 | 0.366 | 0.39 | 0.395 | 0.39 | 0.39 | 0.36 | 0.35 | 0.383 | 0.39914 | 0.37 | 0.37 | 16
| 1 | 4 | 0.39677 | 0.4 | 0.4 | 0.373 | 0.395 | 0.386 | 0.38359 | 0.404 | 0.377 | 0.35594 | 0.386 | 0.387 | 0.38 | 0.367 | 0.395 | 0.402 | 0.402 | 0.379 | 0.386 | 0.39 | 0.385 | 0.387 | 0.39 | 0.35 | 0.35 | 0.348 | 0.40432 | 0.359 | 0.37 | 0.38554
| 1 | 5 | 0.41481 | 0.41 | 0.4 | 0.383 | 0.396 | 0.406 | 0.38637 | 0.396 | 0.382 | 0.36631 | 0.388 | 0.392 | 0.38 | 0.363 | 0.376 | 0.393 | 0.392 | 0.39 | 0.368 | 0.4 | 0.408 | 0.382 | 0.39 | 0.33 | 0.36 | 0.38 | 0.40287 | 0.334 | 0.37 | 7
| 1 | 6 | 0.43264 | 0.4 | 0.39 | 0.385 | 0.399 | 0.396 | 0.39063 | 0.4 | 0.373 | 0.35691 | 0.384 | 0.381 | 0.375 | 0.294 | 0.39 | 0.38 | 0.393 | 0.389 | 0.379 | 0.43 | 0.399 | 0.398 | 0.41 | 0.35 | 0.34 | 0.363 | 0.38405 | 0.355 | 0.39
|  | 7 | 00: 0.40036
|  | 8 | 00: 0.40132
|  | 9 | 00: 0.39367
|  | 10 | 00: 0.43105
|  | 11 | 00: 0.41294
|  | 12 | 00: 0.40344
|  | 13 | 00: 0.4028
|  | 14 | 00: 0.40542
|  | 15 | 00: 0.41304
|  | 16 | 00: 0.4102
|  | 17 | 00: 0.42785
|  | 18 | 00: 0.39686
|  | 19 | 00: 0.39482
|  | 20 | 00: 0.39812
| Mean |  | 0.40783 | 0.405 | 0.39667 | 0.3785 | 0.39533 | 0.39433 | 0.38835 | 0.40117 | 0.372 | 0.35682 | 0.38633 | 0.38617 | 0.379 | 0.34967 | 0.389 | 0.389 | 0.39567 | 0.38633 | 0.37517 | 0.40333 | 0.39633 | 0.39033 | 0.395 | 0.35167 | 0.35167 | 0.37 | 0.39765 | 0.35483 | 0.37667
| Median |  | 0.40443 | 0.405 | 0.4 | 0.3795 | 0.3955 | 0.3955 | 0.38897 | 0.401 | 0.375 | 0.35676 | 0.3865 | 0.388 | 0.38 | 0.359 | 0.3905 | 0.3885 | 0.395 | 0.3865 | 0.3735 | 0.4 | 0.3955 | 0.3905 | 0.39 | 0.35 | 0.35 | 0.373 | 0.40101 | 0.357 | 0.375
| Std Dev. |  | 0.011806 | 0.0054772 | 0.005164 | 0.0059245 | 0.0045898 | 0.0082138 | 0.0031129 | 0.0036009 | 0.010733 | 0.0066809 | 0.001633 | 0.0069402 | 0.0028284 | 0.028507 | 0.0076158 | 0.0080994 | 0.0052026 | 0.0048442 | 0.012497 | 0.015055 | 0.0074207 | 0.0055377 | 0.012247 | 0.013292 | 0.0075277 | 0.013054 | 0.0091552 | 0.012024 | 0.008165
| Rel.Std.Dev. |  | 2.895% | 1.352% | 1.302% | 1.565% | 1.161% | 2.083% | 0.8016% | 0.8976% | 2.885% | 1.872% | 0.4227% | 1.797% | 0.7463% | 8.153% | 1.958% | 2.082% | 1.315% | 1.254% | 3.331% | 3.733% | 1.872% | 1.419% | 3.101% | 3.78% | 2.141% | 3.528% | 2.302% | 3.389% | 2.168%
| PDM3 |  | 5.782% | 5.048% | 2.887% | -1.825% | 2.541% | 2.282% | 0.7305% | 4.054% | -3.511% | -7.448% | 0.2066% | 0.1633% | -1.696% | -9.304% | 0.8982% | 0.8982% | 2.627% | 0.2066% | -2.69% | 4.616% | 2.8% | 1.244% | 2.455% | -8.785% | -8.785% | -4.03% | 3.142% | -7.964% | -2.301%
| Z-Score (Absolute) |  | NA | 1.3 | 0.69 | 0.65 | 0.59 | 0.51 | 0.07 | 1.02 | 1.13 | 2.25 | 0.07 | 0.09 | 0.61 | 2.77 | 0.12 | 0.12 | 0.61 | 0.07 | 0.9 | 1.18 | 0.66 | 0.22 | 0.56 | 2.63 | 2.63 | 1.28 | 0.76 | 2.39 | 0.79
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 606b (Certified Value 0.375 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.23 | 2.24 | 2.26 | 2.27 | 2.28 | ppm
|  |  | INAA | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*OES | 3
|  |  | 0.085g | 30g | 50g | 15g | 15g | 25g | 30g | 25g | 30g | 30g | 10g | 25g | 25g | 25g | 25g | 25g | 10g | 25g | 10g | 50g | 50g | 25g | 50g | 25g | 3
| 1 | 1 | 0.41101 | 0.39 | 0.38838 | 0.3636 | 0.406 | 0.417 | 0.387 | 0.31 | 0.33 | 0.36 | 0.4 | 0.39 | 0.39 | 0.39 | 0.38 | 0.36 | 0.39 | 0.13 | 0.34 | 0.32 | 0.367 | 0.36855 | 0.341 | 0.4 | 1
| 1 | 2 | 0.41182 | 0.39 | 0.39435 | 0.3636 | 0.3935 | 0.397 | 0.366 | 0.32 | 0.3308 | 0.36 | 0.369 | 0.37 | 0.38 | 0.4 | 0.37 | 0.36 | 0.39 | 0.18 | 0.34 | 0.3 | 0.3799 | 0.37992 | 0.349 | 0.45 | 0
| 1 | 3 | 0.39764 | 0.4 | 0.3824 | 0.374 | 0.3948 | 0.397 | 0.383 | 0.31 | 0.347 | 0.2 | 0.426 | 0.46 | 0.39 | 0.4 | 0.39 | 0.36 | 0.36 | 0.16 | 0.32 | 0.4 | 0.379 | 0.37086 | 0.349 | 0.42 | 16
| 1 | 4 | 0.39677 | 0.39 | 0.37344 | 0.376 | 0.3962 | 0.399 | 0.37 | 0.32 | 0.347 | 0.36 | 0.389 | 0.45 | 0.38 | 0.41 | 0.4 | 0.36 | 0.38 | 0.16 | 0.33 | 0.32 | 0.368 | 0.37393 | 0.353 | 0.45 | 0.37491
| 1 | 5 | 0.41481 | 0.4 | 0.39734 | 0.3535 | 0.4072 | 0.404 | 0.38 | 0.32 | 0.319 | 0.35 | 0.433 | 0.38 | 0.38 | 0.41 | 0.4 | 0.36 | 0.39 | 0.07 | 0.34 | 0.32 | 0.394 | 0.37835 | 0.358 | 0.44 | 9
| 1 | 6 | 0.43264 | 0.39 | 0.4063 | 0.359 | 0.3889 | 0.397 | 0.382 | 0.31 | 0.33 | 0.36 | 0.397 | 0.4 | 0.38 | 0.4 | 0.39 | 0.36 | 0.35 | 0.09 | 0.33 | 0.32 | 0.391 | 0.37473 | 0.32 | 0.43
|  | 7 | 00: 0.40036
|  | 8 | 00: 0.40132
|  | 9 | 00: 0.39367
|  | 10 | 00: 0.43105
|  | 11 | 00: 0.41294
|  | 12 | 00: 0.40344
|  | 13 | 00: 0.4028
|  | 14 | 00: 0.40542
|  | 15 | 00: 0.41304
|  | 16 | 00: 0.4102
|  | 17 | 00: 0.42785
|  | 18 | 00: 0.39686
|  | 19 | 00: 0.39482
|  | 20 | 00: 0.39812
| Mean |  | 0.40783 | 0.39333 | 0.39037 | 0.36495 | 0.39777 | 0.40183 | 0.378 | 0.315 | 0.33397 | 0.33167 | 0.40233 | 0.40833 | 0.38333 | 0.40167 | 0.38833 | 0.36 | 0.37667 | 0.13167 | 0.33333 | 0.33 | 0.37982 | 0.37439 | 0.345 | 0.43167
| Median |  | 0.40443 | 0.39 | 0.39136 | 0.3636 | 0.3955 | 0.398 | 0.381 | 0.315 | 0.3304 | 0.36 | 0.3985 | 0.395 | 0.38 | 0.4 | 0.39 | 0.36 | 0.385 | 0.145 | 0.335 | 0.32 | 0.37945 | 0.37433 | 0.349 | 0.435
| Std Dev. |  | 0.011806 | 0.005164 | 0.011596 | 0.0086505 | 0.0072781 | 0.0079099 | 0.0081731 | 0.0054772 | 0.011002 | 0.064627 | 0.023763 | 0.037639 | 0.005164 | 0.0075277 | 0.01169 | 6.0809e-17 | 0.017512 | 0.043551 | 0.008165 | 0.035214 | 0.011232 | 0.0043138 | 0.013461 | 0.019408
| Rel.Std.Dev. |  | 2.895% | 1.313% | 2.971% | 2.37% | 1.83% | 1.968% | 2.162% | 1.739% | 3.294% | 19.49% | 5.906% | 9.218% | 1.347% | 1.874% | 3.01% | 1.689e-14% | 4.649% | 33.08% | 2.449% | 10.67% | 2.957% | 1.152% | 3.902% | 4.496%
| PDM3 |  | 8.78% | 4.913% | 4.122% | -2.657% | 6.096% | 7.18% | 0.8233% | -15.98% | -10.92% | -11.54% | 7.314% | 8.914% | 2.246% | 7.136% | 3.58% | -3.978% | 0.4677% | -64.88% | -11.09% | -11.98% | 1.308% | -0.1396% | -7.979% | 15.14%
| Z-Score (Absolute) |  | NA | 0.44 | 0.35 | 0.37 | 0.56 | 0.68 | 0 | 1.79 | 1.25 | 1.32 | 0.69 | 0.86 | 0.15 | 0.67 | 0.29 | 0.51 | 0.04 | 7.02 | 1.27 | 1.37 | 0.05 | 0.1 | 0.94 | 1.53
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 606b (Indicative Value 0.169 wt.%) | [26 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 2.17 | 2.18 | 2.27 | 3.01 | [21 empty] | wt.%
|  |  | IRC | IRC | IRC | IRC | [21 empty] | 3
3
| 1 | 1 | 0.17 | 0.16 | 0.079 | 0.17 | [21 empty] | 1
| 1 | 2 | 0.17 | 0.17 | 0.048 | 0.17 | [21 empty] | 1
| 1 | 3 | 0.17 | 0.17 | 0.044 | [22 empty] | 16
| 1 | 4 | 0.19 | 0.17 | 0.052 | [22 empty] | 0.16944
| 1 | 5 | 0.16 | 0.17 | 0.069 | [22 empty] | 7
| 1 | 6 | 0.16 | 0.17 | 0.044
| Mean |  | 0.17 | 0.16833 | 0.056 | 0.17
| Median |  | 0.17 | 0.17 | 0.05 | 0.17
| Std Dev. |  | 0.010954 | 0.0040825 | 0.014601 | 0
| Rel.Std.Dev. |  | 6.444% | 2.425% | 26.07% | 0%
| PDM3 |  | 0.3279% | -0.6557% | -66.95% | 0.3279%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 91.57 | 0.67
| Analytical results for C-(Inorganic) in OREAS 606b (Indicative Value 0.127 wt.%) | [26 empty] | Indicative
C-(Inorganic) | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.18 | [24 empty] | wt.%
|  |  | I-550*IRC | [24 empty] | 3
3
| 1 | 1 | 0.12 | [24 empty] | 1
| 1 | 2 | 0.11 | [24 empty] | 1
| 1 | 3 | 0.13 | [24 empty] | 16
| 1 | 4 | 0.14 | [24 empty] | 0.12667
| 1 | 5 | 0.13 | [24 empty] | 7
| 1 | 6 | 0.13
| Mean |  | 0.12667
| Median |  | 0.13
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 8.154%
| PDM3 |  | -2.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for C-(Organic) in OREAS 606b (Indicative Value 0.043 wt.%) | [26 empty] | Indicative
C-(Organic) | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.18 | [24 empty] | wt.%
|  |  | HCl*IRC | [24 empty] | 3
3
| 1 | 1 | 0.04 | [24 empty] | 1
| 1 | 2 | 0.06 | [24 empty] | 1
| 1 | 3 | 0.04 | [24 empty] | 16
| 1 | 4 | 0.04 | [24 empty] | 0.043333
| 1 | 5 | 0.04 | [24 empty] | 7
| 1 | 6 | 0.04
| Mean |  | 0.043333
| Median |  | 0.04
| Std Dev. |  | 0.008165
| Rel.Std.Dev. |  | 18.84%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 606b (Certified Value 0.46 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.5 | 0.552 | 0.46 | 0.46 | 0.44 | 0.43515 | 0.47 | 0.46 | 0.46 | 0.45 | 0.46 | 0.45 | 0.48 | 0.42 | 0.45 | 0.45 | 0.44 | 0.467 | 0.47 | 0.44 | 0.23 | 0.46 | 0.47975 | 0.47 | 0.42 | 1
| 1 | 2 | 0.49 | 0.668 | 0.46 | 0.46 | 0.44 | 0.4451 | 0.46 | 0.48 | 0.46 | 0.46 | 0.46 | 0.44 | 0.49 | 0.43 | 0.46 | 0.45 | 0.44 | 0.465 | 0.48 | 0.43 | 0.21 | 0.45 | 0.49344 | 0.47 | 0.42 | 11
| 1 | 3 | 0.49 | 0.634 | 0.45 | 0.46 | 0.44 | 0.46876 | 0.46 | 0.46 | 0.47 | 0.46 | 0.46 | 0.45 | 0.48 | 0.43 | 0.46 | 0.46 | 0.44 | 0.471 | 0.47 | 0.41 | 0.21 | 0.46 | 0.48769 | 0.47 | 0.35 | 16
| 1 | 4 | 0.49 | 0.591 | 0.46 | 0.47 | 0.45 | 0.45636 | 0.46 | 0.46 | 0.46 | 0.45 | 0.47 | 0.44 | 0.49 | 0.44 | 0.47 | 0.45 | 0.45 | 0.472 | 0.47 | 0.43 | 0.22 | 0.46 | 0.48303 | 0.47 | 0.36 | 0.45984
| 1 | 5 | 0.5 | 0.62 | 0.45 | 0.45 | 0.47 | 0.44377 | 0.47 | 0.45 | 0.46 | 0.46 | 0.47 | 0.43 | 0.49 | 0.44 | 0.47 | 0.47 | 0.44 | 0.479 | 0.48 | 0.44 | 0.22 | 0.45 | 0.48176 | 0.47 | 0.41 | 11
| 1 | 6 | 0.48 | 0.576 | 0.46 | 0.46 | 0.43 | 0.42583 | 0.46 | 0.46 | 0.45 | 0.46 | 0.46 | 0.46 | 0.49 | 0.42 | 0.47 | 0.47 | 0.45 | 0.474 | 0.48 | 0.44 | 0.22 | 0.46 | 0.48999 | 0.47 | 0.36
| Mean |  | 0.49167 | 0.60683 | 0.45667 | 0.46 | 0.445 | 0.44583 | 0.46333 | 0.46167 | 0.46 | 0.45667 | 0.46333 | 0.445 | 0.48667 | 0.43 | 0.46333 | 0.45833 | 0.44333 | 0.47133 | 0.475 | 0.43167 | 0.21833 | 0.45667 | 0.48594 | 0.47 | 0.38667
| Median |  | 0.49 | 0.6055 | 0.46 | 0.46 | 0.44 | 0.44443 | 0.46 | 0.46 | 0.46 | 0.46 | 0.46 | 0.445 | 0.49 | 0.43 | 0.465 | 0.455 | 0.44 | 0.4715 | 0.475 | 0.435 | 0.22 | 0.46 | 0.48536 | 0.47 | 0.385
| Std Dev. |  | 0.0075277 | 0.042097 | 0.005164 | 0.0063246 | 0.013784 | 0.015203 | 0.005164 | 0.0098319 | 0.0063246 | 0.005164 | 0.005164 | 0.010488 | 0.005164 | 0.0089443 | 0.008165 | 0.0098319 | 0.005164 | 0.0050067 | 0.0054772 | 0.01169 | 0.0075277 | 0.005164 | 0.0052919 | 0 | 0.033267
| Rel.Std.Dev. |  | 1.531% | 6.937% | 1.131% | 1.375% | 3.098% | 3.41% | 1.115% | 2.13% | 1.375% | 1.131% | 1.115% | 2.357% | 1.061% | 2.08% | 1.762% | 2.145% | 1.165% | 1.062% | 1.153% | 2.708% | 3.448% | 1.131% | 1.089% | 0% | 8.603%
| PDM3 |  | 6.922% | 31.97% | -0.6897% | 0.03522% | -3.227% | -3.047% | 0.7601% | 0.3977% | 0.03522% | -0.6897% | 0.7601% | -3.227% | 5.834% | -6.489% | 0.7601% | -0.3272% | -3.589% | 2.5% | 3.297% | -6.126% | -52.52% | -0.6897% | 5.677% | 2.21% | -15.91%
| Z-Score (Absolute) |  | 1.51 | 6.99 | 0.16 | 0 | 0.71 | 0.67 | 0.16 | 0.08 | 0 | 0.16 | 0.16 | 0.71 | 1.27 | 1.43 | 0.16 | 0.08 | 0.79 | 0.54 | 0.71 | 1.35 | 11.5 | 0.16 | 1.23 | 0.48 | 3.49
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 606b (Certified Value 1.53 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 1.3 | 1.269 | 1.5 | 1.49 | 1.9738 | 2 | 1.49 | 1.0675 | 1.5 | 1.3 | 1.5 | 1.46 | 1.44 | 1.6 | 1.52 | 1.57 | 1.5 | 1.46 | 1.71 | 1.16 | 1.46 | 1.2638 | 1.2 | 1.564 | < 5 | 1.61 | 1
| 1 | 2 | 1.3 | 1.315 | 1.6 | 1.58 | 1.9693 | 2.3 | 1.56 | 1.095 | 1.6 | 1.2 | 1.5 | 1.54 | 1.5 | 1.54 | 1.52 | 1.58 | 1.5 | 1.54 | 1.79 | 1.16 | 1.48 | 1.2749 | 0.8 | 1.5401 | < 5 | 1.46 | 22
| 1 | 3 | 1.4 | 1.289 | 1.6 | 1.56 | 1.9821 | 2 | 1.56 | 1.1635 | 1.5 | 1.3 | 1.5 | 1.54 | 1.54 | 1.65 | 1.53 | 1.57 | 1.5 | 1.43 | 1.68 | 1.16 | 1.44 | 1.2683 | 0.8 | 1.6294 | < 5 | 1.6 | 16
| 1 | 4 | 1.4 | 1.214 | 1.5 | 1.63 | 1.9456 | 2.2 | 1.52 | 1.142 | 1.5 | 1.2 | 1.5 | 1.54 | 1.51 | 1.54 | 1.52 | 1.63 | 1.5 | 1.44 | 2.09 | 1.13 | 1.43 | 1.3137 | 1.1 | 1.5784 | < 5 | 1.46 | 1.5274
| 1 | 5 | 1.3 | 1.258 | 1.6 | 1.53 | 1.9578 | 2.2 | 1.53 | 1.058 | 1.5 | 1.3 | 1.5 | 1.58 | 1.52 | 1.6 | 1.55 | 1.6 | 1.5 | 1.43 | 1.75 | 1.12 | 1.38 | 1.1941 | 1 | 1.6238 | < 5 | 1.52 | 13
| 1 | 6 | 1.3 | 1.431 | 1.6 | 1.47 | 1.9596 | 2.1 | 1.55 | 1.174 | 1.6 | 1.2 | 1.5 | 1.52 | 1.47 | 1.58 | 1.48 | 1.53 | 1.5 | 1.41 | 1.74 | 1.15 | 1.37 | 1.1714 | 0.8 | 1.5551 | < 5 | 1.64
| Mean |  | 1.3333 | 1.296 | 1.5667 | 1.5433 | 1.9647 | 2.1333 | 1.535 | 1.1167 | 1.5333 | 1.25 | 1.5 | 1.53 | 1.4967 | 1.585 | 1.52 | 1.58 | 1.5 | 1.4517 | 1.7933 | 1.1467 | 1.4267 | 1.2477 | 0.95 | 1.5818 |  | 1.5483
| Median |  | 1.3 | 1.279 | 1.6 | 1.545 | 1.9644 | 2.15 | 1.54 | 1.1185 | 1.5 | 1.25 | 1.5 | 1.54 | 1.505 | 1.59 | 1.52 | 1.575 | 1.5 | 1.435 | 1.745 | 1.155 | 1.435 | 1.266 | 0.9 | 1.5712 |  | 1.56
| Std Dev. |  | 0.05164 | 0.074205 | 0.05164 | 0.059217 | 0.012999 | 0.12111 | 0.027386 | 0.049902 | 0.05164 | 0.054772 | 0 | 0.039497 | 0.036148 | 0.041833 | 0.022804 | 0.033466 | 0 | 0.046224 | 0.15002 | 0.017512 | 0.043665 | 0.053803 | 0.17607 | 0.036901 |  | 0.079099
| Rel.Std.Dev. |  | 3.873% | 5.726% | 3.296% | 3.837% | 0.6616% | 5.677% | 1.784% | 4.469% | 3.368% | 4.382% | 0% | 2.581% | 2.415% | 2.639% | 1.5% | 2.118% | 0% | 3.184% | 8.366% | 1.527% | 3.061% | 4.312% | 18.53% | 2.333% |  | 5.109%
| PDM3 |  | -12.71% | -15.15% | 2.571% | 1.043% | 28.63% | 39.67% | 0.4976% | -26.89% | 0.3885% | -18.16% | -1.794% | 0.1702% | -2.012% | 3.771% | -0.4845% | 3.444% | -1.794% | -4.958% | 17.41% | -24.93% | -6.595% | -18.31% | -37.8% | 3.562% |  | 1.371%
| Z-Score (Absolute) |  | 1.94 | 2.32 | 0.42 | 0.19 | 4.45 | 6.15 | 0.1 | 4.13 | 0.08 | 2.78 | 0.25 | 0.05 | 0.29 | 0.61 | 0.05 | 0.56 | 0.25 | 0.74 | 2.71 | 3.83 | 0.99 | 2.8 | 5.82 | 0.57 | 9.86 | 0.24
| Analytical results for Al in OREAS 606b (Certified Value 7.54 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.72 | 6.74 | 7.65 | 7.6464 | 7.6123 | 7.75 | 7.54 | 7.9189 | 7.67 | 7.1 | 7.8563 | 7.36 | 7.12 | 7.57 | 7.2 | 8.13 | 7.2977 | 5.12 | 8.41 | 7.27 | 7.42 | 7.2489 | 7.73 | 7.6169 | 7.5992 | 6.69 | 1
| 1 | 2 | 7.65 | 7.05 | 7.86 | 7.5502 | 7.6939 | 7.68 | 7.54 | 7.9505 | 7.5 | 7.14 | 7.9054 | 7.18 | 7.2 | 7.4 | 7.5 | 7.96 | 7.1067 | 5.94 | 8.31 | 7.39 | 7.3 | 7.1947 | 7.52 | 7.6719 | 7.6224 | 6.41 | 0
| 1 | 3 | 7.67 | 6.91 | 7.65 | 7.5246 | 7.654 | 7.55 | 7.54 | 7.9505 | 7.78 | 7.17 | 7.8446 | 7.25 | 7.27 | 7.54 | 7.24 | 7.87 | 7.3351 | 5.08 | 8.38 | 7.49 | 7.52 | 7.1829 | 7.57 | 7.4392 | 7.5775 | 6.23 | 16
| 1 | 4 | 7.67 | 6.52 | 7.6 | 7.807 | 7.5449 | 7.69 | 7.52 | 7.9269 | 7.67 | 7.11 | 7.5876 | 7.4 | 7.36 | 7.47 | 7.24 | 8.16 | 7.2552 | 4.94 | 8.13 | 7.25 | 7.45 | 7.3782 | 7.46 | 7.4736 | 7.2162 | 6.42 | 7.5356
| 1 | 5 | 7.69 | 6.7 | 7.58 | 7.6751 | 7.5713 | 7.55 | 7.53 | 7.9882 | 7.73 | 7.01 | 7.6315 | 7.52 | 7.47 | 7.51 | 7.27 | 8.05 | 7.3829 | 5.12 | 8.2 | 7.1 | 7.41 | 7.3425 | 7.52 | 7.6871 | 7.3994 | 6.48 | 14
| 1 | 6 | 7.66 | 8.12 | 7.63 | 7.3553 | 7.6188 | 7.78 | 7.48 | 7.9793 | 7.58 | 7.15 | 8.1085 | 7.22 | 7.31 | 7.61 | 6.2 | 7.97 | 7.2413 | 5.99 | 8.15 | 7.43 | 7.5 | 7.3548 | 7.57 | 7.8159 | 7.7866 | 6.57
| Mean |  | 7.6767 | 7.0067 | 7.6617 | 7.5931 | 7.6159 | 7.6667 | 7.525 | 7.9524 | 7.655 | 7.1133 | 7.8223 | 7.3217 | 7.2883 | 7.5167 | 7.1083 | 8.0233 | 7.2698 | 5.365 | 8.2633 | 7.3217 | 7.4333 | 7.2837 | 7.5617 | 7.6174 | 7.5335 | 6.4667
| Median |  | 7.67 | 6.825 | 7.64 | 7.5983 | 7.6155 | 7.685 | 7.535 | 7.9505 | 7.67 | 7.125 | 7.8505 | 7.305 | 7.29 | 7.525 | 7.24 | 8.01 | 7.2764 | 5.12 | 8.255 | 7.33 | 7.435 | 7.2957 | 7.545 | 7.6444 | 7.5883 | 6.45
| Std Dev. |  | 0.025033 | 0.57493 | 0.10108 | 0.15388 | 0.054033 | 0.097707 | 0.023452 | 0.027515 | 0.10134 | 0.056804 | 0.19077 | 0.12844 | 0.12254 | 0.074744 | 0.45775 | 0.11057 | 0.095353 | 0.46971 | 0.11994 | 0.14261 | 0.078401 | 0.085715 | 0.09196 | 0.14117 | 0.19842 | 0.15629
| Rel.Std.Dev. |  | 0.3261% | 8.205% | 1.319% | 2.027% | 0.7095% | 1.274% | 0.3117% | 0.346% | 1.324% | 0.7986% | 2.439% | 1.754% | 1.681% | 0.9944% | 6.44% | 1.378% | 1.312% | 8.755% | 1.452% | 1.948% | 1.055% | 1.177% | 1.216% | 1.853% | 2.634% | 2.417%
| PDM3 |  | 1.872% | -7.019% | 1.673% | 0.7632% | 1.065% | 1.739% | -0.1405% | 5.531% | 1.585% | -5.603% | 3.805% | -2.839% | -3.281% | -0.2511% | -5.67% | 6.473% | -3.527% | -28.8% | 9.657% | -2.839% | -1.357% | -3.343% | 0.3461% | 1.086% | -0.02726% | -14.18%
| Z-Score (Absolute) |  | 0.48 | 1.7 | 0.43 | 0.21 | 0.28 | 0.45 | 0.01 | 1.37 | 0.41 | 1.35 | 0.95 | 0.67 | 0.78 | 0.04 | 1.37 | 1.6 | 0.84 | 7.03 | 2.38 | 0.67 | 0.31 | 0.8 | 0.1 | 0.29 | 0.01 | 3.45
| Analytical results for As in OREAS 606b (Certified Value 89 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 91 | 95.5 | 61 | 88 | 89.1676 | 63.2 | 89.3 | 86.7 | 94 | 88 | 86.7 | 82 | 93.3 | 92.4 | 89.8 | 86.8 | 90 | 70 | 60.13 | 85 | 106.7868 | 102 | 87.407 | 77.401 | 91 |  | 1
| 1 | 2 | 92 | 91.2 | 62 | 87.1 | 89.7906 | 60.6 | 95.2 | 86.5 | 95 | 88 | 88.4 | 82.3 | 90.4 | 98.1 | 89.2 | 88.9 | 94 | 80 | 61.81 | 82 | 124.6807 | 84 | 86.3277 | 80.641 | 91 |  | 12
| 1 | 3 | 91 | 92.9 | 59 | 86.3 | 90.2232 | 61.9 | 98.1 | 89.1 | 92 | 90 | 89.6 | 80.6 | 93.1 | 96 | 88.1 | 92.3 | 89 | 86 | 60.18 | 83 | 106.3891 | 85 | 84.0292 | 78.554 | 92 |  | 16
| 1 | 4 | 94 | 96.2 | 62 | 89.2 | 90.493 | 64.8 | 93.5 | 87.4 | 93 | 89 | 89.8 | 81.5 | 89.6 | 91.9 | 89.1 | 89.7 | 91 | 88 | 61.07 | 84 | 120.0228 | 103 | 83.3928 | 79.841 | 90 |  | 89.3478
| 1 | 5 | 92 | 94.5 | 60 | 88.2 | 89.2678 | 68.8 | 95.6 | 86.3 | 93 | 89 | 94.7 | 84.4 | 96 | 93.5 | 90 | 94.4 | 92 | 76 | 60.08 | 84 | 101.0508 | 95 | 87.0405 | 76.85 | 94 |  | 15
| 1 | 6 | 93 | 96.5 | 63 | 85.7 | 89.1119 | 63 | 93.5 | 86.3 | 93 | 89 | 86.9 | 75.8 | 92.8 | 91.8 | 88.6 | 85.6 | 94 | 98 | 60.1 | 83 | 107.4713 | 103 | 87.6532 | 82.772 | 91
| Mean |  | 92.1667 | 94.4667 | 61.1667 | 87.4167 | 89.6757 | 63.7167 | 94.2 | 87.05 | 93.3333 | 88.8333 | 89.35 | 81.1 | 92.5333 | 93.95 | 89.1333 | 89.6167 | 91.6667 | 83 | 60.5617 | 83.5 | 111.0669 | 95.3333 | 85.9751 | 79.3432 | 91.5
| Median |  | 92 | 95 | 61.5 | 87.55 | 89.5292 | 63.1 | 94.35 | 86.6 | 93 | 89 | 89 | 81.75 | 92.95 | 92.95 | 89.15 | 89.3 | 91.5 | 83 | 60.155 | 83.5 | 107.1291 | 98.5 | 86.6841 | 79.1975 | 91
| Std Dev. |  | 1.169 | 2.0656 | 1.472 | 1.2983 | 0.58709 | 2.8569 | 2.938 | 1.084 | 1.0328 | 0.75277 | 2.9263 | 2.8858 | 2.2871 | 2.5634 | 0.71461 | 3.3066 | 2.0656 | 9.859 | 0.72026 | 1.0488 | 9.1542 | 8.9144 | 1.821 | 2.2045 | 1.3784
| Rel.Std.Dev. |  | 1.268% | 2.187% | 2.406% | 1.485% | 0.6547% | 4.484% | 3.119% | 1.245% | 1.107% | 0.8474% | 3.275% | 3.558% | 2.472% | 2.728% | 0.8017% | 3.69% | 2.253% | 11.88% | 1.189% | 1.256% | 8.242% | 9.351% | 2.118% | 2.778% | 1.506%
| PDM3 |  | 3.155% | 5.729% | -31.54% | -2.161% | 0.367% | -28.69% | 5.431% | -2.572% | 4.461% | -0.5758% | 0.002456% | -9.231% | 3.565% | 5.151% | -0.24% | 0.3009% | 2.595% | -7.105% | -32.22% | -6.545% | 24.31% | 6.699% | -3.775% | -11.2% | 2.409%
| Z-Score (Absolute) |  | 0.48 | 0.87 | 4.77 | 0.33 | 0.06 | 4.34 | 0.82 | 0.39 | 0.67 | 0.09 | 0 | 1.4 | 0.54 | 0.78 | 0.04 | 0.05 | 0.39 | 1.07 | 4.87 | 0.99 | 3.68 | 1.01 | 0.57 | 1.69 | 0.36
| Analytical results for B in OREAS 606b (Indicative Value 14.2 ppm) | [27 empty] | Indicative
B | Round | Replicate | 2.01: Lab; 2.03: Lab |  | 1
| No | No | 2.01: 2.01; 2.03: 2.23 |  | ppm
|  |  | 2.01: 4A*OES; 2.03: 4A*OES |  | 1
1
| 1 | 1 | 2.01: 15; 2.03: < 0.1 |  | 1
| 1 | 2 | 2.01: 14; 2.03: < 0.1 |  | 3
| 1 | 3 | 2.01: 14; 2.03: < 0.1 |  | 16
| 1 | 4 | 2.01: 14; 2.03: < 0.1 |  | 14.1667
| 1 | 5 | 2.01: 14; 2.03: < 0.1 |  | 9
| 1 | 6 | 2.01: 14; 2.03: < 0.1
| Mean |  | 2.01: 14.1667
| Median |  | 2.01: 14
| Std Dev. |  | 2.01: 0.40825
| Rel.Std.Dev. |  | 2.01: 2.882%
| PDM3 |  | 2.01: -2.331e-13%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.67
| Analytical results for Ba in OREAS 606b (Certified Value 2892 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
0
| 1 | 1 | 2618 | 1312 | 393 | 2831.9 | 2821.8524 | 3160 | 3010 | 3010 | 2900 | 2980 | 2890 | 2780 | 3100 | 2910 | 2836 | 2870 | 2784 | 2898 | 2518.0949 | 2850 | 2855.5666 | 2195.465 | 2738 |  |  |  | 1
| 1 | 2 | 2726 | 1427 | 388 | 2844 | 2888.8342 | 3130 | 3140 | 3100 | 2920 | 3060 | 2820 | 2840 | 3040 | 2816 | 2913 | 2820 | 2805 | 2803 | 2554.6856 | 2770 | 3085.6706 | 2648.192 | 2759 |  |  |  | 24
| 1 | 3 | 2640 | 1268 | 354 | 2828.8 | 2836.606 | 2850 | 3150 | 3060 | 2930 | 3010 | 2870 | 2770 | 3000 | 2906 | 2845 | 2690 | 2804 | 2867 | 2440.7639 | 2800 | 3036.0969 | 2347.098 | 2767 |  |  |  | 16
| 1 | 4 | 2718 | 1391 | 558 | 2956.2 | 2828.1935 | 3140 | 3070 | 2990 | 2960 | 3040 | 2830 | 2760 | 3100 | 2899 | 2798 | 2790 | 2767 | 2916 | 2527.8046 | 2760 | 2985.9837 | 2566.771 | 2817 |  |  |  | 2891.6874
| 1 | 5 | 2687 | 1324 | 465 | 2810.8 | 2838.7821 | 3100 | 3070 | 3040 | 2980 | 3080 | 2880 | 2760 | 3070 | 2937 | 2839 | 2770 | 2819 | 2813 | 2486.0413 | 2780 | 2882.8094 | 2528.05 | 2836 |  |  |  | 16
| 1 | 6 | 2662 | 2435 | 409 | 2835.8 | 2836.5995 | 3140 | 3000 | 3030 | 2860 | 3000 | 2860 | 2400 | 3030 | 2868 | 2820 | 2760 | 2766 | 2873 | 2487.7223 | 2820 | 2939.8045 | 2507.477 | 2708
| Mean |  | 2675.1667 | 1526.1667 | 427.8333 | 2851.25 | 2841.8113 | 3086.6667 | 3073.3333 | 3038.3333 | 2925 | 3028.3333 | 2858.3333 | 2718.3333 | 3056.6667 | 2889.3333 | 2841.8333 | 2783.3333 | 2790.8333 | 2861.6667 | 2502.5188 | 2796.6667 | 2964.3219 | 2465.5088 | 2770.8333
| Median |  | 2674.5 | 1357.5 | 401 | 2833.85 | 2836.6028 | 3135 | 3070 | 3035 | 2925 | 3025 | 2865 | 2765 | 3055 | 2902.5 | 2837.5 | 2780 | 2794 | 2870 | 2502.9086 | 2790 | 2962.8941 | 2517.7635 | 2763
| Std Dev. |  | 42.981 | 448.881 | 73.3905 | 52.5745 | 23.9098 | 117.5868 | 62.8225 | 38.6868 | 42.7785 | 38.1663 | 27.8687 | 158.7976 | 40.332 | 42.1979 | 38.7681 | 60.553 | 21.9036 | 45.2534 | 39.768 | 33.8625 | 88.7913 | 165.0725 | 48.0808
| Rel.Std.Dev. |  | 1.607% | 29.41% | 17.15% | 1.844% | 0.8414% | 3.81% | 2.044% | 1.273% | 1.463% | 1.26% | 0.975% | 5.842% | 1.319% | 1.46% | 1.364% | 2.176% | 0.7848% | 1.581% | 1.589% | 1.211% | 2.995% | 6.695% | 1.735%
| PDM3 |  | -7.488% | -47.22% | -85.2% | -1.398% | -1.725% | 6.743% | 6.282% | 5.071% | 1.152% | 4.725% | -1.153% | -5.995% | 5.705% | -0.08141% | -1.724% | -3.747% | -3.488% | -1.038% | -13.46% | -3.286% | 2.512% | -14.74% | -4.179%
| Z-Score (Absolute) |  | 0.92 | 7.24 | 13.29 | 0.05 | 0 | 1.35 | 1.27 | 1.08 | 0.46 | 1.03 | 0.09 | 0.68 | 1.18 | 0.26 | 0 | 0.32 | 0.28 | 0.11 | 1.87 | 0.25 | 0.67 | 2.07 | 0.39
| Analytical results for Be in OREAS 606b (Certified Value 2.98 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 3.18 | 3 | 3 | 2.89 | 3.0109 | 3 | 2.88 | 3.9799 | 3.1 | 2.9 | 2.9 | 2.9 | 2.95 | 2.87 | 2.8 | 3.09 | 3.02 | 3.21 | 3.1 | 2.92 | 2.91 | 2.4462 | 3.3 | 3.0752 | 2.378 | 3.05 | 1
| 1 | 2 | 3.09 | 3 | 3 | 2.7 | 3.0897 | 2.7 | 2.76 | 3.7937 | 2.8 | 3 | 2.93 | 2.92 | 3 | 2.83 | 2.85 | 3.04 | 2.94 | 3.37 | 3.1 | 3.06 | 3.11 | 2.4739 | 3.3 | 2.926 | 2.409 | 2.9 | 25
| 1 | 3 | 3.31 | 3 | 3 | 2.79 | 3.0875 | 3 | 2.72 | 3.8625 | 2.9 | 3 | 2.93 | 2.93 | 2.98 | 2.86 | 2.78 | 2.98 | 3.02 | 3.25 | 3.1 | 3.01 | 2.98 | 2.4317 | 3.3 | 3.0375 | 2.399 | 2.78 | 16
| 1 | 4 | 3.4 | 3 | 3 | 2.73 | 3.0177 | 3.2 | 2.83 | 3.8988 | 2.9 | 3 | 2.85 | 2.98 | 3.01 | 2.81 | 2.78 | 3.1 | 3.01 | 3.32 | 2.9 | 2.98 | 3.18 | 2.4915 | 3.3 | 3.0631 | 2.367 | 2.73 | 2.9785
| 1 | 5 | 3.2 | 3 | 3 | 2.7 | 2.9906 | 3.4 | 3.05 | 3.8183 | 3 | 2.9 | 2.95 | 2.99 | 3.07 | 2.87 | 2.77 | 3.07 | 3.06 | 3.33 | 3.1 | 3.1 | 2.95 | 2.4925 | 3.3 | 2.9977 | 2.398 | 3.06 | 17
| 1 | 6 | 3.04 | 3 | 3 | 2.69 | 2.9419 | 2.9 | 2.83 | 3.9099 | 2.8 | 2.9 | 2.9 | 2.88 | 2.99 | 2.83 | 2.75 | 3.03 | 2.98 | 3.37 | 2.9 | 3.04 | 3.25 | 2.4921 | 3.4 | 2.9073 | 2.435 | 2.99
| Mean |  | 3.2033 | 3 | 3 | 2.75 | 3.0231 | 3.0333 | 2.845 | 3.8772 | 2.9167 | 2.95 | 2.91 | 2.9333 | 3 | 2.845 | 2.7883 | 3.0517 | 3.005 | 3.3083 | 3.0333 | 3.0183 | 3.0633 | 2.4713 | 3.3167 | 3.0011 | 2.3977 | 2.9183
| Median |  | 3.19 | 3 | 3 | 2.715 | 3.0143 | 3 | 2.83 | 3.8807 | 2.9 | 2.95 | 2.915 | 2.925 | 2.995 | 2.845 | 2.78 | 3.055 | 3.015 | 3.325 | 3.1 | 3.025 | 3.045 | 2.4827 | 3.3 | 3.0176 | 2.3985 | 2.945
| Std Dev. |  | 0.13426 | 0 | 0 | 0.077717 | 0.057287 | 0.24221 | 0.11537 | 0.067438 | 0.1169 | 0.054772 | 0.035214 | 0.043665 | 0.04 | 0.0251 | 0.034303 | 0.04446 | 0.040866 | 0.065243 | 0.10328 | 0.063377 | 0.13706 | 0.026439 | 0.040825 | 0.07087 | 0.02388 | 0.13963
| Rel.Std.Dev. |  | 4.191% | 0% | 0% | 2.826% | 1.895% | 7.985% | 4.055% | 1.739% | 4.008% | 1.857% | 1.21% | 1.489% | 1.333% | 0.8822% | 1.23% | 1.457% | 1.36% | 1.972% | 3.405% | 2.1% | 4.474% | 1.07% | 1.231% | 2.361% | 0.996% | 4.785%
| PDM3 |  | 7.549% | 0.7227% | 0.7227% | -7.671% | 1.497% | 1.842% | -4.481% | 30.17% | -2.075% | -0.9561% | -2.299% | -1.516% | 0.7227% | -4.481% | -6.384% | 2.457% | 0.8905% | 11.07% | 1.842% | 1.338% | 2.849% | -17.03% | 11.35% | 0.7603% | -19.5% | -2.019%
| Z-Score (Absolute) |  | 1.65 | NA | NA | 2.03 | 0.18 | 0.27 | 1.26 | 7.12 | 0.68 | 0.41 | 0.74 | 0.55 | 0 | 1.26 | 1.72 | 0.41 | 0.04 | 2.5 | 0.27 | 0.14 | 0.51 | 4.3 | 2.57 | 0 | 4.89 | 0.67
Results from laboratories 2.03 and 2.04 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 606b (Certified Value 4.83 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 5.57 | 4.41 | 4.8 | 4.98 | 5.0296 | 4.99 | 4.9 | 4.86 | 4.86 | 4.87 | 4.67 | 4.95 | 4.62 | 4.67 | 4.34 | 4.7 | 4.74 | 4.8 | < 5 | 4.46 | < 0.1 | 4.8924 | < 5 | 3.77 |  |  | 1
| 1 | 2 | 5 | 4.76 | 5 | 5.05 | 5.1911 | 5 | 5.09 | 4.95 | 4.92 | 5.04 | 4.92 | 5.12 | 4.42 | 4.87 | 4.39 | 4.47 | 5.17 | 5.01 | < 5 | 4.37 | < 0.1 | 4.8353 | < 5 | 3.75 |  |  | 26
| 1 | 3 | 5.41 | 4.62 | 5 | 4.99 | 5.1308 | 4.89 | 4.98 | 4.86 | 4.95 | 5.07 | 4.8 | 5.02 | 4.55 | 4.71 | 4.2 | 4.7 | 4.7 | 4.8 | < 5 | 4.41 | < 0.1 | 4.6007 | < 5 | 3.65 |  |  | 16
| 1 | 4 | 5.49 | 4.43 | 4.8 | 5.01 | 5.0386 | 4.9 | 4.91 | 4.78 | 4.88 | 4.88 | 4.82 | 5.18 | 4.41 | 4.69 | 4.39 | 4.63 | 4.65 | 4.97 | < 5 | 4.48 | < 0.1 | 4.6949 | < 5 | 3.74 |  |  | 4.8283
| 1 | 5 | 5.27 | 4.84 | 4.9 | 4.94 | 5.0149 | 4.81 | 4.97 | 4.82 | 4.96 | 4.91 | 4.89 | 5.17 | 4.68 | 4.71 | 4.34 | 4.74 | 4.86 | 4.86 | < 5 | 4.47 | < 0.1 | 4.84 | < 5 | 3.69 |  |  | 18
| 1 | 6 | 5.59 | 4.92 | 4.9 | 4.95 | 5.0286 | 4.89 | 5.18 | 4.64 | 4.89 | 4.94 | 4.71 | 5.01 | 4.53 | 4.7 | 4.33 | 4.63 | 5 | 4.88 | < 5 | 4.43 | < 0.1 | 4.5435 | < 5 | 3.73
| Mean |  | 5.3883 | 4.6633 | 4.9 | 4.9867 | 5.0723 | 4.9133 | 5.005 | 4.8183 | 4.91 | 4.9517 | 4.8017 | 5.075 | 4.535 | 4.725 | 4.3317 | 4.645 | 4.8533 | 4.8867 |  | 4.4367 |  | 4.7345 |  | 3.7217
| Median |  | 5.45 | 4.69 | 4.9 | 4.985 | 5.0341 | 4.895 | 4.975 | 4.84 | 4.905 | 4.925 | 4.81 | 5.07 | 4.54 | 4.705 | 4.34 | 4.665 | 4.8 | 4.87 |  | 4.445 |  | 4.7651 |  | 3.735
| Std Dev. |  | 0.22329 | 0.21304 | 0.089443 | 0.040332 | 0.071695 | 0.071181 | 0.10932 | 0.104 | 0.04 | 0.084242 | 0.09786 | 0.094816 | 0.1071 | 0.072595 | 0.06969 | 0.096073 | 0.19977 | 0.087101 |  | 0.041793 |  | 0.14297 |  | 0.044008
| Rel.Std.Dev. |  | 4.144% | 4.568% | 1.825% | 0.8088% | 1.413% | 1.449% | 2.184% | 2.158% | 0.8147% | 1.701% | 2.038% | 1.868% | 2.362% | 1.536% | 1.609% | 2.068% | 4.116% | 1.782% |  | 0.942% |  | 3.02% |  | 1.182%
| PDM3 |  | 11.6% | -3.417% | 1.484% | 3.279% | 5.052% | 1.76% | 3.659% | -0.2071% | 1.691% | 2.554% | -0.5523% | 5.109% | -6.075% | -2.14% | -10.29% | -3.797% | 0.5178% | 1.208% |  | -8.112% |  | -1.944% |  | -22.92%
| Z-Score (Absolute) |  | 2.28 | 0.58 | 0.36 | 0.7 | 1.04 | 0.41 | 0.77 | 0.03 | 0.39 | 0.56 | 0.03 | 1.05 | 1.09 | 0.34 | 1.89 | 0.65 | 0.17 | 0.3 | 9.12 | 1.47 | 18.79 | 0.3 | 9.12 | 4.3
| Analytical results for Ca in OREAS 606b (Certified Value 1.1 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.12 | 1.08 | 1.11 | 1.1132 | 1.1164 | 1.15 | 1.1 | 1.1282 | 1.07 | 1.15 | 1.1194 | 1.13 | 1.05 | 1.08 | 1.05 | 1.14 | 1.1425 | 0.917 | 1.25 | 1.04 | 1.09 | 1.1072 | 1.129 | 1.0996 | 1.0792 | 0.994 | 1
| 1 | 2 | 1.13 | 1.12 | 1.13 | 1.0971 | 1.114 | 1.12 | 1.06 | 1.1225 | 1.11 | 1.19 | 1.1214 | 1.1 | 1.06 | 1.06 | 1.1 | 1.12 | 1.1074 | 0.984 | 1.21 | 1.06 | 1.06 | 1.1025 | 1.093 | 1.1079 | 1.1017 | 0.983 | 0
| 1 | 3 | 1.14 | 1.1 | 1.1 | 1.0938 | 1.1157 | 1.14 | 1.05 | 1.1097 | 1.16 | 1.15 | 1.1323 | 1.11 | 1.06 | 1.08 | 1.06 | 1.1 | 1.1405 | 0.916 | 1.11 | 1.06 | 1.1 | 1.0956 | 1.108 | 1.0851 | 1.0949 | 0.971 | 16
| 1 | 4 | 1.14 | 1.05 | 1.1 | 1.1331 | 1.1085 | 1.13 | 1.04 | 1.1023 | 1.08 | 1.15 | 1.125 | 1.13 | 1.06 | 1.05 | 1.04 | 1.14 | 1.1374 | 0.881 | 1.19 | 1.06 | 1.09 | 1.1321 | 1.086 | 1.0757 | 1.0479 | 0.952 | 1.1
| 1 | 5 | 1.12 | 1.07 | 1.09 | 1.1156 | 1.0986 | 1.11 | 1.05 | 1.1387 | 1.13 | 1.15 | 1.1098 | 1.14 | 1.1 | 1.07 | 1.05 | 1.13 | 1.1579 | 0.88 | 1.2 | 1.06 | 1.07 | 1.1255 | 1.093 | 1.1003 | 1.0348 | 0.985 | 19
| 1 | 6 | 1.12 | 1.24 | 1.1 | 1.0692 | 1.1091 | 1.12 | 1.04 | 1.117 | 1.07 | 1.13 | 1.1097 | 1.1 | 1.07 | 1.08 | 0.96 | 1.11 | 1.1279 | 0.975 | 1.16 | 1.07 | 1.1 | 1.1257 | 1.115 | 1.1194 | 1.1037 | 0.991
| Mean |  | 1.1283 | 1.11 | 1.105 | 1.1037 | 1.1104 | 1.1283 | 1.0567 | 1.1197 | 1.1033 | 1.1533 | 1.1196 | 1.1183 | 1.0667 | 1.07 | 1.0433 | 1.1233 | 1.1356 | 0.9255 | 1.1867 | 1.0583 | 1.085 | 1.1148 | 1.104 | 1.098 | 1.077 | 0.97933
| Median |  | 1.125 | 1.09 | 1.1 | 1.1052 | 1.1116 | 1.125 | 1.05 | 1.1197 | 1.095 | 1.15 | 1.1204 | 1.12 | 1.06 | 1.075 | 1.05 | 1.125 | 1.1389 | 0.9165 | 1.195 | 1.06 | 1.09 | 1.1163 | 1.1005 | 1.1 | 1.087 | 0.984
| Std Dev. |  | 0.0098319 | 0.068118 | 0.013784 | 0.022036 | 0.006671 | 0.01472 | 0.022509 | 0.013036 | 0.036697 | 0.019664 | 0.0088066 | 0.017224 | 0.017512 | 0.012649 | 0.045898 | 0.01633 | 0.016892 | 0.044912 | 0.04761 | 0.0098319 | 0.016432 | 0.014901 | 0.016297 | 0.015666 | 0.029256 | 0.015578
| Rel.Std.Dev. |  | 0.8714% | 6.137% | 1.247% | 1.997% | 0.6008% | 1.305% | 2.13% | 1.164% | 3.326% | 1.705% | 0.7866% | 1.54% | 1.642% | 1.182% | 4.399% | 1.454% | 1.487% | 4.853% | 4.012% | 0.929% | 1.514% | 1.337% | 1.476% | 1.427% | 2.716% | 1.591%
| PDM3 |  | 2.572% | 0.9053% | 0.4507% | 0.3295% | 0.9401% | 2.572% | -3.943% | 1.789% | 0.2992% | 4.844% | 1.778% | 1.663% | -3.034% | -2.731% | -5.155% | 2.117% | 3.232% | -15.87% | 7.875% | -3.792% | -1.367% | 1.339% | 0.3598% | -0.1857% | -2.092% | -10.97%
| Z-Score (Absolute) |  | 0.74 | 0.17 | 0.01 | 0.03 | 0.18 | 0.74 | 1.49 | 0.47 | 0.04 | 1.52 | 0.47 | 0.43 | 1.18 | 1.08 | 1.91 | 0.58 | 0.96 | 5.57 | 2.55 | 1.44 | 0.61 | 0.32 | 0.01 | 0.21 | 0.86 | 3.9
| Analytical results for Cd in OREAS 606b (Certified Value 1.24 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 1.66 | 1.12 | 1.2 | 1.23 | 1.1081 | 1.8 | 1.24 | 3.3184 | 1.42 | 1.23 | 1.29 | 1.26 | 1.19 | 1.38 | 1.18 | 1.18 | 1.19 | 1.14 | 1.2 | 1.26 | 1.7 | 1.1988 | < 1 | 1.2457 | < 5 | 1.14 | 1
| 1 | 2 | 1.71 | 1.37 | 1.2 | 1.2 | 1.1539 | 1.8 | 1.3 | 3.5117 | 1.41 | 1.22 | 1.26 | 1.32 | 1.14 | 1.33 | 1.3 | 1.28 | 1.18 | 1.23 | 1.2 | 1.2 | 1.62 | 1.1445 | < 1 | 1.0661 | < 5 | 1.14 | 27
| 1 | 3 | 1.61 | 1.18 | 1.2 | 1.15 | 1.1746 | 1.7 | 1.28 | 3.7272 | 1.53 | 1.23 | 1.34 | 1.33 | 1.22 | 1.34 | 1.23 | 1.22 | 1.18 | 1.17 | 1.4 | 1.13 | 1.62 | 1.2699 | < 1 | 1.112 | < 5 | 1.27 | 16
| 1 | 4 | 1.5 | 1.12 | 1.2 | 1.37 | 1.1487 | 1.8 | 1.28 | 3.4819 | 1.5 | 1.28 | 1.22 | 1.33 | 1.24 | 1.32 | 1.3 | 1.32 | 1.18 | 1.21 | 1 | 1.48 | 1.71 | 1.3374 | < 1 | 1.219 | < 5 | 1.33 | 1.2405
| 1 | 5 | 1.54 | 1.22 | 1.2 | 1.36 | 1.1635 | 1.8 | 1.34 | 3.6984 | 1.44 | 1.26 | 1.24 | 1.37 | 1.18 | 1.38 | 1.21 | 1.23 | 1.24 | 1.18 | 1 | 1.05 | 1.71 | 1.176 | < 1 | 1.1724 | < 5 | 1.25 | 20
| 1 | 6 | 1.46 | 1.27 | 1.2 | 1.18 | 1.0832 | 1.8 | 1.31 | 3.5252 | 1.38 | 1.22 | 1.25 | 1.3 | 1.11 | 1.34 | 1.23 | 1.27 | 1.2 | 1.22 | 1.2 | 1.42 | 1.55 | 1.3764 | < 1 | 1.1097 | < 5 | 1.19
| Mean |  | 1.58 | 1.2133 | 1.2 | 1.2483 | 1.1387 | 1.7833 | 1.2917 | 3.5438 | 1.4467 | 1.24 | 1.2667 | 1.3183 | 1.18 | 1.3483 | 1.2417 | 1.25 | 1.195 | 1.1917 | 1.1667 | 1.2567 | 1.6517 | 1.2505 |  | 1.1542 |  | 1.22
| Median |  | 1.575 | 1.2 | 1.2 | 1.215 | 1.1513 | 1.8 | 1.29 | 3.5185 | 1.43 | 1.23 | 1.255 | 1.325 | 1.185 | 1.34 | 1.23 | 1.25 | 1.185 | 1.195 | 1.2 | 1.23 | 1.66 | 1.2344 |  | 1.1422 |  | 1.22
| Std Dev. |  | 0.09654 | 0.096264 | 0 | 0.09411 | 0.035372 | 0.040825 | 0.033714 | 0.1507 | 0.057155 | 0.024495 | 0.04274 | 0.03656 | 0.04858 | 0.025626 | 0.048751 | 0.0498 | 0.023452 | 0.034303 | 0.15055 | 0.16645 | 0.065549 | 0.092992 |  | 0.069881 |  | 0.07642
| Rel.Std.Dev. |  | 6.11% | 7.934% | 0% | 7.539% | 3.106% | 2.289% | 2.61% | 4.252% | 3.951% | 1.975% | 3.374% | 2.773% | 4.117% | 1.901% | 3.926% | 3.984% | 1.963% | 2.879% | 12.9% | 13.25% | 3.969% | 7.436% |  | 6.055% |  | 6.264%
| PDM3 |  | 27.37% | -2.187% | -3.262% | 0.6342% | -8.206% | 43.76% | 4.127% | 185.7% | 16.62% | -0.03764% | 2.112% | 6.277% | -4.875% | 8.696% | 0.09672% | 0.7685% | -3.665% | -3.934% | -5.949% | 1.306% | 33.15% | 0.8088% |  | -6.958% |  | -1.65%
| Z-Score (Absolute) |  | 3.52 | 0.38 | 0.52 | 0.01 | 1.18 | 5.69 | 0.45 | 24.43 | 2.1 | 0.1 | 0.19 | 0.74 | 0.74 | 1.06 | 0.08 | 0.01 | 0.58 | 0.61 | 0.88 | 0.08 | 4.29 | 0.01 | 7.98 | 1.01 | 13.32 | 0.31
| Analytical results for Ce in OREAS 606b (Certified Value 89 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.27 | 2.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 0
1
| 1 | 1 | 93 | 85.15 | 92 | 86.8 | 90.1755 | 101 | 95.1 | 110.64 | 91.6 | 92 | 86.3 | 83.1 | 85.3 | 83.8 | 94.6 | 87 | 61.75 | 87.5 | 87 | 85.4196 | 84 | 77.387 | 83.7 |  |  |  | 1
| 1 | 2 | 93 | 87.27 | 95 | 85.58 | 89.6248 | 95.7 | 96.8 | 110.68 | 94 | 93 | 87.8 | 84.2 | 83.8 | 88.8 | 97.3 | 85.4 | 68.96 | 91.6 | 85 | 86.0458 | 82 | 77.781 | 77.54 |  |  |  | 28
| 1 | 3 | 96 | 88.31 | 93 | 84.66 | 90.841 | 96.7 | 94.5 | 110.58 | 91.5 | 93 | 91.7 | 84.6 | 86.6 | 87.8 | 97.6 | 87.4 | 56.58 | 85.6 | 87 | 87.2173 | 82 | 78.032 | 85.03 |  |  |  | 16
| 1 | 4 | 98 | 85.96 | 92 | 89.05 | 89.2828 | 98.9 | 96.2 | 110.23 | 90.2 | 93 | 89 | 85.4 | 81.2 | 88.7 | 100 | 87.4 | 49.28 | 90.6 | 87 | 89.8639 | 82 | 75.097 | 81.81 |  |  |  | 88.8532
| 1 | 5 | 94 | 95.21 | 92 | 85.93 | 90.2427 | 97.3 | 95.5 | 110.45 | 92.4 | 90 | 93.1 | 87.3 | 86.6 | 87.2 | 97.9 | 88.9 | 57.64 | 91.2 | 87 | 91.1912 | 82 | 77.533 | 83.79 |  |  |  | 21
| 1 | 6 | 94 | 92.83 | 90 | 85.51 | 89.1061 | 100 | 92.8 | 110.95 | 91.9 | 93 | 87.1 | 86.6 | 85.3 | 66.8 | 94 | 86.5 | 64.8 | 89.7 | 85 | 87.3118 | 83 | 79.042 | 78.02
| Mean |  | 94.6667 | 89.1217 | 92.3333 | 86.255 | 89.8788 | 98.2667 | 95.15 | 110.5883 | 91.9333 | 92.3333 | 89.1667 | 85.2 | 84.8 | 83.85 | 96.9 | 87.1 | 59.835 | 89.3667 | 86.3333 | 87.8416 | 82.5 | 77.4787 | 81.6483
| Median |  | 94 | 87.79 | 92 | 85.755 | 89.9001 | 98.1 | 95.3 | 110.61 | 91.75 | 93 | 88.4 | 85 | 85.3 | 87.5 | 97.45 | 87.2 | 59.695 | 90.15 | 87 | 87.2645 | 82 | 77.657 | 82.755
| Std Dev. |  | 1.9664 | 4.0169 | 1.633 | 1.5342 | 0.65774 | 2.0422 | 1.4068 | 0.24062 | 1.2485 | 1.2111 | 2.6934 | 1.5608 | 2.0465 | 8.5505 | 2.2343 | 1.1559 | 6.9051 | 2.3517 | 1.0328 | 2.2392 | 0.83666 | 1.3058 | 3.1718
| Rel.Std.Dev. |  | 2.077% | 4.507% | 1.769% | 1.779% | 0.7318% | 2.078% | 1.478% | 0.2176% | 1.358% | 1.312% | 3.021% | 1.832% | 2.413% | 10.2% | 2.306% | 1.327% | 11.54% | 2.632% | 1.196% | 2.549% | 1.014% | 1.685% | 3.885%
| PDM3 |  | 6.543% | 0.3021% | 3.917% | -2.924% | 1.154% | 10.59% | 7.087% | 24.46% | 3.466% | 3.917% | 0.3527% | -4.112% | -4.562% | -5.631% | 9.056% | -1.973% | -32.66% | 0.5778% | -2.836% | -1.139% | -7.15% | -12.8% | -8.109%
| Z-Score (Absolute) |  | 0.95 | 0 | 0.55 | 0.49 | 0.13 | 1.57 | 1.04 | 3.69 | 0.48 | 0.55 | 0.01 | 0.67 | 0.74 | 0.91 | 1.34 | 0.35 | 5.04 | 0.04 | 0.48 | 0.22 | 1.14 | 2 | 1.28
| Analytical results for Co in OREAS 606b (Certified Value 2.25 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.6 | 2.1 | 2.3 | 2.2 | 1.9942 | 2 | 2.3 | 2.2 | 2 | 2.2 | 2.3 | 2.1 | 2.2 | 2.2 | 2.2 | 2.4 | 2.4 | 2.5 | 4.74 | 2.2 | 13.1306 | 1 | 2.1721 | < 5 | 1.9 |  | 1
| 1 | 2 | 2.8 | 2.2 | 2.4 | 2.1 | 2.0949 | 2 | 2.39 | 2.13 | 2 | 2.2 | 2.4 | 2.1 | 2.2 | 2.3 | 2.3 | 2.5 | 2.6 | 2.49 | 4.92 | 2.2 | 16.5683 | 1 | 2.2261 | < 5 | 1.9 |  | 16
| 1 | 3 | 2.7 | 2.2 | 2.3 | 2.1 | 2.0461 | 2 | 2.39 | 2.24 | 2 | 2.3 | 2.4 | 2.1 | 2.2 | 2.3 | 2.2 | 2.4 | 2.5 | 2.3 | 4.7 | 2.2 | 14.8746 | 1 | 2.1119 | < 5 | 1.9 |  | 16
| 1 | 4 | 2.4 | 2.1 | 2.3 | 2.2 | 2.0003 | 2 | 2.31 | 2.22 | 2 | 2.2 | 2.4 | 2.1 | 2.2 | 2.2 | 2.3 | 2.5 | 2.4 | 2.5 | 4.82 | 2.2 | 14.7844 | 1 | 2.1214 | < 5 | 2 |  | 2.2525
| 1 | 5 | 2.7 | 2.2 | 2.3 | 2.2 | 2.0146 | 2 | 2.4 | 2.24 | 2 | 2.2 | 2.5 | 2.2 | 2.3 | 2.2 | 2.3 | 2.4 | 2.5 | 2.47 | 4.71 | 2.2 | 13.7076 | 1 | 2.2809 | < 5 | 1.9 |  | 22
| 1 | 6 | 2.6 | 2.4 | 2.3 | 2.2 | 2.0452 | 2 | 2.35 | 2.14 | 2 | 2.3 | 2.3 | 2.1 | 2.2 | 2.2 | 2.3 | 2.3 | 2.5 | 2.5 | 4.67 | 2.3 | 23.3287 | < 1 | 2.2266 | < 5 | 2
| Mean |  | 2.6333 | 2.2 | 2.3167 | 2.1667 | 2.0325 | 2 | 2.3567 | 2.195 | 2 | 2.2333 | 2.3833 | 2.1167 | 2.2167 | 2.2333 | 2.2667 | 2.4167 | 2.4833 | 2.46 | 4.76 | 2.2167 | 16.0657 | 1 | 2.1898 |  | 1.9333
| Median |  | 2.65 | 2.2 | 2.3 | 2.2 | 2.0299 | 2 | 2.37 | 2.21 | 2 | 2.2 | 2.4 | 2.1 | 2.2 | 2.2 | 2.3 | 2.4 | 2.5 | 2.495 | 4.725 | 2.2 | 14.8295 | 1 | 2.1991 |  | 1.9
| Std Dev. |  | 0.13663 | 0.10954 | 0.040825 | 0.05164 | 0.037582 | 0 | 0.043665 | 0.048888 | 0 | 0.05164 | 0.075277 | 0.040825 | 0.040825 | 0.05164 | 0.05164 | 0.075277 | 0.075277 | 0.079246 | 0.093595 | 0.040825 | 3.7481 | 0 | 0.066378 |  | 0.05164
| Rel.Std.Dev. |  | 5.188% | 4.979% | 1.762% | 2.383% | 1.849% | 0% | 1.853% | 2.227% | 0% | 2.312% | 3.158% | 1.929% | 1.842% | 2.312% | 2.278% | 3.115% | 3.031% | 3.221% | 1.966% | 1.842% | 23.33% | 0% | 3.031% |  | 2.671%
| PDM3 |  | 16.91% | -2.329% | 2.85% | -3.809% | -9.764% | -11.21% | 4.626% | -2.551% | -11.21% | -0.8495% | 5.81% | -6.029% | -1.589% | -0.8495% | 0.6303% | 7.29% | 10.25% | 9.214% | 111.3% | -1.589% | 613.2% | -55.6% | -2.78% |  | -14.17%
| Z-Score (Absolute) |  | 2.19 | 0.18 | 0.46 | 0.36 | 1.1 | NA | 0.67 | 0.21 | 1.28 | 0 | 0.82 | 0.64 | 0.09 | 0 | 0.18 | 1 | 1.37 | 1.24 | 13.81 | 0.09 | 75.63 | NA | 0.24 | 1.46 | 1.64
Results from laboratories 2.07 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 606b (Certified Value 14.7 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 1
1
| 1 | 1 | 17 | 9 | 12 | 18 | 15.4073 | 14 | 13.3 | 14.7 | 15 | 14 | 14 | 14 | 14 | 14 | 17 | 12.9 | 12 | 13 | 20.24 | 18 | 67.037 | 14 | 14.8474 | 14.42 | < 10 |  | 1
| 1 | 2 | 17 | 11 | 12 | 16 | 15.2698 | 15 | 13.2 | 14 | 14 | 16 | 14 | 14 | 14 | 13 | 15 | 15.3 | 13 | 12 | 23.43 | 18 | 64.6556 | 15 | 16.0984 | 17.2 | < 10 |  | 30
| 1 | 3 | 17 | 10 | 12 | 19 | 15.2494 | 15 | 12.6 | 13.7 | 14 | 15 | 15 | 14 | 14 | 14 | 18 | 13.6 | 11 | 14 | 25.32 | 16 | 71.3486 | 15 | 13.2874 | 16.953 | < 10 |  | 16
| 1 | 4 | 17 | 10 | 13 | 17 | 15.1388 | 15 | 12.3 | 15 | 14 | 16 | 14 | 14 | 14 | 14 | 16 | 13.6 | 13 | 14 | 29.99 | 16 | 68.5954 | 15 | 15.1357 | 17.519 | < 10 |  | 14.7162
| 1 | 5 | 18 | 12 | 12 | 17 | 15.4713 | 15 | 13.2 | 14.7 | 14 | 15 | 16 | 14 | 13 | 14 | 15 | 15.5 | 10 | 13 | 25.39 | 17 | 71.3229 | 14 | 15.1997 | 15.734 | 11 |  | 23
| 1 | 6 | 17 | 10 | 14 | 18 | 15.2589 | 15 | 13.9 | 14.3 | 16 | 16 | 14 | 14 | 15 | 13 | 15 | 13.2 | 12 | 16 | 20.7 | 18 | 78.8015 | 15 | 13.2254 | 15.829 | < 10
| Mean |  | 17.1667 | 10.3333 | 12.5 | 17.5 | 15.2993 | 14.8333 | 13.0833 | 14.4 | 14.5 | 15.3333 | 14.5 | 14 | 14 | 13.6667 | 16 | 14.0167 | 11.8333 | 13.6667 | 24.1783 | 17.1667 | 70.2935 | 14.6667 | 14.6324 | 16.2758 | 11
| Median |  | 17 | 10 | 12 | 17.5 | 15.2644 | 15 | 13.2 | 14.5 | 14 | 15.5 | 14 | 14 | 14 | 14 | 15.5 | 13.6 | 12 | 13.5 | 24.375 | 17.5 | 69.9591 | 15 | 14.9916 | 16.391 | 11
| Std Dev. |  | 0.40825 | 1.0328 | 0.83666 | 1.0488 | 0.11998 | 0.40825 | 0.56362 | 0.4899 | 0.83666 | 0.8165 | 0.83666 | 0 | 0.63246 | 0.5164 | 1.2649 | 1.1053 | 1.169 | 1.3663 | 3.5977 | 0.98319 | 4.8956 | 0.5164 | 1.1454 | 1.1658
| Rel.Std.Dev. |  | 2.378% | 9.995% | 6.693% | 5.993% | 0.7843% | 2.752% | 4.308% | 3.402% | 5.77% | 5.325% | 5.77% | 0% | 4.518% | 3.779% | 7.906% | 7.886% | 9.879% | 9.997% | 14.88% | 5.727% | 6.964% | 3.521% | 7.828% | 7.163%
| PDM3 |  | 16.65% | -29.78% | -15.06% | 18.92% | 3.962% | 0.7958% | -11.1% | -2.149% | -1.469% | 4.193% | -1.469% | -4.867% | -4.867% | -7.132% | 8.724% | -4.754% | -19.59% | -7.132% | 64.3% | 16.65% | 377.7% | -0.3368% | -0.5699% | 10.6% | -25.25%
| Z-Score (Absolute) |  | 2.16 | 3.37 | 1.62 | 2.43 | 0.65 | 0.27 | 1.15 | 0.08 | 0 | 0.67 | 0 | 0.4 | 0.4 | 0.67 | 1.21 | 0.39 | 2.16 | 0.67 | 7.83 | 2.16 | 45.15 | 0.13 | 0.11 | 1.44 | 6.88
| Analytical results for Cs in OREAS 606b (Certified Value 6.11 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 6.18 | 5 | 6.18 | 5.7661 | 6.41 | 6.4 | 6.5648 | 5.99 | 6.1 | 6.43 | 6.1 | 5.64 | 5.78 | 6.06 | 6.07 | 5.8 | 4.73 | 6.27 | 6.25 | 6.4249 | 5.22 | [5 empty] | 1
| 1 | 2 | 6.41 | 5.2 | 6.19 | 5.9985 | 6.33 | 6.77 | 6.4867 | 6.01 | 6.1 | 6.49 | 6.17 | 5.81 | 5.66 | 6.35 | 6.27 | 5.8 | 4.96 | 6.63 | 6.12 | 6.3789 | 5.24 | [5 empty] | 31
| 1 | 3 | 6.64 | 5.1 | 6.11 | 5.9645 | 6.48 | 6.84 | 6.7704 | 6.19 | 6.1 | 6.57 | 6.18 | 5.86 | 5.92 | 6.18 | 6.13 | 5.8 | 4.69 | 6.06 | 6.18 | 6.2497 | 5.24 | [5 empty] | 16
| 1 | 4 | 6.77 | 5 | 6.15 | 5.8739 | 6.4 | 6.49 | 6.6351 | 5.93 | 6.1 | 6.47 | 6.23 | 5.9 | 5.53 | 6.11 | 6.39 | 5.8 | 4.55 | 6.47 | 6.24 | 6.1967 | 5.19 | [5 empty] | 6.112
| 1 | 5 | 6.5 | 5.4 | 6.18 | 5.9673 | 6.18 | 6.64 | 6.7963 | 6.16 | 6.2 | 6.3 | 6.52 | 5.97 | 5.94 | 6.19 | 6.27 | 5.9 | 4.74 | 6.74 | 6.19 | 6.43 | 5.24 | [5 empty] | 24
| 1 | 6 | 6.5 | 5.8 | 6.18 | 6.0276 | 6.52 | 6.5 | 6.5945 | 5.93 | 6.2 | 6.37 | 6.13 | 5.82 | 5.76 | 5.86 | 6.12 | 5.7 | 4.95 | 6.49 | 6.17 | 6.504 | 5.22
| Mean |  | 6.5 | 5.25 | 6.165 | 5.933 | 6.3867 | 6.6067 | 6.6413 | 6.035 | 6.1333 | 6.4383 | 6.2217 | 5.8333 | 5.765 | 6.125 | 6.2083 | 5.8 | 4.77 | 6.4433 | 6.1917 | 6.364 | 5.225
| Median |  | 6.5 | 5.15 | 6.18 | 5.9659 | 6.405 | 6.57 | 6.6148 | 6 | 6.1 | 6.45 | 6.175 | 5.84 | 5.77 | 6.145 | 6.2 | 5.8 | 4.735 | 6.48 | 6.185 | 6.4019 | 5.23
| Std Dev. |  | 0.20149 | 0.30822 | 0.030166 | 0.096702 | 0.12094 | 0.17317 | 0.12056 | 0.11345 | 0.05164 | 0.094745 | 0.15276 | 0.1113 | 0.15566 | 0.16282 | 0.12139 | 0.063246 | 0.15862 | 0.24606 | 0.047924 | 0.11742 | 0.019748
| Rel.Std.Dev. |  | 3.1% | 5.871% | 0.4893% | 1.63% | 1.894% | 2.621% | 1.815% | 1.88% | 0.842% | 1.472% | 2.455% | 1.908% | 2.7% | 2.658% | 1.955% | 1.09% | 3.325% | 3.819% | 0.774% | 1.845% | 0.378%
| PDM3 |  | 6.348% | -14.1% | 0.8668% | -2.929% | 4.494% | 8.093% | 8.66% | -1.26% | 0.3487% | 5.339% | 1.794% | -4.56% | -5.678% | 0.2124% | 1.576% | -5.105% | -21.96% | 5.421% | 1.303% | 4.123% | -14.51%
| Z-Score (Absolute) |  | 0.83 | 2.26 | 0 | 0.57 | 0.55 | 1.09 | 1.18 | 0.32 | 0.08 | 0.67 | 0.14 | 0.82 | 0.99 | 0.1 | 0.11 | 0.9 | 3.44 | 0.69 | 0.07 | 0.49 | 2.32
| Analytical results for Cu in OREAS 606b (Certified Value 0.033 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0348 | 0.03561 | 0.03321 | 0.03277 | 0.034298 | 0.0326 | 0.0358 | 0.034123 | 0.0339 | 0.0353 | 0.0337 | 0.0331 | 0.0324 | 0.0321 | 0.0332 | 0.034 | 0.03311 | 0.034 | 0.0344 | 0.03038 | 0.03209 | 0.033297 | 0.0345 | 0.032854 | 0.031273 | 0.0307 | 1
| 1 | 2 | 0.035 | 0.03488 | 0.03334 | 0.03274 | 0.034861 | 0.0326 | 0.0359 | 0.034137 | 0.033 | 0.0358 | 0.0337 | 0.0335 | 0.0334 | 0.0309 | 0.0336 | 0.0335 | 0.03219 | 0.0348 | 0.0354 | 0.03174 | 0.03353 | 0.033615 | 0.0336 | 0.033241 | 0.032129 | 0.0313 | 18
| 1 | 3 | 0.0351 | 0.03587 | 0.03298 | 0.03227 | 0.034544 | 0.0319 | 0.0356 | 0.033592 | 0.035 | 0.0352 | 0.0344 | 0.0336 | 0.0327 | 0.0315 | 0.0329 | 0.0329 | 0.03254 | 0.0344 | 0.0343 | 0.03008 | 0.03193 | 0.032959 | 0.0335 | 0.03272 | 0.03186 | 0.03 | 16
| 1 | 4 | 0.0354 | 0.03642 | 0.03254 | 0.03413 | 0.034172 | 0.0324 | 0.0351 | 0.034149 | 0.0334 | 0.036 | 0.0337 | 0.0337 | 0.0334 | 0.0311 | 0.0329 | 0.0342 | 0.03235 | 0.0339 | 0.0347 | 0.02998 | 0.03278 | 0.03422 | 0.0334 | 0.032592 | 0.031442 | 0.0307 | 0.033391
| 1 | 5 | 0.0342 | 0.0352 | 0.03279 | 0.03281 | 0.03474 | 0.0318 | 0.0355 | 0.034054 | 0.0328 | 0.0349 | 0.0335 | 0.0344 | 0.0336 | 0.0318 | 0.0329 | 0.034 | 0.03352 | 0.035 | 0.035 | 0.03088 | 0.03255 | 0.033464 | 0.0334 | 0.033599 | 0.03197 | 0.0308 | 25
| 1 | 6 | 0.035 | 0.03622 | 0.03327 | 0.03219 | 0.034595 | 0.0326 | 0.0352 | 0.034245 | 0.0329 | 0.0363 | 0.0335 | 0.0328 | 0.0332 | 0.0313 | 0.0328 | 0.0333 | 0.03207 | 0.0346 | 0.0344 | 0.03027 | 0.03197 | 0.033877 | 0.034 | 0.03373 | 0.032416 | 0.0301
| Mean |  | 0.034917 | 0.0357 | 0.033022 | 0.032818 | 0.034535 | 0.032317 | 0.035517 | 0.03405 | 0.0335 | 0.035583 | 0.03375 | 0.033517 | 0.033117 | 0.03145 | 0.03305 | 0.03365 | 0.03263 | 0.03445 | 0.0347 | 0.030555 | 0.032475 | 0.033572 | 0.033733 | 0.033122 | 0.031848 | 0.0306
| Median |  | 0.035 | 0.03574 | 0.033095 | 0.032755 | 0.03457 | 0.0325 | 0.03555 | 0.03413 | 0.0332 | 0.03555 | 0.0337 | 0.03355 | 0.0333 | 0.0314 | 0.0329 | 0.03375 | 0.032445 | 0.0345 | 0.03455 | 0.030325 | 0.03232 | 0.03354 | 0.03355 | 0.033047 | 0.031915 | 0.0307
| Std Dev. |  | 0.00040208 | 0.0005913 | 0.00031173 | 0.00069635 | 0.00026061 | 0.00037103 | 0.00031885 | 0.00023279 | 0.00083905 | 0.00053448 | 0.00033317 | 0.00054924 | 0.00046655 | 0.00044609 | 0.00030166 | 0.000501 | 0.00056854 | 0.00043704 | 0.00042895 | 0.00065999 | 0.00061838 | 0.00044205 | 0.00043665 | 0.00047463 | 0.00042724 | 0.00048166
| Rel.Std.Dev. |  | 1.152% | 1.656% | 0.944% | 2.122% | 0.7546% | 1.148% | 0.8978% | 0.6837% | 2.505% | 1.502% | 0.9872% | 1.639% | 1.409% | 1.418% | 0.9127% | 1.489% | 1.742% | 1.269% | 1.236% | 2.16% | 1.904% | 1.317% | 1.294% | 1.433% | 1.341% | 1.574%
| PDM3 |  | 4.568% | 6.914% | -1.107% | -1.716% | 3.425% | -3.219% | 6.365% | 1.972% | 0.3251% | 6.564% | 1.074% | 0.375% | -0.8229% | -5.814% | -1.023% | 0.7743% | -2.28% | 3.17% | 3.919% | -8.495% | -2.745% | 0.5414% | 1.024% | -0.8055% | -4.621% | -8.36%
| Z-Score (Absolute) |  | 1.04 | 1.62 | 0.36 | 0.51 | 0.76 | 0.88 | 1.49 | 0.4 | 0.01 | 1.54 | 0.18 | 0.01 | 0.29 | 1.53 | 0.34 | 0.1 | 0.65 | 0.7 | 0.88 | 2.19 | 0.77 | 0.05 | 0.17 | 0.29 | 1.23 | 2.16
| Analytical results for Dy in OREAS 606b (Certified Value 3.55 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 3.5; 2.03: 3.53; 2.04: 3.379; 2.05: 3.9; 2.06: 4.8067; 2.07: 3.5; 2.08: 3.08; 2.09: 3.64; 2.10: 3.5087; 2.11: 3.7; 2.12: 3.77 | 1
| 1 | 2 | 2.01: 3.4; 2.03: 3.47; 2.04: 3.5549; 2.05: 3.8; 2.06: 4.6368; 2.07: 3.68; 2.08: 3.28; 2.09: 3.55; 2.10: 3.4505; 2.11: 3.6; 2.12: 3.45 | 33
| 1 | 3 | 2.01: 3.5; 2.03: 3.48; 2.04: 3.4645; 2.05: 3.7; 2.06: 4.9371; 2.07: 3.71; 2.08: 2.91; 2.09: 3.61; 2.10: 3.4349; 2.11: 3.6; 2.12: 3.71 | 16
| 1 | 4 | 2.01: 3.4; 2.03: 3.6; 2.04: 3.2629; 2.05: 3.9; 2.06: 4.8737; 2.07: 3.62; 2.08: 2.81; 2.09: 3.53; 2.10: 3.5526; 2.11: 3.6; 2.12: 3.63 | 3.5543
| 1 | 5 | 2.01: 3.8; 2.03: 3.51; 2.04: 3.3598; 2.05: 3.7; 2.06: 4.5627; 2.07: 3.63; 2.08: 2.96; 2.09: 3.52; 2.10: 3.658; 2.11: 3.5; 2.12: 3.62 | 26
| 1 | 6 | 2.01: 3.8; 2.03: 3.46; 2.04: 3.5063; 2.05: 3.9; 2.06: 4.61; 2.07: 3.66; 2.08: 3.2; 2.09: 3.67; 2.10: 3.553; 2.11: 3.6; 2.12: 3.37
| Mean |  | 2.01: 3.5667; 2.03: 3.5083; 2.04: 3.4212; 2.05: 3.8167; 2.06: 4.7378; 2.07: 3.6333; 2.08: 3.04; 2.09: 3.5867; 2.10: 3.5263; 2.11: 3.6; 2.12: 3.5917
| Median |  | 2.01: 3.5; 2.03: 3.495; 2.04: 3.4218; 2.05: 3.85; 2.06: 4.7218; 2.07: 3.645; 2.08: 3.02; 2.09: 3.58; 2.10: 3.5307; 2.11: 3.6; 2.12: 3.625
| Std Dev. |  | 2.01: 0.18619; 2.03: 0.051929; 2.04: 0.10732; 2.05: 0.098319; 2.06: 0.155; 2.07: 0.073121; 2.08: 0.17944; 2.09: 0.062183; 2.10: 0.081436; 2.11: 0.063246; 2.12: 0.15316
| Rel.Std.Dev. |  | 2.01: 5.22%; 2.03: 1.48%; 2.04: 3.137%; 2.05: 2.576%; 2.06: 3.272%; 2.07: 2.013%; 2.08: 5.903%; 2.09: 1.734%; 2.10: 2.309%; 2.11: 1.757%; 2.12: 4.264%
| PDM3 |  | 2.01: 0.3487%; 2.03: -1.292%; 2.04: -3.743%; 2.05: 7.383%; 2.06: 33.3%; 2.07: 2.224%; 2.08: -14.47%; 2.09: 0.9114%; 2.10: -0.7875%; 2.11: 1.287%; 2.12: 1.052%
| Z-Score (Absolute) |  | 2.01: 0.22; 2.03: 0.87; 2.04: 1.85; 2.05: 2.57; 2.06: 12.86; 2.07: 0.52; 2.08: 6.1; 2.09: 0; 2.10: 0.67; 2.11: 0.15; 2.12: 0.06
| Analytical results for Er in OREAS 606b (Certified Value 1.01 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 0.9; 2.03: 1; 2.04: 0.93392; 2.05: 1.1; 2.06: 1.2825; 2.07: 1.04; 2.08: 0.9; 2.09: 1.09; 2.10: 1.0732; 2.11: < 0.2; 2.12: 1.08 | 1
| 1 | 2 | 2.01: 0.9; 2.03: 0.99; 2.04: 0.99717; 2.05: 1; 2.06: 1.2759; 2.07: 1; 2.08: 0.94; 2.09: 1.06; 2.10: 1.0533; 2.11: < 0.2; 2.12: 1.06 | 34
| 1 | 3 | 2.01: 1; 2.03: 0.98; 2.04: 0.97904; 2.05: 1; 2.06: 1.2346; 2.07: 1.01; 2.08: 0.85; 2.09: 1.06; 2.10: 1.082; 2.11: < 0.2; 2.12: 1.07 | 16
| 1 | 4 | 2.01: 0.9; 2.03: 1.02; 2.04: 0.96564; 2.05: 1; 2.06: 1.274; 2.07: 1.02; 2.08: 0.82; 2.09: 1.05; 2.10: 1.0861; 2.11: < 0.2; 2.12: 1.04 | 1.007
| 1 | 5 | 2.01: 1; 2.03: 1.01; 2.04: 0.95399; 2.05: 1; 2.06: 1.25; 2.07: 1; 2.08: 0.85; 2.09: 1.05; 2.10: 1.1468; 2.11: < 0.2; 2.12: 1.06 | 27
| 1 | 6 | 2.01: 1.1; 2.03: 1.02; 2.04: 0.94722; 2.05: 1.1; 2.06: 1.2527; 2.07: 1.06; 2.08: 0.92; 2.09: 1.02; 2.10: 1.0955; 2.11: < 0.2; 2.12: 1
| Mean |  | 2.01: 0.96667; 2.03: 1.0033; 2.04: 0.96283; 2.05: 1.0333; 2.06: 1.2616; 2.07: 1.0217; 2.08: 0.88; 2.09: 1.055; 2.10: 1.0895; 2.12: 1.0517
| Median |  | 2.01: 0.95; 2.03: 1.005; 2.04: 0.95982; 2.05: 1; 2.06: 1.2633; 2.07: 1.015; 2.08: 0.875; 2.09: 1.055; 2.10: 1.084; 2.12: 1.06
| Std Dev. |  | 2.01: 0.08165; 2.03: 0.01633; 2.04: 0.022837; 2.05: 0.05164; 2.06: 0.018643; 2.07: 0.024014; 2.08: 0.046904; 2.09: 0.022583; 2.10: 0.031506; 2.12: 0.028577
| Rel.Std.Dev. |  | 2.01: 8.447%; 2.03: 1.628%; 2.04: 2.372%; 2.05: 4.997%; 2.06: 1.478%; 2.07: 2.35%; 2.08: 5.33%; 2.09: 2.141%; 2.10: 2.892%; 2.12: 2.717%
| PDM3 |  | 2.01: -4.004%; 2.03: -0.3625%; 2.04: -4.385%; 2.05: 2.617%; 2.06: 25.29%; 2.07: 1.458%; 2.08: -12.61%; 2.09: 4.768%; 2.10: 8.193%; 2.12: 4.437%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.22; 2.04: 0.72; 2.05: 0.14; 2.06: 2.94; 2.07: 0; 2.08: 1.74; 2.09: 0.41; 2.10: 0.83; 2.11: 11.3; 2.12: 0.37
| Analytical results for Eu in OREAS 606b (Certified Value 1.4 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 1.2; 2.03: 1.22; 2.04: 1.357; 2.05: 1.34; 2.06: 2.7725; 2.07: 1.32; 2.08: 1.53; 2.09: 1.56; 2.10: 1.574; 2.11: 1.3; 2.12: 1.38 | 1
| 1 | 2 | 2.01: 1.2; 2.03: 1.24; 2.04: 1.3761; 2.05: 1.29; 2.06: 2.7181; 2.07: 1.31; 2.08: 1.65; 2.09: 1.55; 2.10: 1.581; 2.11: 1.2; 2.12: 1.31 | 35
| 1 | 3 | 2.01: 1.3; 2.03: 1.22; 2.04: 1.3923; 2.05: 1.32; 2.06: 2.7624; 2.07: 1.26; 2.08: 1.48; 2.09: 1.58; 2.10: 1.5656; 2.11: 1.2; 2.12: 1.32 | 16
| 1 | 4 | 2.01: 1.2; 2.03: 1.19; 2.04: 1.3794; 2.05: 1.31; 2.06: 2.7713; 2.07: 1.28; 2.08: 1.42; 2.09: 1.57; 2.10: 1.6057; 2.11: 1.2; 2.12: 1.28 | 1.3977
| 1 | 5 | 2.01: 1.3; 2.03: 1.22; 2.04: 1.3514; 2.05: 1.31; 2.06: 2.7768; 2.07: 1.25; 2.08: 1.51; 2.09: 1.55; 2.10: 1.6641; 2.11: 1.2; 2.12: 1.3 | 28
| 1 | 6 | 2.01: 1.3; 2.03: 1.21; 2.04: 1.3942; 2.05: 1.38; 2.06: 2.7384; 2.07: 1.33; 2.08: 1.62; 2.09: 1.53; 2.10: 1.5815; 2.11: 1.2; 2.12: 1.21
| Mean |  | 2.01: 1.25; 2.03: 1.2167; 2.04: 1.3751; 2.05: 1.325; 2.06: 2.7566; 2.07: 1.2917; 2.08: 1.535; 2.09: 1.5567; 2.10: 1.5953; 2.11: 1.2167; 2.12: 1.3
| Median |  | 2.01: 1.25; 2.03: 1.22; 2.04: 1.3778; 2.05: 1.315; 2.06: 2.7668; 2.07: 1.295; 2.08: 1.52; 2.09: 1.555; 2.10: 1.5813; 2.11: 1.2; 2.12: 1.305
| Std Dev. |  | 2.01: 0.054772; 2.03: 0.01633; 2.04: 0.017711; 2.05: 0.031464; 2.06: 0.023341; 2.07: 0.033116; 2.08: 0.086429; 2.09: 0.017512; 2.10: 0.036256; 2.11: 0.040825; 2.12: 0.055498
| Rel.Std.Dev. |  | 2.01: 4.382%; 2.03: 1.342%; 2.04: 1.288%; 2.05: 2.375%; 2.06: 0.8467%; 2.07: 2.564%; 2.08: 5.631%; 2.09: 1.125%; 2.10: 2.273%; 2.11: 3.355%; 2.12: 4.269%
| PDM3 |  | 2.01: -10.57%; 2.03: -12.95%; 2.04: -1.62%; 2.05: -5.202%; 2.06: 97.22%; 2.07: -7.586%; 2.08: 9.823%; 2.09: 11.37%; 2.10: 14.14%; 2.11: -12.95%; 2.12: -6.99%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.67; 2.04: 0; 2.05: 0.21; 2.06: 5.88; 2.07: 0.36; 2.08: 0.68; 2.09: 0.77; 2.10: 0.94; 2.11: NA; 2.12: 0.32
Results from laboratories 2.03 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 606b (Certified Value 2.33 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.42 | 2.2 | 2.31 | 2.27 | 2.3744 | 2.53 | 2.45 | 2.3278 | 2.38 | 2.4 | 2.34 | 2.25 | 2.23 | 2.22 | 2.14 | 2.4 | 2.2841 | 2.07 | 2.39 | 2.21 | 2.29 | 2.3928 | 2.245 | 2.3098 | 2.4265 | 2.29 | 1
| 1 | 2 | 2.43 | 2.19 | 2.39 | 2.24 | 2.4346 | 2.46 | 2.39 | 2.3109 | 2.46 | 2.45 | 2.36 | 2.22 | 2.29 | 2.16 | 2.2 | 2.33 | 2.2843 | 2.19 | 2.45 | 2.32 | 2.27 | 2.3738 | 2.189 | 2.3451 | 2.5017 | 2.31 | 0
| 1 | 3 | 2.46 | 2.23 | 2.33 | 2.23 | 2.3849 | 2.47 | 2.38 | 2.3373 | 2.44 | 2.38 | 2.38 | 2.24 | 2.27 | 2.2 | 2.14 | 2.33 | 2.3316 | 2.08 | 2.29 | 2.37 | 2.35 | 2.3616 | 2.217 | 2.2909 | 2.3878 | 2.31 | 16
| 1 | 4 | 2.44 | 2.26 | 2.32 | 2.31 | 2.377 | 2.44 | 2.33 | 2.3443 | 2.32 | 2.39 | 2.38 | 2.27 | 2.29 | 2.17 | 2.13 | 2.39 | 2.3113 | 2.07 | 2.43 | 2.31 | 2.34 | 2.4372 | 2.175 | 2.2749 | 2.2908 | 2.25 | 2.3276
| 1 | 5 | 2.45 | 2.2 | 2.31 | 2.28 | 2.371 | 2.41 | 2.36 | 2.3837 | 2.46 | 2.36 | 2.32 | 2.32 | 2.32 | 2.19 | 2.14 | 2.36 | 2.3585 | 2.09 | 2.46 | 2.32 | 2.29 | 2.402 | 2.189 | 2.3505 | 2.3536 | 2.33 | 29
| 1 | 6 | 2.47 | 2.37 | 2.35 | 2.21 | 2.383 | 2.49 | 2.33 | 2.3709 | 2.29 | 2.34 | 2.36 | 2.22 | 2.26 | 2.2 | 2.05 | 2.33 | 2.2945 | 2.16 | 2.42 | 2.34 | 2.31 | 2.4258 | 2.224 | 2.3816 | 2.4406 | 2.25
| Mean |  | 2.445 | 2.2417 | 2.335 | 2.2567 | 2.3875 | 2.4667 | 2.3733 | 2.3458 | 2.3917 | 2.3867 | 2.3567 | 2.2533 | 2.2767 | 2.19 | 2.1333 | 2.3567 | 2.3107 | 2.11 | 2.4067 | 2.3117 | 2.3083 | 2.3989 | 2.2065 | 2.3255 | 2.4002 | 2.29
| Median |  | 2.445 | 2.215 | 2.325 | 2.255 | 2.38 | 2.465 | 2.37 | 2.3408 | 2.41 | 2.385 | 2.36 | 2.245 | 2.28 | 2.195 | 2.14 | 2.345 | 2.3029 | 2.085 | 2.425 | 2.32 | 2.3 | 2.3974 | 2.203 | 2.3275 | 2.4071 | 2.3
| Std Dev. |  | 0.018708 | 0.067946 | 0.030822 | 0.036697 | 0.02368 | 0.041312 | 0.045019 | 0.02714 | 0.073869 | 0.037771 | 0.023381 | 0.037771 | 0.030768 | 0.021909 | 0.048028 | 0.032042 | 0.029628 | 0.051769 | 0.062183 | 0.054191 | 0.031252 | 0.029184 | 0.026471 | 0.040406 | 0.07337 | 0.033466
| Rel.Std.Dev. |  | 0.7652% | 3.031% | 1.32% | 1.626% | 0.9918% | 1.675% | 1.897% | 1.157% | 3.089% | 1.583% | 0.9921% | 1.676% | 1.351% | 1% | 2.251% | 1.36% | 1.282% | 2.453% | 2.584% | 2.344% | 1.354% | 1.217% | 1.2% | 1.738% | 3.057% | 1.461%
| PDM3 |  | 5.044% | -3.692% | 0.318% | -3.047% | 2.573% | 5.975% | 1.965% | 0.7828% | 2.753% | 2.538% | 1.249% | -3.191% | -2.188% | -5.912% | -8.346% | 1.249% | -0.7252% | -9.349% | 3.397% | -0.6844% | -0.8276% | 3.062% | -5.203% | -0.09171% | 3.117% | -1.615%
| Z-Score (Absolute) |  | 1.3 | 1.01 | 0.05 | 0.84 | 0.65 | 1.55 | 0.49 | 0.18 | 0.7 | 0.64 | 0.3 | 0.87 | 0.61 | 1.59 | 2.24 | 0.3 | 0.22 | 2.5 | 0.87 | 0.21 | 0.25 | 0.78 | 1.4 | 0.05 | 0.79 | 0.46
| Analytical results for Ga in OREAS 606b (Certified Value 23 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 23.96 | 21.88 | 21.72 | 23.1073 | 23 | 24.9 | 28.753 | 23.32 | 22.4 | 23.72 | 21.5 | 20.8 | 22 | 21.9 | 22 | 22 | 23.9 | 32 | 57.66 | 23.19 | 17 | 24.1932 | 20.297 | 26.16 |  |  | 1
| 1 | 2 | 24.59 | 21.72 | 22.22 | 23.4687 | 23.5 | 25.6 | 28.557 | 24.28 | 22.6 | 24.14 | 22.1 | 21.1 | 21.6 | 22.7 | 22.4 | 21.05 | 24.31 | 31.4 | 54.31 | 23.2757 | 17 | 24.3259 | 20.775 | 26.15 |  |  | 36
| 1 | 3 | 23.85 | 22.12 | 21.96 | 23.4297 | 23.1 | 26 | 28.887 | 24.3 | 23.1 | 23.85 | 22.4 | 21.4 | 22.3 | 22.1 | 21.8 | 21.76 | 24.5 | 29.2 | 56.85 | 23.8134 | 17 | 23.8656 | 20.321 | 26.73 |  |  | 16
| 1 | 4 | 23.72 | 22.12 | 23.21 | 23.0498 | 24.4 | 25.2 | 28.618 | 23.23 | 22.2 | 23.44 | 22.8 | 21.5 | 21.4 | 22.1 | 22.3 | 21.56 | 24.3 | 30.2 | 57.27 | 24.3769 | 17 | 23.9603 | 20.027 | 25.63 |  |  | 23.044
| 1 | 5 | 24.24 | 21.56 | 22.65 | 23.0952 | 23.7 | 26 | 28.633 | 24 | 22.7 | 23.63 | 23.6 | 21.7 | 22.4 | 22 | 22.3 | 21.94 | 24.65 | 32.5 | 55.67 | 24.5768 | 17 | 23.9416 | 19.284 | 27.53 |  |  | 30
| 1 | 6 | 22.76 | 22.52 | 21.69 | 23.1541 | 23.9 | 25.3 | 28.618 | 22.35 | 23 | 24.01 | 21.9 | 21.1 | 22.2 | 21.6 | 21.8 | 21.76 | 24.88 | 31.2 | 57.12 | 23.7725 | 18 | 24.1715 | 20.473 | 25.93
| Mean |  | 23.8533 | 21.9867 | 22.2417 | 23.2175 | 23.6 | 25.5 | 28.6777 | 23.58 | 22.6667 | 23.7983 | 22.3833 | 21.2667 | 21.9833 | 22.0667 | 22.1 | 21.6783 | 24.4233 | 31.0833 | 56.48 | 23.8342 | 17.1667 | 24.0763 | 20.1962 | 26.355
| Median |  | 23.905 | 22 | 22.09 | 23.1307 | 23.6 | 25.45 | 28.6255 | 23.66 | 22.65 | 23.785 | 22.25 | 21.25 | 22.1 | 22.05 | 22.15 | 21.76 | 24.405 | 31.3 | 56.985 | 23.793 | 17 | 24.0659 | 20.309 | 26.155
| Std Dev. |  | 0.61915 | 0.34192 | 0.59341 | 0.18296 | 0.52154 | 0.44721 | 0.12097 | 0.75997 | 0.34448 | 0.25577 | 0.7414 | 0.3266 | 0.40208 | 0.36148 | 0.26833 | 0.3447 | 0.33697 | 1.2073 | 1.2592 | 0.56163 | 0.40825 | 0.17947 | 0.50953 | 0.67935
| Rel.Std.Dev. |  | 2.596% | 1.555% | 2.668% | 0.788% | 2.21% | 1.754% | 0.4218% | 3.223% | 1.52% | 1.075% | 3.312% | 1.536% | 1.829% | 1.638% | 1.214% | 1.59% | 1.38% | 3.884% | 2.229% | 2.356% | 2.378% | 0.7454% | 2.523% | 2.578%
| PDM3 |  | 3.512% | -4.588% | -3.482% | 0.7527% | 2.413% | 10.66% | 24.45% | 2.326% | -1.638% | 3.273% | -2.867% | -7.713% | -4.603% | -4.241% | -4.097% | -5.926% | 5.986% | 34.89% | 145.1% | 3.429% | -25.5% | 4.48% | -12.36% | 14.37%
| Z-Score (Absolute) |  | 0.23 | 0.72 | 0.59 | 0.09 | 0.1 | 1.08 | 2.71 | 0.09 | 0.38 | 0.2 | 0.52 | 1.09 | 0.73 | 0.68 | 0.67 | 0.88 | 0.53 | 3.94 | 16.96 | 0.22 | 3.19 | 0.35 | 1.64 | 1.52
| Analytical results for Gd in OREAS 606b (Certified Value 5.6 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.09: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.05: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  | 2
3
| 1 | 1 | 2.01: 5.6; 2.03: 5.53; 2.05: 5.2626; 2.06: 6; 2.07: 8.0392; 2.08: 5.78; 2.09: 5.1; 2.10: 5.71; 2.11: 6.174; 2.12: 3; 2.13: 5.3 |  |  | 1
| 1 | 2 | 2.01: 5.6; 2.03: 5.49; 2.05: 5.5948; 2.06: 5.8; 2.07: 8.0216; 2.08: 5.71; 2.09: 5.7; 2.10: 5.66; 2.11: 6.2183; 2.12: 3; 2.13: 5.1 |  |  | 37
| 1 | 3 | 2.01: 5.5; 2.03: 5.45; 2.05: 5.3695; 2.06: 6; 2.07: 7.9957; 2.08: 5.68; 2.09: 4.7; 2.10: 5.74; 2.11: 6.1164; 2.12: 3; 2.13: 5.3 |  |  | 16
| 1 | 4 | 2.01: 5.5; 2.03: 5.39; 2.05: 5.2768; 2.06: 6.1; 2.07: 8.2252; 2.08: 5.69; 2.09: 4.5; 2.10: 5.74; 2.11: 6.3305; 2.12: 3; 2.13: 5.3 |  |  | 5.6028
| 1 | 5 | 2.01: 6; 2.03: 5.54; 2.05: 5.2552; 2.06: 5.7; 2.07: 8.0916; 2.08: 5.58; 2.09: 4.9; 2.10: 5.71; 2.11: 6.5448; 2.12: 3; 2.13: 5.2 |  |  | 31
| 1 | 6 | 2.01: 6; 2.03: 5.51; 2.05: 5.5268; 2.06: 6.1; 2.07: 8.0489; 2.08: 5.76; 2.09: 5.4; 2.10: 5.65; 2.11: 6.2627; 2.12: 3; 2.13: 4.9
| Mean |  | 2.01: 5.7; 2.03: 5.485; 2.05: 5.3809; 2.06: 5.95; 2.07: 8.0704; 2.08: 5.7; 2.09: 5.05; 2.10: 5.7017; 2.11: 6.2744; 2.12: 3; 2.13: 5.1833
| Median |  | 2.01: 5.6; 2.03: 5.5; 2.05: 5.3231; 2.06: 6; 2.07: 8.0441; 2.08: 5.7; 2.09: 5; 2.10: 5.71; 2.11: 6.2405; 2.12: 3; 2.13: 5.25
| Std Dev. |  | 2.01: 0.23664; 2.03: 0.05648; 2.05: 0.14684; 2.06: 0.16432; 2.07: 0.082231; 2.08: 0.070711; 2.09: 0.44609; 2.10: 0.038687; 2.11: 0.15139; 2.12: 0; 2.13: 0.16021
| Rel.Std.Dev. |  | 2.01: 4.152%; 2.03: 1.03%; 2.05: 2.729%; 2.06: 2.762%; 2.07: 1.019%; 2.08: 1.241%; 2.09: 8.834%; 2.10: 0.6785%; 2.11: 2.413%; 2.12: 0%; 2.13: 3.091%
| PDM3 |  | 2.01: 1.734%; 2.03: -2.103%; 2.05: -3.96%; 2.06: 6.196%; 2.07: 44.04%; 2.08: 1.734%; 2.09: -9.867%; 2.10: 1.764%; 2.11: 11.99%; 2.12: -46.46%; 2.13: -7.487%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.51; 2.05: 0.76; 2.06: 0.59; 2.07: 5.62; 2.08: 0; 2.09: 1.54; 2.10: 0; 2.11: 1.36; 2.12: NA; 2.13: 1.22
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 606b (Certified Value 0.19 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 0.27 | 0.08 | 1.3 | < 0.2 | < 0.1 | 0.34 | 0.3093 | 1.3 | 1.3 | 0.16 | 0.16 | 0.16 | 0.15 | 0.16 | 0.05 | < 0.2 | 0.34 | 0.11 | 0.29005 | 1.1 | [6 empty] | 1
| 1 | 2 | 0.3 | 0.13 | 1.3 | < 0.2 | < 0.1 | 0.34 | 0.2329 | 1.3 | 1.3 | 0.16 | 0.13 | 0.19 | 0.17 | 0.16 | 0.22 | < 0.2 | 0.32 | 0.13 | 0.2836 | 1.1 | [6 empty] | 38
| 1 | 3 | 0.29 | < 0.05 | 1.3 | < 0.2 | < 0.1 | 0.35 | 0.3161 | 1.3 | 1.2 | 0.18 | 0.15 | 0.21 | 0.18 | 0.16 | 0.14 | < 0.2 | 0.26 | 0.11 | 0.27326 | 0.9 | [6 empty] | 16
| 1 | 4 | 0.31 | 0.08 | 1.4 | < 0.2 | < 0.1 | 0.41 | 0.2668 | 1.3 | 1.3 | 0.16 | 0.09 | 0.18 | 0.18 | 0.13 | 0.14 | < 0.2 | 0.31 | 0.12 | 0.27376 | 0.9 | [6 empty] | 0.19297
| 1 | 5 | 0.28 | 0.16 | 1.2 | < 0.2 | < 0.1 | 0.37 | 0.2919 | 1.3 | 1.2 | 0.18 | 0.13 | 0.17 | 0.17 | 0.14 | 0.05 | < 0.2 | 0.31 | 0.11 | 0.29415 | 1.1 | [6 empty] | 32
| 1 | 6 | 0.28 | 0.07 | 1.4 | < 0.2 | < 0.1 | 0.35 | 0.2958 | 1.3 | 1.2 | 0.18 | 0.14 | 0.19 | 0.16 | 0.17 | 0.1 | < 0.2 | 0.36 | 0.11 | 0.27222 | 0.9
| Mean |  | 0.28833 | 0.104 | 1.3167 |  |  | 0.36 | 0.28547 | 1.3 | 1.25 | 0.17 | 0.13333 | 0.18333 | 0.16833 | 0.15333 | 0.11667 |  | 0.31667 | 0.115 | 0.28117 | 1
| Median |  | 0.285 | 0.08 | 1.3 |  |  | 0.35 | 0.29385 | 1.3 | 1.25 | 0.17 | 0.135 | 0.185 | 0.17 | 0.16 | 0.12 |  | 0.315 | 0.11 | 0.27868 | 1
| Std Dev. |  | 0.01472 | 0.039115 | 0.075277 |  |  | 0.026833 | 0.030875 | 0 | 0.054772 | 0.010954 | 0.024221 | 0.017512 | 0.01169 | 0.015055 | 0.064704 |  | 0.033862 | 0.0083666 | 0.0094959 | 0.10954
| Rel.Std.Dev. |  | 5.105% | 37.61% | 5.717% |  |  | 7.454% | 10.82% | 0% | 4.382% | 6.444% | 18.17% | 9.552% | 6.945% | 9.819% | 55.46% |  | 10.69% | 7.275% | 3.377% | 10.95%
| PDM3 |  | 49.42% | -46.11% | 582.3% |  |  | 86.56% | 47.93% | 573.7% | 547.8% | -11.9% | -30.9% | -4.994% | -12.77% | -20.54% | -39.54% |  | 64.1% | -40.41% | 45.71% | 418.2%
| Z-Score (Absolute) |  | 1.41 | 0.78 | NA | 0.67 | 1.23 | 2.21 | 1.38 | NA | NA | 0.1 | 0.3 | 0.25 | 0.08 | 0.08 | 0.49 | 0.67 | 1.73 | 0.51 | 1.33 | NA
Results from laboratories 2.05, 2.11, 2.12 and 2.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 606b (Certified Value 6.48 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 4.9 | 6.07 | 6.1 | 6.21 | 6.6819 | 7 | 6.73 | 6.54 | 6.2 | 6.3 | 6.3 | 6.2 | 6.6 | 6.4 | 6.6 | 5.78 | 7.36 | 6.5 | 6.9 | 9 | 6.6481 | 6.53 |  |  |  |  | 1
| 1 | 2 | 5.4 | 6.46 | 6.4 | 6.27 | 6.7211 | 6.7 | 6.9 | 6.7 | 6.2 | 6.4 | 6.7 | 6.4 | 6.4 | 6.7 | 6.7 | 5.83 | 7.72 | 6.84 | 6.8 | 9 | 6.4704 | 6.31 |  |  |  |  | 22
| 1 | 3 | 5.1 | 6.24 | 6.2 | 6.22 | 6.4878 | 6.7 | 6.78 | 6.51 | 6.1 | 6.27 | 6.6 | 6.3 | 6.5 | 6.6 | 6.3 | 5.95 | 7.31 | 6.6 | 6.5 | 9 | 6.5575 | 6.21 |  |  |  |  | 16
| 1 | 4 | 6 | 6.09 | 6.1 | 6.68 | 6.5936 | 7 | 6.73 | 6.66 | 6.2 | 6.36 | 6.6 | 6.3 | 6.1 | 6.4 | 6.8 | 5.93 | 7.31 | 6.9 | 6.9 | 9 | 6.3956 | 6.38 |  |  |  |  | 6.4822
| 1 | 5 | 5.4 | 6.74 | 6.4 | 6.33 | 6.6519 | 6.8 | 6.75 | 6.82 | 6 | 6.33 | 6.9 | 6.3 | 6.6 | 6.6 | 6.6 | 5.85 | 7.49 | 6.93 | 6.5 | 9 | 6.7253 | 6.57 |  |  |  |  | 33
| 1 | 6 | 5.3 | 6.94 | 6.3 | 6.85 | 6.3517 | 6.7 | 6.68 | 6.55 | 6 | 6.38 | 6.4 | 6.1 | 6.4 | 6.5 | 6.8 | 5.86 | 7.56 | 6.8 | 6.8 | 9 | 6.851 | 6.38
| Mean |  | 5.35 | 6.4233 | 6.25 | 6.4267 | 6.5813 | 6.8167 | 6.7617 | 6.63 | 6.1167 | 6.34 | 6.5833 | 6.2667 | 6.4333 | 6.5333 | 6.6333 | 5.8667 | 7.4583 | 6.7617 | 6.7333 | 9 | 6.608 | 6.3967
| Median |  | 5.35 | 6.35 | 6.25 | 6.3 | 6.6228 | 6.75 | 6.74 | 6.605 | 6.15 | 6.345 | 6.6 | 6.3 | 6.45 | 6.55 | 6.65 | 5.855 | 7.425 | 6.82 | 6.8 | 9 | 6.6028 | 6.38
| Std Dev. |  | 0.37283 | 0.35725 | 0.13784 | 0.2709 | 0.13874 | 0.1472 | 0.075211 | 0.119 | 0.098319 | 0.049396 | 0.2137 | 0.10328 | 0.18619 | 0.12111 | 0.18619 | 0.063456 | 0.16339 | 0.17302 | 0.18619 | 0 | 0.16792 | 0.13471
| Rel.Std.Dev. |  | 6.969% | 5.562% | 2.205% | 4.215% | 2.108% | 2.159% | 1.112% | 1.795% | 1.607% | 0.7791% | 3.246% | 1.648% | 2.894% | 1.854% | 2.807% | 1.082% | 2.191% | 2.559% | 2.765% | 0% | 2.541% | 2.106%
| PDM3 |  | -17.47% | -0.9088% | -3.583% | -0.8574% | 1.529% | 5.159% | 4.311% | 2.279% | -5.64% | -2.194% | 1.559% | -3.326% | -0.7546% | 0.7881% | 2.331% | -9.496% | 15.06% | 4.311% | 3.873% | 38.84% | 1.94% | -1.32%
| Z-Score (Absolute) |  | 4.13 | 0.38 | 0.99 | 0.37 | 0.17 | 0.99 | 0.8 | 0.34 | 1.45 | 0.67 | 0.17 | 0.93 | 0.35 | 0 | 0.35 | 2.33 | 3.23 | 0.8 | 0.7 | NA | 0.26 | 0.48
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 606b (Indicative Value 0.023 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab |  |  |  |  | 1
| No | No | 2.01: 2.06; 2.03: 2.18; 2.04: 2.22 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 2.01: < 1; 2.03: 0.02; 2.04: < 2 |  |  |  |  | 1
| 1 | 2 | 2.01: < 1; 2.03: 0.03; 2.04: < 2 |  |  |  |  | 3
| 1 | 3 | 2.01: < 1; 2.03: 0.02; 2.04: < 2 |  |  |  |  | 16
| 1 | 4 | 2.01: < 1; 2.03: 0.03; 2.04: < 2 |  |  |  |  | 0.023333
| 1 | 5 | 2.01: < 1; 2.03: 0.02; 2.04: < 2 |  |  |  |  | 9
| 1 | 6 | 2.01: < 1; 2.03: 0.02; 2.04: < 2
| Mean |  | 2.03: 0.023333
| Median |  | 2.03: 0.02
| Std Dev. |  | 2.03: 0.005164
| Rel.Std.Dev. |  | 2.03: 22.13%
| PDM3 |  | 2.03: 1.554e-13%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.67; 2.04: 0.71
| Analytical results for Ho in OREAS 606b (Certified Value 0.48 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 0.6; 2.03: 0.46; 2.04: 0.45223; 2.05: 0.6; 2.06: 0.5338; 2.07: 0.49; 2.08: 0.45; 2.10: 0.53; 2.11: 0.4218; 2.12: < 1; 2.13: 0.48 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.6; 2.03: 0.49; 2.04: 0.47189; 2.05: 0.5; 2.06: 0.5311; 2.07: 0.53; 2.08: 0.48; 2.10: 0.51; 2.11: 0.4207; 2.12: < 1; 2.13: 0.46 |  |  |  |  | 23
| 1 | 3 | 2.01: 0.6; 2.03: 0.48; 2.04: 0.46383; 2.05: 0.5; 2.06: 0.5467; 2.07: 0.49; 2.08: 0.42; 2.10: 0.53; 2.11: 0.4156; 2.12: < 1; 2.13: 0.45 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.6; 2.03: 0.51; 2.04: 0.46602; 2.05: 0.5; 2.06: 0.5551; 2.07: 0.5; 2.08: 0.41; 2.10: 0.52; 2.11: 0.4297; 2.12: < 1; 2.13: 0.45 |  |  |  |  | 0.48276
| 1 | 5 | 2.01: 0.6; 2.03: 0.47; 2.04: 0.45716; 2.05: 0.5; 2.06: 0.5578; 2.07: 0.51; 2.08: 0.43; 2.10: 0.52; 2.11: 0.4368; 2.12: < 1; 2.13: 0.44 |  |  |  |  | 34
| 1 | 6 | 2.01: 0.6; 2.03: 0.48; 2.04: 0.4429; 2.05: 0.5; 2.06: 0.5492; 2.07: 0.5; 2.08: 0.47; 2.10: 0.5; 2.11: 0.4269; 2.12: < 1; 2.13: 0.43
| Mean |  | 2.01: 0.6; 2.03: 0.48167; 2.04: 0.45901; 2.05: 0.51667; 2.06: 0.54562; 2.07: 0.50333; 2.08: 0.44333; 2.10: 0.51833; 2.11: 0.42525; 2.13: 0.45167
| Median |  | 2.01: 0.6; 2.03: 0.48; 2.04: 0.4605; 2.05: 0.5; 2.06: 0.54795; 2.07: 0.5; 2.08: 0.44; 2.10: 0.52; 2.11: 0.42435; 2.13: 0.45
| Std Dev. |  | 2.01: 0; 2.03: 0.017224; 2.04: 0.010464; 2.05: 0.040825; 2.06: 0.01098; 2.07: 0.015055; 2.08: 0.028048; 2.10: 0.01169; 2.11: 0.0075054; 2.13: 0.017224
| Rel.Std.Dev. |  | 2.01: 0%; 2.03: 3.576%; 2.04: 2.28%; 2.05: 7.902%; 2.06: 2.012%; 2.07: 2.991%; 2.08: 6.327%; 2.10: 2.255%; 2.11: 1.765%; 2.13: 3.813%
| PDM3 |  | 2.01: 24.28%; 2.03: -0.2272%; 2.04: -4.921%; 2.05: 7.023%; 2.06: 13.02%; 2.07: 4.261%; 2.08: -8.168%; 2.10: 7.368%; 2.11: -11.91%; 2.13: -6.441%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.21; 2.04: 0.72; 2.05: 0.59; 2.06: 1.25; 2.07: 0.28; 2.08: 1.08; 2.10: 0.63; 2.11: 1.49; 2.12: 0.21; 2.13: 0.89
Results from laboratory 2.03 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 606b (Certified Value 0.14 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 0.135 | 0.13 | 0.13 | 0.13 | 0.1 | 0.147 | 0.1 | 0.14 | 0.14 | 0.13 | 0.127 | 0.153 | 0.142 | 0.143 | 0.147 | 0.145 | 0.167 | 0.138 | 0.14047 | 0.164 | [6 empty] | 1
| 1 | 2 | 0.149 | 0.13 | 0.13 | 0.14 | 0.1 | 0.147 | 0.1 | 0.13 | 0.15 | 0.134 | 0.13 | 0.151 | 0.147 | 0.15 | 0.148 | 0.15 | 0.161 | 0.127 | 0.13523 | 0.174 | [6 empty] | 24
| 1 | 3 | 0.152 | 0.12 | 0.13 | 0.14 | 0.1 | 0.152 | 0.1 | 0.15 | 0.15 | 0.132 | 0.14 | 0.152 | 0.145 | 0.148 | 0.148 | 0.142 | 0.163 | 0.112 | 0.14648 | 0.137 | [6 empty] | 16
| 1 | 4 | 0.145 | 0.12 | 0.13 | 0.14 | 0.2 | 0.145 | 0.1 | 0.14 | 0.14 | 0.14 | 0.138 | 0.147 | 0.141 | 0.156 | 0.151 | 0.146 | 0.199 | 0.121 | 0.13158 | 0.142 | [6 empty] | 0.14224
| 1 | 5 | 0.143 | 0.14 | 0.13 | 0.15 | 0.1 | 0.15 | 0.1 | 0.14 | 0.14 | 0.139 | 0.136 | 0.156 | 0.149 | 0.157 | 0.152 | 0.139 | 0.152 | 0.118 | 0.13658 | 0.135 | [6 empty] | 35
| 1 | 6 | 0.159 | 0.13 | 0.13 | 0.14 | 0.1 | 0.143 | 0.1 | 0.13 | 0.15 | 0.134 | 0.134 | 0.146 | 0.135 | 0.146 | 0.146 | 0.152 | 0.165 | 0.135 | 0.14309 | 0.155
| Mean |  | 0.14717 | 0.12833 | 0.13 | 0.14 | 0.11667 | 0.14733 | 0.1 | 0.13833 | 0.145 | 0.13483 | 0.13417 | 0.15083 | 0.14317 | 0.15 | 0.14867 | 0.14567 | 0.16783 | 0.12517 | 0.13891 | 0.15117
| Median |  | 0.147 | 0.13 | 0.13 | 0.14 | 0.1 | 0.147 | 0.1 | 0.14 | 0.145 | 0.134 | 0.135 | 0.1515 | 0.1435 | 0.149 | 0.148 | 0.1455 | 0.164 | 0.124 | 0.13853 | 0.1485
| Std Dev. |  | 0.0082077 | 0.0075277 | 0 | 0.0063246 | 0.040825 | 0.003266 | 1.5202e-17 | 0.0075277 | 0.0054772 | 0.00392 | 0.004916 | 0.0037639 | 0.0049967 | 0.0055498 | 0.0023381 | 0.0048442 | 0.01613 | 0.010068 | 0.0054756 | 0.015791
| Rel.Std.Dev. |  | 5.577% | 5.866% | 0% | 4.518% | 34.99% | 2.217% | 1.52e-14% | 5.442% | 3.777% | 2.907% | 3.664% | 2.495% | 3.49% | 3.7% | 1.573% | 3.326% | 9.611% | 8.044% | 3.942% | 10.45%
| PDM3 |  | 3.463% | -9.778% | -8.606% | -1.576% | -17.98% | 3.58% | -29.7% | -2.747% | 1.94% | -5.208% | -5.677% | 6.041% | 0.6508% | 5.455% | 4.517% | 2.408% | 17.99% | -12% | -2.345% | 6.275%
| Z-Score (Absolute) |  | 0.36 | 1.82 | 1.63 | 0.47 | NA | 0.38 | NA | 0.66 | 0.11 | 1.07 | 1.15 | 0.78 | 0.11 | 0.68 | 0.53 | 0.18 | 2.75 | 2.19 | 0.6 | 0.82
Results from laboratories 2.07 and 2.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 606b (Certified Value 3.02 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.15 | 3.21 | 3.23 | 3.0455 | 3.0186 | 3.3 | 3.02 | 2.9262 | 3.19 | 2.97 | 3.0217 | 3.13 | 2.89 | 2.93 | 2.87 | 3.3 | 2.9699 | 3.07 | 2.66 | 2.8 | 2.77 | 2.6789 | 3.013 | 2.9741 | 2.8196 | 3.02 | 1
| 1 | 2 | 3.13 | 3 | 3.32 | 3.0072 | 3.0771 | 3.16 | 3.06 | 2.9957 | 3.31 | 2.95 | 3.0231 | 3.13 | 3 | 2.86 | 2.91 | 3.17 | 2.8982 | 3.14 | 2.46 | 2.84 | 2.74 | 2.673 | 2.939 | 3.0133 | 3.3377 | 3.01 | 0
| 1 | 3 | 3.16 | 3.01 | 3.24 | 3.0055 | 3.0651 | 3.27 | 3.04 | 2.9303 | 3.26 | 2.99 | 3.056 | 3.14 | 2.91 | 2.9 | 2.84 | 3.14 | 2.9855 | 3.07 | 2.58 | 2.86 | 2.79 | 2.6198 | 2.906 | 2.9707 | 3.2314 | 3.01 | 16
| 1 | 4 | 3.15 | 3.16 | 3.25 | 3.1109 | 3.0017 | 3.23 | 3.02 | 2.9039 | 3.1 | 2.96 | 3.0139 | 3.15 | 2.98 | 2.87 | 2.83 | 3.32 | 2.9478 | 3.05 | 2.55 | 2.86 | 2.8 | 2.7418 | 2.922 | 2.9321 | 3.4442 | 3.02 | 3.0169
| 1 | 5 | 3.12 | 3.11 | 3.22 | 3.0613 | 3.0339 | 3.16 | 3.03 | 2.9056 | 3.3 | 2.91 | 2.9918 | 3.19 | 3 | 2.92 | 2.83 | 3.26 | 2.9723 | 3.07 | 2.43 | 2.84 | 2.75 | 2.6912 | 2.947 | 3.0206 | 2.968 | 3.08 | 36
| 1 | 6 | 3.14 | 3.19 | 3.27 | 2.9497 | 3.0586 | 3.14 | 3.01 | 2.9628 | 3.07 | 2.92 | 2.9993 | 3.05 | 2.98 | 2.9 | 2.79 | 3.22 | 2.9321 | 3.07 | 2.62 | 2.86 | 2.82 | 2.6976 | 2.997 | 3.0557 | 2.9435 | 2.95
| Mean |  | 3.1417 | 3.1133 | 3.255 | 3.03 | 3.0425 | 3.21 | 3.03 | 2.9374 | 3.205 | 2.95 | 3.0176 | 3.1317 | 2.96 | 2.8967 | 2.845 | 3.235 | 2.951 | 3.0783 | 2.55 | 2.8433 | 2.7783 | 2.6837 | 2.954 | 2.9944 | 3.1241 | 3.015
| Median |  | 3.145 | 3.135 | 3.245 | 3.0263 | 3.0462 | 3.195 | 3.025 | 2.9283 | 3.225 | 2.955 | 3.0178 | 3.135 | 2.98 | 2.9 | 2.835 | 3.24 | 2.9588 | 3.07 | 2.565 | 2.85 | 2.78 | 2.685 | 2.943 | 2.9937 | 3.0997 | 3.015
| Std Dev. |  | 0.01472 | 0.09048 | 0.036194 | 0.055377 | 0.029249 | 0.066332 | 0.017889 | 0.03564 | 0.10252 | 0.030332 | 0.02252 | 0.045789 | 0.047749 | 0.027325 | 0.040866 | 0.071484 | 0.032114 | 0.031252 | 0.089889 | 0.023381 | 0.030605 | 0.039613 | 0.042266 | 0.043945 | 0.24872 | 0.041352
| Rel.Std.Dev. |  | 0.4685% | 2.906% | 1.112% | 1.828% | 0.9613% | 2.066% | 0.5904% | 1.213% | 3.199% | 1.028% | 0.7463% | 1.462% | 1.613% | 0.9433% | 1.436% | 2.21% | 1.088% | 1.015% | 3.525% | 0.8223% | 1.102% | 1.476% | 1.431% | 1.468% | 7.961% | 1.372%
| PDM3 |  | 4.135% | 3.196% | 7.891% | 0.4341% | 0.8473% | 6.4% | 0.4335% | -2.635% | 6.234% | -2.218% | 0.02361% | 3.803% | -1.887% | -3.986% | -5.699% | 7.229% | -2.186% | 2.036% | -15.48% | -5.754% | -7.908% | -11.04% | -2.086% | -0.7455% | 3.551% | -0.06367%
| Z-Score (Absolute) |  | 0.83 | 0.64 | 1.57 | 0.09 | 0.17 | 1.28 | 0.09 | 0.52 | 1.24 | 0.44 | 0.01 | 0.76 | 0.37 | 0.79 | 1.13 | 1.44 | 0.43 | 0.41 | 3.07 | 1.14 | 1.57 | 2.19 | 0.41 | 0.14 | 0.71 | 0.01
| Analytical results for La in OREAS 606b (Certified Value 45 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 45 | 41.7 | 46.3 | 45.46 | 40.6773 | 48.3 | 50.7 | 50.677 | 46.73 | 43.91 | 42.3 | 40.3 | 43.2 | 43.1 | 43.7 | 45.1 | 31.02 | 50 | 34.75 | 45 | 40.0805 | 46 | 36.618 | 42.04 |  |  | 1
| 1 | 2 | 45 | 41.5 | 46.6 | 45.04 | 41.4559 | 46.9 | 52.2 | 50.218 | 47.9 | 44.79 | 45.6 | 41.4 | 42 | 45.1 | 45.4 | 43.8 | 34.06 | 49 | 38.59 | 45 | 39.6462 | 45 | 36.954 | 38.7 |  |  | 25
| 1 | 3 | 46 | 42.7 | 47.4 | 44.72 | 41.4946 | 47.3 | 51.4 | 50.077 | 47.27 | 42.57 | 45.3 | 39.9 | 42.9 | 44.4 | 45 | 45.4 | 27.83 | 49 | 36.85 | 46 | 41.4616 | 46 | 36.965 | 42.65 |  |  | 16
| 1 | 4 | 47 | 41.1 | 46 | 44.88 | 39.8059 | 48.1 | 49.9 | 50.675 | 45.4 | 44.44 | 44.7 | 42.3 | 40.6 | 44.9 | 46.5 | 45.2 | 24.91 | 48 | 36.74 | 46 | 42.3311 | 44 | 35.751 | 41.12 |  |  | 44.9941
| 1 | 5 | 46 | 47.5 | 46.3 | 45.36 | 40.6379 | 47.7 | 50.1 | 50.443 | 47.07 | 43.69 | 44.8 | 43.1 | 43.4 | 43.7 | 44.8 | 45.7 | 28.71 | 50 | 35.09 | 45 | 42.8194 | 45 | 36.797 | 42.57 |  |  | 37
| 1 | 6 | 45 | 45.1 | 46.2 | 44.93 | 41.0918 | 48.1 | 49.2 | 50.175 | 44.47 | 44.69 | 43.6 | 42 | 42.2 | 31.1 | 43.1 | 44.8 | 32 | 48 | 38.28 | 46 | 41.078 | 46 | 37.605 | 39.17
| Mean |  | 45.6667 | 43.2667 | 46.4667 | 45.065 | 40.8606 | 47.7333 | 50.5833 | 50.3775 | 46.4733 | 44.015 | 44.3833 | 41.5 | 42.3833 | 42.05 | 44.75 | 45 | 29.755 | 49 | 36.7167 | 45.5 | 41.2361 | 45.3333 | 36.7817 | 41.0417
| Median |  | 45.5 | 42.2 | 46.3 | 44.985 | 40.8846 | 47.9 | 50.4 | 50.3305 | 46.9 | 44.175 | 44.75 | 41.7 | 42.55 | 44.05 | 44.9 | 45.15 | 29.865 | 49 | 36.795 | 45.5 | 41.2698 | 45.5 | 36.8755 | 41.58
| Std Dev. |  | 0.8165 | 2.5248 | 0.49666 | 0.28815 | 0.63333 | 0.54283 | 1.087 | 0.26048 | 1.2851 | 0.82964 | 1.2287 | 1.2215 | 1.0323 | 5.4162 | 1.2145 | 0.6603 | 3.2694 | 0.89443 | 1.5802 | 0.54772 | 1.2365 | 0.8165 | 0.60506 | 1.727
| Rel.Std.Dev. |  | 1.788% | 5.835% | 1.069% | 0.6394% | 1.55% | 1.137% | 2.149% | 0.5171% | 2.765% | 1.885% | 2.768% | 2.943% | 2.436% | 12.88% | 2.714% | 1.467% | 10.99% | 1.825% | 4.304% | 1.204% | 2.999% | 1.801% | 1.645% | 4.208%
| PDM3 |  | 1.495% | -3.839% | 3.273% | 0.1576% | -9.187% | 6.088% | 12.42% | 11.96% | 3.288% | -2.176% | -1.357% | -7.766% | -5.802% | -6.543% | -0.5425% | 0.01314% | -33.87% | 8.903% | -18.4% | 1.124% | -8.352% | 0.754% | -18.25% | -8.784%
| Z-Score (Absolute) |  | 0.32 | 0.37 | 0.55 | 0.14 | 1.06 | 0.91 | 1.73 | 1.67 | 0.55 | 0.16 | 0.05 | 0.88 | 0.63 | 0.72 | 0.05 | 0.12 | 4.25 | 1.27 | 2.25 | 0.27 | 0.96 | 0.22 | 2.23 | 1.01
| Analytical results for Li in OREAS 606b (Certified Value 32.1 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 35.5 | 29.4 | 31.2 | 30.8 | 31.7637 | 35 | 29.1 | 40.4714 | 31.7 | 31.9 | 33 | 32.5 | 31 | 31.5 | 31.3 | 34.3 | 33 | 30.2 | 31.6 | 30.67 | 30 | 47.9831 | 32 | 34.0558 | 44.316 | 33.5 | 1
| 1 | 2 | 35.3 | 30.3 | 33 | 31.3 | 32.9481 | 35 | 30 | 40.3762 | 30 | 32.7 | 33 | 33 | 31.2 | 30.8 | 31.6 | 34.4 | 32 | 30.8 | 33.2 | 31.56 | 31 | 47.5808 | 31 | 32.8747 | 46.565 | 33.9 | 45
| 1 | 3 | 34.9 | 30.8 | 33 | 30.5 | 33.1168 | 34 | 30.9 | 40.4776 | 31.5 | 32.6 | 33 | 32.8 | 33.4 | 31.4 | 31 | 33.4 | 33 | 30.5 | 30.9 | 31.42 | 30 | 46.794 | 32 | 33.6004 | 46.267 | 32 | 16
| 1 | 4 | 36.1 | 29.3 | 32.5 | 30.8 | 32.1597 | 34 | 29.4 | 40.3475 | 31.4 | 31.8 | 33 | 32.9 | 30.6 | 30.9 | 30.9 | 34.9 | 33 | 31.2 | 33 | 30.91 | 31 | 48.6344 | 31 | 33.0267 | 45.966 | 32.5 | 32.1337
| 1 | 5 | 34.6 | 31.8 | 32.2 | 30.8 | 32.411 | 34 | 30.4 | 40.5281 | 30.9 | 32.7 | 33 | 33.4 | 32.1 | 31.6 | 31 | 34.3 | 33 | 31 | 33.2 | 31.65 | 30 | 48.6542 | 31 | 34.1644 | 44.619 | 33.9 | 38
| 1 | 6 | 33.2 | 33.9 | 32.2 | 30.6 | 32.0324 | 35 | 31 | 40.5861 | 30.3 | 32.4 | 33 | 32.1 | 32.1 | 31.3 | 30.7 | 33.7 | 33 | 30.4 | 32.8 | 31.31 | 31 | 49.0835 | 32 | 32.3795 | 45.789 | 33
| Mean |  | 34.9333 | 30.9167 | 32.35 | 30.8 | 32.4053 | 34.5 | 30.1333 | 40.4645 | 30.9667 | 32.35 | 33 | 32.7833 | 31.7333 | 31.25 | 31.0833 | 34.1667 | 32.8333 | 30.6833 | 32.45 | 31.2533 | 30.5 | 48.1217 | 31.5 | 33.3502 | 45.587 | 33.1333
| Median |  | 35.1 | 30.55 | 32.35 | 30.8 | 32.2854 | 34.5 | 30.2 | 40.4745 | 31.15 | 32.5 | 33 | 32.85 | 31.65 | 31.35 | 31 | 34.3 | 33 | 30.65 | 32.9 | 31.365 | 30.5 | 48.3088 | 31.5 | 33.3135 | 45.8775 | 33.25
| Std Dev. |  | 0.99331 | 1.7314 | 0.66858 | 0.27568 | 0.5314 | 0.54772 | 0.77889 | 0.090062 | 0.69186 | 0.40373 | 0 | 0.4446 | 1.0152 | 0.32711 | 0.31885 | 0.53541 | 0.40825 | 0.38166 | 0.96695 | 0.38485 | 0.54772 | 0.84215 | 0.54772 | 0.7066 | 0.91168 | 0.77632
| Rel.Std.Dev. |  | 2.843% | 5.6% | 2.067% | 0.8951% | 1.64% | 1.588% | 2.585% | 0.2226% | 2.234% | 1.248% | 0% | 1.356% | 3.199% | 1.047% | 1.026% | 1.567% | 1.243% | 1.244% | 2.98% | 1.231% | 1.796% | 1.75% | 1.739% | 2.119% | 2% | 2.343%
| PDM3 |  | 8.712% | -3.787% | 0.6731% | -4.151% | 0.8451% | 7.364% | -6.225% | 25.93% | -3.632% | 0.6731% | 2.696% | 2.022% | -1.246% | -2.75% | -3.269% | 6.327% | 2.177% | -4.514% | 0.9843% | -2.74% | -5.084% | 49.75% | -1.972% | 3.786% | 41.87% | 3.111%
| Z-Score (Absolute) |  | 1.42 | 0.81 | 0.02 | 0.88 | 0.02 | 1.18 | 1.25 | 4.5 | 0.79 | 0.02 | 0.35 | 0.23 | 0.36 | 0.63 | 0.72 | 1 | 0.25 | 0.94 | 0.04 | 0.63 | 1.05 | 8.77 | 0.49 | 0.54 | 7.36 | 0.42
| Analytical results for Lu in OREAS 606b (Certified Value 0.074 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.12; 2.07: 2.19; 2.08: 2.22; 2.09: 2.23; 2.10: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS | 3
3
| 1 | 1 | 2.01: < 0.1; 2.03: 0.06; 2.04: 0.088985; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0897; 2.10: 0.08 | 1
| 1 | 2 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.092411; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.07; 2.09: 0.1303; 2.10: 0.06 | 26
| 1 | 3 | 2.01: < 0.1; 2.03: 0.06; 2.04: 0.089841; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.08; 2.08: 0.07; 2.09: 0.0908; 2.10: 0.05 | 16
| 1 | 4 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.087852; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.08; 2.08: 0.08; 2.09: 0.0882; 2.10: 0.04 | 0.074113
| 1 | 5 | 2.01: < 0.1; 2.03: 0.06; 2.04: 0.090966; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.08; 2.08: 0.07; 2.09: 0.0935; 2.10: 0.04 | 39
| 1 | 6 | 2.01: < 0.1; 2.03: 0.06; 2.04: 0.087785; 2.05: < 0.1; 2.06: 0.07; 2.07: 0.09; 2.08: 0.07; 2.09: 0.0919; 2.10: 0.06
| Mean |  | 2.03: 0.063333; 2.04: 0.08964; 2.06: 0.061667; 2.07: 0.085; 2.08: 0.073333; 2.09: 0.0974; 2.10: 0.055
| Median |  | 2.03: 0.06; 2.04: 0.089413; 2.06: 0.06; 2.07: 0.085; 2.08: 0.07; 2.09: 0.09135; 2.10: 0.055
| Std Dev. |  | 2.03: 0.005164; 2.04: 0.001819; 2.06: 0.0040825; 2.07: 0.0054772; 2.08: 0.005164; 2.09: 0.016219; 2.10: 0.015166
| Rel.Std.Dev. |  | 2.03: 8.154%; 2.04: 2.029%; 2.06: 6.62%; 2.07: 6.444%; 2.08: 7.042%; 2.09: 16.65%; 2.10: 27.57%
| PDM3 |  | 2.03: -14.55%; 2.04: 20.95%; 2.06: -16.79%; 2.07: 14.69%; 2.08: -1.052%; 2.09: 31.42%; 2.10: -25.79%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0; 2.04: 1.33; 2.05: 0.67; 2.06: 0.08; 2.07: 1.1; 2.08: 0.51; 2.09: 1.72; 2.10: 0.42
| Analytical results for Mg in OREAS 606b (Certified Value 0.127 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.13 | 0.13 | 0.13 | 0.1274 | 0.12862 | 0.14 | 0.13 | 0.13681 | 0.13 | 0.132 | 0.1275 | 0.12 | 0.11 | 0.12 | 0.11 | 0.09 | 0.1295 | 0.122 | 0.13 | 0.13 | 0.12 | 0.1187 | 0.133 | 0.12805 | 0.10298 | 0.115 | 1
| 1 | 2 | 0.13 | 0.12 | 0.14 | 0.1259 | 0.13245 | 0.13 | 0.14 | 0.13849 | 0.13 | 0.134 | 0.128 | 0.12 | 0.12 | 0.12 | 0.12 | 0.09 | 0.1265 | 0.127 | 0.135 | 0.13 | 0.12 | 0.1193 | 0.127 | 0.12983 | 0.12357 | 0.115 | 0
| 1 | 3 | 0.13 | 0.13 | 0.13 | 0.1254 | 0.12852 | 0.13 | 0.13 | 0.1361 | 0.13 | 0.135 | 0.1298 | 0.12 | 0.12 | 0.12 | 0.11 | 0.09 | 0.1288 | 0.123 | 0.13 | 0.13 | 0.12 | 0.1181 | 0.127 | 0.127 | 0.10969 | 0.117 | 16
| 1 | 4 | 0.13 | 0.12 | 0.13 | 0.1307 | 0.12846 | 0.13 | 0.13 | 0.13773 | 0.12 | 0.136 | 0.1312 | 0.12 | 0.12 | 0.12 | 0.11 | 0.09 | 0.1287 | 0.119 | 0.13 | 0.13 | 0.12 | 0.1218 | 0.127 | 0.12776 | 0.119 | 0.117 | 0.12667
| 1 | 5 | 0.13 | 0.12 | 0.14 | 0.1277 | 0.12825 | 0.13 | 0.13 | 0.13267 | 0.14 | 0.134 | 0.1268 | 0.13 | 0.12 | 0.12 | 0.11 | 0.09 | 0.1315 | 0.122 | 0.132 | 0.13 | 0.12 | 0.1206 | 0.127 | 0.13238 | 0.11878 | 0.122 | 40
| 1 | 6 | 0.13 | 0.13 | 0.14 | 0.1227 | 0.12883 | 0.13 | 0.13 | 0.13735 | 0.12 | 0.132 | 0.1271 | 0.12 | 0.12 | 0.12 | 0.1 | 0.09 | 0.1274 | 0.124 | 0.13 | 0.13 | 0.12 | 0.1215 | 0.127 | 0.1321 | 0.11885 | 0.113
| Mean |  | 0.13 | 0.125 | 0.135 | 0.12663 | 0.12919 | 0.13167 | 0.13167 | 0.13653 | 0.12833 | 0.13383 | 0.1284 | 0.12167 | 0.11833 | 0.12 | 0.11 | 0.09 | 0.12873 | 0.12283 | 0.13117 | 0.13 | 0.12 | 0.12 | 0.128 | 0.12952 | 0.11548 | 0.1165
| Median |  | 0.13 | 0.125 | 0.135 | 0.12665 | 0.12857 | 0.13 | 0.13 | 0.13708 | 0.13 | 0.134 | 0.12775 | 0.12 | 0.12 | 0.12 | 0.11 | 0.09 | 0.12875 | 0.1225 | 0.13 | 0.13 | 0.12 | 0.11995 | 0.127 | 0.12894 | 0.11881 | 0.116
| Std Dev. |  | 0 | 0.0054772 | 0.0054772 | 0.0026756 | 0.0016093 | 0.0040825 | 0.0040825 | 0.002054 | 0.0075277 | 0.0016021 | 0.0017332 | 0.0040825 | 0.0040825 | 0 | 0.0063246 | 1.5202e-17 | 0.0017328 | 0.0026394 | 0.0020412 | 0 | 0 | 0.0015258 | 0.0024495 | 0.0023034 | 0.0076131 | 0.0030822
| Rel.Std.Dev. |  | 0% | 4.382% | 4.057% | 2.113% | 1.246% | 3.101% | 3.101% | 1.504% | 5.866% | 1.197% | 1.35% | 3.355% | 3.45% | 0% | 5.75% | 1.689e-14% | 1.346% | 2.149% | 1.556% | 0% | 0% | 1.271% | 1.914% | 1.778% | 6.593% | 2.646%
| PDM3 |  | 2.632% | -1.315% | 6.579% | -0.02585% | 1.993% | 3.948% | 3.948% | 7.786% | 1.316% | 5.658% | 1.369% | -3.947% | -6.579% | -5.263% | -13.16% | -28.95% | 1.632% | -3.026% | 3.553% | 2.632% | -5.263% | -5.263% | 1.053% | 2.254% | -8.831% | -8.026%
| Z-Score (Absolute) |  | 0.31 | 0.46 | 1.07 | 0.21 | 0.18 | 0.56 | 0.56 | 1.3 | 0.05 | 0.89 | 0.06 | 0.97 | 1.47 | 1.22 | 2.75 | 5.8 | 0.11 | 0.79 | 0.48 | 0.31 | 1.22 | 1.22 | 0.05 | 0.23 | 1.91 | 1.75
| Analytical results for Mn in OREAS 606b (Certified Value 0.028 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0275 | 0.0281 | 0.028 | 0.0279 | 0.028812 | 0.0315 | 0.0289 | 0.02766 | 0.0294 | 0.0266 | 0.0293 | 0.028 | 0.0268 | 0.0269 | 0.0265 | 0.0293 | 0.0272 | 0.0222 | 0.0361 | 0.02523 | 0.0274 | < 0.1 | 0.031 | 0.028868 | 0.026352 | 0.0235 | 1
| 1 | 2 | 0.0279 | 0.0279 | 0.029 | 0.0275 | 0.029519 | 0.0304 | 0.0285 | 0.027785 | 0.0308 | 0.0269 | 0.0295 | 0.0279 | 0.0274 | 0.0266 | 0.0275 | 0.0286 | 0.0265 | 0.0233 | 0.0357 | 0.02604 | 0.0272 | < 0.1 | 0.023 | 0.028137 | 0.026823 | 0.0229 | 27
| 1 | 3 | 0.0272 | 0.028 | 0.0282 | 0.0273 | 0.029262 | 0.0306 | 0.0281 | 0.027887 | 0.0304 | 0.0268 | 0.0299 | 0.0283 | 0.0271 | 0.027 | 0.0267 | 0.0286 | 0.0273 | 0.022 | 0.0434 | 0.02554 | 0.0275 | < 0.1 | 0.023 | 0.028172 | 0.026509 | 0.0233 | 16
| 1 | 4 | 0.0277 | 0.028 | 0.0282 | 0.0282 | 0.028898 | 0.0306 | 0.0279 | 0.027198 | 0.0292 | 0.0266 | 0.0297 | 0.0286 | 0.0272 | 0.0269 | 0.0265 | 0.0294 | 0.0274 | 0.0206 | 0.0345 | 0.02567 | 0.0278 | < 0.1 | 0.023 | 0.027862 | 0.026354 | 0.0231 | 0.028045
| 1 | 5 | 0.0273 | 0.0281 | 0.0278 | 0.028 | 0.02887 | 0.0301 | 0.0282 | 0.027522 | 0.0302 | 0.0263 | 0.029 | 0.0291 | 0.0277 | 0.0267 | 0.0264 | 0.0294 | 0.0278 | 0.0215 | 0.035 | 0.02597 | 0.0273 | < 0.1 | 0.023 | 0.028895 | 0.026508 | 0.0246 | 41
| 1 | 6 | 0.0274 | 0.0297 | 0.0282 | 0.0273 | 0.028871 | 0.0305 | 0.0277 | 0.027436 | 0.0288 | 0.0264 | 0.0293 | 0.028 | 0.0272 | 0.0268 | 0.0258 | 0.0286 | 0.0272 | 0.0231 | 0.0347 | 0.02575 | 0.028 | < 0.1 | 0.031 | 0.028829 | 0.026933 | 0.0227
| Mean |  | 0.0275 | 0.0283 | 0.028233 | 0.0277 | 0.029039 | 0.030617 | 0.028217 | 0.027581 | 0.0298 | 0.0266 | 0.02945 | 0.028317 | 0.027233 | 0.026817 | 0.026567 | 0.028983 | 0.027233 | 0.022117 | 0.036567 | 0.0257 | 0.027533 |  | 0.025667 | 0.028461 | 0.02658 | 0.02335
| Median |  | 0.02745 | 0.02805 | 0.0282 | 0.0277 | 0.028885 | 0.03055 | 0.02815 | 0.027591 | 0.0298 | 0.0266 | 0.0294 | 0.02815 | 0.0272 | 0.02685 | 0.0265 | 0.02895 | 0.02725 | 0.0221 | 0.03535 | 0.02571 | 0.02745 |  | 0.023 | 0.028501 | 0.026508 | 0.0232
| Std Dev. |  | 0.00026077 | 0.00068993 | 0.00040825 | 0.00038471 | 0.00028562 | 0.00047081 | 0.00043089 | 0.00025007 | 0.00077974 | 0.00022804 | 0.00032094 | 0.00046224 | 0.00030111 | 0.0001472 | 0.00055015 | 0.00042151 | 0.00042269 | 0.0010068 | 0.0034022 | 0.00029611 | 0.00030768 |  | 0.0041312 | 0.00045524 | 0.00024363 | 0.00067454
| Rel.Std.Dev. |  | 0.9482% | 2.438% | 1.446% | 1.389% | 0.9836% | 1.538% | 1.527% | 0.9067% | 2.617% | 0.8573% | 1.09% | 1.632% | 1.106% | 0.5489% | 2.071% | 1.454% | 1.552% | 4.552% | 9.304% | 1.152% | 1.117% |  | 16.1% | 1.6% | 0.9166% | 2.889%
| PDM3 |  | -1.942% | 0.9105% | 0.6728% | -1.229% | 3.545% | 9.171% | 0.6134% | -1.652% | 6.259% | -5.151% | 5.011% | 0.97% | -2.893% | -4.379% | -5.27% | 3.347% | -2.893% | -21.14% | 30.39% | -8.36% | -1.823% |  | -8.479% | 1.483% | -5.223% | -16.74%
| Z-Score (Absolute) |  | 0.09 | 0.42 | 0.38 | 0.04 | 0.9 | 1.91 | 0.37 | 0.04 | 1.39 | 0.67 | 1.16 | 0.43 | 0.26 | 0.53 | 0.69 | 0.86 | 0.26 | 3.54 | 5.73 | 1.25 | 0.07 | 14.35 | 1.17 | 0.53 | 0.68 | 2.75
| Analytical results for Mo in OREAS 606b (Certified Value 4.38 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.45 | 3.95 | 4.2 | 4.4 | 4.6121 | 5 | 4.15 | 4.48 | 4.5 | 4.2 | 4.06 | 4.3 | 4.48 | 4.34 | 4.41 | 4.3 | 4.5 | 4.6 | 3.93 | 6 | 5.425 | 7 | 4.4603 | < 5 | 4.1 |  | 1
| 1 | 2 | 4.6 | 4.14 | 4.6 | 4.3 | 4.7731 | 5 | 4.39 | 4.37 | 4.3 | 4.3 | 4.21 | 4.86 | 4.36 | 4.59 | 4.35 | 4.4 | 4.8 | 4.68 | 3.73 | 6 | 4.4733 | 6 | 4.3699 | < 5 | 4.4 |  | 28
| 1 | 3 | 4.69 | 4.28 | 4.3 | 4.4 | 4.7308 | 4 | 4.4 | 4.39 | 4.4 | 4.5 | 4.19 | 4.41 | 4.51 | 4.42 | 4.63 | 4.4 | 4.4 | 4.71 | 3.71 | 6 | 5.5232 | 7 | 4.2839 | < 5 | 4 |  | 16
| 1 | 4 | 4.54 | 4 | 4.4 | 4.5 | 4.5274 | 4 | 4.28 | 4.44 | 4.5 | 4.6 | 4.26 | 4.6 | 4.38 | 4.2 | 4.59 | 4.3 | 4.4 | 4.84 | 3.76 | 6 | 5.1127 | 5 | 4.2948 | < 5 | 3.9 |  | 4.3841
| 1 | 5 | 4.52 | 4.16 | 4.3 | 4.2 | 4.7479 | 4 | 4.43 | 4.59 | 4.4 | 4.3 | 4.53 | 4.49 | 4.46 | 4.37 | 4.56 | 4.5 | 4.4 | 5.11 | 3.69 | 6 | 4.5416 | 7 | 4.3396 | < 5 | 4.4 |  | 42
| 1 | 6 | 4.92 | 4.41 | 4.3 | 4.6 | 4.5111 | 5 | 4.26 | 4.26 | 4.4 | 4.3 | 4.15 | 4.43 | 4.44 | 4.18 | 4.21 | 4.4 | 4.5 | 4.93 | 3.7 | 6 | 5.603 | 7 | 4.48 | < 5 | 4
| Mean |  | 4.62 | 4.1567 | 4.35 | 4.4 | 4.6504 | 4.5 | 4.3183 | 4.4217 | 4.4167 | 4.3667 | 4.2333 | 4.515 | 4.4383 | 4.35 | 4.4583 | 4.3833 | 4.5 | 4.8117 | 3.7533 | 6 | 5.1131 | 6.5 | 4.3714 |  | 4.1333
| Median |  | 4.57 | 4.15 | 4.3 | 4.4 | 4.6714 | 4.5 | 4.335 | 4.415 | 4.4 | 4.3 | 4.2 | 4.46 | 4.45 | 4.355 | 4.485 | 4.4 | 4.45 | 4.775 | 3.72 | 6 | 5.2688 | 7 | 4.3547 |  | 4.05
| Std Dev. |  | 0.16769 | 0.17143 | 0.13784 | 0.14142 | 0.11575 | 0.54772 | 0.10722 | 0.11125 | 0.075277 | 0.15055 | 0.16008 | 0.19563 | 0.058109 | 0.15126 | 0.16302 | 0.075277 | 0.15492 | 0.18777 | 0.090037 | 0 | 0.49828 | 0.83666 | 0.082734 |  | 0.21602
| Rel.Std.Dev. |  | 3.63% | 4.124% | 3.169% | 3.214% | 2.489% | 12.17% | 2.483% | 2.516% | 1.704% | 3.448% | 3.781% | 4.333% | 1.309% | 3.477% | 3.657% | 1.717% | 3.443% | 3.902% | 2.399% | 0% | 9.745% | 12.87% | 1.893% |  | 5.226%
| PDM3 |  | 5.38% | -5.188% | -0.7782% | 0.3622% | 6.074% | 2.643% | -1.501% | 0.8564% | 0.7424% | -0.3981% | -3.439% | 2.985% | 1.237% | -0.7782% | 1.693% | -0.01792% | 2.643% | 9.752% | -14.39% | 36.86% | 16.63% | 48.26% | -0.2897% |  | -5.72%
| Z-Score (Absolute) |  | 1.48 | 1.64 | 0.34 | 0 | 1.69 | NA | 0.55 | 0.15 | 0.11 | 0.22 | 1.12 | 0.78 | 0.26 | 0.34 | 0.39 | 0.11 | 0.67 | 2.78 | 4.36 | 10.79 | 4.81 | NA | 0.19 | 12.81 | 1.8
Results from laboratories 2.07 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 606b (Certified Value 2.67 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.57 | 3.436 | 2.812 | 2.6459 | 2.6163 | 2.98 | 2.55 | 2.7893 | 2.82 | 2.72 | 2.7251 | 2.69 | 2.57 | 2.53 | 2.58 | 2.83 | 2.9548 | 2.704 | 2.22 | 2.55 | 2.56 | 2.6634 | 2.7 | 2.4766 | 2.1657 | 2.723 | 1
| 1 | 2 | 2.63 | 3.463 | 2.861 | 2.6108 | 2.6747 | 2.9 | 2.64 | 2.7254 | 2.65 | 2.75 | 2.7245 | 2.74 | 2.62 | 2.53 | 2.63 | 2.81 | 2.8364 | 2.771 | 2.34 | 2.57 | 2.51 | 2.6449 | 2.641 | 2.5064 | 2.6468 | 2.727 | 0
| 1 | 3 | 2.6 | 3.38 | 2.783 | 2.6053 | 2.6441 | 2.88 | 2.59 | 2.7213 | 2.89 | 2.76 | 2.7609 | 2.73 | 2.6 | 2.5 | 2.56 | 2.73 | 2.9075 | 2.723 | 2.11 | 2.59 | 2.54 | 2.6165 | 2.589 | 2.4823 | 2.3142 | 2.718 | 16
| 1 | 4 | 2.61 | 3.49 | 2.794 | 2.7039 | 2.6534 | 2.92 | 2.58 | 2.7589 | 2.75 | 2.79 | 2.7192 | 2.74 | 2.61 | 2.47 | 2.56 | 2.87 | 2.9133 | 2.698 | 2.3 | 2.58 | 2.49 | 2.7143 | 2.619 | 2.471 | 2.5557 | 2.738 | 2.6734
| 1 | 5 | 2.58 | 3.449 | 2.771 | 2.6494 | 2.6215 | 2.94 | 2.57 | 2.7825 | 2.75 | 2.74 | 2.6953 | 2.79 | 2.64 | 2.52 | 2.56 | 2.81 | 2.9839 | 2.713 | 2.31 | 2.49 | 2.61 | 2.6819 | 2.641 | 2.5199 | 2.5418 | 2.784 | 43
| 1 | 6 | 2.61 | 4.436 | 2.803 | 2.5615 | 2.6377 | 2.94 | 2.52 | 2.7423 | 2.71 | 2.73 | 2.7063 | 2.62 | 2.59 | 2.5 | 2.49 | 2.78 | 2.9106 | 2.68 | 2.26 | 2.58 | 2.57 | 2.6964 | 2.685 | 2.524 | 2.479 | 2.667
| Mean |  | 2.6 | 3.609 | 2.804 | 2.6295 | 2.6413 | 2.9267 | 2.575 | 2.7533 | 2.7617 | 2.7483 | 2.7219 | 2.7183 | 2.605 | 2.5083 | 2.5633 | 2.805 | 2.9177 | 2.7148 | 2.2567 | 2.56 | 2.5467 | 2.6696 | 2.6458 | 2.4967 | 2.4506 | 2.7262
| Median |  | 2.605 | 3.456 | 2.7985 | 2.6284 | 2.6409 | 2.93 | 2.575 | 2.7506 | 2.75 | 2.745 | 2.7218 | 2.735 | 2.605 | 2.51 | 2.56 | 2.81 | 2.912 | 2.7085 | 2.28 | 2.575 | 2.55 | 2.6726 | 2.641 | 2.4944 | 2.5104 | 2.725
| Std Dev. |  | 0.021909 | 0.40679 | 0.031445 | 0.048497 | 0.021452 | 0.035024 | 0.040373 | 0.028663 | 0.084004 | 0.024833 | 0.022344 | 0.057764 | 0.02429 | 0.023166 | 0.045019 | 0.047223 | 0.050085 | 0.031109 | 0.083106 | 0.036878 | 0.043205 | 0.035597 | 0.041146 | 0.023034 | 0.17801 | 0.037584
| Rel.Std.Dev. |  | 0.8427% | 11.27% | 1.121% | 1.844% | 0.8122% | 1.197% | 1.568% | 1.041% | 3.042% | 0.9036% | 0.8209% | 2.125% | 0.9324% | 0.9236% | 1.756% | 1.684% | 1.717% | 1.146% | 3.683% | 1.441% | 1.697% | 1.333% | 1.555% | 0.9226% | 7.264% | 1.379%
| PDM3 |  | -2.744% | 35% | 4.887% | -1.642% | -1.2% | 9.475% | -3.679% | 2.989% | 3.303% | 2.804% | 1.815% | 1.682% | -2.557% | -6.173% | -4.116% | 4.924% | 9.142% | 1.551% | -15.59% | -4.24% | -4.739% | -0.1419% | -1.03% | -6.608% | -8.334% | 1.975%
| Z-Score (Absolute) |  | 0.41 | 6.75 | 1.04 | 0.2 | 0.12 | 1.91 | 0.59 | 0.68 | 0.74 | 0.64 | 0.46 | 0.43 | 0.37 | 1.06 | 0.67 | 1.05 | 1.85 | 0.41 | 2.85 | 0.69 | 0.79 | 0.08 | 0.08 | 1.14 | 1.47 | 0.49
| Analytical results for Nb in OREAS 606b (Certified Value 18.5 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 19 | 17.53 | 18.1 | 18.39 | 19.0539 | 20.1 | 19.8 | 18 | 18.1 | 18.74 | 18.6 | 17.9 | 18.2 | 19.1 | 18 | 18.9 | 19 | 18.9 | 16.8 | 17.3581 | 8 | 18.3316 | 16.9 |  |  |  | 1
| 1 | 2 | 19 | 18.07 | 18.8 | 18.5 | 19.1584 | 21.5 | 20 | 17.3 | 18.1 | 19.03 | 18.9 | 18.2 | 17.8 | 20.1 | 18.2 | 16.21 | 19.5 | 21.5 | 16.9 | 16.0366 | 9 | 18.7811 | 17 |  |  |  | 29
| 1 | 3 | 19 | 17.85 | 18.2 | 18.41 | 19.4119 | 19.7 | 20.7 | 19 | 18.2 | 19.44 | 18.8 | 18.3 | 18.4 | 19.4 | 18 | 19.2 | 19.1 | 19.4 | 15.9 | 15.0665 | 10 | 18.0084 | 17.3 |  |  |  | 16
| 1 | 4 | 20 | 17.19 | 18.5 | 18.65 | 19.5551 | 19.8 | 20.2 | 18.1 | 17.8 | 19.46 | 19.3 | 18.4 | 17.8 | 19.6 | 18.3 | 18.6 | 19 | 21.9 | 17.9 | 17.5722 | 10 | 18.7841 | 17.4 |  |  |  | 18.4985
| 1 | 5 | 19 | 19.21 | 18.5 | 18.89 | 19.1765 | 19.6 | 20.1 | 17.6 | 18.1 | 18.81 | 19.9 | 18.4 | 18.7 | 19.2 | 18.2 | 20.25 | 19.1 | 21.3 | 15.6 | 16.5998 | 9 | 18.3202 | 17.5 |  |  |  | 44
| 1 | 6 | 20 | 19.59 | 17.9 | 18.55 | 19.3937 | 19.8 | 19.7 | 18 | 18.1 | 19.06 | 18.8 | 18 | 18.3 | 18.9 | 18.9 | 18.3 | 19.3 | 21.2 | 16.8 | 16.7349 | 10 | 18.4266 | 16.9
| Mean |  | 19.3333 | 18.24 | 18.3333 | 18.565 | 19.2916 | 20.0833 | 20.0833 | 18 | 18.0667 | 19.09 | 19.05 | 18.2 | 18.2 | 19.3833 | 18.2667 | 18.5767 | 19.1667 | 20.7 | 16.65 | 16.5614 | 9.3333 | 18.442 | 17.1667
| Median |  | 19 | 17.96 | 18.35 | 18.525 | 19.2851 | 19.8 | 20.05 | 18 | 18.1 | 19.045 | 18.85 | 18.25 | 18.25 | 19.3 | 18.2 | 18.75 | 19.1 | 21.25 | 16.8 | 16.6673 | 9.5 | 18.3791 | 17.15
| Std Dev. |  | 0.5164 | 0.95404 | 0.3266 | 0.18545 | 0.19069 | 0.71391 | 0.35449 | 0.57619 | 0.13663 | 0.30489 | 0.47645 | 0.20976 | 0.35214 | 0.42622 | 0.33267 | 1.3395 | 0.19664 | 1.2345 | 0.8167 | 0.91631 | 0.8165 | 0.2991 | 0.26583
| Rel.Std.Dev. |  | 2.671% | 5.231% | 1.781% | 0.9989% | 0.9885% | 3.555% | 1.765% | 3.201% | 0.7562% | 1.597% | 2.501% | 1.153% | 1.935% | 2.199% | 1.821% | 7.211% | 1.026% | 5.964% | 4.905% | 5.533% | 8.748% | 1.622% | 1.549%
| PDM3 |  | 4.513% | -1.398% | -0.893% | 0.3594% | 4.287% | 8.567% | 8.567% | -2.695% | -2.335% | 3.197% | 2.981% | -1.614% | -1.614% | 4.783% | -1.253% | 0.4225% | 3.612% | 11.9% | -9.993% | -10.47% | -49.55% | -0.3056% | -7.2%
| Z-Score (Absolute) |  | 0.99 | 0.31 | 0.2 | 0.07 | 0.94 | 1.88 | 1.88 | 0.6 | 0.52 | 0.7 | 0.65 | 0.36 | 0.36 | 1.05 | 0.28 | 0.09 | 0.79 | 2.61 | 2.21 | 2.31 | NA | 0.07 | 1.59
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 606b (Certified Value 37.4 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  | 1
2
| 1 | 1 | 2.01: 36.8; 2.03: 38.47; 2.04: 35.499; 2.05: 38.8; 2.06: 41.257; 2.07: 39.55; 2.08: 30.1; 2.10: 39.4; 2.11: 35.7315; 2.12: 41; 2.13: 34.4 |  |  |  | 1
| 1 | 2 | 2.01: 37.2; 2.03: 38.16; 2.04: 35.7963; 2.05: 37.7; 2.06: 41.545; 2.07: 40.24; 2.08: 33.3; 2.10: 38; 2.11: 35.5816; 2.12: 41; 2.13: 32.4 |  |  |  | 30
| 1 | 3 | 2.01: 37.4; 2.03: 37.65; 2.04: 35.8682; 2.05: 38.7; 2.06: 41.588; 2.07: 39.9; 2.08: 28.2; 2.10: 39.3; 2.11: 35.2436; 2.12: 41; 2.13: 35.4 |  |  |  | 16
| 1 | 4 | 2.01: 36.5; 2.03: 37.7; 2.04: 35.6843; 2.05: 39.5; 2.06: 41.528; 2.07: 39.7; 2.08: 25.8; 2.10: 39.2; 2.11: 36.4117; 2.12: 42; 2.13: 34.2 |  |  |  | 37.4253
| 1 | 5 | 2.01: 41; 2.03: 37.43; 2.04: 35.3795; 2.05: 38.2; 2.06: 41.092; 2.07: 39.36; 2.08: 28.1; 2.10: 38.6; 2.11: 37.4198; 2.12: 40; 2.13: 35.9 |  |  |  | 45
| 1 | 6 | 2.01: 40.1; 2.03: 37.35; 2.04: 35.9668; 2.05: 39.2; 2.06: 41.878; 2.07: 40.69; 2.08: 31.3; 2.10: 38.9; 2.11: 35.8545; 2.12: 41; 2.13: 32.2
| Mean |  | 2.01: 38.1667; 2.03: 37.7933; 2.04: 35.699; 2.05: 38.6833; 2.06: 41.4813; 2.07: 39.9067; 2.08: 29.4667; 2.10: 38.9; 2.11: 36.0404; 2.12: 41; 2.13: 34.0833
| Median |  | 2.01: 37.3; 2.03: 37.675; 2.04: 35.7403; 2.05: 38.75; 2.06: 41.5365; 2.07: 39.8; 2.08: 29.15; 2.10: 39.05; 2.11: 35.793; 2.12: 41; 2.13: 34.3
| Std Dev. |  | 2.01: 1.8938; 2.03: 0.43592; 2.04: 0.22456; 2.05: 0.65549; 2.06: 0.27454; 2.07: 0.48858; 2.08: 2.6598; 2.10: 0.52915; 2.11: 0.77666; 2.12: 0.63246; 2.13: 1.5184
| Rel.Std.Dev. |  | 2.01: 4.962%; 2.03: 1.153%; 2.04: 0.629%; 2.05: 1.695%; 2.06: 0.6618%; 2.07: 1.224%; 2.08: 9.027%; 2.10: 1.36%; 2.11: 2.155%; 2.12: 1.543%; 2.13: 4.455%
| PDM3 |  | 2.01: 1.981%; 2.03: 0.9834%; 2.04: -4.613%; 2.05: 3.361%; 2.06: 10.84%; 2.07: 6.63%; 2.08: -21.27%; 2.10: 3.94%; 2.11: -3.7%; 2.12: 9.552%; 2.13: -8.93%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.12; 2.04: 0.78; 2.05: 0.16; 2.06: 1.05; 2.07: 0.55; 2.08: 2.76; 2.10: 0.23; 2.11: 0.67; 2.12: 0.9; 2.13: 1.29
| Analytical results for Ni in OREAS 606b (Certified Value 3.56 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | < 5 | 3.7 | 3.4 | 3.5 | 3.9832 | 4 | 3.9 | 3.6 | 4 | 3.5 | 3.5 | 3.3 | 3.5 | 3.6 | 5 | 3 | 5 | 3 | 3.9 | 3.6 | 2.656 | 3 | 3.0279 | < 5 | 4 |  | 1
| 1 | 2 | < 5 | 4.9 | 3.6 | 3.4 | 4.0652 | 4 | 3.6 | 3.6 | 4 | 3.8 | 3.6 | 3.4 | 3.4 | 3.7 | 3.6 | 3.6 | 5 | 2.4 | 3.68 | 3.3 | 2.6327 | 2 | 3.2195 | < 5 | 4 |  | 16
| 1 | 3 | < 5 | 4.4 | 3.7 | 3.5 | 4.0861 | 4 | 3.2 | 3.4 | 4 | 3.5 | 3.8 | 3.3 | 3.7 | 3.8 | 3.8 | 3.3 | 5 | 2.5 | 3.35 | 3.9 | 2.8459 | 3 | 3.014 | < 5 | 3 |  | 16
| 1 | 4 | < 5 | 3.8 | 3.6 | 4 | 4.0114 | 4 | 3.1 | 3.5 | 4 | 3.9 | 3.6 | 3.2 | 3.2 | 3.9 | 4.3 | 3.5 | 5 | 2.9 | 3.89 | 3.4 | 3.2422 | 3 | 2.9643 | < 5 | 4 |  | 3.561
| 1 | 5 | 5 | 4.8 | 3.8 | 3.4 | 3.9994 | 4 | 3.4 | 3.4 | 4 | 3.5 | 4.1 | 3.5 | 3.4 | 3.8 | 3.5 | 3.4 | 5 | 3.5 | 3.99 | 3.8 | 1.6468 | 4 | 3.4475 | < 5 | 4 |  | 46
| 1 | 6 | < 5 | 4.5 | 3.7 | 3.9 | 3.9726 | 4 | 3.4 | 3.4 | 4 | 3.9 | 3.6 | 3.3 | 3.5 | 3.7 | 3.7 | 3.1 | 6 | 3.9 | 3.77 | 3.6 | 2.7026 | 4 | 3.3386 | < 5 | 4
| Mean |  | 5 | 4.35 | 3.6333 | 3.6167 | 4.0197 | 4 | 3.4333 | 3.4833 | 4 | 3.6833 | 3.7 | 3.3333 | 3.45 | 3.75 | 3.9833 | 3.3167 | 5.1667 | 3.0333 | 3.7633 | 3.6 | 2.621 | 3.1667 | 3.1686 |  | 3.8333
| Median |  | 5 | 4.45 | 3.65 | 3.5 | 4.0054 | 4 | 3.4 | 3.45 | 4 | 3.65 | 3.6 | 3.3 | 3.45 | 3.75 | 3.75 | 3.35 | 5 | 2.95 | 3.83 | 3.6 | 2.6793 | 3 | 3.1237 |  | 4
| Std Dev. |  |  | 0.501 | 0.13663 | 0.26394 | 0.045856 | 0 | 0.28752 | 0.098319 | 0 | 0.20412 | 0.21909 | 0.10328 | 0.16432 | 0.10488 | 0.57067 | 0.23166 | 0.40825 | 0.5785 | 0.22975 | 0.22804 | 0.52793 | 0.75277 | 0.19736 |  | 0.40825
| Rel.Std.Dev. |  |  | 11.52% | 3.76% | 7.298% | 1.141% | 0% | 8.374% | 2.823% | 0% | 5.542% | 5.921% | 3.098% | 4.763% | 2.797% | 14.33% | 6.985% | 7.902% | 19.07% | 6.105% | 6.334% | 20.14% | 23.77% | 6.228% |  | 10.65%
| PDM3 |  | 40.41% | 22.16% | 2.03% | 1.562% | 12.88% | 12.33% | -3.586% | -2.182% | 12.33% | 3.434% | 3.902% | -6.394% | -3.118% | 5.306% | 11.86% | -6.862% | 45.09% | -14.82% | 5.681% | 1.094% | -26.4% | -11.07% | -11.02% |  | 7.646%
| Z-Score (Absolute) |  | 1.96 | 2.54 | 0.29 | 0.24 | 1.5 | NA | 0.34 | 0.18 | NA | 0.44 | 0.5 | 0.65 | 0.29 | 0.65 | 1.39 | 0.71 | NA | 1.59 | 0.7 | 0.18 | 2.89 | NA | 1.17 | 3.27 | NA
Results from laboratories 2.07, 2.11, 2.19, 2.25 and 2.28 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 606b (Certified Value 0.03 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.0296 | 0.03 | 0.029 | 0.0299 | 0.030747 | 0.029 | 0.0315 | 0.0312 | 0.03 | 0.0301 | 0.03 | 0.028 | 0.032 | 0.03 | 0.033 | 0.03 | 0.028 | 0.032 | 0.03 | 0.0302 | < 0.1 | 0.031 | 0.030077 | 0.026974 | 0.028 |  | 1
| 1 | 2 | 0.0299 | 0.029 | 0.031 | 0.0295 | 0.031553 | 0.029 | 0.0308 | 0.0326 | 0.03 | 0.0302 | 0.03 | 0.029 | 0.03 | 0.03 | 0.032 | 0.0292 | 0.029 | 0.0306 | 0.03 | 0.0305 | < 0.1 | 0.031 | 0.030568 | 0.027541 | 0.032 |  | 0
| 1 | 3 | 0.0295 | 0.03 | 0.029 | 0.0296 | 0.031494 | 0.028 | 0.0303 | 0.0324 | 0.029 | 0.0305 | 0.03 | 0.029 | 0.03 | 0.029 | 0.031 | 0.0297 | 0.03 | 0.033 | 0.03 | 0.0299 | < 0.1 | 0.031 | 0.029849 | 0.027061 | 0.029 |  | 16
| 1 | 4 | 0.0309 | 0.03 | 0.029 | 0.0295 | 0.030861 | 0.029 | 0.03 | 0.0307 | 0.03 | 0.0304 | 0.03 | 0.028 | 0.03 | 0.029 | 0.032 | 0.0296 | 0.029 | 0.031 | 0.03 | 0.0301 | < 0.1 | 0.031 | 0.030077 | 0.026883 | 0.029 |  | 0.029991
| 1 | 5 | 0.0296 | 0.03 | 0.029 | 0.0298 | 0.030916 | 0.028 | 0.0301 | 0.0323 | 0.03 | 0.0301 | 0.031 | 0.029 | 0.031 | 0.029 | 0.032 | 0.03 | 0.028 | 0.0312 | 0.03 | 0.0297 | < 0.1 | 0.031 | 0.03118 | 0.026832 | 0.031 |  | 47
| 1 | 6 | 0.0299 | 0.031 | 0.03 | 0.0297 | 0.030864 | 0.029 | 0.0301 | 0.0307 | 0.03 | 0.0302 | 0.029 | 0.028 | 0.03 | 0.028 | 0.032 | 0.0293 | 0.029 | 0.026 | 0.03 | 0.0308 | < 0.1 | 0.031 | 0.031012 | 0.026268 | 0.026
| Mean |  | 0.0299 | 0.03 | 0.0295 | 0.029667 | 0.031072 | 0.028667 | 0.030467 | 0.03165 | 0.029833 | 0.03025 | 0.03 | 0.0285 | 0.0305 | 0.029167 | 0.032 | 0.029633 | 0.028833 | 0.030633 | 0.03 | 0.0302 |  | 0.031 | 0.03046 | 0.026926 | 0.029167
| Median |  | 0.02975 | 0.03 | 0.029 | 0.02965 | 0.03089 | 0.029 | 0.0302 | 0.03175 | 0.03 | 0.0302 | 0.03 | 0.0285 | 0.03 | 0.029 | 0.032 | 0.02965 | 0.029 | 0.0311 | 0.03 | 0.03015 |  | 0.031 | 0.030322 | 0.026928 | 0.029
| Std Dev. |  | 0.00051769 | 0.00063246 | 0.00083666 | 0.0001633 | 0.00035414 | 0.0005164 | 0.00058195 | 0.00088261 | 0.00040825 | 0.00016432 | 0.00063246 | 0.00054772 | 0.00083666 | 0.00075277 | 0.00063246 | 0.00033862 | 0.00075277 | 0.0024246 | 0 | 0.0004 |  | 0 | 0.00054804 | 0.0004106 | 0.002137
| Rel.Std.Dev. |  | 1.731% | 2.108% | 2.836% | 0.5504% | 1.14% | 1.801% | 1.91% | 2.789% | 1.368% | 0.5432% | 2.108% | 1.922% | 2.743% | 2.581% | 1.976% | 1.143% | 2.611% | 7.915% | 0% | 1.325% |  | 0% | 1.799% | 1.525% | 7.327%
| PDM3 |  | -0.303% | 0.03043% | -1.637% | -1.081% | 3.606% | -4.415% | 1.586% | 5.532% | -0.5253% | 0.864% | 0.03043% | -4.971% | 1.698% | -2.748% | 6.699% | -1.192% | -3.86% | 2.142% | 0.03043% | 0.6973% |  | 3.365% | 1.566% | -10.22% | -2.748%
| Z-Score (Absolute) |  | 0.13 | 0 | 0.67 | 0.45 | 1.45 | 1.8 | 0.63 | 2.23 | 0.22 | 0.34 | 0 | 2.02 | 0.67 | 1.12 | 2.7 | 0.49 | 1.57 | 0.85 | 0 | 0.27 | 26.97 | 0.74 | 0.62 | 4.15 | 1.12
| Analytical results for Pb in OREAS 606b (Certified Value 157 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 151 | 136.08 | 159.3 | 161.3 | 156.3126 | 152 | 154 | 159.742 | 161 | 159 | 162 | 159 | 156 | 160.5 | 157.5 | 166.5 | 168 | 157 | 259 | 122.4 | 150 | 132.9654 | 172 | 151.094 | 141.512 | 144 | 1
| 1 | 2 | 140 | 147.27 | 160.5 | 160.6 | 156.8639 | 150 | 161 | 163.625 | 160 | 160 | 162 | 162.5 | 161.5 | 155.5 | 161 | 163.5 | 168 | 164 | 248 | 128.7 | 149 | 130.4399 | 163 | 158.3458 | 151.249 | 148 | 18
| 1 | 3 | 142 | 144.51 | 157.4 | 160.9 | 155.2146 | 151 | 156 | 161.7 | 162 | 160 | 164 | 161.5 | 159 | 157 | 157.5 | 158 | 172 | 154 | 271 | 127.8 | 155 | 131.0465 | 168 | 151.4502 | 150.225 | 149 | 16
| 1 | 4 | 166 | 135.94 | 158.7 | 160.7 | 150.0186 | 148 | 153 | 167.155 | 160 | 160 | 164 | 162.5 | 159.5 | 154.5 | 157.5 | 167.5 | 169 | 149 | 278 | 132.4 | 152 | 133.2763 | 167 | 153.4724 | 138.925 | 146 | 157.0775
| 1 | 5 | 136 | 148.43 | 162.7 | 160.5 | 157.9074 | 156 | 158 | 164.978 | 158 | 160 | 162 | 166 | 162.5 | 158.5 | 158 | 164 | 170 | 153 | 257 | 121.6 | 150 | 132.0814 | 167 | 158.5024 | 143.103 | 146 | 48
| 1 | 6 | 148 | 157.9 | 156.7 | 161.2 | 151.7998 | 152 | 154 | 163.526 | 158 | 161 | 161 | 158 | 159 | 155.5 | 157.5 | 162.5 | 168 | 159 | 248 | 132 | 151 | 129.5662 | 167 | 161.6589 | 155.178 | 146
| Mean |  | 147.1667 | 145.0217 | 159.2167 | 160.8667 | 154.6861 | 151.5 | 156 | 163.4543 | 159.8333 | 160 | 162.5 | 161.5833 | 159.5833 | 156.9167 | 158.1667 | 163.6667 | 169.1667 | 156 | 260.1667 | 127.4833 | 151.1667 | 131.5626 | 167.3333 | 155.754 | 146.6987 | 146.5
| Median |  | 145 | 145.89 | 159 | 160.8 | 155.7636 | 151.5 | 155 | 163.5755 | 160 | 160 | 162 | 162 | 159.25 | 156.25 | 157.5 | 163.75 | 168.5 | 155.5 | 258 | 128.25 | 150.5 | 131.564 | 167 | 155.9091 | 146.664 | 146
| Std Dev. |  | 10.7036 | 8.3092 | 2.1766 | 0.3266 | 3.1036 | 2.6646 | 3.0332 | 2.5655 | 1.6021 | 0.63246 | 1.2247 | 2.8534 | 2.2675 | 2.2454 | 1.4024 | 3.3566 | 1.6021 | 5.2154 | 12.1888 | 4.6175 | 2.137 | 1.4619 | 2.8752 | 4.3493 | 6.4078 | 1.7607
| Rel.Std.Dev. |  | 7.273% | 5.73% | 1.367% | 0.203% | 2.006% | 1.759% | 1.944% | 1.57% | 1.002% | 0.3953% | 0.7537% | 1.766% | 1.421% | 1.431% | 0.8866% | 2.051% | 0.947% | 3.343% | 4.685% | 3.622% | 1.414% | 1.111% | 1.718% | 2.792% | 4.368% | 1.202%
| PDM3 |  | -6.309% | -7.675% | 1.362% | 2.412% | -1.522% | -3.551% | -0.6859% | 4.06% | 1.754% | 1.861% | 3.452% | 2.869% | 1.595% | -0.1024% | 0.6934% | 4.195% | 7.696% | -0.6859% | 65.63% | -18.84% | -3.763% | -16.24% | 6.529% | -0.8426% | -6.607% | -6.734%
| Z-Score (Absolute) |  | 1.29 | 1.55 | 0.21 | 0.41 | 0.35 | 0.75 | 0.19 | 0.73 | 0.28 | 0.3 | 0.62 | 0.5 | 0.25 | 0.08 | 0.08 | 0.76 | 1.44 | 0.19 | 12.73 | 3.73 | 0.79 | 3.22 | 1.21 | 0.22 | 1.35 | 1.37
| Analytical results for Pr in OREAS 606b (Certified Value 10.1 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 1
2
| 1 | 1 | 2.01: 9.4; 2.03: 9.85; 2.04: 10.401; 2.05: 10.3; 2.06: 11.262; 2.07: 10.53; 2.08: 9.2; 2.09: 10.1; 2.10: 9.7466; 2.11: 9; 2.12: 9.43 | 1
| 1 | 2 | 2.01: 9.5; 2.03: 9.64; 2.04: 10.7571; 2.05: 9.8; 2.06: 11.215; 2.07: 11.01; 2.08: 10.05; 2.09: 10; 2.10: 9.8735; 2.11: 8; 2.12: 8.93 | 33
| 1 | 3 | 2.01: 9.7; 2.03: 9.74; 2.04: 10.7904; 2.05: 9.9; 2.06: 11.279; 2.07: 10.61; 2.08: 8.52; 2.09: 10.11; 2.10: 9.9666; 2.11: 9; 2.12: 9.62 | 16
| 1 | 4 | 2.01: 9.5; 2.03: 9.97; 2.04: 10.4945; 2.05: 10.1; 2.06: 11.311; 2.07: 10.63; 2.08: 7.76; 2.09: 10.14; 2.10: 10.2715; 2.11: 9; 2.12: 9.38 | 10.0749
| 1 | 5 | 2.01: 10.4; 2.03: 9.72; 2.04: 10.6973; 2.05: 10.1; 2.06: 11.662; 2.07: 10.53; 2.08: 8.61; 2.09: 10.1; 2.10: 10.2773; 2.11: 8; 2.12: 9.62 | 49
| 1 | 6 | 2.01: 10.3; 2.03: 9.69; 2.04: 10.0894; 2.05: 10.2; 2.06: 11.235; 2.07: 10.85; 2.08: 9.56; 2.09: 10.12; 2.10: 9.9767; 2.11: 9; 2.12: 8.94
| Mean |  | 2.01: 9.8; 2.03: 9.7683; 2.04: 10.5383; 2.05: 10.0667; 2.06: 11.3273; 2.07: 10.6933; 2.08: 8.95; 2.09: 10.095; 2.10: 10.0187; 2.11: 8.6667; 2.12: 9.32
| Median |  | 2.01: 9.6; 2.03: 9.73; 2.04: 10.5959; 2.05: 10.1; 2.06: 11.2705; 2.07: 10.62; 2.08: 8.905; 2.09: 10.105; 2.10: 9.9717; 2.11: 9; 2.12: 9.405
| Std Dev. |  | 2.01: 0.43818; 2.03: 0.1209; 2.04: 0.26798; 2.05: 0.18619; 2.06: 0.16734; 2.07: 0.19449; 2.08: 0.8199; 2.09: 0.048888; 2.10: 0.21465; 2.11: 0.5164; 2.12: 0.31375
| Rel.Std.Dev. |  | 2.01: 4.471%; 2.03: 1.238%; 2.04: 2.543%; 2.05: 1.85%; 2.06: 1.477%; 2.07: 1.819%; 2.08: 9.161%; 2.09: 0.4843%; 2.10: 2.143%; 2.11: 5.958%; 2.12: 3.366%
| PDM3 |  | 2.01: -2.728%; 2.03: -3.043%; 2.04: 4.6%; 2.05: -0.08145%; 2.06: 12.43%; 2.07: 6.139%; 2.08: -11.17%; 2.09: 0.1998%; 2.10: -0.5576%; 2.11: -13.98%; 2.12: -7.493%
| Z-Score (Absolute) |  | 2.01: 0.43; 2.03: 0.48; 2.04: 0.87; 2.05: 0.04; 2.06: 2.25; 2.07: 1.14; 2.08: 1.91; 2.09: 0.09; 2.10: 0.04; 2.11: NA; 2.12: 1.27
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 606b (Certified Value 139 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 154 | 114.8 | 137.5 | 132.43 | 136.9061 | 142 | 147 | 153.042 | 139 | 127 | 142.72 | 136.5 | 124 | 132.5 | 145 | 144.5 | 150.2 | 145 | 124 | 134.6 | 140.1802 | 126 |  |  |  |  | 1
| 1 | 2 | 158 | 109.4 | 142.6 | 132.46 | 138.7547 | 140 | 151 | 153.1992 | 146 | 128 | 147.05 | 138.5 | 127 | 129.5 | 150 | 144.5 | 149.2 | 151 | 132 | 132.8 | 139.7857 | 125 |  |  |  |  | 34
| 1 | 3 | 161 | 113.4 | 138 | 131.28 | 131.8642 | 142 | 156 | 153.6152 | 144 | 127 | 144.81 | 138.5 | 126 | 133.5 | 145.5 | 144 | 149.6 | 136 | 111 | 132.6 | 139.069 | 126 |  |  |  |  | 16
| 1 | 4 | 159 | 108.5 | 140.3 | 137.11 | 131.7259 | 143 | 152 | 153.33 | 138 | 126 | 141.85 | 140.5 | 128 | 127 | 145.5 | 147 | 148.7 | 128 | 118 | 133.5 | 138.673 | 121 |  |  |  |  | 139.0542
| 1 | 5 | 160 | 113.8 | 140.3 | 132.59 | 137.2184 | 140 | 150 | 153.5224 | 140 | 129 | 142.21 | 147.5 | 130.5 | 134.5 | 145.5 | 144.5 | 151.8 | 135 | 135 | 130.6 | 139.2859 | 131 |  |  |  |  | 50
| 1 | 6 | 147 | 133.2 | 138.9 | 128.87 | 136.8018 | 142 | 149 | 153.87 | 135 | 127 | 142.96 | 136 | 128.5 | 132.5 | 136.5 | 141.5 | 146.8 | 141 | 118 | 134.9 | 140.047 | 122
| Mean |  | 156.5 | 115.5167 | 139.6 | 132.4567 | 135.5452 | 141.5 | 150.8333 | 153.4298 | 140.3333 | 127.3333 | 143.6 | 139.5833 | 127.3333 | 131.5833 | 144.6667 | 144.3333 | 149.3833 | 139.3333 | 123 | 133.1667 | 139.5068 | 125.1667
| Median |  | 158.5 | 113.6 | 139.6 | 132.445 | 136.8539 | 142 | 150.5 | 153.4262 | 139.5 | 127 | 142.84 | 138.5 | 127.5 | 132.5 | 145.5 | 144.5 | 149.4 | 138.5 | 121 | 133.15 | 139.5358 | 125.5
| Std Dev. |  | 5.244 | 9.0254 | 1.8676 | 2.6803 | 2.9889 | 1.2247 | 3.0605 | 0.30008 | 4.0332 | 1.0328 | 1.9771 | 4.2002 | 2.2286 | 2.8003 | 4.4121 | 1.7512 | 1.657 | 8.1158 | 9.1652 | 1.5629 | 0.5931 | 3.5449
| Rel.Std.Dev. |  | 3.351% | 7.813% | 1.338% | 2.024% | 2.205% | 0.8655% | 2.029% | 0.1956% | 2.874% | 0.8111% | 1.377% | 3.009% | 1.75% | 2.128% | 3.05% | 1.213% | 1.109% | 5.825% | 7.451% | 1.174% | 0.4251% | 2.832%
| PDM3 |  | 12.55% | -16.93% | 0.3925% | -4.745% | -2.524% | 1.759% | 8.471% | 10.34% | 0.9199% | -8.429% | 3.269% | 0.3805% | -8.429% | -5.373% | 4.036% | 3.796% | 7.428% | 0.2007% | -11.55% | -4.234% | 0.3255% | -9.987%
| Z-Score (Absolute) |  | 1.7 | 2.41 | 0.01 | 0.71 | 0.4 | 0.2 | 1.13 | 1.39 | 0.08 | 1.22 | 0.41 | 0 | 1.22 | 0.8 | 0.51 | 0.48 | 0.99 | 0.02 | 1.66 | 0.64 | 0 | 1.44
| Analytical results for Re in OREAS 606b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.002 | 0.016 | < 0.005 | < 0.01 | 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 1
| 1 | 2 | 0.022 | 0.017 | < 0.005 | < 0.01 | 0.001 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 35
| 1 | 3 | 0.014 | 0.017 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 16
| 1 | 4 | 0.04 | 0.028 | < 0.005 | < 0.01 | 0.001 | < 0.002 | < 0.05 | < 0.002 | 0.002 | 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | < 0.002
| 1 | 5 | 0.013 | 0.025 | < 0.005 | < 0.01 | 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 51
| 1 | 6 | 0.01 | 0.032 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | 0.002 | < 0.002 | < 0.05 | < 0.1
| Mean |  | 2.01: 0.0198; 2.03: 0.0225; 2.06: 0.001; 2.09: 0.002; 2.10: 0.002; 2.11: 0.002; 2.13: 0.002; 2.16: 0.002
| Median |  | 2.01: 0.014; 2.03: 0.021; 2.06: 0.001; 2.09: 0.002; 2.10: 0.002; 2.11: 0.002; 2.13: 0.002; 2.16: 0.002
| Std Dev. |  | 2.01: 0.012133; 2.03: 0.006775; 2.06: 0; 2.10: 0; 2.11: 0; 2.13: 0; 2.16: 0
| Rel.Std.Dev. |  | 2.01: 61.28%; 2.03: 30.11%; 2.06: 0%; 2.10: 0%; 2.11: 0%; 2.13: 0%; 2.16: 0%
PDM3
| Z-Score (Absolute) |  | 13.41 | 18.66 | 0.67 | 2.92 | 0.82 | 0.67 | 20.9 | 0.52 | 0.22 | 0.37 | 0.67 | 0.22 | 0.22 | 43.38 | 0.37 | 0.67 | 20.9 | 43.38
| Analytical results for S in OREAS 606b (Certified Value 0.468 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.47 | 0.43 | 0.4 | 0.4645 | 0.49667 | 0.46 | 0.51 | 0.45887 | 0.47 | 0.451 | 0.48 | 0.48 | 0.47 | 0.47 | 0.46 | 0.51 | 0.4541 | 0.448 | 0.52 | 0.21 | 0.46 | 0.4894 | 0.479 | 0.46277 | 0.456 | 0.469 | 1
| 1 | 2 | 0.47 | 0.42 | 0.5 | 0.4565 | 0.50083 | 0.46 | 0.5 | 0.45863 | 0.46 | 0.455 | 0.47 | 0.48 | 0.48 | 0.45 | 0.47 | 0.5 | 0.4563 | 0.474 | 0.517 | 0.22 | 0.47 | 0.4877 | 0.463 | 0.46975 | 0.458 | 0.456 | 11
| 1 | 3 | 0.47 | 0.42 | 0.4 | 0.4567 | 0.49176 | 0.43 | 0.5 | 0.45411 | 0.47 | 0.452 | 0.47 | 0.48 | 0.47 | 0.46 | 0.45 | 0.5 | 0.4541 | 0.447 | 0.48 | 0.22 | 0.47 | 0.4697 | 0.469 | 0.45894 | 0.457 | 0.46 | 16
| 1 | 4 | 0.47 | 0.43 | 0.4 | 0.469 | 0.49392 | 0.45 | 0.49 | 0.45452 | 0.47 | 0.443 | 0.45 | 0.48 | 0.48 | 0.45 | 0.45 | 0.51 | 0.4433 | 0.461 | 0.5 | 0.22 | 0.47 | 0.4909 | 0.469 | 0.45491 | 0.447 | 0.474 | 0.4683
| 1 | 5 | 0.46 | 0.43 | 0.4 | 0.47 | 0.49488 | 0.43 | 0.49 | 0.45295 | 0.49 | 0.45 | 0.47 | 0.49 | 0.49 | 0.46 | 0.45 | 0.51 | 0.4639 | 0.464 | 0.508 | 0.22 | 0.47 | 0.4949 | 0.465 | 0.4715 | 0.455 | 0.485 | 52
| 1 | 6 | 0.46 | 0.47 | 0.4 | 0.4466 | 0.49436 | 0.46 | 0.49 | 0.45134 | 0.46 | 0.452 | 0.48 | 0.47 | 0.47 | 0.45 | 0.44 | 0.5 | 0.4461 | 0.464 | 0.49 | 0.21 | 0.46 | 0.5014 | 0.465 | 0.47533 | 0.454 | 0.448
| Mean |  | 0.46667 | 0.43333 | 0.41667 | 0.46055 | 0.4954 | 0.44833 | 0.49667 | 0.45507 | 0.47 | 0.4505 | 0.47 | 0.48 | 0.47667 | 0.45667 | 0.45333 | 0.505 | 0.45297 | 0.45967 | 0.5025 | 0.21667 | 0.46667 | 0.489 | 0.46833 | 0.46553 | 0.4545 | 0.46533
| Median |  | 0.47 | 0.43 | 0.4 | 0.4606 | 0.49462 | 0.455 | 0.495 | 0.45431 | 0.47 | 0.4515 | 0.47 | 0.48 | 0.475 | 0.455 | 0.45 | 0.505 | 0.4541 | 0.4625 | 0.504 | 0.22 | 0.47 | 0.49015 | 0.467 | 0.46626 | 0.4555 | 0.4645
| Std Dev. |  | 0.005164 | 0.018619 | 0.040825 | 0.0089681 | 0.0030944 | 0.01472 | 0.008165 | 0.0030592 | 0.010954 | 0.0040373 | 0.010954 | 0.0063246 | 0.008165 | 0.008165 | 0.010328 | 0.0054772 | 0.0074002 | 0.010405 | 0.015592 | 0.005164 | 0.005164 | 0.010644 | 0.0057504 | 0.0079162 | 0.003937 | 0.013352
| Rel.Std.Dev. |  | 1.107% | 4.297% | 9.798% | 1.947% | 0.6246% | 3.283% | 1.644% | 0.6723% | 2.331% | 0.8962% | 2.331% | 1.318% | 1.713% | 1.788% | 2.278% | 1.085% | 1.634% | 2.264% | 3.103% | 2.383% | 1.107% | 2.177% | 1.228% | 1.7% | 0.8662% | 2.869%
| PDM3 |  | -0.3489% | -7.467% | -11.03% | -1.655% | 5.787% | -4.264% | 6.057% | -2.825% | 0.3629% | -3.801% | 0.3629% | 2.498% | 1.786% | -2.484% | -3.196% | 7.837% | -3.274% | -1.844% | 7.303% | -53.73% | -0.3489% | 4.42% | 0.006971% | -0.5912% | -2.947% | -0.6336%
| Z-Score (Absolute) |  | 0.07 | 1.86 | 2.82 | 0.28 | 1.73 | 0.99 | 1.81 | 0.6 | 0.26 | 0.86 | 0.26 | 0.84 | 0.65 | 0.51 | 0.7 | 2.29 | 0.72 | 0.33 | 2.14 | 14.38 | 0.07 | 1.36 | 0.17 | 0.01 | 0.63 | 0.01
| Analytical results for Sb in OREAS 606b (Certified Value 6.38 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 6.76 | 5.53 | 6 | 6.75 | 6.2053 | 6.8 | 6.7 | 6.7 | 6.7 | 6.65 | 6.31 | 5.78 | 6.62 | 6.41 | 6 | 6.05 | 6.22 | 11 | 5.04 | 5.62 | 13.7316 | 6.588 | 4.55 | 6.8 |  |  | 1
| 1 | 2 | 6.58 | 6.03 | 6.1 |  | 6.5588 | 6.9 | 6.99 | 6.29 | 6.7 | 6.74 | 6.45 | 5.91 | 6.45 | 6.52 | 6.1 | 5.95 | 6.46 | 7 | 5.97 | 5.68 | 13.6951 | 6.172 | 7.18 | 6.81 |  |  | 36
| 1 | 3 | 6.9 | 5.98 | 6.1 | 6.58 | 6.6128 | 6.5 | 7.11 | 6.88 | 6.7 | 7.24 | 6.59 | 5.84 | 6.55 | 6.41 | 5.87 | 6.01 | 6.21 | 7 | 5.59 | 5.69 | 12.219 | 6.4639 | 3.65 | 6.67 |  |  | 16
| 1 | 4 | 7.28 | 5.53 | 6.1 | 6.87 | 6.5055 | 6.8 | 6.79 | 6.62 | 6.7 | 6.78 | 6.6 | 6.07 | 6.33 | 6.44 | 6.35 | 5.99 | 6.22 | 10 | 5.86 | 5.63 | 9.3886 | 6.165 | 5.55 | 7.14 |  |  | 6.3769
| 1 | 5 | 6.25 | 5.97 | 6.2 | 6.59 | 6.736 | 6.4 | 6.97 | 6.37 | 6.6 | 6.8 | 6.91 | 5.94 | 6.71 | 6.43 | 6.17 | 6.1 | 6.33 | 11 | 5.6 | 5.6 | 8.4588 | 6.5321 | 4.42 | 7.05 |  |  | 53
| 1 | 6 | 7.78 | 6.53 | 5.9 | 6.56 | 6.2813 | 6.5 | 6.77 | 6.52 | 6.7 | 6.72 | 6.41 | 5.82 | 6.49 | 6.33 | 5.97 | 6.03 | 6.46 | 8 | < 5 | 5.64 | 9.9075 | 6.6439 | 5.41 | 6.7
| Mean |  | 6.925 | 5.9283 | 6.0667 | 6.67 | 6.4833 | 6.65 | 6.8883 | 6.5633 | 6.6833 | 6.8217 | 6.545 | 5.8933 | 6.525 | 6.4233 | 6.0767 | 6.0217 | 6.3167 | 9 | 5.612 | 5.6433 | 11.2334 | 6.4275 | 5.1267 | 6.8617
| Median |  | 6.83 | 5.975 | 6.1 | 6.59 | 6.5322 | 6.65 | 6.88 | 6.57 | 6.7 | 6.76 | 6.52 | 5.875 | 6.52 | 6.42 | 6.05 | 6.02 | 6.275 | 9 | 5.6 | 5.635 | 11.0633 | 6.498 | 4.98 | 6.805
| Std Dev. |  | 0.54029 | 0.37258 | 0.10328 | 0.13509 | 0.20245 | 0.20736 | 0.15829 | 0.21732 | 0.040825 | 0.21151 | 0.21012 | 0.10463 | 0.13323 | 0.06121 | 0.16967 | 0.051543 | 0.11944 | 1.8974 | 0.35968 | 0.035024 | 2.2861 | 0.20932 | 1.2238 | 0.19094
| Rel.Std.Dev. |  | 7.802% | 6.285% | 1.702% | 2.025% | 3.123% | 3.118% | 2.298% | 3.311% | 0.6108% | 3.101% | 3.21% | 1.775% | 2.042% | 0.9529% | 2.792% | 0.856% | 1.891% | 21.08% | 6.409% | 0.6206% | 20.35% | 3.257% | 23.87% | 2.783%
| PDM3 |  | 8.595% | -7.034% | -4.865% | 4.597% | 1.669% | 4.283% | 8.02% | 2.924% | 4.806% | 6.975% | 2.636% | -7.583% | 2.323% | 0.7285% | -4.708% | -5.57% | -0.9442% | 41.13% | -11.99% | -11.5% | 76.16% | 0.7936% | -19.61% | 7.602%
| Z-Score (Absolute) |  | 0.77 | 1.05 | 0.8 | 0.3 | 0.04 | 0.27 | 0.7 | 0.11 | 0.33 | 0.58 | 0.07 | 1.11 | 0.04 | 0.15 | 0.78 | 0.88 | 0.34 | 4.54 | 2.57 | 1.57 | 8.6 | 0.14 | 2.5 | 0.65
| Analytical results for Sc in OREAS 606b (Certified Value 4.26 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 5.5 | 4 | 4 | 4 | 3.9992 | 4 | 5 | 5.8303 | 4.77 | 4 | 4 | 3.9 | 3.7 | 3.8 | 4.2 | 4.1 | 4.11 | 2.7 | 4.3 | < 5 | 4.8035 | 3 | 4.5715 | 4.7 |  |  | 1
| 1 | 2 | 5.5 | 3.9 | 4 | 4.1 | 4.0204 | 4 | 5.1 | 5.8277 | 4.7 | 4 | 4 | 3.8 | 3.7 | 3.7 | 4.4 | 4.2 | 4.08 | 2.8 | 4.3 | < 5 | 4.937 | 3 | 4.182 | 4.4 |  |  | 37
| 1 | 3 | 5.8 | 4 | 4 | 3.8 | 4.0499 | 4 | 5.2 | 5.9499 | 4.97 | 4 | 4.1 | 3.9 | 3.8 | 3.9 | 4.3 | 4.1 | 4.14 | 3.7 | 4.3 | < 5 | 5.2091 | 3 | 4.1142 | 4.7 |  |  | 16
| 1 | 4 | 5.7 | 3.9 | 4 | 3.9 | 3.981 | 4 | 5 | 5.8562 | 4.77 | 4 | 4.1 | 4 | 3.8 | 3.6 | 4.3 | 4.1 | 4.14 | 7 | 4.2 | < 5 | 5.0805 | 3 | 4.191 | 4.9 |  |  | 4.2592
| 1 | 5 | 5.6 | 4.4 | 4 | 3.6 | 3.9544 | 4 | 5.1 | 5.8757 | 4.83 | 4 | 4 | 4.2 | 3.9 | 3.9 | 4.2 | 4.1 | 4.15 | 2.5 | 4 | < 5 | 5.132 | 3 | 4.3528 | 5.1 |  |  | 54
| 1 | 6 | 6 | 4.2 | 4 | 3.7 | 3.9947 | 4 | 5 | 5.8849 | 4.63 | 4 | 4 | 3.8 | 3.7 | 3.8 | 4 | 4.1 | 4.08 | 3 | 4.3 | < 5 | 4.9889 | 3 | 4.6678 | 4.7
| Mean |  | 5.6833 | 4.0667 | 4 | 3.85 | 3.9999 | 4 | 5.0667 | 5.8708 | 4.7783 | 4 | 4.0333 | 3.9333 | 3.7667 | 3.7833 | 4.2333 | 4.1167 | 4.1167 | 3.6167 | 4.2333 |  | 5.0252 | 3 | 4.3466 | 4.75
| Median |  | 5.65 | 4 | 4 | 3.85 | 3.997 | 4 | 5.05 | 5.8659 | 4.77 | 4 | 4 | 3.9 | 3.75 | 3.8 | 4.25 | 4.1 | 4.125 | 2.9 | 4.3 |  | 5.0347 | 3 | 4.2719 | 4.7
| Std Dev. |  | 0.19408 | 0.19664 | 0 | 0.18708 | 0.032796 | 0 | 0.08165 | 0.045162 | 0.11635 | 0 | 0.05164 | 0.15055 | 0.08165 | 0.1169 | 0.13663 | 0.040825 | 0.031411 | 1.7081 | 0.12111 |  | 0.14597 | 0 | 0.22764 | 0.23452
| Rel.Std.Dev. |  | 3.415% | 4.835% | 0% | 4.859% | 0.8199% | 0% | 1.612% | 0.7693% | 2.435% | 0% | 1.28% | 3.828% | 2.168% | 3.09% | 3.227% | 0.9917% | 0.763% | 47.23% | 2.861% |  | 2.905% | 0% | 5.237% | 4.937%
| PDM3 |  | 33.44% | -4.52% | -6.085% | -9.607% | -6.086% | -6.085% | 18.96% | 37.84% | 12.19% | -6.085% | -5.302% | -7.65% | -11.56% | -11.17% | -0.6065% | -3.346% | -3.346% | -15.09% | -0.6065% |  | 17.98% | -29.56% | 2.052% | 11.52%
| Z-Score (Absolute) |  | 3.52 | 0.11 | NA | 0.6 | 0.26 | NA | 2.14 | 3.94 | 1.49 | NA | 0.19 | 0.41 | 0.79 | 0.75 | 0.26 | 0 | 0 | 1.12 | 0.26 | 3.63 | 2.04 | NA | 0.52 | 1.42
Results from laboratories 2.04, 2.07, 2.11 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 606b (Certified Value 1.8 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2 | 1.9 | 1 | 1.6 | 2.5983 | 2.4 | 1.2 | 2.6598 | 1 | 1.2 | 2 | 2 | 1 | 1 | 1 | 1.32 | 1 | 3 | 3 | < 2 | < 10 | 2 |  |  |  |  | 1
| 1 | 2 | 2 | 1.4 | 1 | 0.9 | 2.6551 | 2.3 | 1.2 | 2.2361 | 1 | 1 | 2 | 2 | 1 | 2 | 1 | 1.28 | 1 | 3.2 | 3 | < 2 | < 10 | 2 |  |  |  |  | 38
| 1 | 3 | 2 | 1.4 | 1 | 1.4 | 2.4989 | 2.2 | 1.3 | 2.7894 | 1 | 1.3 | 2 | 2 | 2 | 1 | 1 | 1.28 | 2 | 2 | 3 | < 2 | < 10 | 2 |  |  |  |  | 16
| 1 | 4 | 2 | 2 | 1 | 1.4 | 2.3545 | 2.1 | 1.3 | 2.5589 | 1 | 1.5 | 2 | 2 | 2 | 2 | 1 | 1.27 | 1 | 2 | 3 | < 2 | < 10 | 2 |  |  |  |  | 1.7977
| 1 | 5 | 2 | 2.2 | 1 | 1.6 | 2.5827 | 2.2 | 1.2 | 2.7434 | 1 | 1.3 | 3 | 3 | 2 | 1 | 1 | 1.36 | 1 | 3.4 | 3 | < 2 | < 10 | 2 |  |  |  |  | 55
| 1 | 6 | 2 | 2.4 | 1 | 0.8 | 2.5991 | 2.2 | 1.2 | 2.5128 | 2 | 1.3 | 2 | 2 | 2 | 1 | 1 | 1.27 | 2 | 3 | 3 | < 2 | < 10 | 2
| Mean |  | 2 | 1.8833 | 1 | 1.2833 | 2.5481 | 2.2333 | 1.2333 | 2.5834 | 1.1667 | 1.2667 | 2.1667 | 2.1667 | 1.6667 | 1.3333 | 1 | 1.2967 | 1.3333 | 2.7667 | 3 |  |  | 2
| Median |  | 2 | 1.95 | 1 | 1.4 | 2.5905 | 2.2 | 1.2 | 2.6094 | 1 | 1.3 | 2 | 2 | 2 | 1 | 1 | 1.28 | 1 | 3 | 3 |  |  | 2
| Std Dev. |  | 0 | 0.41191 | 0 | 0.34881 | 0.10739 | 0.10328 | 0.05164 | 0.20004 | 0.40825 | 0.1633 | 0.40825 | 0.40825 | 0.5164 | 0.5164 | 0 | 0.036148 | 0.5164 | 0.6121 | 0 |  |  | 0
| Rel.Std.Dev. |  | 0% | 21.87% | 0% | 27.18% | 4.214% | 4.624% | 4.187% | 7.743% | 34.99% | 12.89% | 18.84% | 18.84% | 30.98% | 38.73% | 0% | 2.788% | 38.73% | 22.12% | 0% |  |  | 0%
| PDM3 |  | 11.25% | 4.763% | -44.37% | -28.61% | 41.74% | 24.23% | -31.39% | 43.7% | -35.1% | -29.54% | 20.52% | 20.52% | -7.29% | -25.83% | -44.37% | -27.87% | -25.83% | 53.9% | 66.88% |  |  | 11.25%
| Z-Score (Absolute) |  | 0.3 | 0.15 | 1.05 | 0.66 | 1.04 | 0.62 | 0.73 | 1.09 | 0.82 | 0.69 | 0.53 | 0.53 | 0.15 | 0.6 | 1.05 | 0.65 | 0.6 | 1.34 | 1.65 | 1.05 | 4.35 | 0.3
| Analytical results for Sm in OREAS 606b (Certified Value 7.27 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 7.1; 2.03: 7.11; 2.04: 6.873; 2.05: 7.6; 2.06: 8.111; 2.07: 7.4; 2.08: 5.8; 2.09: 7.29; 2.11: 6.6833; 2.12: 6; 2.13: 7.1 |  |  |  |  | 1
| 1 | 2 | 2.01: 7.4; 2.03: 7.08; 2.04: 6.973; 2.05: 7.5; 2.06: 8.0284; 2.07: 7.39; 2.08: 6.3; 2.09: 7.22; 2.11: 6.6647; 2.12: 6; 2.13: 6.6 |  |  |  |  | 22
| 1 | 3 | 2.01: 7.6; 2.03: 7; 2.04: 6.9122; 2.05: 7.5; 2.06: 8.069; 2.07: 7.18; 2.08: 5.5; 2.09: 7.2; 2.11: 6.6855; 2.12: 6; 2.13: 7.2 |  |  |  |  | 16
| 1 | 4 | 2.01: 7.5; 2.03: 7.34; 2.04: 6.925; 2.05: 7.6; 2.06: 8.0117; 2.07: 7.27; 2.08: 5; 2.09: 7.27; 2.11: 6.8368; 2.12: 6; 2.13: 6.9 |  |  |  |  | 7.2675
| 1 | 5 | 2.01: 8; 2.03: 6.94; 2.04: 6.7368; 2.05: 7.7; 2.06: 8.0246; 2.07: 7.14; 2.08: 5.5; 2.09: 7.19; 2.11: 7.044; 2.12: 6; 2.13: 7.4 |  |  |  |  | 56
| 1 | 6 | 2.01: 8.2; 2.03: 6.87; 2.04: 6.8247; 2.05: 7.5; 2.06: 8.0411; 2.07: 7.41; 2.08: 5.9; 2.09: 7.2; 2.11: 6.727; 2.12: 6; 2.13: 6.7
| Mean |  | 2.01: 7.6333; 2.03: 7.0567; 2.04: 6.8741; 2.05: 7.5667; 2.06: 8.0476; 2.07: 7.2983; 2.08: 5.6667; 2.09: 7.2283; 2.11: 6.7736; 2.12: 6; 2.13: 6.9833
| Median |  | 2.01: 7.55; 2.03: 7.04; 2.04: 6.8926; 2.05: 7.55; 2.06: 8.0348; 2.07: 7.33; 2.08: 5.65; 2.09: 7.21; 2.11: 6.7062; 2.12: 6; 2.13: 7
| Std Dev. |  | 2.01: 0.40332; 2.03: 0.16452; 2.04: 0.08379; 2.05: 0.08165; 2.06: 0.036623; 2.07: 0.11923; 2.08: 0.44121; 2.09: 0.041673; 2.11: 0.14633; 2.12: 0; 2.13: 0.30605
| Rel.Std.Dev. |  | 2.01: 5.284%; 2.03: 2.331%; 2.04: 1.219%; 2.05: 1.079%; 2.06: 0.4551%; 2.07: 1.634%; 2.08: 7.786%; 2.09: 0.5765%; 2.11: 2.16%; 2.12: 0%; 2.13: 4.383%
| PDM3 |  | 2.01: 5.033%; 2.03: -2.902%; 2.04: -5.413%; 2.05: 4.116%; 2.06: 10.73%; 2.07: 0.4237%; 2.08: -22.03%; 2.09: -0.5395%; 2.11: -6.797%; 2.12: -17.44%; 2.13: -3.911%
| Z-Score (Absolute) |  | 2.01: 1.04; 2.03: 0.18; 2.04: 0.57; 2.05: 0.9; 2.06: 1.92; 2.07: 0.33; 2.08: 3.12; 2.09: 0.18; 2.11: 0.78; 2.12: NA; 2.13: 0.34
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 606b (Certified Value 3.72 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 3.5 | 3.5 | 3.5 | 3.6 | 3.5709 | 4 | 3.9 | 3.83 | 3.7 | 4 | 3.7 | 3.6 | 3.6 | 3.5 | 3.6 | 3.77 | 3.7 | 3.1 | < 5 | 3.7 | 9 | 3.9711 | < 10 | 4 |  |  | 1
| 1 | 2 | 3.5 | 3.6 | 3.7 | 3.8 | 3.5578 | 4 | 4 | 3.93 | 3.8 | 4 | 3.8 | 3.6 | 3.6 | 3.6 | 3.7 | 3.76 | 3.8 | 3.3 | < 5 | 3.6 | 6 | 3.7914 | < 10 | 3.9 |  |  | 13
| 1 | 3 | 3.4 | 3.5 | 3.6 | 3.6 | 3.6199 | 4 | 4 | 3.9 | 3.8 | 4.1 | 3.8 | 3.7 | 3.7 | 3.6 | 3.7 | 3.82 | 3.5 | 3.2 | < 5 | 3.5 | 3 | 3.8236 | < 10 | 4 |  |  | 16
| 1 | 4 | 3.6 | 3.5 | 3.6 | 3.9 | 3.5575 | 4 | 3.8 | 3.7 | 3.8 | 3.9 | 3.9 | 3.8 | 3.5 | 3.5 | 3.8 | 3.75 | 3.7 | 3.2 | < 5 | 3.4 | 6 | 3.7834 | < 10 | 4 |  |  | 3.7162
| 1 | 5 | 3.5 | 3.8 | 3.6 | 3.6 | 3.3474 | 4 | 4 | 3.8 | 3.8 | 3.9 | 3.9 | 3.6 | 3.8 | 3.6 | 3.8 | 3.85 | 3.6 | 3.1 | < 5 | 3.5 | 4 | 3.766 | < 10 | 4.3 |  |  | 57
| 1 | 6 | 3.6 | 4.1 | 3.4 | 3.6 | 3.5273 | 4 | 3.9 | 3.83 | 3.8 | 3.9 | 3.8 | 3.7 | 3.6 | 3.6 | 3.6 | 3.71 | 3.8 | 3.3 | < 5 | 3.6 | 6 | 3.9862 | < 10 | 4.5
| Mean |  | 3.5167 | 3.6667 | 3.5667 | 3.6833 | 3.5301 | 4 | 3.9333 | 3.8317 | 3.7833 | 3.9667 | 3.8167 | 3.6667 | 3.6333 | 3.5667 | 3.7 | 3.7767 | 3.6833 | 3.2 |  | 3.55 | 5.6667 | 3.8536 |  | 4.1167
| Median |  | 3.5 | 3.55 | 3.6 | 3.6 | 3.5576 | 4 | 3.95 | 3.83 | 3.8 | 3.95 | 3.8 | 3.65 | 3.6 | 3.6 | 3.7 | 3.765 | 3.7 | 3.2 |  | 3.55 | 6 | 3.8075 |  | 4
| Std Dev. |  | 0.075277 | 0.24221 | 0.10328 | 0.13292 | 0.094471 | 0 | 0.08165 | 0.08085 | 0.040825 | 0.08165 | 0.075277 | 0.08165 | 0.10328 | 0.05164 | 0.089443 | 0.050465 | 0.1169 | 0.089443 |  | 0.10488 | 2.0656 | 0.098734 |  | 0.23166
| Rel.Std.Dev. |  | 2.141% | 6.606% | 2.896% | 3.609% | 2.676% | 0% | 2.076% | 2.11% | 1.079% | 2.058% | 1.972% | 2.227% | 2.843% | 1.448% | 2.417% | 1.336% | 3.174% | 2.795% |  | 2.954% | 36.45% | 2.562% |  | 5.627%
| PDM3 |  | -5.368% | -1.332% | -4.023% | -0.8832% | -5.006% | 7.638% | 5.844% | 3.108% | 1.808% | 6.741% | 2.705% | -1.332% | -2.229% | -4.023% | -0.4347% | 1.628% | -0.8832% | -13.89% |  | -4.471% | 52.49% | 3.699% |  | 10.78%
| Z-Score (Absolute) |  | 0.84 | 0.08 | 0.59 | 0 | 0.77 | NA | 1.26 | 0.75 | 0.51 | 1.43 | 0.67 | 0.08 | 0.25 | 0.59 | 0.08 | 0.47 | 0 | 2.44 | 5.98 | 0.67 | NA | 0.86 | 6.66 | 2.19
Results from laboratories 2.07 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 606b (Certified Value 200 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 210 | 177 | 193 | 199.37 | 203.2158 | 210 | 211 | 215.754 | 201 | 198 | 209 | 197.5 | 192 | 197.5 | 195 | 203 | 227 | 179 | 198 | 185.3 | 200.2 | 177.902 | 208 | 208.8192 | 171.73 | 177 | 1
| 1 | 2 | 211 | 184 | 203 | 199.85 | 201.4287 | 204 | 218 | 216.875 | 202 | 199 | 209 | 198 | 198.5 | 192.5 | 198 | 203 | 223 | 190 | 205 | 189.9 | 197.6 | 181.2554 | 202 | 206.1411 | 174.9 | 176 | 14
| 1 | 3 | 213 | 184 | 194 | 198.05 | 201.633 | 208 | 225 | 214.279 | 211 | 198 | 212 | 199.5 | 198.5 | 196 | 194 | 197.5 | 228 | 176 | 198 | 192.6 | 196.3 | 177.6628 | 198 | 200.3443 | 173.98 | 189 | 16
| 1 | 4 | 217 | 181 | 198 | 206.18 | 200.9759 | 208 | 216 | 219.349 | 203 | 197 | 212 | 201 | 199.5 | 190.5 | 193.5 | 205 | 228 | 165 | 205 | 187 | 198.3 | 183.7375 | 201 | 200.0463 | 169.94 | 187 | 199.5368
| 1 | 5 | 211 | 180 | 197 | 202.93 | 200.4915 | 204 | 213 | 216.578 | 206 | 197 | 208 | 204 | 201 | 196 | 193 | 203 | 230 | 175 | 201 | 187.4 | 196.3 | 180.4724 | 203 | 207.6571 | 173 | 189 | 58
| 1 | 6 | 210 | 206 | 191 | 195.55 | 203.7073 | 208 | 212 | 212.013 | 198 | 198 | 208 | 196 | 195.5 | 196.5 | 184 | 199.5 | 225 | 188 | 200 | 190.7 | 200.9 | 180.5918 | 206 | 211.3175 | 176.54 | 179
| Mean |  | 212 | 185.3333 | 196 | 200.3217 | 201.9087 | 207 | 215.8333 | 215.808 | 203.5 | 197.8333 | 209.6667 | 199.3333 | 197.5 | 194.8333 | 192.9167 | 201.8333 | 226.8333 | 178.8333 | 201.1667 | 188.8167 | 198.2667 | 180.2703 | 203 | 205.7209 | 173.3483 | 182.8333
| Median |  | 211 | 182.5 | 195.5 | 199.61 | 201.5308 | 208 | 214.5 | 216.166 | 202.5 | 198 | 209 | 198.75 | 198.5 | 196 | 193.75 | 203 | 227.5 | 177.5 | 200.5 | 188.65 | 197.95 | 180.5321 | 202.5 | 206.8991 | 173.49 | 183
| Std Dev. |  | 2.6833 | 10.4626 | 4.2895 | 3.7465 | 1.2751 | 2.4495 | 5.1929 | 2.4905 | 4.5056 | 0.75277 | 1.8619 | 2.8577 | 3.2404 | 2.7142 | 4.7161 | 2.7689 | 2.4833 | 9.196 | 3.1885 | 2.7096 | 1.9418 | 2.2603 | 3.5777 | 4.6034 | 2.3388 | 6.1455
| Rel.Std.Dev. |  | 1.266% | 5.645% | 2.189% | 1.87% | 0.6315% | 1.183% | 2.406% | 1.154% | 2.214% | 0.3805% | 0.888% | 1.434% | 1.641% | 1.393% | 2.445% | 1.372% | 1.095% | 5.142% | 1.585% | 1.435% | 0.9794% | 1.254% | 1.762% | 2.238% | 1.349% | 3.361%
| PDM3 |  | 6.246% | -7.118% | -1.773% | 0.3933% | 1.189% | 3.74% | 8.167% | 8.154% | 1.986% | -0.8537% | 5.077% | -0.102% | -1.021% | -2.357% | -3.318% | 1.151% | 13.68% | -10.38% | 0.8168% | -5.373% | -0.6365% | -9.656% | 1.736% | 3.099% | -13.12% | -8.371%
| Z-Score (Absolute) |  | 1.28 | 1.53 | 0.4 | 0.05 | 0.22 | 0.76 | 1.69 | 1.68 | 0.39 | 0.21 | 1.04 | 0.05 | 0.25 | 0.53 | 0.73 | 0.21 | 2.84 | 2.21 | 0.14 | 1.16 | 0.16 | 2.06 | 0.33 | 0.62 | 2.79 | 1.79
| Analytical results for Ta in OREAS 606b (Certified Value 1.36 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 1.21 | 1.6 | 1.3 | 1.41 | 1.4255 | 1.6 | 1.28 | 1.61 | 1.39 | 1.42 | 1.38 | 1.34 | 1.36 | 1.34 | 1.36 | 1.85 | 1.73 | 0.45 | 0.94 | < 1 | 1.3847 | 1.3 |  |  |  |  | 1
| 1 | 2 | 1.12 | 1.6 | 1.3 | 1.38 | 1.4705 | 1.6 | 1.32 | 1.65 | 1.38 | 1.43 | 1.41 | 1.4 | 1.32 | 1.39 | 1.38 | 1.88 | 1.82 | 1.26 | 0.95 | 1 | 1.3987 | 1.23 |  |  |  |  | 25
| 1 | 3 | 1.29 | 1.6 | 1.3 | 1.36 | 1.4485 | 1.5 | 1.32 | 1.59 | 1.37 | 1.44 | 1.4 | 1.36 | 1.38 | 1.38 | 1.33 | 1.92 | 1.79 | 1.35 | 0.93 | < 1 | 1.3532 | 1.18 |  |  |  |  | 16
| 1 | 4 | 1.26 | 1.6 | 1.3 | 1.46 | 1.4521 | 1.6 | 1.33 | 1.64 | 1.37 | 1.46 | 1.41 | 1.39 | 1.3 | 1.33 | 1.41 | 1.9 | 1.85 | 1.41 | 0.96 | < 1 | 1.3364 | 1.18 |  |  |  |  | 1.3643
| 1 | 5 | 1.19 | 1.9 | 1.4 | 1.38 | 1.4763 | 1.5 | 1.31 | 1.62 | 1.36 | 1.43 | 1.47 | 1.41 | 1.38 | 1.35 | 1.37 | 1.85 | 1.78 | 1.18 | 0.92 | 2 | 1.3652 | 1.25 |  |  |  |  | 59
| 1 | 6 | 1.33 | 1.7 | 1.3 | 1.41 | 1.4252 | 1.5 | 1.3 | 1.6 | 1.38 | 1.43 | 1.37 | 1.34 | 1.35 | 1.34 | 1.36 | 1.9 | 1.8 | 1.36 | 0.99 | 2 | 1.4147 | 1.18
| Mean |  | 1.2333 | 1.6667 | 1.3167 | 1.4 | 1.4497 | 1.55 | 1.31 | 1.6183 | 1.375 | 1.435 | 1.4067 | 1.3733 | 1.3483 | 1.355 | 1.3683 | 1.8833 | 1.795 | 1.1683 | 0.94833 | 1.6667 | 1.3755 | 1.22
| Median |  | 1.235 | 1.6 | 1.3 | 1.395 | 1.4503 | 1.55 | 1.315 | 1.615 | 1.375 | 1.43 | 1.405 | 1.375 | 1.355 | 1.345 | 1.365 | 1.89 | 1.795 | 1.305 | 0.945 | 2 | 1.3749 | 1.205
| Std Dev. |  | 0.075542 | 0.12111 | 0.040825 | 0.035214 | 0.021587 | 0.054772 | 0.017889 | 0.023166 | 0.010488 | 0.013784 | 0.035024 | 0.030768 | 0.032506 | 0.02429 | 0.026394 | 0.028752 | 0.040373 | 0.3613 | 0.024833 | 0.57735 | 0.0293 | 0.049396
| Rel.Std.Dev. |  | 6.125% | 7.266% | 3.101% | 2.515% | 1.489% | 3.534% | 1.366% | 1.431% | 0.7628% | 0.9606% | 2.49% | 2.24% | 2.411% | 1.793% | 1.929% | 1.527% | 2.249% | 30.92% | 2.619% | 34.64% | 2.13% | 4.049%
| PDM3 |  | -9.6% | 22.16% | -3.492% | 2.617% | 6.259% | 13.61% | -3.98% | 18.62% | 0.7841% | 5.182% | 3.105% | 0.662% | -1.17% | -0.6818% | 0.2955% | 38.04% | 31.57% | -14.36% | -30.49% | 22.16% | 0.8189% | -10.58%
| Z-Score (Absolute) |  | 1.36 | 2.82 | 0.56 | 0.25 | 0.73 | 1.7 | 0.62 | 2.36 | 0.01 | 0.59 | 0.31 | 0.01 | 0.25 | 0.19 | 0.06 | 4.92 | 4.06 | 1.99 | 4.11 | 2.81 | 0.01 | 1.49
| Analytical results for Tb in OREAS 606b (Certified Value 0.73 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.10; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 0.7; 2.03: 0.72; 2.04: 0.77774; 2.05: 0.9; 2.06: 1.0101; 2.07: 0.72; 2.08: 0.69; 2.10: 0.81; 2.11: 0.7506; 2.12: < 1; 2.13: 0.71 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.7; 2.03: 0.68; 2.04: 0.77525; 2.05: 0.8; 2.06: 1.0816; 2.07: 0.74; 2.08: 0.74; 2.10: 0.78; 2.11: 0.7489; 2.12: < 1; 2.13: 0.67 |  |  |  |  | 26
| 1 | 3 | 2.01: 0.7; 2.03: 0.69; 2.04: 0.77757; 2.05: 0.8; 2.06: 1.0235; 2.07: 0.74; 2.08: 0.67; 2.10: 0.79; 2.11: 0.7483; 2.12: < 1; 2.13: 0.67 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.7; 2.03: 0.72; 2.04: 0.77909; 2.05: 0.9; 2.06: 1.039; 2.07: 0.71; 2.08: 0.63; 2.10: 0.8; 2.11: 0.7797; 2.12: < 1; 2.13: 0.66 |  |  |  |  | 0.73014
| 1 | 5 | 2.01: 0.8; 2.03: 0.68; 2.04: 0.76439; 2.05: 0.9; 2.06: 1.053; 2.07: 0.76; 2.08: 0.67; 2.10: 0.8; 2.11: 0.7957; 2.12: < 1; 2.13: 0.67 |  |  |  |  | 60
| 1 | 6 | 2.01: 0.8; 2.03: 0.68; 2.04: 0.76567; 2.05: 0.9; 2.06: 1.0116; 2.07: 0.7; 2.08: 0.73; 2.10: 0.8; 2.11: 0.7571; 2.12: < 1; 2.13: 0.66
| Mean |  | 2.01: 0.73333; 2.03: 0.695; 2.04: 0.77329; 2.05: 0.86667; 2.06: 1.0365; 2.07: 0.72833; 2.08: 0.68833; 2.10: 0.79667; 2.11: 0.76338; 2.13: 0.67333
| Median |  | 2.01: 0.7; 2.03: 0.685; 2.04: 0.77641; 2.05: 0.9; 2.06: 1.0312; 2.07: 0.73; 2.08: 0.68; 2.10: 0.8; 2.11: 0.75385; 2.13: 0.67
| Std Dev. |  | 2.01: 0.05164; 2.03: 0.019748; 2.04: 0.0065246; 2.05: 0.05164; 2.06: 0.027563; 2.07: 0.022286; 2.08: 0.041191; 2.10: 0.010328; 2.11: 0.019752; 2.13: 0.018619
| Rel.Std.Dev. |  | 2.01: 7.042%; 2.03: 2.841%; 2.04: 0.8438%; 2.05: 5.958%; 2.06: 2.659%; 2.07: 3.06%; 2.08: 5.984%; 2.10: 1.296%; 2.11: 2.587%; 2.13: 2.765%
| PDM3 |  | 2.01: 0.4369%; 2.03: -4.813%; 2.04: 5.909%; 2.05: 18.7%; 2.06: 41.95%; 2.07: -0.2479%; 2.08: -5.726%; 2.10: 9.111%; 2.11: 4.553%; 2.13: -7.781%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.5; 2.04: 0.67; 2.05: NA; 2.06: 4.62; 2.07: 0; 2.08: 0.6; 2.10: 1.03; 2.11: 0.53; 2.12: 3.43; 2.13: 0.83
Results from laboratories 2.03 and 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 606b (Certified Value 0.84 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.27 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.11 | 0.79 | 1 | 0.8 | 0.82209 | 0.8 | 0.89 | 0.8 | 0.9 | 0.91 | 0.76 | 0.85 | 0.8 | 0.76 | 0.96 | 1.51 | 0.92 | 0.89 | 1.033 | < 10 | 0.84 | [5 empty] | 1
| 1 | 2 | 1.99 | 0.77 | 1 | 0.9 | 0.86561 | 0.9 | 0.95 | 0.8 | 0.9 | 0.87 | 0.79 | 0.85 | 0.82 | 0.77 | 0.84 | 2.3 | 0.67 | 0.8 | 0.88901 | < 10 | 0.89 | [5 empty] | 27
| 1 | 3 | 2.01 | 0.77 | 0.8 | 0.9 | 0.85369 | 0.8 | 0.87 | 0.8 | 0.9 | 0.84 | 0.75 | 0.85 | 0.82 | 0.83 | 0.79 | 2.94 | 0.64 | 0.82 | 0.81442 | < 10 | 0.8 | [5 empty] | 16
| 1 | 4 | 2.09 | 0.71 | 0.9 | 0.8 | 0.89132 | 0.9 | 0.9 | 0.8 | 0.8 | 0.84 | 0.79 | 0.79 | 0.75 | 0.82 | 0.81 | 3.44 | 1.09 | 0.76 | 0.88056 | < 10 | 0.86 | [5 empty] | 0.83512
| 1 | 5 | 1.89 | 0.86 | 0.9 | 0.8 | 0.81556 | 0.7 | 0.94 | 0.8 | 0.7 | 0.9 | 0.83 | 0.81 | 0.86 | 0.87 | 0.85 | 3.37 | 0.833 | 0.83 | 0.96741 | < 10 | 0.82 | [5 empty] | 61
| 1 | 6 | 1.99 | 0.66 | 0.7 | 0.8 | 0.8259 | 0.9 | 0.83 | 0.8 | 0.9 | 0.78 | 0.75 | 0.8 | 0.77 | 0.79 | 0.9 | 2.57 | 1.28 | 0.85 | 0.97707 | < 10 | 0.85
| Mean |  | 2.0133 | 0.76 | 0.88333 | 0.83333 | 0.84569 | 0.83333 | 0.89667 | 0.8 | 0.85 | 0.85667 | 0.77833 | 0.825 | 0.80333 | 0.80667 | 0.85833 | 2.6883 | 0.9055 | 0.825 | 0.9269 |  | 0.84333
| Median |  | 2 | 0.77 | 0.9 | 0.8 | 0.83979 | 0.85 | 0.895 | 0.8 | 0.9 | 0.855 | 0.775 | 0.83 | 0.81 | 0.805 | 0.845 | 2.755 | 0.8765 | 0.825 | 0.92821 |  | 0.845
| Std Dev. |  | 0.079415 | 0.068702 | 0.1169 | 0.05164 | 0.029662 | 0.08165 | 0.044572 | 1.2162e-16 | 0.083666 | 0.04761 | 0.031252 | 0.028107 | 0.039328 | 0.041312 | 0.062423 | 0.7275 | 0.24727 | 0.044159 | 0.07955 |  | 0.031411
| Rel.Std.Dev. |  | 3.944% | 9.04% | 13.23% | 6.197% | 3.507% | 9.798% | 4.971% | 1.52e-14% | 9.843% | 5.558% | 4.015% | 3.407% | 4.896% | 5.121% | 7.273% | 27.06% | 27.31% | 5.353% | 8.582% |  | 3.725%
| PDM3 |  | 141.1% | -8.995% | 5.774% | -0.2136% | 1.266% | -0.2136% | 7.37% | -4.205% | 1.782% | 2.58% | -6.8% | -1.212% | -3.806% | -3.407% | 2.78% | 221.9% | 8.428% | -1.212% | 10.99% |  | 0.9838%
| Z-Score (Absolute) |  | 16.79 | 1.32 | NA | NA | 0.08 | NA | 0.66 | NA | NA | 0.08 | 1.05 | 0.38 | 0.69 | 0.64 | 0.1 | 26.55 | 0.78 | 0.38 | 1.09 | 59.95 | 0.11
Results from laboratories 2.04, 2.05, 2.07, 2.11 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 606b (Certified Value 14.8 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.98 | 13.8 | 14 | 14.56 | 14.1025 | 15 | 15.2 | 17.13 | 15.65 | 14.7 | 15.12 | 14.1 | 13.9 | 14.7 | 15.1 | 15.8 | 14.31 | 11.26 | < 20 | 13.76 | 15.3484 | 14 | 14.927 | 12.8 |  |  | 1
| 1 | 2 | 14.32 | 14.3 | 14.4 | 14.78 | 14.6115 | 14.9 | 15.8 | 17.575 | 15.45 | 14.8 | 15.2 | 15 | 14.45 | 13.85 | 15.8 | 16.2 | 14.24 | 12.27 | < 20 | 13.04 | 15.2563 | 14 | 14.8686 | 12.6 |  |  | 28
| 1 | 3 | 15.52 | 14.4 | 14.2 | 14.64 | 14.3389 | 14.4 | 15.7 | 17.379 | 15.36 | 14.8 | 15.01 | 14.5 | 14.2 | 14.45 | 15.25 | 15.4 | 14.49 | 10.86 | < 20 | 13.4 | 15.7451 | 14 | 14.56 | 12.72 |  |  | 16
| 1 | 4 | 16.61 | 13.9 | 14.5 | 14.62 | 14.2195 | 14.6 | 15.1 | 17.323 | 14.79 | 14.6 | 15.26 | 14.5 | 14.35 | 14 | 15.5 | 16.9 | 14.56 | 10.2 | < 20 | 14.07 | 15.5506 | 13 | 14.3896 | 12.81 |  |  | 14.794
| 1 | 5 | 15.21 | 15.5 | 14.2 | 14.63 | 14.0811 | 14.4 | 15.8 | 17.26 | 15.99 | 14.8 | 15.3 | 14.55 | 14.55 | 14.4 | 15.3 | 16.35 | 14.57 | 11.15 | < 20 | 12.94 | 16.0625 | 13 | 14.4784 | 12.4 |  |  | 62
| 1 | 6 | 15.85 | 15 | 13.9 | 14.91 | 14.1426 | 14.7 | 15.4 | 17.982 | 15 | 14.9 | 15.27 | 14.55 | 14.35 | 14.2 | 13.25 | 15.7 | 14.31 | 12.02 | < 20 | 13.1 | 15.4579 | 13 | 14.9965 | 12.27
| Mean |  | 15.415 | 14.4833 | 14.2 | 14.69 | 14.2493 | 14.6667 | 15.5 | 17.4415 | 15.3733 | 14.7667 | 15.1933 | 14.5333 | 14.3 | 14.2667 | 15.0333 | 16.0583 | 14.4133 | 11.2933 |  | 13.385 | 15.5701 | 13.5 | 14.7033 | 12.6
| Median |  | 15.365 | 14.35 | 14.2 | 14.635 | 14.1811 | 14.65 | 15.55 | 17.351 | 15.405 | 14.8 | 15.23 | 14.525 | 14.35 | 14.3 | 15.275 | 16 | 14.4 | 11.205 |  | 13.25 | 15.5043 | 13.5 | 14.7143 | 12.66
| Std Dev. |  | 0.78212 | 0.65549 | 0.22804 | 0.12992 | 0.20066 | 0.25033 | 0.30984 | 0.30262 | 0.43417 | 0.10328 | 0.11021 | 0.28577 | 0.22804 | 0.31252 | 0.90646 | 0.53704 | 0.14376 | 0.75983 |  | 0.44908 | 0.29451 | 0.54772 | 0.25801 | 0.22244
| Rel.Std.Dev. |  | 5.074% | 4.526% | 1.606% | 0.8844% | 1.408% | 1.707% | 1.999% | 1.735% | 2.824% | 0.6994% | 0.7254% | 1.966% | 1.595% | 2.191% | 6.03% | 3.344% | 0.9974% | 6.728% |  | 3.355% | 1.891% | 4.057% | 1.755% | 1.765%
| PDM3 |  | 4.198% | -2.1% | -4.015% | -0.7028% | -3.681% | -0.8605% | 4.772% | 17.9% | 3.916% | -0.1845% | 2.7% | -1.762% | -3.339% | -3.564% | 1.618% | 8.546% | -2.573% | -23.66% |  | -9.524% | 5.247% | -8.747% | -0.6126% | -14.83%
| Z-Score (Absolute) |  | 1.08 | 0.26 | 0.67 | 0.03 | 0.6 | 0 | 1.2 | 4.01 | 1.02 | 0.14 | 0.76 | 0.19 | 0.53 | 0.58 | 0.53 | 2.01 | 0.37 | 4.87 | 6.74 | 1.85 | 1.31 | NA | 0.05 | 2.99
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 606b (Certified Value 0.116 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.12 | 0.12 | 0.115 | 0.1148 | 0.1172 | 0.119 | 0.11 | 0.1148 | 0.12 | 0.117 | 0.1139 | 0.112 | 0.112 | 0.111 | 0.106 | 0.12 | 0.126 | 0.115 | 0.123 | 0.11 | 0.119 | < 0.1 | 0.114 | 0.11492 | 0.121 |  | 1
| 1 | 2 | 0.12 | 0.117 | 0.123 | 0.1129 | 0.11894 | 0.114 | 0.11 | 0.11402 | 0.127 | 0.119 | 0.1141 | 0.113 | 0.115 | 0.11 | 0.109 | 0.118 | 0.1285 | 0.118 | 0.127 | 0.11 | 0.117 | < 0.1 | 0.114 | 0.11575 | 0.124 |  | 29
| 1 | 3 | 0.12 | 0.122 | 0.122 | 0.1133 | 0.11902 | 0.118 | 0.11 | 0.1149 | 0.127 | 0.119 | 0.1158 | 0.113 | 0.113 | 0.11 | 0.105 | 0.116 | 0.1312 | 0.115 | 0.115 | 0.11 | 0.117 | < 0.1 | 0.114 | 0.11475 | 0.124 |  | 16
| 1 | 4 | 0.13 | 0.125 | 0.116 | 0.1172 | 0.11615 | 0.115 | 0.11 | 0.1144 | 0.117 | 0.121 | 0.1147 | 0.115 | 0.114 | 0.111 | 0.106 | 0.12 | 0.1305 | 0.115 | 0.125 | 0.11 | 0.117 | < 0.1 | 0.114 | 0.11532 | 0.121 |  | 0.11614
| 1 | 5 | 0.12 | 0.12 | 0.116 | 0.1153 | 0.118 | 0.115 | 0.11 | 0.11422 | 0.127 | 0.119 | 0.1134 | 0.116 | 0.116 | 0.11 | 0.105 | 0.12 | 0.1287 | 0.116 | 0.127 | 0.11 | 0.116 | < 0.1 | 0.114 | 0.11629 | 0.125 |  | 63
| 1 | 6 | 0.12 | 0.124 | 0.118 | 0.111 | 0.11978 | 0.115 | 0.11 | 0.1126 | 0.117 | 0.119 | 0.114 | 0.111 | 0.114 | 0.109 | 0.103 | 0.117 | 0.1298 | 0.115 | 0.127 | 0.11 | 0.119 | < 0.1 | 0.114 | 0.11572 | 0.12
| Mean |  | 0.12167 | 0.12133 | 0.11833 | 0.11408 | 0.11818 | 0.116 | 0.11 | 0.11416 | 0.1225 | 0.119 | 0.11432 | 0.11333 | 0.114 | 0.11017 | 0.10567 | 0.1185 | 0.12912 | 0.11567 | 0.124 | 0.11 | 0.1175 |  | 0.114 | 0.11546 | 0.1225
| Median |  | 0.12 | 0.121 | 0.117 | 0.11405 | 0.11847 | 0.115 | 0.11 | 0.11431 | 0.1235 | 0.119 | 0.11405 | 0.113 | 0.114 | 0.11 | 0.1055 | 0.119 | 0.12925 | 0.115 | 0.126 | 0.11 | 0.117 |  | 0.114 | 0.11552 | 0.1225
| Std Dev. |  | 0.0040825 | 0.0029439 | 0.0033862 | 0.0021554 | 0.0013387 | 0.002 | 0 | 0.00083397 | 0.0050498 | 0.0012649 | 0.00083766 | 0.0018619 | 0.0014142 | 0.00075277 | 0.0019664 | 0.0017607 | 0.0018433 | 0.0012111 | 0.0046904 | 0 | 0.0012247 |  | 0 | 0.00057449 | 0.0020736
| Rel.Std.Dev. |  | 3.355% | 2.426% | 2.862% | 1.889% | 1.133% | 1.724% | 0% | 0.7306% | 4.122% | 1.063% | 0.7328% | 1.643% | 1.241% | 0.6833% | 1.861% | 1.486% | 1.428% | 1.047% | 3.783% | 0% | 1.042% |  | 0% | 0.4976% | 1.693%
| PDM3 |  | 4.759% | 4.472% | 1.889% | -1.77% | 1.757% | -0.1201% | -5.286% | -1.707% | 5.477% | 2.463% | -1.569% | -2.416% | -1.842% | -5.143% | -9.017% | 2.033% | 11.17% | -0.4071% | 6.768% | -5.286% | 1.171% |  | -1.842% | -0.5886% | 5.477%
| Z-Score (Absolute) |  | 1.43 | 1.35 | 0.63 | 0.38 | 0.6 | 0.08 | 1.35 | 0.36 | 1.63 | 0.79 | 0.32 | 0.56 | 0.4 | 1.31 | 2.38 | 0.67 | 3.2 | 0 | 1.98 | 1.35 | 0.44 | 15.63 | 0.42 | 0.05 | 1.63
| Analytical results for Tl in OREAS 606b (Certified Value 0.89 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 0.55 | 0.82 | 0.9 | 0.87 | 0.89899 | 0.88 | 0.9 | 0.96 | 0.89 | 0.94 | 0.88 | 0.79 | 0.87 | 0.87 | 0.86 | 0.858 | 1.01 | 0.94 | 0.65 | 0.88062 | 0.76 | [5 empty] | 1
| 1 | 2 | 0.48 | 0.88 | 1 | 0.87 | 0.90067 | 0.88 | 0.91 | 0.96 | 0.86 | 0.96 | 0.92 | 0.8 | 0.84 | 0.9 | 0.84 | 0.858 | 1.02 | 0.99 | 0.7 | 0.85252 | 0.76 | [5 empty] | 30
| 1 | 3 | 0.5 | 0.81 | 0.9 | 0.91 | 0.89326 | 0.87 | 0.91 | 0.94 | 0.86 | 0.97 | 0.9 | 0.8 | 0.87 | 0.89 | 0.82 | 0.857 | 0.99 | 0.98 | 0.7 | 0.83541 | 0.71 | [5 empty] | 16
| 1 | 4 | 0.59 | 0.8 | 0.9 | 0.9 | 0.8986 | 0.88 | 0.89 | 0.9 | 0.86 | 0.93 | 0.9 | 0.83 | 0.83 | 0.9 | 0.88 | 0.857 | 1 | 0.96 | 0.66 | 0.83904 | 0.71 | [5 empty] | 0.88616
| 1 | 5 | 0.51 | 0.88 | 0.9 | 0.9 | 0.89418 | 0.86 | 0.91 | 0.93 | 0.88 | 0.92 | 0.92 | 0.79 | 0.87 | 0.9 | 0.87 | 0.866 | 0.99 | 0.95 | 0.71 | 0.8926 | 0.71 | [5 empty] | 64
| 1 | 6 | 0.57 | 0.91 | 0.9 | 0.88 | 0.89744 | 0.87 | 0.92 | 0.92 | 0.87 | 0.95 | 0.86 | 0.81 | 0.85 | 0.89 | 0.86 | 0.852 | 0.98 | 0.95 | 0.69 | 0.86668 | 0.71
| Mean |  | 0.53333 | 0.85 | 0.91667 | 0.88833 | 0.89719 | 0.87333 | 0.90667 | 0.935 | 0.87 | 0.945 | 0.89667 | 0.80333 | 0.855 | 0.89167 | 0.855 | 0.858 | 0.99833 | 0.96167 | 0.685 | 0.86114 | 0.72667
| Median |  | 0.53 | 0.85 | 0.9 | 0.89 | 0.89802 | 0.875 | 0.91 | 0.935 | 0.865 | 0.945 | 0.9 | 0.8 | 0.86 | 0.895 | 0.86 | 0.8575 | 0.995 | 0.955 | 0.695 | 0.8596 | 0.71
| Std Dev. |  | 0.043205 | 0.045607 | 0.040825 | 0.017224 | 0.0028952 | 0.008165 | 0.010328 | 0.023452 | 0.012649 | 0.018708 | 0.023381 | 0.015055 | 0.017607 | 0.01169 | 0.021679 | 0.0045166 | 0.01472 | 0.019408 | 0.02429 | 0.022911 | 0.02582
| Rel.Std.Dev. |  | 8.101% | 5.366% | 4.454% | 1.939% | 0.3227% | 0.9349% | 1.139% | 2.508% | 1.454% | 1.98% | 2.608% | 1.874% | 2.059% | 1.311% | 2.536% | 0.5264% | 1.474% | 2.018% | 3.546% | 2.66% | 3.553%
| PDM3 |  | -39.82% | -4.08% | 3.443% | 0.2456% | 1.245% | -1.447% | 2.314% | 5.512% | -1.823% | 6.64% | 1.186% | -9.346% | -3.516% | 0.6218% | -3.516% | -3.177% | 12.66% | 8.521% | -22.7% | -2.823% | -18%
| Z-Score (Absolute) |  | 9.61 | 0.66 | 1.22 | 0.42 | 0.67 | 0 | 0.94 | 1.74 | 0.09 | 2.03 | 0.66 | 1.98 | 0.52 | 0.52 | 0.52 | 0.43 | 3.53 | 2.5 | 5.32 | 0.34 | 4.15
| Analytical results for Tm in OREAS 606b (Certified Value 0.1 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | [5 empty] | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.28 | [5 empty] | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.10288; 2.05: 0.1; 2.06: 0.1201; 2.07: 0.1; 2.08: 0.1; 2.10: 0.1; 2.11: 0.1058; 2.12: 0.13 | [5 empty] | 1
| 1 | 2 | 2.01: < 0.1; 2.03: 0.1; 2.04: 0.10504; 2.05: 0.1; 2.06: 0.1117; 2.07: 0.11; 2.08: 0.11; 2.10: 0.09; 2.11: 0.1062; 2.12: 0.12 | [5 empty] | 5
| 1 | 3 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.1068; 2.05: 0.1; 2.06: 0.1175; 2.07: 0.11; 2.08: 0.1; 2.10: 0.1; 2.11: 0.1054; 2.12: 0.09 | [5 empty] | 16
| 1 | 4 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.10446; 2.05: < 0.1; 2.06: 0.1071; 2.07: 0.11; 2.08: 0.1; 2.10: 0.1; 2.11: 0.1053; 2.12: 0.1 | [5 empty] | 0.10423
| 1 | 5 | 2.01: 0.1; 2.03: 0.11; 2.04: 0.10713; 2.05: 0.1; 2.06: 0.1172; 2.07: 0.11; 2.08: 0.09; 2.10: 0.1; 2.11: 0.1119; 2.12: 0.09 | [5 empty] | 65
| 1 | 6 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.10203; 2.05: 0.1; 2.06: 0.1155; 2.07: 0.11; 2.08: 0.1; 2.10: 0.1; 2.11: 0.1057; 2.12: 0.11
| Mean |  | 2.01: 0.1; 2.03: 0.10833; 2.04: 0.10472; 2.05: 0.1; 2.06: 0.11485; 2.07: 0.10833; 2.08: 0.1; 2.10: 0.098333; 2.11: 0.10672; 2.12: 0.10667
| Median |  | 2.01: 0.1; 2.03: 0.11; 2.04: 0.10475; 2.05: 0.1; 2.06: 0.11635; 2.07: 0.11; 2.08: 0.1; 2.10: 0.1; 2.11: 0.10575; 2.12: 0.105
| Std Dev. |  | 2.03: 0.0040825; 2.04: 0.0020454; 2.05: 0; 2.06: 0.0047014; 2.07: 0.0040825; 2.08: 0.0063246; 2.10: 0.0040825; 2.11: 0.0025592; 2.12: 0.01633
| Rel.Std.Dev. |  | 2.03: 3.768%; 2.04: 1.953%; 2.05: 0%; 2.06: 4.093%; 2.07: 3.768%; 2.08: 6.325%; 2.10: 4.152%; 2.11: 2.398%; 2.12: 15.31%
| PDM3 |  | 2.01: -4.054%; 2.03: 3.941%; 2.04: 0.478%; 2.05: -4.054%; 2.06: 10.19%; 2.07: 3.941%; 2.08: -4.054%; 2.10: -5.653%; 2.11: 2.39%; 2.12: 2.342%
| Z-Score (Absolute) |  | 2.01: 7.66; 2.03: 0.43; 2.04: 0.16; 2.05: 2.27; 2.06: 1.48; 2.07: 0.43; 2.08: 0.92; 2.10: 1.19; 2.11: 0.17; 2.12: 0.16
| Analytical results for U in OREAS 606b (Certified Value 5.57 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 5.4 | 5.1 | 5.8 | 5.67 | 5.7011 | 5.8 | 5.6 | 6.5491 | 5.95 | 5.74 | 5.74 | 5.3 | 5.3 | 5.7 | 5.4 | 5.6 | 5.5 | 5.27 | 6.06 | 5.3 | 5.0716 | 2 | 5.7854 | 4.64 |  |  | 1
| 1 | 2 | 4.8 | 5.3 | 5.9 | 5.77 | 5.9196 | 5.7 | 5.72 | 6.4047 | 5.66 | 5.59 | 5.8 | 5.6 | 5.6 | 5.4 | 5.7 | 5.6 | 5.49 | 5.74 | 6.29 | 5.2 | 5.0602 | 2 | 5.8398 | 4.58 |  |  | 32
| 1 | 3 | 5.1 | 5.2 | 5.7 | 5.64 | 5.8591 | 5.6 | 5.6 | 6.412 | 5.9 | 5.72 | 5.75 | 5.4 | 5.4 | 5.7 | 5.6 | 5.4 | 5.6 | 4.67 | 5.99 | 5.2 | 5.0432 | 2 | 5.4695 | 4.6 |  |  | 16
| 1 | 4 | 5.6 | 5.1 | 5.8 | 5.75 | 5.5551 | 5.6 | 5.54 | 6.498 | 5.69 | 5.59 | 5.86 | 5.5 | 5.5 | 5.5 | 5.6 | 5.8 | 5.59 | 4.23 | 6.26 | 5.2 | 5.1939 | 2 | 5.5129 | 4.55 |  |  | 5.5657
| 1 | 5 | 4.9 | 5.5 | 5.9 | 5.69 | 5.6284 | 5.5 | 5.66 | 6.8366 | 5.93 | 5.79 | 5.9 | 5.4 | 5.6 | 5.9 | 5.5 | 5.6 | 5.6 | 4.78 | 6.18 | 5.2 | 5.2656 | 3 | 5.817 | 4.54 |  |  | 66
| 1 | 6 | 5.2 | 5.7 | 5.7 | 5.76 | 5.5043 | 5.7 | 5.59 | 6.5461 | 5.62 | 5.68 | 5.79 | 5.4 | 5.4 | 5.7 | 5.4 | 5.6 | 5.52 | 5.45 | 6.17 | 5.2 | 5.1129 | < 2 | 5.9085 | 4.54
| Mean |  | 5.1667 | 5.3167 | 5.8 | 5.7133 | 5.6946 | 5.65 | 5.6183 | 6.5411 | 5.7917 | 5.685 | 5.8067 | 5.4333 | 5.4667 | 5.65 | 5.5333 | 5.6 | 5.55 | 5.0233 | 6.1583 | 5.2167 | 5.1246 | 2.2 | 5.7222 | 4.575
| Median |  | 5.15 | 5.25 | 5.8 | 5.72 | 5.6647 | 5.65 | 5.6 | 6.5221 | 5.795 | 5.7 | 5.795 | 5.4 | 5.45 | 5.7 | 5.55 | 5.6 | 5.555 | 5.025 | 6.175 | 5.2 | 5.0922 | 2 | 5.8012 | 4.565
| Std Dev. |  | 0.30111 | 0.24014 | 0.089443 | 0.053914 | 0.16603 | 0.10488 | 0.062743 | 0.15782 | 0.15039 | 0.08167 | 0.062503 | 0.10328 | 0.12111 | 0.17607 | 0.12111 | 0.12649 | 0.052154 | 0.56035 | 0.11514 | 0.040825 | 0.087639 | 0.44721 | 0.18396 | 0.039875
| Rel.Std.Dev. |  | 5.828% | 4.517% | 1.542% | 0.9436% | 2.916% | 1.856% | 1.117% | 2.413% | 2.597% | 1.437% | 1.076% | 1.901% | 2.215% | 3.116% | 2.189% | 2.259% | 0.9397% | 11.15% | 1.87% | 0.7826% | 1.71% | 20.33% | 3.215% | 0.8716%
| PDM3 |  | -7.17% | -4.475% | 4.209% | 2.652% | 2.315% | 1.514% | 0.9454% | 17.52% | 4.06% | 2.143% | 4.329% | -2.379% | -1.78% | 1.514% | -0.5818% | 0.616% | -0.2823% | -9.745% | 10.65% | -6.271% | -7.926% | -60.47% | 2.811% | -17.8%
| Z-Score (Absolute) |  | 1.76 | 1.17 | 0.71 | 0.37 | 0.3 | 0.12 | 0 | 3.59 | 0.67 | 0.26 | 0.73 | 0.72 | 0.59 | 0.12 | 0.33 | 0.07 | 0.27 | 2.31 | 2.1 | 1.56 | 1.92 | NA | 0.4 | 4.06
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 606b (Certified Value 4.97 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 2
2
| 1 | 1 | 5 | 5 | 5 | 5 | 4.8332 | 5 | 4.6 | 4.57 | 6 | 5 | 5 | 4 | 5 | 5 | 5 | 5 | 6 | 5.7 | 10.98 | 4 | < 0.1 | 4 | 6.0903 | < 10 | 6 |  | 1
| 1 | 2 | 5 | 5 | 5 | 5 | 4.8287 | 5 | 4.8 | 4.67 | 6 | 5 | 5 | 4 | 5 | 4 | 4 | 5 | 6 | 5.4 | 12.11 | 4 | < 0.1 | 5 | 5.895 | < 10 | 6 |  | 33
| 1 | 3 | 5 | 5 | 5 | 5 | 4.9831 | 5 | 4.8 | 4.67 | 6 | 5 | 5 | 4 | 5 | 4 | 4 | 5 | 6 | 5.5 | 10.25 | 4 | < 0.1 | 5 | 5.9918 | < 10 | 6 |  | 16
| 1 | 4 | 5 | 5 | 5 | 5 | 4.598 | 5 | 4.7 | 4.57 | 6 | 5 | 5 | 4 | 4 | 4 | 4 | 5 | 6 | 5.4 | 10.85 | 4 | < 0.1 | 4 | 5.9476 | < 10 | 6 |  | 4.9654
| 1 | 5 | 5 | 5 | 5 | 5 | 4.9731 | 5 | 4.8 | 4.73 | 6 | 5 | 5 | 4 | 5 | 4 | 5 | 5 | 6 | 5 | 11.98 | 5 | < 0.1 | 5 | 6.0982 | < 10 | 6 |  | 67
| 1 | 6 | 5 | 5 | 5 | 5 | 4.9788 | 5 | 4.6 | 4.43 | 6 | 5 | 5 | 4 | 5 | 4 | 4 | 5 | 5 | 5.4 | 11.23 | 4 | < 0.1 | 4 | 5.8739 | < 10 | 5
| Mean |  | 5 | 5 | 5 | 5 | 4.8658 | 5 | 4.7167 | 4.6067 | 6 | 5 | 5 | 4 | 4.8333 | 4.1667 | 4.3333 | 5 | 5.8333 | 5.4 | 11.2333 | 4.1667 |  | 4.5 | 5.9828 |  | 5.8333
| Median |  | 5 | 5 | 5 | 5 | 4.9032 | 5 | 4.75 | 4.62 | 6 | 5 | 5 | 4 | 5 | 4 | 4 | 5 | 6 | 5.4 | 11.105 | 4 |  | 4.5 | 5.9697 |  | 6
| Std Dev. |  | 0 | 0 | 0 | 0 | 0.1498 | 0 | 0.098319 | 0.1069 | 0 | 0 | 0 | 0 | 0.40825 | 0.40825 | 0.5164 | 0 | 0.40825 | 0.22804 | 0.70769 | 0.40825 |  | 0.54772 | 0.095681 |  | 0.40825
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 3.079% | 0% | 2.085% | 2.32% | 0% | 0% | 0% | 0% | 8.447% | 9.798% | 11.92% | 0% | 6.999% | 4.223% | 6.3% | 9.798% |  | 12.17% | 1.599% |  | 6.999%
| PDM3 |  | 0.697% | 0.697% | 0.697% | 0.697% | -2.005% | 0.697% | -5.009% | -7.225% | 20.84% | 0.697% | 0.697% | -19.44% | -2.66% | -16.09% | -12.73% | 0.697% | 17.48% | 8.753% | 126.2% | -16.09% |  | -9.373% | 20.49% |  | 17.48%
| Z-Score (Absolute) |  | 0 | 0 | 0 | 0 | 0.23 | 0 | 0.49 | 0.67 | 1.71 | 0 | 0 | 1.71 | 0.29 | 1.43 | 1.14 | 0 | 1.43 | 0.69 | 10.69 | 1.43 | 8.49 | 0.86 | 1.68 | 0 | 1.43
| Analytical results for W in OREAS 606b (Certified Value 1.85 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 1.7 | 1.6 | 1.9 | 1.8 | < 10 | 2 | 1.9 | 1.9 | 2.1 | 1.9 | 1.8 | 1.7 | 1.8 | 1.8 | 1.7 | 1.87 | 2.2 | 2.21 | < 5 | 1.7 |  | < 1 | 1.7572 | 1.7 |  |  | 1
| 1 | 2 | 1.4 | 1.7 | 1.9 | 1.8 | < 10 | 1.9 | 1.9 | 2 | 2.1 | 1.9 | 1.9 | 1.8 | 1.8 | 1.9 | 1.8 | 2.02 | 2.3 | 2.26 | < 5 | 1.6 |  | < 1 | 1.7802 | 1.7 |  |  | 34
| 1 | 3 | 1.4 | 1.7 | 1.9 | 1.8 | < 10 | 2 | 2.2 | 2.03 | 2 | 1.9 | 1.9 | 1.7 | 1.8 | 1.9 | 1.8 | 1.96 | 2.3 | 2.36 | < 5 | 1.4 |  | < 1 | 1.8285 | 1.6 |  |  | 16
| 1 | 4 | 1.5 | 1.6 | 2 | 2 | < 10 | 1.9 | 1.8 | 1.97 | 2.1 | 1.9 | 1.8 | 1.7 | 1.7 | 1.8 | 1.8 | 2.03 | 2.2 | 2.16 | < 5 | 1.5 |  | < 1 | 1.7396 | 1.6 |  |  | 1.8512
| 1 | 5 | 1.3 | 1.8 | 1.8 | 1.9 | < 10 | 1.8 | 1.9 | 2.07 | 2 | 1.9 | 1.9 | 1.8 | 1.8 | 1.8 | 1.8 | 1.92 | 2.3 | 2.12 | < 5 | 1.7 |  | < 1 | 1.8006 | 1.8 |  |  | 68
| 1 | 6 | 1.5 | 1.8 | 1.8 | 1.8 | < 10 | 1.9 | 1.9 | 1.93 | 2.1 | 1.9 | 1.8 | 1.7 | 1.8 | 1.8 | 1.8 | 1.89 | 2.3 | 2.13 | < 5 | 1.6 | 4.3799 | < 1 | 1.7972 | 1.6
| Mean |  | 1.4667 | 1.7 | 1.8833 | 1.85 |  | 1.9167 | 1.9333 | 1.9833 | 2.0667 | 1.9 | 1.85 | 1.7333 | 1.7833 | 1.8333 | 1.7833 | 1.9483 | 2.2667 | 2.2067 |  | 1.5833 | 4.3799 |  | 1.7839 | 1.6667
| Median |  | 1.45 | 1.7 | 1.9 | 1.8 |  | 1.9 | 1.9 | 1.985 | 2.1 | 1.9 | 1.85 | 1.7 | 1.8 | 1.8 | 1.8 | 1.94 | 2.3 | 2.185 |  | 1.6 | 4.3799 |  | 1.7887 | 1.65
| Std Dev. |  | 0.13663 | 0.089443 | 0.075277 | 0.083666 |  | 0.075277 | 0.13663 | 0.06314 | 0.05164 | 2.4324e-16 | 0.054772 | 0.05164 | 0.040825 | 0.05164 | 0.040825 | 0.066758 | 0.05164 | 0.091579 |  | 0.1169 |  |  | 0.032026 | 0.08165
| Rel.Std.Dev. |  | 9.315% | 5.261% | 3.997% | 4.522% |  | 3.928% | 7.067% | 3.184% | 2.499% | 1.28e-14% | 2.961% | 2.979% | 2.289% | 2.817% | 2.289% | 3.426% | 2.278% | 4.15% |  | 7.383% |  |  | 1.795% | 4.899%
| PDM3 |  | -20.77% | -8.168% | 1.736% | -0.06464% |  | 3.537% | 4.437% | 7.138% | 11.64% | 2.636% | -0.06464% | -6.367% | -3.666% | -0.965% | -3.666% | 5.247% | 22.44% | 19.2% |  | -14.47% | 136.6% |  | -3.637% | -9.968%
| Z-Score (Absolute) |  | 2.16 | 0.9 | 0.09 | 0.09 | 16.9 | 0.27 | 0.36 | 0.63 | 1.08 | 0.18 | 0.09 | 0.72 | 0.45 | 0.18 | 0.45 | 0.44 | 2.16 | 1.83 | 3.42 | 1.53 | 13.56 | 7.37 | 0.45 | 1.08
| Analytical results for Y in OREAS 606b (Certified Value 14.8 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 16.2 | 14.6 | 14.9 | 14.39 | 14.3358 | 14.6 | 15.5 | 14.796 | 14.8 | 14.1 | 15.35 | 14.8 | 13.4 | 15 | 15.7 | 13.8 | 15.22 | 11.17 | 15.8 | 15.2 | 11.0517 | 16 | 14.9507 | 12.7 | 14.16 |  | 1
| 1 | 2 | 16.4 | 14.8 | 15 | 14.48 | 14.55 | 14.3 | 16 | 14.4432 | 15.5 | 14.4 | 15.55 | 15.1 | 13.6 | 14.4 | 16.4 | 13.6 | 14.97 | 12.32 | 16.7 | 14.9 | 10.9034 | 16 | 14.9899 | 12.821 | 13.66 |  | 35
| 1 | 3 | 16.4 | 14.9 | 14.9 | 14.31 | 14.5534 | 14.4 | 16.3 | 14.9373 | 15.3 | 14.3 | 15.43 | 15.3 | 13.6 | 14.8 | 15.9 | 13.8 | 15.24 | 10.93 | 15.2 | 15.2 | 10.7545 | 16 | 14.3778 | 12.82 | 14.49 |  | 16
| 1 | 4 | 16.6 | 14.5 | 15.3 | 14.85 | 14.005 | 14.3 | 15.7 | 14.724 | 14.9 | 14 | 15.05 | 15.6 | 14.1 | 14.1 | 15.9 | 13.9 | 15.13 | 10.44 | 16.6 | 15.1 | 11.1315 | 16 | 14.4344 | 12.493 | 14.26 |  | 14.85
| 1 | 5 | 17 | 16.3 | 15.1 | 14.57 | 14.4475 | 14.4 | 16 | 14.2298 | 15 | 14.5 | 14.74 | 15.9 | 14 | 14.7 | 15.6 | 14.5 | 15.37 | 11.06 | 16.4 | 15.2 | 11.3223 | 16 | 14.9653 | 12.765 | 14.61 |  | 69
| 1 | 6 | 15.2 | 15.8 | 15.1 | 14.01 | 14.5358 | 14.4 | 15.7 | 14.6115 | 14.2 | 14.1 | 15.11 | 15.2 | 13.6 | 14.6 | 12.8 | 13.9 | 14.93 | 12.21 | 16 | 15.1 | 11.1427 | 16 | 14.7792 | 13.012 | 13.67
| Mean |  | 16.3 | 15.15 | 15.05 | 14.435 | 14.4046 | 14.4 | 15.8667 | 14.6236 | 14.95 | 14.2333 | 15.205 | 15.3167 | 13.7167 | 14.6 | 15.3833 | 13.9167 | 15.1433 | 11.355 | 16.1167 | 15.1167 | 11.051 | 16 | 14.7495 | 12.7685 | 14.1417
| Median |  | 16.4 | 14.85 | 15.05 | 14.435 | 14.4916 | 14.4 | 15.85 | 14.6677 | 14.95 | 14.2 | 15.23 | 15.25 | 13.6 | 14.65 | 15.8 | 13.85 | 15.175 | 11.115 | 16.2 | 15.15 | 11.0916 | 16 | 14.8649 | 12.7925 | 14.21
| Std Dev. |  | 0.60332 | 0.7287 | 0.15166 | 0.27955 | 0.21294 | 0.10954 | 0.28752 | 0.25533 | 0.45056 | 0.19664 | 0.29663 | 0.38687 | 0.27142 | 0.31623 | 1.2952 | 0.30605 | 0.16872 | 0.74851 | 0.56716 | 0.1169 | 0.19909 | 0 | 0.27683 | 0.17047 | 0.40236
| Rel.Std.Dev. |  | 3.701% | 4.81% | 1.008% | 1.937% | 1.478% | 0.7607% | 1.812% | 1.746% | 3.014% | 1.382% | 1.951% | 2.526% | 1.979% | 2.166% | 8.42% | 2.199% | 1.114% | 6.592% | 3.519% | 0.7733% | 1.802% | 0% | 1.877% | 1.335% | 2.845%
| PDM3 |  | 9.765% | 2.02% | 1.347% | -2.794% | -2.999% | -3.03% | 6.846% | -1.524% | 0.6736% | -4.152% | 2.391% | 3.143% | -7.632% | -1.683% | 3.592% | -6.285% | 1.976% | -23.54% | 8.53% | 1.796% | -25.58% | 7.744% | -0.6763% | -14.02% | -4.77%
| Z-Score (Absolute) |  | 2.22 | 0.64 | 0.5 | 0.35 | 0.39 | 0.39 | 1.62 | 0.09 | 0.36 | 0.62 | 0.71 | 0.87 | 1.33 | 0.12 | 0.96 | 1.06 | 0.63 | 4.58 | 1.96 | 0.59 | 4.99 | NA | 0.09 | 2.63 | 0.75
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 606b (Certified Value 0.61 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.10; 2.08: 2.12; 2.09: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  | 2
2
| 1 | 1 | 2.01: 0.5; 2.03: 0.52; 2.04: 0.67458; 2.05: 0.6; 2.06: 0.6169; 2.07: 0.55; 2.08: 0.57; 2.09: 0.6; 2.10: 0.63; 2.11: 0.7466; 2.12: 0.7; 2.13: 0.5 |  | 1
| 1 | 2 | 2.01: 0.5; 2.03: 0.54; 2.04: 0.69826; 2.05: 0.6; 2.06: 0.6522; 2.07: 0.53; 2.08: 0.56; 2.09: 0.6; 2.10: 0.62; 2.11: 0.6182; 2.12: 0.6; 2.13: 0.6 |  | 36
| 1 | 3 | 2.01: 0.5; 2.03: 0.51; 2.04: 0.69043; 2.05: 0.6; 2.06: 0.6733; 2.07: 0.55; 2.08: 0.58; 2.09: 0.5; 2.10: 0.62; 2.11: 0.6435; 2.12: 0.6; 2.13: 0.5 |  | 16
| 1 | 4 | 2.01: 0.5; 2.03: 0.54; 2.04: 0.6705; 2.05: 0.6; 2.06: 0.6588; 2.07: 0.53; 2.08: 0.57; 2.09: 0.5; 2.10: 0.66; 2.11: 0.6426; 2.12: 0.6; 2.13: 0.5 |  | 0.60919
| 1 | 5 | 2.01: 0.5; 2.03: 0.51; 2.04: 0.68295; 2.05: 0.6; 2.06: 0.6984; 2.07: 0.55; 2.08: 0.57; 2.09: 0.5; 2.10: 0.63; 2.11: 0.6954; 2.12: 0.7; 2.13: 0.5 |  | 70
| 1 | 6 | 2.01: 0.6; 2.03: 0.52; 2.04: 0.66378; 2.05: 0.6; 2.06: 0.6205; 2.07: 0.54; 2.08: 0.58; 2.09: 0.6; 2.10: 0.63; 2.11: 0.6632; 2.12: 0.7; 2.13: 0.5
| Mean |  | 2.01: 0.51667; 2.03: 0.52333; 2.04: 0.68008; 2.05: 0.6; 2.06: 0.65335; 2.07: 0.54167; 2.08: 0.57167; 2.09: 0.55; 2.10: 0.63167; 2.11: 0.66825; 2.12: 0.65; 2.13: 0.51667
| Median |  | 2.01: 0.5; 2.03: 0.52; 2.04: 0.67876; 2.05: 0.6; 2.06: 0.6555; 2.07: 0.545; 2.08: 0.57; 2.09: 0.55; 2.10: 0.63; 2.11: 0.65335; 2.12: 0.65; 2.13: 0.5
| Std Dev. |  | 2.01: 0.040825; 2.03: 0.013663; 2.04: 0.012913; 2.05: 0; 2.06: 0.031197; 2.07: 0.0098319; 2.08: 0.0075277; 2.09: 0.054772; 2.10: 0.01472; 2.11: 0.04621; 2.12: 0.054772; 2.13: 0.040825
| Rel.Std.Dev. |  | 2.01: 7.902%; 2.03: 2.611%; 2.04: 1.899%; 2.05: 0%; 2.06: 4.775%; 2.07: 1.815%; 2.08: 1.317%; 2.09: 9.959%; 2.10: 2.33%; 2.11: 6.915%; 2.12: 8.427%; 2.13: 7.902%
| PDM3 |  | 2.01: -15.19%; 2.03: -14.09%; 2.04: 11.64%; 2.05: -1.509%; 2.06: 7.248%; 2.07: -11.08%; 2.08: -6.16%; 2.09: -9.717%; 2.10: 3.689%; 2.11: 9.694%; 2.12: 6.699%; 2.13: -15.19%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 1.51; 2.04: 0.67; 2.05: NA; 2.06: 0.3; 2.07: 1.25; 2.08: 0.84; 2.09: NA; 2.10: 0; 2.11: 0.51; 2.12: NA; 2.13: NA
Results from laboratories 2.03, 2.07, 2.19, 2.25 and 2.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 606b (Certified Value 316 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 307 | 317.6 | 315 | 304 | 311.2482 | 288 | 314 | 328.2672 | 310 | 326 | 333 | 319 | 313 | 313 | 305 | 351 | 331 | 323 | 383 | 281.8 | 319 | 270.1664 | 325 | 312.8904 | 290.217 | 310 | 1
| 1 | 2 | 306 | 310.6 | 327 | 308 | 322.5553 | 286 | 328 | 331.056 | 325 | 330 | 334 | 321 | 321 | 305 | 315 | 341 | 330 | 326 | 392 | 294.7 | 311 | 266.0419 | 312 | 312.0345 | 291.6 | 314 | 17
| 1 | 3 | 312 | 315.9 | 313 | 302 | 321.9995 | 282 | 333 | 323.2272 | 318 | 329 | 338 | 324 | 315 | 309 | 304 | 342 | 332 | 317 | 387 | 292.9 | 320 | 264.5829 | 310 | 308.7545 | 294.278 | 305 | 16
| 1 | 4 | 311 | 322.9 | 307 | 320 | 315.296 | 286 | 324 | 332.3472 | 305 | 334 | 336 | 329 | 316 | 306 | 304 | 351 | 330 | 311 | 380 | 282 | 315 | 274.2234 | 314 | 308.2558 | 286.215 | 310 | 315.9317
| 1 | 5 | 312 | 313.9 | 319 | 308 | 321.0716 | 281 | 337 | 331.1904 | 319 | 328 | 331 | 335 | 323 | 308 | 302 | 349 | 336 | 350 | 382 | 293.5 | 311 | 273.4558 | 318 | 315.9103 | 292.374 | 309 | 71
| 1 | 6 | 312 | 320.8 | 315 | 300 | 316.6756 | 287 | 330 | 334.6848 | 303 | 325 | 331 | 317 | 317 | 306 | 299 | 342 | 330 | 360 | 382 | 293.9 | 318 | 273.0664 | 325 | 314.6083 | 295.11 | 301
| Mean |  | 310 | 316.95 | 316 | 307 | 318.1411 | 285 | 327.6667 | 330.1288 | 313.3333 | 328.6667 | 333.8333 | 324.1667 | 317.5 | 307.8333 | 304.8333 | 346 | 331.5 | 331.1667 | 384.3333 | 289.8 | 315.6667 | 270.2561 | 317.3333 | 312.0756 | 291.6323 | 308.1667
| Median |  | 311.5 | 316.75 | 315 | 306 | 318.8736 | 286 | 329 | 331.1232 | 314 | 328.5 | 333.5 | 322.5 | 316.5 | 307 | 304 | 345.5 | 330.5 | 324.5 | 382.5 | 293.2 | 316.5 | 271.6164 | 316 | 312.4625 | 291.987 | 309.5
| Std Dev. |  | 2.7568 | 4.5028 | 6.6633 | 7.1274 | 4.4881 | 2.8284 | 8.0166 | 3.9689 | 8.6872 | 3.2042 | 2.7869 | 6.7651 | 3.7815 | 2.9269 | 5.4191 | 4.8166 | 2.3452 | 19.4259 | 4.4121 | 6.1475 | 3.9833 | 4.0946 | 6.5013 | 3.0791 | 3.1936 | 4.535
| Rel.Std.Dev. |  | 0.8893% | 1.421% | 2.109% | 2.322% | 1.411% | 0.9924% | 2.447% | 1.202% | 2.772% | 0.9749% | 0.8348% | 2.087% | 1.191% | 0.9508% | 1.778% | 1.392% | 0.7075% | 5.866% | 1.148% | 2.121% | 1.262% | 1.515% | 2.049% | 0.9867% | 1.095% | 1.472%
| PDM3 |  | -1.878% | 0.3223% | 0.02161% | -2.827% | 0.6993% | -9.791% | 3.714% | 4.494% | -0.8225% | 4.031% | 5.666% | 2.607% | 0.4964% | -2.563% | -3.513% | 9.517% | 4.928% | 4.822% | 21.65% | -8.271% | -0.08389% | -14.46% | 0.4436% | -1.221% | -7.691% | -2.458%
| Z-Score (Absolute) |  | 0.42 | 0.03 | 0.03 | 0.62 | 0.11 | 2.05 | 0.73 | 0.89 | 0.21 | 0.8 | 1.13 | 0.5 | 0.07 | 0.56 | 0.76 | 1.93 | 0.98 | 0.96 | 4.43 | 1.74 | 0.05 | 3.02 | 0.06 | 0.29 | 1.62 | 0.54
| Analytical results for Zr in OREAS 606b (Certified Value 243 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 230 | 232.4 | 225.1 | 235.2 | 246.5423 | 247 | 237 | 208.8125 | 254 | 237 | 239.4 | 241 | 221 | 253 | 236 | 248 | 226.4 | 230 | 251 | 278.9 | 244 | 248.59 | 296 | 241.0498 | 207.3 | 259 | 1
| 1 | 2 | 232 | 228.5 | 237.5 | 236.7 | 247.2641 | 252 | 250 | 203.7465 | 242 | 237 | 256.3 | 247 | 229 | 251 | 245 | 245 | 228.9 | 235 | 263 | 274.8 | 232 | 248.9308 | 280 | 226.6642 | 206.13 | 263 | 18
| 1 | 3 | 237 | 232.9 | 229.1 | 239.5 | 244.1933 | 244 | 252 | 201.2325 | 264 | 238 | 245.4 | 248 | 227 | 260 | 238 | 243 | 233.7 | 224 | 242 | 279.2 | 244 | 254.0788 | 288 | 228.7852 | 205.46 | 259 | 16
| 1 | 4 | 233 | 235.8 | 230.4 | 248.1 | 242.5815 | 250 | 243 | 205.403 | 249 | 236 | 244.3 | 250 | 227 | 252 | 239 | 250 | 232.9 | 216 | 258 | 273.5 | 240 | 251.5662 | 286 | 225.0454 | 200.72 | 256 | 242.9697
| 1 | 5 | 233 | 232.6 | 229.4 | 250 | 246.0621 | 239 | 246 | 204.9055 | 261 | 239 | 249.9 | 258 | 229 | 253 | 237 | 248 | 231.2 | 222 | 261 | 266.3 | 235 | 252.7561 | 286 | 233.6911 | 205.74 | 270 | 72
| 1 | 6 | 237 | 242.7 | 225.2 | 234.7 | 244.0452 | 245 | 241 | 201.5235 | 240 | 236 | 242 | 241 | 226 | 258 | 233 | 244 | 231.6 | 230 | 257 | 262.2 | 243 | 254.7039 | 287 | 238.8288 | 208.07 | 250
| Mean |  | 233.6667 | 234.15 | 229.45 | 240.7 | 245.1147 | 246.1667 | 244.8333 | 204.2706 | 251.6667 | 237.1667 | 246.2167 | 247.5 | 226.5 | 254.5 | 238 | 246.3333 | 230.7833 | 226.1667 | 255.3333 | 272.4833 | 239.6667 | 251.771 | 287.1667 | 232.3441 | 205.57 | 259.5
| Median |  | 233 | 232.75 | 229.25 | 238.1 | 245.1277 | 246 | 244.5 | 204.326 | 251.5 | 237 | 244.85 | 247.5 | 227 | 253 | 237.5 | 246.5 | 231.4 | 227 | 257.5 | 274.15 | 241.5 | 252.1611 | 286.5 | 231.2381 | 205.935 | 259
| Std Dev. |  | 2.8048 | 4.7911 | 4.5346 | 6.7073 | 1.7869 | 4.6224 | 5.6362 | 2.8059 | 9.8116 | 1.169 | 6.0635 | 6.3482 | 2.9496 | 3.6194 | 4 | 2.7325 | 2.7051 | 6.824 | 7.7115 | 6.8799 | 5.0859 | 2.5739 | 5.1543 | 6.6008 | 2.5737 | 6.7157
| Rel.Std.Dev. |  | 1.2% | 2.046% | 1.976% | 2.787% | 0.729% | 1.878% | 2.302% | 1.374% | 3.899% | 0.4929% | 2.463% | 2.565% | 1.302% | 1.422% | 1.681% | 1.109% | 1.172% | 3.017% | 3.02% | 2.525% | 2.122% | 1.022% | 1.795% | 2.841% | 1.252% | 2.588%
| PDM3 |  | -3.829% | -3.63% | -5.564% | -0.9341% | 0.8828% | 1.316% | 0.767% | -15.93% | 3.579% | -2.388% | 1.336% | 1.865% | -6.778% | 4.746% | -2.045% | 1.384% | -5.016% | -6.916% | 5.089% | 12.15% | -1.359% | 3.622% | 18.19% | -4.373% | -15.39% | 6.803%
| Z-Score (Absolute) |  | 0.68 | 0.64 | 0.99 | 0.15 | 0.17 | 0.25 | 0.15 | 2.87 | 0.66 | 0.42 | 0.26 | 0.35 | 1.21 | 0.87 | 0.36 | 0.27 | 0.89 | 1.24 | 0.94 | 2.21 | 0.23 | 0.67 | 3.31 | 0.78 | 2.77 | 1.25
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C952: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 606b (Certified Value 1.49 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 1.2 | 1.304 | 1.5 | 1.43 | 1.5276 | 1.64 | 1.46 | 1.171 | 1.5 | 1.1 | 1.4 | 1.45 | 1.5 | 1.56 | 1.68 | 1.55 | 1.3 | 1.1 | 1.45 | 1.18 | 1.49 | 3.4119 | 1.56 | 1.5697 | < 5 | 1.79 | 1
| 1 | 2 | 1.2 | 1.376 | 1.5 | 1.42 | 1.5471 | 1.55 | 1.48 | 1.1845 | 1.5 | 1.16 | 1.4 | 1.53 | 1.54 | 1.53 | 1.61 | 1.55 | 1.3 | 1.1 | 1.48 | 1.17 | 1.49 | 2.5537 | 1.62 | 1.5041 | < 5 | 1.7 | 22
| 1 | 3 | 1.2 | 1.345 | 1.5 | 1.4 | 1.5667 | 1.53 | 1.46 | 1.0895 | 1.5 | 1.14 | 1.4 | 1.54 | 1.58 | 1.56 | 1.56 | 1.63 | 1.3 | 1.1 | 1.45 | 1.16 | 1.39 | 5.4229 | 1.6 | 1.5522 | < 5 | 1.58 | 16
| 1 | 4 | 1.2 | 1.347 | 1.5 | 1.4 | 1.5774 | 1.59 | 1.47 | 1.215 | 1.6 | 1.21 | 1.4 | 1.42 | 1.58 | 1.54 | 1.54 | 1.55 | 1.3 | 1 | 1.54 | 1.13 | 1.41 | 2.6064 | 1.59 | 1.558 | < 5 | 1.62 | 1.4853
| 1 | 5 | 1.2 | 1.258 | 1.5 | 1.41 | 1.576 | 1.66 | 1.67 | 1.1955 | 1.4 | 1.13 | 1.4 | 1.45 | 1.54 | 1.52 | 1.59 | 1.55 | 1.3 | 1.1 | 1.47 | 1.19 | 1.44 | 284.7763 | 1.61 | 1.5709 | < 5 | 1.61 | 74
| 1 | 6 | 1.2 | 1.346 | 1.5 | 1.39 | 1.5747 | 1.54 | 1.54 | 1.1105 | 1.5 | 1.15 | 1.5 | 1.5 | 1.54 | 1.5 | 1.59 | 1.54 | 1.3 | 1.1 | 1.51 | 1.17 | 1.45 | 95.6139 | 1.6 | 1.5145 | < 5 | 1.5
| Mean |  | 1.2 | 1.3293 | 1.5 | 1.4083 | 1.5616 | 1.585 | 1.5133 | 1.161 | 1.5 | 1.1483 | 1.4167 | 1.4817 | 1.5467 | 1.535 | 1.595 | 1.5617 | 1.3 | 1.0833 | 1.4833 | 1.1667 | 1.445 | 65.7309 | 1.5967 | 1.5449 |  | 1.6333
| Median |  | 1.2 | 1.3455 | 1.5 | 1.405 | 1.5707 | 1.57 | 1.475 | 1.1778 | 1.5 | 1.145 | 1.4 | 1.475 | 1.54 | 1.535 | 1.59 | 1.55 | 1.3 | 1.1 | 1.475 | 1.17 | 1.445 | 4.4174 | 1.6 | 1.5551 |  | 1.615
| Std Dev. |  | 0 | 0.041817 | 0 | 0.01472 | 0.020104 | 0.054681 | 0.082381 | 0.049842 | 0.063246 | 0.03656 | 0.040825 | 0.048751 | 0.030111 | 0.023452 | 0.048477 | 0.033714 | 0 | 0.040825 | 0.03559 | 0.020656 | 0.040866 | 113.4642 | 0.020656 | 0.028679 |  | 0.10033
| Rel.Std.Dev. |  | 0% | 3.146% | 0% | 1.045% | 1.287% | 3.45% | 5.444% | 4.293% | 4.216% | 3.184% | 2.882% | 3.29% | 1.947% | 1.528% | 3.039% | 2.159% | 0% | 3.768% | 2.399% | 1.771% | 2.828% | 172.6% | 1.294% | 1.856% |  | 6.143%
| PDM3 |  | -19.21% | -10.5% | 0.9871% | -5.184% | 5.134% | 6.71% | 1.885% | -21.84% | 0.9871% | -22.69% | -4.623% | -0.2471% | 4.129% | 3.344% | 7.383% | 5.139% | -12.48% | -27.06% | -0.1349% | -21.45% | -2.716% | 4325% | 7.495% | 4.009% |  | 9.964%
| Z-Score (Absolute) |  | 2.29 | 1.3 | 0 | 0.7 | 0.47 | 0.65 | 0.1 | 2.59 | 0 | 2.68 | 0.64 | 0.14 | 0.36 | 0.27 | 0.73 | 0.47 | 1.53 | 3.18 | 0.13 | 2.54 | 0.42 | 490.32 | 0.74 | 0.34 | 7.63 | 1.02
| Analytical results for Al in OREAS 606b (Certified Value 0.804 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.83 | 0.97 | 0.82 | 0.817 | 0.78244 | 0.82 | 0.61 | 0.79 | 0.967 | 1.0241 | 0.78 | 0.83 | 0.79 | 0.94 | 0.84 | 0.869 | 0.773 | 0.73 | 0.85 | 0.7298 | 0.73 | 0.78035 | 0.87503 |  |  |  | 1
| 1 | 2 | 0.82 | 0.78 | 0.82 | 0.8038 | 0.77604 | 0.8 | 0.61 | 0.8 | 0.956 | 1.0258 | 0.79 | 0.81 | 0.79 | 0.91 | 0.83 | 0.87 | 0.78 | 0.73 | 0.88 | 0.657 | 0.693 | 0.75821 | 0.88775 |  |  |  | 0
| 1 | 3 | 0.84 | 1.08 | 0.83 | 0.8153 | 0.78352 | 0.72 | 0.61 | 0.77 | 0.98 | 1.0435 | 0.78 | 0.81 | 0.79 | 0.87 | 0.84 | 0.812 | 0.783 | 0.73 | 0.78 | 0.6757 | 0.704 | 0.77977 | 0.90232 |  |  |  | 16
| 1 | 4 | 0.84 | 1 | 0.82 | 0.8003 | 0.78374 | 0.81 | 0.58 | 0.79 | 0.979 | 0.9956 | 0.78 | 0.86 | 0.79 | 0.85 | 0.8 | 0.81 | 0.72 | 0.74 | 0.76 | 0.6554 | 0.725 | 0.7604 | 0.88219 |  |  |  | 0.80363
| 1 | 5 | 0.82 | 0.83 | 0.82 | 0.8078 | 0.76935 | 0.75 | 0.69 | 0.79 | 0.965 | 1.0402 | 0.8 | 0.83 | 0.79 | 0.88 | 0.82 | 0.825 | 0.79 | 0.71 | 0.8 | 0.6782 | 0.714 | 0.76896 | 0.87649 |  |  |  | 75
| 1 | 6 | 0.84 | 1.06 | 0.82 | 0.8151 | 0.78024 | 0.79 | 0.61 | 0.8 | 0.983 | 1.0222 | 0.8 | 0.84 | 0.8 | 0.91 | 0.8 | 0.178 | 0.78 | 0.72 | 0.81 | 0.7139 | 0.699 | 0.7604 | 0.90919
| Mean |  | 0.83167 | 0.95333 | 0.82167 | 0.80988 | 0.77922 | 0.78167 | 0.61833 | 0.79 | 0.97167 | 1.0252 | 0.78833 | 0.83 | 0.79167 | 0.89333 | 0.82167 | 0.72733 | 0.771 | 0.72667 | 0.81333 | 0.685 | 0.71083 | 0.76801 | 0.88883
| Median |  | 0.835 | 0.985 | 0.82 | 0.81145 | 0.78134 | 0.795 | 0.61 | 0.79 | 0.973 | 1.025 | 0.785 | 0.83 | 0.79 | 0.895 | 0.825 | 0.8185 | 0.78 | 0.73 | 0.805 | 0.67695 | 0.709 | 0.76468 | 0.88497
| Std Dev. |  | 0.0098319 | 0.12258 | 0.0040825 | 0.0069338 | 0.0056173 | 0.038687 | 0.037103 | 0.010954 | 0.010614 | 0.017009 | 0.0098319 | 0.018974 | 0.0040825 | 0.03266 | 0.018348 | 0.27046 | 0.025581 | 0.010328 | 0.044572 | 0.030447 | 0.014716 | 0.010035 | 0.014032
| Rel.Std.Dev. |  | 1.182% | 12.86% | 0.4969% | 0.8561% | 0.7209% | 4.949% | 6.001% | 1.387% | 1.092% | 1.659% | 1.247% | 2.286% | 0.5157% | 3.656% | 2.233% | 37.18% | 3.318% | 1.421% | 5.48% | 4.445% | 2.07% | 1.307% | 1.579%
| PDM3 |  | 3.489% | 18.63% | 2.245% | 0.7785% | -3.037% | -2.733% | -23.06% | -1.696% | 20.91% | 27.58% | -1.903% | 3.282% | -1.488% | 11.16% | 2.245% | -9.494% | -4.06% | -9.577% | 1.208% | -14.76% | -11.55% | -4.432% | 10.6%
| Z-Score (Absolute) |  | 0.7 | 2.84 | 0.53 | 0.32 | 0.22 | 0.18 | 3.05 | 0.03 | 3.17 | 4.11 | 0.06 | 0.67 | 0 | 1.79 | 0.53 | 1.13 | 0.36 | 1.14 | 0.38 | 1.88 | 1.42 | 0.42 | 1.71
| Analytical results for As in OREAS 606b (Certified Value 86 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 1
| 1 | 1 | 91 | 87 | 91.5 | 88 | 88.8189 | 86.7 | 82.5 | 85.8495 | 86.6 | 86 | 81 | 84 | 91.5 | 86.1 | 92.8 | 94.2 | 70.9 | 70 | 89.6 | 76.4 | 86.6 | 89.5732 | 79.2 | 71.5764 | 78.85 | 93.3 | 1
| 1 | 2 | 91 | 81.7 | 91.8 | 87 | 88.2768 | 84.2 | 82.6 | 82.8185 | 85 | 86 | 81 | 84.2 | 93 | 88.4 | 87.8 | 93.7 | 72.1 | 72 | 90.2 | 76.7 | 86.5 | 86.0177 | 82 | 69.5629 | 74.787 | 87.7 | 12
| 1 | 3 | 91 | 89.4 | 91 | 85 | 88.988 | 84.6 | 81.6 | 83.1155 | 83 | 88 | 82 | 83.8 | 92.8 | 88.5 | 84.5 | 97.1 | 70.2 | 68 | 88.9 | 77.7 | 81.4 | 87.6137 | 82.2 | 69.8451 | 77.613 | 82.3 | 16
| 1 | 4 | 91 | 86.2 | 91.1 | 87 | 89.6558 | 86.7 | 82.3 | 84.8315 | 83.4 | 87 | 80 | 83 | 92.1 | 88.3 | 84.2 | 93 | 71 | 67 | 90.1 | 76.6 | 81.1 | 85.6047 | 81.4 | 71.161 | 79 | 83.4 | 85.8337
| 1 | 5 | 90 | 88 | 92.9 | 85 | 87.6656 | 90.6 | 91.9 | 85.604 | 82.9 | 86 | 81 | 86.7 | 92.5 | 89.4 | 85 | 94.2 | 72.1 | 68 | 90.7 | 77.6 | 84.1 | 86.6718 | 83.3 | 70.6576 | 75.572 | 83.6 | 76
| 1 | 6 | 90 | 85.9 | 90.4 | 86 | 89.0302 | 85.3 | 84.1 | 81.572 | 84.1 | 87 | 81 | 84.3 | 93.6 | 88.6 | 86.8 | 92.8 | 70.3 | 69 | 89.8 | 76.5 | 85.6 | 106.8013 | 83.5 | 72.023 | 77.599 | 77.7
| Mean |  | 90.6667 | 86.3667 | 91.45 | 86.3333 | 88.7392 | 86.35 | 84.1667 | 83.9652 | 84.1667 | 86.6667 | 81 | 84.3333 | 92.5833 | 88.2167 | 86.85 | 94.1667 | 71.1 | 69 | 89.8833 | 76.9167 | 84.2167 | 90.3804 | 81.9333 | 70.8043 | 77.2368 | 84.6667
| Median |  | 91 | 86.6 | 91.3 | 86.5 | 88.9035 | 86 | 82.55 | 83.9735 | 83.75 | 86.5 | 81 | 84.1 | 92.65 | 88.45 | 85.9 | 93.95 | 70.95 | 68.5 | 89.95 | 76.65 | 84.85 | 87.1428 | 82.1 | 70.9093 | 77.606 | 83.5
| Std Dev. |  | 0.5164 | 2.6189 | 0.85499 | 1.2111 | 0.68693 | 2.3279 | 3.8759 | 1.7176 | 1.4264 | 0.8165 | 0.63246 | 1.2485 | 0.73052 | 1.1089 | 3.2334 | 1.5526 | 0.83666 | 1.7889 | 0.61128 | 0.57764 | 2.4685 | 8.1676 | 1.5591 | 0.96861 | 1.718 | 5.3031
| Rel.Std.Dev. |  | 0.5696% | 3.032% | 0.9349% | 1.403% | 0.7741% | 2.696% | 4.605% | 2.046% | 1.695% | 0.9421% | 0.7808% | 1.48% | 0.789% | 1.257% | 3.723% | 1.649% | 1.177% | 2.593% | 0.6801% | 0.751% | 2.931% | 9.037% | 1.903% | 1.368% | 2.224% | 6.263%
| PDM3 |  | 5.631% | 0.621% | 6.543% | 0.5821% | 3.385% | 0.6016% | -1.942% | -2.177% | -1.942% | 0.9705% | -5.631% | -1.748% | 7.864% | 2.776% | 1.184% | 9.708% | -17.17% | -19.61% | 4.718% | -10.39% | -1.884% | 5.297% | -4.544% | -17.51% | -10.02% | -1.36%
| Z-Score (Absolute) |  | 1.02 | 0.17 | 1.18 | 0.17 | 0.64 | 0.17 | 0.26 | 0.3 | 0.26 | 0.23 | 0.89 | 0.23 | 1.4 | 0.54 | 0.27 | 1.72 | 2.85 | 3.27 | 0.87 | 1.7 | 0.25 | 0.97 | 0.71 | 2.91 | 1.64 | 0.17
| Analytical results for B in OREAS 606b (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab; 2.15: Lab | 1
| No | No | 2.01: 2.01; 2.03: 2.03; 2.04: 2.04; 2.05: 2.06; 2.06: 2.07; 2.07: 2.08; 2.08: 2.12; 2.09: 2.13; 2.10: 2.14; 2.11: 2.15; 2.12: 2.16; 2.13: 2.17; 2.14: 2.23; 2.15: 2.26 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*OES; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*OES; 2.15: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 15g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 01g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.5g; 2.14: 0.5g; 2.15: 01g | 1
| 1 | 1 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 2; 2.08: 14; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 1.9; 2.08: 16; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 3
| 1 | 3 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 2.1; 2.08: 16; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 1.8; 2.08: 19; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | < 10
| 1 | 5 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 3; 2.07: 2.6; 2.08: 21; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 77
| 1 | 6 | 2.01: < 10; 2.03: < 20; 2.04: 1; 2.05: < 2; 2.06: 3; 2.07: 2.1; 2.08: 14; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10
| Mean |  | 2.04: 1.8333; 2.06: 2.3333; 2.07: 2.0833; 2.08: 16.6667
| Median |  | 2.04: 2; 2.06: 2; 2.07: 2.05; 2.08: 16
| Std Dev. |  | 2.04: 0.40825; 2.06: 0.5164; 2.07: 0.27869; 2.08: 2.8048
| Rel.Std.Dev. |  | 2.04: 22.27%; 2.06: 22.13%; 2.07: 13.38%; 2.08: 16.83%
PDM3
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 2.53; 2.04: 1.6; 2.05: 2.02; 2.06: 1.35; 2.07: 1.48; 2.08: 5.9; 2.09: 0; 2.10: 0; 2.11: 0; 2.12: 0; 2.13: 0; 2.14: 2.5; 2.15: 0
| Analytical results for Ba in OREAS 606b (Certified Value 212 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 0
| 1 | 1 | 229 | 280.7 | 213 | 208 | 213.7092 | 185 | 199 | 231 | 239 | 210 | 220 | 210 | 240 | 220 | 178 | 160 | 209.8 | 227 | 223.3476 | 157 | 188.5713 | 212.171 | 258 |  |  |  | 1
| 1 | 2 | 219 | 206.8 | 210 | 208 | 212.3285 | 124 | 200 | 225 | 239 | 210 | 210 | 210 | 240 | 220 | 180 | 170 | 206.9 | 226 | 199.1667 | 149 | 189.8136 | 210.336 | 248 |  |  |  | 24
| 1 | 3 | 232 | 325.4 | 211 | 206 | 212.2063 | 165 | 197 | 224 | 242 | 210 | 210 | 210 | 230 | 220 | 166 | 183 | 215.6 | 223 | 207.5167 | 156 | 187.1643 | 209.859 | 236 |  |  |  | 16
| 1 | 4 | 229 | 293.4 | 209 | 208 | 207.7556 | 146 | 199 | 227 | 232 | 210 | 230 | 210 | 220 | 210 | 167 | 164 | 209.6 | 225 | 198.9137 | 156 | 190.3682 | 211.076 | 236 |  |  |  | 211.7273
| 1 | 5 | 225 | 221.9 | 212 | 204 | 208.746 | 159 | 227 | 233 | 242 | 220 | 220 | 210 | 230 | 220 | 169 | 170 | 205.2 | 223 | 464.8221 | 159 | 189.3642 | 210.551 | 233 |  |  |  | 78
| 1 | 6 | 239 | 313 | 210 | 205 | 213.4235 | 129 | 203 | 223 | 240 | 210 | 220 | 210 | 240 | 210 | 170 | 160 | 206.9 | 228 | 222.9744 | 153 | 186.3087 | 211.217 | 206
| Mean |  | 228.8333 | 273.5333 | 210.8333 | 206.5 | 211.3615 | 151.3333 | 204.1667 | 227.1667 | 239 | 211.6667 | 218.3333 | 210 | 233.3333 | 216.6667 | 171.6667 | 167.8333 | 209 | 225.3333 | 252.7902 | 155 | 188.5984 | 210.8683 | 236.1667
| Median |  | 229 | 287.05 | 210.5 | 207 | 212.2674 | 152.5 | 199.5 | 226 | 239.5 | 210 | 220 | 210 | 235 | 220 | 169.5 | 167 | 208.25 | 225.5 | 215.2456 | 156 | 188.9678 | 210.8135 | 236
| Std Dev. |  | 6.7057 | 48.6069 | 1.472 | 1.7607 | 2.5003 | 23.0362 | 11.3563 | 4.0208 | 3.6878 | 4.0825 | 7.5277 | 0 | 8.165 | 5.164 | 5.8878 | 8.6814 | 3.6829 | 2.0656 | 104.4425 | 3.5214 | 1.5809 | 0.80802 | 17.5319
| Rel.Std.Dev. |  | 2.93% | 17.77% | 0.6982% | 0.8526% | 1.183% | 15.22% | 5.562% | 1.77% | 1.543% | 1.929% | 3.448% | 0% | 3.499% | 2.383% | 3.43% | 5.173% | 1.762% | 0.9167% | 41.32% | 2.272% | 0.8382% | 0.3832% | 7.424%
| PDM3 |  | 8.079% | 29.19% | -0.4222% | -2.469% | -0.1727% | -28.52% | -3.571% | 7.292% | 12.88% | -0.02862% | 3.12% | -0.8158% | 10.2% | 2.333% | -18.92% | -20.73% | -1.288% | 6.426% | 19.39% | -26.79% | -10.92% | -0.4057% | 11.54%
| Z-Score (Absolute) |  | 0.75 | 2.65 | 0.02 | 0.21 | 0 | 2.56 | 0.31 | 0.67 | 1.18 | 0.01 | 0.3 | 0.06 | 0.94 | 0.23 | 1.69 | 1.86 | 0.1 | 0.6 | 1.77 | 2.4 | 0.97 | 0.02 | 1.06
| Analytical results for Be in OREAS 606b (Certified Value 0.51 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 3
| 1 | 1 | 0.71 | 0.5 | 0.5 | 0.53048 | 0.5 | 0.49 | 0.6294 | 0.6 | 0.6 | 0.49 | 0.53 | 0.49 | 0.57 | 0.52 | 0.43 | 0.42 | < 2 | 0.47 | 0.6531 | 0.53 | 0.49951 | < 5 | 0.57 |  |  |  | 1
| 1 | 2 | 0.69 | 0.4 | 0.52 | 0.52671 | 0.5 | 0.41 | 0.6729 | 0.6 | 0.55 | 0.49 | 0.52 | 0.48 | 0.55 | 0.51 | 0.44 | 0.76 | < 2 | 0.42 | 0.7693 | 0.53 | 0.5066 | < 5 | 0.55 |  |  |  | 25
| 1 | 3 | 0.71 | 0.6 | 0.55 | 0.52057 | 0.5 | 0.54 | 0.6331 | 0.6 | 0.58 | 0.49 | 0.53 | 0.49 | 0.54 | 0.52 | 0.42 | 0.38 | < 2 | 0.4 | 0.6928 | 0.51 | 0.49949 | < 5 | 0.48 |  |  |  | 16
| 1 | 4 | 0.71 | 0.6 | 0.5 | 0.51754 | 0.6 | 0.46 | 0.6575 | 0.6 | 0.55 | 0.49 | 0.55 | 0.49 | 0.52 | 0.5 | 0.42 | 0.37 | < 2 | 0.38 | 0.7143 | 0.49 | 0.50674 | < 5 | 0.55 |  |  |  | 0.51276
| 1 | 5 | 0.73 | 0.6 | 0.49 | 0.50479 | 0.6 | 0.58 | 0.6819 | 0.6 | 0.57 | 0.5 | 0.53 | 0.48 | 0.53 | 0.51 | 0.42 | 0.48 | < 2 | 0.47 | 0.6645 | 0.51 | 0.49961 | < 5 | 0.5 |  |  |  | 79
| 1 | 6 | 0.72 | 0.5 | 0.54 | 0.51948 | 0.5 | 0.58 | 0.6576 | 0.6 | 0.58 | 0.5 | 0.54 | 0.49 | 0.55 | 0.5 | 0.43 | 0.59 | < 2 | 0.49 | 0.6517 | 0.48 | 0.49572 | < 5 | 0.49
| Mean |  | 0.71167 | 0.53333 | 0.51667 | 0.51993 | 0.53333 | 0.51 | 0.6554 | 0.6 | 0.57167 | 0.49333 | 0.53333 | 0.48667 | 0.54333 | 0.51 | 0.42667 | 0.5 |  | 0.43833 | 0.69095 | 0.50833 | 0.50128 |  | 0.52333
| Median |  | 0.71 | 0.55 | 0.51 | 0.52003 | 0.5 | 0.515 | 0.65755 | 0.6 | 0.575 | 0.49 | 0.53 | 0.49 | 0.545 | 0.51 | 0.425 | 0.45 |  | 0.445 | 0.67865 | 0.51 | 0.49956 |  | 0.525
| Std Dev. |  | 0.013292 | 0.08165 | 0.024221 | 0.0088655 | 0.05164 | 0.068702 | 0.020933 | 0 | 0.019408 | 0.005164 | 0.010328 | 0.005164 | 0.017512 | 0.0089443 | 0.008165 | 0.15086 |  | 0.04446 | 0.04552 | 0.020412 | 0.0044302 |  | 0.037771
| Rel.Std.Dev. |  | 1.868% | 15.31% | 4.688% | 1.705% | 9.682% | 13.47% | 3.194% | 0% | 3.395% | 1.047% | 1.936% | 1.061% | 3.223% | 1.754% | 1.914% | 30.17% |  | 10.14% | 6.588% | 4.016% | 0.8838% |  | 7.217%
| PDM3 |  | 38.79% | 4.012% | 0.762% | 1.398% | 4.012% | -0.5381% | 27.82% | 17.01% | 11.49% | -3.789% | 4.012% | -5.089% | 5.963% | -0.5381% | -16.79% | -2.488% |  | -14.51% | 34.75% | -0.8632% | -2.239% |  | 2.062%
| Z-Score (Absolute) |  | 5.22 | NA | 0.04 | 0.04 | NA | 0.22 | 3.7 | NA | 1.44 | 0.67 | 0.41 | 0.85 | 0.68 | 0.22 | 2.47 | 0.49 | 12.99 | 2.16 | 4.66 | 0.27 | 0.46 | 53.45 | 0.14
Results from laboratories 2.03, 2.07 and 2.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 606b (Certified Value 4.96 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 3
| 1 | 1 | 5.24 | 4.94 | 5.4 | 5.19 | 5.0252 | 5.11 | 5.1 | 5.09 | 4.63 | 5.04 | 5.12 | 4.81 | 5.25 | 5.75 | 4.72 | 4.98 | 4.34 | 4.68 | < 5 | 4.96 | 19.6914 | 5.33 | 4.6975 | < 5 | 4.74 |  | 1
| 1 | 2 | 4.92 | 4.67 | 5.2 | 5.2 | 4.9036 | 5.15 | 5.17 | 4.91 | 4.67 | 4.95 | 4.94 | 4.84 | 5.29 | 5.81 | 4.7 | 4.98 | 4.48 | 4.59 | < 5 | 4.97 | 15.3771 | 5.41 | 4.8185 | < 5 | 4.37 |  | 26
| 1 | 3 | 5.1 | 5.07 | 5.3 | 5.06 | 4.9332 | 4.98 | 5.1 | 5.02 | 4.68 | 5.07 | 4.97 | 4.71 | 5.26 | 5.45 | 4.81 | 4.98 | 4.15 | 4.7 | < 5 | 4.7 | 20.696 | 5.38 | 4.6677 | < 5 | 4.19 |  | 16
| 1 | 4 | 5.17 | 4.83 | 5.2 | 5.14 | 4.8918 | 5.25 | 5.15 | 5.12 | 4.67 | 4.89 | 5 | 4.88 | 5.32 | 5.61 | 4.64 | 4.98 | 4.14 | 4.95 | < 5 | 4.67 | 22.58 | 5.31 | 4.8568 | < 5 | 4.18 |  | 4.9574
| 1 | 5 | 4.98 | 4.4 | 5.5 | 5.09 | 5.0311 | 5.37 | 5.88 | 5.08 | 4.59 | 5.15 | 5.13 | 4.85 | 5.37 | 5.53 | 4.67 | 4.98 | 4.28 | 4.72 | < 5 | 4.72 | 16.278 | 5.83 | 4.8903 | < 5 | 4.23 |  | 80
| 1 | 6 | 5.25 | 4.96 | 5.2 | 5.13 | 4.8876 | 5.16 | 5.21 | 4.88 | 4.65 | 4.98 | 5.02 | 4.69 | 5.41 | 5.54 | 4.61 | 4.98 | 4.31 | 4.84 | < 5 | 4.86 | 15.2463 | 5.36 | 4.8235 | < 5 | 3.9
| Mean |  | 5.11 | 4.8117 | 5.3 | 5.135 | 4.9454 | 5.17 | 5.2683 | 5.0167 | 4.6483 | 5.0133 | 5.03 | 4.7967 | 5.3167 | 5.615 | 4.6917 | 4.98 | 4.2833 | 4.7467 |  | 4.8133 | 18.3115 | 5.4367 | 4.7924 |  | 4.2683
| Median |  | 5.135 | 4.885 | 5.25 | 5.135 | 4.9184 | 5.155 | 5.16 | 5.05 | 4.66 | 5.01 | 5.01 | 4.825 | 5.305 | 5.575 | 4.685 | 4.98 | 4.295 | 4.71 |  | 4.79 | 17.9847 | 5.37 | 4.821 |  | 4.21
| Std Dev. |  | 0.13653 | 0.24293 | 0.12649 | 0.054681 | 0.066047 | 0.13161 | 0.30262 | 0.10013 | 0.033714 | 0.092664 | 0.078486 | 0.078401 | 0.06314 | 0.13882 | 0.070261 | 0 | 0.12723 | 0.12801 |  | 0.13441 | 3.0968 | 0.19593 | 0.089382 |  | 0.27709
| Rel.Std.Dev. |  | 2.672% | 5.049% | 2.387% | 1.065% | 1.336% | 2.546% | 5.744% | 1.996% | 0.7253% | 1.848% | 1.56% | 1.634% | 1.188% | 2.472% | 1.498% | 0% | 2.97% | 2.697% |  | 2.792% | 16.91% | 3.604% | 1.865% |  | 6.492%
| PDM3 |  | 3.079% | -2.939% | 6.912% | 3.583% | -0.2408% | 4.289% | 6.273% | 1.196% | -6.234% | 1.129% | 1.465% | -3.242% | 7.248% | 13.27% | -5.36% | 0.4565% | -13.6% | -4.25% |  | -2.905% | 269.4% | 9.668% | -3.329% |  | -13.9%
| Z-Score (Absolute) |  | 0.38 | 0.49 | 0.92 | 0.45 | 0.1 | 0.55 | 0.83 | 0.11 | 0.96 | 0.1 | 0.14 | 0.53 | 0.97 | 1.84 | 0.83 | 0 | 2.01 | 0.67 | 7.17 | 0.48 | 38.53 | 1.32 | 0.54 | 7.17 | 2.06
| Analytical results for Ca in OREAS 606b (Certified Value 0.632 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.65 | 0.62 | 0.63 | 0.62 | 0.6299 | 0.65 | 0.59 | 0.62 | 0.67 | 0.67 | 0.64 | 0.65 | 0.64 | 0.68 | 0.68 | 0.6362 | 0.648 | 0.704 | 0.58 | 0.64 | 0.5945 | 0.643 | 0.62447 | 0.68805 |  |  | 1
| 1 | 2 | 0.64 | 0.58 | 0.62 | 0.61 | 0.62333 | 0.67 | 0.6 | 0.61 | 0.66 | 0.67 | 0.63 | 0.65 | 0.63 | 0.65 | 0.67 | 0.6239 | 0.657 | 0.71 | 0.58 | 0.63 | 0.5794 | 0.622 | 0.58687 | 0.65218 |  |  | 0
| 1 | 3 | 0.64 | 0.64 | 0.63 | 0.62 | 0.62887 | 0.65 | 0.6 | 0.61 | 0.66 | 0.68 | 0.63 | 0.65 | 0.63 | 0.64 | 0.68 | 0.6245 | 0.631 | 0.687 | 0.59 | 0.6 | 0.5916 | 0.636 | 0.61338 | 0.65573 |  |  | 16
| 1 | 4 | 0.65 | 0.61 | 0.63 | 0.6 | 0.63639 | 0.7 | 0.6 | 0.62 | 0.66 | 0.66 | 0.64 | 0.64 | 0.63 | 0.63 | 0.65 | 0.6324 | 0.634 | 0.709 | 0.58 | 0.6 | 0.5723 | 0.643 | 0.62949 | 0.66134 |  |  | 0.63219
| 1 | 5 | 0.64 | 0.63 | 0.63 | 0.62 | 0.61979 | 0.68 | 0.68 | 0.62 | 0.68 | 0.67 | 0.65 | 0.65 | 0.63 | 0.64 | 0.66 | 0.6441 | 0.639 | 0.72 | 0.59 | 0.62 | 0.5866 | 0.643 | 0.61928 | 0.65747 |  |  | 81
| 1 | 6 | 0.64 | 0.6 | 0.62 | 0.61 | 0.63003 | 0.65 | 0.61 | 0.62 | 0.66 | 0.67 | 0.64 | 0.65 | 0.64 | 0.66 | 0.65 | 0.6305 | 0.138 | 0.7 | 0.58 | 0.63 | 0.5828 | 0.622 | 0.57838 | 0.65891
| Mean |  | 0.64333 | 0.61333 | 0.62667 | 0.61333 | 0.62805 | 0.66667 | 0.61333 | 0.61667 | 0.665 | 0.67 | 0.63833 | 0.64833 | 0.63333 | 0.65 | 0.665 | 0.63193 | 0.55783 | 0.705 | 0.58333 | 0.62 | 0.58453 | 0.63483 | 0.60864 | 0.66228
| Median |  | 0.64 | 0.615 | 0.63 | 0.615 | 0.62938 | 0.66 | 0.6 | 0.62 | 0.66 | 0.67 | 0.64 | 0.65 | 0.63 | 0.645 | 0.665 | 0.63145 | 0.6365 | 0.7065 | 0.58 | 0.625 | 0.5847 | 0.6395 | 0.61633 | 0.65819
| Std Dev. |  | 0.005164 | 0.021602 | 0.005164 | 0.008165 | 0.0057977 | 0.020656 | 0.033267 | 0.005164 | 0.0083666 | 0.0063246 | 0.0075277 | 0.0040825 | 0.005164 | 0.017889 | 0.013784 | 0.0075933 | 0.2059 | 0.0111 | 0.005164 | 0.016733 | 0.0081554 | 0.010304 | 0.021027 | 0.012994
| Rel.Std.Dev. |  | 0.8027% | 3.522% | 0.824% | 1.331% | 0.9231% | 3.098% | 5.424% | 0.8374% | 1.258% | 0.944% | 1.179% | 0.6297% | 0.8154% | 2.752% | 2.073% | 1.202% | 36.91% | 1.574% | 0.8853% | 2.699% | 1.395% | 1.623% | 3.455% | 1.962%
| PDM3 |  | 1.763% | -2.982% | -0.8733% | -2.982% | -0.6542% | 5.454% | -2.982% | -2.455% | 5.19% | 5.981% | 0.9722% | 2.554% | 0.1813% | 2.818% | 5.19% | -0.04018% | -11.76% | 11.52% | -7.728% | -1.928% | -7.538% | 0.4185% | -3.724% | 4.76%
| Z-Score (Absolute) |  | 0.37 | 0.67 | 0.21 | 0.67 | 0.16 | 1.19 | 0.67 | 0.56 | 1.13 | 1.31 | 0.2 | 0.55 | 0.02 | 0.61 | 1.13 | 0.02 | 2.61 | 2.53 | 1.72 | 0.44 | 1.68 | 0.08 | 0.84 | 1.04
| Analytical results for Cd in OREAS 606b (Certified Value 1.19 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 1.33 | 1.24 | 1.2 | 1.15 | 1.1221 | 1.15 | 1.21 | 1.27 | 1.14 | 1.17 | 1.16 | 1.14 | 1.19 | 1.33 | 1.23 | 1.11 | 1.18 | 1.1 | 1.21 | 1.3963 | 1.21 | 1.1697 | < 5 | 1.47 |  |  | 1
| 1 | 2 | 1.34 | 1.19 | 1.2 | 1.07 | 1.1202 | 1.17 | 1.2 | 1.25 | 1.12 | 1.19 | 1.27 | 1.18 | 1.22 | 1.28 | 1.25 | 1.18 | 1.23 | 1.1 | 1.24 | 1.3209 | 1.25 | 1.1529 | < 5 | 1.34 |  |  | 27
| 1 | 3 | 1.17 | 1.21 | 1.1 | 1.15 | 1.1299 | 1.14 | 1.24 | 1.2 | 1.13 | 1.15 | 1.21 | 1.15 | 1.27 | 1.21 | 1.22 | 1.09 | 1.04 | 1.1 | 1.16 | 1.2411 | 1.27 | 1.1148 | < 5 | 1.2 |  |  | 16
| 1 | 4 | 1.41 | 1.15 | 1.2 | 1.19 | 1.0874 | 1.19 | 1.25 | 1.29 | 1.19 | 1.23 | 1.1 | 1.63 | 1.25 | 1.21 | 1.26 | 1.11 | 1.22 | 1.1 | 1.16 | 1.1243 | 1.21 | 1.0993 | < 5 | 1.33 |  |  | 1.1946
| 1 | 5 | 1.26 | 1.17 | 1.2 | 1.15 | 1.0969 | 1.18 | 1.33 | 1.14 | 1.12 | 1.22 | 1.19 | 1.2 | 1.25 | 1.26 | 1.22 | 1.12 | 1.2 | 1.1 | 1.2 | 1.2245 | 1.22 | 1.1016 | < 5 | 1.28 |  |  | 82
| 1 | 6 | 1.45 | 1.1 | 1.1 | 1.23 | 1.1272 | 1.21 | 1.26 | 1.23 | 1.15 | 1.21 | 1.23 | 1.17 | 1.22 | 1.24 | 1.22 | 1.18 | 1.1 | 1.1 | 1.23 | 1.2989 | 1.24 | 1.1349 | < 5 | 1.18
| Mean |  | 1.3267 | 1.1767 | 1.1667 | 1.1567 | 1.114 | 1.1733 | 1.2483 | 1.23 | 1.1417 | 1.195 | 1.1933 | 1.245 | 1.2333 | 1.255 | 1.2333 | 1.1317 | 1.1617 | 1.1 | 1.2 | 1.2677 | 1.2333 | 1.1289 |  | 1.3
| Median |  | 1.335 | 1.18 | 1.2 | 1.15 | 1.1212 | 1.175 | 1.245 | 1.24 | 1.135 | 1.2 | 1.2 | 1.175 | 1.235 | 1.25 | 1.225 | 1.115 | 1.19 | 1.1 | 1.205 | 1.27 | 1.23 | 1.1249 |  | 1.305
| Std Dev. |  | 0.10132 | 0.048854 | 0.05164 | 0.053166 | 0.017495 | 0.02582 | 0.046224 | 0.054037 | 0.026394 | 0.030822 | 0.058878 | 0.18982 | 0.028752 | 0.045935 | 0.017512 | 0.038687 | 0.075476 | 2.4324e-16 | 0.034059 | 0.093251 | 0.024221 | 0.028638 |  | 0.10602
| Rel.Std.Dev. |  | 7.638% | 4.152% | 4.426% | 4.597% | 1.571% | 2.201% | 3.703% | 4.393% | 2.312% | 2.579% | 4.934% | 15.25% | 2.331% | 3.66% | 1.42% | 3.419% | 6.497% | 2.211e-14% | 2.838% | 7.356% | 1.964% | 2.537% |  | 8.155%
| PDM3 |  | 11.05% | -1.504% | -2.341% | -3.178% | -6.753% | -1.783% | 4.495% | 2.961% | -4.434% | 0.03092% | -0.1086% | 4.216% | 3.24% | 5.053% | 3.24% | -5.271% | -2.759% | -7.921% | 0.4495% | 6.114% | 3.24% | -5.504% |  | 8.82%
| Z-Score (Absolute) |  | 1.97 | 0.32 | 0.47 | 0.62 | 1.28 | 0.37 | 0.78 | 0.5 | 0.85 | 0.04 | 0.06 | 0.73 | 0.55 | 0.88 | 0.55 | 1.01 | 0.55 | 1.49 | 0.04 | 1.07 | 0.55 | 1.05 | 19.89 | 1.56
| Analytical results for Ce in OREAS 606b (Certified Value 56 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 1
| 1 | 1 | 63 | 53.5 | 57.637 | 49.1467 | 61.6 | 53.8 | 67.052 | 53.63 | 57.9 | 63.49 | 50.4 | 53.1 | 54.7 | 60.6 | 56.9 | 55.78 | 41 | 56.6 | 61.8443 | 38.6 | 58.2075 | 57.837 | 58.58 |  |  |  | 1
| 1 | 2 | 62 | 49.9 | 56.857 | 48.5913 | 61.7 | 54.1 | 67.755 | 53.7 | 57.4 | 62.7 | 57.8 | 53.6 | 57.1 | 58.1 | 55.6 | 56.47 | 42.2 | 56.85 | 56.7838 | 37.3 | 55.6631 | 56.958 | 56.48 |  |  |  | 28
| 1 | 3 | 63 | 52.8 | 56.193 | 49.4736 | 60.1 | 53.6 | 67.88 | 54.2 | 58.8 | 63.42 | 56.4 | 53.3 | 58 | 55.7 | 58 | 50.46 | 39.7 | 52.39 | 55.2078 | 37.4 | 55.553 | 56.934 | 53.31 |  |  |  | 16
| 1 | 4 | 64 | 51.3 | 57.498 | 49.4792 | 63.2 | 52.9 | 67.814 | 53.7 | 58.1 | 60.3 | 49.6 | 53.6 | 56.6 | 53.4 | 55.7 | 50.83 | 41.7 | 52.5 | 51.9683 | 38.1 | 57.5731 | 57.231 | 53.84 |  |  |  | 56.153
| 1 | 5 | 63 | 50.1 | 56.52 | 48.7777 | 65 | 60.4 | 67.019 | 53.8 | 59.1 | 62.91 | 52.5 | 54.1 | 56.7 | 56.5 | 55.6 | 52.07 | 42.7 | 52.77 | 55.5469 | 37.9 | 56.7461 | 57.683 | 52.28 |  |  |  | 83
| 1 | 6 | 64 | 52.5 | 57.459 | 49.4821 | 61.2 | 54.8 | 67.969 | 54.27 | 58.9 | 62.95 | 59.1 | 54.1 | 56.9 | 56.5 | 56.3 | 52.39 | 41.9 | 55.66 | 62.6762 | 37.4 | 57.5895 | 56.614 | 47.5
| Mean |  | 63.1667 | 51.6833 | 57.0273 | 49.1584 | 62.1333 | 54.9333 | 67.5815 | 53.8833 | 58.3667 | 62.6283 | 54.3 | 53.6333 | 56.6667 | 56.8 | 56.35 | 53 | 41.5333 | 54.4617 | 57.3379 | 37.7833 | 56.8887 | 57.2095 | 53.665
| Median |  | 63 | 51.9 | 57.158 | 49.3101 | 61.65 | 53.95 | 67.7845 | 53.75 | 58.45 | 62.93 | 54.45 | 53.6 | 56.8 | 56.5 | 56 | 52.23 | 41.8 | 54.215 | 56.1653 | 37.65 | 57.1596 | 57.0945 | 53.575
| Std Dev. |  | 0.75277 | 1.4865 | 0.59364 | 0.39337 | 1.7224 | 2.7493 | 0.429 | 0.27862 | 0.66833 | 1.1814 | 4.006 | 0.40825 | 1.0857 | 2.4116 | 0.95656 | 2.5364 | 1.0596 | 2.1314 | 4.1398 | 0.51153 | 1.0959 | 0.47163 | 3.799
| Rel.Std.Dev. |  | 1.192% | 2.876% | 1.041% | 0.8002% | 2.772% | 5.005% | 0.6348% | 0.5171% | 1.145% | 1.886% | 7.378% | 0.7612% | 1.916% | 4.246% | 1.698% | 4.786% | 2.551% | 3.914% | 7.22% | 1.354% | 1.926% | 0.8244% | 7.079%
| PDM3 |  | 12.49% | -7.96% | 1.557% | -12.46% | 10.65% | -2.172% | 20.35% | -4.042% | 3.942% | 11.53% | -3.3% | -4.487% | 0.9148% | 1.152% | 0.3509% | -5.615% | -26.04% | -3.012% | 2.11% | -32.71% | 1.31% | 1.882% | -4.431%
| Z-Score (Absolute) |  | 1.86 | 1.28 | 0.19 | 1.97 | 1.58 | 0.39 | 3.07 | 0.67 | 0.55 | 1.72 | 0.56 | 0.74 | 0.09 | 0.12 | 0 | 0.92 | 4.05 | 0.52 | 0.27 | 5.08 | 0.15 | 0.23 | 0.73
| Analytical results for Co in OREAS 606b (Certified Value 2.02 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 2 | 2 | 2.1 | 1.9 | 2.0462 | 2 | 2 | 2.07 | 2.2 | 2 | 2 | 2 | 2 | 2.2 | 1.9 | 2.1 | 2.02 | 2.2 | 2 | 5.1055 | 1.69 | 2.0367 | < 5 | 2.2 |  |  | 1
| 1 | 2 | 2 | 1.9 | 2.1 | 2 | 2.0251 | 1.9 | 2 | 2.03 | 2 | 2 | 2.1 | 2.1 | 2 | 2.1 | 1.9 | 2.1 | 2.1 | 2.3 | 2 |  | 1.73 | 2.1176 | < 5 | 2.1 |  |  | 16
| 1 | 3 | 2 | 2 | 2.1 | 1.9 | 2.0111 | 2 | 1.9 | 2.11 | 2 | 2 | 2.1 | 2 | 2 | 2 | 1.9 | 2 | 2 | 2.2 | 2 | 5.4519 | 1.74 | 1.9678 | < 5 | 2 |  |  | 16
| 1 | 4 | 2 | 1.9 | 2 | 1.9 | 2.0507 | 2 | 2 | 2.03 | 2.1 | 2 | 1.9 | 2 | 2 | 1.9 | 2.1 | 2 | 1.9 | 2.2 | 2 | 4.8559 | 1.69 | 2.0349 | < 5 | 2 |  |  | 2.0231
| 1 | 5 | 2 | 1.9 | 2.1 | 1.9 | 1.9655 | 2.1 | 2.2 | 2.01 | 2.1 | 2 | 2 | 2.1 | 2 | 2 | 1.9 | 2 | 2 | 2.2 | 2 | 1.5059 | 1.7 | 2.1273 | < 5 | 2 |  |  | 84
| 1 | 6 | 2 | 1.9 | 2 | 1.9 | 1.9932 | 1.9 | 1.9 | 2.09 | 2.1 | 2 | 2.1 | 2 | 2 | 2 | 1.9 | 2 | 2 | 2.3 | 2 | 7.2077 | 1.73 | 2.0255 | < 5 | 1.8
| Mean |  | 2 | 1.9333 | 2.0667 | 1.9167 | 2.0153 | 1.9833 | 2 | 2.0567 | 2.0833 | 2 | 2.0333 | 2.0333 | 2 | 2.0333 | 1.9333 | 2.0333 | 2.0033 | 2.2333 | 2 | 4.8254 | 1.7133 | 2.0517 |  | 2.0167
| Median |  | 2 | 1.9 | 2.1 | 1.9 | 2.0181 | 2 | 2 | 2.05 | 2.1 | 2 | 2.05 | 2 | 2 | 2 | 1.9 | 2 | 2 | 2.2 | 2 | 5.1055 | 1.715 | 2.0358 |  | 2
| Std Dev. |  | 0 | 0.05164 | 0.05164 | 0.040825 | 0.032511 | 0.075277 | 0.10954 | 0.039328 | 0.075277 | 0 | 0.08165 | 0.05164 | 0 | 0.10328 | 0.08165 | 0.05164 | 0.06377 | 0.05164 | 0 | 2.0716 | 0.022509 | 0.060482 |  | 0.13292
| Rel.Std.Dev. |  | 0% | 2.671% | 2.499% | 2.13% | 1.613% | 3.795% | 5.477% | 1.912% | 3.613% | 0% | 4.016% | 2.54% | 0% | 5.079% | 4.223% | 2.54% | 3.183% | 2.312% | 0% | 42.93% | 1.314% | 2.948% |  | 6.591%
| PDM3 |  | -1.144% | -4.439% | 2.151% | -5.263% | -0.3885% | -1.968% | -1.144% | 1.657% | 2.975% | -1.144% | 0.5037% | 0.5037% | -1.144% | 0.5037% | -4.439% | 0.5037% | -0.9791% | 10.39% | -1.144% | 138.5% | -15.31% | 1.41% |  | -0.3201%
| Z-Score (Absolute) |  | 0.43 | 2.23 | 1.37 | 2.68 | 0.02 | 0.88 | 0.43 | 1.1 | 1.82 | 0.43 | 0.47 | 0.47 | 0.43 | 0.47 | 2.23 | 0.47 | 0.34 | 5.86 | 0.43 | 75.78 | 8.16 | 0.96 | 13.06 | 0.02
| Analytical results for Cr in OREAS 606b (Certified Value 16.1 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 2
| 1 | 1 | 15 | 16.4 | 18 | 16 | 15.3953 | 14 | 16.4 | 13.385 | 15.7 | 15 | 17 | 16 | 17 | 16 | 18 | 16 | 18 | 16.9 | 17 | 14.7266 | 16 | 14.3202 | 16.212 |  |  |  | 1
| 1 | 2 | 15 | 15 | 18 | 16 | 16.525 | 14 | 16.8 | 14.1515 | 15.7 | 15 | 17 | 16 | 18 | 15 | 17 | 17 | 19 | 17 | 17 | 18.0378 | 15 | 15.1176 | 15.379 |  |  |  | 30
| 1 | 3 | 16 | 16.4 | 18 | 16 | 15.5348 | 16 | 16.5 | 14.031 | 16.3 | 15 | 17 | 16 | 16 | 15 | 17 | 17 | 18 | 16.8 | 15 | 16.3798 | 17 | 14.4828 | 15.418 |  |  |  | 16
| 1 | 4 | 17 | 16.8 | 18 | 15 | 16.4497 | 15 | 16.6 | 14.267 | 15.7 | 15 | 16 | 16 | 17 | 16 | 16 | 16 | 19 | 16.2 | 15 | 15.0423 | 17 | 15.3862 | 15.954 |  |  |  | 16.1418
| 1 | 5 | 18 | 16.1 | 18 | 16 | 15.6074 | 15 | 18.9 | 13.7315 | 16.7 | 15 | 17 | 17 | 17 | 16 | 16 | 16 | 18 | 16 | 15 | 17.5441 | 16 | 15.5152 | 15.898 |  |  |  | 85
| 1 | 6 | 17 | 15.9 | 18 | 16 | 15.8181 | 13 | 16.3 | 13.343 | 16.3 | 15 | 17 | 16 | 17 | 17 | 18 | 17 | 5 | 17 | 16 | 16.0778 | 16 | 15.3192 | 15.002
| Mean |  | 16.3333 | 16.1 | 18 | 15.8333 | 15.8884 | 14.5 | 16.9167 | 13.8182 | 16.0667 | 15 | 16.8333 | 16.1667 | 17 | 15.8333 | 17 | 16.5 | 16.1667 | 16.65 | 15.8333 | 16.3014 | 16.1667 | 15.0236 | 15.6438
| Median |  | 16.5 | 16.25 | 18 | 16 | 15.7127 | 14.5 | 16.55 | 13.8812 | 16 | 15 | 17 | 16 | 17 | 16 | 17 | 16.5 | 18 | 16.85 | 15.5 | 16.2288 | 16 | 15.2184 | 15.658
| Std Dev. |  | 1.2111 | 0.61968 | 0 | 0.40825 | 0.48425 | 1.0488 | 0.98675 | 0.39466 | 0.4274 | 0 | 0.40825 | 0.40825 | 0.63246 | 0.75277 | 0.89443 | 0.54772 | 5.4924 | 0.43704 | 0.98319 | 1.3176 | 0.75277 | 0.50133 | 0.45093
| Rel.Std.Dev. |  | 7.415% | 3.849% | 0% | 2.578% | 3.048% | 7.233% | 5.833% | 2.856% | 2.66% | 0% | 2.425% | 2.525% | 3.72% | 4.754% | 5.261% | 3.32% | 33.97% | 2.625% | 6.21% | 8.083% | 4.656% | 3.337% | 2.882%
| PDM3 |  | 1.187% | -0.2588% | 11.51% | -1.911% | -1.57% | -10.17% | 4.801% | -14.4% | -0.4653% | -7.073% | 4.284% | 0.1542% | 5.317% | -1.911% | 5.317% | 2.219% | 0.1542% | 3.148% | -1.911% | 0.9889% | 0.1542% | -6.928% | -3.085%
| Z-Score (Absolute) |  | 0.34 | 0.13 | 3.71 | 0.67 | 0.56 | 3.37 | 1.52 | 4.75 | 0.2 | 2.36 | 1.35 | 0 | 1.69 | 0.67 | 1.69 | 0.67 | 0 | 0.98 | 0.67 | 0.27 | 0 | 2.31 | 1.06
| Analytical results for Cs in OREAS 606b (Certified Value 1.48 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [5 empty] | 2
| 1 | 1 | 1.65 | 1.54 | 1.53 | 1.4835 | 1.44 | 1.44 | 1.7523 | 1.38 | 1.84 | 1.69 | 1.24 | 1.32 | 1.44 | 1.59 | 1.38 | 1.41 | 1.3 | 1.63 | 1.03 | 1.2852 | 1.57 | [5 empty] | 1
| 1 | 2 | 1.52 | 1.52 | 1.49 | 1.4666 | 1.43 | 1.45 | 1.7282 | 1.45 | 1.74 | 1.76 | 1.37 | 1.33 | 1.5 | 1.54 | 1.35 | 1.43 | 1.23 | 1.57 | 1.11 | 1.2749 | 1.55 | [5 empty] | 31
| 1 | 3 | 1.57 | 1.5 | 1.49 | 1.4976 | 1.38 | 1.43 | 1.8605 | 1.46 | 1.76 | 1.72 | 1.36 | 1.35 | 1.49 | 1.43 | 1.4 | 1.3 | 1.31 | 1.49 | 1.05 | 1.3328 | 1.48 | [5 empty] | 16
| 1 | 4 | 1.66 | 1.4 | 1.53 | 1.4485 | 1.43 | 1.45 | 1.772 | 1.44 | 1.78 | 1.69 | 1.19 | 1.33 | 1.47 | 1.4 | 1.34 | 1.3 | 1.18 | 1.49 | 1 | 1.318 | 1.49 | [5 empty] | 1.4778
| 1 | 5 | 1.59 | 1.48 | 1.51 | 1.4473 | 1.46 | 1.65 | 1.7695 | 1.41 | 1.85 | 1.74 | 1.29 | 1.34 | 1.52 | 1.47 | 1.37 | 1.31 | 1.27 | 1.53 | 0.99 | 1.3074 | 1.44 | [5 empty] | 86
| 1 | 6 | 1.72 | 1.47 | 1.49 | 1.4856 | 1.45 | 1.47 | 1.6518 | 1.41 | 1.73 | 1.75 | 1.39 | 1.34 | 1.48 | 1.48 | 1.34 | 1.35 | 1.2 | 1.58 | 1.01 | 1.2927 | 1.37
| Mean |  | 1.6183 | 1.485 | 1.5067 | 1.4715 | 1.4317 | 1.4817 | 1.7557 | 1.425 | 1.7833 | 1.725 | 1.3067 | 1.335 | 1.4833 | 1.485 | 1.3633 | 1.35 | 1.2483 | 1.5483 | 1.0317 | 1.3018 | 1.4833
| Median |  | 1.62 | 1.49 | 1.5 | 1.4751 | 1.435 | 1.45 | 1.7609 | 1.425 | 1.77 | 1.73 | 1.325 | 1.335 | 1.485 | 1.475 | 1.36 | 1.33 | 1.25 | 1.55 | 1.02 | 1.3 | 1.485
| Std Dev. |  | 0.071949 | 0.048888 | 0.019664 | 0.020813 | 0.027869 | 0.083526 | 0.067827 | 0.030166 | 0.050859 | 0.030166 | 0.080166 | 0.010488 | 0.027325 | 0.070071 | 0.024221 | 0.057619 | 0.053448 | 0.055287 | 0.044008 | 0.021632 | 0.073121
| Rel.Std.Dev. |  | 4.446% | 3.292% | 1.305% | 1.414% | 1.947% | 5.637% | 3.863% | 2.117% | 2.852% | 1.749% | 6.135% | 0.7856% | 1.842% | 4.719% | 1.777% | 4.268% | 4.282% | 3.571% | 4.266% | 1.662% | 4.929%
| PDM3 |  | 9.512% | 0.4892% | 1.955% | -0.4227% | -3.12% | 0.2636% | 18.81% | -3.571% | 20.68% | 16.73% | -11.58% | -9.661% | 0.3764% | 0.4892% | -7.744% | -8.646% | -15.53% | 4.775% | -30.19% | -11.91% | 0.3764%
| Z-Score (Absolute) |  | 0.78 | 0.02 | 0.14 | 0.06 | 0.28 | 0 | 1.56 | 0.32 | 1.72 | 1.39 | 1 | 0.84 | 0.01 | 0.02 | 0.67 | 0.75 | 1.33 | 0.38 | 2.56 | 1.02 | 0.01
| Analytical results for Cu in OREAS 606b (Certified Value 0.034 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 0.0331 | 0.036125 | 0.03407 | 0.03483 | 0.034314 | 0.0351 | 0.0328 | 0.03271 | 0.0344 | 0.0346 | 0.0332 | 0.0333 | 0.0341 | 0.033 | 0.0369 | 0.0339 | 0.0365 | 0.0343 | 0.0354 | 0.02989 | 0.0322 | 0.034584 | 0.0326 | 0.033509 | 0.032133 | 0.0368 | 1
| 1 | 2 | 0.0329 | 0.033048 | 0.03404 | 0.03512 | 0.033583 | 0.0347 | 0.0327 | 0.032458 | 0.034 | 0.0339 | 0.0334 | 0.0334 | 0.0339 | 0.0333 | 0.0356 | 0.034 | 0.0365 | 0.0353 | 0.0341 | 0.0302 | 0.0323 | 0.031905 | 0.0328 | 0.033611 | 0.031867 | 0.0341 | 18
| 1 | 3 | 0.0332 | 0.037173 | 0.03409 | 0.03396 | 0.03351 | 0.0349 | 0.0323 | 0.032469 | 0.0345 | 0.0344 | 0.0335 | 0.0338 | 0.0337 | 0.0331 | 0.035 | 0.0345 | 0.0359 | 0.0333 | 0.0359 | 0.03077 | 0.0322 | 0.0338 | 0.0322 | 0.032981 | 0.032011 | 0.0323 | 16
| 1 | 4 | 0.0329 | 0.035695 | 0.03363 | 0.03488 | 0.033864 | 0.035 | 0.0326 | 0.032675 | 0.0341 | 0.0347 | 0.0327 | 0.0332 | 0.035 | 0.0333 | 0.0346 | 0.0332 | 0.0362 | 0.0336 | 0.0373 | 0.03042 | 0.0325 | 0.031569 | 0.0324 | 0.03354 | 0.032 | 0.0327 | 0.033727
| 1 | 5 | 0.0328 | 0.035906 | 0.03408 | 0.03423 | 0.033434 | 0.0371 | 0.0365 | 0.032979 | 0.034 | 0.0351 | 0.0332 | 0.0342 | 0.0345 | 0.0332 | 0.0344 | 0.0333 | 0.0371 | 0.0337 | 0.0342 | 0.03029 | 0.0322 | 0.033756 | 0.0322 | 0.032797 | 0.03207 | 0.0329 | 87
| 1 | 6 | 0.0332 | 0.036085 | 0.0331 | 0.03467 | 0.033924 | 0.0349 | 0.0338 | 0.032375 | 0.0338 | 0.034 | 0.033 | 0.0334 | 0.0338 | 0.0334 | 0.036 | 0.0333 | 0.0362 | 0.0341 | 0.0351 | 0.02998 | 0.0323 | 0.034727 | 0.0322 | 0.032807 | 0.031729 | 0.0304
| Mean |  | 0.033017 | 0.035672 | 0.033835 | 0.034615 | 0.033772 | 0.035283 | 0.03345 | 0.032611 | 0.034133 | 0.03445 | 0.033167 | 0.03355 | 0.034167 | 0.033217 | 0.035417 | 0.0337 | 0.0364 | 0.03405 | 0.035333 | 0.030258 | 0.032283 | 0.03339 | 0.0324 | 0.033207 | 0.031968 | 0.0332
| Median |  | 0.033 | 0.035995 | 0.034055 | 0.03475 | 0.033724 | 0.03495 | 0.03275 | 0.032572 | 0.03405 | 0.0345 | 0.0332 | 0.0334 | 0.034 | 0.03325 | 0.0353 | 0.0336 | 0.03635 | 0.0339 | 0.03525 | 0.030245 | 0.03225 | 0.033778 | 0.0323 | 0.033245 | 0.032005 | 0.0328
| Std Dev. |  | 0.00017224 | 0.0013834 | 0.00040114 | 0.00043629 | 0.00032965 | 0.00089981 | 0.0015783 | 0.00022306 | 0.00026583 | 0.00045056 | 0.00028752 | 0.00037815 | 0.00049666 | 0.0001472 | 0.00094322 | 0.00051769 | 0.00040988 | 0.00070922 | 0.0011877 | 0.00031796 | 0.0001169 | 0.0013445 | 0.00025298 | 0.00038583 | 0.00014669 | 0.0021335
| Rel.Std.Dev. |  | 0.5217% | 3.878% | 1.186% | 1.26% | 0.9761% | 2.55% | 4.718% | 0.684% | 0.7788% | 1.308% | 0.8669% | 1.127% | 1.454% | 0.4431% | 2.663% | 1.536% | 1.126% | 2.083% | 3.361% | 1.051% | 0.3621% | 4.027% | 0.7808% | 1.162% | 0.4589% | 6.426%
| PDM3 |  | -2.105% | 5.768% | 0.3214% | 2.634% | 0.1333% | 4.616% | -0.8202% | -3.307% | 1.206% | 2.145% | -1.66% | -0.5237% | 1.305% | -1.512% | 5.011% | -0.0789% | 7.927% | 0.9589% | 4.764% | -10.28% | -4.279% | -0.9977% | -3.933% | -1.54% | -5.213% | -1.561%
| Z-Score (Absolute) |  | 0.71 | 2.4 | 0.25 | 1.16 | 0.17 | 1.94 | 0.21 | 1.19 | 0.6 | 0.97 | 0.54 | 0.09 | 0.64 | 0.48 | 2.1 | 0.09 | 3.25 | 0.5 | 2 | 3.95 | 1.57 | 0.28 | 1.44 | 0.49 | 1.94 | 0.5
| Analytical results for Dy in OREAS 606b (Certified Value 2.08 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 3
| 1 | 1 | 2.01: 2.08; 2.03: 2.012; 2.04: 2.131; 2.05: 2.1; 2.06: 2.3475; 2.07: 2.3; 2.08: 1.96; 2.09: 2.1116; 2.10: 1.62; 2.11: 2.72 | 1
| 1 | 2 | 2.01: 1.88; 2.03: 1.986; 2.04: 2.0027; 2.05: 2.1; 2.06: 2.3507; 2.07: 2.24; 2.08: 2.01; 2.09: 2.0059; 2.10: 1.78; 2.11: 2.64 | 33
| 1 | 3 | 2.01: 2.07; 2.03: 1.98; 2.04: 2.1282; 2.05: 2; 2.06: 2.4746; 2.07: 2.24; 2.08: 1.85; 2.09: 1.9357; 2.10: 1.57; 2.11: 2.44 | 16
| 1 | 4 | 2.01: 1.96; 2.03: 2.003; 2.04: 2.0696; 2.05: 2.2; 2.06: 2.3708; 2.07: 2.16; 2.08: 1.85; 2.09: 1.8918; 2.10: 1.58; 2.11: 2.64 | 2.084
| 1 | 5 | 2.01: 1.91; 2.03: 1.998; 2.04: 2.0462; 2.05: 2.1; 2.06: 2.4865; 2.07: 2.26; 2.08: 1.86; 2.09: 1.9814; 2.10: 1.57; 2.11: 2.48 | 88
| 1 | 6 | 2.01: 1.94; 2.03: 2.035; 2.04: 2.0535; 2.05: 2.1; 2.06: 2.2648; 2.07: 2.27; 2.08: 1.85; 2.09: 2.2091; 2.10: 1.57; 2.11: 2.3
| Mean |  | 2.01: 1.9733; 2.03: 2.0023; 2.04: 2.0719; 2.05: 2.1; 2.06: 2.3825; 2.07: 2.245; 2.08: 1.8967; 2.09: 2.0226; 2.10: 1.615; 2.11: 2.5367
| Median |  | 2.01: 1.95; 2.03: 2.0005; 2.04: 2.0616; 2.05: 2.1; 2.06: 2.3607; 2.07: 2.25; 2.08: 1.855; 2.09: 1.9936; 2.10: 1.575; 2.11: 2.56
| Std Dev. |  | 2.01: 0.083347; 2.03: 0.019725; 2.04: 0.049902; 2.05: 0.063246; 2.06: 0.084291; 2.07: 0.047223; 2.08: 0.070333; 2.09: 0.11776; 2.10: 0.083126; 2.11: 0.15718
| Rel.Std.Dev. |  | 2.01: 4.224%; 2.03: 0.9851%; 2.04: 2.409%; 2.05: 3.012%; 2.06: 3.538%; 2.07: 2.103%; 2.08: 3.708%; 2.09: 5.822%; 2.10: 5.147%; 2.11: 6.196%
| PDM3 |  | 2.01: -5.308%; 2.03: -3.916%; 2.04: -0.5797%; 2.05: 0.7702%; 2.06: 14.33%; 2.07: 7.728%; 2.08: -8.987%; 2.09: -2.945%; 2.10: -22.5%; 2.11: 21.72%
| Z-Score (Absolute) |  | 2.01: 0.44; 2.03: 0.27; 2.04: 0.15; 2.05: 0.32; 2.06: 2.01; 2.07: 1.19; 2.08: 0.9; 2.09: 0.15; 2.10: 2.6; 2.11: 2.94
| Analytical results for Er in OREAS 606b (Certified Value 0.51 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 0.49; 2.03: 0.485; 2.04: 0.53443; 2.05: 0.5; 2.06: 0.6298; 2.07: 0.52; 2.08: 0.48; 2.09: 0.5859; 2.10: 0.43; 2.11: 0.6 | 1
| 1 | 2 | 2.01: 0.43; 2.03: 0.484; 2.04: 0.52551; 2.05: 0.5; 2.06: 0.6132; 2.07: 0.51; 2.08: 0.49; 2.09: 0.5375; 2.10: 0.45; 2.11: 0.56 | 34
| 1 | 3 | 2.01: 0.48; 2.03: 0.473; 2.04: 0.53003; 2.05: 0.5; 2.06: 0.6556; 2.07: 0.53; 2.08: 0.45; 2.09: 0.5164; 2.10: 0.42; 2.11: 0.53 | 16
| 1 | 4 | 2.01: 0.49; 2.03: 0.466; 2.04: 0.54034; 2.05: 0.5; 2.06: 0.5999; 2.07: 0.48; 2.08: 0.45; 2.09: 0.5131; 2.10: 0.41; 2.11: 0.65 | 0.51123
| 1 | 5 | 2.01: 0.43; 2.03: 0.458; 2.04: 0.55329; 2.05: 0.6; 2.06: 0.6361; 2.07: 0.53; 2.08: 0.44; 2.09: 0.5294; 2.10: 0.4; 2.11: 0.55 | 89
| 1 | 6 | 2.01: 0.49; 2.03: 0.465; 2.04: 0.51943; 2.05: 0.5; 2.06: 0.604; 2.07: 0.54; 2.08: 0.45; 2.09: 0.5915; 2.10: 0.41; 2.11: 0.47
| Mean |  | 2.01: 0.46833; 2.03: 0.47183; 2.04: 0.53384; 2.05: 0.51667; 2.06: 0.6231; 2.07: 0.51833; 2.08: 0.46; 2.09: 0.54563; 2.10: 0.42; 2.11: 0.56
| Median |  | 2.01: 0.485; 2.03: 0.4695; 2.04: 0.53223; 2.05: 0.5; 2.06: 0.6215; 2.07: 0.525; 2.08: 0.45; 2.09: 0.53345; 2.10: 0.415; 2.11: 0.555
| Std Dev. |  | 2.01: 0.029944; 2.03: 0.010907; 2.04: 0.011938; 2.05: 0.040825; 2.06: 0.021313; 2.07: 0.02137; 2.08: 0.02; 2.09: 0.034548; 2.10: 0.017889; 2.11: 0.061319
| Rel.Std.Dev. |  | 2.01: 6.394%; 2.03: 2.312%; 2.04: 2.236%; 2.05: 7.902%; 2.06: 3.42%; 2.07: 4.123%; 2.08: 4.348%; 2.09: 6.332%; 2.10: 4.259%; 2.11: 10.95%
| PDM3 |  | 2.01: -8.391%; 2.03: -7.706%; 2.04: 4.423%; 2.05: 1.063%; 2.06: 21.88%; 2.07: 1.389%; 2.08: -10.02%; 2.09: 6.729%; 2.10: -17.85%; 2.11: 9.54%
| Z-Score (Absolute) |  | 2.01: 0.73; 2.03: 0.67; 2.04: 0.22; 2.05: NA; 2.06: 1.52; 2.07: 0; 2.08: 0.85; 2.09: 0.4; 2.10: 1.43; 2.11: 0.6
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 606b (Certified Value 0.7 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 3
| 1 | 1 | 2.01: 0.72; 2.03: 0.682; 2.04: 0.59466; 2.05: 0.7; 2.06: 0.9656; 2.07: 0.76; 2.08: 0.75; 2.09: 0.7577; 2.10: 0.57; 2.11: 0.8 | 1
| 1 | 2 | 2.01: 0.64; 2.03: 0.68; 2.04: 0.59345; 2.05: 0.7; 2.06: 0.968; 2.07: 0.75; 2.08: 0.76; 2.09: 0.6997; 2.10: 0.61; 2.11: 0.78 | 35
| 1 | 3 | 2.01: 0.71; 2.03: 0.664; 2.04: 0.59174; 2.05: 0.6; 2.06: 0.9362; 2.07: 0.76; 2.08: 0.69; 2.09: 0.6651; 2.10: 0.56; 2.11: 0.71 | 16
| 1 | 4 | 2.01: 0.67; 2.03: 0.666; 2.04: 0.58862; 2.05: 0.7; 2.06: 0.9925; 2.07: 0.73; 2.08: 0.7; 2.09: 0.6455; 2.10: 0.54; 2.11: 0.78 | 0.69609
| 1 | 5 | 2.01: 0.65; 2.03: 0.681; 2.04: 0.59736; 2.05: 0.7; 2.06: 0.9278; 2.07: 0.79; 2.08: 0.7; 2.09: 0.6898; 2.10: 0.53; 2.11: 0.73 | 90
| 1 | 6 | 2.01: 0.71; 2.03: 0.669; 2.04: 0.59549; 2.05: 0.7; 2.06: 0.9591; 2.07: 0.77; 2.08: 0.71; 2.09: 0.7647; 2.10: 0.54; 2.11: 0.64
| Mean |  | 2.01: 0.68333; 2.03: 0.67367; 2.04: 0.59355; 2.05: 0.68333; 2.06: 0.9582; 2.07: 0.76; 2.08: 0.71833; 2.09: 0.70375; 2.10: 0.55833; 2.11: 0.74
| Median |  | 2.01: 0.69; 2.03: 0.6745; 2.04: 0.59406; 2.05: 0.7; 2.06: 0.96235; 2.07: 0.76; 2.08: 0.705; 2.09: 0.69475; 2.10: 0.55; 2.11: 0.755
| Std Dev. |  | 2.01: 0.034448; 2.03: 0.0082138; 2.04: 0.0030703; 2.05: 0.040825; 2.06: 0.023393; 2.07: 0.02; 2.08: 0.029269; 2.09: 0.048422; 2.10: 0.029269; 2.11: 0.059666
| Rel.Std.Dev. |  | 2.01: 5.041%; 2.03: 1.219%; 2.04: 0.5173%; 2.05: 5.974%; 2.06: 2.441%; 2.07: 2.632%; 2.08: 4.075%; 2.09: 6.881%; 2.10: 5.242%; 2.11: 8.063%
| PDM3 |  | 2.01: -1.833%; 2.03: -3.221%; 2.04: -14.73%; 2.05: -1.833%; 2.06: 37.65%; 2.07: 9.181%; 2.08: 3.195%; 2.09: 1.1%; 2.10: -19.79%; 2.11: 6.308%
| Z-Score (Absolute) |  | 2.01: 0.38; 2.03: 0.56; 2.04: 2.05; 2.05: NA; 2.06: 4.73; 2.07: 1.05; 2.08: 0.27; 2.09: 0; 2.10: 2.7; 2.11: 0.67
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 606b (Certified Value 2 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  | 3
| 1 | 1 | 2.09 | 1.96 | 1.99 | 1.97 | 2.0414 | 2.07 | 2.1 | 2.02 | 2.25 | 2.12 | 1.97 | 1.99 | 1.96 | 2.11 | 2.02 | 1.9722 | 2.06 | 1.93 | 1.87 | 1.83 | 1.9804 | 2.077 | 2.009 | 1.9658 |  |  | 1
| 1 | 2 | 2.09 | 1.93 | 1.99 | 1.93 | 2.0473 | 2.05 | 2.12 | 2.01 | 2.18 | 2.11 | 1.96 | 1.98 | 1.95 | 2.04 | 2 | 1.8703 | 2.09 | 1.95 | 1.87 | 1.8 | 1.9165 | 2.007 | 1.9557 | 1.9155 |  |  | 0
| 1 | 3 | 2.09 | 2.01 | 1.99 | 1.97 | 2.0741 | 2.05 | 2.12 | 1.99 | 2.25 | 2.14 | 1.94 | 2 | 1.96 | 2 | 2.05 | 1.9 | 2 | 1.9 | 1.9 | 1.82 | 1.9454 | 2.049 | 1.9567 | 1.9525 |  |  | 16
| 1 | 4 | 2.12 | 1.97 | 1.99 | 1.91 | 2.0658 | 2.09 | 2.09 | 2.02 | 2.24 | 2.07 | 1.96 | 2.01 | 1.97 | 1.96 | 1.95 | 1.9288 | 2.01 | 1.86 | 1.88 | 1.83 | 1.902 | 2.07 | 2.0133 | 1.9364 |  |  | 1.9992
| 1 | 5 | 2.07 | 2.08 | 1.99 | 1.95 | 2.0284 | 2.2 | 2.4 | 2.02 | 2.31 | 2.12 | 2.01 | 1.99 | 1.95 | 1.98 | 1.98 | 1.9723 | 2.03 | 1.95 | 1.9 | 1.81 | 1.9719 | 2.07 | 2.002 | 1.9338 |  |  | 91
| 1 | 6 | 2.09 | 1.95 | 1.97 | 1.96 | 2.0415 | 2.04 | 2.14 | 2.01 | 2.25 | 2.1 | 1.95 | 2 | 1.98 | 2.06 | 1.95 | 1.9249 | 0.56 | 1.96 | 1.88 | 1.81 | 1.9654 | 2 | 2.0171 | 1.9809
| Mean |  | 2.0917 | 1.9833 | 1.9867 | 1.9483 | 2.0497 | 2.0833 | 2.1617 | 2.0117 | 2.2467 | 2.11 | 1.965 | 1.995 | 1.9617 | 2.025 | 1.9917 | 1.9281 | 1.7917 | 1.925 | 1.8833 | 1.8167 | 1.9469 | 2.0455 | 1.9923 | 1.9475
| Median |  | 2.09 | 1.965 | 1.99 | 1.955 | 2.0444 | 2.06 | 2.12 | 2.015 | 2.25 | 2.115 | 1.96 | 1.995 | 1.96 | 2.02 | 1.99 | 1.9268 | 2.02 | 1.94 | 1.88 | 1.815 | 1.9554 | 2.0595 | 2.0055 | 1.9444
| Std Dev. |  | 0.016021 | 0.054283 | 0.008165 | 0.024014 | 0.01701 | 0.059889 | 0.11805 | 0.01169 | 0.041312 | 0.023664 | 0.02429 | 0.010488 | 0.01169 | 0.055767 | 0.039707 | 0.040099 | 0.6043 | 0.038341 | 0.013663 | 0.012111 | 0.031724 | 0.033934 | 0.02841 | 0.023726
| Rel.Std.Dev. |  | 0.7659% | 2.737% | 0.411% | 1.233% | 0.8299% | 2.875% | 5.461% | 0.5811% | 1.839% | 1.122% | 1.236% | 0.5257% | 0.5959% | 2.754% | 1.994% | 2.08% | 33.73% | 1.992% | 0.7254% | 0.6666% | 1.629% | 1.659% | 1.426% | 1.218%
| PDM3 |  | 4.628% | -0.7914% | -0.6247% | -2.542% | 2.53% | 4.211% | 8.129% | 0.6258% | 12.38% | 5.545% | -1.708% | -0.2079% | -1.875% | 1.293% | -0.3746% | -3.555% | -10.38% | -3.709% | -5.794% | -9.128% | -2.612% | 2.318% | -0.343% | -2.585%
| Z-Score (Absolute) |  | 1.4 | 0.08 | 0.03 | 0.56 | 0.83 | 1.28 | 2.35 | 0.31 | 3.51 | 1.65 | 0.33 | 0.08 | 0.37 | 0.49 | 0.03 | 0.83 | 2.69 | 0.87 | 1.44 | 2.35 | 0.58 | 0.77 | 0.04 | 0.57
| Analytical results for Ga in OREAS 606b (Certified Value 5.9 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 5.78 | 5.6 | 6 | 5.69 | 6.2544 | 6.54 | 6.39 | 6.26 | 5.9 | 6.23 | 5.06 | 6.27 | 5.53 | 5.95 | 5.25 | 5.63 | < 10 | 5.7 | 5.91 | 6.7561 | 4.38 | 5.6429 | 6.463 | 6.33 |  |  | 1
| 1 | 2 | 5.93 | 5.7 | 5 | 5.62 | 6.0896 | 6.53 | 6.49 | 6.45 | 6 | 6.25 | 5.39 | 6.28 | 5.76 | 5.75 | 5.18 | 5.79 | < 10 | 5.8 | 5.95 | 6.0589 | 4.47 | 5.5162 | 6.568 | 6.2 |  |  | 36
| 1 | 3 | 5.7 | 5.7 | 5 | 5.57 | 6.041 | 6.39 | 6.3 | 6.14 | 6 | 5.89 | 5.41 | 6.36 | 5.66 | 5.38 | 5.27 | 5.37 | < 10 | 5.4 | 5.51 | 5.9948 | 4.3 | 5.4733 | 6.572 | 5.74 |  |  | 16
| 1 | 4 | 5.88 | 5.5 | 6 | 5.5 | 6.2325 | 6.58 | 6.43 | 6.15 | 6 | 5.79 | 5.17 | 6.31 | 5.74 | 5.29 | 5.18 | 5.32 | < 10 | 6.8 | 5.52 | 5.8045 | 4.35 | 5.4481 | 6.643 | 6.08 |  |  | 5.8994
| 1 | 5 | 5.72 | 5.6 | 6 | 5.47 | 6.0857 | 6.59 | 7.18 | 6.33 | 5.9 | 6.18 | 5.25 | 6.27 | 5.81 | 5.55 | 5.13 | 5.25 | < 10 | 6.8 | 5.52 | 5.9672 | 4.28 | 5.566 | 7.013 | 5.79 |  |  | 92
| 1 | 6 | 5.85 | 5.9 | 5 | 5.59 | 6.0446 | 6.52 | 6.55 | 6.1 | 5.9 | 6.2 | 5.45 | 6.33 | 5.72 | 5.53 | 5.23 | 5.39 | < 10 | 5.5 | 5.8 | 6.8459 | 4.23 | 5.6219 | 6.374 | 5.46
| Mean |  | 5.81 | 5.6667 | 5.5 | 5.5733 | 6.1246 | 6.525 | 6.5567 | 6.2383 | 5.95 | 6.09 | 5.2883 | 6.3033 | 5.7033 | 5.575 | 5.2067 | 5.4583 |  | 6 | 5.7017 | 6.2379 | 4.335 | 5.5447 | 6.6055 | 5.9333
| Median |  | 5.815 | 5.65 | 5.5 | 5.58 | 6.0877 | 6.535 | 6.46 | 6.205 | 5.95 | 6.19 | 5.32 | 6.295 | 5.73 | 5.54 | 5.205 | 5.38 |  | 5.75 | 5.66 | 6.0268 | 4.325 | 5.5411 | 6.57 | 5.935
| Std Dev. |  | 0.091652 | 0.13663 | 0.54772 | 0.080166 | 0.094457 | 0.071764 | 0.31709 | 0.13438 | 0.054772 | 0.19769 | 0.1542 | 0.036697 | 0.098116 | 0.24214 | 0.052409 | 0.20712 |  | 0.63561 | 0.20856 | 0.44507 | 0.084558 | 0.079111 | 0.22086 | 0.32617
| Rel.Std.Dev. |  | 1.577% | 2.411% | 9.959% | 1.438% | 1.542% | 1.1% | 4.836% | 2.154% | 0.9205% | 3.246% | 2.916% | 0.5822% | 1.72% | 4.343% | 1.007% | 3.794% |  | 10.59% | 3.658% | 7.135% | 1.951% | 1.427% | 3.344% | 5.497%
| PDM3 |  | -1.515% | -3.944% | -6.77% | -5.527% | 3.819% | 10.61% | 11.14% | 5.746% | 0.8584% | 3.231% | -10.36% | 6.848% | -3.323% | -5.498% | -11.74% | -7.476% |  | 1.706% | -3.351% | 5.739% | -26.52% | -6.011% | 11.97% | 0.5758%
| Z-Score (Absolute) |  | 0 | 0.31 | NA | 0.51 | 0.67 | 1.53 | 1.6 | 0.92 | 0.3 | 0.6 | 1.12 | 1.06 | 0.23 | 0.5 | 1.29 | 0.75 | 1.74 | 0.41 | 0.23 | 0.92 | 3.16 | 0.57 | 1.7 | 0.26
Results from laboratory 2.04 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 606b (Certified Value 3.82 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  | 3
| 1 | 1 | 2.01: 3.8; 2.03: 3.653; 2.04: 3.9091; 2.05: 4; 2.06: 5.4511; 2.07: 4.02; 2.08: 3.74; 2.10: 4.7828; 2.11: 3.09; 2.12: 3.85 |  |  | 1
| 1 | 2 | 2.01: 3.6; 2.03: 3.569; 2.04: 3.8814; 2.05: 4; 2.06: 5.4998; 2.07: 3.98; 2.08: 4; 2.10: 4.3083; 2.11: 3.47; 2.12: 3.78 |  |  | 37
| 1 | 3 | 2.01: 3.75; 2.03: 3.505; 2.04: 3.9125; 2.05: 3.9; 2.06: 5.7592; 2.07: 3.96; 2.08: 3.59; 2.10: 4.1771; 2.11: 2.88; 2.12: 3.48 |  |  | 16
| 1 | 4 | 2.01: 3.65; 2.03: 3.671; 2.04: 3.8927; 2.05: 4.1; 2.06: 5.6298; 2.07: 3.84; 2.08: 3.56; 2.10: 4.0739; 2.11: 2.84; 2.12: 3.63 |  |  | 3.8181
| 1 | 5 | 2.01: 3.52; 2.03: 3.629; 2.04: 3.8839; 2.05: 4.2; 2.06: 5.636; 2.07: 3.93; 2.08: 3.69; 2.10: 4.2907; 2.11: 2.81; 2.12: 3.47 |  |  | 93
| 1 | 6 | 2.01: 3.59; 2.03: 3.591; 2.04: 3.8961; 2.05: 3.9; 2.06: 5.3946; 2.07: 3.98; 2.08: 3.63; 2.10: 4.81; 2.11: 2.86; 2.12: 3.22
| Mean |  | 2.01: 3.6517; 2.03: 3.603; 2.04: 3.8959; 2.05: 4.0167; 2.06: 5.5617; 2.07: 3.9517; 2.08: 3.7017; 2.10: 4.4071; 2.11: 2.9917; 2.12: 3.5717
| Median |  | 2.01: 3.625; 2.03: 3.61; 2.04: 3.8944; 2.05: 4; 2.06: 5.5648; 2.07: 3.97; 2.08: 3.66; 2.10: 4.2995; 2.11: 2.87; 2.12: 3.555
| Std Dev. |  | 2.01: 0.10534; 2.03: 0.061162; 2.04: 0.012752; 2.05: 0.1169; 2.06: 0.13649; 2.07: 0.062102; 2.08: 0.16018; 2.10: 0.31328; 2.11: 0.25467; 2.12: 0.23095
| Rel.Std.Dev. |  | 2.01: 2.885%; 2.03: 1.698%; 2.04: 0.3273%; 2.05: 2.91%; 2.06: 2.454%; 2.07: 1.572%; 2.08: 4.327%; 2.10: 7.108%; 2.11: 8.513%; 2.12: 6.466%
| PDM3 |  | 2.01: -4.36%; 2.03: -5.635%; 2.04: 2.038%; 2.05: 5.2%; 2.06: 45.67%; 2.07: 3.497%; 2.08: -3.051%; 2.10: 15.43%; 2.11: -21.65%; 2.12: -6.455%
| Z-Score (Absolute) |  | 2.01: 0.48; 2.03: 0.64; 2.04: 0.32; 2.05: 0.71; 2.06: 5.75; 2.07: 0.5; 2.08: 0.32; 2.10: 1.98; 2.11: 2.63; 2.12: 0.74
| Analytical results for Ge in OREAS 606b (Certified Value 0.12 ppm) | [27 empty] | Certified
Ge | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab; 2.15: Lab; 2.16: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.03: 2.03; 2.04: 2.06; 2.05: 2.07; 2.06: 2.08; 2.07: 2.13; 2.08: 2.14; 2.10: 2.15; 2.11: 2.16; 2.12: 2.17; 2.13: 2.19; 2.14: 2.20; 2.15: 2.22; 2.16: 2.25 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*MS; 2.15: AR*MS; 2.16: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.4g; 2.14: 0.5g; 2.15: 0.2g; 2.16: 0.5g |  |  | 2
| 1 | 1 | 2.01: 0.18; 2.03: 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.25; 2.07: 0.08; 2.08: 0.1; 2.10: 0.11; 2.11: 0.12; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.14; 2.15: 0.08; 2.16: < 0.01 |  |  | 1
| 1 | 2 | 2.01: 0.2; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.23; 2.07: 0.06; 2.08: 0.12; 2.10: 0.11; 2.11: 0.12; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.13; 2.15: 0.09; 2.16: < 0.01 |  |  | 38
| 1 | 3 | 2.01: 0.19; 2.03: 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.22; 2.07: 0.08; 2.08: 0.11; 2.10: 0.11; 2.11: 0.12; 2.12: 0.19; 2.13: < 0.05; 2.14: 0.11; 2.15: 0.08; 2.16: < 0.01 |  |  | 16
| 1 | 4 | 2.01: 0.19; 2.03: 0.2; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.25; 2.07: 0.08; 2.08: 0.11; 2.10: 0.12; 2.11: 0.11; 2.12: 0.16; 2.13: < 0.05; 2.14: 0.11; 2.15: 0.07; 2.16: < 0.01 |  |  | 0.12292
| 1 | 5 | 2.01: 0.19; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.27; 2.07: 0.09; 2.08: 0.11; 2.10: 0.11; 2.11: 0.12; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.11; 2.15: 0.08; 2.16: < 0.01 |  |  | 94
| 1 | 6 | 2.01: 0.18; 2.03: 0.2; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.23; 2.07: 0.09; 2.08: 0.11; 2.10: 0.12; 2.11: 0.12; 2.12: 0.18; 2.13: < 0.05; 2.14: 0.12; 2.15: 0.08; 2.16: < 0.01
| Mean |  | 2.01: 0.18833; 2.03: 0.15; 2.06: 0.24167; 2.07: 0.08; 2.08: 0.11; 2.10: 0.11333; 2.11: 0.11833; 2.12: 0.17333; 2.14: 0.12; 2.15: 0.08
| Median |  | 2.01: 0.19; 2.03: 0.15; 2.06: 0.24; 2.07: 0.08; 2.08: 0.11; 2.10: 0.11; 2.11: 0.12; 2.12: 0.17; 2.14: 0.115; 2.15: 0.08
| Std Dev. |  | 2.01: 0.0075277; 2.03: 0.057735; 2.06: 0.018348; 2.07: 0.010954; 2.08: 0.0063246; 2.10: 0.005164; 2.11: 0.0040825; 2.12: 0.010328; 2.14: 0.012649; 2.15: 0.0063246
| Rel.Std.Dev. |  | 2.01: 3.997%; 2.03: 38.49%; 2.06: 7.592%; 2.07: 13.69%; 2.08: 5.75%; 2.10: 4.556%; 2.11: 3.45%; 2.12: 5.958%; 2.14: 10.54%; 2.15: 7.906%
| PDM3 |  | 2.01: 53.22%; 2.03: 22.03%; 2.06: 96.61%; 2.07: -34.92%; 2.08: -10.51%; 2.10: -7.797%; 2.11: -3.729%; 2.12: 41.02%; 2.14: -2.373%; 2.15: -34.92%
| Z-Score (Absolute) |  | 2.01: 1.76; 2.03: NA; 2.04: 0.22; 2.05: 1.35; 2.06: 2.96; 2.07: 0.67; 2.08: 0; 2.10: 0.07; 2.11: 0.19; 2.12: 1.42; 2.13: 1.91; 2.14: 0.22; 2.15: 0.67; 2.16: 2.36
Results from laboratory 2.03 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 606b (Certified Value 1.22 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 3
| 1 | 1 | 1.59 | 1.26 | 1.23 | 1.2051 | 0.9 | 1.15 | 1.25 | 1.19 | 1.23 | 1.16 | 1.26 | 1.18 | 1.28 | 1.13 | 1.43 | 1.36 | 1.16 | 0.87 | 1.2045 | 1.55 | [6 empty] | 1
| 1 | 2 | 1.56 | 1.17 | 1.26 | 1.2832 | 0.3 | 1.18 | 1.32 | 1.16 | 1.25 | 1.22 | 1.3 | 1.18 | 1.27 | 1.09 | 1.43 | 1.38 | 1.06 | 0.93 | 1.1722 | 1.32 | [6 empty] | 22
| 1 | 3 | 1.55 | 1.27 | 1.23 | 1.2028 | 0.7 | 1.15 | 1.32 | 1.22 | 1.27 | 1.27 | 1.24 | 1.22 | 1.19 | 1.13 | 1.29 | 1.19 | 1.11 | 0.86 | 1.1714 | 0.78 | [6 empty] | 16
| 1 | 4 | 1.57 | 1.18 | 1.23 | 1.1983 | 0.5 | 1.14 | 1.25 | 1.19 | 1.18 | 1.18 | 1.26 | 1.2 | 1.22 | 1.11 | 1.28 | 1.39 | 1.1 | 0.83 | 1.2087 | 1.3 | [6 empty] | 1.2159
| 1 | 5 | 1.46 | 1.12 | 1.24 | 1.1871 | 0.7 | 1.32 | 1.23 | 1.16 | 1.29 | 1.17 | 1.27 | 1.22 | 1.22 | 1.12 | 1.32 | 1.41 | 1.08 | 0.8 | 1.1519 | 0.87 | [6 empty] | 95
| 1 | 6 | 1.66 | 1.23 | 1.22 | 1.1941 | 0.4 | 1.2 | 1.28 | 1.19 | 1.26 | 1.26 | 1.26 | 1.22 | 1.2 | 1.11 | 1.34 | 1.39 | 1.16 | 0.84 | 1.1667 | 1.25
| Mean |  | 1.565 | 1.205 | 1.235 | 1.2118 | 0.58333 | 1.19 | 1.275 | 1.185 | 1.2467 | 1.21 | 1.265 | 1.2033 | 1.23 | 1.115 | 1.3483 | 1.3533 | 1.1117 | 0.855 | 1.1792 | 1.1783
| Median |  | 1.565 | 1.205 | 1.23 | 1.2006 | 0.6 | 1.165 | 1.265 | 1.19 | 1.255 | 1.2 | 1.26 | 1.21 | 1.22 | 1.115 | 1.33 | 1.385 | 1.105 | 0.85 | 1.1718 | 1.275
| Std Dev. |  | 0.06473 | 0.058224 | 0.013784 | 0.035599 | 0.22286 | 0.067528 | 0.038341 | 0.022583 | 0.038297 | 0.047329 | 0.019748 | 0.019664 | 0.036878 | 0.015166 | 0.066758 | 0.08165 | 0.041191 | 0.044159 | 0.022437 | 0.2939
| Rel.Std.Dev. |  | 4.136% | 4.832% | 1.116% | 2.938% | 38.2% | 5.675% | 3.007% | 1.906% | 3.072% | 3.911% | 1.561% | 1.634% | 2.998% | 1.36% | 4.951% | 6.033% | 3.705% | 5.165% | 1.903% | 24.94%
| PDM3 |  | 28.72% | -0.8929% | 1.574% | -0.3358% | -52.02% | -2.127% | 4.864% | -2.538% | 2.534% | -0.4817% | 4.042% | -1.03% | 1.163% | -8.295% | 10.9% | 11.31% | -8.569% | -29.68% | -3.012% | -3.086%
| Z-Score (Absolute) |  | 7.06 | 0.05 | 0.54 | 0.08 | 12.32 | 0.35 | 1.33 | 0.44 | 0.77 | 0.05 | 1.13 | 0.08 | 0.44 | 1.83 | 2.78 | 2.88 | 1.89 | 6.96 | 0.56 | 0.58
| Analytical results for Hg in OREAS 606b (Certified Value 0.048 ppm) | [27 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 0.5g | 15g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 3
| 1 | 1 | 0.07 | 0.068 | 0.05 | < 1 | 0.05 | 0.05 |  | 0.03 | 0.04 | 0.04 | 0.05 | 0.06 | 0.05 | < 0.05 | 0.019 | < 5 | < 2 | < 0.01 | < 0.05 | 0.06 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.048 | 0.04 | < 1 | 0.04 | 0.05 |  | 0.02 | 0.05 | 0.05 | 0.05 | 0.06 | 0.06 | < 0.05 | 0.02 | < 5 | < 2 | < 0.01 | < 0.05 | 0.06 | [6 empty] | 3
| 1 | 3 | < 0.01 | 0.041 | 0.05 | < 1 | 0.04 | 0.05 |  | 0.05 | 0.04 | 0.05 | 0.05 | 0.06 | 0.05 | < 0.05 | 0.024 | < 5 | < 2 | < 0.01 | < 0.05 | 0.06 | [6 empty] | 16
| 1 | 4 | 0.01 | 0.04 | 0.04 | < 1 | 0.05 | 0.05 | 0.05 | 0.03 | 0.05 | 0.06 | 0.04 | 0.05 | 0.05 | < 0.05 | 0.019 | < 5 | < 2 | < 0.01 | < 0.05 | 0.09 | [6 empty] | 0.048397
| 1 | 5 | < 0.01 | 0.05 | 0.05 | < 1 | 0.06 | 0.06 |  | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.05 | < 0.05 | 0.02 | < 5 | < 2 | < 0.01 | < 0.05 | 0.04 | [6 empty] | 96
| 1 | 6 | 0.04 | 0.036 | 0.04 | < 1 | 0.04 | 0.06 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 | < 0.05 | 0.02 | < 5 | < 2 | < 0.01 | < 0.05 | 0.04
| Mean |  | 0.04 | 0.047167 | 0.045 |  | 0.046667 | 0.053333 | 0.05 | 0.036667 | 0.043333 | 0.046667 | 0.046667 | 0.055 | 0.051667 |  | 0.020333 |  |  |  |  | 0.058333
| Median |  | 0.04 | 0.0445 | 0.045 |  | 0.045 | 0.05 | 0.05 | 0.035 | 0.04 | 0.045 | 0.05 | 0.055 | 0.05 |  | 0.02 |  |  |  |  | 0.06
| Std Dev. |  | 0.03 | 0.011462 | 0.0054772 |  | 0.008165 | 0.005164 | 0 | 0.012111 | 0.005164 | 0.008165 | 0.005164 | 0.0054772 | 0.0040825 |  | 0.0018619 |  |  |  |  | 0.018348
| Rel.Std.Dev. |  | 75% | 24.3% | 12.17% |  | 17.5% | 9.682% | 0% | 33.03% | 11.92% | 17.5% | 11.07% | 9.959% | 7.902% |  | 9.157% |  |  |  |  | 31.45%
| PDM3 |  | -17.35% | -2.543% | -7.02% |  | -3.576% | 10.2% | 3.311% | -24.24% | -10.46% | -3.576% | -3.576% | 13.64% | 6.755% |  | -57.99% |  |  |  |  | 20.53%
| Z-Score (Absolute) |  | 1.78 | 0.04 | 0.12 | 33.35 | 0 | 0.49 | 0.25 | 0.74 | 0.25 | 0 | 0 | 0.61 | 0.37 | 1.59 | 1.94 | 180.47 | 70.13 | 3.07 | 1.59 | 0.86
| Analytical results for Ho in OREAS 606b (Certified Value 0.25 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab; 2.15: Lab | [6 empty] | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.12: 2.19; 2.13: 2.23; 2.14: 2.25; 2.15: 2.28 | [6 empty] | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*MS; 2.15: AR*MS | [6 empty] | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.06: 0.5g; 2.07: 0.25g; 2.08: 01g; 2.12: 0.4g; 2.13: 0.5g; 2.14: 0.5g; 2.15: 25g | [6 empty] | 3
| 1 | 1 | 2.01: 0.26; 2.03: 0.261; 2.04: 0.2515; 2.06: 0.3; 2.07: 0.2861; 2.08: 0.26; 2.12: 0.26; 2.13: 0.2369; 2.14: 0.22; 2.15: 0.27 | [6 empty] | 1
| 1 | 2 | 2.01: 0.24; 2.03: 0.255; 2.04: 0.24709; 2.06: 0.3; 2.07: 0.3004; 2.08: 0.26; 2.12: 0.27; 2.13: 0.22; 2.14: 0.24; 2.15: 0.27 | [6 empty] | 23
| 1 | 3 | 2.01: 0.23; 2.03: 0.25; 2.04: 0.2405; 2.06: 0.3; 2.07: 0.3256; 2.08: 0.27; 2.12: 0.25; 2.13: 0.2124; 2.14: 0.22; 2.15: 0.26 | [6 empty] | 16
| 1 | 4 | 2.01: 0.24; 2.03: 0.261; 2.04: 0.24114; 2.06: 0.3; 2.07: 0.3052; 2.08: 0.25; 2.12: 0.25; 2.13: 0.2139; 2.14: 0.21; 2.15: 0.36 | [6 empty] | 0.24504
| 1 | 5 | 2.01: 0.23; 2.03: 0.254; 2.04: 0.23986; 2.06: 0.3; 2.07: 0.3125; 2.08: 0.26; 2.12: 0.25; 2.13: 0.2206; 2.14: 0.21; 2.15: 0.25 | [6 empty] | 97
| 1 | 6 | 2.01: 0.25; 2.03: 0.264; 2.04: 0.25302; 2.06: 0.3; 2.07: 0.2895; 2.08: 0.26; 2.12: 0.25; 2.13: 0.2441; 2.14: 0.22; 2.15: 0.23
| Mean |  | 2.01: 0.24167; 2.03: 0.2575; 2.04: 0.24552; 2.06: 0.3; 2.07: 0.30322; 2.08: 0.26; 2.12: 0.255; 2.13: 0.22465; 2.14: 0.22; 2.15: 0.27333
| Median |  | 2.01: 0.24; 2.03: 0.258; 2.04: 0.24411; 2.06: 0.3; 2.07: 0.3028; 2.08: 0.26; 2.12: 0.25; 2.13: 0.2203; 2.14: 0.22; 2.15: 0.265
| Std Dev. |  | 2.01: 0.01169; 2.03: 0.0053198; 2.04: 0.0058488; 2.06: 0; 2.07: 0.014698; 2.08: 0.0063246; 2.12: 0.0083666; 2.13: 0.0129; 2.14: 0.010954; 2.15: 0.045019
| Rel.Std.Dev. |  | 2.01: 4.837%; 2.03: 2.066%; 2.04: 2.382%; 2.06: 0%; 2.07: 4.847%; 2.08: 2.433%; 2.12: 3.281%; 2.13: 5.742%; 2.14: 4.979%; 2.15: 16.47%
| PDM3 |  | 2.01: -1.377%; 2.03: 5.084%; 2.04: 0.194%; 2.06: 22.43%; 2.07: 23.74%; 2.08: 6.104%; 2.12: 4.064%; 2.13: -8.322%; 2.14: -10.22%; 2.15: 11.55%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.13; 2.04: 0.48; 2.06: NA; 2.07: 2.44; 2.08: 0.25; 2.12: 0; 2.13: 1.53; 2.14: 1.77; 2.15: 0.93
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 606b (Certified Value 0.1 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 01g | 25g | [8 empty] | 3
| 1 | 1 | 0.127 | 0.1 | 0.1 | 0.11 | 0.102 | 0.1 | 0.11 | 0.11 | 0.097 | 0.111 | 0.104 | 0.111 | 0.108 | 0.112 | 0.107 | 0.005 | 0.10443 | 0.103 | [8 empty] | 1
| 1 | 2 | 0.123 | 0.1 | 0.1 | 0.11 | 0.102 | 0.1 | 0.11 | 0.11 | 0.115 | 0.11 | 0.115 | 0.103 | 0.109 | 0.104 | 0.089 | 0.005 | 0.095884 | 0.102 | [8 empty] | 24
| 1 | 3 | 0.132 | 0.1 | 0.1 | 0.1 | 0.105 | 0.1 | 0.1 | 0.11 | 0.1 | 0.108 | 0.11 | 0.098 | 0.112 | 0.102 | 0.105 | 0.007 | 0.098004 | 0.094 | [8 empty] | 16
| 1 | 4 | 0.124 | 0.1 | 0.1 | 0.11 | 0.105 | 0.1 | 0.11 | 0.11 | 0.098 | 0.145 | 0.108 | 0.097 | 0.101 | 0.108 | 0.089 | 0.006 | 0.10326 | 0.096 | [8 empty] | 0.10472
| 1 | 5 | 0.141 | 0.1 | 0.1 | 0.11 | 0.116 | 0.1 | 0.11 | 0.11 | 0.106 | 0.108 | 0.111 | 0.102 | 0.107 | 0.105 | 0.104 | 0.007 | 0.10212 | 0.094 | [8 empty] | 98
| 1 | 6 | 0.127 | 0.1 | 0.1 | 0.1 | 0.111 | 0.1 | 0.1 | 0.1 | 0.113 | 0.112 | 0.107 | 0.102 | 0.111 | 0.102 | 0.113 | 0.006 | 0.098676 | 0.09
| Mean |  | 0.129 | 0.1 | 0.1 | 0.10667 | 0.10683 | 0.1 | 0.10667 | 0.10833 | 0.10483 | 0.11567 | 0.10917 | 0.10217 | 0.108 | 0.1055 | 0.10117 | 0.006 | 0.1004 | 0.0965
| Median |  | 0.127 | 0.1 | 0.1 | 0.11 | 0.105 | 0.1 | 0.11 | 0.11 | 0.103 | 0.1105 | 0.109 | 0.102 | 0.1085 | 0.1045 | 0.1045 | 0.006 | 0.1004 | 0.095
| Std Dev. |  | 0.0066633 | 1.5202e-17 | 1.5202e-17 | 0.005164 | 0.0055648 | 1.5202e-17 | 0.005164 | 0.0040825 | 0.0077825 | 0.014459 | 0.0037639 | 0.0049565 | 0.0038987 | 0.0038859 | 0.0099281 | 0.00089443 | 0.0033632 | 0.0050498
| Rel.Std.Dev. |  | 5.165% | 1.52e-14% | 1.52e-14% | 4.841% | 5.209% | 1.52e-14% | 4.841% | 3.768% | 7.424% | 12.5% | 3.448% | 4.851% | 3.61% | 3.683% | 9.814% | 14.91% | 3.35% | 5.233%
| PDM3 |  | 23.19% | -4.504% | -4.504% | 1.862% | 2.021% | -4.504% | 1.862% | 3.454% | 0.1116% | 10.46% | 4.25% | -2.435% | 3.136% | 0.7482% | -3.39% | -94.27% | -4.124% | -7.846%
| Z-Score (Absolute) |  | 4.42 | 1.17 | NA | 0.11 | 0.14 | NA | 0.11 | 0.43 | 0.24 | 1.85 | 0.59 | 0.75 | 0.37 | 0.11 | 0.95 | 19.28 | 1.1 | 1.85
Results from laboratory 2.05 were removed due to their 0.01 ppm reading resolution. Results from laboratory 2.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 606b (Certified Value 0.235 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.24 | 0.33 | 0.23 | 0.2401 | 0.24321 | 0.26 | 0.22 | 0.24 | 0.318 | 0.3644 | 0.22 | 0.25 | 0.23 | 0.28 | 0.24 | 0.248 | 0.23 | 0.24 | 0.26 | 0.2198 | 0.2017 | 0.23279 | 0.2875 |  |  |  | 1
| 1 | 2 | 0.23 | 0.23 | 0.23 | 0.2391 | 0.23813 | 0.26 | 0.22 | 0.24 | 0.331 | 0.3699 | 0.23 | 0.24 | 0.22 | 0.27 | 0.24 | 0.245 | 0.248 | 0.24 | 0.26 | 0.2035 | 0.1943 | 0.21655 | 0.2894 |  |  |  | 0
| 1 | 3 | 0.24 | 0.37 | 0.23 | 0.242 | 0.23915 | 0.26 | 0.22 | 0.23 | 0.317 | 0.3716 | 0.22 | 0.23 | 0.23 | 0.25 | 0.24 | 0.226 | 0.23 | 0.24 | 0.25 | 0.2059 | 0.1992 | 0.23451 | 0.28887 |  |  |  | 16
| 1 | 4 | 0.24 | 0.34 | 0.23 | 0.2362 | 0.24108 | 0.28 | 0.22 | 0.24 | 0.328 | 0.3569 | 0.22 | 0.28 | 0.22 | 0.25 | 0.23 | 0.225 | 0.23 | 0.25 | 0.26 | 0.1995 | 0.2026 | 0.23963 | 0.28909 |  |  |  | 0.23532
| 1 | 5 | 0.24 | 0.24 | 0.23 | 0.2385 | 0.23638 | 0.26 | 0.25 | 0.24 | 0.325 | 0.3787 | 0.22 | 0.25 | 0.22 | 0.26 | 0.23 | 0.231 | 0.24 | 0.24 | 0.25 | 0.1997 | 0.2017 | 0.23042 | 0.28531 |  |  |  | 99
| 1 | 6 | 0.24 | 0.37 | 0.23 | 0.2389 | 0.24043 | 0.27 | 0.23 | 0.24 | 0.32 | 0.3693 | 0.23 | 0.25 | 0.23 | 0.27 | 0.23 | 0.064 | 0.181 | 0.24 | 0.26 | 0.2106 | 0.1934 | 0.21193 | 0.29209
| Mean |  | 0.23833 | 0.31333 | 0.23 | 0.23913 | 0.23973 | 0.265 | 0.22667 | 0.23833 | 0.32317 | 0.36847 | 0.22333 | 0.25 | 0.225 | 0.26333 | 0.235 | 0.2065 | 0.2265 | 0.24167 | 0.25667 | 0.2065 | 0.19882 | 0.22764 | 0.28871
| Median |  | 0.24 | 0.335 | 0.23 | 0.239 | 0.23979 | 0.26 | 0.22 | 0.24 | 0.3225 | 0.3696 | 0.22 | 0.25 | 0.225 | 0.265 | 0.235 | 0.2285 | 0.23 | 0.24 | 0.26 | 0.2047 | 0.20045 | 0.2316 | 0.28898
| Std Dev. |  | 0.0040825 | 0.062823 | 0 | 0.0019086 | 0.0023898 | 0.0083666 | 0.012111 | 0.0040825 | 0.0057067 | 0.0073159 | 0.005164 | 0.016733 | 0.0054772 | 0.012111 | 0.0054772 | 0.070475 | 0.023458 | 0.0040825 | 0.005164 | 0.007724 | 0.0040207 | 0.010908 | 0.0022394
| Rel.Std.Dev. |  | 1.713% | 20.05% | 0% | 0.7981% | 0.9969% | 3.157% | 5.343% | 1.713% | 1.766% | 1.986% | 2.312% | 6.693% | 2.434% | 4.599% | 2.331% | 34.13% | 10.36% | 1.689% | 2.012% | 3.74% | 2.022% | 4.792% | 0.7757%
| PDM3 |  | 1.28% | 33.15% | -2.261% | 1.62% | 1.874% | 12.61% | -3.677% | 1.28% | 37.33% | 56.58% | -5.094% | 6.238% | -4.386% | 11.9% | -0.1362% | -12.25% | -3.748% | 2.697% | 9.071% | -12.25% | -15.51% | -3.265% | 22.69%
| Z-Score (Absolute) |  | 0 | 3.79 | 0.42 | 0.04 | 0.07 | 1.35 | 0.59 | 0 | 4.29 | 6.58 | 0.76 | 0.59 | 0.67 | 1.26 | 0.17 | 1.61 | 0.6 | 0.17 | 0.93 | 1.61 | 2 | 0.54 | 2.55
| Analytical results for La in OREAS 606b (Certified Value 27.9 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 2
| 1 | 1 | 34.4 | 26.3 | 29 | 28.707 | 25.0636 | 30.6 | 26.7 | 33.183 | 27.3 | 31.1 | 31.51 | 26 | 27.1 | 28 | 31 | 28.6 | 28.3 | 25.1 | 24.9 | 28.3 | 28.4263 | 20 | 26.5621 | 28.822 | 32.97 |  | 1
| 1 | 2 | 32.7 | 26.4 | 28 | 28.666 | 25.0097 | 30.3 | 27.1 | 33.349 | 27.5 | 31.5 | 31.51 | 28.5 | 27.3 | 29 | 30.1 | 28.4 | 28.8 | 25.9 | 25 | 28.2 | 25.9535 | 20 | 26.5314 | 28.562 | 31.84 |  | 25
| 1 | 3 | 33.5 | 26.3 | 28 | 28.225 | 24.7485 | 29.7 | 26.4 | 33.939 | 27.2 | 31.7 | 31.94 | 28 | 27.5 | 29.1 | 27.4 | 29.5 | 25.8 | 23.8 | 24.6 | 26.2 | 25.7782 | 20 | 27.0831 | 28.439 | 30.03 |  | 16
| 1 | 4 | 34.8 | 25.4 | 29 | 28.895 | 25.3735 | 31.6 | 26.9 | 33.098 | 27.3 | 32.1 | 29.94 | 25.9 | 26.9 | 28.7 | 27.6 | 27.9 | 25.9 | 25.4 | 25.2 | 26 | 24.3594 | 20 | 27.1358 | 28.658 | 30.85 |  | 27.9003
| 1 | 5 | 34.2 | 26.3 | 28 | 28.494 | 24.7693 | 31.8 | 30.8 | 33.248 | 27.3 | 32.5 | 31.52 | 26.9 | 27.3 | 28.7 | 28.1 | 28.4 | 26.6 | 25.8 | 25.1 | 26.2 | 25.616 | 21 | 26.7773 | 28.825 | 29.39 |  | 100
| 1 | 6 | 34.9 | 26.4 | 28 | 29.032 | 25.3961 | 30.4 | 28.1 | 33.277 | 27.6 | 31.7 | 31.72 | 28.6 | 27.5 | 29.3 | 28.8 | 28.2 | 26.6 | 24.7 | 24.3 | 27.6 | 29.2688 | 20 | 26.8179 | 28.241 | 26.8
| Mean |  | 34.0833 | 26.1833 | 28.3333 | 28.6698 | 25.0601 | 30.7333 | 27.6667 | 33.349 | 27.3667 | 31.7667 | 31.3567 | 27.3167 | 27.2667 | 28.8 | 28.8333 | 28.5 | 27 | 25.1167 | 24.85 | 27.0833 | 26.567 | 20.1667 | 26.8179 | 28.5912 | 30.3133
| Median |  | 34.3 | 26.3 | 28 | 28.6865 | 25.0366 | 30.5 | 27 | 33.2625 | 27.3 | 31.7 | 31.515 | 27.45 | 27.3 | 28.85 | 28.45 | 28.4 | 26.6 | 25.25 | 24.95 | 26.9 | 25.8658 | 20 | 26.7976 | 28.61 | 30.44
| Std Dev. |  | 0.84242 | 0.38687 | 0.5164 | 0.28707 | 0.28118 | 0.80911 | 1.6403 | 0.30134 | 0.15055 | 0.48442 | 0.71455 | 1.2189 | 0.23381 | 0.45607 | 1.4431 | 0.54406 | 1.257 | 0.78337 | 0.33912 | 1.0704 | 1.8726 | 0.40825 | 0.25319 | 0.22762 | 2.1426
| Rel.Std.Dev. |  | 2.472% | 1.478% | 1.823% | 1.001% | 1.122% | 2.633% | 5.929% | 0.9036% | 0.5501% | 1.525% | 2.279% | 4.462% | 0.8575% | 1.584% | 5.005% | 1.909% | 4.655% | 3.119% | 1.365% | 3.952% | 7.049% | 2.024% | 0.9441% | 0.7961% | 7.068%
| PDM3 |  | 22.16% | -6.154% | 1.552% | 2.758% | -10.18% | 10.15% | -0.8373% | 19.53% | -1.913% | 13.86% | 12.39% | -2.092% | -2.271% | 3.225% | 3.344% | 2.15% | -3.227% | -9.977% | -10.93% | -2.928% | -4.779% | -27.72% | -3.879% | 2.476% | 8.649%
| Z-Score (Absolute) |  | 3.82 | 0.88 | 0.4 | 0.6 | 1.55 | 1.82 | 0 | 3.38 | 0.18 | 2.44 | 2.2 | 0.21 | 0.24 | 0.67 | 0.69 | 0.5 | 0.4 | 1.52 | 1.68 | 0.35 | 0.65 | 4.46 | 0.5 | 0.55 | 1.57
| Analytical results for Li in OREAS 606b (Certified Value 19.9 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 2
| 1 | 1 | 29.4 | 19.4 | 18.9 | 18.7275 | 19.9 | 19.2 | 19.214 | 18.8 | 22.8 | 18.4 | 19.7 | 20 | 19.7 | 22.8 | 20.3 | 23 | 21.2 | 17.9 | 17.9087 | 16 | 20.3988 | 24.089 | 24.7 |  |  |  | 1
| 1 | 2 | 28.3 | 18.6 | 19.1 | 18.6459 | 20.4 | 19.9 | 19.769 | 19.2 | 22.2 | 17.9 | 19.6 | 20 | 19.6 | 21.8 | 20.1 | 24.3 | 22.5 | 18.4 | 19.563 | 15 | 19.6659 | 23.773 | 22.3 |  |  |  | 45
| 1 | 3 | 29.9 | 19.7 | 19 | 19.1802 | 19.2 | 19.7 | 19.788 | 19.5 | 22.4 | 18.5 | 19.4 | 19.9 | 19.5 | 21.3 | 20.4 | 22.1 | 22.2 | 17.9 | 20.2474 | 15 | 19.9175 | 23.716 | 20.7 |  |  |  | 16
| 1 | 4 | 30.5 | 18.7 | 18.5 | 19.3807 | 21 | 19 | 19.734 | 19.4 | 22.1 | 18.4 | 19.6 | 20.5 | 19.6 | 21 | 19.6 | 23.3 | 21 | 17.8 | 19.8998 | 16 | 19.4005 | 23.842 | 21.6 |  |  |  | 19.8592
| 1 | 5 | 28.3 | 18.4 | 18.9 | 18.8907 | 20.2 | 21.1 | 19.67 | 19.3 | 22.7 | 18.7 | 20 | 20.8 | 19.6 | 21.3 | 20.1 | 22.1 | 21.8 | 19.2 | 17.6671 | 15 | 20.2001 | 23.569 | 21.5 |  |  |  | 101
| 1 | 6 | 31.4 | 19 | 19 | 18.8802 | 19.9 | 19.5 | 19.875 | 19.6 | 22.2 | 18.2 | 19.6 | 20 | 19.6 | 21.9 | 19.6 | 23 | 23.1 | 18.2 | 18.037 | 15 | 19.6055 | 23.725 | 19.7
| Mean |  | 29.6333 | 18.9667 | 18.9 | 18.9509 | 20.1 | 19.7333 | 19.675 | 19.3 | 22.4 | 18.35 | 19.65 | 20.2 | 19.6 | 21.6833 | 20.0167 | 22.9667 | 21.9667 | 18.2333 | 18.8872 | 15.3333 | 19.8647 | 23.7857 | 21.75
| Median |  | 29.65 | 18.85 | 18.95 | 18.8854 | 20.05 | 19.6 | 19.7515 | 19.35 | 22.3 | 18.4 | 19.6 | 20 | 19.6 | 21.55 | 20.1 | 23 | 22 | 18.05 | 18.8 | 15 | 19.7917 | 23.749 | 21.55
| Std Dev. |  | 1.2291 | 0.50067 | 0.20976 | 0.2789 | 0.6 | 0.74476 | 0.23564 | 0.28284 | 0.28983 | 0.27386 | 0.19748 | 0.36332 | 0.063246 | 0.64317 | 0.34303 | 0.82381 | 0.79666 | 0.52409 | 1.1403 | 0.5164 | 0.38019 | 0.17368 | 1.6968
| Rel.Std.Dev. |  | 4.148% | 2.64% | 1.11% | 1.472% | 2.985% | 3.774% | 1.198% | 1.466% | 1.294% | 1.492% | 1.005% | 1.799% | 0.3227% | 2.966% | 1.714% | 3.587% | 3.627% | 2.874% | 6.037% | 3.368% | 1.914% | 0.7302% | 7.801%
| PDM3 |  | 49.22% | -4.494% | -4.83% | -4.574% | 1.213% | -0.6337% | -0.9274% | -2.816% | 12.79% | -7.599% | -1.053% | 1.716% | -1.305% | 9.185% | 0.793% | 15.65% | 10.61% | -8.187% | -4.895% | -22.79% | 0.02789% | 19.77% | 9.521%
| Z-Score (Absolute) |  | 8.01 | 0.62 | 0.67 | 0.63 | 0.3 | 0 | 0.05 | 0.35 | 2.16 | 1.12 | 0.07 | 0.38 | 0.11 | 1.58 | 0.23 | 2.62 | 1.81 | 1.21 | 0.68 | 3.56 | 0.11 | 3.28 | 1.63
| Analytical results for Lu in OREAS 606b (Certified Value 0.028 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS |  |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 01g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 25g |  |  |  | 3
| 1 | 1 | 2.01: 0.02; 2.03: 0.027; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.04; 2.09: 0.0344; 2.10: 0.026 |  |  |  | 1
| 1 | 2 | 2.01: 0.02; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.04; 2.09: 0.0297; 2.10: 0.059 |  |  |  | 26
| 1 | 3 | 2.01: 0.03; 2.03: 0.024; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0298; 2.10: 0.024 |  |  |  | 16
| 1 | 4 | 2.01: 0.02; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0276; 2.10: 0.025 |  |  |  | 0.027669
| 1 | 5 | 2.01: 0.02; 2.03: 0.022; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0302; 2.10: 0.025 |  |  |  | 102
| 1 | 6 | 2.01: 0.02; 2.03: 0.024; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0326; 2.10: 0.029
| Mean |  | 2.01: 0.021667; 2.03: 0.0245; 2.07: 0.03; 2.08: 0.033333; 2.09: 0.030717; 2.10: 0.031333
| Median |  | 2.01: 0.02; 2.03: 0.0245; 2.07: 0.03; 2.08: 0.03; 2.09: 0.03; 2.10: 0.0255
| Std Dev. |  | 2.01: 0.0040825; 2.03: 0.0016432; 2.07: 0; 2.08: 0.005164; 2.09: 0.002407; 2.10: 0.013663
| Rel.Std.Dev. |  | 2.01: 18.84%; 2.03: 6.707%; 2.07: 0%; 2.08: 15.49%; 2.09: 7.836%; 2.10: 43.6%
| PDM3 |  | 2.01: -21.69%; 2.03: -11.45%; 2.07: 8.423%; 2.08: 20.47%; 2.09: 11.01%; 2.10: 13.24%
| Z-Score (Absolute) |  | 2.01: 1.41; 2.03: 0.95; 2.05: 0.87; 2.06: 3.18; 2.07: 0.06; 2.08: 0.48; 2.09: 0.06; 2.10: 0.16
| Analytical results for Mg in OREAS 606b (Certified Value 0.088 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.09 | 0.09 | 0.09 | 0.09 | 0.090606 | 0.1 | 0.09 | 0.09 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.09 | 0.103 | 0.089 | 0.09 | 0.09 | < 0.1 | 0.0929 | 0.089734 | 0.089817 |  |  |  | 1
| 1 | 2 | 0.09 | 0.08 | 0.09 | 0.09 | 0.089679 | 0.1 | 0.09 | 0.087 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.104 | 0.09 | 0.09 | 0.09 | < 0.1 | 0.0887 | 0.084201 | 0.089579 |  |  |  | 0
| 1 | 3 | 0.09 | 0.09 | 0.09 | 0.09 | 0.089098 | 0.1 | 0.09 | 0.087 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.09 | 0.1 | 0.091 | 0.09 | 0.08 | < 0.1 | 0.0905 | 0.086883 | 0.089928 |  |  |  | 16
| 1 | 4 | 0.09 | 0.08 | 0.09 | 0.09 | 0.090078 | 0.1 | 0.08 | 0.087 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.1 | 0.085 | 0.09 | 0.08 | < 0.1 | 0.0917 | 0.089982 | 0.089164 |  |  |  | 0.088435
| 1 | 5 | 0.09 | 0.08 | 0.1 | 0.09 | 0.088152 | 0.11 | 0.1 | 0.087 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.102 | 0.09 | 0.09 | 0.09 | < 0.1 | 0.0917 | 0.08849 | 0.089086 |  |  |  | 103
| 1 | 6 | 0.09 | 0.08 | 0.09 | 0.09 | 0.090288 | 0.1 | 0.09 | 0.087 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.029 | 0.09 | 0.09 | 0.09 | < 0.1 | 0.088 | 0.082096 | 0.089523
| Mean |  | 0.09 | 0.083333 | 0.091667 | 0.09 | 0.08965 | 0.10167 | 0.09 | 0.0875 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.083333 | 0.089667 | 0.089167 | 0.09 | 0.086667 |  | 0.090583 | 0.086897 | 0.089516
| Median |  | 0.09 | 0.08 | 0.09 | 0.09 | 0.089879 | 0.1 | 0.09 | 0.087 | 0.1 | 0.09 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.101 | 0.09 | 0.09 | 0.09 |  | 0.0911 | 0.087686 | 0.089551
| Std Dev. |  | 1.5202e-17 | 0.005164 | 0.0040825 | 1.5202e-17 | 0.00090039 | 0.0040825 | 0.0063246 | 0.0012247 | 1.5202e-17 | 1.5202e-17 | 0 | 0 | 0 | 1.5202e-17 | 0.005164 | 0.029764 | 0.002137 | 1.5202e-17 | 0.005164 |  | 0.001902 | 0.0031756 | 0.00033881
| Rel.Std.Dev. |  | 1.689e-14% | 6.197% | 4.454% | 1.689e-14% | 1.004% | 4.016% | 7.027% | 1.4% | 1.52e-14% | 1.689e-14% | 0% | 0% | 0% | 1.689e-14% | 6.197% | 33.19% | 2.397% | 1.689e-14% | 5.958% |  | 2.1% | 3.654% | 0.3785%
| PDM3 |  | 1.77% | -5.769% | 3.654% | 1.77% | 1.374% | 14.96% | 1.77% | -1.057% | 13.08% | 1.77% | -9.538% | -9.538% | -9.538% | 1.77% | -5.769% | 1.393% | 0.8275% | 1.77% | -1.999% |  | 2.429% | -1.738% | 1.223%
| Z-Score (Absolute) |  | 0.18 | 2.75 | 0.92 | 0.18 | 0.03 | 5.33 | 0.18 | 0.92 | NA | 0.18 | 4.22 | 4.22 | 4.22 | 0.18 | 2.75 | 0.04 | 0.18 | 0.18 | 1.29 | 17.45 | 0.43 | 1.18 | 0.03
Results from laboratory 2.11 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Mn in OREAS 606b (Certified Value 0.025 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 0.0259 | 0.0253 | 0.0254 | 0.0249 | 0.025734 | 0.0261 | 0.0246 | 0.0247 | 0.0252 | 0.0262 | 0.0245 | 0.0264 | 0.0248 | 0.0279 | 0.0254 | 0.0243 | 0.026 | 0.02355 | 0.0248 | < 0.1 | 0.0263 | 0.024287 | 0.024169 | 0.0297 |  |  | 1
| 1 | 2 | 0.026 | 0.0254 | 0.0252 | 0.0245 | 0.02519 | 0.0252 | 0.0248 | 0.0248 | 0.0243 | 0.0263 | 0.0242 | 0.0265 | 0.0247 | 0.0269 | 0.0249 | 0.0249 | 0.026 | 0.02384 | 0.0246 | < 0.1 | 0.0248 | 0.023271 | 0.023873 | 0.0269 |  |  | 27
| 1 | 3 | 0.026 | 0.0261 | 0.0254 | 0.025 | 0.025471 | 0.0259 | 0.0251 | 0.0245 | 0.025 | 0.0264 | 0.0245 | 0.0265 | 0.0248 | 0.0267 | 0.0256 | 0.0232 | 0.0295 | 0.02403 | 0.0249 | < 0.1 | 0.0256 | 0.023633 | 0.023696 | 0.0256 |  |  | 16
| 1 | 4 | 0.0261 | 0.0255 | 0.0255 | 0.0247 | 0.025593 | 0.0261 | 0.0248 | 0.0244 | 0.0251 | 0.0257 | 0.0242 | 0.0269 | 0.0248 | 0.026 | 0.0244 | 0.0236 | 0.0246 | 0.02398 | 0.0251 | < 0.1 | 0.0263 | 0.024524 | 0.023929 | 0.0262 |  |  | 0.025177
| 1 | 5 | 0.0259 | 0.0274 | 0.0253 | 0.0249 | 0.025151 | 0.0271 | 0.0282 | 0.0245 | 0.0248 | 0.0266 | 0.0249 | 0.0266 | 0.0248 | 0.0261 | 0.0248 | 0.0238 | 0.0262 | 0.024 | 0.0248 | < 0.1 | 0.0256 | 0.023958 | 0.023931 | 0.0263 |  |  | 104
| 1 | 6 | 0.0261 | 0.0253 | 0.0252 | 0.0248 | 0.025533 | 0.025 | 0.0253 | 0.0249 | 0.0251 | 0.0263 | 0.025 | 0.0266 | 0.0249 | 0.0274 | 0.0246 | 0.024 | 0.0262 | 0.02382 | 0.0247 | < 0.1 | 0.0248 | 0.023794 | 0.023545 | 0.0243
| Mean |  | 0.026 | 0.025833 | 0.025333 | 0.0248 | 0.025445 | 0.0259 | 0.025467 | 0.024633 | 0.024917 | 0.02625 | 0.02455 | 0.026583 | 0.0248 | 0.026833 | 0.02495 | 0.023967 | 0.026417 | 0.02387 | 0.024817 |  | 0.025567 | 0.023911 | 0.023857 | 0.0265
| Median |  | 0.026 | 0.02545 | 0.02535 | 0.02485 | 0.025502 | 0.026 | 0.02495 | 0.0246 | 0.02505 | 0.0263 | 0.0245 | 0.02655 | 0.0248 | 0.0268 | 0.02485 | 0.0239 | 0.0261 | 0.02391 | 0.0248 |  | 0.0256 | 0.023876 | 0.023901 | 0.02625
| Std Dev. |  | 8.9443e-05 | 0.00082381 | 0.00012111 | 0.00017889 | 0.00023032 | 0.000751 | 0.0013619 | 0.00019664 | 0.00033116 | 0.00030166 | 0.00033912 | 0.00017224 | 6.3246e-05 | 0.00073666 | 0.00046368 | 0.00058878 | 0.0016278 | 0.00017911 | 0.00017224 |  | 0.00067132 | 0.00045142 | 0.00021519 | 0.0017989
| Rel.Std.Dev. |  | 0.344% | 3.189% | 0.4781% | 0.7213% | 0.9051% | 2.9% | 5.348% | 0.7983% | 1.329% | 1.149% | 1.381% | 0.6479% | 0.255% | 2.745% | 1.858% | 2.457% | 6.162% | 0.7504% | 0.6941% |  | 2.626% | 1.888% | 0.902% | 6.788%
| PDM3 |  | 3.268% | 2.606% | 0.6204% | -1.498% | 1.066% | 2.871% | 1.15% | -2.16% | -1.035% | 4.261% | -2.491% | 5.585% | -1.498% | 6.578% | -0.9022% | -4.808% | 4.923% | -5.192% | -1.432% |  | 1.547% | -5.028% | -5.243% | 5.254%
| Z-Score (Absolute) |  | 0.6 | 0.44 | 0.06 | 0.58 | 0.06 | 0.5 | 0.08 | 0.75 | 0.47 | 0.85 | 0.83 | 1.18 | 0.58 | 1.42 | 0.43 | 1.4 | 1.01 | 1.5 | 0.57 | 24.29 | 0.17 | 1.46 | 1.51 | 1.1
| Analytical results for Mo in OREAS 606b (Certified Value 4 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 3
| 1 | 1 | 4.02 | 3.83 | 4 | 4.1 | 4.2969 | 4.43 | 3.91 | 4.13 | 4.2 | 3.9 | 4.17 | 3.97 | 4.1 | 4.43 | 3.87 | 4 | 3.46 | 4.35 | 3.8 | 4.38 | 4.549 | 4.02 | 3.9352 | < 5 | 4.35 |  | 1
| 1 | 2 | 3.74 | 3.67 | 4 | 4.1 | 4.1606 | 4.26 | 4.02 | 4.68 | 4.2 | 4 | 4.2 | 4.14 | 4.16 | 4.15 | 3.84 | 3.8 | 3.61 | 3.97 | 3.4 | 4.9 | 3.7582 | 4 | 4.159 | < 5 | 4.15 |  | 28
| 1 | 3 | 3.84 | 4.03 | 3.9 | 4 | 4.3632 | 4.49 | 3.84 | 4.27 | 4.2 | 3.8 | 4.11 | 3.98 | 4.26 | 3.93 | 3.98 | 3.7 | 3.3 | 4.03 | 3.7 | 4.76 | 3.5145 | 3.85 | 3.5442 | < 5 | 3.85 |  | 16
| 1 | 4 | 3.98 | 3.73 | 4.1 | 4.1 | 4.207 | 4.44 | 3.89 | 4.06 | 4.2 | 3.7 | 3.85 | 4.01 | 4.22 | 3.91 | 4.27 | 4 | 3.46 | 4.27 | 3.6 | 4.3 | 3.0789 | 3.69 | 3.629 | < 5 | 4.08 |  | 3.9978
| 1 | 5 | 3.87 | 3.74 | 4 | 4.1 | 4.0898 | 4.54 | 4.42 | 4.29 | 4.1 | 3.8 | 4.07 | 4.06 | 4.22 | 4.04 | 3.91 | 4 | 3.47 | 4.29 | 3.8 | 4.7 | 3.0792 | 3.97 | 3.8465 | < 5 | 4.01 |  | 105
| 1 | 6 | 4 | 3.93 | 3.9 | 4.1 | 4.2465 | 4.21 | 3.88 | 4.31 | 4.1 | 3.7 | 4.22 | 3.97 | 4.17 | 4.08 | 3.84 | 3.6 | 3.26 | 4.17 | 3.5 | 4.21 | 2.977 | 3.78 | 4.0759 | < 5 | 3.72
| Mean |  | 3.9083 | 3.8217 | 3.9833 | 4.0833 | 4.2273 | 4.395 | 3.9933 | 4.29 | 4.1667 | 3.8167 | 4.1033 | 4.0217 | 4.1883 | 4.09 | 3.9517 | 3.85 | 3.4267 | 4.18 | 3.6333 | 4.5417 | 3.4928 | 3.885 | 3.865 |  | 4.0267
| Median |  | 3.925 | 3.785 | 4 | 4.1 | 4.2268 | 4.435 | 3.9 | 4.28 | 4.2 | 3.8 | 4.14 | 3.995 | 4.195 | 4.06 | 3.89 | 3.9 | 3.46 | 4.22 | 3.65 | 4.54 | 3.2969 | 3.91 | 3.8909 |  | 4.045
| Std Dev. |  | 0.10998 | 0.13659 | 0.075277 | 0.040825 | 0.097404 | 0.13096 | 0.21759 | 0.21513 | 0.05164 | 0.1169 | 0.13619 | 0.067355 | 0.056716 | 0.18963 | 0.16461 | 0.17607 | 0.1277 | 0.15218 | 0.1633 | 0.28131 | 0.59915 | 0.13338 | 0.24281 |  | 0.22277
| Rel.Std.Dev. |  | 2.814% | 3.574% | 1.89% | 0.9998% | 2.304% | 2.98% | 5.449% | 5.015% | 1.239% | 3.063% | 3.319% | 1.675% | 1.354% | 4.636% | 4.166% | 4.573% | 3.727% | 3.641% | 4.494% | 6.194% | 17.15% | 3.433% | 6.282% |  | 5.532%
| PDM3 |  | -2.237% | -4.405% | -0.3612% | 2.14% | 5.742% | 9.936% | -0.1111% | 7.31% | 4.225% | -4.53% | 2.64% | 0.5976% | 4.767% | 2.307% | -1.153% | -3.696% | -14.29% | 4.558% | -9.116% | 13.6% | -12.63% | -2.821% | -3.322% |  | 0.7227%
| Z-Score (Absolute) |  | 0.33 | 0.67 | 0.04 | 0.35 | 0.92 | 1.58 | 0 | 1.17 | 0.68 | 0.69 | 0.43 | 0.11 | 0.77 | 0.38 | 0.16 | 0.56 | 2.23 | 0.73 | 1.41 | 2.15 | 1.97 | 0.43 | 0.5 | 5.87 | 0.13
| Analytical results for Na in OREAS 606b (Certified Value 0.075 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.07 | 0.146 | 0.079 | 0.07 | 0.07641 | 0.075 | 0.06 | 0.07 | 0.08 | 0.08 | 0.08 | 0.1 | 0.08 | 0.09 | 0.11 | 0.085 | 0.068 | 0.07 | 0.08 | < 0.1 | 0.074 | 0.072304 | 0.071841 |  |  |  | 1
| 1 | 2 | 0.07 | 0.07 | 0.077 | 0.07 | 0.074643 | 0.079 | 0.06 | 0.073 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.11 | 0.077 | 0.07 | 0.07 | 0.08 | < 0.1 | 0.067 | 0.066277 | 0.072197 |  |  |  | 0
| 1 | 3 | 0.07 | 0.187 | 0.078 | 0.07 | 0.075822 | 0.072 | 0.06 | 0.07 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.11 | 0.07 | 0.069 | 0.07 | 0.08 | < 0.1 | 0.067 | 0.073558 | 0.074147 |  |  |  | 16
| 1 | 4 | 0.07 | 0.157 | 0.078 | 0.07 | 0.075785 | 0.078 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.11 | 0.08 | 0.08 | 0.11 | 0.071 | 0.064 | 0.08 | 0.08 | < 0.1 | 0.074 | 0.073759 | 0.073161 |  |  |  | 0.075404
| 1 | 5 | 0.07 | 0.075 | 0.077 | 0.07 | 0.073974 | 0.08 | 0.07 | 0.073 | 0.08 | 0.09 | 0.08 | 0.09 | 0.08 | 0.08 | 0.11 | 0.075 | 0.07 | 0.07 | 0.08 | < 0.1 | 0.067 | 0.071546 | 0.070994 |  |  |  | 106
| 1 | 6 | 0.07 | 0.184 | 0.077 | 0.07 | 0.076379 | 0.08 | 0.07 | 0.073 | 0.08 | 0.08 | 0.08 | 0.1 | 0.08 | 0.09 | 0.11 | 0.03 | 0.07 | 0.08 | 0.08 | < 0.1 | 0.067 | 0.066612 | 0.074606
| Mean |  | 0.07 | 0.1365 | 0.077667 | 0.07 | 0.075502 | 0.077333 | 0.065 | 0.0715 | 0.08 | 0.081667 | 0.08 | 0.093333 | 0.08 | 0.085 | 0.11 | 0.068 | 0.0685 | 0.073333 | 0.08 |  | 0.069333 | 0.070676 | 0.072824
| Median |  | 0.07 | 0.1515 | 0.0775 | 0.07 | 0.075803 | 0.0785 | 0.065 | 0.0715 | 0.08 | 0.08 | 0.08 | 0.095 | 0.08 | 0.085 | 0.11 | 0.073 | 0.0695 | 0.07 | 0.08 |  | 0.067 | 0.071925 | 0.072679
| Std Dev. |  | 0 | 0.052003 | 0.0008165 | 0 | 0.00098474 | 0.0032042 | 0.0054772 | 0.0016432 | 0 | 0.0040825 | 0 | 0.012111 | 0 | 0.0054772 | 0 | 0.01937 | 0.0023452 | 0.005164 | 0 |  | 0.0036148 | 0.0033789 | 0.0013964
| Rel.Std.Dev. |  | 0% | 38.1% | 1.051% | 0% | 1.304% | 4.143% | 8.427% | 2.298% | 0% | 4.999% | 0% | 12.98% | 0% | 6.444% | 0% | 28.49% | 3.424% | 7.042% | 0% |  | 5.214% | 4.781% | 1.917%
| PDM3 |  | -7.166% | 81.03% | 3.001% | -7.166% | 0.1305% | 2.559% | -13.8% | -5.177% | 6.096% | 8.306% | 6.096% | 23.78% | 6.096% | 12.73% | 45.88% | -9.819% | -9.156% | -2.746% | 6.096% |  | -8.05% | -6.27% | -3.421%
| Z-Score (Absolute) |  | 0.67 | 7.48 | 0.27 | 0.67 | 0 | 0.22 | 1.29 | 0.49 | 0.55 | 0.76 | 0.55 | 2.19 | 0.55 | 1.16 | 4.23 | 0.92 | 0.86 | 0.27 | 0.55 | 3.13 | 0.77 | 0.59 | 0.33
| Analytical results for Nb in OREAS 606b (Certified Value 0.9 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 25g | [6 empty] | 2
| 1 | 1 | 0.53 | 0.79 | 0.79 | 1.038 | 0.6 | 1.73 | 1.3 | 0.85 | 0.69 | 0.85 | 0.92 | 1.14 | 0.84 | 1.31 | 0.76 | 1.08 | 8.5674 | 0.02 | 1.291 | 0.81 | [6 empty] | 1
| 1 | 2 | 0.53 | 0.88 | 0.87 | 1.0272 | 0.6 | 1.67 | 1.2 | 1.18 | 0.74 | 0.96 | 0.83 | 1.13 | 0.77 | 1.2 | 0.72 | 1.1 | 8.4423 | 0.02 | 1.2023 | 0.65 | [6 empty] | 29
| 1 | 3 | 0.55 | 0.75 | 0.85 | 1.0298 | 0.6 | 1.78 | 1.4 | 1 | 0.76 | 0.96 | 0.93 | 0.99 | 0.8 | 1.09 | 0.86 | 1.07 | 10.0973 | 0.02 | 1.2584 | 0.61 | [6 empty] | 16
| 1 | 4 | 0.58 | 0.77 | 0.86 | 1.0548 | 0.6 | 1.56 | 1.3 | 0.97 | 0.65 | 0.98 | 0.84 | 0.94 | 0.79 | 1.05 | 0.77 | 1.06 | 9.5323 | 0.01 | 1.2104 | 0.68 | [6 empty] | 0.90081
| 1 | 5 | 0.59 | 0.88 | 0.8 | 1.043 | 0.7 | 1.66 | 1.3 | 1.07 | 0.67 | 0.94 | 0.87 | 1.02 | 0.81 | 1.06 | 0.7 | 1.09 | 9.7883 | 0.02 | 1.2423 | 0.6 | [6 empty] | 107
| 1 | 6 | 0.61 | 0.72 | 0.77 | 1.0405 | 0.6 | 1.62 | 1.3 | 1.02 | 0.82 | 0.91 | 0.91 | 1.02 | 0.76 | 1.12 | 0.74 | 1.07 | 9.3971 | 0.02 | 1.3153 | 0.67
| Mean |  | 0.565 | 0.79833 | 0.82333 | 1.0389 | 0.61667 | 1.67 | 1.3 | 1.015 | 0.72167 | 0.93333 | 0.88333 | 1.04 | 0.795 | 1.1383 | 0.75833 | 1.0783 | 9.3041 | 0.018333 | 1.2533 | 0.67
| Median |  | 0.565 | 0.78 | 0.825 | 1.0392 | 0.6 | 1.665 | 1.3 | 1.01 | 0.715 | 0.95 | 0.89 | 1.02 | 0.795 | 1.105 | 0.75 | 1.075 | 9.4647 | 0.02 | 1.2503 | 0.66
| Std Dev. |  | 0.033317 | 0.067355 | 0.041793 | 0.0099368 | 0.040825 | 0.077974 | 0.063246 | 0.10932 | 0.063692 | 0.047188 | 0.04274 | 0.079246 | 0.02881 | 0.099883 | 0.056006 | 0.01472 | 0.6648 | 0.0040825 | 0.044391 | 0.075631
| Rel.Std.Dev. |  | 5.897% | 8.437% | 5.076% | 0.9565% | 6.62% | 4.669% | 4.865% | 10.77% | 8.826% | 5.056% | 4.838% | 7.62% | 3.624% | 8.775% | 7.385% | 1.365% | 7.145% | 22.27% | 3.542% | 11.29%
| PDM3 |  | -37.28% | -11.38% | -8.601% | 15.33% | -31.54% | 85.39% | 44.31% | 12.68% | -19.89% | 3.61% | -1.94% | 15.45% | -11.75% | 26.37% | -15.82% | 19.71% | 932.9% | -97.96% | 39.13% | -25.62%
| Z-Score (Absolute) |  | 1.45 | 0.46 | 0.36 | 0.55 | NA | 3.21 | NA | 0.45 | 0.79 | 0.11 | 0.11 | 0.55 | 0.48 | 0.97 | 0.63 | 0.72 | 35.38 | 3.75 | 1.45 | 1
Results from laboratories 2.07 and 2.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 606b (Certified Value 25.5 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  |  | 1
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  |  | 2
| 1 | 1 | 2.01: 25.42; 2.03: 24.924; 2.04: 23.9272; 2.05: 26.2; 2.06: 28.294; 2.07: 28.67; 2.08: 22.61; 2.10: 28.5011; 2.11: 17.7; 2.12: 26.87 |  |  |  | 1
| 1 | 2 | 2.01: 23.46; 2.03: 25; 2.04: 23.8589; 2.05: 26; 2.06: 28.748; 2.07: 28.59; 2.08: 23.44; 2.10: 26.2518; 2.11: 20.3; 2.12: 26.21 |  |  |  | 30
| 1 | 3 | 2.01: 25.09; 2.03: 24.704; 2.04: 23.9488; 2.05: 25.5; 2.06: 28.68; 2.07: 28.97; 2.08: 21.02; 2.10: 25.2323; 2.11: 16.1; 2.12: 24.73 |  |  |  | 16
| 1 | 4 | 2.01: 24; 2.03: 25.002; 2.04: 23.7793; 2.05: 26.8; 2.06: 28.945; 2.07: 27.29; 2.08: 21; 2.10: 24.433; 2.11: 16; 2.12: 24.8 |  |  |  | 25.5332
| 1 | 5 | 2.01: 23.06; 2.03: 24.865; 2.04: 24.4168; 2.05: 27.3; 2.06: 28.135; 2.07: 28.44; 2.08: 21.24; 2.10: 26.1176; 2.11: 16; 2.12: 23.8 |  |  |  | 108
| 1 | 6 | 2.01: 25.1; 2.03: 24.953; 2.04: 24.213; 2.05: 26; 2.06: 28.198; 2.07: 28.47; 2.08: 21.51; 2.10: 28.7779; 2.11: 16.4; 2.12: 21.96
| Mean |  | 2.01: 24.355; 2.03: 24.908; 2.04: 24.024; 2.05: 26.3; 2.06: 28.5; 2.07: 28.405; 2.08: 21.8033; 2.10: 26.5523; 2.11: 17.0833; 2.12: 24.7283
| Median |  | 2.01: 24.545; 2.03: 24.9385; 2.04: 23.938; 2.05: 26.1; 2.06: 28.487; 2.07: 28.53; 2.08: 21.375; 2.10: 26.1847; 2.11: 16.25; 2.12: 24.765
| Std Dev. |  | 2.01: 0.98321; 2.03: 0.11229; 2.04: 0.24172; 2.05: 0.64498; 2.06: 0.3343; 2.07: 0.57833; 2.08: 0.99917; 2.10: 1.7472; 2.11: 1.7034; 2.12: 1.75
| Rel.Std.Dev. |  | 2.01: 4.037%; 2.03: 0.4508%; 2.04: 1.006%; 2.05: 2.452%; 2.06: 1.173%; 2.07: 2.036%; 2.08: 4.583%; 2.10: 6.58%; 2.11: 9.971%; 2.12: 7.077%
| PDM3 |  | 2.01: -4.614%; 2.03: -2.449%; 2.04: -5.911%; 2.05: 3.003%; 2.06: 11.62%; 2.07: 11.25%; 2.08: -14.61%; 2.10: 3.991%; 2.11: -33.09%; 2.12: -3.152%
| Z-Score (Absolute) |  | 2.01: 0.19; 2.03: 0.04; 2.04: 0.33; 2.05: 0.62; 2.06: 1.54; 2.07: 1.5; 2.08: 1.26; 2.10: 0.73; 2.11: 3.24; 2.12: 0.04
| Analytical results for Ni in OREAS 606b (Certified Value 3.41 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 3.3 | 3.3 | 3.4 | 3.5 | 3.5048 | 3 | 5.2 | 3.3 | 3 | 3.4 | 3.2 | 3.5 | 3.2 | 3.8 | 3.3 | 3.8 | 3.3 | 3.21 | 3.3 | 7.8931 | 3.74 | 3.2718 | < 5 | 5 |  |  | 1
| 1 | 2 | 3.7 | 3.2 | 3.6 | 3.3 | 3.4993 | 3 | 3.5 | 3.5 | 3 | 3.4 | 3.5 | 3.7 | 3.2 | 3.5 | 3.6 | 4 | 3.3 | 3.13 | 3.3 | 5.325 | 3.63 | 3.5857 | < 5 | 4.5 |  |  | 16
| 1 | 3 | 3.6 | 3.6 | 3.5 | 3.3 | 3.3987 | 3.2 | 3.6 | 3.3 | 3 | 3.3 | 3.8 | 3.4 | 3.2 | 3.5 | 3.6 | 3.7 | 3.7 | 3.18 | 2.7 | 10.099 | 4.1 | 3.4209 | < 5 | 4.2 |  |  | 16
| 1 | 4 | 3.7 | 3.1 | 3.4 | 3.5 | 3.7716 | 3.2 | 3.5 | 3.3 | 3 | 3.3 | 3.2 | 3.4 | 3.2 | 3.3 | 4.9 | 3.9 | 3.5 | 3.12 | 2.9 | 7.7927 | 3.56 | 3.0537 | < 5 | 4.1 |  |  | 3.4101
| 1 | 5 | 3.9 | 3.1 | 3.5 | 3.4 | 3.5077 | 3.3 | 3.9 | 3.6 | 3 | 3.3 | 3.5 | 3.4 | 3.3 | 3.5 | 3.6 | 3.9 | 3.1 | 3.11 | 2.9 | 7.9426 | 3.74 | 3.0617 | < 5 | 4.1 |  |  | 109
| 1 | 6 | 3.5 | 3.4 | 3.4 | 3.4 | 3.3776 | 3 | 3.5 | 3.4 | 3 | 4.9 | 4.3 | 3.5 | 3.4 | 3.6 | 3.5 | 3.8 | 3.1 | 3.26 | 2.8 | 12.6721 | 3.57 | 3.47 | < 5 | 4
| Mean |  | 3.6167 | 3.2833 | 3.4667 | 3.4 | 3.5099 | 3.1167 | 3.8667 | 3.4 | 3 | 3.6 | 3.5833 | 3.4833 | 3.25 | 3.5333 | 3.75 | 3.85 | 3.3333 | 3.1683 | 2.9833 | 8.6207 | 3.7233 | 3.3106 |  | 4.3167
| Median |  | 3.65 | 3.25 | 3.45 | 3.4 | 3.5021 | 3.1 | 3.55 | 3.35 | 3 | 3.35 | 3.5 | 3.45 | 3.2 | 3.5 | 3.6 | 3.85 | 3.3 | 3.155 | 2.9 | 7.9178 | 3.685 | 3.3463 |  | 4.15
| Std Dev. |  | 0.20412 | 0.19408 | 0.08165 | 0.089443 | 0.14035 | 0.13292 | 0.67132 | 0.12649 | 0 | 0.63875 | 0.41673 | 0.1169 | 0.083666 | 0.1633 | 0.57533 | 0.10488 | 0.23381 | 0.059133 | 0.25626 | 2.4954 | 0.20067 | 0.22034 |  | 0.37639
| Rel.Std.Dev. |  | 5.644% | 5.911% | 2.355% | 2.631% | 3.999% | 4.265% | 17.36% | 3.72% | 0% | 17.74% | 11.63% | 3.356% | 2.574% | 4.622% | 15.34% | 2.724% | 7.014% | 1.866% | 8.59% | 28.95% | 5.389% | 6.655% |  | 8.719%
| PDM3 |  | 6.059% | -3.716% | 1.66% | -0.2951% | 2.929% | -8.604% | 13.39% | -0.2951% | -12.03% | 5.57% | 5.081% | 2.149% | -4.694% | 3.615% | 9.969% | 12.9% | -2.25% | -7.089% | -12.51% | 152.8% | 9.187% | -2.915% |  | 26.59%
| Z-Score (Absolute) |  | 0.45 | 0.67 | 0.06 | 0.28 | 0.09 | 1.24 | 1.29 | 0.28 | NA | 0.39 | 0.34 | 0 | 0.79 | 0.17 | 0.9 | 1.24 | 0.51 | 1.06 | 1.69 | 17.32 | 0.81 | 0.58 | 3.32 | 2.81
Results from laboratory 2.11 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 606b (Certified Value 0.022 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 |  |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.0219 | 0.021 | 0.022 | 0.0222 | 0.022303 | 0.022 | 0.021 | 0.0222 | 0.0222 | 0.0225 | 0.022 | 0.022 | 0.021 | 0.023 | 0.022 | 0.022 | 0.02 | 0.02 | 0.0211 | 0.0179 | 0.02124 | 0.021007 |  |  |  |  | 1
| 1 | 2 | 0.0217 | 0.019 | 0.021 | 0.0219 | 0.022449 | 0.021 | 0.021 | 0.0223 | 0.0221 | 0.0224 | 0.022 | 0.022 | 0.021 | 0.023 | 0.022 | 0.022 | 0.02 | 0.02 | 0.0211 | 0.0179 | 0.019861 | 0.021202 |  |  |  |  | 0
| 1 | 3 | 0.0218 | 0.022 | 0.022 | 0.0222 | 0.022363 | 0.022 | 0.021 | 0.0221 | 0.0219 | 0.0227 | 0.022 | 0.022 | 0.021 | 0.022 | 0.023 | 0.022 | 0.02 | 0.02 | 0.0211 | 0.0183 | 0.020921 | 0.02118 |  |  |  |  | 16
| 1 | 4 | 0.0217 | 0.021 | 0.021 | 0.0218 | 0.022962 | 0.022 | 0.02 | 0.0224 | 0.0222 | 0.0221 | 0.022 | 0.022 | 0.021 | 0.022 | 0.022 | 0.021 | 0.02 | 0.02 | 0.0215 | 0.0183 | 0.021423 | 0.021257 |  |  |  |  | 0.021608
| 1 | 5 | 0.0219 | 0.018 | 0.022 | 0.0221 | 0.02236 | 0.023 | 0.023 | 0.0225 | 0.0222 | 0.0223 | 0.022 | 0.022 | 0.021 | 0.022 | 0.022 | 0.022 | 0.022 | 0.02 | 0.0212 | 0.0175 | 0.021053 | 0.021248 |  |  |  |  | 110
| 1 | 6 | 0.0223 | 0.021 | 0.021 | 0.0221 | 0.0226 | 0.022 | 0.021 | 0.0222 | 0.0221 | 0.0223 | 0.022 | 0.022 | 0.022 | 0.023 | 0.022 | 0.019 | 0.0168 | 0.02 | 0.0211 | 0.0175 | 0.019751 | 0.02102
| Mean |  | 0.021883 | 0.020333 | 0.0215 | 0.02205 | 0.022506 | 0.022 | 0.021167 | 0.022283 | 0.022117 | 0.022383 | 0.022 | 0.022 | 0.021167 | 0.0225 | 0.022167 | 0.021333 | 0.0198 | 0.02 | 0.021183 | 0.0179 | 0.020708 | 0.021153
| Median |  | 0.02185 | 0.021 | 0.0215 | 0.0221 | 0.022406 | 0.022 | 0.021 | 0.02225 | 0.02215 | 0.02235 | 0.022 | 0.022 | 0.021 | 0.0225 | 0.022 | 0.022 | 0.02 | 0.02 | 0.0211 | 0.0179 | 0.020987 | 0.021191
| Std Dev. |  | 0.00022286 | 0.0015055 | 0.00054772 | 0.00016432 | 0.00024607 | 0.00063246 | 0.00098319 | 0.0001472 | 0.0001169 | 0.00020412 | 0 | 0 | 0.00040825 | 0.00054772 | 0.00040825 | 0.0012111 | 0.0016733 | 0 | 0.00016021 | 0.00035777 | 0.00072014 | 0.00011133
| Rel.Std.Dev. |  | 1.018% | 7.404% | 2.548% | 0.7452% | 1.093% | 2.875% | 4.645% | 0.6606% | 0.5286% | 0.9119% | 0% | 0% | 1.929% | 2.434% | 1.842% | 5.677% | 8.451% | 0% | 0.7563% | 1.999% | 3.478% | 0.5263%
| PDM3 |  | 1.274% | -5.899% | -0.4996% | 2.046% | 4.157% | 1.814% | -2.042% | 3.126% | 2.354% | 3.588% | 1.814% | 1.814% | -2.042% | 4.128% | 2.586% | -1.271% | -8.367% | -7.441% | -1.965% | -17.16% | -4.164% | -2.108%
| Z-Score (Absolute) |  | 0.25 | 1.74 | 0.25 | 0.46 | 1.05 | 0.4 | 0.67 | 0.76 | 0.55 | 0.89 | 0.4 | 0.4 | 0.67 | 1.04 | 0.61 | 0.46 | 2.43 | 2.17 | 0.65 | 4.88 | 1.26 | 0.69
| Analytical results for Pb in OREAS 606b (Certified Value 128 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 121 | 126.54 | 134.8 | 131.9 | 130.3799 | 124 | 125 | 118.637 | 127 | 134 | 126.7 | 128 | 127.5 | 131.5 | 135 | 132 | 126 | 109 | 180 | 108.5 | 126 | 125.5315 | 128 | 134.9394 | 121.892 | 143 | 1
| 1 | 2 | 121 | 120.48 | 131.1 | 130.1 | 128.854 | 125 | 126 | 119.729 | 123 | 130 | 125.8 | 128 | 128.5 | 131 | 131.5 | 131.5 | 124 | 118 | 179 | 109.2 | 126 | 122.3943 | 129 | 132.954 | 119.761 | 136 | 18
| 1 | 3 | 121 | 129.53 | 133 | 127.4 | 130.0615 | 122 | 125 | 124.9685 | 125 | 131 | 129.8 | 127.5 | 127.5 | 130 | 127 | 134 | 124 | 107 | 189 | 111.5 | 126 | 125.5653 | 130 | 131.4332 | 120.233 | 129 | 16
| 1 | 4 | 121 | 124.01 | 133.1 | 130.4 | 130.9646 | 127 | 126 | 120.254 | 130 | 131 | 122.6 | 126 | 128 | 131.5 | 126.5 | 130.5 | 126 | 105 | 175 | 132.4 | 127 | 124.6272 | 129 | 135.9423 | 121.408 | 130 | 127.5224
| 1 | 5 | 120 | 126.2 | 137.4 | 128.9 | 129.6377 | 130 | 143 | 124.8035 | 124 | 133 | 129.1 | 131 | 129 | 131 | 128 | 130.5 | 127 | 113 | 179 | 110.5 | 126 | 121.4225 | 127 | 133.4326 | 124.274 | 130 | 111
| 1 | 6 | 120 | 128.07 | 128.9 | 130 | 128.8125 | 124 | 128 | 122.375 | 122 | 129 | 127.2 | 130 | 126.5 | 133 | 129.5 | 129.5 | 124 | 107 | 154 | 108.9 | 126 | 123.9935 | 128 | 131.6068 | 127.977 | 121
| Mean |  | 120.6667 | 125.805 | 133.05 | 129.7833 | 129.785 | 125.3333 | 128.8333 | 121.7945 | 125.1667 | 131.3333 | 126.8667 | 128.4167 | 127.8333 | 131.3333 | 129.5833 | 131.3333 | 125.1667 | 109.8333 | 176 | 113.5 | 126.1667 | 123.9224 | 128.5 | 133.3847 | 122.5908 | 131.5
| Median |  | 121 | 126.37 | 133.05 | 130.05 | 129.8496 | 124.5 | 126 | 121.3145 | 124.5 | 131 | 126.95 | 128 | 127.75 | 131.25 | 128.75 | 131 | 125 | 108 | 179 | 109.85 | 126 | 124.3103 | 128.5 | 133.1933 | 121.65 | 130
| Std Dev. |  | 0.5164 | 3.2025 | 2.9331 | 1.5145 | 0.8551 | 2.8048 | 7.0261 | 2.6861 | 2.9269 | 1.8619 | 2.5719 | 1.8005 | 0.8756 | 0.98319 | 3.2158 | 1.5706 | 1.3292 | 4.8339 | 11.7303 | 9.3259 | 0.40825 | 1.6956 | 1.0488 | 1.7953 | 3.0757 | 7.3959
| Rel.Std.Dev. |  | 0.428% | 2.546% | 2.204% | 1.167% | 0.6589% | 2.238% | 5.454% | 2.205% | 2.338% | 1.418% | 2.027% | 1.402% | 0.685% | 0.7486% | 2.482% | 1.196% | 1.062% | 4.401% | 6.665% | 8.217% | 0.3236% | 1.368% | 0.8162% | 1.346% | 2.509% | 5.624%
| PDM3 |  | -5.376% | -1.347% | 4.335% | 1.773% | 1.774% | -1.717% | 1.028% | -4.492% | -1.847% | 2.988% | -0.5142% | 0.7012% | 0.2438% | 2.988% | 1.616% | 2.988% | -1.847% | -13.87% | 38.01% | -11% | -1.063% | -2.823% | 0.7666% | 4.597% | -3.867% | 3.119%
| Z-Score (Absolute) |  | 1.63 | 0.51 | 1.08 | 0.36 | 0.36 | 0.61 | 0.15 | 1.38 | 0.65 | 0.7 | 0.28 | 0.06 | 0.06 | 0.7 | 0.32 | 0.7 | 0.65 | 4 | 10.47 | 3.2 | 0.43 | 0.92 | 0.08 | 1.15 | 1.21 | 0.74
| Analytical results for Pd in OREAS 606b (Indicative Value 18.3 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.08; 2.04: 2.26 | ppb
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 01g | 1
| 1 | 1 | 2.01: < 10; 2.03: 21; 2.04: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.03: 18; 2.04: < 10 | 5
| 1 | 3 | 2.01: < 10; 2.03: 18; 2.04: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.03: 17; 2.04: < 10 | 18.3333
| 1 | 5 | 2.01: < 10; 2.03: 18; 2.04: < 10 | 11
| 1 | 6 | 2.01: < 10; 2.03: 18; 2.04: < 10
| Mean |  | 2.03: 18.3333
| Median |  | 2.03: 18
| Std Dev. |  | 2.03: 1.3663
| Rel.Std.Dev. |  | 2.03: 7.452%
| PDM3 |  | 2.03: 1.776e-13%
Z-Score (Absolute) | bar: 2.01 to 2.04
| Analytical results for Pr in OREAS 606b (Certified Value 6.49 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 3
| 1 | 1 | 2.01: 6.58; 2.03: 6.611; 2.04: 6.8272; 2.05: 6.4; 2.06: 7.5074; 2.07: 7.59; 2.08: 6.83; 2.09: 7.305; 2.10: 4.92; 2.11: 6.88 | 1
| 1 | 2 | 2.01: 6.06; 2.03: 6.549; 2.04: 6.566; 2.05: 6.4; 2.06: 7.5608; 2.07: 7.45; 2.08: 6.96; 2.09: 6.6931; 2.10: 5.72; 2.11: 6.66 | 33
| 1 | 3 | 2.01: 6.48; 2.03: 6.414; 2.04: 6.5649; 2.05: 6.2; 2.06: 7.5254; 2.07: 7.61; 2.08: 6.2; 2.09: 6.6553; 2.10: 4.51; 2.11: 6.34 | 16
| 1 | 4 | 2.01: 6.28; 2.03: 6.646; 2.04: 6.6639; 2.05: 6.5; 2.06: 7.7817; 2.07: 7.17; 2.08: 6.27; 2.09: 6.2853; 2.10: 4.47; 2.11: 6.48 | 6.4858
| 1 | 5 | 2.01: 5.78; 2.03: 6.472; 2.04: 6.4328; 2.05: 6.7; 2.06: 7.7626; 2.07: 7.41; 2.08: 6.36; 2.09: 6.5978; 2.10: 4.48; 2.11: 6.21 | 112
| 1 | 6 | 2.01: 6.56; 2.03: 6.499; 2.04: 6.4126; 2.05: 6.3; 2.06: 7.3555; 2.07: 7.43; 2.08: 6.4; 2.09: 7.5473; 2.10: 4.57; 2.11: 5.67
| Mean |  | 2.01: 6.29; 2.03: 6.5318; 2.04: 6.5779; 2.05: 6.4167; 2.06: 7.5822; 2.07: 7.4433; 2.08: 6.5033; 2.09: 6.8473; 2.10: 4.7783; 2.11: 6.3733
| Median |  | 2.01: 6.38; 2.03: 6.524; 2.04: 6.5655; 2.05: 6.4; 2.06: 7.5431; 2.07: 7.44; 2.08: 6.38; 2.09: 6.6742; 2.10: 4.54; 2.11: 6.41
| Std Dev. |  | 2.01: 0.31818; 2.03: 0.087337; 2.04: 0.15379; 2.05: 0.17224; 2.06: 0.16311; 2.07: 0.15832; 2.08: 0.31399; 2.09: 0.47712; 2.10: 0.49118; 2.11: 0.41779
| Rel.Std.Dev. |  | 2.01: 5.059%; 2.03: 1.337%; 2.04: 2.338%; 2.05: 2.684%; 2.06: 2.151%; 2.07: 2.127%; 2.08: 4.828%; 2.09: 6.968%; 2.10: 10.28%; 2.11: 6.555%
| PDM3 |  | 2.01: -3.018%; 2.03: 0.7102%; 2.04: 1.421%; 2.05: -1.065%; 2.06: 16.91%; 2.07: 14.76%; 2.08: 0.2708%; 2.09: 5.574%; 2.10: -26.33%; 2.11: -1.734%
| Z-Score (Absolute) |  | 2.01: 0.83; 2.03: 0.05; 2.04: 0.22; 2.05: 0.37; 2.06: 3.86; 2.07: 3.36; 2.08: 0.05; 2.09: 1.2; 2.10: 6.31; 2.11: 0.52
| Analytical results for Pt in OREAS 606b (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.08; 2.04: 2.26 | ppb
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 01g | 2
| 1 | 1 | 2.01: 2; 2.03: < 10; 2.04: < 5 | 1
| 1 | 2 | 2.01: 7; 2.03: < 10; 2.04: < 5 | 5
| 1 | 3 | 2.01: 7; 2.03: < 10; 2.04: < 5 | 16
| 1 | 4 | 2.01: 10; 2.03: < 10; 2.04: < 5 | < 5
| 1 | 5 | 2.01: < 2; 2.03: < 10; 2.04: < 5 | 11
| 1 | 6 | 2.01: 9; 2.03: < 10; 2.04: < 5
| Mean |  | 2.01: 7
| Median |  | 2.01: 7
| Std Dev. |  | 2.01: 3.0822
| Rel.Std.Dev. |  | 2.01: 44.03%
PDM3
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0; 2.04: 1.69
| Analytical results for Rb in OREAS 606b (Certified Value 13.3 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 2
| 1 | 1 | 13.6 | 12.8 | 13.37 | 13.8846 | 14.6 | 12.4 | 13.5 | 15.5 | 19.31 | 11.7 | 12.4 | 12.6 | 14.6 | 12.9 | 14.21 | 10.8 | 15.6 | 11 | 13.271 | 15.94 | [6 empty] | 1
| 1 | 2 | 14.2 | 13.6 | 13.63 | 13.492 | 14.2 | 12.4 | 14.8 | 15 | 19.55 | 12.7 | 12.5 | 13.4 | 14.4 | 12.7 | 14.48 | 11.2 | 15.2 | 12.2 | 12.6942 | 15.59 | [6 empty] | 34
| 1 | 3 | 14.5 | 13 | 13.44 | 13.4754 | 14.1 | 12.3 | 13.9 | 15.4 | 19.16 | 12.6 | 12.6 | 13.2 | 13.2 | 13 | 12.83 | 10.9 | 14.4 | 10.9 | 13.1452 | 14.85 | [6 empty] | 16
| 1 | 4 | 13.8 | 12.7 | 13.67 | 13.6558 | 14.3 | 12.3 | 13.9 | 15.6 | 18.22 | 11.3 | 12.7 | 13.2 | 12.8 | 12.8 | 13.11 | 10.9 | 14.2 | 10.8 | 13.502 | 15.55 | [6 empty] | 13.3302
| 1 | 5 | 13.9 | 12.9 | 13.17 | 13.5123 | 14.8 | 13.9 | 13.8 | 15.7 | 19.47 | 11.9 | 12.9 | 13.6 | 13.4 | 12.8 | 13.03 | 10.5 | 14.5 | 10.5 | 13.1196 | 14.88 | [6 empty] | 113
| 1 | 6 | 15 | 13.1 | 13.11 | 13.3114 | 14.2 | 12.9 | 14.6 | 15.7 | 19.36 | 13.3 | 12.9 | 13 | 13.8 | 12.6 | 13.27 | 10.9 | 14.9 | 11 | 12.8528 | 13.6
| Mean |  | 14.1667 | 13.0167 | 13.3983 | 13.5552 | 14.3667 | 12.7 | 14.0833 | 15.4833 | 19.1783 | 12.25 | 12.6667 | 13.1667 | 13.7 | 12.8 | 13.4883 | 10.8667 | 14.8 | 11.0667 | 13.0975 | 15.0683
| Median |  | 14.05 | 12.95 | 13.405 | 13.5021 | 14.25 | 12.4 | 13.9 | 15.55 | 19.335 | 12.25 | 12.65 | 13.2 | 13.6 | 12.8 | 13.19 | 10.9 | 14.7 | 10.95 | 13.1324 | 15.215
| Std Dev. |  | 0.5164 | 0.31885 | 0.23034 | 0.19507 | 0.27325 | 0.62929 | 0.50365 | 0.26394 | 0.48832 | 0.74229 | 0.20656 | 0.34448 | 0.70143 | 0.14142 | 0.68388 | 0.22509 | 0.53292 | 0.58538 | 0.28949 | 0.83703
| Rel.Std.Dev. |  | 3.645% | 2.45% | 1.719% | 1.439% | 1.902% | 4.955% | 3.576% | 1.705% | 2.546% | 6.06% | 1.631% | 2.616% | 5.12% | 1.105% | 5.07% | 2.071% | 3.601% | 5.29% | 2.21% | 5.555%
| PDM3 |  | 6.275% | -2.352% | 0.5109% | 1.688% | 7.775% | -4.728% | 5.65% | 16.15% | 43.87% | -8.104% | -4.978% | -1.227% | 2.774% | -3.978% | 1.186% | -18.48% | 11.03% | -16.98% | -1.746% | 13.04%
| Z-Score (Absolute) |  | 0.67 | 0.39 | 0.04 | 0.1 | 0.85 | 0.68 | 0.59 | 1.88 | 5.27 | 1.1 | 0.71 | 0.25 | 0.24 | 0.59 | 0.04 | 2.37 | 1.25 | 2.19 | 0.32 | 1.49
| Analytical results for Re in OREAS 606b (Certified Value < 0.05 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 01g | 25g | [9 empty] | 3
| 1 | 1 | 0.046 | 0.002 | 0.002 | < 0.01 | 0.002 | 0.001 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.002 | < 0.05 | < 0.001 | [9 empty] | 1
| 1 | 2 | 0.032 | 0.002 | 0.002 | < 0.01 | 0.002 | 0.001 | < 0.05 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.002 | < 0.05 | 0.031 | [9 empty] | 35
| 1 | 3 | 0.02 | 0.001 |  | < 0.01 | 0.002 | 0.001 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | < 0.001 | [9 empty] | 16
| 1 | 4 | 0.025 | 0.001 |  | < 0.01 | 0.002 | < 0.001 | < 0.05 | 0.001 | 0.002 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | 0.002 | [9 empty] | < 0.05
| 1 | 5 | 0.039 | 0.001 | 0.001 | < 0.01 | 0.002 | 0.002 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | 0.001 | < 0.002 | < 0.05 | 0.002 | [9 empty] | 114
| 1 | 6 | 0.04 | 0.002 | 0.001 | < 0.01 | 0.002 | 0.001 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | 0.008
| Mean |  | 2.01: 0.033667; 2.03: 0.0015; 2.04: 0.0015; 2.06: 0.002; 2.07: 0.0012; 2.09: 0.0011667; 2.10: 0.0011667; 2.11: 0.001; 2.12: 0.001; 2.13: 0.001; 2.15: 0.001; 2.18: 0.01075
| Median |  | 2.01: 0.0355; 2.03: 0.0015; 2.04: 0.0015; 2.06: 0.002; 2.07: 0.001; 2.09: 0.001; 2.10: 0.001; 2.11: 0.001; 2.12: 0.001; 2.13: 0.001; 2.15: 0.001; 2.18: 0.005
| Std Dev. |  | 2.01: 0.0098522; 2.03: 0.00054772; 2.04: 0.00057735; 2.06: 0; 2.07: 0.00044721; 2.09: 0.00040825; 2.10: 0.00040825; 2.11: 0; 2.12: 0; 2.13: 0; 2.15: 0; 2.18: 0.013793
| Rel.Std.Dev. |  | 2.01: 29.26%; 2.03: 36.51%; 2.04: 38.49%; 2.06: 0%; 2.07: 37.27%; 2.09: 34.99%; 2.10: 34.99%; 2.11: 0%; 2.12: 0%; 2.13: 0%; 2.15: 0%; 2.18: 128.3%
PDM3
| Z-Score (Absolute) |  | 43.38 | 0 | 0 | 4.72 | 0.67 | 0.56 | 31.69 | 0.45 | 0.45 | 0.67 | 0.67 | 0.67 | 31.69 | 1.01 | 0.67 | 31.69 | 7.87
| Analytical results for S in OREAS 606b (Certified Value 0.251 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 0.26 | 0.25 | 0.25 | 0.2407 | 0.26038 | 0.289 | 0.24 | 0.25314 | 0.24 | 0.27 | 0.2766 | 0.24 | 0.26 | 0.25 | 0.29 | 0.27 | 0.258 | 0.232 | 0.225 | 0.44 | 0.27 | 0.2548 | 0.253 | 0.24139 | 0.247 | 0.23 | 1
| 1 | 2 | 0.25 | 0.23 | 0.25 | 0.2385 | 0.25971 | 0.29 | 0.24 | 0.25065 | 0.24 | 0.27 | 0.2786 | 0.24 | 0.26 | 0.25 | 0.28 | 0.27 | 0.2549 | 0.235 | 0.23 | 0.43 | 0.27 | 0.2378 | 0.248 | 0.22445 | 0.248 | 0.24 | 11
| 1 | 3 | 0.26 | 0.25 | 0.25 | 0.2403 | 0.26565 | 0.285 | 0.24 | 0.25785 | 0.24 | 0.27 | 0.2786 | 0.24 | 0.26 | 0.25 | 0.27 | 0.27 | 0.2524 | 0.222 | 0.225 | 0.4 | 0.27 | 0.2423 | 0.251 | 0.24112 | 0.246 | 0.24 | 16
| 1 | 4 | 0.26 | 0.25 | 0.25 | 0.2381 | 0.2592 | 0.298 | 0.24 | 0.25484 | 0.24 | 0.28 | 0.2707 | 0.24 | 0.27 | 0.25 | 0.27 | 0.26 | 0.2564 | 0.223 | 0.225 | 0.43 | 0.27 | 0.2412 | 0.255 | 0.24725 | 0.246 | 0.23 | 0.25123
| 1 | 5 | 0.26 | 0.24 | 0.25 | 0.2401 | 0.25816 | 0.295 | 0.27 | 0.25101 | 0.24 | 0.27 | 0.2803 | 0.25 | 0.26 | 0.25 | 0.27 | 0.27 | 0.2586 | 0.224 | 0.23 | 0.41 | 0.27 | 0.2521 | 0.253 | 0.23863 | 0.243 | 0.25 | 115
| 1 | 6 | 0.25 | 0.25 | 0.25 | 0.2398 | 0.26388 | 0.302 | 0.25 | 0.25143 | 0.24 | 0.27 | 0.2765 | 0.24 | 0.27 | 0.25 | 0.28 | 0.26 | 0.2528 | 0.064 | 0.23 | 0.43 | 0.27 | 0.2542 | 0.244 | 0.22208 | 0.248 | 0.22
| Mean |  | 0.25667 | 0.245 | 0.25 | 0.23958 | 0.26116 | 0.29317 | 0.24667 | 0.25315 | 0.24 | 0.27167 | 0.27688 | 0.24167 | 0.26333 | 0.25 | 0.27667 | 0.26667 | 0.25552 | 0.2 | 0.2275 | 0.42333 | 0.27 | 0.24707 | 0.25067 | 0.23582 | 0.24633 | 0.235
| Median |  | 0.26 | 0.25 | 0.25 | 0.23995 | 0.26004 | 0.2925 | 0.24 | 0.25229 | 0.24 | 0.27 | 0.2776 | 0.24 | 0.26 | 0.25 | 0.275 | 0.27 | 0.25565 | 0.2235 | 0.2275 | 0.43 | 0.27 | 0.2472 | 0.252 | 0.23987 | 0.2465 | 0.235
| Std Dev. |  | 0.005164 | 0.0083666 | 0 | 0.0010439 | 0.0029355 | 0.0063061 | 0.012111 | 0.0027826 | 0 | 0.0040825 | 0.0033475 | 0.0040825 | 0.005164 | 0 | 0.008165 | 0.005164 | 0.0026049 | 0.066834 | 0.0027386 | 0.015055 | 0 | 0.0074704 | 0.0040332 | 0.010157 | 0.0018619 | 0.010488
| Rel.Std.Dev. |  | 2.012% | 3.415% | 0% | 0.4357% | 1.124% | 2.151% | 4.91% | 1.099% | 0% | 1.503% | 1.209% | 1.689% | 1.961% | 0% | 2.951% | 1.936% | 1.019% | 33.42% | 1.204% | 3.556% | 0% | 3.024% | 1.609% | 4.307% | 0.7558% | 4.463%
| PDM3 |  | 2.163% | -2.48% | -0.4902% | -4.636% | 3.954% | 16.69% | -1.817% | 0.7652% | -4.471% | 8.134% | 10.21% | -3.807% | 4.817% | -0.4902% | 10.12% | 6.144% | 1.706% | -20.39% | -9.446% | 68.5% | 7.471% | -1.658% | -0.2248% | -6.135% | -1.95% | -6.461%
| Z-Score (Absolute) |  | 0.4 | 0.33 | 0.02 | 0.67 | 0.68 | 2.68 | 0.23 | 0.18 | 0.65 | 1.33 | 1.66 | 0.54 | 0.81 | 0.02 | 1.65 | 1.02 | 0.32 | 3.15 | 1.43 | 10.81 | 1.23 | 0.2 | 0.02 | 0.91 | 0.25 | 0.96
| Analytical results for Sb in OREAS 606b (Certified Value 4.84 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 2
| 1 | 1 | 4.21 | 5.11 | 4.7 | 5.24 | 4.9048 | 5.24 | 4.96 | 4.68 | 3.85 | 4.06 | 4.86 | 5.01 | 4.78 | 5.36 | 5.03 | 4.44 | 5.58 | 6.92 | 5.8 | 5.01 | 18.2218 | 5.28 | 4.4523 |  | 5.71 |  | 1
| 1 | 2 | 4.21 | 5.41 | 4.6 | 5.23 | 4.8019 | 4.74 | 4.97 | 4.85 | 3.9 | 4.26 | 5.26 | 5.06 | 4.9 | 5.16 | 4.92 | 4.33 | 5.65 | 4.31 | 5.5 | 5.16 | 10.4481 | 5.56 | 4.504 | 5.64 | 4.72 |  | 36
| 1 | 3 | 3.89 | 5.35 | 4.5 | 5.24 | 4.892 | 5.08 | 4.92 | 4.75 | 3.91 | 3.9 | 5.24 | 5.04 | 4.94 | 5.02 | 5.14 | 4.27 | 5.32 | 4.84 | 5.6 | 4.69 | 13.424 | 4.09 | 4.4528 |  | 4.44 |  | 16
| 1 | 4 | 4.37 | 4.95 | 4.6 | 5.28 | 4.8757 | 4.8 | 4.98 | 4.64 | 3.83 | 4.05 | 4.88 | 5.08 | 4.91 | 4.9 | 5.05 | 4.5 | 5.32 | 4.85 | 5.7 | 4.67 | 12.2498 | 4.03 | 4.6023 | 5.65 | 4.41 |  | 4.8405
| 1 | 5 | 4.25 | 5.16 | 4.7 | 5.29 | 4.6756 | 5.2 | 5.63 | 4.75 | 3.89 | 3.99 | 4.88 | 5.18 | 4.95 | 5.05 | 5 | 4.52 | 5.39 | 5.71 | 5 | 4.69 | 11.5414 | 4.48 | 4.6418 | 5.27 | 4.35 |  | 116
| 1 | 6 | 4.94 | 5.08 | 4.4 | 5.25 | 5.0035 | 4.87 | 5.03 | 4.89 | 3.86 | 4.02 | 5.48 | 5.06 | 4.91 | 5.05 | 4.93 | 4.15 | 5.34 | 6.98 | 5.1 | 4.97 | 9.7873 | 4 | 4.641 | 6.43 | 4.28
| Mean |  | 4.3117 | 5.1767 | 4.5833 | 5.255 | 4.8589 | 4.9883 | 5.0817 | 4.76 | 3.8733 | 4.0467 | 5.1 | 5.0717 | 4.8983 | 5.09 | 5.0117 | 4.3683 | 5.4333 | 5.6017 | 5.45 | 4.865 | 12.6121 | 4.5733 | 4.549 | 5.7475 | 4.6517
| Median |  | 4.23 | 5.135 | 4.6 | 5.245 | 4.8838 | 4.975 | 4.975 | 4.75 | 3.875 | 4.035 | 5.06 | 5.06 | 4.91 | 5.05 | 5.015 | 4.385 | 5.365 | 5.28 | 5.55 | 4.83 | 11.8956 | 4.285 | 4.5531 | 5.645 | 4.425
| Std Dev. |  | 0.34655 | 0.17317 | 0.1169 | 0.02429 | 0.11068 | 0.21339 | 0.27096 | 0.095917 | 0.031411 | 0.11928 | 0.2623 | 0.058109 | 0.061128 | 0.1562 | 0.081833 | 0.14442 | 0.14473 | 1.1369 | 0.32711 | 0.20897 | 3.0354 | 0.68398 | 0.090048 | 0.48815 | 0.53983
| Rel.Std.Dev. |  | 8.037% | 3.345% | 2.551% | 0.4622% | 2.278% | 4.278% | 5.332% | 2.015% | 0.811% | 2.948% | 5.143% | 1.146% | 1.248% | 3.069% | 1.633% | 3.306% | 2.664% | 20.3% | 6.002% | 4.295% | 24.07% | 14.96% | 1.979% | 8.493% | 11.61%
| PDM3 |  | -10.93% | 6.945% | -5.313% | 8.563% | 0.3804% | 3.054% | 4.982% | -1.663% | -19.98% | -16.4% | 5.361% | 4.776% | 1.195% | 5.154% | 3.536% | -9.754% | 12.25% | 15.73% | 12.59% | 0.5062% | 160.6% | -5.519% | -6.021% | 18.74% | -3.901%
| Z-Score (Absolute) |  | 1.36 | 0.38 | 0.81 | 0.53 | 0.26 | 0 | 0.19 | 0.46 | 2.23 | 1.89 | 0.22 | 0.17 | 0.18 | 0.2 | 0.05 | 1.24 | 0.89 | 1.23 | 0.92 | 0.25 | 15.27 | 0.83 | 0.88 | 1.52 | 0.67
| Analytical results for Sc in OREAS 606b (Certified Value 1.52 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  | 2
| 1 | 1 | 1.9 | 1.4 | 1.4 | 1.5 | 1.5866 | 1.5 | 1.4 | 2.2436 | 1.4 | 2 | 1.7 | 1.5 | 1.7 | 1.7 | 1.8 | 1.6 | 2.3 | 1.31 | 1.3 | 1.6 | 1.8091 | 1.1 | 1.4378 | 1.7 |  |  | 1
| 1 | 2 | 2.1 | 1.5 | 1.4 | 1.5 | 1.5425 | 1.7 | 1.6 | 2.1219 | 1.53 | 2 | 1.7 | 1.6 | 1.7 | 1.7 | 1.7 | 1.7 | 2.3 | 1.7 | 1.3 | 1.4 | 1.3439 | 1.2 | 1.4487 | 1.5 |  |  | 37
| 1 | 3 | 2.1 | 1.5 | 1.4 | 1.5 | 1.5455 | 1.6 | 1.5 | 2.2392 | 1.43 | 2 | 1.8 | 1.5 | 1.7 | 1.8 | 1.6 | 1.7 | 2.1 | 1.22 | 1.3 | 1.4 | 1.4936 | 1.1 | 1.4425 | 1.5 |  |  | 16
| 1 | 4 | 2 | 1.3 | 1.4 | 1.4 | 1.5615 | 1.8 | 1.5 | 2.0728 | 1.33 | 2 | 1.6 | 1.4 | 1.7 | 1.7 | 1.5 | 1.7 | 2.1 | 1.43 | 1.3 | 1.5 | 1.3197 | 1.1 | 1.4228 | 1.9 |  |  | 1.5165
| 1 | 5 | 2.2 | 1.4 | 1.4 | 1.5 | 1.5279 | 1.8 | 1.6 | 2.1074 | 1.4 | 2 | 1.7 | 1.5 | 1.7 | 1.7 | 1.6 | 1.7 | 2.1 | 1.7 | 1.3 | 1.3 | 1.4522 | 1.1 | 1.4312 | 1.6 |  |  | 117
| 1 | 6 | 2 | 1.4 | 1.4 | 1.5 | 1.5701 | 1.8 | 1.5 | 2.1219 | 1.37 | 2 | 1.7 | 1.6 | 1.7 | 1.7 | 1.6 | 1.6 | 2.1 | 1.4 | 1.3 | 1.6 | 1.7949 | 1.1 | 1.4248 | 1.4
| Mean |  | 2.05 | 1.4167 | 1.4 | 1.4833 | 1.5557 | 1.7 | 1.5167 | 2.1511 | 1.41 | 2 | 1.7 | 1.5167 | 1.7 | 1.7167 | 1.6333 | 1.6667 | 2.1667 | 1.46 | 1.3 | 1.4667 | 1.5356 | 1.1167 | 1.4346 | 1.6
| Median |  | 2.05 | 1.4 | 1.4 | 1.5 | 1.5535 | 1.75 | 1.5 | 2.1219 | 1.4 | 2 | 1.7 | 1.5 | 1.7 | 1.7 | 1.6 | 1.7 | 2.1 | 1.415 | 1.3 | 1.45 | 1.4729 | 1.1 | 1.4345 | 1.55
| Std Dev. |  | 0.10488 | 0.075277 | 2.4324e-16 | 0.040825 | 0.021192 | 0.12649 | 0.075277 | 0.072199 | 0.067823 | 0 | 0.063246 | 0.075277 | 0 | 0.040825 | 0.10328 | 0.05164 | 0.10328 | 0.1999 | 0 | 0.12111 | 0.21639 | 0.040825 | 0.010177 | 0.17889
| Rel.Std.Dev. |  | 5.116% | 5.314% | 1.737e-14% | 2.752% | 1.362% | 7.441% | 4.963% | 3.356% | 4.81% | 0% | 3.72% | 4.963% | 0% | 2.378% | 6.323% | 3.098% | 4.767% | 13.69% | 0% | 8.257% | 14.09% | 3.656% | 0.7094% | 11.18%
| PDM3 |  | 35.18% | -6.581% | -7.68% | -2.185% | 2.587% | 12.1% | 0.01359% | 41.85% | -7.02% | 31.89% | 12.1% | 0.01359% | 12.1% | 13.2% | 7.707% | 9.905% | 42.88% | -3.723% | -14.27% | -3.284% | 1.26% | -26.36% | -5.397% | 5.509%
| Z-Score (Absolute) |  | 2.76 | 0.64 | 0.73 | 0.28 | 0.11 | 0.88 | 0.1 | 3.31 | 0.67 | NA | 0.88 | 0.1 | 0.88 | 0.97 | 0.53 | 0.7 | 3.39 | 0.41 | 1.27 | 0.37 | 0 | 2.25 | 0.54 | 0.35
Results from laboratory 2.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 606b (Certified Value 1.28 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 2
| 1 | 1 | < 5 | 1.2 | 1.2 | 1 | < 1.5 | 1.6 | 1.2 | < 1 | 1 | 1 | 1.2 | 1.2 | 1.2 | 1.3 | 1.34 | 1 | 0.9 | 1.4 | 1.1 | 1.1994 | < 10 | 1.4 |  |  |  |  | 1
| 1 | 2 | < 5 | 1.6 | 1.2 | 1 | < 1.5 | 1.4 | 1.2 | < 1 | 1 | 1.2 | 1.2 | 1.2 | 1.3 | 1.4 | 1.34 | 1 | 1.6 | 1.3 | 1.2 | 1.2235 | < 10 | 1.4 |  |  |  |  | 38
| 1 | 3 | < 5 | 1.6 | 1.2 |  | < 1.5 | 1.5 | 1.2 | < 1 | 1 | 1.1 | 1.3 | 1.1 | 1.1 | 1.5 | 1.33 | 1 | 1.5 | 1.5 | 1.2 | 1.2548 | < 10 | 1.3 |  |  |  |  | 16
| 1 | 4 | < 5 | 1.3 | 1.3 | 1 | < 1.5 | 1.4 | 1.2 | < 1 | 1 | 0.9 | 1.3 | 1.4 | 1 | 1.4 | 1.33 | 1 | 0.7 | 1.3 | 1.4 | 1.1283 | < 10 | 1.2 |  |  |  |  | 1.2844
| 1 | 5 | < 5 | 1.3 | 1.2 |  | < 1.5 | 1.7 | 1.4 | < 1 | 1 | 1.4 | 1.3 | 1.3 | 1.1 | 1.3 | 1.34 | 1 | 1.6 | 1.5 | 1.4 | 1.0385 | < 10 | 1.3 |  |  |  |  | 118
| 1 | 6 | < 5 | 1.3 | 1.2 | 1 | < 1.5 | 1.5 | 1.3 | < 1 | 1 | 1 | 1.2 | 1.2 | 1.2 | 1.4 | 1.25 | 1 | 1.1 | 1.3 | 1.4 | 0.99377 | < 10 | 1.2
| Mean |  |  | 1.3833 | 1.2167 | 1 |  | 1.5167 | 1.25 |  | 1 | 1.1 | 1.25 | 1.2333 | 1.15 | 1.3833 | 1.3217 | 1 | 1.2333 | 1.3833 | 1.2833 | 1.1397 |  | 1.3
| Median |  |  | 1.3 | 1.2 | 1 |  | 1.5 | 1.2 |  | 1 | 1.05 | 1.25 | 1.2 | 1.15 | 1.4 | 1.335 | 1 | 1.3 | 1.35 | 1.3 | 1.1639 |  | 1.3
| Std Dev. |  |  | 0.17224 | 0.040825 | 0 |  | 0.1169 | 0.083666 |  | 0 | 0.17889 | 0.054772 | 0.10328 | 0.10488 | 0.075277 | 0.035449 | 0 | 0.38816 | 0.098319 | 0.13292 | 0.10537 |  | 0.089443
| Rel.Std.Dev. |  |  | 12.45% | 3.355% | 0% |  | 7.708% | 6.693% |  | 0% | 16.26% | 4.382% | 8.374% | 9.12% | 5.442% | 2.682% | 0% | 31.47% | 7.107% | 10.36% | 9.245% |  | 6.88%
| PDM3 |  |  | 7.704% | -5.272% | -22.14% |  | 18.09% | -2.677% |  | -22.14% | -14.36% | -2.677% | -3.975% | -10.46% | 7.704% | 2.903% | -22.14% | -3.975% | 7.704% | -0.08162% | -11.26% |  | 1.216%
| Z-Score (Absolute) |  | NA | 0.86 | 0.21 | NA | 3.21 | 1.71 | 0 | 4.81 | NA | 0.96 | 0 | 0.11 | 0.64 | 0.86 | 0.46 | NA | 0.11 | 0.86 | 0.21 | 0.71 | 24.05 | 0.32
Results from laboratories 2.01, 2.05, 2.12 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 606b (Certified Value 4.62 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS |  |  |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 25g |  |  |  |  | 2
| 1 | 1 | 2.01: 4.83; 2.03: 4.69; 2.04: 4.687; 2.05: 4.2; 2.06: 5.3861; 2.07: 5.24; 2.08: 4.27; 2.11: 5.166; 2.12: 3.48; 2.13: 5.97 |  |  |  |  | 1
| 1 | 2 | 2.01: 4.51; 2.03: 4.694; 2.04: 4.3658; 2.05: 4.4; 2.06: 5.3277; 2.07: 5.15; 2.08: 4.38; 2.11: 4.7308; 2.12: 3.95; 2.13: 5.61 |  |  |  |  | 22
| 1 | 3 | 2.01: 4.72; 2.03: 4.512; 2.04: 4.4824; 2.05: 4.1; 2.06: 5.6586; 2.07: 5.47; 2.08: 3.98; 2.11: 4.621; 2.12: 3.21; 2.13: 5.02 |  |  |  |  | 16
| 1 | 4 | 2.01: 4.45; 2.03: 4.684; 2.04: 4.4212; 2.05: 4.6; 2.06: 5.4715; 2.07: 5.05; 2.08: 3.9; 2.11: 4.4649; 2.12: 3.17; 2.13: 5.25 |  |  |  |  | 4.6177
| 1 | 5 | 2.01: 4.37; 2.03: 4.529; 2.04: 4.5031; 2.05: 4.4; 2.06: 5.4426; 2.07: 5.31; 2.08: 3.96; 2.11: 4.7587; 2.12: 3.15; 2.13: 5.13 |  |  |  |  | 119
| 1 | 6 | 2.01: 4.6; 2.03: 4.674; 2.04: 4.5421; 2.05: 4.3; 2.06: 5.4105; 2.07: 5.31; 2.08: 4.18; 2.11: 5.2765; 2.12: 3.21; 2.13: 4.69
| Mean |  | 2.01: 4.58; 2.03: 4.6305; 2.04: 4.5003; 2.05: 4.3333; 2.06: 5.4495; 2.07: 5.255; 2.08: 4.1117; 2.11: 4.8363; 2.12: 3.3617; 2.13: 5.2783
| Median |  | 2.01: 4.555; 2.03: 4.679; 2.04: 4.4927; 2.05: 4.35; 2.06: 5.4265; 2.07: 5.275; 2.08: 4.08; 2.11: 4.7447; 2.12: 3.21; 2.13: 5.19
| Std Dev. |  | 2.01: 0.17228; 2.03: 0.085641; 2.04: 0.11069; 2.05: 0.17512; 2.06: 0.11369; 2.07: 0.14529; 2.08: 0.19333; 2.11: 0.31747; 2.12: 0.31231; 2.13: 0.45257
| Rel.Std.Dev. |  | 2.01: 3.762%; 2.03: 1.849%; 2.04: 2.46%; 2.05: 4.041%; 2.06: 2.086%; 2.07: 2.765%; 2.08: 4.702%; 2.11: 6.564%; 2.12: 9.29%; 2.13: 8.574%
| PDM3 |  | 2.01: -0.8166%; 2.03: 0.277%; 2.04: -2.544%; 2.05: -6.158%; 2.06: 18.01%; 2.07: 13.8%; 2.08: -10.96%; 2.11: 4.734%; 2.12: -27.2%; 2.13: 14.31%
| Z-Score (Absolute) |  | 2.01: 0.04; 2.03: 0.04; 2.04: 0.18; 2.05: 0.48; 2.06: 1.49; 2.07: 1.14; 2.08: 0.87; 2.11: 0.41; 2.12: 2.19; 2.13: 1.19
| Analytical results for Sn in OREAS 606b (Certified Value 1.15 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 2
| 1 | 1 | 1.4 | 1 | 1.16 | 1.2549 | 1.4 | 1 | 1.2 | 1.5 | 1.56 | 1.1 | 1.2 | 1.1 | 1.4 | 1.2 | 1.1 | 0.8 | < 5 | 1.2 | 0.73 | 1.1182 | < 10 | 1.24 |  |  |  |  | 1
| 1 | 2 | 1.2 | 1.2 | 1.14 | 1.2554 | 1.43 | 1 | 1.2 | 1.6 | 1.59 | 1.2 | 1.1 | 1.2 | 1.3 | 1.2 | 1.1 | 0.8 | < 5 | 1.2 | 0.81 | 1.0325 | < 10 | 1.23 |  |  |  |  | 13
| 1 | 3 | 1.2 | 0.9 | 1.13 | 1.2515 | 1.39 | 1 | 1.17 | 1.6 | 1.63 | 1.1 | 1.1 | 1.2 | 1.2 | 1.2 | 1 | 0.9 | < 5 | 1.2 | 0.76 | 1.1915 | < 10 | 1.15 |  |  |  |  | 16
| 1 | 4 | 1.2 | 0.9 | 1.11 | 1.246 | 1.41 | 1 | 1.17 | 1.6 | 1.58 | 1.1 | 1.1 | 1.1 | 1.2 | 1.1 | 1 | 0.5 | < 5 | 1.1 | 0.69 | 1.108 | < 10 | 1.17 |  |  |  |  | 1.147
| 1 | 5 | 1.3 | 1.2 | 1.11 | 1.2448 | 1.49 | 1.1 | 1.37 | 1.5 | 1.57 | 1.1 | 1.2 | 1.2 | 1.2 | 1.2 | 1 | 0.9 | < 5 | 1.1 | 0.73 | 1.1049 | < 10 | 1.13 |  |  |  |  | 120
| 1 | 6 | 1.3 | 1.2 | 1.11 | 1.2185 | 1.39 | 1 | 1.13 | 1.5 | 1.52 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1 | 0.8 | < 5 | 1.1 | 0.7 | 1.2238 | < 10 | 1.06
| Mean |  | 1.2667 | 1.0667 | 1.1267 | 1.2452 | 1.4183 | 1.0167 | 1.2067 | 1.55 | 1.575 | 1.1333 | 1.15 | 1.1667 | 1.25 | 1.1833 | 1.0333 | 0.78333 |  | 1.15 | 0.73667 | 1.1298 |  | 1.1633
| Median |  | 1.25 | 1.1 | 1.12 | 1.2488 | 1.405 | 1 | 1.185 | 1.55 | 1.575 | 1.1 | 1.15 | 1.2 | 1.2 | 1.2 | 1 | 0.8 |  | 1.15 | 0.73 | 1.1131 |  | 1.16
| Std Dev. |  | 0.08165 | 0.15055 | 0.020656 | 0.013814 | 0.038166 | 0.040825 | 0.084063 | 0.054772 | 0.036194 | 0.05164 | 0.054772 | 0.05164 | 0.083666 | 0.040825 | 0.05164 | 0.1472 |  | 0.054772 | 0.043665 | 0.068329 |  | 0.066833
| Rel.Std.Dev. |  | 6.446% | 14.11% | 1.833% | 1.109% | 2.691% | 4.016% | 6.967% | 3.534% | 2.298% | 4.556% | 4.763% | 4.426% | 6.693% | 3.45% | 4.997% | 18.79% |  | 4.763% | 5.927% | 6.048% |  | 5.745%
| PDM3 |  | 10.43% | -7.007% | -1.776% | 8.556% | 23.65% | -11.37% | 5.198% | 35.13% | 37.31% | -1.195% | 0.2581% | 1.711% | 8.976% | 3.164% | -9.913% | -31.71% |  | 0.2581% | -35.78% | -1.502% |  | 1.421%
| Z-Score (Absolute) |  | 0.75 | 0.72 | 0.28 | 0.59 | 1.86 | 1.09 | 0.31 | 2.83 | 3.02 | 0.23 | 0.11 | 0.01 | 0.63 | 0.13 | 0.97 | 2.81 | 9.82 | 0.11 | 3.15 | 0.26 | 28.21 | 0.01
| Analytical results for Sr in OREAS 606b (Certified Value 24.8 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 2
| 1 | 1 | 30.7 | 30.2 | 24 | 25.73 | 24.0638 | 24.8 | 22.7 | 23.6 | 27 | 27.13 | 24.1 | 25.3 | 25.2 | 27.9 | 25.6 | 25.5 | 34.2 | 23.2 | 27 | 23.5171 | 18 | 24.2987 | 24.52 | 27.6 |  |  | 1
| 1 | 2 | 29.6 | 25.9 | 24 | 25.42 | 23.5178 | 24.4 | 22.8 | 23.5 | 25.9 | 26.88 | 26.1 | 25.6 | 26.3 | 26.9 | 25.5 | 25.8 | 34 | 23.3 | 27.5 | 22.2645 | 17 | 23.2856 | 24.3 | 26.6 |  |  | 14
| 1 | 3 | 29.8 | 33 | 24 | 24.97 | 23.9113 | 24.6 | 22.5 | 23.1 | 25.7 | 26.73 | 25.3 | 25.5 | 26.2 | 24.6 | 26.3 | 24.3 | 35.1 | 23.6 | 25.3 | 23.0811 | 18 | 24.2845 | 24.27 | 25.3 |  |  | 16
| 1 | 4 | 29.4 | 33.1 | 24 | 25.23 | 23.8087 | 24.9 | 22.6 | 23.5 | 26.5 | 25.84 | 24.1 | 25.7 | 25.7 | 25.1 | 25.4 | 24.4 | 32.8 | 23.5 | 25.5 | 22.2302 | 18 | 24.8563 | 24.46 | 26.4 |  |  | 24.8338
| 1 | 5 | 30.5 | 27.6 | 25 | 24.78 | 23.2787 | 26.5 | 25.3 | 23.1 | 26.6 | 26.7 | 25 | 25.6 | 25.7 | 25.6 | 25.4 | 24.6 | 34.5 | 23.2 | 25.4 | 27.7335 | 18 | 23.329 | 24.35 | 25.7 |  |  | 121
| 1 | 6 | 30.9 | 32.5 | 24 | 25.33 | 23.6264 | 24.3 | 23 | 23.2 | 25.9 | 26.91 | 25.8 | 25.9 | 25.9 | 26.2 | 25.9 | 6.8 | 34.3 | 22.9 | 26.4 | 23.4492 | 17 | 24.1251 | 24.4 | 23.4
| Mean |  | 30.15 | 30.3833 | 24.1667 | 25.2433 | 23.7011 | 24.9167 | 23.15 | 23.3333 | 26.2667 | 26.6983 | 25.0667 | 25.6 | 25.8333 | 26.05 | 25.6833 | 21.9 | 34.15 | 23.2833 | 26.1833 | 23.7126 | 17.6667 | 24.0299 | 24.3833 | 25.8333
| Median |  | 30.15 | 31.35 | 24 | 25.28 | 23.7175 | 24.7 | 22.75 | 23.35 | 26.2 | 26.805 | 25.15 | 25.6 | 25.8 | 25.9 | 25.55 | 24.5 | 34.25 | 23.25 | 25.95 | 23.2651 | 18 | 24.2048 | 24.375 | 26.05
| Std Dev. |  | 0.62849 | 3.0525 | 0.40825 | 0.33619 | 0.28458 | 0.8085 | 1.0672 | 0.22509 | 0.50859 | 0.44763 | 0.84063 | 0.2 | 0.39833 | 1.2145 | 0.35449 | 7.4227 | 0.76092 | 0.24833 | 0.92826 | 2.0479 | 0.5164 | 0.61226 | 0.095638 | 1.432
| Rel.Std.Dev. |  | 2.085% | 10.05% | 1.689% | 1.332% | 1.201% | 3.245% | 4.61% | 0.9647% | 1.936% | 1.677% | 3.354% | 0.7812% | 1.542% | 4.662% | 1.38% | 33.89% | 2.228% | 1.067% | 3.545% | 8.636% | 2.923% | 2.548% | 0.3922% | 5.543%
| PDM3 |  | 21.41% | 22.35% | -2.686% | 1.649% | -4.561% | 0.3337% | -6.78% | -6.042% | 5.77% | 7.508% | 0.9377% | 3.085% | 4.025% | 4.897% | 3.421% | -11.81% | 37.51% | -6.243% | 5.434% | -4.515% | -28.86% | -3.237% | -1.814% | 4.025%
| Z-Score (Absolute) |  | 3.01 | 3.15 | 0.6 | 0.05 | 0.88 | 0.14 | 1.21 | 1.1 | 0.67 | 0.93 | 0.05 | 0.27 | 0.41 | 0.54 | 0.32 | 1.96 | 5.42 | 1.13 | 0.62 | 0.87 | 4.51 | 0.68 | 0.47 | 0.41
| Analytical results for Ta in OREAS 606b (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [9 empty] | 3
| 1 | 1 | 0.02 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | 1
| 1 | 2 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.01 | [9 empty] | 25
| 1 | 3 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | 16
| 1 | 4 | 0.02 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.09 | [9 empty] | < 0.01
| 1 | 5 | 0.02 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | 122
| 1 | 6 | 0.03 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.01
| Mean |  | 2.01: 0.018333; 2.08: 0.01; 2.10: 0.01; 2.11: 0.01; 2.12: 0.01; 2.13: 0.01; 2.14: 0.01; 2.20: 0.028333
| Median |  | 2.01: 0.02; 2.08: 0.01; 2.10: 0.01; 2.11: 0.01; 2.12: 0.01; 2.13: 0.01; 2.14: 0.01; 2.20: 0.02
| Std Dev. |  | 2.01: 0.0075277; 2.08: 0; 2.10: 0; 2.11: 0; 2.12: 0; 2.13: 0; 2.14: 0; 2.20: 0.030605
| Rel.Std.Dev. |  | 2.01: 41.06%; 2.08: 0%; 2.10: 0%; 2.11: 0%; 2.12: 0%; 2.13: 0%; 2.14: 0%; 2.20: 108%
PDM3
| Z-Score (Absolute) |  | 1.12 | 2.02 | 0.67 | 2.02 | 0.67 | 2.02 | 0 | 0 | 0 | 0 | 0 | 0.22 | 0.67 | 0.67 | 0.67 | 2.02 | 2.47
| Analytical results for Tb in OREAS 606b (Certified Value 0.46 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  | 3
| 1 | 1 | 2.01: 0.49; 2.03: 0.448; 2.04: 0.43517; 2.05: 0.5; 2.06: 0.6426; 2.07: 0.5; 2.08: 0.49; 2.10: 0.5124; 2.11: 0.34; 2.12: 0.458 |  |  | 1
| 1 | 2 | 2.01: 0.43; 2.03: 0.442; 2.04: 0.45552; 2.05: 0.4; 2.06: 0.6509; 2.07: 0.5; 2.08: 0.51; 2.10: 0.468; 2.11: 0.39; 2.12: 0.491 |  |  | 26
| 1 | 3 | 2.01: 0.48; 2.03: 0.434; 2.04: 0.43372; 2.05: 0.4; 2.06: 0.6851; 2.07: 0.49; 2.08: 0.45; 2.10: 0.4572; 2.11: 0.35; 2.12: 0.424 |  |  | 16
| 1 | 4 | 2.01: 0.44; 2.03: 0.45; 2.04: 0.43913; 2.05: 0.5; 2.06: 0.6628; 2.07: 0.48; 2.08: 0.45; 2.10: 0.4402; 2.11: 0.33; 2.12: 0.441 |  |  | 0.46139
| 1 | 5 | 2.01: 0.42; 2.03: 0.447; 2.04: 0.43654; 2.05: 0.5; 2.06: 0.6785; 2.07: 0.5; 2.08: 0.47; 2.10: 0.4672; 2.11: 0.33; 2.12: 0.425 |  |  | 123
| 1 | 6 | 2.01: 0.48; 2.03: 0.447; 2.04: 0.4524; 2.05: 0.5; 2.06: 0.6296; 2.07: 0.5; 2.08: 0.46; 2.10: 0.5261; 2.11: 0.33; 2.12: 0.408
| Mean |  | 2.01: 0.45667; 2.03: 0.44467; 2.04: 0.44208; 2.05: 0.46667; 2.06: 0.65825; 2.07: 0.495; 2.08: 0.47167; 2.10: 0.47852; 2.11: 0.345; 2.12: 0.44117
| Median |  | 2.01: 0.46; 2.03: 0.447; 2.04: 0.43783; 2.05: 0.5; 2.06: 0.65685; 2.07: 0.5; 2.08: 0.465; 2.10: 0.4676; 2.11: 0.335; 2.12: 0.433
| Std Dev. |  | 2.01: 0.030111; 2.03: 0.0058538; 2.04: 0.0094245; 2.05: 0.05164; 2.06: 0.021314; 2.07: 0.0083666; 2.08: 0.024014; 2.10: 0.033387; 2.11: 0.023452; 2.12: 0.029742
| Rel.Std.Dev. |  | 2.01: 6.594%; 2.03: 1.316%; 2.04: 2.132%; 2.05: 11.07%; 2.06: 3.238%; 2.07: 1.69%; 2.08: 5.091%; 2.10: 6.977%; 2.11: 6.798%; 2.12: 6.742%
| PDM3 |  | 2.01: -1.025%; 2.03: -3.626%; 2.04: -4.186%; 2.05: 1.143%; 2.06: 42.67%; 2.07: 7.283%; 2.08: 2.226%; 2.10: 3.711%; 2.11: -25.23%; 2.12: -4.384%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.52; 2.04: 0.63; 2.05: NA; 2.06: 8.77; 2.07: 1.67; 2.08: 0.65; 2.10: 0.95; 2.11: 4.86; 2.12: 0.67
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 606b (Certified Value 0.8 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 3
| 1 | 1 | 1.2 | 0.76 | 0.9 | 0.8 | 0.86768 | 0.82 | 0.95 | 0.76 | 0.82 | 0.72 | 0.74 | 0.78 | 0.85 | 0.82 | 0.84 | 0.64 | 0.112 | 0.76 | 0.7 | 0.82645 | < 10 | 0.38 |  |  |  |  | 1
| 1 | 2 | 1.25 | 0.81 | 0.8 | 0.8 | 0.75023 | 0.81 | 0.86 | 0.81 | 0.81 | 0.78 | 0.77 | 0.82 | 0.8 | 0.79 | 0.79 | 0.49 | 0.7 | 0.77 | 0.72 | 0.6592 | < 10 | 0.59 |  |  |  |  | 27
| 1 | 3 | 1.13 | 0.84 | 0.8 | 0.8 | 0.81484 | 0.82 | 0.84 | 0.77 | 0.85 | 0.83 | 0.77 | 0.79 | 0.77 | 0.86 | 0.78 | 0.39 | 0.112 | 0.76 | 0.73 | 0.69739 | < 10 | 0.25 |  |  |  |  | 16
| 1 | 4 | 1.22 | 0.74 | 0.9 | 0.9 | 0.83899 | 0.82 | 0.93 | 0.8 | 0.78 | 0.68 | 0.77 | 0.82 | 0.77 | 0.82 | 0.79 | 0.33 | 1.19 | 0.79 | 0.67 | 0.6288 | < 10 | 0.51 |  |  |  |  | 0.79864
| 1 | 5 | 1.27 | 0.72 | 0.9 | 0.8 | 0.81681 | 0.9 | 0.98 | 0.79 | 0.77 | 0.7 | 0.78 | 0.81 | 0.79 | 0.85 | 0.86 | 0.5 | 1 | 0.83 | 0.79 | 0.83447 | < 10 | 0.43 |  |  |  |  | 124
| 1 | 6 | 1.33 | 0.72 | 0.9 | 0.8 | 0.79125 | 0.83 | 0.82 | 0.81 | 0.74 | 0.79 | 0.76 | 0.83 | 0.77 | 0.81 | 0.77 | 0.44 | 1.04 | 0.84 | 0.71 | 0.76303 | < 10 | 0.42
| Mean |  | 1.2333 | 0.765 | 0.86667 | 0.81667 | 0.8133 | 0.83333 | 0.89667 | 0.79 | 0.795 | 0.75 | 0.765 | 0.80833 | 0.79167 | 0.825 | 0.805 | 0.465 | 0.69233 | 0.79167 | 0.72 | 0.73489 |  | 0.43
| Median |  | 1.235 | 0.75 | 0.9 | 0.8 | 0.81582 | 0.82 | 0.895 | 0.795 | 0.795 | 0.75 | 0.77 | 0.815 | 0.78 | 0.82 | 0.79 | 0.465 | 0.85 | 0.78 | 0.715 | 0.73021 |  | 0.425
| Std Dev. |  | 0.067725 | 0.049699 | 0.05164 | 0.040825 | 0.040238 | 0.033267 | 0.06532 | 0.020976 | 0.03937 | 0.058652 | 0.013784 | 0.019408 | 0.031252 | 0.025884 | 0.036194 | 0.10672 | 0.47683 | 0.035449 | 0.04 | 0.086582 |  | 0.11576
| Rel.Std.Dev. |  | 5.491% | 6.497% | 5.958% | 4.999% | 4.947% | 3.992% | 7.285% | 2.655% | 4.952% | 7.82% | 1.802% | 2.401% | 3.948% | 3.137% | 4.496% | 22.95% | 68.87% | 4.478% | 5.556% | 11.78% |  | 26.92%
| PDM3 |  | 54.43% | -4.212% | 8.518% | 2.258% | 1.836% | 4.344% | 12.27% | -1.081% | -0.4554% | -6.09% | -4.212% | 1.214% | -0.8728% | 3.301% | 0.7968% | -41.78% | -13.31% | -0.8728% | -9.846% | -7.982% |  | -46.16%
| Z-Score (Absolute) |  | 7.94 | 0.48 | NA | NA | 0.39 | 0.75 | 1.89 | 0.03 | 0.06 | 0.75 | 0.48 | 0.3 | 0 | 0.6 | 0.24 | 5.87 | 1.79 | 0 | 1.29 | 1.02 | 75.67 | 6.5
Results from laboratories 2.04 and 2.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 606b (Certified Value 10.8 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  |  | 2
| 1 | 1 | 13 | 10.1 | 10.8 | 10.62 | 9.4136 | 11.3 | 10.3 | 12.23 | 10.8 | 11.25 | 10.3 | 10.6 | 10.6 | 10.8 | 11 | 9.16 | < 20 | 10.7 | 13.2 | 12.4302 | 9.33 | 10.8043 | 12.51 |  |  |  | 1
| 1 | 2 | 12.5 | 9.1 | 10.5 | 10.48 | 9.5854 | 11.6 | 10.4 | 12.128 | 11.1 | 10.97 | 11.4 | 10.6 | 10.6 | 10.6 | 11 | 9.55 | < 20 | 10.9 | 12 | 11.4639 | 10.1 | 10.6511 | 12.19 |  |  |  | 28
| 1 | 3 | 12.6 | 10.2 | 10.6 | 10.29 | 9.8032 | 11 | 10.1 | 12.761 | 11.2 | 11.26 | 11.2 | 10.5 | 10.7 | 10 | 11.4 | 8.52 | < 20 | 10.7 | 12.3 | 11.6371 | 8.91 | 10.6975 | 11.15 |  |  |  | 16
| 1 | 4 | 12.8 | 9.7 | 10.6 | 10.4 | 9.3651 | 11.7 | 10.2 | 12.341 | 11 | 10.95 | 10.3 | 10.4 | 10.8 | 10.2 | 10.9 | 8.58 | < 20 | 11 | 12.3 | 10.8717 | 8.84 | 10.7673 | 11.67 |  |  |  | 10.8095
| 1 | 5 | 12.6 | 9.1 | 10.9 | 10.25 | 9.403 | 11.9 | 11.7 | 12.155 | 11.4 | 11.25 | 10.5 | 10.2 | 11 | 10.2 | 11.2 | 8.64 | < 20 | 11 | 12 | 11.267 | 8.67 | 10.7506 | 11.34 |  |  |  | 125
| 1 | 6 | 13.1 | 10 | 10.8 | 10.53 | 9.5496 | 11.5 | 10.3 | 12.172 | 11 | 11.05 | 11.3 | 10.2 | 10.9 | 10.2 | 10.9 | 8.79 | < 20 | 10.7 | 13.1 | 12.591 | 8.96 | 10.6916 | 10.49
| Mean |  | 12.7667 | 9.7 | 10.7 | 10.4283 | 9.52 | 11.5 | 10.5 | 12.2978 | 11.0833 | 11.1217 | 10.8333 | 10.4167 | 10.7667 | 10.3333 | 11.0667 | 8.8733 |  | 10.8333 | 12.4833 | 11.7102 | 9.135 | 10.7271 | 11.5583
| Median |  | 12.7 | 9.85 | 10.7 | 10.44 | 9.4816 | 11.55 | 10.3 | 12.201 | 11.05 | 11.15 | 10.85 | 10.45 | 10.75 | 10.2 | 11 | 8.715 |  | 10.8 | 12.3 | 11.5505 | 8.935 | 10.724 | 11.505
| Std Dev. |  | 0.24221 | 0.49396 | 0.15492 | 0.14247 | 0.16392 | 0.31623 | 0.59666 | 0.23919 | 0.20412 | 0.14811 | 0.52026 | 0.18348 | 0.1633 | 0.30111 | 0.19664 | 0.40307 |  | 0.15055 | 0.53448 | 0.67227 | 0.52034 | 0.056665 | 0.73115
| Rel.Std.Dev. |  | 1.897% | 5.092% | 1.448% | 1.366% | 1.722% | 2.75% | 5.682% | 1.945% | 1.842% | 1.332% | 4.802% | 1.761% | 1.517% | 2.914% | 1.777% | 4.543% |  | 1.39% | 4.282% | 5.741% | 5.696% | 0.5282% | 6.326%
| PDM3 |  | 18.11% | -10.26% | -1.013% | -3.526% | -11.93% | 6.388% | -2.863% | 13.77% | 2.534% | 2.888% | 0.2208% | -3.634% | -0.396% | -4.405% | 2.379% | -17.91% |  | 0.2208% | 15.49% | 8.332% | -15.49% | -0.7623% | 6.928%
| Z-Score (Absolute) |  | 3.11 | 1.66 | 0.1 | 0.53 | 1.94 | 1.14 | 0.41 | 2.38 | 0.49 | 0.55 | 0.1 | 0.54 | 0 | 0.67 | 0.47 | 2.95 | 1.19 | 0.1 | 2.67 | 1.47 | 2.54 | 0.06 | 1.23
| Analytical results for Ti in OREAS 606b (Certified Value 0.021 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | [5 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | [5 empty] | 3
| 1 | 1 | 0.02 | 0.02 | 0.019 | 0.0213 | 0.022367 | 0.04 | 0.019 | 0.0327 | 0.0285 | 0.02 | 0.021 | 0.022 | 0.026 | 0.022 | 0.024 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0156 | 0.025447 | [5 empty] | 1
| 1 | 2 | 0.02 | 0.022 | 0.019 | 0.0215 | 0.021825 | 0.04 | 0.02 | 0.0341 | 0.0299 | 0.021 | 0.021 | 0.021 | 0.025 | 0.022 | 0.023 | 0.023 | 0.04 | 0.02 | < 0.1 | 0.015 | 0.022264 | [5 empty] | 29
| 1 | 3 | 0.02 | 0.02 | 0.019 | 0.0218 | 0.021919 | 0.04 | 0.02 | 0.0323 | 0.0299 | 0.021 | 0.022 | 0.022 | 0.024 | 0.022 | 0.022 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0156 | 0.027354 | [5 empty] | 16
| 1 | 4 | 0.02 | 0.02 | 0.019 | 0.0215 | 0.022579 | 0.04 | 0.019 | 0.0335 | 0.0277 | 0.019 | 0.023 | 0.022 | 0.023 | 0.021 | 0.021 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0162 | 0.025898 | [5 empty] | 0.021017
| 1 | 5 | 0.02 | 0.023 | 0.019 | 0.0212 | 0.021864 | 0.04 | 0.022 | 0.0335 | 0.0301 | 0.02 | 0.022 | 0.022 | 0.024 | 0.022 | 0.022 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0162 | 0.022157 | [5 empty] | 126
| 1 | 6 | 0.02 | 0.02 | 0.02 | 0.0216 | 0.021582 | 0.04 | 0.02 | 0.0334 | 0.0291 | 0.022 | 0.022 | 0.022 | 0.025 | 0.021 | 0.022 | 0.012 | 0.04 | 0.02 | < 0.1 | 0.0156 | 0.022357
| Mean |  | 0.02 | 0.020833 | 0.019167 | 0.021483 | 0.022023 | 0.04 | 0.02 | 0.03325 | 0.0292 | 0.0205 | 0.021833 | 0.021833 | 0.0245 | 0.021667 | 0.022333 | 0.019167 | 0.04 | 0.02 |  | 0.0157 | 0.024246
| Median |  | 0.02 | 0.02 | 0.019 | 0.0215 | 0.021892 | 0.04 | 0.02 | 0.03345 | 0.0295 | 0.0205 | 0.022 | 0.022 | 0.0245 | 0.022 | 0.022 | 0.02 | 0.04 | 0.02 |  | 0.0156 | 0.023902
| Std Dev. |  | 0 | 0.0013292 | 0.00040825 | 0.0002137 | 0.00037337 | 0 | 0.0010954 | 0.0006442 | 0.00095289 | 0.0010488 | 0.00075277 | 0.00040825 | 0.0010488 | 0.0005164 | 0.0010328 | 0.0037103 | 0 | 0 |  | 0.00045166 | 0.0022668
| Rel.Std.Dev. |  | 0% | 6.38% | 2.13% | 0.9947% | 1.695% | 0% | 5.477% | 1.937% | 3.263% | 5.116% | 3.448% | 1.87% | 4.281% | 2.383% | 4.624% | 19.36% | 0% | 0% |  | 2.877% | 9.349%
| PDM3 |  | -4.839% | -0.8739% | -8.804% | 2.219% | 4.785% | 90.32% | -4.839% | 58.21% | 38.94% | -2.46% | 3.884% | 3.884% | 16.57% | 3.091% | 6.263% | -8.804% | 90.32% | -4.839% |  | -25.3% | 15.36%
| Z-Score (Absolute) |  | 0.67 | 0.37 | 0.98 | 0.13 | 0.07 | 6.68 | 0.67 | 4.2 | 2.71 | 0.49 | 0 | 0 | 0.98 | 0.06 | 0.18 | 0.98 | 6.68 | 0.67 | 10.36 | 2.26 | 0.89
| Analytical results for Tl in OREAS 606b (Certified Value 0.25 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 01g | 25g |  |  |  |  | 3
| 1 | 1 | 0.2 | 0.25 | 0.3 | 0.25 | 0.25189 | 0.25 | 0.25 | 0.28 | 0.27 | 0.28 | 0.25 | 0.24 | 0.23 | 0.26 | 0.22 | 0.25 | 0.28 | 0.17 | < 5 | 0.18 | 0.2389 | 0.21 |  |  |  |  | 1
| 1 | 2 | 0.18 | 0.2 | 0.3 | 0.24 | 0.2514 | 0.26 | 0.26 | 0.26 | 0.27 | 0.28 | 0.25 | 0.24 | 0.25 | 0.25 | 0.21 | 0.252 | 0.27 | 0.17 | < 5 | 0.19 | 0.23875 | 0.23 |  |  |  |  | 30
| 1 | 3 | 0.18 | 0.26 | 0.3 | 0.23 | 0.25003 | 0.25 | 0.26 | 0.26 | 0.26 | 0.29 | 0.27 | 0.24 | 0.24 | 0.24 | 0.22 | 0.241 | 0.27 | 0.18 | < 5 | 0.16 | 0.23659 | 0.2 |  |  |  |  | 16
| 1 | 4 | 0.2 | 0.24 | 0.3 | 0.25 | 0.24124 | 0.25 | 0.26 | 0.3 | 0.27 | 0.26 | 0.24 | 0.23 | 0.23 | 0.25 | 0.21 | 0.251 | 0.27 | 0.17 | < 5 | 0.17 | 0.24287 | 0.2 |  |  |  |  | 0.25128
| 1 | 5 | 0.19 | 0.22 | 0.3 | 0.25 | 0.23625 | 0.26 | 0.29 | 0.28 | 0.27 | 0.29 | 0.25 | 0.24 | 0.24 | 0.25 | 0.22 | 0.254 | 0.27 | 0.17 | < 5 | 0.2 | 0.24274 | 0.2 |  |  |  |  | 127
| 1 | 6 | 0.2 | 0.25 | 0.3 | 0.25 | 0.24786 | 0.26 | 0.26 | 0.25 | 0.28 | 0.28 | 0.27 | 0.23 | 0.25 | 0.25 | 0.21 | 0.229 | 0.27 | 0.18 | < 5 | 0.16 | 0.23896 | 0.18
| Mean |  | 0.19167 | 0.23667 | 0.3 | 0.245 | 0.24644 | 0.255 | 0.26333 | 0.27167 | 0.27 | 0.28 | 0.255 | 0.23667 | 0.24 | 0.25 | 0.215 | 0.24617 | 0.27167 | 0.17333 |  | 0.17667 | 0.2398 | 0.20333
| Median |  | 0.195 | 0.245 | 0.3 | 0.25 | 0.24894 | 0.255 | 0.26 | 0.27 | 0.27 | 0.28 | 0.25 | 0.24 | 0.24 | 0.25 | 0.215 | 0.2505 | 0.27 | 0.17 |  | 0.175 | 0.23893 | 0.2
| Std Dev. |  | 0.0098319 | 0.022509 | 0 | 0.0083666 | 0.0063236 | 0.0054772 | 0.013663 | 0.018348 | 0.0063246 | 0.010954 | 0.012247 | 0.005164 | 0.0089443 | 0.0063246 | 0.0054772 | 0.0095376 | 0.0040825 | 0.005164 |  | 0.01633 | 0.0024898 | 0.01633
| Rel.Std.Dev. |  | 5.13% | 9.511% | 0% | 3.415% | 2.566% | 2.148% | 5.188% | 6.754% | 2.342% | 3.912% | 4.803% | 2.182% | 3.727% | 2.53% | 2.548% | 3.874% | 1.503% | 2.979% |  | 9.243% | 1.038% | 8.031%
| PDM3 |  | -23.72% | -5.816% | 19.39% | -2.5% | -1.925% | 1.48% | 4.796% | 8.112% | 7.449% | 11.43% | 1.48% | -5.816% | -4.49% | -0.5103% | -14.44% | -2.036% | 8.112% | -31.02% |  | -29.69% | -4.568% | -19.08%
| Z-Score (Absolute) |  | 2.14 | 0.37 | NA | 0.05 | 0.01 | 0.35 | 0.67 | 1 | 0.94 | 1.33 | 0.35 | 0.37 | 0.24 | 0.15 | 1.22 | 0 | 1 | 2.86 | 88.53 | 2.73 | 0.25 | 1.68
Results from laboratory 2.04 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 606b (Indicative Value 0.047 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.12; 2.07: 2.23 |  |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS |  |  |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 01g; 2.07: 0.5g |  |  |  |  | 3
| 1 | 1 | 2.01: 0.05; 2.03: 0.048; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0495 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.05; 2.03: 0.047; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0444 |  |  |  |  | 5
| 1 | 3 | 2.01: 0.05; 2.03: 0.047; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0441 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.04; 2.03: 0.048; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0425 |  |  |  |  | 0.047479
| 1 | 5 | 2.01: 0.04; 2.03: 0.048; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0434 |  |  |  |  | 11
| 1 | 6 | 2.01: 0.05; 2.03: 0.049; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0486
| Mean |  | 2.01: 0.046667; 2.03: 0.047833; 2.06: 0.05; 2.07: 0.045417
| Median |  | 2.01: 0.05; 2.03: 0.048; 2.06: 0.05; 2.07: 0.04425
| Std Dev. |  | 2.01: 0.005164; 2.03: 0.00075277; 2.06: 7.6012e-18; 2.07: 0.0029034
| Rel.Std.Dev. |  | 2.01: 11.07%; 2.03: 1.574%; 2.06: 1.52e-14%; 2.07: 6.393%
| PDM3 |  | 2.01: -1.711%; 2.03: 0.7459%; 2.06: 5.309%; 2.07: -4.344%
| Z-Score (Absolute) |  | 2.01: 0.17; 2.03: 0.17; 2.04: 6.55; 2.05: 0.81; 2.06: 0.81; 2.07: 0.54
| Analytical results for U in OREAS 606b (Certified Value 2.76 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  | 3
| 1 | 1 | 3.07 | 2.63 | 2.8 | 2.78 | 2.6464 | 2.8 | 2.66 | 3.2873 | 2.94 | 3.19 | 3.13 | 2.7 | 2.57 | 2.58 | 3.07 | 2.5 | 3.22 | 2.8 | 2.51 | 2.73 | 2.9102 | 2.13 | 2.7566 | 2.92 |  |  | 1
| 1 | 2 | 2.78 | 2.47 | 2.8 | 2.8 | 2.7535 | 2.9 | 2.71 | 3.3649 | 2.75 | 3.06 | 3.09 | 2.85 | 2.43 | 2.62 | 2.88 | 2.46 | 3.13 | 2.95 | 2.57 | 2.69 | 2.6463 | 2.27 | 2.7375 | 2.75 |  |  | 32
| 1 | 3 | 2.92 | 2.72 | 2.8 | 2.77 | 2.5975 | 2.7 | 2.63 | 3.3236 | 2.75 | 3.07 | 3.16 | 2.81 | 2.53 | 2.69 | 2.79 | 2.55 | 3.08 | 2.59 | 2.38 | 2.54 | 2.579 | 2.09 | 2.7346 | 2.57 |  |  | 16
| 1 | 4 | 3.1 | 2.62 | 2.8 | 2.8 | 2.5702 | 2.9 | 2.64 | 3.3835 | 2.79 | 3.07 | 3.07 | 2.65 | 2.43 | 2.62 | 2.84 | 2.44 | 3.17 | 3.1 | 2.73 | 2.56 | 2.5723 | 2.1 | 2.7566 | 2.69 |  |  | 2.76
| 1 | 5 | 2.91 | 2.42 | 2.8 | 2.73 | 2.664 | 2.9 | 3.01 | 3.4221 | 2.8 | 3.12 | 3.21 | 2.77 | 2.45 | 2.67 | 2.85 | 2.45 | 3.26 | 2.69 | 2.61 | 2.53 | 2.6575 | 2.19 | 2.7227 | 2.61 |  |  | 128
| 1 | 6 | 3.11 | 2.64 | 2.7 | 2.79 | 2.6385 | 2.8 | 2.71 | 3.2285 | 2.84 | 3.03 | 3.1 | 2.86 | 2.49 | 2.66 | 2.83 | 2.44 | 2.93 | 2.74 | 2.57 | 2.67 | 2.9704 | 2.11 | 2.7209 | 2.41
| Mean |  | 2.9817 | 2.5833 | 2.7833 | 2.7783 | 2.645 | 2.8333 | 2.7267 | 3.335 | 2.8117 | 3.09 | 3.1267 | 2.7733 | 2.4833 | 2.64 | 2.8767 | 2.4733 | 3.1317 | 2.8117 | 2.5617 | 2.62 | 2.7226 | 2.1483 | 2.7381 | 2.6583
| Median |  | 2.995 | 2.625 | 2.8 | 2.785 | 2.6425 | 2.85 | 2.685 | 3.3443 | 2.795 | 3.07 | 3.115 | 2.79 | 2.47 | 2.64 | 2.845 | 2.455 | 3.15 | 2.77 | 2.57 | 2.615 | 2.6519 | 2.12 | 2.736 | 2.65
| Std Dev. |  | 0.13258 | 0.11396 | 0.040825 | 0.026394 | 0.063276 | 0.08165 | 0.14292 | 0.070114 | 0.071391 | 0.056921 | 0.05164 | 0.084063 | 0.057504 | 0.040497 | 0.09913 | 0.043665 | 0.11754 | 0.18519 | 0.11531 | 0.086718 | 0.17313 | 0.069402 | 0.015704 | 0.17302
| Rel.Std.Dev. |  | 4.446% | 4.411% | 1.467% | 0.95% | 2.392% | 2.882% | 5.242% | 2.102% | 2.539% | 1.842% | 1.652% | 3.031% | 2.316% | 1.534% | 3.446% | 1.765% | 3.753% | 6.587% | 4.501% | 3.31% | 6.359% | 3.231% | 0.5735% | 6.509%
| PDM3 |  | 8.031% | -6.401% | 0.8452% | 0.6641% | -4.167% | 2.657% | -1.208% | 20.83% | 1.872% | 11.96% | 13.28% | 0.4829% | -10.02% | -4.348% | 4.227% | -10.39% | 13.47% | 1.872% | -7.186% | -5.073% | -1.355% | -22.16% | -0.7921% | -3.684%
| Z-Score (Absolute) |  | 1.29 | 0.98 | 0.16 | 0.13 | 0.63 | 0.44 | 0.17 | 3.3 | 0.32 | 1.9 | 2.11 | 0.1 | 1.55 | 0.66 | 0.69 | 1.61 | 2.14 | 0.32 | 1.11 | 0.77 | 0.19 | 3.46 | 0.1 | 0.56
| Analytical results for V in OREAS 606b (Certified Value 1.97 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 2 | 3 | 2 | 2 | 2.1232 | 2 | 1.6 | 2.53 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 3 | 1.8 | 2 | 1.6194 | 1.4 | 2.13 | < 10 |  |  |  | 1
| 1 | 2 | 2 | 2 | 2 | 2 | 2.0829 | 2 | 1.6 | 2.42 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2.6 | 1.7 | 2 | 1.1079 | 1.5 | 2.0756 | < 10 |  |  |  | 33
| 1 | 3 | 2 | 3 | 2 | 2 | 2.1125 | 2 | 1.5 | 2.49 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 3 | 1.7 | 2 | 1.312 | 1.6 | 2.025 | < 10 |  |  |  | 16
| 1 | 4 | 2 | 3 | 2 |  | 2.1232 | 2 | 1.6 | 2.42 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 1.7 | 2 | 1.4687 | 1.6 | 2.0729 | < 10 |  |  |  | 1.9682
| 1 | 5 | 2 | 2 | 2 | 2 | 2.0943 | 3 | 1.8 | 2.66 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 3 | 1.7 | 2 | 1.2114 | 1.5 | 2.1172 | < 10 |  |  |  | 129
| 1 | 6 | 2 | 3 | 2 | 2 | 2.13 | 3 | 1.6 | 2.35 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | < 1 | 2 | 1.7 | 2 | 1.2235 | 1.6 | 2.0466 | < 10
| Mean |  | 2 | 2.6667 | 2 | 2 | 2.111 | 2.3333 | 1.6167 | 2.4783 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2.6 | 1.7167 | 2 | 1.3238 | 1.5333 | 2.0779
| Median |  | 2 | 3 | 2 | 2 | 2.1178 | 2 | 1.6 | 2.455 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2.8 | 1.7 | 2 | 1.2677 | 1.55 | 2.0743
| Std Dev. |  | 0 | 0.5164 | 0 | 0 | 0.018565 | 0.5164 | 0.098319 | 0.1087 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.4899 | 0.040825 | 0 | 0.18858 | 0.08165 | 0.040162
| Rel.Std.Dev. |  | 0% | 19.36% | 0% | 0% | 0.8794% | 22.13% | 6.082% | 4.386% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 18.84% | 2.378% | 0% | 14.25% | 5.325% | 1.933%
| PDM3 |  | 1.618% | 35.49% | 1.618% | 1.618% | 7.259% | 18.55% | -17.86% | 25.92% | 52.43% | 1.618% | 1.618% | 1.618% | 1.618% | 1.618% | 1.618% | 52.43% | 32.1% | -12.78% | 1.618% | -32.74% | -22.09% | 5.576%
| Z-Score (Absolute) |  | 0 | 4.05 | 0 | 0 | 0.67 | 2.02 | 2.33 | 2.91 | 6.07 | 0 | 0 | 0 | 0 | 0 | 0 | 3.54 | 3.64 | 1.72 | 0 | 4.11 | 2.83 | 0.47 | 18.22
| Analytical results for W in OREAS 606b (Certified Value 0.52 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  |  | 2
| 1 | 1 | 1.44 | 0.66 | 0.4 | 0.56 | < 1 | 0.7 | 0.57 | 0.6 | 0.5 | 0.34 | 0.56 | 0.49 | 0.47 | 0.53 | 0.52 | 0.78 | < 0.05 | < 5 | 0.43 | < 0.1 | 0.1 | 0.49907 | 2 |  |  |  | 1
| 1 | 2 | 1.26 | 0.46 | 0.4 | 0.54 | < 1 | 1.2 | 0.58 | 0.6 | 0.5 | 0.38 | 0.54 | 0.52 | 0.47 | 0.52 | 0.51 | 0.78 | < 0.05 | < 5 | 0.39 | < 0.1 | 0.1 | 0.5392 | 1 |  |  |  | 34
| 1 | 3 | 1.21 | 0.59 | 0.4 | 0.55 | < 1 | 0.6 | 0.59 | 0.57 | 0.6 | 0.34 | 0.54 | 0.51 | 0.49 | 0.5 | 0.53 | 0.75 | < 0.05 | < 5 | 0.34 | < 0.1 | < 0.1 | 0.47395 | 2 |  |  |  | 16
| 1 | 4 | 1.41 | 0.52 | 0.4 | 0.54 | < 1 | 0.8 | 0.59 | 0.57 | 0.6 | 0.34 | 0.51 | 0.49 | 0.48 | 0.5 | 0.51 | 0.72 | < 0.05 | < 5 | 0.41 | < 0.1 | < 0.1 | 0.57857 | 1 |  |  |  | 0.52141
| 1 | 5 | 1.26 | 0.49 | 0.4 | 0.53 | < 1 | 0.8 | 0.66 | 0.6 | 0.5 | 0.35 | 0.49 | 0.51 | 0.47 | 0.5 | 0.55 | 0.74 | < 0.05 | < 5 | 0.49 | < 0.1 | < 0.1 | 0.49772 | 2 |  |  |  | 130
| 1 | 6 | 1.51 | 0.62 | 0.4 | 0.52 | < 1 | 0.9 | 0.57 | 0.6 | 0.5 | 0.32 | 0.53 | 0.59 | 0.49 | 0.5 | 0.48 | 0.82 | < 0.05 | < 5 | 0.32 | < 0.1 | < 0.1 | 0.5645 | 1
| Mean |  | 1.3483 | 0.55667 | 0.4 | 0.54 |  | 0.83333 | 0.59333 | 0.59 | 0.53333 | 0.345 | 0.52833 | 0.51833 | 0.47833 | 0.50833 | 0.51667 | 0.765 |  |  | 0.39667 |  | 0.1 | 0.5255 | 1.5
| Median |  | 1.335 | 0.555 | 0.4 | 0.54 |  | 0.8 | 0.585 | 0.6 | 0.5 | 0.34 | 0.535 | 0.51 | 0.475 | 0.5 | 0.515 | 0.765 |  |  | 0.4 |  | 0.1 | 0.51913 | 1.5
| Std Dev. |  | 0.1209 | 0.078655 | 6.0809e-17 | 0.014142 |  | 0.20656 | 0.033862 | 0.015492 | 0.05164 | 0.019748 | 0.024833 | 0.037103 | 0.0098319 | 0.013292 | 0.023381 | 0.035637 |  |  | 0.06186 |  | 0 | 0.041601 | 0.54772
| Rel.Std.Dev. |  | 8.967% | 14.13% | 1.52e-14% | 2.619% |  | 24.79% | 5.707% | 2.626% | 9.682% | 5.724% | 4.7% | 7.158% | 2.055% | 2.615% | 4.525% | 4.658% |  |  | 15.59% |  | 0% | 7.916% | 36.51%
| PDM3 |  | 158.6% | 6.762% | -23.28% | 3.565% |  | 59.82% | 13.79% | 13.15% | 2.287% | -33.83% | 1.328% | -0.5899% | -8.261% | -2.508% | -0.9096% | 46.72% |  |  | -23.92% |  | -80.82% | 0.7848% | 187.7%
| Z-Score (Absolute) |  | 8.6 | 0.33 | NA | 0.15 | 0.27 | 3.22 | 0.71 | 0.67 | NA | 1.89 | 0.03 | 0.07 | 0.49 | 0.18 | 0.09 | 2.5 | 5.23 | 20.64 | 1.35 | 4.97 | NA | 0 | NA
Results from laboratories 2.04, 2.11 and 2.25 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 606b (Certified Value 7.75 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 3
| 1 | 1 | 8.66 | 6.98 | 7.85 | 7.9902 | 7.64 | 7.13 | 8.252 | 7.6 | 8.7 | 8.28 | 7.09 | 7.96 | 7.56 | 8.77 | 7.66 | 6.96 | 7.32 | 7.42 | 8.3219 | 6.02 | 8.0857 | < 10 | 8.62 |  |  |  | 1
| 1 | 2 | 8.82 | 7.05 | 7.82 | 7.8072 | 7.43 | 7.15 | 8.0564 | 7.6 | 8.53 | 8.24 | 7.71 | 7.93 | 8.17 | 8.16 | 7.58 | 7.24 | 7.21 | 7.26 | 7.6493 | 6.58 | 7.6769 | < 10 | 8.51 |  |  |  | 35
| 1 | 3 | 8.64 | 7.14 | 7.71 | 7.9791 | 7.41 | 7.05 | 8.5453 | 7.4 | 8.48 | 8.16 | 7.55 | 7.85 | 7.97 | 7.71 | 7.76 | 6.57 | 7.03 | 6.94 | 7.5812 | 5.91 | 8.0044 | < 10 | 7.71 |  |  |  | 16
| 1 | 4 | 8.87 | 7.13 | 7.81 | 7.9336 | 7.63 | 7 | 8.3564 | 7.6 | 8.83 | 7.99 | 6.98 | 7.94 | 8.02 | 7.62 | 7.51 | 6.5 | 7.32 | 6.93 | 7.3069 | 5.81 | 7.935 | < 10 | 8.29 |  |  |  | 7.7506
| 1 | 5 | 8.87 | 6.84 | 7.67 | 7.7688 | 7.98 | 7.96 | 8.1894 | 7.7 | 8.7 | 8.03 | 7.38 | 7.99 | 8.08 | 8.04 | 7.61 | 6.54 | 7.48 | 6.94 | 7.489 | 5.77 | 7.8668 | < 10 | 7.78 |  |  |  | 131
| 1 | 6 | 8.8 | 7.21 | 7.74 | 7.9006 | 7.34 | 7.19 | 8.1978 | 7.6 | 8.54 | 8.19 | 7.81 | 8.09 | 7.84 | 7.95 | 7.69 | 6.68 | 7.42 | 7.42 | 8.6725 | 5.87 | 7.7403 | < 10 | 7.36
| Mean |  | 8.7767 | 7.0583 | 7.7667 | 7.8966 | 7.5717 | 7.2467 | 8.2662 | 7.5833 | 8.63 | 8.1483 | 7.42 | 7.96 | 7.94 | 8.0417 | 7.635 | 6.7483 | 7.2967 | 7.1517 | 7.8368 | 5.9933 | 7.8849 |  | 8.045
| Median |  | 8.81 | 7.09 | 7.775 | 7.9171 | 7.53 | 7.14 | 8.2249 | 7.6 | 8.62 | 8.175 | 7.465 | 7.95 | 7.995 | 7.995 | 7.635 | 6.625 | 7.32 | 7.1 | 7.6152 | 5.89 | 7.9009 |  | 8.035
| Std Dev. |  | 0.10211 | 0.13318 | 0.070616 | 0.090866 | 0.23439 | 0.35624 | 0.16783 | 0.098319 | 0.13446 | 0.11548 | 0.33383 | 0.078994 | 0.21643 | 0.41004 | 0.087807 | 0.2926 | 0.16033 | 0.24269 | 0.5359 | 0.30018 | 0.15595 |  | 0.50171
| Rel.Std.Dev. |  | 1.163% | 1.887% | 0.9092% | 1.151% | 3.096% | 4.916% | 2.03% | 1.297% | 1.558% | 1.417% | 4.499% | 0.9924% | 2.726% | 5.099% | 1.15% | 4.336% | 2.197% | 3.393% | 6.838% | 5.009% | 1.978% |  | 6.236%
| PDM3 |  | 13.24% | -8.931% | 0.2078% | 1.884% | -2.308% | -6.501% | 6.653% | -2.158% | 11.35% | 5.132% | -4.265% | 2.702% | 2.444% | 3.756% | -1.491% | -12.93% | -5.856% | -7.727% | 1.113% | -22.67% | 1.733% |  | 3.799%
| Z-Score (Absolute) |  | 1.96 | 1.38 | 0 | 0.25 | 0.38 | 1.01 | 0.97 | 0.36 | 1.68 | 0.74 | 0.67 | 0.38 | 0.34 | 0.53 | 0.26 | 1.98 | 0.91 | 1.2 | 0.14 | 3.45 | 0.23 | 5.38 | 0.54
| Analytical results for Yb in OREAS 606b (Certified Value 0.23 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.09: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.25g; 2.08: 01g; 2.09: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  |  | 2
| 1 | 1 | 2.01: 0.23; 2.03: 0.22; 2.05: 0.23326; 2.06: 0.2; 2.07: 0.3017; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2696; 2.11: 0.18; 2.12: 0.28 |  |  |  | 1
| 1 | 2 | 2.01: 0.2; 2.03: 0.222; 2.05: 0.22189; 2.06: 0.2; 2.07: 0.3239; 2.08: 0.25; 2.09: 0.3; 2.10: 0.2441; 2.11: 0.2; 2.12: 0.26 |  |  |  | 36
| 1 | 3 | 2.01: 0.25; 2.03: 0.226; 2.05: 0.23266; 2.06: 0.2; 2.07: 0.3235; 2.08: 0.23; 2.09: 0.3; 2.10: 0.2426; 2.11: 0.19; 2.12: 0.23 |  |  |  | 16
| 1 | 4 | 2.01: 0.2; 2.03: 0.219; 2.05: 0.23656; 2.06: 0.2; 2.07: 0.3144; 2.08: 0.21; 2.09: 0.2; 2.10: 0.2524; 2.11: 0.18; 2.12: 0.35 |  |  |  | 0.22761
| 1 | 5 | 2.01: 0.2; 2.03: 0.218; 2.05: 0.22229; 2.06: 0.2; 2.07: 0.3059; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2467; 2.11: 0.18; 2.12: 0.24 |  |  |  | 132
| 1 | 6 | 2.01: 0.23; 2.03: 0.221; 2.05: 0.23271; 2.06: 0.2; 2.07: 0.3031; 2.08: 0.24; 2.09: 0.2; 2.10: 0.2747; 2.11: 0.18; 2.12: 0.21
| Mean |  | 2.01: 0.21833; 2.03: 0.221; 2.05: 0.22989; 2.06: 0.2; 2.07: 0.31208; 2.08: 0.235; 2.09: 0.26667; 2.10: 0.25502; 2.11: 0.185; 2.12: 0.26167
| Median |  | 2.01: 0.215; 2.03: 0.2205; 2.05: 0.23269; 2.06: 0.2; 2.07: 0.31015; 2.08: 0.24; 2.09: 0.3; 2.10: 0.24955; 2.11: 0.18; 2.12: 0.25
| Std Dev. |  | 2.01: 0.02137; 2.03: 0.0028284; 2.05: 0.0062183; 2.06: 3.0405e-17; 2.07: 0.010021; 2.08: 0.013784; 2.09: 0.05164; 2.10: 0.01378; 2.11: 0.0083666; 2.12: 0.049565
| Rel.Std.Dev. |  | 2.01: 9.788%; 2.03: 1.28%; 2.05: 2.705%; 2.06: 1.52e-14%; 2.07: 3.211%; 2.08: 5.866%; 2.09: 19.36%; 2.10: 5.404%; 2.11: 4.522%; 2.12: 18.94%
| PDM3 |  | 2.01: -4.074%; 2.03: -2.902%; 2.05: 1.005%; 2.06: -12.13%; 2.07: 37.12%; 2.08: 3.248%; 2.09: 17.16%; 2.10: 12.04%; 2.11: -18.72%; 2.12: 14.96%
| Z-Score (Absolute) |  | 2.01: 0.52; 2.03: 0.42; 2.05: 0.09; 2.06: NA; 2.07: 2.93; 2.08: 0.09; 2.09: NA; 2.10: 0.83; 2.11: 1.74; 2.12: 1.07
Results from laboratories 2.07 and 2.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 606b (Certified Value 282 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 285 | 269.6 | 291 | 287 | 289.3109 | 307 | 280 | 265.532 | 286 | 296 | 275 | 284 | 283 | 281 | 341 | 294 | 309 | 246 | 363 | 246.6 | 274 | 269.3556 | 286 | 293.6832 | 259.359 | 311 | 1
| 1 | 2 | 287 | 253.3 | 290 | 287 | 283.6654 | 309 | 282 | 259.6835 | 282 | 288 | 274 | 280 | 283 | 281 | 327 | 290 | 303 | 256 | 351 | 248.8 | 275 | 256.3042 | 292 | 282.0695 | 251.571 | 286 | 17
| 1 | 3 | 288 | 277.7 | 287 | 282 | 287.6555 | 305 | 281 | 263.5015 | 289 | 295 | 274 | 285 | 281 | 280 | 323 | 294 | 306 | 240 | 390 | 256.4 | 275 | 265.7514 | 287 | 287.2034 | 251.399 | 277 | 16
| 1 | 4 | 285 | 266.8 | 291 | 281 | 284.7105 | 311 | 280 | 266.5925 | 286 | 296 | 265 | 282 | 387 | 280 | 320 | 285 | 308 | 236 | 340 | 251.9 | 277 | 254.7049 | 288 | 290.7281 | 252.183 | 278 | 282.1638
| 1 | 5 | 291 | 272.9 | 295 | 279 | 280.8547 | 326 | 316 | 260.043 | 288 | 304 | 266 | 289 | 286 | 281 | 319 | 286 | 312 | 241 | 349 | 254.4 | 276 | 267.2466 | 283 | 288.8637 | 254.743 | 274 | 133
| 1 | 6 | 273 | 265.1 | 285 | 279 | 288.4437 | 307 | 286 | 269.3015 | 285 | 295 | 274 | 283 | 286 | 280 | 331 | 284 | 305 | 243 | 304 | 249.4 | 276 | 274.7387 | 289 | 281.8574 | 253.563 | 260
| Mean |  | 284.8333 | 267.5667 | 289.8333 | 282.5 | 285.7735 | 310.8333 | 287.5 | 264.109 | 286 | 295.6667 | 271.3333 | 283.8333 | 301 | 280.5 | 326.8333 | 288.8333 | 307.1667 | 243.6667 | 349.5 | 251.25 | 275.5 | 264.6836 | 287.5 | 287.4009 | 253.803 | 281
| Median |  | 286 | 268.2 | 290.5 | 281.5 | 286.183 | 308 | 281.5 | 264.5168 | 286 | 295.5 | 274 | 283.5 | 284.5 | 280.5 | 325 | 288 | 307 | 242 | 350 | 250.65 | 275.5 | 266.499 | 287.5 | 288.0336 | 252.873 | 277.5
| Std Dev. |  | 6.2102 | 8.3123 | 3.4881 | 3.6742 | 3.2543 | 7.705 | 14.1386 | 3.7853 | 2.4495 | 5.0859 | 4.5461 | 3.0605 | 42.1758 | 0.54772 | 8.2563 | 4.4907 | 3.1885 | 6.8896 | 28.233 | 3.6855 | 1.0488 | 7.7515 | 3.0166 | 4.7301 | 3.0065 | 16.9706
| Rel.Std.Dev. |  | 2.18% | 3.107% | 1.203% | 1.301% | 1.139% | 2.479% | 4.918% | 1.433% | 0.8565% | 1.72% | 1.675% | 1.078% | 14.01% | 0.1953% | 2.526% | 1.555% | 1.038% | 2.827% | 8.078% | 1.467% | 0.3807% | 2.929% | 1.049% | 1.646% | 1.185% | 6.039%
| PDM3 |  | 0.9461% | -5.173% | 2.718% | 0.1192% | 1.279% | 10.16% | 1.891% | -6.399% | 1.36% | 4.785% | -3.838% | 0.5917% | 6.676% | -0.5897% | 15.83% | 2.364% | 8.861% | -13.64% | 23.86% | -10.96% | -2.362% | -6.195% | 1.891% | 1.856% | -10.05% | -0.4124%
| Z-Score (Absolute) |  | 0.03 | 1.19 | 0.3 | 0.19 | 0.03 | 1.71 | 0.15 | 1.42 | 0.05 | 0.69 | 0.93 | 0.1 | 1.05 | 0.32 | 2.78 | 0.24 | 1.46 | 2.78 | 4.29 | 2.28 | 0.66 | 1.38 | 0.15 | 0.14 | 2.11 | 0.29
| Analytical results for Zr in OREAS 606b (Certified Value 43 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 1
| 1 | 1 | 42 | 40.3 | 42.6 | 41.7773 | 33.3 | 38.7 | 46.36 | 41.5 | 45 | 48.6 | 40.4 | 41.7 | 39.9 | 45.5 | 41.9 | 38.8 | 44.6 | 52 | 31 | 52.5413 | 27.5 | 38.5352 | 45.9 | 59.3 |  |  | 1
| 1 | 2 | 42 | 39.9 | 42 | 40.5989 | 16.2 | 39.3 | 47.321 | 44.2 | 46.2 | 48.6 | 41.6 | 41.9 | 41.3 | 44.1 | 40.2 | 38.6 | 47.1 | 53.3 | 31 | 47.1006 | 29.6 | 36.3736 | 45.02 | 54.7 |  |  | 18
| 1 | 3 | 42 | 39.7 | 41.7 | 40.9241 | 29.8 | 38.5 | 50.2505 | 42 | 46.5 | 48.5 | 40.7 | 42.1 | 42.2 | 41.6 | 42 | 35.9 | 42 | 55.9 | 30 | 47.9963 | 26.8 | 38.5415 | 45.27 | 38.6 |  |  | 16
| 1 | 4 | 42 | 38.6 | 41.8 | 41.0397 | 20.5 | 38.4 | 46.481 | 40.5 | 46.5 | 46.5 | 39.6 | 42.2 | 41.7 | 41.8 | 40.8 | 35.4 | 43.9 | 55.9 | 30 | 49.1457 | 26.5 | 38.6422 | 45.41 | 52.1 |  |  | 42.9689
| 1 | 5 | 41 | 38.2 | 41 | 41.4414 | 29.8 | 43.7 | 50.7335 | 38.9 | 47.6 | 48.3 | 42.3 | 41.7 | 42.1 | 43.1 | 40.9 | 36.7 | 47.3 | 54.9 | 30 | 49.3217 | 25.1 | 36.8929 | 45.01 | 42.1 |  |  | 134
| 1 | 6 | 42 | 39.5 | 40.8 | 41.1062 | 18.3 | 39.9 | 48.8225 | 42.2 | 46.7 | 48.7 | 41.9 | 42.1 | 41.7 | 42.2 | 41.1 | 32.9 | 45.4 | 52.3 | 31 | 53.9816 | 26.3 | 37.5122 | 44.54 | 50.5
| Mean |  | 41.8333 | 39.3667 | 41.65 | 41.1479 | 24.65 | 39.75 | 48.3281 | 41.55 | 46.4167 | 48.2 | 41.0833 | 41.95 | 41.4833 | 43.05 | 41.15 | 36.3833 | 45.05 | 54.05 | 30.5 | 50.0145 | 26.9667 | 37.7496 | 45.1917 | 49.55
| Median |  | 42 | 39.6 | 41.75 | 41.073 | 25.15 | 39 | 48.0718 | 41.75 | 46.5 | 48.55 | 41.15 | 42 | 41.7 | 42.65 | 41 | 36.3 | 45 | 54.1 | 30.5 | 49.2337 | 26.65 | 38.0237 | 45.145 | 51.3
| Std Dev. |  | 0.40825 | 0.80416 | 0.66257 | 0.41164 | 7.1668 | 2.0157 | 1.8987 | 1.7762 | 0.84242 | 0.8438 | 1.0226 | 0.21679 | 0.84004 | 1.5162 | 0.6892 | 2.1995 | 2.0107 | 1.7547 | 0.54772 | 2.6803 | 1.5095 | 0.97209 | 0.45666 | 7.8025
| Rel.Std.Dev. |  | 0.9759% | 2.043% | 1.591% | 1% | 29.07% | 5.071% | 3.929% | 4.275% | 1.815% | 1.751% | 2.489% | 0.5168% | 2.025% | 3.522% | 1.675% | 6.045% | 4.463% | 3.246% | 1.796% | 5.359% | 5.598% | 2.575% | 1.01% | 15.75%
| PDM3 |  | -2.643% | -8.383% | -3.069% | -4.238% | -42.63% | -7.491% | 12.47% | -3.302% | 8.024% | 12.17% | -4.388% | -2.371% | -3.457% | 0.1888% | -4.233% | -15.33% | 4.843% | 25.79% | -29.02% | 16.4% | -37.24% | -12.15% | 5.173% | 15.32%
| Z-Score (Absolute) |  | 0.04 | 0.43 | 0.01 | 0.09 | 3.25 | 0.35 | 1.29 | 0.01 | 0.92 | 1.26 | 0.1 | 0.07 | 0.02 | 0.28 | 0.09 | 1 | 0.66 | 2.38 | 2.13 | 1.61 | 2.8 | 0.74 | 0.69 | 1.52
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C875: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C893: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ALK
| Analytical results for S-(Sulphate) in OREAS 606b (Indicative Value 0.245 wt.%) | [18 empty] | Indicative
S-(Sulphate) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.25 | 2.26 |  | wt.%
|  |  | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | Na2CO3*IRC | IRC-x*Calc | IRC-x*Calc | Na2CO3*GRAV | IRC-x*Calc | IRC-x*Calc | Na2CO3*GRAV | I-550*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC | Na2CO3*GRAV |  | 3
3
| 1 | 1 | 0.11 | 0.21 | 0.29705 | 0.2663 | 0.33376 | 0.25 | 0.23 | 0.15 | 0.22 | 0.24 | 0.05 | 0.15687 | 0.36 | 0.39 | 0.27 | 0.30079 |  | 1
| 1 | 2 | 0.11 | 0.22 | 0.3104 | 0.26 | 0.3371 | 0.27 | 0.25 | 0.16 | 0.16 | 0.22 | 0.07 | 0.17022 | 0.36 | 0.39 | 0.27 | 0.30909 |  | 7
| 1 | 3 | 0.11 | 0.21 | 0.31374 | 0.25385 | 0.33376 | 0.25 | 0.23 | 0.19 | 0.22 | 0.28 | 0.06 | 0.14018 | 0.37 | 0.4 | 0.28 | 0.30488 |  | 16
| 1 | 4 | 0.12 | 0.21 | 0.29037 | 0.26175 | 0.33043 | 0.26 | 0.27 | 0.23 | 0.24 | 0.24 | 0.07 | 0.15687 | 0.37 | 0.38 | 0.27 | 0.30356 |  | 0.24491
| 1 | 5 | 0.11 | 0.21 | 0.30706 | 0.2446 | 0.3371 | 0.27 | 0.22 | 0.2 | 0.23 | 0.24 | 0.08 | 0.13017 | 0.36 | 0.38 | 0.26 | 0.30765 |  | 13
| 1 | 6 | 0.11 | 0.22 | 0.33376 | 0.2592 | 0.33376 | 0.26 | 0.25 | 0.16 | 0.2 | 0.24 | 0.1 | 0.15687 | 0.36 | 0.39 | 0.28 | 0.3021
| Mean |  | 0.11167 | 0.21333 | 0.30873 | 0.25762 | 0.33432 | 0.26 | 0.24167 | 0.18167 | 0.21167 | 0.24333 | 0.071667 | 0.15186 | 0.36333 | 0.38833 | 0.27167 | 0.30468
| Median |  | 0.11 | 0.21 | 0.30873 | 0.2596 | 0.33376 | 0.26 | 0.24 | 0.175 | 0.22 | 0.24 | 0.07 | 0.15687 | 0.36 | 0.39 | 0.27 | 0.30422
| Std Dev. |  | 0.0040825 | 0.005164 | 0.015038 | 0.0075405 | 0.0025125 | 0.0089443 | 0.018348 | 0.030605 | 0.028577 | 0.019664 | 0.017224 | 0.014278 | 0.005164 | 0.0075277 | 0.0075277 | 0.0032058
| Rel.Std.Dev. |  | 3.656% | 2.421% | 4.871% | 2.927% | 0.7515% | 3.44% | 7.592% | 16.85% | 13.5% | 8.081% | 24.03% | 9.402% | 1.421% | 1.938% | 2.771% | 1.052%
| PDM3 |  | -54.4% | -12.89% | 26.06% | 5.189% | 36.51% | 6.162% | -1.324% | -25.82% | -13.57% | -0.6433% | -70.74% | -37.99% | 48.35% | 58.56% | 10.93% | 24.4%
| Z-Score (Absolute) |  | 1.66 | 0.45 | 0.7 | 0.09 | 1.01 | 0.11 | 0.11 | 0.83 | 0.47 | 0.09 | 2.14 | 1.18 | 1.35 | 1.65 | 0.25 | 0.65
| Analytical results for S-(Sulphide) in OREAS 606b (Indicative Value 0.208 wt.%) | [18 empty] | Indicative
S-(Sulphide) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.25 | 2.26 | wt.%
|  |  | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | IRC | Na2CO3*IRC | IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | HCl*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC | HCl*IRC | 3
3
| 1 | 1 | 0.39 | 0.25 | 0.16 | 0.28 | 0.36005 | 0.13 | 0.2 | 0.15 | 0.33 | 0.2 | 0.18 | 0.39 | 0.07 | 0.11 | 0.05 | 0.18 | 0.18115 | 1
| 1 | 2 | 0.38 | 0.24 | 0.15 | 0.27 | 0.36755 | 0.12 | 0.19 | 0.15 | 0.33 | 0.27 | 0.13 | 0.37 | 0.05 | 0.12 | 0.05 | 0.17 | 0.18 | 7
| 1 | 3 | 0.38 | 0.24 | 0.15 | 0.28 | 0.3856 | 0.12 | 0.21 | 0.16 | 0.29 | 0.21 | 0.15 | 0.38 | 0.06 | 0.1 | < 0.05 | 0.17 | 0.17763 | 16
| 1 | 4 | 0.37 | 0.25 | 0.18 | 0.3 | 0.37925 | 0.13 | 0.19 | 0.15 | 0.26 | 0.2 | 0.12 | 0.38 | 0.11 | 0.1 | 0.05 | 0.18 | 0.18621 | 0.20838
| 1 | 5 | 0.39 | 0.24 | 0.15 | 0.32 | 0.3555 | 0.13 | 0.19 | 0.19 | 0.29 | 0.21 | 0.12 | 0.37 | 0.09 | 0.12 | 0.06 | 0.18 | 0.18514 | 13
| 1 | 6 | 0.37 | 0.24 | 0.13 | 0.27 | 0.36635 | 0.12 | 0.2 | 0.19 | 0.33 | 0.22 | 0.15 | 0.35 | 0.07 | 0.12 | 0.05 | 0.17 | 0.18021
| Mean |  | 0.38 | 0.24333 | 0.15333 | 0.28667 | 0.36905 | 0.125 | 0.19667 | 0.165 | 0.305 | 0.21833 | 0.14167 | 0.37333 | 0.075 | 0.11167 | 0.052 | 0.175 | 0.18172
| Median |  | 0.38 | 0.24 | 0.15 | 0.28 | 0.36695 | 0.125 | 0.195 | 0.155 | 0.31 | 0.21 | 0.14 | 0.375 | 0.07 | 0.115 | 0.05 | 0.175 | 0.18068
| Std Dev. |  | 0.0089443 | 0.005164 | 0.01633 | 0.019664 | 0.01142 | 0.0054772 | 0.008165 | 0.019748 | 0.029496 | 0.026394 | 0.023166 | 0.013663 | 0.021679 | 0.0098319 | 0.0044721 | 0.0054772 | 0.0032888
| Rel.Std.Dev. |  | 2.354% | 2.122% | 10.65% | 6.859% | 3.094% | 4.382% | 4.152% | 11.97% | 9.671% | 12.09% | 16.35% | 3.66% | 28.91% | 8.805% | 8.6% | 3.13% | 1.81%
| PDM3 |  | 82.36% | 16.77% | -26.42% | 37.57% | 77.11% | -40.01% | -5.621% | -20.82% | 46.37% | 4.777% | -32.01% | 79.16% | -64.01% | -46.41% | -75.05% | -16.02% | -12.79%
| Z-Score (Absolute) |  | 2.17 | 0.67 | 0.31 | 1.15 | 2.05 | 0.62 | 0.16 | 0.18 | 1.35 | 0.4 | 0.44 | 2.1 | 1.17 | 0.77 | 1.47 | 0.07 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Acid Leach
| Analytical results for S-(Sulphate) in OREAS 606b (Indicative Value 0.158 wt.%) |  |  | Indicative
S-(Sulphate) | Round | Replicate | Lab | 1
| No | No | 2.05 | wt.%
|  |  | HCl*OES | 3
3
| 1 | 1 | 0.16 | 1
| 1 | 2 | 0.17 | 7
| 1 | 3 | 0.16 | 16
| 1 | 4 | 0.15 | 0.15833
| 1 | 5 | 0.15 | 15
| 1 | 6 | 0.16
| Mean |  | 0.15833
| Median |  | 0.16
| Std Dev. |  | 0.0075277
| Rel.Std.Dev. |  | 4.754%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 606b (Indicative Value 14.86 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.86 | 1
| 1 | 2 | 14.85 | 11
| Mean |  | 14.855 | 16
| Median |  | 14.855 | 14.855
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.0476%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 606b (Indicative Value 80 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 80 | 12
| Mean |  | 80 | 16
| Median |  | 80 | 80
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 606b (Indicative Value 3439 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 3450 | 1
| 1 | 2 | 3428 | 13
| Mean |  | 3439 | 16
| Median |  | 3439 | 3438.82
| Std Dev. |  | 15.5563 | 19
| Rel.Std.Dev. |  | 0.4524%
| PDM3 |  | 0.005234%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 606b (Indicative Value 1.54 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.54 | 14
| Mean |  | 1.535 | 16
| Median |  | 1.535 | 1.535
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.4607%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 606b (Indicative Value 20 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 30 | 15
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 14.1421 | 21
| Rel.Std.Dev. |  | 70.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 606b (Indicative Value < 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 16
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 22
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 606b (Indicative Value 36.5 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 29 | 1
| 1 | 2 | 44 | 17
| Mean |  | 36.5 | 16
| Median |  | 36.5 | 36.5375
| Std Dev. |  | 10.6066 | 23
| Rel.Std.Dev. |  | 29.06%
| PDM3 |  | -0.1026%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 606b (Indicative Value 0.036 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.036 | 1
| 1 | 2 | 0.036 | 18
| Mean |  | 0.036 | 16
| Median |  | 0.036 | 0.036
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 606b (Indicative Value 3.41 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.388 | 1
| 1 | 2 | 3.431 | 55
| Mean |  | 3.4095 | 16
| Median |  | 3.4095 | 3.4098
| Std Dev. |  | 0.030406 | 17
| Rel.Std.Dev. |  | 0.8918%
| PDM3 |  | -0.00981%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 606b (Indicative Value 3.69 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.7 | 1
| 1 | 2 | 3.68 | 12
| Mean |  | 3.69 | 16
| Median |  | 3.69 | 3.69
| Std Dev. |  | 0.014142 | 18
| Rel.Std.Dev. |  | 0.3833%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 606b (Indicative Value 0.24 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.24 | 1
| 1 | 2 | 0.24 | 13
| Mean |  | 0.24 | 16
| Median |  | 0.24 | 0.24
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 606b (Indicative Value 0.034 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.033 | 1
| 1 | 2 | 0.035 | 14
| Mean |  | 0.034 | 16
| Median |  | 0.034 | 0.034
| Std Dev. |  | 0.0014142 | 20
| Rel.Std.Dev. |  | 4.159%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 606b (Indicative Value 3.67 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.66 | 1
| 1 | 2 | 3.67 | 15
| Mean |  | 3.665 | 16
| Median |  | 3.665 | 3.665
| Std Dev. |  | 0.0070711 | 21
| Rel.Std.Dev. |  | 0.1929%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 606b (Indicative Value < 10 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 16
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 22
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 606b (Indicative Value 0.066 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0665 | 1
| 1 | 2 | 0.0665 | 17
| Mean |  | 0.0665 | 16
| Median |  | 0.0665 | 0.066456
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.06561%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 606b (Indicative Value 180 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 180 | 18
| Mean |  | 180 | 16
| Median |  | 180 | 180
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 606b (Indicative Value 0.473 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.473 | 1
| 1 | 2 | 0.473 | 11
| Mean |  | 0.473 | 16
| Median |  | 0.473 | 0.473
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 606b (Indicative Value 69.24 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69.28 | 1
| 1 | 2 | 69.19 | 12
| Mean |  | 69.235 | 16
| Median |  | 69.235 | 69.235
| Std Dev. |  | 0.06364 | 18
| Rel.Std.Dev. |  | 0.09192%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 606b (Indicative Value 55 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 60 | 13
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 7.0711 | 19
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 606b (Indicative Value 200 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 200 | 1
| 1 | 2 | 200 | 14
| Mean |  | 200 | 16
| Median |  | 200 | 200
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 606b (Indicative Value 0.203 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.202 | 1
| 1 | 2 | 0.204 | 15
| Mean |  | 0.203 | 16
| Median |  | 0.203 | 0.203
| Std Dev. |  | 0.0014142 | 21
| Rel.Std.Dev. |  | 0.6967%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 606b (Indicative Value 17.9 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 18 | 16
| Mean |  | 18 | 16
| Median |  | 18 | 17.852
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.829%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 606b (Indicative Value 320 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 320 | 1
| 1 | 2 | 320 | 17
| Mean |  | 320 | 16
| Median |  | 320 | 320
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 606b (Indicative Value 270 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 270 | 1
| 1 | 2 | 270 | 18
| Mean |  | 270 | 16
| Median |  | 270 | 270
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 606b (Indicative Value 1.98 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.99 | 1
| 1 | 2 | 1.96 | 20
| Mean |  | 1.975 | 16
| Median |  | 1.975 | 1.975
| Std Dev. |  | 0.021213 | 26
| Rel.Std.Dev. |  | 1.074%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 606b (Indicative Value 2.15 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2.2 | 22
| Mean |  | 2.15 | 16
| Median |  | 2.15 | 2.15
| Std Dev. |  | 0.070711 | 28
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 606b (Indicative Value 101 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 99.8 | 1
| 1 | 2 | 103 | 12
| Mean |  | 101.4 | 16
| Median |  | 101.4 | 101.4
| Std Dev. |  | 2.2627 | 29
| Rel.Std.Dev. |  | 2.232%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 606b (Indicative Value 2890 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2890 | 1
| 1 | 2 | 2890 | 24
| Mean |  | 2890 | 16
| Median |  | 2890 | 2890
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 606b (Indicative Value 3.3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.2 | 25
| Mean |  | 3.3 | 16
| Median |  | 3.3 | 3.3
| Std Dev. |  | 0.14142 | 31
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 606b (Indicative Value 6.15 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.96 | 1
| 1 | 2 | 6.34 | 26
| Mean |  | 6.15 | 16
| Median |  | 6.15 | 6.15
| Std Dev. |  | 0.2687 | 32
| Rel.Std.Dev. |  | 4.369%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 606b (Indicative Value 1.6 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.6 | 27
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 606b (Indicative Value 87 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 86.3 | 1
| 1 | 2 | 87.2 | 28
| Mean |  | 86.75 | 16
| Median |  | 86.75 | 86.75
| Std Dev. |  | 0.6364 | 34
| Rel.Std.Dev. |  | 0.7336%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 606b (Indicative Value 2.8 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 3 | 16
| Mean |  | 2.8 | 16
| Median |  | 2.8 | 2.8
| Std Dev. |  | 0.28284 | 35
| Rel.Std.Dev. |  | 10.1%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 606b (Indicative Value 18 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17 | 1
| 1 | 2 | 19 | 30
| Mean |  | 18 | 16
| Median |  | 18 | 18
| Std Dev. |  | 1.4142 | 36
| Rel.Std.Dev. |  | 7.857%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 606b (Indicative Value 6.16 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.07 | 1
| 1 | 2 | 6.24 | 31
| Mean |  | 6.155 | 16
| Median |  | 6.155 | 6.155
| Std Dev. |  | 0.12021 | 37
| Rel.Std.Dev. |  | 1.953%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 606b (Indicative Value 0.035 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0352 | 1
| 1 | 2 | 0.0348 | 18
| Mean |  | 0.035 | 16
| Median |  | 0.035 | 0.035
| Std Dev. |  | 0.00028284 | 38
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 606b (Indicative Value 3.48 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.44 | 1
| 1 | 2 | 3.52 | 33
| Mean |  | 3.48 | 16
| Median |  | 3.48 | 3.48
| Std Dev. |  | 0.056569 | 39
| Rel.Std.Dev. |  | 1.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 606b (Indicative Value 0.99 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.99 | 1
| 1 | 2 | 0.98 | 34
| Mean |  | 0.985 | 16
| Median |  | 0.985 | 0.985
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.7179%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 606b (Indicative Value 1.25 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.25 | 35
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 606b (Indicative Value 23.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23.3 | 1
| 1 | 2 | 23.3 | 36
| Mean |  | 23.3 | 16
| Median |  | 23.3 | 23.3
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 606b (Indicative Value 5.61 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.49 | 1
| 1 | 2 | 5.72 | 37
| Mean |  | 5.605 | 16
| Median |  | 5.605 | 5.605
| Std Dev. |  | 0.16263 | 43
| Rel.Std.Dev. |  | 2.902%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 606b (Indicative Value 1.53 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.65 | 38
| Mean |  | 1.525 | 16
| Median |  | 1.525 | 1.525
| Std Dev. |  | 0.17678 | 44
| Rel.Std.Dev. |  | 11.59%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 606b (Indicative Value 7.67 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.62 | 1
| 1 | 2 | 7.72 | 22
| Mean |  | 7.67 | 16
| Median |  | 7.67 | 7.67
| Std Dev. |  | 0.070711 | 28
| Rel.Std.Dev. |  | 0.9219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 606b (Indicative Value 0.52 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.51 | 1
| 1 | 2 | 0.52 | 23
| Mean |  | 0.515 | 16
| Median |  | 0.515 | 0.515
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 1.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 606b (Indicative Value 0.13 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.1 | 24
| Mean |  | 0.125 | 16
| Median |  | 0.125 | 0.125
| Std Dev. |  | 0.035355 | 30
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 606b (Indicative Value 44.9 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45.1 | 1
| 1 | 2 | 44.7 | 25
| Mean |  | 44.9 | 16
| Median |  | 44.9 | 44.9
| Std Dev. |  | 0.28284 | 31
| Rel.Std.Dev. |  | 0.6299%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 606b (Indicative Value 0.07 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.07 | 26
| Mean |  | 0.07 | 16
| Median |  | 0.07 | 0.07
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 606b (Indicative Value 0.029 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0287 | 1
| 1 | 2 | 0.0294 | 27
| Mean |  | 0.02905 | 16
| Median |  | 0.02905 | 0.02905
| Std Dev. |  | 0.00049497 | 33
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 606b (Indicative Value 4.3 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4.4 | 28
| Mean |  | 4.3 | 16
| Median |  | 4.3 | 4.3
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 606b (Indicative Value 18.3 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.2 | 1
| 1 | 2 | 18.3 | 29
| Mean |  | 18.25 | 16
| Median |  | 18.25 | 18.25
| Std Dev. |  | 0.070711 | 35
| Rel.Std.Dev. |  | 0.3875%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 606b (Indicative Value 38.4 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.6 | 1
| 1 | 2 | 38.1 | 30
| Mean |  | 38.35 | 16
| Median |  | 38.35 | 38.35
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 0.9219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 606b (Indicative Value 5 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 6 | 16
| Mean |  | 5 | 16
| Median |  | 5 | 5
| Std Dev. |  | 1.4142 | 37
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 606b (Indicative Value 189 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 188 | 1
| 1 | 2 | 190 | 18
| Mean |  | 189 | 16
| Median |  | 189 | 189
| Std Dev. |  | 1.4142 | 38
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 606b (Indicative Value 10.3 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.1 | 1
| 1 | 2 | 10.5 | 33
| Mean |  | 10.3 | 16
| Median |  | 10.3 | 10.3
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 2.746%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 606b (Indicative Value 137 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 136 | 1
| 1 | 2 | 137 | 34
| Mean |  | 136.5 | 16
| Median |  | 136.5 | 136.5
| Std Dev. |  | 0.70711 | 40
| Rel.Std.Dev. |  | 0.518%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 606b (Indicative Value 0.008 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | < 0.01 | 35
| Mean |  | 0.01 | 16
| Median |  | 0.01 | 0.0075
| Std Dev. |  |  | 41
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 606b (Indicative Value 7.7 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.5 | 1
| 1 | 2 | 7.9 | 36
| Mean |  | 7.7 | 16
| Median |  | 7.7 | 7.7
| Std Dev. |  | 0.28284 | 42
| Rel.Std.Dev. |  | 3.673%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 606b (Indicative Value 3.75 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 4 | 37
| Mean |  | 3.75 | 16
| Median |  | 3.75 | 3.75
| Std Dev. |  | 0.35355 | 43
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 606b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 38
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 606b (Indicative Value 7.33 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.28 | 1
| 1 | 2 | 7.38 | 22
| Mean |  | 7.33 | 16
| Median |  | 7.33 | 7.33
| Std Dev. |  | 0.070711 | 28
| Rel.Std.Dev. |  | 0.9647%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 606b (Indicative Value 4.5 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 5 | 13
| Mean |  | 4.5 | 16
| Median |  | 4.5 | 4.5
| Std Dev. |  | 0.70711 | 29
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 606b (Indicative Value 195 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 195 | 1
| 1 | 2 | 194 | 14
| Mean |  | 194.5 | 16
| Median |  | 194.5 | 194.5
| Std Dev. |  | 0.70711 | 30
| Rel.Std.Dev. |  | 0.3636%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 606b (Indicative Value 1.4 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.37 | 1
| 1 | 2 | 1.42 | 25
| Mean |  | 1.395 | 16
| Median |  | 1.395 | 1.395
| Std Dev. |  | 0.035355 | 31
| Rel.Std.Dev. |  | 2.534%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 606b (Indicative Value 0.75 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.75 | 1
| 1 | 2 | 0.75 | 26
| Mean |  | 0.75 | 16
| Median |  | 0.75 | 0.75
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 606b (Indicative Value 1 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 1.2 | 27
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0.28284 | 33
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 606b (Indicative Value 14.3 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.2 | 1
| 1 | 2 | 14.4 | 28
| Mean |  | 14.3 | 16
| Median |  | 14.3 | 14.3
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 0.989%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 606b (Indicative Value 0.119 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.115 | 1
| 1 | 2 | 0.123 | 29
| Mean |  | 0.119 | 16
| Median |  | 0.119 | 0.119
| Std Dev. |  | 0.0056569 | 35
| Rel.Std.Dev. |  | 4.754%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 606b (Indicative Value 1.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 30
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 606b (Indicative Value 0.12 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.12 | 5
| Mean |  | 0.115 | 16
| Median |  | 0.115 | 0.115
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 606b (Indicative Value 5.76 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.78 | 1
| 1 | 2 | 5.73 | 32
| Mean |  | 5.755 | 16
| Median |  | 5.755 | 5.755
| Std Dev. |  | 0.035355 | 38
| Rel.Std.Dev. |  | 0.6143%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 606b (Indicative Value 6.15 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.9 | 1
| 1 | 2 | 6.4 | 33
| Mean |  | 6.15 | 16
| Median |  | 6.15 | 6.15
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 5.749%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 606b (Indicative Value 1.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.5 | 1
| 1 | 2 | 1.5 | 34
| Mean |  | 1.5 | 16
| Median |  | 1.5 | 1.5
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 606b (Indicative Value 14.5 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.4 | 1
| 1 | 2 | 14.5 | 35
| Mean |  | 14.45 | 16
| Median |  | 14.45 | 14.45
| Std Dev. |  | 0.070711 | 41
| Rel.Std.Dev. |  | 0.4893%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 606b (Indicative Value 0.58 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.55 | 1
| 1 | 2 | 0.6 | 36
| Mean |  | 0.575 | 16
| Median |  | 0.575 | 0.575
| Std Dev. |  | 0.035355 | 42
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 606b (Indicative Value 373 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 370 | 1
| 1 | 2 | 375 | 17
| Mean |  | 372.5 | 16
| Median |  | 372.5 | 372.5
| Std Dev. |  | 3.5355 | 43
| Rel.Std.Dev. |  | 0.9491%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 606b (Indicative Value 275 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 274 | 1
| 1 | 2 | 276 | 18
| Mean |  | 275 | 16
| Median |  | 275 | 275
| Std Dev. |  | 1.4142 | 44
| Rel.Std.Dev. |  | 0.5143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 606b (Indicative Value 1.79 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.85 | 1
| 1 | 2 | 1.81 | 22
| 1 | 3 | 1.73 | 16
| 1 | 4 | 1.8 | 1.785
| 1 | 5 | 1.74 | 46
| 1 | 6 | 1.78
| Mean |  | 1.785
| Median |  | 1.79
| Std Dev. |  | 0.045056
| Rel.Std.Dev. |  | 2.524%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 606b (Indicative Value 14.28 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 14.6 | 1
| 1 | 2 | 14.2 | 11
| 1 | 3 | 14.3 | 16
| 1 | 4 | 14.1 | 14.2833
| 1 | 5 | 14.2 | 47
| 1 | 6 | 14.3
| Mean |  | 14.2833
| Median |  | 14.25
| Std Dev. |  | 0.17224
| Rel.Std.Dev. |  | 1.206%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 606b (Indicative Value 2797 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 2850 | 1
| 1 | 2 | 2770 | 24
| 1 | 3 | 2800 | 16
| 1 | 4 | 2760 | 2796.6667
| 1 | 5 | 2780 | 48
| 1 | 6 | 2820
| Mean |  | 2796.6667
| Median |  | 2790
| Std Dev. |  | 33.8625
| Rel.Std.Dev. |  | 1.211%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 606b (Indicative Value 3 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 3.1 | 25
| 1 | 3 | 2.9 | 16
| 1 | 4 | 3 | 3
| 1 | 5 | 3 | 49
| 1 | 6 | 3
| Mean |  | 3
| Median |  | 3
| Std Dev. |  | 0.063246
| Rel.Std.Dev. |  | 2.108%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 606b (Indicative Value 4.9 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5 | 1
| 1 | 2 | 5 | 26
| 1 | 3 | 4.9 | 16
| 1 | 4 | 4.8 | 4.9
| 1 | 5 | 4.8 | 50
| 1 | 6 | 4.9
| Mean |  | 4.9
| Median |  | 4.9
| Std Dev. |  | 0.089443
| Rel.Std.Dev. |  | 1.825%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 606b (Indicative Value 1.55 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.58 | 1
| 1 | 2 | 1.53 | 14
| 1 | 3 | 1.55 | 16
| 1 | 4 | 1.52 | 1.545
| 1 | 5 | 1.53 | 51
| 1 | 6 | 1.56
| Mean |  | 1.545
| Median |  | 1.54
| Std Dev. |  | 0.022583
| Rel.Std.Dev. |  | 1.462%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 606b (Indicative Value 1.25 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.3 | 27
| 1 | 3 | 1.2 | 16
| 1 | 4 | 1.2 | 1.25
| 1 | 5 | 1.2 | 52
| 1 | 6 | 1.4
| Mean |  | 1.25
| Median |  | 1.2
| Std Dev. |  | 0.083666
| Rel.Std.Dev. |  | 6.693%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 606b (Indicative Value 83 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 84 | 1
| 1 | 2 | 82 | 28
| 1 | 3 | 82 | 16
| 1 | 4 | 82 | 82.5
| 1 | 5 | 82 | 53
| 1 | 6 | 83
| Mean |  | 82.5
| Median |  | 82
| Std Dev. |  | 0.83666
| Rel.Std.Dev. |  | 1.014%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 606b (Indicative Value 2.14 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.18 | 1
| 1 | 2 | 2.21 | 16
| 1 | 3 | 2.12 | 16
| 1 | 4 | 2.14 | 2.1433
| 1 | 5 | 2.09 | 54
| 1 | 6 | 2.12
| Mean |  | 2.1433
| Median |  | 2.13
| Std Dev. |  | 0.044121
| Rel.Std.Dev. |  | 2.059%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 606b (Indicative Value 14.7 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 14 | 1
| 1 | 2 | 15 | 30
| 1 | 3 | 15 | 16
| 1 | 4 | 15 | 14.6667
| 1 | 5 | 14 | 55
| 1 | 6 | 15
| Mean |  | 14.6667
| Median |  | 15
| Std Dev. |  | 0.5164
| Rel.Std.Dev. |  | 3.521%
| PDM3 |  | -2.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 606b (Indicative Value 6 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.2 | 1
| 1 | 2 | 6.1 | 31
| 1 | 3 | 5.9 | 16
| 1 | 4 | 6 | 6
| 1 | 5 | 5.8 | 56
| 1 | 6 | 6
| Mean |  | 6
| Median |  | 6
| Std Dev. |  | 0.14142
| Rel.Std.Dev. |  | 2.357%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 606b (Indicative Value 0.033 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.0342 | 1
| 1 | 2 | 0.0335 | 18
| 1 | 3 | 0.0323 | 16
| 1 | 4 | 0.0332 | 0.032867
| 1 | 5 | 0.0316 | 57
| 1 | 6 | 0.0324
| Mean |  | 0.032867
| Median |  | 0.0328
| Std Dev. |  | 0.00094163
| Rel.Std.Dev. |  | 2.865%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 606b (Indicative Value 3.43 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.49 | 1
| 1 | 2 | 3.52 | 33
| 1 | 3 | 3.4 | 16
| 1 | 4 | 3.39 | 3.4283
| 1 | 5 | 3.37 | 58
| 1 | 6 | 3.4
| Mean |  | 3.4283
| Median |  | 3.4
| Std Dev. |  | 0.061128
| Rel.Std.Dev. |  | 1.783%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 606b (Indicative Value 0.97 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1.02 | 34
| 1 | 3 | 0.97 | 16
| 1 | 4 | 0.96 | 0.97333
| 1 | 5 | 0.95 | 59
| 1 | 6 | 0.94
| Mean |  | 0.97333
| Median |  | 0.965
| Std Dev. |  | 0.030768
| Rel.Std.Dev. |  | 3.161%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 606b (Indicative Value 1.83 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.88 | 1
| 1 | 2 | 1.88 | 35
| 1 | 3 | 1.81 | 16
| 1 | 4 | 1.78 | 1.825
| 1 | 5 | 1.77 | 60
| 1 | 6 | 1.83
| Mean |  | 1.825
| Median |  | 1.82
| Std Dev. |  | 0.047645
| Rel.Std.Dev. |  | 2.611%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 606b (Indicative Value 3.16 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.21 | 1
| 1 | 2 | 3.13 | 55
| 1 | 3 | 3.17 | 16
| 1 | 4 | 3.11 | 3.155
| 1 | 5 | 3.13 | 61
| 1 | 6 | 3.18
| Mean |  | 3.155
| Median |  | 3.15
| Std Dev. |  | 0.037815
| Rel.Std.Dev. |  | 1.199%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 606b (Indicative Value 21.4 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 21.9 | 1
| 1 | 2 | 21.6 | 36
| 1 | 3 | 21.1 | 16
| 1 | 4 | 21.9 | 21.4333
| 1 | 5 | 20.7 | 62
| 1 | 6 | 21.4
| Mean |  | 21.4333
| Median |  | 21.5
| Std Dev. |  | 0.47188
| Rel.Std.Dev. |  | 2.202%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 606b (Indicative Value 5.6 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 5.8 | 37
| 1 | 3 | 5.5 | 16
| 1 | 4 | 5.5 | 5.6
| 1 | 5 | 5.4 | 46
| 1 | 6 | 5.6
| Mean |  | 5.6
| Median |  | 5.55
| Std Dev. |  | 0.16733
| Rel.Std.Dev. |  | 2.988%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 606b (Indicative Value 6.38 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 6.5 | 22
| 1 | 3 | 6.2 | 16
| 1 | 4 | 6.3 | 6.3833
| 1 | 5 | 6.4 | 47
| 1 | 6 | 6.3
| Mean |  | 6.3833
| Median |  | 6.35
| Std Dev. |  | 0.1472
| Rel.Std.Dev. |  | 2.306%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 606b (Indicative Value 0.43 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.44 | 1
| 1 | 2 | 0.44 | 23
| 1 | 3 | 0.41 | 16
| 1 | 4 | 0.43 | 0.42667
| 1 | 5 | 0.42 | 48
| 1 | 6 | 0.42
| Mean |  | 0.42667
| Median |  | 0.425
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 2.838%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 606b (Indicative Value 3.56 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.63 | 1
| 1 | 2 | 3.54 | 12
| 1 | 3 | 3.5 | 16
| 1 | 4 | 3.52 | 3.5583
| 1 | 5 | 3.55 | 49
| 1 | 6 | 3.61
| Mean |  | 3.5583
| Median |  | 3.545
| Std Dev. |  | 0.051153
| Rel.Std.Dev. |  | 1.438%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 606b (Indicative Value 45.3 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 45 | 25
| 1 | 3 | 46 | 16
| 1 | 4 | 44 | 45.3333
| 1 | 5 | 45 | 50
| 1 | 6 | 46
| Mean |  | 45.3333
| Median |  | 45.5
| Std Dev. |  | 0.8165
| Rel.Std.Dev. |  | 1.801%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 606b (Indicative Value 31.5 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 32 | 1
| 1 | 2 | 31 | 45
| 1 | 3 | 32 | 16
| 1 | 4 | 31 | 31.5
| 1 | 5 | 31 | 51
| 1 | 6 | 32
| Mean |  | 31.5
| Median |  | 31.5
| Std Dev. |  | 0.54772
| Rel.Std.Dev. |  | 1.739%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 606b (Indicative Value 0.213 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.215 | 1
| 1 | 2 | 0.213 | 13
| 1 | 3 | 0.212 | 16
| 1 | 4 | 0.212 | 0.213
| 1 | 5 | 0.212 | 52
| 1 | 6 | 0.214
| Mean |  | 0.213
| Median |  | 0.2125
| Std Dev. |  | 0.0012649
| Rel.Std.Dev. |  | 0.5939%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 606b (Indicative Value 0.035 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.036 | 1
| 1 | 2 | 0.034 | 14
| 1 | 3 | 0.035 | 16
| 1 | 4 | 0.034 | 0.034833
| 1 | 5 | 0.035 | 53
| 1 | 6 | 0.035
| Mean |  | 0.034833
| Median |  | 0.035
| Std Dev. |  | 0.00075277
| Rel.Std.Dev. |  | 2.161%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 606b (Indicative Value 4.4 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.42 | 1
| 1 | 2 | 4.43 | 28
| 1 | 3 | 4.45 | 16
| 1 | 4 | 4.48 | 4.3983
| 1 | 5 | 4.29 | 54
| 1 | 6 | 4.32
| Mean |  | 4.3983
| Median |  | 4.425
| Std Dev. |  | 0.075741
| Rel.Std.Dev. |  | 1.722%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 606b (Indicative Value 3.57 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.64 | 1
| 1 | 2 | 3.56 | 15
| 1 | 3 | 3.49 | 16
| 1 | 4 | 3.53 | 3.5667
| 1 | 5 | 3.56 | 55
| 1 | 6 | 3.62
| Mean |  | 3.5667
| Median |  | 3.56
| Std Dev. |  | 0.055737
| Rel.Std.Dev. |  | 1.563%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 606b (Indicative Value 17.6 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 17.9 | 1
| 1 | 2 | 17.9 | 29
| 1 | 3 | 17.4 | 16
| 1 | 4 | 17.9 | 17.6167
| 1 | 5 | 17.2 | 56
| 1 | 6 | 17.4
| Mean |  | 17.6167
| Median |  | 17.65
| Std Dev. |  | 0.31885
| Rel.Std.Dev. |  | 1.81%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 606b (Indicative Value 36.6 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 37.9 | 1
| 1 | 2 | 37.3 | 30
| 1 | 3 | 36.1 | 16
| 1 | 4 | 36.9 | 36.6167
| 1 | 5 | 35.1 | 57
| 1 | 6 | 36.4
| Mean |  | 36.6167
| Median |  | 36.65
| Std Dev. |  | 0.98065
| Rel.Std.Dev. |  | 2.678%
| PDM3 |  | -9.992e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 606b (Indicative Value 4.12 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.7 | 1
| 1 | 2 | 4.2 | 16
| 1 | 3 | 4.1 | 16
| 1 | 4 | 3.9 | 4.1167
| 1 | 5 | 3.8 | 58
| 1 | 6 | 4
| Mean |  | 4.1167
| Median |  | 4.05
| Std Dev. |  | 0.31885
| Rel.Std.Dev. |  | 7.745%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 606b (Indicative Value 0.071 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.071 | 1
| 1 | 2 | 0.069 | 17
| 1 | 3 | 0.071 | 16
| 1 | 4 | 0.071 | 0.0705
| 1 | 5 | 0.071 | 59
| 1 | 6 | 0.07
| Mean |  | 0.0705
| Median |  | 0.071
| Std Dev. |  | 0.00083666
| Rel.Std.Dev. |  | 1.187%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 606b (Indicative Value 156 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 156 | 18
| 1 | 3 | 154 | 16
| 1 | 4 | 154 | 155.8333
| 1 | 5 | 152 | 60
| 1 | 6 | 155
| Mean |  | 155.8333
| Median |  | 154.5
| Std Dev. |  | 4.2151
| Rel.Std.Dev. |  | 2.705%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 606b (Indicative Value 10 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 10.3 | 1
| 1 | 2 | 10.1 | 33
| 1 | 3 | 9.8 | 16
| 1 | 4 | 10.4 | 10.0167
| 1 | 5 | 9.6 | 61
| 1 | 6 | 9.9
| Mean |  | 10.0167
| Median |  | 10
| Std Dev. |  | 0.30605
| Rel.Std.Dev. |  | 3.055%
| PDM3 |  | -3.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 606b (Indicative Value 134 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 137 | 1
| 1 | 2 | 137 | 34
| 1 | 3 | 133 | 16
| 1 | 4 | 135 | 134
| 1 | 5 | 129 | 62
| 1 | 6 | 133
| Mean |  | 134
| Median |  | 134
| Std Dev. |  | 3.0332
| Rel.Std.Dev. |  | 2.264%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 606b (Indicative Value 0.468 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.479 | 1
| 1 | 2 | 0.463 | 11
| 1 | 3 | 0.469 | 16
| 1 | 4 | 0.469 | 0.46833
| 1 | 5 | 0.465 | 46
| 1 | 6 | 0.465
| Mean |  | 0.46833
| Median |  | 0.467
| Std Dev. |  | 0.0057504
| Rel.Std.Dev. |  | 1.228%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 606b (Indicative Value 4.05 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.1 | 1
| 1 | 2 | 4.2 | 37
| 1 | 3 | 4 | 16
| 1 | 4 | 4.1 | 4.05
| 1 | 5 | 3.9 | 47
| 1 | 6 | 4
| Mean |  | 4.05
| Median |  | 4.05
| Std Dev. |  | 0.10488
| Rel.Std.Dev. |  | 2.59%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 606b (Indicative Value 6.97 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 7.2 | 1
| 1 | 2 | 7.2 | 22
| 1 | 3 | 6.8 | 16
| 1 | 4 | 6.9 | 6.9667
| 1 | 5 | 6.7 | 48
| 1 | 6 | 7
| Mean |  | 6.9667
| Median |  | 6.95
| Std Dev. |  | 0.20656
| Rel.Std.Dev. |  | 2.965%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 606b (Indicative Value 3.72 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.27 | 1
| 1 | 2 | 3.72 | 13
| 1 | 3 | 3.59 | 16
| 1 | 4 | 3.63 | 3.7167
| 1 | 5 | 3.44 | 49
| 1 | 6 | 3.65
| Mean |  | 3.7167
| Median |  | 3.64
| Std Dev. |  | 0.28661
| Rel.Std.Dev. |  | 7.712%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 606b (Indicative Value 203 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 208 | 1
| 1 | 2 | 202 | 14
| 1 | 3 | 198 | 16
| 1 | 4 | 201 | 203
| 1 | 5 | 203 | 50
| 1 | 6 | 206
| Mean |  | 203
| Median |  | 202.5
| Std Dev. |  | 3.5777
| Rel.Std.Dev. |  | 1.762%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 606b (Indicative Value 1.31 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.33 | 1
| 1 | 2 | 1.33 | 25
| 1 | 3 | 1.31 | 16
| 1 | 4 | 1.32 | 1.3117
| 1 | 5 | 1.28 | 51
| 1 | 6 | 1.3
| Mean |  | 1.3117
| Median |  | 1.315
| Std Dev. |  | 0.019408
| Rel.Std.Dev. |  | 1.48%
| PDM3 |  | -2.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 606b (Indicative Value 0.68 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.69 | 1
| 1 | 2 | 0.69 | 26
| 1 | 3 | 0.67 | 16
| 1 | 4 | 0.67 | 0.67667
| 1 | 5 | 0.66 | 52
| 1 | 6 | 0.68
| Mean |  | 0.67667
| Median |  | 0.675
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 606b (Indicative Value 13.2 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13.4 | 1
| 1 | 2 | 13.2 | 28
| 1 | 3 | 13 | 16
| 1 | 4 | 13 | 13.1667
| 1 | 5 | 13.2 | 53
| 1 | 6 | 13.2
| Mean |  | 13.1667
| Median |  | 13.2
| Std Dev. |  | 0.15055
| Rel.Std.Dev. |  | 1.143%
| PDM3 |  | -2.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 606b (Indicative Value 0.188 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.191 | 1
| 1 | 2 | 0.187 | 15
| 1 | 3 | 0.188 | 16
| 1 | 4 | 0.186 | 0.18817
| 1 | 5 | 0.187 | 54
| 1 | 6 | 0.19
| Mean |  | 0.18817
| Median |  | 0.1875
| Std Dev. |  | 0.0019408
| Rel.Std.Dev. |  | 1.031%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 606b (Indicative Value 5.54 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.6 | 1
| 1 | 2 | 5.52 | 32
| 1 | 3 | 5.41 | 16
| 1 | 4 | 5.36 | 5.5367
| 1 | 5 | 5.88 | 55
| 1 | 6 | 5.45
| Mean |  | 5.5367
| Median |  | 5.485
| Std Dev. |  | 0.18811
| Rel.Std.Dev. |  | 3.398%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 606b (Indicative Value 4.98 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 4.7 | 1
| 1 | 2 | 5.2 | 33
| 1 | 3 | 5.1 | 16
| 1 | 4 | 4.9 | 4.9833
| 1 | 5 | 5 | 56
| 1 | 6 | 5
| Mean |  | 4.9833
| Median |  | 5
| Std Dev. |  | 0.17224
| Rel.Std.Dev. |  | 3.456%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 606b (Indicative Value 2.02 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 2 | 34
| 1 | 3 | 2 | 16
| 1 | 4 | 2.1 | 2.0167
| 1 | 5 | 2 | 57
| 1 | 6 | 2
| Mean |  | 2.0167
| Median |  | 2
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 2.024%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 606b (Indicative Value 13.9 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 14.2 | 1
| 1 | 2 | 14.2 | 35
| 1 | 3 | 13.7 | 16
| 1 | 4 | 14.1 | 13.9333
| 1 | 5 | 13.5 | 58
| 1 | 6 | 13.9
| Mean |  | 13.9333
| Median |  | 14
| Std Dev. |  | 0.28752
| Rel.Std.Dev. |  | 2.064%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 606b (Indicative Value 0.56 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.58 | 1
| 1 | 2 | 0.56 | 36
| 1 | 3 | 0.57 | 16
| 1 | 4 | 0.57 | 0.56167
| 1 | 5 | 0.55 | 59
| 1 | 6 | 0.54
| Mean |  | 0.56167
| Median |  | 0.565
| Std Dev. |  | 0.01472
| Rel.Std.Dev. |  | 2.621%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 606b (Indicative Value 324 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 360 | 1
| 1 | 2 | 324 | 17
| 1 | 3 | 310 | 16
| 1 | 4 | 322 | 323.8333
| 1 | 5 | 311 | 60
| 1 | 6 | 316
| Mean |  | 323.8333
| Median |  | 319
| Std Dev. |  | 18.5948
| Rel.Std.Dev. |  | 5.742%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 606b (Indicative Value 287 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 296 | 1
| 1 | 2 | 280 | 18
| 1 | 3 | 288 | 16
| 1 | 4 | 286 | 287.1667
| 1 | 5 | 286 | 61
| 1 | 6 | 287
| Mean |  | 287.1667
| Median |  | 286.5
| Std Dev. |  | 5.1543
| Rel.Std.Dev. |  | 1.795%
| PDM3 |  | -1.221e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.