## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 609c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 4.7939 | 4.6931 | 4.8946 | 4.7838 | 4.8039
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 4.9949 | 4.8911 | 5.0987 | 4.9834 | 5.0063
Infrared Combustion
S, Sulphur (wt.%) | 1.6032 | 1.5592 | 1.6472 | 1.5775 | 1.6289
4-Acid Digestion
Ag, Silver (ppm) | 24.0481 | 23.2595 | 24.8367 | 23.5262 | 24.5699
Al, Aluminium (wt.%) | 7.0856 | 6.8294 | 7.3417 | 6.9223 | 7.2488
As, Arsenic (ppm) | 883.7257 | 847.2486 | 920.2029 | 864.0806 | 903.3709
Be, Beryllium (ppm) | 2.6855 | 2.5566 | 2.8144 | 2.6013 | 2.7698
Bi, Bismuth (ppm) | 64.1513 | 61.1796 | 67.1229 | 61.9776 | 66.325
Ca, Calcium (wt.%) | 1.0923 | 1.0547 | 1.1298 | 1.0636 | 1.1209
Cd, Cadmium (ppm) | 5.7087 | 5.4507 | 5.9666 | 5.5498 | 5.8676
Ce, Cerium (ppm) | 78.952 | 73.9687 | 83.9354 | 77.0342 | 80.8699
Co, Cobalt (ppm) | 7.1274 | 6.7659 | 7.4889 | 6.8397 | 7.4152
Cr, Chromium (ppm) | 15.8587 | 14.1869 | 17.5305 | 14.8362 | 16.8811
Cs, Caesium (ppm) | 5.8608 | 5.5987 | 6.1229 | 5.6934 | 6.0283
Cu, Copper (wt.%) | 0.47796 | 0.46256 | 0.49336 | 0.47006 | 0.48586
Dy, Dysprosium (ppm) | 3.2707 | 3.0905 | 3.4508 | 3.1345 | 3.4069
Er, Erbium (ppm) | 0.98526 | 0.92193 | 1.0486 | 0.9469 | 1.0236
Eu, Europium (ppm) | 1.1643 | 1.0915 | 1.2371 | 1.1161 | 1.2126
Fe, Iron (wt.%) | 3.0042 | 2.8875 | 3.1209 | 2.9303 | 3.0781
Ga, Gallium (ppm) | 23.1779 | 22.1033 | 24.2526 | 22.3525 | 24.0034
Gd, Gadolinium (ppm) | 5.1772 | 4.8965 | 5.4579 | 4.971 | 5.3834
Ge, Germanium (ppm) | 0.24655 | 0.18124 | 0.31186 | 0.20652 | 0.28658
Hf, Hafnium (ppm) | 6.2657 | 5.9876 | 6.5438 | 6.0709 | 6.4605
Ho, Holmium (ppm) | 0.45736 | 0.41706 | 0.49765 | 0.44129 | 0.47342
In, Indium (ppm) | 1.3187 | 1.2551 | 1.3822 | 1.2661 | 1.3712
K, Potassium (wt.%) | 2.7604 | 2.6703 | 2.8505 | 2.6972 | 2.8236
La, Lanthanum (ppm) | 38.6767 | 36.0231 | 41.3302 | 37.4148 | 39.9385
Li, Lithium (ppm) | 21.6166 | 20.7877 | 22.4455 | 21.0382 | 22.195
Lu, Lutetium (ppm) | 0.079712 | 0.065384 | 0.094039 | IND | IND
Mg, Magnesium (wt.%) | 0.119 | 0.11197 | 0.12603 | 0.11558 | 0.12241
Mn, Manganese (wt.%) | 0.025097 | 0.024189 | 0.026006 | 0.024446 | 0.025749
Mo, Molybdenum (ppm) | 40.0522 | 38.4916 | 41.6128 | 39.0638 | 41.0406
Na, Sodium (wt.%) | 2.221 | 2.1472 | 2.2949 | 2.1753 | 2.2667
Nb, Niobium (ppm) | 16.4747 | 15.4976 | 17.4518 | 15.8643 | 17.0851
Nd, Neodymium (ppm) | 34.627 | 32.1673 | 37.0867 | 33.7613 | 35.4927
Ni, Nickel (ppm) | 15.6371 | 14.977 | 16.2973 | 15.1638 | 16.1105
P, Phosphorus (wt.%) | 0.029076 | 0.028053 | 0.030099 | 0.028195 | 0.029957
Pb, Lead (ppm) | 366.0398 | 352.1908 | 379.8887 | 357.7989 | 374.2807
Pr, Praseodymium (ppm) | 9.2438 | 8.6405 | 9.847 | 9.0123 | 9.4752
Rb, Rubidium (ppm) | 125.344 | 119.5631 | 131.125 | 122.5575 | 128.1306
Re, Rhenium (ppm) | 0.029018 | 0.025146 | 0.03289 | 0.024915 | 0.033121
S, Sulphur (wt.%) | 1.5773 | 1.5218 | 1.6328 | 1.5427 | 1.6119
Sb, Antimony (ppm) | 152.3775 | 145.6515 | 159.1035 | 147.6196 | 157.1353
Sc, Scandium (ppm) | 4.1816 | 3.8805 | 4.4828 | 4.0402 | 4.3231
Se, Selenium (ppm) | 10.1777 | 9.393 | 10.9624 | 9.6703 | 10.6852
Sm, Samarium (ppm) | 6.5245 | 6.0437 | 7.0053 | 6.3545 | 6.6944
Sn, Tin (ppm) | 9.5251 | 9.0063 | 10.044 | 9.2246 | 9.8256
Sr, Strontium (ppm) | 194.2418 | 186.5361 | 201.9474 | 190.3509 | 198.1326
Ta, Tantalum (ppm) | 1.2212 | 1.1345 | 1.3079 | 1.1599 | 1.2825
Tb, Terbium (ppm) | 0.69555 | 0.63235 | 0.75874 | 0.66649 | 0.72461
Te, Tellurium (ppm) | 11.8941 | 11.0579 | 12.7303 | 11.5368 | 12.2515
Th, Thorium (ppm) | 13.2049 | 12.431 | 13.9788 | 12.8763 | 13.5335
Ti, Titanium (wt.%) | 0.11604 | 0.11109 | 0.12098 | 0.11372 | 0.11835
Tl, Thallium (ppm) | 1.1392 | 1.0775 | 1.2009 | 1.1059 | 1.1726
Tm, Thulium (ppm) | 0.1048 | 0.095389 | 0.11422 | IND | IND
U, Uranium (ppm) | 4.8577 | 4.6594 | 5.056 | 4.7132 | 5.0022
V, Vanadium (ppm) | 8.8476 | 8.1865 | 9.5088 | 8.4939 | 9.2013
W, Tungsten (ppm) | 3.29 | 3.0551 | 3.525 | 3.1062 | 3.4738
Y, Yttrium (ppm) | 13.4217 | 12.7121 | 14.1314 | 13.0385 | 13.8049
Yb, Ytterbium (ppm) | 0.58107 | 0.51722 | 0.64491 | 0.5509 | 0.61123
Zn, Zinc (ppm) | 1078.2733 | 1040.4087 | 1116.1379 | 1057.8193 | 1098.7273
Zr, Zirconium (ppm) | 238.1314 | 227.5655 | 248.6973 | 232.1396 | 244.1232
Aqua Regia Digestion
Ag, Silver (ppm) | 23.9683 | 23.1304 | 24.8062 | 23.4928 | 24.4438
Al, Aluminium (wt.%) | 0.56095 | 0.53176 | 0.59014 | 0.54344 | 0.57845
As, Arsenic (ppm) | 851.5309 | 825.6903 | 877.3714 | 837.5881 | 865.4737
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.61727 | 0.56755 | 0.66698 | 0.59494 | 0.63959
Bi, Bismuth (ppm) | 64.81 | 61.9382 | 67.6818 | 63.1521 | 66.4678
Ca, Calcium (wt.%) | 0.84518 | 0.8185 | 0.87187 | 0.82727 | 0.86309
Cd, Cadmium (ppm) | 5.5719 | 5.3313 | 5.8124 | 5.4206 | 5.7232
Ce, Cerium (ppm) | 37.4229 | 34.3749 | 40.4709 | 36.2813 | 38.5645
Co, Cobalt (ppm) | 6.3638 | 6.0886 | 6.6391 | 6.1558 | 6.5719
Cr, Chromium (ppm) | 15.9239 | 14.7407 | 17.107 | 15.1024 | 16.7453
Cs, Caesium (ppm) | 0.97483 | 0.90536 | 1.0443 | 0.93834 | 1.0113
Cu, Copper (wt.%) | 0.47797 | 0.46805 | 0.48789 | 0.46981 | 0.48613
Dy, Dysprosium (ppm) | 1.3394 | 1.1285 | 1.5502 | 1.2839 | 1.3949
Er, Erbium (ppm) | 0.38202 | 0.32149 | 0.44255 | 0.36329 | 0.40075
Eu, Europium (ppm) | 0.50112 | 0.4084 | 0.59384 | 0.47015 | 0.53209
Fe, Iron (wt.%) | 2.293 | 2.1983 | 2.3876 | 2.2356 | 2.3503
Ga, Gallium (ppm) | 3.9649 | 3.6346 | 4.2952 | 3.8065 | 4.1233
Gd, Gadolinium (ppm) | 2.2985 | 1.9105 | 2.6866 | 2.1484 | 2.4487
Ge, Germanium (ppm) | 0.089074 | 0.061204 | 0.11694 | IND | IND
Hf, Hafnium (ppm) | 1.3597 | 1.2608 | 1.4587 | 1.3092 | 1.4103
Hg, Mercury (ppm) | 0.34372 | 0.31813 | 0.36931 | 0.31967 | 0.36777
Ho, Holmium (ppm) | 0.17728 | 0.15176 | 0.20279 | 0.16397 | 0.19059
In, Indium (ppm) | 1.2599 | 1.1914 | 1.3283 | 1.2238 | 1.2959
K, Potassium (wt.%) | 0.24301 | 0.23125 | 0.25476 | 0.23515 | 0.25087
La, Lanthanum (ppm) | 18.3481 | 16.8762 | 19.82 | 17.7325 | 18.9637
Li, Lithium (ppm) | 5.2089 | 4.7866 | 5.6313 | 5.0077 | 5.4101
Lu, Lutetium (ppm) | 0.026139 | 0.0196 | 0.032677 | IND | IND
Mg, Magnesium (wt.%) | 0.042072 | 0.040355 | 0.043788 | 0.040558 | 0.043586
Mn, Manganese (wt.%) | 0.022768 | 0.022085 | 0.023451 | 0.022273 | 0.023263
Mo, Molybdenum (ppm) | 39.5356 | 38.0051 | 41.066 | 38.6625 | 40.4086
Na, Sodium (wt.%) | 0.063871 | 0.057564 | 0.070178 | 0.062164 | 0.065578
Nb, Niobium (ppm) | 0.59447 | 0.47834 | 0.7106 | 0.55184 | 0.6371
Nd, Neodymium (ppm) | 14.4705 | 11.483 | 17.458 | 13.799 | 15.142
Ni, Nickel (ppm) | 15.5051 | 14.8026 | 16.2077 | 14.9953 | 16.0149
P, Phosphorus (wt.%) | 0.02089 | 0.01999 | 0.02179 | 0.02041 | 0.02137
Pb, Lead (ppm) | 329.9253 | 319.5646 | 340.286 | 323.2355 | 336.615
Pr, Praseodymium (ppm) | 4.0301 | 3.3014 | 4.7587 | 3.8416 | 4.2186
Rb, Rubidium (ppm) | 12.4037 | 11.5157 | 13.2918 | 11.9773 | 12.8302
Re, Rhenium (ppm) | 0.030382 | 0.027798 | 0.032965 | 0.027059 | 0.033705
S, Sulphur (wt.%) | 1.3195 | 1.2782 | 1.3607 | 1.2951 | 1.3439
Sb, Antimony (ppm) | 132.2242 | 125.4551 | 138.9933 | 128.9103 | 135.5381
Sc, Scandium (ppm) | 1.1784 | 1.0293 | 1.3274 | 1.0931 | 1.2636
Se, Selenium (ppm) | 9.9439 | 9.3425 | 10.5453 | 9.6142 | 10.2736
Sm, Samarium (ppm) | 2.7394 | 2.2128 | 3.266 | 2.5455 | 2.9333
Sn, Tin (ppm) | 7.0011 | 6.5868 | 7.4154 | 6.8332 | 7.1691
Sr, Strontium (ppm) | 25.8438 | 24.5237 | 27.164 | 25.1743 | 26.5134
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.31143 | 0.25225 | 0.3706 | 0.29166 | 0.3312
Te, Tellurium (ppm) | 12.2846 | 11.6295 | 12.9396 | 11.8561 | 12.713
Th, Thorium (ppm) | 6.7274 | 6.1802 | 7.2746 | 6.4544 | 7.0004
Ti, Titanium (wt.%) | 0.017716 | 0.015869 | 0.019562 | 0.017114 | 0.018317
Tl, Thallium (ppm) | 0.60816 | 0.57341 | 0.64291 | 0.5835 | 0.63282
U, Uranium (ppm) | 2.2201 | 2.0933 | 2.3469 | 2.1141 | 2.3261
V, Vanadium (ppm) | 3.9829 | 3.799 | 4.1668 | 3.7563 | 4.2096
W, Tungsten (ppm) | 1.574 | 1.4732 | 1.6748 | 1.5029 | 1.6452
Y, Yttrium (ppm) | 5.7168 | 5.3193 | 6.1144 | 5.5042 | 5.9295
Yb, Ytterbium (ppm) | 0.19936 | 0.16161 | 0.23712 | 0.17396 | 0.22477
Zn, Zinc (ppm) | 1022.056 | 991.4359 | 1052.6761 | 1007.3961 | 1036.7158
Zr, Zirconium (ppm) | 51.3432 | 46.4959 | 56.1905 | 49.3925 | 53.2939
Alkaline Leach
S-(Sulphide), Sulphur as S2- (wt.%) | 1.0336 | 0.90254 | 1.1647 | 1.0028 | 1.0644
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 609c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Infrared Combustion
C | wt.% | 0.20917 | C-(Inorganic) | wt.% | 0.20167 | C-(Organic) | wt.% | 0.04
4-Acid Digestion
B | ppm | 19 | Ba | ppm | 1442.9064 | Hg | ppm | 0.12
Aqua Regia Digestion
Ba | ppm | 99.2922 | Pt | ppb | 11.6667
Pd | ppb | 13.6667 | Tm | ppm | 0.039822
Alkaline Leach
S-(Sulphate) | wt.% | 0.54033
Acid Leach
S-(Sulphate) | wt.% | 0.30333
Borate Fusion XRF
Al2O3 | wt.% | 13.865 | Fe2O3 | wt.% | 4.3177 | S | wt.% | 1.6
As | ppm | 820 | K2O | wt.% | 3.34 | SiO2 | wt.% | 67.705
BaO | ppm | 3282.51 | MgO | wt.% | 0.24 | Sn | ppm | 40
CaO | wt.% | 1.515 | MnO | wt.% | 0.03 | Sr | ppm | 200
Cl | ppm | 30 | Na2O | wt.% | 3.025 | TiO2 | wt.% | 0.2065
Co | ppm | < 10 | Ni | ppm | 10 | V2O5 | ppm | 17.852
Cr2O3 | ppm | 25.5763 | P2O5 | wt.% | 0.063019 | Zn | ppm | 1060
Cu | wt.% | 0.4795 | Pb | ppm | 380 | Zr | ppm | 245
Thermogravimetry
LOI1000 | wt.% | 3.38
Laser Ablation ICP-MS
Ag | ppm | 26.2 | Hf | ppm | 7.025 | Sm | ppm | 6.6
As | ppm | 876.5 | Ho | ppm | 0.455 | Sn | ppm | 9.8
Ba | ppm | 2870 | In | ppm | 1.175 | Sr | ppm | 199
Be | ppm | 3 | La | ppm | 42.4 | Ta | ppm | 1.25
Bi | ppm | 67.75 | Lu | ppm | 0.08 | Tb | ppm | 0.7
Cd | ppm | 6.2 | Mn | wt.% | 0.0265 | Te | ppm | 13.5
Ce | ppm | 81.75 | Mo | ppm | 40.1 | Th | ppm | 13.8
Co | ppm | 7.75 | Nb | ppm | 16.85 | Ti | wt.% | 0.123
Cr | ppm | 19.5 | Nd | ppm | 35.9 | Tl | ppm | 1.2
Cs | ppm | 5.775 | Ni | ppm | 17 | Tm | ppm | 0.11
Cu | wt.% | 0.455 | Pb | ppm | 375 | U | ppm | 5.07
Dy | ppm | 3.195 | Pr | ppm | 9.85 | V | ppm | 9.65
Er | ppm | 1.01 | Rb | ppm | 121.5 | W | ppm | 3.25
Eu | ppm | 1.135 | Re | ppm | 0.045 | Y | ppm | 14
Ga | ppm | 21.8 | Sb | ppm | 163.5 | Yb | ppm | 0.55
Gd | ppm | 5.09 | Sc | ppm | 3.9 | Zn | ppm | 1125
Ge | ppm | 2.175 | Se | ppm | < 5 | Zr | ppm | 256
3-Acid Digestion (no HF)
Ag | ppm | 24.8333 | Gd | ppm | 5.15 | S | wt.% | 1.6333
Al2O3 | wt.% | 13.4 | Hf | ppm | 6.15 | Sc | ppm | 4.05
Ba | ppm | 2625 | Ho | ppm | 0.39833 | Sm | ppm | 6.3833
Be | ppm | 2.8333 | K2O | wt.% | 3.255 | Sn | ppm | 9.2583
Bi | ppm | 63.9833 | La | ppm | 42.1667 | Sr | ppm | 204.8333
CaO | wt.% | 1.5267 | Li | ppm | 21 | Ta | ppm | 1.2217
Cd | ppm | 5.5333 | MgO | wt.% | 0.201 | Tb | ppm | 0.61333
Ce | ppm | 77.5 | MnO | wt.% | 0.0315 | Th | ppm | 12.5833
Co | ppm | 7.0233 | Mo | ppm | 40.9833 | TiO2 | wt.% | 0.1925
Cr | ppm | 16.8333 | Na2O | wt.% | 2.9983 | U | ppm | 4.9067
Cs | ppm | 5.7833 | Nb | ppm | 16.3833 | V | ppm | 9.0833
Cu | wt.% | 0.47133 | Nd | ppm | 33.6167 | W | ppm | 3.65
Dy | ppm | 3.1833 | Ni | ppm | 15.9833 | Y | ppm | 13.0833
Er | ppm | 0.93333 | P2O5 | wt.% | 0.082 | Yb | ppm | 0.57667
Eu | ppm | 1.685 | Pb | ppm | 378.6667 | Zn | ppm | 1115
Fe2O3 | wt.% | 4.1133 | Pr | ppm | 9.2833 | Zr | ppm | 268.8333
Ga | ppm | 22.1 | Rb | ppm | 125
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 609c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 4.7939 | 0.28161 | 4.2306 | 5.3571 | 3.949 | 5.6387 | 5.874% | 11.75% | 17.62% | 4.5542 | 5.0335
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 4.9949 | 0.25956 | 4.4757 | 5.514 | 4.2162 | 5.7735 | 5.197% | 10.39% | 15.59% | 4.7451 | 5.2446
Infrared Combustion
S, wt.% | 1.6032 | 0.059234 | 1.4847 | 1.7217 | 1.4255 | 1.7809 | 3.695% | 7.389% | 11.08% | 1.523 | 1.6834
4-Acid Digestion
Ag, ppm | 24.0481 | 0.86612 | 22.3158 | 25.7803 | 21.4497 | 26.6464 | 3.602% | 7.203% | 10.8% | 22.8457 | 25.2505
Al, wt.% | 7.0856 | 0.28052 | 6.5245 | 7.6466 | 6.244 | 7.9271 | 3.959% | 7.918% | 11.88% | 6.7313 | 7.4398
As, ppm | 883.7257 | 45.0625 | 793.6008 | 973.8507 | 748.5383 | 1018.9131 | 5.099% | 10.2% | 15.3% | 839.5395 | 927.912
Be, ppm | 2.6855 | 0.16962 | 2.3463 | 3.0248 | 2.1767 | 3.1944 | 6.316% | 12.63% | 18.95% | 2.5512 | 2.8198
Bi, ppm | 64.1513 | 3.4349 | 57.2815 | 71.0211 | 53.8465 | 74.456 | 5.354% | 10.71% | 16.06% | 60.9437 | 67.3588
Ca, wt.% | 1.0923 | 0.035979 | 1.0203 | 1.1642 | 0.98433 | 1.2002 | 3.294% | 6.588% | 9.882% | 1.0377 | 1.1469
Cd, ppm | 5.7087 | 0.33325 | 5.0422 | 6.3752 | 4.7089 | 6.7084 | 5.838% | 11.68% | 17.51% | 5.4232 | 5.9941
Ce, ppm | 78.952 | 6.2909 | 66.3703 | 91.5338 | 60.0794 | 97.8247 | 7.968% | 15.94% | 23.9% | 75.0044 | 82.8997
Co, ppm | 7.1274 | 0.42604 | 6.2753 | 7.9795 | 5.8493 | 8.4055 | 5.977% | 11.95% | 17.93% | 6.771 | 7.4838
Cr, ppm | 15.8587 | 1.7015 | 12.4558 | 19.2616 | 10.7543 | 20.9631 | 10.73% | 21.46% | 32.19% | 15.0657 | 16.6516
Cs, ppm | 5.8608 | 0.31095 | 5.2389 | 6.4827 | 4.928 | 6.7937 | 5.305% | 10.61% | 15.92% | 5.5678 | 6.1539
Cu, wt.% | 0.47796 | 0.014732 | 0.4485 | 0.50742 | 0.43376 | 0.52216 | 3.082% | 6.165% | 9.247% | 0.45406 | 0.50186
Dy, ppm | 3.2707 | 0.10131 | 3.068 | 3.4733 | 2.9667 | 3.5746 | 3.098% | 6.195% | 9.293% | 3.1071 | 3.4342
Er, ppm | 0.98526 | 0.044502 | 0.89625 | 1.0743 | 0.85175 | 1.1188 | 4.517% | 9.034% | 13.55% | 0.93599 | 1.0345
Eu, ppm | 1.1643 | 0.037806 | 1.0887 | 1.2399 | 1.0509 | 1.2777 | 3.247% | 6.494% | 9.741% | 1.1061 | 1.2225
Fe, wt.% | 3.0042 | 0.095945 | 2.8123 | 3.1961 | 2.7164 | 3.292 | 3.194% | 6.387% | 9.581% | 2.854 | 3.1544
Ga, ppm | 23.1779 | 1.5154 | 20.1472 | 26.2087 | 18.6318 | 27.7241 | 6.538% | 13.08% | 19.61% | 22.019 | 24.3368
Gd, ppm | 5.1772 | 0.23417 | 4.7089 | 5.6455 | 4.4747 | 5.8797 | 4.523% | 9.046% | 13.57% | 4.9183 | 5.4361
Ge, ppm | 0.24655 | 0.098897 | 0.048759 | 0.44435 | 0 | 0.54324 | 40.11% | 80.22% | 120.3% | 0.23423 | 0.25888
Hf, ppm | 6.2657 | 0.24144 | 5.7828 | 6.7486 | 5.5414 | 6.99 | 3.853% | 7.707% | 11.56% | 5.9524 | 6.579
Ho, ppm | 0.45736 | 0.030916 | 0.39552 | 0.51919 | 0.36461 | 0.55011 | 6.76% | 13.52% | 20.28% | 0.43449 | 0.48022
In, ppm | 1.3187 | 0.07666 | 1.1654 | 1.472 | 1.0887 | 1.5487 | 5.813% | 11.63% | 17.44% | 1.2527 | 1.3846
K, wt.% | 2.7604 | 0.10258 | 2.5553 | 2.9656 | 2.4527 | 3.0681 | 3.716% | 7.432% | 11.15% | 2.6224 | 2.8984
La, ppm | 38.6767 | 3.7586 | 31.1594 | 46.1939 | 27.4007 | 49.9526 | 9.718% | 19.44% | 29.15% | 36.7428 | 40.6105
Li, ppm | 21.6166 | 0.80007 | 20.0165 | 23.2167 | 19.2164 | 24.0168 | 3.701% | 7.402% | 11.1% | 20.5358 | 22.6974
Lu, ppm | 0.079712 | 0.0097413 | 0.060229 | 0.099194 | 0.050488 | 0.10894 | 12.22% | 24.44% | 36.66% | 0.075726 | 0.083697
Mg, wt.% | 0.119 | 0.0063967 | 0.1062 | 0.13179 | 0.099805 | 0.13819 | 5.376% | 10.75% | 16.13% | 0.11305 | 0.12495
Mn, wt.% | 0.025097 | 0.0010317 | 0.023034 | 0.027161 | 0.022002 | 0.028193 | 4.111% | 8.222% | 12.33% | 0.023843 | 0.026352
Mo, ppm | 40.0522 | 1.748 | 36.5563 | 43.5482 | 34.8083 | 45.2961 | 4.364% | 8.728% | 13.09% | 38.0496 | 42.0548
Na, wt.% | 2.221 | 0.076947 | 2.0671 | 2.3749 | 1.9902 | 2.4518 | 3.465% | 6.929% | 10.39% | 2.11 | 2.3321
Nb, ppm | 16.4747 | 1.2382 | 13.9983 | 18.9511 | 12.7601 | 20.1893 | 7.516% | 15.03% | 22.55% | 15.651 | 17.2984
Nd, ppm | 34.627 | 2.5098 | 29.6075 | 39.6465 | 27.0977 | 42.1563 | 7.248% | 14.5% | 21.74% | 32.8956 | 36.3583
Ni, ppm | 15.6371 | 0.79308 | 14.051 | 17.2233 | 13.2579 | 18.0164 | 5.072% | 10.14% | 15.22% | 14.8553 | 16.419
P, wt.% | 0.029076 | 0.001035 | 0.027006 | 0.031146 | 0.025971 | 0.032181 | 3.56% | 7.119% | 10.68% | 0.027622 | 0.03053
Pb, ppm | 366.0398 | 23.1958 | 319.6481 | 412.4314 | 296.4523 | 435.6272 | 6.337% | 12.67% | 19.01% | 347.7378 | 384.3417
Pr, ppm | 9.2438 | 0.65312 | 7.9375 | 10.55 | 7.2844 | 11.2031 | 7.066% | 14.13% | 21.2% | 8.7816 | 9.706
Rb, ppm | 125.344 | 5.4185 | 114.5071 | 136.181 | 109.0886 | 141.5995 | 4.323% | 8.646% | 12.97% | 119.0768 | 131.6112
Re, ppm | 0.029018 | 0.0023394 | 0.024339 | 0.033697 | 0.022 | 0.036037 | 8.062% | 16.12% | 24.19% | 0.027567 | 0.030469
S, wt.% | 1.5773 | 0.06379 | 1.4498 | 1.7049 | 1.386 | 1.7687 | 4.044% | 8.088% | 12.13% | 1.4985 | 1.6562
Sb, ppm | 152.3775 | 11.7446 | 128.8882 | 175.8667 | 117.1436 | 187.6114 | 7.708% | 15.42% | 23.12% | 144.7586 | 159.9963
Sc, ppm | 4.1816 | 0.43324 | 3.3152 | 5.0481 | 2.8819 | 5.4814 | 10.36% | 20.72% | 31.08% | 3.9726 | 4.3907
Se, ppm | 10.1777 | 0.49395 | 9.1898 | 11.1657 | 8.6959 | 11.6596 | 4.853% | 9.707% | 14.56% | 9.6689 | 10.6866
Sm, ppm | 6.5245 | 0.44565 | 5.6332 | 7.4158 | 5.1875 | 7.8614 | 6.83% | 13.66% | 20.49% | 6.1982 | 6.8507
Sn, ppm | 9.5251 | 0.52205 | 8.481 | 10.5692 | 7.959 | 11.0913 | 5.481% | 10.96% | 16.44% | 9.0489 | 10.0014
Sr, ppm | 194.2418 | 13.4694 | 167.303 | 221.1805 | 153.8336 | 234.6499 | 6.934% | 13.87% | 20.8% | 184.5297 | 203.9538
Ta, ppm | 1.2212 | 0.11402 | 0.99317 | 1.4492 | 0.87915 | 1.5632 | 9.336% | 18.67% | 28.01% | 1.1601 | 1.2823
Tb, ppm | 0.69555 | 0.074081 | 0.54739 | 0.84371 | 0.47331 | 0.91779 | 10.65% | 21.3% | 31.95% | 0.66077 | 0.73033
Te, ppm | 11.8941 | 0.93472 | 10.0247 | 13.7636 | 9.0899 | 14.6983 | 7.859% | 15.72% | 23.58% | 11.2994 | 12.4888
Th, ppm | 13.2049 | 1.0667 | 11.0714 | 15.3384 | 10.0047 | 16.4051 | 8.078% | 16.16% | 24.24% | 12.5447 | 13.8652
Ti, wt.% | 0.11604 | 0.0048447 | 0.10635 | 0.12572 | 0.1015 | 0.13057 | 4.175% | 8.35% | 12.53% | 0.11023 | 0.12184
Tl, ppm | 1.1392 | 0.057138 | 1.0249 | 1.2535 | 0.9678 | 1.3106 | 5.016% | 10.03% | 15.05% | 1.0823 | 1.1962
Tm, ppm | 0.1048 | 0.0042293 | 0.096344 | 0.11326 | 0.092115 | 0.11749 | 4.035% | 8.071% | 12.11% | 0.099563 | 0.11004
U, ppm | 4.8577 | 0.17857 | 4.5006 | 5.2148 | 4.322 | 5.3934 | 3.676% | 7.352% | 11.03% | 4.6148 | 5.1006
V, ppm | 8.8476 | 0.63552 | 7.5766 | 10.1186 | 6.9411 | 10.7542 | 7.183% | 14.37% | 21.55% | 8.4052 | 9.29
W, ppm | 3.29 | 0.23482 | 2.8204 | 3.7596 | 2.5856 | 3.9945 | 7.137% | 14.27% | 21.41% | 3.1255 | 3.4545
Y, ppm | 13.4217 | 0.65248 | 12.1168 | 14.7267 | 11.4643 | 15.3792 | 4.861% | 9.723% | 14.58% | 12.7506 | 14.0928
Yb, ppm | 0.58107 | 0.053906 | 0.47326 | 0.68888 | 0.41935 | 0.74279 | 9.277% | 18.55% | 27.83% | 0.55202 | 0.61012
Zn, ppm | 1078.2733 | 49.371 | 979.5314 | 1177.0152 | 930.1604 | 1226.3862 | 4.579% | 9.157% | 13.74% | 1024.3596 | 1132.187
Zr, ppm | 238.1314 | 14.3583 | 209.4148 | 266.848 | 195.0566 | 281.2063 | 6.03% | 12.06% | 18.09% | 226.2248 | 250.038
Aqua Regia Digestion
Ag, ppm | 23.9683 | 1.0721 | 21.8241 | 26.1125 | 20.7519 | 27.1847 | 4.473% | 8.946% | 13.42% | 22.7699 | 25.1667
Al, wt.% | 0.56095 | 0.040816 | 0.47932 | 0.64258 | 0.4385 | 0.6834 | 7.276% | 14.55% | 21.83% | 0.5329 | 0.58899
As, ppm | 851.5309 | 47.7301 | 756.0706 | 946.9912 | 708.3405 | 994.7213 | 5.605% | 11.21% | 16.82% | 808.9543 | 894.1074
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.61727 | 0.057296 | 0.50267 | 0.73186 | 0.44538 | 0.78915 | 9.282% | 18.56% | 27.85% | 0.5864 | 0.64813
Bi, ppm | 64.81 | 4.8577 | 55.0945 | 74.5254 | 50.2368 | 79.3832 | 7.495% | 14.99% | 22.49% | 61.5695 | 68.0505
Ca, wt.% | 0.84518 | 0.030795 | 0.78359 | 0.90677 | 0.7528 | 0.93757 | 3.644% | 7.287% | 10.93% | 0.80292 | 0.88744
Cd, ppm | 5.5719 | 0.26323 | 5.0454 | 6.0984 | 4.7822 | 6.3616 | 4.724% | 9.449% | 14.17% | 5.2933 | 5.8505
Ce, ppm | 37.4229 | 4.2346 | 28.9537 | 45.8921 | 24.7191 | 50.1266 | 11.32% | 22.63% | 33.95% | 35.5517 | 39.294
Co, ppm | 6.3638 | 0.36519 | 5.6335 | 7.0942 | 5.2683 | 7.4594 | 5.739% | 11.48% | 17.22% | 6.0456 | 6.682
Cr, ppm | 15.9239 | 0.86162 | 14.2006 | 17.6471 | 13.339 | 18.5087 | 5.411% | 10.82% | 16.23% | 15.1277 | 16.7201
Cs, ppm | 0.97483 | 0.097486 | 0.77985 | 1.1698 | 0.68237 | 1.2673 | 10% | 20% | 30% | 0.92608 | 1.0236
Cu, wt.% | 0.47797 | 0.012803 | 0.45236 | 0.50358 | 0.43956 | 0.51638 | 2.679% | 5.357% | 8.036% | 0.45407 | 0.50187
Dy, ppm | 1.3394 | 0.18407 | 0.97123 | 1.7075 | 0.78716 | 1.8916 | 13.74% | 27.49% | 41.23% | 1.2724 | 1.4063
Er, ppm | 0.38202 | 0.053034 | 0.27595 | 0.48809 | 0.22292 | 0.54112 | 13.88% | 27.76% | 41.65% | 0.36292 | 0.40112
Eu, ppm | 0.50112 | 0.089865 | 0.32139 | 0.68085 | 0.23153 | 0.77072 | 17.93% | 35.87% | 53.8% | 0.47606 | 0.52618
Fe, wt.% | 2.293 | 0.14576 | 2.0015 | 2.5845 | 1.8557 | 2.7303 | 6.357% | 12.71% | 19.07% | 2.1783 | 2.4076
Ga, ppm | 3.9649 | 0.54628 | 2.8723 | 5.0574 | 2.3261 | 5.6037 | 13.78% | 27.56% | 41.33% | 3.7666 | 4.1631
Gd, ppm | 2.2985 | 0.39818 | 1.5022 | 3.0949 | 1.104 | 3.4931 | 17.32% | 34.65% | 51.97% | 2.1836 | 2.4135
Ge, ppm | 0.089074 | 0.026128 | 0.036819 | 0.14133 | 0.010691 | 0.16746 | 29.33% | 58.66% | 88% | 0.08462 | 0.093528
Hf, ppm | 1.3597 | 0.13636 | 1.087 | 1.6325 | 0.95066 | 1.7688 | 10.03% | 20.06% | 30.09% | 1.2918 | 1.4277
Hg, ppm | 0.34372 | 0.020642 | 0.30243 | 0.385 | 0.28179 | 0.40565 | 6.006% | 12.01% | 18.02% | 0.32653 | 0.36091
Ho, ppm | 0.17728 | 0.024297 | 0.12868 | 0.22587 | 0.10439 | 0.25017 | 13.71% | 27.41% | 41.12% | 0.16841 | 0.18614
In, ppm | 1.2599 | 0.07369 | 1.1125 | 1.4073 | 1.0388 | 1.4809 | 5.849% | 11.7% | 17.55% | 1.1969 | 1.3229
K, wt.% | 0.24301 | 0.018953 | 0.2051 | 0.28091 | 0.18615 | 0.29987 | 7.799% | 15.6% | 23.4% | 0.23086 | 0.25516
La, ppm | 18.3481 | 2.2304 | 13.8873 | 22.809 | 11.6568 | 25.0394 | 12.16% | 24.31% | 36.47% | 17.4307 | 19.2655
Li, ppm | 5.2089 | 0.61766 | 3.9736 | 6.4442 | 3.356 | 7.0619 | 11.86% | 23.72% | 35.57% | 4.9485 | 5.4694
Lu, ppm | 0.026139 | 0.0043254 | 0.017488 | 0.03479 | 0.013163 | 0.039115 | 16.55% | 33.1% | 49.64% | 0.024832 | 0.027446
Mg, wt.% | 0.042072 | 0.0035364 | 0.034999 | 0.049145 | 0.031463 | 0.052681 | 8.406% | 16.81% | 25.22% | 0.039968 | 0.044175
Mn, wt.% | 0.022768 | 0.00097543 | 0.020817 | 0.024719 | 0.019842 | 0.025694 | 4.284% | 8.568% | 12.85% | 0.02163 | 0.023906
Mo, ppm | 39.5356 | 1.9542 | 35.6272 | 43.444 | 33.673 | 45.3981 | 4.943% | 9.886% | 14.83% | 37.5588 | 41.5124
Na, wt.% | 0.063871 | 0.0058679 | 0.052135 | 0.075607 | 0.046268 | 0.081475 | 9.187% | 18.37% | 27.56% | 0.060678 | 0.067065
Nb, ppm | 0.59447 | 0.17556 | 0.24336 | 0.94558 | 0.067801 | 1.1211 | 29.53% | 59.06% | 88.59% | 0.56475 | 0.62419
Nd, ppm | 14.4705 | 3.5183 | 7.4338 | 21.5071 | 3.9155 | 25.0255 | 24.31% | 48.63% | 72.94% | 13.747 | 15.194
Ni, ppm | 15.5051 | 0.91576 | 13.6736 | 17.3367 | 12.7579 | 18.2524 | 5.906% | 11.81% | 17.72% | 14.7299 | 16.2804
P, wt.% | 0.02089 | 0.00092323 | 0.019044 | 0.022736 | 0.01812 | 0.02366 | 4.419% | 8.839% | 13.26% | 0.019846 | 0.021935
Pb, ppm | 329.9253 | 13.3447 | 303.2359 | 356.6147 | 289.8912 | 369.9594 | 4.045% | 8.09% | 12.13% | 313.429 | 346.4216
Pr, ppm | 4.0301 | 0.72453 | 2.581 | 5.4791 | 1.8565 | 6.2037 | 17.98% | 35.96% | 53.93% | 3.8286 | 4.2316
Rb, ppm | 12.4037 | 1.4184 | 9.567 | 15.2404 | 8.1487 | 16.6588 | 11.43% | 22.87% | 34.3% | 11.7836 | 13.0239
Re, ppm | 0.030382 | 0.0019786 | 0.026425 | 0.034339 | 0.024446 | 0.036318 | 6.512% | 13.02% | 19.54% | 0.028863 | 0.031901
S, wt.% | 1.3195 | 0.066277 | 1.1869 | 1.452 | 1.1206 | 1.5183 | 5.023% | 10.05% | 15.07% | 1.2535 | 1.3855
Sb, ppm | 132.2242 | 10.7904 | 110.6433 | 153.8051 | 99.8529 | 164.5955 | 8.161% | 16.32% | 24.48% | 125.613 | 138.8354
Sc, ppm | 1.1784 | 0.19232 | 0.79371 | 1.563 | 0.60139 | 1.7553 | 16.32% | 32.64% | 48.96% | 1.1194 | 1.2373
Se, ppm | 9.9439 | 0.46362 | 9.0167 | 10.8712 | 8.5531 | 11.3348 | 4.662% | 9.325% | 13.99% | 9.4467 | 10.4411
Sm, ppm | 2.7394 | 0.56142 | 1.6165 | 3.8622 | 1.0551 | 4.4236 | 20.49% | 40.99% | 61.48% | 2.6024 | 2.8763
Sn, ppm | 7.0011 | 0.45949 | 6.0822 | 7.9201 | 5.6227 | 8.3796 | 6.563% | 13.13% | 19.69% | 6.6511 | 7.3512
Sr, ppm | 25.8438 | 1.5561 | 22.7316 | 28.9561 | 21.1755 | 30.5122 | 6.021% | 12.04% | 18.06% | 24.5517 | 27.136
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.31143 | 0.063002 | 0.18542 | 0.43743 | 0.12242 | 0.50044 | 20.23% | 40.46% | 60.69% | 0.29586 | 0.327
Te, ppm | 12.2846 | 0.75581 | 10.7729 | 13.7962 | 10.0171 | 14.552 | 6.153% | 12.31% | 18.46% | 11.6703 | 12.8988
Th, ppm | 6.7274 | 1.0146 | 4.6981 | 8.7566 | 3.6835 | 9.7712 | 15.08% | 30.16% | 45.25% | 6.391 | 7.0637
Ti, wt.% | 0.017716 | 0.0029355 | 0.011845 | 0.023586 | 0.0089092 | 0.026522 | 16.57% | 33.14% | 49.71% | 0.01683 | 0.018601
Tl, ppm | 0.60816 | 0.039261 | 0.52964 | 0.68668 | 0.49037 | 0.72594 | 6.456% | 12.91% | 19.37% | 0.57775 | 0.63856
U, ppm | 2.2201 | 0.12776 | 1.9646 | 2.4756 | 1.8368 | 2.6034 | 5.755% | 11.51% | 17.26% | 2.1091 | 2.3311
V, ppm | 3.9829 | 0.14069 | 3.7015 | 4.2643 | 3.5609 | 4.405 | 3.532% | 7.065% | 10.6% | 3.7838 | 4.1821
W, ppm | 1.574 | 0.10863 | 1.3567 | 1.7913 | 1.2481 | 1.8999 | 6.902% | 13.8% | 20.7% | 1.4953 | 1.6527
Y, ppm | 5.7168 | 0.66923 | 4.3784 | 7.0553 | 3.7092 | 7.7245 | 11.71% | 23.41% | 35.12% | 5.431 | 6.0027
Yb, ppm | 0.19936 | 0.030062 | 0.13924 | 0.25949 | 0.10918 | 0.28955 | 15.08% | 30.16% | 45.24% | 0.18939 | 0.20933
Zn, ppm | 1022.056 | 35.5632 | 950.9295 | 1093.1824 | 915.3662 | 1128.7457 | 3.48% | 6.959% | 10.44% | 970.9532 | 1073.1588
Zr, ppm | 51.3432 | 8.8696 | 33.6039 | 69.0824 | 24.7343 | 77.952 | 17.28% | 34.55% | 51.83% | 48.776 | 53.9103
Alkaline Leach
S-(Sulphide), wt.% | 1.0336 | 0.23207 | 0.56948 | 1.4978 | 0.33742 | 1.7298 | 22.45% | 44.9% | 67.36% | 0.98194 | 1.0853
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 609c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
HCl*IRC | hydrochloric acid digestion with infrared combustion
HCl*OES | hydrochloric acid digest with inductively coupled plasma optical emission spectroscopy
I-550*IRC | ignition @ 550 degrees Celsius with infrared combustion
INAA | instrumental neutron activation analysis
IRC | infrared combustion
IRC-x*Calc | calculation by difference from IRC total value
LOI*TGA | loss on ignition by thermogravimetric analysis
Na2CO3*GRAV | sodium carbonate solution with Gravimetric finish
Na2CO3*IRC | sodium carbonate leach with infrared combustion
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 609c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek Genalysis, Perth, WA, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Intertek Testing Services, Townsville, QLD, Australia
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Mineral Services, Townsville, QLD, Australia
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 609c (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
1 | 1 | ANSLu | 1 | 9 | 9 | 261740 | 0.0836 | 5.2147 | 5.2298
1 | 1 | ANSLu | 1 | 11 | 11 | 261741 | 0.085144 | 5.2104 | 5.2296
1 | 1 | ANSLu | 1 | 17 | 6 | 261742 | 0.087558 | 5.3234 | 5.2356
1 | 1 | ANSLu | 1 | 14 | 2 | 261743 | 0.082141 | 5.2306 | 5.2307
1 | 1 | ANSLu | 1 | 1 | 1 | 261744 | 0.086906 | 5.3293 | 5.2359
1 | 1 | ANSLu | 1 | 3 | 3 | 261745 | 0.08348 | 5.18 | 5.228
1 | 1 | ANSLu | 1 | 13 | 2 | 261746 | 0.085206 | 5.3148 | 5.2352
1 | 1 | ANSLu | 1 | 10 | 9 | 261747 | 0.086096 | 5.1436 | 5.2261
1 | 1 | ANSLu | 1 | 15 | 4 | 261748 | 0.085803 | 5.2094 | 5.2296
1 | 1 | ANSLu | 1 | 7 | 7 | 261749 | 0.082744 | 5.1899 | 5.2285
1 | 1 | ANSLu | 1 | 18 | 6 | 261750 | 0.084235 | 5.2749 | 5.233
1 | 1 | ANSLu | 1 | 4 | 3 | 261751 | 0.087285 | 5.1195 | 5.2248
1 | 1 | ANSLu | 1 | 12 | 11 | 261752 | 0.084475 | 5.16 | 5.2269
1 | 1 | ANSLu | 1 | 20 | 8 | 261753 | 0.08262 | 5.2308 | 5.2307
1 | 1 | ANSLu | 1 | 5 | 5 | 261754 | 0.08559 | 5.1529 | 5.2265
1 | 1 | ANSLu | 1 | 8 | 7 | 261755 | 0.082944 | 5.2726 | 5.2329
1 | 1 | ANSLu | 1 | 16 | 4 | 261756 | 0.087263 | 5.2882 | 5.2338
1 | 1 | ANSLu | 1 | 19 | 8 | 261757 | 0.087412 | 5.3324 | 5.2361
1 | 1 | ANSLu | 1 | 2 | 1 | 261758 | 0.084768 | 5.2338 | 5.2309
1 | 1 | ANSLu | 1 | 6 | 5 | 261759 | 0.083878 | 5.2026 | 5.2292
Mean | [6 empty] | 0.084957 | 5.2307 | 5.2307
Median | [7 empty] | 5.2226 | 5.2303
SD | [7 empty] | 0.064847 | 0.0034509
RSD | [7 empty] | 1.24% | 0.06597%
RSD Upscale Factor | [7 empty] | 0.053216
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 609c (Certified Value 4.79 ppm) | [30 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | 2
|  |  | 0.085g | 40g | 50g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 50g | 50g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 30g | 30g | 15g | 50g | 3
| 1 | 1 | 5.3293 | 5.07 | 4.9725 | 4.711 | 5.085 | 5.095 | 4.7645 | 4.94 | 4.56 | 5.246 | 4.924 | 5.37 | 4.811 | 4.99 | 4.67 | 4.93 | 4.8 | 4.5 | 5.009 | 4.21 | 4.619 | 4.229 | 4.96 | 5.01 | 4.67 | 4.27 | 4.8426 | 4.68 | 3.22 | 1
| 1 | 2 | 5.2338 | 5.03 | 4.77 | 4.451 | 5.062 | 4.858 | 4.8187 | 4.66 | 4.31 | 5.2268 | 5.057 | 5.22 | 4.58 | 4.53 | 4.65 | 4.99 | 4.7 | 4.54 | 3.299 | 4.36 | 4.621 | 4.198 | 4.92 | 5.15 | 4.6 | 3.58 | 4.8218 | 4.816 | 4.81 | 0
| 1 | 3 | 5.18 | 5.04 | 5.08 | 4.477 | 5.144 | 5.075 | 4.8053 | 4.83 | 4.26 | 5.2369 | 4.988 | 5.15 | 4.521 | 4.88 | 5.01 | 4.97 | 5.53 | 4.35 | 4.999 | 4.36 | 4.702 | 4.276 | 4.9 | 5.09 | 4.61 | 4.58 | 4.9013 | 4.59 | 4.58 | 16
| 1 | 4 | 5.1195 | 5.03 | 4.8125 | 4.483 | 5.007 | 5.062 | 4.854 | 4.76 | 4.47 | 5.2369 | 4.926 | 5.16 | 4.548 | 4.38 | 4.92 | 5.07 | 4.76 | 4.41 | 4.889 | 4.4 | 4.876 | 4.232 | 4.94 | 4.96 | 4.68 | 4.84 | 5.006 | 4.634 | 4.61 | 4.7939
| 1 | 5 | 5.1529 | 5.04 | 5.06 | 4.52 | 5.239 | 4.993 | 4.9885 | 4.77 | 4.65 | 5.3288 | 4.971 | 5.19 | 5.003 | 4.92 | 4.91 | 5 | 5 | 4.48 | 4.889 | 4.2 | 4.939 | 4.25 | 5 | 4.74 | 4.61 | 4.06 | 4.9597 | 4.466 | 4.59 | 7
| 1 | 6 | 5.2026 | 5.01 | 5.0663 | 4.385 | 5.19 | 5.014 | 4.8974 | 4.78 | 4.96 | 5.2257 | 5.017 | 5.19 | 4.989 | 4.5 | 4.88 | 4.98 | 4.8 | 4.41 | 5.039 | 4.39 | > 5 | 4.233 | 4.94 | 5.1 | 4.56 | 4.54 | 4.9452 | 4.508 | 4.57
|  | 7 | 00: 5.1899
|  | 8 | 00: 5.2726
|  | 9 | 00: 5.2147
|  | 10 | 00: 5.1436
|  | 11 | 00: 5.2104
|  | 12 | 00: 5.16
|  | 13 | 00: 5.3148
|  | 14 | 00: 5.2306
|  | 15 | 00: 5.2094
|  | 16 | 00: 5.2882
|  | 17 | 00: 5.3234
|  | 18 | 00: 5.2749
|  | 19 | 00: 5.3324
|  | 20 | 00: 5.2308
| Mean |  | 5.2307 | 5.0367 | 4.9602 | 4.5045 | 5.1212 | 5.0162 | 4.8547 | 4.79 | 4.535 | 5.2502 | 4.9805 | 5.2133 | 4.742 | 4.7 | 4.84 | 4.99 | 4.9317 | 4.4483 | 4.6873 | 4.32 | 4.7514 | 4.2363 | 4.9433 | 5.0083 | 4.6217 | 4.3117 | 4.9128 | 4.6157 | 4.3967
| Median |  | 5.2226 | 5.035 | 5.0163 | 4.48 | 5.1145 | 5.038 | 4.8363 | 4.775 | 4.515 | 5.2369 | 4.9795 | 5.19 | 4.6955 | 4.705 | 4.895 | 4.985 | 4.8 | 4.445 | 4.944 | 4.36 | 4.702 | 4.2325 | 4.94 | 5.05 | 4.61 | 4.405 | 4.9233 | 4.612 | 4.585
| Std Dev. |  | 0.064847 | 0.019664 | 0.1369 | 0.11067 | 0.086008 | 0.086412 | 0.07949 | 0.092087 | 0.25493 | 0.039237 | 0.051949 | 0.080664 | 0.22209 | 0.25931 | 0.14615 | 0.046043 | 0.30999 | 0.070261 | 0.68309 | 0.090554 | 0.14807 | 0.025742 | 0.034448 | 0.14798 | 0.04535 | 0.44804 | 0.0711 | 0.1259 | 0.58343
| Rel.Std.Dev. |  | 1.24% | 0.3904% | 2.76% | 2.457% | 1.679% | 1.723% | 1.637% | 1.922% | 5.621% | 0.7473% | 1.043% | 1.547% | 4.683% | 5.517% | 3.02% | 0.9227% | 6.286% | 1.579% | 14.57% | 2.096% | 3.116% | 0.6077% | 0.6969% | 2.955% | 0.9813% | 10.39% | 1.447% | 2.728% | 13.27%
| PDM3 |  | 9.112% | 5.065% | 3.47% | -6.036% | 6.828% | 4.637% | 1.27% | -0.08041% | -5.4% | 9.519% | 3.893% | 8.75% | -1.082% | -1.958% | 0.9626% | 4.092% | 2.875% | -7.208% | -2.222% | -9.885% | -0.8856% | -11.63% | 3.118% | 4.474% | -3.592% | -10.06% | 2.481% | -3.717% | -8.285%
| Z-Score (Absolute) |  | NA | 0.77 | 0.51 | 1.08 | 1.07 | 0.7 | 0.14 | 0.09 | 0.98 | 1.52 | 0.58 | 1.39 | 0.25 | 0.4 | 0.09 | 0.61 | 0.41 | 1.28 | 0.45 | 1.73 | 0.22 | 2.02 | 0.45 | 0.67 | 0.67 | 1.76 | 0.34 | 0.7 | 1.46
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 609c (Certified Value 4.99 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.23 | 2.24 | 2.26 | 2.27 | 2.28 | ppm
|  |  | INAA | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*OES | 2
|  |  | 0.085g | 30g | 50g | 15g | 15g | 25g | 30g | 25g | 30g | 30g | 10g | 25g | 25g | 25g | 25g | 25g | 10g | 25g | 10g | 50g | 50g | 25g | 50g | 25g | 3
| 1 | 1 | 5.3293 | 4.96 | 4.907 | 5.0925 | 4.7849 | 4.93 | > 2 | 5.21 | 5.262 | 4.94 | > 2 | 4.66 | 5.15 | 5.16 | 4.95 | 4.44 | 4.81 | 5.08 | 4.09 | 4.88 | 5.35 | 4.8211 | 4.857 | 4.89 | 1
| 1 | 2 | 5.2338 | 4.98 | 4.8651 | 4.9567 | 5.0825 | 5.11 | > 2 |  | 5.385 | 4.99 | > 2 | 4.74 | 5.23 | 5.23 | 5.1 | 4.5 | 4.78 | 5.02 | 4.09 | 4.72 | 5.4 | 4.8232 | 4.907 | 5.05 | 0
| 1 | 3 | 5.18 | 4.96 | 4.7232 | 5.0052 | 5.0331 | 4.96 | > 2 |  | 5.254 | 4.96 | > 2 | 4.23 | 5.08 | 5.3 | 5.07 | 4.37 | 4.53 | 5.36 | 4.02 | 4.8 | 5.62 | 4.844 | 4.757 | 5.2 | 16
| 1 | 4 | 5.1195 | 4.98 | 4.8308 | 4.9955 | 4.9854 |  | > 2 |  | 5.274 | 4.91 | > 2 | 5.73 | 5.2 | 5.23 | 4.94 | 4.55 | 4.97 | 4.93 | 4.07 | 4.8 | 5.38 | 4.9107 | 4.791 | 5.15 | 4.9949
| 1 | 5 | 5.1529 | 4.94 | 4.9906 | 5.0149 | 4.9835 | 4.9 | > 2 |  | 5.306 | 4.94 |  | 4.38 | 5.31 | 5.15 | 4.92 | 4.42 | 4.63 | 5.22 | 4.09 | 4.8 | 5.4 | 4.8406 | 4.325 | 5.49 | 9
| 1 | 6 | 5.2026 | 4.97 | 5.4462 | 5.1119 | 5.092 | 5.05 | > 2 |  | 5.252 | 4.91 | > 2 | 4.91 | 5.05 | 5.26 | 5.11 | 4.41 | 4.81 | 5.45 | 4.08 | 4.78 | 5.36 | 4.9113 | 4.957 | 5.22
|  | 7 | 00: 5.1899
|  | 8 | 00: 5.2726
|  | 9 | 00: 5.2147
|  | 10 | 00: 5.1436
|  | 11 | 00: 5.2104
|  | 12 | 00: 5.16
|  | 13 | 00: 5.3148
|  | 14 | 00: 5.2306
|  | 15 | 00: 5.2094
|  | 16 | 00: 5.2882
|  | 17 | 00: 5.3234
|  | 18 | 00: 5.2749
|  | 19 | 00: 5.3324
|  | 20 | 00: 5.2308
| Mean |  | 5.2307 | 4.965 | 4.9605 | 5.0294 | 4.9936 | 4.99 |  | 5.21 | 5.2888 | 4.9417 |  | 4.775 | 5.17 | 5.2217 | 5.015 | 4.4483 | 4.755 | 5.1767 | 4.0733 | 4.7967 | 5.4183 | 4.8585 | 4.7657 | 5.1667
| Median |  | 5.2226 | 4.965 | 4.8861 | 5.01 | 5.0092 | 4.96 |  | 5.21 | 5.268 | 4.94 |  | 4.7 | 5.175 | 5.23 | 5.01 | 4.43 | 4.795 | 5.15 | 4.085 | 4.8 | 5.39 | 4.8423 | 4.824 | 5.175
| Std Dev. |  | 0.064847 | 0.015166 | 0.25373 | 0.06003 | 0.11213 | 0.087464 |  |  | 0.051094 | 0.030605 |  | 0.52872 | 0.096954 | 0.057764 | 0.08735 | 0.065549 | 0.15437 | 0.20245 | 0.027325 | 0.051251 | 0.10088 | 0.041678 | 0.22796 | 0.19947
| Rel.Std.Dev. |  | 1.24% | 0.3055% | 5.115% | 1.194% | 2.246% | 1.753% |  |  | 0.9661% | 0.6193% |  | 11.07% | 1.875% | 1.106% | 1.742% | 1.474% | 3.246% | 3.911% | 0.6708% | 1.068% | 1.862% | 0.8578% | 4.783% | 3.861%
| PDM3 |  | 4.722% | -0.5978% | -0.688% | 0.6926% | -0.02584% | -0.09725% |  | 4.307% | 5.886% | -1.065% |  | -4.402% | 3.506% | 4.541% | 0.4033% | -10.94% | -4.802% | 3.64% | -18.45% | -3.968% | 8.478% | -2.73% | -4.589% | 3.44%
| Z-Score (Absolute) |  | NA | 0.09 | 0.11 | 0.14 | 0.01 | 0 | NA | 0.79 | 1.08 | 0.17 | NA | 0.78 | 0.65 | 0.84 | 0.09 | 1.96 | 0.85 | 0.67 | 3.31 | 0.7 | 1.55 | 0.48 | 0.81 | 0.64
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 609c (Indicative Value 0.209 wt.%) | [26 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 2.17 | 2.18 | 2.27 | 3.01 | [21 empty] | wt.%
|  |  | IRC | IRC | IRC | IRC | [21 empty] | 3
3
| 1 | 1 | 0.21 | 0.24 | 0.13 | 0.24 | [21 empty] | 1
| 1 | 2 | 0.21 | 0.24 | 0.14 | 0.25 | [21 empty] | 1
| 1 | 3 | 0.21 | 0.25 | 0.15 | [22 empty] | 16
| 1 | 4 | 0.21 | 0.24 | 0.15 | [22 empty] | 0.20917
| 1 | 5 | 0.21 | 0.24 | 0.14 | [22 empty] | 7
| 1 | 6 | 0.21 | 0.25 | 0.12
| Mean |  | 0.21 | 0.24333 | 0.13833 | 0.245
| Median |  | 0.21 | 0.24 | 0.14 | 0.245
| Std Dev. |  | 0 | 0.005164 | 0.01169 | 0.0070711
| Rel.Std.Dev. |  | 0% | 2.122% | 8.451% | 2.886%
| PDM3 |  | 0.3984% | 16.33% | -33.86% | 17.13%
| Z-Score (Absolute) |  | 0.64 | 0.64 | 3.4 | 0.71
| Analytical results for C-(Inorganic) in OREAS 609c (Indicative Value 0.202 wt.%) | [26 empty] | Indicative
C-(Inorganic) | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.18 | [24 empty] | wt.%
|  |  | I-550*IRC | [24 empty] | 3
3
| 1 | 1 | 0.2 | [24 empty] | 1
| 1 | 2 | 0.2 | [24 empty] | 1
| 1 | 3 | 0.21 | [24 empty] | 16
| 1 | 4 | 0.2 | [24 empty] | 0.20167
| 1 | 5 | 0.2 | [24 empty] | 7
| 1 | 6 | 0.2
| Mean |  | 0.20167
| Median |  | 0.2
| Std Dev. |  | 0.0040825
| Rel.Std.Dev. |  | 2.024%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for C-(Organic) in OREAS 609c (Indicative Value 0.04 wt.%) | [26 empty] | Indicative
C-(Organic) | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.18 | [24 empty] | wt.%
|  |  | HCl*IRC | [24 empty] | 3
3
| 1 | 1 | 0.04 | [24 empty] | 1
| 1 | 2 | 0.04 | [24 empty] | 1
| 1 | 3 | 0.03 | [24 empty] | 16
| 1 | 4 | 0.05 | [24 empty] | 0.04
| 1 | 5 | 0.04 | [24 empty] | 7
| 1 | 6 | 0.04
| Mean |  | 0.04
| Median |  | 0.04
| Std Dev. |  | 0.0063246
| Rel.Std.Dev. |  | 15.81%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 609c (Certified Value 1.6 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 1.62 | 1.84 | 1.69 | 1.59 | 1.59 | 1.5684 | 1.59 | 1.7 | 1.55 | 1.79 | 1.5 | 1.64 | 1.62 | 1.58 | 1.62 | 1.58 | 1.64 | 1.548 | 1.63 | 1.56 | 1.7 | 1.64 | 1.5438 | 1.61 | 1.63 | 1
| 1 | 2 | 1.62 | 1.83 | 1.7 | 1.57 | 1.58 | 1.5681 | 1.57 | 1.7 | 1.57 | 1.79 | 1.58 | 1.58 | 1.62 | 1.46 | 1.66 | 1.48 | 1.66 | 1.523 | 1.65 | 1.57 | 1.7 | 3.7 | 1.5548 | 1.62 | 1.55 | 12
| 1 | 3 | 1.62 | 1.77 | 1.68 | 1.63 | 1.56 | 1.5809 | 1.63 | 1.69 | 1.55 | 1.78 | 1.59 | 1.58 | 1.63 | 1.47 | 1.63 | 1.64 | 1.63 | 1.543 | 1.65 | 1.58 | 1.65 | 1.61 | 1.5558 | 1.62 | 1.48 | 16
| 1 | 4 | 1.63 | 1.7 | 1.68 | 1.6 | 1.56 | 1.5659 | 1.6 | 1.7 | 1.54 | 1.78 | 1.58 | 1.57 | 1.57 | 1.55 | 1.64 | 1.55 | 1.66 | 1.727 | 1.65 | 1.55 | 1.68 | 1.6 | 1.5522 | 1.61 | 1.57 | 1.6032
| 1 | 5 | 1.63 | 1.77 | 1.68 | 1.59 | 1.58 | 1.5289 | 1.63 | 1.7 | 1.58 | 1.78 | 1.5 | 1.58 | 1.63 | 1.6 | 1.62 | 1.66 | 1.64 | 1.509 | 1.67 | 1.55 | 1.73 | 1.64 | 1.5374 | 1.61 | 1.53 | 11
| 1 | 6 | 1.63 | 1.62 | 1.69 | 1.56 | 1.62 | 1.5513 | 1.57 | 1.72 | 1.56 | 1.79 | 1.58 | 1.57 | 1.62 | 1.56 | 1.63 | 1.54 | 1.72 | 1.704 | 1.59 | 1.54 | 1.77 | 1.6 | 1.5352 | 1.61 | 1.49
| Mean |  | 1.625 | 1.755 | 1.6867 | 1.59 | 1.5817 | 1.5606 | 1.5983 | 1.7017 | 1.5583 | 1.785 | 1.555 | 1.5867 | 1.615 | 1.5367 | 1.6333 | 1.575 | 1.6583 | 1.5923 | 1.64 | 1.5583 | 1.705 | 1.965 | 1.5465 | 1.6133 | 1.5417
| Median |  | 1.625 | 1.77 | 1.685 | 1.59 | 1.58 | 1.567 | 1.595 | 1.7 | 1.555 | 1.785 | 1.58 | 1.58 | 1.62 | 1.555 | 1.63 | 1.565 | 1.65 | 1.5455 | 1.65 | 1.555 | 1.7 | 1.625 | 1.548 | 1.61 | 1.54
| Std Dev. |  | 0.0054772 | 0.083126 | 0.008165 | 0.024495 | 0.022286 | 0.018177 | 0.027142 | 0.0098319 | 0.01472 | 0.0054772 | 0.042778 | 0.026583 | 0.022583 | 0.058195 | 0.015055 | 0.066858 | 0.032506 | 0.096701 | 0.027568 | 0.01472 | 0.041352 | 0.85017 | 0.0090007 | 0.005164 | 0.055287
| Rel.Std.Dev. |  | 0.3371% | 4.737% | 0.4841% | 1.541% | 1.409% | 1.165% | 1.698% | 0.5778% | 0.9446% | 0.3068% | 2.751% | 1.675% | 1.398% | 3.787% | 0.9218% | 4.245% | 1.96% | 6.073% | 1.681% | 0.9446% | 2.425% | 43.27% | 0.582% | 0.3201% | 3.586%
| PDM3 |  | 1.36% | 9.469% | 5.207% | -0.8229% | -1.343% | -2.658% | -0.3031% | 6.142% | -2.798% | 11.34% | -3.006% | -1.031% | 0.7365% | -4.15% | 1.88% | -1.759% | 3.439% | -0.6774% | 2.296% | -2.798% | 6.35% | 22.57% | -3.534% | 0.6325% | -3.838%
| Z-Score (Absolute) |  | 0.45 | 2.64 | 1.49 | 0.14 | 0.28 | 0.64 | 0 | 1.74 | 0.67 | 3.15 | 0.73 | 0.2 | 0.28 | 1.04 | 0.59 | 0.39 | 1.01 | 0.1 | 0.7 | 0.67 | 1.8 | 6.18 | 0.87 | 0.25 | 0.96
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 609c (Certified Value 24 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 24.4 | 25.08 | 24.8 | 23.89 | 24.5631 | 24 | 22.9 | 24.7775 | 24.6 | 24.2 | 23.36 | 24.5 | 23.9 | 23.7 | 23.6 | 24.8 | 22.7 | 24.4 | 24.6 | 23.61 | 24 | 22.8839 | 22.3 | 25.3744 | 21.55 | 26.7 | 1
| 1 | 2 | 24.1 | 25.346 | 23.7 | 23.64 | 24.7107 | 23.1 | 23.1 | 24.0845 | 25 | 24.5 | 23.18 | 23.9 | 24 | 23.4 | 24.4 | 25.4 | 22.6 | 24.6 | 24.3 | 23.42 | 24.1 | 22.3141 | 22.4 | 24.843 | 21.95 | 27.4 | 23
| 1 | 3 | 23.9 | 25.612 | 24.9 | 24.01 | 24.6479 | 23.5 | 24 | 24.4835 | 24.9 | 24.4 | 24.18 | 24.7 | 24.2 | 24.2 | 23.7 | 25 | 22.1 | 24.1 | 24.1 | 23.39 | 23.9 | 22.0296 | 22.2 | 25.5182 | 21.33 | 27 | 16
| 1 | 4 | 24.6 | 25.802 | 25 | 23.45 | 25.0173 | 23.8 | 23.6 | 24.6515 | 24.8 | 24.5 | 23.45 | 23.2 | 23.9 | 24.9 | 24.4 | 25.4 | 21.6 | 24 | 24.1 | 23.47 | 24.2 | 22.3236 | 23.3 | 24.4204 | 25.08 | 28.6 | 24.0481
| 1 | 5 | 24.2 | 25.118 | 24.6 | 23.61 | 25.2819 | 23.4 | 23.2 | 24.935 | 24.4 | 24.6 | 23.18 | 24.5 | 23.9 | 25 | 23.9 | 25.3 | 22.3 | 24.6 | 24.2 | 24.91 | 23.7 | 22.8379 | 23.2 | 24.4992 | 21.52 | 27 | 13
| 1 | 6 | 24.3 | 24.89 | 25.1 | 23.86 | 24.7307 | 24.3 | 23.7 | 24.851 | 24.5 | 24.5 | 23.73 | 22.8 | 24.2 | 24.7 | 23.3 | 25.1 | 22.2 | 23.9 | 23.8 | 23.89 | 24.1 | 22.3882 | 22.5 | 25.5775 | 21.77 | 27.9
| Mean |  | 24.25 | 25.308 | 24.6833 | 23.7433 | 24.8253 | 23.6833 | 23.4167 | 24.6305 | 24.7 | 24.45 | 23.5133 | 23.9333 | 24.0167 | 24.3167 | 23.8833 | 25.1667 | 22.25 | 24.2667 | 24.1833 | 23.7817 | 24 | 22.4629 | 22.65 | 25.0388 | 22.2 | 27.4333
| Median |  | 24.25 | 25.232 | 24.85 | 23.75 | 24.7207 | 23.65 | 23.4 | 24.7145 | 24.7 | 24.5 | 23.405 | 24.2 | 23.95 | 24.45 | 23.8 | 25.2 | 22.25 | 24.25 | 24.15 | 23.54 | 24.05 | 22.3559 | 22.45 | 25.1087 | 21.66 | 27.2
| Std Dev. |  | 0.2429 | 0.34661 | 0.51153 | 0.21011 | 0.27116 | 0.43551 | 0.41673 | 0.31084 | 0.23664 | 0.13784 | 0.385 | 0.78145 | 0.1472 | 0.66156 | 0.4446 | 0.24221 | 0.3937 | 0.30768 | 0.26394 | 0.58235 | 0.17889 | 0.33249 | 0.47645 | 0.51857 | 1.4271 | 0.70616
| Rel.Std.Dev. |  | 1.002% | 1.37% | 2.072% | 0.8849% | 1.092% | 1.839% | 1.78% | 1.262% | 0.9581% | 0.5638% | 1.637% | 3.265% | 0.6129% | 2.721% | 1.862% | 0.9624% | 1.769% | 1.268% | 1.091% | 2.449% | 0.7454% | 1.48% | 2.104% | 2.071% | 6.429% | 2.574%
| PDM3 |  | 0.8397% | 5.239% | 2.642% | -1.267% | 3.232% | -1.517% | -2.626% | 2.422% | 2.711% | 1.671% | -2.224% | -0.4771% | -0.1306% | 1.117% | -0.685% | 4.651% | -7.477% | 0.909% | 0.5625% | -1.108% | -0.1999% | -6.592% | -5.814% | 4.12% | -7.685% | 14.08%
| Z-Score (Absolute) |  | 0.18 | 1.46 | 0.71 | 0.43 | 0.88 | 0.5 | 0.83 | 0.64 | 0.73 | 0.42 | 0.71 | 0.2 | 0.1 | 0.26 | 0.26 | 1.29 | 2.24 | 0.2 | 0.1 | 0.39 | 0.12 | 1.98 | 1.76 | 1.14 | 2.3 | 4.04
| Analytical results for Al in OREAS 609c (Certified Value 7.09 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.3 | 5.88 | 7.24 | 7.114 | 7.0154 | 7.1 | 6.83 | 7.4769 | 7.46 | 6.8 | 7.2985 | 6.98 | 7.18 | 6.76 | 7.29 | 7.39 | 6.7685 | 6.59 | 7.68 | 6.98 | 7.25 | 6.9554 | 7.04 | 7.2498 | 6.721 | 6.86 | 1
| 1 | 2 | 7.27 | 6.9 | 7.17 | 7.2669 | 7.1699 | 6.84 | 6.64 | 7.4507 | 7.3 | 6.98 | 7.4414 | 6.88 | 7.07 | 6.96 | 7.52 | 7.45 | 6.8029 | 5.88 | 7.57 | 7.1 | 7.35 | 6.7331 | 6.99 | 6.9809 | 7.2343 | 6.65 | 0
| 1 | 3 | 7.22 | 6.34 | 7.24 | 7.2906 | 7.0237 | 7.15 | 6.88 | 7.4293 | 7.3 | 6.82 | 7.5069 | 6.94 | 7.07 | 7.14 | 7.19 | 7.49 | 6.8455 | 5.86 | 7.58 | 7.01 | 7.29 | 6.7401 | 6.93 | 7.2331 | 6.7807 | 6.3 | 16
| 1 | 4 | 7.28 | 6.92 | 7.08 | 7.1628 | 7.1145 | 7.1 | 6.95 | 7.4728 | 7.23 | 6.87 | 7.721 | 6.74 | 6.96 | 7.12 | 7.49 | 7.54 | 6.6174 | 6.38 | 7.58 | 7.14 | 7.25 | 6.727 | 7.46 | 6.8814 | 7.1012 | 7.66 | 7.0856
| 1 | 5 | 7.26 | 6.76 | 7.23 | 7.383 | 7.1549 | 6.94 | 6.8 | 7.4903 | 7.16 | 6.66 | 7.278 | 7.13 | 7.16 | 7.12 | 7.32 | 7.57 | 6.6912 | 6.56 | 7.35 | 6.99 | 7.18 | 6.9334 | 7.2 | 6.8195 | 7.0061 | 6.54 | 14
| 1 | 6 | 7.26 | 6.51 | 7.24 | 7.1262 | 7.1718 | 7.25 | 6.84 | 7.4808 | 7.11 | 6.62 | 7.6 | 6.81 | 7.02 | 7.03 | 7.21 | 7.48 | 6.7145 | 5.85 | 7.42 | 6.99 | 7.14 | 6.821 | 6.93 | 6.9674 | 6.6938 | 6.71
| Mean |  | 7.265 | 6.5517 | 7.2 | 7.2239 | 7.1084 | 7.0633 | 6.8233 | 7.4668 | 7.26 | 6.7917 | 7.4743 | 6.9133 | 7.0767 | 7.0217 | 7.3367 | 7.4867 | 6.74 | 6.1867 | 7.53 | 7.035 | 7.2433 | 6.8183 | 7.0917 | 7.022 | 6.9229 | 6.7867
| Median |  | 7.265 | 6.635 | 7.235 | 7.2149 | 7.1347 | 7.1 | 6.835 | 7.4749 | 7.265 | 6.81 | 7.4741 | 6.91 | 7.07 | 7.075 | 7.305 | 7.485 | 6.7415 | 6.13 | 7.575 | 7 | 7.25 | 6.7805 | 7.015 | 6.9741 | 6.8934 | 6.68
| Std Dev. |  | 0.026646 | 0.39952 | 0.064807 | 0.10674 | 0.07185 | 0.14841 | 0.10367 | 0.022608 | 0.12377 | 0.13363 | 0.17219 | 0.13706 | 0.083106 | 0.14538 | 0.13938 | 0.064083 | 0.082415 | 0.36153 | 0.12133 | 0.06775 | 0.075277 | 0.10369 | 0.20624 | 0.17996 | 0.22322 | 0.46698
| Rel.Std.Dev. |  | 0.3668% | 6.098% | 0.9001% | 1.478% | 1.011% | 2.101% | 1.519% | 0.3028% | 1.705% | 1.968% | 2.304% | 1.983% | 1.174% | 2.071% | 1.9% | 0.856% | 1.223% | 5.844% | 1.611% | 0.963% | 1.039% | 1.521% | 2.908% | 2.563% | 3.224% | 6.881%
| PDM3 |  | 2.532% | -7.535% | 1.615% | 1.953% | 0.322% | -0.3137% | -3.701% | 5.381% | 2.462% | -4.148% | 5.486% | -2.431% | -0.1255% | -0.9018% | 3.544% | 5.661% | -4.877% | -12.69% | 6.272% | -0.7136% | 2.227% | -3.771% | 0.08617% | -0.8968% | -2.296% | -4.218%
| Z-Score (Absolute) |  | 0.68 | 1.82 | 0.46 | 0.54 | 0.13 | 0.02 | 0.86 | 1.39 | 0.67 | 0.97 | 1.42 | 0.55 | 0.02 | 0.17 | 0.93 | 1.46 | 1.16 | 3.09 | 1.61 | 0.12 | 0.61 | 0.88 | 0.08 | 0.17 | 0.52 | 0.99
| Analytical results for As in OREAS 609c (Certified Value 884 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 898 | 922.5 | 522 | 893.5 | 863.3257 | 762 | 900 | 864 | 900 | 855 | 899 | 916 | 835 | 925 | 949 | 859 | 924 | 892 | 709.8 | 842 | 800.6847 | 875 | 865.1872 | 847.884 | 911 |  | 1
| 1 | 2 | 887 | 932 | 525 | 891.3 | 857.4108 | 773 | 923 | 867 | 900 | 870 | 908 | 917 | 852 | 944 | 948 | 868.7 | 945 | 859 | 708.6 | 838 | 696.7658 | 857 | 819.9174 | 796.43 | 903 |  | 13
| 1 | 3 | 891 | 921.6 | 560 | 903.6 | 856.0199 | 805 | 945 | 882 | 909 | 877 | 893 | 919 | 879 | 917 | 958 | 860.5 | 922 | 600 | 728.8 | 833 | 690.9305 | 865 | 869.0434 | 758.975 | 938 |  | 16
| 1 | 4 | 901 | 938.9 | 488 | 874.9 | 858.5015 | 789 | 936 | 877 | 906 | 902 | 860 | 904 | 869 | 948 | 958 | 819.3 | 922 | 885 | 701.1 | 847 | 755.4495 | 904 | 832.5563 | 790.481 | 955 |  | 883.7257
| 1 | 5 | 898 | 922.5 | 502 | 894.1 | 876.2181 | 768 | 901 | 852 | 898 | 857 | 912 | 921 | 871 | 934 | 971 | 833.9 | 947 | 911 | 715.9 | 837 | 800.247 | 887 | 837.8302 | 821.022 | 897 |  | 15
| 1 | 6 | 898 | 917.6 | 518 | 886.3 | 855.1795 | 809 | 938 | 867 | 917 | 878 | 874 | 913 | 862 | 916 | 955 | 845.8 | 934 | 891 | 708.6 | 841 | 661.5742 | 846 | 870.2209 | 803.995 | 963
| Mean |  | 895.5 | 925.85 | 519.1667 | 890.6167 | 861.1093 | 784.3333 | 923.8333 | 868.1667 | 905 | 873.1667 | 891 | 915 | 861.3333 | 930.6667 | 956.5 | 847.8667 | 932.3333 | 839.6667 | 712.1333 | 839.6667 | 734.2753 | 872.3333 | 849.1259 | 803.1312 | 927.8333
| Median |  | 898 | 922.5 | 520 | 892.4 | 857.9561 | 781 | 929.5 | 867 | 903 | 873.5 | 896 | 916.5 | 865.5 | 929.5 | 956.5 | 852.4 | 929 | 888 | 709.2 | 839.5 | 726.1077 | 870 | 851.5087 | 800.2125 | 924.5
| Std Dev. |  | 5.3198 | 7.9586 | 24.4001 | 9.5407 | 7.9339 | 19.7552 | 19.4259 | 10.496 | 7.2111 | 17.1513 | 20.2386 | 6.0332 | 15.7818 | 13.5892 | 8.3126 | 18.6171 | 11.5007 | 118.6013 | 9.4246 | 4.8028 | 59.6356 | 21.0111 | 21.7023 | 29.9277 | 28.0315
| Rel.Std.Dev. |  | 0.5941% | 0.8596% | 4.7% | 1.071% | 0.9214% | 2.519% | 2.103% | 1.209% | 0.7968% | 1.964% | 2.271% | 0.6594% | 1.832% | 1.46% | 0.8691% | 2.196% | 1.234% | 14.12% | 1.323% | 0.572% | 8.122% | 2.409% | 2.556% | 3.726% | 3.021%
| PDM3 |  | 1.332% | 4.767% | -41.25% | 0.7798% | -2.559% | -11.25% | 4.538% | -1.761% | 2.407% | -1.195% | 0.8231% | 3.539% | -2.534% | 5.312% | 8.235% | -4.058% | 5.5% | -4.986% | -19.42% | -4.986% | -16.91% | -1.289% | -3.915% | -9.12% | 4.991%
| Z-Score (Absolute) |  | 0.48 | 1.1 | 7.29 | 0.38 | 0.23 | 1.82 | 1.06 | 0.09 | 0.67 | 0.02 | 0.39 | 0.88 | 0.23 | 1.2 | 1.74 | 0.51 | 1.24 | 0.67 | 3.31 | 0.67 | 2.85 | 0 | 0.48 | 1.43 | 1.15
| Analytical results for B in OREAS 609c (Indicative Value 19 ppm) | [27 empty] | Indicative
B | Round | Replicate | 2.01: Lab; 2.03: Lab |  | 1
| No | No | 2.01: 2.01; 2.03: 2.23 |  | ppm
|  |  | 2.01: 4A*OES; 2.03: 4A*OES |  | 1
1
| 1 | 1 | 2.01: 19; 2.03: < 0.1 |  | 1
| 1 | 2 | 2.01: 20; 2.03: < 0.1 |  | 3
| 1 | 3 | 2.01: 18; 2.03: < 0.1 |  | 16
| 1 | 4 | 2.01: 20; 2.03: < 0.1 |  | 19
| 1 | 5 | 2.01: 18; 2.03: < 0.1 |  | 9
| 1 | 6 | 2.01: 19; 2.03: < 0.1
| Mean |  | 2.01: 19
| Median |  | 2.01: 19
| Std Dev. |  | 2.01: 0.89443
| Rel.Std.Dev. |  | 2.01: 4.708%
PDM3 | bar: 2.01
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.67
| Analytical results for Ba in OREAS 609c (Indicative Value 1443 ppm) | [27 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 1567 | 212 | 39 | 2721.3 | 466.3068 | 250 | 1000 | 2900 | 850 | 800 | 1780 | 730 | 260 | 2795 | 1130 | 1970 | 2726 | 2359 | 1417.4084 | 2680 | 2151.1215 | 2522.995 | 239 |  |  |  | 1
| 1 | 2 | 1598 | 182 | 40 | 2725.4 | 463.2441 | 231 | 1190 | 2990 | 840 | 650 | 1260 | 1240 | 330 | 2804 | 866 | 1940 | 2736 | 2242 | 1714.6129 | 2570 | 2284.7589 | 2460.474 | 302 |  |  |  | 5
| 1 | 3 | 1580 | 182 | 39 | 2766.5 | 434.3427 | 289 | 1010 | 3010 | 800 | 580 | 930 | 550 | 240 | 2816 | 804 | 1930 | 2744 | 2358 | 1243.0101 | 2600 | 2227.9781 | 2355.608 | 992 |  |  |  | 16
| 1 | 4 | 1586 | 182 | 43 | 2699.1 | 472.9202 | 284 | 847 | 3010 | 670 | 650 | 1050 | 1060 | 250 | 2737 | 873 | 1950 | 2748 | 2380 | 1793.8822 | 2630 | 2122.9257 | 2413.145 | 366 |  |  |  | 1442.9064
| 1 | 5 | 1564 | 202 | 38 | 2723.1 | 466.5023 | 222 | 876 | 2980 | 740 | 660 | 1380 | 450 | 330 | 2756 | 716 | 1940 | 2711 | 2228 | 1662.9922 | 2660 | 2246.076 | 2507.473 | 1239 |  |  |  | 9
| 1 | 6 | 1572 | 182 | 37 | 2750.1 | 471.9608 | 225 | 1160 | 2900 | 830 | 570 | 1170 | 640 | 370 | 2779 | 1076 | 1900 | 2684 | 2361 | 1482.5586 | 2610 | 2213.2074 | 2504.232 | 802
| Mean |  | 1577.8333 | 190.3333 | 39.3333 | 2730.9167 | 462.5462 | 250.1667 | 1013.8333 | 2965 | 788.3333 | 651.6667 | 1261.6667 | 778.3333 | 296.6667 | 2781.1667 | 910.8333 | 1938.3333 | 2724.8333 | 2321.3333 | 1552.4107 | 2625 | 2207.678 | 2460.6545 | 656.6667
| Median |  | 1576 | 182 | 39 | 2724.25 | 466.4046 | 240.5 | 1005 | 2985 | 815 | 650 | 1215 | 685 | 295 | 2787 | 869.5 | 1940 | 2731 | 2358.5 | 1572.7754 | 2620 | 2220.5928 | 2482.353 | 584
| Std Dev. |  | 12.8128 | 13.2916 | 2.0656 | 23.7928 | 14.2967 | 29.8223 | 141.0226 | 51.672 | 70.2614 | 82.3205 | 298.6246 | 307.8582 | 53.5413 | 30.0494 | 160.0805 | 23.1661 | 24.0368 | 67.5031 | 207.6921 | 40.3733 | 60.3989 | 65.2216 | 414.1022
| Rel.Std.Dev. |  | 0.812% | 6.983% | 5.252% | 0.8712% | 3.091% | 11.92% | 13.91% | 1.743% | 8.913% | 12.63% | 23.67% | 39.55% | 18.05% | 1.08% | 17.58% | 1.195% | 0.8821% | 2.908% | 13.38% | 1.538% | 2.736% | 2.651% | 63.06%
| PDM3 |  | 9.351% | -86.81% | -97.27% | 89.26% | -67.94% | -82.66% | -29.74% | 105.5% | -45.36% | -54.84% | -12.56% | -46.06% | -79.44% | 92.75% | -36.88% | 34.34% | 88.84% | 60.88% | 7.589% | 81.92% | 53% | 70.53% | -54.49%
| Z-Score (Absolute) |  | 0.23 | 0.76 | 0.87 | 1.05 | 0.57 | 0.72 | 0.18 | 1.21 | 0.34 | 0.43 | 0 | 0.34 | 0.69 | 1.08 | 0.25 | 0.48 | 1.04 | 0.76 | 0.21 | 0.97 | 0.67 | 0.85 | 0.43
| Analytical results for Be in OREAS 609c (Certified Value 2.69 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | < 5 | 3 | 3 | 2.62 | 2.6869 | 2.5 | 2.37 | 3.1719 | 2.5 | 2.8 | 2.76 | 2.69 | 2.81 | 2.61 | 2.66 | 2.74 | 2.69 | 3.41 | 2.16 | 2.84 | 2.75 | 2.2454 | 3 | 2.6745 | 2.122 | 2.45 | 1
| 1 | 2 | < 5 | 3 | 3 | 2.59 | 2.6916 | 2.8 | 2.27 | 3.1503 | 2.5 | 2.8 | 2.79 | 2.69 | 2.79 | 2.57 | 2.73 | 2.77 | 2.71 | 3.49 | 2.9 | 2.84 | 2.69 | 2.2376 | 3 | 2.5454 | 2.161 | 2.59 | 26
| 1 | 3 | < 5 | 2 | 3 | 2.59 | 2.7124 | 3.3 | 2.49 | 3.1494 | 2.6 | 2.8 | 2.8 | 2.64 | 2.81 | 2.62 | 2.63 | 2.79 | 2.71 | 3.35 | 2.51 | 2.78 | 2.89 | 2.1897 | 3 | 2.7833 | 2.092 | 2.61 | 16
| 1 | 4 | < 5 | 3 | 3 | 2.66 | 2.7202 | 2.3 | 2.33 | 3.2 | 2.6 | 2.8 | 2.83 | 2.58 | 2.75 | 2.63 | 2.73 | 2.8 | 2.64 | 3.42 | 2.6 | 2.77 | 2.91 | 2.1851 | 3.1 | 2.5413 | 2.182 | 2.61 | 2.6855
| 1 | 5 | < 5 | 2 | 3 | 2.67 | 2.7295 | 2.7 | 2.33 | 3.1794 | 2.6 | 2.8 | 2.87 | 2.73 | 2.8 | 2.62 | 2.69 | 2.83 | 2.67 | 3.45 | 2.54 | 2.77 | 2.65 | 2.2659 | 3 | 2.7072 | 2.115 | 2.68 | 16
| 1 | 6 | < 5 | 3 | 3 | 2.49 | 2.7042 | 3.1 | 2.38 | 3.1811 | 2.6 | 2.7 | 2.85 | 2.6 | 2.77 | 2.59 | 2.64 | 2.78 | 2.68 | 3.38 | 2.18 | 2.71 | 2.83 | 2.1986 | 3 | 2.6448 | 2.132 | 2.9
| Mean |  |  | 2.6667 | 3 | 2.6033 | 2.7075 | 2.7833 | 2.3617 | 3.172 | 2.5667 | 2.7833 | 2.8167 | 2.655 | 2.7883 | 2.6067 | 2.68 | 2.785 | 2.6833 | 3.4167 | 2.4817 | 2.785 | 2.7867 | 2.2204 | 3.0167 | 2.6494 | 2.134 | 2.64
| Median |  |  | 3 | 3 | 2.605 | 2.7083 | 2.75 | 2.35 | 3.1756 | 2.6 | 2.8 | 2.815 | 2.665 | 2.795 | 2.615 | 2.675 | 2.785 | 2.685 | 3.415 | 2.525 | 2.775 | 2.79 | 2.2181 | 3 | 2.6596 | 2.127 | 2.61
| Std Dev. |  |  | 0.5164 | 0 | 0.065013 | 0.016458 | 0.37103 | 0.073869 | 0.019527 | 0.05164 | 0.040825 | 0.040825 | 0.058224 | 0.024014 | 0.022509 | 0.043818 | 0.030166 | 0.026583 | 0.049666 | 0.27845 | 0.049295 | 0.1069 | 0.03363 | 0.040825 | 0.09425 | 0.032576 | 0.14805
| Rel.Std.Dev. |  |  | 19.36% | 0% | 2.497% | 0.6079% | 13.33% | 3.128% | 0.6156% | 2.012% | 1.467% | 1.449% | 2.193% | 0.8612% | 0.8635% | 1.635% | 1.083% | 0.9907% | 1.454% | 11.22% | 1.77% | 3.836% | 1.515% | 1.353% | 3.557% | 1.527% | 5.608%
| PDM3 |  |  | -0.702% | 11.71% | -3.06% | 0.8176% | 3.642% | -12.06% | 18.12% | -4.426% | 3.642% | 4.883% | -1.136% | 3.828% | -2.936% | -0.2055% | 3.704% | -0.0814% | 27.23% | -7.591% | 3.704% | 3.766% | -17.32% | 12.33% | -1.345% | -20.54% | -1.695%
| Z-Score (Absolute) |  | 1.18 | NA | NA | 0.51 | 0.17 | 0.66 | 2.07 | 3.17 | 0.74 | 0.66 | 0.87 | 0.17 | 0.69 | 0.49 | 0.01 | 0.67 | 0.01 | 4.76 | 1.29 | 0.67 | 0.68 | 2.99 | 2.17 | 0.21 | 3.55 | 0.27
Results from laboratories 2.03 and 2.04 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 609c (Certified Value 64 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
1
| 1 | 1 | 66 | 61.96 | 66.4 | 66.6 | 65.2049 | 64.8 | 62.7 | 63.42 | 66.6 | 65.59 | 68.5 | 73.9 | 63.6 | 63.9 | 60.5 | 61.12 | 69.35 | 60 | 58.81 | 59.57 | 49.3498 | 65.5468 | 50.412 | 60.91 |  |  | 1
| 1 | 2 | 64 | 60.63 | 64.1 | 66.49 | 66.6939 | 66 | 63.2 | 62.65 | 65.3 | 64.64 | 67.5 | 72.4 | 60.9 | 66.4 | 61.5 | 61.1 | 68.26 | 59.6 | 62.7 | 61.42 | 38.9641 | 61.8081 | 50.631 | 61.36 |  |  | 27
| 1 | 3 | 65 | 57.68 | 69.1 | 67.03 | 65.6444 | 67.2 | 68 | 64.67 | 67 | 65.95 | 66.9 | 72 | 62 | 64 | 62.1 | 60 | 67.46 | 59.3 | 61.6 | 59.65 | 39.2167 | 65.0391 | 49.585 | 55.35 |  |  | 16
| 1 | 4 | 64 | 61.96 | 65.8 | 66.19 | 65.7882 | 63.3 | 64.3 | 63.23 | 67.3 | 66.72 | 63.2 | 70.2 | 63.6 | 65.4 | 62.3 | 58.22 | 67.35 | 59.6 | 60.1 | 58.79 | 40.271 | 61.983 | 50.99 | 60.64 |  |  | 64.1513
| 1 | 5 | 63 | 58.96 | 66.5 | 66.18 | 67.7873 | 63.8 | 62.9 | 62.13 | 67.2 | 63.87 | 71.1 | 72.2 | 63.9 | 64.5 | 62.8 | 59.94 | 70 | 62.8 | 61.41 | 60.51 | 44.3223 | 62.125 | 50.024 | 63.63 |  |  | 17
| 1 | 6 | 66 | 61.54 | 67.8 | 65.97 | 66.1457 | 68.6 | 65.9 | 65.03 | 68.9 | 64.6 | 64.3 | 71.1 | 64.2 | 64.1 | 61.4 | 60.41 | 64.97 | 61.1 | 58 | 61.16 | 43.9625 | 65.5629 | 52.163 | 57.98
| Mean |  | 64.6667 | 60.455 | 66.6167 | 66.41 | 66.2107 | 65.6167 | 64.5 | 63.5217 | 67.05 | 65.2283 | 66.9167 | 71.9667 | 63.0333 | 64.7167 | 61.7667 | 60.1317 | 67.8983 | 60.4 | 60.4367 | 60.1833 | 42.6811 | 63.6775 | 50.6342 | 59.9783
| Median |  | 64.5 | 61.085 | 66.45 | 66.34 | 65.967 | 65.4 | 63.75 | 63.325 | 67.1 | 65.115 | 67.2 | 72.1 | 63.6 | 64.3 | 61.8 | 60.205 | 67.86 | 59.8 | 60.755 | 60.08 | 42.1168 | 63.5821 | 50.5215 | 60.775
| Std Dev. |  | 1.2111 | 1.7705 | 1.7105 | 0.37995 | 0.92025 | 2.0439 | 2.0852 | 1.1301 | 1.164 | 1.0452 | 2.8639 | 1.2533 | 1.2941 | 0.99079 | 0.80911 | 1.0678 | 1.7733 | 1.3342 | 1.7956 | 1.019 | 4.0108 | 1.8804 | 0.8929 | 2.897
| Rel.Std.Dev. |  | 1.873% | 2.929% | 2.568% | 0.5721% | 1.39% | 3.115% | 3.233% | 1.779% | 1.736% | 1.602% | 4.28% | 1.741% | 2.053% | 1.531% | 1.31% | 1.776% | 2.612% | 2.209% | 2.971% | 1.693% | 9.397% | 2.953% | 1.763% | 4.83%
| PDM3 |  | 0.8034% | -5.762% | 3.843% | 3.521% | 3.21% | 2.284% | 0.5436% | -0.9815% | 4.519% | 1.679% | 4.311% | 12.18% | -1.743% | 0.8813% | -3.717% | -6.266% | 5.841% | -5.848% | -5.79% | -6.185% | -33.47% | -0.7385% | -21.07% | -6.505%
| Z-Score (Absolute) |  | 0.15 | 0.91 | 0.64 | 0.58 | 0.53 | 0.38 | 0.1 | 0.14 | 0.75 | 0.29 | 0.71 | 1.98 | 0.27 | 0.16 | 0.58 | 1 | 0.96 | 0.93 | 0.92 | 0.98 | 5.39 | 0.1 | 3.39 | 1.04
| Analytical results for Ca in OREAS 609c (Certified Value 1.09 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.13 | 0.97 | 1.1 | 1.0951 | 1.0527 | 1.07 | 1.08 | 1.1201 | 1.08 | 1.03 | 1.0709 | 1.11 | 1.11 | 1.06 | 1.14 | 1.12 | 1.1148 | 1.08 | 1.12 | 1.05 | 1.12 | 1.1266 | 1.086 | 1.091 | 1.0153 | 1.051 | 1
| 1 | 2 | 1.12 | 1.14 | 1.09 | 1.115 | 1.0557 | 1.08 | 1.06 | 1.1591 | 1.07 | 1.03 | 1.0931 | 1.09 | 1.1 | 1.05 | 1.18 | 1.13 | 1.1244 | 1.008 | 1.12 | 1.07 | 1.09 | 1.0937 | 1.072 | 1.0527 | 1.0842 | 1.103 | 0
| 1 | 3 | 1.11 | 1.07 | 1.1 | 1.1227 | 1.0517 | 1.1 | 1.09 | 1.1557 | 1.1 | 1.04 | 1.1018 | 1.11 | 1.1 | 1.08 | 1.12 | 1.14 | 1.1303 | 0.99 | 1.1 | 1.07 | 1.09 | 1.1108 | 1.072 | 1.0878 | 1.0468 | 1.05 | 16
| 1 | 4 | 1.13 | 1.13 | 1.08 | 1.1023 | 1.0659 | 1.07 | 1.1 | 1.1132 | 1.08 | 1.07 | 1.1312 | 1.08 | 1.08 | 1.08 | 1.16 | 1.14 | 1.0929 | 1.057 | 1.12 | 1.08 | 1.05 | 1.0837 | 1.144 | 1.0664 | 1.071 | 1.164 | 1.0923
| 1 | 5 | 1.14 | 1.13 | 1.09 | 1.1363 | 1.0817 | 1.14 | 1.06 | 1.1147 | 1.11 | 1.04 | 1.0673 | 1.14 | 1.11 | 1.07 | 1.14 | 1.16 | 1.1082 | 1.09 | 1.12 | 1.06 | 1.09 | 1.1224 | 1.108 | 1.0533 | 1.0379 | 1.061 | 18
| 1 | 6 | 1.13 | 1.08 | 1.1 | 1.0963 | 1.0674 | 1.19 | 1.1 | 1.1251 | 1.11 | 1.03 | 1.1133 | 1.09 | 1.1 | 1.05 | 1.12 | 1.14 | 1.1158 | 0.981 | 1.14 | 1.05 | 1.11 | 1.1022 | 1.065 | 1.0665 | 1.0128 | 1.066
| Mean |  | 1.1267 | 1.0867 | 1.0933 | 1.1113 | 1.0625 | 1.1083 | 1.0817 | 1.1313 | 1.0917 | 1.04 | 1.0963 | 1.1033 | 1.1 | 1.065 | 1.1433 | 1.1383 | 1.1144 | 1.0343 | 1.12 | 1.0633 | 1.0917 | 1.1066 | 1.0912 | 1.0696 | 1.0447 | 1.0825
| Median |  | 1.13 | 1.105 | 1.095 | 1.1086 | 1.0608 | 1.09 | 1.085 | 1.1226 | 1.09 | 1.035 | 1.0975 | 1.1 | 1.1 | 1.065 | 1.14 | 1.14 | 1.1153 | 1.0325 | 1.12 | 1.065 | 1.09 | 1.1065 | 1.079 | 1.0664 | 1.0424 | 1.0635
| Std Dev. |  | 0.010328 | 0.064083 | 0.008165 | 0.016353 | 0.011524 | 0.047924 | 0.018348 | 0.020657 | 0.017224 | 0.015492 | 0.024632 | 0.021602 | 0.010954 | 0.013784 | 0.023381 | 0.013292 | 0.013081 | 0.047331 | 0.012649 | 0.012111 | 0.024014 | 0.016598 | 0.030069 | 0.016488 | 0.028918 | 0.044375
| Rel.Std.Dev. |  | 0.9167% | 5.897% | 0.7468% | 1.472% | 1.085% | 4.324% | 1.696% | 1.826% | 1.578% | 1.49% | 2.247% | 1.958% | 0.9959% | 1.294% | 2.045% | 1.168% | 1.174% | 4.576% | 1.129% | 1.139% | 2.2% | 1.5% | 2.756% | 1.542% | 2.768% | 4.099%
| PDM3 |  | 3.15% | -0.5126% | 0.09776% | 1.741% | -2.724% | 1.471% | -0.9704% | 3.574% | -0.05483% | -4.785% | 0.3663% | 1.013% | 0.7081% | -2.496% | 4.675% | 4.218% | 2.026% | -5.304% | 2.539% | -2.649% | -0.05483% | 1.309% | -0.1006% | -2.074% | -4.358% | -0.8941%
| Z-Score (Absolute) |  | 1.13 | 0.19 | 0.03 | 0.62 | 0.99 | 0.52 | 0.36 | 1.29 | 0.03 | 1.74 | 0.12 | 0.36 | 0.25 | 0.91 | 1.69 | 1.52 | 0.73 | 1.93 | 0.91 | 0.97 | 0.03 | 0.47 | 0.05 | 0.76 | 1.59 | 0.33
| Analytical results for Cd in OREAS 609c (Certified Value 5.71 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 6.68 | 5.81 | 5.3 | 5.72 | 5.5094 | 6.5 | 5.38 | 7.8061 | 5.86 | 5.84 | 5.55 | 6.22 | 5.58 | 5.36 | 5.54 | 5.58 | 5.51 | 5.49 | 5.78 | 6.12 | 6.24 | 5.2971 | < 1 | 5.2855 | 5.226 | 5.51 | 1
| 1 | 2 | 6.35 | 5.65 | 5.4 | 5.79 | 5.5482 | 6.4 | 5.47 | 7.8158 | 5.75 | 5.74 | 5.6 | 6.28 | 5.73 | 5.82 | 5.87 | 5.71 | 5.43 | 5.51 | 5.59 | 6.19 | 6.36 | 5.0469 | < 1 | 5.1092 | 5.457 | 5.29 | 28
| 1 | 3 | 6.04 | 5.21 | 5.4 | 5.74 | 5.573 | 6.4 | 5.68 | 7.9072 | 5.84 | 5.77 | 6.01 | 6.15 | 5.71 | 5.8 | 5.63 | 5.66 | 5.4 | 5.5 | 5.63 | 6.13 | 6.15 | 5.0764 | < 1 | 5.2977 | 5.362 | 5.82 | 16
| 1 | 4 | 6.24 | 5.94 | 5.5 | 5.66 | 5.5703 | 6.4 | 5.54 | 7.9554 | 5.66 | 5.83 | 5.89 | 5.94 | 5.62 | 6.16 | 5.87 | 5.75 | 5.12 | 5.68 | 6.17 | 6.34 | 6.23 | 5.2698 | < 1 | 5.2092 | 5.42 | 6 | 5.7087
| 1 | 5 | 6.19 | 5.47 | 5.3 | 5.66 | 5.6423 | 6.4 | 5.41 | 7.7822 | 5.53 | 5.83 | 5.5 | 6.09 | 5.6 | 6.19 | 5.72 | 5.75 | 5.46 | 5.73 | 5.9 | 6.31 | 6.3 | 5.2887 | < 1 | 5.1159 | 5.362 | 5.88 | 19
| 1 | 6 | 6.68 | 5.8 | 5.4 | 5.66 | 5.6199 | 6.6 | 5.66 | 7.8034 | 5.61 | 5.75 | 5.68 | 5.82 | 5.78 | 6.16 | 5.67 | 5.68 | 5.39 | 5.59 | 5.84 | 6.09 | 6.26 | 5.0207 | < 1 | 5.3761 | 5.505 | 5.94
| Mean |  | 6.3633 | 5.6467 | 5.3833 | 5.705 | 5.5772 | 6.45 | 5.5233 | 7.845 | 5.7083 | 5.7933 | 5.705 | 6.0833 | 5.67 | 5.915 | 5.7167 | 5.6883 | 5.385 | 5.5833 | 5.8183 | 6.1967 | 6.2567 | 5.1666 |  | 5.2323 | 5.3887 | 5.74
| Median |  | 6.295 | 5.725 | 5.4 | 5.69 | 5.5716 | 6.4 | 5.505 | 7.811 | 5.705 | 5.8 | 5.64 | 6.12 | 5.665 | 5.99 | 5.695 | 5.695 | 5.415 | 5.55 | 5.81 | 6.16 | 6.25 | 5.1731 |  | 5.2474 | 5.391 | 5.85
| Std Dev. |  | 0.26478 | 0.26763 | 0.075277 | 0.054314 | 0.048106 | 0.083666 | 0.12628 | 0.069422 | 0.13106 | 0.045019 | 0.20246 | 0.1742 | 0.080994 | 0.32433 | 0.13261 | 0.064317 | 0.13693 | 0.10191 | 0.20971 | 0.10501 | 0.070616 | 0.13141 |  | 0.10679 | 0.097023 | 0.27893
| Rel.Std.Dev. |  | 4.161% | 4.74% | 1.398% | 0.952% | 0.8626% | 1.297% | 2.286% | 0.8849% | 2.296% | 0.7771% | 3.549% | 2.864% | 1.428% | 5.483% | 2.32% | 1.131% | 2.543% | 1.825% | 3.604% | 1.695% | 1.129% | 2.543% |  | 2.041% | 1.801% | 4.859%
| PDM3 |  | 11.47% | -1.086% | -5.699% | -0.06426% | -2.304% | 12.99% | -3.247% | 37.42% | -0.005872% | 1.483% | -0.06426% | 6.563% | -0.6774% | 3.614% | 0.1401% | -0.3562% | -5.67% | -2.196% | 1.921% | 8.548% | 9.599% | -9.496% |  | -8.345% | -5.606% | 0.5488%
| Z-Score (Absolute) |  | 2.27 | 0.2 | 1.11 | 0 | 0.44 | 2.57 | 0.63 | 7.37 | 0.01 | 0.3 | 0 | 1.3 | 0.12 | 0.72 | 0.04 | 0.06 | 1.1 | 0.42 | 0.39 | 1.69 | 1.9 | 1.85 | 17.92 | 1.63 | 1.09 | 0.12
| Analytical results for Ce in OREAS 609c (Certified Value 79 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.27 | 2.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 0
1
| 1 | 1 | 80 | 73.99 | 75 | 84.06 | 81.312 | 80 | 84.5 | 98.224 | 77.3 | 89.48 | 87.2 | 79.8 | 80.5 | 77.7 | 87.7 | 66.6 | 73.09 | 81.1 | 79.54 | 68.7553 | 76 | 63.771 | 73.86 |  |  |  | 1
| 1 | 2 | 78 | 73.93 | 75 | 83.99 | 80.4678 | 80.1 | 85.9 | 98.626 | 77.7 | 88.96 | 80.9 | 83.2 | 73.6 | 82.5 | 89.3 | 66.2 | 65.27 | 79.4 | 81.25 | 75.705 | 77 | 60.455 | 73.4 |  |  |  | 29
| 1 | 3 | 79 | 73.53 | 79 | 84.19 | 81.3762 | 84.5 | 89.4 | 98.604 | 77.9 | 91.4 | 87.7 | 80.1 | 72.8 | 79.2 | 88.1 | 64.9 | 65.85 | 79.1 | 79.06 | 69.7143 | 77 | 50.15 | 77.43 |  |  |  | 16
| 1 | 4 | 78 | 78.36 | 74 | 84.01 | 80.5489 | 78.9 | 88.5 | 98.29 | 77.8 | 92.45 | 82.8 | 80.8 | 77 | 85.4 | 89.7 | 64 | 69.14 | 81.3 | 78.55 | 78.2967 | 81 | 63.331 | 77.28 |  |  |  | 78.952
| 1 | 5 | 70 | 74.68 | 76 | 83.7 | 81.5555 | 80 | 85.8 | 98.1 | 75.2 | 87.5 | 79.9 | 81 | 79 | 80.4 | 90 | 64 | 70.66 | 84.7 | 80.66 | 74.0109 | 78 | 50.606 | 71.13 |  |  |  | 20
| 1 | 6 | 81 | 74.92 | 76 | 84.6 | 81.1975 | 80.5 | 87 | 98.54 | 76.8 | 84.83 | 79 | 80.9 | 77 | 80.1 | 87.9 | 64 | 63.85 | 80.7 | 81.99 | 76.9479 | 76 | 59.481 | 80.23
| Mean |  | 77.6667 | 74.9017 | 75.8333 | 84.0917 | 81.0763 | 80.6667 | 86.85 | 98.3973 | 77.1167 | 89.1033 | 82.9167 | 80.9667 | 76.65 | 80.8833 | 88.7833 | 64.95 | 67.9767 | 81.05 | 80.175 | 73.905 | 77.5 | 57.9657 | 75.555
| Median |  | 78.5 | 74.335 | 75.5 | 84.035 | 81.2547 | 80.05 | 86.45 | 98.415 | 77.5 | 89.22 | 81.85 | 80.85 | 77 | 80.25 | 88.7 | 64.45 | 67.495 | 80.9 | 80.1 | 74.858 | 77 | 59.968 | 75.57
| Std Dev. |  | 3.9328 | 1.7701 | 1.7224 | 0.29647 | 0.45567 | 1.9521 | 1.8317 | 0.22151 | 1.0226 | 2.7375 | 3.7349 | 1.1944 | 2.9905 | 2.714 | 1.0008 | 1.1828 | 3.5659 | 2.0017 | 1.3396 | 3.8956 | 1.8708 | 6.103 | 3.33
| Rel.Std.Dev. |  | 5.064% | 2.363% | 2.271% | 0.3526% | 0.562% | 2.42% | 2.109% | 0.2251% | 1.326% | 3.072% | 4.504% | 1.475% | 3.901% | 3.355% | 1.127% | 1.821% | 5.246% | 2.47% | 1.671% | 5.271% | 2.414% | 10.53% | 4.407%
| PDM3 |  | -1.628% | -5.13% | -3.95% | 6.51% | 2.691% | 2.172% | 10% | 24.63% | -2.325% | 12.86% | 5.022% | 2.552% | -2.916% | 2.446% | 12.45% | -17.73% | -13.9% | 2.657% | 1.549% | -6.393% | -1.839% | -26.58% | -4.303%
| Z-Score (Absolute) |  | 0.43 | 0.91 | 0.75 | 0.67 | 0.16 | 0.08 | 1.15 | 3.14 | 0.53 | 1.54 | 0.47 | 0.14 | 0.61 | 0.12 | 1.48 | 2.62 | 2.1 | 0.15 | 0 | 1.08 | 0.46 | 3.82 | 0.8
| Analytical results for Co in OREAS 609c (Certified Value 7.13 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 8.3 | 7 | 7.5 | 7.1 | 6.7931 | 6 | 6.87 | 6.98 | 7 | 7.2 | 7.6 | 6.9 | 7.2 | 7.3 | 7.1 | 6.3 | 7.7 | 7.3 | 9.12 | 7.4 | 17.1498 | 3 | 7.1425 | 6.545 | 5.8 |  | 1
| 1 | 2 | 7.8 | 6.9 | 7.2 | 7 | 6.796 | 6 | 7.06 | 6.77 | 7 | 7.2 | 7.4 | 7.1 | 6.5 | 7.6 | 7 | 6.3 | 7.8 | 7.4 | 9.21 | 7.7 | 17.6283 | 4 | 6.6905 | 6.588 | 6 |  | 17
| 1 | 3 | 7.8 | 6.4 | 7.4 | 7.1 | 6.7826 | 6 | 7.26 | 6.93 | 7 | 7.3 | 7.6 | 7.1 | 6.6 | 7.2 | 7.1 | 6.2 | 7.7 | 7.2 | 8.95 | 7.5 | 18.8948 | 4 | 7.0781 | 6.654 | 5.2 |  | 16
| 1 | 4 | 8 | 7 | 7.5 | 7 | 6.836 | 5 | 7.12 | 6.92 | 7 | 7.3 | 7.1 | 7 | 6.9 | 7.5 | 7.2 | 6 | 7.8 | 7.5 | 9.02 | 7.5 | 16.0211 | 4 | 6.5943 | 7.006 | 6.1 |  | 7.1274
| 1 | 5 | 7.5 | 6.9 | 7.4 | 7.1 | 6.8802 | 6 | 7 | 6.67 | 7 | 7.4 | 7.3 | 7.1 | 6.9 | 7.4 | 7 | 6.2 | 8 | 7.6 | 9.19 | 7.4 | 13.4757 | 5 | 6.6876 | 6.471 | 5.9 |  | 21
| 1 | 6 | 8.1 | 7.1 | 7.6 | 6.9 | 6.8516 | 6 | 7.15 | 6.84 | 7 | 7.2 | 7.1 | 7.1 | 6.9 | 7.2 | 6.8 | 6.3 | 7.7 | 7.5 | 8.67 | 7.7 | 18.0156 | 3 | 7.2166 | 6.513 | 5.6
| Mean |  | 7.9167 | 6.8833 | 7.4333 | 7.0333 | 6.8233 | 5.8333 | 7.0767 | 6.8517 | 7 | 7.2667 | 7.35 | 7.05 | 6.8333 | 7.3667 | 7.0333 | 6.2167 | 7.7833 | 7.4167 | 9.0267 | 7.5333 | 16.8642 | 3.8333 | 6.9016 | 6.6295 | 5.7667
| Median |  | 7.9 | 6.95 | 7.45 | 7.05 | 6.816 | 6 | 7.09 | 6.88 | 7 | 7.25 | 7.35 | 7.1 | 6.9 | 7.35 | 7.05 | 6.25 | 7.75 | 7.45 | 9.07 | 7.5 | 17.3891 | 4 | 6.8843 | 6.5665 | 5.85
| Std Dev. |  | 0.27869 | 0.24833 | 0.13663 | 0.08165 | 0.038758 | 0.40825 | 0.13397 | 0.11583 | 0 | 0.08165 | 0.22583 | 0.083666 | 0.25033 | 0.1633 | 0.13663 | 0.1169 | 0.1169 | 0.1472 | 0.20106 | 0.13663 | 1.9135 | 0.75277 | 0.27323 | 0.19489 | 0.3266
| Rel.Std.Dev. |  | 3.52% | 3.608% | 1.838% | 1.161% | 0.568% | 6.999% | 1.893% | 1.691% | 0% | 1.124% | 3.073% | 1.187% | 3.663% | 2.217% | 1.943% | 1.881% | 1.502% | 1.985% | 2.227% | 1.814% | 11.35% | 19.64% | 3.959% | 2.94% | 5.664%
| PDM3 |  | 11.07% | -3.424% | 4.292% | -1.32% | -4.267% | -18.16% | -0.7119% | -3.869% | -1.788% | 1.954% | 3.123% | -1.086% | -4.126% | 3.357% | -1.32% | -12.78% | 9.203% | 4.058% | 26.65% | 5.695% | 136.6% | -46.22% | -3.168% | -6.986% | -19.09%
| Z-Score (Absolute) |  | 1.95 | 0.41 | 0.85 | 0.07 | 0.55 | NA | 0.03 | 0.48 | NA | 0.46 | 0.66 | 0.03 | 0.53 | 0.69 | 0.07 | 1.94 | 1.65 | 0.81 | 4.49 | 1.07 | 22.4 | NA | 0.37 | 0.99 | 2.96
Results from laboratories 2.07, 2.11 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 609c (Certified Value 15.9 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 1
1
| 1 | 1 | 17 | 17 | 16 | 19 | 15.9528 | 18 | 12.8 | 16.3 | 14 | 15 | 16 | 17 | 15 | 15 | 17 | 15.1 | 15 | 14 | 29.58 | 19 | 52.8337 | 17 | 13.7819 | 17.004 | 12 |  | 1
| 1 | 2 | 17 | 15 | 14 | 17 | 15.8566 | 22 | 13.7 | 15.7 | 15 | 16 | 17 | 16 | 15 | 13 | 17 | 14.8 | 13 | 14 | 27.36 | 22 | 45.7865 | 17 | 15.2831 | 16.049 | 15 |  | 31
| 1 | 3 | 17 | 15 | 15 | 19 | 15.7491 | 21 | 13 | 17.3 | 15 | 15 | 16 | 17 | 16 | 15 | 18 | 14.7 | 15 | 15 | 26.78 | 18 | 60.2568 | 17 | 14.9477 | 16.888 | 11 |  | 16
| 1 | 4 | 18 | 15 | 15 | 18 | 16.0839 | 15 | 12.6 | 16 | 14 | 18 | 15 | 17 | 16 | 15 | 16 | 15 | 16 | 14 | 29.03 | 20 | 50.1843 | 17 | 19.6473 | 18.197 | 13 |  | 15.8587
| 1 | 5 | 17 | 16 | 15 | 18 | 15.686 | 24 | 12.1 | 16.3 | 14 | 16 | 17 | 18 | 15 | 15 | 17 | 14.8 | 14 | 15 | 27.42 | 18 | 58.4417 | 17 | 17.348 | 16.922 | 13 |  | 22
| 1 | 6 | 18 | 16 | 17 | 18 | 15.904 | 21 | 13.1 | 16 | 16 | 16 | 16 | 16 | 15 | 17 | 17 | 14.7 | 15 | 15 | 33.09 | 21 | 43.6182 | 16 | 19.4471 | 17.052 | 14
| Mean |  | 17.3333 | 15.6667 | 15.3333 | 18.1667 | 15.8721 | 20.1667 | 12.8833 | 16.2667 | 14.6667 | 16 | 16.1667 | 16.8333 | 15.3333 | 15 | 17 | 14.85 | 14.6667 | 14.5 | 28.8767 | 19.6667 | 51.8535 | 16.8333 | 16.7425 | 17.0187 | 13
| Median |  | 17 | 15.5 | 15 | 18 | 15.8803 | 21 | 12.9 | 16.15 | 14.5 | 16 | 16 | 17 | 15 | 15 | 17 | 14.8 | 15 | 14.5 | 28.225 | 19.5 | 51.509 | 17 | 16.3155 | 16.963 | 13
| Std Dev. |  | 0.5164 | 0.8165 | 1.0328 | 0.75277 | 0.14312 | 3.1885 | 0.53448 | 0.55377 | 0.8165 | 1.0954 | 0.75277 | 0.75277 | 0.5164 | 1.2649 | 0.63246 | 0.16432 | 1.0328 | 0.54772 | 2.3276 | 1.633 | 6.6693 | 0.40825 | 2.4591 | 0.68651 | 1.4142
| Rel.Std.Dev. |  | 2.979% | 5.212% | 6.736% | 4.144% | 0.9017% | 15.81% | 4.149% | 3.404% | 5.567% | 6.847% | 4.656% | 4.472% | 3.368% | 8.433% | 3.72% | 1.107% | 7.042% | 3.777% | 8.06% | 8.303% | 12.86% | 2.425% | 14.69% | 4.034% | 10.88%
| PDM3 |  | 9.299% | -1.211% | -3.313% | 14.55% | 0.08448% | 27.16% | -18.76% | 2.573% | -7.516% | 0.8911% | 1.942% | 6.146% | -3.313% | -5.415% | 7.197% | -6.36% | -7.516% | -8.567% | 82.09% | 24.01% | 227% | 6.146% | 5.573% | 7.315% | -18.03%
| Z-Score (Absolute) |  | 0.67 | 0.29 | 0.48 | 1.16 | 0.17 | 2.31 | 1.9 | 0.06 | 0.87 | 0.1 | 0 | 0.39 | 0.48 | 0.67 | 0.48 | 0.76 | 0.87 | 0.96 | 7.35 | 2.02 | 20.63 | 0.39 | 0.33 | 0.49 | 1.83
| Analytical results for Cs in OREAS 609c (Certified Value 5.86 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 5.73 | 5.5 | 5.96 | 5.7668 | 5.87 | 5.88 | 6.0202 | 5.55 | 6.1 | 6.18 | 6.17 | 5.75 | 5.31 | 5.72 | 5.91 | 5.3 | 5.08 | 6.12 | 6.08 | 6.1405 | 5.06 | [5 empty] | 1
| 1 | 2 | 5.63 | 5.2 | 5.99 | 5.8025 | 5.77 | 6.06 | 6.1095 | 5.62 | 6 | 6.11 | 5.9 | 5.85 | 5.22 | 5.91 | 6.02 | 5.4 | 4.87 | 6.2 | 6.18 | 5.9035 | 4.94 | [5 empty] | 32
| 1 | 3 | 5.58 | 5.1 | 6.04 | 5.8178 | 6.04 | 6.24 | 6.1598 | 5.65 | 6 | 6.21 | 6.2 | 5.88 | 5.21 | 5.78 | 5.98 | 5.3 | 4.88 | 6.25 | 5.99 | 6.2467 | 4.52 | [5 empty] | 16
| 1 | 4 | 5.58 | 5.4 | 5.92 | 6.0081 | 5.82 | 6.22 | 6.1175 | 5.65 | 6.2 | 6.13 | 5.9 | 5.77 | 5.49 | 5.94 | 6.08 | 5.2 | 4.87 | 6.44 | 5.97 | 5.9433 | 5.43 | [5 empty] | 5.8608
| 1 | 5 | 5.31 | 5.2 | 5.96 | 6.0596 | 5.9 | 6.1 | 6.0123 | 5.68 | 6 | 6.25 | 6.02 | 5.82 | 5.49 | 5.79 | 6.01 | 5.3 | 5.14 | 6.49 | 6.06 | 6.0067 | 5.14 | [5 empty] | 23
| 1 | 6 | 5.38 | 5.3 | 6.07 | 5.7894 | 5.95 | 6.2 | 6.0418 | 5.76 | 6.1 | 6.01 | 5.68 | 5.82 | 5.48 | 5.73 | 5.96 | 5.3 | 4.74 | 6.25 | 6.18 | 6.2886 | 4.7
| Mean |  | 5.535 | 5.2833 | 5.99 | 5.8741 | 5.8917 | 6.1167 | 6.0769 | 5.6517 | 6.0667 | 6.1483 | 5.9783 | 5.815 | 5.3667 | 5.8117 | 5.9933 | 5.3 | 4.93 | 6.2917 | 6.0767 | 6.0882 | 4.965
| Median |  | 5.58 | 5.25 | 5.975 | 5.8101 | 5.885 | 6.15 | 6.0757 | 5.65 | 6.05 | 6.155 | 5.96 | 5.82 | 5.395 | 5.785 | 5.995 | 5.3 | 4.875 | 6.25 | 6.07 | 6.0736 | 5
| Std Dev. |  | 0.15859 | 0.1472 | 0.055857 | 0.12598 | 0.095795 | 0.13589 | 0.060338 | 0.069113 | 0.08165 | 0.084951 | 0.19436 | 0.048477 | 0.13604 | 0.092394 | 0.05785 | 0.063246 | 0.14993 | 0.14331 | 0.090037 | 0.16109 | 0.32396
| Rel.Std.Dev. |  | 2.865% | 2.786% | 0.9325% | 2.145% | 1.626% | 2.222% | 0.9929% | 1.223% | 1.346% | 1.382% | 3.251% | 0.8337% | 2.535% | 1.59% | 0.9652% | 1.193% | 3.041% | 2.278% | 1.482% | 2.646% | 6.525%
| PDM3 |  | -5.559% | -9.854% | 2.204% | 0.2256% | 0.5261% | 4.365% | 3.686% | -3.569% | 3.512% | 4.905% | 2.005% | -0.782% | -8.432% | -0.8389% | 2.261% | -9.569% | -15.88% | 7.351% | 3.683% | 3.88% | -15.29%
| Z-Score (Absolute) |  | 1.22 | 2.09 | 0.34 | 0.06 | 0 | 0.77 | 0.64 | 0.82 | 0.6 | 0.88 | 0.3 | 0.26 | 1.8 | 0.27 | 0.35 | 2.03 | 3.3 | 1.37 | 0.63 | 0.67 | 3.18
| Analytical results for Cu in OREAS 609c (Certified Value 0.478 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.4885 | 0.49734 | 0.46995 | 0.48437 | 0.4733 | 0.445 | 0.482 | 0.4711 | 0.4946 | 0.481 | 0.4611 | 0.477 | 0.482 | 0.45 | 0.473 | 0.483 | 0.47426 | 0.4933 | 0.475 | 0.4772 | 0.4452 | 0.44638 | 0.495 | 0.47832 | 0.47426 | 0.4427 | 1
| 1 | 2 | 0.5034 | 0.47439 | 0.46883 | 0.48845 | 0.47272 | 0.448 | 0.483 | 0.47013 | 0.4947 | 0.485 | 0.4645 | 0.476 | 0.478 | 0.442 | 0.487 | 0.48 | 0.47469 | 0.4922 | 0.476 | 0.4882 | 0.444 | 0.44006 | 0.487 | 0.48886 | 0.46805 | 0.4373 | 19
| 1 | 3 | 0.4964 | 0.47365 | 0.47594 | 0.48871 | 0.47837 | 0.463 | 0.494 | 0.47417 | 0.4953 | 0.487 | 0.4717 | 0.467 | 0.476 | 0.458 | 0.463 | 0.489 | 0.46975 | 0.4863 | 0.469 | 0.4892 | 0.4401 | 0.43681 | 0.486 | 0.48299 | 0.44899 | 0.4473 | 16
| 1 | 4 | 0.4919 | 0.47722 | 0.46366 | 0.48751 | 0.47746 | 0.452 | 0.497 | 0.47461 | 0.4948 | 0.485 | 0.4852 | 0.461 | 0.468 | 0.454 | 0.483 | 0.49 | 0.45802 | 0.4931 | 0.498 | 0.4888 | 0.4462 | 0.42493 | 0.518 | 0.47348 | 0.46967 | 0.4637 | 0.47796
| 1 | 5 | 0.4932 | 0.48912 | 0.47133 | 0.48825 | 0.47903 | 0.459 | 0.481 | 0.47108 | 0.4962 | 0.487 | 0.459 | 0.487 | 0.478 | 0.459 | 0.472 | 0.495 | 0.46832 | 0.5065 | 0.499 | 0.4758 | 0.4364 | 0.45065 | 0.505 | 0.47816 | 0.47366 | 0.4386 | 24
| 1 | 6 | 0.4906 | 0.49272 | 0.47424 | 0.48301 | 0.47388 | 0.469 | 0.503 | 0.47418 | 0.4973 | 0.493 | 0.4786 | 0.462 | 0.479 | 0.452 | 0.471 | 0.483 | 0.47181 | 0.4885 | 0.486 | 0.4802 | 0.4439 | 0.44613 | 0.481 | 0.48978 | 0.48075 | 0.4495
| Mean |  | 0.494 | 0.48407 | 0.47066 | 0.48672 | 0.47579 | 0.456 | 0.49 | 0.47254 | 0.49548 | 0.48633 | 0.47002 | 0.47167 | 0.47683 | 0.4525 | 0.47483 | 0.48667 | 0.46947 | 0.49332 | 0.48383 | 0.48323 | 0.44263 | 0.44083 | 0.49533 | 0.48193 | 0.46923 | 0.44652
| Median |  | 0.49255 | 0.48317 | 0.47064 | 0.48788 | 0.47567 | 0.4555 | 0.4885 | 0.47263 | 0.49505 | 0.486 | 0.4681 | 0.4715 | 0.478 | 0.453 | 0.4725 | 0.486 | 0.47078 | 0.49265 | 0.481 | 0.4842 | 0.44395 | 0.4431 | 0.491 | 0.48066 | 0.47167 | 0.445
| Std Dev. |  | 0.00531 | 0.010253 | 0.0043358 | 0.0024168 | 0.0028007 | 0.009252 | 0.009252 | 0.001983 | 0.0010685 | 0.0039328 | 0.010376 | 0.010113 | 0.0047504 | 0.0061887 | 0.0087274 | 0.0056095 | 0.0061347 | 0.007036 | 0.012608 | 0.0061986 | 0.0036893 | 0.0092165 | 0.013924 | 0.0064709 | 0.010852 | 0.009664
| Rel.Std.Dev. |  | 1.075% | 2.118% | 0.9212% | 0.4966% | 0.5886% | 2.029% | 1.888% | 0.4196% | 0.2156% | 0.8087% | 2.208% | 2.144% | 0.9962% | 1.368% | 1.838% | 1.153% | 1.307% | 1.426% | 2.606% | 1.283% | 0.8335% | 2.091% | 2.811% | 1.343% | 2.313% | 2.164%
| PDM3 |  | 3.356% | 1.279% | -1.528% | 1.832% | -0.4537% | -4.594% | 2.519% | -1.133% | 3.666% | 1.752% | -1.662% | -1.317% | -0.2357% | -5.327% | -0.6541% | 1.822% | -1.775% | 3.213% | 1.229% | 1.103% | -7.391% | -7.769% | 3.635% | 0.8308% | -1.826% | -6.579%
| Z-Score (Absolute) |  | 1.34 | 0.59 | 0.43 | 0.79 | 0.04 | 1.54 | 1.04 | 0.29 | 1.45 | 0.76 | 0.48 | 0.35 | 0.04 | 1.81 | 0.11 | 0.79 | 0.52 | 1.29 | 0.57 | 0.52 | 2.55 | 2.69 | 1.44 | 0.43 | 0.54 | 2.26
| Analytical results for Dy in OREAS 609c (Certified Value 3.27 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 3.3; 2.03: 3.29; 2.04: 3.166; 2.05: 3.3; 2.06: 3.3485; 2.07: 3.3; 2.08: 3.27; 2.09: 3.29; 2.10: 2.9055; 2.11: 3.2; 2.12: 3.21 | 1
| 1 | 2 | 2.01: 3.2; 2.03: 3.33; 2.04: 3.1887; 2.05: 3.3; 2.06: 3.3738; 2.07: 3.34; 2.08: 3.13; 2.09: 3.39; 2.10: 3.2931; 2.11: 3.2; 2.12: 3.19 | 34
| 1 | 3 | 2.01: 3.1; 2.03: 3.3; 2.04: 3.1225; 2.05: 3.4; 2.06: 3.4203; 2.07: 3.43; 2.08: 3.08; 2.09: 3.29; 2.10: 3.1718; 2.11: 3.2; 2.12: 3.33 | 16
| 1 | 4 | 2.01: 3.3; 2.03: 3.26; 2.04: 3.128; 2.05: 3.3; 2.06: 3.4633; 2.07: 3.42; 2.08: 3.25; 2.09: 3.2; 2.10: 3.1553; 2.11: 3.4; 2.12: 3.39 | 3.2707
| 1 | 5 | 2.01: 3.2; 2.03: 3.25; 2.04: 3.2731; 2.05: 3.3; 2.06: 3.3317; 2.07: 3.39; 2.08: 3.32; 2.09: 3.25; 2.10: 3.0971; 2.11: 3.3; 2.12: 3.06 | 25
| 1 | 6 | 2.01: 3.3; 2.03: 3.19; 2.04: 3.2777; 2.05: 3.3; 2.06: 3.4107; 2.07: 3.29; 2.08: 3.01; 2.09: 3.37; 2.10: 3.432; 2.11: 3.2; 2.12: 3.36
| Mean |  | 2.01: 3.2333; 2.03: 3.27; 2.04: 3.1927; 2.05: 3.3167; 2.06: 3.3914; 2.07: 3.3617; 2.08: 3.1767; 2.09: 3.2983; 2.10: 3.1758; 2.11: 3.25; 2.12: 3.2567
| Median |  | 2.01: 3.25; 2.03: 3.275; 2.04: 3.1773; 2.05: 3.3; 2.06: 3.3923; 2.07: 3.365; 2.08: 3.19; 2.09: 3.29; 2.10: 3.1635; 2.11: 3.2; 2.12: 3.27
| Std Dev. |  | 2.01: 0.08165; 2.03: 0.04858; 2.04: 0.068578; 2.05: 0.040825; 2.06: 0.049176; 2.07: 0.06047; 2.08: 0.1216; 2.09: 0.071671; 2.10: 0.17838; 2.11: 0.083666; 2.12: 0.1258
| Rel.Std.Dev. |  | 2.01: 2.525%; 2.03: 1.486%; 2.04: 2.148%; 2.05: 1.231%; 2.06: 1.45%; 2.07: 1.799%; 2.08: 3.828%; 2.09: 2.173%; 2.10: 5.617%; 2.11: 2.574%; 2.12: 3.863%
| PDM3 |  | 2.01: -1.141%; 2.03: -0.02012%; 2.04: -2.385%; 2.05: 1.407%; 2.06: 3.691%; 2.07: 2.783%; 2.08: -2.874%; 2.09: 0.8462%; 2.10: -2.9%; 2.11: -0.6316%; 2.12: -0.4278%
| Z-Score (Absolute) |  | 2.01: 0.26; 2.03: 0.15; 2.04: 0.72; 2.05: 0.67; 2.06: 1.51; 2.07: 1.18; 2.08: 0.9; 2.09: 0.47; 2.10: 0.91; 2.11: 0.07; 2.12: 0
| Analytical results for Er in OREAS 609c (Certified Value 0.99 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 1; 2.03: 0.96; 2.04: 0.93012; 2.05: 0.9; 2.06: 1.0827; 2.07: 0.99; 2.08: 0.96; 2.09: 0.99; 2.10: 0.9681; 2.11: < 0.2; 2.12: 0.95 | 1
| 1 | 2 | 2.01: 0.9; 2.03: 0.96; 2.04: 0.92227; 2.05: 0.9; 2.06: 1.0565; 2.07: 0.96; 2.08: 0.95; 2.09: 1.02; 2.10: 1.0338; 2.11: < 0.2; 2.12: 0.96 | 35
| 1 | 3 | 2.01: 0.9; 2.03: 0.96; 2.04: 0.92589; 2.05: 1; 2.06: 1.0529; 2.07: 0.98; 2.08: 0.95; 2.09: 1; 2.10: 1.03; 2.11: < 0.2; 2.12: 0.97 | 16
| 1 | 4 | 2.01: 0.9; 2.03: 0.93; 2.04: 0.96659; 2.05: 0.9; 2.06: 1.0593; 2.07: 0.99; 2.08: 0.97; 2.09: 0.99; 2.10: 1.0142; 2.11: < 0.2; 2.12: 1.01 | 0.98526
| 1 | 5 | 2.01: 0.9; 2.03: 0.94; 2.04: 0.96818; 2.05: 0.9; 2.06: 1.0818; 2.07: 0.97; 2.08: 0.99; 2.09: 1.01; 2.10: 1.0113; 2.11: < 0.2; 2.12: 0.91 | 26
| 1 | 6 | 2.01: 0.9; 2.03: 0.98; 2.04: 0.96385; 2.05: 0.9; 2.06: 1.0701; 2.07: 0.97; 2.08: 0.89; 2.09: 1.01; 2.10: 1.0047; 2.11: < 0.2; 2.12: 1.03
| Mean |  | 2.01: 0.91667; 2.03: 0.955; 2.04: 0.94615; 2.05: 0.91667; 2.06: 1.0672; 2.07: 0.97667; 2.08: 0.95167; 2.09: 1.0033; 2.10: 1.0104; 2.12: 0.97167
| Median |  | 2.01: 0.9; 2.03: 0.96; 2.04: 0.94698; 2.05: 0.9; 2.06: 1.0647; 2.07: 0.975; 2.08: 0.955; 2.09: 1.005; 2.10: 1.0128; 2.12: 0.965
| Std Dev. |  | 2.01: 0.040825; 2.03: 0.017607; 2.04: 0.022155; 2.05: 0.040825; 2.06: 0.012986; 2.07: 0.012111; 2.08: 0.033714; 2.09: 0.012111; 2.10: 0.023532; 2.12: 0.043089
| Rel.Std.Dev. |  | 2.01: 4.454%; 2.03: 1.844%; 2.04: 2.342%; 2.05: 4.454%; 2.06: 1.217%; 2.07: 1.24%; 2.08: 3.543%; 2.09: 1.207%; 2.10: 2.329%; 2.12: 4.435%
| PDM3 |  | 2.01: -6.962%; 2.03: -3.071%; 2.04: -3.969%; 2.05: -6.962%; 2.06: 8.319%; 2.07: -0.8718%; 2.08: -3.409%; 2.09: 1.835%; 2.10: 2.547%; 2.12: -1.379%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.44; 2.04: 0.67; 2.05: NA; 2.06: 2.53; 2.07: 0.13; 2.08: 0.53; 2.09: 0.84; 2.10: 1.02; 2.11: 23.04; 2.12: 0
Results from laboratories 2.03 and 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 609c (Certified Value 1.16 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 1.2; 2.03: 1.16; 2.04: 1.1487; 2.05: 1.18; 2.06: 2.3314; 2.07: 1.19; 2.08: 1.23; 2.09: 1.46; 2.10: 1.2288; 2.11: 1.2; 2.12: 1.12 | 1
| 1 | 2 | 2.01: 1.2; 2.03: 1.18; 2.04: 1.1695; 2.05: 1.16; 2.06: 2.3079; 2.07: 1.2; 2.08: 1.11; 2.09: 1.39; 2.10: 1.3643; 2.11: 1.2; 2.12: 1.12 | 36
| 1 | 3 | 2.01: 1.1; 2.03: 1.18; 2.04: 1.1683; 2.05: 1.21; 2.06: 2.4667; 2.07: 1.19; 2.08: 1.08; 2.09: 1.45; 2.10: 1.2643; 2.11: 1.2; 2.12: 1.21 | 16
| 1 | 4 | 2.01: 1.2; 2.03: 1.12; 2.04: 1.1739; 2.05: 1.13; 2.06: 2.4268; 2.07: 1.24; 2.08: 1.12; 2.09: 1.43; 2.10: 1.3771; 2.11: 1.2; 2.12: 1.21 | 1.1643
| 1 | 5 | 2.01: 1.1; 2.03: 1.12; 2.04: 1.1586; 2.05: 1.18; 2.06: 2.4479; 2.07: 1.17; 2.08: 1.14; 2.09: 1.41; 2.10: 1.3312; 2.11: 1.2; 2.12: 1.1 | 27
| 1 | 6 | 2.01: 1.2; 2.03: 1.18; 2.04: 1.1544; 2.05: 1.18; 2.06: 2.4354; 2.07: 1.15; 2.08: 1.12; 2.09: 1.44; 2.10: 1.321; 2.11: 1.2; 2.12: 1.15
| Mean |  | 2.01: 1.1667; 2.03: 1.1567; 2.04: 1.1622; 2.05: 1.1733; 2.06: 2.4027; 2.07: 1.19; 2.08: 1.1333; 2.09: 1.43; 2.10: 1.3144; 2.11: 1.2; 2.12: 1.1517
| Median |  | 2.01: 1.2; 2.03: 1.17; 2.04: 1.1634; 2.05: 1.18; 2.06: 2.4311; 2.07: 1.19; 2.08: 1.12; 2.09: 1.435; 2.10: 1.3261; 2.11: 1.2; 2.12: 1.135
| Std Dev. |  | 2.01: 0.05164; 2.03: 0.029439; 2.04: 0.0098494; 2.05: 0.026583; 2.06: 0.066121; 2.07: 0.030332; 2.08: 0.051251; 2.09: 0.026077; 2.10: 0.057592; 2.11: 0; 2.12: 0.047924
| Rel.Std.Dev. |  | 2.01: 4.426%; 2.03: 2.545%; 2.04: 0.8475%; 2.05: 2.266%; 2.06: 2.752%; 2.07: 2.549%; 2.08: 4.522%; 2.09: 1.824%; 2.10: 4.381%; 2.11: 0%; 2.12: 4.161%
| PDM3 |  | 2.01: 0.2016%; 2.03: -0.6573%; 2.04: -0.1798%; 2.05: 0.7741%; 2.06: 106.4%; 2.07: 2.206%; 2.08: -2.661%; 2.09: 22.82%; 2.10: 12.89%; 2.11: 3.064%; 2.12: -1.087%
| Z-Score (Absolute) |  | 2.01: 0.21; 2.03: 0.52; 2.04: 0.35; 2.05: 0; 2.06: 38.26; 2.07: 0.52; 2.08: 1.24; 2.09: 7.99; 2.10: 4.39; 2.11: 0.83; 2.12: 0.67
| Analytical results for Fe in OREAS 609c (Certified Value 3 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.07 | 2.88 | 3.06 | 2.92 | 3.0502 | 3.09 | 2.99 | 3.0135 | 3.22 | 2.97 | 2.96 | 2.93 | 3.04 | 2.78 | 2.89 | 3.01 | 2.992 | 2.95 | 2.99 | 3.02 | 2.96 | 3.0631 | 2.875 | 3.0295 | 2.9487 | 3.05 | 1
| 1 | 2 | 3.06 | 3.05 | 2.93 | 2.97 | 3.0661 | 3.08 | 3.15 | 3.022 | 3.15 | 2.92 | 2.99 | 2.96 | 3.02 | 2.79 | 2.96 | 3.03 | 3.0025 | 2.87 | 2.99 | 3.09 | 3.13 | 3.0331 | 2.826 | 2.9093 | 3.0893 | 3.1 | 0
| 1 | 3 | 3 | 2.84 | 3.02 | 2.99 | 3.0591 | 3.05 | 3.22 | 3.0333 | 3.13 | 2.9 | 3.01 | 2.93 | 3.04 | 2.87 | 2.85 | 4.01 | 2.9454 | 2.84 | 2.98 | 3.06 | 3.1 | 2.9893 | 2.812 | 3.0478 | 2.8672 | 3.1 | 16
| 1 | 4 | 3.08 | 3.04 | 3.07 | 2.93 | 3.0809 | 3.1 | 3.15 | 3.0015 | 3.12 | 2.99 | 3.14 | 2.86 | 2.99 | 2.85 | 2.96 | 3.08 | 2.9161 | 2.93 | 3.21 | 3.06 | 3.04 | 2.9999 | 3.008 | 2.8284 | 2.9827 | 3.25 | 3.0042
| 1 | 5 | 3.14 | 2.98 | 3 | 3.03 | 3.0946 | 3.13 | 3.08 | 3.0368 | 3.06 | 3.03 | 2.94 | 3.01 | 3.03 | 2.86 | 2.92 | 3.09 | 2.95 | 3.02 | 3.15 | 3.01 | 3.11 | 3.1273 | 2.945 | 2.8811 | 2.8072 | 3.07 | 28
| 1 | 6 | 3.07 | 2.96 | 3.05 | 2.92 | 3.0258 | 3.22 | 3.23 | 3.0894 | 3.04 | 2.89 | 3.07 | 2.88 | 3 | 2.82 | 2.86 | 3.05 | 2.9379 | 2.84 | 3.08 | 3.01 | 3.11 | 3.0317 | 2.798 | 2.9538 | 2.9184 | 3.11
| Mean |  | 3.07 | 2.9583 | 3.0217 | 2.96 | 3.0628 | 3.1117 | 3.1367 | 3.0328 | 3.12 | 2.95 | 3.0183 | 2.9283 | 3.02 | 2.8283 | 2.9067 | 3.2117 | 2.9573 | 2.9083 | 3.0667 | 3.0417 | 3.075 | 3.0407 | 2.8773 | 2.9416 | 2.9356 | 3.1133
| Median |  | 3.07 | 2.97 | 3.035 | 2.95 | 3.0626 | 3.095 | 3.15 | 3.0277 | 3.125 | 2.945 | 3 | 2.93 | 3.025 | 2.835 | 2.905 | 3.065 | 2.9477 | 2.9 | 3.035 | 3.04 | 3.105 | 3.0324 | 2.8505 | 2.9315 | 2.9335 | 3.1
| Std Dev. |  | 0.044721 | 0.084479 | 0.051929 | 0.044721 | 0.024061 | 0.059133 | 0.090259 | 0.030641 | 0.064807 | 0.055498 | 0.074677 | 0.054191 | 0.020976 | 0.037639 | 0.048028 | 0.39225 | 0.033214 | 0.071391 | 0.096885 | 0.033116 | 0.064109 | 0.049893 | 0.083493 | 0.085647 | 0.097451 | 0.070616
| Rel.Std.Dev. |  | 1.457% | 2.856% | 1.719% | 1.511% | 0.7856% | 1.9% | 2.878% | 1.01% | 2.077% | 1.881% | 2.474% | 1.851% | 0.6946% | 1.331% | 1.652% | 12.21% | 1.123% | 2.455% | 3.159% | 1.089% | 2.085% | 1.641% | 2.902% | 2.912% | 3.32% | 2.268%
| PDM3 |  | 2.191% | -1.526% | 0.5818% | -1.471% | 1.951% | 3.578% | 4.41% | 0.9509% | 3.855% | -1.804% | 0.4709% | -2.525% | 0.5264% | -5.854% | -3.246% | 6.906% | -1.56% | -3.191% | 2.08% | 1.248% | 2.357% | 1.217% | -4.223% | -2.082% | -2.284% | 3.633%
| Z-Score (Absolute) |  | 0.49 | 0.63 | 0.01 | 0.61 | 0.42 | 0.91 | 1.16 | 0.12 | 1 | 0.71 | 0.03 | 0.93 | 0.01 | 1.93 | 1.15 | 1.92 | 0.64 | 1.13 | 0.46 | 0.21 | 0.54 | 0.2 | 1.44 | 0.79 | 0.86 | 0.93
| Analytical results for Ga in OREAS 609c (Certified Value 23.2 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 23.77 | 21.36 | 22.44 | 22.5679 | 22.8 | 22.8 | 25.551 | 23.52 | 23.4 | 24.66 | 22.9 | 22.3 | 22.6 | 21.9 | 22.2 | 21.78 | 25.54 | 31 | 57.04 | 22.9839 | 17 | 23.8424 | 20.651 | 26.21 |  |  | 1
| 1 | 2 | 23.77 | 21.6 | 23.22 | 23.3079 | 23.1 | 23.6 | 25.632 | 22.83 | 23.2 | 24.59 | 22.3 | 22.8 | 20.8 | 22.6 | 22.5 | 21.61 | 25.67 | 30 | 49.65 | 24.4681 | 18 | 22.8162 | 20.126 | 25.17 |  |  | 37
| 1 | 3 | 23.07 | 21.2 | 23.19 | 22.96 | 23.9 | 24.3 | 25.803 | 22.98 | 23.5 | 24.82 | 23 | 22.7 | 20.9 | 21.9 | 22.5 | 21.33 | 24.97 | 28 | 50.69 | 23.6408 | 18 | 24.4274 | 19.44 | 26.18 |  |  | 16
| 1 | 4 | 23.55 | 21.84 | 22.38 | 23.3732 | 22.6 | 23.8 | 25.307 | 22.99 | 23.5 | 25.16 | 21.7 | 22.3 | 21.8 | 22.5 | 23 | 20.89 | 25.49 | 30 | 57.59 | 23.8189 | 18 | 23.0426 | 20.745 | 26.04 |  |  | 23.1779
| 1 | 5 | 22.39 | 21.12 | 22.87 | 23.4686 | 23.3 | 23.7 | 25.717 | 22.48 | 23.6 | 24.21 | 22.7 | 22.7 | 22.2 | 22.2 | 22.9 | 21.32 | 26.17 | 31 | 46.73 | 23.9098 | 18 | 23.5025 | 20.129 | 25.19 |  |  | 29
| 1 | 6 | 23.98 | 21.52 | 22.72 | 23.1851 | 23.8 | 24.3 | 25.862 | 22.51 | 23.4 | 24.78 | 21.1 | 22.6 | 21.9 | 21.8 | 22.6 | 21.25 | 25.69 | 30 | 49.33 | 24.4015 | 17 | 24.3167 | 20.226 | 27.44
| Mean |  | 23.4217 | 21.44 | 22.8033 | 23.1438 | 23.25 | 23.75 | 25.6453 | 22.885 | 23.4333 | 24.7033 | 22.2833 | 22.5667 | 21.7 | 22.15 | 22.6167 | 21.3633 | 25.5883 | 30 | 51.8383 | 23.8705 | 17.6667 | 23.658 | 20.2195 | 26.0383
| Median |  | 23.66 | 21.44 | 22.795 | 23.2465 | 23.2 | 23.75 | 25.6745 | 22.905 | 23.45 | 24.72 | 22.5 | 22.65 | 21.85 | 22.05 | 22.55 | 21.325 | 25.605 | 30 | 50.17 | 23.8644 | 18 | 23.6725 | 20.1775 | 26.11
| Std Dev. |  | 0.59311 | 0.26773 | 0.35937 | 0.33234 | 0.5244 | 0.55408 | 0.20024 | 0.38235 | 0.13663 | 0.31181 | 0.74944 | 0.21602 | 0.71554 | 0.33912 | 0.29269 | 0.308 | 0.38722 | 1.0954 | 4.4415 | 0.54445 | 0.5164 | 0.65901 | 0.46625 | 0.83521
| Rel.Std.Dev. |  | 2.532% | 1.249% | 1.576% | 1.436% | 2.256% | 2.333% | 0.7808% | 1.671% | 0.583% | 1.262% | 3.363% | 0.9573% | 3.297% | 1.531% | 1.294% | 1.442% | 1.513% | 3.651% | 8.568% | 2.281% | 2.923% | 2.786% | 2.306% | 3.208%
| PDM3 |  | 1.052% | -7.498% | -1.616% | -0.1473% | 0.3109% | 2.468% | 10.65% | -1.264% | 1.102% | 6.581% | -3.86% | -2.637% | -6.377% | -4.435% | -2.422% | -7.829% | 10.4% | 29.43% | 123.7% | 2.988% | -23.78% | 2.071% | -12.76% | 12.34%
| Z-Score (Absolute) |  | 0.12 | 1.26 | 0.31 | 0.07 | 0 | 0.35 | 1.67 | 0.25 | 0.13 | 1.01 | 0.67 | 0.48 | 1.08 | 0.77 | 0.44 | 1.32 | 1.63 | 4.71 | 19.94 | 0.43 | NA | 0.28 | 2.11 | 1.95
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 609c (Certified Value 5.18 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.09: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.05: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  | 2
3
| 1 | 1 | 2.01: 5.4; 2.03: 5.18; 2.05: 5.0222; 2.06: 5.1; 2.07: 7.9092; 2.08: 5.1; 2.09: 5.6; 2.10: 5.43; 2.11: 4.9822; 2.12: 3; 2.13: 4.6 |  |  | 1
| 1 | 2 | 2.01: 5; 2.03: 5.11; 2.05: 5.0255; 2.06: 5.2; 2.07: 7.8308; 2.08: 5.14; 2.09: 5.3; 2.10: 5.42; 2.11: 5.3965; 2.12: 3; 2.13: 4.6 |  |  | 38
| 1 | 3 | 2.01: 5.1; 2.03: 5.24; 2.05: 5.0083; 2.06: 5.6; 2.07: 7.1301; 2.08: 5.21; 2.09: 5.2; 2.10: 5.29; 2.11: 5.32; 2.12: 3; 2.13: 4.8 |  |  | 16
| 1 | 4 | 2.01: 5.5; 2.03: 5.13; 2.05: 4.9991; 2.06: 5.1; 2.07: 7.8577; 2.08: 5.39; 2.09: 5.4; 2.10: 5.27; 2.11: 5.5563; 2.12: 3; 2.13: 4.9 |  |  | 5.1772
| 1 | 5 | 2.01: 5; 2.03: 5.03; 2.05: 5.2198; 2.06: 5.2; 2.07: 7.6758; 2.08: 5.17; 2.09: 5.6; 2.10: 5.38; 2.11: 5.3082; 2.12: 3; 2.13: 4.5 |  |  | 30
| 1 | 6 | 2.01: 5.3; 2.03: 5.12; 2.05: 5.0258; 2.06: 5.3; 2.07: 7.6855; 2.08: 5.11; 2.09: 5; 2.10: 5.3; 2.11: 5.3926; 2.12: 3; 2.13: 5
| Mean |  | 2.01: 5.2167; 2.03: 5.135; 2.05: 5.0501; 2.06: 5.25; 2.07: 7.6815; 2.08: 5.1867; 2.09: 5.35; 2.10: 5.3483; 2.11: 5.326; 2.12: 3; 2.13: 4.7333
| Median |  | 2.01: 5.2; 2.03: 5.125; 2.05: 5.0238; 2.06: 5.2; 2.07: 7.7581; 2.08: 5.155; 2.09: 5.35; 2.10: 5.34; 2.11: 5.3563; 2.12: 3; 2.13: 4.7
| Std Dev. |  | 2.01: 0.2137; 2.03: 0.07064; 2.05: 0.083821; 2.06: 0.18708; 2.07: 0.2861; 2.08: 0.10746; 2.09: 0.23452; 2.10: 0.070261; 2.11: 0.19026; 2.12: 0; 2.13: 0.19664
| Rel.Std.Dev. |  | 2.01: 4.096%; 2.03: 1.376%; 2.05: 1.66%; 2.06: 3.563%; 2.07: 3.725%; 2.08: 2.072%; 2.09: 4.384%; 2.10: 1.314%; 2.11: 3.572%; 2.12: 0%; 2.13: 4.154%
| PDM3 |  | 2.01: 0.7623%; 2.03: -0.8152%; 2.05: -2.455%; 2.06: 1.406%; 2.07: 48.37%; 2.08: 0.1828%; 2.09: 3.338%; 2.10: 3.305%; 2.11: 2.873%; 2.12: -42.05%; 2.13: -8.574%
| Z-Score (Absolute) |  | 2.01: 0.11; 2.03: 0.62; 2.05: 1.16; 2.06: 0.11; 2.07: 15.48; 2.08: 0.3; 2.09: 0.74; 2.10: 0.73; 2.11: 0.59; 2.12: NA; 2.13: 3.16
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 609c (Certified Value 0.25 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 0.27 | 0.44 | 2.3 | < 0.2 | 0.1 | 0.3 | 0.2697 | 2.4 | 2.3 | 0.18 | 0.19 | 0.17 | 0.09 | 0.18 | 0.26 | < 0.2 | 0.35 | 0.15 | 0.44278 | 1 | [6 empty] | 1
| 1 | 2 | 0.29 | 0.26 | 2.5 | < 0.2 | 0.2 | 0.29 | 0.261 | 2.4 | 2.4 | 0.17 | 0.23 | 0.16 | 0.1 | 0.2 | 0.48 | < 0.2 | 0.33 | 0.14 | 0.42457 | 1 | [6 empty] | 39
| 1 | 3 | 0.27 | 0.19 | 2.4 | < 0.2 | 0.2 | 0.31 | 0.3151 | 2.3 | 2.3 | 0.18 | 0.18 | 0.15 | 0.1 | 0.17 | 0.12 | < 0.2 | 0.25 | 0.13 | 0.41738 | 1.6 | [6 empty] | 16
| 1 | 4 | 0.31 | 0.24 | 2.3 | < 0.2 | 0.1 | 0.32 | 0.2909 | 2.3 | 2.4 | 0.18 | 0.22 | 0.19 | 0.12 | 0.2 | 0.27 | < 0.2 | 0.35 | 0.16 | 0.43258 | 1.2 | [6 empty] | 0.24655
| 1 | 5 | 0.28 | 0.25 | 2.6 | < 0.2 | 0.2 | 0.3 | 0.2619 | 2.3 | 2.5 | 0.17 | 0.24 | 0.18 | 0.11 | 0.18 | 0.5 | < 0.2 | 0.35 | 0.14 | 0.44312 | 1.3 | [6 empty] | 31
| 1 | 6 | 0.28 | 0.27 | 2.5 | < 0.2 | 0.2 | 0.45 | 0.288 | 2.4 | 2.3 | 0.16 | 0.19 | 0.2 | 0.11 | 0.19 | 0.37 | < 0.2 | 0.31 | 0.14 | 0.4301 | 1.6
| Mean |  | 0.28333 | 0.275 | 2.4333 |  | 0.16667 | 0.32833 | 0.2811 | 2.35 | 2.3667 | 0.17333 | 0.20833 | 0.175 | 0.105 | 0.18667 | 0.33333 |  | 0.32333 | 0.14333 | 0.43176 | 1.2833
| Median |  | 0.28 | 0.255 | 2.45 |  | 0.2 | 0.305 | 0.27885 | 2.35 | 2.35 | 0.175 | 0.205 | 0.175 | 0.105 | 0.185 | 0.32 |  | 0.34 | 0.14 | 0.43134 | 1.25
| Std Dev. |  | 0.015055 | 0.085499 | 0.12111 |  | 0.05164 | 0.06047 | 0.020987 | 0.054772 | 0.08165 | 0.008165 | 0.024833 | 0.018708 | 0.010488 | 0.012111 | 0.14528 |  | 0.039328 | 0.010328 | 0.010123 | 0.27142
| Rel.Std.Dev. |  | 5.314% | 31.09% | 4.977% |  | 30.98% | 18.42% | 7.466% | 2.331% | 3.45% | 4.711% | 11.92% | 10.69% | 9.989% | 6.488% | 43.58% |  | 12.16% | 7.206% | 2.345% | 21.15%
| PDM3 |  | 14.92% | 11.54% | 886.9% |  | -32.4% | 33.17% | 14.01% | 853.1% | 859.9% | -29.7% | -15.5% | -29.02% | -57.41% | -24.29% | 35.2% |  | 31.14% | -41.87% | 75.12% | 420.5%
| Z-Score (Absolute) |  | 0.67 | 0.6 | NA | 0.97 | NA | 1.08 | 0.65 | NA | NA | 0.31 | 0 | 0.3 | 0.93 | 0.19 | 1.12 | 0.97 | 1.03 | 0.58 | 2.01 | NA
Results from laboratories 2.05, 2.07, 2.11, 2.12 and 2.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 609c (Certified Value 6.27 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 5.3 | 6.38 | 6.1 | 6.17 | 6.4368 | 6.3 | 6.24 | 6.16 | 6 | 6.19 | 6.6 | 6.3 | 5.8 | 6 | 6.3 | 6.01 | 7.71 | 6.5 | 6.7 | 9 | 6.5784 | 6.11 |  |  |  |  | 1
| 1 | 2 | 4.4 | 6.26 | 6 | 6.3 | 6.3123 | 6.6 | 6.35 | 5.95 | 6 | 6.16 | 6.5 | 6.3 | 5.9 | 6.1 | 6.5 | 6 | 7.75 | 6.3 | 6.6 | 11 | 6.3017 | 6.1 |  |  |  |  | 23
| 1 | 3 | 5.1 | 5.93 | 6.3 | 6.11 | 6.4378 | 6.6 | 6.62 | 5.81 | 6 | 6.14 | 6.4 | 6.3 | 6.1 | 6 | 6.5 | 5.8 | 7.6 | 6.3 | 6.4 | 9 | 6.1708 | 6.23 |  |  |  |  | 16
| 1 | 4 | 4.9 | 6.44 | 6 | 6.25 | 6.4284 | 6.5 | 6.61 | 6.06 | 6.1 | 6.33 | 6.2 | 6.2 | 6.2 | 6.2 | 6.6 | 5.78 | 7.69 | 6.4 | 6.7 | 10 | 6.1786 | 6.67 |  |  |  |  | 6.2657
| 1 | 5 | 4.8 | 5.97 | 6.2 | 6.17 | 6.6265 | 6.3 | 6.5 | 5.89 | 6.2 | 6.11 | 6.8 | 6.3 | 6.3 | 6 | 6.6 | 5.82 | 7.78 | 6.7 | 6.5 | 10 | 6.2338 | 6.19 |  |  |  |  | 32
| 1 | 6 | 5.1 | 6.44 | 6.1 | 6.27 | 6.4301 | 6.5 | 6.6 | 6.15 | 5.9 | 6.06 | 6.1 | 6.2 | 6.3 | 6 | 6.5 | 5.83 | 7.6 | 6.5 | 6.6 | 10 | 6.5813 | 6.63
| Mean |  | 4.9333 | 6.2367 | 6.1167 | 6.2117 | 6.4453 | 6.4667 | 6.4867 | 6.0033 | 6.0333 | 6.165 | 6.4333 | 6.2667 | 6.1 | 6.05 | 6.5 | 5.8733 | 7.6883 | 6.45 | 6.5833 | 9.8333 | 6.3408 | 6.3217
| Median |  | 5 | 6.32 | 6.1 | 6.21 | 6.4335 | 6.5 | 6.55 | 6.005 | 6 | 6.15 | 6.45 | 6.3 | 6.15 | 6 | 6.5 | 5.825 | 7.7 | 6.45 | 6.6 | 10 | 6.2678 | 6.21
| Std Dev. |  | 0.31411 | 0.23192 | 0.1169 | 0.072778 | 0.10118 | 0.13663 | 0.15845 | 0.14306 | 0.10328 | 0.09225 | 0.2582 | 0.05164 | 0.20976 | 0.083666 | 0.10954 | 0.10347 | 0.075211 | 0.15166 | 0.1169 | 0.75277 | 0.19106 | 0.25926
| Rel.Std.Dev. |  | 6.367% | 3.719% | 1.911% | 1.172% | 1.57% | 2.113% | 2.443% | 2.383% | 1.712% | 1.496% | 4.013% | 0.824% | 3.439% | 1.383% | 1.685% | 1.762% | 0.9782% | 2.351% | 1.776% | 7.655% | 3.013% | 4.101%
| PDM3 |  | -21.26% | -0.4633% | -2.378% | -0.8623% | 2.867% | 3.208% | 3.527% | -4.187% | -3.708% | -1.607% | 2.676% | 0.01553% | -2.644% | -3.442% | 3.74% | -6.262% | 22.71% | 2.942% | 5.07% | 56.94% | 1.198% | 0.8933%
| Z-Score (Absolute) |  | 4.9 | 0.11 | 0.55 | 0.2 | 0.66 | 0.74 | 0.81 | 0.97 | 0.86 | 0.37 | 0.61 | 0 | 0.61 | 0.8 | 0.86 | 1.45 | 5.23 | 0.67 | 1.16 | NA | 0.27 | 0.2
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 609c (Indicative Value 0.12 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab |  |  |  |  | 1
| No | No | 2.01: 2.06; 2.03: 2.18; 2.04: 2.22 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 2.01: < 1; 2.03: 0.12; 2.04: < 2 |  |  |  |  | 1
| 1 | 2 | 2.01: < 1; 2.03: 0.13; 2.04: < 2 |  |  |  |  | 3
| 1 | 3 | 2.01: < 1; 2.03: 0.12; 2.04: < 2 |  |  |  |  | 16
| 1 | 4 | 2.01: < 1; 2.03: 0.11; 2.04: < 2 |  |  |  |  | 0.12
| 1 | 5 | 2.01: < 1; 2.03: 0.12; 2.04: < 2 |  |  |  |  | 9
| 1 | 6 | 2.01: < 1; 2.03: 0.12; 2.04: < 2
| Mean |  | 2.03: 0.12
| Median |  | 2.03: 0.12
| Std Dev. |  | 2.03: 0.0063246
| Rel.Std.Dev. |  | 2.03: 5.27%
PDM3 | bar: 2.03
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.67; 2.04: 0.89
| Analytical results for Ho in OREAS 609c (Certified Value 0.46 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 0.6; 2.03: 0.45; 2.04: 0.43657; 2.05: 0.5; 2.06: 0.4505; 2.07: 0.46; 2.08: 0.48; 2.10: 0.49; 2.11: 0.3631; 2.12: < 1; 2.13: 0.39 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.5; 2.03: 0.46; 2.04: 0.43842; 2.05: 0.5; 2.06: 0.4712; 2.07: 0.47; 2.08: 0.47; 2.10: 0.51; 2.11: 0.3941; 2.12: < 1; 2.13: 0.39 |  |  |  |  | 24
| 1 | 3 | 2.01: 0.5; 2.03: 0.45; 2.04: 0.43983; 2.05: 0.5; 2.06: 0.453; 2.07: 0.46; 2.08: 0.47; 2.10: 0.49; 2.11: 0.3868; 2.12: < 1; 2.13: 0.41 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.6; 2.03: 0.45; 2.04: 0.43205; 2.05: 0.5; 2.06: 0.4754; 2.07: 0.48; 2.08: 0.47; 2.10: 0.5; 2.11: 0.3842; 2.12: < 1; 2.13: 0.42 |  |  |  |  | 0.45736
| 1 | 5 | 2.01: 0.5; 2.03: 0.45; 2.04: 0.44386; 2.05: 0.5; 2.06: 0.5056; 2.07: 0.47; 2.08: 0.5; 2.10: 0.49; 2.11: 0.3837; 2.12: < 1; 2.13: 0.38 |  |  |  |  | 33
| 1 | 6 | 2.01: 0.5; 2.03: 0.46; 2.04: 0.43741; 2.05: 0.5; 2.06: 0.4751; 2.07: 0.46; 2.08: 0.45; 2.10: 0.5; 2.11: 0.3833; 2.12: < 1; 2.13: 0.42
| Mean |  | 2.01: 0.53333; 2.03: 0.45333; 2.04: 0.43802; 2.05: 0.5; 2.06: 0.4718; 2.07: 0.46667; 2.08: 0.47333; 2.10: 0.49667; 2.11: 0.38253; 2.13: 0.40167
| Median |  | 2.01: 0.5; 2.03: 0.45; 2.04: 0.43792; 2.05: 0.5; 2.06: 0.47315; 2.07: 0.465; 2.08: 0.47; 2.10: 0.495; 2.11: 0.38395; 2.13: 0.4
| Std Dev. |  | 2.01: 0.05164; 2.03: 0.005164; 2.04: 0.0038885; 2.05: 0; 2.06: 0.019868; 2.07: 0.008165; 2.08: 0.01633; 2.10: 0.008165; 2.11: 0.010338; 2.13: 0.017224
| Rel.Std.Dev. |  | 2.01: 9.682%; 2.03: 1.139%; 2.04: 0.8877%; 2.05: 0%; 2.06: 4.211%; 2.07: 1.75%; 2.08: 3.45%; 2.10: 1.644%; 2.11: 2.703%; 2.13: 4.288%
| PDM3 |  | 2.01: 16.61%; 2.03: -0.8795%; 2.04: -4.227%; 2.05: 9.324%; 2.06: 3.158%; 2.07: 2.036%; 2.08: 3.493%; 2.10: 8.595%; 2.11: -16.36%; 2.13: -12.18%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.31; 2.04: 0.67; 2.05: NA; 2.06: 0.12; 2.07: 0; 2.08: 0.16; 2.10: 0.71; 2.11: 1.98; 2.12: 0.78; 2.13: 1.53
Results from laboratories 2.03 and 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 609c (Certified Value 1.32 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.436 | 1.32 | 1.24 | 1.29 | 1.3 | 1.31 | 1.3 | 1.21 | 1.38 | 1.32 | 1.295 | 1.245 | 1.285 | 1.405 | 1.237 | 1.424 | 1.3 | 1.356 | 1.2841 | 1.245 | [6 empty] | 1
| 1 | 2 | 1.514 | 1.19 | 1.21 | 1.3 | 1.3 | 1.34 | 1.2 | 1.25 | 1.29 | 1.285 | 1.355 | 1.365 | 1.35 | 1.435 | 1.209 | 1.423 | 1.34 | 1.378 | 1.224 | 1.238 | [6 empty] | 25
| 1 | 3 | 1.473 | 1.24 | 1.26 | 1.33 | 1.3 | 1.37 | 1.2 | 1.27 | 1.33 | 1.33 | 1.32 | 1.35 | 1.295 | 1.385 | 1.208 | 1.399 | 1.33 | 1.354 | 1.3272 | 1.126 | [6 empty] | 16
| 1 | 4 | 1.392 | 1.28 | 1.24 | 1.3 | 1.2 | 1.35 | 1.2 | 1.27 | 1.37 | 1.25 | 1.315 | 1.43 | 1.385 | 1.44 | 1.161 | 1.412 | 1.42 | 1.414 | 1.2654 | 1.308 | [6 empty] | 1.3187
| 1 | 5 | 1.354 | 1.21 | 1.22 | 1.32 | 1.2 | 1.33 | 1.1 | 1.25 | 1.33 | 1.305 | 1.31 | 1.42 | 1.335 | 1.41 | 1.202 | 1.447 | 1.32 | 1.419 | 1.2468 | 1.284 | [6 empty] | 34
| 1 | 6 | 1.48 | 1.32 | 1.25 | 1.3 | 1.3 | 1.37 | 1.2 | 1.28 | 1.36 | 1.245 | 1.32 | 1.445 | 1.28 | 1.42 | 1.202 | 1.381 | 1.29 | 1.396 | 1.3154 | 1.176
| Mean |  | 1.4415 | 1.26 | 1.2367 | 1.3067 | 1.2667 | 1.345 | 1.2 | 1.255 | 1.3433 | 1.2892 | 1.3192 | 1.3758 | 1.3217 | 1.4158 | 1.2032 | 1.4143 | 1.3333 | 1.3862 | 1.2771 | 1.2295
| Median |  | 1.4545 | 1.26 | 1.24 | 1.3 | 1.3 | 1.345 | 1.2 | 1.26 | 1.345 | 1.295 | 1.3175 | 1.3925 | 1.315 | 1.415 | 1.205 | 1.4175 | 1.325 | 1.387 | 1.2747 | 1.2415
| Std Dev. |  | 0.059779 | 0.055498 | 0.018619 | 0.015055 | 0.05164 | 0.023452 | 0.063246 | 0.0251 | 0.033267 | 0.035695 | 0.019854 | 0.074257 | 0.041913 | 0.020351 | 0.024425 | 0.022748 | 0.046332 | 0.028145 | 0.039744 | 0.067852
| Rel.Std.Dev. |  | 4.147% | 4.405% | 1.506% | 1.152% | 4.077% | 1.744% | 5.27% | 2% | 2.476% | 2.769% | 1.505% | 5.397% | 3.171% | 1.437% | 2.03% | 1.608% | 3.475% | 2.03% | 3.112% | 5.519%
| PDM3 |  | 9.314% | -4.45% | -6.219% | -0.9106% | -3.944% | 1.996% | -9% | -4.829% | 1.87% | -2.238% | 0.03731% | 4.335% | 0.2269% | 7.368% | -8.759% | 7.254% | 1.112% | 5.118% | -3.149% | -6.762%
| Z-Score (Absolute) |  | 1.41 | 0.7 | 0.98 | 0.16 | NA | 0.29 | NA | 0.76 | 0.27 | 0.36 | 0.01 | 0.65 | 0.01 | 1.11 | 1.37 | 1.09 | 0.15 | 0.77 | 0.5 | 1.06
Results from laboratories 2.07 and 2.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 609c (Certified Value 2.76 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.85 | 2.86 | 2.99 | 2.7577 | 2.7415 | 2.8 | 2.67 | 2.7223 | 2.91 | 2.71 | 2.6347 | 2.82 | 2.75 | 2.58 | 2.67 | 2.85 | 2.6987 | 2.85 | 2.55 | 2.59 | 2.64 | 2.4468 | 2.69 | 2.7338 | 3.1255 | 2.83 | 1
| 1 | 2 | 2.83 | 2.89 | 2.87 | 2.8114 | 2.7256 | 2.87 | 2.69 | 2.7396 | 2.84 | 2.68 | 2.6982 | 2.86 | 2.74 | 2.6 | 2.77 | 2.87 | 2.7054 | 2.83 | 2.63 | 2.63 | 2.78 | 2.3839 | 2.665 | 2.6216 | 3.1373 | 2.58 | 0
| 1 | 3 | 2.87 | 2.87 | 2.94 | 2.8177 | 2.7312 | 2.89 | 2.75 | 2.7794 | 2.85 | 2.68 | 2.7288 | 2.83 | 2.71 | 2.67 | 2.63 | 2.87 | 2.7 | 2.81 | 2.65 | 2.61 | 2.79 | 2.3824 | 2.665 | 2.7428 | 3.0633 | 2.69 | 16
| 1 | 4 | 2.86 | 2.97 | 2.97 | 2.7662 | 2.72 | 2.81 | 2.78 | 2.7092 | 2.82 | 2.77 | 2.805 | 2.74 | 2.69 | 2.66 | 2.74 | 2.89 | 2.6449 | 2.85 | 3.05 | 2.65 | 2.72 | 2.3759 | 2.798 | 2.5479 | 2.6413 | 2.89 | 2.7604
| 1 | 5 | 2.85 | 2.85 | 2.94 | 2.8447 | 2.7628 | 2.96 | 2.7 | 2.722 | 2.81 | 2.77 | 2.6337 | 2.91 | 2.73 | 2.65 | 2.67 | 2.95 | 2.6815 | 2.94 | 2.81 | 2.59 | 2.77 | 2.4634 | 2.748 | 2.6078 | 3.0843 | 2.65 | 35
| 1 | 6 | 2.87 | 2.83 | 2.99 | 2.756 | 2.7493 | 3.08 | 2.65 | 2.7006 | 2.86 | 2.65 | 2.7553 | 2.77 | 2.71 | 2.63 | 2.66 | 2.89 | 2.6656 | 2.81 | 2.82 | 2.59 | 2.77 | 2.4072 | 2.648 | 2.6613 | 3.0511 | 2.79
| Mean |  | 2.855 | 2.8783 | 2.95 | 2.7923 | 2.7384 | 2.9017 | 2.7067 | 2.7288 | 2.8483 | 2.71 | 2.7093 | 2.8217 | 2.7217 | 2.6317 | 2.69 | 2.8867 | 2.6827 | 2.8483 | 2.7517 | 2.61 | 2.745 | 2.4099 | 2.7023 | 2.6525 | 3.0171 | 2.7383
| Median |  | 2.855 | 2.865 | 2.955 | 2.7888 | 2.7363 | 2.88 | 2.695 | 2.7222 | 2.845 | 2.695 | 2.7135 | 2.825 | 2.72 | 2.64 | 2.67 | 2.88 | 2.6901 | 2.84 | 2.73 | 2.6 | 2.77 | 2.3956 | 2.6775 | 2.6414 | 3.0738 | 2.74
| Std Dev. |  | 0.015166 | 0.04916 | 0.045166 | 0.037288 | 0.015964 | 0.10496 | 0.049261 | 0.028081 | 0.035449 | 0.0502 | 0.067909 | 0.061128 | 0.022286 | 0.035449 | 0.053292 | 0.034448 | 0.023636 | 0.048339 | 0.18049 | 0.025298 | 0.056833 | 0.036931 | 0.058531 | 0.075786 | 0.18719 | 0.11771
| Rel.Std.Dev. |  | 0.5312% | 1.708% | 1.531% | 1.335% | 0.583% | 3.617% | 1.82% | 1.029% | 1.245% | 1.852% | 2.507% | 2.166% | 0.8188% | 1.347% | 1.981% | 1.193% | 0.8811% | 1.697% | 6.559% | 0.9693% | 2.07% | 1.532% | 2.166% | 2.857% | 6.204% | 4.299%
| PDM3 |  | 3.427% | 4.272% | 6.868% | 1.155% | -0.7973% | 5.117% | -1.947% | -1.144% | 3.185% | -1.826% | -1.852% | 2.219% | -1.403% | -4.664% | -2.551% | 4.574% | -2.816% | 3.185% | -0.3167% | -5.449% | -0.5582% | -12.7% | -2.104% | -3.908% | 9.3% | -0.7997%
| Z-Score (Absolute) |  | 1.13 | 1.36 | 2.05 | 0.52 | 0 | 1.58 | 0.31 | 0.09 | 1.07 | 0.28 | 0.28 | 0.81 | 0.16 | 1.04 | 0.47 | 1.44 | 0.54 | 1.07 | 0.13 | 1.25 | 0.06 | 3.19 | 0.35 | 0.83 | 2.7 | 0
| Analytical results for La in OREAS 609c (Certified Value 38.7 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 39.7 | 33.7 | 33.9 | 43.81 | 37.3984 | 38.1 | 45.8 | 43.531 | 38.27 | 42.28 | 42.2 | 39.6 | 41.6 | 39.1 | 39.5 | 34.2 | 36.01 | 55.1 | 35.95 | 39.7 | 32.0065 | 42 | 28.969 | 36.88 |  |  | 1
| 1 | 2 | 39.4 | 33.3 | 35.3 | 43.71 | 37.0258 | 38.4 | 45.1 | 42.445 | 38.37 | 41.85 | 39.4 | 39.7 | 35.6 | 41.1 | 40.6 | 33.7 | 32.08 | 53.5 | 36.11 | 40 | 34.7912 | 41 | 27.409 | 36.62 |  |  | 26
| 1 | 3 | 41.3 | 33.5 | 35.7 | 44.06 | 37.3327 | 40.1 | 46.8 | 43.672 | 37.73 | 43.07 | 42.7 | 38.7 | 34.9 | 39.3 | 39.7 | 33.6 | 31.81 | 53.8 | 36.57 | 39.1 | 32.1951 | 41 | 22.309 | 38.78 |  |  | 16
| 1 | 4 | 39.8 | 35.6 | 33.8 | 43.57 | 37.0447 | 37.9 | 46.8 | 43.622 | 37.7 | 43.65 | 41.4 | 39.7 | 36.6 | 42 | 39.6 | 33.2 | 33.83 | 55.6 | 37.27 | 38.6 | 35.9367 | 45 | 29.094 | 38.54 |  |  | 38.6767
| 1 | 5 | 36 | 34.5 | 34 | 43.37 | 36.9769 | 38 | 46.2 | 43.252 | 36.17 | 41.56 | 44.1 | 39.3 | 37.8 | 40.2 | 40.9 | 33.1 | 34.47 | 57.3 | 33.47 | 39.4 | 41.6573 | 43 | 22.73 | 35.4 |  |  | 36
| 1 | 6 | 42 | 34 | 33.9 | 43.71 | 37.5179 | 37.9 | 47.1 | 43.681 | 36.85 | 38.49 | 39.4 | 38.2 | 36.7 | 39.7 | 39.4 | 33 | 31.25 | 54.9 | 34.7 | 40.2 | 35.4995 | 41 | 27.346 | 39.52
| Mean |  | 39.7 | 34.1 | 34.4333 | 43.705 | 37.2161 | 38.4 | 46.3 | 43.3672 | 37.515 | 41.8167 | 41.5333 | 39.2 | 37.2 | 40.2333 | 39.95 | 33.4667 | 33.2417 | 55.0333 | 35.6783 | 39.5 | 35.3477 | 42.1667 | 26.3095 | 37.6233
| Median |  | 39.75 | 33.85 | 33.95 | 43.71 | 37.1887 | 38.05 | 46.5 | 43.5765 | 37.715 | 42.065 | 41.8 | 39.45 | 36.65 | 39.95 | 39.65 | 33.4 | 32.955 | 55 | 36.03 | 39.55 | 35.1454 | 41.5 | 27.3775 | 37.71
| Std Dev. |  | 2.0804 | 0.84617 | 0.83825 | 0.23132 | 0.22834 | 0.85323 | 0.75366 | 0.47893 | 0.85238 | 1.8045 | 1.8715 | 0.61968 | 2.374 | 1.1237 | 0.63482 | 0.45461 | 1.8376 | 1.3677 | 1.3728 | 0.5933 | 3.5098 | 1.6021 | 3.0307 | 1.567
| Rel.Std.Dev. |  | 5.24% | 2.481% | 2.434% | 0.5293% | 0.6136% | 2.222% | 1.628% | 1.104% | 2.272% | 4.315% | 4.506% | 1.581% | 6.382% | 2.793% | 1.589% | 1.358% | 5.528% | 2.485% | 3.848% | 1.502% | 9.929% | 3.799% | 11.52% | 4.165%
| PDM3 |  | 2.646% | -11.83% | -10.97% | 13% | -3.776% | -0.7153% | 19.71% | 12.13% | -3.003% | 8.119% | 7.386% | 1.353% | -3.818% | 4.025% | 3.292% | -13.47% | -14.05% | 42.29% | -7.752% | 2.129% | -8.607% | 9.024% | -31.98% | -2.723%
| Z-Score (Absolute) |  | 0.2 | 1.03 | 0.96 | 1.08 | 0.35 | 0.09 | 1.65 | 1 | 0.28 | 0.66 | 0.6 | 0.09 | 0.35 | 0.31 | 0.25 | 1.17 | 1.22 | 3.57 | 0.69 | 0.15 | 0.76 | 0.74 | 2.74 | 0.26
| Analytical results for Li in OREAS 609c (Certified Value 21.6 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 17.1 | 22.1 | 22.9 | 21.5 | 21.3408 | 23 | 17.9 | 24.1589 | 21 | 21.9 | 22.2 | 22.1 | 20.4 | 20.6 | 20.9 | 21.7 | 22 | 22.7 | 20.9 | 21.42 | 18.1 | 36.9684 | 21 | 22.4539 | 28.304 | 21.7 | 1
| 1 | 2 | 16 | 21.5 | 22.3 | 21 | 21.6478 | 22 | 18.5 | 24.0219 | 20 | 22.2 | 21.6 | 22.4 | 20.4 | 21 | 21.4 | 22.1 | 22 | 22.7 | 20.6 | 22.29 | 18.9 | 35.5388 | 21 | 21.0971 | 28.942 | 21.4 | 46
| 1 | 3 | 15.5 | 20.8 | 22.8 | 21.3 | 21.4839 | 23 | 19.6 | 24.7854 | 20.4 | 21.7 | 22 | 22 | 22 | 21.6 | 20.6 | 21.9 | 21 | 22.1 | 20.8 | 20.99 | 18 | 36.6669 | 21 | 22.2915 | 28.943 | 22.7 | 16
| 1 | 4 | 15.8 | 23.4 | 23 | 21.3 | 21.5861 | 23 | 18.9 | 24.8291 | 20 | 22.3 | 22.4 | 21.4 | 20 | 21.4 | 21.5 | 22.2 | 21 | 22.8 | 20.4 | 21.34 | 20.6 | 35.1713 | 22 | 21.2459 | 28.389 | 22.4 | 21.6166
| 1 | 5 | 17 | 20.7 | 22.5 | 21.3 | 21.8189 | 23 | 18.1 | 24.8498 | 20.3 | 21.7 | 21.9 | 22.7 | 20.3 | 21.3 | 21.2 | 22.1 | 22 | 23.1 | 21.6 | 21.93 | 18.8 | 36.9051 | 21 | 21.4185 | 28.738 | 22 | 37
| 1 | 6 | 17.1 | 22.4 | 23 | 21.2 | 21.8321 | 24 | 19.3 | 24.4702 | 19.8 | 22.1 | 22.2 | 21.6 | 20.4 | 21.2 | 20.8 | 21.9 | 22 | 23.2 | 21.8 | 21.28 | 17.7 | 36.0284 | 20 | 21.7264 | 30.663 | 22.7
| Mean |  | 16.4167 | 21.8167 | 22.75 | 21.2667 | 21.6183 | 23 | 18.7167 | 24.5192 | 20.25 | 21.9833 | 22.05 | 22.0333 | 20.5833 | 21.1833 | 21.0667 | 21.9833 | 21.6667 | 22.7667 | 21.0167 | 21.5417 | 18.6833 | 36.2131 | 21 | 21.7056 | 28.9965 | 22.15
| Median |  | 16.5 | 21.8 | 22.85 | 21.3 | 21.6169 | 23 | 18.7 | 24.6278 | 20.15 | 22 | 22.1 | 22.05 | 20.4 | 21.25 | 21.05 | 22 | 22 | 22.75 | 20.85 | 21.38 | 18.45 | 36.3477 | 21 | 21.5725 | 28.84 | 22.2
| Std Dev. |  | 0.73052 | 1.0304 | 0.2881 | 0.1633 | 0.19126 | 0.63246 | 0.67057 | 0.36217 | 0.42778 | 0.25626 | 0.28107 | 0.48442 | 0.7111 | 0.34881 | 0.3559 | 0.18348 | 0.5164 | 0.38816 | 0.56006 | 0.4773 | 1.0496 | 0.75225 | 0.63246 | 0.55997 | 0.86006 | 0.53944
| Rel.Std.Dev. |  | 4.45% | 4.723% | 1.266% | 0.7679% | 0.8847% | 2.75% | 3.583% | 1.477% | 2.113% | 1.166% | 1.275% | 2.199% | 3.455% | 1.647% | 1.689% | 0.8347% | 2.383% | 1.705% | 2.665% | 2.216% | 5.618% | 2.077% | 3.012% | 2.58% | 2.966% | 2.435%
| PDM3 |  | -24.06% | 0.9255% | 5.243% | -1.619% | 0.007733% | 6.4% | -13.42% | 13.43% | -6.322% | 1.697% | 2.005% | 1.928% | -4.78% | -2.004% | -2.544% | 1.697% | 0.2316% | 5.32% | -2.775% | -0.3466% | -13.57% | 67.52% | -2.852% | 0.4115% | 34.14% | 2.468%
| Z-Score (Absolute) |  | 5.58 | 0.12 | 1.1 | 0.46 | 0.09 | NA | 3.15 | 2.97 | 1.54 | 0.29 | 0.36 | 0.35 | 1.18 | 0.55 | 0.67 | 0.29 | NA | 1.12 | 0.73 | 0.17 | 3.19 | 15.31 | NA | 0 | 7.7 | 0.47
Results from laboratories 2.07, 2.18 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 609c (Certified Value 0.08 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.12; 2.07: 2.19; 2.08: 2.22; 2.09: 2.23; 2.10: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS | 3
3
| 1 | 1 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.085073; 2.05: < 0.1; 2.06: 0.07; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0839; 2.10: 0.05 | 1
| 1 | 2 | 2.01: < 0.1; 2.03: 0.08; 2.04: 0.08676; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0918; 2.10: 0.05 | 27
| 1 | 3 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.082233; 2.05: < 0.1; 2.06: 0.07; 2.07: 0.08; 2.08: 0.09; 2.09: 0.0885; 2.10: 0.05 | 16
| 1 | 4 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.084416; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0868; 2.10: 0.05 | 0.079712
| 1 | 5 | 2.01: < 0.1; 2.03: 0.08; 2.04: 0.086759; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0881; 2.10: 0.04 | 38
| 1 | 6 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.087459; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.08; 2.08: 0.08; 2.09: 0.1021; 2.10: 0.05
| Mean |  | 2.03: 0.073333; 2.04: 0.08545; 2.06: 0.063333; 2.07: 0.086667; 2.08: 0.081667; 2.09: 0.0902; 2.10: 0.048333
| Median |  | 2.03: 0.07; 2.04: 0.085916; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0883; 2.10: 0.05
| Std Dev. |  | 2.03: 0.005164; 2.04: 0.0019508; 2.06: 0.005164; 2.07: 0.005164; 2.08: 0.0040825; 2.09: 0.0063674; 2.10: 0.0040825
| Rel.Std.Dev. |  | 2.03: 7.042%; 2.04: 2.283%; 2.06: 8.154%; 2.07: 5.958%; 2.08: 4.999%; 2.09: 7.059%; 2.10: 8.447%
| PDM3 |  | 2.03: -8.002%; 2.04: 7.199%; 2.06: -20.55%; 2.07: 8.725%; 2.08: 2.453%; 2.09: 13.16%; 2.10: -39.36%
| Z-Score (Absolute) |  | 2.01: 1.18; 2.03: 0; 2.04: 0.61; 2.05: 1.18; 2.06: 0.51; 2.07: 0.67; 2.08: 0.42; 2.09: 0.85; 2.10: 1.26
| Analytical results for Mg in OREAS 609c (Certified Value 0.119 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.12 | 0.11 | 0.13 | 0.1194 | 0.11762 | 0.12 | 0.12 | 0.13151 | 0.12 | 0.124 | 0.117 | 0.12 | 0.11 | 0.11 | 0.11 | 0.09 | 0.1192 | 0.124 | 0.12 | 0.12 | 0.11 | 0.115 | 0.121 | 0.12523 | 0.11185 | 0.117 | 1
| 1 | 2 | 0.12 | 0.12 | 0.12 | 0.1224 | 0.11957 | 0.12 | 0.12 | 0.13678 | 0.12 | 0.12 | 0.1188 | 0.12 | 0.11 | 0.11 | 0.12 | 0.09 | 0.119 | 0.12 | 0.12 | 0.13 | 0.12 | 0.112 | 0.121 | 0.11677 | 0.1097 | 0.11 | 0
| 1 | 3 | 0.12 | 0.11 | 0.13 | 0.1224 | 0.11823 | 0.13 | 0.12 | 0.13079 | 0.12 | 0.124 | 0.1196 | 0.12 | 0.11 | 0.11 | 0.11 | 0.09 | 0.1217 | 0.121 | 0.12 | 0.12 | 0.11 | 0.1136 | 0.121 | 0.12315 | 0.10502 | 0.108 | 16
| 1 | 4 | 0.13 | 0.12 | 0.13 | 0.1205 | 0.11976 | 0.12 | 0.12 | 0.13255 | 0.12 | 0.123 | 0.1233 | 0.11 | 0.11 | 0.11 | 0.11 | 0.09 | 0.1169 | 0.123 | 0.13 | 0.13 | 0.11 | 0.1108 | 0.127 | 0.11639 | 0.10998 | 0.122 | 0.119
| 1 | 5 | 0.12 | 0.11 | 0.13 | 0.1238 | 0.11929 | 0.13 | 0.12 | 0.13348 | 0.12 | 0.129 | 0.1163 | 0.12 | 0.11 | 0.11 | 0.11 | 0.09 | 0.1182 | 0.126 | 0.13 | 0.12 | 0.12 | 0.1154 | 0.121 | 0.11798 | 0.11203 | 0.108 | 39
| 1 | 6 | 0.12 | 0.11 | 0.13 | 0.1197 | 0.11766 | 0.13 | 0.12 | 0.13175 | 0.12 | 0.126 | 0.1213 | 0.11 | 0.11 | 0.11 | 0.11 | 0.08 | 0.119 | 0.119 | 0.12 | 0.12 | 0.12 | 0.1138 | 0.121 | 0.12232 | 0.11183 | 0.114
| Mean |  | 0.12167 | 0.11333 | 0.12833 | 0.12137 | 0.11869 | 0.125 | 0.12 | 0.13281 | 0.12 | 0.12433 | 0.11938 | 0.11667 | 0.11 | 0.11 | 0.11167 | 0.088333 | 0.119 | 0.12217 | 0.12333 | 0.12333 | 0.115 | 0.11343 | 0.122 | 0.12031 | 0.11007 | 0.11317
| Median |  | 0.12 | 0.11 | 0.13 | 0.12145 | 0.11876 | 0.125 | 0.12 | 0.13215 | 0.12 | 0.124 | 0.1192 | 0.12 | 0.11 | 0.11 | 0.11 | 0.09 | 0.119 | 0.122 | 0.12 | 0.12 | 0.115 | 0.1137 | 0.121 | 0.12015 | 0.1109 | 0.112
| Std Dev. |  | 0.0040825 | 0.005164 | 0.0040825 | 0.001758 | 0.00096952 | 0.0054772 | 0 | 0.0021539 | 0 | 0.0030111 | 0.0026301 | 0.005164 | 0 | 0 | 0.0040825 | 0.0040825 | 0.0015735 | 0.0026394 | 0.005164 | 0.005164 | 0.0054772 | 0.0017592 | 0.0024495 | 0.0037321 | 0.0026723 | 0.0056006
| Rel.Std.Dev. |  | 3.355% | 4.556% | 3.181% | 1.449% | 0.8169% | 4.382% | 0% | 1.622% | 0% | 2.422% | 2.203% | 4.426% | 0% | 0% | 3.656% | 4.622% | 1.322% | 2.161% | 4.187% | 4.187% | 4.763% | 1.551% | 2.008% | 3.102% | 2.428% | 4.949%
| PDM3 |  | 2.245% | -4.758% | 7.847% | 1.993% | -0.2584% | 5.046% | 0.8442% | 11.61% | 0.8442% | 4.486% | 0.326% | -1.957% | -7.559% | -7.559% | -6.159% | -25.77% | 0.003815% | 2.665% | 3.645% | 3.645% | -3.358% | -4.674% | 2.525% | 1.101% | -7.502% | -4.898%
| Z-Score (Absolute) |  | 0.32 | 1.04 | 1.41 | 0.27 | 0.16 | 0.86 | 0.05 | 2.14 | 0.05 | 0.76 | 0.05 | 0.49 | 1.58 | 1.58 | 1.31 | 5.11 | 0.11 | 0.4 | 0.59 | 0.59 | 0.76 | 1.02 | 0.31 | 0.1 | 1.57 | 1.06
| Analytical results for Mn in OREAS 609c (Certified Value 0.025 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0254 | 0.0247 | 0.0262 | 0.0249 | 0.025912 | 0.0268 | 0.0249 | 0.02471 | 0.0266 | 0.0232 | 0.0257 | 0.0254 | 0.0251 | 0.0233 | 0.0247 | 0.026 | 0.0244 | 0.0254 | 0.0299 | 0.02338 | 0.0251 | < 0.1 | 0.023 | 0.026113 | 0.024452 | 0.0215 | 1
| 1 | 2 | 0.0254 | 0.0261 | 0.0249 | 0.0254 | 0.026004 | 0.0268 | 0.026 | 0.024876 | 0.0257 | 0.0239 | 0.0261 | 0.0253 | 0.025 | 0.0238 | 0.0254 | 0.026 | 0.0242 | 0.024 | 0.0303 | 0.02367 | 0.0252 | < 0.1 | 0.023 | 0.024928 | 0.024571 | 0.0214 | 28
| 1 | 3 | 0.0249 | 0.0252 | 0.0257 | 0.0256 | 0.025785 | 0.0267 | 0.0265 | 0.024385 | 0.0262 | 0.0237 | 0.0261 | 0.0253 | 0.0252 | 0.0245 | 0.0242 | 0.0296 | 0.0244 | 0.0234 | 0.0299 | 0.02353 | 0.0247 | < 0.1 | 0.023 | 0.025779 | 0.024239 | 0.0229 | 16
| 1 | 4 | 0.0257 | 0.0261 | 0.0265 | 0.025 | 0.026132 | 0.0264 | 0.0263 | 0.024861 | 0.0263 | 0.024 | 0.0273 | 0.0246 | 0.0245 | 0.0246 | 0.0252 | 0.0266 | 0.0236 | 0.0249 | 0.0322 | 0.02341 | 0.0242 | < 0.1 | 0.023 | 0.024751 | 0.024719 | 0.022 | 0.025097
| 1 | 5 | 0.0254 | 0.0256 | 0.0256 | 0.0259 | 0.025808 | 0.0274 | 0.0255 | 0.025061 | 0.0247 | 0.0232 | 0.0254 | 0.026 | 0.0252 | 0.0247 | 0.0247 | 0.0268 | 0.024 | 0.0257 | 0.0313 | 0.0232 | 0.0248 | < 0.1 | 0.023 | 0.0252 | 0.024031 | 0.0212 | 40
| 1 | 6 | 0.0259 | 0.0253 | 0.026 | 0.025 | 0.026006 | 0.0283 | 0.0267 | 0.024561 | 0.026 | 0.0235 | 0.0267 | 0.025 | 0.0251 | 0.0241 | 0.0246 | 0.0262 | 0.0243 | 0.0242 | 0.0312 | 0.02309 | 0.0251 | < 0.1 | 0.023 | 0.026368 | 0.024512 | 0.0225
| Mean |  | 0.02545 | 0.0255 | 0.025817 | 0.0253 | 0.025941 | 0.027067 | 0.025983 | 0.024742 | 0.025917 | 0.023583 | 0.026217 | 0.025267 | 0.025017 | 0.024167 | 0.0248 | 0.026867 | 0.02415 | 0.0246 | 0.0308 | 0.02338 | 0.02485 |  | 0.023 | 0.025523 | 0.024421 | 0.021917
| Median |  | 0.0254 | 0.02545 | 0.02585 | 0.0252 | 0.025958 | 0.0268 | 0.02615 | 0.024785 | 0.0261 | 0.0236 | 0.0261 | 0.0253 | 0.0251 | 0.0243 | 0.0247 | 0.0264 | 0.02425 | 0.02455 | 0.03075 | 0.023395 | 0.02495 |  | 0.023 | 0.025489 | 0.024482 | 0.02175
| Std Dev. |  | 0.00033912 | 0.00054772 | 0.00055648 | 0.0004 | 0.00013256 | 0.00068605 | 0.00067651 | 0.00024294 | 0.00066758 | 0.00034303 | 0.00068823 | 0.00046332 | 0.00026394 | 0.00054283 | 0.00043359 | 0.0013779 | 0.00030822 | 0.00088318 | 0.00092087 | 0.00021166 | 0.00037283 |  | 3.8006e-18 | 0.00066039 | 0.00024754 | 0.00067355
| Rel.Std.Dev. |  | 1.332% | 2.148% | 2.155% | 1.581% | 0.511% | 2.535% | 2.604% | 0.9819% | 2.576% | 1.455% | 2.625% | 1.834% | 1.055% | 2.246% | 1.748% | 5.129% | 1.276% | 3.59% | 2.99% | 0.9053% | 1.5% |  | 1.652e-14% | 2.587% | 1.014% | 3.073%
| PDM3 |  | 1.405% | 1.604% | 2.866% | 0.8071% | 3.362% | 7.846% | 3.53% | -1.415% | 3.264% | -6.033% | 4.46% | 0.6743% | -0.3218% | -3.709% | -1.185% | 7.049% | -3.775% | -1.982% | 22.72% | -6.843% | -0.9859% |  | -8.357% | 1.696% | -2.696% | -12.67%
| Z-Score (Absolute) |  | 0.16 | 0.21 | 0.52 | 0.02 | 0.64 | 1.74 | 0.68 | 0.53 | 0.62 | 1.66 | 0.91 | 0.02 | 0.26 | 1.09 | 0.47 | 1.54 | 1.1 | 0.67 | 5.38 | 1.86 | 0.42 | 24.1 | 2 | 0.23 | 0.84 | 3.28
| Analytical results for Mo in OREAS 609c (Certified Value 40.1 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 41.07 | 39.07 | 38.6 | 41.5 | 39.7699 | 41 | 39.1 | 39.24 | 40.9 | 40.6 | 41.1 | 42.6 | 40.6 | 39.6 | 41.3 | 40.8 | 41.6 | 41.4 | 37.31 | 47.47 | 38.2848 | 38 | 40.8106 | 37.89 | 37.4 |  | 1
| 1 | 2 | 45.5 | 38.11 | 37.8 | 41.3 | 40.1047 | 42 | 40 | 39.13 | 40.2 | 40.7 | 40 | 43.4 | 38.7 | 41 | 41.8 | 41.1 | 43.3 | 42.2 | 38.47 | 47.79 | 36.7987 | 39 | 38.6905 | 38.76 | 36.3 |  | 29
| 1 | 3 | 41.88 | 35.85 | 38.7 | 42.1 | 39.4633 | 42 | 41 | 40.03 | 40.3 | 41.3 | 40.2 | 43.6 | 39.1 | 39.8 | 41 | 39.3 | 42.8 | 41.7 | 36.61 | 48.1 | 36.8852 | 37 | 38.464 | 37.56 | 38.3 |  | 16
| 1 | 4 | 42.55 | 39.89 | 38.8 | 40.7 | 40.1035 | 41 | 40.2 | 39.73 | 40.6 | 41.7 | 38.4 | 42.5 | 41.1 | 41.4 | 41.5 | 39 | 42.8 | 41.7 | 38.63 | 47.36 | 35.9723 | 42 | 38.0739 | 37.93 | 37.6 |  | 40.0522
| 1 | 5 | 41.59 | 37.73 | 38.7 | 41.3 | 40.3857 | 41 | 39.8 | 38.13 | 40 | 42.1 | 39.5 | 42.8 | 41.2 | 40.4 | 41.4 | 40 | 43.1 | 41.7 | 38.03 | 47.81 | 37.9273 | 41 | 39.1442 | 37.18 | 38.6 |  | 41
| 1 | 6 | 41.74 | 39.66 | 39.4 | 41.2 | 40.1668 | 43 | 40.6 | 39.4 | 40.8 | 40.5 | 38.6 | 43.2 | 41.3 | 39.8 | 41.1 | 40.6 | 42.1 | 41.1 | 38.09 | 50.16 | 37.2074 | 37 | 41.2501 | 38.09 | 39.9
| Mean |  | 42.3883 | 38.385 | 38.6667 | 41.35 | 39.999 | 41.6667 | 40.1167 | 39.2767 | 40.4667 | 41.15 | 39.6333 | 43.0167 | 40.3333 | 40.3333 | 41.35 | 40.1333 | 42.6167 | 41.6333 | 37.8567 | 48.115 | 37.1793 | 39 | 39.4055 | 37.9017 | 38.0167
| Median |  | 41.81 | 38.59 | 38.7 | 41.3 | 40.1041 | 41.5 | 40.1 | 39.32 | 40.45 | 41 | 39.75 | 43 | 40.85 | 40.1 | 41.35 | 40.3 | 42.8 | 41.7 | 38.06 | 47.8 | 37.0463 | 38.5 | 38.9173 | 37.91 | 37.95
| Std Dev. |  | 1.5977 | 1.5027 | 0.51251 | 0.45497 | 0.32839 | 0.8165 | 0.65853 | 0.65206 | 0.3559 | 0.65651 | 1.0211 | 0.44907 | 1.1431 | 0.73394 | 0.2881 | 0.84774 | 0.64317 | 0.36697 | 0.76312 | 1.036 | 0.83314 | 2.0976 | 1.3127 | 0.53117 | 1.2222
| Rel.Std.Dev. |  | 3.769% | 3.915% | 1.325% | 1.1% | 0.821% | 1.96% | 1.642% | 1.66% | 0.8795% | 1.595% | 2.576% | 1.044% | 2.834% | 1.82% | 0.6967% | 2.112% | 1.509% | 0.8814% | 2.016% | 2.153% | 2.241% | 5.379% | 3.331% | 1.401% | 3.215%
| PDM3 |  | 5.833% | -4.163% | -3.459% | 3.24% | -0.1329% | 4.031% | 0.1609% | -1.936% | 1.035% | 2.741% | -1.046% | 7.401% | 0.7019% | 0.7019% | 3.24% | 0.2025% | 6.403% | 3.948% | -5.482% | 20.13% | -7.173% | -2.627% | -1.615% | -5.369% | -5.082%
| Z-Score (Absolute) |  | 1.25 | 0.97 | 0.81 | 0.67 | 0.07 | 0.85 | 0.01 | 0.47 | 0.18 | 0.56 | 0.28 | 1.6 | 0.11 | 0.11 | 0.67 | 0 | 1.38 | 0.83 | 1.26 | 4.42 | 1.64 | 0.63 | 0.4 | 1.24 | 1.17
| Analytical results for Na in OREAS 609c (Certified Value 2.22 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.19 | 2.261 | 2.338 | 2.2223 | 2.2076 | 2.32 | 2.08 | 2.3574 | 2.36 | 2.18 | 2.1915 | 2.26 | 2.27 | 2.08 | 2.16 | 2.17 | 2.4256 | 2.33 | 1.85 | 2.16 | 2.23 | 2.2709 | 2.226 | 2.12 | 2.1188 | 2.348 | 1
| 1 | 2 | 2.19 | 2.253 | 2.317 | 2.2586 | 2.201 | 2.42 | 2.18 | 2.3635 | 2.31 | 2.11 | 2.2443 | 2.29 | 2.25 | 2.13 | 2.22 | 2.21 | 2.464 | 2.315 | 1.85 | 2.19 | 2.19 | 2.2128 | 2.196 | 2.0277 | 2.0824 | 2.137 | 0
| 1 | 3 | 2.19 | 2.258 | 2.331 | 2.267 | 2.2128 | 2.46 | 2.21 | 2.3263 | 2.31 | 2.13 | 2.2658 | 2.24 | 2.26 | 2.17 | 2.14 | 2.2 | 2.4765 | 2.297 | 1.83 | 2.18 | 2.25 | 2.209 | 2.188 | 2.1262 | 1.9974 | 2.238 | 16
| 1 | 4 | 2.21 | 2.316 | 2.278 | 2.2245 | 2.2377 | 2.44 | 2.22 | 2.3496 | 2.28 | 2.16 | 2.3373 | 2.18 | 2.21 | 2.17 | 2.22 | 2.21 | 2.383 | 2.329 | 2.02 | 2.18 | 2.17 | 2.2216 | 2.3 | 1.9779 | 2.0628 | 2.397 | 2.221
| 1 | 5 | 2.2 | 2.233 | 2.318 | 2.2948 | 2.2624 | 2.54 | 2.16 | 2.3498 | 2.28 | 2.15 | 2.1887 | 2.3 | 2.25 | 2.17 | 2.19 | 2.22 | 2.4033 | 2.408 | 1.92 | 2.16 | 2.25 | 2.2969 | 2.263 | 2.0213 | 2.1211 | 2.203 | 42
| 1 | 6 | 2.2 | 2.214 | 2.342 | 2.2169 | 2.2186 | 2.63 | 2.23 | 2.3472 | 2.24 | 2.12 | 2.2911 | 2.2 | 2.24 | 2.14 | 2.15 | 2.21 | 2.4415 | 2.314 | 1.87 | 2.16 | 2.21 | 2.2241 | 2.174 | 2.0723 | 2.1276 | 2.316
| Mean |  | 2.1967 | 2.2558 | 2.3207 | 2.2473 | 2.2234 | 2.4683 | 2.18 | 2.349 | 2.2967 | 2.1417 | 2.2531 | 2.245 | 2.2467 | 2.1433 | 2.18 | 2.2033 | 2.4323 | 2.3322 | 1.89 | 2.1717 | 2.2167 | 2.2392 | 2.2245 | 2.0576 | 2.085 | 2.2732
| Median |  | 2.195 | 2.2555 | 2.3245 | 2.2416 | 2.2157 | 2.45 | 2.195 | 2.3497 | 2.295 | 2.14 | 2.2551 | 2.25 | 2.25 | 2.155 | 2.175 | 2.21 | 2.4335 | 2.322 | 1.86 | 2.17 | 2.22 | 2.2229 | 2.211 | 2.05 | 2.1006 | 2.277
| Std Dev. |  | 0.008165 | 0.034441 | 0.023252 | 0.031116 | 0.022859 | 0.10629 | 0.055498 | 0.012634 | 0.040332 | 0.026394 | 0.057817 | 0.048062 | 0.020656 | 0.03559 | 0.035214 | 0.017512 | 0.035664 | 0.039056 | 0.070711 | 0.013292 | 0.03266 | 0.036003 | 0.048743 | 0.058952 | 0.049851 | 0.097438
| Rel.Std.Dev. |  | 0.3717% | 1.527% | 1.002% | 1.385% | 1.028% | 4.306% | 2.546% | 0.5379% | 1.756% | 1.232% | 2.566% | 2.141% | 0.9194% | 1.661% | 1.615% | 0.7948% | 1.466% | 1.675% | 3.741% | 0.612% | 1.473% | 1.608% | 2.191% | 2.865% | 2.391% | 4.286%
| PDM3 |  | -1.096% | 1.568% | 4.487% | 1.186% | 0.1058% | 11.14% | -1.846% | 5.761% | 3.407% | -3.572% | 1.446% | 1.08% | 1.155% | -3.497% | -1.846% | -0.7957% | 9.514% | 5.005% | -14.9% | -2.222% | -0.1954% | 0.8199% | 0.1573% | -7.359% | -6.122% | 2.348%
| Z-Score (Absolute) |  | 0.46 | 0.31 | 1.16 | 0.2 | 0.11 | 3.08 | 0.67 | 1.53 | 0.85 | 1.17 | 0.28 | 0.17 | 0.19 | 1.15 | 0.67 | 0.37 | 2.61 | 1.31 | 4.45 | 0.78 | 0.2 | 0.1 | 0.1 | 2.27 | 1.91 | 0.54
| Analytical results for Nb in OREAS 609c (Certified Value 16.5 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 12 | 16.79 | 15 | 17.1 | 17.2492 | 16 | 16.2 | 16.8 | 14.9 | 17.19 | 18 | 17 | 16.8 | 15.5 | 17.6 | 10.81 | 18 | 18.3 | 15.7 | 13.3247 | 11 | 15.0754 | 15.8 |  |  |  | 1
| 1 | 2 | 12 | 16.66 | 14.4 | 16.99 | 17.3126 | 16.7 | 16.4 | 16.1 | 14.6 | 16.92 | 17.4 | 17.2 | 15.7 | 15.8 | 17.4 | 11.26 | 18.3 | 19.2 | 15.3 | 15.8309 | 11 | 15.6313 | 15.8 |  |  |  | 30
| 1 | 3 | 12 | 15.59 | 15 | 17.29 | 17.4419 | 16.5 | 17.1 | 16.3 | 14.6 | 16.8 | 18 | 17.4 | 16 | 15.9 | 17.3 | 8.57 | 18 | 18.5 | 16.8 | 12.5069 | 11 | 15.578 | 14.6 |  |  |  | 16
| 1 | 4 | 12 | 16.98 | 15.4 | 16.99 | 17.0342 | 15.9 | 16.5 | 16.5 | 14.7 | 17.44 | 17.2 | 17.1 | 16.6 | 15.8 | 16.9 | 12.17 | 17.9 | 19.6 | 15 | 12.8705 | 10 | 15.3026 | 16 |  |  |  | 16.4747
| 1 | 5 | 12 | 15.79 | 14.6 | 17.04 | 17.549 | 16.8 | 16.4 | 15.8 | 14.4 | 16.82 | 16.9 | 17.2 | 17.2 | 16.1 | 17.6 | 13.81 | 18.5 | 19.9 | 16.1 | 14.3934 | 11 | 14.967 | 16 |  |  |  | 43
| 1 | 6 | 12 | 16.75 | 14.8 | 17.26 | 17.6041 | 17.2 | 17 | 16.1 | 14.7 | 17.07 | 16.3 | 17.4 | 16.6 | 15.2 | 16.9 | 10.21 | 18.2 | 19.7 | 16.6 | 14.8946 | 11 | 15.7325 | 15.1
| Mean |  | 12 | 16.4267 | 14.8667 | 17.1117 | 17.3652 | 16.5167 | 16.6 | 16.2667 | 14.65 | 17.04 | 17.3 | 17.2167 | 16.4833 | 15.7167 | 17.2833 | 11.1383 | 18.15 | 19.2 | 15.9167 | 13.9702 | 10.8333 | 15.3811 | 15.55
| Median |  | 12 | 16.705 | 14.9 | 17.07 | 17.3773 | 16.6 | 16.45 | 16.2 | 14.65 | 16.995 | 17.3 | 17.2 | 16.6 | 15.8 | 17.35 | 11.035 | 18.1 | 19.4 | 15.9 | 13.859 | 11 | 15.4403 | 15.8
| Std Dev. |  | 0 | 0.58353 | 0.35024 | 0.13318 | 0.21093 | 0.49565 | 0.36332 | 0.35024 | 0.16432 | 0.2465 | 0.65727 | 0.16021 | 0.54559 | 0.31885 | 0.31885 | 1.7767 | 0.22583 | 0.66332 | 0.71391 | 1.2855 | 0.40825 | 0.31495 | 0.57184
| Rel.Std.Dev. |  | 0% | 3.552% | 2.356% | 0.7783% | 1.215% | 3.001% | 2.189% | 2.153% | 1.122% | 1.447% | 3.799% | 0.9305% | 3.31% | 2.029% | 1.845% | 15.95% | 1.244% | 3.455% | 4.485% | 9.202% | 3.768% | 2.048% | 3.677%
| PDM3 |  | -27.16% | -0.2916% | -9.761% | 3.866% | 5.405% | 0.2547% | 0.7605% | -1.263% | -11.08% | 3.431% | 5.009% | 4.504% | 0.05236% | -4.601% | 4.908% | -32.39% | 10.17% | 16.54% | -3.387% | -15.2% | -34.24% | -6.638% | -5.613%
| Z-Score (Absolute) |  | 3.59 | 0.02 | 1.28 | 0.53 | 0.73 | 0.05 | 0.12 | 0.15 | 1.45 | 0.47 | 0.68 | 0.61 | 0.02 | 0.6 | 0.67 | 4.28 | 1.37 | 2.21 | 0.43 | 2 | NA | 0.87 | 0.73
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 609c (Certified Value 34.6 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  | 1
2
| 1 | 1 | 2.01: 35; 2.03: 36.37; 2.04: 33.8627; 2.05: 33.7; 2.06: 35.524; 2.07: 37.63; 2.08: 34; 2.10: 36.7; 2.11: 28.5867; 2.12: 38; 2.13: 30.9 |  |  |  | 1
| 1 | 2 | 2.01: 33.2; 2.03: 36.29; 2.04: 34.3048; 2.05: 34; 2.06: 35.814; 2.07: 37.33; 2.08: 31.6; 2.10: 36.8; 2.11: 31.3527; 2.12: 38; 2.13: 30.9 |  |  |  | 31
| 1 | 3 | 2.01: 33.4; 2.03: 36.42; 2.04: 33.5024; 2.05: 35.8; 2.06: 35.426; 2.07: 38.23; 2.08: 31.3; 2.10: 36.2; 2.11: 30.1638; 2.12: 38; 2.13: 32.5 |  |  |  | 16
| 1 | 4 | 2.01: 35.1; 2.03: 36.44; 2.04: 34.1503; 2.05: 33.3; 2.06: 35.899; 2.07: 38.99; 2.08: 32.5; 2.10: 35.8; 2.11: 32.1919; 2.12: 39; 2.13: 32.7 |  |  |  | 34.627
| 1 | 5 | 2.01: 33.5; 2.03: 36.49; 2.04: 34.3783; 2.05: 33.7; 2.06: 35.814; 2.07: 37.45; 2.08: 34.3; 2.10: 37; 2.11: 30.4372; 2.12: 38; 2.13: 29.9 |  |  |  | 44
| 1 | 6 | 2.01: 34.7; 2.03: 36.45; 2.04: 33.742; 2.05: 34.1; 2.06: 36.683; 2.07: 35.43; 2.08: 30.6; 2.10: 37.2; 2.11: 30.5692; 2.12: 37; 2.13: 33.1
| Mean |  | 2.01: 34.15; 2.03: 36.41; 2.04: 33.9901; 2.05: 34.1; 2.06: 35.86; 2.07: 37.51; 2.08: 32.3833; 2.10: 36.6167; 2.11: 30.5502; 2.12: 38; 2.13: 31.6667
| Median |  | 2.01: 34.1; 2.03: 36.43; 2.04: 34.0065; 2.05: 33.85; 2.06: 35.814; 2.07: 37.54; 2.08: 32.05; 2.10: 36.75; 2.11: 30.5032; 2.12: 38; 2.13: 31.7
| Std Dev. |  | 2.01: 0.8735; 2.03: 0.070711; 2.04: 0.34383; 2.05: 0.87864; 2.06: 0.4437; 2.07: 1.1904; 2.08: 1.5012; 2.10: 0.52313; 2.11: 1.213; 2.12: 0.63246; 2.13: 1.2738
| Rel.Std.Dev. |  | 2.01: 2.558%; 2.03: 0.1942%; 2.04: 1.012%; 2.05: 2.577%; 2.06: 1.237%; 2.07: 3.174%; 2.08: 4.636%; 2.10: 1.429%; 2.11: 3.971%; 2.12: 1.664%; 2.13: 4.023%
| PDM3 |  | 2.01: -1.378%; 2.03: 5.149%; 2.04: -1.839%; 2.05: -1.522%; 2.06: 3.561%; 2.07: 8.326%; 2.08: -6.48%; 2.10: 5.746%; 2.11: -11.77%; 2.12: 9.741%; 2.13: -8.549%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.67; 2.04: 0.05; 2.05: 0.01; 2.06: 0.51; 2.07: 1; 2.08: 0.53; 2.10: 0.74; 2.11: 1.07; 2.12: 1.15; 2.13: 0.74
| Analytical results for Ni in OREAS 609c (Certified Value 15.6 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 17.8 | 17 | 17 | 15.9 | 15.252 | 17 | 14.6 | 15.5 | 16 | 16.7 | 17.1 | 15.7 | 15.4 | 15 | 15.9 | 14.8 | 19 | 15.6 | 14.38 | 15.8 | 15.657 | 16 | 15.548 | 13.036 | 14 |  | 1
| 1 | 2 | 18 | 16 | 15.8 | 15.6 | 15.1047 | 16 | 15.5 | 14.7 | 16 | 16.5 | 17.5 | 15.9 | 13.8 | 15.8 | 16.4 | 14.8 | 19 | 16.1 | 14.87 | 16.3 | 13.8322 | 15 | 15.0595 | 13.008 | 15 |  | 17
| 1 | 3 | 17.9 | 15.3 | 16.8 | 15.9 | 15.3654 | 15 | 15.7 | 15.1 | 16 | 16.8 | 17 | 15.9 | 14.3 | 15.2 | 16.2 | 14.8 | 19 | 15.4 | 15.19 | 16 | 13.6922 | 15 | 15.5086 | 13.267 | 15 |  | 16
| 1 | 4 | 16.7 | 16.5 | 16.6 | 15.8 | 15.2586 | 17 | 15.4 | 15.1 | 16 | 17 | 16.2 | 15.5 | 15 | 15.8 | 16 | 14.4 | 19 | 14.8 | 14.55 | 15.8 | 13.0025 | 17 | 15.0891 | 14.171 | 15 |  | 15.6371
| 1 | 5 | 18.4 | 15.7 | 16.6 | 16.1 | 15.7245 | 16 | 15.5 | 15 | 16 | 17 | 16.8 | 16 | 14.8 | 15.4 | 16 | 14.9 | 19 | 15.7 | 14.69 | 16.5 | 15.1799 | 15 | 15.0899 | 13.008 | 15 |  | 45
| 1 | 6 | 18.2 | 17.1 | 16.8 | 15.6 | 15.493 | 16 | 16 | 15.6 | 16 | 16.8 | 16.3 | 16 | 14.8 | 15.3 | 16 | 14.7 | 18 | 15.6 | 14.23 | 16 | 14.6392 | 15 | 15.2262 | 12.971 | 15
| Mean |  | 17.8333 | 16.2667 | 16.6 | 15.8167 | 15.3664 | 16.1667 | 15.45 | 15.1667 | 16 | 16.8 | 16.8167 | 15.8333 | 14.6833 | 15.4167 | 16.0833 | 14.7333 | 18.8333 | 15.5333 | 14.6517 | 16.0667 | 14.3338 | 15.5 | 15.2535 | 13.2435 | 14.8333
| Median |  | 17.95 | 16.25 | 16.7 | 15.85 | 15.312 | 16 | 15.5 | 15.1 | 16 | 16.8 | 16.9 | 15.9 | 14.8 | 15.35 | 16 | 14.8 | 19 | 15.6 | 14.62 | 16 | 14.2357 | 15 | 15.158 | 13.022 | 15
| Std Dev. |  | 0.59554 | 0.72296 | 0.41952 | 0.19408 | 0.21792 | 0.75277 | 0.46797 | 0.33267 | 0 | 0.18974 | 0.49565 | 0.19664 | 0.56006 | 0.32506 | 0.18348 | 0.17512 | 0.40825 | 0.4274 | 0.34667 | 0.28048 | 0.99956 | 0.83666 | 0.2209 | 0.4667 | 0.40825
| Rel.Std.Dev. |  | 3.339% | 4.444% | 2.527% | 1.227% | 1.418% | 4.656% | 3.029% | 2.193% | 0% | 1.129% | 2.947% | 1.242% | 3.814% | 2.109% | 1.141% | 1.189% | 2.168% | 2.751% | 2.366% | 1.746% | 6.973% | 5.398% | 1.448% | 3.524% | 2.752%
| PDM3 |  | 14.04% | 4.026% | 6.158% | 1.148% | -1.732% | 3.386% | -1.197% | -3.009% | 2.321% | 7.437% | 7.543% | 1.255% | -6.1% | -1.41% | 2.853% | -5.78% | 20.44% | -0.6638% | -6.302% | 2.747% | -8.335% | -0.8769% | -2.453% | -15.31% | -5.14%
| Z-Score (Absolute) |  | 2.34 | 0.77 | 1.11 | 0.32 | 0.13 | NA | 0.04 | 0.32 | NA | 1.31 | 1.32 | 0.34 | 0.81 | 0.07 | 0.59 | 0.76 | NA | 0.04 | 0.84 | 0.57 | 1.16 | NA | 0.24 | 2.25 | NA
Results from laboratories 2.07, 2.11, 2.19, 2.25 and 2.28 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 609c (Certified Value 0.029 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.0305 | 0.029 | 0.029 | 0.0296 | 0.029808 | 0.027 | 0.0301 | 0.0306 | 0.029 | 0.0287 | 0.029 | 0.028 | 0.028 | 0.029 | 0.031 | 0.0276 | 0.03 | 0.028 | 0.03 | 0.0286 | < 0.1 | 0.035 | 0.029362 | 0.028151 | 0.028 |  | 1
| 1 | 2 | 0.0304 | 0.029 | 0.028 | 0.0297 | 0.0296 | 0.027 | 0.0322 | 0.0299 | 0.029 | 0.0292 | 0.03 | 0.028 | 0.029 | 0.029 | 0.03 | 0.028 | 0.031 | 0.029 | 0.03 | 0.0285 | < 0.1 | 0.035 | 0.028166 | 0.03067 | 0.029 |  | 0
| 1 | 3 | 0.03 | 0.029 | 0.029 | 0.0297 | 0.029322 | 0.028 | 0.0323 | 0.03 | 0.028 | 0.0296 | 0.029 | 0.028 | 0.03 | 0.029 | 0.031 | 0.0279 | 0.029 | 0.028 | 0.03 | 0.0295 | < 0.1 | 0.035 | 0.029445 | 0.028992 | 0.028 |  | 16
| 1 | 4 | 0.0299 | 0.029 | 0.029 | 0.0295 | 0.029862 | 0.027 | 0.0322 | 0.0299 | 0.029 | 0.03 | 0.028 | 0.028 | 0.029 | 0.029 | 0.031 | 0.027 | 0.028 | 0.038 | 0.03 | 0.0285 | < 0.1 | 0.035 | 0.027547 | 0.029839 | 0.036 |  | 0.029076
| 1 | 5 | 0.0273 | 0.029 | 0.029 | 0.0295 | 0.030241 | 0.027 | 0.0313 | 0.0297 | 0.029 | 0.0289 | 0.03 | 0.028 | 0.029 | 0.029 | 0.031 | 0.0272 | 0.029 | 0.029 | 0.03 | 0.0294 | < 0.1 | 0.035 | 0.028058 | 0.029666 | 0.032 |  | 46
| 1 | 6 | 0.0308 | 0.028 | 0.03 | 0.0293 | 0.030046 | 0.028 | 0.0326 | 0.0293 | 0.029 | 0.0298 | 0.029 | 0.028 | 0.029 | 0.028 | 0.031 | 0.0276 | 0.028 | 0.026 | 0.03 | 0.0289 | < 0.1 | 0.035 | 0.028592 | 0.02997 | 0.029
| Mean |  | 0.029817 | 0.028833 | 0.029 | 0.02955 | 0.029813 | 0.027333 | 0.031783 | 0.0299 | 0.028833 | 0.029367 | 0.029167 | 0.028 | 0.029 | 0.028833 | 0.030833 | 0.02755 | 0.029167 | 0.029667 | 0.03 | 0.0289 |  | 0.035 | 0.028528 | 0.029548 | 0.030333
| Median |  | 0.0302 | 0.029 | 0.029 | 0.02955 | 0.029835 | 0.027 | 0.0322 | 0.0299 | 0.029 | 0.0294 | 0.029 | 0.028 | 0.029 | 0.029 | 0.031 | 0.0276 | 0.029 | 0.0285 | 0.03 | 0.02875 |  | 0.035 | 0.028379 | 0.029752 | 0.029
| Std Dev. |  | 0.0012766 | 0.00040825 | 0.00063246 | 0.00015166 | 0.00032439 | 0.0005164 | 0.00093256 | 0.00042426 | 0.00040825 | 0.0005164 | 0.00075277 | 0 | 0.00063246 | 0.00040825 | 0.00040825 | 0.00038859 | 0.001169 | 0.0042269 | 0 | 0.00045166 |  | 0 | 0.00075547 | 0.00087114 | 0.0031411
| Rel.Std.Dev. |  | 4.281% | 1.416% | 2.181% | 0.5132% | 1.088% | 1.889% | 2.934% | 1.419% | 1.416% | 1.758% | 2.581% | 0% | 2.181% | 1.416% | 1.324% | 1.41% | 4.008% | 14.25% | 0% | 1.563% |  | 0% | 2.648% | 2.948% | 10.36%
| PDM3 |  | 2.547% | -0.8352% | -0.262% | 1.63% | 2.535% | -5.994% | 9.311% | 2.833% | -0.8352% | 0.999% | 0.3112% | -3.701% | -0.262% | -0.8352% | 6.043% | -5.249% | 0.3112% | 2.031% | 3.177% | -0.6059% |  | 20.37% | -1.884% | 1.622% | 4.324%
| Z-Score (Absolute) |  | 0.57 | 0.67 | 0.46 | 0.23 | 0.56 | 2.57 | 3.06 | 0.67 | 0.67 | 0 | 0.25 | 1.73 | 0.46 | 0.67 | 1.85 | 2.3 | 0.25 | 0.38 | 0.8 | 0.59 | 26.09 | 7.01 | 1.06 | 0.23 | 1.22
| Analytical results for Pb in OREAS 609c (Certified Value 366 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 400 | 360.64 | 360.8 | 391 | 365.2273 | 363 | 341 | 380.863 | 353 | 393 | 374 | 373 | 376 | 356 | 374 | 381 | 401 | 343 | 571 | 323 | 352.3 | 320.1546 | 377 | 365.7513 | 327.252 | 267 | 1
| 1 | 2 | 406 | 369.16 | 358.2 | 386.7 | 366.4482 | 358 | 350 | 383.78 | 352 | 384 | 377 | 379 | 378 | 354 | 385 | 382 | 407 | 338 | 546 | 325.4 | 360.7 | 308.8937 | 388 | 360.1808 | 329.64 | 251 | 19
| 1 | 3 | 399 | 361.84 | 378.7 | 389.4 | 364.4271 | 370 | 364 | 384.78 | 358 | 387 | 382 | 373 | 381 | 368 | 369 | 385 | 397 | 340 | 544 | 327.8 | 352.9 | 312.8986 | 378 | 366.8118 | 324.518 | 261 | 16
| 1 | 4 | 398 | 370.16 | 364.8 | 384.7 | 361.9239 | 355 | 362 | 387.354 | 356 | 393 | 391 | 359 | 371 | 365 | 381 | 390 | 391 | 337 | 535 | 325.7 | 356.9 | 309.4785 | 401 | 359.4511 | 332.942 | 247 | 366.0398
| 1 | 5 | 411 | 365.2 | 362.9 | 385.1 | 368.8901 | 405 | 351 | 381.798 | 354 | 394 | 372 | 378 | 375 | 362 | 377 | 390 | 398 | 304 | 564 | 318.4 | 355 | 320.1352 | 392 | 365.2198 | 325.891 | 317 | 47
| 1 | 6 | 400 | 362.88 | 378.6 | 384.3 | 367.9336 | 364 | 367 | 389.455 | 353 | 393 | 382 | 363 | 373 | 360 | 367 | 380 | 399 | 320 | 528 | 319.5 | 358.8 | 310.8765 | 368 | 357.5635 | 333.106 | 264
| Mean |  | 402.3333 | 364.98 | 367.3333 | 386.8667 | 365.8084 | 369.1667 | 355.8333 | 384.6717 | 354.3333 | 390.6667 | 379.6667 | 370.8333 | 375.6667 | 360.8333 | 375.5 | 384.6667 | 398.8333 | 330.3333 | 548 | 323.3 | 356.1 | 313.7395 | 384 | 362.4964 | 328.8915 | 267.8333
| Median |  | 400 | 364.04 | 363.85 | 385.9 | 365.8378 | 363.5 | 356.5 | 384.28 | 353.5 | 393 | 379.5 | 373 | 375.5 | 361 | 375.5 | 383.5 | 398.5 | 337.5 | 545 | 324.2 | 355.95 | 311.8875 | 383 | 362.7003 | 328.446 | 262.5
| Std Dev. |  | 5.0859 | 3.9359 | 9.0367 | 2.7544 | 2.5202 | 18.3021 | 10.0681 | 3.2755 | 2.2509 | 4.1312 | 6.8896 | 8.1097 | 3.559 | 5.3072 | 6.921 | 4.4572 | 5.2313 | 15.214 | 16.5771 | 3.7138 | 3.316 | 5.1498 | 11.9164 | 3.8885 | 3.6207 | 25.2857
| Rel.Std.Dev. |  | 1.264% | 1.078% | 2.46% | 0.712% | 0.6889% | 4.958% | 2.829% | 0.8515% | 0.6353% | 1.057% | 1.815% | 2.187% | 0.9474% | 1.471% | 1.843% | 1.159% | 1.312% | 4.606% | 3.025% | 1.149% | 0.9312% | 1.641% | 3.103% | 1.073% | 1.101% | 9.441%
| PDM3 |  | 9.915% | -0.2895% | 0.3534% | 5.69% | -0.06321% | 0.8543% | -2.788% | 5.09% | -3.198% | 6.728% | 3.723% | 1.31% | 2.63% | -1.422% | 2.584% | 5.089% | 8.959% | -9.755% | 49.71% | -11.68% | -2.715% | -14.29% | 4.907% | -0.968% | -10.15% | -26.83%
| Z-Score (Absolute) |  | 1.55 | 0.15 | 0.04 | 0.85 | 0.11 | 0.04 | 0.56 | 0.75 | 0.63 | 1.02 | 0.52 | 0.12 | 0.34 | 0.34 | 0.33 | 0.75 | 1.39 | 1.72 | 8.17 | 2.04 | 0.55 | 2.48 | 0.72 | 0.26 | 1.79 | 4.56
| Analytical results for Pr in OREAS 609c (Certified Value 9.24 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 8.8; 2.03: 9.45; 2.04: 9.7664; 2.05: 8.7; 2.06: 9.7692; 2.07: 10.13; 2.08: 10.54; 2.09: 9.43; 2.10: 7.7326; 2.11: 8; 2.12: 8.44 | 1
| 1 | 2 | 2.01: 8.6; 2.03: 9.43; 2.04: 9.5912; 2.05: 8.7; 2.06: 9.6023; 2.07: 10.1; 2.08: 9.55; 2.09: 9.52; 2.10: 8.5354; 2.11: 8; 2.12: 8.3 | 34
| 1 | 3 | 2.01: 8.6; 2.03: 9.35; 2.04: 9.6443; 2.05: 9.1; 2.06: 9.6765; 2.07: 10.29; 2.08: 9.53; 2.09: 9.33; 2.10: 7.9009; 2.11: 8; 2.12: 8.76 | 16
| 1 | 4 | 2.01: 8.9; 2.03: 9.42; 2.04: 9.7675; 2.05: 8.6; 2.06: 9.8134; 2.07: 10.47; 2.08: 10.1; 2.09: 9.3; 2.10: 8.873; 2.11: 8; 2.12: 8.73 | 9.2438
| 1 | 5 | 2.01: 8.5; 2.03: 9.43; 2.04: 9.8942; 2.05: 8.6; 2.06: 9.6834; 2.07: 9.94; 2.08: 10.32; 2.09: 9.41; 2.10: 8.3543; 2.11: 8; 2.12: 8.02 | 48
| 1 | 6 | 2.01: 8.7; 2.03: 9.45; 2.04: 9.6302; 2.05: 8.8; 2.06: 9.8172; 2.07: 9.51; 2.08: 9.04; 2.09: 9.43; 2.10: 8.6439; 2.11: 7; 2.12: 9.03
| Mean |  | 2.01: 8.6833; 2.03: 9.4217; 2.04: 9.7156; 2.05: 8.75; 2.06: 9.727; 2.07: 10.0733; 2.08: 9.8467; 2.09: 9.4033; 2.10: 8.34; 2.11: 7.8333; 2.12: 8.5467
| Median |  | 2.01: 8.65; 2.03: 9.43; 2.04: 9.7053; 2.05: 8.7; 2.06: 9.7263; 2.07: 10.115; 2.08: 9.825; 2.09: 9.42; 2.10: 8.4449; 2.11: 8; 2.12: 8.585
| Std Dev. |  | 2.01: 0.1472; 2.03: 0.037103; 2.04: 0.11402; 2.05: 0.18708; 2.06: 0.086467; 2.07: 0.32965; 2.08: 0.56709; 2.09: 0.078909; 2.10: 0.44194; 2.11: 0.40825; 2.12: 0.36396
| Rel.Std.Dev. |  | 2.01: 1.695%; 2.03: 0.3938%; 2.04: 1.174%; 2.05: 2.138%; 2.06: 0.8889%; 2.07: 3.272%; 2.08: 5.759%; 2.09: 0.8392%; 2.10: 5.299%; 2.11: 5.212%; 2.12: 4.258%
| PDM3 |  | 2.01: -6.063%; 2.03: 1.925%; 2.04: 5.105%; 2.05: -5.342%; 2.06: 5.228%; 2.07: 8.974%; 2.08: 6.522%; 2.09: 1.726%; 2.10: -9.777%; 2.11: -15.26%; 2.12: -7.541%
| Z-Score (Absolute) |  | 2.01: 0.9; 2.03: 0.01; 2.04: 0.37; 2.05: 0.82; 2.06: 0.39; 2.07: 0.81; 2.08: 0.53; 2.09: 0.01; 2.10: 1.32; 2.11: NA; 2.12: 1.07
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 609c (Certified Value 125 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 132 | 106.3 | 125.2 | 123.91 | 123.2664 | 123 | 126 | 133.938 | 126 | 118 | 133.13 | 128.5 | 121.5 | 122.5 | 119.5 | 130.5 | 130.7 | 141 | 110 | 123.9 | 127.4229 | 120 |  |  |  |  | 1
| 1 | 2 | 128 | 119.5 | 117 | 123.09 | 126.7271 | 122 | 129 | 133.2572 | 125 | 120 | 131.66 | 123.5 | 123.5 | 114.5 | 124 | 134.5 | 130.2 | 141 | 115 | 125.7 | 124.8713 | 116 |  |  |  |  | 35
| 1 | 3 | 134 | 121.9 | 127 | 125.94 | 125.1153 | 124 | 131 | 133.1912 | 126 | 119 | 133.18 | 128.5 | 122 | 115 | 120.5 | 131.5 | 129.2 | 137 | 117 | 122.9 | 132.9789 | 121 |  |  |  |  | 16
| 1 | 4 | 130 | 126.1 | 122 | 122.53 | 127.8387 | 124 | 133 | 133.4092 | 125 | 121 | 131.91 | 122.5 | 120 | 121 | 125.5 | 138.5 | 126 | 142 | 125 | 123.1 | 125.9107 | 124 |  |  |  |  | 125.344
| 1 | 5 | 134 | 113 | 125.8 | 124.64 | 124.6986 | 124 | 133 | 133.548 | 119 | 120 | 135.6 | 130.5 | 122.5 | 121.5 | 122.5 | 133.5 | 127.8 | 145 | 121 | 124.6 | 126.1092 | 116 |  |  |  |  | 49
| 1 | 6 | 135 | 125.2 | 125.3 | 125.13 | 125.5464 | 130 | 125 | 133.2684 | 125 | 117 | 127.42 | 120 | 123.5 | 119 | 120 | 131 | 128.7 | 139 | 121 | 124.2 | 131.6095 | 123
| Mean |  | 132.1667 | 118.6667 | 123.7167 | 124.2067 | 125.5321 | 124.5 | 129.5 | 133.4353 | 124.3333 | 119.1667 | 132.15 | 125.5833 | 122.1667 | 118.9167 | 122 | 133.25 | 128.7667 | 140.8333 | 118.1667 | 124.0667 | 128.1504 | 120
| Median |  | 133 | 120.7 | 125.25 | 124.275 | 125.3309 | 124 | 130 | 133.3388 | 125 | 119.5 | 132.52 | 126 | 122.25 | 120 | 121.5 | 132.5 | 128.95 | 141 | 119 | 124.05 | 126.766 | 120.5
| Std Dev. |  | 2.7142 | 7.6699 | 3.6848 | 1.2799 | 1.5969 | 2.8107 | 3.4496 | 0.27758 | 2.6583 | 1.472 | 2.7053 | 4.1523 | 1.3292 | 3.4266 | 2.4083 | 2.9958 | 1.7072 | 2.7142 | 5.3072 | 1.0289 | 3.339 | 3.4059
| Rel.Std.Dev. |  | 2.054% | 6.463% | 2.978% | 1.03% | 1.272% | 2.258% | 2.664% | 0.208% | 2.138% | 1.235% | 2.047% | 3.306% | 1.088% | 2.882% | 1.974% | 2.248% | 1.326% | 1.927% | 4.491% | 0.8293% | 2.606% | 2.838%
| PDM3 |  | 5.443% | -5.327% | -1.298% | -0.9074% | 0.15% | -0.6734% | 3.316% | 6.455% | -0.8063% | -4.928% | 5.43% | 0.1909% | -2.535% | -5.128% | -2.668% | 6.307% | 2.731% | 12.36% | -5.726% | -1.019% | 2.239% | -4.263%
| Z-Score (Absolute) |  | 1.19 | 0.88 | 0.11 | 0.03 | 0.17 | 0.01 | 0.78 | 1.39 | 0.01 | 0.81 | 1.19 | 0.18 | 0.35 | 0.85 | 0.37 | 1.36 | 0.67 | 2.53 | 0.96 | 0.05 | 0.57 | 0.68
| Analytical results for Re in OREAS 609c (Certified Value 0.029 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 3
3
| 1 | 1 | < 0.002 | 0.318 | 0.028 | 0.024 | 0.032944 | 0.026 | 0.03 | < 0.05 | 0.032 | 0.031 | 0.032 | 0.03 | 0.03 | 0.031 | 0.007 | < 0.1 | 0.03 | 0.008 | < 0.05 | < 0.1 | [6 empty] | 1
| 1 | 2 | 0.033 | 0.266 | 0.026 | 0.026 | 0.031735 | 0.027 | 0.029 | < 0.05 | 0.025 | 0.028 | 0.028 | 0.029 | 0.026 | 0.032 | 0.007 | < 0.1 | 0.025 | 0.008 | < 0.05 | < 0.1 | [6 empty] | 36
| 1 | 3 | 0.056 | 0.229 | 0.03 | 0.028 | 0.031495 | 0.025 | 0.029 | < 0.05 | 0.029 | 0.029 | 0.032 | 0.028 | 0.028 | 0.031 | 0.006 | < 0.1 | 0.03 | 0.009 | < 0.05 | < 0.1 | [6 empty] | 16
| 1 | 4 | 0.032 | 0.253 | 0.026 | 0.029 | 0.029527 | 0.027 | 0.029 | < 0.05 | 0.03 | 0.027 | 0.029 | 0.036 | 0.028 | 0.033 | 0.006 | < 0.1 | 0.029 | 0.008 | < 0.05 | < 0.1 | [6 empty] | 0.029018
| 1 | 5 | 0.046 | 0.256 | 0.028 | 0.03 | 0.030928 | 0.026 | 0.028 | < 0.05 | 0.028 | 0.031 | 0.026 | 0.032 | 0.03 | 0.033 | 0.007 | < 0.1 | 0.026 | 0.008 | < 0.05 | < 0.1 | [6 empty] | 50
| 1 | 6 | 0.011 | 0.31 | 0.03 | 0.028 | 0.033885 | 0.027 | 0.03 | < 0.05 | 0.024 | 0.03 | 0.032 | 0.03 | 0.027 | 0.032 | 0.007 | < 0.1 | 0.03 | 0.01 | < 0.05 | < 0.1
| Mean |  | 0.0356 | 0.272 | 0.028 | 0.0275 | 0.031752 | 0.026333 | 0.029167 |  | 0.028 | 0.029333 | 0.029833 | 0.030833 | 0.028167 | 0.032 | 0.0066667 |  | 0.028333 | 0.0085
| Median |  | 0.033 | 0.261 | 0.028 | 0.028 | 0.031615 | 0.0265 | 0.029 |  | 0.0285 | 0.0295 | 0.0305 | 0.03 | 0.028 | 0.032 | 0.007 |  | 0.0295 | 0.008
| Std Dev. |  | 0.01695 | 0.03482 | 0.0017889 | 0.0021679 | 0.0015274 | 0.0008165 | 0.00075277 |  | 0.0030332 | 0.001633 | 0.0025626 | 0.0028577 | 0.0016021 | 0.00089443 | 0.0005164 |  | 0.0022509 | 0.00083666
| Rel.Std.Dev. |  | 47.61% | 12.8% | 6.389% | 7.883% | 4.811% | 3.101% | 2.581% |  | 10.83% | 5.567% | 8.59% | 9.268% | 5.688% | 2.795% | 7.746% |  | 7.944% | 9.843%
| PDM3 |  | 22.68% | 837.3% | -3.509% | -5.232% | 9.422% | -9.252% | 0.5115% |  | -3.509% | 1.086% | 2.809% | 6.255% | -2.935% | 10.28% | -77.03% |  | -2.36% | -70.71%
| Z-Score (Absolute) |  | 0.32 | 72.9 | 0.22 | 0.37 | 0.9 | 0.72 | 0.12 | 1.12 | 0.22 | 0.17 | 0.32 | 0.62 | 0.17 | 0.97 | 6.62 | 6.37 | 0.12 | 6.07 | 1.12 | 6.37
| Analytical results for S in OREAS 609c (Certified Value 1.58 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.59 | 1.51 | 1.5 | 1.5934 | 1.6105 | 1.45 | 1.53 | 1.5008 | 1.63 | 1.65 | 1.5688 | 1.59 | 1.6 | 1.5 | 1.56 | 1.68 | 1.4868 | 1.635 | 1.63 | 1.17 | 1.55 | 1.5314 | 1.63 | 1.566 | 1.542 | 1.649 | 1
| 1 | 2 | 1.62 | 1.53 | 1.4 | 1.6084 | 1.6277 | 1.43 | 1.62 | 1.5956 | 1.6 | 1.53 | 1.5825 | 1.61 | 1.59 | 1.52 | 1.6 | 1.68 | 1.5001 | 1.622 | 1.62 | 1.15 | 1.62 | 1.5025 | 1.63 | 1.5201 | 1.551 | 1.575 | 12
| 1 | 3 | 1.63 | 1.51 | 1.5 | 1.6201 | 1.6006 | 1.45 | 1.63 | 1.5427 | 1.6 | 1.63 | 1.6012 | 1.58 | 1.59 | 1.55 | 1.54 | 1.69 | 1.5039 | 1.601 | 1.61 | 1.17 | 1.57 | 1.4925 | 1.61 | 1.5629 | 1.528 | 1.594 | 16
| 1 | 4 | 1.6 | 1.55 | 1.5 | 1.5915 | 1.6005 | 1.43 | 1.61 | 1.5452 | 1.61 | 1.64 | 1.6428 | 1.53 | 1.57 | 1.53 | 1.6 | 1.7 | 1.4458 | 1.628 | 1.65 | 1.17 | 1.62 | 1.4949 | 1.7 | 1.5171 | 1.536 | 1.729 | 1.5773
| 1 | 5 | 1.63 | 1.52 | 1.5 | 1.628 | 1.6212 | 1.43 | 1.58 | 1.5982 | 1.59 | 1.69 | 1.568 | 1.61 | 1.6 | 1.55 | 1.56 | 1.72 | 1.4462 | 1.656 | 1.63 | 1.17 | 1.56 | 1.5687 | 1.66 | 1.5054 | 1.558 | 1.581 | 51
| 1 | 6 | 1.61 | 1.51 | 1.5 | 1.5856 | 1.6105 | 1.49 | 1.65 | 1.551 | 1.61 | 1.63 | 1.604 | 1.56 | 1.58 | 1.52 | 1.54 | 1.69 | 1.465 | 1.606 | 1.65 | 1.15 | 1.64 | 1.5186 | 1.57 | 1.533 | 1.552 | 1.671
| Mean |  | 1.6133 | 1.5217 | 1.4833 | 1.6045 | 1.6118 | 1.4467 | 1.6033 | 1.5556 | 1.6067 | 1.6283 | 1.5945 | 1.58 | 1.5883 | 1.5283 | 1.5667 | 1.6933 | 1.4746 | 1.6247 | 1.6317 | 1.1633 | 1.5933 | 1.5181 | 1.6333 | 1.5341 | 1.5445 | 1.6332
| Median |  | 1.615 | 1.515 | 1.5 | 1.6009 | 1.6105 | 1.44 | 1.615 | 1.5481 | 1.605 | 1.635 | 1.5918 | 1.585 | 1.59 | 1.525 | 1.56 | 1.69 | 1.4759 | 1.625 | 1.63 | 1.17 | 1.595 | 1.5106 | 1.63 | 1.5266 | 1.5465 | 1.6215
| Std Dev. |  | 0.01633 | 0.016021 | 0.040825 | 0.017089 | 0.010932 | 0.023381 | 0.04274 | 0.036661 | 0.013663 | 0.053072 | 0.028195 | 0.030984 | 0.01169 | 0.019408 | 0.027325 | 0.015055 | 0.026022 | 0.020077 | 0.016021 | 0.010328 | 0.037771 | 0.028894 | 0.044121 | 0.025124 | 0.011238 | 0.060836
| Rel.Std.Dev. |  | 1.012% | 1.053% | 2.752% | 1.065% | 0.6782% | 1.616% | 2.666% | 2.357% | 0.8504% | 3.259% | 1.768% | 1.961% | 0.736% | 1.27% | 1.744% | 0.8891% | 1.765% | 1.236% | 0.9819% | 0.8878% | 2.371% | 1.903% | 2.701% | 1.638% | 0.7276% | 3.725%
| PDM3 |  | 2.281% | -3.53% | -5.96% | 1.721% | 2.187% | -8.285% | 1.648% | -1.379% | 1.859% | 3.232% | 1.091% | 0.1682% | 0.6965% | -3.107% | -0.6771% | 7.353% | -6.512% | 3% | 3.444% | -26.25% | 1.014% | -3.756% | 3.549% | -2.741% | -2.082% | 3.539%
| Z-Score (Absolute) |  | 0.39 | 1.19 | 1.85 | 0.24 | 0.36 | 2.48 | 0.22 | 0.61 | 0.27 | 0.65 | 0.06 | 0.19 | 0.04 | 1.08 | 0.42 | 1.76 | 2 | 0.58 | 0.7 | 7.36 | 0.04 | 1.25 | 0.73 | 0.98 | 0.8 | 0.73
| Analytical results for Sb in OREAS 609c (Certified Value 152 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 158 | 152.41 | 92.2 | 161.51 | 161.2706 | 124 | 160 | 157 | 146 | 159 | 161.5 | 163 | 148 | 160 | 168 | 146.91 | 154 | 143 | 153.9 | 145 | 143.1275 | 118.4204 | 147.18 | 120 |  |  | 1
| 1 | 2 | 160 | 151.16 | 102.3 | 159.58 | 160.483 | 151 | 170 | 152 | 149 | 161 | 163 | 164 | 149 | 164.5 | 170.5 | 147.27 | 156 | 141 | 152.9 | 145 | 141.2423 | 142.3099 | 150.61 | 109 |  |  | 37
| 1 | 3 | 157 | 151.63 | 93.7 | 161.92 | 160.5365 | 153 | 177 | 156 | 148 | 163 | 163 | 164 | 154 | 160 | 164 | 143.38 | 154 | 136 | 146.3 | 141 | 141.7994 | 116.5567 | 145.12 | 138 |  |  | 16
| 1 | 4 | 164 | 154.7 | 96.9 | 161.64 | 159.5591 | 129 | 172 | 155 | 152 | 166 | 156 | 160.5 | 156.5 | 160.5 | 172.5 | 141.18 | 152 | 146 | 150.8 | 144 | 133.5154 | 125.9452 | 148.01 | 116 |  |  | 152.3775
| 1 | 5 | 156 | 151.13 | 98.3 | 162.66 | 159.0018 | 148 | 172 | 151 | 150 | 159 | 167 | 161.5 | 157 | 161.5 | 172.5 | 145.34 | 159 | 140 | 150.1 | 143 | 144.1862 | 118.529 | 145.56 | 152 |  |  | 52
| 1 | 6 | 156 | 149.96 | 93.6 | 161.01 | 160.9611 | 139 | 172 | 154 | 150 | 162 | 158.5 | 161 | 154 | 159.5 | 168.5 | 145.32 | 156 | 143 | 149.3 | 146 | 144.0498 | 142.8418 | 148.6 | 140
| Mean |  | 158.5 | 151.8317 | 96.1667 | 161.3867 | 160.302 | 140.6667 | 170.5 | 154.1667 | 149.1667 | 161.6667 | 161.5 | 162.3333 | 153.0833 | 161 | 169.3333 | 144.9 | 155.1667 | 141.5 | 150.55 | 144 | 141.3201 | 127.4338 | 147.5133 | 129.1667
| Median |  | 157.5 | 151.395 | 95.3 | 161.575 | 160.5098 | 143.5 | 172 | 154.5 | 149.5 | 161.5 | 162.25 | 162.25 | 154 | 160.25 | 169.5 | 145.33 | 155 | 142 | 150.45 | 144.5 | 142.4634 | 122.2371 | 147.595 | 129
| Std Dev. |  | 3.0822 | 1.6155 | 3.7713 | 1.0384 | 0.86058 | 12.0775 | 5.648 | 2.3166 | 2.0412 | 2.6583 | 3.8471 | 1.5384 | 3.7738 | 1.8439 | 3.2352 | 2.288 | 2.4014 | 3.3912 | 2.7068 | 1.7889 | 4.002 | 12.1636 | 2.0333 | 16.6183
| Rel.Std.Dev. |  | 1.945% | 1.064% | 3.922% | 0.6434% | 0.5368% | 8.586% | 3.313% | 1.503% | 1.368% | 1.644% | 2.382% | 0.9477% | 2.465% | 1.145% | 1.911% | 1.579% | 1.548% | 2.397% | 1.798% | 1.242% | 2.832% | 9.545% | 1.378% | 12.87%
| PDM3 |  | 4.018% | -0.3582% | -36.89% | 5.912% | 5.201% | -7.685% | 11.89% | 1.174% | -2.107% | 6.096% | 5.987% | 6.534% | 0.4632% | 5.659% | 11.13% | -4.907% | 1.83% | -7.139% | -1.199% | -5.498% | -7.257% | -16.37% | -3.192% | -15.23%
| Z-Score (Absolute) |  | 0.47 | 0.05 | 4.35 | 0.69 | 0.61 | 0.91 | 1.39 | 0.13 | 0.25 | 0.71 | 0.7 | 0.76 | 0.05 | 0.66 | 1.3 | 0.58 | 0.21 | 0.85 | 0.15 | 0.65 | 0.86 | 1.93 | 0.38 | 1.8
| Analytical results for Sc in OREAS 609c (Certified Value 4.18 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 5.1 | 4.4 | 4 | 3.9 | 3.8165 | 4 | 4.4 | 4.8318 | 4.87 | 4 | 3.7 | 3.9 | 4 | 3.5 | 3.9 | 4.2 | 3.97 | 3.5 | 4.2 | 2.7 | 4.6336 | 3 | 4.4647 | 4.5 |  |  | 1
| 1 | 2 | 4.7 | 4 | 4 | 3.8 | 3.8488 | 4 | 4.5 | 4.8863 | 4.57 | 4 | 3.8 | 3.8 | 4.1 | 3.5 | 4.1 | 4.2 | 3.98 | 3.3 | 4.1 | 2.7 | 4.9657 | 3 | 4.0443 | 4.4 |  |  | 38
| 1 | 3 | 5 | 4.1 | 4 | 4.4 | 3.8164 | 4 | 4.7 | 4.8719 | 4.53 | 4 | 3.8 | 3.9 | 4.1 | 3.4 | 3.9 | 4.2 | 3.96 | 3.5 | 4.1 | 2.6 | 4.8544 | 3 | 4.4873 | 4.8 |  |  | 16
| 1 | 4 | 5 | 4.2 | 4 | 4 | 3.876 | 4 | 4.6 | 4.9292 | 4.67 | 4 | 4 | 3.7 | 3.9 | 3.7 | 4.2 | 4.2 | 3.89 | 3.5 | 4.2 | 2.5 | 4.8437 | 3 | 3.928 | 4.6 |  |  | 4.1816
| 1 | 5 | 4.6 | 4 | 4 | 4.1 | 3.9199 | 4 | 4.6 | 4.8473 | 4.5 | 4 | 3.7 | 3.8 | 4 | 3.7 | 4 | 4.2 | 3.94 | 3.5 | 3.8 | 2.5 | 4.8414 | 3 | 4.23 | 4.1 |  |  | 53
| 1 | 6 | 4.5 | 4.2 | 4 | 3.9 | 3.8484 | 4 | 4.6 | 4.8849 | 4.65 | 4 | 3.9 | 3.7 | 4 | 3.7 | 3.9 | 4.1 | 3.92 | 3.1 | 4 | 2.8 | 4.9648 | 3 | 4.4619 | 4.9
| Mean |  | 4.8167 | 4.15 | 4 | 4.0167 | 3.8544 | 4 | 4.5667 | 4.8752 | 4.6317 | 4 | 3.8167 | 3.8 | 4.0167 | 3.5833 | 4 | 4.1833 | 3.9433 | 3.4 | 4.0667 | 2.6333 | 4.8506 | 3 | 4.2694 | 4.55
| Median |  | 4.85 | 4.15 | 4 | 3.95 | 3.8486 | 4 | 4.6 | 4.8784 | 4.61 | 4 | 3.8 | 3.8 | 4 | 3.6 | 3.95 | 4.2 | 3.95 | 3.5 | 4.1 | 2.65 | 4.8491 | 3 | 4.3459 | 4.55
| Std Dev. |  | 0.24833 | 0.15166 | 0 | 0.2137 | 0.039269 | 0 | 0.10328 | 0.034065 | 0.13423 | 0 | 0.1169 | 0.089443 | 0.075277 | 0.13292 | 0.12649 | 0.040825 | 0.033862 | 0.16733 | 0.15055 | 0.12111 | 0.12126 | 0 | 0.24145 | 0.2881
| Rel.Std.Dev. |  | 5.156% | 3.654% | 0% | 5.32% | 1.019% | 0% | 2.262% | 0.6987% | 2.898% | 0% | 3.063% | 2.354% | 1.874% | 3.709% | 3.162% | 0.9759% | 0.8587% | 4.922% | 3.702% | 4.599% | 2.5% | 0% | 5.655% | 6.332%
| PDM3 |  | 15.19% | -0.7567% | -4.344% | -3.945% | -7.827% | -4.344% | 9.207% | 16.59% | 10.76% | -4.344% | -8.728% | -9.127% | -3.945% | -14.31% | -4.344% | 0.04044% | -5.699% | -18.69% | -2.75% | -37.03% | 16% | -28.26% | 2.098% | 8.809%
| Z-Score (Absolute) |  | 2.23 | 0.31 | NA | 0.07 | 0.54 | NA | 1.51 | 2.4 | 1.7 | NA | 0.65 | 0.69 | 0.07 | 1.32 | 0.12 | 0.41 | 0.28 | 1.84 | 0.07 | 4.05 | 2.32 | NA | 0.65 | 1.46
Results from laboratories 2.04, 2.07, 2.11 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 609c (Certified Value 10.2 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 10 | 7.6 | 10 | 10.8 | 10.8756 | 10.4 | 8.4 | 8.1379 | 10 | 10.5 | 10 | 11 | 10 | 10 | 10 | 10.14 | 10 | 5 | 14 | 9.4683 | 14.286 | 10 |  |  |  |  | 1
| 1 | 2 | 10 | 8.8 | 10 | 10.8 | 10.7373 | 9.9 | 10.1 | 8.4608 | 10 | 10.6 | 10 | 13 | 11 | 10 | 11 | 10.17 | 10 | 7 | 13 | 9.3389 | 11.813 | 10 |  |  |  |  | 39
| 1 | 3 | 10 | 9.1 | 10 | 9.7 | 10.5638 | 10.3 | 10.1 | 8.12 | 9 | 10.2 | 11 | 12 | 11 | 10 | 8 | 9.9 | 10 | 6 | 14 | 9.5135 | 14.021 | 10 |  |  |  |  | 16
| 1 | 4 | 11 | 8.3 | 10 | 10.3 | 10.2205 | 10.2 | 9.8 | 8.105 | 10 | 10.3 | 10 | 12 | 11 | 11 | 12 | 9.72 | 9 | 7 | 13 | 9.5694 | 10.791 | 11 |  |  |  |  | 10.1777
| 1 | 5 | 10 | 9.7 | 10 | 11.1 | 10.9358 | 10 | 9.7 | 8.2911 | 10 | 10.8 | 10 | 12 | 11 | 10 | 11 | 10.02 | 10 | 6 | 13 | 9.3466 | 13.088 | 10 |  |  |  |  | 54
| 1 | 6 | 10 | 10 | 11 | 10 | 10.5961 | 9.9 | 10.2 | 8.1988 | 10 | 10.6 | 10 | 11 | 10 | 9 | 11 | 9.89 | 10 | 7 | 14 | 9.2774 | 13.202 | 10
| Mean |  | 10.1667 | 8.9167 | 10.1667 | 10.45 | 10.6549 | 10.1167 | 9.7167 | 8.2189 | 9.8333 | 10.5 | 10.1667 | 11.8333 | 10.6667 | 10 | 10.5 | 9.9733 | 9.8333 | 6.3333 | 13.5 | 9.419 | 12.8668 | 10.1667
| Median |  | 10 | 8.95 | 10 | 10.55 | 10.6667 | 10.1 | 9.95 | 8.1684 | 10 | 10.55 | 10 | 12 | 11 | 10 | 11 | 9.96 | 10 | 6.5 | 13.5 | 9.4075 | 13.145 | 10
| Std Dev. |  | 0.40825 | 0.88863 | 0.40825 | 0.53944 | 0.25878 | 0.2137 | 0.67355 | 0.13671 | 0.40825 | 0.21909 | 0.40825 | 0.75277 | 0.5164 | 0.63246 | 1.3784 | 0.17037 | 0.40825 | 0.8165 | 0.54772 | 0.11463 | 1.3362 | 0.40825
| Rel.Std.Dev. |  | 4.016% | 9.966% | 4.016% | 5.162% | 2.429% | 2.112% | 6.932% | 1.663% | 4.152% | 2.087% | 4.016% | 6.361% | 4.841% | 6.325% | 13.13% | 1.708% | 4.152% | 12.89% | 4.057% | 1.217% | 10.38% | 4.016%
| PDM3 |  | -0.1088% | -12.39% | -0.1088% | 2.675% | 4.688% | -0.6001% | -4.53% | -19.25% | -3.384% | 3.166% | -0.1088% | 16.27% | 4.804% | -1.746% | 3.166% | -2.008% | -3.384% | -37.77% | 32.64% | -7.455% | 26.42% | -0.1088%
| Z-Score (Absolute) |  | 0 | 2.53 | 0 | 0.57 | 0.99 | 0.1 | 0.91 | 3.94 | 0.67 | 0.67 | 0 | 3.37 | 1.01 | 0.34 | 0.67 | 0.39 | 0.67 | 7.75 | 6.74 | 1.51 | 5.46 | 0
| Analytical results for Sm in OREAS 609c (Certified Value 6.52 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 7.2; 2.03: 6.68; 2.04: 6.4126; 2.05: 6.4; 2.06: 6.592; 2.07: 6.95; 2.08: 6.3; 2.09: 6.6; 2.11: 5.3338; 2.12: 5; 2.13: 6.4 |  |  |  |  | 1
| 1 | 2 | 2.01: 6.9; 2.03: 6.58; 2.04: 6.3193; 2.05: 6.7; 2.06: 6.6527; 2.07: 7.23; 2.08: 5.9; 2.09: 6.67; 2.11: 5.8617; 2.12: 5; 2.13: 6.3 |  |  |  |  | 23
| 1 | 3 | 2.01: 6.8; 2.03: 6.51; 2.04: 6.2782; 2.05: 6.7; 2.06: 6.6528; 2.07: 7.58; 2.08: 5.9; 2.09: 6.56; 2.11: 5.6749; 2.12: 5; 2.13: 6.4 |  |  |  |  | 16
| 1 | 4 | 2.01: 7.2; 2.03: 6.59; 2.04: 6.306; 2.05: 6.6; 2.06: 6.5066; 2.07: 7.28; 2.08: 6.1; 2.09: 6.59; 2.11: 6.0566; 2.12: 6; 2.13: 6.6 |  |  |  |  | 6.5245
| 1 | 5 | 2.01: 6.8; 2.03: 6.56; 2.04: 6.3561; 2.05: 6.6; 2.06: 6.4928; 2.07: 7.41; 2.08: 6.4; 2.09: 6.55; 2.11: 5.6995; 2.12: 5; 2.13: 6.2 |  |  |  |  | 55
| 1 | 6 | 2.01: 6.9; 2.03: 6.59; 2.04: 6.3765; 2.05: 7.1; 2.06: 6.8488; 2.07: 7.16; 2.08: 5.7; 2.09: 6.58; 2.11: 5.8967; 2.12: 5; 2.13: 7.1
| Mean |  | 2.01: 6.9667; 2.03: 6.585; 2.04: 6.3415; 2.05: 6.6833; 2.06: 6.6243; 2.07: 7.2683; 2.08: 6.05; 2.09: 6.5917; 2.11: 5.7539; 2.12: 5.1667; 2.13: 6.5
| Median |  | 2.01: 6.9; 2.03: 6.585; 2.04: 6.3377; 2.05: 6.65; 2.06: 6.6223; 2.07: 7.255; 2.08: 6; 2.09: 6.585; 2.11: 5.7806; 2.12: 5; 2.13: 6.4
| Std Dev. |  | 2.01: 0.18619; 2.03: 0.055408; 2.04: 0.049515; 2.05: 0.23166; 2.06: 0.12978; 2.07: 0.21535; 2.08: 0.26646; 2.09: 0.042622; 2.11: 0.24873; 2.12: 0.40825; 2.13: 0.32249
| Rel.Std.Dev. |  | 2.01: 2.673%; 2.03: 0.8414%; 2.04: 0.7808%; 2.05: 3.466%; 2.06: 1.959%; 2.07: 2.963%; 2.08: 4.404%; 2.09: 0.6466%; 2.11: 4.323%; 2.12: 7.902%; 2.13: 4.961%
| PDM3 |  | 2.01: 6.778%; 2.03: 0.9279%; 2.04: -2.805%; 2.05: 2.435%; 2.06: 1.53%; 2.07: 11.4%; 2.08: -7.272%; 2.09: 1.03%; 2.11: -11.81%; 2.12: -20.81%; 2.13: -0.3749%
| Z-Score (Absolute) |  | 2.01: 1.49; 2.03: 0.01; 2.04: 0.97; 2.05: 0.37; 2.06: 0.14; 2.07: 2.68; 2.08: 2.12; 2.09: 0.01; 2.11: 3.29; 2.12: NA; 2.13: 0.35
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 609c (Certified Value 9.53 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 9.8 | 9.8 | 8.8 | 9.6 | 9.5092 | 10 | 9.5 | 9.63 | 9.4 | 10.1 | 10 | 9.9 | 8.6 | 8.6 | 9.3 | 8.92 | 9.8 | 7.2 | 12.32 | < 20 | 7 | 10.2187 | < 10 | 10.6 |  |  | 1
| 1 | 2 | 9.3 | 9.3 | 8.6 | 9.6 | 9.3496 | 10 | 9.7 | 9.3 | 9.3 | 9.8 | 9.9 | 10 | 8.9 | 9 | 9.5 | 9.04 | 10 | 7.1 | 10.38 | < 20 | 7 | 9.4068 | < 10 | 10.5 |  |  | 14
| 1 | 3 | 9.2 | 8.9 | 8.9 | 9.8 | 9.5734 | 10 | 10.1 | 9.23 | 9.4 | 10 | 10 | 9.9 | 9.1 | 8.6 | 9.4 | 8.69 | 9.9 | 7.4 | 12.35 | < 20 | 5 | 9.8693 | < 10 | 11 |  |  | 16
| 1 | 4 | 9.9 | 10.1 | 8.8 | 9.3 | 9.5482 | 10 | 9.8 | 9.5 | 9.6 | 10.1 | 9.5 | 9.9 | 9.6 | 8.8 | 9.7 | 8.49 | 9.8 | 7.4 | 12.11 | < 20 | 6 | 9.5779 | < 10 | 10.6 |  |  | 9.5251
| 1 | 5 | 8.4 | 9.2 | 8.7 | 9.6 | 9.6621 | 10 | 9.6 | 9.13 | 9.5 | 10.1 | 9.9 | 10 | 9.5 | 8.9 | 9.6 | 8.84 | 10.2 | 7.4 | 12.21 | < 20 | 10 | 9.7605 | < 10 | 10.4 |  |  | 56
| 1 | 6 | 9.6 | 9.2 | 9.1 | 9.5 | 9.5022 | 10 | 10 | 9.1 | 9.4 | 10 | 9.3 | 9.9 | 9.4 | 8.3 | 9.5 | 8.73 | 9.9 | 7.3 | 12.69 | < 20 | 7 | 10.3357 | < 10 | 10.2
| Mean |  | 9.3667 | 9.4167 | 8.8167 | 9.5667 | 9.5241 | 10 | 9.7833 | 9.315 | 9.4333 | 10.0167 | 9.7667 | 9.9333 | 9.1833 | 8.7 | 9.5 | 8.785 | 9.9333 | 7.3 | 12.01 |  | 7 | 9.8615 |  | 10.55
| Median |  | 9.45 | 9.25 | 8.8 | 9.6 | 9.5287 | 10 | 9.75 | 9.265 | 9.4 | 10.05 | 9.9 | 9.9 | 9.25 | 8.7 | 9.5 | 8.785 | 9.9 | 7.35 | 12.265 |  | 7 | 9.8149 |  | 10.55
| Std Dev. |  | 0.5465 | 0.4446 | 0.17224 | 0.1633 | 0.10315 | 0 | 0.23166 | 0.2104 | 0.10328 | 0.1169 | 0.29439 | 0.05164 | 0.38687 | 0.25298 | 0.14142 | 0.19254 | 0.15055 | 0.12649 | 0.82231 |  | 1.6733 | 0.36052 |  | 0.26646
| Rel.Std.Dev. |  | 5.835% | 4.721% | 1.954% | 1.707% | 1.083% | 0% | 2.368% | 2.259% | 1.095% | 1.167% | 3.014% | 0.5199% | 4.213% | 2.908% | 1.489% | 2.192% | 1.516% | 1.733% | 6.847% |  | 23.9% | 3.656% |  | 2.526%
| PDM3 |  | -1.664% | -1.139% | -7.438% | 0.4361% | -0.0106% | 4.985% | 2.711% | -2.206% | -0.9637% | 5.16% | 2.536% | 4.286% | -3.588% | -8.663% | -0.2638% | -7.77% | 4.286% | -23.36% | 26.09% |  | -26.51% | 3.531% |  | 10.76%
| Z-Score (Absolute) |  | 0.25 | 0.17 | 1.22 | 0.1 | 0.02 | NA | 0.47 | 0.34 | 0.14 | 0.88 | 0.45 | 0.74 | 0.58 | 1.42 | 0.02 | 1.27 | 0.74 | 3.87 | 4.37 | 0.85 | NA | 0.61 | 7.9 | 1.82
Results from laboratories 2.07 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 609c (Certified Value 194 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 206 | 157 | 171 | 204.98 | 194.5847 | 189 | 196 | 196.371 | 192 | 203 | 203 | 195 | 197.5 | 191.5 | 187 | 187.5 | 222 | 215 | 196 | 190.6 | 191 | 178.8222 | 205 | 208.3611 | 167.22 | 175 | 1
| 1 | 2 | 207 | 183 | 166 | 204.93 | 195.3559 | 188 | 205 | 196.102 | 189 | 204 | 205 | 193.5 | 198 | 187 | 191.5 | 190.5 | 224 | 206 | 196 | 194 | 190 | 177.3697 | 202 | 201.2508 | 172.24 | 173 | 15
| 1 | 3 | 205 | 165 | 174 | 206.42 | 194.7122 | 193 | 208 | 196.485 | 190 | 202 | 209 | 192 | 198.5 | 190.5 | 185 | 189.5 | 222 | 201 | 196 | 192.1 | 187 | 173.0769 | 201 | 212.7647 | 164.52 | 160 | 16
| 1 | 4 | 207 | 185 | 169 | 201.18 | 194.152 | 191 | 210 | 196.003 | 192 | 207 | 214 | 187.5 | 195 | 190.5 | 192.5 | 191 | 217 | 215 | 207 | 195.4 | 192 | 176.007 | 211 | 200.2606 | 172.97 | 178 | 194.2418
| 1 | 5 | 208 | 181 | 170 | 206.55 | 197.2672 | 191 | 203 | 196.815 | 190 | 208 | 203 | 197.5 | 198 | 193 | 188.5 | 193.5 | 220 | 218 | 206 | 189.1 | 187 | 181.0249 | 208 | 204.8349 | 167.32 | 176 | 57
| 1 | 6 | 209 | 173 | 171 | 204.06 | 194.5578 | 197 | 213 | 196.947 | 190 | 201 | 211 | 190.5 | 197.5 | 190.5 | 186 | 190 | 220 | 206 | 200 | 188.3 | 190 | 177.8912 | 202 | 211.1364 | 168.27 | 170
| Mean |  | 207 | 174 | 170.1667 | 204.6867 | 195.105 | 191.5 | 205.8333 | 196.4538 | 190.5 | 204.1667 | 207.5 | 192.6667 | 197.4167 | 190.5 | 188.4167 | 190.3333 | 220.8333 | 210.1667 | 200.1667 | 191.5833 | 189.5 | 177.3653 | 204.8333 | 206.4347 | 168.7567 | 172
| Median |  | 207 | 177 | 170.5 | 204.955 | 194.6485 | 191 | 206.5 | 196.428 | 190 | 203.5 | 207 | 192.75 | 197.75 | 190.5 | 187.75 | 190.25 | 221 | 210.5 | 198 | 191.35 | 190 | 177.6304 | 203.5 | 206.598 | 167.795 | 174
| Std Dev. |  | 1.4142 | 11.1535 | 2.6394 | 1.9653 | 1.1288 | 3.2094 | 5.9805 | 0.37644 | 1.2247 | 2.7869 | 4.5497 | 3.5024 | 1.2416 | 1.9748 | 3.0235 | 1.9664 | 2.4014 | 6.7355 | 5.1543 | 2.7781 | 2.0736 | 2.6843 | 3.9707 | 5.166 | 3.2403 | 6.4807
| Rel.Std.Dev. |  | 0.6832% | 6.41% | 1.551% | 0.9601% | 0.5786% | 1.676% | 2.906% | 0.1916% | 0.6429% | 1.365% | 2.193% | 1.818% | 0.6289% | 1.037% | 1.605% | 1.033% | 1.087% | 3.205% | 2.575% | 1.45% | 1.094% | 1.513% | 1.939% | 2.502% | 1.92% | 3.768%
| PDM3 |  | 6.568% | -10.42% | -12.39% | 5.377% | 0.4444% | -1.412% | 5.968% | 1.139% | -1.926% | 5.11% | 6.826% | -0.8109% | 1.635% | -1.926% | -2.999% | -2.012% | 13.69% | 8.199% | 3.05% | -1.369% | -2.441% | -8.688% | 5.453% | 6.277% | -13.12% | -11.45%
| Z-Score (Absolute) |  | 0.84 | 1.27 | 1.52 | 0.69 | 0.08 | 0.15 | 0.76 | 0.16 | 0.22 | 0.66 | 0.87 | 0.08 | 0.23 | 0.22 | 0.35 | 0.23 | 1.72 | 1.04 | 0.4 | 0.15 | 0.28 | 1.06 | 0.7 | 0.8 | 1.61 | 1.4
| Analytical results for Ta in OREAS 609c (Certified Value 1.22 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.77 | 1.4 | 1 | 1.26 | 1.2856 | 1.3 | 1.08 | 1.26 | 1.06 | 1.28 | 1.32 | 1.28 | 1.18 | 1.16 | 1.23 | 1.45 | 1.63 | 1.21 | 1.3 | 4 | 1.1185 | 1.05 |  |  |  |  | 1
| 1 | 2 | 0.69 | 1.4 | 1 | 1.24 | 1.2506 | 1.3 | 1.06 | 1.23 | 1.07 | 1.29 | 1.32 | 1.26 | 1.16 | 1.2 | 1.3 | 1.46 | 1.69 | 1.26 | 1.4 | < 1 | 1.1988 | 1.04 |  |  |  |  | 26
| 1 | 3 | 0.71 | 1.3 | 1.1 | 1.28 | 1.268 | 1.3 | 1.15 | 1.25 | 1.05 | 1.27 | 1.31 | 1.27 | 1.2 | 1.16 | 1.24 | 1.14 | 1.75 | 1.23 | 1.3 | < 1 | 1.0092 | 1.05 |  |  |  |  | 16
| 1 | 4 | 0.63 | 1.4 | 1 | 1.3 | 1.24 | 1.2 | 1.14 | 1.24 | 1.08 | 1.35 | 1.26 | 1.24 | 1.27 | 1.16 | 1.28 | 1.45 | 1.67 | 1.24 | 1.35 | < 1 | 0.97261 | 1.08 |  |  |  |  | 1.2212
| 1 | 5 | 0.7 | 1.4 | 1 | 1.25 | 1.2723 | 1.3 | 1.06 | 1.14 | 1.01 | 1.21 | 1.35 | 1.27 | 1.24 | 1.15 | 1.29 | 1.38 | 1.68 | 1.3 | 1.33 | < 1 | 1.1237 | 1.08 |  |  |  |  | 58
| 1 | 6 | 0.7 | 1.4 | 1 | 1.3 | 1.2756 | 1.3 | 1.17 | 1.16 | 1.02 | 1.26 | 1.23 | 1.28 | 1.25 | 1.19 | 1.27 | 1.25 | 1.6 | 1.3 | 1.32 | < 1 | 1.1919 | 1.1
| Mean |  | 0.7 | 1.3833 | 1.0167 | 1.2717 | 1.2653 | 1.2833 | 1.11 | 1.2133 | 1.0483 | 1.2767 | 1.2983 | 1.2667 | 1.2167 | 1.17 | 1.2683 | 1.355 | 1.67 | 1.2567 | 1.3333 | 4 | 1.1024 | 1.0667
| Median |  | 0.7 | 1.4 | 1 | 1.27 | 1.2701 | 1.3 | 1.11 | 1.235 | 1.055 | 1.275 | 1.315 | 1.27 | 1.22 | 1.16 | 1.275 | 1.415 | 1.675 | 1.25 | 1.325 | 4 | 1.1211 | 1.065
| Std Dev. |  | 0.044721 | 0.040825 | 0.040825 | 0.025626 | 0.016918 | 0.040825 | 0.04899 | 0.050465 | 0.027869 | 0.045461 | 0.04446 | 0.015055 | 0.043205 | 0.02 | 0.027869 | 0.13187 | 0.051769 | 0.037238 | 0.037771 |  | 0.093329 | 0.023381
| Rel.Std.Dev. |  | 6.389% | 2.951% | 4.016% | 2.015% | 1.337% | 3.181% | 4.413% | 4.159% | 2.658% | 3.561% | 3.424% | 1.189% | 3.551% | 1.709% | 2.197% | 9.732% | 3.1% | 2.963% | 2.833% |  | 8.466% | 2.192%
| PDM3 |  | -42.68% | 13.28% | -16.75% | 4.133% | 3.615% | 5.088% | -9.106% | -0.6441% | -14.16% | 4.542% | 6.316% | 3.723% | -0.3711% | -4.193% | 3.86% | 10.96% | 36.75% | 2.904% | 9.182% | 227.5% | -9.725% | -12.65%
| Z-Score (Absolute) |  | 5.61 | 1.17 | 2.47 | 0.06 | 0 | 0.18 | 1.54 | 0.52 | 2.15 | 0.11 | 0.33 | 0.01 | 0.48 | 0.95 | 0.03 | 0.89 | 4.01 | 0.09 | 0.67 | NA | 1.62 | 1.97
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 609c (Certified Value 0.7 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.10; 2.08: 2.12; 2.10: 2.19; 2.11: 2.22; 2.12: 2.23; 2.13: 2.25; 2.14: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*MS; 2.13: 4A*OES/MS; 2.14: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2.01: 0.7; 2.03: 0.65; 2.04: 0.71527; 2.05: 0.7; 2.06: 0.8428; 2.07: 0.64; 2.08: 0.68; 2.10: 0.75; 2.11: 0.74; 2.12: 0.6097; 2.13: < 1; 2.14: 0.59 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.6; 2.03: 0.63; 2.04: 0.71691; 2.05: 0.7; 2.06: 0.8463; 2.07: 0.61; 2.08: 0.69; 2.10: 0.72; 2.11: 0.76; 2.12: 0.6642; 2.13: < 1; 2.14: 0.61 |  |  |  |  | 27
| 1 | 3 | 2.01: 0.7; 2.03: 0.64; 2.04: 0.70213; 2.05: 0.8; 2.06: 0.8764; 2.07: 0.64; 2.08: 0.7; 2.10: 0.7; 2.11: 0.73; 2.12: 0.6953; 2.13: < 1; 2.14: 0.6 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.7; 2.03: 0.64; 2.04: 0.71687; 2.05: 0.7; 2.06: 0.8508; 2.07: 0.64; 2.08: 0.69; 2.10: 0.73; 2.11: 0.72; 2.12: 0.681; 2.13: < 1; 2.14: 0.64 |  |  |  |  | 0.69555
| 1 | 5 | 2.01: 0.7; 2.03: 0.64; 2.04: 0.73722; 2.05: 0.7; 2.06: 0.8559; 2.07: 0.62; 2.08: 0.68; 2.10: 0.77; 2.11: 0.74; 2.12: 0.6489; 2.13: < 1; 2.14: 0.58 |  |  |  |  | 59
| 1 | 6 | 2.01: 0.7; 2.03: 0.63; 2.04: 0.72573; 2.05: 0.8; 2.06: 0.8662; 2.07: 0.62; 2.08: 0.67; 2.10: 0.68; 2.11: 0.74; 2.12: 0.658; 2.13: < 1; 2.14: 0.64
| Mean |  | 2.01: 0.68333; 2.03: 0.63833; 2.04: 0.71902; 2.05: 0.73333; 2.06: 0.8564; 2.07: 0.62833; 2.08: 0.685; 2.10: 0.725; 2.11: 0.73833; 2.12: 0.65952; 2.14: 0.61
| Median |  | 2.01: 0.7; 2.03: 0.64; 2.04: 0.71689; 2.05: 0.7; 2.06: 0.85335; 2.07: 0.63; 2.08: 0.685; 2.10: 0.725; 2.11: 0.74; 2.12: 0.6611; 2.14: 0.605
| Std Dev. |  | 2.01: 0.040825; 2.03: 0.0075277; 2.04: 0.011702; 2.05: 0.05164; 2.06: 0.012764; 2.07: 0.013292; 2.08: 0.010488; 2.10: 0.032711; 2.11: 0.013292; 2.12: 0.029532; 2.14: 0.025298
| Rel.Std.Dev. |  | 2.01: 5.974%; 2.03: 1.179%; 2.04: 1.628%; 2.05: 7.042%; 2.06: 1.49%; 2.07: 2.115%; 2.08: 1.531%; 2.10: 4.512%; 2.11: 1.8%; 2.12: 4.478%; 2.14: 4.147%
| PDM3 |  | 2.01: -1.756%; 2.03: -8.226%; 2.04: 3.375%; 2.05: 5.432%; 2.06: 23.13%; 2.07: -9.664%; 2.08: -1.517%; 2.10: 4.234%; 2.11: 6.151%; 2.12: -5.18%; 2.14: -12.3%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.46; 2.04: 0.63; 2.05: NA; 2.06: 2.5; 2.07: 0.6; 2.08: 0.17; 2.10: 0.71; 2.11: 0.9; 2.12: 0.17; 2.13: 2.33; 2.14: 0.84
Results from laboratories 2.03 and 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 609c (Certified Value 11.9 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.27 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 12.21 | 11.6 | 10.6 | 12.8 | 12.4456 | 10.1 | 11.8 | 11.4 | 12.3 | 12.9 | 12.25 | 11.7 | 11.6 | 11.6 | 11.88 | 10.33 | 14.1 | 11.03 | 13.395 | 13.07 | 10.42 | [5 empty] | 1
| 1 | 2 | 12.14 | 10.54 | 10.9 | 12.4 | 12.3908 | 11.4 | 12.3 | 11.4 | 12.6 | 12.85 | 12.2 | 11.65 | 12.15 | 12 | 11.58 | 10.85 | 15 | 11.41 | 12.2476 | 14.67 | 9.88 | [5 empty] | 28
| 1 | 3 | 11.53 | 10.49 | 10.8 | 12.8 | 12.2068 | 11.6 | 12.3 | 11.5 | 12.8 | 13 | 12.2 | 12.15 | 11.6 | 11.65 | 11.82 | 10.88 | 11.6 | 11.11 | 11.8654 | 15.1 | 9.91 | [5 empty] | 16
| 1 | 4 | 11.43 | 11.75 | 10.5 | 12.8 | 12.1415 | 10.5 | 12.2 | 11.4 | 13 | 12.25 | 12.3 | 12.3 | 12.05 | 11.9 | 11.22 | 10.59 | 19.1 | 10.64 | 11.6628 | 13.36 | 10.57 | [5 empty] | 11.8941
| 1 | 5 | 11.36 | 10.52 | 10.4 | 12.5 | 12.6466 | 11 | 12.4 | 11.5 | 13 | 13.15 | 12.2 | 12.35 | 11.85 | 12.15 | 11.96 | 11.17 | 13.1 | 11.26 | 12.5268 | 14.09 | 12.25 | [5 empty] | 60
| 1 | 6 | 12.4 | 11.42 | 10.5 | 12.7 | 12.5479 | 11.1 | 12.4 | 11.5 | 12.2 | 11.95 | 12.3 | 12.45 | 11.3 | 11.65 | 12.1 | 10.46 | 14.1 | 11.37 | 13.2957 | 13.25 | 10.21
| Mean |  | 11.845 | 11.0533 | 10.6167 | 12.6667 | 12.3965 | 10.95 | 12.2333 | 11.45 | 12.65 | 12.6833 | 12.2417 | 12.1 | 11.7583 | 11.825 | 11.76 | 10.7133 | 14.5 | 11.1367 | 12.4989 | 13.9233 | 10.54
| Median |  | 11.835 | 10.98 | 10.55 | 12.75 | 12.4182 | 11.05 | 12.3 | 11.45 | 12.7 | 12.875 | 12.225 | 12.225 | 11.725 | 11.775 | 11.85 | 10.72 | 14.1 | 11.185 | 12.3872 | 13.725 | 10.315
| Std Dev. |  | 0.45496 | 0.59732 | 0.19408 | 0.17512 | 0.19435 | 0.56125 | 0.22509 | 0.054772 | 0.34496 | 0.47293 | 0.04916 | 0.34351 | 0.31846 | 0.22528 | 0.31547 | 0.31027 | 2.5338 | 0.28395 | 0.72139 | 0.83291 | 0.881
| Rel.Std.Dev. |  | 3.841% | 5.404% | 1.828% | 1.383% | 1.568% | 5.126% | 1.84% | 0.4784% | 2.727% | 3.729% | 0.4016% | 2.839% | 2.708% | 1.905% | 2.683% | 2.896% | 17.47% | 2.55% | 5.772% | 5.982% | 8.359%
| PDM3 |  | -0.4129% | -7.069% | -10.74% | 6.495% | 4.224% | -7.938% | 2.852% | -3.734% | 6.355% | 6.635% | 2.922% | 1.731% | -1.142% | -0.5811% | -1.128% | -9.927% | 21.91% | -6.368% | 5.085% | 17.06% | -11.38%
| Z-Score (Absolute) |  | 0 | 0.75 | 1.17 | 0.78 | 0.53 | 0.85 | 0.37 | 0.38 | 0.77 | 0.8 | 0.38 | 0.24 | 0.08 | 0.02 | 0.08 | 1.08 | 2.53 | 0.67 | 0.62 | 1.98 | 1.24
| Analytical results for Th in OREAS 609c (Certified Value 13.2 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 12.23 | 12.5 | 12 | 13.93 | 13.434 | 11.8 | 13.5 | 15.342 | 12.77 | 14.2 | 14.01 | 13.75 | 13.2 | 12.85 | 12.95 | 14.05 | 13.13 | 12.53 | < 20 | 11.77 | 13.3699 | 13 | 14.0152 | 11.6 |  |  | 1
| 1 | 2 | 12.58 | 12.8 | 12.1 | 13.89 | 13.4436 | 12.2 | 13.7 | 15.457 | 12.47 | 14 | 13.96 | 13.15 | 13.2 | 12.25 | 13.4 | 14.55 | 13.22 | 11.65 | < 20 | 12.36 | 14.4089 | 13 | 13.412 | 11.36 |  |  | 29
| 1 | 3 | 12.51 | 12.6 | 12.7 | 14.16 | 13.5411 | 12.4 | 14.8 | 15.644 | 12.66 | 14.3 | 14.38 | 13.9 | 12.75 | 12.6 | 12.9 | 14.5 | 12.96 | 11.58 | < 20 | 12.36 | 15.3451 | 13 | 13.7523 | 10.39 |  |  | 16
| 1 | 4 | 11.83 | 13.1 | 12.1 | 13.79 | 13.5187 | 12.1 | 14.6 | 15.705 | 12.53 | 14.5 | 14.52 | 13.05 | 12.75 | 12.9 | 13 | 14.55 | 12.69 | 12.34 | < 20 | 12.05 | 14.2608 | 13 | 13.1459 | 11.33 |  |  | 13.2049
| 1 | 5 | 11.33 | 12.7 | 12.2 | 13.86 | 13.5572 | 11.9 | 14.1 | 15.508 | 12.14 | 14.4 | 14.22 | 13.45 | 13 | 12.9 | 13.05 | 14.55 | 12.86 | 12.54 | < 20 | 12.67 | 14.0233 | 13 | 13.4816 | 11.79 |  |  | 61
| 1 | 6 | 12.83 | 13.1 | 12.4 | 13.86 | 13.8777 | 12.3 | 14.6 | 15.344 | 12.77 | 14.7 | 13.5 | 13.25 | 12.65 | 13.15 | 12.95 | 14.55 | 12.83 | 11.35 | < 20 | 12.4 | 14.1185 | 13 | 13.9921 | 10.65
| Mean |  | 12.2183 | 12.8 | 12.25 | 13.915 | 13.562 | 12.1167 | 14.2167 | 15.5 | 12.5567 | 14.35 | 14.0983 | 13.425 | 12.925 | 12.775 | 13.0417 | 14.4583 | 12.9483 | 11.9983 |  | 12.2683 | 14.2544 | 13 | 13.6332 | 11.1867
| Median |  | 12.37 | 12.75 | 12.15 | 13.875 | 13.5299 | 12.15 | 14.35 | 15.4825 | 12.595 | 14.35 | 14.115 | 13.35 | 12.875 | 12.875 | 12.975 | 14.55 | 12.91 | 11.995 |  | 12.36 | 14.1897 | 13 | 13.6169 | 11.345
| Std Dev. |  | 0.55319 | 0.25298 | 0.25884 | 0.12849 | 0.16272 | 0.23166 | 0.53448 | 0.15098 | 0.23813 | 0.2429 | 0.36246 | 0.34022 | 0.24238 | 0.31105 | 0.1828 | 0.20104 | 0.19773 | 0.53094 |  | 0.31352 | 0.64308 | 0 | 0.34591 | 0.54921
| Rel.Std.Dev. |  | 4.528% | 1.976% | 2.113% | 0.9234% | 1.2% | 1.912% | 3.76% | 0.9741% | 1.896% | 1.693% | 2.571% | 2.534% | 1.875% | 2.435% | 1.402% | 1.39% | 1.527% | 4.425% |  | 2.556% | 4.511% | 0% | 2.537% | 4.909%
| PDM3 |  | -7.471% | -3.066% | -7.231% | 5.377% | 2.705% | -8.241% | 7.662% | 17.38% | -4.909% | 8.672% | 6.766% | 1.667% | -2.12% | -3.256% | -1.236% | 9.492% | -1.943% | -9.137% |  | -7.093% | 7.948% | -1.552% | 3.243% | -15.28%
| Z-Score (Absolute) |  | 0.67 | 0.14 | 0.65 | 0.89 | 0.57 | 0.77 | 1.17 | 2.36 | 0.36 | 1.29 | 1.06 | 0.44 | 0.02 | 0.16 | 0.09 | 1.39 | 0 | 0.88 | 2.72 | 0.63 | 1.21 | NA | 0.63 | 1.63
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 609c (Certified Value 0.116 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.1 | 0.126 | 0.115 | 0.1152 | 0.11653 | 0.108 | 0.13 | 0.11244 | 0.123 | 0.112 | 0.1109 | 0.115 | 0.118 | 0.108 | 0.116 | 0.119 | 0.1163 | 0.119 | 0.122 | 0.11 | 0.12 | < 0.1 | 0.114 | 0.11359 | 0.123 |  | 1
| 1 | 2 | 0.1 | 0.125 | 0.11 | 0.1168 | 0.1175 | 0.112 | 0.11 | 0.11129 | 0.117 | 0.108 | 0.1144 | 0.114 | 0.117 | 0.11 | 0.12 | 0.12 | 0.1153 | 0.119 | 0.12 | 0.11 | 0.12 | < 0.1 | 0.114 | 0.11057 | 0.123 |  | 30
| 1 | 3 | 0.1 | 0.118 | 0.114 | 0.1176 | 0.11661 | 0.111 | 0.11 | 0.11194 | 0.12 | 0.111 | 0.1138 | 0.114 | 0.118 | 0.113 | 0.114 | 0.121 | 0.1029 | 0.118 | 0.121 | 0.11 | 0.11 | < 0.1 | 0.114 | 0.11481 | 0.126 |  | 16
| 1 | 4 | 0.1 | 0.126 | 0.116 | 0.1163 | 0.11759 | 0.108 | 0.11 | 0.11264 | 0.12 | 0.112 | 0.119 | 0.112 | 0.116 | 0.113 | 0.119 | 0.121 | 0.1139 | 0.119 | 0.133 | 0.11 | 0.12 | < 0.1 | 0.12 | 0.11155 | 0.128 |  | 0.11604
| 1 | 5 | 0.11 | 0.122 | 0.116 | 0.1191 | 0.11939 | 0.112 | 0.11 | 0.11307 | 0.117 | 0.114 | 0.11 | 0.118 | 0.117 | 0.113 | 0.117 | 0.122 | 0.1148 | 0.123 | 0.127 | 0.11 | 0.11 | < 0.1 | 0.12 | 0.11051 | 0.123 |  | 62
| 1 | 6 | 0.1 | 0.123 | 0.118 | 0.1155 | 0.11884 | 0.113 | 0.11 | 0.11231 | 0.115 | 0.113 | 0.1164 | 0.112 | 0.118 | 0.112 | 0.116 | 0.12 | 0.1072 | 0.118 | 0.125 | 0.11 | 0.12 | < 0.1 | 0.114 | 0.11366 | 0.126
| Mean |  | 0.10167 | 0.12333 | 0.11483 | 0.11675 | 0.11774 | 0.11067 | 0.11333 | 0.11228 | 0.11867 | 0.11167 | 0.11408 | 0.11417 | 0.11733 | 0.1115 | 0.117 | 0.1205 | 0.11173 | 0.11933 | 0.12467 | 0.11 | 0.11667 |  | 0.116 | 0.11245 | 0.12483
| Median |  | 0.1 | 0.124 | 0.1155 | 0.11655 | 0.11754 | 0.1115 | 0.11 | 0.11237 | 0.1185 | 0.112 | 0.1141 | 0.114 | 0.1175 | 0.1125 | 0.1165 | 0.1205 | 0.11435 | 0.119 | 0.1235 | 0.11 | 0.12 |  | 0.114 | 0.11257 | 0.1245
| Std Dev. |  | 0.0040825 | 0.0030768 | 0.0027142 | 0.0014433 | 0.0011605 | 0.0021602 | 0.008165 | 0.0006117 | 0.0028752 | 0.0020656 | 0.0033624 | 0.0022286 | 0.0008165 | 0.0020736 | 0.0021909 | 0.0010488 | 0.0054084 | 0.0018619 | 0.0048442 | 0 | 0.005164 |  | 0.0030984 | 0.0018112 | 0.002137
| Rel.Std.Dev. |  | 4.016% | 2.495% | 2.364% | 1.236% | 0.9856% | 1.952% | 7.204% | 0.5448% | 2.423% | 1.85% | 2.947% | 1.952% | 0.6959% | 1.86% | 1.873% | 0.8704% | 4.84% | 1.56% | 3.886% | 0% | 4.426% |  | 2.671% | 1.611% | 1.712%
| PDM3 |  | -12.38% | 6.29% | -1.036% | 0.6161% | 1.473% | -4.627% | -2.328% | -3.235% | 2.268% | -3.765% | -1.682% | -1.61% | 1.119% | -3.908% | 0.8315% | 3.848% | -3.707% | 2.842% | 7.439% | -5.201% | 0.5443% |  | -0.03026% | -3.091% | 7.582%
| Z-Score (Absolute) |  | 2.86 | 1.85 | 0 | 0.42 | 0.63 | 0.91 | 0.33 | 0.56 | 0.83 | 0.69 | 0.16 | 0.15 | 0.54 | 0.73 | 0.47 | 1.23 | 0.67 | 0.98 | 2.14 | 1.05 | 0.4 | 14.1 | 0.23 | 0.52 | 2.18
| Analytical results for Tl in OREAS 609c (Certified Value 1.14 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 0.56 | 1.1 | 1.2 | 1.17 | 1.135 | 1.11 | 1.11 | 1.17 | 1.16 | 1.16 | 1.19 | 1.1 | 1.04 | 1.07 | 1.12 | 1.066 | 1.24 | 1.17 | 0.81 | 1.1743 | 0.93 | [5 empty] | 1
| 1 | 2 | 0.56 | 1.09 | 1.1 | 1.15 | 1.1612 | 1.11 | 1.12 | 1.09 | 1.14 | 1.21 | 1.17 | 1.07 | 1.06 | 1.1 | 1.14 | 1.072 | 1.29 | 1.15 | 0.86 | 1.0637 | 0.92 | [5 empty] | 31
| 1 | 3 | 0.56 | 1.02 | 1.2 | 1.19 | 1.1401 | 1.12 | 1.17 | 1.17 | 1.14 | 1.2 | 1.17 | 1.06 | 1.06 | 1.08 | 1.12 | 1.054 | 1.28 | 1.17 | 0.82 | 1.1475 | 0.84 | [5 empty] | 16
| 1 | 4 | 0.54 | 1.1 | 1.2 | 1.13 | 1.175 | 1.1 | 1.16 | 1.15 | 1.16 | 1.24 | 1.11 | 1.06 | 1.14 | 1.09 | 1.14 | 1.03 | 1.25 | 1.14 | 0.8 | 1.0772 | 0.93 | [5 empty] | 1.1392
| 1 | 5 | 0.59 | 1.13 | 1.2 | 1.15 | 1.1648 | 1.11 | 1.14 | 1.13 | 1.16 | 1.22 | 1.24 | 1.07 | 1.12 | 1.09 | 1.13 | 1.055 | 1.28 | 1.21 | 0.79 | 1.0873 | 0.97 | [5 empty] | 63
| 1 | 6 | 0.61 | 1.13 | 1.2 | 1.2 | 1.1735 | 1.13 | 1.18 | 1.14 | 1.15 | 1.18 | 1.07 | 1.02 | 1.12 | 1.09 | 1.12 | 1.055 | 1.27 | 1.18 | 0.85 | 1.1341 | 0.86
| Mean |  | 0.57 | 1.095 | 1.1833 | 1.165 | 1.1583 | 1.1133 | 1.1467 | 1.1417 | 1.1517 | 1.2017 | 1.1583 | 1.0633 | 1.09 | 1.0867 | 1.1283 | 1.0553 | 1.2683 | 1.17 | 0.82167 | 1.114 | 0.90833
| Median |  | 0.56 | 1.1 | 1.2 | 1.16 | 1.163 | 1.11 | 1.15 | 1.145 | 1.155 | 1.205 | 1.17 | 1.065 | 1.09 | 1.09 | 1.125 | 1.055 | 1.275 | 1.17 | 0.815 | 1.1107 | 0.925
| Std Dev. |  | 0.025298 | 0.040373 | 0.040825 | 0.026646 | 0.016948 | 0.010328 | 0.028048 | 0.029944 | 0.0098319 | 0.028577 | 0.060139 | 0.02582 | 0.041473 | 0.010328 | 0.0098319 | 0.01439 | 0.019408 | 0.024495 | 0.027869 | 0.044174 | 0.048751
| Rel.Std.Dev. |  | 4.438% | 3.687% | 3.45% | 2.287% | 1.463% | 0.9277% | 2.446% | 2.623% | 0.8537% | 2.378% | 5.192% | 2.428% | 3.805% | 0.9504% | 0.8714% | 1.364% | 1.53% | 2.094% | 3.392% | 3.965% | 5.367%
| PDM3 |  | -49.97% | -3.881% | 3.872% | 2.263% | 1.672% | -2.272% | 0.6538% | 0.2149% | 1.093% | 5.482% | 1.678% | -6.661% | -4.32% | -4.613% | -0.9555% | -7.363% | 11.33% | 2.702% | -27.87% | -2.213% | -20.27%
| Z-Score (Absolute) |  | 9.82 | 0.59 | 0.97 | 0.64 | 0.53 | 0.26 | 0.32 | 0.23 | 0.41 | 1.29 | 0.53 | 1.14 | 0.67 | 0.73 | 0 | 1.28 | 2.46 | 0.73 | 5.39 | 0.25 | 3.87
| Analytical results for Tm in OREAS 609c (Certified Value 0.1 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | [5 empty] | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.28 | [5 empty] | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.01: 0.1; 2.03: 0.11; 2.04: 0.10805; 2.05: < 0.1; 2.06: 0.1048; 2.07: 0.1; 2.08: 0.12; 2.10: 0.1; 2.11: 0.1004; 2.12: 0.09 | [5 empty] | 1
| 1 | 2 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.10731; 2.05: 0.1; 2.06: < 0.1; 2.07: 0.1; 2.08: 0.11; 2.10: 0.11; 2.11: 0.11; 2.12: 0.08 | [5 empty] | 6
| 1 | 3 | 2.01: < 0.1; 2.03: 0.1; 2.04: 0.10449; 2.05: < 0.1; 2.06: 0.1074; 2.07: 0.11; 2.08: 0.11; 2.10: 0.1; 2.11: 0.1082; 2.12: 0.09 | [5 empty] | 16
| 1 | 4 | 2.01: < 0.1; 2.03: 0.1; 2.04: 0.10463; 2.05: < 0.1; 2.06: 0.1025; 2.07: 0.11; 2.08: 0.11; 2.10: 0.1; 2.11: 0.1056; 2.12: 0.1 | [5 empty] | 0.1048
| 1 | 5 | 2.01: < 0.1; 2.03: 0.1; 2.04: 0.10839; 2.05: 0.1; 2.06: 0.1048; 2.07: 0.1; 2.08: 0.11; 2.10: 0.11; 2.11: 0.1041; 2.12: 0.09 | [5 empty] | 64
| 1 | 6 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.10905; 2.05: < 0.1; 2.06: 0.1025; 2.07: 0.1; 2.08: 0.1; 2.10: 0.11; 2.11: 0.1059; 2.12: 0.1
| Mean |  | 2.01: 0.1; 2.03: 0.105; 2.04: 0.10699; 2.05: 0.1; 2.06: 0.1044; 2.07: 0.10333; 2.08: 0.11; 2.10: 0.105; 2.11: 0.1057; 2.12: 0.091667
| Median |  | 2.01: 0.1; 2.03: 0.105; 2.04: 0.10768; 2.05: 0.1; 2.06: 0.1048; 2.07: 0.1; 2.08: 0.11; 2.10: 0.105; 2.11: 0.10575; 2.12: 0.09
| Std Dev. |  | 2.03: 0.0054772; 2.04: 0.0019613; 2.05: 0; 2.06: 0.0020335; 2.07: 0.005164; 2.08: 0.0063246; 2.10: 0.0054772; 2.11: 0.0033299; 2.12: 0.0075277
| Rel.Std.Dev. |  | 2.03: 5.216%; 2.04: 1.833%; 2.05: 0%; 2.06: 1.948%; 2.07: 4.997%; 2.08: 5.75%; 2.10: 5.216%; 2.11: 3.15%; 2.12: 8.212%
| PDM3 |  | 2.01: -4.583%; 2.03: 0.1882%; 2.04: 2.086%; 2.05: -4.583%; 2.06: -0.3843%; 2.07: -1.402%; 2.08: 4.959%; 2.10: 0.1882%; 2.11: 0.8561%; 2.12: -12.53%
| Z-Score (Absolute) |  | 2.01: 15.82; 2.03: 0; 2.04: 0.67; 2.05: NA; 2.06: 3.28; 2.07: 0.57; 2.08: 1.7; 2.10: 0; 2.11: 0.24; 2.12: 4.52
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 609c (Certified Value 4.86 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 3.3 | 4.7 | 4.9 | 5.01 | 5.2886 | 4.8 | 4.58 | 5.37 | 4.72 | 4.93 | 4.96 | 5 | 4.8 | 4.6 | 4.8 | 4.8 | 4.89 | 5.52 | < 10 | < 10 | 4.1805 | < 2 | 5.0977 | 4.2 |  |  | 1
| 1 | 2 | 3.4 | 4.6 | 4.8 | 5.09 | 5.2734 | 4.8 | 4.63 | 5.2596 | 4.79 | 4.82 | 4.99 | 4.9 | 4.7 | 4.6 | 5 | 4.8 | 4.91 | 5.3 | < 10 | < 10 | 4.5309 | 2 | 4.9851 | 4.12 |  |  | 33
| 1 | 3 | 3.3 | 4.8 | 5.1 | 5.11 | 5.0994 | 4.9 | 4.9 | 5.6659 | 4.88 | 4.84 | 5.18 | 4.9 | 4.8 | 4.7 | 4.8 | 4.8 | 4.85 | 5.26 | < 10 | < 10 | 4.5418 | 2 | 5.0195 | 3.8 |  |  | 16
| 1 | 4 | 3.3 | 4.8 | 4.8 | 4.95 | 5.1981 | 4.8 | 4.88 | 5.5344 | 4.85 | 5 | 5.12 | 4.7 | 4.7 | 4.9 | 5 | 4.8 | 4.84 | 5.45 | < 10 | < 10 | 4.5001 | 2 | 4.9648 | 4.18 |  |  | 4.8577
| 1 | 5 | 3.2 | 4.6 | 4.8 | 5.08 | 5.177 | 4.8 | 4.71 | 5.4847 | 4.71 | 4.98 | 5.07 | 5 | 4.7 | 5.1 | 4.8 | 4.9 | 4.85 | 5.63 | < 10 | < 10 | 4.4224 | 2 | 4.9506 | 4.29 |  |  | 65
| 1 | 6 | 3.4 | 4.7 | 5.1 | 4.96 | 5.044 | 4.9 | 4.83 | 5.5205 | 4.73 | 4.81 | 4.92 | 4.6 | 4.7 | 5 | 4.9 | 4.8 | 4.79 | 5.28 | < 10 | < 10 | 4.3894 | 2 | 5.0656 | 3.89
| Mean |  | 3.3167 | 4.7 | 4.9167 | 5.0333 | 5.1801 | 4.8333 | 4.755 | 5.4725 | 4.78 | 4.8967 | 5.04 | 4.85 | 4.7333 | 4.8167 | 4.8833 | 4.8167 | 4.855 | 5.4067 |  |  | 4.4275 | 2 | 5.0139 | 4.08
| Median |  | 3.3 | 4.7 | 4.85 | 5.045 | 5.1876 | 4.8 | 4.77 | 5.5026 | 4.76 | 4.885 | 5.03 | 4.9 | 4.7 | 4.8 | 4.85 | 4.8 | 4.85 | 5.375 |  |  | 4.4612 | 2 | 5.0023 | 4.15
| Std Dev. |  | 0.075277 | 0.089443 | 0.1472 | 0.069474 | 0.095734 | 0.05164 | 0.13457 | 0.14108 | 0.072111 | 0.084063 | 0.1002 | 0.16432 | 0.05164 | 0.2137 | 0.098319 | 0.040825 | 0.041833 | 0.15069 |  |  | 0.13529 | 0 | 0.058252 | 0.19215
| Rel.Std.Dev. |  | 2.27% | 1.903% | 2.994% | 1.38% | 1.848% | 1.068% | 2.83% | 2.578% | 1.509% | 1.717% | 1.988% | 3.388% | 1.091% | 4.437% | 2.013% | 0.8476% | 0.8616% | 2.787% |  |  | 3.056% | 0% | 1.162% | 4.709%
| PDM3 |  | -31.72% | -3.246% | 1.214% | 3.616% | 6.637% | -0.5016% | -2.114% | 12.66% | -1.6% | 0.8022% | 3.753% | -0.1585% | -2.56% | -0.8447% | 0.5277% | -0.8447% | -0.05558% | 11.3% |  |  | -8.856% | -58.83% | 3.215% | -16.01%
| Z-Score (Absolute) |  | 7.15 | 0.72 | 0.29 | 0.83 | 1.51 | 0.1 | 0.47 | 2.87 | 0.35 | 0.19 | 0.86 | 0.02 | 0.57 | 0.18 | 0.13 | 0.18 | 0 | 2.57 | 0.67 | 0.67 | 1.99 | NA | 0.74 | 3.6
| Analytical results for V in OREAS 609c (Certified Value 8.85 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 9 | 9 | 9 | 9 | 8.6462 | 9 | 7.9 | 8.63 | 9 | 9 | 9 | 8 | 8 | 8 | 8 | 9 | 9 | 9.5 | 13.67 | 9 | < 0.1 | 8 | 9.9736 | < 10 | 10 |  | 1
| 1 | 2 | 9 | 9 | 8 | 9 | 8.7786 | 9 | 8.5 | 8.3 | 9 | 9 | 9 | 8 | 8 | 9 | 8 | 9 | 10 | 9.3 | 13.1 | 10 | < 0.1 | 9 | 9.8423 | < 10 | 10 |  | 34
| 1 | 3 | 9 | 9 | 9 | 9 | 8.6146 | 9 | 8.4 | 8.33 | 9 | 9 | 8 | 8 | 8 | 8 | 9 | 9 | 10 | 9.1 | 12.44 | 9 | < 0.1 | 8 | 9.9881 | < 10 | 10 |  | 16
| 1 | 4 | 9 | 9 | 9 | 8 | 8.8908 | 9 | 8.2 | 8.5 | 9 | 9 | 8 | 8 | 8 | 8 | 8 | 9 | 10 | 9.4 | 12.79 | 10 | < 0.1 | 9 | 9.8394 | < 10 | 10 |  | 8.8476
| 1 | 5 | 8 | 9 | 9 | 9 | 8.9092 | 9 | 8.2 | 8.25 | 9 | 9 | 8 | 8 | 8 | 8 | 9 | 9 | 10 | 9.3 | 12.45 | 9 | < 0.1 | 9 | 9.9556 | < 10 | 9 |  | 66
| 1 | 6 | 9 | 9 | 9 | 9 | 8.6817 | 9 | 8.5 | 8.35 | 9 | 9 | 8 | 8 | 8 | 8 | 8 | 9 | 10 | 9.1 | 11.44 | 10 | < 0.1 | 8 | 10.0039 | < 10 | 10
| Mean |  | 8.8333 | 9 | 8.8333 | 8.8333 | 8.7535 | 9 | 8.2833 | 8.3933 | 9 | 9 | 8.3333 | 8 | 8 | 8.1667 | 8.3333 | 9 | 9.8333 | 9.2833 | 12.6483 | 9.5 |  | 8.5 | 9.9338 |  | 9.8333
| Median |  | 9 | 9 | 9 | 9 | 8.7302 | 9 | 8.3 | 8.34 | 9 | 9 | 8 | 8 | 8 | 8 | 8 | 9 | 10 | 9.3 | 12.62 | 9.5 |  | 8.5 | 9.9646 |  | 10
| Std Dev. |  | 0.40825 | 0 | 0.40825 | 0.40825 | 0.12624 | 0 | 0.23166 | 0.1432 | 0 | 0 | 0.5164 | 0 | 0 | 0.40825 | 0.5164 | 0 | 0.40825 | 0.16021 | 0.74976 | 0.54772 |  | 0.54772 | 0.073745 |  | 0.40825
| Rel.Std.Dev. |  | 4.622% | 0% | 4.622% | 4.622% | 1.442% | 0% | 2.797% | 1.706% | 0% | 0% | 6.197% | 0% | 0% | 4.999% | 6.197% | 0% | 4.152% | 1.726% | 5.928% | 5.766% |  | 6.444% | 0.7424% |  | 4.152%
| PDM3 |  | -0.1613% | 1.722% | -0.1613% | -0.1613% | -1.063% | 1.722% | -6.378% | -5.134% | 1.722% | 1.722% | -5.813% | -9.58% | -9.58% | -7.696% | -5.813% | 1.722% | 11.14% | 4.925% | 42.96% | 7.374% |  | -3.929% | 12.28% |  | 11.14%
| Z-Score (Absolute) |  | 0 | 0.22 | 0 | 0 | 0.11 | 0.22 | 0.74 | 0.59 | 0.22 | 0.22 | 0.67 | 1.12 | 1.12 | 0.9 | 0.67 | 0.22 | 1.35 | 0.61 | 5.14 | 0.9 | 11.85 | 0.45 | 1.48 | 5.17 | 1.35
| Analytical results for W in OREAS 609c (Certified Value 3.29 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | < 10 | 3.2 | 3.2 | 3.5 | < 10 | 3.3 | 3.4 | 3.33 | 3 | 3.5 | 3.5 | 3.4 | 3.3 | 3.2 | 3.3 | 3.93 | 4.3 | 3.58 | < 5 | 2.7 |  | 2 | 2.8936 | 2.9 |  |  | 1
| 1 | 2 | < 10 | 3.6 | 3.4 | 3.4 | < 10 | 3.4 | 3.3 | 3.47 | 3.2 | 3.5 | 3.6 | 3.3 | 3.2 | 3.5 | 3.5 | 3.46 | 4.3 | 3.6 | < 5 | 2.9 | 10.4255 | < 1 | 3.0485 | 3 |  |  | 35
| 1 | 3 | < 10 | 3.1 | 3.2 | 3.4 | < 10 | 3.6 | 3.5 | 3.37 | 3 | 3.5 | 3.5 | 3.3 | 3.3 | 3.2 | 3.5 | 3.31 | 4.3 | 3.56 | < 5 | 2.9 | 3.7243 | 4 | 2.8264 | 3.1 |  |  | 16
| 1 | 4 | < 10 | 3.2 | 3.2 | 3.5 | < 10 | 3.4 | 3.4 | 3.3 | 3 | 3.6 | 3.3 | 3.2 | 3.5 | 3.3 | 3.4 | 3.35 | 4.3 | 3.72 | < 5 | 2.7 |  | 3 | 2.7157 | 2.9 |  |  | 3.29
| 1 | 5 | < 10 | 3.1 | 3 | 3.6 | < 10 | 3.5 | 3.3 | 3.1 | 3.2 | 3.5 | 3.6 | 3.4 | 3.3 | 3.4 | 3.4 | 3.54 | 4.4 | 3.81 | < 5 | 3 | 2.2356 | 2 | 2.9466 | 3 |  |  | 67
| 1 | 6 | < 10 | 3.2 | 3.2 | 3.4 | < 10 | 3.5 | 3.8 | 3.05 | 3.2 | 3.4 | 3.3 | 3.3 | 3.3 | 3.4 | 3.5 | 3.26 | 4.3 | 3.75 | < 5 | 2.9 | 2.2263 | 3 | 3.1065 | 3.1
| Mean |  |  | 3.2333 | 3.2 | 3.4667 |  | 3.45 | 3.45 | 3.27 | 3.1 | 3.5 | 3.4667 | 3.3167 | 3.3167 | 3.3333 | 3.4333 | 3.475 | 4.3167 | 3.67 |  | 2.85 | 4.6529 | 2.8 | 2.9229 | 3
| Median |  |  | 3.2 | 3.2 | 3.45 |  | 3.45 | 3.4 | 3.315 | 3.1 | 3.5 | 3.5 | 3.3 | 3.3 | 3.35 | 3.45 | 3.405 | 4.3 | 3.66 |  | 2.9 | 2.98 | 3 | 2.9201 | 3
| Std Dev. |  |  | 0.18619 | 0.12649 | 0.08165 |  | 0.10488 | 0.18708 | 0.16236 | 0.10954 | 0.063246 | 0.13663 | 0.075277 | 0.098319 | 0.12111 | 0.08165 | 0.24517 | 0.040825 | 0.10354 |  | 0.12247 | 3.9122 | 0.83666 | 0.14364 | 0.089443
| Rel.Std.Dev. |  |  | 5.758% | 3.953% | 2.355% |  | 3.04% | 5.423% | 4.965% | 3.534% | 1.807% | 3.941% | 2.27% | 2.964% | 3.633% | 2.378% | 7.055% | 0.9457% | 2.821% |  | 4.297% | 84.08% | 29.88% | 4.914% | 2.981%
| PDM3 |  |  | -1.723% | -2.736% | 5.369% |  | 4.863% | 4.863% | -0.6083% | -5.775% | 6.383% | 5.369% | 0.8101% | 0.8101% | 1.317% | 4.356% | 5.623% | 31.21% | 11.55% |  | -13.37% | 41.43% | -14.89% | -11.16% | -8.815%
| Z-Score (Absolute) |  | 5.28 | 0.67 | 0.79 | 0.11 | 5.28 | 0.06 | 0.06 | 0.55 | 1.12 | 0.22 | 0.11 | 0.39 | 0.39 | 0.34 | 0 | 0.14 | 2.98 | 0.8 | 3.15 | 1.97 | 4.11 | NA | 1.72 | 1.46
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 609c (Certified Value 13.4 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 14.4 | 14.6 | 13.3 | 13.45 | 13.2058 | 12.6 | 13.4 | 12.3309 | 13.3 | 13.3 | 14.03 | 14.6 | 13.1 | 13.5 | 13.1 | 13.6 | 13.39 | 12.78 | 14.1 | 14.3 | 12.9809 | 14 | 14.0123 | 11.346 | 13.56 |  | 1
| 1 | 2 | 13.7 | 13.6 | 13 | 13.6 | 13.4589 | 12.5 | 13.8 | 12.2319 | 13.3 | 13.1 | 13.79 | 13.7 | 13.4 | 12.6 | 13.5 | 13.8 | 13.29 | 11.95 | 14.1 | 14.2 | 13.8694 | 14 | 13.2315 | 11.421 | 12.87 |  | 36
| 1 | 3 | 14.1 | 14 | 13.3 | 13.47 | 13.1632 | 12.7 | 14.2 | 12.2871 | 13.3 | 12.9 | 14.25 | 14.6 | 13.2 | 12.6 | 13.2 | 12.8 | 13.14 | 11.49 | 14.2 | 13.8 | 13.4226 | 14 | 13.968 | 10.947 | 12.11 |  | 16
| 1 | 4 | 14 | 14.3 | 13.3 | 13.1 | 13.3991 | 12.5 | 14 | 12.3912 | 13.5 | 13.2 | 13.84 | 13.8 | 13.2 | 13.5 | 13.6 | 13 | 12.9 | 12.52 | 15.1 | 13.9 | 13.6293 | 15 | 13.4112 | 11.446 | 13.62 |  | 13.4217
| 1 | 5 | 14.2 | 13.5 | 13.4 | 13.63 | 13.5496 | 13 | 13.8 | 13.2615 | 12.9 | 13.2 | 14.39 | 14.5 | 13.3 | 13.6 | 13.4 | 13.4 | 12.99 | 12.78 | 14.9 | 14 | 13.555 | 15 | 13.5394 | 11.018 | 12.54 |  | 68
| 1 | 6 | 14.9 | 14.1 | 13.4 | 13.39 | 13.414 | 13.2 | 14.1 | 12.2819 | 13.5 | 13 | 13.62 | 13.2 | 13.2 | 13.2 | 13 | 12.6 | 13.13 | 11.68 | 14.8 | 14.2 | 13.8743 | 14 | 14.0265 | 11.303 | 12.56
| Mean |  | 14.2167 | 14.0167 | 13.2833 | 13.44 | 13.3651 | 12.75 | 13.8833 | 12.4641 | 13.3 | 13.1167 | 13.9867 | 14.0667 | 13.2333 | 13.1667 | 13.3 | 13.2 | 13.14 | 12.2 | 14.5333 | 14.0667 | 13.5552 | 14.3333 | 13.6981 | 11.2468 | 12.8767
| Median |  | 14.15 | 14.05 | 13.3 | 13.46 | 13.4066 | 12.65 | 13.9 | 12.309 | 13.3 | 13.15 | 13.935 | 14.15 | 13.2 | 13.35 | 13.3 | 13.2 | 13.135 | 12.235 | 14.5 | 14.1 | 13.5922 | 14 | 13.7537 | 11.3245 | 12.715
| Std Dev. |  | 0.40702 | 0.41673 | 0.1472 | 0.19016 | 0.15001 | 0.2881 | 0.28577 | 0.3943 | 0.21909 | 0.1472 | 0.29276 | 0.58538 | 0.10328 | 0.45898 | 0.23664 | 0.47329 | 0.18177 | 0.56784 | 0.45019 | 0.19664 | 0.3327 | 0.5164 | 0.34775 | 0.21225 | 0.60348
| Rel.Std.Dev. |  | 2.863% | 2.973% | 1.108% | 1.415% | 1.122% | 2.26% | 2.058% | 3.163% | 1.647% | 1.122% | 2.093% | 4.161% | 0.7805% | 3.486% | 1.779% | 3.586% | 1.383% | 4.654% | 3.098% | 1.398% | 2.454% | 3.603% | 2.539% | 1.887% | 4.687%
| PDM3 |  | 5.923% | 4.433% | -1.031% | 0.1361% | -0.4218% | -5.005% | 3.439% | -7.135% | -0.907% | -2.273% | 4.209% | 4.805% | -1.404% | -1.9% | -0.907% | -1.652% | -2.099% | -9.103% | 8.282% | 4.805% | 0.9948% | 6.792% | 2.059% | -16.2% | -4.061%
| Z-Score (Absolute) |  | 1.5 | 1.18 | 0.03 | 0.23 | 0.11 | 0.9 | 0.96 | 1.37 | 0 | 0.3 | 1.13 | 1.26 | 0.11 | 0.22 | 0 | 0.16 | 0.26 | 1.81 | 2.02 | 1.26 | 0.42 | NA | 0.65 | 3.37 | 0.69
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 609c (Certified Value 0.58 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.10; 2.08: 2.12; 2.09: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 0.6; 2.03: 0.55; 2.04: 0.55086; 2.05: 0.6; 2.06: 0.6431; 2.07: 0.55; 2.08: 0.56; 2.09: 0.7; 2.10: 0.65; 2.11: 0.589; 2.12: 0.7; 2.13: 0.5 |  | 1
| 1 | 2 | 2.01: 0.5; 2.03: 0.55; 2.04: 0.56785; 2.05: 0.6; 2.06: 0.6126; 2.07: 0.51; 2.08: 0.56; 2.09: 0.6; 2.10: 0.63; 2.11: 0.6536; 2.12: 0.7; 2.13: 0.5 |  | 37
| 1 | 3 | 2.01: 0.5; 2.03: 0.98; 2.04: 0.56862; 2.05: 0.6; 2.06: 0.6258; 2.07: 0.52; 2.08: 0.57; 2.09: 0.6; 2.10: 0.63; 2.11: 0.6352; 2.12: 0.7; 2.13: 0.5 |  | 16
| 1 | 4 | 2.01: 0.6; 2.03: 0.55; 2.04: 0.57999; 2.05: 0.6; 2.06: 0.6624; 2.07: 0.52; 2.08: 0.6; 2.09: 0.6; 2.10: 0.63; 2.11: 0.637; 2.12: 0.7; 2.13: 0.5 |  | 0.58107
| 1 | 5 | 2.01: 0.5; 2.03: 0.54; 2.04: 0.60648; 2.05: 0.5; 2.06: 0.6743; 2.07: 0.5; 2.08: 0.59; 2.09: 0.7; 2.10: 0.62; 2.11: 0.6311; 2.12: 0.7; 2.13: 0.5 |  | 69
| 1 | 6 | 2.01: 0.6; 2.03: 0.53; 2.04: 0.57712; 2.05: 0.6; 2.06: 0.6601; 2.07: 0.52; 2.08: 0.56; 2.09: 0.6; 2.10: 0.64; 2.11: 0.6415; 2.12: 0.7; 2.13: 0.6
| Mean |  | 2.01: 0.55; 2.03: 0.61667; 2.04: 0.57515; 2.05: 0.58333; 2.06: 0.64638; 2.07: 0.52; 2.08: 0.57333; 2.09: 0.63333; 2.10: 0.63333; 2.11: 0.63123; 2.12: 0.7; 2.13: 0.51667
| Median |  | 2.01: 0.55; 2.03: 0.55; 2.04: 0.57287; 2.05: 0.6; 2.06: 0.6516; 2.07: 0.52; 2.08: 0.565; 2.09: 0.6; 2.10: 0.63; 2.11: 0.6361; 2.12: 0.7; 2.13: 0.5
| Std Dev. |  | 2.01: 0.054772; 2.03: 0.17818; 2.04: 0.018409; 2.05: 0.040825; 2.06: 0.023663; 2.07: 0.016733; 2.08: 0.017512; 2.09: 0.05164; 2.10: 0.010328; 2.11: 0.022083; 2.12: 1.2162e-16; 2.13: 0.040825
| Rel.Std.Dev. |  | 2.01: 9.959%; 2.03: 28.89%; 2.04: 3.201%; 2.05: 6.999%; 2.06: 3.661%; 2.07: 3.218%; 2.08: 3.054%; 2.09: 8.154%; 2.10: 1.631%; 2.11: 3.498%; 2.12: 1.737e-14%; 2.13: 7.902%
| PDM3 |  | 2.01: -5.347%; 2.03: 6.126%; 2.04: -1.018%; 2.05: 0.3897%; 2.06: 11.24%; 2.07: -10.51%; 2.08: -1.331%; 2.09: 8.995%; 2.10: 8.995%; 2.11: 8.633%; 2.12: 20.47%; 2.13: -11.08%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.38; 2.04: 0.38; 2.05: NA; 2.06: 0.94; 2.07: 1.41; 2.08: 0.42; 2.09: NA; 2.10: 0.69; 2.11: 0.65; 2.12: NA; 2.13: 1.47
Results from laboratories 2.03, 2.07, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 609c (Certified Value 1078 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 1096 | 1123.2 | 1126 | 1079 | 1066 | 965 | 1000 | 1125.192 | 1066 | 1090 | 1107 | 1110 | 1105 | 1020 | 1080 | 1180 | 1163 | 1065 | 1230 | 1008 | 1078 | 922.2698 | 1050 | 1083.1517 | 1015.16 | 1077 | 1
| 1 | 2 | 1096 | 1141.2 | 1085 | 1093 | 1069.0484 | 935 | 1020 | 1128.9936 | 1048 | 1040 | 1128 | 1125 | 1115 | 1045 | 1100 | 1170 | 1159 | 1045 | 1240 | 1000.7 | 1077 | 889.5557 | 1040 | 1029.4495 | 998.609 | 1078 | 18
| 1 | 3 | 1086 | 1079.7 | 1121 | 1087 | 1059.1044 | 977 | 1050 | 1125.0992 | 1069 | 1060 | 1134 | 1100 | 1125 | 1080 | 1055 | 1235 | 1139 | 1047 | 1220 | 1013 | 1063 | 903.1578 | 1020 | 1086.922 | 932.729 | 1084 | 16
| 1 | 4 | 1096 | 1135.4 | 1129 | 1081 | 1069.7942 | 957 | 1040 | 1120.936 | 1056 | 1060 | 1169 | 1085 | 1085 | 1080 | 1090 | 1200 | 1108 | 1059 | 1300 | 1041 | 1082 | 881.3279 | 1120 | 1041.9251 | 1002.578 | 1094 | 1078.2733
| 1 | 5 | 1111 | 1118.3 | 1112 | 1079 | 1092.9727 | 962 | 1030 | 1134.264 | 1048 | 1070 | 1094 | 1130 | 1115 | 1080 | 1070 | 1210 | 1116 | 1064 | 1300 | 1010 | 1059 | 926.2187 | 1050 | 1058.4018 | 1010.332 | 1067 | 70
| 1 | 6 | 1093 | 1118.8 | 1133 | 1077 | 1065.675 | 998 | 1050 | 1115.448 | 1048 | 1050 | 1144 | 1075 | 1115 | 1060 | 1065 | 1190 | 1151 | 1056 | 1260 | 1000.5 | 1072 | 903.9171 | 1030 | 1106.406 | 1022.596 | 1113
| Mean |  | 1096.3333 | 1119.4333 | 1117.6667 | 1082.6667 | 1070.4324 | 965.6667 | 1031.6667 | 1124.9888 | 1055.8333 | 1061.6667 | 1129.3333 | 1104.1667 | 1110 | 1060.8333 | 1076.6667 | 1197.5 | 1139.3333 | 1056 | 1258.3333 | 1012.2 | 1071.8333 | 904.4078 | 1051.6667 | 1067.7093 | 997.0007 | 1085.5
| Median |  | 1096 | 1121 | 1123.5 | 1080 | 1067.5242 | 963.5 | 1035 | 1125.1456 | 1052 | 1060 | 1131 | 1105 | 1115 | 1070 | 1075 | 1195 | 1145 | 1057.5 | 1250 | 1009 | 1074.5 | 903.5375 | 1045 | 1070.7767 | 1006.455 | 1081
| Std Dev. |  | 8.165 | 21.5561 | 17.5689 | 6.121 | 11.6704 | 20.9921 | 19.4079 | 6.4742 | 9.6003 | 17.224 | 26.6808 | 21.7754 | 13.784 | 24.5798 | 16.6333 | 23.184 | 22.8444 | 8.438 | 34.8807 | 14.978 | 9.0646 | 17.6059 | 35.4495 | 29.3949 | 32.6409 | 16.1338
| Rel.Std.Dev. |  | 0.7448% | 1.926% | 1.572% | 0.5654% | 1.09% | 2.174% | 1.881% | 0.5755% | 0.9093% | 1.622% | 2.363% | 1.972% | 1.242% | 2.317% | 1.545% | 1.936% | 2.005% | 0.7991% | 2.772% | 1.48% | 0.8457% | 1.947% | 3.371% | 2.753% | 3.274% | 1.486%
| PDM3 |  | 1.675% | 3.817% | 3.653% | 0.4074% | -0.7272% | -10.44% | -4.322% | 4.332% | -2.081% | -1.54% | 4.735% | 2.401% | 2.942% | -1.617% | -0.149% | 11.06% | 5.663% | -2.066% | 16.7% | -6.128% | -0.5972% | -16.12% | -2.468% | -0.9797% | -7.537% | 0.6702%
| Z-Score (Absolute) |  | 0.45 | 0.93 | 0.89 | 0.17 | 0.08 | 2.23 | 0.87 | 1.04 | 0.38 | 0.26 | 1.13 | 0.61 | 0.73 | 0.28 | 0.05 | 2.53 | 1.34 | 0.37 | 3.78 | 1.27 | 0.05 | 3.49 | 0.46 | 0.13 | 1.59 | 0.23
| Analytical results for Zr in OREAS 609c (Certified Value 238 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 233 | 224.4 | 230.3 | 233.9 | 237.6844 | 235 | 227 | 198.861 | 246 | 235 | 238.1 | 244 | 227 | 228 | 217 | 239 | 235.6 | 246 | 247 | 265.2 | 233 | 253.5321 | 271 | 225.4758 | 209.01 | 254 | 1
| 1 | 2 | 218 | 228.6 | 217 | 236.5 | 240.3559 | 228 | 232 | 194.9115 | 235 | 236 | 236.2 | 241 | 229 | 236 | 223 | 240 | 236.3 | 251 | 246 | 260.8 | 243 | 253.4366 | 262 | 216.7039 | 211.14 | 254 | 19
| 1 | 3 | 226 | 225.2 | 226.1 | 238.5 | 236.5015 | 230 | 244 | 193.382 | 235 | 233 | 247.4 | 241 | 228 | 238 | 215 | 243 | 229.6 | 245 | 250 | 262.6 | 241.4 | 244.4839 | 261 | 218.2725 | 206.14 | 263 | 16
| 1 | 4 | 230 | 231.8 | 227.4 | 231.9 | 239.7198 | 231 | 235 | 199.3955 | 239 | 238 | 240.6 | 232 | 225 | 246 | 226 | 242 | 230.7 | 252 | 256 | 267.5 | 242 | 242.0533 | 286 | 213.8296 | 210.3 | 269 | 238.1314
| 1 | 5 | 236 | 224.6 | 220.7 | 236.1 | 238.5598 | 238 | 238 | 195.0065 | 233 | 242 | 244.7 | 242 | 228 | 248 | 218 | 245 | 232.9 | 251 | 261 | 271.1 | 246.5 | 252.0765 | 271 | 218.5218 | 204.24 | 254 | 71
| 1 | 6 | 234 | 223 | 229.4 | 235.3 | 242.1514 | 238 | 240 | 198.3445 | 249 | 233 | 240.7 | 231 | 229 | 243 | 214 | 241 | 228.8 | 252 | 252 | 268.4 | 240.5 | 250.4984 | 262 | 223.7448 | 210.54 | 272
| Mean |  | 229.5 | 226.2667 | 225.15 | 235.3667 | 239.1621 | 233.3333 | 236 | 196.6502 | 239.5 | 236.1667 | 241.2833 | 238.5 | 227.6667 | 239.8333 | 218.8333 | 241.6667 | 232.3167 | 249.5 | 252 | 265.9333 | 241.0667 | 249.3468 | 268.8333 | 219.4248 | 208.5617 | 261
| Median |  | 231.5 | 224.9 | 226.75 | 235.7 | 239.1398 | 233 | 236.5 | 196.6755 | 237 | 235.5 | 240.65 | 241 | 228 | 240.5 | 217.5 | 241.5 | 231.8 | 251 | 251 | 266.35 | 241.7 | 251.2875 | 266.5 | 218.3972 | 209.655 | 258.5
| Std Dev. |  | 6.6257 | 3.2904 | 5.2302 | 2.2722 | 2.0148 | 4.274 | 6.0332 | 2.518 | 6.5651 | 3.4303 | 4.1441 | 5.5408 | 1.5055 | 7.3869 | 4.7081 | 2.1602 | 3.1416 | 3.1464 | 5.6921 | 3.8271 | 4.4639 | 4.8962 | 9.5795 | 4.3847 | 2.7686 | 8.1976
| Rel.Std.Dev. |  | 2.887% | 1.454% | 2.323% | 0.9654% | 0.8424% | 1.832% | 2.556% | 1.28% | 2.741% | 1.452% | 1.718% | 2.323% | 0.6613% | 3.08% | 2.151% | 0.8939% | 1.352% | 1.261% | 2.259% | 1.439% | 1.852% | 1.964% | 3.563% | 1.998% | 1.327% | 3.141%
| PDM3 |  | -3.625% | -4.982% | -5.451% | -1.161% | 0.4328% | -2.015% | -0.8951% | -17.42% | 0.5747% | -0.8251% | 1.324% | 0.1548% | -4.395% | 0.7147% | -8.104% | 1.485% | -2.442% | 4.774% | 5.824% | 11.68% | 1.233% | 4.71% | 12.89% | -7.856% | -12.42% | 9.603%
| Z-Score (Absolute) |  | 0.6 | 0.85 | 0.94 | 0.15 | 0.14 | 0.31 | 0.1 | 3.14 | 0.17 | 0.09 | 0.3 | 0.09 | 0.74 | 0.19 | 1.43 | 0.33 | 0.39 | 0.94 | 1.13 | 2.2 | 0.29 | 0.93 | 2.43 | 1.38 | 2.22 | 1.82
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C897: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1120: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1138: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 609c (Certified Value 24 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 2
| 1 | 1 | 23.5 | 23.56 | 24 | 23.4 | 46.4597 | 23.1 | 24 | 25.229 | 23.9 | 23.9 | 22.2 | 23.7 | 24 | 25.1 | 25.1 | 26.3 | 22.8 | 23.4 | 23.9 | 25.4 | 23.74 | 23.0981 | 3.37 | 23.951 | 21.63 | 22.71 | 1
| 1 | 2 | 23.5 | 19.637 | 24 | 23.21 | 24.0848 | 23.6 | 24.3 | 24.6725 | 24 | 23.8 | 23 | 23.4 | 24.7 | 23.9 | 25 | 26.7 | 22.9 | 23.2 | 22.5 | 25.03 | 25.02 | 381.25 | 8.21 | 24.721 | 21.87 | 23.59 | 23
| 1 | 3 | 23.2 | 23.08 | 23.9 | 23.54 | 24.4343 | 23.3 | 24.1 | 24.9455 | 24 | 23.2 | 22 | 23.3 | 24.9 | 24.5 | 24.9 | 26.6 | 22.4 | 23 | 23.1 | 25.92 | 23.68 | 89.5256 | 3.9 | 24.049 | 22.24 | 24.37 | 16
| 1 | 4 | 23.2 | 22.44 | 23.9 | 23.37 | 23.7694 | 23.3 | 24 | 25.3655 | 24.1 | 23.9 | 22.8 | 23.1 | 24.1 | 24 | 25.3 | 25.5 | 23.3 | 23 | 23.5 | 25.36 | 24.46 | 166.8863 | 3.34 | 24.92 | 21.86 | 25.38 | 23.9683
| 1 | 5 | 23.3 | 23.4 | 23.9 | 23.03 | 24.1319 | 24 | 23.8 | 25.8905 | 23.8 | 23.8 | 21.9 | 24.1 | 25.4 | 25 | 24 | 26.5 | 23.1 | 23.3 | 23.2 | 26.56 | 23.56 | 353.781 | 3.04 | 24.6442 | 22 | 26.14 | 73
| 1 | 6 | 23.5 | 23.9 | 24.1 | 23.14 | 24.2596 | 23.4 | 24.1 | 25.754 | 24.3 | 23.5 | 23.4 | 23.3 | 24.8 | 24.4 | 24.9 | 26.2 | 22.9 | 23.5 | 23.6 | 25.98 | 24.69 | 43.5579 | 2.47 | 24.3705 | 21.79 | 24.94
| Mean |  | 23.3667 | 22.6695 | 23.9667 | 23.2817 | 27.8566 | 23.45 | 24.05 | 25.3095 | 24.0167 | 23.6833 | 22.55 | 23.4833 | 24.65 | 24.4833 | 24.8667 | 26.3 | 22.9 | 23.2333 | 23.3 | 25.7083 | 24.1917 | 176.3498 | 4.055 | 24.4426 | 21.8983 | 24.5217
| Median |  | 23.4 | 23.24 | 23.95 | 23.29 | 24.1958 | 23.35 | 24.05 | 25.2972 | 24 | 23.8 | 22.5 | 23.35 | 24.75 | 24.45 | 24.95 | 26.4 | 22.9 | 23.25 | 23.35 | 25.66 | 24.1 | 128.206 | 3.355 | 24.5074 | 21.865 | 24.655
| Std Dev. |  | 0.15055 | 1.5658 | 0.08165 | 0.18819 | 9.1163 | 0.31464 | 0.16432 | 0.46541 | 0.17224 | 0.27869 | 0.60581 | 0.36009 | 0.5244 | 0.49565 | 0.45019 | 0.43359 | 0.30332 | 0.20656 | 0.4858 | 0.55145 | 0.61177 | 156.3417 | 2.0886 | 0.38674 | 0.20644 | 1.2413
| Rel.Std.Dev. |  | 0.6443% | 6.907% | 0.3407% | 0.8083% | 32.73% | 1.342% | 0.6832% | 1.839% | 0.7172% | 1.177% | 2.686% | 1.533% | 2.127% | 2.024% | 1.81% | 1.649% | 1.325% | 0.8891% | 2.085% | 2.145% | 2.529% | 88.65% | 51.51% | 1.582% | 0.9427% | 5.062%
| PDM3 |  | -2.51% | -5.419% | -0.0068% | -2.865% | 16.22% | -2.162% | 0.3409% | 5.596% | 0.2018% | -1.189% | -5.917% | -2.023% | 2.844% | 2.149% | 3.748% | 9.728% | -4.457% | -3.066% | -2.788% | 7.26% | 0.9319% | 635.8% | -83.08% | 1.979% | -8.636% | 2.309%
| Z-Score (Absolute) |  | 0.6 | 1.27 | 0.02 | 0.68 | 3.72 | 0.52 | 0.06 | 1.27 | 0.02 | 0.3 | 1.39 | 0.49 | 0.63 | 0.47 | 0.84 | 2.22 | 1.05 | 0.73 | 0.67 | 1.65 | 0.19 | 146.59 | 19.18 | 0.43 | 2.01 | 0.51
| Analytical results for Al in OREAS 609c (Certified Value 0.561 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.58 | 0.56 | 0.58 | 0.5566 | 0.73231 | 0.58 | 0.4 | 0.54 | 0.844 | 0.7819 | 0.53 | 0.56 | 0.56 | 0.63 | 0.55 | 0.553 | 0.51 | 0.49 | 0.59 | 0.5579 | 0.497 | 0.5372 | 0.62728 |  |  |  | 1
| 1 | 2 | 0.61 | 0.54 | 0.6 | 0.5731 | 0.54858 | 0.57 | 0.4 | 0.53 | 0.832 | 0.8073 | 0.51 | 0.55 | 0.56 | 0.64 | 0.57 | 0.535 | 0.443 | 0.49 | 0.59 | 0.5918 | 0.656 | 0.54296 | 0.64364 |  |  |  | 0
| 1 | 3 | 0.58 | 0.52 | 0.6 | 0.5651 | 0.56881 | 0.6 | 0.41 | 0.54 | 0.851 | 0.7795 | 0.51 | 0.55 | 0.54 | 0.63 | 0.57 | 0.533 | 0.53 | 0.5 | 0.57 | 0.6043 | 0.497 | 0.54394 | 0.63132 |  |  |  | 16
| 1 | 4 | 0.58 | 0.5 | 0.6 | 0.5791 | 0.55542 | 0.6 | 0.4 | 0.53 | 0.828 | 0.805 | 0.52 | 0.65 | 0.55 | 0.62 | 0.54 | 0.567 | 0.545 | 0.49 | 0.6 | 0.6147 | 0.476 | 0.57009 | 0.64257 |  |  |  | 0.56095
| 1 | 5 | 0.58 | 0.55 | 0.6 | 0.5542 | 0.55569 | 0.6 | 0.39 | 0.53 | 0.84 | 0.7793 | 0.52 | 0.56 | 0.57 | 0.62 | 0.56 | 0.581 | 0.525 | 0.49 | 0.6 | 0.634 | 0.466 | 0.56832 | 0.58177 |  |  |  | 74
| 1 | 6 | 0.57 | 0.57 | 0.59 | 0.5665 | 0.56347 | 0.63 | 0.4 | 0.55 | 0.858 | 0.8334 | 0.51 | 0.55 | 0.55 | 0.62 | 0.55 | 0.579 | 0.54 | 0.48 | 0.6 | 0.5852 | 0.476 | 0.5483 | 0.64188
| Mean |  | 0.58333 | 0.54 | 0.595 | 0.56577 | 0.58738 | 0.59667 | 0.4 | 0.53667 | 0.84217 | 0.79773 | 0.51667 | 0.57 | 0.555 | 0.62667 | 0.55667 | 0.558 | 0.5155 | 0.49 | 0.59167 | 0.59798 | 0.51133 | 0.5518 | 0.62808
| Median |  | 0.58 | 0.545 | 0.6 | 0.5658 | 0.55958 | 0.6 | 0.4 | 0.535 | 0.842 | 0.79345 | 0.515 | 0.555 | 0.555 | 0.625 | 0.555 | 0.56 | 0.5275 | 0.49 | 0.595 | 0.59805 | 0.4865 | 0.54612 | 0.6366
| Std Dev. |  | 0.013663 | 0.026077 | 0.0083666 | 0.0094906 | 0.071349 | 0.020656 | 0.0063246 | 0.008165 | 0.011321 | 0.02163 | 0.008165 | 0.039497 | 0.010488 | 0.008165 | 0.012111 | 0.021119 | 0.03757 | 0.0063246 | 0.01169 | 0.026172 | 0.07196 | 0.013949 | 0.023658
| Rel.Std.Dev. |  | 2.342% | 4.829% | 1.406% | 1.677% | 12.15% | 3.462% | 1.581% | 1.521% | 1.344% | 2.711% | 1.58% | 6.929% | 1.89% | 1.303% | 2.176% | 3.785% | 7.288% | 1.291% | 1.976% | 4.377% | 14.07% | 2.528% | 3.767%
| PDM3 |  | 3.991% | -3.734% | 6.071% | 0.8591% | 4.712% | 6.368% | -28.69% | -4.329% | 50.13% | 42.21% | -7.894% | 1.614% | -1.06% | 11.72% | -0.7632% | -0.5255% | -8.102% | -12.65% | 5.476% | 6.602% | -8.845% | -1.63% | 11.97%
| Z-Score (Absolute) |  | 0.41 | 0.59 | 0.67 | 0 | 0.5 | 0.71 | 3.82 | 0.67 | 6.38 | 5.35 | 1.13 | 0.1 | 0.25 | 1.4 | 0.21 | 0.18 | 1.16 | 1.75 | 0.6 | 0.74 | 1.25 | 0.32 | 1.44
| Analytical results for As in OREAS 609c (Certified Value 852 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 889 | 841.1 | 879.6 | 897 | 2722.5323 | 772 | 875 | 902.714 | 851 | 875 | 869 | 865 | 893 | 845 | 929 | 894 | 804.2 | 689 | 927 | 770.2 | 862.6 | 784.6227 | 821 | 727.1693 | 837.52 | 698 | 1
| 1 | 2 | 880 | 808 | 886.5 | 898 | 877.9234 | 788 | 871 | 904.0515 | 844 | 867 | 889 | 859 | 886 | 854 | 934 | 894 | 798.3 | 685 | 933 | 764.5 | 895.7 | 787.7712 | 2710 | 751.9917 | 810.886 | 720 | 13
| 1 | 3 | 883 | 821.5 | 883.2 | 889 | 870.9527 | 778 | 858 | 909.367 | 851 | 890 | 840 | 860 | 880 | 842 | 920 | 882 | 785 | 683 | 922 | 785.6 | 837.3 | 785.7211 | 805 | 734.1295 | 808.176 | 734 | 16
| 1 | 4 | 884 | 801.4 | 878.4 | 887 | 860.346 | 779 | 850 | 911.156 | 853 | 877 | 880 | 854 | 881 | 842 | 921 | 859 | 806.5 | 678 | 927 | 774.9 | 901.4 | 775.0382 | 808 | 734.8845 | 796.374 | 765 | 851.5309
| 1 | 5 | 885 | 837.6 | 879.9 | 881 | 851.9926 | 778 | 833 | 911.9645 | 849 | 873 | 817 | 871 | 894 | 874 | 904 | 887 | 804 | 674 | 934 | 791.1 | 858.2 | 785.1856 | 828 | 750.1283 | 816.92 | 815 | 75
| 1 | 6 | 895 | 851.8 | 888.1 | 889 | 871.8237 | 789 | 859 | 905.9415 | 861 | 874 | 882 | 864 | 885 | 842 | 912 | 879 | 797.6 | 682 | 927 | 778.5 | 875 | 780.7813 | 828 | 735.2364 | 835.31 | 749
| Mean |  | 886 | 826.9 | 882.6167 | 890.1667 | 1175.9285 | 780.6667 | 857.6667 | 907.5324 | 851.5 | 876 | 862.8333 | 862.1667 | 886.5 | 849.8333 | 920 | 882.5 | 799.2667 | 681.8333 | 928.3333 | 777.4667 | 871.7 | 783.1867 | 1133.3333 | 738.9233 | 817.531 | 746.8333
| Median |  | 884.5 | 829.55 | 881.55 | 889 | 871.3882 | 778.5 | 858.5 | 907.6543 | 851 | 874.5 | 874.5 | 862 | 885.5 | 843.5 | 920.5 | 884.5 | 801.15 | 682.5 | 927 | 776.7 | 868.8 | 784.9041 | 824.5 | 735.0604 | 813.903 | 741.5
| Std Dev. |  | 5.2915 | 19.8633 | 3.9937 | 6.4005 | 757.7344 | 6.5625 | 15.1482 | 3.8472 | 5.5767 | 7.642 | 28.3084 | 5.8452 | 5.8907 | 12.7188 | 10.9362 | 13.0346 | 7.827 | 5.2694 | 4.4572 | 9.8042 | 24.1553 | 4.5974 | 772.4676 | 9.873 | 16.0954 | 40.6222
| Rel.Std.Dev. |  | 0.5972% | 2.402% | 0.4525% | 0.719% | 64.44% | 0.8406% | 1.766% | 0.4239% | 0.6549% | 0.8724% | 3.281% | 0.678% | 0.6645% | 1.497% | 1.189% | 1.477% | 0.9793% | 0.7728% | 0.4801% | 1.261% | 2.771% | 0.587% | 68.16% | 1.336% | 1.969% | 5.439%
| PDM3 |  | 4.048% | -2.893% | 3.651% | 4.537% | 38.1% | -8.322% | 0.7206% | 6.577% | -0.003626% | 2.874% | 1.327% | 1.249% | 4.107% | -0.1994% | 8.041% | 3.637% | -6.138% | -19.93% | 9.019% | -8.698% | 2.369% | -8.026% | 33.09% | -13.22% | -3.993% | -12.3%
| Z-Score (Absolute) |  | 0.39 | 0.6 | 0.34 | 0.46 | 5.25 | 1.37 | 0.08 | 0.75 | 0.18 | 0.23 | 0.01 | 0.01 | 0.4 | 0.21 | 0.96 | 0.33 | 1.06 | 3.02 | 1.1 | 1.42 | 0.15 | 1.33 | 4.53 | 2.07 | 0.75 | 1.94
| Analytical results for B in OREAS 609c (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab | 1
| No | No | 2.01: 2.01; 2.03: 2.03; 2.04: 2.04; 2.05: 2.06; 2.06: 2.07; 2.07: 2.08; 2.08: 2.13; 2.09: 2.14; 2.10: 2.15; 2.11: 2.16; 2.12: 2.17; 2.13: 2.23; 2.14: 2.26 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*OES; 2.14: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 15g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.5g; 2.14: 01g | 1
| 1 | 1 | 2.01: < 10; 2.03: < 20; 2.04: 1; 2.05: < 2; 2.06: 2; 2.07: 2.1; 2.08: < 10; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 0.1; 2.14: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 1.8; 2.08: < 10; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 0.1; 2.14: < 10 | 3
| 1 | 3 | 2.01: < 10; 2.03: < 20; 2.04: 1; 2.05: < 2; 2.06: 2; 2.07: 1.9; 2.08: < 10; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 0.1; 2.14: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.03: < 20; 2.04: 1; 2.05: < 2; 2.06: 2; 2.07: 1.8; 2.08: < 10; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 0.1; 2.14: < 10 | < 10
| 1 | 5 | 2.01: < 10; 2.03: < 20; 2.04: 1; 2.05: < 2; 2.06: 2; 2.07: 1.8; 2.08: < 10; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 0.1; 2.14: < 10 | 76
| 1 | 6 | 2.01: < 10; 2.03: < 20; 2.04: 1; 2.05: < 2; 2.06: 2; 2.07: 1.7; 2.08: < 10; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 0.1; 2.14: < 10
| Mean |  | 2.04: 1.1667; 2.06: 2; 2.07: 1.85
| Median |  | 2.04: 1; 2.06: 2; 2.07: 1.8
| Std Dev. |  | 2.04: 0.40825; 2.06: 0; 2.07: 0.13784
| Rel.Std.Dev. |  | 2.04: 34.99%; 2.06: 0%; 2.07: 7.451%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.14
| Analytical results for Ba in OREAS 609c (Indicative Value 99 ppm) | [27 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 0
| 1 | 1 | 50 | 127.3 | 64 | 176 | 155.3225 | 16.5 | 458 | 519 | 540 | 130 | 80 | 160 | 330 | 110 | 169 | 46 | 481.6 | 510 | 507.3511 | 23 | 124.7293 | 486.919 | 43.4 |  |  |  | 1
| 1 | 2 | 51 | 140.8 | 87 | 102 | 460.1448 | 14.9 | 480 | 532 | 566 | 130 | 80 | 210 | 320 | 90 | 165 | 50 | 477.5 | 503 | 525.0993 | 9 | 108.6606 | 485.446 | 44.5 |  |  |  | 5
| 1 | 3 | 47 | 114.8 | 103 | 120 | 446.9178 | 14.7 | 489 | 537 | 537 | 120 | 130 | 130 | 360 | 90 | 186 | 57 | 485.1 | 508 | 526.353 | 23 | 128.9497 | 495.73 | 45 |  |  |  | 16
| 1 | 4 | 49 | 120.7 | 58 | 76 | 458.1959 | 15.4 | 481 | 530 | 559 | 120 | 160 | 110 | 210 | 150 | 160 | 50 | 478.3 | 518 | 530.8086 | 25 | 104.8195 | 487.082 | 46.6 |  |  |  | 99.2922
| 1 | 5 | 48 | 133.6 | 71 | 142 | 457.5891 | 16.5 | 460 | 525 | 529 | 100 | 80 | 170 | 280 | 100 | 160 | 50 | 479.9 | 515 | 540.4534 | 25 | 121.0311 | 490.308 | 42.3 |  |  |  | 11
| 1 | 6 | 49 | 129.2 | 94 | 130 | 463.5079 | 13.1 | 476 | 526 | 565 | 140 | 80 | 110 | 230 | 80 | 163 | 50 | 473.2 | 521 | 526.7456 | 24 | 111.4053 | 482.745 | 46.8
| Mean |  | 49 | 127.7333 | 79.5 | 124.3333 | 406.9463 | 15.1833 | 474 | 528.1667 | 549.3333 | 123.3333 | 101.6667 | 148.3333 | 288.3333 | 103.3333 | 167.1667 | 50.5 | 479.2667 | 512.5 | 526.1352 | 21.5 | 116.5993 | 488.0383 | 44.7667
| Median |  | 49 | 128.25 | 79 | 125 | 457.8925 | 15.15 | 478 | 528 | 549.5 | 125 | 80 | 145 | 300 | 95 | 164 | 50 | 479.1 | 512.5 | 526.5493 | 23.5 | 116.2182 | 487.0005 | 44.75
| Std Dev. |  | 1.4142 | 9.2112 | 17.8522 | 34.2559 | 123.3959 | 1.2781 | 12.3774 | 6.2423 | 15.9332 | 13.6626 | 34.8807 | 39.2003 | 59.1326 | 25.0333 | 9.8268 | 3.5637 | 4.0183 | 6.7157 | 10.7813 | 6.1887 | 9.6647 | 4.4993 | 1.7648
| Rel.Std.Dev. |  | 2.886% | 7.211% | 22.46% | 27.55% | 30.32% | 8.418% | 2.611% | 1.182% | 2.9% | 11.08% | 34.31% | 26.43% | 20.51% | 24.23% | 5.878% | 7.057% | 0.8384% | 1.31% | 2.049% | 28.78% | 8.289% | 0.9219% | 3.942%
| PDM3 |  | -50.65% | 28.64% | -19.93% | 25.22% | 309.8% | -84.71% | 377.4% | 431.9% | 453.2% | 24.21% | 2.391% | 49.39% | 190.4% | 4.07% | 68.36% | -49.14% | 382.7% | 416.2% | 429.9% | -78.35% | 17.43% | 391.5% | -54.91%
| Z-Score (Absolute) |  | 0.64 | 0 | 0.39 | 0.03 | 2.27 | 0.91 | 2.81 | 3.25 | 3.43 | 0.04 | 0.21 | 0.17 | 1.31 | 0.2 | 0.32 | 0.63 | 2.86 | 3.13 | 3.24 | 0.86 | 0.09 | 2.93 | 0.67
| Analytical results for Be in OREAS 609c (Certified Value 0.62 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 3
| 1 | 1 | 0.65 | 0.8 | 0.61 | < 0.5 | 0.6 | 0.64 | 0.326 | 0.7 | 0.68 | 0.61 | 0.65 | 0.6 | 0.69 | 0.62 | 0.58 | 0.74 | < 2 | 0.53 | 0.7998 | 0.55 | 0.59393 | < 5 | 0.48 |  |  |  | 1
| 1 | 2 | 0.64 | 0.7 | 0.62 | 0.6538 | 0.7 | 0.64 | 0.3377 | 0.8 | 0.73 | 0.6 | 0.64 | 0.61 | 0.7 | 0.64 | 0.57 | 0.82 | < 2 | 0.53 | 0.8064 | 0.33 | 0.64456 | < 5 | 0.51 |  |  |  | 26
| 1 | 3 | 0.68 | 0.8 | 0.64 | 0.64449 | 0.6 | 0.6 | 0.3529 | 0.8 | 0.73 | 0.6 | 0.64 | 0.59 | 0.68 | 0.63 | 0.59 | 0.53 | < 2 | 0.59 | 0.7587 | 0.59 | 0.62972 | < 5 | 0.5 |  |  |  | 16
| 1 | 4 | 0.67 | 0.8 | 0.65 | 0.64769 | 0.6 | 0.68 | 0.3445 | 0.7 | 0.71 | 0.6 | 0.67 | 0.6 | 0.68 | 0.61 | 0.59 | 0.53 | < 2 | 0.57 | 0.8228 | 0.52 | 0.58586 | < 5 | 0.51 |  |  |  | 0.61727
| 1 | 5 | 0.71 | 0.7 | 0.64 | 0.64352 | 0.6 | 0.71 | 0.3916 | 0.8 | 0.69 | 0.6 | 0.65 | 0.62 | 0.67 | 0.63 | 0.58 | 0.59 | < 2 | 0.52 | 0.8367 | 0.55 | 0.63604 | < 5 | 0.53 |  |  |  | 77
| 1 | 6 | 0.67 | 0.8 | 0.62 | 0.65289 | 0.6 | 0.61 | 0.3985 | 0.8 | 0.73 | 0.6 | 0.64 | 0.59 | 0.67 | 0.62 | 0.57 | 0.63 | < 2 | 0.59 | 0.8389 | 0.56 | 0.56853 | < 5 | 0.52
| Mean |  | 0.67 | 0.76667 | 0.63 | 0.64848 | 0.61667 | 0.64667 | 0.35853 | 0.76667 | 0.71167 | 0.60167 | 0.64833 | 0.60167 | 0.68167 | 0.625 | 0.58 | 0.64 |  | 0.555 | 0.81055 | 0.51667 | 0.60977 |  | 0.50833
| Median |  | 0.67 | 0.8 | 0.63 | 0.64769 | 0.6 | 0.64 | 0.3487 | 0.8 | 0.72 | 0.6 | 0.645 | 0.6 | 0.68 | 0.625 | 0.58 | 0.61 |  | 0.55 | 0.8146 | 0.55 | 0.61182 |  | 0.51
| Std Dev. |  | 0.024495 | 0.05164 | 0.015492 | 0.0047153 | 0.040825 | 0.041793 | 0.029705 | 0.05164 | 0.022286 | 0.0040825 | 0.01169 | 0.01169 | 0.01169 | 0.010488 | 0.0089443 | 0.11764 |  | 0.032094 | 0.029862 | 0.094163 | 0.031053 |  | 0.017224
| Rel.Std.Dev. |  | 3.656% | 6.736% | 2.459% | 0.7271% | 6.62% | 6.463% | 8.285% | 6.736% | 3.132% | 0.6785% | 1.803% | 1.943% | 1.715% | 1.678% | 1.542% | 18.38% |  | 5.783% | 3.684% | 18.23% | 5.093% |  | 3.388%
| PDM3 |  | 8.543% | 24.2% | 2.063% | 5.057% | -0.09703% | 4.763% | -41.92% | 24.2% | 15.29% | -2.527% | 5.033% | -2.527% | 10.43% | 1.253% | -6.037% | 3.683% |  | -10.09% | 31.31% | -16.3% | -1.214% |  | -17.65%
| Z-Score (Absolute) |  | 0.62 | NA | 0.04 | 0.66 | NA | 0.28 | 3.9 | NA | 1.22 | 0.37 | 0.3 | 0.37 | 0.79 | 0.04 | 0.69 | 0.18 | 5.41 | 1.05 | 2.66 | 1.61 | 0.26 | 27.17 | 1.73
Results from laboratories 2.03, 2.07 and 2.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 609c (Certified Value 65 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 1
| 1 | 1 | 64 | 61.39 | 69.4 | 68.78 | 206.4618 | 57.6 | 67.8 | 65.98 | 62.3 | 67.42 | 68.7 | 62.9 | 71.4 | 69.6 | 67.7 | 67.98 | 55.24 | 67.5 | 61.7 | 64.1 | 56.4152 | 68.5 | 61.0005 | 54.661 | 49.65 |  | 1
| 1 | 2 | 67 | 57.88 | 72 | 68.99 | 66.1292 | 59.2 | 67.4 | 65.72 | 63.9 | 68.61 | 71.1 | 65 | 69.8 | 69.8 | 67.7 | 68.98 | 55.73 | 63.8 | 60.9 | 65.52 | 50.7571 | 220 | 63.8807 | 55.43 | 51.88 |  | 27
| 1 | 3 | 66 | 65.13 | 72.9 | 69.58 | 67.5396 | 58 | 69.3 | 65.52 | 61.9 | 66.24 | 70 | 62.2 | 71.5 | 68 | 68.3 | 66.98 | 56.21 | 65.6 | 61.4 | 62.26 | 55.4146 | 68.4 | 60.4501 | 56.462 | 55.27 |  | 16
| 1 | 4 | 66 | 64.14 | 70.9 | 68 | 66.1269 | 58.1 | 68 | 66.98 | 63.1 | 69.07 | 69.6 | 62 | 69.5 | 69.8 | 64.5 | 68.98 | 55.58 | 66.8 | 60.8 | 65.46 | 54.0181 | 66.8 | 63.2087 | 55.897 | 56.22 |  | 64.81
| 1 | 5 | 66 | 62.92 | 70.3 | 68.62 | 67.806 | 59.3 | 67.8 | 65.81 | 62 | 64.42 | 68.7 | 62.9 | 72.5 | 67.1 | 68.5 | 69.98 | 54.17 | 66 | 62 | 62.56 | 52.0646 | 69.2 | 63.6965 | 56.701 | 58.4 |  | 78
| 1 | 6 | 64 | 65.19 | 72.9 | 68.94 | 67.1923 | 59.3 | 67.2 | 66.95 | 62 | 69.91 | 67 | 62.8 | 70.2 | 70 | 67.8 | 68.98 | 54.82 | 67.2 | 61.2 | 64.27 | 45.111 | 68.8 | 60.4083 | 55.247 | 55.99
| Mean |  | 65.5 | 62.775 | 71.4 | 68.8183 | 90.2093 | 58.5833 | 67.9167 | 66.16 | 62.5333 | 67.6117 | 69.1833 | 62.9667 | 70.8167 | 69.05 | 67.4167 | 68.6467 | 55.2917 | 66.15 | 61.3333 | 64.0283 | 52.2968 | 93.6167 | 62.1075 | 55.733 | 54.5683
| Median |  | 66 | 63.53 | 71.45 | 68.86 | 67.366 | 58.65 | 67.8 | 65.895 | 62.15 | 68.015 | 69.15 | 62.85 | 70.8 | 69.7 | 67.75 | 68.98 | 55.41 | 66.4 | 61.3 | 64.185 | 53.0414 | 68.65 | 62.1046 | 55.6635 | 55.63
| Std Dev. |  | 1.2247 | 2.7979 | 1.4367 | 0.5167 | 56.9562 | 0.7679 | 0.73869 | 0.64103 | 0.80166 | 2.0254 | 1.3963 | 1.0671 | 1.1652 | 1.2029 | 1.4675 | 1.0328 | 0.72137 | 1.3561 | 0.46332 | 1.3872 | 4.0916 | 61.9203 | 1.6578 | 0.77104 | 3.2044
| Rel.Std.Dev. |  | 1.87% | 4.457% | 2.012% | 0.7508% | 63.14% | 1.311% | 1.088% | 0.9689% | 1.282% | 2.996% | 2.018% | 1.695% | 1.645% | 1.742% | 2.177% | 1.505% | 1.305% | 2.05% | 0.7554% | 2.167% | 7.824% | 66.14% | 2.669% | 1.383% | 5.872%
| PDM3 |  | 1.065% | -3.14% | 10.17% | 6.185% | 39.19% | -9.608% | 4.794% | 2.083% | -3.513% | 4.323% | 6.748% | -2.844% | 9.268% | 6.542% | 4.022% | 5.92% | -14.69% | 2.068% | -5.364% | -1.206% | -19.31% | 44.45% | -4.17% | -14.01% | -15.8%
| Z-Score (Absolute) |  | 0.13 | 0.67 | 1.05 | 0.53 | 4.81 | 1.51 | 0.35 | 0 | 0.72 | 0.29 | 0.61 | 0.64 | 0.93 | 0.58 | 0.25 | 0.5 | 2.17 | 0 | 0.96 | 0.42 | 2.77 | 5.49 | 0.81 | 2.08 | 2.31
| Analytical results for Ca in OREAS 609c (Certified Value 0.845 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.83 | 0.83 | 0.83 | 0.78 | 0.43893 | 0.83 | 0.81 | 0.83 | 0.95 | 0.89 | 0.86 | 0.86 | 0.85 | 0.92 | 0.89 | 0.8569 | 0.87 | 0.938 | 0.78 | 0.85 | 0.818 | 0.836 | 0.84554 | 0.87372 |  |  | 1
| 1 | 2 | 0.87 | 0.82 | 0.85 | 0.81 | 0.83267 | 0.83 | 0.81 | 0.81 | 0.94 | 0.93 | 0.85 | 0.86 | 0.86 | 0.92 | 0.9 | 0.859 | 0.859 | 0.94 | 0.79 | 0.86 | 0.8143 | 0.436 | 0.85214 | 0.87232 |  |  | 0
| 1 | 3 | 0.85 | 0.81 | 0.84 | 0.8 | 0.83896 | 0.83 | 0.81 | 0.83 | 0.94 | 0.88 | 0.85 | 0.85 | 0.84 | 0.89 | 0.88 | 0.8395 | 0.86 | 0.916 | 0.78 | 0.86 | 0.8304 | 0.85 | 0.85662 | 0.85492 |  |  | 16
| 1 | 4 | 0.85 | 0.8 | 0.83 | 0.84 | 0.82496 | 0.82 | 0.8 | 0.83 | 0.93 | 0.92 | 0.84 | 0.84 | 0.84 | 0.9 | 0.87 | 0.8576 | 0.854 | 0.93 | 0.8 | 0.92 | 0.7941 | 0.815 | 0.84735 | 0.85765 |  |  | 0.84518
| 1 | 5 | 0.86 | 0.84 | 0.83 | 0.79 | 0.83191 | 0.84 | 0.8 | 0.82 | 0.95 | 0.87 | 0.85 | 0.87 | 0.88 | 0.88 | 0.89 | 0.8615 | 0.873 | 0.93 | 0.81 | 0.87 | 0.8118 | 0.85 | 0.86163 | 0.89015 |  |  | 79
| 1 | 6 | 0.83 | 0.85 | 0.84 | 0.8 | 0.82844 | 0.86 | 0.81 | 0.83 | 0.95 | 0.95 | 0.86 | 0.86 | 0.84 | 0.89 | 0.88 | 0.8565 | 0.873 | 0.92 | 0.78 | 0.84 | 0.8185 | 0.858 | 0.84146 | 0.87977
| Mean |  | 0.84833 | 0.825 | 0.83667 | 0.80333 | 0.76598 | 0.835 | 0.80667 | 0.825 | 0.94333 | 0.90667 | 0.85167 | 0.85667 | 0.85167 | 0.9 | 0.885 | 0.85517 | 0.86483 | 0.929 | 0.79 | 0.86667 | 0.81452 | 0.77417 | 0.85079 | 0.87142
| Median |  | 0.85 | 0.825 | 0.835 | 0.8 | 0.83018 | 0.83 | 0.81 | 0.83 | 0.945 | 0.905 | 0.85 | 0.86 | 0.845 | 0.895 | 0.885 | 0.85725 | 0.865 | 0.93 | 0.785 | 0.86 | 0.81615 | 0.843 | 0.84974 | 0.87302
| Std Dev. |  | 0.016021 | 0.018708 | 0.008165 | 0.020656 | 0.16029 | 0.013784 | 0.005164 | 0.0083666 | 0.008165 | 0.031411 | 0.0075277 | 0.010328 | 0.016021 | 0.016733 | 0.010488 | 0.0078861 | 0.0081833 | 0.0095289 | 0.012649 | 0.028048 | 0.011871 | 0.16636 | 0.0074812 | 0.013331
| Rel.Std.Dev. |  | 1.889% | 2.268% | 0.9759% | 2.571% | 20.93% | 1.651% | 0.6402% | 1.014% | 0.8655% | 3.464% | 0.8839% | 1.206% | 1.881% | 1.859% | 1.185% | 0.9222% | 0.9462% | 1.026% | 1.601% | 3.236% | 1.457% | 21.49% | 0.8793% | 1.53%
| PDM3 |  | 0.3728% | -2.388% | -1.008% | -4.952% | -9.371% | -1.205% | -4.557% | -2.388% | 11.61% | 7.275% | 0.7672% | 1.359% | 0.7672% | 6.486% | 4.711% | 1.181% | 2.325% | 9.917% | -6.529% | 2.542% | -3.628% | -8.402% | 0.6631% | 3.104%
| Z-Score (Absolute) |  | 0.07 | 0.67 | 0.37 | 1.23 | 2.19 | 0.42 | 1.15 | 0.67 | 2.37 | 1.43 | 0.01 | 0.14 | 0.01 | 1.25 | 0.87 | 0.1 | 0.35 | 2 | 1.57 | 0.4 | 0.94 | 1.98 | 0.01 | 0.52
| Analytical results for Cd in OREAS 609c (Certified Value 5.57 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 6.15 | 5.84 | 5.4 | 5.41 | 16.2475 | 5.25 | 5.66 | 5.41 | 5.48 | 5.4 | 5.59 | 5.27 | 5.64 | 5.94 | 5.83 | 5.35 | 5.62 | 6.5 | 5.77 | 4.9637 | 5.7 | 5.3556 | 5.553 | 5.25 |  |  | 1
| 1 | 2 | 6.12 | 5.33 | 5.4 | 5.55 | 5.662 | 5.29 | 5.69 | 5.51 | 5.45 | 5.7 | 5.57 | 5.48 | 5.73 | 6.01 | 5.98 | 5.38 | 5.09 | 6.4 | 6.05 | 5.4764 | 17.1 | 5.2951 | 5.526 | 5.3 |  |  | 28
| 1 | 3 | 5.76 | 6.06 | 5.3 | 5.44 | 5.6149 | 5.29 | 5.73 | 5.24 | 5.65 | 5.4 | 5.51 | 5.4 | 5.65 | 5.89 | 5.89 | 5.42 | 5.23 | 6.5 | 5.77 | 5.0433 | 5.61 | 5.2459 | 5.532 | 5.34 |  |  | 16
| 1 | 4 | 6.16 | 5.73 | 5.4 | 5.45 | 5.4915 | 5.28 | 5.57 | 5.59 | 5.43 | 5.6 | 5.63 | 5.26 | 5.53 | 6 | 5.55 | 5.48 | 5.36 | 6.5 | 5.84 | 5.0677 | 5.73 | 5.2303 | 5.55 | 5.65 |  |  | 5.5719
| 1 | 5 | 6.03 | 5.6 | 5.3 | 5.43 | 5.5822 | 5.24 | 5.56 | 5.38 | 5.53 | 5.3 | 5.81 | 5.54 | 5.8 | 5.61 | 5.88 | 5.32 | 5.1 | 6.6 | 5.6 | 5.3451 | 5.84 | 5.3202 | 5.709 | 6.17 |  |  | 80
| 1 | 6 | 6.13 | 5.8 | 5.5 | 5.49 | 5.6596 | 5.37 | 5.76 | 5.71 | 5.35 | 5.7 | 5.62 | 5.38 | 5.56 | 5.97 | 5.91 | 5.43 | 5.35 | 6.4 | 6.12 | 5.1072 | 5.76 | 5.2848 | 5.717 | 5.84
| Mean |  | 6.0583 | 5.7267 | 5.3833 | 5.4617 | 7.3763 | 5.2867 | 5.6617 | 5.4733 | 5.4817 | 5.5167 | 5.6217 | 5.3883 | 5.6517 | 5.9033 | 5.84 | 5.3967 | 5.2917 | 6.4833 | 5.8583 | 5.1672 | 7.6233 | 5.2886 | 5.5978 | 5.5917
| Median |  | 6.125 | 5.765 | 5.4 | 5.445 | 5.6372 | 5.285 | 5.675 | 5.46 | 5.465 | 5.5 | 5.605 | 5.39 | 5.645 | 5.955 | 5.885 | 5.4 | 5.29 | 6.5 | 5.805 | 5.0874 | 5.745 | 5.2899 | 5.5515 | 5.495
| Std Dev. |  | 0.15329 | 0.24606 | 0.075277 | 0.050761 | 4.3464 | 0.045898 | 0.082321 | 0.16621 | 0.10167 | 0.17224 | 0.10167 | 0.11143 | 0.10147 | 0.15016 | 0.1502 | 0.058195 | 0.19854 | 0.075277 | 0.19385 | 0.19875 | 4.6432 | 0.046397 | 0.089834 | 0.3646
| Rel.Std.Dev. |  | 2.53% | 4.297% | 1.398% | 0.9294% | 58.92% | 0.8682% | 1.454% | 3.037% | 1.855% | 3.122% | 1.809% | 2.068% | 1.795% | 2.544% | 2.572% | 1.078% | 3.752% | 1.161% | 3.309% | 3.846% | 60.91% | 0.8773% | 1.605% | 6.521%
| PDM3 |  | 8.73% | 2.778% | -3.384% | -1.978% | 32.38% | -5.119% | 1.611% | -1.769% | -1.619% | -0.9911% | 0.8934% | -3.294% | 1.432% | 5.949% | 4.812% | -3.145% | -5.029% | 16.36% | 5.141% | -7.262% | 36.82% | -5.084% | 0.4657% | 0.355%
| Z-Score (Absolute) |  | 1.5 | 0.43 | 0.68 | 0.43 | 5.75 | 0.99 | 0.22 | 0.39 | 0.36 | 0.25 | 0.09 | 0.67 | 0.18 | 1 | 0.79 | 0.64 | 0.98 | 2.87 | 0.85 | 1.38 | 6.55 | 0.99 | 0.01 | 0.01
| Analytical results for Ce in OREAS 609c (Certified Value 37.4 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 1
| 1 | 1 | 35.08 | 35.8 | 38.132 | 32.5335 | 38.9 | 35.5 | 34.739 | 35.43 | 47 | 60.79 | 33.1 | 35.7 | 37.4 | 43.3 | 35 | 30.51 | 21.7 | 38.72 | 49.2224 | 19.6 | 36.5921 | 48.368 | 23.22 |  |  |  | 1
| 1 | 2 | 36.58 | 34.3 | 37.184 | 36.0568 | 39.1 | 35 | 34.799 | 36.27 | 45.6 | 59.56 | 31.7 | 35.9 | 38.7 | 43.7 | 36.8 | 29.52 | 20.9 | 39.9 | 46.3409 | 21.3 | 37.846 | 48.019 | 23.9 |  |  |  | 29
| 1 | 3 | 33.48 | 36.4 | 38.291 | 36.6737 | 39 | 36.4 | 34.273 | 36.1 | 43.1 | 57.86 | 31.4 | 36.9 | 36.7 | 44 | 37 | 30.36 | 21.6 | 36.51 | 44.2386 | 19.1 | 35.7105 | 49.024 | 22.69 |  |  |  | 16
| 1 | 4 | 33.17 | 35.2 | 35.178 | 35.9292 | 38.7 | 35.9 | 34.348 | 36.13 | 48.3 | 57.52 | 32.6 | 34.9 | 35.9 | 42.6 | 35.2 | 32.38 | 21.9 | 41.24 | 45.4952 | 19.5 | 37.8741 | 47.816 | 23.4 |  |  |  | 37.4229
| 1 | 5 | 34.15 | 36.6 | 38.091 | 35.2071 | 40.3 | 34.3 | 34.202 | 35.97 | 45 | 57.69 | 33.1 | 36.2 | 39.3 | 40.8 | 37.1 | 31.62 | 21.6 | 39.01 | 47.1548 | 19.2 | 38.4658 | 47.514 | 22.33 |  |  |  | 81
| 1 | 6 | 33.65 | 38.7 | 37.306 | 35.8693 | 39.9 | 34.9 | 34.86 | 36.33 | 46.7 | 62.68 | 31.4 | 34.9 | 35 | 41.8 | 35.3 | 32.33 | 21.7 | 40.14 | 48.1937 | 18.8 | 36.2492 | 48.625 | 26.59
| Mean |  | 34.3517 | 36.1667 | 37.3637 | 35.3783 | 39.3167 | 35.3333 | 34.5368 | 36.0383 | 45.95 | 59.35 | 32.2167 | 35.75 | 37.1667 | 42.7 | 36.0667 | 31.12 | 21.5667 | 39.2533 | 46.7743 | 19.5833 | 37.1229 | 48.2277 | 23.6883
| Median |  | 33.9 | 36.1 | 37.6985 | 35.8992 | 39.05 | 35.25 | 34.5435 | 36.115 | 46.15 | 58.71 | 32.15 | 35.8 | 37.05 | 42.95 | 36.05 | 31.065 | 21.65 | 39.455 | 46.7478 | 19.35 | 37.2191 | 48.1935 | 23.31
| Std Dev. |  | 1.2795 | 1.4976 | 1.1655 | 1.47 | 0.63377 | 0.75542 | 0.29374 | 0.32425 | 1.8075 | 2.0751 | 0.81343 | 0.77395 | 1.6415 | 1.2231 | 0.99532 | 1.1672 | 0.34448 | 1.6137 | 1.8114 | 0.88863 | 1.0891 | 0.5539 | 1.5237
| Rel.Std.Dev. |  | 3.725% | 4.141% | 3.119% | 4.155% | 1.612% | 2.138% | 0.8505% | 0.8997% | 3.934% | 3.496% | 2.525% | 2.165% | 4.417% | 2.864% | 2.76% | 3.751% | 1.597% | 4.111% | 3.873% | 4.538% | 2.934% | 1.149% | 6.432%
| PDM3 |  | -8.207% | -3.357% | -0.1583% | -5.464% | 5.06% | -5.584% | -7.712% | -3.7% | 22.79% | 58.59% | -13.91% | -4.47% | -0.6847% | 14.1% | -3.624% | -16.84% | -42.37% | 4.891% | 24.99% | -47.67% | -0.8015% | 28.87% | -36.7%
| Z-Score (Absolute) |  | 0.36 | 0.02 | 0.27 | 0.15 | 0.69 | 0.16 | 0.32 | 0.01 | 2.09 | 4.93 | 0.81 | 0.07 | 0.23 | 1.4 | 0 | 1.05 | 3.07 | 0.67 | 2.27 | 3.49 | 0.22 | 2.57 | 2.62
| Analytical results for Co in OREAS 609c (Certified Value 6.36 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 7 | 6.4 | 6.4 | 6.3 | 7.6566 | 6.5 | 6.2 | 6.58 | 6.5 | 6.3 | 6.3 | 6.5 | 6.2 | 7.2 | 5.8 | 6.6 | 6.5 | 6.1 | 6.1 | 6.2684 | 5.86 | 6.0793 | 6.037 | 5.3 |  |  | 1
| 1 | 2 | 7.5 | 6 | 6.4 | 6.3 | 6.4034 | 6.6 | 6.3 | 6.55 | 6.7 | 6.6 | 6.2 | 6.6 | 6.2 | 7.2 | 6 | 6.5 | 6.5 | 6.1 | 6.7 | 8.3569 | 7.25 | 6.2871 | 6.064 | 5.4 |  |  | 17
| 1 | 3 | 7.2 | 7 | 6.3 | 6.3 | 6.3276 | 6.6 | 6.2 | 6.54 | 6.8 | 6.2 | 6.3 | 6.7 | 6.1 | 6.9 | 5.8 | 6.6 | 6 | 6.1 | 6.2 | 29.4631 | 5.76 | 6.2017 | 6.18 | 5.6 |  |  | 16
| 1 | 4 | 7.1 | 6.8 | 6.5 | 6.3 | 6.3222 | 6.6 | 6.2 | 6.66 | 6.6 | 6.5 | 6.1 | 6.5 | 6.1 | 7.2 | 5.6 | 6.4 | 6.4 | 6.1 | 6.4 | 6.006 | 5.85 | 6.3176 | 6.142 | 5.7 |  |  | 6.3638
| 1 | 5 | 7.1 | 6.6 | 6.4 | 6.3 | 6.2752 | 6.5 | 6.1 | 6.75 | 6.6 | 6.1 | 6.4 | 6.8 | 6.4 | 6.7 | 5.9 | 6.4 | 6.7 | 6.1 | 6.1 | 6.9649 | 5.9 | 6.3808 | 6.294 | 6.2 |  |  | 82
| 1 | 6 | 7 | 6.7 | 6.5 | 6.2 | 6.3811 | 6.7 | 6.3 | 6.69 | 6.7 | 6.6 | 6 | 6.6 | 6 | 6.9 | 5.8 | 6.4 | 7 | 6.1 | 6.5 | 6.4494 | 5.93 | 6.1136 | 6.101 | 5.7
| Mean |  | 7.15 | 6.5833 | 6.4167 | 6.2833 | 6.561 | 6.5833 | 6.2167 | 6.6283 | 6.65 | 6.3833 | 6.2167 | 6.6167 | 6.1667 | 7.0167 | 5.8167 | 6.4833 | 6.5167 | 6.1 | 6.3333 | 10.5848 | 6.0917 | 6.23 | 6.1363 | 5.65
| Median |  | 7.1 | 6.65 | 6.4 | 6.3 | 6.3544 | 6.6 | 6.2 | 6.62 | 6.65 | 6.4 | 6.25 | 6.6 | 6.15 | 7.05 | 5.8 | 6.45 | 6.5 | 6.1 | 6.3 | 6.7072 | 5.88 | 6.2444 | 6.1215 | 5.65
| Std Dev. |  | 0.18708 | 0.34881 | 0.075277 | 0.040825 | 0.53866 | 0.075277 | 0.075277 | 0.084715 | 0.10488 | 0.2137 | 0.1472 | 0.1169 | 0.13663 | 0.2137 | 0.13292 | 0.098319 | 0.33116 | 9.7295e-16 | 0.24221 | 9.2861 | 0.57038 | 0.11896 | 0.092893 | 0.31464
| Rel.Std.Dev. |  | 2.617% | 5.298% | 1.173% | 0.6497% | 8.21% | 1.143% | 1.211% | 1.278% | 1.577% | 3.348% | 2.368% | 1.767% | 2.216% | 3.046% | 2.285% | 1.516% | 5.082% | 1.595e-14% | 3.824% | 87.73% | 9.363% | 1.909% | 1.514% | 5.569%
| PDM3 |  | 12.35% | 3.449% | 0.8302% | -1.265% | 3.098% | 3.449% | -2.313% | 4.156% | 4.497% | 0.3064% | -2.313% | 3.973% | -3.098% | 10.26% | -8.598% | 1.878% | 2.402% | -4.146% | -0.4793% | 66.33% | -4.277% | -2.103% | -3.575% | -11.22%
| Z-Score (Absolute) |  | 2.53 | 0.62 | 0.06 | 0.39 | 0.54 | 0.62 | 0.62 | 0.77 | 0.84 | 0.06 | 0.62 | 0.73 | 0.79 | 2.08 | 1.97 | 0.28 | 0.39 | 1.01 | 0.22 | 14.11 | 1.04 | 0.57 | 0.89 | 2.53
| Analytical results for Cr in OREAS 609c (Certified Value 15.9 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 2
| 1 | 1 | 15 | 15.8 | 17 | 16 | 22.2358 | 14 | 16.5 | 16.472 | 16 | 14 | 17 | 16 | 17 | 17 | 17 | 17 | 18 | 16.8 | 15 | 15.7674 | 16 | 15.3868 | 15.206 |  |  |  | 1
| 1 | 2 | 16 | 15.5 | 18 | 16 | 16.0224 | 15 | 16.1 | 15.989 | 15.7 | 15 | 17 | 16 | 16 | 17 | 18 | 17 | 17 | 14 | 16 | 15.7827 | 22 | 15.7282 | 15.333 |  |  |  | 31
| 1 | 3 | 16 | 15.7 | 17 | 16 | 15.5877 | 15 | 15.7 | 15.9785 | 16.3 | 14 | 17 | 16 | 17 | 16 | 18 | 17 | 17 | 15.4 | 14 | 15.8341 | 17 | 14.9799 | 15.952 |  |  |  | 16
| 1 | 4 | 16 | 14.6 | 17 | 17 | 15.6388 | 16 | 16.1 | 15.758 | 16 | 15 | 18 | 16 | 17 | 15 | 18 | 16 | 17 | 14 | 15 | 14.7379 | 16 | 16.2948 | 15.578 |  |  |  | 15.9239
| 1 | 5 | 16 | 15.4 | 17 | 16 | 15.6712 | 15 | 15.7 | 14.9705 | 17 | 15 | 17 | 16 | 17 | 17 | 17 | 17 | 19 | 15 | 15 | 16.7848 | 15 | 16.3528 | 15.824 |  |  |  | 83
| 1 | 6 | 15 | 15.4 | 18 | 16 | 16.6027 | 15 | 15.7 | 15.8945 | 17 | 15 | 19 | 16 | 17 | 16 | 17 | 16 | 16 | 15 | 17 | 15.0834 | 15 | 15.0011 | 15.082
| Mean |  | 15.6667 | 15.4 | 17.3333 | 16.1667 | 16.9597 | 15 | 15.9667 | 15.8438 | 16.3333 | 14.6667 | 17.5 | 16 | 16.8333 | 16.3333 | 17.5 | 16.6667 | 17.3333 | 15.0333 | 15.3333 | 15.6651 | 16.8333 | 15.6239 | 15.4958
| Median |  | 16 | 15.45 | 17 | 16 | 15.8468 | 15 | 15.9 | 15.9365 | 16.15 | 15 | 17 | 16 | 17 | 16.5 | 17.5 | 17 | 17 | 15 | 15 | 15.7751 | 16 | 15.5575 | 15.4555
| Std Dev. |  | 0.5164 | 0.42426 | 0.5164 | 0.40825 | 2.6127 | 0.63246 | 0.3266 | 0.49121 | 0.55015 | 0.5164 | 0.83666 | 0 | 0.40825 | 0.8165 | 0.54772 | 0.5164 | 1.0328 | 1.0386 | 1.0328 | 0.70777 | 2.6394 | 0.60845 | 0.34765
| Rel.Std.Dev. |  | 3.296% | 2.755% | 2.979% | 2.525% | 15.41% | 4.216% | 2.046% | 3.1% | 3.368% | 3.521% | 4.781% | 0% | 2.425% | 4.999% | 3.13% | 3.098% | 5.958% | 6.909% | 6.736% | 4.518% | 15.68% | 3.894% | 2.244%
| PDM3 |  | -1.615% | -3.29% | 8.851% | 1.525% | 6.505% | -5.802% | 0.2688% | -0.5031% | 2.571% | -7.895% | 9.898% | 0.4781% | 5.711% | 2.571% | 9.898% | 4.665% | 8.851% | -5.592% | -3.708% | -1.625% | 5.711% | -1.884% | -2.688%
| Z-Score (Absolute) |  | 0.34 | 0.61 | 1.35 | 0.17 | 0.97 | 1.01 | 0.03 | 0.16 | 0.34 | 1.35 | 1.52 | 0 | 0.84 | 0.34 | 1.52 | 0.67 | 1.35 | 0.98 | 0.67 | 0.34 | 0.84 | 0.38 | 0.51
| Analytical results for Cs in OREAS 609c (Certified Value 0.97 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [5 empty] | 2
| 1 | 1 | 1.05 | 1.08 | 1.07 | 1.1123 | 1 | 1.01 | 0.8653 | 0.98 | 1.28 | 1.31 | 0.86 | 0.92 | 1.02 | 1.06 | 0.95 | 0.94 | 0.86 | 1.15 | 0.66 | 0.8988 | 0.77 | [5 empty] | 1
| 1 | 2 | 1.05 | 1.05 | 1.08 | 1.0553 | 1.01 | 1 | 0.8368 | 0.94 | 1.43 | 1.37 | 0.83 | 0.93 | 1.04 | 1.06 | 0.96 | 0.92 | 0.82 | 1.19 | 0.86 | 0.92255 | 0.85 | [5 empty] | 32
| 1 | 3 | 0.99 | 1.14 | 1.1 | 1.0593 | 1.01 | 1.02 | 0.8401 | 1.01 | 1.32 | 1.31 | 0.83 | 0.94 | 1 | 1.06 | 0.96 | 0.95 | 0.79 | 1.09 | 0.68 | 0.9262 | 0.81 | [5 empty] | 16
| 1 | 4 | 1.01 | 1.08 | 1.05 | 1.0662 | 0.98 | 1.01 | 0.8647 | 1.02 | 1.39 | 1.32 | 0.85 | 0.9 | 0.98 | 1.05 | 0.92 | 0.95 | 0.84 | 1.19 | 0.67 | 0.89114 | 0.84 | [5 empty] | 0.97483
| 1 | 5 | 0.95 | 1.08 | 1.09 | 1.0166 | 1.08 | 0.98 | 0.8786 | 0.99 | 1.34 | 1.31 | 0.85 | 0.93 | 1.04 | 1.01 | 0.96 | 0.95 | 0.81 | 1.15 | 0.66 | 0.91086 | 0.77 | [5 empty] | 84
| 1 | 6 | 1.08 | 1.19 | 1.07 | 1.0235 | 1.02 | 1 | 0.8907 | 1.02 | 1.31 | 1.35 | 0.84 | 0.91 | 0.98 | 1.04 | 0.92 | 0.98 | 0.88 | 1.18 | 0.68 | 0.87825 | 0.91
| Mean |  | 1.0217 | 1.1033 | 1.0767 | 1.0555 | 1.0167 | 1.0033 | 0.8627 | 0.99333 | 1.345 | 1.3283 | 0.84333 | 0.92167 | 1.01 | 1.0467 | 0.945 | 0.94833 | 0.83333 | 1.1583 | 0.70167 | 0.90463 | 0.825
| Median |  | 1.03 | 1.08 | 1.075 | 1.0573 | 1.01 | 1.005 | 0.865 | 1 | 1.33 | 1.315 | 0.845 | 0.925 | 1.01 | 1.055 | 0.955 | 0.95 | 0.83 | 1.165 | 0.675 | 0.90483 | 0.825
| Std Dev. |  | 0.047504 | 0.05164 | 0.017512 | 0.034316 | 0.033862 | 0.013663 | 0.021115 | 0.030768 | 0.055408 | 0.025626 | 0.012111 | 0.01472 | 0.027568 | 0.019664 | 0.019748 | 0.019408 | 0.033267 | 0.038166 | 0.078081 | 0.018639 | 0.053572
| Rel.Std.Dev. |  | 4.65% | 4.68% | 1.626% | 3.251% | 3.331% | 1.362% | 2.448% | 3.097% | 4.12% | 1.929% | 1.436% | 1.597% | 2.73% | 1.879% | 2.09% | 2.047% | 3.992% | 3.295% | 11.13% | 2.06% | 6.494%
| PDM3 |  | 4.805% | 13.18% | 10.45% | 8.278% | 4.292% | 2.924% | -11.5% | 1.899% | 37.97% | 36.26% | -13.49% | -5.453% | 3.608% | 7.37% | -3.06% | -2.718% | -14.51% | 18.82% | -28.02% | -7.2% | -15.37%
| Z-Score (Absolute) |  | 0.15 | 0.83 | 0.61 | 0.43 | 0.11 | 0 | 1.16 | 0.08 | 2.82 | 2.68 | 1.32 | 0.67 | 0.06 | 0.36 | 0.48 | 0.45 | 1.4 | 1.28 | 2.49 | 0.81 | 1.47
| Analytical results for Cu in OREAS 609c (Certified Value 0.478 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 0.486 | 0.46666 | 0.47303 | 0.48781 | 0.94955 | 0.474 | 0.457 | 0.49525 | 0.4896 | 0.474 | 0.4699 | 0.487 | 0.484 | 0.473 | 0.5 | 0.484 | 0.5184 | > 0.5 | 0.524 | 0.4592 | 0.47798 | 0.44117 | 0.456 | 0.48022 | 0.4565 | 0.4613 | 1
| 1 | 2 | 0.4998 | 0.44571 | 0.4779 | 0.48286 | 0.48535 | 0.481 | 0.462 | 0.49701 | 0.4788 | 0.483 | 0.4871 | 0.482 | 0.477 | 0.482 | 0.503 | 0.481 | 0.5222 | 0.4957 | 0.494 | 0.463 | 0.50325 | 0.45035 | 0.903 | 0.47634 | 0.45526 | 0.4701 | 19
| 1 | 3 | 0.4912 | 0.45774 | 0.47428 | 0.48619 | 0.47592 | 0.48 | 0.453 | 0.49911 | 0.4857 | 0.473 | 0.4646 | 0.484 | 0.483 | 0.47 | 0.485 | 0.475 | 0.5108 | 0.4949 | 0.504 | 0.4569 | 0.46882 | 0.44368 | 0.455 | 0.47394 | 0.45555 | 0.4821 | 16
| 1 | 4 | 0.4912 | 0.44947 | 0.47371 | 0.4825 | 0.47619 | 0.478 | 0.462 | 0.49697 | 0.4875 | 0.484 | 0.4967 | 0.478 | 0.473 | 0.468 | 0.494 | 0.468 | 0.5178 | 0.4893 | 0.506 | 0.4634 | 0.47368 | 0.44997 | 0.456 | 0.47453 | 0.44879 | 0.4964 | 0.47797
| 1 | 5 | 0.4903 | 0.45846 | 0.47385 | 0.47875 | 0.4759 | 0.477 | 0.453 | 0.49805 | 0.4839 | 0.479 | 0.4651 | 0.484 | 0.487 | 0.491 | 0.481 | 0.476 | 0.5221 | > 0.5 | 0.497 | 0.4556 | 0.45666 | 0.45099 | 0.464 | 0.48092 | 0.4551 | > 0.5 | 85
| 1 | 6 | 0.4822 | 0.46848 | 0.47467 | 0.48392 | 0.47787 | 0.485 | 0.456 | 0.4921 | 0.4878 | 0.476 | 0.5024 | 0.483 | 0.481 | 0.471 | 0.491 | 0.474 | 0.5211 | > 0.5 | 0.51 | 0.4592 | 0.48965 | 0.45411 | 0.467 | 0.48004 | 0.45345 | 0.489
| Mean |  | 0.49012 | 0.45775 | 0.47457 | 0.48367 | 0.5568 | 0.47917 | 0.45717 | 0.49641 | 0.48555 | 0.47817 | 0.48097 | 0.483 | 0.48083 | 0.47583 | 0.49233 | 0.47633 | 0.51873 | 0.4933 | 0.50583 | 0.45955 | 0.47834 | 0.44838 | 0.5335 | 0.47767 | 0.45411 | 0.47978
| Median |  | 0.49075 | 0.4581 | 0.47407 | 0.48339 | 0.47703 | 0.479 | 0.4565 | 0.49699 | 0.4866 | 0.4775 | 0.4785 | 0.4835 | 0.482 | 0.472 | 0.4925 | 0.4755 | 0.51975 | 0.4949 | 0.505 | 0.4592 | 0.47583 | 0.45016 | 0.46 | 0.47819 | 0.45518 | 0.4821
| Std Dev. |  | 0.0059291 | 0.0090409 | 0.0017212 | 0.0031554 | 0.19244 | 0.0037639 | 0.0040702 | 0.002474 | 0.003836 | 0.0046224 | 0.016653 | 0.0029665 | 0.0050761 | 0.0088863 | 0.0084774 | 0.0056095 | 0.0043098 | 0.0034871 | 0.010666 | 0.0031494 | 0.016315 | 0.0049005 | 0.18108 | 0.0031048 | 0.0027843 | 0.014158
| Rel.Std.Dev. |  | 1.21% | 1.975% | 0.3627% | 0.6524% | 34.56% | 0.7855% | 0.8903% | 0.4984% | 0.79% | 0.9667% | 3.462% | 0.6142% | 1.056% | 1.868% | 1.722% | 1.178% | 0.8308% | 0.7069% | 2.109% | 0.6853% | 3.411% | 1.093% | 33.94% | 0.65% | 0.6131% | 2.951%
| PDM3 |  | 2.541% | -4.229% | -0.7105% | 1.193% | 16.49% | 0.2505% | -4.352% | 3.859% | 1.586% | 0.04125% | 0.6271% | 1.052% | 0.5992% | -0.4469% | 3.005% | -0.3423% | 8.529% | 3.207% | 5.83% | -3.854% | 0.07752% | -6.191% | 11.62% | -0.06355% | -4.992% | 0.3788%
| Z-Score (Absolute) |  | 0.85 | 1.96 | 0.5 | 0.29 | 6.64 | 0.1 | 2.01 | 1.4 | 0.45 | 0.19 | 0.06 | 0.23 | 0.05 | 0.39 | 1.04 | 0.34 | 3.33 | 1.13 | 2.21 | 1.8 | 0.17 | 2.77 | 4.62 | 0.23 | 2.27 | 0.05
| Analytical results for Dy in OREAS 609c (Certified Value 1.34 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 1.39; 2.03: 1.417; 2.04: 1.527; 2.05: 1.5; 2.06: 1.1733; 2.07: 1.86; 2.08: 1.27; 2.09: 1.5753; 2.10: 1.01; 2.11: 1.29 | 1
| 1 | 2 | 2.01: 1.31; 2.03: 1.439; 2.04: 1.5424; 2.05: 1.5; 2.06: 1.1424; 2.07: 1.9; 2.08: 1.25; 2.09: 1.6349; 2.10: 1.17; 2.11: 1.34 | 34
| 1 | 3 | 2.01: 1.46; 2.03: 1.44; 2.04: 1.5972; 2.05: 1.5; 2.06: 1.185; 2.07: 1.81; 2.08: 1.22; 2.09: 1.3659; 2.10: 0.99; 2.11: 1.37 | 16
| 1 | 4 | 2.01: 1.42; 2.03: 1.439; 2.04: 1.563; 2.05: 1.4; 2.06: 1.0853; 2.07: 1.94; 2.08: 1.34; 2.09: 1.4472; 2.10: 1; 2.11: 1.4 | 1.3394
| 1 | 5 | 2.01: 1.43; 2.03: 1.455; 2.04: 1.5454; 2.05: 1.5; 2.06: 1.1103; 2.07: 1.81; 2.08: 1.29; 2.09: 1.4944; 2.10: 1; 2.11: 1.27 | 86
| 1 | 6 | 2.01: 1.53; 2.03: 1.445; 2.04: 1.4783; 2.05: 1.5; 2.06: 1.091; 2.07: 1.99; 2.08: 1.35; 2.09: 1.6246; 2.10: 1.01; 2.11: 1.53
| Mean |  | 2.01: 1.4233; 2.03: 1.4392; 2.04: 1.5422; 2.05: 1.4833; 2.06: 1.1312; 2.07: 1.885; 2.08: 1.2867; 2.09: 1.5237; 2.10: 1.03; 2.11: 1.3667
| Median |  | 2.01: 1.425; 2.03: 1.4395; 2.04: 1.5439; 2.05: 1.5; 2.06: 1.1263; 2.07: 1.88; 2.08: 1.28; 2.09: 1.5349; 2.10: 1.005; 2.11: 1.355
| Std Dev. |  | 2.01: 0.073121; 2.03: 0.012465; 2.04: 0.039433; 2.05: 0.040825; 2.06: 0.042313; 2.07: 0.072319; 2.08: 0.050859; 2.09: 0.1066; 2.10: 0.068993; 2.11: 0.093524
| Rel.Std.Dev. |  | 2.01: 5.137%; 2.03: 0.8661%; 2.04: 2.557%; 2.05: 2.752%; 2.06: 3.74%; 2.07: 3.837%; 2.08: 3.953%; 2.09: 6.996%; 2.10: 6.698%; 2.11: 6.843%
| PDM3 |  | 2.01: 6.269%; 2.03: 7.451%; 2.04: 15.14%; 2.05: 10.75%; 2.06: -15.54%; 2.07: 40.74%; 2.08: -3.935%; 2.09: 13.76%; 2.10: -23.1%; 2.11: 2.038%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.09; 2.04: 0.67; 2.05: NA; 2.06: 1.66; 2.07: 2.62; 2.08: 0.78; 2.09: 0.57; 2.10: 2.23; 2.11: 0.32
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 609c (Certified Value 0.38 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 0.39; 2.03: 0.382; 2.04: 0.40831; 2.05: 0.4; 2.06: 0.3472; 2.07: 0.46; 2.08: 0.34; 2.09: 0.4391; 2.10: 0.29; 2.11: 0.33 | 1
| 1 | 2 | 2.01: 0.37; 2.03: 0.39; 2.04: 0.43894; 2.05: 0.4; 2.06: 0.3338; 2.07: 0.46; 2.08: 0.35; 2.09: 0.4996; 2.10: 0.32; 2.11: 0.38 | 35
| 1 | 3 | 2.01: 0.37; 2.03: 0.385; 2.04: 0.45496; 2.05: 0.4; 2.06: 0.3485; 2.07: 0.45; 2.08: 0.34; 2.09: 0.3945; 2.10: 0.29; 2.11: 0.35 | 16
| 1 | 4 | 2.01: 0.38; 2.03: 0.39; 2.04: 0.42849; 2.05: 0.4; 2.06: 0.295; 2.07: 0.46; 2.08: 0.36; 2.09: 0.4136; 2.10: 0.29; 2.11: 0.38 | 0.38202
| 1 | 5 | 2.01: 0.39; 2.03: 0.389; 2.04: 0.43898; 2.05: 0.4; 2.06: 0.3711; 2.07: 0.45; 2.08: 0.35; 2.09: 0.4313; 2.10: 0.29; 2.11: 0.34 | 87
| 1 | 6 | 2.01: 0.4; 2.03: 0.385; 2.04: 0.41596; 2.05: 0.4; 2.06: 0.2976; 2.07: 0.49; 2.08: 0.38; 2.09: 0.4631; 2.10: 0.3; 2.11: 0.4
| Mean |  | 2.01: 0.38333; 2.03: 0.38683; 2.04: 0.43094; 2.05: 0.4; 2.06: 0.3322; 2.07: 0.46167; 2.08: 0.35333; 2.09: 0.4402; 2.10: 0.29667; 2.11: 0.36333
| Median |  | 2.01: 0.385; 2.03: 0.387; 2.04: 0.43371; 2.05: 0.4; 2.06: 0.3405; 2.07: 0.46; 2.08: 0.35; 2.09: 0.4352; 2.10: 0.29; 2.11: 0.365
| Std Dev. |  | 2.01: 0.012111; 2.03: 0.0033116; 2.04: 0.017018; 2.05: 6.0809e-17; 2.06: 0.03029; 2.07: 0.01472; 2.08: 0.015055; 2.09: 0.037227; 2.10: 0.012111; 2.11: 0.027325
| Rel.Std.Dev. |  | 2.01: 3.159%; 2.03: 0.8561%; 2.04: 3.949%; 2.05: 1.52e-14%; 2.06: 9.118%; 2.07: 3.188%; 2.08: 4.261%; 2.09: 8.457%; 2.10: 4.082%; 2.11: 7.521%
| PDM3 |  | 2.01: 0.3439%; 2.03: 1.26%; 2.04: 12.81%; 2.05: 4.707%; 2.06: -13.04%; 2.07: 20.85%; 2.08: -7.509%; 2.09: 15.23%; 2.10: -22.34%; 2.11: -4.891%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.05; 2.04: 0.67; 2.05: NA; 2.06: 0.72; 2.07: 1.11; 2.08: 0.42; 2.09: 0.81; 2.10: 1.23; 2.11: 0.28
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 609c (Certified Value 0.5 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 0.49; 2.03: 0.522; 2.04: 0.40773; 2.05: 0.4; 2.06: 0.5162; 2.07: 0.75; 2.08: 0.5; 2.09: 0.6863; 2.10: 0.35; 2.11: 0.39 | 1
| 1 | 2 | 2.01: 0.48; 2.03: 0.51; 2.04: 0.52353; 2.05: 0.5; 2.06: 0.5926; 2.07: 0.74; 2.08: 0.5; 2.09: 0.6625; 2.10: 0.43; 2.11: 0.42 | 36
| 1 | 3 | 2.01: 0.51; 2.03: 0.534; 2.04: 0.51218; 2.05: 0.5; 2.06: 0.5854; 2.07: 0.74; 2.08: 0.51; 2.09: 0.6124; 2.10: 0.33; 2.11: 0.4 | 16
| 1 | 4 | 2.01: 0.5; 2.03: 0.468; 2.04: 0.5077; 2.05: 0.5; 2.06: 0.5576; 2.07: 0.74; 2.08: 0.51; 2.09: 0.6318; 2.10: 0.35; 2.11: 0.42 | 0.50112
| 1 | 5 | 2.01: 0.52; 2.03: 0.496; 2.04: 0.51118; 2.05: 0.5; 2.06: 0.5276; 2.07: 0.74; 2.08: 0.51; 2.09: 0.6596; 2.10: 0.34; 2.11: 0.39 | 88
| 1 | 6 | 2.01: 0.53; 2.03: 0.514; 2.04: 0.50333; 2.05: 0.5; 2.06: 0.5892; 2.07: 0.8; 2.08: 0.53; 2.09: 0.6971; 2.10: 0.34; 2.11: 0.46
| Mean |  | 2.01: 0.505; 2.03: 0.50733; 2.04: 0.49427; 2.05: 0.48333; 2.06: 0.56143; 2.07: 0.75167; 2.08: 0.51; 2.09: 0.65828; 2.10: 0.35667; 2.11: 0.41333
| Median |  | 2.01: 0.505; 2.03: 0.512; 2.04: 0.50944; 2.05: 0.5; 2.06: 0.5715; 2.07: 0.74; 2.08: 0.51; 2.09: 0.66105; 2.10: 0.345; 2.11: 0.41
| Std Dev. |  | 2.01: 0.018708; 2.03: 0.023036; 2.04: 0.04293; 2.05: 0.040825; 2.06: 0.033233; 2.07: 0.024014; 2.08: 0.010954; 2.09: 0.031993; 2.10: 0.036697; 2.11: 0.026583
| Rel.Std.Dev. |  | 2.01: 3.705%; 2.03: 4.541%; 2.04: 8.685%; 2.05: 8.447%; 2.06: 5.919%; 2.07: 3.195%; 2.08: 2.148%; 2.09: 4.86%; 2.10: 10.29%; 2.11: 6.431%
| PDM3 |  | 2.01: 0.7741%; 2.03: 1.24%; 2.04: -1.366%; 2.05: -3.55%; 2.06: 12.04%; 2.07: 50%; 2.08: 1.772%; 2.09: 31.36%; 2.10: -28.83%; 2.11: -17.52%
| Z-Score (Absolute) |  | 2.01: 0.03; 2.03: 0; 2.04: 0.16; 2.05: NA; 2.06: 0.67; 2.07: 3.05; 2.08: 0.03; 2.09: 1.88; 2.10: 1.88; 2.11: 1.17
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 609c (Certified Value 2.29 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  | 3
| 1 | 1 | 2.32 | 2.14 | 2.22 | 2.18 | 3.6588 | 2.45 | 2.39 | 2.34 | 2.77 | 2.54 | 2.24 | 2.3 | 2.29 | 2.41 | 2.24 | 1.9884 | 2.23 | 2.23 | 2.37 | 1.9 | 2.2859 | 2.21 | 2.3366 | 2.3822 |  |  | 1
| 1 | 2 | 2.35 | 2.07 | 2.27 | 2.26 | 2.4571 | 2.51 | 2.39 | 2.3 | 2.88 | 2.61 | 2.19 | 2.29 | 2.3 | 2.45 | 2.28 | 1.9978 | 2.16 | 2 | 2.38 | 2.02 | 2.273 | 3.735 | 2.4305 | 2.4286 |  |  | 0
| 1 | 3 | 2.38 | 2.07 | 2.26 | 2.24 | 2.4512 | 2.51 | 2.4 | 2.34 | 2.88 | 2.51 | 2.19 | 2.3 | 2.26 | 2.4 | 2.27 | 1.958 | 2.16 | 2.24 | 2.37 | 2.01 | 2.3137 | 2.217 | 2.4256 | 2.3845 |  |  | 16
| 1 | 4 | 2.38 | 2.02 | 2.23 | 2.31 | 2.4283 | 2.49 | 2.37 | 2.33 | 2.86 | 2.64 | 2.21 | 2.26 | 2.24 | 2.39 | 2.2 | 2.0117 | 2.22 | 2.31 | 2.39 | 1.96 | 2.2438 | 2.14 | 2.479 | 2.357 |  |  | 2.293
| 1 | 5 | 2.36 | 2.11 | 2.24 | 2.18 | 2.4873 | 2.46 | 2.33 | 2.32 | 2.81 | 2.47 | 2.23 | 2.31 | 2.34 | 2.37 | 2.29 | 2.0235 | 2.31 | 2.27 | 2.43 | 1.83 | 2.3199 | 2.196 | 2.512 | 2.1811 |  |  | 89
| 1 | 6 | 2.32 | 2.16 | 2.27 | 2.22 | 2.4646 | 2.53 | 2.38 | 2.35 | 2.95 | 2.72 | 2.19 | 2.3 | 2.25 | 2.37 | 2.23 | 1.9909 | 2.27 | 2.34 | 2.33 | 2.03 | 2.2036 | 2.224 | 2.3831 | 2.3633
| Mean |  | 2.3517 | 2.095 | 2.2483 | 2.2317 | 2.6579 | 2.4917 | 2.3767 | 2.33 | 2.8583 | 2.5817 | 2.2083 | 2.2933 | 2.28 | 2.3983 | 2.2517 | 1.995 | 2.225 | 2.2317 | 2.3783 | 1.9583 | 2.2733 | 2.4537 | 2.4278 | 2.3495
| Median |  | 2.355 | 2.09 | 2.25 | 2.23 | 2.4609 | 2.5 | 2.385 | 2.335 | 2.87 | 2.575 | 2.2 | 2.3 | 2.275 | 2.395 | 2.255 | 1.9944 | 2.225 | 2.255 | 2.375 | 1.985 | 2.2795 | 2.2135 | 2.4281 | 2.3728
| Std Dev. |  | 0.027142 | 0.051672 | 0.02137 | 0.049967 | 0.49072 | 0.031252 | 0.025033 | 0.017889 | 0.062423 | 0.092394 | 0.022286 | 0.017512 | 0.037417 | 0.029944 | 0.034303 | 0.022487 | 0.059582 | 0.1209 | 0.032506 | 0.079352 | 0.044025 | 0.62845 | 0.063308 | 0.086238
| Rel.Std.Dev. |  | 1.154% | 2.466% | 0.9505% | 2.239% | 18.46% | 1.254% | 1.053% | 0.7677% | 2.184% | 3.579% | 1.009% | 0.7636% | 1.641% | 1.249% | 1.523% | 1.127% | 2.678% | 5.417% | 1.367% | 4.052% | 1.937% | 25.61% | 2.608% | 3.671%
| PDM3 |  | 2.56% | -8.634% | -1.947% | -2.674% | 15.91% | 8.665% | 3.65% | 1.615% | 24.66% | 12.59% | -3.691% | 0.01574% | -0.5657% | 4.595% | -1.801% | -12.99% | -2.964% | -2.674% | 3.723% | -14.59% | -0.8572% | 7.008% | 5.881% | 2.463%
| Z-Score (Absolute) |  | 0.32 | 1.75 | 0.51 | 0.65 | 2.8 | 1.46 | 0.53 | 0.15 | 4.42 | 2.18 | 0.84 | 0.15 | 0.26 | 0.7 | 0.49 | 2.56 | 0.7 | 0.65 | 0.54 | 2.86 | 0.31 | 1.15 | 0.94 | 0.31
| Analytical results for Ga in OREAS 609c (Certified Value 3.96 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 2
| 1 | 1 | 3.76 | 3.9 | 4 | 3.85 | 7.7596 | 4.47 | 3.98 | 4.45 | 4.4 | 4.76 | 3.34 | 3.98 | 3.75 | 4.01 | 3.27 | 3.29 | < 10 | < 5 | 8.82 | 4.3948 | 2.44 | 3.515 | 4.613 | 2.87 |  |  | 1
| 1 | 2 | 3.78 | 3.7 | 4 | 3.79 | 4.0969 | 4.66 | 4.1 | 4.16 | 4.6 | 4.83 | 3.21 | 4.09 | 3.8 | 4.1 | 3.42 | 3.27 | < 10 | < 5 | 9.03 | 5.0137 | 6.02 | 3.7581 | 5.133 | 3.04 |  |  | 37
| 1 | 3 | 3.67 | 4.1 | 4 | 3.72 | 4.0786 | 4.6 | 4.09 | 4.26 | 4.6 | 4.66 | 3.21 | 4.08 | 3.66 | 4.03 | 3.46 | 3.22 | < 10 | < 5 | 8.87 | 4.0861 | 2.4 | 3.5346 | 5.571 | 2.96 |  |  | 16
| 1 | 4 | 3.63 | 3.8 | 4 | 3.67 | 4.1318 | 4.48 | 4.05 | 4.42 | 4.6 | 4.78 | 3.24 | 3.91 | 3.59 | 4.05 | 3.2 | 3.45 | < 10 | < 5 | 8.99 | 4.2012 | 2.44 | 3.5512 | 5.146 | 3.04 |  |  | 3.9649
| 1 | 5 | 3.51 | 4 | 4 | 3.68 | 3.9578 | 4.66 | 3.89 | 4.28 | 4.5 | 4.74 | 3.33 | 4.23 | 3.84 | 3.88 | 3.37 | 3.32 | < 10 | < 5 | 8.82 | 4.1591 | 2.46 | 3.6544 | 5.133 | 3.06 |  |  | 90
| 1 | 6 | 3.53 | 4.2 | 4 | 3.84 | 4.2121 | 4.73 | 4 | 4.26 | 4.5 | 4.93 | 3.22 | 4.07 | 3.58 | 4.06 | 3.28 | 3.38 | < 10 | < 5 | 9.05 | 4.3875 | 2.47 | 3.7608 | 4.37 | 3.31
| Mean |  | 3.6467 | 3.95 | 4 | 3.7583 | 4.7062 | 4.6 | 4.0183 | 4.305 | 4.5333 | 4.7833 | 3.2583 | 4.06 | 3.7033 | 4.0217 | 3.3333 | 3.3217 |  |  | 8.93 | 4.3737 | 3.0383 | 3.629 | 4.9943 | 3.0467
| Median |  | 3.65 | 3.95 | 4 | 3.755 | 4.1144 | 4.63 | 4.025 | 4.27 | 4.55 | 4.77 | 3.23 | 4.075 | 3.705 | 4.04 | 3.325 | 3.305 |  |  | 8.93 | 4.2943 | 2.45 | 3.6028 | 5.133 | 3.04
| Std Dev. |  | 0.1129 | 0.18708 | 0 | 0.079352 | 1.4982 | 0.10526 | 0.078846 | 0.1095 | 0.08165 | 0.090921 | 0.06047 | 0.10881 | 0.10967 | 0.075741 | 0.099532 | 0.082321 |  |  | 0.10564 | 0.33727 | 1.4609 | 0.1119 | 0.43109 | 0.1472
| Rel.Std.Dev. |  | 3.096% | 4.736% | 0% | 2.111% | 31.83% | 2.288% | 1.962% | 2.544% | 1.801% | 1.901% | 1.856% | 2.68% | 2.961% | 1.883% | 2.986% | 2.478% |  |  | 1.183% | 7.711% | 48.08% | 3.083% | 8.632% | 4.831%
| PDM3 |  | -8.026% | -0.3755% | 0.8856% | -5.21% | 18.7% | 16.02% | 1.348% | 8.578% | 14.34% | 20.64% | -17.82% | 2.399% | -6.597% | 1.432% | -15.93% | -16.22% |  |  | 125.2% | 10.31% | -23.37% | -8.471% | 25.96% | -23.16%
| Z-Score (Absolute) |  | 0.43 | 0.08 | NA | 0.3 | 0.8 | 0.67 | 0 | 0.33 | 0.6 | 0.89 | 0.88 | 0.05 | 0.37 | 0 | 0.79 | 0.81 | 1.14 | 1.76 | 5.69 | 0.41 | 1.14 | 0.45 | 1.13 | 1.13
Results from laboratory 2.04 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 609c (Certified Value 2.3 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  | 2
| 1 | 1 | 2.01: 2.46; 2.03: 2.435; 2.04: 2.7734; 2.05: 2.4; 2.06: 2.2754; 2.07: 3.41; 2.08: 2.26; 2.10: 3.4855; 2.11: 1.68; 2.12: 1.64 |  |  | 1
| 1 | 2 | 2.01: 2.32; 2.03: 2.404; 2.04: 2.8153; 2.05: 2.6; 2.06: 2.6174; 2.07: 3.42; 2.08: 2.2; 2.10: 3.4179; 2.11: 2.1; 2.12: 1.72 |  |  | 38
| 1 | 3 | 2.01: 2.49; 2.03: 2.462; 2.04: 2.9877; 2.05: 2.5; 2.06: 2.5523; 2.07: 3.35; 2.08: 2.21; 2.10: 3.1087; 2.11: 1.61; 2.12: 1.68 |  |  | 16
| 1 | 4 | 2.01: 2.47; 2.03: 2.342; 2.04: 2.9453; 2.05: 2.5; 2.06: 2.5585; 2.07: 3.43; 2.08: 2.36; 2.10: 3.2526; 2.11: 1.63; 2.12: 1.71 |  |  | 2.2985
| 1 | 5 | 2.01: 2.44; 2.03: 2.292; 2.04: 2.7686; 2.05: 2.5; 2.06: 2.8754; 2.07: 3.28; 2.08: 2.34; 2.10: 3.398; 2.11: 1.64; 2.12: 1.63 |  |  | 91
| 1 | 6 | 2.01: 2.55; 2.03: 2.401; 2.04: 2.7432; 2.05: 2.6; 2.06: 2.6326; 2.07: 3.63; 2.08: 2.34; 2.10: 3.5812; 2.11: 1.65; 2.12: 1.89
| Mean |  | 2.01: 2.455; 2.03: 2.3893; 2.04: 2.8389; 2.05: 2.5167; 2.06: 2.5853; 2.07: 3.42; 2.08: 2.285; 2.10: 3.374; 2.11: 1.7183; 2.12: 1.7117
| Median |  | 2.01: 2.465; 2.03: 2.4025; 2.04: 2.7944; 2.05: 2.5; 2.06: 2.588; 2.07: 3.415; 2.08: 2.3; 2.10: 3.408; 2.11: 1.645; 2.12: 1.695
| Std Dev. |  | 2.01: 0.076092; 2.03: 0.06234; 2.04: 0.10236; 2.05: 0.075277; 2.06: 0.19249; 2.07: 0.1173; 2.08: 0.070922; 2.10: 0.16911; 2.11: 0.18841; 2.12: 0.094534
| Rel.Std.Dev. |  | 2.01: 3.099%; 2.03: 2.609%; 2.04: 3.606%; 2.05: 2.991%; 2.06: 7.446%; 2.07: 3.43%; 2.08: 3.104%; 2.10: 5.012%; 2.11: 10.96%; 2.12: 5.523%
| PDM3 |  | 2.01: 6.808%; 2.03: 3.951%; 2.04: 23.51%; 2.05: 9.491%; 2.06: 12.48%; 2.07: 48.79%; 2.08: -0.5884%; 2.10: 46.79%; 2.11: -25.24%; 2.12: -25.53%
| Z-Score (Absolute) |  | 2.01: 0.08; 2.03: 0.23; 2.04: 0.86; 2.05: 0.08; 2.06: 0.24; 2.07: 2.27; 2.08: 0.49; 2.10: 2.16; 2.11: 1.87; 2.12: 1.88
| Analytical results for Ge in OREAS 609c (Certified Value 0.089 ppm) | [27 empty] | Certified
Ge | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab; 2.15: Lab; 2.16: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.03: 2.03; 2.04: 2.06; 2.05: 2.07; 2.06: 2.08; 2.07: 2.13; 2.08: 2.14; 2.10: 2.15; 2.11: 2.16; 2.12: 2.17; 2.13: 2.19; 2.14: 2.20; 2.15: 2.22; 2.16: 2.25 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*MS; 2.15: AR*MS; 2.16: AR*MS |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.4g; 2.14: 0.5g; 2.15: 0.2g; 2.16: 0.5g |  |  | 3
| 1 | 1 | 2.01: 0.11; 2.03: < 0.1; 2.04: 0.38465; 2.05: < 0.1; 2.06: 0.17; 2.07: 0.1; 2.08: 0.12; 2.10: 0.13; 2.11: 0.08; 2.12: 0.18; 2.13: 0.08; 2.14: 0.06; 2.15: 0.09; 2.16: < 0.01 |  |  | 1
| 1 | 2 | 2.01: 0.11; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.19; 2.07: 0.09; 2.08: 0.11; 2.10: 0.13; 2.11: 0.1; 2.12: 0.17; 2.13: 0.09; 2.14: 0.06; 2.15: 0.1; 2.16: 0.23 |  |  | 39
| 1 | 3 | 2.01: 0.11; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.19; 2.07: 0.09; 2.08: 0.12; 2.10: 0.12; 2.11: 0.09; 2.12: 0.18; 2.13: 0.12; 2.14: 0.07; 2.15: 0.09; 2.16: < 0.01 |  |  | 16
| 1 | 4 | 2.01: 0.11; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.18; 2.07: 0.1; 2.08: 0.11; 2.10: 0.12; 2.11: 0.1; 2.12: 0.2; 2.13: 0.13; 2.14: 0.07; 2.15: 0.09; 2.16: 0.01 |  |  | 0.089074
| 1 | 5 | 2.01: 0.11; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.18; 2.07: 0.09; 2.08: 0.12; 2.10: 0.13; 2.11: 0.09; 2.12: 0.17; 2.13: 0.1; 2.14: 0.09; 2.15: 0.08; 2.16: < 0.01 |  |  | 92
| 1 | 6 | 2.01: 0.1; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.18; 2.07: 0.09; 2.08: 0.11; 2.10: 0.13; 2.11: 0.09; 2.12: 0.18; 2.13: 0.05; 2.14: 0.05; 2.15: 0.09; 2.16: 0.02
| Mean |  | 2.01: 0.10833; 2.04: 0.38465; 2.06: 0.18167; 2.07: 0.093333; 2.08: 0.115; 2.10: 0.12667; 2.11: 0.091667; 2.12: 0.18; 2.13: 0.095; 2.14: 0.066667; 2.15: 0.09; 2.16: 0.086667
| Median |  | 2.01: 0.11; 2.04: 0.38465; 2.06: 0.18; 2.07: 0.09; 2.08: 0.115; 2.10: 0.13; 2.11: 0.09; 2.12: 0.18; 2.13: 0.095; 2.14: 0.065; 2.15: 0.09; 2.16: 0.02
| Std Dev. |  | 2.01: 0.0040825; 2.06: 0.0075277; 2.07: 0.005164; 2.08: 0.0054772; 2.10: 0.005164; 2.11: 0.0075277; 2.12: 0.010954; 2.13: 0.02881; 2.14: 0.013663; 2.15: 0.0063246; 2.16: 0.12423
| Rel.Std.Dev. |  | 2.01: 3.768%; 2.06: 4.144%; 2.07: 5.533%; 2.08: 4.763%; 2.10: 4.077%; 2.11: 8.212%; 2.12: 6.086%; 2.13: 30.33%; 2.14: 20.49%; 2.15: 7.027%; 2.16: 143.3%
| PDM3 |  | 2.01: 21.62%; 2.04: 331.8%; 2.06: 104%; 2.07: 4.782%; 2.08: 29.11%; 2.10: 42.2%; 2.11: 2.911%; 2.12: 102.1%; 2.13: 6.653%; 2.14: -25.16%; 2.15: 1.04%; 2.16: -2.703%
| Z-Score (Absolute) |  | 2.01: 0.32; 2.03: 0.99; 2.04: 1.2; 2.05: 0.99; 2.06: 1.97; 2.07: 0.02; 2.08: 0.47; 2.10: 0.73; 2.11: 0.06; 2.12: 1.93; 2.13: 0.02; 2.14: 0.62; 2.15: 0.09; 2.16: 1.09
| Analytical results for Hf in OREAS 609c (Certified Value 1.36 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 2
| 1 | 1 | 1.42 | 1.54 | 1.42 | 0.92578 | 0.7 | 1.17 | 1.32 | 1.6 | 1.5 | 1.26 | 1.44 | 1.36 | 1.41 | 1.23 | 1.31 | 1.18 | 1.32 | 0.67 | 1.3038 | 1.11 | [6 empty] | 1
| 1 | 2 | 1.51 | 1.45 | 1.42 | 1.3539 | 0.6 | 1.19 | 1.36 | 1.73 | 1.53 | 1.31 | 1.46 | 1.35 | 1.41 | 1.22 | 1.3 | 1.09 | 1.27 | 0.65 | 1.3978 | 1.09 | [6 empty] | 23
| 1 | 3 | 1.53 | 1.61 | 1.43 | 1.4028 | 0.5 | 1.22 | 1.28 | 1.56 | 1.47 | 1.24 | 1.43 | 1.33 | 1.42 | 1.24 | 1.3 | 1.14 | 1.33 | 0.7 | 1.3817 | 1.19 | [6 empty] | 16
| 1 | 4 | 1.34 | 1.62 | 1.4 | 1.4015 | 0.5 | 1.21 | 1.34 | 1.6 | 1.55 | 1.32 | 1.36 | 1.3 | 1.42 | 1.18 | 1.42 | 1.23 | 1.35 | 0.62 | 1.3609 | 1.2 | [6 empty] | 1.3597
| 1 | 5 | 1.58 | 1.58 | 1.4 | 1.3179 | 0.5 | 1.12 | 1.34 | 1.62 | 1.43 | 1.28 | 1.42 | 1.42 | 1.38 | 1.28 | 1.47 | 1.15 | 1.36 | 0.62 | 1.4181 | 1.07 | [6 empty] | 93
| 1 | 6 | 1.49 | 1.58 | 1.4 | 1.3389 | 0.4 | 1.17 | 1.35 | 1.61 | 1.55 | 1.24 | 1.38 | 1.32 | 1.41 | 1.17 | 1.47 | 1.18 | 1.32 | 0.68 | 1.2867 | 1.08
| Mean |  | 1.4783 | 1.5633 | 1.4117 | 1.2901 | 0.53333 | 1.18 | 1.3317 | 1.62 | 1.505 | 1.275 | 1.415 | 1.3467 | 1.4083 | 1.22 | 1.3783 | 1.1617 | 1.325 | 0.65667 | 1.3582 | 1.1233
| Median |  | 1.5 | 1.58 | 1.41 | 1.3464 | 0.5 | 1.18 | 1.34 | 1.605 | 1.515 | 1.27 | 1.425 | 1.34 | 1.41 | 1.225 | 1.365 | 1.165 | 1.325 | 0.66 | 1.3713 | 1.1
| Std Dev. |  | 0.085654 | 0.062183 | 0.013292 | 0.18169 | 0.10328 | 0.035777 | 0.028577 | 0.057619 | 0.048062 | 0.034496 | 0.037815 | 0.041793 | 0.01472 | 0.040497 | 0.084242 | 0.047081 | 0.031464 | 0.03266 | 0.052524 | 0.057155
| Rel.Std.Dev. |  | 5.794% | 3.978% | 0.9416% | 14.08% | 19.36% | 3.032% | 2.146% | 3.557% | 3.194% | 2.706% | 2.672% | 3.103% | 1.045% | 3.319% | 6.112% | 4.053% | 2.375% | 4.974% | 3.867% | 5.088%
| PDM3 |  | 8.721% | 14.97% | 3.818% | -5.121% | -60.78% | -13.22% | -2.065% | 19.14% | 10.68% | -6.233% | 4.063% | -0.9621% | 3.573% | -10.28% | 1.367% | -14.57% | -2.556% | -51.71% | -0.1165% | -17.39%
| Z-Score (Absolute) |  | 0.96 | 1.55 | 0.5 | 0.34 | 5.57 | 1.1 | 0.05 | 1.94 | 1.15 | 0.44 | 0.52 | 0.05 | 0.48 | 0.82 | 0.27 | 1.23 | 0.1 | 4.72 | 0.13 | 1.49
| Analytical results for Hg in OREAS 609c (Certified Value 0.34 ppm) | [27 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 25g | [5 empty] | 3
| 1 | 1 | 0.2 | 0.264 | 0.34 | 0.4 | < 1 | 0.31 | 0.36 | 0.34 | 0.31 | 0.34 | 0.33 | 0.36 | 0.38 | 0.34 | 0.32 | 0.345 | < 5 | < 2 | < 0.01 | 0.35762 | 0.31 | [5 empty] | 1
| 1 | 2 | 0.16 | 0.333 | 0.36 | 0.4 | < 1 | 0.31 | 0.36 | 0.36 | 0.3 | 0.35 | 0.38 | 0.36 | 0.38 | 0.33 | 0.36 | 0.31 | < 5 | < 2 | 0.28 | 0.36686 | 0.32 | [5 empty] | 3
| 1 | 3 | 0.19 | 0.323 | 0.36 | 0.4 | < 1 | 0.33 | 0.35 | 0.34 | 0.31 | 0.35 | 0.33 | 0.37 | 0.38 | 0.32 | 0.35 | 0.33 | < 5 | < 2 | < 0.01 | 0.36661 | 0.3 | [5 empty] | 16
| 1 | 4 | 0.33 | 0.338 | 0.36 | 0.4 | < 1 | 0.32 | 0.35 | 0.35 | 0.36 | 0.35 | 0.36 | 0.37 | 0.36 | 0.35 | 0.35 | 0.33 | < 5 | < 2 | < 0.01 | 0.36533 | 0.35 | [5 empty] | 0.34372
| 1 | 5 | 0.26 | 0.352 | 0.34 | 0.4 | < 1 | 0.34 | 0.36 | 0.34 | 0.34 | 0.35 | 0.38 | 0.35 | 0.39 | 0.31 | 0.35 | 0.34 | < 5 | < 2 | < 0.01 | 0.35768 | 0.34 | [5 empty] | 94
| 1 | 6 | 0.11 | 0.318 | 0.37 | 0.4 | < 1 | 0.34 | 0.34 | 0.36 | 0.31 | 0.34 | 0.34 | 0.35 | 0.38 | 0.33 | 0.33 | 0.33 | < 5 | < 2 | < 0.01 | 0.35483 | 0.32
| Mean |  | 0.20833 | 0.32133 | 0.355 | 0.4 |  | 0.325 | 0.35333 | 0.34833 | 0.32167 | 0.34667 | 0.35333 | 0.36 | 0.37833 | 0.33 | 0.34333 | 0.33083 |  |  | 0.28 | 0.36149 | 0.32333
| Median |  | 0.195 | 0.328 | 0.36 | 0.4 |  | 0.325 | 0.355 | 0.345 | 0.31 | 0.35 | 0.35 | 0.36 | 0.38 | 0.33 | 0.35 | 0.33 |  |  | 0.28 | 0.3615 | 0.32
| Std Dev. |  | 0.077309 | 0.030513 | 0.012247 | 6.0809e-17 |  | 0.013784 | 0.008165 | 0.0098319 | 0.023166 | 0.005164 | 0.023381 | 0.0089443 | 0.0098319 | 0.014142 | 0.015055 | 0.012007 |  |  |  | 0.0053606 | 0.018619
| Rel.Std.Dev. |  | 37.11% | 9.496% | 3.45% | 1.52e-14% |  | 4.241% | 2.311% | 2.823% | 7.202% | 1.49% | 6.617% | 2.485% | 2.599% | 4.285% | 4.385% | 3.629% |  |  |  | 1.483% | 5.758%
| PDM3 |  | -39.39% | -6.513% | 3.282% | 16.37% |  | -5.446% | 2.797% | 1.342% | -6.416% | 0.8575% | 2.797% | 4.737% | 10.07% | -3.991% | -0.1123% | -3.749% |  |  | -18.54% | 5.17% | -5.931%
| Z-Score (Absolute) |  | 4.69 | 0.88 | 0.25 | NA | 5.14 | 0.76 | 0.2 | 0.03 | 0.87 | 0.03 | 0.2 | 0.42 | 1.04 | 0.59 | 0.14 | 0.56 | 72.57 | 22 | 10 | 0.47 | 0.81
Results from laboratory 2.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ho in OREAS 609c (Certified Value 0.18 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab | [5 empty] | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.12: 2.23; 2.13: 2.25; 2.14: 2.28 | [5 empty] | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*MS | [5 empty] | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.12: 0.5g; 2.13: 0.5g; 2.14: 25g | [5 empty] | 2
| 1 | 1 | 2.01: 0.18; 2.03: 0.193; 2.04: 0.195; 2.05: 0.2; 2.06: 0.1256; 2.07: 0.22; 2.08: 0.18; 2.12: 0.1739; 2.13: 0.15; 2.14: 0.15 | [5 empty] | 1
| 1 | 2 | 2.01: 0.18; 2.03: 0.191; 2.04: 0.19273; 2.05: 0.2; 2.06: 0.1417; 2.07: 0.22; 2.08: 0.18; 2.12: 0.1924; 2.13: 0.16; 2.14: 0.16 | [5 empty] | 24
| 1 | 3 | 2.01: 0.18; 2.03: 0.201; 2.04: 0.18911; 2.05: 0.2; 2.06: 0.159; 2.07: 0.22; 2.08: 0.18; 2.12: 0.1543; 2.13: 0.14; 2.14: 0.15 | [5 empty] | 16
| 1 | 4 | 2.01: 0.18; 2.03: 0.193; 2.04: 0.19569; 2.05: 0.2; 2.06: 0.1375; 2.07: 0.22; 2.08: 0.19; 2.12: 0.1623; 2.13: 0.15; 2.14: 0.18 | [5 empty] | 0.17728
| 1 | 5 | 2.01: 0.18; 2.03: 0.195; 2.04: 0.1853; 2.05: 0.2; 2.06: 0.1631; 2.07: 0.21; 2.08: 0.18; 2.12: 0.167; 2.13: 0.14; 2.14: 0.14 | [5 empty] | 95
| 1 | 6 | 2.01: 0.19; 2.03: 0.198; 2.04: 0.18668; 2.05: 0.2; 2.06: 0.1507; 2.07: 0.23; 2.08: 0.19; 2.12: 0.18; 2.13: 0.15; 2.14: 0.17
| Mean |  | 2.01: 0.18167; 2.03: 0.19517; 2.04: 0.19075; 2.05: 0.2; 2.06: 0.14627; 2.07: 0.22; 2.08: 0.18333; 2.12: 0.17165; 2.13: 0.14833; 2.14: 0.15833
| Median |  | 2.01: 0.18; 2.03: 0.194; 2.04: 0.19092; 2.05: 0.2; 2.06: 0.1462; 2.07: 0.22; 2.08: 0.18; 2.12: 0.17045; 2.13: 0.15; 2.14: 0.155
| Std Dev. |  | 2.01: 0.0040825; 2.03: 0.0037103; 2.04: 0.0043661; 2.05: 3.0405e-17; 2.06: 0.014071; 2.07: 0.0063246; 2.08: 0.005164; 2.12: 0.013531; 2.13: 0.0075277; 2.14: 0.01472
| Rel.Std.Dev. |  | 2.01: 2.247%; 2.03: 1.901%; 2.04: 2.289%; 2.05: 1.52e-14%; 2.06: 9.62%; 2.07: 2.875%; 2.08: 2.817%; 2.12: 7.883%; 2.13: 5.075%; 2.14: 9.297%
| PDM3 |  | 2.01: 2.476%; 2.03: 10.09%; 2.04: 7.6%; 2.05: 12.82%; 2.06: -17.49%; 2.07: 24.1%; 2.08: 3.416%; 2.12: -3.175%; 2.13: -16.33%; 2.14: -10.69%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.67; 2.04: 0.45; 2.05: NA; 2.06: 1.77; 2.07: 1.91; 2.08: 0.08; 2.12: 0.5; 2.13: 1.66; 2.14: 1.17
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 609c (Certified Value 1.26 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 01g | 25g | [8 empty] | 3
| 1 | 1 | 1.398 | 1.33 | 1.24 | 1.13 | 1.24 | 1.2 | 1.21 | 1.24 | 1.26 | 1.315 | 1.25 | 1.32 | 1.285 | 1.314 | 1.21 | 1.267 | 1.2137 | 1.043 | [8 empty] | 1
| 1 | 2 | 1.452 | 1.25 | 1.25 | 1.15 | 1.26 | 1.2 | 1.2 | 1.29 | 1.255 | 1.34 | 1.275 | 1.325 | 1.33 | 1.308 | 1.16 | 1.306 | 1.2187 | 1.105 | [8 empty] | 25
| 1 | 3 | 1.292 | 1.43 | 1.24 | 1.14 | 1.24 | 1.2 | 1.18 | 1.24 | 1.25 | 1.31 | 1.26 | 1.3 | 1.28 | 1.344 | 1.2 | 1.346 | 1.2177 | 1.116 | [8 empty] | 16
| 1 | 4 | 1.335 | 1.37 | 1.25 | 1.15 | 1.22 | 1.1 | 1.17 | 1.29 | 1.235 | 1.3 | 1.24 | 1.345 | 1.205 | 1.305 | 1.16 | 1.256 | 1.2321 | 1.145 | [8 empty] | 1.2599
| 1 | 5 | 1.331 | 1.33 | 1.23 | 1.16 | 1.22 | 1.2 | 1.18 | 1.22 | 1.275 | 1.355 | 1.3 | 1.255 | 1.29 | 1.299 | 1.2 | 1.348 | 1.2326 | 1.205 | [8 empty] | 96
| 1 | 6 | 1.362 | 1.41 | 1.23 | 1.15 | 1.25 | 1.2 | 1.18 | 1.29 | 1.22 | 1.32 | 1.24 | 1.32 | 1.275 | 1.324 | 1.29 | 1.354 | 1.2176 | 1.156
| Mean |  | 1.3617 | 1.3533 | 1.24 | 1.1467 | 1.2383 | 1.1833 | 1.1867 | 1.2617 | 1.2492 | 1.3233 | 1.2608 | 1.3108 | 1.2775 | 1.3157 | 1.2033 | 1.3128 | 1.2221 | 1.1283
| Median |  | 1.3485 | 1.35 | 1.24 | 1.15 | 1.24 | 1.2 | 1.18 | 1.265 | 1.2525 | 1.3175 | 1.255 | 1.32 | 1.2825 | 1.311 | 1.2 | 1.326 | 1.2182 | 1.1305
| Std Dev. |  | 0.056543 | 0.065013 | 0.0089443 | 0.010328 | 0.016021 | 0.040825 | 0.015055 | 0.031885 | 0.019343 | 0.020412 | 0.023327 | 0.03089 | 0.040589 | 0.016281 | 0.04761 | 0.043379 | 0.0081455 | 0.054588
| Rel.Std.Dev. |  | 4.152% | 4.804% | 0.7213% | 0.9007% | 1.294% | 3.45% | 1.269% | 2.527% | 1.549% | 1.542% | 1.85% | 2.356% | 3.177% | 1.237% | 3.956% | 3.304% | 0.6665% | 4.838%
| PDM3 |  | 8.079% | 7.418% | -1.578% | -8.986% | -1.71% | -6.075% | -5.811% | 0.1421% | -0.8501% | 5.037% | 0.07592% | 4.045% | 1.399% | 4.428% | -4.488% | 4.203% | -3.001% | -10.44%
| Z-Score (Absolute) |  | 1.31 | 1.2 | 0.27 | 1.48 | 0.29 | NA | 0.96 | 0.01 | 0.15 | 0.81 | 0 | 0.65 | 0.22 | 0.71 | 0.75 | 0.67 | 0.5 | 1.72
Results from laboratory 2.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 609c (Certified Value 0.243 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.24 | 0.26 | 0.24 | 0.2384 | 0.19425 | 0.27 | 0.21 | 0.23 | 0.37 | 0.3652 | 0.23 | 0.23 | 0.23 | 0.27 | 0.23 | 0.243 | 0.246 | 0.23 | 0.27 | 0.2442 | 0.2109 | 0.23721 | 0.27192 |  |  |  | 1
| 1 | 2 | 0.25 | 0.25 | 0.25 | 0.246 | 0.23706 | 0.28 | 0.22 | 0.23 | 0.355 | 0.3796 | 0.22 | 0.23 | 0.23 | 0.27 | 0.24 | 0.236 | 0.217 | 0.23 | 0.27 | 0.2558 | 0.1586 | 0.23929 | 0.26966 |  |  |  | 0
| 1 | 3 | 0.24 | 0.24 | 0.25 | 0.2418 | 0.24041 | 0.28 | 0.22 | 0.23 | 0.36 | 0.3638 | 0.22 | 0.23 | 0.23 | 0.27 | 0.24 | 0.237 | 0.25 | 0.23 | 0.26 | 0.2631 | 0.21 | 0.23747 | 0.27031 |  |  |  | 16
| 1 | 4 | 0.24 | 0.23 | 0.25 | 0.2465 | 0.2388 | 0.27 | 0.22 | 0.23 | 0.352 | 0.3758 | 0.22 | 0.3 | 0.23 | 0.27 | 0.23 | 0.247 | 0.257 | 0.23 | 0.27 | 0.2696 | 0.2009 | 0.24164 | 0.2627 |  |  |  | 0.24301
| 1 | 5 | 0.24 | 0.25 | 0.25 | 0.2394 | 0.23692 | 0.29 | 0.21 | 0.23 | 0.367 | 0.3639 | 0.22 | 0.23 | 0.24 | 0.27 | 0.24 | 0.252 | 0.248 | 0.23 | 0.27 | 0.2799 | 0.2067 | 0.24128 | 0.26344 |  |  |  | 97
| 1 | 6 | 0.24 | 0.26 | 0.25 | 0.2415 | 0.2371 | 0.29 | 0.22 | 0.23 | 0.371 | 0.391 | 0.22 | 0.23 | 0.23 | 0.27 | 0.23 | 0.252 | 0.25 | 0.23 | 0.27 | 0.258 | 0.2092 | 0.23736 | 0.27091
| Mean |  | 0.24167 | 0.24833 | 0.24833 | 0.24227 | 0.23075 | 0.28 | 0.21667 | 0.23 | 0.3625 | 0.37322 | 0.22167 | 0.24167 | 0.23167 | 0.27 | 0.235 | 0.2445 | 0.24467 | 0.23 | 0.26833 | 0.26177 | 0.19938 | 0.23904 | 0.26815
| Median |  | 0.24 | 0.25 | 0.25 | 0.24165 | 0.23708 | 0.28 | 0.22 | 0.23 | 0.3635 | 0.3705 | 0.22 | 0.23 | 0.23 | 0.27 | 0.235 | 0.245 | 0.249 | 0.23 | 0.27 | 0.26055 | 0.20795 | 0.23838 | 0.26998
| Std Dev. |  | 0.0040825 | 0.01169 | 0.0040825 | 0.0033417 | 0.017938 | 0.0089443 | 0.005164 | 0 | 0.0080187 | 0.010986 | 0.0040825 | 0.028577 | 0.0040825 | 0 | 0.0054772 | 0.007064 | 0.014052 | 0 | 0.0040825 | 0.01225 | 0.020302 | 0.0020231 | 0.0040165
| Rel.Std.Dev. |  | 1.689% | 4.708% | 1.644% | 1.379% | 7.774% | 3.194% | 2.383% | 0% | 2.212% | 2.943% | 1.842% | 11.83% | 1.762% | 0% | 2.331% | 2.889% | 5.743% | 0% | 1.521% | 4.68% | 10.18% | 0.8463% | 1.498%
| PDM3 |  | -0.5524% | 2.191% | 2.191% | -0.3055% | -5.043% | 15.22% | -10.84% | -5.353% | 49.17% | 53.58% | -8.783% | -0.5524% | -4.667% | 11.11% | -3.296% | 0.6135% | 0.6821% | -5.353% | 10.42% | 7.719% | -17.95% | -1.633% | 10.35%
| Z-Score (Absolute) |  | 0.03 | 0.33 | 0.33 | 0 | 0.63 | 2.07 | 1.41 | 0.67 | 6.61 | 7.2 | 1.13 | 0.03 | 0.58 | 1.52 | 0.4 | 0.12 | 0.13 | 0.67 | 1.43 | 1.07 | 2.36 | 0.18 | 1.42
| Analytical results for La in OREAS 609c (Certified Value 18.3 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 1
| 1 | 1 | 18.7 | 17.6 | 16 | 19.201 | 16.5156 | 18.6 | 17.4 | 16.67 | 17.9 | 23.8 | 31 | 15.7 | 17.8 | 19 | 21 | 18 | 15.3 | 13.2 | 20.4 | 19.1 | 22.701 | 10 | 18.0885 | 23.792 | 12.61 |  | 1
| 1 | 2 | 19.4 | 16.7 | 18 | 18.638 | 18.4968 | 18.6 | 17.5 | 16.59 | 18.2 | 23 | 32 | 15.2 | 17.8 | 19.8 | 21.8 | 18.6 | 14.9 | 13.3 | 20.2 | 20 | 21.6265 | 10 | 18.2215 | 23.537 | 13.09 |  | 26
| 1 | 3 | 17.5 | 17.9 | 18 | 19.072 | 18.4373 | 18.4 | 17.9 | 16.477 | 18.4 | 21.5 | 30 | 15.1 | 18.6 | 18.6 | 21.3 | 18.5 | 15.2 | 13.3 | 20.1 | 18.5 | 20.3848 | 9 | 17.4167 | 24.227 | 12.41 |  | 16
| 1 | 4 | 17.4 | 17.4 | 17 | 18.481 | 18.3995 | 18.3 | 17.7 | 16.668 | 18.3 | 24 | 31 | 15.7 | 17 | 18.1 | 20.9 | 17.7 | 16 | 13.5 | 20 | 20.4 | 21.0249 | 9 | 18.1926 | 23.542 | 12.78 |  | 18.3481
| 1 | 5 | 18.2 | 18.1 | 17 | 18.958 | 18.4678 | 19.1 | 16.8 | 16.804 | 18.4 | 22.2 | 30 | 15.4 | 18 | 20.1 | 20.5 | 18.5 | 15.8 | 13.5 | 20.3 | 19.4 | 21.5911 | 9 | 18.2118 | 23.483 | 11.99 |  | 98
| 1 | 6 | 18 | 18.8 | 18 | 18.732 | 18.7923 | 19 | 17.4 | 16.158 | 18.5 | 23.8 | 33 | 15.1 | 17.2 | 17.9 | 21 | 17.7 | 16.2 | 13.3 | 20.6 | 20.4 | 22.4104 | 9 | 17.4446 | 24.208 | 14.42
| Mean |  | 18.2 | 17.75 | 17.3333 | 18.847 | 18.1849 | 18.6667 | 17.45 | 16.5612 | 18.2833 | 23.05 | 31.1667 | 15.3667 | 17.7333 | 18.9167 | 21.0833 | 18.1667 | 15.5667 | 13.35 | 20.2667 | 19.6333 | 21.6231 | 9.3333 | 17.9293 | 23.7982 | 12.8833
| Median |  | 18.1 | 17.75 | 17.5 | 18.845 | 18.4526 | 18.6 | 17.45 | 16.629 | 18.35 | 23.4 | 31 | 15.3 | 17.8 | 18.8 | 21 | 18.25 | 15.55 | 13.3 | 20.25 | 19.7 | 21.6088 | 9 | 18.1405 | 23.667 | 12.695
| Std Dev. |  | 0.75631 | 0.7064 | 0.8165 | 0.27536 | 0.82977 | 0.32042 | 0.37283 | 0.2248 | 0.2137 | 1.0154 | 1.169 | 0.28048 | 0.57504 | 0.89312 | 0.43551 | 0.41793 | 0.50859 | 0.12247 | 0.21602 | 0.76594 | 0.8572 | 0.5164 | 0.38925 | 0.34207 | 0.83804
| Rel.Std.Dev. |  | 4.156% | 3.98% | 4.711% | 1.461% | 4.563% | 1.717% | 2.137% | 1.357% | 1.169% | 4.405% | 3.751% | 1.825% | 3.243% | 4.721% | 2.066% | 2.301% | 3.267% | 0.9174% | 1.066% | 3.901% | 3.964% | 5.533% | 2.171% | 1.437% | 6.505%
| PDM3 |  | -0.8073% | -3.26% | -5.531% | 2.719% | -0.8897% | 1.736% | -4.895% | -9.739% | -0.3532% | 25.63% | 69.86% | -16.25% | -3.351% | 3.099% | 14.91% | -0.989% | -15.16% | -27.24% | 10.46% | 7.005% | 17.85% | -49.13% | -2.283% | 29.7% | -29.78%
| Z-Score (Absolute) |  | 0 | 0.21 | NA | 0.3 | 0.01 | 0.22 | 0.35 | 0.77 | 0.04 | 2.28 | 6.1 | 1.33 | 0.22 | 0.34 | 1.36 | 0.02 | 1.24 | 2.28 | 0.97 | 0.67 | 1.61 | NA | 0.13 | 2.63 | 2.5
Results from laboratories 2.04 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 609c (Certified Value 5.21 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 2
| 1 | 1 | 5.2 | 5.7 | 4.9 | 15.7212 | 5.8 | 4.9 | 3.5 | 5.3 | 6.9 | 5.7 | 5.1 | 5.1 | 4.8 | 5.7 | 5.3 | 6.2 | 4.37 | 4.8 | 11.018 | 4 | 5.0608 | 6.025 | 4.4 |  |  |  | 1
| 1 | 2 | 5 | 5.5 | 5 | 5.2647 | 6 | 4.6 | 3.9479 | 5 | 7.1 | 6 | 5 | 5.1 | 5 | 5.7 | 5.4 | 5.9 | 3.8 | 5.7 | 12.1167 | 12 | 5.5721 | 5.95 | 4.5 |  |  |  | 46
| 1 | 3 | 5.2 | 6 | 5 | 5.3676 | 5.9 | 4.8 | 3.4702 | 5.2 | 7 | 5.7 | 5 | 5.1 | 4.9 | 5.8 | 5.4 | 5.9 | 4.36 | 5.3 | 18.0412 | 4 | 5.3024 | 6.144 | 4.4 |  |  |  | 16
| 1 | 4 | 5 | 5.9 | 5 | 5.1574 | 5.8 | 5 | 3.8743 | 5.2 | 7.1 | 6.1 | 5 | 5.5 | 4.8 | 5.7 | 5.1 | 6.1 | 4.7 | 4.6 | 11.4008 | 3 | 5.1408 | 5.992 | 4.8 |  |  |  | 5.2089
| 1 | 5 | 5.1 | 5.8 | 4.9 | 5.2004 | 6.3 | 4.7 | 3.211 | 5.2 | 7.2 | 5.8 | 5 | 5.2 | 5 | 5.7 | 5.3 | 6.1 | 4.4 | 4.7 | 11.7421 | 3 | 5.3184 | 5.97 | 4.6 |  |  |  | 99
| 1 | 6 | 5.4 | 6.2 | 5 | 5.1319 | 6.2 | 4.6 | 3.902 | 5.2 | 7.1 | 6 | 5 | 5.1 | 4.9 | 5.7 | 5.2 | 6.1 | 3.8 | 6.1 | 10.0287 | 3 | 5.5003 | 5.965 | 4.7
| Mean |  | 5.15 | 5.85 | 4.9667 | 6.9739 | 6 | 4.7667 | 3.6509 | 5.1833 | 7.0667 | 5.8833 | 5.0167 | 5.1833 | 4.9 | 5.7167 | 5.2833 | 6.05 | 4.2383 | 5.2 | 12.3912 | 4.8333 | 5.3158 | 6.0077 | 4.5667
| Median |  | 5.15 | 5.85 | 5 | 5.2326 | 5.95 | 4.75 | 3.6871 | 5.2 | 7.1 | 5.9 | 5 | 5.1 | 4.9 | 5.7 | 5.3 | 6.1 | 4.365 | 5.05 | 11.5715 | 3.5 | 5.3104 | 5.981 | 4.55
| Std Dev. |  | 0.15166 | 0.2429 | 0.05164 | 4.2861 | 0.20976 | 0.1633 | 0.30003 | 0.098319 | 0.10328 | 0.17224 | 0.040825 | 0.16021 | 0.089443 | 0.040825 | 0.1169 | 0.12247 | 0.36213 | 0.60663 | 2.8589 | 3.5449 | 0.19776 | 0.07169 | 0.1633
| Rel.Std.Dev. |  | 2.945% | 4.152% | 1.04% | 61.46% | 3.496% | 3.426% | 8.218% | 1.897% | 1.462% | 2.928% | 0.8138% | 3.091% | 1.825% | 0.7141% | 2.213% | 2.024% | 8.544% | 11.67% | 23.07% | 73.34% | 3.72% | 1.193% | 3.576%
| PDM3 |  | -1.131% | 12.31% | -4.651% | 33.88% | 15.19% | -8.49% | -29.91% | -0.4912% | 35.66% | 12.95% | -3.691% | -0.4912% | -5.931% | 9.748% | 1.429% | 16.15% | -18.63% | -0.1712% | 137.9% | -7.21% | 2.052% | 15.33% | -12.33%
| Z-Score (Absolute) |  | 0.11 | 0.76 | 0.34 | 2.16 | 0.94 | 0.59 | 1.98 | 0.07 | 2.27 | 0.8 | 0.28 | 0.07 | 0.43 | 0.59 | 0.05 | 1.01 | 1.25 | 0.05 | 8.9 | NA | 0.09 | 0.95 | 0.84
| Analytical results for Lu in OREAS 609c (Certified Value 0.026 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS |  |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 01g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 25g |  |  |  | 3
| 1 | 1 | 2.01: 0.02; 2.03: 0.021; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0312; 2.10: 0.021 |  |  |  | 1
| 1 | 2 | 2.01: 0.02; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0304; 2.10: 0.031 |  |  |  | 27
| 1 | 3 | 2.01: 0.03; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.028; 2.10: 0.021 |  |  |  | 16
| 1 | 4 | 2.01: 0.02; 2.03: 0.023; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0285; 2.10: 0.023 |  |  |  | 0.026139
| 1 | 5 | 2.01: 0.02; 2.03: 0.026; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0272; 2.10: 0.018 |  |  |  | 100
| 1 | 6 | 2.01: 0.02; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0309; 2.10: 0.022
| Mean |  | 2.01: 0.021667; 2.03: 0.024167; 2.07: 0.03; 2.08: 0.03; 2.09: 0.029367; 2.10: 0.022667
| Median |  | 2.01: 0.02; 2.03: 0.025; 2.07: 0.03; 2.08: 0.03; 2.09: 0.02945; 2.10: 0.0215
| Std Dev. |  | 2.01: 0.0040825; 2.03: 0.0018348; 2.07: 0; 2.08: 0; 2.09: 0.0016789; 2.10: 0.0044121
| Rel.Std.Dev. |  | 2.01: 18.84%; 2.03: 7.592%; 2.07: 0%; 2.08: 0%; 2.09: 5.717%; 2.10: 19.47%
| PDM3 |  | 2.01: -17.11%; 2.03: -7.545%; 2.07: 14.77%; 2.08: 14.77%; 2.09: 12.35%; 2.10: -13.28%
| Z-Score (Absolute) |  | 2.01: 1.28; 2.03: 0.7; 2.05: 0.5; 2.06: 5.28; 2.07: 0.65; 2.08: 0.65; 2.09: 0.5; 2.10: 1.04
| Analytical results for Mg in OREAS 609c (Certified Value 0.042 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.04 | 0.04 | 0.04 | 0.04 | 0.066339 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.046 | 0.04 | 0.04 | 0.04 | < 0.1 | 0.0446 | 0.040475 | 0.04137 |  |  |  | 1
| 1 | 2 | 0.04 | 0.04 | 0.04 | 0.04 | 0.042692 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.045 | 0.036 | 0.04 | 0.04 | < 0.1 | 0.0706 | 0.04076 | 0.041082 |  |  |  | 0
| 1 | 3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.043216 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.045 | 0.04 | 0.04 | 0.04 | < 0.1 | 0.044 | 0.041483 | 0.042451 |  |  |  | 16
| 1 | 4 | 0.04 | 0.04 | 0.04 | 0.04 | 0.042163 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.046 | 0.042 | 0.04 | 0.04 | < 0.1 | 0.0428 | 0.041958 | 0.041836 |  |  |  | 0.042072
| 1 | 5 | 0.04 | 0.04 | 0.04 | 0.04 | 0.042604 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.047 | 0.041 | 0.04 | 0.04 | < 0.1 | 0.0428 | 0.041955 | 0.041661 |  |  |  | 101
| 1 | 6 | 0.04 | 0.04 | 0.04 | 0.04 | 0.042264 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.047 | 0.04 | 0.04 | 0.04 | < 0.1 | 0.0434 | 0.040964 | 0.041104
| Mean |  | 0.04 | 0.04 | 0.04 | 0.04 | 0.046546 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.046 | 0.039833 | 0.04 | 0.04 |  | 0.048033 | 0.041266 | 0.041584
| Median |  | 0.04 | 0.04 | 0.04 | 0.04 | 0.042648 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.046 | 0.04 | 0.04 | 0.04 |  | 0.0437 | 0.041223 | 0.041516
| Std Dev. |  | 0 | 0 | 0 | 0 | 0.0097034 | 7.6012e-18 | 0 | 0 | 7.6012e-18 | 7.6012e-18 | 0 | 0 | 0 | 0 | 0 | 0.00089443 | 0.0020412 | 0 | 0 |  | 0.011077 | 0.00062823 | 0.00051934
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 20.85% | 1.52e-14% | 0% | 0% | 1.52e-14% | 1.52e-14% | 0% | 0% | 0% | 0% | 0% | 1.944% | 5.124% | 0% | 0% |  | 23.06% | 1.522% | 1.249%
| PDM3 |  | -4.924% | -4.924% | -4.924% | -4.924% | 10.64% | 18.84% | -4.924% | -4.924% | 18.84% | 18.84% | -4.924% | -4.924% | -4.924% | -4.924% | -4.924% | 9.337% | -5.321% | -4.924% | -4.924% |  | 14.17% | -1.916% | -1.159%
Z-Score (Absolute) | bar: 2.01 to 2.27
| Analytical results for Mn in OREAS 609c (Certified Value 0.023 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 0.0239 | 0.0227 | 0.0231 | 0.0217 | 0.020033 | 0.0226 | 0.023 | 0.0225 | 0.0229 | 0.0236 | 0.0224 | 0.0239 | 0.0226 | 0.0242 | 0.0226 | 0.0239 | 0.0235 | 0.02115 | 0.0218 | < 0.1 | 0.0232 | 0.021447 | 0.021263 | 0.02 |  |  | 1
| 1 | 2 | 0.0246 | 0.022 | 0.023 | 0.0224 | 0.023488 | 0.0234 | 0.0227 | 0.0226 | 0.023 | 0.0245 | 0.0219 | 0.0237 | 0.0225 | 0.0245 | 0.0231 | 0.0233 | 0.0221 | 0.02123 | 0.0232 | < 0.1 | 0.0194 | 0.021856 | 0.021285 | 0.0207 |  |  | 28
| 1 | 3 | 0.0242 | 0.0224 | 0.0231 | 0.0223 | 0.02319 | 0.0234 | 0.0225 | 0.0225 | 0.0233 | 0.0233 | 0.0219 | 0.0237 | 0.022 | 0.0239 | 0.0226 | 0.0239 | 0.0235 | 0.02104 | 0.0229 | < 0.1 | 0.0232 | 0.021665 | 0.021789 | 0.0208 |  |  | 16
| 1 | 4 | 0.0242 | 0.0218 | 0.0231 | 0.0231 | 0.022892 | 0.0233 | 0.0223 | 0.0223 | 0.0224 | 0.0243 | 0.0217 | 0.0238 | 0.0221 | 0.024 | 0.0223 | 0.0236 | 0.024 | 0.02119 | 0.0234 | < 0.1 | 0.0225 | 0.021838 | 0.021482 | 0.0216 |  |  | 0.022768
| 1 | 5 | 0.0241 | 0.0227 | 0.0231 | 0.0218 | 0.023413 | 0.023 | 0.0223 | 0.0234 | 0.0232 | 0.023 | 0.0218 | 0.0241 | 0.0229 | 0.0239 | 0.0228 | 0.0232 | 0.0241 | 0.02122 | 0.0218 | < 0.1 | 0.0225 | 0.022083 | 0.021659 | 0.023 |  |  | 102
| 1 | 6 | 0.0239 | 0.0231 | 0.0231 | 0.0222 | 0.023499 | 0.0233 | 0.0225 | 0.0229 | 0.0228 | 0.0253 | 0.0221 | 0.0236 | 0.0223 | 0.0239 | 0.0225 | 0.023 | 0.0244 | 0.02078 | 0.0233 | < 0.1 | 0.0232 | 0.021398 | 0.021254 | 0.0212
| Mean |  | 0.02415 | 0.02245 | 0.023083 | 0.02225 | 0.022752 | 0.023167 | 0.02255 | 0.0227 | 0.022933 | 0.024 | 0.021967 | 0.0238 | 0.0224 | 0.024067 | 0.02265 | 0.023483 | 0.0236 | 0.021102 | 0.022733 |  | 0.022333 | 0.021715 | 0.021455 | 0.021217
| Median |  | 0.02415 | 0.02255 | 0.0231 | 0.02225 | 0.023301 | 0.0233 | 0.0225 | 0.02255 | 0.02295 | 0.02395 | 0.0219 | 0.02375 | 0.0224 | 0.02395 | 0.0226 | 0.02345 | 0.02375 | 0.02117 | 0.02305 |  | 0.02285 | 0.021752 | 0.021383 | 0.021
| Std Dev. |  | 0.00025884 | 0.00048477 | 4.0825e-05 | 0.000501 | 0.0013522 | 0.00031411 | 0.00026646 | 0.00039497 | 0.00032042 | 0.0008579 | 0.00025033 | 0.00017889 | 0.00033466 | 0.00024221 | 0.00027386 | 0.00037639 | 0.00081486 | 0.00017198 | 0.00074207 |  | 0.0014774 | 0.00026249 | 0.00022795 | 0.0010245
| Rel.Std.Dev. |  | 1.072% | 2.159% | 0.1769% | 2.252% | 5.943% | 1.356% | 1.182% | 1.74% | 1.397% | 3.575% | 1.14% | 0.7516% | 1.494% | 1.006% | 1.209% | 1.603% | 3.453% | 0.815% | 3.264% |  | 6.615% | 1.209% | 1.062% | 4.829%
| PDM3 |  | 6.07% | -1.397% | 1.385% | -2.275% | -0.06916% | 1.751% | -0.9577% | -0.2988% | 0.726% | 5.411% | -3.52% | 4.532% | -1.616% | 5.704% | -0.5184% | 3.142% | 3.654% | -7.319% | -0.1524% |  | -1.909% | -4.627% | -5.766% | -6.814%
| Z-Score (Absolute) |  | 1.59 | 0.3 | 0.41 | 0.52 | 0.04 | 0.5 | 0.18 | 0.02 | 0.24 | 1.42 | 0.83 | 1.2 | 0.35 | 1.5 | 0.07 | 0.85 | 0.98 | 1.79 | 0.02 | 30.24 | 0.43 | 1.11 | 1.4 | 1.66
| Analytical results for Mo in OREAS 609c (Certified Value 39.5 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 2
| 1 | 1 | 40.17 | 37.94 | 40.9 | 40.7 | 6.8778 | 40.2 | 38.4 | 40.85 | 40.2 | 38.5 | 39.5 | 40.2 | 43 | 41.4 | 40.8 | 38 | 34.26 | 41.5 | 35.8 | 39.15 | 34.9106 | 36.4 | 37.4417 | 36.93 | 33.73 |  | 1
| 1 | 2 | 42.9 | 35.39 | 40.7 | 40.6 | 40.6917 | 40.5 | 38.8 | 40.83 | 40.2 | 39.9 | 38.8 | 41.3 | 43.3 | 41.8 | 42.3 | 38.2 | 33.85 | 39.3 | 36.5 | 39.96 | 34.9707 | 6.51 | 38.8795 | 36.4 | 35.61 |  | 29
| 1 | 3 | 40.83 | 40.79 | 41.5 | 40.4 | 40.8903 | 41 | 38.3 | 41.18 | 40.2 | 38 | 39.3 | 40.9 | 42 | 41.1 | 41 | 37.2 | 33.75 | 40 | 35.5 | 38.91 | 35.0632 | 36.6 | 37.3911 | 37.91 | 36.63 |  | 16
| 1 | 4 | 41.43 | 39.66 | 42.1 | 40.8 | 40.6287 | 40.3 | 38.5 | 40.26 | 40.8 | 39.7 | 38.5 | 40.1 | 42.8 | 41.8 | 39.1 | 39.8 | 33.14 | 41.2 | 35.8 | 40.5 | 34.9936 | 36.4 | 37.6147 | 36.46 | 37.94 |  | 39.5356
| 1 | 5 | 41.05 | 38.37 | 41.4 | 40.8 | 40.2155 | 41.7 | 37.9 | 41.19 | 40.9 | 36.9 | 40.6 | 41.6 | 43.6 | 39.5 | 40.6 | 36.9 | 33.51 | 40.5 | 36.2 | 39.94 | 36.3617 | 36.2 | 38.7546 | 38.16 | 39.26 |  | 103
| 1 | 6 | 40.51 | 39.86 | 41.9 | 39.9 | 40.1006 | 40.9 | 38.6 | 40.66 | 40.7 | 40.5 | 38.5 | 40.7 | 42.1 | 41 | 40.6 | 37.9 | 33.52 | 41.7 | 35.4 | 41.35 | 34.8734 | 37.1 | 37.9787 | 36.78 | 37.77
| Mean |  | 41.1483 | 38.6683 | 41.4167 | 40.5333 | 34.9008 | 40.7667 | 38.4167 | 40.8283 | 40.5 | 38.9167 | 39.2 | 40.8 | 42.8 | 41.1 | 40.7333 | 38 | 33.6717 | 40.7 | 35.8667 | 39.9683 | 35.1955 | 31.535 | 38.01 | 37.1067 | 36.8233
| Median |  | 40.94 | 39.015 | 41.45 | 40.65 | 40.4221 | 40.7 | 38.45 | 40.84 | 40.45 | 39.1 | 39.05 | 40.8 | 42.9 | 41.25 | 40.7 | 37.95 | 33.635 | 40.85 | 35.8 | 39.95 | 34.9821 | 36.4 | 37.7967 | 36.855 | 37.2
| Std Dev. |  | 0.96159 | 1.9116 | 0.54559 | 0.34448 | 13.7316 | 0.55737 | 0.30605 | 0.34822 | 0.33466 | 1.3541 | 0.8 | 0.5933 | 0.64187 | 0.85323 | 1.0231 | 1.0139 | 0.37839 | 0.93595 | 0.41793 | 0.89244 | 0.57511 | 12.2635 | 0.65933 | 0.74971 | 1.9575
| Rel.Std.Dev. |  | 2.337% | 4.944% | 1.317% | 0.8499% | 39.34% | 1.367% | 0.7967% | 0.8529% | 0.8263% | 3.48% | 2.041% | 1.454% | 1.5% | 2.076% | 2.512% | 2.668% | 1.124% | 2.3% | 1.165% | 2.233% | 1.634% | 38.89% | 1.735% | 2.02% | 5.316%
| PDM3 |  | 4.079% | -2.194% | 4.758% | 2.524% | -11.72% | 3.114% | -2.83% | 3.27% | 2.439% | -1.565% | -0.8488% | 3.198% | 8.257% | 3.957% | 3.03% | -3.884% | -14.83% | 2.945% | -9.28% | 1.095% | -10.98% | -20.24% | -3.859% | -6.144% | -6.86%
| Z-Score (Absolute) |  | 0.82 | 0.22 | 0.93 | 0.56 | 1.81 | 0.66 | 0.33 | 0.69 | 0.55 | 0.12 | 0 | 0.67 | 1.52 | 0.8 | 0.65 | 0.51 | 2.33 | 0.63 | 1.4 | 0.32 | 1.69 | 3.23 | 0.5 | 0.88 | 1
| Analytical results for Na in OREAS 609c (Certified Value 0.064 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.06 | 0.058 | 0.064 | 0.06 | 0.057733 | 0.06 | 0.06 | 0.067 | 0.06 | 0.07 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | 0.064 | 0.058 | 0.06 | 0.07 | < 0.1 | 0.059 | 0.06486 | 0.064319 |  |  |  | 1
| 1 | 2 | 0.07 | 0.056 | 0.064 | 0.06 | 0.065273 | 0.062 | 0.06 | 0.067 | 0.07 | 0.07 | 0.06 | 0.06 | 0.08 | 0.07 | 0.08 | 0.065 | 0.05 | 0.06 | 0.07 | 0.1237 | 0.052 | 0.065366 | 0.062783 |  |  |  | 0
| 1 | 3 | 0.06 | 0.056 | 0.065 | 0.06 | 0.065309 | 0.066 | 0.06 | 0.063 | 0.07 | 0.07 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | 0.066 | 0.054 | 0.07 | 0.07 | < 0.1 | 0.059 | 0.066428 | 0.062616 |  |  |  | 16
| 1 | 4 | 0.06 | 0.054 | 0.064 | 0.06 | 0.064647 | 0.062 | 0.06 | 0.063 | 0.06 | 0.07 | 0.06 | 0.1 | 0.07 | 0.07 | 0.09 | 0.063 | 0.06 | 0.07 | 0.08 | 0.1206 | 0.059 | 0.065512 | 0.062204 |  |  |  | 0.063871
| 1 | 5 | 0.06 | 0.058 | 0.064 | 0.06 | 0.064286 | 0.061 | 0.06 | 0.067 | 0.06 | 0.07 | 0.06 | 0.06 | 0.08 | 0.07 | 0.08 | 0.064 | 0.05 | 0.06 | 0.08 | 0.1301 | 0.059 | 0.066116 | 0.06239 |  |  |  | 104
| 1 | 6 | 0.06 | 0.059 | 0.064 | 0.06 | 0.064703 | 0.066 | 0.06 | 0.067 | 0.07 | 0.08 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | 0.064 | 0.06 | 0.09 | 0.08 | < 0.1 | 0.059 | 0.065655 | 0.062796
| Mean |  | 0.061667 | 0.056833 | 0.064167 | 0.06 | 0.063658 | 0.062833 | 0.06 | 0.065667 | 0.065 | 0.071667 | 0.06 | 0.066667 | 0.073333 | 0.07 | 0.081667 | 0.064333 | 0.055333 | 0.068333 | 0.075 | 0.1248 | 0.057833 | 0.065656 | 0.062851
| Median |  | 0.06 | 0.057 | 0.064 | 0.06 | 0.064675 | 0.062 | 0.06 | 0.067 | 0.065 | 0.07 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | 0.064 | 0.056 | 0.065 | 0.075 | 0.1237 | 0.059 | 0.065584 | 0.062699
| Std Dev. |  | 0.0040825 | 0.0018348 | 0.00040825 | 0 | 0.0029291 | 0.0025626 | 0 | 0.0020656 | 0.0054772 | 0.0040825 | 0 | 0.01633 | 0.005164 | 0 | 0.0040825 | 0.0010328 | 0.0046762 | 0.01169 | 0.0054772 | 0.0048446 | 0.0028577 | 0.00055614 | 0.000755
| Rel.Std.Dev. |  | 6.62% | 3.228% | 0.6362% | 0% | 4.601% | 4.078% | 0% | 3.146% | 8.427% | 5.696% | 0% | 24.49% | 7.042% | 0% | 4.999% | 1.605% | 8.451% | 17.11% | 7.303% | 3.882% | 4.941% | 0.8471% | 1.201%
| PDM3 |  | -3.452% | -11.02% | 0.4626% | -6.061% | -0.3331% | -1.625% | -6.061% | 2.811% | 1.767% | 12.2% | -6.061% | 4.377% | 14.81% | 9.596% | 27.86% | 0.7236% | -13.37% | 6.986% | 17.42% | 95.39% | -9.453% | 2.795% | -1.597%
| Z-Score (Absolute) |  | 0.41 | 1.17 | 0.03 | 0.67 | 0.11 | 0.23 | 0.67 | 0.21 | 0.1 | 1.14 | 0.67 | 0.36 | 1.4 | 0.88 | 2.7 | 0 | 1.4 | 0.62 | 1.66 | 3.59 | 0.97 | 0.21 | 0.23
| Analytical results for Nb in OREAS 609c (Certified Value 0.59 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 25g | [6 empty] | 2
| 1 | 1 | 0.27 | 0.57 | 0.66 | 0.87322 | 0.4 | 0.98 | 0.9 | 0.67 | 0.45 | 0.65 | 0.6 | 0.62 | 0.51 | 0.71 | 0.78 | < 1 | 7.215 | < 0.01 | 0.79343 | 0.27 | [6 empty] | 1
| 1 | 2 | 0.29 | 0.53 | 0.59 | 0.61568 | 0.4 | 0.95 | 1 | 0.59 | 0.42 | 0.65 | 0.61 | 0.62 | 0.5 | 0.66 | 0.62 | < 1 | 9.7566 | 0.03 | 0.85658 | 0.3 | [6 empty] | 30
| 1 | 3 | 0.3 | 0.57 | 0.62 | 0.63184 | 0.4 | 0.97 | 0.9 | 0.64 | 0.41 | 0.73 | 0.56 | 0.66 | 0.5 | 0.71 | 0.64 | < 1 | 6.4917 | 0.01 | 0.83993 | 0.35 | [6 empty] | 16
| 1 | 4 | 0.31 | 0.56 | 0.57 | 0.61468 | 0.4 | 0.99 | 0.9 | 0.62 | 0.42 | 0.65 | 0.53 | 0.58 | 0.52 | 0.74 | 0.57 | < 1 | 8.2047 | < 0.01 | 0.85711 | 0.32 | [6 empty] | 0.59447
| 1 | 5 | 0.3 | 0.59 | 0.6 | 0.6017 | 0.4 | 0.9 | 0.9 | 0.64 | 0.4 | 0.65 | 0.63 | 0.55 | 0.53 | 0.82 | 0.62 | < 1 | 8.7148 | < 0.01 | 0.84181 | 0.25 | [6 empty] | 105
| 1 | 6 | 0.3 | 0.59 | 0.63 | 0.59189 | 0.4 | 0.97 | 0.9 | 0.68 | 0.4 | 0.63 | 0.51 | 0.58 | 0.46 | 0.76 | 0.67 | < 1 | 11.392 | < 0.01 | 0.78635 | 0.33
| Mean |  | 0.295 | 0.56833 | 0.61167 | 0.65484 | 0.4 | 0.96 | 0.91667 | 0.64 | 0.41667 | 0.66 | 0.57333 | 0.60167 | 0.50333 | 0.73333 | 0.65 |  | 8.6291 | 0.02 | 0.8292 | 0.30333
| Median |  | 0.3 | 0.57 | 0.61 | 0.61518 | 0.4 | 0.97 | 0.9 | 0.64 | 0.415 | 0.65 | 0.58 | 0.6 | 0.505 | 0.725 | 0.63 |  | 8.4597 | 0.02 | 0.84087 | 0.31
| Std Dev. |  | 0.013784 | 0.022286 | 0.031885 | 0.10785 | 6.0809e-17 | 0.032249 | 0.040825 | 0.032863 | 0.018619 | 0.035214 | 0.04761 | 0.0392 | 0.024221 | 0.054283 | 0.071554 |  | 1.7698 | 0.014142 | 0.031362 | 0.037771
| Rel.Std.Dev. |  | 4.673% | 3.921% | 5.213% | 16.47% | 1.52e-14% | 3.359% | 4.454% | 5.135% | 4.469% | 5.335% | 8.304% | 6.515% | 4.812% | 7.402% | 11.01% |  | 20.51% | 70.71% | 3.782% | 12.45%
| PDM3 |  | -50.38% | -4.396% | 2.893% | 10.15% | -32.71% | 61.49% | 54.2% | 7.659% | -29.91% | 11.02% | -3.555% | 1.211% | -15.33% | 23.36% | 9.341% |  | 1352% | -96.64% | 39.49% | -48.97%
| Z-Score (Absolute) |  | 2 | 0.25 | 0.03 | 0.31 | NA | 2.27 | NA | 0.21 | 1.22 | 0.34 | 0.21 | 0.03 | 0.66 | 0.81 | 0.28 | 0.69 | 51.52 | 3.83 | 1.43 | 1.95
Results from laboratories 2.07 and 2.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 609c (Certified Value 14.5 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  |  | 1
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  |  | 1
| 1 | 1 | 2.01: 15.3; 2.03: 16.591; 2.04: 16.1185; 2.05: 15.7; 2.06: 13.61; 2.07: 25.73; 2.08: 12.31; 2.10: 21.8438; 2.11: 8.53; 2.12: 10.72 |  |  |  | 1
| 1 | 2 | 2.01: 14.26; 2.03: 15.969; 2.04: 16.9664; 2.05: 15.5; 2.06: 13.837; 2.07: 26.13; 2.08: 11.97; 2.10: 20.691; 2.11: 11.9; 2.12: 10.78 |  |  |  | 31
| 1 | 3 | 2.01: 15.59; 2.03: 16.454; 2.04: 17.472; 2.05: 15.6; 2.06: 13.586; 2.07: 25.14; 2.08: 12.08; 2.10: 19.4532; 2.11: 8.25; 2.12: 10.5 |  |  |  | 16
| 1 | 4 | 2.01: 15.31; 2.03: 15.813; 2.04: 17.0155; 2.05: 15.5; 2.06: 13.06; 2.07: 25.15; 2.08: 13; 2.10: 19.9458; 2.11: 8.28; 2.12: 10.43 |  |  |  | 14.4705
| 1 | 5 | 2.01: 15.26; 2.03: 16.064; 2.04: 17.1054; 2.05: 16.2; 2.06: 13.129; 2.07: 24.56; 2.08: 12.94; 2.10: 21.33; 2.11: 8.26; 2.12: 10.42 |  |  |  | 106
| 1 | 6 | 2.01: 16.56; 2.03: 15.854; 2.04: 17.2753; 2.05: 16; 2.06: 13.793; 2.07: 27.06; 2.08: 12.95; 2.10: 22.083; 2.11: 8.22; 2.12: 12.12
| Mean |  | 2.01: 15.38; 2.03: 16.1242; 2.04: 16.9922; 2.05: 15.75; 2.06: 13.5025; 2.07: 25.6283; 2.08: 12.5417; 2.10: 20.8911; 2.11: 8.9067; 2.12: 10.8283
| Median |  | 2.01: 15.305; 2.03: 16.0165; 2.04: 17.0605; 2.05: 15.65; 2.06: 13.598; 2.07: 25.44; 2.08: 12.625; 2.10: 21.0105; 2.11: 8.27; 2.12: 10.61
| Std Dev. |  | 2.01: 0.7371; 2.03: 0.3238; 2.04: 0.46642; 2.05: 0.2881; 2.06: 0.3317; 2.07: 0.88547; 2.08: 0.4752; 2.10: 1.0509; 2.11: 1.4708; 2.12: 0.65049
| Rel.Std.Dev. |  | 2.01: 4.793%; 2.03: 2.008%; 2.04: 2.745%; 2.05: 1.829%; 2.06: 2.457%; 2.07: 3.455%; 2.08: 3.789%; 2.10: 5.03%; 2.11: 16.51%; 2.12: 6.007%
| PDM3 |  | 2.01: 6.285%; 2.03: 11.43%; 2.04: 17.43%; 2.05: 8.842%; 2.06: -6.689%; 2.07: 77.11%; 2.08: -13.33%; 2.10: 44.37%; 2.11: -38.45%; 2.12: -25.17%
| Z-Score (Absolute) |  | 2.01: 0.05; 2.03: 0.15; 2.04: 0.38; 2.05: 0.05; 2.06: 0.55; 2.07: 2.67; 2.08: 0.8; 2.10: 1.41; 2.11: 1.77; 2.12: 1.26
| Analytical results for Ni in OREAS 609c (Certified Value 15.5 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 2
| 1 | 1 | 16.6 | 16.2 | 15.4 | 16 | 7.9952 | 14 | 15.5 | 15.1 | 15 | 15.4 | 15.1 | 15.4 | 15.6 | 15.6 | 15.8 | 15.8 | 17.4 | 13.53 | 14.7 | 22.2267 | 15.7 | 14.313 | 14.024 | 15.5 |  |  | 1
| 1 | 2 | 17.8 | 15.5 | 15.6 | 16 | 15.1896 | 14.9 | 16.6 | 15.2 | 16 | 16 | 14.8 | 15.6 | 15.2 | 15.8 | 16 | 15.8 | 16.8 | 13.63 | 15.6 | 27.171 | 7.85 | 14.8916 | 13.952 | 16.5 |  |  | 17
| 1 | 3 | 16.1 | 17.8 | 15.3 | 16 | 15.285 | 14.5 | 15.6 | 15.9 | 16 | 15.1 | 15 | 15.6 | 14.8 | 16 | 16 | 16.2 | 17.2 | 13.45 | 14.8 | 133.7924 | 15.6 | 14.4235 | 14.7 | 16.1 |  |  | 16
| 1 | 4 | 16.3 | 16.8 | 15.5 | 16.2 | 15.2059 | 14.7 | 15.5 | 15.3 | 16 | 16.2 | 14.5 | 15.2 | 14.7 | 16.1 | 14.9 | 15.6 | 17.3 | 13.74 | 15 | 25.0523 | 15.7 | 14.7932 | 14.015 | 16.8 |  |  | 15.5051
| 1 | 5 | 16.9 | 16.7 | 15.5 | 16.1 | 14.9071 | 14.9 | 15.3 | 15.7 | 16 | 15.2 | 15.2 | 15.7 | 15.8 | 14.8 | 16.2 | 15.5 | 17.2 | 13.96 | 14.5 | 24.1202 | 16 | 14.8392 | 14.548 | 18.2 |  |  | 107
| 1 | 6 | 16.3 | 17.6 | 15.5 | 16.1 | 15.0655 | 14.3 | 15.7 | 15.4 | 16 | 16.4 | 14.6 | 15.4 | 14.6 | 15.4 | 15.6 | 15.6 | 17.5 | 13.99 | 15.6 | 28.6832 | 15.9 | 14.5158 | 14.02 | 16.9
| Mean |  | 16.6667 | 16.7667 | 15.4667 | 16.0667 | 13.9414 | 14.55 | 15.7 | 15.4333 | 15.8333 | 15.7167 | 14.8667 | 15.4833 | 15.1167 | 15.6167 | 15.75 | 15.75 | 17.2333 | 13.7167 | 15.0333 | 43.5076 | 14.4583 | 14.6294 | 14.2098 | 16.6667
| Median |  | 16.45 | 16.75 | 15.5 | 16.05 | 15.1275 | 14.6 | 15.55 | 15.35 | 16 | 15.7 | 14.9 | 15.5 | 15 | 15.7 | 15.9 | 15.7 | 17.25 | 13.685 | 14.9 | 26.1117 | 15.7 | 14.6545 | 14.022 | 16.65
| Std Dev. |  | 0.62183 | 0.85946 | 0.10328 | 0.08165 | 2.916 | 0.35637 | 0.46043 | 0.30768 | 0.40825 | 0.55287 | 0.28048 | 0.18348 | 0.49967 | 0.47504 | 0.46368 | 0.251 | 0.24221 | 0.22268 | 0.46762 | 44.2885 | 3.2407 | 0.2429 | 0.32546 | 0.90921
| Rel.Std.Dev. |  | 3.731% | 5.126% | 0.6678% | 0.5082% | 20.92% | 2.449% | 2.933% | 1.994% | 2.578% | 3.518% | 1.887% | 1.185% | 3.305% | 3.042% | 2.944% | 1.594% | 1.405% | 1.623% | 3.111% | 101.8% | 22.41% | 1.66% | 2.29% | 5.455%
| PDM3 |  | 7.491% | 8.136% | -0.2482% | 3.622% | -10.09% | -6.16% | 1.257% | -0.4631% | 2.117% | 1.364% | -4.118% | -0.1407% | -2.505% | 0.7193% | 1.579% | 1.579% | 11.15% | -11.53% | -3.043% | 180.6% | -6.751% | -5.648% | -8.354% | 7.491%
| Z-Score (Absolute) |  | 1.29 | 1.4 | 0.02 | 0.64 | 1.69 | 1.02 | 0.24 | 0.05 | NA | 0.26 | 0.67 | 0 | 0.4 | 0.15 | 0.29 | 0.29 | 1.91 | 1.93 | 0.49 | 30.64 | 1.12 | 0.93 | 1.39 | 1.29
| Analytical results for P in OREAS 609c (Certified Value 0.021 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 |  |  |  |  | wt.%
|  |  | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.0229 | 0.02 | 0.02 | 0.0199 | 0.018069 | 0.021 | 0.021 | 0.0209 | 0.0223 | 0.0222 | 0.021 | 0.02 | 0.02 | 0.022 | 0.021 | 0.02 | 0.0293 | 0.02 | 0.0201 | < 0.0009 | 0.020681 | 0.025208 |  |  |  |  | 1
| 1 | 2 | 0.0222 | 0.02 | 0.021 | 0.0204 | 0.021749 | 0.021 | 0.021 | 0.0209 | 0.0219 | 0.0231 | 0.02 | 0.02 | 0.02 | 0.023 | 0.021 | 0.021 | 0.028 | 0.02 | 0.02 | < 0.0009 | 0.020705 | 0.025577 |  |  |  |  | 0
| 1 | 3 | 0.022 | 0.02 | 0.02 | 0.0204 | 0.021971 | 0.021 | 0.021 | 0.021 | 0.022 | 0.0219 | 0.021 | 0.02 | 0.02 | 0.022 | 0.021 | 0.021 | 0.0271 | 0.02 | 0.0199 | < 0.0009 | 0.020729 | 0.025435 |  |  |  |  | 16
| 1 | 4 | 0.0216 | 0.019 | 0.02 | 0.0209 | 0.02154 | 0.02 | 0.02 | 0.021 | 0.0217 | 0.0228 | 0.02 | 0.021 | 0.019 | 0.022 | 0.02 | 0.02 | 0.027 | 0.02 | 0.0201 | < 0.0009 | 0.020894 | 0.025733 |  |  |  |  | 0.02089
| 1 | 5 | 0.0217 | 0.02 | 0.021 | 0.0201 | 0.021422 | 0.021 | 0.02 | 0.0207 | 0.0221 | 0.0218 | 0.021 | 0.02 | 0.02 | 0.022 | 0.021 | 0.022 | 0.028 | 0.02 | 0.0201 | < 0.0009 | 0.021167 | 0.024492 |  |  |  |  | 108
| 1 | 6 | 0.0221 | 0.021 | 0.02 | 0.0205 | 0.021788 | 0.021 | 0.021 | 0.0211 | 0.0221 | 0.0235 | 0.021 | 0.02 | 0.019 | 0.022 | 0.021 | 0.022 | 0.027 | 0.02 | 0.02 | < 0.0009 | 0.02122 | 0.024444
| Mean |  | 0.022083 | 0.02 | 0.020333 | 0.020367 | 0.02109 | 0.020833 | 0.020667 | 0.020933 | 0.022017 | 0.02255 | 0.020667 | 0.020167 | 0.019667 | 0.022167 | 0.020833 | 0.021 | 0.027733 | 0.02 | 0.020033 |  | 0.020899 | 0.025148
| Median |  | 0.02205 | 0.02 | 0.02 | 0.0204 | 0.021645 | 0.021 | 0.021 | 0.02095 | 0.02205 | 0.0225 | 0.021 | 0.02 | 0.02 | 0.022 | 0.021 | 0.021 | 0.02755 | 0.02 | 0.02005 |  | 0.020811 | 0.025321
| Std Dev. |  | 0.00046224 | 0.00063246 | 0.0005164 | 0.00034448 | 0.0014923 | 0.00040825 | 0.0005164 | 0.00013663 | 0.00020412 | 0.0006892 | 0.0005164 | 0.00040825 | 0.0005164 | 0.00040825 | 0.00040825 | 0.00089443 | 0.00090259 | 0 | 8.165e-05 |  | 0.00024016 | 0.00055461
| Rel.Std.Dev. |  | 2.093% | 3.162% | 2.54% | 1.691% | 7.076% | 1.96% | 2.499% | 0.6527% | 0.9271% | 3.056% | 2.499% | 2.024% | 2.626% | 1.842% | 1.96% | 4.259% | 3.255% | 0% | 0.4076% |  | 1.149% | 2.205%
| PDM3 |  | 5.712% | -4.26% | -2.665% | -2.505% | 0.9573% | -0.2713% | -1.069% | 0.2074% | 5.393% | 7.946% | -1.069% | -3.463% | -5.856% | 6.111% | -0.2713% | 0.5266% | 32.76% | -4.26% | -4.101% |  | 0.04436% | 20.38%
| Z-Score (Absolute) |  | 1.15 | 0.77 | 0.46 | 0.43 | 0.24 | 0 | 0.15 | 0.09 | 1.09 | 1.58 | 0.15 | 0.61 | 1.07 | 1.23 | 0 | 0.15 | 6.34 | 0.77 | 0.74 | 18.76 | 0.06 | 3.97
| Analytical results for Pb in OREAS 609c (Certified Value 330 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 311 | 316.58 | 326.9 | 342 | 679.1915 | 297 | 327 | 335.2415 | 336 | 340 | 327 | 334 | 332 | 334 | 348 | 330 | 330 | 318 | 500 | 297.3 | 321.4 | 308.0052 | 333 | 343.1713 | 312.682 | 318 | 1
| 1 | 2 | 319 | 306.6 | 343.3 | 339.5 | 346.7409 | 305 | 322 | 329.9705 | 337 | 365 | 339 | 330 | 333 | 332 | 353 | 333 | 327 | 323 | 476 | 293.7 | 338.7 | 309.5392 | 661 | 349.1523 | 314.619 | 329 | 19
| 1 | 3 | 314 | 308.49 | 350.3 | 342.5 | 343.8697 | 302 | 330 | 335.525 | 331 | 341 | 322 | 329 | 334 | 328 | 344 | 330 | 321 | 327 | 492 | 295.3 | 325.3 | 307.4752 | 332 | 339.161 | 320.392 | 331 | 16
| 1 | 4 | 314 | 302.4 | 339.1 | 339.9 | 344.6624 | 303 | 319 | 335.788 | 339 | 357 | 335 | 328 | 330 | 327 | 346 | 322 | 329 | 320 | 495 | 293.6 | 336.2 | 311.8623 | 330 | 346.3048 | 316.942 | 345 | 329.9253
| 1 | 5 | 314 | 316.58 | 339 | 338.6 | 348.8734 | 304 | 320 | 337.1945 | 343 | 345 | 320 | 335 | 334 | 339 | 338 | 331 | 331 | 310 | 486 | 294.8 | 319.8 | 310.1845 | 339 | 353.6446 | 317.197 | 357 | 109
| 1 | 6 | 310 | 322.98 | 348.2 | 339.3 | 336.8116 | 310 | 323 | 334.3175 | 339 | 345 | 344 | 332 | 335 | 329 | 342 | 329 | 330 | 319 | 497 | 292.9 | 330.1 | 305.5367 | 336 | 335.9329 | 312.523 | 349
| Mean |  | 313.6667 | 312.2717 | 341.1333 | 340.3 | 400.0249 | 303.5 | 323.5 | 334.6728 | 337.5 | 348.8333 | 331.1667 | 331.3333 | 333 | 331.5 | 345.1667 | 329.1667 | 328 | 319.5 | 491 | 294.6 | 328.5833 | 308.7672 | 388.5 | 344.5611 | 315.7258 | 338.1667
| Median |  | 314 | 312.535 | 341.2 | 339.7 | 345.7016 | 303.5 | 322.5 | 335.3832 | 338 | 345 | 331 | 331 | 333.5 | 330.5 | 345 | 330 | 329.5 | 319.5 | 493.5 | 294.25 | 327.7 | 308.7722 | 334.5 | 344.738 | 315.7805 | 338
| Std Dev. |  | 3.1411 | 7.6904 | 8.3658 | 1.5761 | 136.824 | 4.2308 | 4.2308 | 2.4856 | 3.9875 | 9.9683 | 9.6626 | 2.8048 | 1.7889 | 4.5056 | 5.1543 | 3.7639 | 3.6878 | 5.6833 | 8.7636 | 1.5824 | 7.7742 | 2.2303 | 133.5346 | 6.5141 | 3.0386 | 14.5659
| Rel.Std.Dev. |  | 1.001% | 2.463% | 2.452% | 0.4631% | 34.2% | 1.394% | 1.308% | 0.7427% | 1.181% | 2.858% | 2.918% | 0.8465% | 0.5372% | 1.359% | 1.493% | 1.143% | 1.124% | 1.779% | 1.785% | 0.5371% | 2.366% | 0.7223% | 34.37% | 1.891% | 0.9624% | 4.307%
| PDM3 |  | -4.928% | -5.351% | 3.397% | 3.145% | 21.25% | -8.009% | -1.947% | 1.439% | 2.296% | 5.731% | 0.3763% | 0.4268% | 0.9319% | 0.4773% | 4.62% | -0.2299% | -0.5836% | -3.16% | 48.82% | -10.71% | -0.4067% | -6.413% | 17.75% | 4.436% | -4.304% | 2.498%
| Z-Score (Absolute) |  | 1.11 | 1.19 | 0.61 | 0.55 | 4.28 | 1.74 | 0.49 | 0.2 | 0.38 | 1.09 | 0.02 | 0.01 | 0.1 | 0.01 | 0.86 | 0.14 | 0.21 | 0.74 | 9.95 | 2.3 | 0.18 | 1.41 | 3.56 | 0.82 | 0.98 | 0.42
| Analytical results for Pd in OREAS 609c (Indicative Value 13.7 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.08; 2.04: 2.26 | ppb
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 01g | 1
| 1 | 1 | 2.01: < 10; 2.03: 14; 2.04: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.03: 17; 2.04: < 10 | 6
| 1 | 3 | 2.01: < 10; 2.03: 13; 2.04: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.03: 13; 2.04: < 10 | 13.6667
| 1 | 5 | 2.01: < 10; 2.03: 13; 2.04: < 10 | 12
| 1 | 6 | 2.01: < 10; 2.03: 12; 2.04: < 10
| Mean |  | 2.03: 13.6667
| Median |  | 2.03: 13
| Std Dev. |  | 2.03: 1.7512
| Rel.Std.Dev. |  | 2.03: 12.81%
| PDM3 |  | 2.03: -2.442e-13%
Z-Score (Absolute) | bar: 2.01 to 2.04
| Analytical results for Pr in OREAS 609c (Certified Value 4.03 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 4.11; 2.03: 4.317; 2.04: 4.4161; 2.05: 3.9; 2.06: 3.4705; 2.07: 6.9; 2.08: 3.65; 2.09: 5.6065; 2.10: 2.34; 2.11: 2.73 | 1
| 1 | 2 | 2.01: 3.8; 2.03: 4.147; 2.04: 4.7886; 2.05: 3.9; 2.06: 3.7944; 2.07: 6.94; 2.08: 3.54; 2.09: 5.3663; 2.10: 3.35; 2.11: 2.76 | 34
| 1 | 3 | 2.01: 4.07; 2.03: 4.295; 2.04: 4.7132; 2.05: 3.9; 2.06: 3.4176; 2.07: 6.68; 2.08: 3.6; 2.09: 5.0428; 2.10: 2.26; 2.11: 2.7 | 16
| 1 | 4 | 2.01: 4.09; 2.03: 3.937; 2.04: 4.8818; 2.05: 3.8; 2.06: 3.5833; 2.07: 6.59; 2.08: 3.81; 2.09: 5.1817; 2.10: 2.27; 2.11: 2.73 | 4.0301
| 1 | 5 | 2.01: 4.15; 2.03: 4.284; 2.04: 4.4821; 2.05: 4; 2.06: 3.6728; 2.07: 6.61; 2.08: 3.76; 2.09: 5.3493; 2.10: 2.26; 2.11: 2.65 | 110
| 1 | 6 | 2.01: 4.46; 2.03: 4.153; 2.04: 4.5583; 2.05: 4; 2.06: 3.7647; 2.07: 7.2; 2.08: 3.86; 2.09: 5.5365; 2.10: 2.28; 2.11: 3.11
| Mean |  | 2.01: 4.1133; 2.03: 4.1888; 2.04: 4.64; 2.05: 3.9167; 2.06: 3.6172; 2.07: 6.82; 2.08: 3.7033; 2.09: 5.3472; 2.10: 2.46; 2.11: 2.78
| Median |  | 2.01: 4.1; 2.03: 4.2185; 2.04: 4.6357; 2.05: 3.9; 2.06: 3.628; 2.07: 6.79; 2.08: 3.705; 2.09: 5.3578; 2.10: 2.275; 2.11: 2.73
| Std Dev. |  | 2.01: 0.21078; 2.03: 0.14367; 2.04: 0.18313; 2.05: 0.075277; 2.06: 0.15412; 2.07: 0.2374; 2.08: 0.12596; 2.09: 0.21142; 2.10: 0.43704; 2.11: 0.16589
| Rel.Std.Dev. |  | 2.01: 5.124%; 2.03: 3.43%; 2.04: 3.947%; 2.05: 1.922%; 2.06: 4.261%; 2.07: 3.481%; 2.08: 3.401%; 2.09: 3.954%; 2.10: 17.77%; 2.11: 5.967%
| PDM3 |  | 2.01: 2.066%; 2.03: 3.939%; 2.04: 15.13%; 2.05: -2.814%; 2.06: -10.24%; 2.07: 69.23%; 2.08: -8.108%; 2.09: 32.68%; 2.10: -38.96%; 2.11: -31.02%
| Z-Score (Absolute) |  | 2.01: 0.13; 2.03: 0.23; 2.04: 0.82; 2.05: 0.13; 2.06: 0.52; 2.07: 3.7; 2.08: 0.41; 2.09: 1.76; 2.10: 2.05; 2.11: 1.63
| Analytical results for Pt in OREAS 609c (Indicative Value 11.7 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.08; 2.04: 2.26 | ppb
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 01g | 1
| 1 | 1 | 2.01: 9; 2.03: < 10; 2.04: < 5 | 1
| 1 | 2 | 2.01: 13; 2.03: < 10; 2.04: < 5 | 5
| 1 | 3 | 2.01: 15; 2.03: < 10; 2.04: < 5 | 16
| 1 | 4 | 2.01: 6; 2.03: < 10; 2.04: < 5 | 11.6667
| 1 | 5 | 2.01: 15; 2.03: < 10; 2.04: < 5 | 11
| 1 | 6 | 2.01: 12; 2.03: < 10; 2.04: < 5
| Mean |  | 2.01: 11.6667
| Median |  | 2.01: 12.5
| Std Dev. |  | 2.01: 3.559
| Rel.Std.Dev. |  | 2.01: 30.51%
| PDM3 |  | 2.01: -2.887e-13%
| Z-Score (Absolute) |  | 2.01: 1.8; 2.03: 0; 2.04: 0.67
| Analytical results for Rb in OREAS 609c (Certified Value 12.4 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 1
| 1 | 1 | 12.4 | 13.1 | 13.51 | 9.9737 | 14.2 | 11.1 | 12.4 | 14.4 | 18.23 | 11.3 | 12.2 | 12.3 | 12.8 | 11.8 | 13.15 | 10.3 | 14.6 | 9.37 | 11.8508 | 11.45 | [6 empty] | 1
| 1 | 2 | 13.4 | 12.5 | 13.27 | 12.3717 | 14.7 | 11.4 | 12.2 | 14.4 | 18.96 | 10.9 | 12.4 | 12.3 | 13 | 12.2 | 12.61 | 9.6 | 15.2 | 7.99 | 12.1353 | 12.07 | [6 empty] | 35
| 1 | 3 | 12.3 | 13.5 | 13.27 | 12.5203 | 14.8 | 11.4 | 12.3 | 13.6 | 18.52 | 10.7 | 12.4 | 12.6 | 12.8 | 12.3 | 13 | 10 | 14.2 | 9.29 | 11.8002 | 12.02 | [6 empty] | 16
| 1 | 4 | 11.8 | 13.1 | 13.35 | 12.5199 | 14.4 | 11.4 | 12.1 | 14.4 | 18.48 | 11.1 | 11.9 | 12.3 | 12.7 | 10.7 | 13.57 | 10.2 | 15.3 | 9.25 | 12.1741 | 12.22 | [6 empty] | 12.4037
| 1 | 5 | 13 | 13.3 | 13.15 | 12.1617 | 15.4 | 10.9 | 12 | 13.9 | 17.92 | 11.2 | 12.3 | 12.7 | 12.2 | 12.2 | 13.52 | 9.4 | 14.5 | 9.18 | 12.1628 | 12.16 | [6 empty] | 111
| 1 | 6 | 12.8 | 14 | 13.05 | 12.1973 | 15 | 11.3 | 12.1 | 14.5 | 18.83 | 10.5 | 12.1 | 12.4 | 12.6 | 12.1 | 13.57 | 10 | 15.1 | 9.43 | 11.9899 | 13.22
| Mean |  | 12.6167 | 13.25 | 13.2667 | 11.9574 | 14.75 | 11.25 | 12.1833 | 14.2 | 18.49 | 10.95 | 12.2167 | 12.4333 | 12.6833 | 11.8833 | 13.2367 | 9.9167 | 14.8167 | 9.085 | 12.0188 | 12.19
| Median |  | 12.6 | 13.2 | 13.27 | 12.2845 | 14.75 | 11.35 | 12.15 | 14.4 | 18.5 | 11 | 12.25 | 12.35 | 12.75 | 12.15 | 13.335 | 10 | 14.85 | 9.27 | 12.0626 | 12.115
| Std Dev. |  | 0.56716 | 0.49699 | 0.15921 | 0.98381 | 0.42778 | 0.20736 | 0.1472 | 0.36332 | 0.3821 | 0.30822 | 0.19408 | 0.17512 | 0.27142 | 0.6047 | 0.38955 | 0.34881 | 0.4446 | 0.54361 | 0.1645 | 0.57508
| Rel.Std.Dev. |  | 4.495% | 3.751% | 1.2% | 8.228% | 2.9% | 1.843% | 1.208% | 2.559% | 2.067% | 2.815% | 1.589% | 1.408% | 2.14% | 5.089% | 2.943% | 3.517% | 3.001% | 5.984% | 1.369% | 4.718%
| PDM3 |  | 1.717% | 6.823% | 6.957% | -3.598% | 18.92% | -9.302% | -1.777% | 14.48% | 49.07% | -11.72% | -1.508% | 0.2386% | 2.254% | -4.196% | 6.715% | -20.05% | 19.45% | -26.76% | -3.103% | -1.723%
| Z-Score (Absolute) |  | 0.21 | 0.68 | 0.69 | 0.27 | 1.78 | 0.79 | 0.1 | 1.38 | 4.53 | 1.01 | 0.08 | 0.08 | 0.26 | 0.32 | 0.67 | 1.77 | 1.83 | 2.38 | 0.22 | 0.1
| Analytical results for Re in OREAS 609c (Certified Value 0.03 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 01g | 25g | [8 empty] | 3
| 1 | 1 | 0.039 | 0.301 | 0.033 | < 0.01 | 0.03 | 0.029 | < 0.05 | 0.028 | 0.032 | 0.03 | 0.032 | 0.031 | 0.032 | < 0.05 | 0.034 | 0.008 | < 0.05 | 0.028 | [8 empty] | 1
| 1 | 2 | 0.055 | 0.239 | 0.03 | 0.023023 | 0.03 | 0.029 | < 0.05 | 0.029 | 0.031 | 0.03 | 0.029 | 0.031 | 0.03 | < 0.05 | 0.028 | 0.01 | < 0.05 | 0.032 | [8 empty] | 36
| 1 | 3 | 0.055 | 0.301 | 0.03 | 0.023641 | 0.03 | 0.029 | < 0.05 | 0.031 | 0.029 | 0.032 | 0.033 | 0.032 | 0.03 | < 0.05 | 0.025 | 0.008 | < 0.05 | 0.031 | [8 empty] | 16
| 1 | 4 | 0.045 | 0.333 | 0.033 | 0.022773 | 0.03 | 0.029 | < 0.05 | 0.029 | 0.031 | 0.031 | 0.033 | 0.032 | 0.032 | < 0.05 | 0.029 | 0.007 | < 0.05 | 0.034 | [8 empty] | 0.030382
| 1 | 5 | 0.035 | 0.267 | 0.032 | 0.024353 | 0.031 | 0.028 | < 0.05 | 0.026 | 0.029 | 0.03 | 0.035 | 0.029 | 0.032 | < 0.05 | 0.027 | 0.007 | < 0.05 | 0.029 | [8 empty] | 112
| 1 | 6 | 0.083 | 0.331 | 0.032 | 0.023068 | 0.03 | 0.028 | < 0.05 | 0.03 | 0.028 | 0.029 | 0.033 | 0.032 | 0.032 | < 0.05 | 0.029 | 0.01 | < 0.05 | 0.027
| Mean |  | 2.01: 0.052; 2.03: 0.29533; 2.04: 0.031667; 2.05: 0.023372; 2.06: 0.030167; 2.07: 0.028667; 2.09: 0.028833; 2.10: 0.03; 2.11: 0.030333; 2.12: 0.0325; 2.13: 0.031167; 2.14: 0.031333; 2.16: 0.028667; 2.17: 0.0083333; 2.19: 0.030167
| Median |  | 2.01: 0.05; 2.03: 0.301; 2.04: 0.032; 2.05: 0.023068; 2.06: 0.03; 2.07: 0.029; 2.09: 0.029; 2.10: 0.03; 2.11: 0.03; 2.12: 0.033; 2.13: 0.0315; 2.14: 0.032; 2.16: 0.0285; 2.17: 0.008; 2.19: 0.03
| Std Dev. |  | 2.01: 0.017239; 2.03: 0.036692; 2.04: 0.0013663; 2.05: 0.00063374; 2.06: 0.00040825; 2.07: 0.0005164; 2.09: 0.0017224; 2.10: 0.0015492; 2.11: 0.0010328; 2.12: 0.0019748; 2.13: 0.001169; 2.14: 0.0010328; 2.16: 0.0030111; 2.17: 0.0013663; 2.19: 0.0026394
| Rel.Std.Dev. |  | 2.01: 33.15%; 2.03: 12.42%; 2.04: 4.315%; 2.05: 2.712%; 2.06: 1.353%; 2.07: 1.801%; 2.09: 5.974%; 2.10: 5.164%; 2.11: 3.405%; 2.12: 6.076%; 2.13: 3.751%; 2.14: 3.296%; 2.16: 10.5%; 2.17: 16.4%; 2.19: 8.75%
| PDM3 |  | 2.01: 71.15%; 2.03: 872.1%; 2.04: 4.229%; 2.05: -23.07%; 2.06: -0.7082%; 2.07: -5.645%; 2.09: -5.097%; 2.10: -1.257%; 2.11: -0.1596%; 2.12: 6.972%; 2.13: 2.583%; 2.14: 3.132%; 2.16: -5.645%; 2.17: -72.57%; 2.19: -0.7082%
| Z-Score (Absolute) |  | 9.85 | 119.24 | 0.71 | 4.39 | 0.04 | 0.64 | 2.29 | 0.56 | 0.04 | 0.11 | 1.09 | 0.49 | 0.56 | 2.29 | 0.64 | 9.78 | 2.29 | 0.04
| Analytical results for S in OREAS 609c (Certified Value 1.32 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 1.27 | 1.24 | 1.3 | 1.2191 | 3.1702 | 1.427 | 1.28 | 1.5787 | 1.28 | 1.43 | 1.3113 | 1.28 | 1.36 | 1.28 | 1.37 | 1.34 | 1.3657 | 1.29 | 1.31 | 1.42 | 1.55 | 1.2736 | 1.4 | 1.3247 | 1.23 | 1.49 | 1
| 1 | 2 | 1.31 | 1.19 | 1.3 | 1.2672 | 1.3331 | 1.436 | 1.26 | 1.5656 | 1.26 | 1.42 | 1.3641 | 1.27 | 1.36 | 1.3 | 1.38 | 1.36 | 1.358 | 1.273 | 1.32 | 1.42 | 1.48 | 1.2322 | 3.02 | 1.3303 | 1.224 | 1.3 | 12
| 1 | 3 | 1.3 | 1.21 | 1.31 | 1.2439 | 1.3259 | 1.448 | 1.25 | 1.5821 | 1.28 | 1.44 | 1.2944 | 1.26 | 1.36 | 1.27 | 1.36 | 1.34 | 1.3314 | 1.277 | 1.29 | 1.45 | 1.52 | 1.2565 | 1.4 | 1.3217 | 1.262 | 1.38 | 16
| 1 | 4 | 1.29 | 1.18 | 1.29 | 1.2915 | 1.3102 | 1.436 | 1.24 | 1.5082 | 1.28 | 1.41 | 1.3633 | 1.26 | 1.37 | 1.26 | 1.36 | 1.3 | 1.3582 | 1.259 | 1.31 | 1.44 | 1.58 | 1.2171 | 1.36 | 1.3427 | 1.231 | 1.37 | 1.3195
| 1 | 5 | 1.29 | 1.23 | 1.29 | 1.2293 | 1.3437 | 1.434 | 1.23 | 1.5884 | 1.27 | 1.44 | 1.2743 | 1.28 | 1.36 | 1.32 | 1.33 | 1.35 | 1.3537 | 1.295 | 1.31 | 1.42 | 1.5 | 1.2337 | 1.37 | 1.3393 | 1.233 | 1.41 | 113
| 1 | 6 | 1.26 | 1.25 | 1.31 | 1.247 | 1.3295 | 1.453 | 1.25 | 1.5766 | 1.29 | 1.45 | 1.3992 | 1.28 | 1.36 | 1.26 | 1.34 | 1.33 | 1.3655 | 1.302 | 1.29 | 1.42 | 1.53 | 1.2384 | 1.37 | 1.33 | 1.227 | 1.36
| Mean |  | 1.2867 | 1.2167 | 1.3 | 1.2497 | 1.6354 | 1.439 | 1.2517 | 1.5666 | 1.2767 | 1.4317 | 1.3344 | 1.2717 | 1.3617 | 1.2817 | 1.3567 | 1.3367 | 1.3554 | 1.2827 | 1.305 | 1.4283 | 1.5267 | 1.2419 | 1.6533 | 1.3314 | 1.2345 | 1.385
| Median |  | 1.29 | 1.22 | 1.3 | 1.2454 | 1.3313 | 1.436 | 1.25 | 1.5776 | 1.28 | 1.435 | 1.3373 | 1.275 | 1.36 | 1.275 | 1.36 | 1.34 | 1.3581 | 1.2835 | 1.31 | 1.42 | 1.525 | 1.2361 | 1.385 | 1.3302 | 1.2305 | 1.375
| Std Dev. |  | 0.018619 | 0.028048 | 0.0089443 | 0.026256 | 0.75197 | 0.0096333 | 0.017224 | 0.029563 | 0.010328 | 0.01472 | 0.048295 | 0.0098319 | 0.0040825 | 0.024014 | 0.018619 | 0.020656 | 0.012662 | 0.015908 | 0.012247 | 0.013292 | 0.03559 | 0.020033 | 0.66974 | 0.0081712 | 0.013838 | 0.062849
| Rel.Std.Dev. |  | 1.447% | 2.305% | 0.688% | 2.101% | 45.98% | 0.6694% | 1.376% | 1.887% | 0.809% | 1.028% | 3.619% | 0.7732% | 0.2998% | 1.874% | 1.372% | 1.545% | 0.9342% | 1.24% | 0.9385% | 0.9306% | 2.331% | 1.613% | 40.51% | 0.6137% | 1.121% | 4.538%
| PDM3 |  | -2.487% | -7.792% | -1.476% | -5.291% | 23.95% | 9.058% | -5.139% | 18.73% | -3.245% | 8.502% | 1.133% | -3.624% | 3.197% | -2.866% | 2.818% | 1.303% | 2.724% | -2.79% | -1.097% | 8.25% | 15.7% | -5.878% | 25.3% | 0.907% | -6.44% | 4.966%
| Z-Score (Absolute) |  | 0.53 | 1.33 | 0.38 | 0.96 | 3.47 | 1.22 | 0.93 | 2.68 | 0.65 | 1.13 | 0.02 | 0.7 | 0.33 | 0.59 | 0.27 | 0.04 | 0.26 | 0.58 | 0.32 | 1.09 | 2.22 | 1.04 | 3.68 | 0.02 | 1.13 | 0.6
| Analytical results for Sb in OREAS 609c (Certified Value 132 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 0
| 1 | 1 | 132 | 138.96 | 122.4 | 144.11 | 266.6335 | 122 | 143 | 130 | 122 | 109.15 | 135.5 | 138 | 126.5 | 148 | 134.5 | 128.89 | 146 | 140 | 142.4 | 140.42 | 137.3909 | 83.6 | 134.5925 | 114.86 | 97.67 |  | 1
| 1 | 2 | 134 | 133.04 | 125.2 | 143.89 | 140.4863 | 120 | 139 | 135 | 116 | 114.73 | 135.5 | 140.5 | 127.5 | 148.5 | 134 | 129.58 | 146 | 125 | 140.9 | 148.65 | 142.8524 | 249 | 134.6607 | 119.42 | 104 |  | 37
| 1 | 3 | 134 | 137.56 | 124.7 | 144.17 | 140.2736 | 124 | 135 | 131 | 112 | 108.68 | 134.5 | 140.5 | 126 | 146.5 | 132 | 123.03 | 147 | 139 | 141.4 | 138.87 | 132.9763 | 96 | 133.2937 | 117.1 | 115 |  | 16
| 1 | 4 | 134 | 133.72 | 121.1 | 144.78 | 135.6585 | 118 | 136 | 131 | 118 | 109.96 | 134 | 140 | 124 | 148 | 131.5 | 143.48 | 144 | 139 | 142.9 | 146.95 | 132.7781 | 81.7 | 132.804 | 115.39 | 121 |  | 132.2242
| 1 | 5 | 133 | 138.2 | 121.8 | 144.14 | 138.8162 | 126 | 134 | 132 | 116 | 105.65 | 137.5 | 141.5 | 130.5 | 142 | 133 | 124.98 | 144 | 140 | 143.1 | 139.76 | 133.0692 | 66.2 | 139.0374 | 114.46 | 122 |  | 114
| 1 | 6 | 132 | 140.92 | 128.1 | 143.9 | 139.9944 | 111 | 140 | 131 | 117 | 114.19 | 132.5 | 138.5 | 123 | 144.5 | 131.5 | 129.57 | 147 | 141 | 141.2 | 146.6 | 137.7333 | 62.8 | 138.6625 | 113.47 | 109
| Mean |  | 133.1667 | 137.0667 | 123.8833 | 144.165 | 160.3104 | 120.1667 | 137.8333 | 131.6667 | 116.8333 | 110.3933 | 134.9167 | 139.8333 | 126.25 | 146.25 | 132.75 | 129.9217 | 145.6667 | 137.3333 | 141.9833 | 143.5417 | 136.1334 | 106.55 | 135.5085 | 115.7833 | 111.445
| Median |  | 133.5 | 137.88 | 123.55 | 144.125 | 140.134 | 121 | 137.5 | 131 | 116.5 | 109.555 | 135 | 140.25 | 126.25 | 147.25 | 132.5 | 129.23 | 146 | 139.5 | 141.9 | 143.51 | 135.2301 | 82.65 | 134.6266 | 115.125 | 112
| Std Dev. |  | 0.98319 | 3.0781 | 2.6256 | 0.32488 | 52.1182 | 5.3072 | 3.4303 | 1.7512 | 3.2506 | 3.4753 | 1.6857 | 1.3292 | 2.6599 | 2.5446 | 1.2942 | 7.1701 | 1.3663 | 6.0882 | 0.93684 | 4.3113 | 3.9975 | 70.8374 | 2.6899 | 2.148 | 9.6592
| Rel.Std.Dev. |  | 0.7383% | 2.246% | 2.119% | 0.2254% | 32.51% | 4.417% | 2.489% | 1.33% | 2.782% | 3.148% | 1.249% | 0.9505% | 2.107% | 1.74% | 0.9749% | 5.519% | 0.9379% | 4.433% | 0.6598% | 3.004% | 2.936% | 66.48% | 1.985% | 1.855% | 8.667%
| PDM3 |  | 0.7128% | 3.662% | -6.308% | 9.031% | 21.24% | -9.119% | 4.242% | -0.4217% | -11.64% | -16.51% | 2.036% | 5.755% | -4.518% | 10.61% | 0.3977% | -1.741% | 10.17% | 3.864% | 7.381% | 8.559% | 2.956% | -19.42% | 2.484% | -12.43% | -15.72%
| Z-Score (Absolute) |  | 0.14 | 0.17 | 0.86 | 0.72 | 1.99 | 1.15 | 0.23 | 0.25 | 1.41 | 1.92 | 0 | 0.38 | 0.68 | 0.89 | 0.17 | 0.39 | 0.84 | 0.19 | 0.55 | 0.67 | 0.1 | 2.22 | 0.05 | 1.5 | 1.84
| Analytical results for Sc in OREAS 609c (Certified Value 1.18 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  | 2
| 1 | 1 | 1.5 | 1.2 | 1 | 1 | 1.2678 | 1.2 | 1.1 | 0.9247 | 1.17 | 1 | 1.4 | 1 | 1.2 | 1.2 | 1.3 | 1.1 | 1.2 | 0.77 | 1.1 | 1.5 | 1.4024 | 0.6 | 0.92284 | 0.8 |  |  | 1
| 1 | 2 | 1.5 | 1.1 | 1 | 1 | 1.1577 | 1.3 | 1.1 | 0.9089 | 1.13 | 1 | 1.4 | 1 | 1.3 | 1.3 | 1.3 | 1.2 | 1.2 | 0.7 | 1 | 1.5 | 1.6608 | 0.9 | 1.173 | 0.9 |  |  | 38
| 1 | 3 | 1.4 | 1.2 | 1 | 1 | 1.1704 | 1.2 | 1.2 | 0.8999 | 1.2 | 1 | 1.4 | 1 | 1.3 | 1.2 | 1.3 | 1.2 | 1.2 | 0.94 | 1.1 | 1.4 | 1.3731 | 0.6 | 1.0779 | 0.7 |  |  | 16
| 1 | 4 | 1.3 | 1.1 | 1 | 1 | 1.16 | 1.3 | 1.1 | 0.8654 | 1.1 | 1 | 1.5 | 1 | 1.2 | 1.2 | 1.3 | 1.1 | 1.2 | 0.9 | 1 | 1.5 | 1.477 | 0.6 | 0.96915 | 0.9 |  |  | 1.1784
| 1 | 5 | 1.3 | 1.1 | 1 | 1 | 1.142 | 1.3 | 1.1 | 0.9329 | 1.2 | 1 | 1.4 | 1 | 1.3 | 1.3 | 1.3 | 1.2 | 1.2 | 0.8 | 1 | 1.5 | 1.3148 | 0.6 | 1.0886 | 0.8 |  |  | 115
| 1 | 6 | 1.3 | 1.2 | 1 | 1 | 1.1599 | 1.4 | 1.1 | 0.952 | 1.3 | 1 | 1.4 | 1 | 1.2 | 1.2 | 1.3 | 1.1 | 1.3 | 0.9 | 1 | 1.6 | 1.3678 | 0.6 | 1.0372 | 0.9
| Mean |  | 1.3833 | 1.15 | 1 | 1 | 1.1763 | 1.2833 | 1.1167 | 0.91397 | 1.1833 | 1 | 1.4167 | 1 | 1.25 | 1.2333 | 1.3 | 1.15 | 1.2167 | 0.835 | 1.0333 | 1.5 | 1.4326 | 0.65 | 1.0448 | 0.83333
| Median |  | 1.35 | 1.15 | 1 | 1 | 1.1599 | 1.3 | 1.1 | 0.9168 | 1.185 | 1 | 1.4 | 1 | 1.25 | 1.2 | 1.3 | 1.15 | 1.2 | 0.85 | 1 | 1.5 | 1.3878 | 0.6 | 1.0575 | 0.85
| Std Dev. |  | 0.098319 | 0.054772 | 0 | 0 | 0.04576 | 0.075277 | 0.040825 | 0.030008 | 0.069474 | 0 | 0.040825 | 0 | 0.054772 | 0.05164 | 0 | 0.054772 | 0.040825 | 0.092898 | 0.05164 | 0.063246 | 0.12375 | 0.12247 | 0.089542 | 0.08165
| Rel.Std.Dev. |  | 7.107% | 4.763% | 0% | 0% | 3.89% | 5.866% | 3.656% | 3.283% | 5.871% | 0% | 2.882% | 0% | 4.382% | 4.187% | 0% | 4.763% | 3.355% | 11.13% | 4.997% | 4.216% | 8.638% | 18.84% | 8.571% | 9.798%
| PDM3 |  | 17.4% | -2.406% | -15.14% | -15.14% | -0.1743% | 8.909% | -5.235% | -22.44% | 0.4225% | -15.14% | 20.22% | -15.14% | 6.08% | 4.666% | 10.32% | -2.406% | 3.251% | -29.14% | -12.31% | 27.3% | 21.58% | -44.84% | -11.34% | -29.28%
| Z-Score (Absolute) |  | 1.08 | 0.16 | NA | NA | 0.02 | 0.55 | 0.33 | 1.4 | 0.02 | NA | 1.25 | NA | 0.37 | 0.28 | 0.64 | 0.16 | 0.19 | 1.82 | 0.77 | 1.69 | 1.34 | 2.8 | 0.71 | 1.83
Results from laboratories 2.04, 2.05 and 2.13 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 609c (Certified Value 9.94 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 3
| 1 | 1 | 10 | 7.9 | 9.9 | 9 | 27.4853 | 10 | 10.2 | 8 | 10 | 10.1 | 9.9 | 9.9 | 10.2 | 10.4 | 10.75 | 10 | 3.9 | 10.1 | 9.2 | 10.1425 | < 10 | 8.2 |  |  |  |  | 1
| 1 | 2 | 10 | 7.9 | 10 | 10 | 9.6593 | 10.5 | 10.3 | 9 | 10 | 10.2 | 10 | 10.3 | 10.9 | 10.4 | 10.74 | 10 | 2.8 | 9.8 | 27.3 | 10.0365 | < 10 | 8.5 |  |  |  |  | 39
| 1 | 3 | 10 | 8.1 | 9.9 | 10 | 9.5601 | 10.3 | 10.2 | 9 | 9 | 9.3 | 9.6 | 10.1 | 10.9 | 10.6 | 10.33 | 10 | 3.7 | 10 | 8.7 | 9.6068 | < 10 | 8.9 |  |  |  |  | 16
| 1 | 4 | 10 | 7.2 | 10 | 10 | 9.5337 | 10.2 | 10.2 | 10 | 10 | 10.2 | 9.6 | 9.7 | 10.5 | 10 | 11.19 | 10 | 3.8 | 10 | 9.2 | 9.9349 | < 10 | 9.3 |  |  |  |  | 9.9439
| 1 | 5 | 10 | 7.9 | 10 | 10 | 9.4442 | 10.6 | 10.1 | 8 | 9 | 9.5 | 10.4 | 10 | 9.7 | 10.5 | 10.33 | 10 | 3.1 | 9.5 | 9.2 | 10.2975 | < 10 | 9.5 |  |  |  |  | 116
| 1 | 6 | 10 | 8.7 | 10.1 | 10 | 9.6413 | 10.4 | 10.2 | 8 | 10 | 10 | 9.9 | 10 | 10.2 | 10.6 | 10.67 | 10 | 3 | 10.2 | 9.6 | 9.5279 | < 10 | 9.3
| Mean |  | 10 | 7.95 | 9.9833 | 9.8333 | 12.554 | 10.3333 | 10.2 | 8.6667 | 9.6667 | 9.8833 | 9.9 | 10 | 10.4 | 10.4167 | 10.6683 | 10 | 3.3833 | 9.9333 | 12.2 | 9.9243 |  | 8.95
| Median |  | 10 | 7.9 | 10 | 10 | 9.6007 | 10.35 | 10.2 | 8.5 | 10 | 10.05 | 9.9 | 10 | 10.35 | 10.45 | 10.705 | 10 | 3.4 | 10 | 9.2 | 9.9857 |  | 9.1
| Std Dev. |  | 0 | 0.48062 | 0.075277 | 0.40825 | 7.3152 | 0.21602 | 0.063246 | 0.8165 | 0.5164 | 0.38687 | 0.29665 | 0.2 | 0.46476 | 0.22286 | 0.32028 | 0 | 0.47081 | 0.25033 | 7.403 | 0.3025 |  | 0.51284
| Rel.Std.Dev. |  | 0% | 6.046% | 0.754% | 4.152% | 58.27% | 2.091% | 0.6201% | 9.421% | 5.342% | 3.914% | 2.996% | 2% | 4.469% | 2.139% | 3.002% | 0% | 13.92% | 2.52% | 60.68% | 3.048% |  | 5.73%
| PDM3 |  | 0.564% | -20.05% | 0.3964% | -1.112% | 26.25% | 3.916% | 2.575% | -12.84% | -2.788% | -0.6092% | -0.4416% | 0.564% | 4.587% | 4.754% | 7.285% | 0.564% | -65.98% | -0.1064% | 22.69% | -0.1968% |  | -9.995%
| Z-Score (Absolute) |  | 0.08 | 4.06 | 0.05 | 0.25 | 5.25 | 0.76 | 0.49 | 2.61 | 0.59 | 0.15 | 0.12 | 0.08 | 0.89 | 0.93 | 1.44 | 0.08 | 13.3 | 0.05 | 4.53 | 0.07 | 10.03 | 2.04
| Analytical results for Sm in OREAS 609c (Certified Value 2.74 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS |  |  |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 25g |  |  |  |  | 2
| 1 | 1 | 2.01: 2.93; 2.03: 3.094; 2.04: 3.1035; 2.05: 2.5; 2.06: 2.6298; 2.07: 4.67; 2.08: 2.32; 2.11: 3.903; 2.12: 1.74; 2.13: 2.36 |  |  |  |  | 1
| 1 | 2 | 2.01: 2.71; 2.03: 2.991; 2.04: 3.153; 2.05: 2.6; 2.06: 2.6437; 2.07: 4.66; 2.08: 2.32; 2.11: 3.7419; 2.12: 2.38; 2.13: 2.26 |  |  |  |  | 23
| 1 | 3 | 2.01: 2.93; 2.03: 3.149; 2.04: 3.3459; 2.05: 2.8; 2.06: 2.6058; 2.07: 4.56; 2.08: 2.3; 2.11: 3.4808; 2.12: 1.67; 2.13: 2.45 |  |  |  |  | 16
| 1 | 4 | 2.01: 2.87; 2.03: 2.851; 2.04: 3.2826; 2.05: 2.6; 2.06: 2.3985; 2.07: 4.54; 2.08: 2.5; 2.11: 3.6097; 2.12: 1.7; 2.13: 2.27 |  |  |  |  | 2.7394
| 1 | 5 | 2.01: 2.94; 2.03: 3.023; 2.04: 3.1541; 2.05: 2.7; 2.06: 2.8319; 2.07: 4.39; 2.08: 2.47; 2.11: 3.8627; 2.12: 1.68; 2.13: 2.18 |  |  |  |  | 117
| 1 | 6 | 2.01: 3.06; 2.03: 2.961; 2.04: 3.1098; 2.05: 2.8; 2.06: 2.5662; 2.07: 4.83; 2.08: 2.48; 2.11: 4.0083; 2.12: 1.66; 2.13: 2.7
| Mean |  | 2.01: 2.9067; 2.03: 3.0115; 2.04: 3.1915; 2.05: 2.6667; 2.06: 2.6126; 2.07: 4.6083; 2.08: 2.3983; 2.11: 3.7677; 2.12: 1.805; 2.13: 2.37
| Median |  | 2.01: 2.93; 2.03: 3.007; 2.04: 3.1535; 2.05: 2.65; 2.06: 2.6178; 2.07: 4.61; 2.08: 2.395; 2.11: 3.8023; 2.12: 1.69; 2.13: 2.315
| Std Dev. |  | 2.01: 0.11466; 2.03: 0.10448; 2.04: 0.099434; 2.05: 0.12111; 2.06: 0.13957; 2.07: 0.14851; 2.08: 0.093897; 2.11: 0.19655; 2.12: 0.28311; 2.13: 0.18633
| Rel.Std.Dev. |  | 2.01: 3.945%; 2.03: 3.469%; 2.04: 3.116%; 2.05: 4.541%; 2.06: 5.342%; 2.07: 3.223%; 2.08: 3.915%; 2.11: 5.217%; 2.12: 15.68%; 2.13: 7.862%
| PDM3 |  | 2.01: 6.107%; 2.03: 9.934%; 2.04: 16.5%; 2.05: -2.654%; 2.06: -4.626%; 2.07: 68.23%; 2.08: -12.45%; 2.11: 37.54%; 2.12: -34.11%; 2.13: -13.48%
| Z-Score (Absolute) |  | 2.01: 0.2; 2.03: 0.38; 2.04: 0.69; 2.05: 0.2; 2.06: 0.3; 2.07: 3.1; 2.08: 0.66; 2.11: 1.67; 2.12: 1.67; 2.13: 0.71
| Analytical results for Sn in OREAS 609c (Certified Value 7 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 3
| 1 | 1 | 7.5 | 7.6 | 6.97 | 17.482 | 6.75 | 6.6 | 6.83 | 7.3 | 7.12 | 6.7 | 7 | 6.9 | 7.1 | 7.1 | 6.9 | 4.7 | 5.3 | 7.1 | 5.68 | 7.0598 | < 10 | 6.16 |  |  |  |  | 1
| 1 | 2 | 6.8 | 7.3 | 6.91 | 7.2359 | 6.85 | 6.6 | 7 | 7.7 | 7.47 | 6.8 | 7.2 | 6.9 | 7.4 | 7.2 | 7 | 4.5 | 5.1 | 7.8 | 15.9 | 7.3769 | < 10 | 6.32 |  |  |  |  | 14
| 1 | 3 | 7.4 | 8 | 6.86 | 7.2732 | 6.84 | 6.6 | 6.83 | 7.6 | 7.14 | 6.9 | 7.2 | 6.9 | 7.2 | 7.1 | 6.9 | 4.6 | 5.1 | 7.1 | 5.67 | 7.2047 | < 10 | 6.34 |  |  |  |  | 16
| 1 | 4 | 7.2 | 8 | 6.84 | 7.1205 | 6.84 | 6.6 | 6.87 | 7.4 | 7.56 | 6.8 | 7 | 6.8 | 7.5 | 6.7 | 6.8 | 4.4 | 5.1 | 7.4 | 5.73 | 7.1897 | < 10 | 6.55 |  |  |  |  | 7.0011
| 1 | 5 | 7.7 | 7.8 | 6.84 | 7.1821 | 7.07 | 6.4 | 6.83 | 7.2 | 6.99 | 7 | 7.4 | 7.2 | 7.2 | 7.1 | 6.9 | 4.6 | 5.3 | 7.1 | 5.7 | 7.3503 | < 10 | 6.87 |  |  |  |  | 118
| 1 | 6 | 7.2 | 8.2 | 6.85 | 7.2701 | 6.89 | 6.6 | 6.9 | 7.4 | 7.53 | 6.5 | 7 | 6.8 | 7.2 | 6.9 | 7.1 | 4.6 | 5.3 | 7.6 | 5.72 | 6.9566 | < 10 | 6.77
| Mean |  | 7.3 | 7.8167 | 6.8783 | 8.9273 | 6.8733 | 6.5667 | 6.8767 | 7.4333 | 7.3017 | 6.7833 | 7.1333 | 6.9167 | 7.2667 | 7.0167 | 6.9333 | 4.5667 | 5.2 | 7.35 | 7.4 | 7.1897 |  | 6.5017
| Median |  | 7.3 | 7.9 | 6.855 | 7.253 | 6.845 | 6.6 | 6.85 | 7.4 | 7.305 | 6.8 | 7.1 | 6.9 | 7.2 | 7.1 | 6.9 | 4.6 | 5.2 | 7.25 | 5.71 | 7.1972 |  | 6.445
| Std Dev. |  | 0.30984 | 0.32506 | 0.051929 | 4.1913 | 0.10671 | 0.08165 | 0.066833 | 0.18619 | 0.24637 | 0.17224 | 0.1633 | 0.1472 | 0.15055 | 0.18348 | 0.10328 | 0.10328 | 0.10954 | 0.30166 | 4.1642 | 0.16267 |  | 0.27781
| Rel.Std.Dev. |  | 4.244% | 4.159% | 0.755% | 46.95% | 1.552% | 1.243% | 0.9719% | 2.505% | 3.374% | 2.539% | 2.289% | 2.128% | 2.072% | 2.615% | 1.49% | 2.262% | 2.107% | 4.104% | 56.27% | 2.263% |  | 4.273%
| PDM3 |  | 4.269% | 11.65% | -1.754% | 27.51% | -1.826% | -6.206% | -1.778% | 6.173% | 4.293% | -3.111% | 1.888% | -1.207% | 3.793% | 0.2218% | -0.9685% | -34.77% | -25.73% | 4.983% | 5.697% | 2.693% |  | -7.134%
| Z-Score (Absolute) |  | 0.67 | 1.74 | 0.2 | 4.04 | 0.21 | 0.85 | 0.2 | 0.95 | 0.68 | 0.4 | 0.33 | 0.12 | 0.6 | 0.09 | 0.09 | 4.98 | 3.67 | 0.78 | 0.88 | 0.44 | 4.09 | 0.98
| Analytical results for Sr in OREAS 609c (Certified Value 25.8 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 2
| 1 | 1 | 26 | 25.7 | 24 | 26.5 | 29.2967 | 23.8 | 24.3 | 25.1 | 29.6 | 30.61 | 25.4 | 25.8 | 27.4 | 29.9 | 26.2 | 25.3 | 31.8 | 25.4 | 29.2 | 27.9574 | 17 | 24.4012 | 27.01 | 21.9 |  |  | 1
| 1 | 2 | 27 | 24.6 | 25 | 26.31 | 25.0703 | 25.2 | 24.7 | 24.7 | 29.1 | 31.69 | 25 | 26.7 | 27.3 | 30 | 27.1 | 24.7 | 30 | 25.5 | 31.5 | 28.9529 | 16 | 26.2114 | 26.53 | 22.8 |  |  | 15
| 1 | 3 | 26 | 24.5 | 25 | 26.36 | 25.2481 | 24.8 | 24.6 | 25.5 | 29.9 | 30.65 | 24.4 | 27.2 | 26.4 | 29.7 | 26.8 | 24.8 | 30.5 | 25.7 | 29.2 | 28.7474 | 17 | 25.2638 | 27.41 | 22.7 |  |  | 16
| 1 | 4 | 26 | 24 | 25 | 26.34 | 25.3026 | 24.4 | 24.6 | 24.8 | 29.4 | 30.93 | 24.9 | 26 | 26.8 | 29.3 | 25.5 | 25 | 32.3 | 25.7 | 30.3 | 30.238 | 17 | 25.5026 | 26.59 | 23.2 |  |  | 25.8438
| 1 | 5 | 26 | 25.3 | 25 | 26.33 | 25.1162 | 25.1 | 23.9 | 25.3 | 29.9 | 29.86 | 25.2 | 27 | 28.2 | 28.3 | 26.5 | 25.4 | 30.6 | 26 | 29.2 | 29.7737 | 17 | 26.1379 | 26.67 | 23.9 |  |  | 119
| 1 | 6 | 25 | 25.8 | 24 | 26.32 | 25.7152 | 25.9 | 24.7 | 25.5 | 29.7 | 32.14 | 24.8 | 26 | 26.3 | 28.8 | 25.9 | 25.6 | 30.5 | 25.1 | 30.8 | 28.5037 | 17 | 24.8557 | 26.64 | 24.6
| Mean |  | 26 | 24.9833 | 24.6667 | 26.36 | 25.9582 | 24.8667 | 24.4667 | 25.15 | 29.6 | 30.98 | 24.95 | 26.45 | 27.0667 | 29.3333 | 26.3333 | 25.1333 | 30.95 | 25.5667 | 30.0333 | 29.0288 | 16.8333 | 25.3954 | 26.8083 | 23.1833
| Median |  | 26 | 24.95 | 25 | 26.335 | 25.2754 | 24.95 | 24.6 | 25.2 | 29.65 | 30.79 | 24.95 | 26.35 | 27.05 | 29.5 | 26.35 | 25.15 | 30.55 | 25.6 | 29.75 | 28.8501 | 17 | 25.3832 | 26.655 | 23
| Std Dev. |  | 0.63246 | 0.72503 | 0.5164 | 0.070711 | 1.6514 | 0.72019 | 0.31411 | 0.34496 | 0.30984 | 0.81883 | 0.34496 | 0.59245 | 0.71461 | 0.67132 | 0.58878 | 0.3559 | 0.89163 | 0.30768 | 0.98928 | 0.83972 | 0.40825 | 0.71073 | 0.3392 | 0.95376
| Rel.Std.Dev. |  | 2.433% | 2.902% | 2.094% | 0.2682% | 6.362% | 2.896% | 1.284% | 1.372% | 1.047% | 2.643% | 1.383% | 2.24% | 2.64% | 2.289% | 2.236% | 1.416% | 2.881% | 1.203% | 3.294% | 2.893% | 2.425% | 2.799% | 1.265% | 4.114%
| PDM3 |  | 0.6042% | -3.33% | -4.555% | 1.997% | 0.4424% | -3.781% | -5.329% | -2.685% | 14.53% | 19.87% | -3.459% | 2.345% | 4.732% | 13.5% | 1.894% | -2.749% | 19.76% | -1.073% | 16.21% | 12.32% | -34.87% | -1.735% | 3.732% | -10.29%
| Z-Score (Absolute) |  | 0.01 | 0.63 | 0.84 | 0.24 | 0.01 | 0.71 | 0.96 | 0.53 | 2.31 | 3.19 | 0.66 | 0.3 | 0.69 | 2.14 | 0.23 | 0.54 | 3.17 | 0.26 | 2.58 | 1.94 | 5.83 | 0.37 | 0.53 | 1.78
| Analytical results for Ta in OREAS 609c (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [9 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | [9 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | [9 empty] | 26
| 1 | 3 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | 16
| 1 | 4 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.03 | [9 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | [9 empty] | 120
| 1 | 6 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01
| Mean |  | 2.01: 0.01; 2.20: 0.025
| Median |  | 2.01: 0.01; 2.20: 0.025
| Std Dev. |  | 2.01: 0; 2.20: 0.0070711
| Rel.Std.Dev. |  | 2.01: 0%; 2.20: 28.28%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.20
| Analytical results for Tb in OREAS 609c (Certified Value 0.31 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  | 2
| 1 | 1 | 2.01: 0.31; 2.03: 0.315; 2.04: 0.31145; 2.05: 0.3; 2.06: 0.3615; 2.07: 0.42; 2.08: 0.29; 2.10: 0.3726; 2.11: 0.21; 2.12: 0.218 |  |  | 1
| 1 | 2 | 2.01: 0.29; 2.03: 0.308; 2.04: 0.32451; 2.05: 0.3; 2.06: 0.3292; 2.07: 0.42; 2.08: 0.3; 2.10: 0.3763; 2.11: 0.25; 2.12: 0.232 |  |  | 27
| 1 | 3 | 2.01: 0.32; 2.03: 0.319; 2.04: 0.31594; 2.05: 0.3; 2.06: 0.3588; 2.07: 0.41; 2.08: 0.3; 2.10: 0.3313; 2.11: 0.2; 2.12: 0.219 |  |  | 16
| 1 | 4 | 2.01: 0.3; 2.03: 0.302; 2.04: 0.31482; 2.05: 0.3; 2.06: 0.3947; 2.07: 0.42; 2.08: 0.3; 2.10: 0.345; 2.11: 0.2; 2.12: 0.228 |  |  | 0.31143
| 1 | 5 | 2.01: 0.32; 2.03: 0.316; 2.04: 0.3288; 2.05: 0.3; 2.06: 0.3363; 2.07: 0.4; 2.08: 0.31; 2.10: 0.361; 2.11: 0.2; 2.12: 0.215 |  |  | 121
| 1 | 6 | 2.01: 0.32; 2.03: 0.312; 2.04: 0.32264; 2.05: 0.3; 2.06: 0.3217; 2.07: 0.44; 2.08: 0.31; 2.10: 0.3886; 2.11: 0.2; 2.12: 0.246
| Mean |  | 2.01: 0.31; 2.03: 0.312; 2.04: 0.31969; 2.05: 0.3; 2.06: 0.35037; 2.07: 0.41833; 2.08: 0.30167; 2.10: 0.36247; 2.11: 0.21; 2.12: 0.22633
| Median |  | 2.01: 0.315; 2.03: 0.3135; 2.04: 0.31929; 2.05: 0.3; 2.06: 0.34755; 2.07: 0.42; 2.08: 0.3; 2.10: 0.3668; 2.11: 0.2; 2.12: 0.2235
| Std Dev. |  | 2.01: 0.012649; 2.03: 0.0061644; 2.04: 0.0066405; 2.05: 0; 2.06: 0.026938; 2.07: 0.013292; 2.08: 0.0075277; 2.10: 0.02124; 2.11: 0.02; 2.12: 0.011605
| Rel.Std.Dev. |  | 2.01: 4.08%; 2.03: 1.976%; 2.04: 2.077%; 2.05: 0%; 2.06: 7.688%; 2.07: 3.177%; 2.08: 2.495%; 2.10: 5.86%; 2.11: 9.524%; 2.12: 5.127%
| PDM3 |  | 2.01: -0.4589%; 2.03: 0.1833%; 2.04: 2.654%; 2.05: -3.67%; 2.06: 12.5%; 2.07: 34.33%; 2.08: -3.135%; 2.10: 16.39%; 2.11: -32.57%; 2.12: -27.32%
| Z-Score (Absolute) |  | 2.01: 0.04; 2.03: 0; 2.04: 0.14; 2.05: NA; 2.06: 0.67; 2.07: 1.87; 2.08: 0.18; 2.10: 0.89; 2.11: 1.79; 2.12: 1.51
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 609c (Certified Value 12.3 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 2
| 1 | 1 | 10.54 | 12.86 | 12.5 | 12.7 | 42.8085 | 12.3 | 12.3 | 10.8 | 11.91 | 11.5 | 11.8 | 12.25 | 13.3 | 13.25 | 13.13 | 11.39 | 12.5 | 11.31 | 9.63 | 12.1791 | 12.26 | 8.1 |  |  |  |  | 1
| 1 | 2 | 11.92 | 12.3 | 12.5 | 12.2 | 12.1157 | 12.3 | 12.5 | 10.9 | 12.65 | 11.5 | 12.2 | 12.2 | 13.4 | 13.2 | 13.15 | 11.39 | 11.5 | 12.26 | 36 | 12.4278 | 15.48 | 8.77 |  |  |  |  | 28
| 1 | 3 | 11.02 | 14.05 | 12.6 | 12.7 | 12.2224 | 12.3 | 12.7 | 10.7 | 12.01 | 11.6 | 12.15 | 12.25 | 13.2 | 13.1 | 12.91 | 11.25 | 13.1 | 11.36 | 9.44 | 11.8645 | 15.44 | 8.88 |  |  |  |  | 16
| 1 | 4 | 11.16 | 13.59 | 12.5 | 12.4 | 12.4376 | 12.3 | 11.8 | 11.1 | 12.34 | 11.25 | 11.85 | 11.8 | 13.6 | 12.3 | 13.59 | 11.43 | 12.5 | 12.47 | 9.55 | 12.3014 | 12.97 | 8.54 |  |  |  |  | 12.2846
| 1 | 5 | 10.82 | 13.34 | 12.6 | 12.1 | 12.1248 | 12.7 | 12 | 10.8 | 11.85 | 11.7 | 12.25 | 12.7 | 12.7 | 13 | 13.29 | 11.11 | 13.4 | 11.75 | 9.52 | 12.0597 | 15.92 | 8.85 |  |  |  |  | 122
| 1 | 6 | 12.11 | 13.89 | 12.7 | 12.5 | 12.7431 | 12.3 | 12.5 | 10.7 | 12.63 | 11.35 | 12.05 | 11.9 | 13.15 | 13.2 | 13.3 | 11.33 | 12.9 | 12.46 | 9.49 | 12.0422 | 12.96 | 8.84
| Mean |  | 11.2617 | 13.3383 | 12.5667 | 12.4333 | 17.4087 | 12.3667 | 12.3 | 10.8333 | 12.2317 | 11.4833 | 12.05 | 12.1833 | 13.225 | 13.0083 | 13.2283 | 11.3167 | 12.65 | 11.935 | 13.9383 | 12.1458 | 14.1717 | 8.6633
| Median |  | 11.09 | 13.465 | 12.55 | 12.45 | 12.33 | 12.3 | 12.4 | 10.8 | 12.175 | 11.5 | 12.1 | 12.225 | 13.25 | 13.15 | 13.22 | 11.36 | 12.7 | 12.005 | 9.535 | 12.1194 | 14.205 | 8.805
| Std Dev. |  | 0.62252 | 0.66047 | 0.08165 | 0.25033 | 12.4456 | 0.1633 | 0.34059 | 0.15055 | 0.3588 | 0.1633 | 0.18708 | 0.31728 | 0.3029 | 0.35835 | 0.22666 | 0.11911 | 0.66257 | 0.53347 | 10.8082 | 0.20107 | 1.609 | 0.3023
| Rel.Std.Dev. |  | 5.528% | 4.952% | 0.6497% | 2.013% | 71.49% | 1.32% | 2.769% | 1.39% | 2.933% | 1.422% | 1.553% | 2.604% | 2.29% | 2.755% | 1.713% | 1.052% | 5.238% | 4.47% | 77.54% | 1.655% | 11.35% | 3.489%
| PDM3 |  | -8.327% | 8.578% | 2.296% | 1.211% | 41.71% | 0.6684% | 0.1257% | -11.81% | -0.4305% | -6.522% | -1.909% | -0.824% | 7.655% | 5.892% | 7.683% | -7.879% | 2.975% | -2.846% | 13.46% | -1.13% | 15.36% | -29.48%
| Z-Score (Absolute) |  | 0.95 | 0.89 | 0.21 | 0.09 | 4.49 | 0.03 | 0.03 | 1.33 | 0.09 | 0.75 | 0.25 | 0.13 | 0.79 | 0.6 | 0.79 | 0.9 | 0.28 | 0.35 | 1.42 | 0.17 | 1.63 | 3.25
| Analytical results for Th in OREAS 609c (Certified Value 6.73 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  |  | 2
| 1 | 1 | 7.4 | 6.5 | 6.6 | 6.85 | 6.9972 | 7.2 | 6.2 | 5.7389 | 8.27 | 8.55 | 6.5 | 7 | 6.8 | 7.1 | 6.1 | 5.48 | 5.2 | 8.1 | 7.2 | 7.8155 | 4.45 | 6.7747 | 5.79 |  |  |  | 1
| 1 | 2 | 8 | 6.2 | 6.8 | 6.71 | 6.5031 | 7.5 | 6.2 | 5.5996 | 8.69 | 8.6 | 6.5 | 7.2 | 6.6 | 7.1 | 6.3 | 5.21 | 4.9 | 7.9 | 7 | 8.3589 | 5.84 | 6.9962 | 5.97 |  |  |  | 29
| 1 | 3 | 7.4 | 6.9 | 6.8 | 6.83 | 6.9181 | 7.4 | 6.4 | 5.3913 | 8.43 | 8.35 | 6.5 | 6.9 | 6.6 | 7 | 7.1 | 5.33 | 5.1 | 8.1 | 7 | 7.8552 | 4.56 | 6.6071 | 5.69 |  |  |  | 16
| 1 | 4 | 7.3 | 6.6 | 6.6 | 6.59 | 6.4854 | 7.2 | 6.3 | 5.3645 | 8.82 | 8.55 | 6.7 | 6.9 | 6.4 | 7.1 | 6 | 5.73 | 5.2 | 7.9 | 7.4 | 7.3102 | 4.38 | 6.9368 | 5.88 |  |  |  | 6.7274
| 1 | 5 | 7.4 | 6.7 | 6.7 | 6.73 | 6.3645 | 7.5 | 6.1 | 5.048 | 8.52 | 8.1 | 6.5 | 6.9 | 6.8 | 6.8 | 6.4 | 5.5 | 5.1 | 7.8 | 8.1 | 7.5646 | 4.38 | 7.0008 | 5.74 |  |  |  | 123
| 1 | 6 | 7.4 | 6.9 | 6.9 | 6.74 | 6.4107 | 7.6 | 6.3 | 5.5775 | 8.7 | 8.8 | 6.2 | 6.8 | 6.6 | 7.1 | 6.5 | 5.61 | 5.2 | 7.9 | 8.2 | 8.166 | 4.44 | 6.6868 | 6.57
| Mean |  | 7.4833 | 6.6333 | 6.7333 | 6.7417 | 6.6132 | 7.4 | 6.25 | 5.4533 | 8.5717 | 8.4917 | 6.4833 | 6.95 | 6.6333 | 7.0333 | 6.4 | 5.4767 | 5.1167 | 7.95 | 7.4833 | 7.8451 | 4.675 | 6.8337 | 5.94
| Median |  | 7.4 | 6.65 | 6.75 | 6.735 | 6.4943 | 7.45 | 6.25 | 5.4844 | 8.605 | 8.55 | 6.5 | 6.9 | 6.6 | 7.1 | 6.35 | 5.49 | 5.15 | 7.9 | 7.3 | 7.8354 | 4.445 | 6.8557 | 5.835
| Std Dev. |  | 0.25626 | 0.26583 | 0.12111 | 0.09347 | 0.27266 | 0.16733 | 0.10488 | 0.24265 | 0.20292 | 0.23962 | 0.16021 | 0.13784 | 0.15055 | 0.12111 | 0.38987 | 0.18715 | 0.1169 | 0.12247 | 0.53821 | 0.38282 | 0.57452 | 0.16816 | 0.32447
| Rel.Std.Dev. |  | 3.424% | 4.008% | 1.799% | 1.386% | 4.123% | 2.261% | 1.678% | 4.45% | 2.367% | 2.822% | 2.471% | 1.983% | 2.27% | 1.722% | 6.092% | 3.417% | 2.285% | 1.541% | 7.192% | 4.88% | 12.29% | 2.461% | 5.462%
| PDM3 |  | 11.24% | -1.398% | 0.08857% | 0.2124% | -1.697% | 9.998% | -7.096% | -18.94% | 27.41% | 26.23% | -3.628% | 3.309% | -1.398% | 4.548% | -4.866% | -18.59% | -23.94% | 18.17% | 11.24% | 16.61% | -30.51% | 1.581% | -11.7%
| Z-Score (Absolute) |  | 0.76 | 0.1 | 0 | 0.01 | 0.12 | 0.67 | 0.49 | 1.29 | 1.86 | 1.78 | 0.25 | 0.22 | 0.1 | 0.3 | 0.34 | 1.27 | 1.64 | 1.23 | 0.76 | 1.12 | 2.08 | 0.1 | 0.8
| Analytical results for Ti in OREAS 609c (Certified Value 0.018 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | [5 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | [5 empty] | 3
| 1 | 1 | 0.01 | 0.018 | 0.015 | 0.0167 | 0.011356 | 0.03 | 0.017 | 0.0282 | 0.0217 | 0.017 | 0.018 | 0.019 | 0.022 | 0.018 | 0.019 | 0.017 | 0.03 | 0.013 | < 0.1 | 0.0126 | 0.021166 | [5 empty] | 1
| 1 | 2 | 0.02 | 0.017 | 0.017 | 0.0169 | 0.017929 | 0.03 | 0.017 | 0.0273 | 0.0215 | 0.016 | 0.018 | 0.019 | 0.022 | 0.018 | 0.019 | 0.012 | 0.03 | 0.013 | < 0.1 | 0.0078 | 0.021787 | [5 empty] | 30
| 1 | 3 | 0.02 | 0.018 | 0.016 | 0.0167 | 0.01795 | 0.03 | 0.018 | 0.0283 | 0.0215 | 0.016 | 0.018 | 0.018 | 0.022 | 0.018 | 0.019 | 0.02 | 0.03 | 0.012 | < 0.1 | 0.012 | 0.021084 | [5 empty] | 16
| 1 | 4 | 0.02 | 0.018 | 0.016 | 0.0167 | 0.017818 | 0.03 | 0.017 | 0.0272 | 0.0219 | 0.016 | 0.02 | 0.018 | 0.021 | 0.018 | 0.021 | 0.02 | 0.03 | 0.014 | < 0.1 | 0.012 | 0.022176 | [5 empty] | 0.017716
| 1 | 5 | 0.02 | 0.018 | 0.016 | 0.0164 | 0.017582 | 0.03 | 0.016 | 0.0282 | 0.0214 | 0.016 | 0.018 | 0.019 | 0.021 | 0.018 | 0.022 | 0.017 | 0.03 | 0.013 | < 0.1 | 0.012 | 0.022131 | [5 empty] | 124
| 1 | 6 | 0.01 | 0.02 | 0.017 | 0.0168 | 0.017675 | 0.03 | 0.017 | 0.0288 | 0.0231 | 0.016 | 0.018 | 0.018 | 0.021 | 0.018 | 0.021 | 0.02 | 0.03 | 0.013 | < 0.1 | 0.012 | 0.021443
| Mean |  | 0.016667 | 0.018167 | 0.016167 | 0.0167 | 0.016718 | 0.03 | 0.017 | 0.028 | 0.02185 | 0.016167 | 0.018333 | 0.0185 | 0.0215 | 0.018 | 0.020167 | 0.017667 | 0.03 | 0.013 |  | 0.0114 | 0.021631
| Median |  | 0.02 | 0.018 | 0.016 | 0.0167 | 0.017746 | 0.03 | 0.017 | 0.0282 | 0.0216 | 0.016 | 0.018 | 0.0185 | 0.0215 | 0.018 | 0.02 | 0.0185 | 0.03 | 0.013 |  | 0.012 | 0.021615
| Std Dev. |  | 0.005164 | 0.00098319 | 0.00075277 | 0.00016733 | 0.0026306 | 0 | 0.00063246 | 0.0006229 | 0.00063797 | 0.00040825 | 0.0008165 | 0.00054772 | 0.00054772 | 0 | 0.0013292 | 0.0031411 | 0 | 0.00063246 |  | 0.0017799 | 0.00047364
| Rel.Std.Dev. |  | 30.98% | 5.412% | 4.656% | 1.002% | 15.73% | 0% | 3.72% | 2.225% | 2.92% | 2.525% | 4.454% | 2.961% | 2.548% | 0% | 6.591% | 17.78% | 0% | 4.865% |  | 15.61% | 2.19%
| PDM3 |  | -5.921% | 2.546% | -8.743% | -5.733% | -5.63% | 69.34% | -4.039% | 58.05% | 23.34% | -8.743% | 3.487% | 4.428% | 21.36% | 1.606% | 13.84% | -0.276% | 69.34% | -26.62% |  | -35.65% | 22.1%
| Z-Score (Absolute) |  | 0.51 | 0 | 0.67 | 0.49 | 0.49 | 3.99 | 0.39 | 3.32 | 1.24 | 0.67 | 0.06 | 0.11 | 1.12 | 0.06 | 0.67 | 0.17 | 3.99 | 1.74 | 10.73 | 2.28 | 1.17
| Analytical results for Tl in OREAS 609c (Certified Value 0.61 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 01g | 25g |  |  |  |  | 3
| 1 | 1 | 0.53 | 0.57 | 0.6 | 0.64 | 1.5086 | 0.59 | 0.65 | 0.62 | 0.67 | 0.68 | 0.6 | 0.62 | 0.59 | 0.61 | 0.55 | 0.597 | 0.62 | 0.57 | < 5 | 0.4 | 0.6432 | 0.47 |  |  |  |  | 1
| 1 | 2 | 0.56 | 0.54 | 0.7 | 0.61 | 0.61502 | 0.6 | 0.63 | 0.58 | 0.64 | 0.67 | 0.62 | 0.63 | 0.58 | 0.61 | 0.56 | 0.605 | 0.65 | 0.54 | < 5 | 0.41 | 0.65 | 0.49 |  |  |  |  | 31
| 1 | 3 | 0.53 | 0.62 | 0.7 | 0.63 | 0.6408 | 0.6 | 0.64 | 0.59 | 0.65 | 0.66 | 0.59 | 0.59 | 0.58 | 0.6 | 0.55 | 0.562 | 0.66 | 0.55 | < 5 | 0.44 | 0.62732 | 0.49 |  |  |  |  | 16
| 1 | 4 | 0.52 | 0.63 | 0.6 | 0.62 | 0.60111 | 0.59 | 0.64 | 0.61 | 0.68 | 0.68 | 0.6 | 0.61 | 0.58 | 0.61 | 0.53 | 0.622 | 0.62 | 0.58 | < 5 | 0.43 | 0.64972 | 0.51 |  |  |  |  | 0.60816
| 1 | 5 | 0.5 | 0.6 | 0.6 | 0.62 | 0.59653 | 0.61 | 0.63 | 0.61 | 0.64 | 0.64 | 0.64 | 0.6 | 0.59 | 0.59 | 0.56 | 0.58 | 0.61 | 0.55 | < 5 | 0.4 | 0.63504 | 0.52 |  |  |  |  | 125
| 1 | 6 | 0.55 | 0.62 | 0.6 | 0.62 | 0.62736 | 0.62 | 0.62 | 0.61 | 0.67 | 0.7 | 0.59 | 0.6 | 0.58 | 0.62 | 0.55 | 0.594 | 0.64 | 0.59 | < 5 | 0.45 | 0.62762 | 0.52
| Mean |  | 0.53167 | 0.59667 | 0.63333 | 0.62333 | 0.7649 | 0.60167 | 0.635 | 0.60333 | 0.65833 | 0.67167 | 0.60667 | 0.60833 | 0.58333 | 0.60667 | 0.55 | 0.59333 | 0.63333 | 0.56333 |  | 0.42167 | 0.63882 | 0.5
| Median |  | 0.53 | 0.61 | 0.6 | 0.62 | 0.62119 | 0.6 | 0.635 | 0.61 | 0.66 | 0.675 | 0.6 | 0.605 | 0.58 | 0.61 | 0.55 | 0.5955 | 0.63 | 0.56 |  | 0.42 | 0.63912 | 0.5
| Std Dev. |  | 0.02137 | 0.035024 | 0.05164 | 0.010328 | 0.3647 | 0.01169 | 0.010488 | 0.015055 | 0.017224 | 0.020412 | 0.019664 | 0.01472 | 0.005164 | 0.010328 | 0.010954 | 0.020646 | 0.019664 | 0.019664 |  | 0.02137 | 0.010344 | 0.02
| Rel.Std.Dev. |  | 4.019% | 5.87% | 8.154% | 1.657% | 47.68% | 1.943% | 1.652% | 2.495% | 2.616% | 3.039% | 3.241% | 2.42% | 0.8853% | 1.702% | 1.992% | 3.48% | 3.105% | 3.491% |  | 5.068% | 1.619% | 4%
| PDM3 |  | -12.58% | -1.889% | 4.14% | 2.495% | 25.77% | -1.067% | 4.414% | -0.7932% | 8.251% | 10.44% | -0.2451% | 0.02899% | -4.082% | -0.2451% | -9.563% | -2.437% | 4.14% | -7.37% |  | -30.66% | 5.042% | -17.78%
| Z-Score (Absolute) |  | 1.84 | 0.25 | 0.65 | 0.41 | 3.88 | 0.12 | 0.69 | 0.08 | 1.27 | 1.59 | 0 | 0.04 | 0.57 | 0 | 1.39 | 0.33 | 0.65 | 1.06 | 46.43 | 4.54 | 0.79 | 2.62
| Analytical results for Tm in OREAS 609c (Indicative Value 0.04 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.12; 2.07: 2.23 |  |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS |  |  |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 01g; 2.07: 0.5g |  |  |  |  | 3
| 1 | 1 | 2.01: 0.04; 2.03: 0.04; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0398 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.04; 2.03: 0.04; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.046 |  |  |  |  | 6
| 1 | 3 | 2.01: 0.04; 2.03: 0.042; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0335 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.04; 2.03: 0.039; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0359 |  |  |  |  | 0.039822
| 1 | 5 | 2.01: 0.04; 2.03: 0.038; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0392 |  |  |  |  | 12
| 1 | 6 | 2.01: 0.04; 2.03: 0.043; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0404
| Mean |  | 2.01: 0.04; 2.03: 0.040333; 2.06: 0.05; 2.07: 0.039133
| Median |  | 2.01: 0.04; 2.03: 0.04; 2.06: 0.05; 2.07: 0.0395
| Std Dev. |  | 2.01: 0; 2.03: 0.0018619; 2.06: 7.6012e-18; 2.07: 0.0042754
| Rel.Std.Dev. |  | 2.01: 0%; 2.03: 4.616%; 2.06: 1.52e-14%; 2.07: 10.93%
| PDM3 |  | 2.01: 0.4464%; 2.03: 1.283%; 2.06: 25.56%; 2.07: -1.73%
| Z-Score (Absolute) |  | 2.01: 0.02; 2.03: 0.02; 2.04: 1.88; 2.05: 1.22; 2.06: 1.22; 2.07: 0.13
| Analytical results for U in OREAS 609c (Certified Value 2.22 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  | 3
| 1 | 1 | 2.21 | 2.27 | 2.3 | 2.36 | 2.0056 | 2.3 | 2.06 | 1.4988 | 2.31 | 2.63 | 2.8 | 2.15 | 2.1 | 2.26 | 2.4 | 2.03 | 2.15 | 2.22 | 1.81 | 2.15 | 2.201 | 1.37 | 2.0878 | 1.73 |  |  | 1
| 1 | 2 | 2.18 | 2.11 | 2.4 | 2.3 | 2.2834 | 2.4 | 2.05 | 1.8002 | 2.28 | 2.96 | 2.78 | 2.14 | 2.16 | 2.2 | 2.33 | 2.02 | 2.11 | 2.17 | 1.71 | 2.16 | 2.5324 | 1.44 | 2.3669 | 1.79 |  |  | 33
| 1 | 3 | 2.07 | 2.31 | 2.4 | 2.36 | 2.3316 | 2.4 | 2.11 | 1.8587 | 2.31 | 2.71 | 2.7 | 2.2 | 1.98 | 2.27 | 2.29 | 2.08 | 2.03 | 2.14 | 1.74 | 2.16 | 1.9537 | 1.59 | 2.1819 | 1.71 |  |  | 16
| 1 | 4 | 2.14 | 2.33 | 2.4 | 2.26 | 2.3759 | 2.3 | 2.2 | 1.7246 | 2.26 | 2.52 | 2.81 | 2.21 | 1.95 | 2.15 | 2.33 | 1.91 | 2.2 | 2.31 | 1.81 | 2.22 | 2.0787 | 1.31 | 2.3891 | 1.77 |  |  | 2.2201
| 1 | 5 | 2.13 | 2.26 | 2.4 | 2.26 | 2.3203 | 2.4 | 2.01 | 1.9825 | 2.23 | 2.76 | 2.63 | 2.2 | 2.05 | 2.28 | 2.26 | 2.08 | 2.1 | 2.27 | 1.76 | 2.13 | 2.1972 | 1.28 | 2.3864 | 1.69 |  |  | 126
| 1 | 6 | 2.18 | 2.42 | 2.4 | 2.31 | 2.3181 | 2.4 | 2.06 | 1.7702 | 2.32 | 2.79 | 2.88 | 2.06 | 2.1 | 2.16 | 2.29 | 2 | 2.22 | 2.37 | 1.8 | 2.18 | 2.3091 | 1.32 | 2.1815 | 1.99
| Mean |  | 2.1517 | 2.2833 | 2.3833 | 2.3083 | 2.2725 | 2.3667 | 2.0817 | 1.7725 | 2.285 | 2.7283 | 2.7667 | 2.16 | 2.0567 | 2.22 | 2.3167 | 2.02 | 2.135 | 2.2467 | 1.7717 | 2.1667 | 2.212 | 1.385 | 2.2656 | 1.78
| Median |  | 2.16 | 2.29 | 2.4 | 2.305 | 2.3192 | 2.4 | 2.06 | 1.7852 | 2.295 | 2.735 | 2.79 | 2.175 | 2.075 | 2.23 | 2.31 | 2.025 | 2.13 | 2.245 | 1.78 | 2.16 | 2.1991 | 1.345 | 2.2744 | 1.75
| Std Dev. |  | 0.049565 | 0.10231 | 0.040825 | 0.044907 | 0.13411 | 0.05164 | 0.066156 | 0.16094 | 0.035071 | 0.14959 | 0.088468 | 0.056921 | 0.079666 | 0.057619 | 0.048854 | 0.062929 | 0.070071 | 0.086872 | 0.041673 | 0.030768 | 0.19845 | 0.11502 | 0.13101 | 0.10936
| Rel.Std.Dev. |  | 2.304% | 4.481% | 1.713% | 1.945% | 5.901% | 2.182% | 3.178% | 9.08% | 1.535% | 5.483% | 3.198% | 2.635% | 3.874% | 2.595% | 2.109% | 3.115% | 3.282% | 3.867% | 2.352% | 1.42% | 8.971% | 8.305% | 5.782% | 6.144%
| PDM3 |  | -3.082% | 2.849% | 7.353% | 3.975% | 2.36% | 6.603% | -6.235% | -20.16% | 2.924% | 22.89% | 24.62% | -2.706% | -7.361% | -0.003665% | 4.351% | -9.012% | -3.832% | 1.197% | -20.2% | -2.406% | -0.3633% | -37.61% | 2.049% | -19.82%
| Z-Score (Absolute) |  | 0.45 | 0.47 | 1.17 | 0.65 | 0.39 | 1.05 | 0.94 | 3.1 | 0.48 | 3.58 | 3.85 | 0.39 | 1.11 | 0.03 | 0.7 | 1.37 | 0.57 | 0.21 | 3.11 | 0.34 | 0.03 | 5.81 | 0.35 | 3.05
| Analytical results for V in OREAS 609c (Certified Value 3.98 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 4 | 4 | 4 | 4 | 6.5317 | 4 | 3.1 | 4.19 | 6 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3.6 | 4 | 3.1356 | 3 | 3.7888 | < 10 |  |  |  | 1
| 1 | 2 | 4 | 4 | 4 | 4 | 3.8793 | 4 | 3.2 | 4.19 | 6 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3.6 | 3.5 | 4 | 2.7419 | 6 | 4.0215 | < 10 |  |  |  | 34
| 1 | 3 | 4 | 3 | 4 | 4 | 3.8244 | 4 | 3.2 | 4.36 | 6 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3.6 | 4 | 2.9204 | 3 | 3.7271 | < 10 |  |  |  | 16
| 1 | 4 | 4 | 3 | 4 | 4 | 3.7985 | 4 | 3.2 | 4.3 | 5 | 4 | 4 | 4 | 4 | 4 | 3 | 4 | 4.4 | 3.9 | 4 | 2.9283 | 3 | 3.857 | < 10 |  |  |  | 3.9829
| 1 | 5 | 4 | 4 | 4 | 4 | 3.818 | 5 | 3.1 | 4.36 | 5 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4.1 | 3.9 | 4 | 2.7606 | 3 | 4.0305 | < 10 |  |  |  | 127
| 1 | 6 | 4 | 4 | 4 | 4 | 3.9793 | 5 | 3.2 | 4.33 | 6 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3.7 | 4 | 2.6354 | 3 | 3.747 | < 10
| Mean |  | 4 | 3.6667 | 4 | 4 | 4.3052 | 4.3333 | 3.1667 | 4.2883 | 5.6667 | 4 | 4 | 4 | 4 | 4 | 3.8333 | 4 | 4.0167 | 3.7 | 4 | 2.8537 | 3.5 | 3.862
| Median |  | 4 | 4 | 4 | 4 | 3.8518 | 4 | 3.2 | 4.315 | 6 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3.65 | 4 | 2.8405 | 3 | 3.8229
| Std Dev. |  | 0 | 0.5164 | 0 | 0 | 1.0927 | 0.5164 | 0.05164 | 0.079352 | 0.5164 | 0 | 0 | 0 | 0 | 0 | 0.40825 | 0 | 0.25626 | 0.16733 | 0 | 0.17793 | 1.2247 | 0.13463
| Rel.Std.Dev. |  | 0% | 14.08% | 0% | 0% | 25.38% | 11.92% | 1.631% | 1.85% | 9.113% | 0% | 0% | 0% | 0% | 0% | 10.65% | 0% | 6.38% | 4.522% | 0% | 6.235% | 34.99% | 3.486%
| PDM3 |  | 0.4285% | -7.94% | 0.4285% | 0.4285% | 8.091% | 8.798% | -20.49% | 7.668% | 42.27% | 0.4285% | 0.4285% | 0.4285% | 0.4285% | 0.4285% | -3.756% | 0.4285% | 0.847% | -7.104% | 0.4285% | -28.35% | -12.13% | -3.036%
| Z-Score (Absolute) |  | 0 | 1.63 | 0 | 0 | 1.49 | 1.63 | 4.07 | 1.41 | 8.14 | 0 | 0 | 0 | 0 | 0 | 0.81 | 0 | 0.08 | 1.47 | 0 | 5.6 | 2.44 | 0.67 | 4.89
| Analytical results for W in OREAS 609c (Certified Value 1.57 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  |  | 3
| 1 | 1 | 1.77 | 1.49 | 1.4 | 1.56 | 2.8052 | 1.6 | 1.83 | 1.5 | 1.4 | 1.1 | 1.64 | 1.73 | 1.5 | 1.63 | 1.59 | 2.17 | 1.03 | < 5 | 1.33 | < 0.1 | 0.2 | 1.4337 | 2 |  |  |  | 1
| 1 | 2 | 1.67 | 1.41 | 1.4 | 1.59 | 1.5352 | 1.7 | 1.71 | 1.53 | 1.5 | 1.13 | 1.64 | 1.68 | 1.5 | 1.62 | 1.54 | 2.19 | 0.92 | < 5 | 1.41 | < 0.1 | 1.1 | 1.4702 | 2 |  |  |  | 35
| 1 | 3 | 1.57 | 1.63 | 1.4 | 1.56 | 1.6136 | 1.6 | 1.72 | 1.77 | 1.4 | 1.11 | 1.6 | 1.66 | 1.5 | 1.61 | 1.53 | 2.18 | 0.95 | < 5 | 1.45 | < 0.1 | 0.4 | 1.5081 | 3 |  |  |  | 16
| 1 | 4 | 1.62 | 1.58 | 1.4 | 1.59 | 1.5845 | 1.5 | 1.8 | 1.6 | 1.4 | 1.12 | 1.54 | 1.76 | 1.43 | 1.53 | 1.56 | 2.29 | 1 | < 5 | 1.42 | < 0.1 | 0.3 | 1.47 | 3 |  |  |  | 1.574
| 1 | 5 | 1.58 | 1.5 | 1.4 | 1.58 | 1.5452 | 1.7 | 1.76 | 1.7 | 1.4 | 1.02 | 1.65 | 1.62 | 1.55 | 1.53 | 1.57 | 2.16 | 0.96 | < 5 | 1.32 | < 0.1 | 0.3 | 1.5135 | 2 |  |  |  | 128
| 1 | 6 | 1.61 | 1.82 | 1.5 | 1.59 | 1.535 | 1.6 | 1.71 | 1.57 | 1.4 | 1.18 | 1.59 | 1.68 | 1.43 | 1.6 | 1.5 | 2.2 | 1.01 | < 5 | 1.35 | < 0.1 | 0.3 | 1.4514 | 2
| Mean |  | 1.6367 | 1.5717 | 1.4167 | 1.5783 | 1.7698 | 1.6167 | 1.755 | 1.6117 | 1.4167 | 1.11 | 1.61 | 1.6883 | 1.485 | 1.5867 | 1.5483 | 2.1983 | 0.97833 |  | 1.38 |  | 0.43333 | 1.4745 | 2.3333
| Median |  | 1.615 | 1.54 | 1.4 | 1.585 | 1.5649 | 1.6 | 1.74 | 1.585 | 1.4 | 1.115 | 1.62 | 1.68 | 1.5 | 1.605 | 1.55 | 2.185 | 0.98 |  | 1.38 |  | 0.3 | 1.4701 | 2
| Std Dev. |  | 0.074207 | 0.14359 | 0.040825 | 0.01472 | 0.50822 | 0.075277 | 0.050892 | 0.10381 | 0.040825 | 0.052154 | 0.041952 | 0.049967 | 0.046797 | 0.045019 | 0.031885 | 0.047081 | 0.041673 |  | 0.053666 |  | 0.33267 | 0.031264 | 0.5164
| Rel.Std.Dev. |  | 4.534% | 9.136% | 2.882% | 0.9326% | 28.72% | 4.656% | 2.9% | 6.441% | 2.882% | 4.699% | 2.606% | 2.96% | 3.151% | 2.837% | 2.059% | 2.142% | 4.26% |  | 3.889% |  | 76.77% | 2.12% | 22.13%
| PDM3 |  | 3.98% | -0.1491% | -9.997% | 0.2745% | 12.44% | 2.71% | 11.5% | 2.392% | -9.997% | -29.48% | 2.286% | 7.263% | -5.655% | 0.8039% | -1.631% | 39.66% | -37.84% |  | -12.33% |  | -72.47% | -6.325% | 48.24%
| Z-Score (Absolute) |  | 0.34 | 0.07 | NA | 0.03 | 1.18 | NA | 1.09 | 0.18 | NA | 2.98 | 0.17 | 0.67 | 0.61 | 0.03 | 0.22 | 3.88 | 3.81 | 5.79 | 1.28 | 9.66 | NA | 0.68 | NA
Results from laboratories 2.04, 2.07, 2.11 and 2.25 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 609c (Certified Value 5.72 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 2
| 1 | 1 | 6.36 | 6.19 | 5.93 | 5.8946 | 5.68 | 4.92 | 3.1796 | 5.5 | 6.88 | 6.83 | 5.3 | 5.87 | 5.69 | 5.92 | 5.09 | 4.84 | 4.34 | 5.48 | 6.1751 | 3.73 | 5.7791 | < 10 | 4.49 |  |  |  | 1
| 1 | 2 | 6.63 | 5.72 | 6.03 | 6.0034 | 5.78 | 5.1 | 3.7941 | 5.5 | 7.78 | 7.02 | 5.12 | 6.06 | 5.73 | 6.01 | 5.35 | 4.66 | 4.41 | 5.84 | 6.8351 | 4.07 | 6.0505 | < 10 | 4.67 |  |  |  | 36
| 1 | 3 | 5.97 | 6.3 | 5.97 | 6.0681 | 5.92 | 5.06 | 3.3162 | 5.6 | 7.18 | 6.83 | 5.19 | 6.08 | 5.67 | 5.95 | 5.32 | 4.69 | 4.5 | 5.44 | 5.4605 | 3.69 | 5.742 | < 10 | 4.56 |  |  |  | 16
| 1 | 4 | 6.05 | 6.18 | 5.87 | 6.0813 | 5.79 | 5.07 | 3.6156 | 5.6 | 6.8 | 6.95 | 5.2 | 5.87 | 5.6 | 6.04 | 4.96 | 5.06 | 4.55 | 5.67 | 5.7564 | 3.72 | 5.9092 | < 10 | 4.61 |  |  |  | 5.7168
| 1 | 5 | 6.34 | 6.38 | 5.88 | 5.8972 | 5.9 | 4.85 | 3.0898 | 5.6 | 7.05 | 6.73 | 5.31 | 6.08 | 5.89 | 5.65 | 5.32 | 5.08 | 4.42 | 5.51 | 5.8279 | 3.64 | 5.9801 | < 10 | 4.63 |  |  |  | 129
| 1 | 6 | 6.16 | 6.4 | 5.76 | 6.078 | 5.97 | 5.04 | 3.6966 | 5.6 | 7.08 | 7.23 | 5.06 | 5.91 | 5.52 | 5.91 | 5.17 | 5.18 | 4.56 | 5.85 | 6.194 | 3.75 | 5.7534 | < 10 | 5.21
| Mean |  | 6.2517 | 6.195 | 5.9067 | 6.0038 | 5.84 | 5.0067 | 3.4487 | 5.5667 | 7.1283 | 6.9317 | 5.1967 | 5.9783 | 5.6833 | 5.9133 | 5.2017 | 4.9183 | 4.4633 | 5.6317 | 6.0415 | 3.7667 | 5.8691 |  | 4.695
| Median |  | 6.25 | 6.245 | 5.905 | 6.0357 | 5.845 | 5.05 | 3.4659 | 5.6 | 7.065 | 6.89 | 5.195 | 5.985 | 5.68 | 5.935 | 5.245 | 4.95 | 4.46 | 5.59 | 6.0015 | 3.725 | 5.8441 |  | 4.62
| Std Dev. |  | 0.24128 | 0.25026 | 0.093095 | 0.088244 | 0.10826 | 0.098725 | 0.29237 | 0.05164 | 0.34793 | 0.17804 | 0.098116 | 0.10534 | 0.12549 | 0.13866 | 0.15613 | 0.21894 | 0.087331 | 0.1828 | 0.47652 | 0.15345 | 0.12997 |  | 0.25983
| Rel.Std.Dev. |  | 3.859% | 4.04% | 1.576% | 1.47% | 1.854% | 1.972% | 8.478% | 0.9277% | 4.881% | 2.568% | 1.888% | 1.762% | 2.208% | 2.345% | 3.002% | 4.452% | 1.957% | 3.246% | 7.888% | 4.074% | 2.215% |  | 5.534%
| PDM3 |  | 9.355% | 8.364% | 3.32% | 5.019% | 2.154% | -12.42% | -39.68% | -2.627% | 24.69% | 21.25% | -9.099% | 4.574% | -0.5863% | 3.437% | -9.012% | -13.97% | -21.93% | -1.49% | 5.679% | -34.11% | 2.662% |  | -17.87%
| Z-Score (Absolute) |  | 0.79 | 0.71 | 0.31 | 0.44 | 0.22 | 0.94 | 3.1 | 0.16 | 2 | 1.73 | 0.67 | 0.41 | 0 | 0.32 | 0.67 | 1.06 | 1.69 | 0.07 | 0.5 | 2.66 | 0.26 | 0.95 | 1.37
| Analytical results for Yb in OREAS 609c (Certified Value 0.2 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.09: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.25g; 2.08: 01g; 2.09: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  |  | 2
| 1 | 1 | 2.01: 0.2; 2.03: 0.202; 2.05: 0.19226; 2.06: 0.2; 2.07: 0.1641; 2.08: 0.23; 2.09: 0.2; 2.10: 0.2356; 2.11: 0.15; 2.12: 0.19 |  |  |  | 1
| 1 | 2 | 2.01: 0.18; 2.03: 0.204; 2.05: 0.23555; 2.06: 0.2; 2.07: 0.1552; 2.08: 0.24; 2.09: 0.2; 2.10: 0.2613; 2.11: 0.16; 2.12: 0.17 |  |  |  | 37
| 1 | 3 | 2.01: 0.21; 2.03: 0.194; 2.05: 0.2238; 2.06: 0.2; 2.07: 0.1756; 2.08: 0.23; 2.09: 0.2; 2.10: 0.2115; 2.11: 0.15; 2.12: 0.19 |  |  |  | 16
| 1 | 4 | 2.01: 0.19; 2.03: 0.205; 2.05: 0.22903; 2.06: 0.2; 2.07: 0.1774; 2.08: 0.22; 2.09: 0.2; 2.10: 0.2117; 2.11: 0.15; 2.12: 0.21 |  |  |  | 0.19936
| 1 | 5 | 2.01: 0.22; 2.03: 0.199; 2.05: 0.21922; 2.06: 0.2; 2.07: 0.1661; 2.08: 0.22; 2.09: 0.2; 2.10: 0.2206; 2.11: 0.15; 2.12: 0.18 |  |  |  | 130
| 1 | 6 | 2.01: 0.2; 2.03: 0.215; 2.05: 0.21464; 2.06: 0.2; 2.07: 0.1591; 2.08: 0.24; 2.09: 0.2; 2.10: 0.2577; 2.11: 0.15; 2.12: 0.21
| Mean |  | 2.01: 0.2; 2.03: 0.20317; 2.05: 0.21908; 2.06: 0.2; 2.07: 0.16625; 2.08: 0.23; 2.09: 0.2; 2.10: 0.23307; 2.11: 0.15167; 2.12: 0.19167
| Median |  | 2.01: 0.2; 2.03: 0.203; 2.05: 0.22151; 2.06: 0.2; 2.07: 0.1651; 2.08: 0.23; 2.09: 0.2; 2.10: 0.2281; 2.11: 0.15; 2.12: 0.19
| Std Dev. |  | 2.01: 0.014142; 2.03: 0.0070261; 2.05: 0.015042; 2.06: 3.0405e-17; 2.07: 0.0088274; 2.08: 0.0089443; 2.09: 3.0405e-17; 2.10: 0.022304; 2.11: 0.0040825; 2.12: 0.016021
| Rel.Std.Dev. |  | 2.01: 7.071%; 2.03: 3.458%; 2.05: 6.866%; 2.06: 1.52e-14%; 2.07: 5.31%; 2.08: 3.889%; 2.09: 1.52e-14%; 2.10: 9.57%; 2.11: 2.692%; 2.12: 8.359%
| PDM3 |  | 2.01: 0.3199%; 2.03: 1.908%; 2.05: 9.891%; 2.06: 0.3199%; 2.07: -16.61%; 2.08: 15.37%; 2.09: 0.3199%; 2.10: 16.91%; 2.11: -23.92%; 2.12: -3.86%
| Z-Score (Absolute) |  | 2.01: 0.05; 2.03: 0.05; 2.05: 0.51; 2.06: NA; 2.07: 1.04; 2.08: 0.83; 2.09: NA; 2.10: 0.92; 2.11: 1.47; 2.12: 0.29
Results from laboratories 2.07 and 2.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 609c (Certified Value 1022 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 1042 | 989 | 1017 | 1039 | 2762.1717 | 1010 | 1060 | 1006.1285 | 1023 | 1070 | 1006 | 987 | 1040 | 965 | 1055 | 999 | 1106 | 880 | 1240 | 999.3 | 975 | 896.5926 | 1010 | 1070.9599 | 1006.548 | 843 | 1
| 1 | 2 | 1097 | 959.2 | 1013 | 1046 | 1035.865 | 1020 | 1050 | 1006.882 | 1015 | 1020 | 1034 | 979 | 1040 | 968 | 1065 | 1005 | 1103 | 870 | 1180 | 999.9 | 990 | 910.737 | 2840 | 1069.8553 | 1006.307 | 867 | 18
| 1 | 3 | 1073 | 966.1 | 1011 | 1037 | 1044.5703 | 1030 | 1040 | 1001.4245 | 1019 | 1020 | 992 | 974 | 1035 | 956 | 1035 | 1000 | 1097 | 877 | 1230 | 1004 | 977 | 925.6893 | 1010 | 1054.4872 | 1005.639 | 879 | 16
| 1 | 4 | 1104 | 947.5 | 1013 | 1034 | 1024.5461 | 1020 | 1040 | 1008.9805 | 1020 | 1040 | 1024 | 973 | 1030 | 955 | 1045 | 972 | 1100 | 861 | 1240 | 1006 | 1001 | 906.2902 | 1010 | 1046.8534 | 1002.967 | 906 | 1022.056
| 1 | 5 | 1050 | 991.6 | 1013 | 1034 | 1021.3164 | 996 | 1050 | 1004.165 | 1016 | 1080 | 982 | 998 | 1045 | 983 | 1030 | 1000 | 1128 | 851 | 1250 | 999.2 | 976 | 915.6807 | 1030 | 1065.7084 | 1000.147 | 974 | 131
| 1 | 6 | 1040 | 1008.4 | 1011 | 1045 | 1032.3601 | 1040 | 1060 | 1006.1325 | 1028 | 1040 | 1056 | 982 | 1040 | 950 | 1035 | 993 | 1113 | 866 | 1260 | 999.8 | 1007 | 905.952 | 1030 | 1056.5756 | 996.236 | 901
| Mean |  | 1067.6667 | 976.9667 | 1013 | 1039.1667 | 1320.1383 | 1019.3333 | 1050 | 1005.6188 | 1020.1667 | 1045 | 1015.6667 | 982.1667 | 1038.3333 | 962.8333 | 1044.1667 | 994.8333 | 1107.8333 | 867.5 | 1233.3333 | 1001.3667 | 987.6667 | 910.157 | 1321.6667 | 1060.74 | 1002.974 | 895
| Median |  | 1061.5 | 977.55 | 1013 | 1038 | 1034.1125 | 1020 | 1050 | 1006.1305 | 1019.5 | 1040 | 1015 | 980.5 | 1040 | 960.5 | 1040 | 999.5 | 1104.5 | 868 | 1240 | 999.85 | 983.5 | 908.5136 | 1020 | 1061.142 | 1004.303 | 890
| Std Dev. |  | 28.0903 | 23.0174 | 2.1909 | 5.2694 | 706.4977 | 15.3188 | 8.9443 | 2.5739 | 4.7924 | 25.0998 | 27.6381 | 9.3256 | 5.164 | 11.9234 | 13.5708 | 11.8223 | 11.3034 | 10.6724 | 28.0476 | 2.8974 | 13.9092 | 9.8866 | 743.8929 | 9.6073 | 4.108 | 45.0378
| Rel.Std.Dev. |  | 2.631% | 2.356% | 0.2163% | 0.5071% | 53.52% | 1.503% | 0.8518% | 0.256% | 0.4698% | 2.402% | 2.721% | 0.9495% | 0.4973% | 1.238% | 1.3% | 1.188% | 1.02% | 1.23% | 2.274% | 0.2893% | 1.408% | 1.086% | 56.28% | 0.9057% | 0.4096% | 5.032%
| PDM3 |  | 4.463% | -4.412% | -0.8861% | 1.674% | 29.16% | -0.2664% | 2.734% | -1.608% | -0.1849% | 2.245% | -0.6251% | -3.903% | 1.593% | -5.794% | 2.163% | -2.664% | 8.393% | -15.12% | 20.67% | -2.024% | -3.365% | -10.95% | 29.31% | 3.785% | -1.867% | -12.43%
| Z-Score (Absolute) |  | 1.09 | 0.88 | 0.1 | 0.47 | 6.55 | 0.04 | 0.7 | 0.26 | 0.06 | 0.59 | 0.04 | 0.76 | 0.45 | 1.18 | 0.58 | 0.49 | 1.95 | 3.25 | 4.67 | 0.35 | 0.65 | 2.32 | 6.58 | 0.94 | 0.31 | 2.65
| Analytical results for Zr in OREAS 609c (Certified Value 51 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 0
| 1 | 1 | 51 | 54.3 | 52.2 | 31.7643 | 32.9 | 41.9 | 61.412 | 47.8 | 59.3 | 61.6 | 50.4 | 51.5 | 50.7 | 57.5 | 47.7 | 39.8 | 40.6 | 70.5 | 30.6 | 66.357 | 23.4 | 43.7734 | 55.07 | 46.2 |  |  | 1
| 1 | 2 | 53 | 52.6 | 52.1 | 51.165 | 29.6 | 43.7 | 61.412 | 49.4 | 61.2 | 62.5 | 47.8 | 52.2 | 50.5 | 56.9 | 50.2 | 38.1 | 37.8 | 71.3 | 28.5 | 56.3346 | 20.5 | 48.3034 | 54.82 | 45.3 |  |  | 19
| 1 | 3 | 51 | 52 | 51.6 | 51.9376 | 25.8 | 43.7 | 60.824 | 49.1 | 60.4 | 61.1 | 49.4 | 51.8 | 49.8 | 56.8 | 49.5 | 37.3 | 38.3 | 71.6 | 31 | 60.7907 | 23.6 | 49.106 | 55.74 | 44.4 |  |  | 16
| 1 | 4 | 51 | 49.7 | 50.9 | 51.6776 | 28.1 | 43.7 | 63.764 | 49.2 | 59 | 62.9 | 50.3 | 49 | 48.6 | 58.2 | 46.5 | 42.8 | 41.1 | 71.5 | 34.4 | 60.8303 | 21.3 | 46.083 | 54.21 | 44.8 |  |  | 51.3432
| 1 | 5 | 50 | 53.8 | 51.2 | 50.1409 | 28.4 | 40.9 | 64.2995 | 49.8 | 59.3 | 60 | 49.5 | 53.3 | 52 | 53.8 | 50 | 44.6 | 39.1 | 70.8 | 32.9 | 64.7597 | 21.4 | 49.796 | 54.03 | 44.3 |  |  | 132
| 1 | 6 | 50 | 55.9 | 51.1 | 50.1027 | 21.6 | 43.3 | 63.1865 | 48.3 | 60.5 | 65 | 46.5 | 51.4 | 48.3 | 57 | 48.4 | 43.9 | 40 | 70.7 | 29.2 | 52.4808 | 23.4 | 44.1267 | 54.42 | 45.6
| Mean |  | 51 | 53.05 | 51.5167 | 47.798 | 27.7333 | 42.8667 | 62.483 | 48.9333 | 59.95 | 62.1833 | 48.9833 | 51.5333 | 49.9833 | 56.7 | 48.7167 | 41.0833 | 39.4833 | 71.0667 | 31.1 | 60.2589 | 22.2667 | 46.8648 | 54.715 | 45.1
| Median |  | 51 | 53.2 | 51.4 | 50.653 | 28.25 | 43.5 | 62.2993 | 49.15 | 59.85 | 62.05 | 49.45 | 51.65 | 50.15 | 56.95 | 48.95 | 41.3 | 39.55 | 71.05 | 30.8 | 60.8105 | 22.4 | 47.1932 | 54.62 | 45.05
| Std Dev. |  | 1.0954 | 2.1342 | 0.54191 | 7.8918 | 3.7977 | 1.1894 | 1.4479 | 0.74207 | 0.8735 | 1.7221 | 1.5329 | 1.4194 | 1.3877 | 1.5126 | 1.4497 | 3.1019 | 1.3045 | 0.45898 | 2.2253 | 5.176 | 1.353 | 2.5826 | 0.63229 | 0.73756
| Rel.Std.Dev. |  | 2.148% | 4.023% | 1.052% | 16.51% | 13.69% | 2.775% | 2.317% | 1.516% | 1.457% | 2.769% | 3.129% | 2.754% | 2.776% | 2.668% | 2.976% | 7.55% | 3.304% | 0.6459% | 7.155% | 8.59% | 6.076% | 5.511% | 1.156% | 1.635%
| PDM3 |  | -0.6684% | 3.324% | 0.3379% | -6.905% | -45.98% | -16.51% | 21.7% | -4.694% | 16.76% | 21.11% | -4.596% | 0.3704% | -2.649% | 10.43% | -5.116% | -19.98% | -23.1% | 38.42% | -39.43% | 17.36% | -56.63% | -8.723% | 6.567% | -12.16%
| Z-Score (Absolute) |  | 0.17 | 0.41 | 0.23 | 0.19 | 2.48 | 0.75 | 1.48 | 0.06 | 1.19 | 1.45 | 0.06 | 0.23 | 0.06 | 0.82 | 0.09 | 0.96 | 1.14 | 2.46 | 2.09 | 1.23 | 3.1 | 0.3 | 0.6 | 0.5
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C965: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1055: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ALK
| Analytical results for S-(Sulphate) in OREAS 609c (Indicative Value 0.54 wt.%) | [18 empty] | Indicative
S-(Sulphate) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.25 | 2.26 |  | wt.%
|  |  | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | Na2CO3*IRC | IRC-x*Calc | IRC-x*Calc | Na2CO3*GRAV | IRC-x*Calc | IRC-x*Calc | Na2CO3*GRAV | I-550*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC | Na2CO3*GRAV |  | 3
3
| 1 | 1 | 0.22 | 0.54 | 0.61746 | 0.41425 | 0.80437 | 0.81 | 0.42 | 0.39 | 0.68 | 0.51 | 0.29 | 0.33376 | 0.88 | 0.87 | 0.45 | 0.497 |  | 1
| 1 | 2 | 0.22 | 0.49 | 0.61079 | 0.44025 | 0.79102 | 0.82 | 0.43 | 0.32 | 0.63 | 0.47 | 0.29 | 0.35045 | 0.88 | 0.93 | 0.27 | 0.49 |  | 8
| 1 | 3 | 0.22 | 0.54 | 0.62414 | 0.43715 | 0.7977 | 0.79 | 0.43 | 0.39 | 0.66 | 0.48 | 0.23 | 0.35713 | 0.9 | 0.9 | 0.46 | 0.483 |  | 16
| 1 | 4 | 0.22 | 0.54 | 0.65084 | 0.4416 | 0.80437 | 0.8 | 0.42 | 0.27 | 0.53 | 0.48 | 0.3 | 0.46059 | 0.84 | 0.9 | 0.45 | 0.496 |  | 0.54033
| 1 | 5 | 0.22 | 0.55 | 0.63081 | 0.4275 | 0.79436 | 0.84 | 0.42 | 0.34 | 0.71 | 0.49 | 0.21 | 0.3371 | 0.86 | 0.88 | 0.45 | 0.482 |  | 14
| 1 | 6 | 0.22 | 0.54 | 0.64083 | 0.4507 | 0.81438 | 0.81 | 0.44 | 0.35 | 0.57 | 0.47 | 0.34 | 0.37048 | 0.84 | 0.86 | 0.44 | 0.49
| Mean |  | 0.22 | 0.53333 | 0.62914 | 0.43524 | 0.80103 | 0.81167 | 0.42667 | 0.34333 | 0.63 | 0.48333 | 0.27667 | 0.36825 | 0.86667 | 0.89 | 0.42 | 0.48967
| Median |  | 0.22 | 0.54 | 0.62748 | 0.4387 | 0.80103 | 0.81 | 0.425 | 0.345 | 0.645 | 0.48 | 0.29 | 0.35379 | 0.87 | 0.89 | 0.45 | 0.49
| Std Dev. |  | 0 | 0.021602 | 0.014889 | 0.012718 | 0.0084436 | 0.017224 | 0.008165 | 0.045461 | 0.068411 | 0.015055 | 0.048028 | 0.047186 | 0.024221 | 0.025298 | 0.073756 | 0.0062823
| Rel.Std.Dev. |  | 0% | 4.05% | 2.367% | 2.922% | 1.054% | 2.122% | 1.914% | 13.24% | 10.86% | 3.115% | 17.36% | 12.81% | 2.795% | 2.842% | 17.56% | 1.283%
| PDM3 |  | -59.28% | -1.294% | 16.44% | -19.45% | 48.25% | 50.22% | -21.04% | -36.46% | 16.6% | -10.55% | -48.8% | -31.85% | 60.4% | 64.72% | -22.27% | -9.376%
| Z-Score (Absolute) |  | 1.26 | 0.22 | 0.67 | 0.24 | 1.48 | 1.53 | 0.28 | 0.68 | 0.68 | 0.01 | 0.99 | 0.56 | 1.79 | 1.9 | 0.31 | 0.01
| Analytical results for S-(Sulphide) in OREAS 609c (Certified Value 1.03 wt.%) | [18 empty] | Certified
S-(Sulphide) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.25 | 2.26 | wt.%
|  |  | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | IRC | Na2CO3*IRC | IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | HCl*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC | HCl*IRC | 2
2
| 1 | 1 | 1.4 | 1.15 | 0.97 | 1.3 | 1.2421 | 0.78 | 0.98 | 1.07 | 1.23 | 0.9 | 0.91 | 1.34 | 0.71 | 0.75 | 0.69 | 1.17 | 0.9197 | 1
| 1 | 2 | 1.4 | 1.21 | 0.97 | 1.26 | 1.2404 | 0.78 | 0.97 | 1.06 | 1.3 | 0.83 | 0.98 | 1.38 | 0.66 | 0.77 | 0.64 | 3.42 | 0.89137 | 0
| 1 | 3 | 1.4 | 1.14 | 1.01 | 1.26 | 1.2486 | 0.84 | 0.99 | 1.1 | 1.24 | 0.81 | 0.98 | 1.4 | 0.63 | 0.75 | 0.68 | 1.13 | 0.94887 | 16
| 1 | 4 | 1.41 | 1.14 | 0.95 | 1.25 | 1.2239 | 0.79 | 0.98 | 1.1 | 1.3 | 1.02 | 1.02 | 1.35 | 0.5 | 0.81 | 0.65 | 1.13 | 0.92473 | 1.0336
| 1 | 5 | 1.41 | 1.13 | 0.96 | 1.28 | 1.2145 | 0.84 | 0.94 | 1.09 | 1.29 | 0.89 | 0.91 | 1.43 | 0.66 | 0.81 | 0.67 | 1.17 | 0.88439 | 134
| 1 | 6 | 1.41 | 1.15 | 0.92 | 1.32 | 1.2646 | 0.76 | 0.98 | 1.07 | 1.27 | 0.99 | 0.97 | 1.38 | 0.63 | 0.75 | 0.68 | 1.14 | 0.92806
| Mean |  | 1.405 | 1.1533 | 0.96333 | 1.2783 | 1.239 | 0.79833 | 0.97333 | 1.0817 | 1.2717 | 0.90667 | 0.96167 | 1.38 | 0.63167 | 0.77333 | 0.66833 | 1.5267 | 0.91619
| Median |  | 1.405 | 1.145 | 0.965 | 1.27 | 1.2412 | 0.785 | 0.98 | 1.08 | 1.28 | 0.895 | 0.975 | 1.38 | 0.645 | 0.76 | 0.675 | 1.155 | 0.92222
| Std Dev. |  | 0.0054772 | 0.028752 | 0.029439 | 0.027142 | 0.017812 | 0.033714 | 0.017512 | 0.017224 | 0.030605 | 0.084063 | 0.043551 | 0.032863 | 0.070828 | 0.029439 | 0.019408 | 0.92772 | 0.02417
| Rel.Std.Dev. |  | 0.3898% | 2.493% | 3.056% | 2.123% | 1.438% | 4.223% | 1.799% | 1.592% | 2.407% | 9.272% | 4.529% | 2.381% | 11.21% | 3.807% | 2.904% | 60.77% | 2.638%
| PDM3 |  | 35.93% | 11.58% | -6.8% | 23.68% | 19.87% | -22.76% | -5.832% | 4.649% | 23.03% | -12.28% | -6.961% | 33.51% | -38.89% | -25.18% | -35.34% | 47.7% | -11.36%
| Z-Score (Absolute) |  | 1.46 | 0.61 | 0.03 | 1.03 | 0.9 | 0.59 | 0 | 0.37 | 1.01 | 0.22 | 0.04 | 1.37 | 1.15 | 0.67 | 1.03 | 1.87 | 0.19
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Acid Leach
| Analytical results for S-(Sulphate) in OREAS 609c (Indicative Value 0.303 wt.%) |  |  | Indicative
S-(Sulphate) | Round | Replicate | Lab | 1
| No | No | 2.05 | wt.%
|  |  | HCl*OES | 3
3
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.32 | 8
| 1 | 3 | 0.3 | 16
| 1 | 4 | 0.31 | 0.30333
| 1 | 5 | 0.3 | 16
| 1 | 6 | 0.3
| Mean |  | 0.30333
| Median |  | 0.3
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 3.405%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 609c (Indicative Value 13.87 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.87 | 1
| 1 | 2 | 13.86 | 12
| Mean |  | 13.865 | 16
| Median |  | 13.865 | 13.865
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.051%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 609c (Indicative Value 820 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 820 | 1
| 1 | 2 | 820 | 13
| Mean |  | 820 | 16
| Median |  | 820 | 820
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 609c (Indicative Value 3283 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 3305 | 1
| 1 | 2 | 3260 | 14
| Mean |  | 3282.5 | 16
| Median |  | 3282.5 | 3282.51
| Std Dev. |  | 31.8198 | 20
| Rel.Std.Dev. |  | 0.9694%
| PDM3 |  | -0.0003046%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 609c (Indicative Value 1.52 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.52 | 1
| 1 | 2 | 1.51 | 15
| Mean |  | 1.515 | 16
| Median |  | 1.515 | 1.515
| Std Dev. |  | 0.0070711 | 21
| Rel.Std.Dev. |  | 0.4667%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 609c (Indicative Value 30 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 40 | 16
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 14.1421 | 22
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 609c (Indicative Value < 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 17
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 23
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 609c (Indicative Value 25.6 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 15 | 1
| 1 | 2 | 44 | 18
| Mean |  | 44 | 16
| Median |  | 44 | 25.5763
| Std Dev. |  |  | 24
Rel.Std.Dev.
| PDM3 |  | 72.03%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 609c (Indicative Value 0.48 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.479 | 1
| 1 | 2 | 0.48 | 19
| Mean |  | 0.4795 | 16
| Median |  | 0.4795 | 0.4795
| Std Dev. |  | 0.00070711 | 25
| Rel.Std.Dev. |  | 0.1475%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 609c (Indicative Value 4.32 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.303 | 1
| 1 | 2 | 4.332 | 56
| Mean |  | 4.3175 | 16
| Median |  | 4.3175 | 4.3177
| Std Dev. |  | 0.020506 | 18
| Rel.Std.Dev. |  | 0.475%
| PDM3 |  | -0.004493%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 609c (Indicative Value 3.34 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.35 | 1
| 1 | 2 | 3.33 | 13
| Mean |  | 3.34 | 16
| Median |  | 3.34 | 3.34
| Std Dev. |  | 0.014142 | 19
| Rel.Std.Dev. |  | 0.4234%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 609c (Indicative Value 0.24 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.24 | 1
| 1 | 2 | 0.24 | 14
| Mean |  | 0.24 | 16
| Median |  | 0.24 | 0.24
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 609c (Indicative Value 0.03 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.029 | 1
| 1 | 2 | 0.031 | 15
| Mean |  | 0.03 | 16
| Median |  | 0.03 | 0.03
| Std Dev. |  | 0.0014142 | 21
| Rel.Std.Dev. |  | 4.714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 609c (Indicative Value 3.03 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.02 | 1
| 1 | 2 | 3.03 | 16
| Mean |  | 3.025 | 16
| Median |  | 3.025 | 3.025
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 0.2338%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 609c (Indicative Value 10 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 17
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 609c (Indicative Value 0.063 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0642 | 1
| 1 | 2 | 0.0619 | 18
| Mean |  | 0.06305 | 16
| Median |  | 0.06305 | 0.063019
| Std Dev. |  | 0.0016263 | 24
| Rel.Std.Dev. |  | 2.579%
| PDM3 |  | 0.04919%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 609c (Indicative Value 380 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 380 | 1
| 1 | 2 | 380 | 19
| Mean |  | 380 | 16
| Median |  | 380 | 380
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 609c (Indicative Value 1.6 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.61 | 1
| 1 | 2 | 1.59 | 12
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0.014142 | 18
| Rel.Std.Dev. |  | 0.8839%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 609c (Indicative Value 67.71 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.69 | 1
| 1 | 2 | 67.72 | 13
| Mean |  | 67.705 | 16
| Median |  | 67.705 | 67.705
| Std Dev. |  | 0.021213 | 19
| Rel.Std.Dev. |  | 0.03133%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 609c (Indicative Value 40 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 30 | 14
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 14.1421 | 20
| Rel.Std.Dev. |  | 35.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 609c (Indicative Value 200 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 200 | 1
| 1 | 2 | 200 | 15
| Mean |  | 200 | 16
| Median |  | 200 | 200
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 609c (Indicative Value 0.207 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.208 | 1
| 1 | 2 | 0.205 | 16
| Mean |  | 0.2065 | 16
| Median |  | 0.2065 | 0.2065
| Std Dev. |  | 0.0021213 | 22
| Rel.Std.Dev. |  | 1.027%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 609c (Indicative Value 17.9 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 18 | 17
| Mean |  | 18 | 16
| Median |  | 18 | 17.852
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.829%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 609c (Indicative Value 1060 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1060 | 1
| 1 | 2 | 1060 | 18
| Mean |  | 1060 | 16
| Median |  | 1060 | 1060
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 609c (Indicative Value 245 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 240 | 1
| 1 | 2 | 250 | 19
| Mean |  | 245 | 16
| Median |  | 245 | 245
| Std Dev. |  | 7.0711 | 25
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 609c (Indicative Value 3.38 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.37 | 1
| 1 | 2 | 3.39 | 21
| Mean |  | 3.38 | 16
| Median |  | 3.38 | 3.38
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 0.4184%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 609c (Indicative Value 26.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 26.4 | 23
| Mean |  | 26.2 | 16
| Median |  | 26.2 | 26.2
| Std Dev. |  | 0.28284 | 29
| Rel.Std.Dev. |  | 1.08%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 609c (Indicative Value 877 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 875 | 1
| 1 | 2 | 878 | 13
| Mean |  | 876.5 | 16
| Median |  | 876.5 | 876.5
| Std Dev. |  | 2.1213 | 30
| Rel.Std.Dev. |  | 0.242%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 609c (Indicative Value 2870 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2850 | 1
| 1 | 2 | 2890 | 5
| Mean |  | 2870 | 16
| Median |  | 2870 | 2870
| Std Dev. |  | 28.2843 | 31
| Rel.Std.Dev. |  | 0.9855%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 609c (Indicative Value 3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3.2 | 26
| Mean |  | 3 | 16
| Median |  | 3 | 3
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 609c (Indicative Value 68 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 68.2 | 1
| 1 | 2 | 67.3 | 27
| Mean |  | 67.75 | 16
| Median |  | 67.75 | 67.75
| Std Dev. |  | 0.6364 | 33
| Rel.Std.Dev. |  | 0.9393%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 609c (Indicative Value 6.2 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.3 | 1
| 1 | 2 | 6.1 | 28
| Mean |  | 6.2 | 16
| Median |  | 6.2 | 6.2
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 2.281%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 609c (Indicative Value 82 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 82.6 | 1
| 1 | 2 | 80.9 | 29
| Mean |  | 81.75 | 16
| Median |  | 81.75 | 81.75
| Std Dev. |  | 1.2021 | 35
| Rel.Std.Dev. |  | 1.47%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 609c (Indicative Value 7.75 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8 | 1
| 1 | 2 | 7.5 | 17
| Mean |  | 7.75 | 16
| Median |  | 7.75 | 7.75
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 609c (Indicative Value 19.5 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 19 | 31
| Mean |  | 19.5 | 16
| Median |  | 19.5 | 19.5
| Std Dev. |  | 0.70711 | 37
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 609c (Indicative Value 5.78 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.84 | 1
| 1 | 2 | 5.71 | 32
| Mean |  | 5.775 | 16
| Median |  | 5.775 | 5.775
| Std Dev. |  | 0.091924 | 38
| Rel.Std.Dev. |  | 1.592%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 609c (Indicative Value 0.455 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.444 | 1
| 1 | 2 | 0.466 | 19
| Mean |  | 0.455 | 16
| Median |  | 0.455 | 0.455
| Std Dev. |  | 0.015556 | 39
| Rel.Std.Dev. |  | 3.419%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 609c (Indicative Value 3.2 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.2 | 1
| 1 | 2 | 3.19 | 34
| Mean |  | 3.195 | 16
| Median |  | 3.195 | 3.195
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.2213%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 609c (Indicative Value 1.01 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 1 | 35
| Mean |  | 1.01 | 16
| Median |  | 1.01 | 1.01
| Std Dev. |  | 0.014142 | 41
| Rel.Std.Dev. |  | 1.4%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 609c (Indicative Value 1.14 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.13 | 1
| 1 | 2 | 1.14 | 36
| Mean |  | 1.135 | 16
| Median |  | 1.135 | 1.135
| Std Dev. |  | 0.0070711 | 42
| Rel.Std.Dev. |  | 0.623%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 609c (Indicative Value 21.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.4 | 1
| 1 | 2 | 22.2 | 37
| Mean |  | 21.8 | 16
| Median |  | 21.8 | 21.8
| Std Dev. |  | 0.56569 | 43
| Rel.Std.Dev. |  | 2.595%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 609c (Indicative Value 5.09 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.99 | 1
| 1 | 2 | 5.19 | 38
| Mean |  | 5.09 | 16
| Median |  | 5.09 | 5.09
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 2.778%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 609c (Indicative Value 2.18 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.15 | 39
| Mean |  | 2.175 | 16
| Median |  | 2.175 | 2.175
| Std Dev. |  | 0.035355 | 45
| Rel.Std.Dev. |  | 1.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 609c (Indicative Value 7.03 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.04 | 1
| 1 | 2 | 7.01 | 23
| Mean |  | 7.025 | 16
| Median |  | 7.025 | 7.025
| Std Dev. |  | 0.021213 | 29
| Rel.Std.Dev. |  | 0.302%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 609c (Indicative Value 0.46 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.45 | 1
| 1 | 2 | 0.46 | 24
| Mean |  | 0.455 | 16
| Median |  | 0.455 | 0.455
| Std Dev. |  | 0.0070711 | 30
| Rel.Std.Dev. |  | 1.554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 609c (Indicative Value 1.18 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.15 | 25
| Mean |  | 1.175 | 16
| Median |  | 1.175 | 1.175
| Std Dev. |  | 0.035355 | 31
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 609c (Indicative Value 42.4 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 42.7 | 1
| 1 | 2 | 42.1 | 26
| Mean |  | 42.4 | 16
| Median |  | 42.4 | 42.4
| Std Dev. |  | 0.42426 | 32
| Rel.Std.Dev. |  | 1.001%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 609c (Indicative Value 0.08 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.08 | 27
| Mean |  | 0.08 | 16
| Median |  | 0.08 | 0.08
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 609c (Indicative Value 0.027 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0254 | 1
| 1 | 2 | 0.0276 | 28
| Mean |  | 0.0265 | 16
| Median |  | 0.0265 | 0.0265
| Std Dev. |  | 0.0015556 | 34
| Rel.Std.Dev. |  | 5.87%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 609c (Indicative Value 40.1 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.2 | 1
| 1 | 2 | 40 | 29
| Mean |  | 40.1 | 16
| Median |  | 40.1 | 40.1
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 0.3527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 609c (Indicative Value 16.9 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.5 | 1
| 1 | 2 | 17.2 | 30
| Mean |  | 16.85 | 16
| Median |  | 16.85 | 16.85
| Std Dev. |  | 0.49497 | 36
| Rel.Std.Dev. |  | 2.938%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 609c (Indicative Value 35.9 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.2 | 1
| 1 | 2 | 35.6 | 31
| Mean |  | 35.9 | 16
| Median |  | 35.9 | 35.9
| Std Dev. |  | 0.42426 | 37
| Rel.Std.Dev. |  | 1.182%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 609c (Indicative Value 17 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 16 | 17
| Mean |  | 17 | 16
| Median |  | 17 | 17
| Std Dev. |  | 1.4142 | 38
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 609c (Indicative Value 375 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 379 | 1
| 1 | 2 | 371 | 19
| Mean |  | 375 | 16
| Median |  | 375 | 375
| Std Dev. |  | 5.6569 | 39
| Rel.Std.Dev. |  | 1.508%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 609c (Indicative Value 9.85 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.96 | 1
| 1 | 2 | 9.74 | 34
| Mean |  | 9.85 | 16
| Median |  | 9.85 | 9.85
| Std Dev. |  | 0.15556 | 40
| Rel.Std.Dev. |  | 1.579%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 609c (Indicative Value 122 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 121 | 1
| 1 | 2 | 122 | 35
| Mean |  | 121.5 | 16
| Median |  | 121.5 | 121.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 0.582%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 609c (Indicative Value 0.045 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.04 | 36
| Mean |  | 0.045 | 16
| Median |  | 0.045 | 0.045
| Std Dev. |  | 0.0070711 | 42
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 609c (Indicative Value 164 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 163 | 37
| Mean |  | 163.5 | 16
| Median |  | 163.5 | 163.5
| Std Dev. |  | 0.70711 | 43
| Rel.Std.Dev. |  | 0.4325%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 609c (Indicative Value 3.9 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 4 | 38
| Mean |  | 3.9 | 16
| Median |  | 3.9 | 3.9
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 609c (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 39
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 45
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 609c (Indicative Value 6.6 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.67 | 1
| 1 | 2 | 6.53 | 23
| Mean |  | 6.6 | 16
| Median |  | 6.6 | 6.6
| Std Dev. |  | 0.098995 | 29
| Rel.Std.Dev. |  | 1.5%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 609c (Indicative Value 9.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.8 | 1
| 1 | 2 | 9.8 | 14
| Mean |  | 9.8 | 16
| Median |  | 9.8 | 9.8
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 609c (Indicative Value 199 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 198 | 1
| 1 | 2 | 200 | 15
| Mean |  | 199 | 16
| Median |  | 199 | 199
| Std Dev. |  | 1.4142 | 31
| Rel.Std.Dev. |  | 0.7107%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 609c (Indicative Value 1.25 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.26 | 1
| 1 | 2 | 1.24 | 26
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 609c (Indicative Value 0.7 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.69 | 1
| 1 | 2 | 0.71 | 27
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0.014142 | 33
| Rel.Std.Dev. |  | 2.02%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 609c (Indicative Value 13.5 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.4 | 1
| 1 | 2 | 13.6 | 28
| Mean |  | 13.5 | 16
| Median |  | 13.5 | 13.5
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 1.048%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 609c (Indicative Value 13.8 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.9 | 1
| 1 | 2 | 13.7 | 29
| Mean |  | 13.8 | 16
| Median |  | 13.8 | 13.8
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 1.025%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 609c (Indicative Value 0.123 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.123 | 1
| 1 | 2 | 0.123 | 30
| Mean |  | 0.123 | 16
| Median |  | 0.123 | 0.123
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 609c (Indicative Value 1.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 31
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 609c (Indicative Value 0.11 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.11 | 6
| Mean |  | 0.11 | 16
| Median |  | 0.11 | 0.11
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 609c (Indicative Value 5.07 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.16 | 1
| 1 | 2 | 4.98 | 33
| Mean |  | 5.07 | 16
| Median |  | 5.07 | 5.07
| Std Dev. |  | 0.12728 | 39
| Rel.Std.Dev. |  | 2.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 609c (Indicative Value 9.65 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.7 | 1
| 1 | 2 | 9.6 | 34
| Mean |  | 9.65 | 16
| Median |  | 9.65 | 9.65
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 609c (Indicative Value 3.25 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3 | 35
| Mean |  | 3.25 | 16
| Median |  | 3.25 | 3.25
| Std Dev. |  | 0.35355 | 41
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 609c (Indicative Value 14 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.9 | 1
| 1 | 2 | 14.1 | 36
| Mean |  | 14 | 16
| Median |  | 14 | 14
| Std Dev. |  | 0.14142 | 42
| Rel.Std.Dev. |  | 1.01%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 609c (Indicative Value 0.55 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.54 | 1
| 1 | 2 | 0.56 | 37
| Mean |  | 0.55 | 16
| Median |  | 0.55 | 0.55
| Std Dev. |  | 0.014142 | 43
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 609c (Indicative Value 1125 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1130 | 1
| 1 | 2 | 1120 | 18
| Mean |  | 1125 | 16
| Median |  | 1125 | 1125
| Std Dev. |  | 7.0711 | 44
| Rel.Std.Dev. |  | 0.6285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 609c (Indicative Value 256 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 255 | 1
| 1 | 2 | 257 | 19
| Mean |  | 256 | 16
| Median |  | 256 | 256
| Std Dev. |  | 1.4142 | 45
| Rel.Std.Dev. |  | 0.5524%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 609c (Indicative Value 24.8 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 25.5 | 1
| 1 | 2 | 25.1 | 23
| 1 | 3 | 24.6 | 16
| 1 | 4 | 24.6 | 24.8333
| 1 | 5 | 24.6 | 47
| 1 | 6 | 24.6
| Mean |  | 24.8333
| Median |  | 24.6
| Std Dev. |  | 0.38297
| Rel.Std.Dev. |  | 1.542%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 609c (Indicative Value 13.4 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 13.3 | 1
| 1 | 2 | 13.2 | 12
| 1 | 3 | 13.1 | 16
| 1 | 4 | 14.1 | 13.4
| 1 | 5 | 13.6 | 48
| 1 | 6 | 13.1
| Mean |  | 13.4
| Median |  | 13.25
| Std Dev. |  | 0.38987
| Rel.Std.Dev. |  | 2.909%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 609c (Indicative Value 2625 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 2680 | 1
| 1 | 2 | 2570 | 5
| 1 | 3 | 2600 | 16
| 1 | 4 | 2630 | 2625
| 1 | 5 | 2660 | 49
| 1 | 6 | 2610
| Mean |  | 2625
| Median |  | 2620
| Std Dev. |  | 40.3733
| Rel.Std.Dev. |  | 1.538%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 609c (Indicative Value 2.83 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.9 | 1
| 1 | 2 | 2.9 | 26
| 1 | 3 | 2.8 | 16
| 1 | 4 | 2.8 | 2.8333
| 1 | 5 | 2.8 | 50
| 1 | 6 | 2.8
| Mean |  | 2.8333
| Median |  | 2.8
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 1.823%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 609c (Indicative Value 64 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 65.1 | 1
| 1 | 2 | 64.9 | 27
| 1 | 3 | 64.4 | 16
| 1 | 4 | 62 | 63.9833
| 1 | 5 | 64.2 | 51
| 1 | 6 | 63.3
| Mean |  | 63.9833
| Median |  | 64.3
| Std Dev. |  | 1.1583
| Rel.Std.Dev. |  | 1.81%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 609c (Indicative Value 1.53 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.52 | 1
| 1 | 2 | 1.5 | 15
| 1 | 3 | 1.5 | 16
| 1 | 4 | 1.6 | 1.5267
| 1 | 5 | 1.55 | 52
| 1 | 6 | 1.49
| Mean |  | 1.5267
| Median |  | 1.51
| Std Dev. |  | 0.041793
| Rel.Std.Dev. |  | 2.738%
| PDM3 |  | -2.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 609c (Indicative Value 5.53 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.7 | 1
| 1 | 2 | 5.6 | 28
| 1 | 3 | 5.5 | 16
| 1 | 4 | 5.5 | 5.5333
| 1 | 5 | 5.5 | 53
| 1 | 6 | 5.4
| Mean |  | 5.5333
| Median |  | 5.5
| Std Dev. |  | 0.10328
| Rel.Std.Dev. |  | 1.866%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 609c (Indicative Value 78 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 76 | 1
| 1 | 2 | 77 | 29
| 1 | 3 | 77 | 16
| 1 | 4 | 81 | 77.5
| 1 | 5 | 78 | 54
| 1 | 6 | 76
| Mean |  | 77.5
| Median |  | 77
| Std Dev. |  | 1.8708
| Rel.Std.Dev. |  | 2.414%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 609c (Indicative Value 7.02 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 7.26 | 1
| 1 | 2 | 7.07 | 17
| 1 | 3 | 6.97 | 16
| 1 | 4 | 6.81 | 7.0233
| 1 | 5 | 7.03 | 55
| 1 | 6 | 7
| Mean |  | 7.0233
| Median |  | 7.015
| Std Dev. |  | 0.14638
| Rel.Std.Dev. |  | 2.084%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 609c (Indicative Value 16.8 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 17 | 1
| 1 | 2 | 17 | 31
| 1 | 3 | 17 | 16
| 1 | 4 | 17 | 16.8333
| 1 | 5 | 17 | 56
| 1 | 6 | 16
| Mean |  | 16.8333
| Median |  | 17
| Std Dev. |  | 0.40825
| Rel.Std.Dev. |  | 2.425%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 609c (Indicative Value 5.78 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 5.8 | 32
| 1 | 3 | 5.8 | 16
| 1 | 4 | 5.7 | 5.7833
| 1 | 5 | 5.8 | 57
| 1 | 6 | 5.8
| Mean |  | 5.7833
| Median |  | 5.8
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 0.7059%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 609c (Indicative Value 0.471 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.481 | 1
| 1 | 2 | 0.48 | 19
| 1 | 3 | 0.47 | 16
| 1 | 4 | 0.46 | 0.47133
| 1 | 5 | 0.471 | 58
| 1 | 6 | 0.466
| Mean |  | 0.47133
| Median |  | 0.4705
| Std Dev. |  | 0.0080911
| Rel.Std.Dev. |  | 1.717%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 609c (Indicative Value 3.18 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.18 | 1
| 1 | 2 | 3.2 | 34
| 1 | 3 | 3.18 | 16
| 1 | 4 | 3.16 | 3.1833
| 1 | 5 | 3.22 | 59
| 1 | 6 | 3.16
| Mean |  | 3.1833
| Median |  | 3.18
| Std Dev. |  | 0.023381
| Rel.Std.Dev. |  | 0.7345%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 609c (Indicative Value 0.93 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.91 | 35
| 1 | 3 | 0.95 | 16
| 1 | 4 | 0.92 | 0.93333
| 1 | 5 | 0.95 | 60
| 1 | 6 | 0.92
| Mean |  | 0.93333
| Median |  | 0.935
| Std Dev. |  | 0.018619
| Rel.Std.Dev. |  | 1.995%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 609c (Indicative Value 1.69 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.69 | 1
| 1 | 2 | 1.72 | 36
| 1 | 3 | 1.69 | 16
| 1 | 4 | 1.6 | 1.685
| 1 | 5 | 1.71 | 61
| 1 | 6 | 1.7
| Mean |  | 1.685
| Median |  | 1.695
| Std Dev. |  | 0.043243
| Rel.Std.Dev. |  | 2.566%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 609c (Indicative Value 4.11 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 4.11 | 1
| 1 | 2 | 4.04 | 56
| 1 | 3 | 4.02 | 16
| 1 | 4 | 4.3 | 4.1133
| 1 | 5 | 4.21 | 62
| 1 | 6 | 4
| Mean |  | 4.1133
| Median |  | 4.075
| Std Dev. |  | 0.11928
| Rel.Std.Dev. |  | 2.9%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 609c (Indicative Value 22.1 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 23 | 1
| 1 | 2 | 22.3 | 37
| 1 | 3 | 22 | 16
| 1 | 4 | 21.6 | 22.1
| 1 | 5 | 22.1 | 63
| 1 | 6 | 21.6
| Mean |  | 22.1
| Median |  | 22.05
| Std Dev. |  | 0.52154
| Rel.Std.Dev. |  | 2.36%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 609c (Indicative Value 5.15 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.1 | 1
| 1 | 2 | 5 | 38
| 1 | 3 | 5.2 | 16
| 1 | 4 | 5.2 | 5.15
| 1 | 5 | 5.2 | 47
| 1 | 6 | 5.2
| Mean |  | 5.15
| Median |  | 5.2
| Std Dev. |  | 0.083666
| Rel.Std.Dev. |  | 1.625%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 609c (Indicative Value 6.15 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.2 | 1
| 1 | 2 | 6.2 | 23
| 1 | 3 | 6.1 | 16
| 1 | 4 | 6.2 | 6.15
| 1 | 5 | 6.1 | 48
| 1 | 6 | 6.1
| Mean |  | 6.15
| Median |  | 6.15
| Std Dev. |  | 0.054772
| Rel.Std.Dev. |  | 0.8906%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 609c (Indicative Value 0.4 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 24
| 1 | 3 | 0.39 | 16
| 1 | 4 | 0.4 | 0.39833
| 1 | 5 | 0.4 | 49
| 1 | 6 | 0.4
| Mean |  | 0.39833
| Median |  | 0.4
| Std Dev. |  | 0.0040825
| Rel.Std.Dev. |  | 1.025%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 609c (Indicative Value 3.26 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.24 | 1
| 1 | 2 | 3.21 | 13
| 1 | 3 | 3.21 | 16
| 1 | 4 | 3.37 | 3.255
| 1 | 5 | 3.31 | 50
| 1 | 6 | 3.19
| Mean |  | 3.255
| Median |  | 3.225
| Std Dev. |  | 0.070356
| Rel.Std.Dev. |  | 2.161%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 609c (Indicative Value 42.2 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 42 | 1
| 1 | 2 | 41 | 26
| 1 | 3 | 41 | 16
| 1 | 4 | 45 | 42.1667
| 1 | 5 | 43 | 51
| 1 | 6 | 41
| Mean |  | 42.1667
| Median |  | 41.5
| Std Dev. |  | 1.6021
| Rel.Std.Dev. |  | 3.799%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 609c (Indicative Value 21 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 21 | 1
| 1 | 2 | 21 | 46
| 1 | 3 | 21 | 16
| 1 | 4 | 22 | 21
| 1 | 5 | 21 | 52
| 1 | 6 | 20
| Mean |  | 21
| Median |  | 21
| Std Dev. |  | 0.63246
| Rel.Std.Dev. |  | 3.012%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 609c (Indicative Value 0.201 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.199 | 14
| 1 | 3 | 0.197 | 16
| 1 | 4 | 0.21 | 0.201
| 1 | 5 | 0.203 | 53
| 1 | 6 | 0.197
| Mean |  | 0.201
| Median |  | 0.1995
| Std Dev. |  | 0.0049396
| Rel.Std.Dev. |  | 2.458%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 609c (Indicative Value 0.032 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.031 | 1
| 1 | 2 | 0.031 | 15
| 1 | 3 | 0.031 | 16
| 1 | 4 | 0.033 | 0.0315
| 1 | 5 | 0.032 | 54
| 1 | 6 | 0.031
| Mean |  | 0.0315
| Median |  | 0.031
| Std Dev. |  | 0.00083666
| Rel.Std.Dev. |  | 2.656%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 609c (Indicative Value 41 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 42.1 | 1
| 1 | 2 | 41.7 | 29
| 1 | 3 | 41.4 | 16
| 1 | 4 | 39.5 | 40.9833
| 1 | 5 | 41 | 55
| 1 | 6 | 40.2
| Mean |  | 40.9833
| Median |  | 41.2
| Std Dev. |  | 0.97451
| Rel.Std.Dev. |  | 2.378%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 609c (Indicative Value 3 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 2.96 | 16
| 1 | 3 | 2.95 | 16
| 1 | 4 | 3.1 | 2.9983
| 1 | 5 | 3.05 | 56
| 1 | 6 | 2.93
| Mean |  | 2.9983
| Median |  | 2.98
| Std Dev. |  | 0.065549
| Rel.Std.Dev. |  | 2.186%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 609c (Indicative Value 16.4 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 16.6 | 1
| 1 | 2 | 16.7 | 30
| 1 | 3 | 16.2 | 16
| 1 | 4 | 16.3 | 16.3833
| 1 | 5 | 16.3 | 57
| 1 | 6 | 16.2
| Mean |  | 16.3833
| Median |  | 16.3
| Std Dev. |  | 0.2137
| Rel.Std.Dev. |  | 1.304%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 609c (Indicative Value 33.6 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 33.8 | 1
| 1 | 2 | 33.9 | 31
| 1 | 3 | 33.7 | 16
| 1 | 4 | 33.1 | 33.6167
| 1 | 5 | 34.3 | 58
| 1 | 6 | 32.9
| Mean |  | 33.6167
| Median |  | 33.75
| Std Dev. |  | 0.52313
| Rel.Std.Dev. |  | 1.556%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 609c (Indicative Value 16 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 16.8 | 1
| 1 | 2 | 16.4 | 17
| 1 | 3 | 15.8 | 16
| 1 | 4 | 15.6 | 15.9833
| 1 | 5 | 15.7 | 59
| 1 | 6 | 15.6
| Mean |  | 15.9833
| Median |  | 15.75
| Std Dev. |  | 0.49967
| Rel.Std.Dev. |  | 3.126%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 609c (Indicative Value 0.082 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.084 | 18
| 1 | 3 | 0.08 | 16
| 1 | 4 | 0.083 | 0.082
| 1 | 5 | 0.081 | 60
| 1 | 6 | 0.084
| Mean |  | 0.082
| Median |  | 0.082
| Std Dev. |  | 0.0018974
| Rel.Std.Dev. |  | 2.314%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 609c (Indicative Value 379 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 388 | 1
| 1 | 2 | 382 | 19
| 1 | 3 | 380 | 16
| 1 | 4 | 371 | 378.6667
| 1 | 5 | 378 | 61
| 1 | 6 | 373
| Mean |  | 378.6667
| Median |  | 379
| Std Dev. |  | 6.186
| Rel.Std.Dev. |  | 1.634%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 609c (Indicative Value 9.28 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 9.3 | 1
| 1 | 2 | 9.4 | 34
| 1 | 3 | 9.3 | 16
| 1 | 4 | 9.1 | 9.2833
| 1 | 5 | 9.4 | 62
| 1 | 6 | 9.2
| Mean |  | 9.2833
| Median |  | 9.3
| Std Dev. |  | 0.1169
| Rel.Std.Dev. |  | 1.259%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 609c (Indicative Value 125 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 129 | 1
| 1 | 2 | 127 | 35
| 1 | 3 | 124 | 16
| 1 | 4 | 122 | 125
| 1 | 5 | 125 | 63
| 1 | 6 | 123
| Mean |  | 125
| Median |  | 124.5
| Std Dev. |  | 2.6077
| Rel.Std.Dev. |  | 2.086%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 609c (Indicative Value 1.63 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.63 | 1
| 1 | 2 | 1.63 | 12
| 1 | 3 | 1.61 | 16
| 1 | 4 | 1.7 | 1.6333
| 1 | 5 | 1.66 | 47
| 1 | 6 | 1.57
| Mean |  | 1.6333
| Median |  | 1.63
| Std Dev. |  | 0.044121
| Rel.Std.Dev. |  | 2.701%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 609c (Indicative Value 4.05 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.1 | 1
| 1 | 2 | 4.1 | 38
| 1 | 3 | 4 | 16
| 1 | 4 | 4 | 4.05
| 1 | 5 | 4.1 | 48
| 1 | 6 | 4
| Mean |  | 4.05
| Median |  | 4.05
| Std Dev. |  | 0.054772
| Rel.Std.Dev. |  | 1.352%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 609c (Indicative Value 6.38 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.4 | 1
| 1 | 2 | 6.4 | 23
| 1 | 3 | 6.3 | 16
| 1 | 4 | 6.3 | 6.3833
| 1 | 5 | 6.5 | 49
| 1 | 6 | 6.4
| Mean |  | 6.3833
| Median |  | 6.4
| Std Dev. |  | 0.075277
| Rel.Std.Dev. |  | 1.179%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 609c (Indicative Value 9.26 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 9.55 | 1
| 1 | 2 | 9.31 | 14
| 1 | 3 | 9.16 | 16
| 1 | 4 | 9.19 | 9.2583
| 1 | 5 | 9.24 | 50
| 1 | 6 | 9.1
| Mean |  | 9.2583
| Median |  | 9.215
| Std Dev. |  | 0.15968
| Rel.Std.Dev. |  | 1.725%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 609c (Indicative Value 205 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 205 | 1
| 1 | 2 | 202 | 15
| 1 | 3 | 201 | 16
| 1 | 4 | 211 | 204.8333
| 1 | 5 | 208 | 51
| 1 | 6 | 202
| Mean |  | 204.8333
| Median |  | 203.5
| Std Dev. |  | 3.9707
| Rel.Std.Dev. |  | 1.939%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 609c (Indicative Value 1.22 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.22 | 1
| 1 | 2 | 1.24 | 26
| 1 | 3 | 1.22 | 16
| 1 | 4 | 1.21 | 1.2217
| 1 | 5 | 1.24 | 52
| 1 | 6 | 1.2
| Mean |  | 1.2217
| Median |  | 1.22
| Std Dev. |  | 0.016021
| Rel.Std.Dev. |  | 1.311%
| PDM3 |  | -2.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 609c (Indicative Value 0.61 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.63 | 1
| 1 | 2 | 0.6 | 27
| 1 | 3 | 0.62 | 16
| 1 | 4 | 0.61 | 0.61333
| 1 | 5 | 0.61 | 53
| 1 | 6 | 0.61
| Mean |  | 0.61333
| Median |  | 0.61
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 1.684%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 609c (Indicative Value 12.6 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 12.4 | 29
| 1 | 3 | 12.7 | 16
| 1 | 4 | 12.9 | 12.5833
| 1 | 5 | 12.6 | 54
| 1 | 6 | 12.3
| Mean |  | 12.5833
| Median |  | 12.6
| Std Dev. |  | 0.2137
| Rel.Std.Dev. |  | 1.698%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 609c (Indicative Value 0.193 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.192 | 1
| 1 | 2 | 0.189 | 16
| 1 | 3 | 0.188 | 16
| 1 | 4 | 0.202 | 0.1925
| 1 | 5 | 0.196 | 55
| 1 | 6 | 0.188
| Mean |  | 0.1925
| Median |  | 0.1905
| Std Dev. |  | 0.0055767
| Rel.Std.Dev. |  | 2.897%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 609c (Indicative Value 4.91 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.92 | 1
| 1 | 2 | 4.83 | 33
| 1 | 3 | 4.77 | 16
| 1 | 4 | 5.4 | 4.9067
| 1 | 5 | 4.81 | 56
| 1 | 6 | 4.71
| Mean |  | 4.9067
| Median |  | 4.82
| Std Dev. |  | 0.25145
| Rel.Std.Dev. |  | 5.125%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 609c (Indicative Value 9.08 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 8.8 | 1
| 1 | 2 | 9.1 | 34
| 1 | 3 | 8.9 | 16
| 1 | 4 | 9.4 | 9.0833
| 1 | 5 | 9.4 | 57
| 1 | 6 | 8.9
| Mean |  | 9.0833
| Median |  | 9
| Std Dev. |  | 0.26394
| Rel.Std.Dev. |  | 2.906%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 609c (Indicative Value 3.65 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.7 | 1
| 1 | 2 | 3.6 | 35
| 1 | 3 | 3.4 | 16
| 1 | 4 | 3.5 | 3.65
| 1 | 5 | 4.3 | 58
| 1 | 6 | 3.4
| Mean |  | 3.65
| Median |  | 3.55
| Std Dev. |  | 0.33912
| Rel.Std.Dev. |  | 9.291%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 609c (Indicative Value 13.1 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13.1 | 1
| 1 | 2 | 13.3 | 36
| 1 | 3 | 13.1 | 16
| 1 | 4 | 13 | 13.0833
| 1 | 5 | 13.2 | 59
| 1 | 6 | 12.8
| Mean |  | 13.0833
| Median |  | 13.1
| Std Dev. |  | 0.17224
| Rel.Std.Dev. |  | 1.316%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 609c (Indicative Value 0.58 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.59 | 37
| 1 | 3 | 0.58 | 16
| 1 | 4 | 0.57 | 0.57667
| 1 | 5 | 0.57 | 60
| 1 | 6 | 0.56
| Mean |  | 0.57667
| Median |  | 0.575
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 2.1%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 609c (Indicative Value 1115 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 1150 | 1
| 1 | 2 | 1140 | 18
| 1 | 3 | 1120 | 16
| 1 | 4 | 1080 | 1115
| 1 | 5 | 1100 | 61
| 1 | 6 | 1100
| Mean |  | 1115
| Median |  | 1110
| Std Dev. |  | 26.6458
| Rel.Std.Dev. |  | 2.39%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 609c (Indicative Value 269 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 271 | 1
| 1 | 2 | 262 | 19
| 1 | 3 | 261 | 16
| 1 | 4 | 286 | 268.8333
| 1 | 5 | 271 | 62
| 1 | 6 | 262
| Mean |  | 268.8333
| Median |  | 266.5
| Std Dev. |  | 9.5795
| Rel.Std.Dev. |  | 3.563%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.