## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 600c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.22116 | 0.21797 | 0.22434 | 0.21958 | 0.22274
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.21028 | 0.20452 | 0.21603 | 0.20863 | 0.21192
Infrared Combustion
S, Sulphur (wt.%) | 0.23166 | 0.22154 | 0.24178 | 0.22106 | 0.24225
4-Acid Digestion
Ag, Silver (ppm) | 23.5996 | 22.7745 | 24.4248 | 23.1035 | 24.0958
Al, Aluminium (wt.%) | 7.5602 | 7.3331 | 7.7874 | 7.3757 | 7.7448
As, Arsenic (ppm) | 39.4596 | 37.4175 | 41.5016 | 37.6959 | 41.2233
Ba, Barium (ppm) | 3038.8424 | 2931.8109 | 3145.8739 | 2952.1778 | 3125.507
Be, Beryllium (ppm) | 3.1439 | 2.9515 | 3.3362 | 3.0536 | 3.2341
Bi, Bismuth (ppm) | 0.77235 | 0.72925 | 0.81546 | 0.73978 | 0.80493
Ca, Calcium (wt.%) | 1.2465 | 1.2124 | 1.2807 | 1.2189 | 1.2741
Cd, Cadmium (ppm) | 6.3328 | 6.0263 | 6.6392 | 6.1508 | 6.5147
Ce, Cerium (ppm) | 90.9624 | 86.7192 | 95.2055 | 88.2597 | 93.665
Co, Cobalt (ppm) | 2.2161 | 2.0866 | 2.3456 | 2.1191 | 2.3131
Cr, Chromium (ppm) | 12.9696 | 11.4148 | 14.5244 | 11.6848 | 14.2544
Cs, Caesium (ppm) | 6.548 | 6.2606 | 6.8354 | 6.3593 | 6.7367
Cu, Copper (wt.%) | 0.057852 | 0.056184 | 0.059519 | 0.056737 | 0.058967
Dy, Dysprosium (ppm) | 3.7686 | 3.5953 | 3.9419 | 3.6115 | 3.9258
Er, Erbium (ppm) | 1.0662 | 0.99904 | 1.1333 | 1.0241 | 1.1082
Eu, Europium (ppm) | 1.4689 | 1.3001 | 1.6376 | 1.4178 | 1.5199
Fe, Iron (wt.%) | 2.4995 | 2.4299 | 2.5691 | 2.4418 | 2.5572
Ga, Gallium (ppm) | 22.7749 | 21.7054 | 23.8445 | 22.3189 | 23.231
Gd, Gadolinium (ppm) | 5.9607 | 5.5795 | 6.342 | 5.7353 | 6.1862
Ge, Germanium (ppm) | 0.19766 | 0.14442 | 0.2509 | 0.16955 | 0.22576
Hf, Hafnium (ppm) | 6.8578 | 6.5819 | 7.1338 | 6.6587 | 7.057
Ho, Holmium (ppm) | 0.51442 | 0.46857 | 0.56027 | 0.48839 | 0.54045
In, Indium (ppm) | 0.17783 | 0.16456 | 0.19109 | 0.16807 | 0.18758
K, Potassium (wt.%) | 3.1143 | 3.0151 | 3.2136 | 3.0459 | 3.1828
La, Lanthanum (ppm) | 45.475 | 43.0507 | 47.8993 | 44.0822 | 46.8678
Li, Lithium (ppm) | 30.6456 | 29.3693 | 31.9218 | 29.9713 | 31.3199
Lu, Lutetium (ppm) | 0.075468 | 0.052837 | 0.098099 | IND | IND
Mg, Magnesium (wt.%) | 0.14876 | 0.14167 | 0.15585 | 0.14521 | 0.15231
Mn, Manganese (wt.%) | 0.039361 | 0.038105 | 0.040616 | 0.038459 | 0.040263
Mo, Molybdenum (ppm) | 3.9733 | 3.7625 | 4.1841 | 3.7997 | 4.1469
Na, Sodium (wt.%) | 2.7891 | 2.709 | 2.8691 | 2.731 | 2.8471
Nb, Niobium (ppm) | 19.2936 | 18.4242 | 20.163 | 18.6518 | 19.9354
Nd, Neodymium (ppm) | 40.2995 | 37.7989 | 42.8001 | 38.7332 | 41.8659
Ni, Nickel (ppm) | 5.8329 | 5.4044 | 6.2614 | 5.5102 | 6.1556
P, Phosphorus (wt.%) | 0.036538 | 0.035307 | 0.037769 | 0.035581 | 0.037494
Pb, Lead (ppm) | 1128.9188 | 1097.7818 | 1160.0557 | 1110.9854 | 1146.8521
Pr, Praseodymium (ppm) | 10.3215 | 9.7874 | 10.8556 | 9.9209 | 10.7222
Rb, Rubidium (ppm) | 147.7862 | 141.1211 | 154.4513 | 143.8987 | 151.6737
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.23761 | 0.22727 | 0.24794 | 0.23064 | 0.24457
Sb, Antimony (ppm) | 64.4672 | 61.9141 | 67.0203 | 63.0667 | 65.8678
Sc, Scandium (ppm) | 4.3476 | 4.0686 | 4.6265 | 4.1968 | 4.4983
Se, Selenium (ppm) | 1.1691 | 0.38408 | 1.954 | IND | IND
Sm, Samarium (ppm) | 7.6657 | 7.2113 | 8.12 | 7.3505 | 7.9809
Sn, Tin (ppm) | 3.8692 | 3.6855 | 4.0528 | 3.6838 | 4.0545
Sr, Strontium (ppm) | 186.533 | 181.1492 | 191.9168 | 182.9394 | 190.1266
Ta, Tantalum (ppm) | 1.4376 | 1.3607 | 1.5144 | 1.3939 | 1.4812
Tb, Terbium (ppm) | 0.77134 | 0.7163 | 0.82639 | 0.73778 | 0.8049
Te, Tellurium (ppm) | 0.19047 | 0.1345 | 0.24644 | 0.14716 | 0.23377
Th, Thorium (ppm) | 15.1416 | 14.5288 | 15.7544 | 14.6264 | 15.6568
Ti, Titanium (wt.%) | 0.11312 | 0.10933 | 0.11691 | 0.11155 | 0.11469
Tl, Thallium (ppm) | 1.2055 | 1.1549 | 1.2561 | 1.1739 | 1.2372
Tm, Thulium (ppm) | 0.10967 | 0.093882 | 0.12546 | IND | IND
U, Uranium (ppm) | 5.7859 | 5.5755 | 5.9962 | 5.5833 | 5.9884
V, Vanadium (ppm) | 3.8054 | 3.3121 | 4.2987 | 3.5897 | 4.0211
W, Tungsten (ppm) | 1.6193 | 1.4825 | 1.756 | 1.5346 | 1.704
Y, Yttrium (ppm) | 15.6674 | 15.1542 | 16.1806 | 15.2544 | 16.0804
Yb, Ytterbium (ppm) | 0.62221 | 0.52848 | 0.71594 | 0.58406 | 0.66036
Zn, Zinc (ppm) | 2105.9993 | 2034.6323 | 2177.3662 | 2062.8553 | 2149.1433
Zr, Zirconium (ppm) | 259.0134 | 249.9205 | 268.1064 | 253.6429 | 264.3839
Aqua Regia Digestion
Ag, Silver (ppm) | 23.4009 | 22.6375 | 24.1643 | 22.9253 | 23.8765
Al, Aluminium (wt.%) | 0.71256 | 0.67164 | 0.75347 | 0.686 | 0.73912
As, Arsenic (ppm) | 36.7246 | 35.087 | 38.3621 | 35.6263 | 37.8228
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 200.8378 | 190.0005 | 211.675 | 191.7281 | 209.9474
Be, Beryllium (ppm) | 0.58621 | 0.53127 | 0.64115 | 0.55981 | 0.61261
Bi, Bismuth (ppm) | 0.78201 | 0.72743 | 0.83658 | 0.72366 | 0.84035
Ca, Calcium (wt.%) | 0.79395 | 0.77226 | 0.81565 | 0.77844 | 0.80947
Cd, Cadmium (ppm) | 6.1367 | 5.9028 | 6.3706 | 5.9415 | 6.3318
Ce, Cerium (ppm) | 57.4104 | 54.3161 | 60.5047 | 55.4331 | 59.3877
Co, Cobalt (ppm) | 1.819 | 1.7036 | 1.9343 | 1.7216 | 1.9163
Cr, Chromium (ppm) | 13.3849 | 12.3263 | 14.4434 | 12.526 | 14.2437
Cs, Caesium (ppm) | 1.4653 | 1.3612 | 1.5694 | 1.3988 | 1.5319
Cu, Copper (wt.%) | 0.058486 | 0.056344 | 0.060629 | 0.05744 | 0.059532
Dy, Dysprosium (ppm) | 2.1109 | 1.8289 | 2.3929 | 1.9852 | 2.2366
Er, Erbium (ppm) | 0.53565 | 0.45898 | 0.61233 | 0.50843 | 0.56288
Eu, Europium (ppm) | 0.80436 | 0.68171 | 0.927 | 0.76619 | 0.84252
Fe, Iron (wt.%) | 2.0316 | 1.9694 | 2.0938 | 1.9964 | 2.0669
Ga, Gallium (ppm) | 5.2924 | 5.0102 | 5.5745 | 5.0968 | 5.4879
Gd, Gadolinium (ppm) | 3.9817 | 3.6604 | 4.303 | 3.7606 | 4.2027
Ge, Germanium (ppm) | 0.11143 | 0.079616 | 0.14324 | IND | IND
Hf, Hafnium (ppm) | 1.4146 | 1.3217 | 1.5075 | 1.368 | 1.4612
Hg, Mercury (ppm) | 0.14378 | 0.11466 | 0.17289 | IND | IND
Ho, Holmium (ppm) | 0.25791 | 0.22745 | 0.28837 | 0.24389 | 0.27194
In, Indium (ppm) | 0.13672 | 0.12802 | 0.14542 | 0.12768 | 0.14576
K, Potassium (wt.%) | 0.23339 | 0.21939 | 0.24739 | 0.22092 | 0.24586
La, Lanthanum (ppm) | 28.7172 | 27.1332 | 30.3013 | 27.7052 | 29.7292
Li, Lithium (ppm) | 18.1454 | 17.1351 | 19.1557 | 17.4737 | 18.8171
Lu, Lutetium (ppm) | 0.028973 | 0.019676 | 0.038271 | IND | IND
Mg, Magnesium (wt.%) | 0.094498 | 0.087999 | 0.101 | 0.091733 | 0.097262
Mn, Manganese (wt.%) | 0.034159 | 0.033034 | 0.035283 | 0.033332 | 0.034985
Mo, Molybdenum (ppm) | 3.6283 | 3.3871 | 3.8696 | 3.4754 | 3.7813
Na, Sodium (wt.%) | 0.072862 | 0.069269 | 0.076455 | 0.067798 | 0.077927
Nb, Niobium (ppm) | 0.98628 | 0.83118 | 1.1414 | 0.8712 | 1.1013
Nd, Neodymium (ppm) | 26.5984 | 24.5916 | 28.6052 | 25.4621 | 27.7347
Ni, Nickel (ppm) | 5.5912 | 5.2099 | 5.9726 | 5.2501 | 5.9324
P, Phosphorus (wt.%) | 0.029566 | 0.028473 | 0.030659 | 0.028974 | 0.030158
Pb, Lead (ppm) | 1097.7478 | 1069.9811 | 1125.5144 | 1082.653 | 1112.8425
Pr, Praseodymium (ppm) | 6.6982 | 6.3903 | 7.006 | 6.4189 | 6.9774
Rb, Rubidium (ppm) | 14.1589 | 13.1397 | 15.1781 | 13.5557 | 14.7621
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.24046 | 0.2315 | 0.24942 | 0.23344 | 0.24748
Sb, Antimony (ppm) | 55.6883 | 53.3258 | 58.0508 | 54.0718 | 57.3047
Sc, Scandium (ppm) | 1.5699 | 1.4469 | 1.6928 | 1.466 | 1.6738
Se, Selenium (ppm) | 0.32606 | 0.19996 | 0.45216 | IND | IND
Sm, Samarium (ppm) | 4.8007 | 4.0393 | 5.562 | 4.5156 | 5.0858
Sn, Tin (ppm) | 1.2 | 1.0842 | 1.3157 | 1.1296 | 1.2703
Sr, Strontium (ppm) | 21.6444 | 20.6139 | 22.675 | 20.96 | 22.3288
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.47512 | 0.41605 | 0.53419 | 0.4424 | 0.50784
Te, Tellurium (ppm) | 0.15765 | 0.13182 | 0.18348 | IND | IND
Th, Thorium (ppm) | 10.6206 | 9.9348 | 11.3063 | 10.3117 | 10.9295
Ti, Titanium (wt.%) | 0.022589 | 0.020668 | 0.024511 | 0.021269 | 0.02391
Tl, Thallium (ppm) | 0.46517 | 0.4296 | 0.50074 | 0.43938 | 0.49096
U, Uranium (ppm) | 2.8277 | 2.6186 | 3.0368 | 2.7055 | 2.9499
V, Vanadium (ppm) | 1.9818 | 1.6239 | 2.3398 | 1.8096 | 2.1541
W, Tungsten (ppm) | 0.45696 | 0.39383 | 0.52009 | 0.424 | 0.48992
Y, Yttrium (ppm) | 8.0702 | 7.6693 | 8.4712 | 7.783 | 8.3575
Yb, Ytterbium (ppm) | 0.23727 | 0.20691 | 0.26763 | 0.21837 | 0.25617
Zn, Zinc (ppm) | 2052.4771 | 2003.8824 | 2101.0718 | 2015.9429 | 2089.0113
Zr, Zirconium (ppm) | 49.9236 | 46.7427 | 53.1045 | 48.253 | 51.5941
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 600c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Infrared Combustion
C | wt.% | 0.196 | C-(Inorganic) | wt.% | 0.195 | C-(Organic) | wt.% | 0.038333
4-Acid Digestion
B | ppm | 14.8333 | Hg | ppm | 0.043333
Aqua Regia Digestion
Pd | ppb | 11.9167 | Pt | ppb | 5.5 | Tm | ppm | 0.048435
Alkaline Leach
S-(Sulphate) | wt.% | 0.0657 | S-(Sulphide) | wt.% | 0.14305
Acid Leach
S-(Sulphate) | wt.% | 0.015
Borate Fusion XRF
Al2O3 | wt.% | 14.845 | Fe2O3 | wt.% | 3.5814 | S | wt.% | 0.239
As | ppm | 60 | K2O | wt.% | 3.775 | SiO2 | wt.% | 69.03
BaO | ppm | 3561.635 | MgO | wt.% | 0.28 | Sn | ppm | 55
CaO | wt.% | 1.725 | MnO | wt.% | 0.0475 | Sr | ppm | 205
Cl | ppm | 25 | Na2O | wt.% | 3.745 | TiO2 | wt.% | 0.2
Co | ppm | < 10 | Ni | ppm | 10 | V2O5 | ppm | 13.389
Cr2O3 | ppm | 29.23 | P2O5 | wt.% | 0.082498 | Zn | ppm | 2055
Cu | wt.% | 0.062 | Pb | ppm | 1120 | Zr | ppm | 290
Thermogravimetry
LOI1000 | wt.% | 1.56
Laser Ablation ICP-MS
Ag | ppm | 28.55 | Hf | ppm | 7.905 | Sm | ppm | 7.735
As | ppm | 40.5 | Ho | ppm | 0.54 | Sn | ppm | 4.2
Ba | ppm | 3070 | In | ppm | 0.15 | Sr | ppm | 183
Be | ppm | 3.3 | La | ppm | 47.45 | Ta | ppm | 1.435
Bi | ppm | 0.88 | Lu | ppm | 0.065 | Tb | ppm | 0.805
Cd | ppm | 7.5 | Mn | wt.% | 0.0401 | Te | ppm | 0.3
Ce | ppm | 90.5 | Mo | ppm | 3.9 | Th | ppm | 14.85
Co | ppm | 2.75 | Nb | ppm | 18.9 | Ti | wt.% | 0.1175
Cr | ppm | 17 | Nd | ppm | 41.15 | Tl | ppm | 1.1
Cs | ppm | 6.59 | Ni | ppm | 8 | Tm | ppm | 0.125
Cu | wt.% | 0.0592 | Pb | ppm | 1195 | U | ppm | 5.905
Dy | ppm | 3.67 | Pr | ppm | 11.15 | V | ppm | 4.05
Er | ppm | 1.075 | Rb | ppm | 144 | W | ppm | 1.75
Eu | ppm | 1.43 | Re | ppm | 0.0075 | Y | ppm | 15.7
Ga | ppm | 21.8 | Sb | ppm | 70.95 | Yb | ppm | 0.59
Gd | ppm | 6.09 | Sc | ppm | 4.3 | Zn | ppm | 2275
Ge | ppm | 1.325 | Se | ppm | < 5 | Zr | ppm | 288
3-Acid Digestion (no HF)
Ag | ppm | 24.5333 | Gd | ppm | 5.9167 | S | wt.% | 0.22483
Al2O3 | wt.% | 14.25 | Hf | ppm | 6.7333 | Sc | ppm | 4.1667
Ba | ppm | 2908.3333 | Ho | ppm | 0.45 | Sm | ppm | 7.3667
Be | ppm | 3.2 | K2O | wt.% | 3.6417 | Sn | ppm | 3.925
Bi | ppm | 1.0333 | La | ppm | 46.8333 | Sr | ppm | 187.3333
CaO | wt.% | 1.75 | Li | ppm | 30.3333 | Ta | ppm | 1.3683
Cd | ppm | 6.2167 | MgO | wt.% | 0.25017 | Tb | ppm | 0.71667
Ce | ppm | 85.5 | MnO | wt.% | 0.049 | Th | ppm | 13.5333
Co | ppm | 2.1517 | Mo | ppm | 4.1717 | TiO2 | wt.% | 0.18417
Cr | ppm | 13.8333 | Na2O | wt.% | 3.725 | U | ppm | 5.705
Cs | ppm | 6.4 | Nb | ppm | 18.6 | V | ppm | 3.8
Cu | wt.% | 0.057467 | Nd | ppm | 37.9333 | W | ppm | 1.8
Dy | ppm | 3.65 | Ni | ppm | 6.4333 | Y | ppm | 14.9
Er | ppm | 1.0133 | P2O5 | wt.% | 0.084167 | Yb | ppm | 0.56333
Eu | ppm | 1.9933 | Pb | ppm | 1100 | Zn | ppm | 2081.6667
Fe2O3 | wt.% | 3.3667 | Pr | ppm | 10.3833 | Zr | ppm | 300.8333
Ga | ppm | 21.5833 | Rb | ppm | 145.6667
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 600c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.22116 | 0.0093752 | 0.20241 | 0.23991 | 0.19303 | 0.24928 | 4.239% | 8.478% | 12.72% | 0.2101 | 0.23222
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.21028 | 0.014567 | 0.18114 | 0.23941 | 0.16658 | 0.25398 | 6.927% | 13.85% | 20.78% | 0.19976 | 0.22079
Infrared Combustion
S, wt.% | 0.23166 | 0.0083131 | 0.21503 | 0.24828 | 0.20672 | 0.2566 | 3.589% | 7.177% | 10.77% | 0.22008 | 0.24324
4-Acid Digestion
Ag, ppm | 23.5996 | 0.88131 | 21.837 | 25.3623 | 20.9557 | 26.2436 | 3.734% | 7.469% | 11.2% | 22.4197 | 24.7796
Al, wt.% | 7.5602 | 0.32164 | 6.917 | 8.2035 | 6.5953 | 8.5252 | 4.254% | 8.509% | 12.76% | 7.1822 | 7.9383
As, ppm | 39.4596 | 2.869 | 33.7215 | 45.1976 | 30.8525 | 48.0667 | 7.271% | 14.54% | 21.81% | 37.4866 | 41.4325
Ba, ppm | 3038.8424 | 135.5288 | 2767.7847 | 3309.9 | 2632.2559 | 3445.4289 | 4.46% | 8.92% | 13.38% | 2886.9003 | 3190.7845
Be, ppm | 3.1439 | 0.23457 | 2.6747 | 3.613 | 2.4401 | 3.8476 | 7.461% | 14.92% | 22.38% | 2.9867 | 3.3011
Bi, ppm | 0.77235 | 0.045335 | 0.68168 | 0.86302 | 0.63635 | 0.90836 | 5.87% | 11.74% | 17.61% | 0.73374 | 0.81097
Ca, wt.% | 1.2465 | 0.032822 | 1.1809 | 1.3122 | 1.148 | 1.345 | 2.633% | 5.266% | 7.899% | 1.1842 | 1.3088
Cd, ppm | 6.3328 | 0.37889 | 5.575 | 7.0906 | 5.1961 | 7.4695 | 5.983% | 11.97% | 17.95% | 6.0161 | 6.6494
Ce, ppm | 90.9624 | 5.8794 | 79.2035 | 102.7212 | 73.324 | 108.6007 | 6.464% | 12.93% | 19.39% | 86.4142 | 95.5105
Co, ppm | 2.2161 | 0.14312 | 1.9299 | 2.5023 | 1.7867 | 2.6455 | 6.458% | 12.92% | 19.37% | 2.1053 | 2.3269
Cr, ppm | 12.9696 | 1.7666 | 9.4363 | 16.5028 | 7.6697 | 18.2694 | 13.62% | 27.24% | 40.86% | 12.3211 | 13.6181
Cs, ppm | 6.548 | 0.42874 | 5.6905 | 7.4055 | 5.2618 | 7.8342 | 6.548% | 13.1% | 19.64% | 6.2206 | 6.8754
Cu, wt.% | 0.057852 | 0.0017584 | 0.054335 | 0.061368 | 0.052577 | 0.063127 | 3.039% | 6.079% | 9.118% | 0.054959 | 0.060744
Dy, ppm | 3.7686 | 0.12766 | 3.5133 | 4.0239 | 3.3856 | 4.1516 | 3.387% | 6.775% | 10.16% | 3.5802 | 3.957
Er, ppm | 1.0662 | 0.048857 | 0.96844 | 1.1639 | 0.91958 | 1.2127 | 4.583% | 9.165% | 13.75% | 1.0128 | 1.1195
Eu, ppm | 1.4689 | 0.11195 | 1.2449 | 1.6928 | 1.133 | 1.8047 | 7.622% | 15.24% | 22.87% | 1.3954 | 1.5423
Fe, wt.% | 2.4995 | 0.082984 | 2.3335 | 2.6654 | 2.2505 | 2.7484 | 3.32% | 6.64% | 9.96% | 2.3745 | 2.6244
Ga, ppm | 22.7749 | 1.5694 | 19.6362 | 25.9137 | 18.0669 | 27.483 | 6.891% | 13.78% | 20.67% | 21.6362 | 23.9137
Gd, ppm | 5.9607 | 0.45433 | 5.0521 | 6.8694 | 4.5978 | 7.3237 | 7.622% | 15.24% | 22.87% | 5.6627 | 6.2588
Ge, ppm | 0.19766 | 0.071947 | 0.053763 | 0.34155 | 0 | 0.4135 | 36.4% | 72.8% | 109.2% | 0.18777 | 0.20754
Hf, ppm | 6.8578 | 0.27995 | 6.2979 | 7.4177 | 6.018 | 7.6977 | 4.082% | 8.165% | 12.25% | 6.5149 | 7.2007
Ho, ppm | 0.51442 | 0.053045 | 0.40833 | 0.62051 | 0.35528 | 0.67355 | 10.31% | 20.62% | 30.93% | 0.4887 | 0.54014
In, ppm | 0.17783 | 0.011614 | 0.1546 | 0.20105 | 0.14298 | 0.21267 | 6.531% | 13.06% | 19.59% | 0.16893 | 0.18672
K, wt.% | 3.1143 | 0.14042 | 2.8335 | 3.3952 | 2.6931 | 3.5356 | 4.509% | 9.018% | 13.53% | 2.9586 | 3.2701
La, ppm | 45.475 | 4.0851 | 37.3048 | 53.6453 | 33.2196 | 57.7304 | 8.983% | 17.97% | 26.95% | 43.2013 | 47.7488
Li, ppm | 30.6456 | 1.3699 | 27.9057 | 33.3854 | 26.5358 | 34.7553 | 4.47% | 8.94% | 13.41% | 29.1133 | 32.1778
Lu, ppm | 0.075468 | 0.018627 | 0.038213 | 0.11272 | 0.019586 | 0.13135 | 24.68% | 49.36% | 74.05% | 0.071695 | 0.079241
Mg, wt.% | 0.14876 | 0.0088246 | 0.13111 | 0.16641 | 0.12229 | 0.17523 | 5.932% | 11.86% | 17.8% | 0.14132 | 0.1562
Mn, wt.% | 0.039361 | 0.0018556 | 0.03565 | 0.043072 | 0.033794 | 0.044928 | 4.714% | 9.429% | 14.14% | 0.037393 | 0.041329
Mo, ppm | 3.9733 | 0.15624 | 3.6608 | 4.2858 | 3.5046 | 4.442 | 3.932% | 7.864% | 11.8% | 3.7746 | 4.172
Na, wt.% | 2.7891 | 0.11765 | 2.5538 | 3.0244 | 2.4361 | 3.142 | 4.218% | 8.436% | 12.65% | 2.6496 | 2.9285
Nb, ppm | 19.2936 | 0.9522 | 17.3892 | 21.198 | 16.437 | 22.1502 | 4.935% | 9.871% | 14.81% | 18.3289 | 20.2583
Nd, ppm | 40.2995 | 2.8033 | 34.693 | 45.9061 | 31.8897 | 48.7094 | 6.956% | 13.91% | 20.87% | 38.2845 | 42.3145
Ni, ppm | 5.8329 | 0.42734 | 4.9782 | 6.6876 | 4.5509 | 7.1149 | 7.326% | 14.65% | 21.98% | 5.5412 | 6.1245
P, wt.% | 0.036538 | 0.0012585 | 0.034021 | 0.039055 | 0.032762 | 0.040313 | 3.445% | 6.889% | 10.33% | 0.034711 | 0.038365
Pb, ppm | 1128.9188 | 33.7641 | 1061.3906 | 1196.4469 | 1027.6265 | 1230.211 | 2.991% | 5.982% | 8.972% | 1072.4728 | 1185.3647
Pr, ppm | 10.3215 | 0.59474 | 9.132 | 11.511 | 8.5373 | 12.1057 | 5.762% | 11.52% | 17.29% | 9.8054 | 10.8376
Rb, ppm | 147.7862 | 10.7031 | 126.38 | 169.1923 | 115.6769 | 179.8954 | 7.242% | 14.48% | 21.73% | 140.3969 | 155.1755
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.23761 | 0.011807 | 0.21399 | 0.26122 | 0.20219 | 0.27303 | 4.969% | 9.938% | 14.91% | 0.22573 | 0.24949
Sb, ppm | 64.4672 | 3.7308 | 57.0056 | 71.9289 | 53.2748 | 75.6597 | 5.787% | 11.57% | 17.36% | 61.2439 | 67.6906
Sc, ppm | 4.3476 | 0.44264 | 3.4623 | 5.2328 | 3.0196 | 5.6755 | 10.18% | 20.36% | 30.54% | 4.1302 | 4.5649
Se, ppm | 1.1691 | 0.41934 | 0.33038 | 2.0077 | 0 | 2.4271 | 35.87% | 71.74% | 107.6% | 1.1106 | 1.2275
Sm, ppm | 7.6657 | 0.44921 | 6.7673 | 8.5641 | 6.3181 | 9.0133 | 5.86% | 11.72% | 17.58% | 7.2824 | 8.049
Sn, ppm | 3.8692 | 0.17503 | 3.5191 | 4.2192 | 3.344 | 4.3943 | 4.524% | 9.048% | 13.57% | 3.6757 | 4.0626
Sr, ppm | 186.533 | 6.8164 | 172.9002 | 200.1658 | 166.0838 | 206.9822 | 3.654% | 7.309% | 10.96% | 177.2063 | 195.8596
Ta, ppm | 1.4376 | 0.095722 | 1.2461 | 1.629 | 1.1504 | 1.7247 | 6.659% | 13.32% | 19.98% | 1.3657 | 1.5094
Tb, ppm | 0.77134 | 0.059706 | 0.65193 | 0.89076 | 0.59222 | 0.95046 | 7.741% | 15.48% | 23.22% | 0.73278 | 0.80991
Te, ppm | 0.19047 | 0.045375 | 0.099716 | 0.28122 | 0.054341 | 0.32659 | 23.82% | 47.65% | 71.47% | 0.18094 | 0.19999
Th, ppm | 15.1416 | 0.66376 | 13.8141 | 16.4691 | 13.1503 | 17.1329 | 4.384% | 8.767% | 13.15% | 14.3845 | 15.8987
Ti, wt.% | 0.11312 | 0.0048542 | 0.10341 | 0.12283 | 0.098558 | 0.12768 | 4.291% | 8.582% | 12.87% | 0.10746 | 0.11878
Tl, ppm | 1.2055 | 0.053089 | 1.0993 | 1.3117 | 1.0463 | 1.3648 | 4.404% | 8.808% | 13.21% | 1.1453 | 1.2658
Tm, ppm | 0.10967 | 0.0086907 | 0.092291 | 0.12705 | 0.0836 | 0.13574 | 7.924% | 15.85% | 23.77% | 0.10419 | 0.11516
U, ppm | 5.7859 | 0.20802 | 5.3698 | 6.2019 | 5.1618 | 6.4099 | 3.595% | 7.191% | 10.79% | 5.4966 | 6.0752
V, ppm | 3.8054 | 0.69504 | 2.4153 | 5.1955 | 1.7203 | 5.8905 | 18.26% | 36.53% | 54.79% | 3.6151 | 3.9957
W, ppm | 1.6193 | 0.082963 | 1.4534 | 1.7852 | 1.3704 | 1.8682 | 5.123% | 10.25% | 15.37% | 1.5383 | 1.7002
Y, ppm | 15.6674 | 0.54716 | 14.5731 | 16.7617 | 14.0259 | 17.3089 | 3.492% | 6.985% | 10.48% | 14.884 | 16.4508
Yb, ppm | 0.62221 | 0.075062 | 0.47209 | 0.77233 | 0.39702 | 0.8474 | 12.06% | 24.13% | 36.19% | 0.5911 | 0.65332
Zn, ppm | 2105.9993 | 90.3085 | 1925.3823 | 2286.6163 | 1835.0738 | 2376.9248 | 4.288% | 8.576% | 12.86% | 2000.6993 | 2211.2992
Zr, ppm | 259.0134 | 10.3874 | 238.2387 | 279.7881 | 227.8513 | 290.1755 | 4.01% | 8.021% | 12.03% | 246.0627 | 271.9641
Aqua Regia Digestion
Ag, ppm | 23.4009 | 1.145 | 21.111 | 25.6909 | 19.966 | 26.8358 | 4.893% | 9.786% | 14.68% | 22.2309 | 24.571
Al, wt.% | 0.71256 | 0.050099 | 0.61236 | 0.81276 | 0.56226 | 0.86286 | 7.031% | 14.06% | 21.09% | 0.67693 | 0.74819
As, ppm | 36.7246 | 1.8273 | 33.0701 | 40.3791 | 31.2428 | 42.2063 | 4.976% | 9.951% | 14.93% | 34.8883 | 38.5608
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 200.8378 | 18.3077 | 164.2223 | 237.4532 | 145.9145 | 255.761 | 9.116% | 18.23% | 27.35% | 190.7959 | 210.8796
Be, ppm | 0.58621 | 0.05169 | 0.48283 | 0.68959 | 0.43114 | 0.74128 | 8.818% | 17.64% | 26.45% | 0.5569 | 0.61552
Bi, ppm | 0.78201 | 0.054461 | 0.67308 | 0.89093 | 0.61862 | 0.94539 | 6.964% | 13.93% | 20.89% | 0.74291 | 0.82111
Ca, wt.% | 0.79395 | 0.031642 | 0.73067 | 0.85724 | 0.69903 | 0.88888 | 3.985% | 7.971% | 11.96% | 0.75426 | 0.83365
Cd, ppm | 6.1367 | 0.27212 | 5.5924 | 6.6809 | 5.3203 | 6.953 | 4.434% | 8.869% | 13.3% | 5.8298 | 6.4435
Ce, ppm | 57.4104 | 3.9162 | 49.5779 | 65.2429 | 45.6617 | 69.1591 | 6.821% | 13.64% | 20.46% | 54.5399 | 60.2809
Co, ppm | 1.819 | 0.072198 | 1.6746 | 1.9634 | 1.6024 | 2.0356 | 3.969% | 7.938% | 11.91% | 1.728 | 1.9099
Cr, ppm | 13.3849 | 0.86315 | 11.6586 | 15.1112 | 10.7954 | 15.9743 | 6.449% | 12.9% | 19.35% | 12.7156 | 14.0541
Cs, ppm | 1.4653 | 0.15809 | 1.1491 | 1.7815 | 0.99105 | 1.9396 | 10.79% | 21.58% | 32.37% | 1.3921 | 1.5386
Cu, wt.% | 0.058486 | 0.002237 | 0.054012 | 0.06296 | 0.051775 | 0.065197 | 3.825% | 7.65% | 11.47% | 0.055562 | 0.061411
Dy, ppm | 2.1109 | 0.28032 | 1.5503 | 2.6715 | 1.2699 | 2.9519 | 13.28% | 26.56% | 39.84% | 2.0054 | 2.2164
Er, ppm | 0.53565 | 0.073143 | 0.38937 | 0.68194 | 0.31622 | 0.75508 | 13.65% | 27.31% | 40.96% | 0.50887 | 0.56244
Eu, ppm | 0.80436 | 0.11234 | 0.57967 | 1.029 | 0.46732 | 1.1414 | 13.97% | 27.93% | 41.9% | 0.76414 | 0.84457
Fe, wt.% | 2.0316 | 0.080073 | 1.8715 | 2.1918 | 1.7914 | 2.2719 | 3.941% | 7.883% | 11.82% | 1.9301 | 2.1332
Ga, ppm | 5.2924 | 0.45068 | 4.391 | 6.1937 | 3.9403 | 6.6444 | 8.516% | 17.03% | 25.55% | 5.0277 | 5.557
Gd, ppm | 3.9817 | 0.29101 | 3.3997 | 4.5637 | 3.1086 | 4.8547 | 7.309% | 14.62% | 21.93% | 3.7826 | 4.1808
Ge, ppm | 0.11143 | 0.029002 | 0.053425 | 0.16943 | 0.024424 | 0.19843 | 26.03% | 52.05% | 78.08% | 0.10586 | 0.117
Hf, ppm | 1.4146 | 0.084042 | 1.2465 | 1.5827 | 1.1625 | 1.6667 | 5.941% | 11.88% | 17.82% | 1.3439 | 1.4853
Hg, ppm | 0.14378 | 0.017273 | 0.10923 | 0.17832 | 0.091957 | 0.19559 | 12.01% | 24.03% | 36.04% | 0.13659 | 0.15096
Ho, ppm | 0.25791 | 0.023604 | 0.2107 | 0.30512 | 0.1871 | 0.32872 | 9.152% | 18.3% | 27.46% | 0.24502 | 0.27081
In, ppm | 0.13672 | 0.0076689 | 0.12138 | 0.15206 | 0.11371 | 0.15973 | 5.609% | 11.22% | 16.83% | 0.12988 | 0.14356
K, wt.% | 0.23339 | 0.017776 | 0.19784 | 0.26894 | 0.18006 | 0.28672 | 7.616% | 15.23% | 22.85% | 0.22172 | 0.24506
La, ppm | 28.7172 | 2.1923 | 24.3325 | 33.1019 | 22.1402 | 35.2943 | 7.634% | 15.27% | 22.9% | 27.2813 | 30.1531
Li, ppm | 18.1454 | 1.6542 | 14.8371 | 21.4537 | 13.183 | 23.1079 | 9.116% | 18.23% | 27.35% | 17.2382 | 19.0527
Lu, ppm | 0.028973 | 0.0051597 | 0.018654 | 0.039293 | 0.013494 | 0.044452 | 17.81% | 35.62% | 53.43% | 0.027525 | 0.030422
Mg, wt.% | 0.094498 | 0.0042885 | 0.085921 | 0.10307 | 0.081632 | 0.10736 | 4.538% | 9.076% | 13.61% | 0.089773 | 0.099222
Mn, wt.% | 0.034159 | 0.0014425 | 0.031274 | 0.037044 | 0.029831 | 0.038486 | 4.223% | 8.446% | 12.67% | 0.032451 | 0.035866
Mo, ppm | 3.6283 | 0.21253 | 3.2033 | 4.0534 | 2.9907 | 4.2659 | 5.858% | 11.72% | 17.57% | 3.4469 | 3.8097
Na, wt.% | 0.072862 | 0.0059842 | 0.060894 | 0.084831 | 0.05491 | 0.090815 | 8.213% | 16.43% | 24.64% | 0.069219 | 0.076505
Nb, ppm | 0.98628 | 0.17864 | 0.62899 | 1.3436 | 0.45035 | 1.5222 | 18.11% | 36.23% | 54.34% | 0.93696 | 1.0356
Nd, ppm | 26.5984 | 1.7659 | 23.0665 | 30.1303 | 21.3006 | 31.8963 | 6.639% | 13.28% | 19.92% | 25.2685 | 27.9284
Ni, ppm | 5.5912 | 0.34609 | 4.899 | 6.2834 | 4.5529 | 6.6295 | 6.19% | 12.38% | 18.57% | 5.3117 | 5.8708
P, wt.% | 0.029566 | 0.00090424 | 0.027757 | 0.031374 | 0.026853 | 0.032279 | 3.058% | 6.117% | 9.175% | 0.028088 | 0.031044
Pb, ppm | 1097.7478 | 31.9227 | 1033.9023 | 1161.5932 | 1001.9796 | 1193.5159 | 2.908% | 5.816% | 8.724% | 1042.8604 | 1152.6352
Pr, ppm | 6.6982 | 0.37291 | 5.9524 | 7.444 | 5.5794 | 7.8169 | 5.567% | 11.13% | 16.7% | 6.3633 | 7.0331
Rb, ppm | 14.1589 | 1.5746 | 11.0097 | 17.3081 | 9.4351 | 18.8826 | 11.12% | 22.24% | 33.36% | 13.4509 | 14.8668
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.24046 | 0.011375 | 0.21771 | 0.26321 | 0.20633 | 0.27458 | 4.731% | 9.461% | 14.19% | 0.22844 | 0.25248
Sb, ppm | 55.6883 | 4.1999 | 47.2885 | 64.088 | 43.0887 | 68.2879 | 7.542% | 15.08% | 22.63% | 52.9039 | 58.4727
Sc, ppm | 1.5699 | 0.15848 | 1.2529 | 1.8868 | 1.0944 | 2.0453 | 10.1% | 20.19% | 30.29% | 1.4914 | 1.6484
Se, ppm | 0.32606 | 0.12257 | 0.080921 | 0.5712 | 0 | 0.69377 | 37.59% | 75.18% | 112.8% | 0.30976 | 0.34236
Sm, ppm | 4.8007 | 0.82312 | 3.1544 | 6.4469 | 2.3313 | 7.27 | 17.15% | 34.29% | 51.44% | 4.5606 | 5.0407
Sn, ppm | 1.2 | 0.12595 | 0.94807 | 1.4519 | 0.82212 | 1.5778 | 10.5% | 20.99% | 31.49% | 1.14 | 1.26
Sr, ppm | 21.6444 | 1.242 | 19.1605 | 24.1284 | 17.9185 | 25.3704 | 5.738% | 11.48% | 17.21% | 20.5622 | 22.7267
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.47512 | 0.025579 | 0.42396 | 0.52628 | 0.39838 | 0.55186 | 5.384% | 10.77% | 16.15% | 0.45136 | 0.49888
Te, ppm | 0.15765 | 0.021567 | 0.11451 | 0.20078 | 0.092946 | 0.22235 | 13.68% | 27.36% | 41.04% | 0.14976 | 0.16553
Th, ppm | 10.6206 | 0.63806 | 9.3445 | 11.8967 | 8.7064 | 12.5348 | 6.008% | 12.02% | 18.02% | 10.0896 | 11.1516
Ti, wt.% | 0.022589 | 0.002858 | 0.016873 | 0.028305 | 0.014015 | 0.031163 | 12.65% | 25.3% | 37.96% | 0.02146 | 0.023719
Tl, ppm | 0.46517 | 0.04091 | 0.38335 | 0.54699 | 0.34244 | 0.5879 | 8.795% | 17.59% | 26.38% | 0.44191 | 0.48843
U, ppm | 2.8277 | 0.1577 | 2.5123 | 3.1431 | 2.3546 | 3.3008 | 5.577% | 11.15% | 16.73% | 2.6863 | 2.9691
V, ppm | 1.9818 | 0.44808 | 1.0857 | 2.878 | 0.63762 | 3.3261 | 22.61% | 45.22% | 67.83% | 1.8828 | 2.0809
W, ppm | 0.45696 | 0.038888 | 0.37918 | 0.53474 | 0.3403 | 0.57362 | 8.51% | 17.02% | 25.53% | 0.43411 | 0.47981
Y, ppm | 8.0702 | 0.58367 | 6.9029 | 9.2376 | 6.3192 | 9.8212 | 7.232% | 14.46% | 21.7% | 7.6667 | 8.4738
Yb, ppm | 0.23727 | 0.013268 | 0.21073 | 0.26381 | 0.19747 | 0.27708 | 5.592% | 11.18% | 16.78% | 0.22541 | 0.24913
Zn, ppm | 2052.4771 | 70.5622 | 1911.3527 | 2193.6015 | 1840.7905 | 2264.1637 | 3.438% | 6.876% | 10.31% | 1949.8532 | 2155.101
Zr, ppm | 49.9236 | 4.2126 | 41.4983 | 58.3488 | 37.2857 | 62.5615 | 8.438% | 16.88% | 25.31% | 47.4274 | 52.4197
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 600c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
HCl*IRC | hydrochloric acid digestion with infrared combustion
HCl*OES | hydrochloric acid digest with inductively coupled plasma optical emission spectroscopy
I-550*IRC | ignition @ 550 degrees Celsius with infrared combustion
INAA | instrumental neutron activation analysis
IRC | infrared combustion
IRC-x*Calc | calculation by difference from IRC total value
LOI*TGA | loss on ignition by thermogravimetric analysis
Na2CO3*GRAV | sodium carbonate solution with Gravimetric finish
Na2CO3*IRC | sodium carbonate leach with infrared combustion
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 600c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek Genalysis, Perth, WA, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Intertek Testing Services, Townsville, QLD, Australia
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Mineral Services, Townsville, QLD, Australia
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 600c (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
1 | 1 | ANSLu | 1 | 9 | 9 | 261700 | 0.085955 | 0.22212 | 0.23286
1 | 1 | ANSLu | 1 | 4 | 3 | 261701 | 0.08779 | 0.23279 | 0.23342
1 | 1 | ANSLu | 1 | 8 | 7 | 261702 | 0.084627 | 0.22568 | 0.23305
1 | 1 | ANSLu | 1 | 18 | 6 | 261703 | 0.0858 | 0.22533 | 0.23303
1 | 1 | ANSLu | 1 | 3 | 3 | 261704 | 0.085504 | 0.23368 | 0.23347
1 | 1 | ANSLu | 1 | 6 | 5 | 261705 | 0.086356 | 0.25457 | 0.23459
1 | 1 | ANSLu | 1 | 13 | 2 | 261706 | 0.087058 | 0.23681 | 0.23364
1 | 1 | ANSLu | 1 | 2 | 1 | 261707 | 0.082578 | 0.22495 | 0.23301
1 | 1 | ANSLu | 1 | 11 | 11 | 261708 | 0.087689 | 0.23698 | 0.23365
1 | 1 | ANSLu | 1 | 19 | 8 | 261709 | 0.088152 | 0.22057 | 0.23277
1 | 1 | ANSLu | 1 | 15 | 4 | 261710 | 0.083031 | 0.23872 | 0.23374
1 | 1 | ANSLu | 1 | 17 | 6 | 261711 | 0.082217 | 0.23909 | 0.23376
1 | 1 | ANSLu | 1 | 5 | 5 | 261712 | 0.086258 | 0.25028 | 0.23436
1 | 1 | ANSLu | 1 | 10 | 9 | 261713 | 0.082713 | 0.22318 | 0.23291
1 | 1 | ANSLu | 1 | 1 | 1 | 261714 | 0.083699 | 0.23627 | 0.23361
1 | 1 | ANSLu | 1 | 16 | 4 | 261715 | 0.087918 | 0.24243 | 0.23394
1 | 1 | ANSLu | 1 | 7 | 7 | 261716 | 0.087323 | 0.22041 | 0.23276
1 | 1 | ANSLu | 1 | 14 | 2 | 261717 | 0.083973 | 0.24101 | 0.23386
1 | 1 | ANSLu | 1 | 20 | 8 | 261718 | 0.086366 | 0.2397 | 0.23379
1 | 1 | ANSLu | 1 | 12 | 11 | 261719 | 0.082406 | 0.22464 | 0.23299
Mean | [6 empty] | 0.085371 | 0.23346 | 0.23346
Median | [7 empty] | 0.23498 | 0.23354
SD | [7 empty] | 0.0098514 | 0.00052552
RSD | [7 empty] | 4.22% | 0.2251%
RSD Upscale Factor | [7 empty] | 0.053345
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 600c (Certified Value 0.221 ppm) | [30 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | 3
|  |  | 0.085g | 40g | 50g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 50g | 50g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 30g | 30g | 15g | 50g | 3
| 1 | 1 | 0.23627 | 0.21 | 0.23 | 0.213 | 0.223 | 0.219 | 0.22135 | 0.221 | 0.237 | 0.22201 | 0.22 | 0.222 | 0.206 | 0.198 | 0.218 | 0.219 | 0.225 | 0.22 | 0.219 | 0.23 | 0.222 | 0.212 | 0.21 | 0.24 | 0.2 | 0.206 | 0.22493 | 0.191 | 0.22 | 1
| 1 | 2 | 0.22495 | 0.21 | 0.235 | 0.212 | 0.22 | 0.226 | 0.21942 | 0.224 | 0.229 | 0.23462 | 0.221 | 0.229 | 0.207 | 0.227 | 0.221 | 0.224 | 0.228 | 0.222 | 0.208 | 0.23 | 0.22 | 0.229 | 0.2 | 0.23 | 0.2 | 0.217 | 0.23244 | 0.208 | 0.22 | 0
| 1 | 3 | 0.23368 | 0.21 | 0.2325 | 0.21 | 0.222 | 0.222 | 0.21837 | 0.23 | 0.233 | 0.22007 | 0.221 | 0.227 | 0.216 | 0.212 | 0.218 | 0.222 | 0.221 | 0.222 | 0.213 | 0.22 | 0.231 | 0.218 | 0.2 | 0.24 | 0.24 | 0.21 | 0.22191 | 0.201 | 0.22 | 16
| 1 | 4 | 0.23279 | 0.21 | 0.24 | 0.214 | 0.225 | 0.217 | 0.22242 | 0.22 | 0.237 | 0.21619 | 0.222 | 0.237 | 0.212 | 0.212 | 0.209 | 0.222 | 0.226 | 0.223 | 0.217 | 0.23 | 0.224 | 0.221 | 0.21 | 0.24 | 0.22 | 0.206 | 0.23093 | 0.202 | 0.21 | 0.22116
| 1 | 5 | 0.25028 | 0.22 | 0.23 | 0.205 | 0.221 | 0.238 | 0.21712 | 0.233 | 0.235 | 0.22685 | 0.221 | 0.231 | 0.211 | 0.229 | 0.219 | 0.222 | 0.227 | 0.22 | 0.208 | 0.22 | 0.222 | 0.224 | 0.22 | 0.23 | 0.21 | 0.206 | 0.22136 | 0.205 | 0.23 | 7
| 1 | 6 | 0.25457 | 0.22 | 0.2325 | 0.214 | 0.216 | 0.226 | 0.21903 | 0.229 | 0.241 | 0.22783 | 0.215 | 0.234 | 0.212 | 0.209 | 0.217 | 0.233 | 0.222 | 0.226 | 0.211 | 0.22 | 0.22 | 0.21 | 0.22 | 0.24 | 0.23 | 0.206 | 0.22193 | 0.204 | 0.23
|  | 7 | 00: 0.22041
|  | 8 | 00: 0.22568
|  | 9 | 00: 0.22212
|  | 10 | 00: 0.22318
|  | 11 | 00: 0.23698
|  | 12 | 00: 0.22464
|  | 13 | 00: 0.23681
|  | 14 | 00: 0.24101
|  | 15 | 00: 0.23872
|  | 16 | 00: 0.24243
|  | 17 | 00: 0.23909
|  | 18 | 00: 0.22533
|  | 19 | 00: 0.22057
|  | 20 | 00: 0.2397
| Mean |  | 0.23346 | 0.21333 | 0.23333 | 0.21133 | 0.22117 | 0.22467 | 0.21962 | 0.22617 | 0.23533 | 0.22459 | 0.22 | 0.23 | 0.21067 | 0.2145 | 0.217 | 0.22367 | 0.22483 | 0.22217 | 0.21267 | 0.225 | 0.22317 | 0.219 | 0.21 | 0.23667 | 0.21667 | 0.2085 | 0.22558 | 0.20183 | 0.22167
| Median |  | 0.23498 | 0.21 | 0.2325 | 0.2125 | 0.2215 | 0.224 | 0.21922 | 0.2265 | 0.236 | 0.22443 | 0.221 | 0.23 | 0.2115 | 0.212 | 0.218 | 0.222 | 0.2255 | 0.222 | 0.212 | 0.225 | 0.222 | 0.2195 | 0.21 | 0.24 | 0.215 | 0.206 | 0.22343 | 0.203 | 0.22
| Std Dev. |  | 0.0098514 | 0.005164 | 0.0037639 | 0.0034448 | 0.0030605 | 0.0074744 | 0.0019506 | 0.0052694 | 0.0040825 | 0.0065399 | 0.0025298 | 0.0052915 | 0.0036697 | 0.011675 | 0.0041473 | 0.0048442 | 0.0027869 | 0.0022286 | 0.0045898 | 0.0054772 | 0.0041191 | 0.0072111 | 0.0089443 | 0.005164 | 0.01633 | 0.0044609 | 0.0049144 | 0.0058452 | 0.0075277
| Rel.Std.Dev. |  | 4.22% | 2.421% | 1.613% | 1.63% | 1.384% | 3.327% | 0.8882% | 2.33% | 1.735% | 2.912% | 1.15% | 2.301% | 1.742% | 5.443% | 1.911% | 2.166% | 1.24% | 1.003% | 2.158% | 2.434% | 1.846% | 3.293% | 4.259% | 2.182% | 7.537% | 2.14% | 2.179% | 2.896% | 3.396%
| PDM3 |  | 5.562% | -3.539% | 5.505% | -4.443% | 0.00343% | 1.586% | -0.6974% | 2.264% | 6.409% | 1.554% | -0.5241% | 3.998% | -4.744% | -3.011% | -1.881% | 1.134% | 1.661% | 0.4556% | -3.84% | 1.737% | 0.9078% | -0.9763% | -5.046% | 7.012% | -2.031% | -5.724% | 2.001% | -8.738% | 0.2295%
| Z-Score (Absolute) |  | NA | 1.19 | 1.75 | 1.48 | 0.04 | 0.48 | 0.26 | 0.7 | 2.05 | 0.47 | 0.21 | 1.26 | 1.58 | 1.02 | 0.65 | 0.33 | 0.5 | 0.11 | 1.29 | 0.53 | 0.26 | 0.36 | 1.68 | 2.24 | 0.7 | 1.9 | 0.61 | 2.88 | 0.04
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 600c (Certified Value 0.21 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.23 | 2.24 | 2.26 | 2.27 | 2.28 | ppm
|  |  | INAA | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*OES | 3
|  |  | 0.085g | 30g | 50g | 15g | 15g | 25g | 30g | 25g | 30g | 30g | 10g | 25g | 25g | 25g | 25g | 25g | 10g | 25g | 10g | 50g | 50g | 25g | 50g | 25g | 3
| 1 | 1 | 0.23627 | 0.21 | 0.20315 | 0.208 | 0.2081 | 0.222 | 0.2 | 0.15 | 0.219 | 0.2 | 0.214 | 0.24 | 0.21 | 0.22 | 0.22 | 0.21 | 0.21 | 0.35 | 0.2 | 0.18 | 0.244 | 0.19289 | 0.17 | 0.26 | 1
| 1 | 2 | 0.22495 | 0.2 | 0.20016 | 0.1919 | 0.2119 | 0.224 | 0.199 | 0.17 | 0.218 | 0.2 | 0.229 | 0.24 | 0.21 | 0.22 | 0.21 | 0.22 | 0.19 | < 0.01 | 0.2 | 0.19 | 0.2396 | 0.18476 | 0.154 | 0.24 | 0
| 1 | 3 | 0.23368 | 0.2 | 0.24199 | 0.221 | 0.2122 | 0.222 | 0.207 | 0.14 | 0.213 | 0.2 | 0.193 | 0.19 | 0.21 | 0.23 | 0.22 | 0.21 | 0.18 | 0.19 | 0.2 | 0.2 | 0.2478 | 0.19112 | 0.183 | 0.25 | 16
| 1 | 4 | 0.23279 | 0.2 | 0.21211 | 0.2 | 0.2131 | 0.225 | 0.202 | 0.15 | 0.221 | 0.2 | 0.214 | 0.23 | 0.21 | 0.23 | 0.22 | 0.22 | 0.19 | 0.22 | 0.2 | 0.2 | 0.227 | 0.20748 | 0.158 | 0.25 | 0.21028
| 1 | 5 | 0.25028 | 0.21 | 0.20016 | 0.2121 | 0.2156 | 0.225 | 0.197 | 0.14 | 0.223 | 0.2 | 0.178 | 0.11 | 0.21 | 0.22 | 0.22 | 0.21 | 0.14 | 0.21 | 0.2 | 0.18 | 0.238 | 0.20598 | 0.158 | 0.26 | 9
| 1 | 6 | 0.25457 | 0.2 | 0.22107 | 0.204 | 0.2156 | 0.22 | 0.207 | 0.15 | 0.232 | 0.2 | 0.212 | 0.23 | 0.21 | 0.23 | 0.21 | 0.21 | 0.2 | 0.05 | 0.2 | 0.2 | 0.231 | 0.19551 | 0.191 | 0.23
|  | 7 | 00: 0.22041
|  | 8 | 00: 0.22568
|  | 9 | 00: 0.22212
|  | 10 | 00: 0.22318
|  | 11 | 00: 0.23698
|  | 12 | 00: 0.22464
|  | 13 | 00: 0.23681
|  | 14 | 00: 0.24101
|  | 15 | 00: 0.23872
|  | 16 | 00: 0.24243
|  | 17 | 00: 0.23909
|  | 18 | 00: 0.22533
|  | 19 | 00: 0.22057
|  | 20 | 00: 0.2397
| Mean |  | 0.23346 | 0.20333 | 0.21311 | 0.20617 | 0.21275 | 0.223 | 0.202 | 0.15 | 0.221 | 0.2 | 0.20667 | 0.20667 | 0.21 | 0.225 | 0.21667 | 0.21333 | 0.185 | 0.204 | 0.2 | 0.19167 | 0.2379 | 0.19629 | 0.169 | 0.24833
| Median |  | 0.23498 | 0.2 | 0.20763 | 0.206 | 0.21265 | 0.223 | 0.201 | 0.15 | 0.22 | 0.2 | 0.213 | 0.23 | 0.21 | 0.225 | 0.22 | 0.21 | 0.19 | 0.21 | 0.2 | 0.195 | 0.2388 | 0.1942 | 0.164 | 0.25
| Std Dev. |  | 0.0098514 | 0.005164 | 0.016327 | 0.010048 | 0.0027934 | 0.002 | 0.0041952 | 0.010954 | 0.0063561 | 3.0405e-17 | 0.018129 | 0.050859 | 0 | 0.0054772 | 0.005164 | 0.005164 | 0.02429 | 0.10668 | 3.0405e-17 | 0.0098319 | 0.0078044 | 0.0088401 | 0.015153 | 0.01169
| Rel.Std.Dev. |  | 4.22% | 2.54% | 7.661% | 4.874% | 1.313% | 0.8969% | 2.077% | 7.303% | 2.876% | 1.52e-14% | 8.772% | 24.61% | 0% | 2.434% | 2.383% | 2.421% | 13.13% | 52.29% | 1.52e-14% | 5.13% | 3.281% | 4.504% | 8.966% | 4.708%
| PDM3 |  | 11.02% | -3.302% | 1.346% | -1.955% | 1.176% | 6.05% | -3.936% | -28.67% | 5.099% | -4.888% | -1.717% | -1.717% | -0.1319% | 7.002% | 3.038% | 1.453% | -12.02% | -2.985% | -4.888% | -8.851% | 13.14% | -6.652% | -19.63% | 18.1%
| Z-Score (Absolute) |  | NA | 0.22 | 0.43 | 0.03 | 0.41 | 1.1 | 0.31 | 3.82 | 0.97 | 0.45 | 0 | 0 | 0.22 | 1.24 | 0.67 | 0.45 | 1.46 | 2.42 | 0.45 | 1.01 | 2.11 | 0.7 | 2.54 | 2.81
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 600c (Indicative Value 0.196 wt.%) | [26 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 2.17 | 2.18 | 2.27 | 3.01 | [21 empty] | wt.%
|  |  | IRC | IRC | IRC | IRC | [21 empty] | 3
3
| 1 | 1 | 0.26 | 0.23 | 0.12 | 0.24 | [21 empty] | 1
| 1 | 2 | 0.2 | 0.23 | 0.11 | 0.23 | [21 empty] | 1
| 1 | 3 | 0.21 | 0.23 | 0.11 | [22 empty] | 16
| 1 | 4 | 0.2 | 0.23 | 0.11 | [22 empty] | 0.196
| 1 | 5 | 0.21 | 0.23 | 0.1 | [22 empty] | 7
| 1 | 6 | 0.2 | 0.24 | 0.13
| Mean |  | 0.21333 | 0.23167 | 0.11333 | 0.235
| Median |  | 0.205 | 0.23 | 0.11 | 0.235
| Std Dev. |  | 0.023381 | 0.0040825 | 0.010328 | 0.0070711
| Rel.Std.Dev. |  | 10.96% | 1.762% | 9.113% | 3.009%
| PDM3 |  | 8.844% | 18.2% | -42.18% | 19.9%
| Z-Score (Absolute) |  | 0.57 | 0.57 | 6.79 | 0.78
| Analytical results for C-(Inorganic) in OREAS 600c (Indicative Value 0.195 wt.%) | [26 empty] | Indicative
C-(Inorganic) | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.18 | [24 empty] | wt.%
|  |  | I-550*IRC | [24 empty] | 3
3
| 1 | 1 | 0.19 | [24 empty] | 1
| 1 | 2 | 0.2 | [24 empty] | 1
| 1 | 3 | 0.19 | [24 empty] | 16
| 1 | 4 | 0.2 | [24 empty] | 0.195
| 1 | 5 | 0.19 | [24 empty] | 7
| 1 | 6 | 0.2
| Mean |  | 0.195
| Median |  | 0.195
| Std Dev. |  | 0.0054772
| Rel.Std.Dev. |  | 2.809%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for C-(Organic) in OREAS 600c (Indicative Value 0.038 wt.%) | [26 empty] | Indicative
C-(Organic) | Round | Replicate | Lab | [24 empty] | 1
| No | No | 2.18 | [24 empty] | wt.%
|  |  | HCl*IRC | [24 empty] | 3
3
| 1 | 1 | 0.04 | [24 empty] | 1
| 1 | 2 | 0.03 | [24 empty] | 1
| 1 | 3 | 0.04 | [24 empty] | 16
| 1 | 4 | 0.04 | [24 empty] | 0.038333
| 1 | 5 | 0.04 | [24 empty] | 7
| 1 | 6 | 0.04
| Mean |  | 0.038333
| Median |  | 0.04
| Std Dev. |  | 0.0040825
| Rel.Std.Dev. |  | 10.65%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 600c (Certified Value 0.232 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.31 | 0.304 | 0.24 | 0.24 | 0.23 | 0.23307 | 0.23 | 0.22 | 0.24 | 0.23 | 0.23 | 0.24 | 0.24 | 0.22 | 0.24 | 0.23 | 0.23 | 0.219 | 0.24 | 0.18 | 0.56 | 0.22 | 0.23877 | 0.24 | 0.17 | 1
| 1 | 2 | 0.3 | 0.31 | 0.24 | 0.24 | 0.23 | 0.21517 | 0.23 | 0.23 | 0.23 | 0.24 | 0.23 | 0.24 | 0.24 | 0.21 | 0.24 | 0.23 | 0.23 | 0.224 | 0.24 | 0.19 | 0.58 | 0.23 | 0.24434 | 0.24 | 0.15 | 11
| 1 | 3 | 0.32 | 0.3 | 0.23 | 0.23 | 0.23 | 0.23069 | 0.23 | 0.23 | 0.23 | 0.23 | 0.23 | 0.24 | 0.25 | 0.21 | 0.24 | 0.23 | 0.23 | 0.222 | 0.24 | 0.18 | 0.54 | 0.22 | 0.23686 | 0.24 | 0.16 | 16
| 1 | 4 | 0.3 | 0.293 | 0.23 | 0.24 | 0.22 | 0.22076 | 0.23 | 0.22 | 0.23 | 0.23 | 0.23 | 0.23 | 0.24 | 0.22 | 0.24 | 0.24 | 0.23 | 0.222 | 0.23 | 0.19 | 0.59 | 0.23 | 0.22669 | 0.24 | 0.16 | 0.23166
| 1 | 5 | 0.31 | 0.324 | 0.23 | 0.23 | 0.22 | 0.21875 | 0.23 | 0.23 | 0.24 | 0.24 | 0.24 | 0.23 | 0.23 | 0.22 | 0.24 | 0.23 | 0.24 | 0.226 | 0.23 | 0.19 | 0.57 | 0.22 | 0.24343 | 0.24 | 0.16 | 11
| 1 | 6 | 0.31 | 0.301 | 0.23 | 0.24 | 0.22 | 0.24579 | 0.23 | 0.22 | 0.24 | 0.23 | 0.23 | 0.25 | 0.24 | 0.21 | 0.24 | 0.24 | 0.23 | 0.224 | 0.23 | 0.18 | 0.59 | 0.22 | 0.22774 | 0.24 | 0.16
| Mean |  | 0.30833 | 0.30533 | 0.23333 | 0.23667 | 0.225 | 0.22737 | 0.23 | 0.225 | 0.235 | 0.23333 | 0.23167 | 0.23833 | 0.24 | 0.215 | 0.24 | 0.23333 | 0.23167 | 0.22283 | 0.235 | 0.185 | 0.57167 | 0.22333 | 0.2363 | 0.24 | 0.16
| Median |  | 0.31 | 0.3025 | 0.23 | 0.24 | 0.225 | 0.22572 | 0.23 | 0.225 | 0.235 | 0.23 | 0.23 | 0.24 | 0.24 | 0.215 | 0.24 | 0.23 | 0.23 | 0.223 | 0.235 | 0.185 | 0.575 | 0.22 | 0.23781 | 0.24 | 0.16
| Std Dev. |  | 0.0075277 | 0.01069 | 0.005164 | 0.005164 | 0.0054772 | 0.011397 | 0 | 0.0054772 | 0.0054772 | 0.005164 | 0.0040825 | 0.0075277 | 0.0063246 | 0.0054772 | 0 | 0.005164 | 0.0040825 | 0.0024014 | 0.0054772 | 0.0054772 | 0.019408 | 0.005164 | 0.0075836 | 0 | 0.0063246
| Rel.Std.Dev. |  | 2.441% | 3.501% | 2.213% | 2.182% | 2.434% | 5.013% | 0% | 2.434% | 2.331% | 2.213% | 1.762% | 3.158% | 2.635% | 2.548% | 0% | 2.213% | 1.762% | 1.078% | 2.331% | 2.961% | 3.395% | 2.312% | 3.209% | 0% | 3.953%
| PDM3 |  | 33.1% | 31.8% | 0.7229% | 2.162% | -2.874% | -1.852% | -0.716% | -2.874% | 1.442% | 0.7229% | 0.003451% | 2.881% | 3.601% | -7.191% | 3.601% | 0.7229% | 0.003451% | -3.81% | 1.442% | -20.14% | 146.8% | -3.594% | 2.005% | 3.601% | -30.93%
| Z-Score (Absolute) |  | 7.59 | 7.28 | 0 | 0.34 | 0.84 | 0.6 | 0.34 | 0.84 | 0.17 | 0 | 0.17 | 0.51 | 0.67 | 1.85 | 0.67 | 0 | 0.17 | 1.06 | 0.17 | 4.89 | 34.22 | 1.01 | 0.3 | 0.67 | 7.42
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 600c (Certified Value 23.6 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 24.3 | 25.954 | 24.4 | 22.87 | 23.4936 | 23.7 | 23.8 | 24.998 | 23.8 | 24.5 | 23.34 | 23.3 | 23.8 | 24.2 | 24.3 | 24 | 21.8 | 23.4 | 23.9 | 23.55 | 24.41 | 22.5641 | 22.2 | 24.6129 | 20.86 | 27.3 | 1
| 1 | 2 | 24.4 | 25.878 | 24.3 | 22.94 | 23.8118 | 23.8 | 23.8 | 24.221 | 23.9 | 24.2 | 22.89 | 23.8 | 23.4 | 24.5 | 22.8 | 24.3 | 21.6 | 23.2 | 24.6 | 23.49 | 23.25 | 21.8506 | 22.7 | 25.1305 | 20.98 | 26.8 | 22
| 1 | 3 | 24.5 | 26.182 | 23.7 | 23.61 | 23.8763 | 23.3 | 23.2 | 24.116 | 24.7 | 23.8 | 23.48 | 23.5 | 23.1 | 24.4 | 23.2 | 24.4 | 21.6 | 23.7 | 24.2 | 22.11 | 23.87 | 21.9151 | 22.5 | 25.2266 | 20.89 | 26.8 | 16
| 1 | 4 | 24.3 | 25.916 | 24.9 | 23.64 | 23.9301 | 24 | 22.9 | 24.6305 | 23 | 24.7 | 23.48 | 23.6 | 23.2 | 24.1 | 23.2 | 24.7 | 22 | 23.1 | 24.3 | 22.49 | 24.13 | 21.3483 | 21.5 | 24.8727 | 21.04 | 27 | 23.5996
| 1 | 5 | 24.4 | 26.6 | 24.5 | 23.35 | 23.6245 | 23.2 | 22.8 | 24.473 | 24.8 | 24.1 | 23.27 | 23.2 | 24 | 24.1 | 23.1 | 24.2 | 21.8 | 23.5 | 23.8 | 22.43 | 23.44 | 22.2179 | 22.6 | 24.7237 | 20.95 | 25.5 | 13
| 1 | 6 | 24.3 | 25.156 | 24.1 | 22.7 | 23.1302 | 24.4 | 23.8 | 23.612 | 24.3 | 23.7 | 24.23 | 23.3 | 22.7 | 25.2 | 23 | 24.5 | 21.8 | 24.2 | 23.2 | 24.2 | 23.42 | 21.7608 | 22.2 | 25.2213 | 21.5 | 26.7
| Mean |  | 24.3667 | 25.9477 | 24.3167 | 23.185 | 23.6444 | 23.7333 | 23.3833 | 24.3418 | 24.0833 | 24.1667 | 23.4483 | 23.45 | 23.3667 | 24.4167 | 23.2667 | 24.35 | 21.7667 | 23.5167 | 24 | 23.045 | 23.7533 | 21.9428 | 22.2833 | 24.9646 | 21.0367 | 26.6833
| Median |  | 24.35 | 25.935 | 24.35 | 23.145 | 23.7182 | 23.75 | 23.5 | 24.347 | 24.1 | 24.15 | 23.41 | 23.4 | 23.3 | 24.3 | 23.15 | 24.35 | 21.8 | 23.45 | 24.05 | 22.99 | 23.655 | 21.8829 | 22.35 | 25.0016 | 20.965 | 26.8
| Std Dev. |  | 0.08165 | 0.47182 | 0.40208 | 0.40233 | 0.30032 | 0.44572 | 0.47504 | 0.47526 | 0.66758 | 0.38816 | 0.44006 | 0.22583 | 0.4761 | 0.41673 | 0.52789 | 0.2429 | 0.15055 | 0.39707 | 0.4858 | 0.81824 | 0.45811 | 0.41413 | 0.43551 | 0.26541 | 0.23585 | 0.61779
| Rel.Std.Dev. |  | 0.3351% | 1.818% | 1.654% | 1.735% | 1.27% | 1.878% | 2.032% | 1.952% | 2.772% | 1.606% | 1.877% | 0.963% | 2.037% | 1.707% | 2.269% | 0.9975% | 0.6917% | 1.688% | 2.024% | 3.551% | 1.929% | 1.887% | 1.954% | 1.063% | 1.121% | 2.315%
| PDM3 |  | 3.25% | 9.949% | 3.038% | -1.757% | 0.1897% | 0.5665% | -0.9166% | 3.145% | 2.05% | 2.403% | -0.6412% | -0.6341% | -0.9872% | 3.462% | -1.411% | 3.18% | -7.767% | -0.3516% | 1.696% | -2.35% | 0.6512% | -7.021% | -5.578% | 5.784% | -10.86% | 13.07%
| Z-Score (Absolute) |  | 0.81 | 2.69 | 0.75 | 0.6 | 0.05 | 0.05 | 0.36 | 0.78 | 0.47 | 0.57 | 0.29 | 0.28 | 0.38 | 0.87 | 0.5 | 0.79 | 2.29 | 0.21 | 0.37 | 0.77 | 0.08 | 2.08 | 1.67 | 1.52 | 3.16 | 3.57
| Analytical results for Al in OREAS 600c (Certified Value 7.56 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 7.75 | 7.4 | 7.68 | 7.6895 | 7.5024 | 7.75 | 7.46 | 7.9372 | 7.66 | 7.1 | 8.1677 | 7.41 | 7.4 | 7.55 | 7.28 | 7.86 | 7.2536 | 6.47 | 8.11 | 7.5 | 7.78 | 7.3146 | 7.62 | 7.5285 | 7.4911 | 7.24 | 1
| 1 | 2 | 7.69 | 7.34 | 7.78 | 7.6624 | 7.4465 | 7.68 | 7.43 | 7.9339 | 7.64 | 7.05 | 8.0346 | 7.41 | 7.38 | 7.5 | 7.12 | 8.01 | 7.2228 | 7.21 | 8.39 | 7.56 | 7.58 | 6.9121 | 7.73 | 7.8601 | 7.6378 | 6.76 | 0
| 1 | 3 | 7.85 | 7.25 | 7.63 | 7.5837 | 7.5833 | 7.69 | 7.26 | 7.9817 | 7.78 | 7.06 | 7.942 | 7.45 | 7.45 | 7.56 | 7.23 | 7.92 | 7.2785 | 6.25 | 8.23 | 7.57 | 7.88 | 6.8772 | 7.46 | 7.7232 | 7.5767 | 6.92 | 16
| 1 | 4 | 7.76 | 7.19 | 7.8 | 7.5735 | 7.6081 | 7.78 | 7.23 | 7.9302 | 7.73 | 7.24 | 8.147 | 7.31 | 7.32 | 7.45 | 7.36 | 8.2 | 7.3322 | 4.93 | 8.34 | 7.51 | 7.89 | 7.0819 | 7.41 | 7.8474 | 7.5393 | 7.07 | 7.5602
| 1 | 5 | 7.83 | 7.66 | 7.82 | 7.825 | 7.6835 | 7.5 | 7.21 | 7.9667 | 7.76 | 7.12 | 7.983 | 7.31 | 7.61 | 7.48 | 7.36 | 7.84 | 7.1509 | 5.31 | 8.13 | 7.57 | 7.69 | 7.0917 | 7.52 | 7.6025 | 7.6201 | 6.71 | 14
| 1 | 6 | 7.69 | 6.85 | 7.74 | 7.5029 | 7.5166 | 7.78 | 7.26 | 7.9158 | 7.83 | 7.2 | 8.0803 | 7.28 | 7.29 | 7.7 | 7.3 | 8.01 | 7.325 | 5.97 | 7.95 | 7.58 | 7.41 | 7.0667 | 7.52 | 7.5089 | 7.5472 | 7.57
| Mean |  | 7.7617 | 7.2817 | 7.7417 | 7.6395 | 7.5567 | 7.6967 | 7.3083 | 7.9442 | 7.7333 | 7.1283 | 8.0591 | 7.3617 | 7.4083 | 7.54 | 7.275 | 7.9733 | 7.2605 | 6.0233 | 8.1917 | 7.5483 | 7.705 | 7.0574 | 7.5433 | 7.6784 | 7.5687 | 7.045
| Median |  | 7.755 | 7.295 | 7.76 | 7.6231 | 7.55 | 7.72 | 7.26 | 7.9355 | 7.745 | 7.11 | 8.0574 | 7.36 | 7.39 | 7.525 | 7.29 | 7.965 | 7.266 | 6.11 | 8.18 | 7.565 | 7.735 | 7.0743 | 7.52 | 7.6628 | 7.562 | 6.995
| Std Dev. |  | 0.067651 | 0.26694 | 0.073869 | 0.11275 | 0.085005 | 0.10558 | 0.10797 | 0.024805 | 0.072572 | 0.076529 | 0.089525 | 0.069976 | 0.11409 | 0.088769 | 0.090719 | 0.13231 | 0.06798 | 0.8205 | 0.16229 | 0.034303 | 0.1862 | 0.15595 | 0.11535 | 0.15526 | 0.054443 | 0.32353
| Rel.Std.Dev. |  | 0.8716% | 3.666% | 0.9542% | 1.476% | 1.125% | 1.372% | 1.477% | 0.3122% | 0.9384% | 1.074% | 1.111% | 0.9505% | 1.54% | 1.177% | 1.247% | 1.659% | 0.9363% | 13.62% | 1.981% | 0.4544% | 2.417% | 2.21% | 1.529% | 2.022% | 0.7193% | 4.592%
| PDM3 |  | 2.664% | -3.685% | 2.4% | 1.048% | -0.04675% | 1.804% | -3.332% | 5.079% | 2.289% | -5.713% | 6.598% | -2.627% | -2.009% | -0.2678% | -3.773% | 5.464% | -3.965% | -20.33% | 8.352% | -0.1576% | 1.915% | -6.652% | -0.2238% | 1.563% | 0.1118% | -6.815%
| Z-Score (Absolute) |  | 0.71 | 0.91 | 0.64 | 0.29 | 0.01 | 0.49 | 0.82 | 1.32 | 0.61 | 1.43 | 1.71 | 0.64 | 0.49 | 0.04 | 0.94 | 1.42 | 0.98 | 5.16 | 2.15 | 0.01 | 0.51 | 1.67 | 0.04 | 0.42 | 0.05 | 1.71
| Analytical results for As in OREAS 600c (Certified Value 39.5 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 42 | 41.9 | 24 | 38.1 | 38.9039 | 30.5 | 41.8 | 38.2 | 42 | 39 | 39.4 | 38 | 40.8 | 42.9 | 37.6 | 37.7 | 42 | 33 | 28.49 | 37.7 | 43.4709 | 58 | 38.5367 | 31.492 | 41 |  | 1
| 1 | 2 | 42 | 42.7 | 25 | 39 | 40.4695 | 30.8 | 43 | 38.9 | 42 | 38 | 39 | 36.7 | 41.6 | 39.9 | 37.8 | 40.8 | 42 | 39.8 | 27.53 | 36.5 | 34.3911 | 52 | 38.3273 | 30.206 | 43 |  | 12
| 1 | 3 | 43 | 42.6 | 25 | 39.9 | 40.8003 | 28.6 | 42.2 | 39.1 | 42 | 38 | 38.5 | 37.4 | 41.9 | 39.5 | 38.4 | 39.8 | 41 | 39.1 | 29.41 | 35.8 | 19.8321 | 49 | 38.3268 | 32.137 | 42 |  | 16
| 1 | 4 | 41 | 42.1 | 24 | 39.4 | 39.9393 | 26.9 | 40.2 | 38.6 | 42 | 38 | 39.1 | 36.8 | 40.9 | 40.3 | 38.9 | 41.2 | 39 | 37.7 | 30.22 | 36.8 | 48.9508 | 44 | 38.5367 | 27.005 | 42 |  | 39.4596
| 1 | 5 | 42 | 42.9 | 25 | 39.8 | 41.3039 | 29.9 | 41.2 | 39.1 | 42 | 39 | 38.9 | 37.5 | 38.9 | 41.1 | 38.5 | 41.1 | 42 | 44.5 | 29.09 | 35.5 | 34.8473 | 49 | 38.2391 | 31.308 | 43 |  | 15
| 1 | 6 | 41 | 41 | 25 | 38.6 | 39.4327 | 29.5 | 43 | 38.9 | 41 | 38 | 39.6 | 35.4 | 43.3 | 39.9 | 38.3 | 38.1 | 44 | 37.9 | 28.76 | 35.6 | 60.2529 | 44 | 38.6183 | 31.641 | 42
| Mean |  | 41.8333 | 42.2 | 24.6667 | 39.1333 | 40.1416 | 29.3667 | 41.9 | 38.8 | 41.8333 | 38.3333 | 39.0833 | 36.9667 | 41.2333 | 40.6 | 38.25 | 39.7833 | 41.6667 | 38.6667 | 28.9167 | 36.3167 | 40.2908 | 49.3333 | 38.4308 | 30.6315 | 42.1667
| Median |  | 42 | 42.35 | 25 | 39.2 | 40.2044 | 29.7 | 42 | 38.9 | 42 | 38 | 39.05 | 37.1 | 41.25 | 40.1 | 38.35 | 40.3 | 42 | 38.5 | 28.925 | 36.15 | 39.1591 | 49 | 38.432 | 31.4 | 42
| Std Dev. |  | 0.75277 | 0.69857 | 0.5164 | 0.70333 | 0.89082 | 1.4362 | 1.0863 | 0.34641 | 0.40825 | 0.5164 | 0.38687 | 0.9048 | 1.4556 | 1.2506 | 0.47645 | 1.5459 | 1.633 | 3.7184 | 0.90571 | 0.85186 | 13.9039 | 5.2789 | 0.1522 | 1.8877 | 0.75277
| Rel.Std.Dev. |  | 1.799% | 1.655% | 2.094% | 1.797% | 2.219% | 4.891% | 2.593% | 0.8928% | 0.9759% | 1.347% | 0.9899% | 2.448% | 3.53% | 3.08% | 1.246% | 3.886% | 3.919% | 9.617% | 3.132% | 2.346% | 34.51% | 10.7% | 0.396% | 6.162% | 1.785%
| PDM3 |  | 6.016% | 6.945% | -37.49% | -0.8268% | 1.728% | -25.58% | 6.185% | -1.672% | 6.016% | -2.854% | -0.9535% | -6.318% | 4.495% | 2.89% | -3.065% | 0.8205% | 5.593% | -2.009% | -26.72% | -7.965% | 2.107% | 25.02% | -2.607% | -22.37% | 6.86%
| Z-Score (Absolute) |  | 0.84 | 0.95 | 4.5 | 0 | 0.31 | 3.04 | 0.86 | 0.1 | 0.84 | 0.25 | 0.02 | 0.67 | 0.65 | 0.46 | 0.27 | 0.2 | 0.79 | 0.15 | 3.18 | 0.88 | 0.36 | 3.17 | 0.22 | 2.65 | 0.94
| Analytical results for B in OREAS 600c (Indicative Value 14.8 ppm) | [27 empty] | Indicative
B | Round | Replicate | 2.01: Lab; 2.03: Lab |  | 1
| No | No | 2.01: 2.01; 2.03: 2.23 |  | ppm
|  |  | 2.01: 4A*OES; 2.03: 4A*OES |  | 1
1
| 1 | 1 | 2.01: 15; 2.03: < 0.1 |  | 1
| 1 | 2 | 2.01: 15; 2.03: < 0.1 |  | 3
| 1 | 3 | 2.01: 15; 2.03: < 0.1 |  | 16
| 1 | 4 | 2.01: 15; 2.03: < 0.1 |  | 14.8333
| 1 | 5 | 2.01: 15; 2.03: < 0.1 |  | 9
| 1 | 6 | 2.01: 14; 2.03: < 0.1
| Mean |  | 2.01: 14.8333
| Median |  | 2.01: 15
| Std Dev. |  | 2.01: 0.40825
| Rel.Std.Dev. |  | 2.01: 2.752%
| PDM3 |  | 2.01: 2.22e-13%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.67
| Analytical results for Ba in OREAS 600c (Certified Value 3039 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
0
| 1 | 1 | 2864 | 2563 | 2217 | 2911.7 | 2914.9724 | 3220 | 3290 | 3180 | 3050 | 3220 | 2990 | 2900 | 3140 | 2991 | 3041 | 2910 | 2926 | 3111 | 2597.7884 | 2920 | 3178.1984 | 2506.358 | 2955 |  |  |  | 1
| 1 | 2 | 2928 | 2204 | 2555 | 2906.1 | 2929.3232 | 3270 | 3350 | 3190 | 3030 | 3190 | 3000 | 2860 | 3200 | 2987 | 3055 | 2840 | 2955 | 3065 | 2535.0514 | 2970 | 3270.4632 | 2725.915 | 2948 |  |  |  | 24
| 1 | 3 | 2939 | 2536 | 3035 | 2943.7 | 2991.0418 | 3250 | 3100 | 3150 | 3050 | 3210 | 3010 | 2890 | 3180 | 2977 | 2998 | 2940 | 2926 | 3023 | 2543.6252 | 2900 | 3276.6195 | 2670.168 | 2920 |  |  |  | 16
| 1 | 4 | 2841 | 2466 | 1689 | 2999.2 | 2962.2003 | 3340 | 3190 | 3180 | 3020 | 3180 | 2990 | 2910 | 3300 | 3015 | 2901 | 2890 | 2945 | 2979 | 2519.3099 | 2850 | 3214.3062 | 2634.085 | 2879 |  |  |  | 3038.8424
| 1 | 5 | 2833 | 2053 | 2844 | 3010.4 | 3063.4654 | 3260 | 3210 | 3190 | 3020 | 3240 | 2970 | 2900 | 3140 | 2955 | 2981 | 2890 | 2924 | 3062 | 2560.5792 | 2910 | 3180.2177 | 2744.438 | 2911 |  |  |  | 16
| 1 | 6 | 2901 | 2334 | 3043 | 3033.2 | 2880.8906 | 3240 | 3300 | 3130 | 3010 | 3150 | 3070 | 2890 | 3200 | 3002 | 3017 | 2940 | 2928 | 3121 | 2469.7205 | 2900 | 3216.0324 | 2596.321 | 2932
| Mean |  | 2884.3333 | 2359.3333 | 2563.8333 | 2967.3833 | 2956.9823 | 3263.3333 | 3240 | 3170 | 3030 | 3198.3333 | 3005 | 2891.6667 | 3193.3333 | 2987.8333 | 2998.8333 | 2901.6667 | 2934 | 3060.1667 | 2537.6791 | 2908.3333 | 3222.6396 | 2646.2142 | 2924.1667
| Median |  | 2882.5 | 2400 | 2699.5 | 2971.45 | 2945.7618 | 3255 | 3250 | 3180 | 3025 | 3200 | 2995 | 2895 | 3190 | 2989 | 3007.5 | 2900 | 2927 | 3063.5 | 2539.3383 | 2905 | 3215.1693 | 2652.1265 | 2926
| Std Dev. |  | 44.9429 | 201.3352 | 532.1272 | 54.0588 | 64.5689 | 41.3118 | 90.7744 | 24.4949 | 16.7332 | 31.8852 | 34.4964 | 17.224 | 58.8784 | 20.7115 | 55.0506 | 37.6386 | 12.853 | 53.4431 | 42.7448 | 38.6868 | 42.6349 | 88.0569 | 27.6074
| Rel.Std.Dev. |  | 1.558% | 8.534% | 20.76% | 1.822% | 2.184% | 1.266% | 2.802% | 0.7727% | 0.5523% | 0.9969% | 1.148% | 0.5956% | 1.844% | 0.6932% | 1.836% | 1.297% | 0.4381% | 1.746% | 1.684% | 1.33% | 1.323% | 3.328% | 0.9441%
| PDM3 |  | -5.084% | -22.36% | -15.63% | -2.352% | -2.694% | 7.387% | 6.62% | 4.316% | -0.291% | 5.248% | -1.114% | -4.843% | 5.084% | -1.679% | -1.317% | -4.514% | -3.45% | 0.7017% | -16.49% | -4.295% | 6.048% | -12.92% | -3.774%
| Z-Score (Absolute) |  | 0.67 | 4.94 | 3.28 | 0 | 0.08 | 2.4 | 2.21 | 1.65 | 0.51 | 1.88 | 0.31 | 0.61 | 1.83 | 0.17 | 0.26 | 0.53 | 0.27 | 0.75 | 3.49 | 0.48 | 2.07 | 2.61 | 0.35
| Analytical results for Be in OREAS 600c (Certified Value 3.14 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 3.37 | 3 | 3 | 3.01 | 3.1291 | 3.4 | 2.95 | 4.2329 | 2.8 | 3.2 | 3.33 | 3.09 | 3.2 | 3 | 2.93 | 3.17 | 3.16 | 3.61 | 3.1 | 3.18 | 3.48 | 2.6223 | 3.5 | 3.0629 | 2.454 | 2.91 | 1
| 1 | 2 | 3.37 | 3 | 3 | 2.89 | 3.1065 | 3.1 | 2.94 | 4.2463 | 3 | 3.2 | 3.15 | 3.04 | 3.17 | 3.01 | 2.89 | 3.22 | 3.17 | 3.71 | 3.15 | 3.28 | 3.39 | 2.6788 | 3.6 | 3.3196 | 2.501 | 2.83 | 25
| 1 | 3 | 3.18 | 3 | 3 | 3.03 | 3.1454 | 3.4 | 2.76 | 4.2225 | 3.1 | 3.2 | 3.22 | 3.11 | 3.22 | 3.02 | 2.94 | 3.17 | 3.13 | 3.55 | 3.43 | 3.22 | 3.24 | 2.6014 | 3.5 | 3.1853 | 2.516 | 3.15 | 16
| 1 | 4 | 3.15 | 3 | 3 | 3.03 | 3.1146 | 3.7 | 2.76 | 4.2768 | 2.9 | 3.2 | 3.18 | 3.04 | 3.19 | 2.99 | 2.97 | 3.28 | 3.21 | 3.22 | 3.06 | 3.28 | 3.5 | 2.5443 | 3.4 | 3.3563 | 2.527 | 3.06 | 3.1439
| 1 | 5 | 3.1 | 3 | 3 | 3.13 | 3.1607 | 3.5 | 2.85 | 4.1993 | 2.9 | 3.2 | 3.22 | 3.06 | 3.25 | 2.97 | 2.94 | 3.16 | 3.14 | 3.51 | 2.92 | 3.24 | 3.32 | 2.6715 | 3.4 | 3.1098 | 2.501 | 2.86 | 17
| 1 | 6 | 3.04 | 3 | 3 | 2.8 | 3.0961 | 3.7 | 2.9 | 4.3303 | 2.9 | 3.2 | 3.11 | 3.03 | 3.14 | 3.08 | 2.93 | 3.21 | 3.19 | 3.6 | 3.79 | 3.33 | 3.46 | 2.5039 | 3.4 | 3.1039 | 2.507 | 3.09
| Mean |  | 3.2017 | 3 | 3 | 2.9817 | 3.1254 | 3.4667 | 2.86 | 4.2514 | 2.9333 | 3.2 | 3.2017 | 3.0617 | 3.195 | 3.0117 | 2.9333 | 3.2017 | 3.1667 | 3.5333 | 3.2417 | 3.255 | 3.3983 | 2.6037 | 3.4667 | 3.1896 | 2.501 | 2.9833
| Median |  | 3.165 | 3 | 3 | 3.02 | 3.1219 | 3.45 | 2.875 | 4.2396 | 2.9 | 3.2 | 3.2 | 3.05 | 3.195 | 3.005 | 2.935 | 3.19 | 3.165 | 3.575 | 3.125 | 3.26 | 3.425 | 2.6119 | 3.45 | 3.1476 | 2.504 | 2.985
| Std Dev. |  | 0.13877 | 0 | 0 | 0.11737 | 0.02446 | 0.22509 | 0.085088 | 0.046428 | 0.10328 | 4.8648e-16 | 0.075741 | 0.031885 | 0.038341 | 0.037639 | 0.02582 | 0.04535 | 0.030111 | 0.16765 | 0.31657 | 0.05282 | 0.10206 | 0.069392 | 0.08165 | 0.12202 | 0.025084 | 0.13352
| Rel.Std.Dev. |  | 4.334% | 0% | 0% | 3.937% | 0.7826% | 6.493% | 2.975% | 1.092% | 3.521% | 1.52e-14% | 2.366% | 1.041% | 1.2% | 1.25% | 0.8802% | 1.416% | 0.9509% | 4.745% | 9.766% | 1.623% | 3.003% | 2.665% | 2.355% | 3.826% | 1.003% | 4.475%
| PDM3 |  | 1.838% | -4.576% | -4.576% | -5.159% | -0.5875% | 10.27% | -9.029% | 35.23% | -6.697% | 1.785% | 1.838% | -2.615% | 1.626% | -4.205% | -6.697% | 1.838% | 0.7252% | 12.39% | 3.111% | 3.535% | 8.094% | -17.18% | 10.27% | 1.456% | -20.45% | -5.106%
| Z-Score (Absolute) |  | 0.03 | NA | NA | 0.73 | 0.23 | 0.96 | 1.16 | 3.69 | 0.9 | 0.03 | 0.03 | 0.46 | 0.01 | 0.63 | 0.9 | 0.03 | 0.09 | 1.19 | 0.17 | 0.22 | 0.72 | 2.05 | 0.96 | 0.01 | 2.41 | 0.73
Results from laboratories 2.03 and 2.04 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 600c (Certified Value 0.77 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 0.72 | 0.79 | 0.8 | 0.83 | 0.77021 | 0.79 | 0.79 | 0.74 | 0.84 | 0.77 | 0.77 | 0.82 | 0.74 | 0.76 | 0.69 | 0.78 | 0.82 | 1.04 | < 5 | 0.73 | < 0.1 | 0.75636 | < 5 | 0.57 |  |  | 1
| 1 | 2 | 0.69 | 0.67 | 0.8 | 0.83 | 0.77237 | 0.8 | 0.79 | 0.76 | 0.8 | 0.77 | 0.78 | 0.82 | 0.77 | 0.72 | 0.69 | 0.77 | 0.85 | 0.88 | < 5 | 0.66 | < 0.1 | 0.80476 | < 5 | 0.55 |  |  | 26
| 1 | 3 | 0.74 | 0.77 | 0.8 | 0.82 | 0.79122 | 0.8 | 0.79 | 0.73 | 0.84 | 0.8 | 0.77 | 0.81 | 0.77 | 0.72 | 0.7 | 0.82 | 0.82 | 0.96 | < 5 | 0.69 | < 0.1 | 0.80327 | < 5 | 0.58 |  |  | 16
| 1 | 4 | 0.7 | 0.79 | 0.8 | 0.85 | 0.77078 | 0.83 | 0.78 | 0.77 | 0.85 | 0.78 | 0.78 | 0.81 | 0.76 | 0.73 | 0.71 | 0.92 | 0.76 | 0.92 | < 5 | 0.69 | < 0.1 | 0.78714 | < 5 | 0.58 |  |  | 0.77235
| 1 | 5 | 0.7 | 0.76 | 0.7 | 0.9 | 0.79625 | 0.8 | 0.83 | 0.74 | 0.82 | 0.8 | 0.76 | 0.8 | 0.76 | 0.74 | 0.72 | 0.79 | 0.76 | 0.97 | < 5 | 0.7 | < 0.1 | 0.81318 | < 5 | 0.56 |  |  | 18
| 1 | 6 | 0.78 | 0.72 | 0.8 | 0.77 | 0.78649 | 0.83 | 1.05 | 0.74 | 0.83 | 0.77 | 0.79 | 0.8 | 0.76 | 0.74 | 0.7 | 0.78 | 0.79 | 0.92 | < 5 | 0.71 | < 0.1 | 0.75815 | < 5 | 0.55
| Mean |  | 0.72167 | 0.75 | 0.78333 | 0.83333 | 0.78122 | 0.80833 | 0.83833 | 0.74667 | 0.83 | 0.78167 | 0.775 | 0.81 | 0.76 | 0.735 | 0.70167 | 0.81 | 0.8 | 0.94833 |  | 0.69667 |  | 0.78714 |  | 0.565
| Median |  | 0.71 | 0.765 | 0.8 | 0.83 | 0.77943 | 0.8 | 0.79 | 0.74 | 0.835 | 0.775 | 0.775 | 0.81 | 0.76 | 0.735 | 0.7 | 0.785 | 0.805 | 0.94 |  | 0.695 |  | 0.79521 |  | 0.565
| Std Dev. |  | 0.033714 | 0.046904 | 0.040825 | 0.042269 | 0.011509 | 0.017224 | 0.10515 | 0.015055 | 0.017889 | 0.01472 | 0.010488 | 0.0089443 | 0.010954 | 0.015166 | 0.01169 | 0.056569 | 0.036332 | 0.055287 |  | 0.023381 |  | 0.024641 |  | 0.013784
| Rel.Std.Dev. |  | 4.672% | 6.254% | 5.212% | 5.072% | 1.473% | 2.131% | 12.54% | 2.016% | 2.155% | 1.883% | 1.353% | 1.104% | 1.441% | 2.063% | 1.666% | 6.984% | 4.541% | 5.83% |  | 3.356% |  | 3.13% |  | 2.44%
| PDM3 |  | -6.563% | -2.894% | 1.422% | 7.895% | 1.148% | 4.658% | 8.543% | -3.326% | 7.464% | 1.206% | 0.3427% | 4.874% | -1.599% | -4.836% | -9.152% | 4.874% | 3.58% | 22.78% |  | -9.8% |  | 1.915% |  | -26.85%
| Z-Score (Absolute) |  | 0.87 | 0.46 | NA | 0.75 | 0.01 | 0.39 | 0.82 | 0.51 | 0.7 | 0 | 0.1 | 0.41 | 0.31 | 0.67 | 1.16 | 0.41 | 0.26 | 2.41 | 24.83 | 1.23 | 10.57 | 0.08 | 24.83 | 3.13
Results from laboratory 2.04 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 600c (Certified Value 1.25 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.29 | 1.31 | 1.25 | 1.2558 | 1.2345 | 1.27 | 1.26 | 1.2862 | 1.2 | 1.24 | 1.2821 | 1.28 | 1.22 | 1.22 | 1.21 | 1.25 | 1.2783 | 1.171 | 1.29 | 1.21 | 1.21 | 1.2634 | 1.258 | 1.1856 | 1.2145 | 1.246 | 1
| 1 | 2 | 1.29 | 1.3 | 1.27 | 1.2583 | 1.2107 | 1.25 | 1.24 | 1.2604 | 1.24 | 1.21 | 1.2616 | 1.28 | 1.22 | 1.22 | 1.17 | 1.27 | 1.2742 | 1.241 | 1.27 | 1.23 | 1.22 | 1.2122 | 1.279 | 1.2909 | 1.2213 | 1.178 | 0
| 1 | 3 | 1.3 | 1.3 | 1.25 | 1.2445 | 1.2407 | 1.31 | 1.21 | 1.2495 | 1.21 | 1.22 | 1.2449 | 1.28 | 1.24 | 1.23 | 1.22 | 1.26 | 1.275 | 1.167 | 1.3 | 1.21 | 1.2 | 1.2408 | 1.244 | 1.2376 | 1.1947 | 1.19 | 16
| 1 | 4 | 1.29 | 1.31 | 1.27 | 1.2413 | 1.2267 | 1.31 | 1.23 | 1.2838 | 1.19 | 1.25 | 1.2782 | 1.25 | 1.22 | 1.2 | 1.24 | 1.31 | 1.2932 | 1.029 | 1.3 | 1.26 | 1.2 | 1.2516 | 1.229 | 1.2882 | 1.2031 | 1.227 | 1.2465
| 1 | 5 | 1.33 | 1.36 | 1.27 | 1.2876 | 1.2454 | 1.26 | 1.23 | 1.2734 | 1.23 | 1.23 | 1.2506 | 1.26 | 1.26 | 1.21 | 1.21 | 1.24 | 1.2611 | 1.077 | 1.31 | 1.24 | 1.19 | 1.2591 | 1.251 | 1.2389 | 1.2011 | 1.182 | 19
| 1 | 6 | 1.29 | 1.23 | 1.26 | 1.2323 | 1.229 | 1.28 | 1.21 | 1.244 | 1.21 | 1.2 | 1.2658 | 1.25 | 1.22 | 1.26 | 1.21 | 1.26 | 1.2866 | 1.139 | 1.29 | 1.22 | 1.2 | 1.2547 | 1.244 | 1.2218 | 1.2237 | 1.266
| Mean |  | 1.2983 | 1.3017 | 1.2617 | 1.2533 | 1.2312 | 1.28 | 1.23 | 1.2662 | 1.2133 | 1.225 | 1.2639 | 1.2667 | 1.23 | 1.2233 | 1.21 | 1.265 | 1.2781 | 1.1373 | 1.2933 | 1.2283 | 1.2033 | 1.247 | 1.2508 | 1.2438 | 1.2097 | 1.2148
| Median |  | 1.29 | 1.305 | 1.265 | 1.2502 | 1.2318 | 1.275 | 1.23 | 1.2669 | 1.21 | 1.225 | 1.2637 | 1.27 | 1.22 | 1.22 | 1.21 | 1.26 | 1.2767 | 1.153 | 1.295 | 1.225 | 1.2 | 1.2531 | 1.2475 | 1.2382 | 1.2088 | 1.2085
| Std Dev. |  | 0.016021 | 0.041673 | 0.0098319 | 0.01934 | 0.012244 | 0.025298 | 0.018974 | 0.017715 | 0.018619 | 0.018708 | 0.014714 | 0.015055 | 0.016733 | 0.020656 | 0.022804 | 0.02429 | 0.011083 | 0.075019 | 0.013663 | 0.019408 | 0.010328 | 0.018683 | 0.016822 | 0.040329 | 0.011811 | 0.036848
| Rel.Std.Dev. |  | 1.234% | 3.202% | 0.7793% | 1.543% | 0.9945% | 1.976% | 1.543% | 1.399% | 1.535% | 1.527% | 1.164% | 1.189% | 1.36% | 1.688% | 1.885% | 1.92% | 0.8671% | 6.596% | 1.056% | 1.58% | 0.8583% | 1.498% | 1.345% | 3.242% | 0.9763% | 3.033%
| PDM3 |  | 4.158% | 4.425% | 1.216% | 0.5451% | -1.231% | 2.687% | -1.324% | 1.58% | -2.661% | -1.725% | 1.393% | 1.617% | -1.324% | -1.859% | -2.929% | 1.484% | 2.532% | -8.758% | 3.757% | -1.458% | -3.463% | 0.03701% | 0.3472% | -0.215% | -2.95% | -2.541%
| Z-Score (Absolute) |  | 1.7 | 1.8 | 0.52 | 0.25 | 0.46 | 1.11 | 0.49 | 0.67 | 1.03 | 0.65 | 0.59 | 0.68 | 0.49 | 0.71 | 1.13 | 0.63 | 1.05 | 3.46 | 1.54 | 0.55 | 1.35 | 0.05 | 0.17 | 0.05 | 1.14 | 0.98
| Analytical results for Cd in OREAS 600c (Certified Value 6.33 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 7.35 | 6.94 | 5.9 | 6.08 | 6.054 | 7 | 6.53 | 9.3552 | 6.41 | 6.4 | 6.55 | 6.84 | 6.52 | 6.83 | 6.54 | 6.39 | 5.99 | 6.03 | 6.5 | 5.95 | 6.6 | 5.7123 | 5 | 5.834 | 5.545 | 6.65 | 1
| 1 | 2 | 6.63 | 6.09 | 6.3 | 6.41 | 6.2365 | 7.1 | 6.62 | 9.3472 | 6.67 | 6.44 | 6.45 | 6.93 | 6.25 | 7 | 6.19 | 6.25 | 5.94 | 6.22 | 6.7 | 5.84 | 6.51 | 5.5741 | 6 | 5.8262 | 5.671 | 6.11 | 27
| 1 | 3 | 6.68 | 6.48 | 6 | 6.34 | 6.3835 | 7 | 6.36 | 9.3889 | 6.71 | 6.33 | 6.52 | 6.87 | 6.33 | 6.94 | 6.11 | 6.45 | 5.95 | 6.13 | 6.5 | 6.32 | 6.75 | 5.8762 | 5 | 6.0248 | 5.783 | 5.95 | 16
| 1 | 4 | 6.32 | 6.26 | 6 | 6.47 | 6.2773 | 7.2 | 6.32 | 9.4371 | 6.42 | 6.55 | 6.59 | 6.92 | 6.38 | 6.81 | 6.27 | 6.57 | 6.05 | 6.05 | 6.7 | 5.81 | 6.66 | 5.3388 | 5 | 6.0946 | 5.701 | 6.16 | 6.3328
| 1 | 5 | 6.42 | 6.43 | 6.1 | 6.65 | 6.3367 | 6.9 | 6.39 | 9.2734 | 6.28 | 6.31 | 6.41 | 6.82 | 6.44 | 6.9 | 6.4 | 6.32 | 6.02 | 6.05 | 6.5 | 5.78 | 6.57 | 5.5895 | 5 | 6.0115 | 5.653 | 6.14 | 20
| 1 | 6 | 6.5 | 6.06 | 6.3 | 6.35 | 6.0594 | 7.3 | 6.49 | 9.333 | 6.55 | 6.43 | 6.75 | 6.75 | 6.09 | 7.21 | 6.23 | 6.36 | 6.11 | 6.13 | 6.1 | 6.12 | 6.48 | 5.5367 | 5 | 6.1494 | 5.703 | 6.29
| Mean |  | 6.65 | 6.3767 | 6.1 | 6.3833 | 6.2246 | 7.0833 | 6.4517 | 9.3558 | 6.5067 | 6.41 | 6.545 | 6.855 | 6.335 | 6.9483 | 6.29 | 6.39 | 6.01 | 6.1017 | 6.5 | 5.97 | 6.595 | 5.6046 | 5.1667 | 5.9901 | 5.676 | 6.2167
| Median |  | 6.565 | 6.345 | 6.05 | 6.38 | 6.2569 | 7.05 | 6.44 | 9.3512 | 6.485 | 6.415 | 6.535 | 6.855 | 6.355 | 6.92 | 6.25 | 6.375 | 6.005 | 6.09 | 6.5 | 5.895 | 6.585 | 5.5818 | 5 | 6.0181 | 5.686 | 6.15
| Std Dev. |  | 0.36759 | 0.32463 | 0.16733 | 0.18673 | 0.13939 | 0.1472 | 0.11444 | 0.054935 | 0.16621 | 0.086487 | 0.11996 | 0.067157 | 0.15136 | 0.14607 | 0.15556 | 0.11082 | 0.064187 | 0.072226 | 0.21909 | 0.21166 | 0.099348 | 0.17988 | 0.40825 | 0.13355 | 0.078126 | 0.23864
| Rel.Std.Dev. |  | 5.528% | 5.091% | 2.743% | 2.925% | 2.239% | 2.078% | 1.774% | 0.5872% | 2.554% | 1.349% | 1.833% | 0.9797% | 2.389% | 2.102% | 2.473% | 1.734% | 1.068% | 1.184% | 3.371% | 3.545% | 1.506% | 3.21% | 7.902% | 2.23% | 1.376% | 3.839%
| PDM3 |  | 5.009% | 0.6929% | -3.676% | 0.7982% | -1.709% | 11.85% | 1.877% | 47.74% | 2.746% | 1.219% | 3.351% | 8.246% | 0.03494% | 9.72% | -0.6756% | 0.9034% | -5.097% | -3.65% | 2.64% | -5.729% | 4.141% | -11.5% | -18.41% | -5.412% | -10.37% | -1.834%
| Z-Score (Absolute) |  | 0.85 | 0.02 | 0.9 | 0 | 0.51 | 2.23 | 0.22 | 9.47 | 0.39 | 0.08 | 0.52 | 1.5 | 0.15 | 1.8 | 0.3 | 0.02 | 1.19 | 0.9 | 0.37 | 1.32 | 0.67 | 2.48 | NA | 1.25 | 2.25 | 0.53
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 600c (Certified Value 91 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.27 | 2.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 0
1
| 1 | 1 | 95 | 96.97 | 96 | 88.52 | 89.3872 | 104 | 100 | 115.49 | 89.8 | 94 | 88.2 | 89.7 | 86.7 | 91 | 99.1 | 84.8 | 73.49 | 94.6 | 89 | 87.4352 | 87 | 78.098 | 84.45 |  |  |  | 1
| 1 | 2 | 92 | 89.98 | 98 | 90.34 | 89.253 | 101 | 99.4 | 115.94 | 88.9 | 94 | 92.4 | 88.2 | 89.4 | 86.6 | 96.9 | 84.8 | 80.31 | 91 | 91 | 86.4351 | 88 | 78.627 | 86.58 |  |  |  | 28
| 1 | 3 | 98 | 96.31 | 94 | 90.56 | 90.2716 | 100 | 101 | 115.8 | 89.2 | 93 | 86.5 | 89.3 | 89.7 | 86.9 | 99.6 | 84.7 | 70.53 | 95.4 | 89 | 86.9635 | 86 | 80.098 | 81.22 |  |  |  | 16
| 1 | 4 | 93 | 95.15 | 99 | 91.99 | 90.3338 | 103 | 99.1 | 115.7 | 90.4 | 98 | 90 | 88.1 | 87.5 | 92.9 | 102 | 86.5 | 46.82 | 93.1 | 89 | 88.5848 | 84 | 76.361 | 83.72 |  |  |  | 90.9624
| 1 | 5 | 88 | 91.23 | 97 | 93.34 | 90.3063 | 102 | 103 | 115.62 | 89.5 | 95 | 86.6 | 91 | 86.6 | 91.6 | 99.5 | 85.8 | 56.17 | 94.1 | 89 | 90.9838 | 84 | 79.742 | 82.42 |  |  |  | 21
| 1 | 6 | 96 | 86.21 | 95 | 89.03 | 89.2544 | 102 | 101 | 115.06 | 91.3 | 96 | 88.9 | 87.1 | 92.7 | 88.5 | 98.2 | 85.2 | 63.51 | 94.8 | 89 | 86.9046 | 84 | 80.697 | 83.02
| Mean |  | 93.6667 | 92.6417 | 96.5 | 90.63 | 89.8011 | 102 | 100.5833 | 115.6017 | 89.85 | 95 | 88.7667 | 88.9 | 88.7667 | 89.5833 | 99.2167 | 85.3 | 65.1383 | 93.8333 | 89.3333 | 87.8845 | 85.5 | 78.9372 | 83.5683
| Median |  | 94 | 93.19 | 96.5 | 90.45 | 89.8294 | 102 | 100.5 | 115.66 | 89.65 | 94.5 | 88.55 | 88.75 | 88.45 | 89.75 | 99.3 | 85 | 67.02 | 94.35 | 89 | 87.1993 | 85 | 79.1845 | 83.37
| Std Dev. |  | 3.5024 | 4.2173 | 1.8708 | 1.8053 | 0.55336 | 1.4142 | 1.4233 | 0.30649 | 0.87807 | 1.7889 | 2.2313 | 1.3842 | 2.3372 | 2.621 | 1.6964 | 0.71554 | 12.2282 | 1.5883 | 0.8165 | 1.6857 | 1.7607 | 1.583 | 1.8442
| Rel.Std.Dev. |  | 3.739% | 4.552% | 1.939% | 1.992% | 0.6162% | 1.386% | 1.415% | 0.2651% | 0.9773% | 1.883% | 2.514% | 1.557% | 2.633% | 2.926% | 1.71% | 0.8389% | 18.77% | 1.693% | 0.914% | 1.918% | 2.059% | 2.005% | 2.207%
| PDM3 |  | 2.973% | 1.846% | 6.088% | -0.3654% | -1.277% | 12.13% | 10.58% | 27.09% | -1.223% | 4.439% | -2.414% | -2.267% | -2.414% | -1.516% | 9.074% | -6.225% | -28.39% | 3.156% | -1.791% | -3.384% | -6.005% | -13.22% | -8.129%
| Z-Score (Absolute) |  | 0.65 | 0.48 | 1.12 | 0.14 | 0 | 2.04 | 1.8 | 4.31 | 0.01 | 0.87 | 0.17 | 0.15 | 0.17 | 0.04 | 1.57 | 0.75 | 4.12 | 0.67 | 0.08 | 0.32 | 0.72 | 1.82 | 1.04
| Analytical results for Co in OREAS 600c (Certified Value 2.22 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.7 | 2.4 | 2.3 | 2.1 | 1.9793 | 1 | 2.34 | 2.12 | 2 | 2.2 | 2.3 | 2.2 | 2.1 | 2.3 | 2.2 | 2.1 | 2.4 | 2.5 | 4.88 | 2.4 | 9.6515 | < 1 | 2.2749 | < 5 | 2 |  | 1
| 1 | 2 | 2.6 | 2.1 | 2.3 | 2.1 | 1.9451 | 2 | 2.31 | 2.26 | 2 | 2.2 | 2.4 | 2.2 | 2.2 | 2.2 | 2.2 | 2.1 | 2.5 | 2.3 | 4.91 | 2.3 | 17.2246 | < 1 | 2.1483 | < 5 | 1.9 |  | 16
| 1 | 3 | 2.8 | 2.2 | 2.2 | 2.2 | 1.9784 | 2 | 2.25 | 2.17 | 2 | 2.2 | 2.4 | 2.2 | 2.2 | 2.2 | 2.3 | 2.2 | 2.4 | 2.5 | 4.7 | 2.4 | 12.944 | < 1 | 2.238 | < 5 | 2 |  | 16
| 1 | 4 | 2.4 | 2.1 | 2.3 | 2.2 | 1.9633 | 2 | 2.28 | 2.14 | 2 | 2.3 | 2.4 | 2.2 | 2.2 | 2.2 | 2.4 | 2.2 | 2.1 | 2.4 | 4.97 | 2.3 | 26.7741 | < 1 | 2.1494 | < 5 | 1.9 |  | 2.2161
| 1 | 5 | 2.4 | 2.2 | 2.3 | 2.2 | 1.9734 | 2 | 2.26 | 2.05 | 2 | 2.2 | 2.3 | 2.2 | 2.2 | 2.2 | 2.3 | 2.1 | 2.3 | 2.4 | 4.81 | 2.3 | 18.9004 | < 1 | 2.1322 | < 5 | 1.9 |  | 22
| 1 | 6 | 2.8 | 2 | 2.2 | 2 | 1.9385 | 2 | 2.3 | 2.12 | 2 | 2.2 | 2.4 | 2.1 | 2.3 | 2.2 | 2.5 | 2.2 | 2.4 | 2.5 | 4.8 | 2.2 | 21.0737 | < 1 | 2.2173 | < 5 | 1.9
| Mean |  | 2.6167 | 2.1667 | 2.2667 | 2.1333 | 1.963 | 1.8333 | 2.29 | 2.1433 | 2 | 2.2167 | 2.3667 | 2.1833 | 2.2 | 2.2167 | 2.3167 | 2.15 | 2.35 | 2.4333 | 4.845 | 2.3167 | 17.7614 |  | 2.1933 |  | 1.9333
| Median |  | 2.65 | 2.15 | 2.3 | 2.15 | 1.9683 | 2 | 2.29 | 2.13 | 2 | 2.2 | 2.4 | 2.2 | 2.2 | 2.2 | 2.3 | 2.15 | 2.4 | 2.45 | 4.845 | 2.3 | 18.0625 |  | 2.1833 |  | 1.9
| Std Dev. |  | 0.18348 | 0.13663 | 0.05164 | 0.08165 | 0.017502 | 0.40825 | 0.033466 | 0.069474 | 0 | 0.040825 | 0.05164 | 0.040825 | 0.063246 | 0.040825 | 0.1169 | 0.054772 | 0.13784 | 0.08165 | 0.095237 | 0.075277 | 6.0459 |  | 0.058166 |  | 0.05164
| Rel.Std.Dev. |  | 7.012% | 6.306% | 2.278% | 3.827% | 0.8916% | 22.27% | 1.461% | 3.241% | 0% | 1.842% | 2.182% | 1.87% | 2.875% | 1.842% | 5.046% | 2.548% | 5.866% | 3.355% | 1.966% | 3.249% | 34.04% |  | 2.652% |  | 2.671%
| PDM3 |  | 18.08% | -2.23% | 2.282% | -3.734% | -11.42% | -17.27% | 3.335% | -3.283% | -9.751% | 0.02593% | 6.795% | -1.478% | -0.7261% | 0.02593% | 4.538% | -2.982% | 6.043% | 9.803% | 118.6% | 4.538% | 701.5% |  | -1.027% |  | -12.76%
| Z-Score (Absolute) |  | 2.7 | 0.34 | 0.34 | 0.56 | 1.71 | NA | 0.49 | 0.49 | NA | 0 | 1.01 | 0.22 | 0.11 | 0 | 0.67 | 0.45 | 0.9 | 1.46 | 17.72 | 0.67 | 104.82 | 11.58 | 0.16 | 1.91 | 1.91
Results from laboratories 2.07 and 2.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 600c (Certified Value 13 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 1
1
| 1 | 1 | 16 | 10 | 11 | 16 | 13.6229 | 14 | 11.5 | 14.3 | 16 | 13 | 12 | 12 | 12 | 13 | 14 | 11 | < 10 | 11 | 17.83 | 15 | 47.3504 | 14 | 11.6288 | 13.139 | < 10 |  | 1
| 1 | 2 | 16 | 9 | 13 | 15 | 13.4724 | 24 | 11.6 | 13.3 | 14 | 12 | 13 | 13 | 12 | 12 | 13 | 11.9 | 11 | 12 | 24.27 | 16 | 44.7989 | 14 | 9.9888 | 17.494 | < 10 |  | 30
| 1 | 3 | 15 | 12 | 12 | 15 | 13.196 | 25 | 12.6 | 12.3 | 14 | 16 | 13 | 13 | 11 | 12 | 13 | 10.7 | < 10 | 12 | 22.02 | 17 | 35.4779 | 16 | 13.0515 | 13.896 | < 10 |  | 16
| 1 | 4 | 15 | 10 | 11 | 14 | 12.9192 | 17 | 11.8 | 12.7 | 14 | 13 | 13 | 13 | 11 | 16 | 14 | 11.7 | 11 | 13 | 23.48 | 15 | 43.2907 | 12 | 15.8025 | 16.432 | < 10 |  | 12.9696
| 1 | 5 | 16 | 10 | 13 | 16 | 13.8798 | 13 | 11.8 | 12.7 | 14 | 12 | 15 | 14 | 12 | 12 | 13 | 10.8 | 11 | 13 | 29.1 | 16 | 43.3972 | 14 | 17.2339 | 13.624 | 12 |  | 23
| 1 | 6 | 15 | 10 | 11 | 15 | 13.0869 | 18 | 9.6 | 12.7 | 14 | 13 | 12 | 14 | 11 | 12 | 13 | 10.9 | 11 | 12 | 21.99 | 15 | 44.9359 | 13 | 15.6157 | 14.501 | < 10
| Mean |  | 15.5 | 10.1667 | 11.8333 | 15.1667 | 13.3629 | 18.5 | 11.4833 | 13 | 14.3333 | 13.1667 | 13 | 13.1667 | 11.5 | 12.8333 | 13.3333 | 11.1667 | 11 | 12.1667 | 23.115 | 15.6667 | 43.2085 | 13.8333 | 13.8869 | 14.8477 | 12
| Median |  | 15.5 | 10 | 11.5 | 15 | 13.3342 | 17.5 | 11.7 | 12.7 | 14 | 13 | 13 | 13 | 11.5 | 12 | 13 | 10.95 | 11 | 12 | 22.75 | 15.5 | 44.0981 | 14 | 14.3336 | 14.1985 | 12
| Std Dev. |  | 0.54772 | 0.98319 | 0.98319 | 0.75277 | 0.35989 | 5.01 | 1.0008 | 0.71274 | 0.8165 | 1.472 | 1.0954 | 0.75277 | 0.54772 | 1.6021 | 0.5164 | 0.50465 | 0 | 0.75277 | 3.6792 | 0.8165 | 4.0614 | 1.3292 | 2.7876 | 1.7295
| Rel.Std.Dev. |  | 3.534% | 9.671% | 8.309% | 4.963% | 2.693% | 27.08% | 8.716% | 5.483% | 5.696% | 11.18% | 8.427% | 5.717% | 4.763% | 12.48% | 3.873% | 4.519% | 0% | 6.187% | 15.92% | 5.212% | 9.399% | 9.608% | 20.07% | 11.65%
| PDM3 |  | 19.51% | -21.61% | -8.761% | 16.94% | 3.033% | 42.64% | -11.46% | 0.2345% | 10.52% | 1.52% | 0.2345% | 1.52% | -11.33% | -1.051% | 2.805% | -13.9% | -15.19% | -6.191% | 78.22% | 20.8% | 233.2% | 6.66% | 7.073% | 14.48% | -7.476%
| Z-Score (Absolute) |  | 0.94 | 1.21 | 0.54 | 0.81 | 0.08 | 2.16 | 0.68 | 0.07 | 0.47 | 0 | 0.07 | 0 | 0.67 | 0.13 | 0.07 | 0.81 | 1.69 | 0.4 | 4.02 | 1.01 | 12.15 | 0.27 | 0.29 | 0.68 | 2.83
| Analytical results for Cs in OREAS 600c (Certified Value 6.55 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 6.75 | 6 | 6.55 | 6.1693 | 6.77 | 7.13 | 7.1329 | 6.18 | 6.7 | 6.96 | 6.68 | 6.45 | 6.13 | 6.75 | 6.43 | 6 | 5.43 | 7.13 | 6.89 | 6.6389 | 5.87 | [5 empty] | 1
| 1 | 2 | 6.68 | 5.5 | 6.58 | 6.3298 | 6.78 | 7.25 | 7.1953 | 6.44 | 6.7 | 6.63 | 6.75 | 6.35 | 6.12 | 6.4 | 6.48 | 5.9 | 5.77 | 6.6 | 6.67 | 6.8551 | 5.73 | [5 empty] | 31
| 1 | 3 | 7.17 | 5.9 | 6.5 | 6.6046 | 6.89 | 6.85 | 7.3209 | 6.37 | 6.7 | 6.9 | 6.7 | 6.28 | 6.16 | 6.48 | 6.65 | 6 | 5.3 | 7.27 | 6.91 | 6.9162 | 5.7 | [5 empty] | 16
| 1 | 4 | 6.53 | 5.8 | 6.68 | 6.5175 | 7.02 | 6.94 | 7.1653 | 6.32 | 6.7 | 6.69 | 6.65 | 6.39 | 6.11 | 6.52 | 6.8 | 6.1 | 5.09 | 7.11 | 6.9 | 6.804 | 5.71 | [5 empty] | 6.548
| 1 | 5 | 6.19 | 5.6 | 6.74 | 6.3782 | 6.81 | 6.95 | 7.1175 | 6.17 | 6.9 | 6.8 | 6.59 | 6.49 | 6.07 | 6.57 | 6.54 | 6 | 4.99 | 6.84 | 6.77 | 6.6103 | 5.67 | [5 empty] | 24
| 1 | 6 | 7.02 | 5.3 | 6.47 | 6.0991 | 7.09 | 7 | 7.1337 | 6.31 | 6.7 | 6.89 | 6.64 | 6.26 | 6.31 | 6.51 | 6.69 | 6.1 | 5.35 | 7.25 | 6.84 | 6.8996 | 5.91
| Mean |  | 6.7233 | 5.6833 | 6.5867 | 6.3498 | 6.8933 | 7.02 | 7.1776 | 6.2983 | 6.7333 | 6.8117 | 6.6683 | 6.37 | 6.15 | 6.5383 | 6.5983 | 6.0167 | 5.3217 | 7.0333 | 6.83 | 6.7874 | 5.765
| Median |  | 6.715 | 5.7 | 6.565 | 6.354 | 6.85 | 6.975 | 7.1495 | 6.315 | 6.7 | 6.845 | 6.665 | 6.37 | 6.125 | 6.515 | 6.595 | 6 | 5.325 | 7.12 | 6.865 | 6.8296 | 5.72
| Std Dev. |  | 0.34984 | 0.26394 | 0.10463 | 0.19478 | 0.13397 | 0.14533 | 0.075547 | 0.1061 | 0.08165 | 0.12952 | 0.054924 | 0.091433 | 0.083666 | 0.11788 | 0.13963 | 0.075277 | 0.27484 | 0.26204 | 0.094021 | 0.13226 | 0.099549
| Rel.Std.Dev. |  | 5.203% | 4.644% | 1.588% | 3.067% | 1.943% | 2.07% | 1.053% | 1.685% | 1.213% | 1.902% | 0.8237% | 1.435% | 1.36% | 1.803% | 2.116% | 1.251% | 5.165% | 3.726% | 1.377% | 1.949% | 1.727%
| PDM3 |  | 2.677% | -13.21% | 0.5902% | -3.028% | 5.274% | 7.208% | 9.615% | -3.813% | 2.83% | 4.026% | 1.837% | -2.719% | -6.078% | -0.1479% | 0.7684% | -8.115% | -18.73% | 7.412% | 4.306% | 3.655% | -11.96%
| Z-Score (Absolute) |  | 0.34 | 2.48 | 0.03 | 0.67 | 0.8 | 1.14 | 1.57 | 0.81 | 0.37 | 0.58 | 0.19 | 0.62 | 1.22 | 0.16 | 0 | 1.58 | 3.46 | 1.18 | 0.63 | 0.51 | 2.26
| Analytical results for Cu in OREAS 600c (Certified Value 0.058 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0602 | 0.06335 | 0.05779 | 0.05579 | 0.058013 | 0.0562 | 0.0635 | 0.058884 | 0.0566 | 0.0611 | 0.0586 | 0.0575 | 0.0595 | 0.0543 | 0.0572 | 0.0581 | 0.05666 | 0.0586 | 0.0591 | 0.053 | 0.05689 | 0.059232 | 0.0592 | 0.058635 | 0.054389 | 0.0537 | 1
| 1 | 2 | 0.06 | 0.06327 | 0.05857 | 0.05684 | 0.058535 | 0.0556 | 0.0613 | 0.058055 | 0.0585 | 0.0599 | 0.0574 | 0.0568 | 0.0578 | 0.0548 | 0.0571 | 0.0586 | 0.05639 | 0.0588 | 0.0567 | 0.05355 | 0.0544 | 0.057564 | 0.0597 | 0.058565 | 0.054274 | 0.056 | 18
| 1 | 3 | 0.0604 | 0.0641 | 0.05708 | 0.05771 | 0.059552 | 0.0556 | 0.0609 | 0.058513 | 0.058 | 0.0589 | 0.0574 | 0.0583 | 0.0583 | 0.0548 | 0.0563 | 0.0582 | 0.05703 | 0.0582 | 0.0595 | 0.05318 | 0.05602 | 0.057437 | 0.0589 | 0.059188 | 0.055332 | 0.058 | 16
| 1 | 4 | 0.0609 | 0.06221 | 0.05858 | 0.05765 | 0.058736 | 0.0568 | 0.0617 | 0.058058 | 0.0569 | 0.0596 | 0.0594 | 0.0567 | 0.0581 | 0.0544 | 0.0569 | 0.0605 | 0.05799 | 0.0534 | 0.0587 | 0.05431 | 0.05643 | 0.056711 | 0.0575 | 0.058205 | 0.054742 | 0.0567 | 0.057852
| 1 | 5 | 0.0608 | 0.0646 | 0.05959 | 0.05816 | 0.059183 | 0.054 | 0.0614 | 0.058083 | 0.0571 | 0.0604 | 0.0576 | 0.0573 | 0.0597 | 0.0549 | 0.0574 | 0.0578 | 0.05621 | 0.0578 | 0.0592 | 0.05377 | 0.05634 | 0.058338 | 0.0586 | 0.059205 | 0.054024 | 0.059 | 25
| 1 | 6 | 0.0598 | 0.06141 | 0.05739 | 0.05611 | 0.058445 | 0.0571 | 0.0605 | 0.058477 | 0.0575 | 0.0615 | 0.0579 | 0.0561 | 0.056 | 0.0554 | 0.0567 | 0.0596 | 0.05727 | 0.0588 | 0.0606 | 0.05286 | 0.05552 | 0.056755 | 0.0583 | 0.058422 | 0.055572 | 0.0593
| Mean |  | 0.06035 | 0.063157 | 0.058167 | 0.057043 | 0.058744 | 0.055883 | 0.06155 | 0.058345 | 0.057433 | 0.060233 | 0.05805 | 0.057117 | 0.058233 | 0.054767 | 0.056933 | 0.0588 | 0.056925 | 0.0576 | 0.058967 | 0.053445 | 0.055933 | 0.057673 | 0.0587 | 0.058703 | 0.054722 | 0.057117
| Median |  | 0.0603 | 0.06331 | 0.05818 | 0.057245 | 0.058635 | 0.0559 | 0.06135 | 0.05828 | 0.0573 | 0.06015 | 0.05775 | 0.05705 | 0.0582 | 0.0548 | 0.057 | 0.0584 | 0.056845 | 0.0584 | 0.05915 | 0.053365 | 0.05618 | 0.0575 | 0.05875 | 0.0586 | 0.054566 | 0.05735
| Std Dev. |  | 0.00043704 | 0.0011807 | 0.00092578 | 0.00095318 | 0.00055005 | 0.0011071 | 0.0010426 | 0.00033801 | 0.00071461 | 0.00096678 | 0.00079937 | 0.00076004 | 0.0013382 | 0.00039328 | 0.00039328 | 0.0010412 | 0.00065299 | 0.0020938 | 0.0012832 | 0.0005432 | 0.00087826 | 0.00097039 | 0.00076158 | 0.00040943 | 0.00061554 | 0.0021066
| Rel.Std.Dev. |  | 0.7242% | 1.869% | 1.592% | 1.671% | 0.9363% | 1.981% | 1.694% | 0.5793% | 1.244% | 1.605% | 1.377% | 1.331% | 2.298% | 0.7181% | 0.6908% | 1.771% | 1.147% | 3.635% | 2.176% | 1.016% | 1.57% | 1.683% | 1.297% | 0.6974% | 1.125% | 3.688%
| PDM3 |  | 4.319% | 9.17% | 0.5446% | -1.397% | 1.543% | -3.402% | 6.393% | 0.8527% | -0.723% | 4.117% | 0.3429% | -1.27% | 0.6598% | -5.333% | -1.587% | 1.639% | -1.602% | -0.435% | 1.927% | -7.617% | -3.316% | -0.3089% | 1.466% | 1.472% | -5.409% | -1.27%
| Z-Score (Absolute) |  | 1.85 | 3.94 | 0.23 | 0.61 | 0.66 | 1.47 | 2.75 | 0.36 | 0.32 | 1.77 | 0.14 | 0.55 | 0.28 | 2.31 | 0.69 | 0.7 | 0.7 | 0.19 | 0.82 | 3.29 | 1.44 | 0.14 | 0.62 | 0.63 | 2.34 | 0.55
| Analytical results for Dy in OREAS 600c (Certified Value 3.77 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 4; 2.03: 3.74; 2.04: 3.4709; 2.05: 4; 2.06: 4.1895; 2.07: 3.75; 2.08: 3.58; 2.09: 3.87; 2.10: 3.7761; 2.11: 3.8; 2.12: 3.66 | 1
| 1 | 2 | 2.01: 3.6; 2.03: 3.7; 2.04: 3.6886; 2.05: 3.9; 2.06: 4.1627; 2.07: 3.81; 2.08: 3.88; 2.09: 3.76; 2.10: 3.5784; 2.11: 3.9; 2.12: 3.83 | 33
| 1 | 3 | 2.01: 3.9; 2.03: 3.76; 2.04: 3.6912; 2.05: 3.8; 2.06: 4.3403; 2.07: 3.9; 2.08: 3.56; 2.09: 3.87; 2.10: 3.7287; 2.11: 3.8; 2.12: 3.67 | 16
| 1 | 4 | 2.01: 4; 2.03: 3.81; 2.04: 3.5982; 2.05: 3.9; 2.06: 4.1189; 2.07: 3.77; 2.08: 3.01; 2.09: 3.9; 2.10: 3.7324; 2.11: 3.8; 2.12: 3.58 | 3.7686
| 1 | 5 | 2.01: 3.6; 2.03: 3.89; 2.04: 3.6604; 2.05: 3.9; 2.06: 4.1178; 2.07: 3.77; 2.08: 3.05; 2.09: 3.69; 2.10: 3.791; 2.11: 3.8; 2.12: 3.76 | 26
| 1 | 6 | 2.01: 3.6; 2.03: 3.66; 2.04: 3.4376; 2.05: 4; 2.06: 4.162; 2.07: 3.89; 2.08: 3.29; 2.09: 3.78; 2.10: 3.6811; 2.11: 3.8; 2.12: 3.75
| Mean |  | 2.01: 3.7833; 2.03: 3.76; 2.04: 3.5912; 2.05: 3.9167; 2.06: 4.1819; 2.07: 3.815; 2.08: 3.395; 2.09: 3.8117; 2.10: 3.7146; 2.11: 3.8167; 2.12: 3.7083
| Median |  | 2.01: 3.75; 2.03: 3.75; 2.04: 3.6293; 2.05: 3.9; 2.06: 4.1623; 2.07: 3.79; 2.08: 3.425; 2.09: 3.825; 2.10: 3.7306; 2.11: 3.8; 2.12: 3.71
| Std Dev. |  | 2.01: 0.20412; 2.03: 0.081731; 2.04: 0.11169; 2.05: 0.075277; 2.06: 0.082451; 2.07: 0.065038; 2.08: 0.3391; 2.09: 0.081343; 2.10: 0.077211; 2.11: 0.040825; 2.12: 0.088863
| Rel.Std.Dev. |  | 2.01: 5.395%; 2.03: 2.174%; 2.04: 3.11%; 2.05: 1.922%; 2.06: 1.972%; 2.07: 1.705%; 2.08: 9.988%; 2.09: 2.134%; 2.10: 2.079%; 2.11: 1.07%; 2.12: 2.396%
| PDM3 |  | 2.01: 0.3908%; 2.03: -0.2283%; 2.04: -4.709%; 2.05: 3.929%; 2.06: 10.97%; 2.07: 1.231%; 2.08: -9.914%; 2.09: 1.143%; 2.10: -1.433%; 2.11: 1.275%; 2.12: -1.599%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.23; 2.04: 1.89; 2.05: 1.31; 2.06: 3.91; 2.07: 0.31; 2.08: 3.81; 2.09: 0.28; 2.10: 0.67; 2.11: 0.33; 2.12: 0.74
| Analytical results for Er in OREAS 600c (Certified Value 1.07 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 1.1; 2.03: 1.06; 2.04: 0.99285; 2.05: 1.1; 2.06: 1.2807; 2.07: 1.06; 2.08: 1.05; 2.09: 1.1; 2.10: 1.1319; 2.11: < 0.2; 2.12: 1.05 | 1
| 1 | 2 | 2.01: 1; 2.03: 1.06; 2.04: 1.0233; 2.05: 1; 2.06: 1.184; 2.07: 1.06; 2.08: 1.09; 2.09: 1.09; 2.10: 1.0696; 2.11: < 0.2; 2.12: 1.07 | 34
| 1 | 3 | 2.01: 1.2; 2.03: 1.04; 2.04: 1.0066; 2.05: 1.1; 2.06: 1.3386; 2.07: 1.07; 2.08: 1.03; 2.09: 1.12; 2.10: 1.1207; 2.11: < 0.2; 2.12: 1.03 | 16
| 1 | 4 | 2.01: 1.1; 2.03: 1.06; 2.04: 0.99894; 2.05: 1.1; 2.06: 1.2852; 2.07: 1.05; 2.08: 0.86; 2.09: 1.12; 2.10: 1.1412; 2.11: < 0.2; 2.12: 1.08 | 1.0662
| 1 | 5 | 2.01: 1; 2.03: 1.1; 2.04: 1.0102; 2.05: 1.1; 2.06: 1.2861; 2.07: 1.08; 2.08: 0.93; 2.09: 1.04; 2.10: 1.1314; 2.11: < 0.2; 2.12: 1.03 | 27
| 1 | 6 | 2.01: 1; 2.03: 1.05; 2.04: 0.96794; 2.05: 1; 2.06: 1.2915; 2.07: 1.08; 2.08: 0.95; 2.09: 1.06; 2.10: 1.1212; 2.11: < 0.2; 2.12: 1.04
| Mean |  | 2.01: 1.0667; 2.03: 1.0617; 2.04: 0.99996; 2.05: 1.0667; 2.06: 1.2777; 2.07: 1.0667; 2.08: 0.985; 2.09: 1.0883; 2.10: 1.1193; 2.12: 1.05
| Median |  | 2.01: 1.05; 2.03: 1.06; 2.04: 1.0028; 2.05: 1.1; 2.06: 1.2856; 2.07: 1.065; 2.08: 0.99; 2.09: 1.095; 2.10: 1.1263; 2.12: 1.045
| Std Dev. |  | 2.01: 0.08165; 2.03: 0.020412; 2.04: 0.018818; 2.05: 0.05164; 2.06: 0.050626; 2.07: 0.012111; 2.08: 0.086197; 2.09: 0.032506; 2.10: 0.025536; 2.12: 0.020976
| Rel.Std.Dev. |  | 2.01: 7.655%; 2.03: 1.923%; 2.04: 1.882%; 2.05: 4.841%; 2.06: 3.962%; 2.07: 1.135%; 2.08: 8.751%; 2.09: 2.987%; 2.10: 2.281%; 2.12: 1.998%
| PDM3 |  | 2.01: 0.04798%; 2.03: -0.421%; 2.04: -6.209%; 2.05: 0.04798%; 2.06: 19.84%; 2.07: 0.04798%; 2.08: -7.612%; 2.09: 2.08%; 2.10: 4.988%; 2.12: -1.515%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.16; 2.04: 2.08; 2.05: 0; 2.06: 6.57; 2.07: 0; 2.08: 2.54; 2.09: 0.67; 2.10: 1.64; 2.11: 30.08; 2.12: 0.52
| Analytical results for Eu in OREAS 600c (Certified Value 1.47 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 2
3
| 1 | 1 | 2.01: 1.5; 2.03: 1.39; 2.04: 1.4248; 2.05: 1.53; 2.06: 3.0415; 2.07: 1.46; 2.08: 1.85; 2.09: 1.82; 2.10: 1.7626; 2.11: 1.4; 2.12: 1.41 | 1
| 1 | 2 | 2.01: 1.4; 2.03: 1.42; 2.04: 1.5377; 2.05: 1.43; 2.06: 3.1362; 2.07: 1.45; 2.08: 1.92; 2.09: 1.77; 2.10: 1.7196; 2.11: 1.4; 2.12: 1.45 | 35
| 1 | 3 | 2.01: 1.6; 2.03: 1.39; 2.04: 1.5745; 2.05: 1.38; 2.06: 3.0991; 2.07: 1.45; 2.08: 1.79; 2.09: 1.72; 2.10: 1.7367; 2.11: 1.4; 2.12: 1.38 | 16
| 1 | 4 | 2.01: 1.5; 2.03: 1.38; 2.04: 1.5502; 2.05: 1.48; 2.06: 3.0146; 2.07: 1.38; 2.08: 1.48; 2.09: 1.76; 2.10: 1.7606; 2.11: 1.3; 2.12: 1.43 | 1.4689
| 1 | 5 | 2.01: 1.4; 2.03: 1.44; 2.04: 1.5421; 2.05: 1.44; 2.06: 3.1644; 2.07: 1.46; 2.08: 1.63; 2.09: 1.76; 2.10: 1.7561; 2.11: 1.3; 2.12: 1.4 | 28
| 1 | 6 | 2.01: 1.4; 2.03: 1.39; 2.04: 1.4615; 2.05: 1.43; 2.06: 3.0377; 2.07: 1.45; 2.08: 1.73; 2.09: 1.7; 2.10: 1.7174; 2.11: 1.4; 2.12: 1.39
| Mean |  | 2.01: 1.4667; 2.03: 1.4017; 2.04: 1.5151; 2.05: 1.4483; 2.06: 3.0823; 2.07: 1.4417; 2.08: 1.7333; 2.09: 1.755; 2.10: 1.7422; 2.11: 1.3667; 2.12: 1.41
| Median |  | 2.01: 1.45; 2.03: 1.39; 2.04: 1.5399; 2.05: 1.435; 2.06: 3.0703; 2.07: 1.45; 2.08: 1.76; 2.09: 1.76; 2.10: 1.7464; 2.11: 1.4; 2.12: 1.405
| Std Dev. |  | 2.01: 0.08165; 2.03: 0.023166; 2.04: 0.058365; 2.05: 0.051153; 2.06: 0.060274; 2.07: 0.030605; 2.08: 0.15908; 2.09: 0.041833; 2.10: 0.020513; 2.11: 0.05164; 2.12: 0.026077
| Rel.Std.Dev. |  | 2.01: 5.567%; 2.03: 1.653%; 2.04: 3.852%; 2.05: 3.532%; 2.06: 1.956%; 2.07: 2.123%; 2.08: 9.178%; 2.09: 2.384%; 2.10: 1.177%; 2.11: 3.779%; 2.12: 1.849%
| PDM3 |  | 2.01: -0.1487%; 2.03: -4.574%; 2.04: 3.152%; 2.05: -1.397%; 2.06: 109.8%; 2.07: -1.851%; 2.08: 18.01%; 2.09: 19.48%; 2.10: 18.61%; 2.11: -6.957%; 2.12: -4.007%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.67; 2.04: 0.5; 2.05: 0.19; 2.06: 16.76; 2.07: 0.26; 2.08: 2.77; 2.09: 2.99; 2.10: 2.86; 2.11: 1.04; 2.12: 0.59
| Analytical results for Fe in OREAS 600c (Certified Value 2.5 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.66 | 2.49 | 2.47 | 2.41 | 2.5249 | 2.61 | 2.67 | 2.5354 | 2.45 | 2.38 | 2.52 | 2.41 | 2.47 | 2.34 | 2.31 | 2.47 | 2.427 | 2.36 | 2.54 | 2.49 | 2.49 | 2.5298 | 2.371 | 2.3249 | 2.4769 | 2.57 | 1
| 1 | 2 | 2.63 | 2.46 | 2.52 | 2.41 | 2.5062 | 2.65 | 2.62 | 2.5074 | 2.51 | 2.43 | 2.48 | 2.42 | 2.46 | 2.33 | 2.24 | 2.51 | 2.4402 | 2.41 | 2.48 | 2.56 | 2.4 | 2.4802 | 2.406 | 2.6874 | 2.4977 | 2.5 | 0
| 1 | 3 | 2.62 | 2.55 | 2.45 | 2.4 | 2.5436 | 2.67 | 2.61 | 2.5047 | 2.52 | 2.42 | 2.49 | 2.42 | 2.46 | 2.33 | 2.28 | 2.49 | 2.4022 | 2.34 | 2.6 | 2.53 | 2.46 | 2.5169 | 2.364 | 2.4824 | 2.5193 | 2.49 | 16
| 1 | 4 | 2.6 | 2.51 | 2.52 | 2.37 | 2.5315 | 2.71 | 2.61 | 2.5782 | 2.47 | 2.42 | 2.54 | 2.39 | 2.46 | 2.32 | 2.33 | 2.59 | 2.4662 | 2.19 | 2.59 | 2.55 | 2.48 | 2.4833 | 2.308 | 2.6721 | 2.4667 | 2.51 | 2.4995
| 1 | 5 | 2.67 | 2.52 | 2.54 | 2.4 | 2.5853 | 2.61 | 2.58 | 2.5639 | 2.48 | 2.47 | 2.47 | 2.4 | 2.51 | 2.33 | 2.29 | 2.46 | 2.3899 | 2.27 | 2.54 | 2.57 | 2.41 | 2.5298 | 2.343 | 2.4824 | 2.4647 | 2.49 | 29
| 1 | 6 | 2.61 | 2.37 | 2.5 | 2.36 | 2.5321 | 2.67 | 2.57 | 2.5267 | 2.51 | 2.42 | 2.49 | 2.38 | 2.44 | 2.4 | 2.29 | 2.51 | 2.4258 | 2.33 | 2.6 | 2.54 | 2.5 | 2.4388 | 2.336 | 2.417 | 2.5366 | 2.57
| Mean |  | 2.6317 | 2.4833 | 2.5 | 2.3917 | 2.5373 | 2.6533 | 2.61 | 2.536 | 2.49 | 2.4233 | 2.4983 | 2.4033 | 2.4667 | 2.3417 | 2.29 | 2.505 | 2.4252 | 2.3167 | 2.5583 | 2.54 | 2.4567 | 2.4965 | 2.3547 | 2.511 | 2.4937 | 2.5217
| Median |  | 2.625 | 2.5 | 2.51 | 2.4 | 2.5318 | 2.66 | 2.61 | 2.5311 | 2.495 | 2.42 | 2.49 | 2.405 | 2.46 | 2.33 | 2.29 | 2.5 | 2.4264 | 2.335 | 2.565 | 2.545 | 2.47 | 2.5001 | 2.3535 | 2.4824 | 2.4873 | 2.505
| Std Dev. |  | 0.027869 | 0.06314 | 0.034059 | 0.02137 | 0.02653 | 0.038816 | 0.035214 | 0.029823 | 0.027568 | 0.028752 | 0.026394 | 0.01633 | 0.023381 | 0.029269 | 0.030332 | 0.046368 | 0.027147 | 0.076855 | 0.047504 | 0.028284 | 0.042269 | 0.035786 | 0.033619 | 0.14295 | 0.029562 | 0.038166
| Rel.Std.Dev. |  | 1.059% | 2.543% | 1.362% | 0.8935% | 1.046% | 1.463% | 1.349% | 1.176% | 1.107% | 1.186% | 1.056% | 0.6795% | 0.9479% | 1.25% | 1.325% | 1.851% | 1.119% | 3.317% | 1.857% | 1.114% | 1.721% | 1.433% | 1.428% | 5.693% | 1.185% | 1.514%
| PDM3 |  | 5.289% | -0.6456% | 0.02118% | -4.313% | 1.513% | 6.156% | 4.422% | 1.463% | -0.3789% | -3.046% | -0.0455% | -3.846% | -1.312% | -6.313% | -8.381% | 0.2212% | -2.971% | -7.314% | 2.355% | 1.622% | -1.713% | -0.1202% | -5.793% | 0.4622% | -0.2328% | 0.888%
| Z-Score (Absolute) |  | 2.11 | 0.18 | 0.08 | 1.6 | 0.65 | 2.45 | 1.78 | 0.63 | 0.08 | 1.11 | 0.05 | 1.42 | 0.44 | 2.37 | 3.17 | 0.15 | 1.08 | 2.76 | 0.98 | 0.7 | 0.59 | 0.02 | 2.17 | 0.25 | 0.02 | 0.41
| Analytical results for Ga in OREAS 600c (Certified Value 22.8 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 24.77 | 21.64 | 21.76 | 22.7049 | 23.9 | 25.3 | 28.203 | 22.75 | 22.9 | 24.9 | 22 | 22 | 21.4 | 23.1 | 21.7 | 21.32 | 24.15 | 29 | 21 | 23.0192 | 18 | 22.7192 | 19.87 | 26.05 |  |  | 1
| 1 | 2 | 24.86 | 21.72 | 21.24 | 22.5253 | 23.5 | 25.8 | 28.904 | 23.4 | 22.9 | 23.73 | 22.5 | 21.9 | 21.8 | 21.4 | 21.7 | 21.23 | 24.95 | 32 | 21 | 22.6437 | 17 | 23.7481 | 19.545 | 26.96 |  |  | 36
| 1 | 3 | 23.8 | 22.08 | 21.56 | 22.9269 | 23.4 | 24.2 | 28.765 | 23.17 | 23 | 24.31 | 22.2 | 21.3 | 22.1 | 21.5 | 22 | 21.28 | 24.46 | 29.2 | 21 | 23.6936 | 17 | 23.5331 | 19.86 | 26.5 |  |  | 16
| 1 | 4 | 24.02 | 21.72 | 21.51 | 22.6228 | 24.4 | 24.6 | 28.84 | 23.29 | 22.9 | 24.45 | 22.3 | 21.7 | 22 | 21.6 | 22.1 | 21.58 | 21.93 | 30.6 | 22 | 23.9204 | 17 | 23.3088 | 19.738 | 27.13 |  |  | 22.7749
| 1 | 5 | 23.23 | 22.48 | 21.59 | 23.0217 | 22.9 | 24.7 | 28.43 | 22.27 | 22.8 | 23.88 | 21.7 | 21.9 | 21.1 | 21.9 | 21.9 | 21.15 | 23.77 | 28.4 | 21 | 24.25 | 17 | 22.8617 | 19.705 | 27.45 |  |  | 30
| 1 | 6 | 24 | 20.8 | 21.65 | 22.3369 | 24.1 | 25 | 28.57 | 22.86 | 22.9 | 24.57 | 22.1 | 21.1 | 22.4 | 21.6 | 22.2 | 21.4 | 24.41 | 29.8 | 22 | 23.2757 | 17 | 23.6819 | 20.519 | 26.34
| Mean |  | 24.1133 | 21.74 | 21.5517 | 22.6897 | 23.7 | 24.9333 | 28.6187 | 22.9567 | 22.9 | 24.3067 | 22.1333 | 21.65 | 21.8 | 21.85 | 21.9333 | 21.3267 | 23.945 | 29.8333 | 21.3333 | 23.4671 | 17.1667 | 23.3088 | 19.8728 | 26.7383
| Median |  | 24.01 | 21.72 | 21.575 | 22.6638 | 23.7 | 24.85 | 28.6675 | 23.015 | 22.9 | 24.38 | 22.15 | 21.8 | 21.9 | 21.6 | 21.95 | 21.3 | 24.28 | 29.5 | 21 | 23.4846 | 17 | 23.421 | 19.799 | 26.73
| Std Dev. |  | 0.61461 | 0.55814 | 0.1752 | 0.25404 | 0.54037 | 0.56451 | 0.26928 | 0.41874 | 0.063246 | 0.43739 | 0.27325 | 0.36742 | 0.47749 | 0.63482 | 0.20656 | 0.14989 | 1.0606 | 1.2987 | 0.5164 | 0.59749 | 0.40825 | 0.43129 | 0.33809 | 0.52883
| Rel.Std.Dev. |  | 2.549% | 2.567% | 0.813% | 1.12% | 2.28% | 2.264% | 0.9409% | 1.824% | 0.2762% | 1.799% | 1.235% | 1.697% | 2.19% | 2.905% | 0.9418% | 0.7028% | 4.429% | 4.353% | 2.421% | 2.546% | 2.378% | 1.85% | 1.701% | 1.978%
| PDM3 |  | 5.877% | -4.544% | -5.371% | -0.3741% | 4.062% | 9.477% | 25.66% | 0.7979% | 0.5491% | 6.725% | -2.817% | -4.939% | -4.281% | -4.061% | -3.695% | -6.359% | 5.137% | 30.99% | -6.33% | 3.039% | -24.62% | 2.344% | -12.74% | 17.4%
| Z-Score (Absolute) |  | 0.77 | 0.62 | 0.73 | 0.06 | 0.53 | 1.26 | 3.42 | 0.1 | 0.06 | 0.89 | 0.39 | 0.67 | 0.58 | 0.56 | 0.51 | 0.86 | 0.68 | 4.14 | 0.86 | 0.4 | 3.31 | 0.3 | 1.72 | 2.32
| Analytical results for Gd in OREAS 600c (Certified Value 5.96 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.09: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.05: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  | 2
3
| 1 | 1 | 2.01: 6.6; 2.03: 5.81; 2.05: 5.3528; 2.06: 6.4; 2.07: 8.5957; 2.08: 5.93; 2.09: 6; 2.10: 6.27; 2.11: 6.6756; 2.12: 4; 2.13: 5.4 |  |  | 1
| 1 | 2 | 2.01: 6; 2.03: 5.78; 2.05: 5.7544; 2.06: 6.2; 2.07: 8.445; 2.08: 6.08; 2.09: 6.6; 2.10: 6.09; 2.11: 6.6068; 2.12: 4; 2.13: 5.5 |  |  | 37
| 1 | 3 | 2.01: 6.4; 2.03: 5.8; 2.05: 5.6758; 2.06: 6.1; 2.07: 8.6732; 2.08: 5.94; 2.09: 5.9; 2.10: 6.24; 2.11: 6.5408; 2.12: 4; 2.13: 5.3 |  |  | 16
| 1 | 4 | 2.01: 6.5; 2.03: 6.02; 2.05: 5.7155; 2.06: 6.3; 2.07: 8.6223; 2.08: 5.91; 2.09: 4.6; 2.10: 6.19; 2.11: 6.6508; 2.12: 3; 2.13: 5.3 |  |  | 5.9607
| 1 | 5 | 2.01: 6; 2.03: 5.89; 2.05: 5.6432; 2.06: 6.1; 2.07: 8.9654; 2.08: 6.05; 2.09: 5.1; 2.10: 6.06; 2.11: 6.6394; 2.12: 3; 2.13: 5.3 |  |  | 31
| 1 | 6 | 2.01: 5.8; 2.03: 5.84; 2.05: 5.4331; 2.06: 6.4; 2.07: 8.5802; 2.08: 6.23; 2.09: 5.5; 2.10: 6.11; 2.11: 6.4479; 2.12: 3; 2.13: 5.4
| Mean |  | 2.01: 6.2167; 2.03: 5.8567; 2.05: 5.5958; 2.06: 6.25; 2.07: 8.647; 2.08: 6.0233; 2.09: 5.6167; 2.10: 6.16; 2.11: 6.5935; 2.12: 3.5; 2.13: 5.3667
| Median |  | 2.01: 6.2; 2.03: 5.825; 2.05: 5.6595; 2.06: 6.25; 2.07: 8.609; 2.08: 5.995; 2.09: 5.7; 2.10: 6.15; 2.11: 6.6231; 2.12: 3.5; 2.13: 5.35
| Std Dev. |  | 2.01: 0.32506; 2.03: 0.088694; 2.05: 0.16351; 2.06: 0.13784; 2.07: 0.17354; 2.08: 0.12258; 2.09: 0.70828; 2.10: 0.08579; 2.11: 0.085187; 2.12: 0.54772; 2.13: 0.08165
| Rel.Std.Dev. |  | 2.01: 5.229%; 2.03: 1.514%; 2.05: 2.922%; 2.06: 2.205%; 2.07: 2.007%; 2.08: 2.035%; 2.09: 12.61%; 2.10: 1.393%; 2.11: 1.292%; 2.12: 15.65%; 2.13: 1.521%
| PDM3 |  | 2.01: 4.293%; 2.03: -1.746%; 2.05: -6.122%; 2.06: 4.853%; 2.07: 45.07%; 2.08: 1.05%; 2.09: -5.772%; 2.10: 3.343%; 2.11: 10.62%; 2.12: -41.28%; 2.13: -9.967%
| Z-Score (Absolute) |  | 2.01: 0.24; 2.03: 0.45; 2.05: 0.94; 2.06: 0.3; 2.07: 4.85; 2.08: 0.13; 2.09: 0.9; 2.10: 0.13; 2.11: 0.95; 2.12: NA; 2.13: 1.38
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 600c (Certified Value 0.2 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 0.27 | 0.21 | 1.3 | < 0.2 | 0.1 | 0.36 | 0.2542 | 1.3 | 1.4 | 0.18 | 0.2 | 0.18 | 0.18 | 0.16 | 0.09 | < 0.2 | 0.35 | 0.12 | 0.28005 | 0.9 | [6 empty] | 1
| 1 | 2 | 0.28 | 0.14 | 1.4 | < 0.2 | < 0.1 | 0.5 | 0.2938 | 1.3 | 1.2 | 0.18 | 0.18 | 0.19 | 0.17 | 0.18 | 0.08 | < 0.2 | 0.33 | 0.11 | 0.28102 | 1.1 | [6 empty] | 38
| 1 | 3 | 0.26 | 0.16 | 1.3 | < 0.2 | 0.1 | 0.52 | 0.2277 | 1.4 | 1.4 | 0.17 | 0.16 | 0.22 | 0.17 | 0.16 | 0.05 | < 0.2 | 0.32 | 0.14 | 0.27006 | 1.3 | [6 empty] | 16
| 1 | 4 | 0.27 | 0.15 | 1.2 | < 0.2 | 0.1 | 0.35 | 0.2687 | 1.3 | 1.5 | 0.16 | 0.17 | 0.19 | 0.17 | 0.17 | 0.06 | 0.4 | 0.33 | 0.12 | 0.25669 | 1.2 | [6 empty] | 0.19766
| 1 | 5 | 0.23 | 0.13 | 1.4 | < 0.2 | 0.1 | 0.36 | 0.2987 | 1.3 | 1.4 | 0.17 | 0.2 | 0.2 | 0.19 | 0.18 | 0.07 | < 0.2 | 0.35 | 0.13 | 0.25008 | 1.1 | [6 empty] | 32
| 1 | 6 | 0.27 | 0.14 | 1.3 | < 0.2 | < 0.1 | 0.51 | 0.2639 | 1.3 | 1.4 | 0.16 | 0.17 | 0.21 | 0.19 | 0.18 | 0.08 | < 0.2 | 0.3 | 0.12 | 0.26244 | 0.9
| Mean |  | 0.26333 | 0.155 | 1.3167 |  | 0.1 | 0.43333 | 0.26783 | 1.3167 | 1.3833 | 0.17 | 0.18 | 0.19833 | 0.17833 | 0.17167 | 0.071667 | 0.4 | 0.33 | 0.12333 | 0.26672 | 1.0833
| Median |  | 0.27 | 0.145 | 1.3 |  | 0.1 | 0.43 | 0.2663 | 1.3 | 1.4 | 0.17 | 0.175 | 0.195 | 0.175 | 0.175 | 0.075 | 0.4 | 0.33 | 0.12 | 0.26625 | 1.1
| Std Dev. |  | 0.017512 | 0.02881 | 0.075277 |  | 0 | 0.084301 | 0.026229 | 0.040825 | 0.098319 | 0.0089443 | 0.016733 | 0.01472 | 0.0098319 | 0.0098319 | 0.01472 |  | 0.018974 | 0.010328 | 0.012563 | 0.16021
| Rel.Std.Dev. |  | 6.65% | 18.59% | 5.717% |  | 0% | 19.45% | 9.793% | 3.101% | 7.107% | 5.261% | 9.296% | 7.422% | 5.513% | 5.727% | 20.54% |  | 5.75% | 8.374% | 4.71% | 14.79%
| PDM3 |  | 33.23% | -21.58% | 566.1% |  | -49.41% | 119.2% | 35.5% | 566.1% | 599.9% | -13.99% | -8.933% | 0.3419% | -9.777% | -13.15% | -63.74% | 102.4% | 66.96% | -37.6% | 34.94% | 448.1%
| Z-Score (Absolute) |  | 0.7 | 0.21 | NA | 0.67 | NA | 2.14 | 0.74 | NA | 10.14 | 0.08 | 0 | 0.15 | 0.01 | 0.07 | 0.91 | NA | 1.26 | 0.48 | 0.73 | NA
Results from laboratories 2.05, 2.07, 2.11, 2.19 and 2.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 600c (Certified Value 6.86 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 5.1 | 7.28 | 6.4 | 6.73 | 6.6304 | 7.2 | 7.24 | 6.57 | 6.4 | 6.58 | 6.9 | 6.8 | 6.7 | 7.2 | 6.9 | 6.29 | 8.26 | 7.1 | 6.9 | 10 | 6.9897 | 6.8 |  |  |  |  | 1
| 1 | 2 | 5.1 | 6.59 | 6.6 | 6.67 | 6.8177 | 7.2 | 7.34 | 6.83 | 6.7 | 6.72 | 7.2 | 6.7 | 6.9 | 6.8 | 7.2 | 6.25 | 8.35 | 7.1 | 6.6 | 10 | 7.0526 | 6.96 |  |  |  |  | 22
| 1 | 3 | 5.4 | 7.04 | 6.3 | 6.77 | 6.7785 | 7 | 7.25 | 6.57 | 6.5 | 6.8 | 6.9 | 6.6 | 7 | 6.9 | 7.1 | 5.94 | 8.13 | 7.3 | 7.1 | 9 | 7.1115 | 6.66 |  |  |  |  | 16
| 1 | 4 | 5.4 | 6.92 | 6.6 | 6.84 | 6.7418 | 7.4 | 7.18 | 6.62 | 6.7 | 6.76 | 7.1 | 6.7 | 6.8 | 6.9 | 7.3 | 6.37 | 8.12 | 7.4 | 7 | 9 | 7.1754 | 6.65 |  |  |  |  | 6.8578
| 1 | 5 | 4.9 | 6.76 | 6.6 | 6.85 | 6.7779 | 7 | 7.42 | 6.64 | 6.5 | 6.68 | 6.9 | 6.9 | 6.8 | 7 | 7 | 6.3 | 7.83 | 7.1 | 6.8 | 9 | 7.1546 | 6.64 |  |  |  |  | 33
| 1 | 6 | 5.3 | 6.45 | 6.7 | 7.13 | 6.6512 | 7.4 | 7.22 | 6.82 | 6.4 | 6.76 | 7 | 6.7 | 7 | 7 | 7.2 | 6.36 | 8.04 | 7.4 | 7 | 11 | 7.0576 | 6.73
| Mean |  | 5.2 | 6.84 | 6.5333 | 6.8317 | 6.7329 | 7.2 | 7.275 | 6.675 | 6.5333 | 6.7167 | 7 | 6.7333 | 6.8667 | 6.9667 | 7.1167 | 6.2517 | 8.1217 | 7.2333 | 6.9 | 9.6667 | 7.0902 | 6.74
| Median |  | 5.2 | 6.84 | 6.6 | 6.805 | 6.7598 | 7.2 | 7.245 | 6.63 | 6.5 | 6.74 | 6.95 | 6.7 | 6.85 | 6.95 | 7.15 | 6.295 | 8.125 | 7.2 | 6.95 | 9.5 | 7.0846 | 6.695
| Std Dev. |  | 0.2 | 0.30364 | 0.15055 | 0.16105 | 0.075567 | 0.17889 | 0.088487 | 0.11946 | 0.13663 | 0.078401 | 0.12649 | 0.10328 | 0.12111 | 0.13663 | 0.1472 | 0.15917 | 0.1806 | 0.15055 | 0.17889 | 0.8165 | 0.069913 | 0.12377
| Rel.Std.Dev. |  | 3.846% | 4.439% | 2.304% | 2.357% | 1.122% | 2.485% | 1.216% | 1.79% | 2.091% | 1.167% | 1.807% | 1.534% | 1.764% | 1.961% | 2.068% | 2.546% | 2.224% | 2.081% | 2.593% | 8.447% | 0.986% | 1.836%
| PDM3 |  | -24.17% | -0.26% | -4.732% | -0.3815% | -1.822% | 4.989% | 6.083% | -2.666% | -4.732% | -2.058% | 2.073% | -1.815% | 0.1288% | 1.587% | 3.774% | -8.839% | 18.43% | 5.476% | 0.6149% | 40.96% | 3.389% | -1.718%
| Z-Score (Absolute) |  | 6.7 | 0 | 1.25 | 0.03 | 0.44 | 1.47 | 1.78 | 0.67 | 1.25 | 0.5 | 0.65 | 0.44 | 0.11 | 0.52 | 1.13 | 2.4 | 5.24 | 1.61 | 0.25 | NA | 1.02 | 0.41
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 600c (Indicative Value 0.043 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab |  |  |  |  | 1
| No | No | 2.01: 2.06; 2.03: 2.18; 2.04: 2.22 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 2.01: < 1; 2.03: 0.05; 2.04: < 2 |  |  |  |  | 1
| 1 | 2 | 2.01: < 1; 2.03: 0.04; 2.04: < 2 |  |  |  |  | 3
| 1 | 3 | 2.01: < 1; 2.03: 0.05; 2.04: < 2 |  |  |  |  | 16
| 1 | 4 | 2.01: < 1; 2.03: 0.04; 2.04: < 2 |  |  |  |  | 0.043333
| 1 | 5 | 2.01: < 1; 2.03: 0.04; 2.04: < 2 |  |  |  |  | 9
| 1 | 6 | 2.01: < 1; 2.03: 0.04; 2.04: < 2
| Mean |  | 2.03: 0.043333
| Median |  | 2.03: 0.04
| Std Dev. |  | 2.03: 0.005164
| Rel.Std.Dev. |  | 2.03: 11.92%
| PDM3 |  | 2.03: 8.882e-14%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.67; 2.04: 0.74
| Analytical results for Ho in OREAS 600c (Certified Value 0.51 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2.01: 0.6; 2.03: 0.5; 2.04: 0.46667; 2.05: 0.6; 2.06: 0.5729; 2.07: 0.52; 2.08: 0.52; 2.10: 0.56; 2.11: 0.4623; 2.12: < 1; 2.13: 0.46 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.6; 2.03: 0.52; 2.04: 0.45931; 2.05: 0.6; 2.06: 0.5713; 2.07: 0.55; 2.08: 0.56; 2.10: 0.56; 2.11: 0.4284; 2.12: < 1; 2.13: 0.47 |  |  |  |  | 23
| 1 | 3 | 2.01: 0.6; 2.03: 0.51; 2.04: 0.48963; 2.05: 0.6; 2.06: 0.615; 2.07: 0.53; 2.08: 0.51; 2.10: 0.56; 2.11: 0.45; 2.12: < 1; 2.13: 0.45 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.6; 2.03: 0.52; 2.04: 0.489; 2.05: 0.6; 2.06: 0.5464; 2.07: 0.51; 2.08: 0.44; 2.10: 0.58; 2.11: 0.4465; 2.12: < 1; 2.13: 0.47 |  |  |  |  | 0.51442
| 1 | 5 | 2.01: 0.6; 2.03: 0.51; 2.04: 0.49797; 2.05: 0.6; 2.06: 0.5726; 2.07: 0.52; 2.08: 0.43; 2.10: 0.55; 2.11: 0.4502; 2.12: < 1; 2.13: 0.46 |  |  |  |  | 34
| 1 | 6 | 2.01: 0.6; 2.03: 0.5; 2.04: 0.45808; 2.05: 0.6; 2.06: 0.563; 2.07: 0.52; 2.08: 0.49; 2.10: 0.55; 2.11: 0.4391; 2.12: < 1; 2.13: 0.45
| Mean |  | 2.01: 0.6; 2.03: 0.51; 2.04: 0.47678; 2.05: 0.6; 2.06: 0.57353; 2.07: 0.525; 2.08: 0.49167; 2.10: 0.56; 2.11: 0.44608; 2.13: 0.46
| Median |  | 2.01: 0.6; 2.03: 0.51; 2.04: 0.47784; 2.05: 0.6; 2.06: 0.57195; 2.07: 0.52; 2.08: 0.5; 2.10: 0.56; 2.11: 0.44825; 2.13: 0.46
| Std Dev. |  | 2.01: 0; 2.03: 0.0089443; 2.04: 0.01744; 2.05: 0; 2.06: 0.022684; 2.07: 0.013784; 2.08: 0.049565; 2.10: 0.010954; 2.11: 0.011462; 2.13: 0.0089443
| Rel.Std.Dev. |  | 2.01: 0%; 2.03: 1.754%; 2.04: 3.658%; 2.05: 0%; 2.06: 3.955%; 2.07: 2.626%; 2.08: 10.08%; 2.10: 1.956%; 2.11: 2.57%; 2.13: 1.944%
| PDM3 |  | 2.01: 16.64%; 2.03: -0.8589%; 2.04: -7.317%; 2.05: 16.64%; 2.06: 11.49%; 2.07: 2.057%; 2.08: -4.423%; 2.10: 8.861%; 2.11: -13.28%; 2.13: -10.58%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.09; 2.04: 0.52; 2.05: 1.75; 2.06: 1.26; 2.07: 0.37; 2.08: 0.25; 2.10: 1.01; 2.11: 1.09; 2.12: 0.09; 2.13: 0.83
Results from laboratory 2.03 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 600c (Certified Value 0.18 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 0.169 | 0.17 | 0.16 | 0.18 | 0.2 | 0.19 | 0.2 | 0.19 | 0.18 | 0.181 | 0.173 | 0.181 | 0.184 | 0.186 | 0.173 | 0.195 | 0.193 | 0.151 | 0.17338 | 0.184 | [6 empty] | 1
| 1 | 2 | 0.162 | 0.15 | 0.17 | 0.18 | 0.2 | 0.192 | 0.2 | 0.18 | 0.17 | 0.18 | 0.178 | 0.193 | 0.182 | 0.189 | 0.168 | 0.192 | 0.189 | 0.143 | 0.16336 | 0.172 | [6 empty] | 24
| 1 | 3 | 0.164 | 0.17 | 0.15 | 0.17 | 0.2 | 0.174 | 0.2 | 0.18 | 0.18 | 0.167 | 0.165 | 0.192 | 0.185 | 0.186 | 0.175 | 0.192 | 0.204 | 0.145 | 0.16986 | 0.204 | [6 empty] | 16
| 1 | 4 | 0.172 | 0.17 | 0.16 | 0.17 | 0.2 | 0.184 | 0.2 | 0.19 | 0.19 | 0.169 | 0.178 | 0.195 | 0.179 | 0.186 | 0.177 | 0.188 | 0.195 | 0.131 | 0.16033 | 0.185 | [6 empty] | 0.17783
| 1 | 5 | 0.15 | 0.16 | 0.16 | 0.17 | 0.2 | 0.182 | 0.1 | 0.18 | 0.17 | 0.169 | 0.182 | 0.188 | 0.174 | 0.18 | 0.173 | 0.183 | 0.182 | 0.128 | 0.16827 | 0.165 | [6 empty] | 35
| 1 | 6 | 0.168 | 0.15 | 0.17 | 0.18 | 0.2 | 0.183 | 0.1 | 0.19 | 0.18 | 0.177 | 0.168 | 0.195 | 0.179 | 0.195 | 0.172 | 0.19 | 0.218 | 0.135 | 0.1768 | 0.175
| Mean |  | 0.16417 | 0.16167 | 0.16167 | 0.175 | 0.2 | 0.18417 | 0.16667 | 0.185 | 0.17833 | 0.17383 | 0.174 | 0.19067 | 0.1805 | 0.187 | 0.173 | 0.19 | 0.19683 | 0.13883 | 0.16867 | 0.18083
| Median |  | 0.166 | 0.165 | 0.16 | 0.175 | 0.2 | 0.1835 | 0.2 | 0.185 | 0.18 | 0.173 | 0.1755 | 0.1925 | 0.1805 | 0.186 | 0.173 | 0.191 | 0.194 | 0.139 | 0.16907 | 0.1795
| Std Dev. |  | 0.0078081 | 0.0098319 | 0.0075277 | 0.0054772 | 3.0405e-17 | 0.0064005 | 0.05164 | 0.0054772 | 0.0075277 | 0.0062102 | 0.0065422 | 0.0053914 | 0.0040373 | 0.004899 | 0.0030332 | 0.0041473 | 0.01264 | 0.0089088 | 0.0061251 | 0.013615
| Rel.Std.Dev. |  | 4.756% | 6.082% | 4.656% | 3.13% | 1.52e-14% | 3.475% | 30.98% | 2.961% | 4.221% | 3.573% | 3.76% | 2.828% | 2.237% | 2.62% | 1.753% | 2.183% | 6.422% | 6.417% | 3.631% | 7.529%
| PDM3 |  | -7.681% | -9.087% | -9.087% | -1.589% | 12.47% | 3.566% | -6.275% | 4.035% | 0.2855% | -2.245% | -2.151% | 7.221% | 1.504% | 5.159% | -2.714% | 6.846% | 10.69% | -21.93% | -5.15% | 1.691%
| Z-Score (Absolute) |  | 1.03 | 1.24 | 1.24 | 0.14 | NA | 0.62 | NA | 0.69 | 0.14 | 0.23 | 0.22 | 1.16 | 0.32 | 0.85 | 0.3 | 1.1 | 1.67 | 3.12 | 0.66 | 0.34
Results from laboratories 2.07 and 2.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 600c (Certified Value 3.11 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.21 | 2.91 | 3.35 | 3.1342 | 3.0929 | 3.23 | 3.15 | 3.0139 | 3.11 | 3.02 | 3.1178 | 3.22 | 3.04 | 2.97 | 2.91 | 3.25 | 2.999 | 3.2 | 3.17 | 2.92 | 3.1 | 2.7422 | 3.03 | 3.201 | 2.9882 | 3.15 | 1
| 1 | 2 | 3.19 | 3.24 | 3.39 | 3.1236 | 3.0958 | 3.28 | 3.08 | 3.0352 | 3.24 | 3.07 | 3.0788 | 3.14 | 3.02 | 2.98 | 2.88 | 3.27 | 3.0029 | 3.22 | 2.57 | 2.94 | 2.93 | 2.7013 | 3.096 | 3.1488 | 3.4045 | 3.16 | 0
| 1 | 3 | 3.26 | 3.14 | 3.29 | 3.0975 | 3.0861 | 3.36 | 2.78 | 3.0218 | 3.22 | 3.08 | 3.0339 | 3.21 | 3.06 | 2.98 | 2.89 | 3.23 | 2.9903 | 3.19 | 2.74 | 2.95 | 3.11 | 2.7426 | 2.997 | 3.0534 | 3.1271 | 3.19 | 16
| 1 | 4 | 3.21 | 3.21 | 3.38 | 3.0899 | 3.0858 | 3.44 | 3.04 | 3.0068 | 3.2 | 3 | 3.1133 | 3.19 | 3.04 | 2.97 | 2.9 | 3.38 | 3.034 | 3.08 | 2.6 | 2.94 | 3.12 | 2.6634 | 2.98 | 3.027 | 3.514 | 3.02 | 3.1143
| 1 | 5 | 3.23 | 3.46 | 3.4 | 3.1824 | 3.099 | 3.35 | 3.05 | 3.029 | 3.2 | 2.95 | 3.045 | 3.13 | 3.08 | 2.94 | 2.9 | 3.23 | 2.9618 | 3.17 | 2.62 | 2.97 | 3.02 | 2.7367 | 3.005 | 3.0491 | 3.4993 | 3.11 | 36
| 1 | 6 | 3.19 | 3.09 | 3.38 | 3.0528 | 3.0313 | 3.36 | 3.04 | 3.0354 | 3.16 | 3.05 | 3.0978 | 3.15 | 3.02 | 3.04 | 2.91 | 3.35 | 3.0323 | 3.19 | 2.48 | 2.95 | 2.83 | 2.5833 | 3.03 | 2.932 | 3.6094 | 3.12
| Mean |  | 3.215 | 3.175 | 3.365 | 3.1134 | 3.0818 | 3.3367 | 3.0233 | 3.0237 | 3.1883 | 3.0283 | 3.0811 | 3.1733 | 3.0433 | 2.98 | 2.8983 | 3.285 | 3.0034 | 3.175 | 2.6967 | 2.945 | 3.0183 | 2.6949 | 3.023 | 3.0686 | 3.3571 | 3.125
| Median |  | 3.21 | 3.175 | 3.38 | 3.1105 | 3.0895 | 3.355 | 3.045 | 3.0254 | 3.2 | 3.035 | 3.0883 | 3.17 | 3.04 | 2.975 | 2.9 | 3.26 | 3.001 | 3.19 | 2.61 | 2.945 | 3.06 | 2.719 | 3.0175 | 3.0513 | 3.4519 | 3.135
| Std Dev. |  | 0.026646 | 0.18185 | 0.040373 | 0.044216 | 0.025302 | 0.072847 | 0.12628 | 0.011693 | 0.046655 | 0.048751 | 0.035211 | 0.038297 | 0.023381 | 0.032863 | 0.01169 | 0.06442 | 0.027178 | 0.049295 | 0.24663 | 0.016432 | 0.1172 | 0.062845 | 0.040684 | 0.094843 | 0.24486 | 0.058907
| Rel.Std.Dev. |  | 0.8288% | 5.728% | 1.2% | 1.42% | 0.821% | 2.183% | 4.177% | 0.3867% | 1.463% | 1.61% | 1.143% | 1.207% | 0.7683% | 1.103% | 0.4034% | 1.961% | 0.9049% | 1.553% | 9.146% | 0.558% | 3.883% | 2.332% | 1.346% | 3.091% | 7.294% | 1.885%
| PDM3 |  | 3.232% | 1.948% | 8.048% | -0.03043% | -1.045% | 7.139% | -2.922% | -2.911% | 2.376% | -2.762% | -1.068% | 1.894% | -2.28% | -4.314% | -6.936% | 5.48% | -3.563% | 1.948% | -13.41% | -5.438% | -3.083% | -13.47% | -2.933% | -1.47% | 7.794% | 0.342%
| Z-Score (Absolute) |  | 0.98 | 0.7 | 2.02 | 0.27 | 0.05 | 1.83 | 0.36 | 0.36 | 0.79 | 0.32 | 0.04 | 0.69 | 0.22 | 0.66 | 1.23 | 1.47 | 0.5 | 0.7 | 2.64 | 0.91 | 0.39 | 2.65 | 0.36 | 0.04 | 1.97 | 0.35
| Analytical results for La in OREAS 600c (Certified Value 45.5 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 48 | 46 | 49.5 | 46.5 | 41.0073 | 49.2 | 53.9 | 52.918 | 44.93 | 44.26 | 43.4 | 44.6 | 43.6 | 43.9 | 44.8 | 44.6 | 36.91 | 45 | 37.57 | 46 | 40.711 | 48 | 36.944 | 41.63 |  |  | 1
| 1 | 2 | 49 | 44.2 | 49.6 | 47.58 | 41.4173 | 48.6 | 53.9 | 52.127 | 47.97 | 46.68 | 44.6 | 42 | 43.9 | 43.7 | 45.4 | 44.5 | 40.58 | 51 | 38.05 | 47 | 40.5626 | 48 | 37.059 | 44.04 |  |  | 25
| 1 | 3 | 48 | 46.7 | 49.4 | 47.82 | 42.1296 | 48.3 | 52.8 | 52.181 | 46.3 | 45.64 | 42.3 | 43.9 | 44.2 | 43.9 | 46.5 | 43.9 | 35.42 | 50 | 37.44 | 46 | 40.6326 | 47 | 37.743 | 40.47 |  |  | 16
| 1 | 4 | 48 | 45.2 | 50.8 | 48.49 | 42.0779 | 49.9 | 53 | 52.699 | 45.87 | 42.6 | 43.1 | 42.2 | 43.9 | 47.1 | 47.1 | 44.9 | 23.56 | 51 | 38.8 | 46 | 41.954 | 45 | 36.009 | 42.33 |  |  | 45.475
| 1 | 5 | 48 | 42.8 | 49.2 | 48.73 | 41.6402 | 48.4 | 53.8 | 52.335 | 45.23 | 46.36 | 43.6 | 44.4 | 42.8 | 46.6 | 45.7 | 44.1 | 28.33 | 50 | 38.6 | 46 | 42.6847 | 47 | 37.475 | 41.1 |  |  | 37
| 1 | 6 | 48 | 41.4 | 49.5 | 46.57 | 41.1998 | 49.5 | 53.8 | 52.451 | 46.97 | 46.29 | 44.3 | 42.4 | 45.5 | 45.6 | 46.3 | 44.8 | 31.84 | 50 | 38.28 | 46 | 40.6115 | 46 | 38.223 | 42.55
| Mean |  | 48.1667 | 44.3833 | 49.6667 | 47.615 | 41.5787 | 48.9833 | 53.5333 | 52.4518 | 46.2117 | 45.305 | 43.55 | 43.25 | 43.9833 | 45.1333 | 45.9667 | 44.4667 | 32.7733 | 49.5 | 38.1233 | 46.1667 | 41.1927 | 46.8333 | 37.2422 | 42.02
| Median |  | 48 | 44.7 | 49.5 | 47.7 | 41.5288 | 48.9 | 53.8 | 52.393 | 46.085 | 45.965 | 43.5 | 43.15 | 43.9 | 44.75 | 46 | 44.55 | 33.63 | 50 | 38.165 | 46 | 40.6718 | 47 | 37.267 | 41.98
| Std Dev. |  | 0.40825 | 2.0044 | 0.57155 | 0.93673 | 0.45883 | 0.64936 | 0.49666 | 0.30689 | 1.1312 | 1.5808 | 0.83126 | 1.1794 | 0.88412 | 1.5055 | 0.82865 | 0.39328 | 6.1735 | 2.2583 | 0.54555 | 0.40825 | 0.90401 | 1.169 | 0.76292 | 1.2537
| Rel.Std.Dev. |  | 0.8476% | 4.516% | 1.151% | 1.967% | 1.104% | 1.326% | 0.9277% | 0.5851% | 2.448% | 3.489% | 1.909% | 2.727% | 2.01% | 3.336% | 1.803% | 0.8844% | 18.84% | 4.562% | 1.431% | 0.8843% | 2.195% | 2.496% | 2.049% | 2.983%
| PDM3 |  | 5.919% | -2.401% | 9.217% | 4.706% | -8.568% | 7.715% | 17.72% | 15.34% | 1.62% | -0.3739% | -4.233% | -4.893% | -3.28% | -0.7514% | 1.081% | -2.217% | -27.93% | 8.851% | -16.17% | 1.521% | -9.417% | 2.987% | -18.1% | -7.598%
| Z-Score (Absolute) |  | 0.74 | 0.21 | 1.12 | 0.6 | 0.92 | 0.95 | 2.1 | 1.83 | 0.25 | 0.02 | 0.42 | 0.5 | 0.31 | 0.02 | 0.19 | 0.19 | 3.14 | 1.08 | 1.79 | 0.24 | 1.02 | 0.41 | 2.01 | 0.81
| Analytical results for Li in OREAS 600c (Certified Value 30.6 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 30.2 | 32.5 | 31.4 | 29.6 | 30.6714 | 33 | 28.2 | 39.2795 | 28.8 | 31.3 | 31.5 | 31.4 | 31 | 29.9 | 29.8 | 31.9 | 32 | 28.7 | 31.7 | 30.23 | 28.7 | 51.8531 | 31 | 31.7692 | 41.194 | 31.3 | 1
| 1 | 2 | 31.3 | 29.5 | 32.2 | 29.6 | 30.4903 | 33 | 28.9 | 39.7649 | 29.6 | 31.4 | 30.8 | 30.8 | 30.5 | 30 | 29.4 | 32.1 | 32 | 30.3 | 31.4 | 30.12 | 30.2 | 51.0931 | 31 | 33.4728 | 43.151 | 30.9 | 45
| 1 | 3 | 31.5 | 30.7 | 30.5 | 30.4 | 30.6002 | 33 | 27.5 | 39.9493 | 29.5 | 31.4 | 31.3 | 31.5 | 30.8 | 29.9 | 29.3 | 31.9 | 30 | 28.6 | 31.1 | 30.19 | 28.2 | 51.5822 | 30 | 33.481 | 41.772 | 30.8 | 16
| 1 | 4 | 30.2 | 30.4 | 31.5 | 30.3 | 30.2929 | 33 | 27.6 | 39.1106 | 29.7 | 31.3 | 31.4 | 30.9 | 30.9 | 29.9 | 29.4 | 32.5 | 32 | 25.8 | 31.8 | 31.09 | 28.9 | 53.8817 | 30 | 32.552 | 43.317 | 30.6 | 30.6456
| 1 | 5 | 28.9 | 29.3 | 31.5 | 30.6 | 30.5448 | 32 | 27 | 39.1058 | 29.1 | 31.6 | 30.3 | 31.1 | 32.5 | 29.8 | 29.7 | 31.9 | 31 | 28.8 | 31.6 | 31.22 | 26.6 | 51.4461 | 30 | 32.3554 | 42.853 | 30.6 | 38
| 1 | 6 | 31 | 28.1 | 32.2 | 29.6 | 30.3258 | 33 | 27.6 | 39.2255 | 29.2 | 31.3 | 30.4 | 30.9 | 32.9 | 30.6 | 29.7 | 32.8 | 32 | 30 | 30.9 | 31.09 | 29.3 | 57.8455 | 30 | 32.1377 | 43.314 | 30.9
| Mean |  | 30.5167 | 30.0833 | 31.55 | 30.0167 | 30.4876 | 32.8333 | 27.8 | 39.4059 | 29.3167 | 31.3833 | 30.95 | 31.1 | 31.4333 | 30.0167 | 29.55 | 32.1833 | 31.5 | 28.7 | 31.4167 | 30.6567 | 28.65 | 52.9503 | 30.3333 | 32.628 | 42.6002 | 30.85
| Median |  | 30.6 | 29.95 | 31.5 | 29.95 | 30.5176 | 33 | 27.6 | 39.2525 | 29.35 | 31.35 | 31.05 | 31 | 30.95 | 29.9 | 29.55 | 32 | 32 | 28.75 | 31.5 | 30.66 | 28.8 | 51.7176 | 30 | 32.4537 | 43.002 | 30.85
| Std Dev. |  | 0.96212 | 1.4972 | 0.62849 | 0.46655 | 0.1509 | 0.40825 | 0.6603 | 0.36053 | 0.34303 | 0.1169 | 0.5244 | 0.28983 | 1.0033 | 0.29269 | 0.20736 | 0.38166 | 0.83666 | 1.5925 | 0.35449 | 0.5255 | 1.2079 | 2.593 | 0.5164 | 0.70699 | 0.90043 | 0.25884
| Rel.Std.Dev. |  | 3.153% | 4.977% | 1.992% | 1.554% | 0.495% | 1.243% | 2.375% | 0.9149% | 1.17% | 0.3725% | 1.694% | 0.9319% | 3.192% | 0.9751% | 0.7017% | 1.186% | 2.656% | 5.549% | 1.128% | 1.714% | 4.216% | 4.897% | 1.702% | 2.167% | 2.114% | 0.839%
| PDM3 |  | -0.4206% | -1.835% | 2.951% | -2.052% | -0.5156% | 7.139% | -9.285% | 28.59% | -4.336% | 2.407% | 0.9934% | 1.483% | 2.571% | -2.052% | -3.575% | 5.018% | 2.788% | -6.349% | 2.516% | 0.03625% | -6.512% | 72.78% | -1.019% | 6.469% | 39.01% | 0.6671%
| Z-Score (Absolute) |  | 0.31 | 0.63 | 0.43 | 0.67 | 0.33 | 1.36 | 2.28 | 6.11 | 1.18 | 0.31 | 0 | 0.11 | 0.35 | 0.67 | 1.01 | 0.89 | 0.4 | 1.63 | 0.34 | 0.21 | 1.66 | 15.89 | NA | 1.21 | 8.42 | 0.07
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 600c (Certified Value 0.075 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.12; 2.07: 2.19; 2.08: 2.22; 2.09: 2.23; 2.10: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS | 3
3
| 1 | 1 | 2.01: < 0.1; 2.03: 0.06; 2.04: 0.091415; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.07; 2.09: 0.1019; 2.10: 0.05 | 1
| 1 | 2 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.091506; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.1; 2.08: 0.07; 2.09: 0.0936; 2.10: 0.05 | 26
| 1 | 3 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.092635; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.0996; 2.10: 0.05 | 16
| 1 | 4 | 2.01: < 0.1; 2.03: 0.06; 2.04: 0.090141; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.07; 2.08: 0.07; 2.09: 0.1051; 2.10: 0.06 | 0.075468
| 1 | 5 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.09314; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.08; 2.08: 0.07; 2.09: 0.1021; 2.10: 0.03 | 39
| 1 | 6 | 2.01: < 0.1; 2.03: 0.07; 2.04: 0.089715; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.09; 2.08: 0.08; 2.09: 0.1088; 2.10: 0.05
| Mean |  | 2.03: 0.066667; 2.04: 0.091425; 2.06: 0.06; 2.07: 0.086667; 2.08: 0.073333; 2.09: 0.10185; 2.10: 0.048333
| Median |  | 2.03: 0.07; 2.04: 0.09146; 2.06: 0.06; 2.07: 0.09; 2.08: 0.07; 2.09: 0.102; 2.10: 0.05
| Std Dev. |  | 2.03: 0.005164; 2.04: 0.0013406; 2.06: 0; 2.07: 0.010328; 2.08: 0.005164; 2.09: 0.0051392; 2.10: 0.0098319
| Rel.Std.Dev. |  | 2.03: 7.746%; 2.04: 1.466%; 2.06: 0%; 2.07: 11.92%; 2.08: 7.042%; 2.09: 5.046%; 2.10: 20.34%
| PDM3 |  | 2.03: -11.66%; 2.04: 21.14%; 2.06: -20.5%; 2.07: 14.84%; 2.08: -2.828%; 2.09: 34.96%; 2.10: -35.96%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0; 2.04: 1; 2.05: 0.67; 2.06: 0.27; 2.07: 0.81; 2.08: 0.27; 2.09: 1.42; 2.10: 0.74
| Analytical results for Mg in OREAS 600c (Certified Value 0.149 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.16 | 0.15 | 0.16 | 0.1495 | 0.14977 | 0.15 | 0.16 | 0.16448 | 0.15 | 0.158 | 0.1512 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11 | 0.1489 | 0.15 | 0.16 | 0.15 | 0.15 | 0.1404 | 0.151 | 0.15295 | 0.13175 | 0.141 | 1
| 1 | 2 | 0.16 | 0.15 | 0.16 | 0.1497 | 0.14923 | 0.15 | 0.16 | 0.1609 | 0.15 | 0.155 | 0.1497 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11 | 0.1484 | 0.153 | 0.15 | 0.16 | 0.14 | 0.1368 | 0.157 | 0.15948 | 0.14208 | 0.141 | 0
| 1 | 3 | 0.16 | 0.15 | 0.16 | 0.1484 | 0.14991 | 0.16 | 0.15 | 0.16897 | 0.15 | 0.154 | 0.1476 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11 | 0.1498 | 0.148 | 0.16 | 0.16 | 0.15 | 0.1389 | 0.151 | 0.15891 | 0.13978 | 0.136 | 16
| 1 | 4 | 0.16 | 0.14 | 0.16 | 0.1472 | 0.14884 | 0.16 | 0.15 | 0.16654 | 0.15 | 0.151 | 0.1508 | 0.14 | 0.14 | 0.14 | 0.13 | 0.12 | 0.1519 | 0.139 | 0.15 | 0.16 | 0.15 | 0.1378 | 0.145 | 0.15295 | 0.13673 | 0.137 | 0.14876
| 1 | 5 | 0.16 | 0.15 | 0.16 | 0.1528 | 0.15016 | 0.16 | 0.15 | 0.16018 | 0.15 | 0.156 | 0.1481 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11 | 0.148 | 0.145 | 0.15 | 0.16 | 0.14 | 0.1404 | 0.151 | 0.15131 | 0.14235 | 0.137 | 40
| 1 | 6 | 0.16 | 0.14 | 0.16 | 0.1466 | 0.14934 | 0.16 | 0.15 | 0.16281 | 0.15 | 0.158 | 0.1496 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11 | 0.1512 | 0.147 | 0.16 | 0.16 | 0.14 | 0.1354 | 0.151 | 0.14991 | 0.13413 | 0.141
| Mean |  | 0.16 | 0.14667 | 0.16 | 0.14903 | 0.14954 | 0.15667 | 0.15333 | 0.16398 | 0.15 | 0.15533 | 0.1495 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11167 | 0.1497 | 0.147 | 0.155 | 0.15833 | 0.145 | 0.13828 | 0.151 | 0.15425 | 0.1378 | 0.13883
| Median |  | 0.16 | 0.15 | 0.16 | 0.14895 | 0.14955 | 0.16 | 0.15 | 0.16365 | 0.15 | 0.1555 | 0.14965 | 0.14 | 0.14 | 0.14 | 0.13 | 0.11 | 0.14935 | 0.1475 | 0.155 | 0.16 | 0.145 | 0.13835 | 0.151 | 0.15295 | 0.13825 | 0.139
| Std Dev. |  | 0 | 0.005164 | 0 | 0.0022151 | 0.00048831 | 0.005164 | 0.005164 | 0.0033758 | 0 | 0.0026583 | 0.0014283 | 0 | 0 | 0 | 0 | 0.0040825 | 0.0015697 | 0.0047749 | 0.0054772 | 0.0040825 | 0.0054772 | 0.0020044 | 0.0037947 | 0.0039964 | 0.0043406 | 0.0024014
| Rel.Std.Dev. |  | 0% | 3.521% | 0% | 1.486% | 0.3265% | 3.296% | 3.368% | 2.059% | 0% | 1.711% | 0.9554% | 0% | 0% | 0% | 0% | 3.656% | 1.049% | 3.248% | 3.534% | 2.578% | 3.777% | 1.449% | 2.513% | 2.591% | 3.15% | 1.73%
| PDM3 |  | 7.555% | -1.408% | 7.555% | 0.1831% | 0.5245% | 5.314% | 3.074% | 10.23% | 0.8329% | 4.418% | 0.4968% | -5.889% | -5.889% | -5.889% | -12.61% | -24.94% | 0.6312% | -1.184% | 4.194% | 6.435% | -2.528% | -7.043% | 1.505% | 3.693% | -7.368% | -6.674%
| Z-Score (Absolute) |  | 1.09 | 0.3 | 1.09 | 0.05 | 0 | 0.74 | 0.4 | 1.5 | 0.05 | 0.6 | 0 | 0.99 | 0.99 | 0.99 | 2.03 | 3.94 | 0.02 | 0.26 | 0.57 | 0.92 | 0.47 | 1.17 | 0.13 | 0.49 | 1.22 | 1.11
| Analytical results for Mn in OREAS 600c (Certified Value 0.039 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0372 | 0.0414 | 0.0401 | 0.0392 | 0.039631 | 0.043 | 0.0418 | 0.038546 | 0.041 | 0.0378 | 0.0419 | 0.0394 | 0.039 | 0.0375 | 0.0376 | 0.0405 | 0.0374 | 0.0338 | 0.0484 | 0.0363 | 0.0395 | < 0.1 | 0.039 | 0.039236 | 0.035632 | 0.034 | 1
| 1 | 2 | 0.0376 | 0.0406 | 0.0407 | 0.0393 | 0.039622 | 0.0434 | 0.0409 | 0.038395 | 0.0408 | 0.0386 | 0.041 | 0.0395 | 0.0388 | 0.038 | 0.0371 | 0.0412 | 0.0376 | 0.0364 | 0.0492 | 0.03715 | 0.0376 | < 0.1 | 0.039 | 0.040972 | 0.035995 | 0.0339 | 27
| 1 | 3 | 0.0378 | 0.0416 | 0.0395 | 0.039 | 0.040587 | 0.0442 | 0.0409 | 0.038072 | 0.0416 | 0.0391 | 0.0409 | 0.0399 | 0.039 | 0.0382 | 0.0376 | 0.0408 | 0.0373 | 0.0339 | 0.0489 | 0.03672 | 0.0399 | < 0.1 | 0.039 | 0.041077 | 0.036049 | 0.0325 | 16
| 1 | 4 | 0.0378 | 0.0414 | 0.0407 | 0.0384 | 0.040307 | 0.0448 | 0.0407 | 0.038421 | 0.0417 | 0.038 | 0.0418 | 0.0394 | 0.0389 | 0.0376 | 0.0381 | 0.0425 | 0.0381 | 0.0275 | 0.0487 | 0.03719 | 0.0387 | < 0.1 | 0.039 | 0.040398 | 0.035913 | 0.0327 | 0.039361
| 1 | 5 | 0.0379 | 0.042 | 0.0409 | 0.0402 | 0.040382 | 0.0435 | 0.0402 | 0.038887 | 0.0411 | 0.0377 | 0.0408 | 0.0399 | 0.0398 | 0.0377 | 0.038 | 0.0406 | 0.0372 | 0.0312 | 0.0496 | 0.03701 | 0.0396 | < 0.1 | 0.039 | 0.03954 | 0.035777 | 0.034 | 41
| 1 | 6 | 0.0373 | 0.0395 | 0.0402 | 0.0383 | 0.039885 | 0.0443 | 0.0402 | 0.038551 | 0.0408 | 0.0385 | 0.0412 | 0.0392 | 0.0383 | 0.0384 | 0.0374 | 0.0412 | 0.0378 | 0.0333 | 0.049 | 0.03682 | 0.0386 | < 0.1 | 0.039 | 0.040596 | 0.036537 | 0.0335
| Mean |  | 0.0376 | 0.041083 | 0.04035 | 0.039067 | 0.040069 | 0.043867 | 0.040783 | 0.038479 | 0.041167 | 0.038283 | 0.041267 | 0.03955 | 0.038967 | 0.0379 | 0.037633 | 0.041133 | 0.037567 | 0.032683 | 0.048967 | 0.036865 | 0.038983 |  | 0.039 | 0.040303 | 0.035984 | 0.033433
| Median |  | 0.0377 | 0.0414 | 0.04045 | 0.0391 | 0.040096 | 0.04385 | 0.0408 | 0.038483 | 0.04105 | 0.03825 | 0.0411 | 0.03945 | 0.03895 | 0.03785 | 0.0376 | 0.041 | 0.0375 | 0.03355 | 0.04895 | 0.036915 | 0.0391 |  | 0.039 | 0.040497 | 0.035954 | 0.0337
| Std Dev. |  | 0.00028983 | 0.00089981 | 0.00052058 | 0.00069186 | 0.00041193 | 0.00067429 | 0.00059133 | 0.00026541 | 0.00039328 | 0.00054191 | 0.00047188 | 0.0002881 | 0.00048442 | 0.00035777 | 0.00037238 | 0.00073121 | 0.00033862 | 0.0030328 | 0.00041312 | 0.00033171 | 0.00085186 |  | 0 | 0.00075666 | 0.00031046 | 0.00067429
| Rel.Std.Dev. |  | 0.7708% | 2.19% | 1.29% | 1.771% | 1.028% | 1.537% | 1.45% | 0.6898% | 0.9553% | 1.416% | 1.143% | 0.7284% | 1.243% | 0.944% | 0.9895% | 1.778% | 0.9014% | 9.279% | 0.8437% | 0.8998% | 2.185% |  | 0% | 1.877% | 0.8628% | 2.017%
| PDM3 |  | -4.474% | 4.376% | 2.513% | -0.7474% | 1.799% | 11.45% | 3.614% | -2.241% | 4.588% | -2.738% | 4.842% | 0.4806% | -1.001% | -3.711% | -4.389% | 4.503% | -4.558% | -16.96% | 24.4% | -6.341% | -0.9591% |  | -0.9168% | 2.394% | -8.58% | -15.06%
| Z-Score (Absolute) |  | 0.67 | 0.96 | 0.62 | 0.02 | 0.49 | 2.26 | 0.82 | 0.26 | 1 | 0.35 | 1.05 | 0.25 | 0.03 | 0.53 | 0.65 | 0.99 | 0.68 | 2.97 | 4.65 | 1.01 | 0.02 | 5.13 | 0.14 | 0.6 | 1.42 | 2.62
| Analytical results for Mo in OREAS 600c (Certified Value 3.97 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.82 | 4.22 | 3.9 | 3.9 | 4.1703 | 4 | 3.97 | 3.84 | 4.1 | 4.1 | 3.88 | 4.08 | 4.09 | 4.29 | 3.84 | 3.9 | 4 | 4.5 | 3.6 | 4.48 | 5.2004 | 6 | 3.8828 | < 5 | 3.7 |  | 1
| 1 | 2 | 3.44 | 3.68 | 4 | 4 | 4.2203 | 4 | 4.03 | 3.89 | 4 | 4 | 3.91 | 4.03 | 3.9 | 3.72 | 3.82 | 4 | 4.1 | 4.3 | 3.57 | 4.34 | 5.0119 | 8 | 4.005 | < 5 | 3.9 |  | 28
| 1 | 3 | 4.07 | 4.21 | 3.9 | 3.9 | 4.1446 | 5 | 3.92 | 3.85 | 4.2 | 4.2 | 4.31 | 4.13 | 4.07 | 3.79 | 3.93 | 3.9 | 4.1 | 3.9 | 3.48 | 4.3 | 4.5282 | 9 | 3.9908 | < 5 | 3.7 |  | 16
| 1 | 4 | 3.59 | 3.94 | 4 | 3.8 | 4.1882 | 4 | 3.84 | 3.9 | 4.1 | 4 | 4.08 | 4.1 | 4.16 | 4.17 | 4.75 | 4.1 | 4 | 4.7 | 3.65 | 4.65 | 5.0989 | 6 | 3.9508 | < 5 | 3.9 |  | 3.9733
| 1 | 5 | 3.39 | 3.82 | 4.2 | 4.1 | 4.2402 | 4 | 3.86 | 3.71 | 4.1 | 4 | 3.77 | 4.34 | 3.92 | 3.81 | 3.92 | 3.9 | 4 | 3.7 | 3.61 | 4.52 | 4.7451 | 8 | 3.873 | < 5 | 4 |  | 42
| 1 | 6 | 4 | 3.57 | 3.8 | 3.9 | 4.0358 | 4 | 4.06 | 3.79 | 4.1 | 4 | 3.79 | 4.16 | 4.39 | 3.83 | 3.87 | 3.9 | 4.1 | 5 | 3.49 | 4.56 | 5.2787 | 7 | 3.9574 | < 5 | 3.9
| Mean |  | 3.7183 | 3.9067 | 3.9667 | 3.9333 | 4.1666 | 4.1667 | 3.9467 | 3.83 | 4.1 | 4.05 | 3.9567 | 4.14 | 4.0883 | 3.935 | 4.0217 | 3.95 | 4.05 | 4.35 | 3.5667 | 4.475 | 4.9772 | 7.3333 | 3.9433 |  | 3.85
| Median |  | 3.705 | 3.88 | 3.95 | 3.9 | 4.1793 | 4 | 3.945 | 3.845 | 4.1 | 4 | 3.895 | 4.115 | 4.08 | 3.82 | 3.895 | 3.9 | 4.05 | 4.4 | 3.585 | 4.5 | 5.0554 | 7.5 | 3.9541 |  | 3.9
| Std Dev. |  | 0.28813 | 0.26964 | 0.13663 | 0.10328 | 0.072647 | 0.40825 | 0.089368 | 0.070711 | 0.063246 | 0.083666 | 0.2053 | 0.10752 | 0.17904 | 0.23458 | 0.35941 | 0.083666 | 0.054772 | 0.48888 | 0.068313 | 0.13323 | 0.28713 | 1.2111 | 0.054641 |  | 0.12247
| Rel.Std.Dev. |  | 7.749% | 6.902% | 3.444% | 2.626% | 1.744% | 9.798% | 2.264% | 1.846% | 1.543% | 2.066% | 5.189% | 2.597% | 4.379% | 5.961% | 8.937% | 2.118% | 1.352% | 11.24% | 1.915% | 2.977% | 5.769% | 16.51% | 1.386% |  | 3.181%
| PDM3 |  | -6.417% | -1.677% | -0.1667% | -1.006% | 4.865% | 4.867% | -0.67% | -3.606% | 3.189% | 1.931% | -0.4184% | 4.196% | 2.895% | -0.9637% | 1.218% | -0.5861% | 1.931% | 9.481% | -10.23% | 12.63% | 25.27% | 84.57% | -0.7549% |  | -3.103%
| Z-Score (Absolute) |  | 1.51 | 0.32 | 0.06 | 0.15 | 1.33 | NA | 0.06 | 0.8 | 0.91 | 0.59 | 0 | 1.16 | 0.83 | 0.14 | 0.41 | 0.04 | 0.59 | 2.49 | 2.47 | 3.28 | 6.45 | NA | 0.08 | 9.21 | 0.67
Results from laboratories 2.07 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 600c (Certified Value 2.79 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.72 | 4.524 | 2.945 | 2.7641 | 2.8021 | 3 | 2.77 | 2.8081 | 2.85 | 2.88 | 2.8603 | 2.85 | 2.78 | 2.61 | 2.67 | 2.86 | 3.0282 | 2.873 | 2.38 | 2.68 | 2.72 | 2.8587 | 2.76 | 2.5356 | 2.4796 | 2.893 | 1
| 1 | 2 | 2.71 | 4.273 | 3 | 2.7554 | 2.7111 | > 3 | 2.7 | 2.8459 | 2.93 | 2.78 | 2.8153 | 2.81 | 2.74 | 2.61 | 2.64 | 2.86 | 3.012 | 2.865 | 2.3 | 2.71 | 2.65 | 2.8097 | 2.841 | 2.5948 | 2.7525 | 2.895 | 0
| 1 | 3 | 2.72 | 4.505 | 2.92 | 2.728 | 2.7386 | > 3 | 2.7 | 2.8126 | 2.94 | 2.76 | 2.7752 | 2.84 | 2.75 | 2.62 | 2.65 | 2.86 | 3.0089 | 2.842 | 2.39 | 2.72 | 2.67 | 2.8586 | 2.745 | 2.5784 | 2.6333 | 2.931 | 16
| 1 | 4 | 2.71 | 3.98 | 2.97 | 2.7211 | 2.7686 | > 3 | 2.68 | 2.8651 | 2.89 | 2.75 | 2.8541 | 2.83 | 2.72 | 2.61 | 2.67 | 2.91 | 3.0753 | 2.771 | 2.39 | 2.69 | 2.72 | 2.8319 | 2.73 | 2.5386 | 2.6064 | 2.773 | 2.7891
| 1 | 5 | 2.76 | 4.269 | 2.987 | 2.8145 | 2.7617 | > 3 | 2.65 | 2.8851 | 2.88 | 2.82 | 2.791 | 2.81 | 2.8 | 2.66 | 2.66 | 2.84 | 2.9723 | 2.849 | 2.34 | 2.72 | 2.67 | 2.8567 | 2.737 | 2.5808 | 2.7456 | 2.852 | 43
| 1 | 6 | 2.71 | 4.307 | 2.968 | 2.6982 | 2.7304 | > 3 | 2.68 | 2.8644 | 2.94 | 2.88 | 2.8349 | 2.78 | 2.7 | 2.67 | 2.66 | 2.89 | 3.056 | 2.876 | 2.41 | 2.7 | 2.84 | 2.7749 | 2.767 | 2.5507 | 2.5734 | 2.859
| Mean |  | 2.7217 | 4.3097 | 2.965 | 2.7469 | 2.7521 | 3 | 2.6967 | 2.8469 | 2.905 | 2.8117 | 2.8218 | 2.82 | 2.7483 | 2.63 | 2.6583 | 2.87 | 3.0255 | 2.846 | 2.3683 | 2.7033 | 2.7117 | 2.8317 | 2.7633 | 2.5632 | 2.6318 | 2.8672
| Median |  | 2.715 | 4.29 | 2.969 | 2.7417 | 2.7501 | 3 | 2.69 | 2.8551 | 2.91 | 2.8 | 2.8251 | 2.82 | 2.745 | 2.615 | 2.66 | 2.86 | 3.0201 | 2.857 | 2.385 | 2.705 | 2.695 | 2.8443 | 2.7525 | 2.5646 | 2.6198 | 2.876
| Std Dev. |  | 0.019408 | 0.19788 | 0.028872 | 0.040823 | 0.032256 |  | 0.040332 | 0.030901 | 0.037283 | 0.058109 | 0.034233 | 0.025298 | 0.037103 | 0.027568 | 0.01169 | 0.025298 | 0.036621 | 0.03909 | 0.040702 | 0.01633 | 0.069113 | 0.034023 | 0.04049 | 0.024756 | 0.10462 | 0.054186
| Rel.Std.Dev. |  | 0.7131% | 4.592% | 0.9738% | 1.486% | 1.172% |  | 1.496% | 1.085% | 1.283% | 2.067% | 1.213% | 0.8971% | 1.35% | 1.048% | 0.4398% | 0.8815% | 1.21% | 1.373% | 1.719% | 0.6041% | 2.549% | 1.201% | 1.465% | 0.9658% | 3.975% | 1.89%
| PDM3 |  | -2.416% | 54.52% | 6.308% | -1.512% | -1.325% | 7.563% | -3.313% | 2.073% | 4.157% | 0.8106% | 1.174% | 1.109% | -1.46% | -5.703% | -4.687% | 2.902% | 8.476% | 2.042% | -15.08% | -3.074% | -2.775% | 1.531% | -0.9224% | -8.099% | -5.639% | 2.8%
| Z-Score (Absolute) |  | 0.55 | 12.66 | 1.48 | 0.34 | 0.29 | 1.77 | 0.76 | 0.49 | 0.98 | 0.2 | 0.29 | 0.27 | 0.33 | 1.31 | 1.07 | 0.69 | 1.98 | 0.49 | 3.49 | 0.7 | 0.63 | 0.37 | 0.2 | 1.87 | 1.3 | 0.66
| Analytical results for Nb in OREAS 600c (Certified Value 19.3 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 18.6 | 20.29 | 18.2 | 19.28 | 19.8628 | 20.3 | 21.1 | 18.8 | 18.7 | 19.96 | 19.6 | 20.4 | 19.2 | 20.8 | 18.7 | 18.09 | 20.3 | 21.8 | 17.9 | 19.6462 | 10 | 18.5624 | 18.8 |  |  |  | 1
| 1 | 2 | 18.2 | 18.2 | 18.1 | 19.42 | 19.5807 | 20.6 | 21.5 | 19.3 | 18.7 | 19.35 | 20.3 | 20 | 19 | 19.8 | 20 | 18.38 | 20.6 | 21.3 | 16.8 | 18.705 | 11 | 19.2557 | 18.2 |  |  |  | 29
| 1 | 3 | 19.7 | 19.42 | 18.2 | 19.52 | 19.9711 | 20.4 | 20.6 | 19 | 18.6 | 20.06 | 19.8 | 19.2 | 18.8 | 19.6 | 20.9 | 14.64 | 20.4 | 23.5 | 18.6 | 18.1336 | 11 | 19.335 | 18.4 |  |  |  | 16
| 1 | 4 | 17.6 | 18.93 | 18.4 | 19.85 | 19.3347 | 21 | 19.7 | 18.7 | 18.7 | 19.9 | 20.2 | 19.2 | 19.2 | 20.1 | 21.6 | 18.83 | 19.8 | 21.8 | 17.7 | 17.9473 | 11 | 18.9425 | 18.1 |  |  |  | 19.2936
| 1 | 5 | 18.5 | 19.04 | 18.5 | 20.11 | 19.7006 | 20.3 | 20.1 | 18.2 | 19.1 | 19.53 | 19.6 | 19.6 | 19 | 20.2 | 20.2 | 17.04 | 20.2 | 22.1 | 18.2 | 17.5948 | 10 | 18.625 | 18.3 |  |  |  | 44
| 1 | 6 | 19.8 | 17.92 | 18.5 | 19.21 | 19.642 | 20.6 | 21 | 18.9 | 18.8 | 20.25 | 19.7 | 19 | 19.5 | 20 | 20.9 | 17.48 | 20.7 | 22.5 | 18.8 | 19.6942 | 10 | 18.9345 | 18.6
| Mean |  | 18.7333 | 18.9667 | 18.3167 | 19.565 | 19.682 | 20.5333 | 20.6667 | 18.8167 | 18.7667 | 19.8417 | 19.8667 | 19.5667 | 19.1167 | 20.0833 | 20.3833 | 17.41 | 20.3333 | 22.1667 | 18 | 18.6202 | 10.5 | 18.9425 | 18.4
| Median |  | 18.55 | 18.985 | 18.3 | 19.47 | 19.6713 | 20.5 | 20.8 | 18.85 | 18.7 | 19.93 | 19.75 | 19.4 | 19.1 | 20.05 | 20.55 | 17.785 | 20.35 | 21.95 | 18.05 | 18.4193 | 10.5 | 18.9385 | 18.35
| Std Dev. |  | 0.86178 | 0.85402 | 0.17224 | 0.34887 | 0.22318 | 0.26583 | 0.67132 | 0.3656 | 0.17512 | 0.33784 | 0.30768 | 0.54283 | 0.24014 | 0.41191 | 1.0028 | 1.499 | 0.32042 | 0.76333 | 0.71833 | 0.88931 | 0.54772 | 0.31543 | 0.26077
| Rel.Std.Dev. |  | 4.6% | 4.503% | 0.9403% | 1.783% | 1.134% | 1.295% | 3.248% | 1.943% | 0.9331% | 1.703% | 1.549% | 2.774% | 1.256% | 2.051% | 4.92% | 8.61% | 1.576% | 3.444% | 3.991% | 4.776% | 5.216% | 1.665% | 1.417%
| PDM3 |  | -2.904% | -1.694% | -5.063% | 1.407% | 2.013% | 6.426% | 7.117% | -2.472% | -2.731% | 2.841% | 2.97% | 1.415% | -0.917% | 4.093% | 5.648% | -9.763% | 5.389% | 14.89% | -6.705% | -3.49% | -45.58% | -1.82% | -4.632%
| Z-Score (Absolute) |  | 0.62 | 0.38 | 1.04 | 0.23 | 0.35 | 1.21 | 1.35 | 0.53 | 0.58 | 0.51 | 0.53 | 0.23 | 0.23 | 0.75 | 1.06 | 1.96 | 1.01 | 2.87 | 1.36 | 0.73 | NA | 0.4 | 0.96
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 600c (Certified Value 40.3 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  |  |  | 1
2
| 1 | 1 | 2.01: 42.7; 2.03: 39.61; 2.04: 37.251; 2.05: 40.6; 2.06: 44.602; 2.07: 41.39; 2.08: 35; 2.10: 42.1; 2.11: 38.569; 2.12: 44; 2.13: 35.2 |  |  |  | 1
| 1 | 2 | 2.01: 38.5; 2.03: 40.37; 2.04: 37.0395; 2.05: 40.3; 2.06: 44.795; 2.07: 41.7; 2.08: 37.3; 2.10: 40.7; 2.11: 37.0369; 2.12: 44; 2.13: 36.7 |  |  |  | 30
| 1 | 3 | 2.01: 42.4; 2.03: 40.36; 2.04: 38.6577; 2.05: 39.4; 2.06: 44.652; 2.07: 42.38; 2.08: 33.6; 2.10: 41.2; 2.11: 37.6575; 2.12: 43; 2.13: 34.1 |  |  |  | 16
| 1 | 4 | 2.01: 42.3; 2.03: 41.47; 2.04: 37.7512; 2.05: 41; 2.06: 44.598; 2.07: 41.13; 2.08: 24.6; 2.10: 42.1; 2.11: 37.6784; 2.12: 42; 2.13: 34.2 |  |  |  | 40.2995
| 1 | 5 | 2.01: 39.5; 2.03: 41.92; 2.04: 38.203; 2.05: 39.8; 2.06: 44.764; 2.07: 41.06; 2.08: 27.8; 2.10: 41.4; 2.11: 38.1702; 2.12: 43; 2.13: 35.2 |  |  |  | 45
| 1 | 6 | 2.01: 38.5; 2.03: 39.93; 2.04: 37.4101; 2.05: 40.6; 2.06: 44.857; 2.07: 42.94; 2.08: 31; 2.10: 40.9; 2.11: 36.719; 2.12: 42; 2.13: 35.9
| Mean |  | 2.01: 40.65; 2.03: 40.61; 2.04: 37.7187; 2.05: 40.2833; 2.06: 44.7113; 2.07: 41.7667; 2.08: 31.55; 2.10: 41.4; 2.11: 37.6385; 2.12: 43; 2.13: 35.2167
| Median |  | 2.01: 40.9; 2.03: 40.365; 2.04: 37.5806; 2.05: 40.45; 2.06: 44.708; 2.07: 41.545; 2.08: 32.3; 2.10: 41.3; 2.11: 37.6679; 2.12: 43; 2.13: 35.2
| Std Dev. |  | 2.01: 2.0276; 2.03: 0.89869; 2.04: 0.61523; 2.05: 0.58793; 2.06: 0.10891; 2.07: 0.74886; 2.08: 4.7294; 2.10: 0.5933; 2.11: 0.68672; 2.12: 0.89443; 2.13: 0.99482
| Rel.Std.Dev. |  | 2.01: 4.988%; 2.03: 2.213%; 2.04: 1.631%; 2.05: 1.459%; 2.06: 0.2436%; 2.07: 1.793%; 2.08: 14.99%; 2.10: 1.433%; 2.11: 1.825%; 2.12: 2.08%; 2.13: 2.825%
| PDM3 |  | 2.01: 0.8697%; 2.03: 0.7704%; 2.04: -6.404%; 2.05: -0.04017%; 2.06: 10.95%; 2.07: 3.641%; 2.08: -21.71%; 2.10: 2.731%; 2.11: -6.603%; 2.12: 6.701%; 2.13: -12.61%
| Z-Score (Absolute) |  | 2.01: 0.01; 2.03: 0; 2.04: 0.82; 2.05: 0.09; 2.06: 1.16; 2.07: 0.33; 2.08: 2.56; 2.10: 0.22; 2.11: 0.84; 2.12: 0.67; 2.13: 1.52
| Analytical results for Ni in OREAS 600c (Certified Value 5.83 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 7.7 | 9.3 | 5.8 | 5.6 | 6.0765 | 6 | 5.9 | 5.8 | 6 | 6 | 6.1 | 5.7 | 5.4 | 6.7 | 5.7 | 5.5 | 8 | 4.6 | 5.56 | 6.3 | 5.9471 | 5 | 5.4996 | < 5 | 5 |  | 1
| 1 | 2 | 7.8 | 6 | 6.2 | 5.7 | 5.9176 | 6 | 5.8 | 5.9 | 6 | 5.8 | 6.4 | 6 | 5.4 | 6.3 | 6 | 5.4 | 8 | 5.9 | 5.9 | 6.3 | 5.2091 | 5 | 5.5543 | < 5 | 5 |  | 16
| 1 | 3 | 7.7 | 6.6 | 5.8 | 5.9 | 6.11 | 6 | 5.9 | 5.7 | 6 | 6.3 | 6 | 5.8 | 5.4 | 5.9 | 5.8 | 4.9 | 7 | 4.7 | 5.74 | 6.5 | 5.4808 | 5 | 5.5595 | < 5 | 6 |  | 16
| 1 | 4 | 7.3 | 6.3 | 6 | 5.6 | 6.0589 | 6 | 6 | 5.8 | 6 | 5.9 | 6.3 | 5.6 | 5.5 | 6.3 | 6.9 | 5.2 | 6 | 6.3 | 5.86 | 6.6 | 4.6938 | 5 | 5.5595 | < 5 | 6 |  | 5.8329
| 1 | 5 | 7.4 | 6.3 | 6.2 | 6.1 | 6.1945 | 9 | 5.5 | 5.3 | 6 | 5.8 | 6.3 | 5.8 | 5.5 | 6 | 5.8 | 5.2 | 8 | 5.7 | 6.19 | 6.1 | 5.901 | 5 | 5.5739 | < 5 | 6 |  | 46
| 1 | 6 | 7.6 | 6.6 | 5.7 | 5.7 | 5.9519 | 6 | 5.4 | 5.6 | 6 | 6.1 | 6 | 5.6 | 5.8 | 6 | 6.9 | 5.2 | 8 | 5.2 | 5.47 | 5.9 | 4.4295 | 4 | 5.498 | < 5 | 6
| Mean |  | 7.5833 | 6.85 | 5.95 | 5.7667 | 6.0516 | 6.5 | 5.75 | 5.6833 | 6 | 5.9833 | 6.1833 | 5.75 | 5.5 | 6.2 | 6.1833 | 5.2333 | 7.5 | 5.4 | 5.7867 | 6.2833 | 5.2769 | 4.8333 | 5.5408 |  | 5.6667
| Median |  | 7.65 | 6.45 | 5.9 | 5.7 | 6.0677 | 6 | 5.85 | 5.75 | 6 | 5.95 | 6.2 | 5.75 | 5.45 | 6.15 | 5.9 | 5.2 | 8 | 5.45 | 5.8 | 6.3 | 5.3449 | 5 | 5.5569 |  | 6
| Std Dev. |  | 0.19408 | 1.2211 | 0.21679 | 0.19664 | 0.10235 | 1.2247 | 0.2429 | 0.2137 | 0 | 0.19408 | 0.17224 | 0.15166 | 0.15492 | 0.29665 | 0.56362 | 0.20656 | 0.83666 | 0.68118 | 0.25874 | 0.25626 | 0.62337 | 0.40825 | 0.033179 |  | 0.5164
| Rel.Std.Dev. |  | 2.559% | 17.83% | 3.644% | 3.41% | 1.691% | 18.84% | 4.224% | 3.76% | 0% | 3.244% | 2.786% | 2.638% | 2.817% | 4.785% | 9.115% | 3.947% | 11.16% | 12.61% | 4.471% | 4.078% | 11.81% | 8.447% | 0.5988% |  | 9.113%
| PDM3 |  | 30.01% | 17.44% | 2.008% | -1.135% | 3.749% | 11.44% | -1.421% | -2.564% | 2.865% | 2.579% | 6.008% | -1.421% | -5.707% | 6.294% | 6.008% | -10.28% | 28.58% | -7.422% | -0.7925% | 7.722% | -9.532% | -17.14% | -5.008% |  | -2.85%
| Z-Score (Absolute) |  | 3.13 | 1.85 | 0.28 | 0.03 | 0.46 | NA | 0.06 | 0.18 | NA | 0.34 | 0.69 | 0.06 | 0.5 | 0.72 | 0.69 | 0.96 | 2.99 | 0.67 | 0 | 0.87 | 0.89 | NA | 0.43 | 5.73 | NA
| Analytical results for P in OREAS 600c (Certified Value 0.037 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.0368 | 0.037 | 0.036 | 0.0362 | 0.037288 | 0.035 | 0.0389 | 0.0368 | 0.036 | 0.0371 | 0.036 | 0.036 | 0.037 | 0.035 | 0.039 | 0.0357 | 0.034 | 0.041 | 0.04 | 0.0364 | < 0.1 | 0.035 | 0.036245 | 0.03279 | 0.041 |  | 1
| 1 | 2 | 0.0368 | 0.037 | 0.037 | 0.0363 | 0.037551 | 0.035 | 0.0385 | 0.0381 | 0.036 | 0.0362 | 0.036 | 0.036 | 0.037 | 0.035 | 0.039 | 0.0358 | 0.036 | 0.034 | 0.04 | 0.0352 | < 0.1 | 0.039 | 0.038324 | 0.032966 | 0.038 |  | 0
| 1 | 3 | 0.0364 | 0.037 | 0.036 | 0.0363 | 0.037866 | 0.035 | 0.038 | 0.0379 | 0.036 | 0.0362 | 0.037 | 0.036 | 0.036 | 0.036 | 0.038 | 0.0355 | 0.036 | 0.039 | 0.04 | 0.0368 | < 0.1 | 0.039 | 0.0368 | 0.034004 | 0.039 |  | 16
| 1 | 4 | 0.0367 | 0.037 | 0.037 | 0.0362 | 0.037232 | 0.036 | 0.038 | 0.0372 | 0.036 | 0.0369 | 0.036 | 0.035 | 0.035 | 0.036 | 0.039 | 0.0369 | 0.036 | 0.039 | 0.04 | 0.0355 | < 0.1 | 0.035 | 0.036245 | 0.033385 | 0.037 |  | 0.036538
| 1 | 5 | 0.0366 | 0.037 | 0.037 | 0.0361 | 0.037902 | 0.034 | 0.0377 | 0.0373 | 0.037 | 0.0367 | 0.036 | 0.037 | 0.037 | 0.035 | 0.038 | 0.0355 | 0.037 | 0.033 | 0.04 | 0.0361 | < 0.1 | 0.035 | 0.036519 | 0.033015 | 0.035 |  | 47
| 1 | 6 | 0.0364 | 0.035 | 0.036 | 0.036 | 0.037161 | 0.036 | 0.0375 | 0.0376 | 0.036 | 0.0366 | 0.036 | 0.035 | 0.038 | 0.036 | 0.039 | 0.0362 | 0.036 | 0.033 | 0.04 | 0.0362 | < 0.1 | 0.035 | 0.038031 | 0.034246 | 0.038
| Mean |  | 0.036617 | 0.036667 | 0.0365 | 0.036183 | 0.0375 | 0.035167 | 0.0381 | 0.037483 | 0.036167 | 0.036617 | 0.036167 | 0.035833 | 0.036667 | 0.0355 | 0.038667 | 0.035933 | 0.035833 | 0.0365 | 0.04 | 0.036033 |  | 0.036333 | 0.037027 | 0.033401 | 0.038
| Median |  | 0.03665 | 0.037 | 0.0365 | 0.0362 | 0.03742 | 0.035 | 0.038 | 0.03745 | 0.036 | 0.03665 | 0.036 | 0.036 | 0.037 | 0.0355 | 0.039 | 0.03575 | 0.036 | 0.0365 | 0.04 | 0.03615 |  | 0.035 | 0.036659 | 0.0332 | 0.038
| Std Dev. |  | 0.00018348 | 0.0008165 | 0.00054772 | 0.0001169 | 0.0003256 | 0.00075277 | 0.00051769 | 0.00047924 | 0.00040825 | 0.0003656 | 0.00040825 | 0.00075277 | 0.0010328 | 0.00054772 | 0.0005164 | 0.00053914 | 0.00098319 | 0.0035637 | 0 | 0.00058878 |  | 0.0020656 | 0.00091904 | 0.00059824 | 0.002
| Rel.Std.Dev. |  | 0.5011% | 2.227% | 1.501% | 0.3231% | 0.8683% | 2.141% | 1.359% | 1.279% | 1.129% | 0.9985% | 1.129% | 2.101% | 2.817% | 1.543% | 1.336% | 1.5% | 2.744% | 9.764% | 0% | 1.634% |  | 5.685% | 2.482% | 1.791% | 5.263%
| PDM3 |  | 0.2158% | 0.3526% | -0.1035% | -0.9702% | 2.634% | -3.753% | 4.276% | 2.588% | -1.016% | 0.2158% | -1.016% | -1.928% | 0.3526% | -2.84% | 5.826% | -1.654% | -1.928% | -0.1035% | 9.476% | -1.381% |  | -0.5596% | 1.339% | -8.584% | 4.002%
| Z-Score (Absolute) |  | 0.14 | 0.2 | 0 | 0.38 | 1.19 | 1.59 | 1.9 | 1.17 | 0.4 | 0.14 | 0.4 | 0.79 | 0.2 | 1.19 | 2.58 | 0.67 | 0.79 | 0 | 4.16 | 0.56 | 16.06 | 0.16 | 0.63 | 3.69 | 1.78
| Analytical results for Pb in OREAS 600c (Certified Value 1129 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 1230 | 1129.84 | 1160.6 | 1113.2 | 1037.771 | 1100 | 1090 | 1170.51 | 1132 | 1130 | 1143 | 1125 | 1145 | 1080 | 1110 | 1130 | 1184 | 1071 | 1610 | 1171 | 1088.9 | 934.3683 | 1160 | 1142.3219 | 967.944 | 1038 | 1
| 1 | 2 | 1225 | 1124.4 | 1153.5 | 1137.4 | 1073.1068 | 1100 | 1080 | 1155.22 | 1136 | 1160 | 1125 | 1120 | 1135 | 1085 | 1095 | 1150 | 1180 | 1041 | 1650 | 1189 | 1078.6 | 925.8082 | 1170 | 1141.6367 | 1029.966 | 1010 | 18
| 1 | 3 | 1244 | 1140.84 | 1144.9 | 1155.8 | 1089.9697 | 1110 | 1080 | 1143.34 | 1146 | 1150 | 1120 | 1125 | 1130 | 1085 | 1115 | 1140 | 1186 | 1057 | 1580 | 1152 | 1078.1 | 937.852 | 1160 | 1107.2234 | 1026.849 | 994 | 16
| 1 | 4 | 1227 | 1126.4 | 1174.8 | 1147.5 | 1075.4893 | 1100 | 1070 | 1165.34 | 1139 | 1160 | 1135 | 1110 | 1130 | 1080 | 1125 | 1185 | 1194 | 1005 | 1620 | 1198 | 1079.4 | 914.6212 | 1140 | 1116.5029 | 999.902 | 1036 | 1128.9188
| 1 | 5 | 1247 | 1145.96 | 1178.8 | 1164.2 | 1082.3852 | 1120 | 1100 | 1165.34 | 1157 | 1160 | 1135 | 1110 | 1155 | 1080 | 1110 | 1120 | 1180 | 957 | 1580 | 1164 | 1083.8 | 933.2554 | 1130 | 1102.6905 | 1053.31 | 1004 | 48
| 1 | 6 | 1227 | 1088 | 1133.3 | 1114.4 | 1062.0964 | 1100 | 1110 | 1122.11 | 1147 | 1150 | 1134 | 1110 | 1110 | 1120 | 1110 | 1145 | 1190 | 1020 | 1610 | 1188 | 1090.2 | 894.4445 | 1160 | 1091.8696 | 1008.522 | 1025
| Mean |  | 1233.3333 | 1125.9067 | 1157.65 | 1138.75 | 1070.1364 | 1105 | 1088.3333 | 1153.6433 | 1142.8333 | 1151.6667 | 1132 | 1116.6667 | 1134.1667 | 1088.3333 | 1110.8333 | 1145 | 1185.6667 | 1025.1667 | 1608.3333 | 1177 | 1083.1667 | 923.3916 | 1153.3333 | 1117.0408 | 1014.4155 | 1017.8333
| Median |  | 1228.5 | 1128.12 | 1157.05 | 1142.45 | 1074.298 | 1100 | 1085 | 1160.28 | 1142.5 | 1155 | 1134.5 | 1115 | 1132.5 | 1082.5 | 1110 | 1142.5 | 1185 | 1030.5 | 1610 | 1179.5 | 1081.6 | 929.5318 | 1160 | 1111.8632 | 1017.6855 | 1017.5
| Std Dev. |  | 9.6056 | 20.3963 | 17.452 | 21.2721 | 18.4012 | 8.3666 | 14.7196 | 18.2269 | 9.0203 | 11.6905 | 8.1976 | 7.5277 | 15.3025 | 15.7056 | 9.704 | 22.3607 | 5.5737 | 41.097 | 26.3944 | 17.5271 | 5.3564 | 16.4169 | 15.0555 | 20.882 | 29.374 | 17.9379
| Rel.Std.Dev. |  | 0.7788% | 1.812% | 1.508% | 1.868% | 1.72% | 0.7572% | 1.352% | 1.58% | 0.7893% | 1.015% | 0.7242% | 0.6741% | 1.349% | 1.443% | 0.8736% | 1.953% | 0.4701% | 4.009% | 1.641% | 1.489% | 0.4945% | 1.778% | 1.305% | 1.869% | 2.896% | 1.762%
| PDM3 |  | 9.249% | -0.2668% | 2.545% | 0.8709% | -5.207% | -2.119% | -3.595% | 2.19% | 1.233% | 2.015% | 0.2729% | -1.085% | 0.4649% | -3.595% | -1.602% | 1.424% | 5.027% | -9.19% | 42.47% | 4.259% | -4.053% | -18.21% | 2.163% | -1.052% | -10.14% | -9.84%
| Z-Score (Absolute) |  | 2.64 | 0.08 | 0.72 | 0.25 | 1.49 | 0.61 | 1.03 | 0.62 | 0.35 | 0.57 | 0.08 | 0.31 | 0.13 | 1.03 | 0.46 | 0.41 | 1.43 | 2.62 | 12.11 | 1.21 | 1.16 | 5.19 | 0.62 | 0.3 | 2.89 | 2.81
| Analytical results for Pr in OREAS 600c (Certified Value 10.3 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS | 1
2
| 1 | 1 | 2.01: 10.8; 2.03: 10.23; 2.04: 10.1683; 2.05: 10.5; 2.06: 11.79; 2.07: 10.95; 2.08: 10.49; 2.09: 10.67; 2.10: 9.955; 2.11: 9; 2.12: 9.71 | 1
| 1 | 2 | 2.01: 10; 2.03: 10.35; 2.04: 10.6722; 2.05: 10.3; 2.06: 11.572; 2.07: 11.24; 2.08: 11.34; 2.09: 10.59; 2.10: 9.8025; 2.11: 9; 2.12: 9.85 | 33
| 1 | 3 | 2.01: 10.9; 2.03: 10.36; 2.04: 10.5926; 2.05: 10.2; 2.06: 11.711; 2.07: 11.16; 2.08: 10.05; 2.09: 10.76; 2.10: 9.9314; 2.11: 9; 2.12: 9.24 | 16
| 1 | 4 | 2.01: 10.8; 2.03: 10.59; 2.04: 10.8695; 2.05: 10.6; 2.06: 11.859; 2.07: 10.75; 2.08: 7.18; 2.09: 10.65; 2.10: 10.1791; 2.11: 8; 2.12: 9.38 | 10.3215
| 1 | 5 | 2.01: 10.2; 2.03: 10.56; 2.04: 10.5001; 2.05: 10.3; 2.06: 11.962; 2.07: 11.26; 2.08: 8.31; 2.09: 10.65; 2.10: 10.4258; 2.11: 9; 2.12: 9.4 | 49
| 1 | 6 | 2.01: 9.7; 2.03: 10.24; 2.04: 10.4116; 2.05: 10.5; 2.06: 11.747; 2.07: 11.41; 2.08: 9.31; 2.09: 10.63; 2.10: 9.913; 2.11: 9; 2.12: 9.58
| Mean |  | 2.01: 10.4; 2.03: 10.3883; 2.04: 10.5357; 2.05: 10.4; 2.06: 11.7735; 2.07: 11.1283; 2.08: 9.4467; 2.09: 10.6583; 2.10: 10.0345; 2.11: 8.8333; 2.12: 9.5267
| Median |  | 2.01: 10.5; 2.03: 10.355; 2.04: 10.5463; 2.05: 10.4; 2.06: 11.7685; 2.07: 11.2; 2.08: 9.68; 2.09: 10.65; 2.10: 9.9432; 2.11: 9; 2.12: 9.49
| Std Dev. |  | 2.01: 0.502; 2.03: 0.15459; 2.04: 0.23873; 2.05: 0.15492; 2.06: 0.13291; 2.07: 0.23878; 2.08: 1.5156; 2.09: 0.056716; 2.10: 0.22784; 2.11: 0.40825; 2.12: 0.22836
| Rel.Std.Dev. |  | 2.01: 4.827%; 2.03: 1.488%; 2.04: 2.266%; 2.05: 1.49%; 2.06: 1.129%; 2.07: 2.146%; 2.08: 16.04%; 2.09: 0.5321%; 2.10: 2.271%; 2.11: 4.622%; 2.12: 2.397%
| PDM3 |  | 2.01: 0.7605%; 2.03: 0.6474%; 2.04: 2.075%; 2.05: 0.7605%; 2.06: 14.07%; 2.07: 7.817%; 2.08: -8.476%; 2.09: 3.263%; 2.10: -2.781%; 2.11: -14.42%; 2.12: -7.701%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 0.03; 2.04: 0.29; 2.05: 0; 2.06: 2.97; 2.07: 1.57; 2.08: 2.06; 2.09: 0.56; 2.10: 0.79; 2.11: NA; 2.12: 1.89
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 600c (Certified Value 148 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 176 | 118.3 | 143.1 | 141.25 | 140.6372 | 149 | 164 | 165.864 | 140 | 142 | 155.64 | 146.5 | 141.5 | 141 | 158.5 | 150 | 153.6 | 158 | 131 | 149.8 | 147.8948 | 138 |  |  |  |  | 1
| 1 | 2 | 166 | 135.4 | 146 | 142.41 | 143.6562 | 151 | 162 | 165.8196 | 151 | 140 | 153.58 | 147 | 140 | 142 | 151.5 | 151.5 | 153.9 | 166 | 119 | 143.8 | 151.8511 | 141 |  |  |  |  | 34
| 1 | 3 | 172 | 119.5 | 143.5 | 144.23 | 144.4322 | 152 | 140 | 165.6764 | 148 | 139 | 158.66 | 145.5 | 137 | 143.5 | 152.5 | 156 | 155.1 | 157 | 126 | 146.2 | 151.6996 | 136 |  |  |  |  | 16
| 1 | 4 | 163 | 132.6 | 147.1 | 145.13 | 144.7665 | 153 | 157 | 165.7188 | 144 | 139 | 149.25 | 147 | 137.5 | 142.5 | 153.5 | 160 | 157 | 140 | 127 | 149.3 | 147.8948 | 138 |  |  |  |  | 147.7862
| 1 | 5 | 160 | 143.9 | 148.3 | 145.79 | 148.2632 | 150 | 159 | 165.2404 | 146 | 141 | 154.04 | 145 | 139.5 | 137.5 | 155 | 152.5 | 153.9 | 148 | 129 | 143.4 | 148.1998 | 138 |  |  |  |  | 50
| 1 | 6 | 167 | 128.1 | 147.9 | 141.07 | 140.351 | 154 | 162 | 165.0812 | 143 | 137 | 158.26 | 147 | 134.5 | 145 | 152.5 | 156.5 | 155 | 156 | 124 | 146.1 | 148.8187 | 142
| Mean |  | 167.3333 | 129.6333 | 145.9833 | 143.3133 | 143.6844 | 151.5 | 157.3333 | 165.5667 | 145.3333 | 139.6667 | 154.905 | 146.3333 | 138.3333 | 141.9167 | 153.9167 | 154.4167 | 154.75 | 154.1667 | 126 | 146.4333 | 149.3932 | 138.8333
| Median |  | 166.5 | 130.35 | 146.55 | 143.32 | 144.0442 | 151.5 | 160.5 | 165.6976 | 145 | 139.5 | 154.84 | 146.75 | 138.5 | 142.25 | 153 | 154.25 | 154.45 | 156.5 | 126.5 | 146.15 | 148.5093 | 138
| Std Dev. |  | 5.8538 | 9.7885 | 2.2257 | 2.0189 | 2.9358 | 1.8708 | 8.8468 | 0.32548 | 3.8816 | 1.7512 | 3.475 | 0.8756 | 2.5033 | 2.5577 | 2.538 | 3.7339 | 1.2661 | 8.9981 | 4.1952 | 2.6778 | 1.8765 | 2.2286
| Rel.Std.Dev. |  | 3.498% | 7.551% | 1.525% | 1.409% | 2.043% | 1.235% | 5.623% | 0.1966% | 2.671% | 1.254% | 2.243% | 0.5984% | 1.81% | 1.802% | 1.649% | 2.418% | 0.8182% | 5.837% | 3.33% | 1.829% | 1.256% | 1.605%
| PDM3 |  | 13.23% | -12.28% | -1.22% | -3.027% | -2.776% | 2.513% | 6.46% | 12.03% | -1.66% | -5.494% | 4.817% | -0.9831% | -6.396% | -3.972% | 4.148% | 4.487% | 4.712% | 4.317% | -14.74% | -0.9154% | 1.087% | -6.058%
| Z-Score (Absolute) |  | 1.87 | 1.5 | 0.04 | 0.27 | 0.24 | 0.46 | 0.98 | 1.72 | 0.09 | 0.6 | 0.76 | 0 | 0.72 | 0.4 | 0.67 | 0.72 | 0.75 | 0.7 | 1.82 | 0 | 0.27 | 0.68
| Analytical results for Re in OREAS 600c (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [8 empty] | 3
3
| 1 | 1 | 0.01 | 0.017 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 1
| 1 | 2 | < 0.002 | < 0.002 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 35
| 1 | 3 | 0.011 | 0.044 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 16
| 1 | 4 | < 0.002 | 0.016 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | < 0.002
| 1 | 5 | 0.028 | 0.016 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1 | [8 empty] | 51
| 1 | 6 | < 0.002 | 0.024 | < 0.005 | < 0.01 | < 0.001 | < 0.002 | < 0.05 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.003 | < 0.1 | < 0.002 | < 0.002 | < 0.05 | < 0.1
| Mean |  | 2.01: 0.016333; 2.03: 0.0234; 2.13: 0.002; 2.16: 0.002
| Median |  | 2.01: 0.011; 2.03: 0.017; 2.13: 0.002; 2.16: 0.002
| Std Dev. |  | 2.01: 0.010116; 2.03: 0.011992; 2.13: 0
| Rel.Std.Dev. |  | 2.01: 61.93%; 2.03: 51.25%; 2.13: 0%
PDM3
| Z-Score (Absolute) |  | 6.44 | 16.33 | 0.9 | 3.15 | 0.9 | 0.45 | 21.13 | 0.45 | 0.45 | 0.45 | 0.45 | 0 | 0 | 43.61 | 0.3 | 0.45 | 21.13 | 43.61
| Analytical results for S in OREAS 600c (Certified Value 0.238 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.3 | 0.23 | 0.2 | 0.235 | 0.25397 | 0.22 | 0.26 | 0.14288 | 0.24 | 0.232 | 0.2442 | 0.24 | 0.25 | 0.24 | 0.23 | 0.24 | 0.2219 | 0.24 | 0.25 | 0.13 | 0.24 | 0.2682 | 0.224 | 0.23456 | 0.216 | 0.246 | 1
| 1 | 2 | 0.3 | 0.22 | 0.2 | 0.2344 | 0.25155 | 0.23 | 0.26 | 0.1495 | 0.24 | 0.236 | 0.2393 | 0.24 | 0.25 | 0.23 | 0.23 | 0.25 | 0.2288 | 0.249 | 0.24 | 0.13 | 0.24 | 0.2773 | 0.23 | 0.23456 | 0.22 | 0.24 | 11
| 1 | 3 | 0.3 | 0.23 | 0.2 | 0.2369 | 0.25865 | 0.22 | 0.26 | 0.14976 | 0.25 | 0.227 | 0.2391 | 0.24 | 0.25 | 0.23 | 0.23 | 0.24 | 0.2393 | 0.236 | 0.25 | 0.12 | 0.25 | 0.2777 | 0.228 | 0.23449 | 0.22 | 0.232 | 16
| 1 | 4 | 0.29 | 0.22 | 0.2 | 0.233 | 0.25831 | 0.23 | 0.26 | 0.14289 | 0.25 | 0.232 | 0.2414 | 0.24 | 0.25 | 0.23 | 0.23 | 0.26 | 0.2328 | 0.226 | 0.25 | 0.13 | 0.25 | 0.2936 | 0.223 | 0.23197 | 0.214 | 0.233 | 0.23761
| 1 | 5 | 0.3 | 0.22 | 0.2 | 0.2387 | 0.25977 | 0.22 | 0.26 | 0.14549 | 0.24 | 0.228 | 0.2424 | 0.24 | 0.25 | 0.23 | 0.23 | 0.24 | 0.2254 | 0.229 | 0.25 | 0.11 | 0.25 | 0.2825 | 0.219 | 0.23027 | 0.213 | 0.242 | 52
| 1 | 6 | 0.3 | 0.21 | 0.2 | 0.2332 | 0.25387 | 0.23 | 0.26 | 0.14542 | 0.25 | 0.236 | 0.2408 | 0.24 | 0.24 | 0.24 | 0.23 | 0.25 | 0.2303 | 0.242 | 0.25 | 0.12 | 0.24 | 0.2664 | 0.225 | 0.23045 | 0.219 | 0.242
| Mean |  | 0.29833 | 0.22167 | 0.2 | 0.2352 | 0.25602 | 0.225 | 0.26 | 0.14599 | 0.245 | 0.23183 | 0.2412 | 0.24 | 0.24833 | 0.23333 | 0.23 | 0.24667 | 0.22975 | 0.237 | 0.24833 | 0.12333 | 0.245 | 0.27762 | 0.22483 | 0.23272 | 0.217 | 0.23917
| Median |  | 0.3 | 0.22 | 0.2 | 0.2347 | 0.25614 | 0.225 | 0.26 | 0.14546 | 0.245 | 0.232 | 0.2411 | 0.24 | 0.25 | 0.23 | 0.23 | 0.245 | 0.22955 | 0.238 | 0.25 | 0.125 | 0.245 | 0.2775 | 0.2245 | 0.23323 | 0.2175 | 0.241
| Std Dev. |  | 0.0040825 | 0.0075277 | 3.0405e-17 | 0.0022208 | 0.0033188 | 0.0054772 | 0 | 0.0030479 | 0.0054772 | 0.0038166 | 0.0019318 | 0 | 0.0040825 | 0.005164 | 0 | 0.008165 | 0.0060375 | 0.0085323 | 0.0040825 | 0.008165 | 0.0054772 | 0.0099389 | 0.0038687 | 0.0020799 | 0.0030984 | 0.0055287
| Rel.Std.Dev. |  | 1.368% | 3.396% | 1.52e-14% | 0.9442% | 1.296% | 2.434% | 0% | 2.088% | 2.236% | 1.646% | 0.8009% | 0% | 1.644% | 2.213% | 0% | 3.31% | 2.628% | 3.6% | 1.644% | 6.62% | 2.236% | 3.58% | 1.721% | 0.8937% | 1.428% | 2.312%
| PDM3 |  | 25.56% | -6.708% | -15.83% | -1.012% | 7.749% | -5.305% | 9.425% | -38.56% | 3.112% | -2.429% | 1.513% | 1.008% | 4.515% | -1.798% | -3.201% | 3.813% | -3.306% | -0.2549% | 4.515% | -48.09% | 3.112% | 16.84% | -5.375% | -2.057% | -8.672% | 0.6569%
| Z-Score (Absolute) |  | 4.28 | 1.07 | NA | 0.13 | 1.33 | 0.84 | 1.6 | 6.35 | 0.56 | 0.36 | 0.29 | 0.21 | 0.79 | 0.26 | 0.49 | 0.67 | 0.51 | 0 | 0.79 | 7.93 | 0.56 | 2.83 | 0.85 | 0.3 | 1.4 | 0.15
Results from laboratory 2.04 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 600c (Certified Value 64 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
1
| 1 | 1 | 67 | 64.95 | 61.9 | 65.55 | 66.1329 | 65.7 | 70.9 | 63.68 | 65.3 | 68 | 67.7 | 65.6 | 68.5 | 67.1 | 64.9 | 60.25 | 63.93 | 61.1 | 62.77 | 57.48 | 57.459 | 63.9491 | 59.53 | 72.2 |  |  | 1
| 1 | 2 | 68 | 58.31 | 63.9 | 66.05 | 68.2623 | 65.8 | 70.3 | 67.21 | 65 | 67 | 69.9 | 63.6 | 70.5 | 64.4 | 65.4 | 59.4 | 63.92 | 63.5 | 60.47 | 55.35 | 59.308 | 66.134 | 59.73 | 70.12 |  |  | 36
| 1 | 3 | 66 | 61.68 | 61.3 | 67.09 | 69.1694 | 64.8 | 67.7 | 65.91 | 64.5 | 66 | 68.3 | 62.5 | 69.1 | 63.9 | 65.6 | 59.77 | 63.32 | 60.5 | 60.41 | 57.17 | 59.0099 | 66.5475 | 59.6 | 70.42 |  |  | 16
| 1 | 4 | 67 | 60.17 | 63.5 | 66.31 | 66.0768 | 68 | 67.6 | 64.83 | 65.8 | 67 | 68.7 | 64.4 | 69.9 | 65.4 | 67.6 | 61.05 | 65.51 | 61.2 | 61.74 | 57.5 | 56.3366 | 65.1915 | 60.71 | 70.55 |  |  | 64.4672
| 1 | 5 | 67 | 60.16 | 63.7 | 67.12 | 69.1742 | 65.9 | 68.5 | 64.15 | 65.6 | 68 | 66.5 | 66.3 | 67.3 | 65.2 | 65.2 | 59.58 | 62.71 | 216 | 59.65 | 56.88 | 58.5234 | 63.3618 | 57.6 | 70.88 |  |  | 53
| 1 | 6 | 68 | 56.43 | 62.3 | 65.17 | 65.9447 | 67.1 | 69 | 65.13 | 65.7 | 67 | 67.9 | 62.2 | 70.8 | 64 | 66 | 60.15 | 64.9 | 61.7 | 61.7 | 57.62 | 57.6994 | 65.9653 | 60 | 72.65
| Mean |  | 67.1667 | 60.2833 | 62.7667 | 66.215 | 67.46 | 66.2167 | 69 | 65.1517 | 65.3167 | 67.1667 | 68.1667 | 64.1 | 69.35 | 65 | 65.7833 | 60.0333 | 64.0483 | 87.3333 | 61.1233 | 57 | 58.056 | 65.1915 | 59.5283 | 71.1367
| Median |  | 67 | 60.165 | 62.9 | 66.18 | 67.1976 | 65.85 | 68.75 | 64.98 | 65.45 | 67 | 68.1 | 64 | 69.5 | 64.8 | 65.5 | 59.96 | 63.925 | 61.45 | 61.085 | 57.325 | 58.1114 | 65.5784 | 59.665 | 70.715
| Std Dev. |  | 0.75277 | 2.9154 | 1.0783 | 0.79437 | 1.5795 | 1.1409 | 1.3565 | 1.2716 | 0.49565 | 0.75277 | 1.129 | 1.6492 | 1.3202 | 1.1983 | 0.96419 | 0.595 | 1.0212 | 63.0418 | 1.1429 | 0.85213 | 1.1078 | 1.282 | 1.0376 | 1.0371
| Rel.Std.Dev. |  | 1.121% | 4.836% | 1.718% | 1.2% | 2.341% | 1.723% | 1.966% | 1.952% | 0.7588% | 1.121% | 1.656% | 2.573% | 1.904% | 1.844% | 1.466% | 0.9911% | 1.594% | 72.19% | 1.87% | 1.495% | 1.908% | 1.966% | 1.743% | 1.458%
| PDM3 |  | 4.187% | -6.49% | -2.638% | 2.711% | 4.642% | 2.714% | 7.031% | 1.062% | 1.318% | 4.187% | 5.738% | -0.5696% | 7.574% | 0.8264% | 2.042% | -6.878% | -0.6498% | 35.47% | -5.187% | -11.58% | -9.945% | 1.124% | -7.661% | 10.35%
| Z-Score (Absolute) |  | 0.55 | 1.43 | 0.71 | 0.28 | 0.63 | 0.28 | 1.08 | 0.03 | 0.02 | 0.55 | 0.84 | 0.33 | 1.18 | 0.07 | 0.15 | 1.5 | 0.35 | 6.34 | 1.19 | 2.37 | 2.07 | 0.02 | 1.65 | 1.69
| Analytical results for Sc in OREAS 600c (Certified Value 4.35 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 6.4 | 4.7 | 4 | 4.3 | 4.0208 | 4 | 5.2 | 5.9692 | 4.77 | 4 | 4.2 | 4.1 | 4 | 3.8 | 4.6 | 4.3 | 4.17 | 4 | 4.4 | < 5 | 4.96 | 3 | 4.1668 | 4.7 |  |  | 1
| 1 | 2 | 6.2 | 4.1 | 4 | 4.3 | 3.9788 | 4 | 5.2 | 5.9431 | 5.2 | 4 | 4.2 | 4.2 | 4.1 | 3.9 | 4.3 | 4.3 | 4.17 | 3.7 | 4.5 | < 5 | 4.8987 | 3 | 4.3981 | 5.1 |  |  | 37
| 1 | 3 | 6.6 | 4.6 | 4 | 4.2 | 4.1204 | 4 | 5 | 5.9147 | 5 | 4 | 4.1 | 4.1 | 4 | 3.9 | 4.3 | 4.4 | 4.14 | 3.2 | 4.4 | < 5 | 5.1037 | 3 | 4.4159 | 4.9 |  |  | 16
| 1 | 4 | 5.5 | 4.5 | 4 | 4.3 | 4.0771 | 4 | 5.1 | 5.96 | 5.03 | 4 | 4.2 | 4.1 | 4 | 3.9 | 4.3 | 4.5 | 4.23 | 35 | 4.4 | < 5 | 5.1441 | 3 | 4.4078 | 4.6 |  |  | 4.3476
| 1 | 5 | 6 | 4.2 | 4 | 4.2 | 4.0684 | 4 | 5.1 | 5.9393 | 4.87 | 4 | 4.1 | 4 | 4.1 | 3.8 | 4.3 | 4.3 | 4.13 | 2.9 | 4.3 | < 5 | 5.2659 | 3 | 4.3759 | 5 |  |  | 54
| 1 | 6 | 6 | 4 | 4 | 3.8 | 3.9904 | 4 | 5.1 | 5.9575 | 4.93 | 4 | 4.2 | 4.1 | 3.9 | 4 | 4.3 | 4.4 | 4.21 | 3 | 4 | < 5 | 5.0023 | 3 | 4.1956 | 5.1
| Mean |  | 6.1167 | 4.35 | 4 | 4.1833 | 4.0426 | 4 | 5.1167 | 5.9473 | 4.9667 | 4 | 4.1667 | 4.1 | 4.0167 | 3.8833 | 4.35 | 4.3667 | 4.175 | 8.6333 | 4.3333 |  | 5.0625 | 3 | 4.3267 | 4.9
| Median |  | 6.1 | 4.35 | 4 | 4.25 | 4.0446 | 4 | 5.1 | 5.9503 | 4.965 | 4 | 4.2 | 4.1 | 4 | 3.9 | 4.3 | 4.35 | 4.17 | 3.45 | 4.4 |  | 5.053 | 3 | 4.387 | 4.95
| Std Dev. |  | 0.38166 | 0.2881 | 0 | 0.19408 | 0.055113 | 0 | 0.075277 | 0.019441 | 0.1476 | 0 | 0.05164 | 0.063246 | 0.075277 | 0.075277 | 0.12247 | 0.08165 | 0.038859 | 12.9239 | 0.17512 |  | 0.13468 | 0 | 0.11387 | 0.20976
| Rel.Std.Dev. |  | 6.24% | 6.623% | 0% | 4.639% | 1.363% | 0% | 1.471% | 0.3269% | 2.972% | 0% | 1.239% | 1.543% | 1.874% | 1.938% | 2.816% | 1.87% | 0.9307% | 149.7% | 4.041% |  | 2.66% | 0% | 2.632% | 4.281%
| PDM3 |  | 40.69% | 0.05622% | -7.994% | -3.777% | -7.013% | -7.994% | 17.69% | 36.8% | 14.24% | -7.994% | -4.161% | -5.694% | -7.611% | -10.68% | 0.05622% | 0.4396% | -3.969% | 98.58% | -0.3271% |  | 16.44% | -31% | -0.4802% | 12.71%
| Z-Score (Absolute) |  | 3.84 | 0.02 | NA | 0.34 | 0.65 | NA | 1.67 | 3.47 | 1.35 | NA | 0.38 | 0.52 | 0.7 | 0.99 | 0.02 | 0.05 | 0.36 | 9.28 | 0.02 | 3.98 | 1.56 | NA | 0.03 | 1.21
Results from laboratories 2.04, 2.07, 2.11 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 600c (Certified Value 1.17 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.27 | 2.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | [6 empty] | 2
2
| 1 | 1 | < 1 | 1.3 | < 1 | < 1.5 | 1.2 | < 0.5 | 1.6137 | < 1 | 2 | 2 | 1 | < 1 | < 1 | 0.23 | < 1 | 2 | 2 | < 2 | < 10 | 2 | [6 empty] | 1
| 1 | 2 | < 1 | 0.5 | < 1 | < 1.5 | 1.2 | < 0.5 | 1.1757 | < 1 | < 1 | 1 | < 1 | < 1 | < 1 | 0.21 | < 1 | 3 | 2 | < 2 | < 10 | 2 | [6 empty] | 38
| 1 | 3 | < 1 | 0.6 | < 1 | < 1.5 | 1.3 | < 0.5 | 1.0287 | < 1 | 1 | 1 | 1 | < 1 | < 1 | 0.22 | < 1 | 2 | 2 | < 2 | < 10 | 2 | [6 empty] | 16
| 1 | 4 | < 1 | 0.7 | < 1 | < 1.5 | 1.2 | < 0.5 | 1.4523 | < 1 | 1 | 1 | < 1 | < 1 | 1 | 0.23 | < 1 | 2 | 2 | < 2 | < 10 | 2 | [6 empty] | 1.1691
| 1 | 5 | < 1 | 1.1 | < 1 | < 1.5 | 1.2 | < 0.5 | 1.4062 | < 1 | 1 | 2 | 1 | 1 | < 1 | 0.22 | < 1 | 2 | 2 | < 2 | < 10 | 1 | [6 empty] | 55
| 1 | 6 | < 1 | 0.6 | < 1 | < 1.5 | 1.2 | < 0.5 | 1.1526 | < 1 | 1 | 1 | < 1 | < 1 | < 1 | 0.23 | < 1 | 3 | 2 | < 2 | < 10 | 1
| Mean |  | 2.03: 0.8; 2.06: 1.2167; 2.08: 1.3049; 2.10: 1.2; 2.11: 1.3333; 2.12: 1; 2.13: 1; 2.14: 1; 2.15: 0.22333; 2.17: 2.3333; 2.18: 2; 2.22: 1.6667
| Median |  | 2.03: 0.65; 2.06: 1.2; 2.08: 1.291; 2.10: 1; 2.11: 1; 2.12: 1; 2.13: 1; 2.14: 1; 2.15: 0.225; 2.17: 2; 2.18: 2; 2.22: 2
| Std Dev. |  | 2.03: 0.32249; 2.06: 0.040825; 2.08: 0.22069; 2.10: 0.44721; 2.11: 0.5164; 2.12: 0; 2.15: 0.008165; 2.17: 0.5164; 2.18: 0; 2.22: 0.5164
| Rel.Std.Dev. |  | 2.03: 40.31%; 2.06: 3.355%; 2.08: 16.91%; 2.10: 37.27%; 2.11: 38.73%; 2.12: 0%; 2.15: 3.656%; 2.17: 22.13%; 2.18: 0%; 2.22: 30.98%
| PDM3 |  | 2.03: -31.57%; 2.06: 4.072%; 2.08: 11.62%; 2.10: 2.647%; 2.11: 14.05%; 2.12: -14.46%; 2.13: -14.46%; 2.14: -14.46%; 2.15: -80.9%; 2.17: 99.59%; 2.18: 71.08%; 2.22: 42.56%
| Z-Score (Absolute) |  | 0.64 | 0.06 | 0.64 | 0.06 | 1.02 | 1.21 | 1.23 | 0.64 | 0.71 | 1.29 | 0.06 | 0.44 | 0.44 | 1.28 | 0.64 | 3.6 | 2.83 | 0.52 | 9.77 | 2.06
| Analytical results for Sm in OREAS 600c (Certified Value 7.67 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.22; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*OES/MS; 2.12: 4A*MS |  |  | 2
3
| 1 | 1 | 2.01: 8.6; 2.03: 7.31; 2.04: 7.1828; 2.05: 7.8; 2.06: 8.1199; 2.07: 7.82; 2.08: 6.6; 2.09: 7.73; 2.10: 7.2354; 2.11: 6; 2.12: 7.4 |  |  | 1
| 1 | 2 | 2.01: 7.9; 2.03: 7.42; 2.04: 7.0649; 2.05: 7.8; 2.06: 8.2574; 2.07: 7.79; 2.08: 7.1; 2.09: 7.68; 2.10: 7.0295; 2.11: 6; 2.12: 7.6 |  |  | 22
| 1 | 3 | 2.01: 8.7; 2.03: 7.32; 2.04: 7.3368; 2.05: 7.7; 2.06: 8.3791; 2.07: 7.78; 2.08: 6.5; 2.09: 7.68; 2.10: 7.0714; 2.11: 6; 2.12: 7.1 |  |  | 16
| 1 | 4 | 2.01: 8.6; 2.03: 7.64; 2.04: 7.2086; 2.05: 8.1; 2.06: 8.4389; 2.07: 7.6; 2.08: 4.9; 2.09: 7.77; 2.10: 7.1738; 2.11: 6; 2.12: 7.2 |  |  | 7.6657
| 1 | 5 | 2.01: 8.2; 2.03: 7.52; 2.04: 7.4741; 2.05: 8.1; 2.06: 8.4376; 2.07: 7.67; 2.08: 5.5; 2.09: 7.68; 2.10: 7.2607; 2.11: 6; 2.12: 7.2 |  |  | 56
| 1 | 6 | 2.01: 7.8; 2.03: 7.28; 2.04: 7.0722; 2.05: 7.6; 2.06: 8.1492; 2.07: 7.88; 2.08: 6; 2.09: 7.7; 2.10: 6.9846; 2.11: 6; 2.12: 7.4
| Mean |  | 2.01: 8.3; 2.03: 7.415; 2.04: 7.2232; 2.05: 7.85; 2.06: 8.297; 2.07: 7.7567; 2.08: 6.1; 2.09: 7.7067; 2.10: 7.1259; 2.11: 6; 2.12: 7.3167
| Median |  | 2.01: 8.4; 2.03: 7.37; 2.04: 7.1957; 2.05: 7.8; 2.06: 8.3183; 2.07: 7.785; 2.08: 6.25; 2.09: 7.69; 2.10: 7.1226; 2.11: 6; 2.12: 7.3
| Std Dev. |  | 2.01: 0.38987; 2.03: 0.14139; 2.04: 0.15848; 2.05: 0.20736; 2.06: 0.14242; 2.07: 0.10289; 2.08: 0.8025; 2.09: 0.036697; 2.10: 0.11374; 2.11: 0; 2.12: 0.18348
| Rel.Std.Dev. |  | 2.01: 4.697%; 2.03: 1.907%; 2.04: 2.194%; 2.05: 2.642%; 2.06: 1.717%; 2.07: 1.326%; 2.08: 13.16%; 2.09: 0.4762%; 2.10: 1.596%; 2.11: 0%; 2.12: 2.508%
| PDM3 |  | 2.01: 8.275%; 2.03: -3.27%; 2.04: -5.772%; 2.05: 2.404%; 2.06: 8.236%; 2.07: 1.187%; 2.08: -20.42%; 2.09: 0.5347%; 2.10: -7.042%; 2.11: -21.73%; 2.12: -4.553%
| Z-Score (Absolute) |  | 2.01: 1.59; 2.03: 0.31; 2.04: 0.73; 2.05: 0.62; 2.06: 1.58; 2.07: 0.42; 2.08: 3.14; 2.09: 0.31; 2.10: 0.94; 2.11: NA; 2.12: 0.53
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 600c (Certified Value 3.87 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 3.8 | 4.3 | 3.6 | 3.8 | 3.5904 | 4 | 4.2 | 3.8 | 3.8 | 4.1 | 3.8 | 3.9 | 3.8 | 3.9 | 3.7 | 3.86 | 3.8 | 3.3 | < 5 | 3.7 | 5 | 4.0333 | < 10 | 4.3 |  |  | 1
| 1 | 2 | 3.5 | 3.6 | 3.7 | 3.7 | 3.8367 | 4 | 4.2 | 3.9 | 3.9 | 4 | 4 | 4 | 3.8 | 3.7 | 3.9 | 3.8 | 4 | 3.1 | < 5 | 3.5 | 6 | 4.073 | < 10 | 4.5 |  |  | 13
| 1 | 3 | 3.5 | 4.2 | 3.5 | 3.9 | 3.9802 | 4 | 4 | 4.07 | 3.9 | 4.2 | 3.9 | 3.9 | 3.9 | 3.7 | 3.8 | 3.86 | 3.9 | 3.2 | < 5 | 3.7 | 7 | 4.0461 | < 10 | 4.1 |  |  | 16
| 1 | 4 | 3.2 | 4.1 | 3.7 | 3.8 | 3.7369 | 4 | 4 | 3.87 | 4 | 4 | 4 | 4 | 3.9 | 3.8 | 4 | 3.96 | 3.9 | 3.3 | < 5 | 3.7 | 4 | 4.0507 | < 10 | 3.6 |  |  | 3.8692
| 1 | 5 | 3.5 | 3.9 | 3.6 | 3.9 | 3.8986 | 4 | 4 | 3.8 | 3.8 | 4.1 | 3.8 | 4.1 | 3.8 | 3.8 | 3.8 | 3.86 | 3.8 | 3.4 | < 5 | 3.7 | 8 | 4.1536 | < 10 | 4.2 |  |  | 57
| 1 | 6 | 4 | 3.6 | 3.7 | 3.9 | 3.4639 | 4 | 4.2 | 3.77 | 3.9 | 4 | 3.9 | 3.9 | 4 | 3.8 | 3.8 | 3.85 | 4 | 3.1 | < 5 | 3.9 | 6 | 4.1451 | < 10 | 4.7
| Mean |  | 3.5833 | 3.95 | 3.6333 | 3.8333 | 3.7511 | 4 | 4.1 | 3.8683 | 3.8833 | 4.0667 | 3.9 | 3.9667 | 3.8667 | 3.7833 | 3.8333 | 3.865 | 3.9 | 3.2333 |  | 3.7 | 6 | 4.0837 |  | 4.2333
| Median |  | 3.5 | 4 | 3.65 | 3.85 | 3.7868 | 4 | 4.1 | 3.835 | 3.9 | 4.05 | 3.9 | 3.95 | 3.85 | 3.8 | 3.8 | 3.86 | 3.9 | 3.25 |  | 3.7 | 6 | 4.0619 |  | 4.25
| Std Dev. |  | 0.27869 | 0.30166 | 0.08165 | 0.08165 | 0.19497 | 0 | 0.10954 | 0.11017 | 0.075277 | 0.08165 | 0.089443 | 0.08165 | 0.08165 | 0.075277 | 0.10328 | 0.052058 | 0.089443 | 0.12111 |  | 0.12649 | 1.4142 | 0.052559 |  | 0.37771
| Rel.Std.Dev. |  | 7.777% | 7.637% | 2.247% | 2.13% | 5.198% | 0% | 2.672% | 2.848% | 1.938% | 2.008% | 2.293% | 2.058% | 2.112% | 1.99% | 2.694% | 1.347% | 2.293% | 3.746% |  | 3.419% | 23.57% | 1.287% |  | 8.922%
| PDM3 |  | -7.387% | 2.09% | -6.095% | -0.9258% | -3.051% | 3.382% | 5.966% | -0.02119% | 0.3665% | 5.105% | 0.7972% | 2.52% | -0.06427% | -2.218% | -0.9258% | -0.1073% | 0.7972% | -16.43% |  | -4.372% | 55.07% | 5.544% |  | 9.412%
| Z-Score (Absolute) |  | 1.78 | 0.52 | 1.47 | 0.21 | 0.73 | NA | 1.45 | 0.01 | 0.1 | 1.25 | 0.2 | 0.62 | 0.01 | 0.53 | 0.21 | 0.02 | 0.2 | 3.97 | 8.56 | 1.05 | NA | 1.35 | 7.09 | 2.29
Results from laboratories 2.07 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 600c (Certified Value 187 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 198 | 185 | 180 | 183.09 | 182.0418 | 189 | 202 | 195.034 | 179 | 190 | 197 | 185 | 188.5 | 181 | 180.5 | 180.5 | 208 | 180 | 186 | 175.9 | 190 | 163.4776 | 187 | 179.7078 | 157.11 | 174 | 1
| 1 | 2 | 198 | 183 | 181 | 185.28 | 183.5576 | 189 | 205 | 195.973 | 189 | 189 | 191 | 184.5 | 186.5 | 181 | 176.5 | 184 | 208 | 189 | 181 | 179.6 | 184 | 161.7954 | 192 | 196.6821 | 158.32 | 178 | 14
| 1 | 3 | 201 | 188 | 181 | 187.21 | 186.8932 | 191 | 198 | 195.537 | 185 | 189 | 192 | 186.5 | 185.5 | 181.5 | 176.5 | 183 | 207 | 177 | 189 | 178.9 | 187.8 | 163.0894 | 186 | 180.6676 | 158.61 | 172 | 16
| 1 | 4 | 199 | 184 | 184 | 188.41 | 183.9068 | 195 | 197 | 195.073 | 185 | 190 | 196 | 182.5 | 186 | 178.5 | 179.5 | 186.5 | 210 | 162 | 189 | 178.8 | 192.5 | 162.4736 | 185 | 190.9771 | 158.36 | 172 | 186.533
| 1 | 5 | 202 | 187 | 185 | 190.14 | 186.2451 | 190 | 197 | 195.083 | 186 | 190 | 191 | 183 | 189 | 178.5 | 180.5 | 182.5 | 206 | 163 | 184 | 179.2 | 185.6 | 164.1768 | 186 | 188.9343 | 158.78 | 174 | 58
| 1 | 6 | 197 | 175 | 185 | 183.73 | 181.6127 | 196 | 202 | 195.594 | 184 | 190 | 193 | 183.5 | 181.5 | 185 | 178.5 | 188 | 210 | 178 | 190 | 179 | 186.6 | 159.2982 | 188 | 193.8726 | 161.06 | 175
| Mean |  | 199.1667 | 183.6667 | 182.6667 | 186.31 | 184.0429 | 191.6667 | 200.1667 | 195.3823 | 184.6667 | 189.6667 | 193.3333 | 184.1667 | 186.1667 | 180.9167 | 178.6667 | 184.0833 | 208.1667 | 174.8333 | 186.5 | 178.5667 | 187.75 | 162.3852 | 187.3333 | 188.4736 | 158.7067 | 174.1667
| Median |  | 198.5 | 184.5 | 182.5 | 186.245 | 183.7322 | 190.5 | 200 | 195.31 | 185 | 190 | 192.5 | 184 | 186.25 | 181 | 179 | 183.5 | 208 | 177.5 | 187.5 | 178.95 | 187.2 | 162.7815 | 186.5 | 189.9557 | 158.485 | 174
| Std Dev. |  | 1.9408 | 4.6332 | 2.2509 | 2.7559 | 2.1512 | 3.0768 | 3.3116 | 0.38059 | 3.266 | 0.5164 | 2.582 | 1.472 | 2.6771 | 2.3962 | 1.8348 | 2.7462 | 1.6021 | 10.4579 | 3.5071 | 1.3367 | 3.0866 | 1.7196 | 2.5033 | 6.9395 | 1.294 | 2.2286
| Rel.Std.Dev. |  | 0.9745% | 2.523% | 1.232% | 1.479% | 1.169% | 1.605% | 1.654% | 0.1948% | 1.769% | 0.2723% | 1.336% | 0.7993% | 1.438% | 1.324% | 1.027% | 1.492% | 0.7696% | 5.982% | 1.881% | 0.7486% | 1.644% | 1.059% | 1.336% | 3.682% | 0.8153% | 1.28%
| PDM3 |  | 6.773% | -1.537% | -2.073% | -0.1195% | -1.335% | 2.752% | 7.309% | 4.744% | -1.001% | 1.68% | 3.646% | -1.269% | -0.1964% | -3.011% | -4.217% | -1.313% | 11.6% | -6.272% | -0.01769% | -4.271% | 0.6524% | -12.95% | 0.4291% | 1.04% | -14.92% | -6.63%
| Z-Score (Absolute) |  | 2.12 | 0.27 | 0.42 | 0.14 | 0.21 | 0.96 | 2.27 | 1.54 | 0.12 | 0.66 | 1.22 | 0.19 | 0.12 | 0.69 | 1.04 | 0.21 | 3.51 | 1.63 | 0.17 | 1.06 | 0.36 | 3.55 | 0.3 | 0.47 | 4.12 | 1.73
| Analytical results for Ta in OREAS 600c (Certified Value 1.44 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 1.01 | 1.8 | 1.3 | 1.44 | 1.4029 | 1.6 | 1.37 | 1.55 | 1.38 | 1.47 | 1.45 | 1.44 | 1.45 | 1.48 | 1.41 | 2.21 | 1.91 | 1.42 | 0.88 | < 1 | 1.4058 | 1.23 |  |  |  |  | 1
| 1 | 2 | 0.98 | 1.5 | 1.3 | 1.45 | 1.4022 | 1.6 | 1.41 | 1.62 | 1.38 | 1.47 | 1.48 | 1.43 | 1.43 | 1.37 | 1.45 | 1.88 | 1.89 | 1.4 | 0.86 | < 1 | 1.4617 | 1.26 |  |  |  |  | 25
| 1 | 3 | 1.11 | 1.7 | 1.3 | 1.45 | 1.4212 | 1.6 | 1.48 | 1.47 | 1.37 | 1.46 | 1.46 | 1.43 | 1.44 | 1.37 | 1.44 | 1.77 | 1.8 | 1.57 | 0.9 | < 1 | 1.448 | 1.24 |  |  |  |  | 16
| 1 | 4 | 1.09 | 1.6 | 1.3 | 1.48 | 1.4054 | 1.7 | 1.26 | 1.53 | 1.38 | 1.5 | 1.49 | 1.42 | 1.4 | 1.41 | 1.49 | 1.93 | 2.26 | 1.43 | 0.87 | < 1 | 1.429 | 1.24 |  |  |  |  | 1.4376
| 1 | 5 | 0.98 | 1.6 | 1.3 | 1.49 | 1.4284 | 1.6 | 1.32 | 1.5 | 1.37 | 1.48 | 1.44 | 1.48 | 1.41 | 1.42 | 1.41 | 1.88 | 1.8 | 1.46 | 0.9 | < 1 | 1.4019 | 1.23 |  |  |  |  | 59
| 1 | 6 | 0.94 | 1.5 | 1.3 | 1.44 | 1.4234 | 1.6 | 1.42 | 1.52 | 1.39 | 1.51 | 1.47 | 1.39 | 1.43 | 1.38 | 1.47 | 1.88 | 1.84 | 1.51 | 0.92 | < 1 | 1.4221 | 1.21
| Mean |  | 1.0183 | 1.6167 | 1.3 | 1.4583 | 1.4139 | 1.6167 | 1.3767 | 1.5317 | 1.3783 | 1.4817 | 1.465 | 1.4317 | 1.4267 | 1.405 | 1.445 | 1.925 | 1.9167 | 1.465 | 0.88833 |  | 1.4281 | 1.235
| Median |  | 0.995 | 1.6 | 1.3 | 1.45 | 1.4133 | 1.6 | 1.39 | 1.525 | 1.38 | 1.475 | 1.465 | 1.43 | 1.43 | 1.395 | 1.445 | 1.88 | 1.865 | 1.445 | 0.89 |  | 1.4255 | 1.235
| Std Dev. |  | 0.067355 | 0.1169 | 0 | 0.02137 | 0.01169 | 0.040825 | 0.078145 | 0.051153 | 0.0075277 | 0.019408 | 0.018708 | 0.029269 | 0.018619 | 0.042308 | 0.032094 | 0.14923 | 0.1742 | 0.064109 | 0.022286 |  | 0.023438 | 0.016432
| Rel.Std.Dev. |  | 6.614% | 7.231% | 0% | 1.465% | 0.8268% | 2.525% | 5.676% | 3.34% | 0.5461% | 1.31% | 1.277% | 2.044% | 1.305% | 3.011% | 2.221% | 7.752% | 9.089% | 4.376% | 2.509% |  | 1.641% | 1.331%
| PDM3 |  | -29.16% | 12.46% | -9.569% | 1.444% | -1.647% | 12.46% | -4.236% | 6.546% | -4.12% | 3.068% | 1.908% | -0.4105% | -0.7583% | -2.265% | 0.517% | 33.91% | 33.33% | 1.908% | -38.21% |  | -0.6588% | -14.09%
| Z-Score (Absolute) |  | 5.29 | 2.4 | 1.67 | 0.37 | 0.21 | 2.4 | 0.68 | 1.31 | 0.66 | 0.67 | 0.45 | 0.02 | 0.04 | 0.32 | 0.19 | 6.36 | 6.25 | 0.45 | 6.96 | 11.94 | 0.02 | 2.5
| Analytical results for Tb in OREAS 600c (Certified Value 0.77 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.10; 2.08: 2.12; 2.10: 2.19; 2.11: 2.22; 2.12: 2.23; 2.13: 2.25; 2.14: 2.28 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*MS; 2.13: 4A*OES/MS; 2.14: 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 0.9; 2.03: 0.71; 2.04: 0.79809; 2.05: 0.9; 2.06: 1.0812; 2.07: 0.73; 2.08: 0.77; 2.10: 0.84; 2.11: 0.88; 2.12: 0.8142; 2.13: < 1; 2.14: 0.69 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.8; 2.03: 0.73; 2.04: 0.79501; 2.05: 0.9; 2.06: 1.0229; 2.07: 0.77; 2.08: 0.81; 2.10: 0.89; 2.11: 0.84; 2.12: 0.7747; 2.13: < 1; 2.14: 0.69 |  |  |  |  | 26
| 1 | 3 | 2.01: 0.8; 2.03: 0.73; 2.04: 0.84197; 2.05: 0.9; 2.06: 1.1228; 2.07: 0.74; 2.08: 0.8; 2.10: 0.82; 2.11: 0.85; 2.12: 0.8196; 2.13: < 1; 2.14: 0.67 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.9; 2.03: 0.74; 2.04: 0.82489; 2.05: 0.9; 2.06: 1.0809; 2.07: 0.72; 2.08: 0.77; 2.10: 0.65; 2.11: 0.84; 2.12: 0.8087; 2.13: < 1; 2.14: 0.68 |  |  |  |  | 0.77134
| 1 | 5 | 2.01: 0.8; 2.03: 0.75; 2.04: 0.84514; 2.05: 0.9; 2.06: 1.0631; 2.07: 0.73; 2.08: 0.8; 2.10: 0.69; 2.11: 0.83; 2.12: 0.8092; 2.13: < 1; 2.14: 0.68 |  |  |  |  | 60
| 1 | 6 | 2.01: 0.8; 2.03: 0.73; 2.04: 0.80502; 2.05: 0.9; 2.06: 1.0836; 2.07: 0.73; 2.08: 0.82; 2.10: 0.74; 2.11: 0.83; 2.12: 0.7899; 2.13: < 1; 2.14: 0.69
| Mean |  | 2.01: 0.83333; 2.03: 0.73167; 2.04: 0.81836; 2.05: 0.9; 2.06: 1.0757; 2.07: 0.73667; 2.08: 0.795; 2.10: 0.77167; 2.11: 0.845; 2.12: 0.80272; 2.14: 0.68333
| Median |  | 2.01: 0.8; 2.03: 0.73; 2.04: 0.81495; 2.05: 0.9; 2.06: 1.0811; 2.07: 0.73; 2.08: 0.8; 2.10: 0.78; 2.11: 0.84; 2.12: 0.80895; 2.14: 0.685
| Std Dev. |  | 2.01: 0.05164; 2.03: 0.013292; 2.04: 0.022144; 2.05: 0; 2.06: 0.032511; 2.07: 0.017512; 2.08: 0.020736; 2.10: 0.093256; 2.11: 0.018708; 2.12: 0.016994; 2.14: 0.008165
| Rel.Std.Dev. |  | 2.01: 6.197%; 2.03: 1.817%; 2.04: 2.706%; 2.05: 0%; 2.06: 3.022%; 2.07: 2.377%; 2.08: 2.608%; 2.10: 12.09%; 2.11: 2.214%; 2.12: 2.117%; 2.14: 1.195%
| PDM3 |  | 2.01: 8.037%; 2.03: -5.144%; 2.04: 6.095%; 2.05: 16.68%; 2.06: 39.46%; 2.07: -4.495%; 2.08: 3.067%; 2.10: 0.04205%; 2.11: 9.549%; 2.12: 4.068%; 2.14: -11.41%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 0.71; 2.04: 0.48; 2.05: NA; 2.06: 4.01; 2.07: 0.64; 2.08: 0.16; 2.10: 0.16; 2.11: 0.85; 2.12: 0.27; 2.13: 3.89; 2.14: 1.37
Results from laboratories 2.03 and 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 600c (Certified Value 0.19 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.27 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 0.59 | 0.15 | < 0.5 |  | < 0.2 | 0.2 | 0.23 | 0.2 |  | 0.11 | 0.13 | 0.19 | 0.16 | 0.16 | 0.21 | < 0.05 | 0.16 | 0.21 | 0.17339 | < 10 | 0.19 | [5 empty] | 1
| 1 | 2 | 0.67 | 0.15 | < 0.5 | 0.2 | < 0.2 | 0.2 | 0.24 | 0.2 |  | 0.16 | 0.18 | 0.17 | 0.15 | 0.16 | 0.18 | 0.58 | 0.31 | 0.22 | 0.22274 | < 10 | 0.17 | [5 empty] | 27
| 1 | 3 | 0.66 | 0.22 | < 0.5 |  | < 0.2 | 0.2 | 0.21 | 0.2 | 0.2 | 0.19 | 0.14 | 0.22 | 0.15 | 0.15 | 0.17 | 0.69 | 0.31 | 0.21 | 0.22233 | < 10 | 0.22 | [5 empty] | 16
| 1 | 4 | 0.66 | 0.2 | < 0.5 | 0.2 | < 0.2 | 0.2 | 0.19 | 0.2 | 0.2 | 0.18 | 0.19 | 0.2 | 0.14 | 0.19 | 0.21 | < 0.05 | 0.16 | 0.21 | 0.203 | < 10 | 0.19 | [5 empty] | 0.19047
| 1 | 5 | 0.63 | 0.15 | < 0.5 | 0.2 | < 0.2 | 0.1 | 0.21 | 0.2 |  | 0.19 | 0.12 | 0.18 | 0.16 | 0.16 | 0.18 | 0.22 | 0.32 | 0.22 | 0.28843 | < 10 | 0.18 | [5 empty] | 61
| 1 | 6 | 0.55 | 0.1 | < 0.5 |  | < 0.2 | 0.3 | 0.2 | 0.2 |  | 0.15 | 0.18 | 0.21 | 0.16 | 0.15 | 0.19 | < 0.05 | 0.31 | 0.24 | 0.2677 | < 10 | 0.17
| Mean |  | 0.62667 | 0.16167 |  | 0.2 |  | 0.2 | 0.21333 | 0.2 | 0.2 | 0.16333 | 0.15667 | 0.195 | 0.15333 | 0.16167 | 0.19 | 0.49667 | 0.26167 | 0.21833 | 0.2296 |  | 0.18667
| Median |  | 0.645 | 0.15 |  | 0.2 |  | 0.2 | 0.21 | 0.2 | 0.2 | 0.17 | 0.16 | 0.195 | 0.155 | 0.16 | 0.185 | 0.58 | 0.31 | 0.215 | 0.22254 |  | 0.185
| Std Dev. |  | 0.04761 | 0.042622 |  | 3.3993e-17 |  | 0.063246 | 0.018619 | 3.0405e-17 | 0 | 0.030768 | 0.030111 | 0.018708 | 0.008165 | 0.01472 | 0.016733 | 0.24583 | 0.078846 | 0.01169 | 0.042143 |  | 0.018619
| Rel.Std.Dev. |  | 7.597% | 26.36% |  | 1.7e-14% |  | 31.62% | 8.728% | 1.52e-14% | 0% | 18.84% | 19.22% | 9.594% | 5.325% | 9.105% | 8.807% | 49.5% | 30.13% | 5.354% | 18.36% |  | 9.974%
| PDM3 |  | 229% | -15.12% |  | 5.005% |  | 5.005% | 12.01% | 5.005% | 5.005% | -14.25% | -17.75% | 2.38% | -19.5% | -15.12% | -0.2449% | 160.8% | 37.38% | 14.63% | 20.55% |  | -1.995%
| Z-Score (Absolute) |  | 8.41 | 0.65 | 1.07 | NA | 1.85 | NA | 0.36 | NA | NA | 0.62 | 0.75 | 0 | 0.81 | 0.65 | 0.1 | 1.28 | 1.3 | 0.45 | 0.67 | 93.65 | 0.16
Results from laboratories 2.05, 2.07, 2.11 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 600c (Certified Value 15.1 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.34 | 16.4 | 14.9 | 14.79 | 14.2433 | 15.4 | 16.4 | 18.298 | 14.9 | 15.4 | 15.62 | 14.35 | 14.65 | 14.75 | 15.8 | 16.25 | 14.42 | 12.7 | < 20 | 12.67 | 15.4916 | 14 | 14.829 | 12.89 |  |  | 1
| 1 | 2 | 13.66 | 14.6 | 15.1 | 15.06 | 14.7418 | 15.3 | 16.1 | 18.849 | 15.53 | 15.1 | 15.32 | 14.85 | 14.35 | 14.55 | 15.25 | 16.75 | 14.43 | 13.88 | < 20 | 12.52 | 14.644 | 14 | 15.3377 | 12.83 |  |  | 28
| 1 | 3 | 14.75 | 16.5 | 14.9 | 15.01 | 14.9832 | 14.7 | 16 | 18.919 | 15.06 | 15.5 | 15.56 | 14.15 | 14.75 | 15.35 | 15.25 | 16.55 | 14.28 | 12.77 | < 20 | 12.61 | 15.309 | 14 | 15.3575 | 12.31 |  |  | 16
| 1 | 4 | 14.05 | 15.8 | 15.2 | 15.25 | 14.4957 | 15.1 | 15.9 | 18.951 | 14.57 | 15.6 | 14.99 | 14.75 | 14.45 | 14.9 | 16 | 17.15 | 14.51 | 10.71 | < 20 | 12.82 | 15.4058 | 14 | 14.7139 | 12.29 |  |  | 15.1416
| 1 | 5 | 12.64 | 14.9 | 14.9 | 14.93 | 14.9258 | 14.7 | 16.4 | 18.627 | 15.12 | 15.5 | 15.5 | 14.95 | 14.65 | 14.65 | 15.5 | 16.05 | 14.36 | 10.71 | < 20 | 12.54 | 15.1649 | 14 | 14.8104 | 12.74 |  |  | 62
| 1 | 6 | 14.38 | 14.5 | 16 | 15.02 | 14.3311 | 15.1 | 16.1 | 18.405 | 15.43 | 15 | 15.86 | 14.75 | 14.6 | 14.95 | 15.85 | 16.6 | 14.57 | 11.65 | < 20 | 12.93 | 16.4329 | 14 | 15.169 | 12.63
| Mean |  | 13.97 | 15.45 | 15.1667 | 15.01 | 14.6202 | 15.05 | 16.15 | 18.6748 | 15.1017 | 15.35 | 15.475 | 14.6333 | 14.575 | 14.8583 | 15.6083 | 16.5583 | 14.4283 | 12.07 |  | 12.6817 | 15.408 | 14 | 15.0362 | 12.615
| Median |  | 14.195 | 15.35 | 15 | 15.015 | 14.6187 | 15.1 | 16.1 | 18.738 | 15.09 | 15.45 | 15.53 | 14.75 | 14.625 | 14.825 | 15.65 | 16.575 | 14.425 | 12.175 |  | 12.64 | 15.3574 | 14 | 14.999 | 12.685
| Std Dev. |  | 0.7463 | 0.90056 | 0.4274 | 0.15166 | 0.31019 | 0.29496 | 0.20736 | 0.27688 | 0.35119 | 0.2429 | 0.29541 | 0.31252 | 0.14748 | 0.28358 | 0.32158 | 0.38525 | 0.10342 | 1.268 |  | 0.16266 | 0.58481 | 0 | 0.28621 | 0.25936
| Rel.Std.Dev. |  | 5.342% | 5.829% | 2.818% | 1.01% | 2.122% | 1.96% | 1.284% | 1.483% | 2.326% | 1.582% | 1.909% | 2.136% | 1.012% | 1.909% | 2.06% | 2.327% | 0.7168% | 10.51% |  | 1.283% | 3.795% | 0% | 1.903% | 2.056%
| PDM3 |  | -7.738% | 2.037% | 0.1656% | -0.8691% | -3.444% | -0.6049% | 6.66% | 23.33% | -0.2637% | 1.376% | 2.202% | -3.357% | -3.742% | -1.871% | 3.082% | 9.357% | -4.711% | -20.29% |  | -16.25% | 1.76% | -7.539% | -0.6958% | -16.69%
| Z-Score (Absolute) |  | 1.64 | 0.64 | 0.2 | 0.04 | 0.64 | 0.02 | 1.71 | 5.59 | 0.1 | 0.48 | 0.67 | 0.62 | 0.71 | 0.27 | 0.88 | 2.34 | 0.93 | 4.56 | 7.74 | 3.62 | 0.57 | NA | 0 | 3.72
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 600c (Certified Value 0.113 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.12 | 0.122 | 0.117 | 0.1128 | 0.11513 | 0.112 | 0.11 | 0.11149 | 0.113 | 0.114 | 0.1134 | 0.11 | 0.114 | 0.108 | 0.103 | 0.115 | 0.1263 | 0.115 | 0.119 | 0.11 | 0.11 | < 0.1 | 0.108 | 0.11465 | 0.123 |  | 1
| 1 | 2 | 0.12 | 0.121 | 0.117 | 0.113 | 0.11458 | 0.113 | 0.11 | 0.11149 | 0.113 | 0.112 | 0.1115 | 0.108 | 0.112 | 0.108 | 0.103 | 0.116 | 0.1272 | 0.115 | 0.117 | 0.11 | 0.11 | < 0.1 | 0.114 | 0.11467 | 0.122 |  | 29
| 1 | 3 | 0.12 | 0.126 | 0.115 | 0.1117 | 0.11664 | 0.111 | 0.11 | 0.11116 | 0.113 | 0.11 | 0.1103 | 0.112 | 0.113 | 0.108 | 0.103 | 0.115 | 0.1191 | 0.114 | 0.123 | 0.11 | 0.11 | < 0.1 | 0.108 | 0.11267 | 0.121 |  | 16
| 1 | 4 | 0.12 | 0.124 | 0.117 | 0.1113 | 0.11495 | 0.113 | 0.11 | 0.1119 | 0.113 | 0.11 | 0.1136 | 0.109 | 0.113 | 0.107 | 0.104 | 0.119 | 0.1285 | 0.111 | 0.122 | 0.1 | 0.11 | < 0.1 | 0.108 | 0.11216 | 0.119 |  | 0.11312
| 1 | 5 | 0.12 | 0.122 | 0.117 | 0.1123 | 0.11826 | 0.109 | 0.11 | 0.11122 | 0.117 | 0.114 | 0.1101 | 0.109 | 0.114 | 0.107 | 0.104 | 0.114 | 0.1271 | 0.114 | 0.121 | 0.11 | 0.11 | < 0.1 | 0.108 | 0.11261 | 0.12 |  | 63
| 1 | 6 | 0.12 | 0.117 | 0.117 | 0.1098 | 0.11278 | 0.112 | 0.11 | 0.1113 | 0.113 | 0.115 | 0.1129 | 0.108 | 0.11 | 0.109 | 0.103 | 0.116 | 0.1279 | 0.115 | 0.122 | 0.11 | 0.11 | < 0.1 | 0.108 | 0.11129 | 0.12
| Mean |  | 0.12 | 0.122 | 0.11667 | 0.11182 | 0.11539 | 0.11167 | 0.11 | 0.11142 | 0.11367 | 0.1125 | 0.11197 | 0.10933 | 0.11267 | 0.10783 | 0.10333 | 0.11583 | 0.12602 | 0.114 | 0.12067 | 0.10833 | 0.11 |  | 0.109 | 0.11301 | 0.12083
| Median |  | 0.12 | 0.122 | 0.117 | 0.112 | 0.11504 | 0.112 | 0.11 | 0.11139 | 0.113 | 0.113 | 0.1122 | 0.109 | 0.113 | 0.108 | 0.103 | 0.1155 | 0.12715 | 0.1145 | 0.1215 | 0.11 | 0.11 |  | 0.108 | 0.11264 | 0.1205
| Std Dev. |  | 0 | 0.0030332 | 0.0008165 | 0.0011788 | 0.0018711 | 0.0015055 | 0 | 0.00027166 | 0.001633 | 0.0021679 | 0.0015539 | 0.0015055 | 0.0015055 | 0.00075277 | 0.0005164 | 0.0017224 | 0.0034701 | 0.0015492 | 0.0022509 | 0.0040825 | 0 |  | 0.0024495 | 0.001372 | 0.001472
| Rel.Std.Dev. |  | 0% | 2.486% | 0.6999% | 1.054% | 1.622% | 1.348% | 0% | 0.2438% | 1.437% | 1.927% | 1.388% | 1.377% | 1.336% | 0.6981% | 0.4997% | 1.487% | 2.754% | 1.359% | 1.865% | 3.768% | 0% |  | 2.247% | 1.214% | 1.218%
| PDM3 |  | 6.081% | 7.849% | 3.135% | -1.153% | 2.006% | -1.286% | -2.759% | -1.499% | 0.4825% | -0.5488% | -1.02% | -3.348% | -0.4015% | -4.674% | -8.652% | 2.398% | 11.4% | 0.7772% | 6.671% | -4.232% | -2.759% |  | -3.643% | -0.09941% | 6.818%
| Z-Score (Absolute) |  | 1.6 | 2.02 | 0.89 | 0.15 | 0.62 | 0.18 | 0.53 | 0.23 | 0.25 | 0 | 0.11 | 0.67 | 0.04 | 0.99 | 1.95 | 0.71 | 2.88 | 0.32 | 1.74 | 0.89 | 0.53 | 13.31 | 0.77 | 0.11 | 1.77
| Analytical results for Tl in OREAS 600c (Certified Value 1.21 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 0.69 | 1.23 | 1.2 | 1.22 | 1.1735 | 1.21 | 1.25 | 1.19 | 1.17 | 1.26 | 1.21 | 1.12 | 1.14 | 1.23 | 1.17 | 1.131 | 1.4 | 1.31 | 0.85 | 1.1558 | 0.99 | [5 empty] | 1
| 1 | 2 | 0.63 | 1.12 | 1.3 | 1.22 | 1.1863 | 1.21 | 1.25 | 1.26 | 1.19 | 1.3 | 1.25 | 1.11 | 1.2 | 1.2 | 1.2 | 1.14 | 1.37 | 1.27 | 0.84 | 1.1864 | 0.99 | [5 empty] | 30
| 1 | 3 | 0.62 | 1.19 | 1.2 | 1.25 | 1.1711 | 1.17 | 1.26 | 1.24 | 1.16 | 1.3 | 1.2 | 1.11 | 1.21 | 1.22 | 1.2 | 1.148 | 1.37 | 1.3 | 0.89 | 1.1941 | 0.97 | [5 empty] | 16
| 1 | 4 | 0.7 | 1.16 | 1.3 | 1.24 | 1.1925 | 1.21 | 1.24 | 1.17 | 1.22 | 1.26 | 1.22 | 1.12 | 1.16 | 1.21 | 1.21 | 1.121 | 1.37 | 1.29 | 0.9 | 1.1745 | 0.98 | [5 empty] | 1.2055
| 1 | 5 | 0.66 | 1.13 | 1.2 | 1.27 | 1.1861 | 1.17 | 1.28 | 1.2 | 1.22 | 1.3 | 1.22 | 1.14 | 1.18 | 1.21 | 1.15 | 1.122 | 1.33 | 1.3 | 0.89 | 1.176 | 0.98 | [5 empty] | 64
| 1 | 6 | 0.65 | 1.12 | 1.3 | 1.22 | 1.1692 | 1.19 | 1.27 | 1.19 | 1.22 | 1.27 | 1.22 | 1.12 | 1.22 | 1.2 | 1.19 | 1.136 | 1.35 | 1.33 | 0.85 | 1.1603 | 0.97
| Mean |  | 0.65833 | 1.1583 | 1.25 | 1.2367 | 1.1798 | 1.1933 | 1.2583 | 1.2083 | 1.1967 | 1.2817 | 1.22 | 1.12 | 1.185 | 1.2117 | 1.1867 | 1.133 | 1.365 | 1.3 | 0.87 | 1.1745 | 0.98
| Median |  | 0.655 | 1.145 | 1.25 | 1.23 | 1.1798 | 1.2 | 1.255 | 1.195 | 1.205 | 1.285 | 1.22 | 1.12 | 1.19 | 1.21 | 1.195 | 1.1335 | 1.37 | 1.3 | 0.87 | 1.1752 | 0.98
| Std Dev. |  | 0.031885 | 0.04446 | 0.054772 | 0.020656 | 0.0097134 | 0.019664 | 0.01472 | 0.034303 | 0.027325 | 0.020412 | 0.016733 | 0.010954 | 0.030822 | 0.01169 | 0.022509 | 0.010507 | 0.023452 | 0.02 | 0.026077 | 0.014681 | 0.0089443
| Rel.Std.Dev. |  | 4.843% | 3.838% | 4.382% | 1.67% | 0.8233% | 1.648% | 1.17% | 2.839% | 2.283% | 1.593% | 1.372% | 0.9781% | 2.601% | 0.9648% | 1.897% | 0.9274% | 1.718% | 1.538% | 2.997% | 1.25% | 0.9127%
| PDM3 |  | -45.39% | -3.915% | 3.689% | 2.583% | -2.135% | -1.012% | 4.38% | 0.2327% | -0.735% | 6.316% | 1.2% | -7.095% | -1.703% | 0.5092% | -1.565% | -6.016% | 13.23% | 7.837% | -27.83% | -2.572% | -18.71%
| Z-Score (Absolute) |  | 8.33 | 0.54 | 0.88 | 0.67 | 0.21 | 0 | 1.01 | 0.23 | 0.05 | 1.37 | 0.41 | 1.14 | 0.13 | 0.29 | 0.1 | 0.94 | 2.67 | 1.66 | 5.03 | 0.29 | 3.32
| Analytical results for Tm in OREAS 600c (Certified Value 0.11 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | [5 empty] | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.28 | [5 empty] | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.01: 0.1; 2.03: 0.11; 2.04: 0.10308; 2.05: 0.1; 2.06: 0.1106; 2.07: 0.11; 2.08: 0.12; 2.10: 0.11; 2.11: 0.1118; 2.12: 0.1 | [5 empty] | 1
| 1 | 2 | 2.01: < 0.1; 2.03: 0.11; 2.04: 0.10763; 2.05: 0.1; 2.06: < 0.1; 2.07: 0.11; 2.08: 0.12; 2.10: 0.1; 2.11: 0.1189; 2.12: 0.1 | [5 empty] | 5
| 1 | 3 | 2.01: 0.1; 2.03: 0.11; 2.04: 0.10918; 2.05: 0.1; 2.06: 0.1236; 2.07: 0.11; 2.08: 0.12; 2.10: 0.11; 2.11: 0.1097; 2.12: 0.1 | [5 empty] | 16
| 1 | 4 | 2.01: 0.1; 2.03: 0.12; 2.04: 0.10593; 2.05: 0.1; 2.06: 0.1323; 2.07: 0.11; 2.08: 0.09; 2.10: 0.1; 2.11: 0.1105; 2.12: 0.11 | [5 empty] | 0.10967
| 1 | 5 | 2.01: < 0.1; 2.03: 0.12; 2.04: 0.10628; 2.05: 0.1; 2.06: 0.1262; 2.07: 0.11; 2.08: 0.1; 2.10: 0.1; 2.11: 0.1136; 2.12: 0.1 | [5 empty] | 65
| 1 | 6 | 2.01: < 0.1; 2.03: 0.12; 2.04: 0.10342; 2.05: 0.1; 2.06: 0.1077; 2.07: 0.11; 2.08: 0.11; 2.10: 0.1; 2.11: 0.1138; 2.12: 0.09
| Mean |  | 2.01: 0.1; 2.03: 0.115; 2.04: 0.10592; 2.05: 0.1; 2.06: 0.12008; 2.07: 0.11; 2.08: 0.11; 2.10: 0.10333; 2.11: 0.11305; 2.12: 0.1
| Median |  | 2.01: 0.1; 2.03: 0.115; 2.04: 0.10611; 2.05: 0.1; 2.06: 0.1236; 2.07: 0.11; 2.08: 0.115; 2.10: 0.1; 2.11: 0.1127; 2.12: 0.1
| Std Dev. |  | 2.01: 1.6997e-17; 2.03: 0.0054772; 2.04: 0.002366; 2.05: 1.5202e-17; 2.06: 0.010516; 2.07: 0; 2.08: 0.012649; 2.10: 0.005164; 2.11: 0.0032977; 2.12: 0.0063246
| Rel.Std.Dev. |  | 2.01: 1.7e-14%; 2.03: 4.763%; 2.04: 2.234%; 2.05: 1.52e-14%; 2.06: 8.757%; 2.07: 0%; 2.08: 11.5%; 2.10: 4.997%; 2.11: 2.917%; 2.12: 6.325%
| PDM3 |  | 2.01: -8.82%; 2.03: 4.857%; 2.04: -3.423%; 2.05: -8.82%; 2.06: 9.489%; 2.07: 0.2985%; 2.08: 0.2985%; 2.10: -5.78%; 2.11: 3.079%; 2.12: -8.82%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 1.1; 2.04: 0.62; 2.05: NA; 2.06: 0.15; 2.07: 0.15; 2.08: 0.15; 2.10: 1.11; 2.11: 0.73; 2.12: 1.74
Results from laboratories 2.03 and 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 600c (Certified Value 5.79 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 4.5 | 6.3 | 5.8 | 5.9 | 5.3839 | 6 | 5.83 | 6.9715 | 5.68 | 5.92 | 6.02 | 5.4 | 5.7 | 5.8 | 5.8 | 5.7 | 5.68 | 5.86 | 6.47 | 5.5 | 5.292 | 2 | 5.7552 | 4.77 |  |  | 1
| 1 | 2 | 4.7 | 5.3 | 5.9 | 5.98 | 5.6305 | 6 | 5.8 | 6.7337 | 5.95 | 6.03 | 5.88 | 5.6 | 5.5 | 5.9 | 5.6 | 5.9 | 5.69 | 6.21 | 6.24 | 5.4 | 5.0264 | 3 | 6.0936 | 4.66 |  |  | 32
| 1 | 3 | 4.7 | 5.8 | 5.7 | 5.87 | 5.9253 | 5.8 | 5.8 | 6.8906 | 5.83 | 5.97 | 5.96 | 5.3 | 5.7 | 6 | 5.5 | 5.8 | 5.59 | 5.77 | 6.51 | 5.5 | 5.1653 | 2 | 6.0721 | 4.63 |  |  | 16
| 1 | 4 | 4.8 | 5.7 | 6.1 | 6.02 | 5.9907 | 5.8 | 5.81 | 6.0602 | 5.67 | 6.1 | 5.84 | 5.5 | 5.6 | 5.8 | 5.8 | 6 | 5.74 | 4.14 | 6.39 | 5.5 | 5.238 | 2 | 5.7965 | 4.66 |  |  | 5.7859
| 1 | 5 | 4.2 | 5.6 | 5.9 | 6.1 | 6.0028 | 5.8 | 5.96 | 6.7011 | 5.88 | 5.96 | 5.99 | 5.3 | 5.7 | 5.8 | 5.8 | 5.8 | 5.67 | 4.24 | 6.5 | 5.4 | 5.2514 | 2 | 5.8904 | 4.71 |  |  | 66
| 1 | 6 | 4.5 | 5.2 | 5.9 | 5.84 | 5.7077 | 5.9 | 5.82 | 6.7797 | 5.92 | 5.84 | 6.15 | 5.5 | 5.7 | 5.9 | 5.8 | 5.8 | 5.74 | 4.82 | 6.54 | 5.5 | 5.1761 | 2 | 5.9189 | 4.7
| Mean |  | 4.5667 | 5.65 | 5.8833 | 5.9517 | 5.7735 | 5.8833 | 5.8367 | 6.6895 | 5.8217 | 5.97 | 5.9733 | 5.4333 | 5.65 | 5.8667 | 5.7167 | 5.8333 | 5.685 | 5.1733 | 6.4417 | 5.4667 | 5.1915 | 2.1667 | 5.9211 | 4.6883
| Median |  | 4.6 | 5.65 | 5.9 | 5.94 | 5.8165 | 5.85 | 5.815 | 6.7567 | 5.855 | 5.965 | 5.975 | 5.45 | 5.7 | 5.85 | 5.8 | 5.8 | 5.685 | 5.295 | 6.485 | 5.5 | 5.207 | 2 | 5.9046 | 4.68
| Std Dev. |  | 0.21602 | 0.3937 | 0.13292 | 0.099281 | 0.24466 | 0.098319 | 0.061536 | 0.32441 | 0.12057 | 0.089443 | 0.10985 | 0.12111 | 0.083666 | 0.08165 | 0.13292 | 0.10328 | 0.055408 | 0.89061 | 0.11125 | 0.05164 | 0.093843 | 0.40825 | 0.13897 | 0.049565
| Rel.Std.Dev. |  | 4.73% | 6.968% | 2.259% | 1.668% | 4.238% | 1.671% | 1.054% | 4.85% | 2.071% | 1.498% | 1.839% | 2.229% | 1.481% | 1.392% | 2.325% | 1.771% | 0.9746% | 17.22% | 1.727% | 0.9446% | 1.808% | 18.84% | 2.347% | 1.057%
| PDM3 |  | -21.07% | -2.348% | 1.685% | 2.866% | -0.2137% | 1.685% | 0.8781% | 15.62% | 0.6189% | 3.183% | 3.24% | -6.093% | -2.348% | 1.397% | -1.196% | 0.8205% | -1.743% | -10.59% | 11.33% | -5.517% | -10.27% | -62.55% | 2.338% | -18.97%
| Z-Score (Absolute) |  | 5.71 | 0.78 | 0.28 | 0.59 | 0.22 | 0.28 | 0.07 | 3.94 | 0 | 0.67 | 0.69 | 1.77 | 0.78 | 0.2 | 0.48 | 0.05 | 0.62 | 2.95 | 2.82 | 1.61 | 2.86 | NA | 0.45 | 5.15
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 600c (Certified Value 3.81 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 2
2
| 1 | 1 | 4 | 4 | 4 | 4 | 3.3877 | 4 | 3.6 | 3.57 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 3 | 5 | 4.3 | 11.06 | 4 | < 0.1 | 3 | 4.956 | < 10 | 5 |  | 1
| 1 | 2 | 4 | 4 | 4 | 4 | 3.6001 | 4 | 3.7 | 3.73 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 3 | 5 | 4.4 | 11.84 | 4 | < 0.1 | 3 | 5.0775 | < 10 | 6 |  | 33
| 1 | 3 | 4 | 4 | 4 | 4 | 3.6434 | 4 | 3.6 | 3.63 | 5 | 4 | 4 | 3 | 3 | 3 | 3 | 3 | 5 | 4.2 | 11.71 | 4 | < 0.1 | 3 | 5.0459 | < 10 | 5 |  | 16
| 1 | 4 | 4 | 4 | 4 | 4 | 3.6177 | 4 | 3.4 | 3.6 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 3 | 5 | 4.3 | 12.5 | 3 | < 0.1 | 3 | 4.9773 | < 10 | 5 |  | 3.8054
| 1 | 5 | 3 | 4 | 4 | 4 | 3.6666 | 4 | 3.4 | 3.53 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 3 | 4 | 4 | 11.76 | 4 | < 0.1 | 3 | 4.9421 | < 10 | 5 |  | 67
| 1 | 6 | 4 | 4 | 4 | 4 | 3.404 | 4 | 3.5 | 3.47 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 3 | 5 | 4 | 12.22 | 4 | < 0.1 | 4 | 5.0648 | < 10 | 5
| Mean |  | 3.8333 | 4 | 4 | 4 | 3.5532 | 4 | 3.5333 | 3.5883 | 5 | 4 | 3.1667 | 3 | 3 | 3 | 3 | 3 | 4.8333 | 4.2 | 11.8483 | 3.8333 |  | 3.1667 | 5.0106 |  | 5.1667
| Median |  | 4 | 4 | 4 | 4 | 3.6089 | 4 | 3.55 | 3.585 | 5 | 4 | 3 | 3 | 3 | 3 | 3 | 3 | 5 | 4.25 | 11.8 | 4 |  | 3 | 5.0116 |  | 5
| Std Dev. |  | 0.40825 | 0 | 0 | 0 | 0.1241 | 0 | 0.12111 | 0.089088 | 0 | 0 | 0.40825 | 0 | 0 | 0 | 0 | 0 | 0.40825 | 0.16733 | 0.4922 | 0.40825 |  | 0.40825 | 0.059072 |  | 0.40825
| Rel.Std.Dev. |  | 10.65% | 0% | 0% | 0% | 3.493% | 0% | 3.428% | 2.483% | 0% | 0% | 12.89% | 0% | 0% | 0% | 0% | 0% | 8.447% | 3.984% | 4.154% | 10.65% |  | 12.89% | 1.179% |  | 7.902%
| PDM3 |  | 0.734% | 5.114% | 5.114% | 5.114% | -6.627% | 5.114% | -7.15% | -5.704% | 31.39% | 5.114% | -16.78% | -21.16% | -21.16% | -21.16% | -21.16% | -21.16% | 27.01% | 10.37% | 211.4% | 0.734% |  | -16.78% | 31.67% |  | 35.77%
| Z-Score (Absolute) |  | 0 | 0.17 | 0.17 | 0.17 | 0.28 | 0.17 | 0.3 | 0.25 | 1.18 | 0.17 | 0.67 | 0.84 | 0.84 | 0.84 | 0.84 | 0.84 | 1.01 | 0.37 | 8.11 | 0 | 3.83 | 0.67 | 1.19 | 1.18 | 1.35
| Analytical results for W in OREAS 600c (Certified Value 1.62 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 1 | 1.6 | 1.6 | 1.6 | < 10 | 1.6 | 1.6 | 1.67 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6 | 1.7 | 1.6 | 1.56 | 2 | 1.94 | < 5 | 0.9 |  | < 1 | 1.5668 | 1.5 |  |  | 1
| 1 | 2 | 0.9 | 1.7 | 1.7 | 1.6 | < 10 | 1.7 | 1.7 | 1.67 | 1.7 | 1.6 | 1.7 | 1.5 | 1.7 | 1.6 | 1.6 | 1.66 | 2 | 1.86 | < 5 | 1 | 2.6414 | < 1 | 1.6017 | 1.4 |  |  | 34
| 1 | 3 | 0.9 | 1.5 | 1.7 | 1.6 | < 10 | 1.7 | 1.7 | 1.83 | 1.6 | 1.6 | 1.7 | 1.6 | 1.7 | 1.6 | 1.6 | 1.52 | 2.2 | 2.05 | < 5 | 1 | 6.036 | < 1 | 1.5815 | 1.6 |  |  | 16
| 1 | 4 | 0.9 | 1.5 | 1.6 | 1.6 | < 10 | 1.7 | 1.6 | 1.7 | 1.8 | 1.7 | 1.7 | 1.6 | 1.6 | 1.6 | 1.6 | 1.66 | 2 | 1.91 | < 5 | 1 | 4.1897 | < 1 | 1.5948 | 1.5 |  |  | 1.6193
| 1 | 5 | 0.8 | 1.7 | 1.6 | 1.8 | < 10 | 1.7 | 1.6 | 1.63 | 1.8 | 1.8 | 1.7 | 1.6 | 1.6 | 1.6 | 1.6 | 1.61 | 2 | 1.91 | < 5 | 1.1 |  | < 1 | 1.5623 | 1.4 |  |  | 68
| 1 | 6 | 1 | 1.4 | 1.6 | 1.6 | < 10 | 1.5 | 1.7 | 1.7 | 1.6 | 1.8 | 1.6 | 1.5 | 1.6 | 1.6 | 1.6 | 1.62 | 2.1 | 2 | < 5 | 0.9 |  | < 1 | 1.5697 | 1.4
| Mean |  | 0.91667 | 1.5667 | 1.6333 | 1.6333 |  | 1.65 | 1.65 | 1.7 | 1.6833 | 1.6833 | 1.6667 | 1.5667 | 1.6333 | 1.6167 | 1.6 | 1.605 | 2.05 | 1.945 |  | 0.98333 | 4.289 |  | 1.5795 | 1.4667
| Median |  | 0.9 | 1.55 | 1.6 | 1.6 |  | 1.7 | 1.65 | 1.685 | 1.65 | 1.65 | 1.7 | 1.6 | 1.6 | 1.6 | 1.6 | 1.615 | 2 | 1.925 |  | 1 | 4.1897 |  | 1.5756 | 1.45
| Std Dev. |  | 0.075277 | 0.12111 | 0.05164 | 0.08165 |  | 0.083666 | 0.054772 | 0.068702 | 0.098319 | 0.098319 | 0.05164 | 0.05164 | 0.05164 | 0.040825 | 2.4324e-16 | 0.055767 | 0.083666 | 0.06892 |  | 0.075277 | 1.6995 |  | 0.016029 | 0.08165
| Rel.Std.Dev. |  | 8.212% | 7.73% | 3.162% | 4.999% |  | 5.071% | 3.32% | 4.041% | 5.841% | 5.841% | 3.098% | 3.296% | 3.162% | 2.525% | 1.52e-14% | 3.475% | 4.081% | 3.543% |  | 7.655% | 39.62% |  | 1.015% | 5.567%
| PDM3 |  | -43.39% | -3.249% | 0.8679% | 0.8679% |  | 1.897% | 1.897% | 4.985% | 3.956% | 3.956% | 2.926% | -3.249% | 0.8679% | -0.1613% | -1.191% | -0.8818% | 26.6% | 20.12% |  | -39.27% | 164.9% |  | -2.459% | -9.425%
| Z-Score (Absolute) |  | 8.02 | 0.75 | 0 | 0 | 37.67 | 0.19 | 0.19 | 0.75 | 0.56 | 0.56 | 0.37 | 0.75 | 0 | 0.19 | 0.37 | 0.32 | 4.66 | 3.49 | 9.7 | 7.27 | 29.71 | 12.68 | 0.6 | 1.86
| Analytical results for Y in OREAS 600c (Certified Value 15.7 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 17.3 | 17.4 | 15.6 | 15.3 | 15.1568 | 15.2 | 17 | 15.5997 | 15.1 | 15.1 | 16.24 | 15.8 | 15.1 | 15.4 | 16.9 | 14 | 15.51 | 13.22 | 17 | 16.2 | 15.4247 | 17 | 15.593 | 13.092 | 15.5 |  | 1
| 1 | 2 | 16.8 | 15.8 | 15.7 | 15.2 | 15.051 | 15.4 | 17.2 | 15.0732 | 15.6 | 15.3 | 16.06 | 15.9 | 14.6 | 15.6 | 16.6 | 15.4 | 15.38 | 14.64 | 16.6 | 16.1 | 15.1523 | 17 | 16.1643 | 13.217 | 15.39 |  | 35
| 1 | 3 | 17.7 | 17.1 | 15.2 | 15.22 | 15.3039 | 14.8 | 16.3 | 15.7445 | 15.5 | 15.4 | 16.6 | 15.6 | 14.8 | 15.5 | 16.1 | 15.8 | 15.43 | 13.32 | 17.7 | 16.3 | 15.6702 | 17 | 16.1703 | 13.335 | 14.94 |  | 16
| 1 | 4 | 17.2 | 16.4 | 15.4 | 15.25 | 15.1813 | 15.3 | 16.5 | 15.75 | 15.5 | 15.4 | 15.93 | 15.6 | 14.7 | 15.4 | 16.7 | 16.2 | 15.75 | 10.39 | 16.7 | 16.6 | 15.93 | 16 | 15.5373 | 13.215 | 15.19 |  | 15.6674
| 1 | 5 | 17.5 | 15.8 | 15.6 | 15.21 | 15.1989 | 15.5 | 16.4 | 15.8346 | 15 | 15.5 | 16.19 | 15.5 | 15.2 | 15.1 | 16.7 | 15.4 | 15.56 | 11.52 | 17 | 16.2 | 16.1263 | 17 | 15.7289 | 13.231 | 14.96 |  | 69
| 1 | 6 | 17.9 | 15.1 | 15.7 | 15.19 | 15.1417 | 15.5 | 16.8 | 15.7392 | 15.3 | 15.1 | 16.41 | 15.8 | 14.5 | 16 | 16.4 | 15.5 | 15.59 | 12.44 | 17.3 | 16.1 | 15.643 | 17 | 16.0641 | 13.511 | 15.35
| Mean |  | 17.4 | 16.2667 | 15.5333 | 15.2283 | 15.1723 | 15.2833 | 16.7 | 15.6235 | 15.3333 | 15.3 | 16.2383 | 15.7 | 14.8167 | 15.5 | 16.5667 | 15.3833 | 15.5367 | 12.5883 | 17.05 | 16.25 | 15.6578 | 16.8333 | 15.8763 | 13.2668 | 15.2217
| Median |  | 17.4 | 16.1 | 15.6 | 15.215 | 15.169 | 15.35 | 16.65 | 15.7418 | 15.4 | 15.35 | 16.215 | 15.7 | 14.75 | 15.45 | 16.65 | 15.45 | 15.535 | 12.83 | 17 | 16.2 | 15.6566 | 17 | 15.8965 | 13.224 | 15.27
| Std Dev. |  | 0.38987 | 0.87101 | 0.19664 | 0.040702 | 0.082473 | 0.26394 | 0.35777 | 0.28 | 0.24221 | 0.16733 | 0.24062 | 0.15492 | 0.27869 | 0.29665 | 0.28048 | 0.74409 | 0.13079 | 1.4926 | 0.40373 | 0.18708 | 0.34751 | 0.40825 | 0.29034 | 0.14233 | 0.23284
| Rel.Std.Dev. |  | 2.241% | 5.355% | 1.266% | 0.2673% | 0.5436% | 1.727% | 2.142% | 1.792% | 1.58% | 1.094% | 1.482% | 0.9867% | 1.881% | 1.914% | 1.693% | 4.837% | 0.8418% | 11.86% | 2.368% | 1.151% | 2.219% | 2.425% | 1.829% | 1.073% | 1.53%
| PDM3 |  | 11.06% | 3.825% | -0.8558% | -2.802% | -3.16% | -2.451% | 6.591% | -0.2801% | -2.132% | -2.345% | 3.644% | 0.208% | -5.43% | -1.069% | 5.74% | -1.813% | -0.8345% | -19.65% | 8.825% | 3.718% | -0.06167% | 7.442% | 1.333% | -15.32% | -2.845%
| Z-Score (Absolute) |  | 3.84 | 1.51 | 0 | 0.63 | 0.75 | 0.52 | 2.4 | 0.18 | 0.42 | 0.48 | 1.45 | 0.34 | 1.48 | 0.07 | 2.13 | 0.31 | 0 | 6.07 | 3.12 | 1.47 | 0.25 | NA | 0.7 | 4.67 | 0.65
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 600c (Certified Value 0.62 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.10; 2.08: 2.12; 2.09: 2.19; 2.10: 2.22; 2.11: 2.23; 2.12: 2.25; 2.13: 2.28 |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*MS |  | 2
2
| 1 | 1 | 2.01: 0.6; 2.03: 0.51; 2.04: 0.66972; 2.05: 0.6; 2.06: 0.7195; 2.07: 0.51; 2.08: 0.58; 2.09: 0.6; 2.10: 0.62; 2.11: 0.6854; 2.12: 0.7; 2.13: 0.5 |  | 1
| 1 | 2 | 2.01: 0.5; 2.03: 0.53; 2.04: 0.69132; 2.05: 0.6; 2.06: 0.6984; 2.07: 0.58; 2.08: 0.59; 2.09: 0.6; 2.10: 0.64; 2.11: 0.654; 2.12: 0.7; 2.13: 0.5 |  | 36
| 1 | 3 | 2.01: 0.6; 2.03: 0.58; 2.04: 0.69473; 2.05: 0.6; 2.06: 0.7773; 2.07: 0.53; 2.08: 0.59; 2.09: 0.6; 2.10: 0.67; 2.11: 0.6787; 2.12: 0.7; 2.13: 0.5 |  | 16
| 1 | 4 | 2.01: 0.6; 2.03: 0.54; 2.04: 0.68535; 2.05: 0.6; 2.06: 0.742; 2.07: 0.52; 2.08: 0.57; 2.09: 0.6; 2.10: 0.65; 2.11: 0.685; 2.12: 0.7; 2.13: 0.5 |  | 0.62221
| 1 | 5 | 2.01: 0.6; 2.03: 0.54; 2.04: 0.67278; 2.05: 0.6; 2.06: 0.6997; 2.07: 0.49; 2.08: 0.58; 2.09: 0.6; 2.10: 0.66; 2.11: 0.6818; 2.12: 0.7; 2.13: 0.5 |  | 70
| 1 | 6 | 2.01: 0.5; 2.03: 0.48; 2.04: 0.6524; 2.05: 0.6; 2.06: 0.6311; 2.07: 0.54; 2.08: 0.58; 2.09: 0.6; 2.10: 0.63; 2.11: 0.7036; 2.12: 0.7; 2.13: 0.5
| Mean |  | 2.01: 0.56667; 2.03: 0.53; 2.04: 0.67772; 2.05: 0.6; 2.06: 0.71133; 2.07: 0.52833; 2.08: 0.58167; 2.09: 0.6; 2.10: 0.645; 2.11: 0.68142; 2.12: 0.7; 2.13: 0.5
| Median |  | 2.01: 0.6; 2.03: 0.535; 2.04: 0.67907; 2.05: 0.6; 2.06: 0.7096; 2.07: 0.525; 2.08: 0.58; 2.09: 0.6; 2.10: 0.645; 2.11: 0.6834; 2.12: 0.7; 2.13: 0.5
| Std Dev. |  | 2.01: 0.05164; 2.03: 0.033466; 2.04: 0.015888; 2.05: 0; 2.06: 0.049191; 2.07: 0.030605; 2.08: 0.0075277; 2.09: 0; 2.10: 0.018708; 2.11: 0.016; 2.12: 1.2162e-16; 2.13: 0
| Rel.Std.Dev. |  | 2.01: 9.113%; 2.03: 6.314%; 2.04: 2.344%; 2.05: 0%; 2.06: 6.915%; 2.07: 5.793%; 2.08: 1.294%; 2.09: 0%; 2.10: 2.901%; 2.11: 2.348%; 2.12: 1.737e-14%; 2.13: 0%
| PDM3 |  | 2.01: -8.927%; 2.03: -14.82%; 2.04: 8.921%; 2.05: -3.569%; 2.06: 14.32%; 2.07: -15.09%; 2.08: -6.516%; 2.09: -3.569%; 2.10: 3.663%; 2.11: 9.516%; 2.12: 12.5%; 2.13: -19.64%
| Z-Score (Absolute) |  | 2.01: NA; 2.03: 1.22; 2.04: 0.35; 2.05: NA; 2.06: 0.71; 2.07: 1.24; 2.08: 0.67; 2.09: NA; 2.10: 0; 2.11: 0.39; 2.12: NA; 2.13: NA
Results from laboratories 2.03, 2.07, 2.19, 2.25 and 2.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 600c (Certified Value 2106 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 2203 | 2151.5 | 2100 | 2048 | 2056.6842 | 1910 | 2090 | 2123.576 | 2035 | 2200 | 2232 | 2130 | 2140 | 2050 | 2040 | 2260 | 2236 | 2219 | 2450 | 2201 | 2208 | 1738.9182 | 2070 | 2072.3507 | 1886.672 | 2090 | 1
| 1 | 2 | 2197 | 2122.1 | 2127 | 2050 | 2061.6559 | 1900 | 2120 | 2157.16 | 2018 | 2150 | 2196 | 2140 | 2120 | 2050 | 2020 | 2320 | 2194 | 1979 | 2550 | 2187 | 2107 | 1698.3012 | 2050 | 2206.3654 | 2027.744 | 2068 | 17
| 1 | 3 | 2171 | 2144.8 | 2094 | 2105 | 2081.3257 | 1900 | 1990 | 2128.712 | 2075 | 2100 | 2185 | 2140 | 2130 | 2080 | 2040 | 2280 | 2200 | 1990 | 2450 | 2180 | 2153 | 1720.895 | 1970 | 2168.2421 | 1994.23 | 1998 | 16
| 1 | 4 | 2183 | 2130.6 | 2133 | 2101 | 2072.8525 | 1930 | 2050 | 2115.96 | 2043 | 2170 | 2223 | 2120 | 2120 | 2040 | 2060 | 2350 | 2246 | 1990 | 2530 | 2135 | 2188 | 1700.1073 | 2030 | 2148.7784 | 1947.422 | 2059 | 2105.9993
| 1 | 5 | 2179 | 2188.5 | 2152 | 2113 | 2079.3694 | 1840 | 2050 | 2144.168 | 2060 | 2140 | 2180 | 2140 | 2170 | 2050 | 2050 | 2260 | 2151 | 1942 | 2450 | 2188 | 2129 | 1720.4104 | 2040 | 2091.1932 | 2033.509 | 2084 | 71
| 1 | 6 | 2191 | 2050.4 | 2117 | 2026 | 2024.2945 | 1950 | 2060 | 2151 | 2067 | 2180 | 2209 | 2110 | 2080 | 2100 | 2020 | 2310 | 2216 | 2441 | 2380 | 2198 | 2132 | 1684.6321 | 2060 | 2205.7405 | 1900.591 | 2066
| Mean |  | 2187.3333 | 2131.3167 | 2120.5 | 2073.8333 | 2062.697 | 1905 | 2060 | 2136.7627 | 2049.6667 | 2156.6667 | 2204.1667 | 2130 | 2126.6667 | 2061.6667 | 2038.3333 | 2296.6667 | 2207.1667 | 2093.5 | 2468.3333 | 2181.5 | 2152.8333 | 1710.544 | 2036.6667 | 2148.7784 | 1965.028 | 2060.8333
| Median |  | 2187 | 2137.7 | 2122 | 2075.5 | 2067.2542 | 1905 | 2055 | 2136.44 | 2051.5 | 2160 | 2202.5 | 2135 | 2125 | 2050 | 2040 | 2295 | 2208 | 1990 | 2450 | 2187.5 | 2142.5 | 1710.2588 | 2045 | 2158.5102 | 1970.826 | 2067
| Std Dev. |  | 11.894 | 45.8084 | 21.5662 | 36.7881 | 21.1666 | 37.2827 | 43.8178 | 16.4015 | 21.5004 | 35.0238 | 20.8367 | 12.6491 | 29.4392 | 23.1661 | 16.0208 | 36.1478 | 34.0495 | 196.9739 | 62.1021 | 24.0396 | 38.4261 | 19.6968 | 35.5903 | 56.7398 | 63.3679 | 32.9145
| Rel.Std.Dev. |  | 0.5438% | 2.149% | 1.017% | 1.774% | 1.026% | 1.957% | 2.127% | 0.7676% | 1.049% | 1.624% | 0.9453% | 0.5939% | 1.384% | 1.124% | 0.786% | 1.574% | 1.543% | 9.409% | 2.516% | 1.102% | 1.785% | 1.151% | 1.747% | 2.641% | 3.225% | 1.597%
| PDM3 |  | 3.862% | 1.202% | 0.6885% | -1.527% | -2.056% | -9.544% | -2.184% | 1.461% | -2.675% | 2.406% | 4.661% | 1.14% | 0.9814% | -2.105% | -3.213% | 9.054% | 4.804% | -0.5935% | 17.2% | 3.585% | 2.224% | -18.78% | -3.292% | 2.031% | -6.694% | -2.145%
| Z-Score (Absolute) |  | 0.69 | 0.08 | 0.03 | 0.54 | 0.66 | 2.36 | 0.69 | 0.14 | 0.8 | 0.36 | 0.87 | 0.07 | 0.03 | 0.67 | 0.92 | 1.87 | 0.9 | 0.33 | 3.73 | 0.63 | 0.32 | 4.47 | 0.94 | 0.27 | 1.72 | 0.68
| Analytical results for Zr in OREAS 600c (Certified Value 259 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 251 | 253.1 | 238.5 | 255.4 | 253.999 | 263 | 268 | 219.336 | 262 | 260 | 268.3 | 263 | 251 | 268 | 257 | 262 | 249.4 | 255 | 273 | 294.4 | 257.5 | 272.9006 | 305 | 244.9874 | 213.22 | 284 | 1
| 1 | 2 | 250 | 252 | 241.1 | 251.7 | 258.3268 | 260 | 266 | 215.137 | 270 | 253 | 257.5 | 263 | 246 | 269 | 250 | 262 | 253.6 | 256 | 269 | 290.6 | 246.1 | 267.4547 | 308 | 250.9756 | 219.89 | 287 | 18
| 1 | 3 | 254 | 254.6 | 240.6 | 255.3 | 258.1966 | 262 | 258 | 216.049 | 270 | 254 | 276.8 | 264 | 248 | 267 | 250 | 263 | 237.6 | 254 | 279 | 299.1 | 255.5 | 263.9598 | 294 | 251.9396 | 220.15 | 279 | 16
| 1 | 4 | 251 | 252.1 | 242.9 | 257.7 | 259.7005 | 265 | 259 | 211.755 | 263 | 260 | 267.5 | 261 | 246 | 268 | 252 | 271 | 257.1 | 248 | 279 | 289 | 254.6 | 274.1072 | 298 | 248.1878 | 218.71 | 281 | 259.0134
| 1 | 5 | 261 | 259 | 243.7 | 260.2 | 258.1898 | 259 | 255 | 219.0985 | 255 | 261 | 264.3 | 259 | 252 | 265 | 256 | 259 | 246.8 | 248 | 273 | 299.3 | 246.3 | 265.1556 | 300 | 245.0402 | 220.35 | 286 | 72
| 1 | 6 | 255 | 244 | 244.5 | 265.1 | 254.9373 | 263 | 265 | 212.116 | 267 | 265 | 267.5 | 260 | 241 | 271 | 250 | 263 | 252.8 | 259 | 280 | 298.2 | 247.5 | 320.1759 | 300 | 247.9962 | 222.8 | 284
| Mean |  | 253.6667 | 252.4667 | 241.8833 | 257.5667 | 257.225 | 262 | 261.8333 | 215.5819 | 264.5 | 258.8333 | 266.9833 | 261.6667 | 247.3333 | 268 | 252.5 | 263.3333 | 249.55 | 253.3333 | 275.5 | 295.1 | 251.25 | 277.2923 | 300.8333 | 248.1878 | 219.1867 | 283.5
| Median |  | 252.5 | 252.6 | 242 | 256.55 | 258.1932 | 262.5 | 262 | 215.593 | 265 | 260 | 267.5 | 262 | 247 | 268 | 251 | 262.5 | 251.1 | 254.5 | 276 | 296.3 | 251.05 | 270.1777 | 300 | 248.092 | 220.02 | 284
| Std Dev. |  | 4.0825 | 4.8923 | 2.2311 | 4.6466 | 2.2297 | 2.1909 | 5.1929 | 3.2721 | 5.7533 | 4.535 | 6.2573 | 1.9664 | 3.9833 | 2 | 3.2094 | 4.0332 | 6.8462 | 4.4572 | 4.4609 | 4.4989 | 5.1659 | 21.4022 | 4.9967 | 2.8995 | 3.2144 | 3.0166
| Rel.Std.Dev. |  | 1.609% | 1.938% | 0.9224% | 1.804% | 0.8668% | 0.8362% | 1.983% | 1.518% | 2.175% | 1.752% | 2.344% | 0.7515% | 1.61% | 0.7463% | 1.271% | 1.532% | 2.743% | 1.759% | 1.619% | 1.525% | 2.056% | 7.718% | 1.661% | 1.168% | 1.466% | 1.064%
| PDM3 |  | -2.064% | -2.528% | -6.614% | -0.5586% | -0.6905% | 1.153% | 1.089% | -16.77% | 2.118% | -0.06952% | 3.077% | 1.024% | -4.509% | 3.47% | -2.515% | 1.668% | -3.654% | -2.193% | 6.365% | 13.93% | -2.997% | 7.057% | 16.15% | -4.18% | -15.38% | 9.454%
| Z-Score (Absolute) |  | 0.39 | 0.5 | 1.41 | 0.05 | 0.08 | 0.33 | 0.31 | 3.68 | 0.54 | 0.05 | 0.76 | 0.3 | 0.94 | 0.85 | 0.49 | 0.44 | 0.75 | 0.42 | 1.5 | 3.19 | 0.6 | 1.65 | 3.69 | 0.87 | 3.37 | 2.19
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C973: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C991: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1047: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1065: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1083: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1102: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 600c (Certified Value 23.4 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 2
| 1 | 1 | 24.14 | 22.98 | 23.7 | 22.52 | 22.8655 | 23.3 | 23.4 | 23.276 | 23.4 | 22.5 | 22.3 | 23 | 24.8 | 24.5 | 26.2 | 26.3 | 22.3 | 22.5 | 23 | 24.94 | 22.78 | 21.9114 | 7.62 | 23.5133 | 21.51 | 25.75 | 1
| 1 | 2 | 24.42 | 24.42 | 23.8 | 22.57 | 23.4013 | 23.8 | 23.4 | 24.742 | 23.6 | 22.8 | 21.5 | 23.3 | 24.3 | 23.8 | 25.2 | 26.5 | 22.1 | 22.5 | 24.3 | 26.25 | 23.87 | 21.7018 | 3.95 | 23.8441 | 21.38 | 26.89 | 22
| 1 | 3 | 24.02 | 22.64 | 23.5 | 22.77 | 23.6866 | 23.3 | 23.7 | 23.3705 | 23.6 | 23 | 21.4 | 22.5 | 24.5 | 24.1 | 23.6 | 25 | 22.5 | 22.6 | 23.4 | 26.74 | 23.89 | 21.6085 | 1.75 | 23.6048 | 21.47 | 27.37 | 16
| 1 | 4 | 24.05 | 23.46 | 23.6 | 22.97 | 23.2146 | 23.8 | 23.6 | 25.1975 | 23.7 | 22.4 | 22 | 22.6 | 24.1 | 23.5 | 25.1 | 25.3 | 22.5 | 23.2 | 22.8 | 26.21 | 23.68 | 21.5822 | 8.98 | 23.8318 | 20.99 | 28.4 | 23.4009
| 1 | 5 | 24.24 | 24.4 | 23.7 | 23.14 | 23.1185 | 23.6 | 23.7 | 26.195 | 23.9 | 22.8 | 22.3 | 22.6 | 24.8 | 24.6 | 25.7 | 26.2 | 22.1 | 22.7 | 23.1 | 26.26 | 23.18 | 21.7835 | 9.87 | 23.6756 | 21.52 | 27.54 | 74
| 1 | 6 | 24.1 | 24.04 | 23.6 | 23.02 | 23.3978 | 22.9 | 23.7 | 22.583 | 23.6 | 22.8 | 21.7 | 23.9 | 24.2 | 24.9 | 26.3 | 25.9 | 22.6 | 22.7 | 22.9 | 25.71 | 23.05 | 21.8683 | 9.52 | 23.5914 | 21.3 | 25.74
| Mean |  | 24.1617 | 23.6567 | 23.65 | 22.8317 | 23.2807 | 23.45 | 23.5833 | 24.2273 | 23.6333 | 22.7167 | 21.8667 | 22.9833 | 24.45 | 24.2333 | 25.35 | 25.8667 | 22.35 | 22.7 | 23.25 | 26.0183 | 23.4083 | 21.7426 | 6.9483 | 23.6768 | 21.3617 | 26.9483
| Median |  | 24.12 | 23.75 | 23.65 | 22.87 | 23.3062 | 23.45 | 23.65 | 24.0562 | 23.6 | 22.8 | 21.85 | 22.8 | 24.4 | 24.3 | 25.45 | 26.05 | 22.4 | 22.65 | 23.05 | 26.23 | 23.43 | 21.7427 | 8.3 | 23.6402 | 21.425 | 27.13
| Std Dev. |  | 0.14811 | 0.74998 | 0.10488 | 0.25262 | 0.28146 | 0.35071 | 0.1472 | 1.3727 | 0.1633 | 0.22286 | 0.39328 | 0.54191 | 0.30166 | 0.52789 | 0.98944 | 0.59554 | 0.21679 | 0.26077 | 0.55408 | 0.62088 | 0.46782 | 0.13516 | 3.3389 | 0.13507 | 0.20054 | 1.0521
| Rel.Std.Dev. |  | 0.613% | 3.17% | 0.4435% | 1.106% | 1.209% | 1.496% | 0.6242% | 5.666% | 0.691% | 0.981% | 1.799% | 2.358% | 1.234% | 2.178% | 3.903% | 2.302% | 0.97% | 1.149% | 2.383% | 2.386% | 1.999% | 0.6216% | 48.05% | 0.5705% | 0.9388% | 3.904%
| PDM3 |  | 3.251% | 1.093% | 1.064% | -2.433% | -0.5136% | 0.2098% | 0.7796% | 3.532% | 0.9932% | -2.924% | -6.556% | -1.784% | 4.483% | 3.557% | 8.329% | 10.54% | -4.491% | -2.995% | -0.6449% | 11.19% | 0.03174% | -7.086% | -70.31% | 1.179% | -8.714% | 15.16%
| Z-Score (Absolute) |  | 0.61 | 0.13 | 0.13 | 0.65 | 0.22 | 0.06 | 0.06 | 0.67 | 0.11 | 0.76 | 1.56 | 0.5 | 0.88 | 0.68 | 1.73 | 2.22 | 1.1 | 0.77 | 0.25 | 2.36 | 0.1 | 1.68 | 15.65 | 0.15 | 2.04 | 3.24
| Analytical results for Al in OREAS 600c (Certified Value 0.713 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.76 | 0.65 | 0.73 | 0.7083 | 0.67321 | 0.7 | 0.54 | 0.69 | 0.897 | 0.9148 | 0.67 | 0.71 | 0.7 | 0.81 | 0.71 | 0.749 | 0.74 | 0.63 | 0.67 | 0.8119 | 0.63 | 0.66836 | 0.80381 |  |  |  | 1
| 1 | 2 | 0.73 | 0.99 | 0.74 | 0.7224 | 0.69573 | 0.67 | 0.54 | 0.7 | 0.925 | 0.8884 | 0.68 | 0.7 | 0.7 | 0.77 | 0.73 | 0.759 | 0.7 | 0.66 | 0.72 | 0.7747 | 0.63 | 0.68337 | 0.76836 |  |  |  | 0
| 1 | 3 | 0.74 | 0.64 | 0.73 | 0.7174 | 0.68208 | 0.67 | 0.55 | 0.69 | 0.896 | 0.8787 | 0.68 | 0.75 | 0.69 | 0.83 | 0.68 | 0.734 | 0.75 | 0.61 | 0.74 | 0.7307 | 0.625 | 0.6843 | 0.82557 |  |  |  | 16
| 1 | 4 | 0.72 | 0.69 | 0.74 | 0.6919 | 0.68695 | 0.74 | 0.54 | 0.69 | 0.906 | 0.9125 | 0.66 | 0.76 | 0.7 | 0.79 | 0.71 | 0.761 | 0.71 | 0.63 | 0.76 | 0.8615 | 0.635 | 0.67025 | 0.74352 |  |  |  | 0.71256
| 1 | 5 | 0.74 | 0.74 | 0.74 | 0.7324 | 0.67644 | 0.69 | 0.54 | 0.71 | 0.92 | 0.9124 | 0.66 | 0.72 | 0.71 | 0.82 | 0.73 | 0.773 | 0.737 | 0.63 | 0.73 | 0.768 | 0.619 | 0.6816 | 0.77578 |  |  |  | 75
| 1 | 6 | 0.72 | 0.96 | 0.73 | 0.7069 | 0.68013 | 0.7 | 0.55 | 0.69 | 0.907 | 0.8855 | 0.68 | 0.74 | 0.7 | 0.81 | 0.71 | 0.762 | 0.72 | 0.63 | 0.67 | 0.78 | 0.635 | 0.66346 | 0.77349
| Mean |  | 0.735 | 0.77833 | 0.735 | 0.71322 | 0.68242 | 0.695 | 0.54333 | 0.695 | 0.9085 | 0.89872 | 0.67167 | 0.73 | 0.7 | 0.805 | 0.71167 | 0.75633 | 0.72617 | 0.63167 | 0.715 | 0.7878 | 0.629 | 0.67522 | 0.78175
| Median |  | 0.735 | 0.715 | 0.735 | 0.71285 | 0.6811 | 0.695 | 0.54 | 0.69 | 0.9065 | 0.9004 | 0.675 | 0.73 | 0.7 | 0.81 | 0.71 | 0.76 | 0.7285 | 0.63 | 0.725 | 0.77735 | 0.63 | 0.67593 | 0.77463
| Std Dev. |  | 0.015166 | 0.15664 | 0.0054772 | 0.014059 | 0.0080443 | 0.025884 | 0.005164 | 0.0083666 | 0.011845 | 0.016234 | 0.0098319 | 0.023664 | 0.0063246 | 0.021679 | 0.018348 | 0.013352 | 0.019292 | 0.016021 | 0.037283 | 0.044491 | 0.0061644 | 0.0089416 | 0.028813
| Rel.Std.Dev. |  | 2.063% | 20.13% | 0.7452% | 1.971% | 1.179% | 3.724% | 0.9504% | 1.204% | 1.304% | 1.806% | 1.464% | 3.242% | 0.9035% | 2.693% | 2.578% | 1.765% | 2.657% | 2.536% | 5.214% | 5.648% | 0.98% | 1.324% | 3.686%
| PDM3 |  | 3.149% | 9.231% | 3.149% | 0.09242% | -4.229% | -2.464% | -23.75% | -2.464% | 27.5% | 26.13% | -5.739% | 2.448% | -1.762% | 12.97% | -0.1251% | 6.143% | 1.91% | -11.35% | 0.3427% | 10.56% | -11.73% | -5.24% | 9.711%
| Z-Score (Absolute) |  | 0.34 | 1.07 | 0.34 | 0.03 | 0.55 | 0.34 | 2.91 | 0.34 | 3.28 | 3.11 | 0.73 | 0.25 | 0.25 | 1.53 | 0.06 | 0.7 | 0.19 | 1.41 | 0 | 1.23 | 1.46 | 0.67 | 1.13
| Analytical results for As in OREAS 600c (Certified Value 36.7 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 2
| 1 | 1 | 43.4 | 36.3 | 36.4 | 37 | 36.7638 | 36.3 | 35 | 35.141 | 37.1 | 37 | 36 | 36.2 | 41.6 | 37.9 | 38.8 | 40.5 | 29.6 | 29 | 36.5 | 34.2 | 36.3 | 40.8376 | 34 | 29.9782 | 31.205 | 35.2 | 1
| 1 | 2 | 43.6 | 37.9 | 37 | 37 | 38.0125 | 36.5 | 34.8 | 35.687 | 35.6 | 37 | 34 | 36.7 | 41.1 | 37.5 | 37.1 | 40.3 | 28.9 | 29 | 39.1 | 33.7 | 37.6 | 48.8675 | 33 | 31.7079 | 27.661 | 35.6 | 12
| 1 | 3 | 43.4 | 35.6 | 36.4 | 37 | 38.5034 | 36.4 | 35.4 | 36.128 | 36.5 | 37 | 36 | 35.7 | 41.1 | 37.6 | 35.8 | 37.9 | 29.3 | 29 | 36.2 | 34.6 | 38.2 | 39.8258 | 32 | 29.8769 | 32.202 | 39 | 16
| 1 | 4 | 43.6 | 37 | 36.5 | 38 | 37.6433 | 36.7 | 34.8 | 35.8655 | 35.3 | 37 | 35 | 36.3 | 39.7 | 37.5 | 37 | 38.2 | 30.2 | 31 | 38.2 | 34.3 | 37.6 | 40.2308 | 33 | 29.7137 | 32.254 | 39.5 | 36.7246
| 1 | 5 | 43.1 | 37.8 | 36.5 | 38 | 37.4485 | 37.1 | 35.5 | 35.4245 | 34.5 | 38 | 36 | 35.3 | 39.5 | 38.1 | 38.1 | 40.3 | 29.1 | 29 | 37.1 | 34 | 36.8 | 48.9182 | 32 | 30.6799 | 27.65 | 37.6 | 76
| 1 | 6 | 41.2 | 37.1 | 36.2 | 37 | 37.6401 | 35.1 | 35.4 | 37.304 | 34.8 | 38 | 34 | 38.9 | 39.7 | 38.5 | 38.1 | 38.9 | 29.8 | 29 | 37.7 | 34.5 | 38.1 | 45.1393 | 34 | 31.7861 | 32.506 | 36.4
| Mean |  | 43.05 | 36.95 | 36.5 | 37.3333 | 37.6686 | 36.35 | 35.15 | 35.925 | 35.6333 | 37.3333 | 35.1667 | 36.5167 | 40.45 | 37.85 | 37.4833 | 39.35 | 29.4833 | 29.3333 | 37.4667 | 34.2167 | 37.4333 | 43.9699 | 33 | 30.6238 | 30.5797 | 37.2167
| Median |  | 43.4 | 37.05 | 36.45 | 37 | 37.6417 | 36.45 | 35.2 | 35.7762 | 35.45 | 37 | 35.5 | 36.25 | 40.4 | 37.75 | 37.6 | 39.6 | 29.45 | 29 | 37.4 | 34.25 | 37.6 | 42.9885 | 33 | 30.329 | 31.7035 | 37
| Std Dev. |  | 0.92466 | 0.88261 | 0.26833 | 0.5164 | 0.58029 | 0.67454 | 0.32094 | 0.75739 | 0.99933 | 0.5164 | 0.98319 | 1.2656 | 0.91597 | 0.39875 | 1.0685 | 1.1623 | 0.47924 | 0.8165 | 1.0893 | 0.33116 | 0.74476 | 4.2616 | 0.89443 | 0.93083 | 2.3082 | 1.7826
| Rel.Std.Dev. |  | 2.148% | 2.389% | 0.7351% | 1.383% | 1.541% | 1.856% | 0.913% | 2.108% | 2.804% | 1.383% | 2.796% | 3.466% | 2.264% | 1.053% | 2.851% | 2.954% | 1.625% | 2.784% | 2.907% | 0.9678% | 1.99% | 9.692% | 2.71% | 3.04% | 7.548% | 4.79%
| PDM3 |  | 17.22% | 0.6139% | -0.6115% | 1.658% | 2.571% | -1.02% | -4.287% | -2.177% | -2.971% | 1.658% | -4.242% | -0.5661% | 10.14% | 3.065% | 2.066% | 7.149% | -19.72% | -20.13% | 2.021% | -6.829% | 1.93% | 19.73% | -10.14% | -16.61% | -16.73% | 1.34%
| Z-Score (Absolute) |  | 3.84 | 0.13 | 0.14 | 0.37 | 0.57 | 0.23 | 0.96 | 0.49 | 0.67 | 0.37 | 0.95 | 0.13 | 2.26 | 0.68 | 0.46 | 1.59 | 4.41 | 4.5 | 0.45 | 1.53 | 0.43 | 4.4 | 2.27 | 3.72 | 3.74 | 0.29
| Analytical results for B in OREAS 600c (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab; 2.15: Lab | 1
| No | No | 2.01: 2.01; 2.03: 2.03; 2.04: 2.04; 2.05: 2.06; 2.06: 2.07; 2.07: 2.08; 2.08: 2.12; 2.09: 2.13; 2.10: 2.14; 2.11: 2.15; 2.12: 2.16; 2.13: 2.17; 2.14: 2.23; 2.15: 2.26 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*OES; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*OES; 2.15: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 15g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 01g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.5g; 2.14: 0.5g; 2.15: 01g | 1
| 1 | 1 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 3; 2.07: 1.8; 2.08: 11; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 3; 2.07: 2.2; 2.08: 11; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 3
| 1 | 3 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 3; 2.07: 1.9; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 3; 2.07: 2.1; 2.08: 12; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | < 10
| 1 | 5 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 2.4; 2.08: 13; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10 | 77
| 1 | 6 | 2.01: < 10; 2.03: < 20; 2.04: 2; 2.05: < 2; 2.06: 2; 2.07: 2; 2.08: 11; 2.09: < 10; 2.10: < 10; 2.11: < 10; 2.12: < 10; 2.13: < 10; 2.14: < 0.1; 2.15: < 10
| Mean |  | 2.04: 2; 2.06: 2.6667; 2.07: 2.0667; 2.08: 11.6
| Median |  | 2.04: 2; 2.06: 3; 2.07: 2.05; 2.08: 11
| Std Dev. |  | 2.04: 0; 2.06: 0.5164; 2.07: 0.21602; 2.08: 0.89443
| Rel.Std.Dev. |  | 2.04: 0%; 2.06: 19.36%; 2.07: 10.45%; 2.08: 7.711%
PDM3
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 2.89; 2.04: 1.73; 2.05: 2.31; 2.06: 1.35; 2.07: 1.7; 2.08: 3.81; 2.09: 0; 2.10: 0; 2.11: 0; 2.12: 0; 2.13: 0; 2.14: 2.86; 2.15: 0
| Analytical results for Ba in OREAS 600c (Certified Value 201 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 0
| 1 | 1 | 225 | 195.3 | 197 | 196 | 203.1196 | 171 | 194 | 206 | 238 | 200 | 200 | 200 | 230 | 210 | 168 | 160 | 217.2 | 215 | 297.5868 | 148 | 171.4152 | 203.97 | 237 |  |  |  | 1
| 1 | 2 | 205 | 315.5 | 203 | 199 | 209.8653 | 183 | 185 | 206 | 230 | 200 | 200 | 200 | 220 | 210 | 167 | 170 | 206.4 | 226 | 281.7626 | 152 | 184.5772 | 202.052 | 252 |  |  |  | 24
| 1 | 3 | 224 | 195.5 | 201 | 201 | 203.4135 | 174 | 198 | 204 | 234 | 200 | 220 | 200 | 260 | 200 | 165 | 169 | 199.8 | 225 | 254.4989 | 150 | 181.0409 | 200.308 | 251 |  |  |  | 16
| 1 | 4 | 207 | 204.9 | 207 | 198 | 205.9044 | 160 | 195 | 203 | 233 | 200 | 220 | 200 | 230 | 210 | 173 | 160 | 199.8 | 221 | 320.7285 | 143 | 176.2452 | 196.239 | 261 |  |  |  | 200.8378
| 1 | 5 | 217 | 213.1 | 202 | 204 | 203.8506 | 190 | 192 | 205 | 240 | 200 | 200 | 200 | 230 | 210 | 164 | 159 | 199 | 218 | 264.775 | 150 | 194.3758 | 200.11 | 257 |  |  |  | 78
| 1 | 6 | 206 | 306.7 | 204 | 197 | 206.5594 | 181 | 195 | 208 | 232 | 200 | 220 | 200 | 230 | 210 | 169 | 165 | 199.3 | 214 | 285.7619 | 148 | 191.0192 | 202.306 | 215
| Mean |  | 214 | 238.5 | 202.3333 | 199.1667 | 205.4521 | 176.5 | 193.1667 | 205.3333 | 234.5 | 200 | 210 | 200 | 233.3333 | 208.3333 | 167.6667 | 163.8333 | 203.5833 | 219.8333 | 284.1856 | 148.5 | 183.1122 | 200.8308 | 245.5
| Median |  | 212 | 209 | 202.5 | 198.5 | 204.8775 | 177.5 | 194.5 | 205.5 | 233.5 | 200 | 210 | 200 | 230 | 210 | 167.5 | 162.5 | 199.8 | 219.5 | 283.7622 | 149 | 182.8091 | 201.18 | 251.5
| Std Dev. |  | 9.2087 | 56.6914 | 3.3267 | 2.9269 | 2.5717 | 10.5214 | 4.446 | 1.7512 | 3.7815 | 0 | 10.9545 | 0 | 13.6626 | 4.0825 | 3.2042 | 4.8751 | 7.2295 | 5.0365 | 23.5856 | 3.0822 | 8.7155 | 2.6608 | 17.0147
| Rel.Std.Dev. |  | 4.303% | 23.77% | 1.644% | 1.47% | 1.252% | 5.961% | 2.302% | 0.8529% | 1.613% | 0% | 5.216% | 0% | 5.855% | 1.96% | 1.911% | 2.976% | 3.551% | 2.291% | 8.299% | 2.076% | 4.76% | 1.325% | 6.931%
| PDM3 |  | 6.554% | 18.75% | 0.7447% | -0.8321% | 2.298% | -12.12% | -3.82% | 2.238% | 16.76% | -0.4171% | 4.562% | -0.4171% | 16.18% | 3.732% | -16.52% | -18.43% | 1.367% | 9.458% | 41.5% | -26.06% | -8.826% | -0.003449% | 22.24%
| Z-Score (Absolute) |  | 0.67 | 2.26 | 0.08 | 0.29 | 0.12 | 1.75 | 0.67 | 0.11 | 2 | 0.23 | 0.42 | 0.23 | 1.93 | 0.31 | 2.33 | 2.57 | 0 | 1.05 | 5.22 | 3.57 | 1.33 | 0.18 | 2.71
| Analytical results for Be in OREAS 600c (Certified Value 0.59 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 3
| 1 | 1 | 0.74 | 0.8 | 0.59 | 0.59427 | 0.6 | 0.57 | 0.6305 | 0.7 | 0.65 | 0.56 | 0.6 | 0.56 | 0.64 | 0.59 | 0.49 | 0.52 | < 2 | 0.5 | 0.9182 | 0.58 | 0.58434 | < 5 | 0.58 |  |  |  | 1
| 1 | 2 | 0.74 | 0.6 | 0.61 | 0.60855 | 0.6 | 0.57 | 0.5856 | 0.7 | 0.67 | 0.56 | 0.6 | 0.56 | 0.62 | 0.6 | 0.5 | 0.81 | < 2 | 0.52 | 0.9169 | 0.52 | 0.60221 | < 5 | 0.63 |  |  |  | 25
| 1 | 3 | 0.8 | 0.7 | 0.61 | 0.60602 | 0.6 | 0.56 | 0.668 | 0.7 | 0.68 | 0.56 | 0.62 | 0.56 | 0.64 | 0.56 | 0.5 | 0.82 | < 2 | 0.53 | 0.8578 | 0.55 | 0.62184 | < 5 | 0.61 |  |  |  | 16
| 1 | 4 | 0.77 | 0.8 | 0.6 | 0.60897 | 0.6 | 0.49 | 0.5841 | 0.7 | 0.65 | 0.56 | 0.63 | 0.57 | 0.63 | 0.58 | 0.51 | 0.69 | < 2 | 0.55 | 0.9669 | 0.57 | 0.58274 | < 5 | 0.66 |  |  |  | 0.58621
| 1 | 5 | 0.76 | 0.6 | 0.61 | 0.57287 | 0.6 | 0.63 | 0.5338 | 0.7 | 0.69 | 0.55 | 0.61 | 0.56 | 0.65 | 0.6 | 0.5 | 0.48 | < 2 | 0.49 | 0.9595 | 0.57 | 0.60485 | < 5 | 0.63 |  |  |  | 79
| 1 | 6 | 0.75 | 0.7 | 0.61 | 0.60682 | 0.6 | 0.54 | 0.6744 | 0.7 | 0.66 | 0.56 | 0.62 | 0.56 | 0.63 | 0.59 | 0.51 | 0.67 | < 2 | 0.46 | 0.9164 | 0.62 | 0.56433 | < 5 | 0.57
| Mean |  | 0.76 | 0.7 | 0.605 | 0.59958 | 0.6 | 0.56 | 0.61273 | 0.7 | 0.66667 | 0.55833 | 0.61333 | 0.56167 | 0.635 | 0.58667 | 0.50167 | 0.665 |  | 0.50833 | 0.92262 | 0.56833 | 0.59339 |  | 0.61333
| Median |  | 0.755 | 0.7 | 0.61 | 0.60642 | 0.6 | 0.565 | 0.60805 | 0.7 | 0.665 | 0.56 | 0.615 | 0.56 | 0.635 | 0.59 | 0.5 | 0.68 |  | 0.51 | 0.91755 | 0.57 | 0.59328 |  | 0.62
| Std Dev. |  | 0.022804 | 0.089443 | 0.0083666 | 0.01417 | 0 | 0.045607 | 0.054698 | 1.2162e-16 | 0.01633 | 0.0040825 | 0.012111 | 0.0040825 | 0.010488 | 0.015055 | 0.0075277 | 0.14209 |  | 0.031885 | 0.039022 | 0.033116 | 0.02029 |  | 0.033862
| Rel.Std.Dev. |  | 3% | 12.78% | 1.383% | 2.363% | 0% | 8.144% | 8.927% | 1.737e-14% | 2.449% | 0.7312% | 1.975% | 0.7269% | 1.652% | 2.566% | 1.501% | 21.37% |  | 6.273% | 4.229% | 5.827% | 3.419% |  | 5.521%
| PDM3 |  | 29.65% | 19.41% | 3.205% | 2.281% | 2.352% | -4.471% | 4.524% | 19.41% | 13.72% | -4.756% | 4.627% | -4.187% | 8.323% | 0.07769% | -14.42% | 13.44% |  | -13.28% | 57.39% | -3.05% | 1.224% |  | 4.627%
| Z-Score (Absolute) |  | 2.12 | NA | 0.05 | 0.13 | NA | 0.69 | 0.05 | NA | 0.81 | 0.71 | 0.06 | 0.66 | 0.37 | 0.31 | 1.5 | 0.79 | 5.49 | 1.41 | 4.4 | 0.57 | 0.22 | 26.55 | 0.06
Results from laboratories 2.03, 2.07 and 2.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 600c (Certified Value 0.78 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 3
| 1 | 1 | 0.71 | 0.76 | 0.8 | 0.82 | 0.7393 | 0.83 | 0.82 | 0.78 | 0.74 | 0.82 | 0.81 | 0.73 | 0.83 | 0.82 | 0.75 | 0.98 | 0.66 | < 2 | < 5 | 0.71 |  | 0.83 | 0.77962 | < 5 | 0.71 |  | 1
| 1 | 2 | 0.64 | 0.78 | 0.8 | 0.8 | 0.78303 | 0.81 | 0.79 | 0.83 | 0.72 | 0.81 | 0.8 | 0.74 | 0.85 | 0.81 | 0.77 | < 0.02 | 0.64 | 2 | < 5 | 0.76 | 13.9039 | 0.84 | 0.80933 | < 5 | 0.73 |  | 26
| 1 | 3 | 0.72 | 0.77 | 0.8 | 0.84 | 0.80159 | 0.83 | 0.82 | 0.79 | 0.72 | 0.82 | 0.79 | 0.71 | 0.82 | 0.81 | 0.71 | 0.98 | 0.64 | < 2 | < 5 | 0.74 | 11.5157 | 0.92 | 0.81979 | < 5 | 0.77 |  | 16
| 1 | 4 | 0.65 | 0.83 | 0.8 | 0.84 | 0.82153 | 0.84 | 0.84 | 0.77 | 0.71 | 0.82 | 0.8 | 0.68 | 0.83 | 0.82 | 0.72 | 0.98 | 0.67 | < 2 | < 5 | 0.74 | 11.3549 | 0.86 | 0.76654 | < 5 | 0.79 |  | 0.78201
| 1 | 5 | 0.68 | 0.73 | 0.8 | 0.83 | 0.77392 | 0.82 | 0.83 | 0.8 | 0.71 | 0.85 | 0.78 | 0.71 | 0.81 | 0.85 | 0.76 | < 0.02 | 0.65 | < 2 | < 5 | 0.75 | 13.1263 | 0.83 | 0.80784 | < 5 | 0.73 |  | 80
| 1 | 6 | 0.66 | 0.76 | 0.8 | 0.82 | 0.73725 | 0.8 | 0.81 | 0.81 | 0.74 | 0.81 | 0.82 | 0.67 | 0.82 | 0.85 | 0.74 | 0.98 | 0.64 | < 2 | < 5 | 0.78 | 8.445 | 0.94 | 0.77097 | < 5 | 0.71
| Mean |  | 0.67667 | 0.77167 | 0.8 | 0.825 | 0.7761 | 0.82167 | 0.81833 | 0.79667 | 0.72333 | 0.82167 | 0.8 | 0.70667 | 0.82667 | 0.82667 | 0.74167 | 0.98 | 0.65 | 2 |  | 0.74667 | 11.6692 | 0.87 | 0.79235 |  | 0.74
| Median |  | 0.67 | 0.765 | 0.8 | 0.825 | 0.77848 | 0.825 | 0.82 | 0.795 | 0.72 | 0.82 | 0.8 | 0.71 | 0.825 | 0.82 | 0.745 | 0.98 | 0.645 | 2 |  | 0.745 | 11.5157 | 0.85 | 0.79373 |  | 0.73
| Std Dev. |  | 0.03266 | 0.033116 | 1.2162e-16 | 0.015166 | 0.033558 | 0.01472 | 0.017224 | 0.021602 | 0.013663 | 0.01472 | 0.014142 | 0.027325 | 0.013663 | 0.018619 | 0.023166 | 0 | 0.012649 |  |  | 0.023381 | 2.0997 | 0.048166 | 0.022657 |  | 0.032863
| Rel.Std.Dev. |  | 4.827% | 4.291% | 1.52e-14% | 1.838% | 4.324% | 1.791% | 2.105% | 2.712% | 1.889% | 1.791% | 1.768% | 3.867% | 1.653% | 2.252% | 3.124% | 0% | 1.946% |  |  | 3.131% | 17.99% | 5.536% | 2.859% |  | 4.441%
| PDM3 |  | -13.47% | -1.322% | 2.301% | 5.498% | -0.7549% | 5.072% | 4.645% | 1.875% | -7.503% | 5.072% | 2.301% | -9.634% | 5.711% | 5.711% | -5.159% | 25.32% | -16.88% | 155.8% |  | -4.519% | 1392% | 11.25% | 1.322% |  | -5.372%
| Z-Score (Absolute) |  | 1.51 | 0.33 | NA | 0.33 | 0.28 | 0.29 | 0.25 | 0.02 | 0.93 | 0.29 | 0.02 | 1.14 | 0.35 | 0.35 | 0.71 | 1.76 | 1.85 | 4.59 | 21.18 | 0.64 | 135.33 | 0.89 | 0.07 | 21.18 | 0.73
Results from laboratory 2.04 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 600c (Certified Value 0.794 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.81 | 0.78 | 0.77 | 0.76 | 0.77349 | 0.79 | 0.75 | 0.75 | 0.84 | 0.86 | 0.79 | 0.81 | 0.79 | 0.83 | 0.83 | 0.795 | 0.802 | 0.85 | 0.74 | 0.81 | 0.7259 | 0.8 | 0.77011 | 0.83597 |  |  | 1
| 1 | 2 | 0.8 | 0.77 | 0.78 | 0.78 | 0.79313 | 0.8 | 0.75 | 0.77 | 0.87 | 0.84 | 0.79 | 0.8 | 0.8 | 0.81 | 0.84 | 0.7922 | 0.799 | 0.9 | 0.75 | 0.81 | 0.7307 | 0.829 | 0.79631 | 0.81148 |  |  | 0
| 1 | 3 | 0.8 | 0.77 | 0.78 | 0.77 | 0.79835 | 0.8 | 0.77 | 0.77 | 0.87 | 0.83 | 0.79 | 0.8 | 0.79 | 0.79 | 0.79 | 0.8046 | 0.799 | 0.865 | 0.74 | 0.84 | 0.742 | 0.8 | 0.78391 | 0.83859 |  |  | 16
| 1 | 4 | 0.81 | 0.8 | 0.78 | 0.74 | 0.77887 | 0.87 | 0.75 | 0.76 | 0.86 | 0.85 | 0.79 | 0.82 | 0.8 | 0.81 | 0.83 | 0.8111 | 0.823 | 0.868 | 0.75 | 0.8 | 0.7357 | 0.815 | 0.7902 | 0.77944 |  |  | 0.79395
| 1 | 5 | 0.8 | 0.77 | 0.78 | 0.79 | 0.78131 | 0.81 | 0.76 | 0.77 | 0.87 | 0.85 | 0.78 | 0.81 | 0.8 | 0.82 | 0.83 | 0.7888 | 0.808 | 0.889 | 0.74 | 0.8 | 0.7418 | 0.8 | 0.79271 | 0.81818 |  |  | 81
| 1 | 6 | 0.79 | 0.76 | 0.77 | 0.75 | 0.78264 | 0.77 | 0.77 | 0.77 | 0.86 | 0.84 | 0.8 | 0.81 | 0.79 | 0.82 | 0.83 | 0.8055 | 0.802 | 0.897 | 0.74 | 0.77 | 0.7395 | 0.815 | 0.77499 | 0.81424
| Mean |  | 0.80167 | 0.775 | 0.77667 | 0.765 | 0.78463 | 0.80667 | 0.75833 | 0.765 | 0.86167 | 0.845 | 0.79 | 0.80833 | 0.795 | 0.81333 | 0.825 | 0.79953 | 0.8055 | 0.87817 | 0.74333 | 0.805 | 0.73593 | 0.80983 | 0.7847 | 0.81632
| Median |  | 0.8 | 0.77 | 0.78 | 0.765 | 0.78198 | 0.8 | 0.755 | 0.77 | 0.865 | 0.845 | 0.79 | 0.81 | 0.795 | 0.815 | 0.83 | 0.7998 | 0.802 | 0.8785 | 0.74 | 0.805 | 0.7376 | 0.8075 | 0.78706 | 0.81621
| Std Dev. |  | 0.0075277 | 0.013784 | 0.005164 | 0.018708 | 0.0093028 | 0.033862 | 0.0098319 | 0.0083666 | 0.01169 | 0.010488 | 0.0063246 | 0.0075277 | 0.0054772 | 0.013663 | 0.017607 | 0.0087703 | 0.0091815 | 0.020094 | 0.005164 | 0.022583 | 0.0065123 | 0.011923 | 0.010363 | 0.021322
| Rel.Std.Dev. |  | 0.939% | 1.779% | 0.6649% | 2.446% | 1.186% | 4.198% | 1.297% | 1.094% | 1.357% | 1.241% | 0.8006% | 0.9313% | 0.689% | 1.68% | 2.134% | 1.097% | 1.14% | 2.288% | 0.6947% | 2.805% | 0.8849% | 1.472% | 1.321% | 2.612%
| PDM3 |  | 0.9714% | -2.387% | -2.177% | -3.647% | -1.174% | 1.601% | -4.486% | -3.647% | 8.529% | 6.429% | -0.498% | 1.811% | 0.1318% | 2.441% | 3.91% | 0.7028% | 1.454% | 10.61% | -6.376% | 1.391% | -7.308% | 2% | -1.165% | 2.817%
| Z-Score (Absolute) |  | 0.05 | 1.08 | 1.01 | 1.51 | 0.68 | 0.26 | 1.79 | 1.51 | 2.58 | 1.88 | 0.45 | 0.33 | 0.24 | 0.54 | 1.03 | 0.05 | 0.21 | 3.28 | 2.42 | 0.19 | 2.74 | 0.4 | 0.67 | 0.67
| Analytical results for Cd in OREAS 600c (Certified Value 6.14 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 6.4 | 6.64 | 6 | 6.03 | 6.0346 | 6.03 | 6.22 | 6.26 | 6.09 | 6.2 | 6.25 | 6.31 | 6.5 | 6.71 | 6.67 | 5.77 | 6.1 | 5.7 | 5.89 | 5.5826 | 6.38 | 5.8828 | 6.23 | 6.77 |  |  | 1
| 1 | 2 | 6.54 | 6.23 | 6 | 5.94 | 6.1937 | 6.22 | 6.36 | 6.31 | 6.02 | 6.1 | 6.22 | 6.17 | 6.41 | 6.44 | 6.78 | 5.67 | 6.1 | 5.7 | 6.12 | 5.6739 | 6.21 | 5.9627 | 6.113 | 7.03 |  |  | 27
| 1 | 3 | 6.18 | 6.07 | 5.9 | 6.01 | 6.2596 | 5.94 | 6.32 | 5.98 | 6.02 | 6 | 6.06 | 6.3 | 6.42 | 6.39 | 6.3 | 5.85 | 6.1 | 5.7 | 6.24 | 5.796 | 6.29 | 5.8351 | 6.201 | 7.48 |  |  | 16
| 1 | 4 | 6.25 | 7.1 | 6 | 6.29 | 6.1444 | 6.03 | 6.31 | 6.39 | 6.03 | 6.2 | 6.32 | 6.16 | 6.45 | 6.57 | 6.29 | 5.93 | 6.1 | 5.6 | 6.2 | 5.6493 | 6.54 | 5.767 | 6.095 | 7.45 |  |  | 6.1367
| 1 | 5 | 6.26 | 6.11 | 6 | 6.19 | 6.0729 | 6.1 | 6.43 | 6.12 | 5.92 | 6.3 | 6.26 | 6.18 | 6.54 | 6.62 | 6.6 | 5.69 | 6 | 5.6 | 6.1 | 5.354 | 6.26 | 5.9987 | 6.184 | 7.01 |  |  | 82
| 1 | 6 | 6.34 | 5.83 | 5.8 | 6.21 | 6.133 | 5.99 | 6.4 | 6.23 | 6.18 | 6.1 | 6.53 | 6.02 | 6.57 | 6.72 | 6.45 | 5.79 | 5.9 | 5.7 | 6.05 | 5.5751 | 6.49 | 5.7253 | 6.222 | 7.12
| Mean |  | 6.3283 | 6.33 | 5.95 | 6.1117 | 6.1397 | 6.0517 | 6.34 | 6.215 | 6.0433 | 6.15 | 6.2733 | 6.19 | 6.4817 | 6.575 | 6.515 | 5.7833 | 6.05 | 5.6667 | 6.1 | 5.6052 | 6.3617 | 5.8619 | 6.1742 | 7.1433
| Median |  | 6.3 | 6.17 | 6 | 6.11 | 6.1387 | 6.03 | 6.34 | 6.245 | 6.025 | 6.15 | 6.255 | 6.175 | 6.475 | 6.595 | 6.525 | 5.78 | 6.1 | 5.7 | 6.11 | 5.6159 | 6.335 | 5.859 | 6.1925 | 7.075
| Std Dev. |  | 0.12875 | 0.46152 | 0.083666 | 0.13717 | 0.081091 | 0.09786 | 0.074565 | 0.14571 | 0.08641 | 0.10488 | 0.15306 | 0.10658 | 0.065549 | 0.13693 | 0.20127 | 0.097707 | 0.083666 | 0.05164 | 0.12377 | 0.14669 | 0.13197 | 0.10742 | 0.056975 | 0.27493
| Rel.Std.Dev. |  | 2.035% | 7.291% | 1.406% | 2.244% | 1.321% | 1.617% | 1.176% | 2.344% | 1.43% | 1.705% | 2.44% | 1.722% | 1.011% | 2.083% | 3.089% | 1.689% | 1.383% | 0.9113% | 2.029% | 2.617% | 2.074% | 1.833% | 0.9228% | 3.849%
| PDM3 |  | 3.123% | 3.15% | -3.042% | -0.4076% | 0.04936% | -1.385% | 3.313% | 1.276% | -1.521% | 0.2171% | 2.227% | 0.8689% | 5.622% | 7.143% | 6.165% | -5.758% | -1.412% | -7.659% | -0.5977% | -8.662% | 3.666% | -4.477% | 0.6109% | 16.4%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 0.86 | 0.2 | 0.09 | 0.45 | 0.72 | 0.21 | 0.48 | 0.05 | 0.45 | 0.11 | 1.29 | 1.67 | 1.42 | 1.53 | 0.45 | 2 | 0.25 | 2.25 | 0.81 | 1.21 | 0.05 | 3.96
| Analytical results for Ce in OREAS 600c (Certified Value 57 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 0
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 1
| 1 | 1 | 62 | 55.4 | 57.812 | 48.7793 | 63.3 | 54.9 | 73.103 | 54.13 | 64.6 | 73.08 | 51.8 | 59.3 | 58.9 | 60.7 | 59.2 | 53.57 | 39.3 | 54.16 | 44.2541 | 36.7 | 56.974 | 61.612 | 56.22 |  |  |  | 1
| 1 | 2 | 59 | 56.3 | 57.502 | 51.1977 | 65.4 | 52.6 | 73.835 | 55.63 | 62.2 | 68.88 | 51.4 | 56.2 | 59 | 57.4 | 59.4 | 54.33 | 43.2 | 58.6 | 62.5738 | 36.8 | 55.4502 | 61.005 | 59.51 |  |  |  | 28
| 1 | 3 | 60 | 53 | 59.608 | 50.9807 | 63.5 | 56.5 | 73.49 | 55.13 | 64.9 | 70.63 | 51.4 | 56.1 | 57.8 | 58.3 | 54.7 | 51.51 | 42 | 60.03 | 64.2652 | 36.6 | 58.9181 | 61.532 | 58.58 |  |  |  | 16
| 1 | 4 | 58 | 60.6 | 58.804 | 49.6249 | 64.5 | 55.4 | 73.729 | 55.33 | 62.6 | 71.53 | 50.5 | 56.6 | 58.6 | 58.3 | 55.5 | 52.94 | 40.9 | 59.39 | 56.2403 | 36.5 | 56.4384 | 60.432 | 63.32 |  |  |  | 57.4104
| 1 | 5 | 59 | 51.1 | 60.266 | 49.8565 | 63.2 | 55.9 | 73.656 | 55.37 | 63.8 | 72.33 | 50.2 | 57.7 | 58.5 | 58.2 | 57.9 | 53.43 | 40.9 | 56.96 | 60.3584 | 36.6 | 58.4193 | 60.949 | 61.19 |  |  |  | 83
| 1 | 6 | 58 | 54.4 | 58.656 | 51.6091 | 61.2 | 56 | 73.913 | 55.17 | 65.3 | 69.36 | 53.5 | 55.2 | 58.1 | 60.4 | 56.3 | 53.96 | 42.1 | 54.06 | 53.0905 | 37.3 | 54.9049 | 61.259 | 51
| Mean |  | 59.3333 | 55.1333 | 58.7747 | 50.3414 | 63.5167 | 55.2167 | 73.621 | 55.1267 | 63.9 | 70.9683 | 51.4667 | 56.85 | 58.4833 | 58.8833 | 57.1667 | 53.29 | 41.4 | 57.2 | 56.797 | 36.75 | 56.8508 | 61.1315 | 58.3033
| Median |  | 59 | 54.9 | 58.73 | 50.4186 | 63.4 | 55.65 | 73.6925 | 55.25 | 64.2 | 71.08 | 51.4 | 56.4 | 58.55 | 58.3 | 57.1 | 53.5 | 41.45 | 57.78 | 58.2993 | 36.65 | 56.7062 | 61.132 | 59.045
| Std Dev. |  | 1.5055 | 3.2457 | 1.0469 | 1.0897 | 1.4162 | 1.3934 | 0.29306 | 0.51945 | 1.2681 | 1.6548 | 1.1656 | 1.4488 | 0.46224 | 1.3378 | 1.9633 | 0.99191 | 1.3416 | 2.6048 | 7.3919 | 0.2881 | 1.5917 | 0.435 | 4.3061
| Rel.Std.Dev. |  | 2.537% | 5.887% | 1.781% | 2.165% | 2.23% | 2.524% | 0.3981% | 0.9423% | 1.984% | 2.332% | 2.265% | 2.548% | 0.7904% | 2.272% | 3.434% | 1.861% | 3.241% | 4.554% | 13.01% | 0.7839% | 2.8% | 0.7116% | 7.386%
| PDM3 |  | 3.349% | -3.966% | 2.376% | -12.31% | 10.64% | -3.821% | 28.24% | -3.978% | 11.3% | 23.62% | -10.35% | -0.9761% | 1.869% | 2.566% | -0.4245% | -7.177% | -27.89% | -0.3665% | -1.068% | -35.99% | -0.9747% | 6.482% | 1.555%
| Z-Score (Absolute) |  | 0.72 | 0.67 | 0.53 | 2.26 | 2.1 | 0.64 | 5.44 | 0.67 | 2.23 | 4.56 | 1.88 | 0.1 | 0.44 | 0.57 | 0 | 1.28 | 5.21 | 0.01 | 0.12 | 6.75 | 0.1 | 1.31 | 0.38
| Analytical results for Co in OREAS 600c (Certified Value 1.82 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 2 | 1.9 | 1.8 | 1.8 | 1.7471 | 1.8 | 1.8 | 1.74 | 1.9 | 2 | 1.8 | 1.9 | 1.9 | 1.8 | 1.8 | 1.8 | 1.8 | 2 | 1.7 | 1.8496 | 1.45 | 1.7905 | < 5 | 1.7 |  |  | 1
| 1 | 2 | 2 | 1.8 | 1.8 | 1.7 | 1.8416 | 1.9 | 1.8 | 1.85 | 2 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.9 | 2.1 | 1.8 | 2.3066 | 1.43 | 1.9272 | < 5 | 1.9 |  |  | 16
| 1 | 3 | 2 | 1.9 | 1.8 | 1.8 | 1.7878 | 1.8 | 1.8 | 1.88 | 1.9 | 1.8 | 1.8 | 1.8 | 1.8 | 1.9 | 1.6 | 1.8 | 1.8 | 2.1 | 1.8 | 2.0402 | 1.5 | 1.8313 | < 5 | 1.9 |  |  | 16
| 1 | 4 | 2 | 2.1 | 1.8 | 1.7 | 1.8476 | 1.9 | 1.8 | 1.89 | 1.9 | 1.8 | 1.8 | 1.8 | 1.8 | 1.9 | 1.7 | 1.8 | 1.78 | 2.1 | 1.8 | 1.3676 | 1.51 | 2.0791 | < 5 | 1.9 |  |  | 1.819
| 1 | 5 | 2 | 1.8 | 1.8 | 1.8 | 1.8033 | 1.9 | 1.8 | 1.87 | 1.9 | 1.8 | 1.7 | 1.9 | 1.8 | 1.9 | 1.7 | 1.8 | 1.79 | 2 | 1.7 | 4.0829 | 1.44 | 1.9023 | < 5 | 1.9 |  |  | 84
| 1 | 6 | 2 | 1.7 | 1.8 | 1.7 | 1.8274 | 1.8 | 1.8 | 1.84 | 1.9 | 1.8 | 1.9 | 1.8 | 1.8 | 1.9 | 1.7 | 1.8 | 1.72 | 2.1 | 1.7 | 6.6764 | 1.51 | 2.0505 | < 5 | 1.7
| Mean |  | 2 | 1.8667 | 1.8 | 1.75 | 1.8091 | 1.85 | 1.8 | 1.845 | 1.9167 | 1.8333 | 1.8 | 1.8333 | 1.8167 | 1.8667 | 1.7167 | 1.8 | 1.7983 | 2.0667 | 1.75 | 3.0539 | 1.4733 | 1.9301 |  | 1.8333
| Median |  | 2 | 1.85 | 1.8 | 1.75 | 1.8153 | 1.85 | 1.8 | 1.86 | 1.9 | 1.8 | 1.8 | 1.8 | 1.8 | 1.9 | 1.7 | 1.8 | 1.795 | 2.1 | 1.75 | 2.1734 | 1.475 | 1.9147 |  | 1.9
| Std Dev. |  | 0 | 0.13663 | 0 | 0.054772 | 0.037963 | 0.054772 | 0 | 0.054681 | 0.040825 | 0.08165 | 0.063246 | 0.05164 | 0.040825 | 0.05164 | 0.075277 | 0 | 0.058109 | 0.05164 | 0.054772 | 2.0031 | 0.037238 | 0.11552 |  | 0.10328
| Rel.Std.Dev. |  | 0% | 7.319% | 0% | 3.13% | 2.098% | 2.961% | 0% | 2.964% | 2.13% | 4.454% | 3.514% | 2.817% | 2.247% | 2.766% | 4.385% | 0% | 3.231% | 2.499% | 3.13% | 65.59% | 2.527% | 5.985% |  | 5.633%
| PDM3 |  | 9.952% | 2.622% | -1.043% | -3.792% | -0.5404% | 1.706% | -1.043% | 1.431% | 5.371% | 0.7894% | -1.043% | 0.7894% | -0.1268% | 2.622% | -5.624% | -1.043% | -1.135% | 13.62% | -3.792% | 67.89% | -19% | 6.112% |  | 0.7894%
| Z-Score (Absolute) |  | 3.37 | 0.67 | 0.67 | 1.69 | 0.49 | 0.34 | 0.67 | 0.24 | 1.69 | 0 | 0.67 | 0 | 0.34 | 0.67 | 2.36 | 0.67 | 0.71 | 4.72 | 1.69 | 24.69 | 7.28 | 1.96 | 13.49 | 0
| Analytical results for Cr in OREAS 600c (Certified Value 13.4 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 2
| 1 | 1 | 13 | 12.5 | 14 | 14 | 12.855 | 12 | 13.3 | 11.8205 | 14 | 13 | 20 | 12 | 14 | 13 | 14 | 15 | 14 | 13 | 14 | 11.3367 | 13 | 13.1756 | 11.715 |  |  |  | 1
| 1 | 2 | 13 | 13.5 | 14 | 14 | 12.0427 | 15 | 13.7 | 13.049 | 13 | 14 | 14 | 13 | 14 | 13 | 14 | 14 | 14 | 13 | 14 | 13.8148 | 13 | 12.6788 | 12.766 |  |  |  | 30
| 1 | 3 | 12 | 12.5 | 14 | 14 | 12.608 | 12 | 14.5 | 11.8415 | 14 | 14 | 13 | 12 | 13 | 15 | 17 | 13 | 15 | 13 | 15 | 13.8534 | 14 | 12.684 | 11.884 |  |  |  | 16
| 1 | 4 | 13 | 12.4 | 14 | 13 | 12.4474 | 13 | 14.4 | 12.23 | 13 | 14 | 14 | 13 | 14 | 14 | 13 | 14 | 15 | 13.3 | 13 | 13.6517 | 14 | 13.1109 | 12.435 |  |  |  | 13.3849
| 1 | 5 | 13 | 14.1 | 15 | 13 | 12.8118 | 15 | 14.1 | 12.2745 | 13.7 | 14 | 14 | 13 | 14 | 15 | 18 | 13 | 14 | 12.3 | 12 | 15.2877 | 12 | 13.0669 | 12.876 |  |  |  | 85
| 1 | 6 | 13 | 12.8 | 14 | 14 | 12.6396 | 13 | 14.5 | 12.0095 | 14 | 13 | 14 | 13 | 13 | 14 | 14 | 13 | 15 | 12.7 | 14 | 19.6446 | 13 | 12.5989 | 12.658
| Mean |  | 12.8333 | 12.9667 | 14.1667 | 13.6667 | 12.5674 | 13.3333 | 14.0833 | 12.2042 | 13.6167 | 13.6667 | 14.8333 | 12.6667 | 13.6667 | 14 | 15 | 13.6667 | 14.5 | 12.8833 | 13.6667 | 14.5982 | 13.1667 | 12.8859 | 12.389
| Median |  | 13 | 12.65 | 14 | 14 | 12.6238 | 13 | 14.25 | 12.1197 | 13.85 | 14 | 14 | 13 | 14 | 14 | 14 | 13.5 | 14.5 | 13 | 14 | 13.8341 | 13 | 12.8755 | 12.5465
| Std Dev. |  | 0.40825 | 0.68605 | 0.40825 | 0.5164 | 0.29632 | 1.3663 | 0.4916 | 0.45519 | 0.4916 | 0.5164 | 2.5626 | 0.5164 | 0.5164 | 0.89443 | 2 | 0.8165 | 0.54772 | 0.34303 | 1.0328 | 2.78 | 0.75277 | 0.2582 | 0.4823
| Rel.Std.Dev. |  | 3.181% | 5.291% | 2.882% | 3.779% | 2.358% | 10.25% | 3.491% | 3.73% | 3.61% | 3.779% | 17.28% | 4.077% | 3.779% | 6.389% | 13.33% | 5.974% | 3.777% | 2.663% | 7.557% | 19.04% | 5.717% | 2.004% | 3.893%
| PDM3 |  | -4.121% | -3.124% | 5.841% | 2.105% | -6.107% | -0.385% | 5.218% | -8.821% | 1.732% | 2.105% | 10.82% | -5.366% | 2.105% | 4.596% | 12.07% | 2.105% | 8.331% | -3.747% | 2.105% | 9.065% | -1.63% | -3.728% | -7.44%
| Z-Score (Absolute) |  | 0.8 | 0.67 | 0.48 | 0 | 1.06 | 0.32 | 0.4 | 1.41 | 0.05 | 0 | 1.12 | 0.96 | 0 | 0.32 | 1.28 | 0 | 0.8 | 0.75 | 0 | 0.9 | 0.48 | 0.75 | 1.23
| Analytical results for Cs in OREAS 600c (Certified Value 1.47 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [5 empty] | 2
| 1 | 1 | 1.58 | 1.56 | 1.47 | 1.4015 | 1.41 | 1.45 | 1.6887 | 1.37 | 1.64 | 1.77 | 1.2 | 1.38 | 1.49 | 1.48 | 1.4 | 1.31 | 1.25 | 1.49 | 0.99 | 1.2507 | 1.54 | [5 empty] | 1
| 1 | 2 | 1.55 | 1.52 | 1.46 | 1.4543 | 1.49 | 1.39 | 1.7874 | 1.47 | 1.65 | 1.72 | 1.21 | 1.32 | 1.48 | 1.41 | 1.4 | 1.33 | 1.27 | 1.6 | 0.91 | 1.3003 | 1.62 | [5 empty] | 31
| 1 | 3 | 1.51 | 1.51 | 1.5 | 1.4847 | 1.45 | 1.47 | 1.7942 | 1.41 | 1.67 | 1.76 | 1.19 | 1.32 | 1.44 | 1.53 | 1.3 | 1.28 | 1.26 | 1.65 | 0.93 | 1.2914 | 1.66 | [5 empty] | 16
| 1 | 4 | 1.53 | 1.65 | 1.52 | 1.4762 | 1.43 | 1.44 | 1.5996 | 1.47 | 1.61 | 1.77 | 1.23 | 1.33 | 1.48 | 1.45 | 1.31 | 1.37 | 1.26 | 1.6 | 1.02 | 1.2439 | 1.76 | [5 empty] | 1.4653
| 1 | 5 | 1.48 | 1.52 | 1.53 | 1.4535 | 1.46 | 1.43 | 1.7818 | 1.43 | 1.58 | 1.8 | 1.21 | 1.32 | 1.51 | 1.45 | 1.36 | 1.27 | 1.15 | 1.56 | 0.99 | 1.3258 | 1.68 | [5 empty] | 86
| 1 | 6 | 1.52 | 1.43 | 1.49 | 1.4253 | 1.38 | 1.46 | 1.8441 | 1.38 | 1.65 | 1.71 | 1.26 | 1.29 | 1.48 | 1.51 | 1.33 | 1.31 | 1.15 | 1.52 | 1.08 | 1.2799 | 1.43
| Mean |  | 1.5283 | 1.5317 | 1.495 | 1.4493 | 1.4367 | 1.44 | 1.7493 | 1.4217 | 1.6333 | 1.755 | 1.2167 | 1.3267 | 1.48 | 1.4717 | 1.35 | 1.3117 | 1.2233 | 1.57 | 0.98667 | 1.282 | 1.615
| Median |  | 1.525 | 1.52 | 1.495 | 1.4539 | 1.44 | 1.445 | 1.7846 | 1.42 | 1.645 | 1.765 | 1.21 | 1.32 | 1.48 | 1.465 | 1.345 | 1.31 | 1.255 | 1.58 | 0.99 | 1.2857 | 1.64
| Std Dev. |  | 0.034303 | 0.071949 | 0.027386 | 0.031267 | 0.038816 | 0.028284 | 0.088987 | 0.043089 | 0.03266 | 0.033912 | 0.025033 | 0.029439 | 0.022804 | 0.044008 | 0.043818 | 0.036009 | 0.057155 | 0.058652 | 0.061536 | 0.030909 | 0.11589
| Rel.Std.Dev. |  | 2.244% | 4.697% | 1.832% | 2.157% | 2.702% | 1.964% | 5.087% | 3.031% | 2% | 1.932% | 2.058% | 2.219% | 1.541% | 2.99% | 3.246% | 2.745% | 4.672% | 3.736% | 6.237% | 2.411% | 7.176%
| PDM3 |  | 4.3% | 4.527% | 2.025% | -1.097% | -1.956% | -1.728% | 19.38% | -2.979% | 11.47% | 19.77% | -16.97% | -9.463% | 1.002% | 0.4328% | -7.87% | -10.49% | -16.51% | 7.144% | -32.67% | -12.51% | 10.21%
| Z-Score (Absolute) |  | 0.44 | 0.46 | 0.26 | 0 | 0.07 | 0.05 | 1.68 | 0.15 | 1.03 | 1.71 | 1.3 | 0.68 | 0.17 | 0.13 | 0.55 | 0.77 | 1.26 | 0.67 | 2.58 | 0.93 | 0.93
| Analytical results for Cu in OREAS 600c (Certified Value 0.058 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 0.0573 | 0.058384 | 0.05872 | 0.05917 | 0.05784 | 0.0597 | 0.0558 | 0.05918 | 0.0588 | 0.0585 | 0.0595 | 0.0575 | 0.0597 | 0.0577 | 0.0626 | 0.059 | 0.0629 | 0.0643 | 0.0617 | 0.05202 | 0.05753 | 0.053852 | 0.0564 | 0.057377 | 0.055609 | 0.0633 | 1
| 1 | 2 | 0.0568 | 0.063019 | 0.05876 | 0.05902 | 0.058849 | 0.0591 | 0.0552 | 0.05928 | 0.0593 | 0.059 | 0.0575 | 0.0578 | 0.0598 | 0.0595 | 0.0606 | 0.0595 | 0.0626 | 0.0621 | 0.0606 | 0.05284 | 0.05606 | 0.053739 | 0.0546 | 0.057611 | 0.055119 | 0.0672 | 18
| 1 | 3 | 0.0563 | 0.058648 | 0.05813 | 0.0605 | 0.058441 | 0.0599 | 0.0574 | 0.059022 | 0.0595 | 0.0573 | 0.0568 | 0.0579 | 0.0598 | 0.0576 | 0.0585 | 0.0566 | 0.0637 | 0.0638 | 0.063 | 0.05208 | 0.05873 | 0.054105 | 0.055 | 0.05714 | 0.05551 | 0.0686 | 16
| 1 | 4 | 0.0569 | 0.06038 | 0.05901 | 0.05944 | 0.057758 | 0.061 | 0.056 | 0.059655 | 0.0592 | 0.057 | 0.0583 | 0.0572 | 0.0594 | 0.0576 | 0.061 | 0.0581 | 0.0647 | 0.0664 | 0.0624 | 0.0521 | 0.0576 | 0.054543 | 0.0577 | 0.058627 | 0.054074 | 0.071 | 0.058486
| 1 | 5 | 0.0564 | 0.061996 | 0.05841 | 0.06073 | 0.056745 | 0.0585 | 0.0578 | 0.059405 | 0.0598 | 0.0586 | 0.0586 | 0.0575 | 0.0588 | 0.0577 | 0.063 | 0.0588 | 0.0626 | 0.0628 | 0.0614 | 0.05248 | 0.05896 | 0.055198 | 0.056 | 0.056957 | 0.058056 | 0.0683 | 87
| 1 | 6 | 0.0549 | 0.062405 | 0.0582 | 0.06026 | 0.057551 | 0.0563 | 0.0565 | 0.059403 | 0.0594 | 0.059 | 0.0577 | 0.0583 | 0.0591 | 0.0574 | 0.0627 | 0.0577 | 0.0638 | 0.065 | 0.0609 | 0.05217 | 0.05598 | 0.053911 | 0.0585 | 0.058423 | 0.057478 | 0.0639
| Mean |  | 0.056433 | 0.060805 | 0.058538 | 0.059853 | 0.057864 | 0.059083 | 0.05645 | 0.059324 | 0.059333 | 0.058233 | 0.058067 | 0.0577 | 0.059433 | 0.057917 | 0.0614 | 0.058283 | 0.063383 | 0.064067 | 0.061667 | 0.052282 | 0.057477 | 0.054225 | 0.056367 | 0.057689 | 0.055975 | 0.06705
| Median |  | 0.0566 | 0.061188 | 0.058565 | 0.05985 | 0.057799 | 0.0594 | 0.05625 | 0.059341 | 0.05935 | 0.05855 | 0.058 | 0.05765 | 0.05955 | 0.05765 | 0.0618 | 0.05845 | 0.0633 | 0.06405 | 0.06155 | 0.052135 | 0.057565 | 0.054008 | 0.0562 | 0.057494 | 0.05556 | 0.06775
| Std Dev. |  | 0.00083347 | 0.0019786 | 0.00034707 | 0.00073271 | 0.00072993 | 0.0016005 | 0.00099146 | 0.00021753 | 0.00033267 | 0.00086872 | 0.00094375 | 0.00038471 | 0.00041312 | 0.00078337 | 0.0017263 | 0.0010458 | 0.00083287 | 0.0015436 | 0.00090701 | 0.00031802 | 0.0012678 | 0.00055415 | 0.0015135 | 0.00068684 | 0.0015025 | 0.0029521
| Rel.Std.Dev. |  | 1.477% | 3.254% | 0.5929% | 1.224% | 1.261% | 2.709% | 1.756% | 0.3667% | 0.5607% | 1.492% | 1.625% | 0.6667% | 0.6951% | 1.353% | 2.812% | 1.794% | 1.314% | 2.409% | 1.471% | 0.6083% | 2.206% | 1.022% | 2.685% | 1.191% | 2.684% | 4.403%
| PDM3 |  | -3.51% | 3.965% | 0.08905% | 2.337% | -1.064% | 1.021% | -3.482% | 1.433% | 1.448% | -0.4324% | -0.7174% | -1.344% | 1.619% | -0.9739% | 4.982% | -0.3469% | 8.373% | 9.541% | 5.438% | -10.61% | -1.726% | -7.286% | -3.624% | -1.363% | -4.294% | 14.64%
| Z-Score (Absolute) |  | 0.89 | 1.24 | 0.14 | 0.78 | 0.19 | 0.4 | 0.88 | 0.52 | 0.52 | 0.01 | 0.09 | 0.27 | 0.57 | 0.17 | 1.53 | 0.01 | 2.5 | 2.83 | 1.66 | 2.91 | 0.38 | 1.96 | 0.92 | 0.28 | 1.11 | 4.28
| Analytical results for Dy in OREAS 600c (Certified Value 2.11 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 2.07; 2.03: 2.047; 2.04: 1.9848; 2.05: 2.1; 2.06: 2.3469; 2.07: 2.51; 2.08: 1.95; 2.09: 1.5079; 2.10: 1.63; 2.11: 2.65 | 1
| 1 | 2 | 2.01: 2.18; 2.03: 2.075; 2.04: 2.1966; 2.05: 2.2; 2.06: 2.6148; 2.07: 2.38; 2.08: 1.94; 2.09: 2.1261; 2.10: 1.47; 2.11: 2.93 | 33
| 1 | 3 | 2.01: 2.04; 2.03: 2.122; 2.04: 2.1472; 2.05: 2.2; 2.06: 2.5503; 2.07: 2.39; 2.08: 1.91; 2.09: 2.2424; 2.10: 1.53; 2.11: 2.86 | 16
| 1 | 4 | 2.01: 2.28; 2.03: 2.111; 2.04: 2.1153; 2.05: 2.1; 2.06: 2.755; 2.07: 2.46; 2.08: 1.97; 2.09: 1.8795; 2.10: 1.72; 2.11: 2.92 | 2.1109
| 1 | 5 | 2.01: 1.99; 2.03: 2.124; 2.04: 2.099; 2.05: 2.2; 2.06: 2.4558; 2.07: 2.45; 2.08: 1.88; 2.09: 2.0802; 2.10: 1.62; 2.11: 2.87 | 88
| 1 | 6 | 2.01: 2.13; 2.03: 2.063; 2.04: 2.1064; 2.05: 2.1; 2.06: 2.6688; 2.07: 2.38; 2.08: 1.95; 2.09: 1.9899; 2.10: 1.85; 2.11: 2.55
| Mean |  | 2.01: 2.115; 2.03: 2.0903; 2.04: 2.1082; 2.05: 2.15; 2.06: 2.5653; 2.07: 2.4283; 2.08: 1.9333; 2.09: 1.971; 2.10: 1.6367; 2.11: 2.7967
| Median |  | 2.01: 2.1; 2.03: 2.093; 2.04: 2.1109; 2.05: 2.15; 2.06: 2.5825; 2.07: 2.42; 2.08: 1.945; 2.09: 2.0351; 2.10: 1.625; 2.11: 2.865
| Std Dev. |  | 2.01: 0.10483; 2.03: 0.032934; 2.04: 0.070291; 2.05: 0.054772; 2.06: 0.14769; 2.07: 0.053448; 2.08: 0.03266; 2.09: 0.258; 2.10: 0.1356; 2.11: 0.15795
| Rel.Std.Dev. |  | 2.01: 4.957%; 2.03: 1.576%; 2.04: 3.334%; 2.05: 2.548%; 2.06: 5.757%; 2.07: 2.201%; 2.08: 1.689%; 2.09: 13.09%; 2.10: 8.285%; 2.11: 5.648%
| PDM3 |  | 2.01: 0.1941%; 2.03: -0.9745%; 2.04: -0.1282%; 2.05: 1.852%; 2.06: 21.52%; 2.07: 15.04%; 2.08: -8.412%; 2.09: -6.628%; 2.10: -22.47%; 2.11: 32.49%
| Z-Score (Absolute) |  | 2.01: 0.01; 2.03: 0.09; 2.04: 0.01; 2.05: 0.16; 2.06: 1.92; 2.07: 1.34; 2.08: 0.75; 2.09: 0.59; 2.10: 2.01; 2.11: 2.9
| Analytical results for Er in OREAS 600c (Certified Value 0.54 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 0.45; 2.03: 0.494; 2.04: 0.53147; 2.05: 0.5; 2.06: 0.6248; 2.07: 0.57; 2.08: 0.48; 2.09: 0.3862; 2.10: 0.42; 2.11: 0.6 | 1
| 1 | 2 | 2.01: 0.52; 2.03: 0.497; 2.04: 0.55835; 2.05: 0.5; 2.06: 0.68; 2.07: 0.56; 2.08: 0.48; 2.09: 0.5679; 2.10: 0.38; 2.11: 0.65 | 34
| 1 | 3 | 2.01: 0.47; 2.03: 0.494; 2.04: 0.54305; 2.05: 0.5; 2.06: 0.6621; 2.07: 0.55; 2.08: 0.46; 2.09: 0.5957; 2.10: 0.4; 2.11: 0.62 | 16
| 1 | 4 | 2.01: 0.51; 2.03: 0.5; 2.04: 0.55562; 2.05: 0.5; 2.06: 0.6492; 2.07: 0.55; 2.08: 0.49; 2.09: 0.5033; 2.10: 0.44; 2.11: 0.68 | 0.53565
| 1 | 5 | 2.01: 0.47; 2.03: 0.505; 2.04: 0.54714; 2.05: 0.5; 2.06: 0.6431; 2.07: 0.57; 2.08: 0.48; 2.09: 0.5569; 2.10: 0.42; 2.11: 0.66 | 89
| 1 | 6 | 2.01: 0.51; 2.03: 0.502; 2.04: 0.54021; 2.05: 0.5; 2.06: 0.6298; 2.07: 0.54; 2.08: 0.48; 2.09: 0.5333; 2.10: 0.48; 2.11: 0.57
| Mean |  | 2.01: 0.48833; 2.03: 0.49867; 2.04: 0.54597; 2.05: 0.5; 2.06: 0.64817; 2.07: 0.55667; 2.08: 0.47833; 2.09: 0.52388; 2.10: 0.42333; 2.11: 0.63
| Median |  | 2.01: 0.49; 2.03: 0.4985; 2.04: 0.54509; 2.05: 0.5; 2.06: 0.64615; 2.07: 0.555; 2.08: 0.48; 2.09: 0.5451; 2.10: 0.42; 2.11: 0.635
| Std Dev. |  | 2.01: 0.028577; 2.03: 0.0044572; 2.04: 0.01; 2.05: 0; 2.06: 0.020579; 2.07: 0.012111; 2.08: 0.0098319; 2.09: 0.07437; 2.10: 0.034448; 2.11: 0.040988
| Rel.Std.Dev. |  | 2.01: 5.852%; 2.03: 0.8938%; 2.04: 1.832%; 2.05: 0%; 2.06: 3.175%; 2.07: 2.176%; 2.08: 2.055%; 2.09: 14.2%; 2.10: 8.137%; 2.11: 6.506%
| PDM3 |  | 2.01: -8.834%; 2.03: -6.905%; 2.04: 1.926%; 2.05: -6.656%; 2.06: 21%; 2.07: 3.923%; 2.08: -10.7%; 2.09: -2.198%; 2.10: -20.97%; 2.11: 17.61%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.48; 2.04: 0.42; 2.05: NA; 2.06: 2.36; 2.07: 0.62; 2.08: 0.86; 2.09: 0; 2.10: 1.91; 2.11: 2.01
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 600c (Certified Value 0.8 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 2
| 1 | 1 | 2.01: 0.77; 2.03: 0.788; 2.04: 0.66796; 2.05: 0.7; 2.06: 1.2053; 2.07: 0.96; 2.08: 0.83; 2.09: 0.59; 2.10: 0.63; 2.11: 0.93 | 1
| 1 | 2 | 2.01: 0.83; 2.03: 0.805; 2.04: 0.70695; 2.05: 0.8; 2.06: 1.2146; 2.07: 0.91; 2.08: 0.82; 2.09: 0.8731; 2.10: 0.58; 2.11: 0.97 | 35
| 1 | 3 | 2.01: 0.74; 2.03: 0.806; 2.04: 0.70707; 2.05: 0.8; 2.06: 1.1967; 2.07: 0.94; 2.08: 0.8; 2.09: 0.9257; 2.10: 0.6; 2.11: 0.91 | 16
| 1 | 4 | 2.01: 0.83; 2.03: 0.8; 2.04: 0.70952; 2.05: 0.8; 2.06: 1.2143; 2.07: 0.95; 2.08: 0.82; 2.09: 0.7739; 2.10: 0.66; 2.11: 1.01 | 0.80436
| 1 | 5 | 2.01: 0.77; 2.03: 0.774; 2.04: 0.69634; 2.05: 0.8; 2.06: 1.1605; 2.07: 0.97; 2.08: 0.8; 2.09: 0.858; 2.10: 0.64; 2.11: 1.02 | 90
| 1 | 6 | 2.01: 0.77; 2.03: 0.806; 2.04: 0.71265; 2.05: 0.7; 2.06: 1.2369; 2.07: 0.93; 2.08: 0.84; 2.09: 0.7633; 2.10: 0.73; 2.11: 0.82
| Mean |  | 2.01: 0.785; 2.03: 0.7965; 2.04: 0.70008; 2.05: 0.76667; 2.06: 1.2047; 2.07: 0.94333; 2.08: 0.81833; 2.09: 0.79733; 2.10: 0.64; 2.11: 0.94333
| Median |  | 2.01: 0.77; 2.03: 0.8025; 2.04: 0.70701; 2.05: 0.8; 2.06: 1.2098; 2.07: 0.945; 2.08: 0.82; 2.09: 0.81595; 2.10: 0.635; 2.11: 0.95
| Std Dev. |  | 2.01: 0.036742; 2.03: 0.012988; 2.04: 0.016667; 2.05: 0.05164; 2.06: 0.025476; 2.07: 0.021602; 2.08: 0.016021; 2.09: 0.11882; 2.10: 0.052536; 2.11: 0.074207
| Rel.Std.Dev. |  | 2.01: 4.681%; 2.03: 1.631%; 2.04: 2.381%; 2.05: 6.736%; 2.06: 2.115%; 2.07: 2.29%; 2.08: 1.958%; 2.09: 14.9%; 2.10: 8.209%; 2.11: 7.866%
| PDM3 |  | 2.01: -2.406%; 2.03: -0.9766%; 2.04: -12.96%; 2.05: -4.686%; 2.06: 49.77%; 2.07: 17.28%; 2.08: 1.738%; 2.09: -0.873%; 2.10: -20.43%; 2.11: 17.28%
| Z-Score (Absolute) |  | 2.01: 0.09; 2.03: 0.01; 2.04: 0.67; 2.05: NA; 2.06: 2.82; 2.07: 1.01; 2.08: 0.15; 2.09: 0; 2.10: 1.09; 2.11: 1.01
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 600c (Certified Value 2.03 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  | 3
| 1 | 1 | 2.12 | 1.85 | 2 | 1.98 | 2.0642 | 2.16 | 2.12 | 2.01 | 2.37 | 2.22 | 1.95 | 2.02 | 2 | 2.09 | 2.03 | 1.8794 | 2 | 2.13 | 1.98 | 1.77 | 1.9209 | 2.028 | 2.0324 | 2.0317 |  |  | 1
| 1 | 2 | 2.11 | 2.01 | 2.02 | 2.03 | 2.1031 | 2.16 | 2.11 | 2.03 | 2.36 | 2.17 | 1.97 | 2 | 1.99 | 2.05 | 2.04 | 1.8892 | 2.01 | 2.02 | 1.98 | 1.79 | 1.9363 | 2.091 | 2.0939 | 1.9823 |  |  | 0
| 1 | 3 | 2.1 | 1.82 | 1.97 | 2 | 2.1028 | 2.13 | 2.19 | 2.02 | 2.41 | 2.13 | 1.97 | 2 | 2 | 2.05 | 1.94 | 1.8951 | 2 | 2.14 | 1.96 | 1.82 | 1.9783 | 2.035 | 2.0353 | 2.1771 |  |  | 16
| 1 | 4 | 2.12 | 1.9 | 2.01 | 1.93 | 2.1199 | 2.21 | 2.12 | 2.01 | 2.4 | 2.19 | 1.94 | 2.04 | 2 | 2.04 | 2 | 1.8962 | 2.04 | 2.05 | 1.97 | 1.81 | 1.9333 | 2.063 | 2.0416 | 1.9445 |  |  | 2.0316
| 1 | 5 | 2.11 | 2.08 | 2 | 2.04 | 2.0828 | 2.15 | 2.16 | 2.03 | 2.42 | 2.2 | 1.96 | 2.02 | 2.01 | 2.14 | 2.03 | 1.8632 | 2.03 | 2.09 | 1.98 | 1.79 | 1.9655 | 2.028 | 2.1138 | 1.9895 |  |  | 91
| 1 | 6 | 2.1 | 1.97 | 2 | 1.95 | 2.0797 | 2.08 | 2.19 | 2.01 | 2.37 | 2.17 | 1.99 | 2 | 2 | 2.09 | 2.01 | 1.9064 | 2.03 | 2.07 | 1.95 | 1.78 | 1.9228 | 2.07 | 1.9994 | 1.9945
| Mean |  | 2.11 | 1.9383 | 2 | 1.9883 | 2.0921 | 2.1483 | 2.1483 | 2.0183 | 2.3883 | 2.18 | 1.9633 | 2.0133 | 2 | 2.0767 | 2.0083 | 1.8882 | 2.0183 | 2.0833 | 1.97 | 1.7933 | 1.9428 | 2.0525 | 2.0527 | 2.0199
| Median |  | 2.11 | 1.935 | 2 | 1.99 | 2.0928 | 2.155 | 2.14 | 2.015 | 2.385 | 2.18 | 1.965 | 2.01 | 2 | 2.07 | 2.02 | 1.8921 | 2.02 | 2.08 | 1.975 | 1.79 | 1.9348 | 2.049 | 2.0385 | 1.992
| Std Dev. |  | 0.0089443 | 0.099482 | 0.016733 | 0.043551 | 0.020128 | 0.042622 | 0.03656 | 0.0098319 | 0.024833 | 0.030984 | 0.017512 | 0.01633 | 0.0063246 | 0.037771 | 0.03656 | 0.015141 | 0.017224 | 0.046332 | 0.012649 | 0.018619 | 0.023612 | 0.026098 | 0.04268 | 0.081889
| Rel.Std.Dev. |  | 0.4239% | 5.132% | 0.8367% | 2.19% | 0.9621% | 1.984% | 1.702% | 0.4871% | 1.04% | 1.421% | 0.8919% | 0.8111% | 0.3162% | 1.819% | 1.82% | 0.8019% | 0.8534% | 2.224% | 0.6421% | 1.038% | 1.215% | 1.272% | 2.079% | 4.054%
| PDM3 |  | 3.857% | -4.592% | -1.557% | -2.131% | 2.976% | 5.744% | 5.744% | -0.6548% | 17.56% | 7.303% | -3.362% | -0.9009% | -1.557% | 2.216% | -1.147% | -7.058% | -0.6548% | 2.545% | -3.034% | -11.73% | -4.37% | 1.027% | 1.039% | -0.5753%
| Z-Score (Absolute) |  | 1.09 | 0.95 | 0.22 | 0.36 | 0.88 | 1.55 | 1.55 | 0 | 4.4 | 1.92 | 0.65 | 0.06 | 0.22 | 0.69 | 0.12 | 1.55 | 0 | 0.77 | 0.58 | 2.68 | 0.9 | 0.41 | 0.41 | 0.02
| Analytical results for Ga in OREAS 600c (Certified Value 5.29 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 5.51 | 5.3 | 5 | 5.11 | 5.3783 | 5.93 | 5.91 | 5.82 | 5.3 | 5.68 | 4.54 | 5.75 | 5.1 | 5.04 | 4.78 | 4.54 | < 10 | < 5 | 5.36 | 4.1779 | 3.7 | 4.8447 | 6.186 | 5.39 |  |  | 1
| 1 | 2 | 5.18 | 5.3 | 5 | 4.97 | 5.6431 | 5.7 | 5.73 | 5.97 | 5.8 | 5.51 | 4.6 | 5.63 | 5.05 | 4.84 | 4.8 | 4.62 | < 10 | 5.5 | 5.55 | 5.7044 | 3.53 | 5.1217 | 4.976 | 5.82 |  |  | 36
| 1 | 3 | 5.19 | 5.2 | 5 | 5.03 | 5.4446 | 5.93 | 5.94 | 5.55 | 5.5 | 5.87 | 4.48 | 5.57 | 5.05 | 4.84 | 4.38 | 4.59 | < 10 | < 5 | 5.44 | 5.9101 | 3.64 | 4.8966 | 6.356 | 5.8 |  |  | 16
| 1 | 4 | 5.18 | 5.7 | 5 | 5.07 | 5.4181 | 5.95 | 5.86 | 5.93 | 5.5 | 5.76 | 4.46 | 5.45 | 5.13 | 5.01 | 4.5 | 4.75 | < 10 | 5.1 | 5.5 | 5.1191 | 3.83 | 4.9508 | 5.661 | 6.1 |  |  | 5.2924
| 1 | 5 | 5.11 | 5.3 | 5 | 5.07 | 5.3767 | 5.93 | 5.9 | 6.68 | 5.5 | 5.8 | 4.53 | 5.66 | 5.17 | 5.1 | 4.74 | 4.61 | < 10 | 5 | 5.25 | 5.5695 | 3.68 | 5.0816 | 5.579 | 5.86 |  |  | 92
| 1 | 6 | 5.16 | 5.2 | 5 | 5.05 | 5.503 | 5.63 | 5.93 | 5.65 | 5.2 | 5.55 | 4.83 | 5.41 | 5.12 | 5.04 | 4.54 | 4.65 | < 10 | 5.8 | 5.22 | 4.9256 | 3.94 | 4.8445 | 5.704 | 5.12
| Mean |  | 5.2217 | 5.3333 | 5 | 5.05 | 5.4606 | 5.845 | 5.8783 | 5.9333 | 5.4667 | 5.695 | 4.5733 | 5.5783 | 5.1033 | 4.9783 | 4.6233 | 4.6267 |  | 5.35 | 5.3867 | 5.2344 | 3.72 | 4.9567 | 5.7437 | 5.6817
| Median |  | 5.18 | 5.3 | 5 | 5.06 | 5.4314 | 5.93 | 5.905 | 5.875 | 5.5 | 5.72 | 4.535 | 5.6 | 5.11 | 5.025 | 4.64 | 4.615 |  | 5.3 | 5.4 | 5.3443 | 3.69 | 4.9237 | 5.6825 | 5.81
| Std Dev. |  | 0.14414 | 0.18619 | 0 | 0.047329 | 0.10097 | 0.14139 | 0.077825 | 0.39973 | 0.20656 | 0.14237 | 0.13501 | 0.12937 | 0.047188 | 0.11107 | 0.17363 | 0.070616 |  | 0.36968 | 0.13382 | 0.63476 | 0.14491 | 0.11966 | 0.48897 | 0.35791
| Rel.Std.Dev. |  | 2.76% | 3.491% | 0% | 0.9372% | 1.849% | 2.419% | 1.324% | 6.737% | 3.779% | 2.5% | 2.952% | 2.319% | 0.9246% | 2.231% | 3.755% | 1.526% |  | 6.91% | 2.484% | 12.13% | 3.896% | 2.414% | 8.513% | 6.299%
| PDM3 |  | -1.336% | 0.7744% | -5.524% | -4.579% | 3.18% | 10.44% | 11.07% | 12.11% | 3.294% | 7.608% | -13.59% | 5.404% | -3.572% | -5.933% | -12.64% | -12.58% |  | 1.089% | 1.782% | -1.094% | -29.71% | -6.343% | 8.528% | 7.356%
| Z-Score (Absolute) |  | 0.01 | 0.23 | 0.49 | 0.39 | 0.5 | 1.34 | 1.41 | 1.53 | 0.52 | 1.01 | 1.42 | 0.76 | 0.27 | 0.54 | 1.31 | 1.3 | 0.49 | 1.8 | 0.34 | 0.01 | 3.27 | 0.59 | 1.12 | 0.98
| Analytical results for Gd in OREAS 600c (Certified Value 3.98 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  | 3
| 1 | 1 | 2.01: 3.8; 2.03: 3.659; 2.04: 3.7594; 2.05: 3.9; 2.06: 6.1508; 2.07: 4.5; 2.08: 3.78; 2.10: 3.2215; 2.11: 2.9; 2.12: 3.77 |  |  | 1
| 1 | 2 | 2.01: 3.95; 2.03: 3.661; 2.04: 4.1887; 2.05: 4; 2.06: 6.1403; 2.07: 4.26; 2.08: 3.89; 2.10: 4.7576; 2.11: 2.6; 2.12: 3.99 |  |  | 37
| 1 | 3 | 2.01: 3.59; 2.03: 3.714; 2.04: 4.0687; 2.05: 3.9; 2.06: 6.1042; 2.07: 4.35; 2.08: 3.67; 2.10: 4.9937; 2.11: 2.69; 2.12: 3.98 |  |  | 16
| 1 | 4 | 2.01: 4.11; 2.03: 3.748; 2.04: 3.969; 2.05: 4.1; 2.06: 6.0721; 2.07: 4.3; 2.08: 3.83; 2.10: 4.19; 2.11: 3.12; 2.12: 4.23 |  |  | 3.9817
| 1 | 5 | 2.01: 3.71; 2.03: 3.727; 2.04: 3.9678; 2.05: 4; 2.06: 6.2431; 2.07: 4.47; 2.08: 3.79; 2.10: 4.5752; 2.11: 3; 2.12: 4.02 |  |  | 93
| 1 | 6 | 2.01: 3.75; 2.03: 3.712; 2.04: 4.0354; 2.05: 3.9; 2.06: 6.2842; 2.07: 4.32; 2.08: 3.93; 2.10: 4.1908; 2.11: 3.56; 2.12: 3.43
| Mean |  | 2.01: 3.8183; 2.03: 3.7035; 2.04: 3.9982; 2.05: 3.9667; 2.06: 6.1658; 2.07: 4.3667; 2.08: 3.815; 2.10: 4.3215; 2.11: 2.9783; 2.12: 3.9033
| Median |  | 2.01: 3.775; 2.03: 3.713; 2.04: 4.0022; 2.05: 3.95; 2.06: 6.1456; 2.07: 4.335; 2.08: 3.81; 2.10: 4.383; 2.11: 2.95; 2.12: 3.985
| Std Dev. |  | 2.01: 0.18509; 2.03: 0.03606; 2.04: 0.14242; 2.05: 0.08165; 2.06: 0.081784; 2.07: 0.096678; 2.08: 0.091597; 2.10: 0.62462; 2.11: 0.34388; 2.12: 0.27406
| Rel.Std.Dev. |  | 2.01: 4.847%; 2.03: 0.9737%; 2.04: 3.562%; 2.05: 2.058%; 2.06: 1.326%; 2.07: 2.214%; 2.08: 2.401%; 2.10: 14.45%; 2.11: 11.55%; 2.12: 7.021%
| PDM3 |  | 2.01: -4.102%; 2.03: -6.986%; 2.04: 0.4147%; 2.05: -0.3768%; 2.06: 54.85%; 2.07: 9.669%; 2.08: -4.186%; 2.10: 8.534%; 2.11: -25.2%; 2.12: -1.967%
| Z-Score (Absolute) |  | 2.01: 0.45; 2.03: 0.89; 2.04: 0.24; 2.05: 0.12; 2.06: 8.56; 2.07: 1.66; 2.08: 0.46; 2.10: 1.48; 2.11: 3.67; 2.12: 0.12
| Analytical results for Ge in OREAS 600c (Certified Value 0.11 ppm) | [27 empty] | Certified
Ge | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab; 2.15: Lab; 2.16: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.03: 2.03; 2.04: 2.06; 2.05: 2.07; 2.06: 2.08; 2.07: 2.13; 2.08: 2.14; 2.10: 2.15; 2.11: 2.16; 2.12: 2.17; 2.13: 2.19; 2.14: 2.20; 2.15: 2.22; 2.16: 2.25 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*MS; 2.15: AR*MS; 2.16: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.4g; 2.14: 0.5g; 2.15: 0.2g; 2.16: 0.5g |  |  | 2
| 1 | 1 | 2.01: 0.21; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.24; 2.07: 0.09; 2.08: 0.1; 2.10: 0.1; 2.11: 0.12; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.13; 2.15: 0.07; 2.16: < 0.01 |  |  | 1
| 1 | 2 | 2.01: 0.19; 2.03: 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.22; 2.07: 0.09; 2.08: 0.1; 2.10: 0.1; 2.11: 0.11; 2.12: 0.18; 2.13: < 0.05; 2.14: 0.12; 2.15: 0.08; 2.16: < 0.01 |  |  | 38
| 1 | 3 | 2.01: 0.19; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.24; 2.07: 0.06; 2.08: 0.11; 2.10: 0.11; 2.11: 0.1; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.12; 2.15: 0.08; 2.16: < 0.01 |  |  | 16
| 1 | 4 | 2.01: 0.17; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.24; 2.07: 0.09; 2.08: 0.11; 2.10: 0.11; 2.11: 0.11; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.1; 2.15: 0.08; 2.16: < 0.01 |  |  | 0.11143
| 1 | 5 | 2.01: 0.19; 2.03: < 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.24; 2.07: 0.08; 2.08: 0.1; 2.10: 0.12; 2.11: 0.12; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.12; 2.15: 0.08; 2.16: < 0.01 |  |  | 94
| 1 | 6 | 2.01: 0.19; 2.03: 0.1; 2.04: < 0.2; 2.05: < 0.1; 2.06: 0.23; 2.07: 0.1; 2.08: 0.1; 2.10: 0.11; 2.11: 0.11; 2.12: 0.17; 2.13: < 0.05; 2.14: 0.12; 2.15: 0.07; 2.16: < 0.01
| Mean |  | 2.01: 0.19; 2.03: 0.1; 2.06: 0.235; 2.07: 0.085; 2.08: 0.10333; 2.10: 0.10833; 2.11: 0.11167; 2.12: 0.17167; 2.14: 0.11833; 2.15: 0.076667
| Median |  | 2.01: 0.19; 2.03: 0.1; 2.06: 0.24; 2.07: 0.09; 2.08: 0.1; 2.10: 0.11; 2.11: 0.11; 2.12: 0.17; 2.14: 0.12; 2.15: 0.08
| Std Dev. |  | 2.01: 0.012649; 2.03: 0; 2.06: 0.0083666; 2.07: 0.013784; 2.08: 0.005164; 2.10: 0.0075277; 2.11: 0.0075277; 2.12: 0.0040825; 2.14: 0.0098319; 2.15: 0.005164
| Rel.Std.Dev. |  | 2.01: 6.657%; 2.03: 0%; 2.06: 3.56%; 2.07: 16.22%; 2.08: 4.997%; 2.10: 6.949%; 2.11: 6.741%; 2.12: 2.378%; 2.14: 8.309%; 2.15: 6.736%
| PDM3 |  | 2.01: 70.51%; 2.03: -10.26%; 2.06: 110.9%; 2.07: -23.72%; 2.08: -7.265%; 2.10: -2.778%; 2.11: 0.2137%; 2.12: 54.06%; 2.14: 6.197%; 2.15: -31.2%
| Z-Score (Absolute) |  | 2.01: 2.19; 2.03: NA; 2.04: 0.08; 2.05: 1.35; 2.06: 3.33; 2.07: 0.46; 2.08: 0; 2.10: 0.13; 2.11: 0.21; 2.12: 1.73; 2.13: 1.98; 2.14: 0.38; 2.15: 0.67; 2.16: 2.49
Results from laboratory 2.03 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 600c (Certified Value 1.41 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 3
| 1 | 1 | 1.54 | 1.56 | 1.41 | 1.3108 | 0.2 | 1.34 | 1.43 | 1.38 | 1.5 | 1.34 | 1.5 | 1.36 | 1.38 | 1.3 | 1.53 | 1.4 | 1.41 | 0.92 | 1.2392 | 0.76 | [6 empty] | 1
| 1 | 2 | 1.41 | 1.4 | 1.4 | 1.3998 | 0.2 | 1.28 | 1.51 | 1.39 | 1.41 | 1.32 | 1.45 | 1.4 | 1.37 | 1.32 | 1.53 | 1.63 | 1.43 | 0.84 | 1.382 | 0.66 | [6 empty] | 22
| 1 | 3 | 1.56 | 1.51 | 1.44 | 1.39 | 0.1 | 1.34 | 1.48 | 1.4 | 1.43 | 1.44 | 1.47 | 1.38 | 1.36 | 1.22 | 1.51 | 1.57 | 1.53 | 0.87 | 1.381 | 0.76 | [6 empty] | 16
| 1 | 4 | 1.49 | 1.69 | 1.41 | 1.3884 | 0.1 | 1.36 | 1.44 | 1.4 | 1.4 | 1.32 | 1.47 | 1.38 | 1.38 | 1.22 | 1.5 | 1.51 | 1.44 | 0.97 | 1.3122 | 1.05 | [6 empty] | 1.4146
| 1 | 5 | 1.55 | 1.33 | 1.42 | 1.3589 | 0.3 | 1.35 | 1.49 | 1.39 | 1.47 | 1.36 | 1.44 | 1.42 | 1.43 | 1.27 | 1.52 | 1.56 | 1.46 | 0.91 | 1.3711 | 0.83 | [6 empty] | 95
| 1 | 6 | 1.37 | 1.36 | 1.38 | 1.4048 | 0.2 | 1.33 | 1.52 | 1.43 | 1.41 | 1.45 | 1.4 | 1.38 | 1.37 | 1.25 | 1.54 | 1.62 | 1.43 | 1.03 | 1.2416 | 1.2
| Mean |  | 1.4867 | 1.475 | 1.41 | 1.3755 | 0.18333 | 1.3333 | 1.4783 | 1.3983 | 1.4367 | 1.3717 | 1.455 | 1.3867 | 1.3817 | 1.2633 | 1.5217 | 1.5483 | 1.45 | 0.92333 | 1.3212 | 0.87667
| Median |  | 1.515 | 1.455 | 1.41 | 1.3892 | 0.2 | 1.34 | 1.485 | 1.395 | 1.42 | 1.35 | 1.46 | 1.38 | 1.375 | 1.26 | 1.525 | 1.565 | 1.435 | 0.915 | 1.3416 | 0.795
| Std Dev. |  | 0.079666 | 0.13751 | 0.02 | 0.035465 | 0.075277 | 0.028048 | 0.03656 | 0.017224 | 0.039833 | 0.058793 | 0.033912 | 0.020656 | 0.024833 | 0.041312 | 0.01472 | 0.084715 | 0.042426 | 0.068605 | 0.067685 | 0.20539
| Rel.Std.Dev. |  | 5.359% | 9.323% | 1.418% | 2.578% | 41.06% | 2.104% | 2.473% | 1.232% | 2.773% | 4.286% | 2.331% | 1.49% | 1.797% | 3.27% | 0.9673% | 5.471% | 2.926% | 7.43% | 5.123% | 23.43%
| PDM3 |  | 5.094% | 4.27% | -0.3253% | -2.767% | -87.04% | -5.745% | 4.505% | -1.15% | 1.56% | -3.035% | 2.856% | -1.975% | -2.328% | -10.69% | 7.569% | 9.454% | 2.502% | -34.73% | -6.605% | -38.03%
| Z-Score (Absolute) |  | 0.95 | 0.83 | 0.18 | 0.17 | 12.18 | 0.6 | 0.86 | 0.06 | 0.45 | 0.21 | 0.63 | 0.06 | 0.11 | 1.3 | 1.3 | 1.57 | 0.58 | 4.73 | 0.72 | 5.2
| Analytical results for Hg in OREAS 600c (Certified Value 0.14 ppm) | [27 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 25g | [5 empty] | 2
| 1 | 1 | 0.14 | 0.245 | 0.15 | 0.2 | < 1 | 0.14 | 0.16 | 0.16 | 0.17 | 0.15 | 0.13 | 0.15 | 0.16 | 0.15 | 0.07 | 0.12 | < 5 | < 2 | < 0.01 | 0.15621 | 0.13 | [5 empty] | 1
| 1 | 2 | 0.09 | 0.126 | 0.15 | 0.2 | < 1 | 0.14 | 0.14 | 0.15 | 0.13 | 0.16 | 0.16 | 0.13 | 0.16 | 0.14 | 0.07 | 0.12 | < 5 | < 2 | < 0.01 | 0.15232 | 0.18 | [5 empty] | 3
| 1 | 3 | 0.08 | 0.22 | 0.14 | 0.2 | < 1 | 0.15 | 0.14 | 0.14 | 0.11 | 0.16 | 0.13 | 0.13 | 0.15 | 0.13 | 0.09 | 0.119 | < 5 | < 2 | < 0.01 | 0.15725 | 0.15 | [5 empty] | 16
| 1 | 4 | 0.15 | 0.238 | 0.15 | 0.2 | < 1 | 0.15 | 0.15 | 0.15 | 0.14 | 0.14 | 0.14 | 0.15 | 0.17 | 0.14 | 0.1 | 0.105 | < 5 | < 2 | < 0.01 | 0.14661 | 0.16 | [5 empty] | 0.14378
| 1 | 5 | 0.12 | 0.122 | 0.15 | 0.2 | < 1 | 0.16 | 0.16 | 0.18 | 0.12 | 0.15 | 0.14 | 0.16 | 0.16 | 0.11 | 0.08 | 0.116 | < 5 | < 2 | < 0.01 | 0.15475 | 0.16 | [5 empty] | 96
| 1 | 6 | 0.13 | 0.114 | 0.15 | 0.2 | < 1 | 0.14 | 0.15 | 0.16 | 0.14 | 0.12 | 0.16 | 0.15 | 0.16 | 0.15 | 0.07 | 0.106 | < 5 | < 2 | < 0.01 | 0.14601 | 0.14
| Mean |  | 0.11833 | 0.1775 | 0.14833 | 0.2 |  | 0.14667 | 0.15 | 0.15667 | 0.135 | 0.14667 | 0.14333 | 0.145 | 0.16 | 0.13667 | 0.08 | 0.11433 |  |  |  | 0.15219 | 0.15333
| Median |  | 0.125 | 0.173 | 0.15 | 0.2 |  | 0.145 | 0.15 | 0.155 | 0.135 | 0.15 | 0.14 | 0.15 | 0.16 | 0.14 | 0.075 | 0.1175 |  |  |  | 0.15354 | 0.155
| Std Dev. |  | 0.027869 | 0.062909 | 0.0040825 | 3.0405e-17 |  | 0.008165 | 0.0089443 | 0.013663 | 0.020736 | 0.015055 | 0.013663 | 0.012247 | 0.0063246 | 0.015055 | 0.012649 | 0.0070048 |  |  |  | 0.004852 | 0.017512
| Rel.Std.Dev. |  | 23.55% | 35.44% | 2.752% | 1.52e-14% |  | 5.567% | 5.963% | 8.721% | 15.36% | 10.27% | 9.532% | 8.447% | 3.953% | 11.02% | 15.81% | 6.127% |  |  |  | 3.188% | 11.42%
| PDM3 |  | -17.7% | 23.46% | 3.17% | 39.11% |  | 2.011% | 4.329% | 8.966% | -6.104% | 2.011% | -0.3076% | 0.8516% | 11.28% | -4.944% | -44.36% | -20.48% |  |  |  | 5.854% | 6.648%
| Z-Score (Absolute) |  | 1.69 | 1.73 | 0.05 | NA | 20.37 | 0.05 | 0.14 | 0.53 | 0.72 | 0.05 | 0.24 | 0.14 | 0.72 | 0.63 | 3.9 | 1.92 | 135.97 | 49.27 | 8.24 | 0.27 | 0.34
| Analytical results for Ho in OREAS 600c (Certified Value 0.26 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab | [5 empty] | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.12: 2.23; 2.13: 2.25; 2.14: 2.28 | [5 empty] | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.12: AR*MS; 2.13: AR*MS; 2.14: AR*MS | [5 empty] | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.12: 0.5g; 2.13: 0.5g; 2.14: 25g | [5 empty] | 3
| 1 | 1 | 2.01: 0.25; 2.03: 0.266; 2.04: 0.24831; 2.05: 0.3; 2.06: 0.3065; 2.07: 0.29; 2.08: 0.26; 2.12: 0.164; 2.13: 0.22; 2.14: 0.27 | [5 empty] | 1
| 1 | 2 | 2.01: 0.27; 2.03: 0.266; 2.04: 0.2566; 2.05: 0.3; 2.06: 0.3314; 2.07: 0.28; 2.08: 0.27; 2.12: 0.2369; 2.13: 0.2; 2.14: 0.29 | [5 empty] | 23
| 1 | 3 | 2.01: 0.24; 2.03: 0.266; 2.04: 0.26814; 2.05: 0.3; 2.06: 0.3231; 2.07: 0.28; 2.08: 0.27; 2.12: 0.2453; 2.13: 0.21; 2.14: 0.29 | [5 empty] | 16
| 1 | 4 | 2.01: 0.27; 2.03: 0.267; 2.04: 0.25674; 2.05: 0.3; 2.06: 0.3305; 2.07: 0.28; 2.08: 0.26; 2.12: 0.2081; 2.13: 0.23; 2.14: 0.3 | [5 empty] | 0.25791
| 1 | 5 | 2.01: 0.25; 2.03: 0.272; 2.04: 0.25385; 2.05: 0.3; 2.06: 0.316; 2.07: 0.28; 2.08: 0.26; 2.12: 0.232; 2.13: 0.22; 2.14: 0.3 | [5 empty] | 97
| 1 | 6 | 2.01: 0.26; 2.03: 0.267; 2.04: 0.26261; 2.05: 0.3; 2.06: 0.3422; 2.07: 0.27; 2.08: 0.27; 2.12: 0.219; 2.13: 0.25; 2.14: 0.27
| Mean |  | 2.01: 0.25667; 2.03: 0.26733; 2.04: 0.25771; 2.05: 0.3; 2.06: 0.32495; 2.07: 0.28; 2.08: 0.265; 2.12: 0.21755; 2.13: 0.22167; 2.14: 0.28667
| Median |  | 2.01: 0.255; 2.03: 0.2665; 2.04: 0.25667; 2.05: 0.3; 2.06: 0.3268; 2.07: 0.28; 2.08: 0.265; 2.12: 0.2255; 2.13: 0.22; 2.14: 0.29
| Std Dev. |  | 2.01: 0.012111; 2.03: 0.0023381; 2.04: 0.0069027; 2.05: 0; 2.06: 0.012601; 2.07: 0.0063246; 2.08: 0.0054772; 2.12: 0.029368; 2.13: 0.017224; 2.14: 0.013663
| Rel.Std.Dev. |  | 2.01: 4.718%; 2.03: 0.8746%; 2.04: 2.678%; 2.05: 0%; 2.06: 3.878%; 2.07: 2.259%; 2.08: 2.067%; 2.12: 13.5%; 2.13: 7.77%; 2.14: 4.766%
| PDM3 |  | 2.01: -0.4831%; 2.03: 3.653%; 2.04: -0.07971%; 2.05: 16.32%; 2.06: 25.99%; 2.07: 8.564%; 2.08: 2.748%; 2.12: -15.65%; 2.13: -14.05%; 2.14: 11.15%
| Z-Score (Absolute) |  | 2.01: 0.37; 2.03: 0.1; 2.04: 0.33; 2.05: NA; 2.06: 2.69; 2.07: 0.67; 2.08: 0; 2.12: 2.13; 2.13: 1.95; 2.14: 0.97
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 600c (Certified Value 0.14 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 01g | 25g | [8 empty] | 3
| 1 | 1 | 0.133 | 0.15 | 0.13 | 0.13 | 0.13 | 0.1 | 0.14 | 0.14 | 0.128 | 0.147 | 0.141 | 0.135 | 0.137 | 0.134 | 0.128 | 0.016 | 0.12792 | 0.133 | [8 empty] | 1
| 1 | 2 | 0.161 | 0.15 | 0.13 | 0.13 | 0.13 | 0.1 | 0.14 | 0.14 | 0.134 | 0.144 | 0.135 | 0.128 | 0.14 | 0.132 | 0.128 | 0.017 | 0.1384 | 0.137 | [8 empty] | 24
| 1 | 3 | 0.157 | 0.14 | 0.13 | 0.13 | 0.134 | 0.1 | 0.13 | 0.15 | 0.119 | 0.144 | 0.144 | 0.139 | 0.135 | 0.142 | 0.131 | 0.014 | 0.13612 | 0.142 | [8 empty] | 16
| 1 | 4 | 0.172 | 0.16 | 0.13 | 0.14 | 0.135 | 0.1 | 0.13 | 0.14 | 0.133 | 0.144 | 0.14 | 0.133 | 0.134 | 0.141 | 0.12 | 0.014 | 0.13171 | 0.146 | [8 empty] | 0.13672
| 1 | 5 | 0.149 | 0.12 | 0.14 | 0.13 | 0.135 | 0.1 | 0.14 | 0.14 | 0.129 | 0.141 | 0.14 | 0.134 | 0.145 | 0.136 | 0.125 | 0.015 | 0.13452 | 0.139 | [8 empty] | 98
| 1 | 6 | 0.145 | 0.13 | 0.13 | 0.13 | 0.141 | 0.1 | 0.14 | 0.14 | 0.13 | 0.139 | 0.138 | 0.135 | 0.136 | 0.139 | 0.141 | 0.013 | 0.13045 | 0.132
| Mean |  | 0.15283 | 0.14167 | 0.13167 | 0.13167 | 0.13417 | 0.1 | 0.13667 | 0.14167 | 0.12883 | 0.14317 | 0.13967 | 0.134 | 0.13783 | 0.13733 | 0.12883 | 0.014833 | 0.13319 | 0.13817
| Median |  | 0.153 | 0.145 | 0.13 | 0.13 | 0.1345 | 0.1 | 0.14 | 0.14 | 0.1295 | 0.144 | 0.14 | 0.1345 | 0.1365 | 0.1375 | 0.128 | 0.0145 | 0.13311 | 0.138
| Std Dev. |  | 0.013571 | 0.01472 | 0.0040825 | 0.0040825 | 0.0040702 | 1.5202e-17 | 0.005164 | 0.0040825 | 0.0053448 | 0.0027869 | 0.0030111 | 0.0035777 | 0.0040702 | 0.0039833 | 0.0070261 | 0.001472 | 0.0038702 | 0.0053448
| Rel.Std.Dev. |  | 8.879% | 10.39% | 3.101% | 3.101% | 3.034% | 1.52e-14% | 3.779% | 2.882% | 4.149% | 1.947% | 2.156% | 2.67% | 2.953% | 2.9% | 5.454% | 9.923% | 2.906% | 3.868%
| PDM3 |  | 11.79% | 3.618% | -3.696% | -3.696% | -1.868% | -26.86% | -0.03896% | 3.618% | -5.768% | 4.715% | 2.155% | -1.989% | 0.8144% | 0.4486% | -5.768% | -89.15% | -2.585% | 1.058%
| Z-Score (Absolute) |  | 2.18 | 0.67 | 0.67 | 0.67 | 0.34 | NA | 0 | 0.67 | 1.06 | 0.88 | 0.4 | 0.36 | 0.16 | 0.09 | 1.06 | 16.43 | 0.47 | 0.2
Results from laboratory 2.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 600c (Certified Value 0.233 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.23 | 0.23 | 0.23 | 0.2356 | 0.23104 | 0.26 | 0.22 | 0.23 | 0.31 | 0.359 | 0.21 | 0.23 | 0.22 | 0.26 | 0.23 | 0.237 | 0.23 | 0.22 | 0.24 | 0.3645 | 0.1967 | 0.22017 | 0.26567 |  |  |  | 1
| 1 | 2 | 0.23 | 0.37 | 0.23 | 0.2362 | 0.24026 | 0.25 | 0.21 | 0.23 | 0.327 | 0.3465 | 0.21 | 0.22 | 0.22 | 0.25 | 0.23 | 0.24 | 0.22 | 0.24 | 0.25 | 0.3372 | 0.1976 | 0.23365 | 0.27241 |  |  |  | 0
| 1 | 3 | 0.24 | 0.23 | 0.23 | 0.2384 | 0.23213 | 0.27 | 0.22 | 0.24 | 0.315 | 0.3448 | 0.22 | 0.27 | 0.22 | 0.29 | 0.22 | 0.232 | 0.24 | 0.22 | 0.26 | 0.2873 | 0.1959 | 0.23029 | 0.2685 |  |  |  | 16
| 1 | 4 | 0.23 | 0.25 | 0.23 | 0.2278 | 0.23154 | 0.28 | 0.21 | 0.23 | 0.328 | 0.3603 | 0.21 | 0.27 | 0.22 | 0.25 | 0.22 | 0.241 | 0.23 | 0.23 | 0.25 | 0.4124 | 0.1976 | 0.22364 | 0.2495 |  |  |  | 0.23339
| 1 | 5 | 0.23 | 0.23 | 0.23 | 0.2399 | 0.23121 | 0.25 | 0.22 | 0.23 | 0.332 | 0.3606 | 0.21 | 0.23 | 0.22 | 0.26 | 0.23 | 0.247 | 0.23 | 0.23 | 0.25 | 0.3148 | 0.1951 | 0.22827 | 0.26466 |  |  |  | 99
| 1 | 6 | 0.22 | 0.36 | 0.23 | 0.2288 | 0.23531 | 0.26 | 0.22 | 0.24 | 0.315 | 0.3473 | 0.21 | 0.25 | 0.22 | 0.26 | 0.23 | 0.239 | 0.23 | 0.23 | 0.24 | 0.3767 | 0.1992 | 0.22049 | 0.2755
| Mean |  | 0.23 | 0.27833 | 0.23 | 0.23445 | 0.23358 | 0.26167 | 0.21667 | 0.23333 | 0.32117 | 0.35308 | 0.21167 | 0.245 | 0.22 | 0.26167 | 0.22667 | 0.23933 | 0.23 | 0.22833 | 0.24833 | 0.34882 | 0.19702 | 0.22609 | 0.26604
| Median |  | 0.23 | 0.24 | 0.23 | 0.2359 | 0.23183 | 0.26 | 0.22 | 0.23 | 0.321 | 0.35315 | 0.21 | 0.24 | 0.22 | 0.26 | 0.23 | 0.2395 | 0.23 | 0.23 | 0.25 | 0.35085 | 0.19715 | 0.22596 | 0.26708
| Std Dev. |  | 0.0063246 | 0.067651 | 0 | 0.0050167 | 0.0036321 | 0.01169 | 0.005164 | 0.005164 | 0.0089312 | 0.0076025 | 0.0040825 | 0.021679 | 0 | 0.01472 | 0.005164 | 0.0049261 | 0.0063246 | 0.0075277 | 0.0075277 | 0.045039 | 0.001447 | 0.0055105 | 0.0090746
| Rel.Std.Dev. |  | 2.75% | 24.31% | 0% | 2.14% | 1.555% | 4.468% | 2.383% | 2.213% | 2.781% | 2.153% | 1.929% | 8.849% | 0% | 5.625% | 2.278% | 2.058% | 2.75% | 3.297% | 3.031% | 12.91% | 0.7344% | 2.437% | 3.411%
| PDM3 |  | -1.453% | 19.26% | -1.453% | 0.4532% | 0.08112% | 12.11% | -7.166% | -0.02522% | 37.61% | 51.28% | -9.309% | 4.974% | -5.738% | 12.11% | -2.882% | 2.546% | -1.453% | -2.168% | 6.402% | 49.46% | -15.59% | -3.13% | 13.99%
| Z-Score (Absolute) |  | 0.18 | 2.22 | 0.18 | 0.04 | 0 | 1.39 | 0.84 | 0.01 | 4.35 | 5.93 | 1.09 | 0.57 | 0.67 | 1.39 | 0.34 | 0.29 | 0.18 | 0.26 | 0.73 | 5.72 | 1.82 | 0.37 | 1.61
| Analytical results for La in OREAS 600c (Certified Value 28.7 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 2
| 1 | 1 | 33.4 | 27.7 | 28 | 29.281 | 24.7985 | 30.7 | 28.3 | 35.131 | 27.6 | 32 | 36.64 | 27 | 30.4 | 30.1 | 30.1 | 30.1 | 27.3 | 23.8 | 26.6 | 27 | 20.2533 | 20 | 26.9484 | 31.043 | 31.69 |  | 1
| 1 | 2 | 31 | 28.5 | 29 | 28.855 | 25.8186 | 31.4 | 26.6 | 35.442 | 28.2 | 31 | 34.54 | 27.5 | 28.5 | 30.2 | 29.4 | 30 | 27.7 | 25.7 | 27.5 | 29.3 | 28.6137 | 20 | 28.6234 | 30.671 | 33.56 |  | 25
| 1 | 3 | 32.8 | 26.3 | 28 | 30.133 | 25.3183 | 30.9 | 28.8 | 35.072 | 27.9 | 32.1 | 35.41 | 26.8 | 28.6 | 29.2 | 29 | 27.9 | 26.2 | 24.7 | 26.6 | 30 | 29.5687 | 19 | 28.3722 | 30.943 | 33.14 |  | 16
| 1 | 4 | 31.3 | 30.2 | 29 | 29.746 | 25.1209 | 31.7 | 28.3 | 35.208 | 28.4 | 31.1 | 35.63 | 26.4 | 28.5 | 30.1 | 29.7 | 27.8 | 27 | 24.8 | 27.5 | 29.4 | 26.0344 | 20 | 27.0002 | 30.393 | 35.48 |  | 28.7172
| 1 | 5 | 31.9 | 27.2 | 28 | 30.348 | 24.9601 | 31.3 | 27.4 | 34.997 | 28.2 | 31.8 | 36.08 | 25.7 | 28.8 | 29.8 | 29.4 | 29.4 | 27 | 24.8 | 27.4 | 28.4 | 27.6066 | 19 | 28.826 | 30.586 | 34.27 |  | 100
| 1 | 6 | 31.1 | 27 | 28 | 29.474 | 26.1068 | 30.3 | 28.8 | 35.16 | 28.1 | 31.6 | 34.9 | 27.1 | 27.8 | 29.5 | 30.8 | 28.4 | 27.4 | 25.9 | 27.5 | 27.2 | 24.3608 | 20 | 26.8054 | 30.709 | 28.59
| Mean |  | 31.9167 | 27.8167 | 28.3333 | 29.6395 | 25.3539 | 31.05 | 28.0333 | 35.1683 | 28.0667 | 31.6 | 35.5333 | 26.75 | 28.7667 | 29.8167 | 29.7333 | 28.9333 | 27.1 | 24.95 | 27.1833 | 28.55 | 26.0729 | 19.6667 | 27.7626 | 30.7242 | 32.7883
| Median |  | 31.6 | 27.45 | 28 | 29.61 | 25.2196 | 31.1 | 28.3 | 35.1455 | 28.15 | 31.7 | 35.52 | 26.9 | 28.55 | 29.95 | 29.55 | 28.9 | 27.15 | 24.8 | 27.45 | 28.85 | 26.8205 | 20 | 27.6862 | 30.69 | 33.35
| Std Dev. |  | 0.98675 | 1.379 | 0.5164 | 0.55285 | 0.51028 | 0.51284 | 0.86872 | 0.15269 | 0.28048 | 0.46043 | 0.76628 | 0.62849 | 0.86872 | 0.39707 | 0.6377 | 1.0347 | 0.51381 | 0.76092 | 0.4535 | 1.2357 | 3.4009 | 0.5164 | 0.93851 | 0.23722 | 2.408
| Rel.Std.Dev. |  | 3.092% | 4.957% | 1.823% | 1.865% | 2.013% | 1.652% | 3.099% | 0.4342% | 0.9993% | 1.457% | 2.157% | 2.349% | 3.02% | 1.332% | 2.145% | 3.576% | 1.896% | 3.05% | 1.668% | 4.328% | 13.04% | 2.626% | 3.38% | 0.7721% | 7.344%
| PDM3 |  | 11.14% | -3.136% | -1.337% | 3.212% | -11.71% | 8.123% | -2.381% | 22.46% | -2.265% | 10.04% | 23.74% | -6.85% | 0.1722% | 3.829% | 3.538% | 0.7526% | -5.632% | -13.12% | -5.341% | -0.5823% | -9.208% | -31.52% | -3.324% | 6.989% | 14.18%
| Z-Score (Absolute) |  | 1.45 | 0.37 | 0.14 | 0.44 | 1.47 | 1.06 | 0.28 | 2.89 | 0.26 | 1.31 | 3.06 | 0.85 | 0.05 | 0.51 | 0.48 | 0.12 | 0.69 | 1.65 | 0.66 | 0.05 | 1.15 | NA | 0.4 | 0.92 | 1.84
| Analytical results for Li in OREAS 600c (Certified Value 18.1 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 2
| 1 | 1 | 23.2 | 19.6 | 16.8 | 16.8551 | 18 | 17.1 | 15.336 | 16.8 | 21.9 | 16.8 | 17.2 | 17.8 | 17.6 | 19.6 | 17.8 | 19.9 | 18.2 | 15.2 | 19.9444 | 13 | 18.9535 | 21.27 | 18.6 |  |  |  | 1
| 1 | 2 | 22.7 | 16.9 | 16.7 | 16.8722 | 17.7 | 17.4 | 15.97 | 17 | 21.3 | 16.3 | 17.6 | 17.5 | 17.4 | 19.2 | 18 | 19.4 | 19.9 | 14.9 | 19.4044 | 14 | 18.891 | 21.029 | 19.9 |  |  |  | 45
| 1 | 3 | 24.2 | 19.4 | 17.3 | 16.5127 | 18.4 | 16.6 | 15.884 | 17.1 | 20.7 | 16.4 | 17.3 | 17.8 | 17.1 | 18.3 | 17.2 | 19.7 | 18.8 | 15.9 | 18.5792 | 13 | 18.1486 | 20.96 | 20.3 |  |  |  | 16
| 1 | 4 | 23.5 | 21.5 | 16.9 | 16.2267 | 18.3 | 16.9 | 15.887 | 17.1 | 20.4 | 16.2 | 17.2 | 18 | 17.2 | 19.2 | 17.4 | 20.8 | 19 | 15.8 | 21.587 | 14 | 18.3026 | 20.576 | 21.1 |  |  |  | 18.1454
| 1 | 5 | 23.5 | 17.3 | 17.3 | 16.154 | 18.7 | 17.9 | 15.372 | 17 | 20.7 | 17.3 | 17.3 | 17.8 | 17.5 | 19.8 | 18 | 20 | 17.8 | 16.7 | 20.1861 | 13 | 19.0097 | 21.047 | 20.5 |  |  |  | 101
| 1 | 6 | 22.7 | 17 | 17.5 | 16.7272 | 17.4 | 17 | 15.915 | 16.4 | 20.7 | 16.7 | 17.6 | 17.6 | 17.4 | 19.7 | 17.7 | 20 | 17.5 | 15.6 | 22.1468 | 14 | 18.4315 | 21.185 | 18.8
| Mean |  | 23.3 | 18.6167 | 17.0833 | 16.558 | 18.0833 | 17.15 | 15.7273 | 16.9 | 20.95 | 16.6167 | 17.3667 | 17.75 | 17.3667 | 19.3 | 17.6833 | 19.9667 | 18.5333 | 15.6833 | 20.308 | 13.5 | 18.6228 | 21.0112 | 19.8667
| Median |  | 23.35 | 18.35 | 17.1 | 16.6199 | 18.15 | 17.05 | 15.8855 | 17 | 20.7 | 16.55 | 17.3 | 17.8 | 17.4 | 19.4 | 17.75 | 19.95 | 18.5 | 15.7 | 20.0652 | 13.5 | 18.6612 | 21.038 | 20.1
| Std Dev. |  | 0.56921 | 1.8541 | 0.32506 | 0.31326 | 0.47924 | 0.45056 | 0.29105 | 0.26833 | 0.55045 | 0.40702 | 0.18619 | 0.17607 | 0.18619 | 0.55136 | 0.32506 | 0.46762 | 0.88015 | 0.62423 | 1.3394 | 0.54772 | 0.37282 | 0.24117 | 0.98522
| Rel.Std.Dev. |  | 2.443% | 9.959% | 1.903% | 1.892% | 2.65% | 2.627% | 1.851% | 1.588% | 2.627% | 2.449% | 1.072% | 0.9919% | 1.072% | 2.857% | 1.838% | 2.342% | 4.749% | 3.98% | 6.595% | 4.057% | 2.002% | 1.148% | 4.959%
| PDM3 |  | 28.41% | 2.597% | -5.853% | -8.748% | -0.3422% | -5.486% | -13.33% | -6.864% | 15.46% | -8.425% | -4.292% | -2.179% | -4.292% | 6.363% | -2.547% | 10.04% | 2.138% | -13.57% | 11.92% | -25.6% | 2.631% | 15.79% | 9.486%
| Z-Score (Absolute) |  | 3.13 | 0.41 | 0.49 | 0.79 | 0.1 | 0.45 | 1.27 | 0.59 | 1.77 | 0.76 | 0.32 | 0.1 | 0.32 | 0.81 | 0.14 | 1.19 | 0.36 | 1.3 | 1.39 | NA | 0.41 | 1.8 | 1.14
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 600c (Certified Value 0.029 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS |  |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 01g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 25g |  |  |  | 3
| 1 | 1 | 2.01: 0.02; 2.03: 0.024; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.04; 2.09: 0.0189; 2.10: 0.027 |  |  |  | 1
| 1 | 2 | 2.01: 0.02; 2.03: 0.024; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0295; 2.10: 0.03 |  |  |  | 26
| 1 | 3 | 2.01: 0.02; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.04; 2.09: 0.0306; 2.10: 0.029 |  |  |  | 16
| 1 | 4 | 2.01: 0.03; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.04; 2.09: 0.0261; 2.10: 0.032 |  |  |  | 0.028973
| 1 | 5 | 2.01: 0.03; 2.03: 0.028; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.03; 2.09: 0.0282; 2.10: 0.03 |  |  |  | 102
| 1 | 6 | 2.01: 0.03; 2.03: 0.025; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.03; 2.08: 0.04; 2.09: 0.0293; 2.10: 0.025
| Mean |  | 2.01: 0.025; 2.03: 0.025167; 2.07: 0.03; 2.08: 0.036667; 2.09: 0.0271; 2.10: 0.028833
| Median |  | 2.01: 0.025; 2.03: 0.025; 2.07: 0.03; 2.08: 0.04; 2.09: 0.02875; 2.10: 0.0295
| Std Dev. |  | 2.01: 0.0054772; 2.03: 0.001472; 2.07: 0; 2.08: 0.005164; 2.09: 0.0042965; 2.10: 0.0024833
| Rel.Std.Dev. |  | 2.01: 21.91%; 2.03: 5.849%; 2.07: 0%; 2.08: 14.08%; 2.09: 15.85%; 2.10: 8.613%
| PDM3 |  | 2.01: -13.71%; 2.03: -13.14%; 2.07: 3.543%; 2.08: 26.55%; 2.09: -6.466%; 2.10: -0.4832%
| Z-Score (Absolute) |  | 2.01: 0.69; 2.03: 0.65; 2.05: 0.69; 2.06: 5.15; 2.07: 0.48; 2.08: 2.03; 2.09: 0.2; 2.10: 0.2
| Analytical results for Mg in OREAS 600c (Certified Value 0.094 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.09 | 0.09 | 0.1 | 0.1 | 0.095861 | 0.11 | 0.09 | 0.093 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.111 | 0.09 | 0.09 | 0.09 | < 0.1 | 0.0983 | 0.095365 | 0.093844 |  |  |  | 1
| 1 | 2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.099014 | 0.11 | 0.1 | 0.093 | 0.12 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.111 | 0.1 | 0.1 | 0.09 | < 0.1 | 0.1013 | 0.099021 | 0.095731 |  |  |  | 0
| 1 | 3 | 0.1 | 0.09 | 0.1 | 0.1 | 0.098057 | 0.11 | 0.1 | 0.093 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.109 | 0.097 | 0.09 | 0.1 | < 0.1 | 0.0977 | 0.097471 | 0.094919 |  |  |  | 16
| 1 | 4 | 0.1 | 0.09 | 0.1 | 0.09 | 0.097809 | 0.12 | 0.09 | 0.093 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.113 | 0.097 | 0.1 | 0.1 | < 0.1 | 0.0983 | 0.09626 | 0.089002 |  |  |  | 0.094498
| 1 | 5 | 0.1 | 0.09 | 0.1 | 0.1 | 0.095658 | 0.12 | 0.1 | 0.093 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.111 | 0.094 | 0.1 | 0.1 | < 0.1 | 0.0983 | 0.099317 | 0.09431 |  |  |  | 103
| 1 | 6 | 0.09 | 0.09 | 0.1 | 0.1 | 0.097956 | 0.11 | 0.1 | 0.093 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.111 | 0.096 | 0.1 | 0.09 | < 0.1 | 0.0995 | 0.096165 | 0.09727
| Mean |  | 0.096667 | 0.091667 | 0.1 | 0.098333 | 0.097393 | 0.11333 | 0.096667 | 0.093 | 0.11167 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.111 | 0.095667 | 0.096667 | 0.095 |  | 0.0989 | 0.097267 | 0.094179
| Median |  | 0.1 | 0.09 | 0.1 | 0.1 | 0.097882 | 0.11 | 0.1 | 0.093 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.111 | 0.0965 | 0.1 | 0.095 |  | 0.0983 | 0.096866 | 0.094615
| Std Dev. |  | 0.005164 | 0.0040825 | 1.5202e-17 | 0.0040825 | 0.0013352 | 0.005164 | 0.005164 | 1.5202e-17 | 0.0040825 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 0.0012649 | 0.0033862 | 0.005164 | 0.0054772 |  | 0.0013145 | 0.0016228 | 0.0028087
| Rel.Std.Dev. |  | 5.342% | 4.454% | 1.52e-14% | 4.152% | 1.371% | 4.556% | 5.342% | 1.635e-14% | 3.656% | 1.52e-14% | 1.689e-14% | 1.689e-14% | 1.689e-14% | 1.52e-14% | 1.689e-14% | 1.14% | 3.54% | 5.342% | 5.766% |  | 1.329% | 1.668% | 2.982%
| PDM3 |  | 2.295% | -2.996% | 5.823% | 4.059% | 3.064% | 19.93% | 2.295% | -1.585% | 18.17% | 5.823% | -4.759% | -4.759% | -4.759% | 5.823% | -4.759% | 17.46% | 1.237% | 2.295% | 0.5317% |  | 4.659% | 2.93% | -0.3367%
| Z-Score (Absolute) |  | 0.11 | 1.03 | NA | 0.5 | 0.28 | 3.92 | 0.11 | 0.72 | 3.54 | NA | 1.41 | 1.41 | 1.41 | NA | 1.41 | 3.39 | 0.11 | 0.11 | 0.27 | 10.55 | 0.63 | 0.25 | 0.45
Results from laboratories 2.04, 2.12 and 2.16 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Mn in OREAS 600c (Certified Value 0.034 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 0.0347 | 0.0334 | 0.0347 | 0.0338 | 0.033525 | 0.0368 | 0.0338 | 0.0333 | 0.0353 | 0.0369 | 0.0322 | 0.0357 | 0.0331 | 0.0366 | 0.0339 | 0.0304 | 0.0336 | 0.03097 | 0.0336 | < 0.1 | 0.0341 | 0.032968 | 0.032535 | 0.0344 |  |  | 1
| 1 | 2 | 0.034 | 0.0353 | 0.0347 | 0.0343 | 0.034826 | 0.036 | 0.0327 | 0.0332 | 0.0353 | 0.0357 | 0.0325 | 0.035 | 0.0336 | 0.0358 | 0.0343 | 0.0302 | 0.0325 | 0.03205 | 0.034 | < 0.1 | 0.0356 | 0.034065 | 0.032336 | 0.0373 |  |  | 27
| 1 | 3 | 0.0339 | 0.0328 | 0.0343 | 0.0339 | 0.034273 | 0.0353 | 0.0336 | 0.0335 | 0.0339 | 0.0353 | 0.0327 | 0.0357 | 0.0332 | 0.0349 | 0.0323 | 0.0308 | 0.0339 | 0.03128 | 0.0343 | < 0.1 | 0.0341 | 0.034215 | 0.032217 | 0.0379 |  |  | 16
| 1 | 4 | 0.0338 | 0.0344 | 0.035 | 0.0326 | 0.03465 | 0.0379 | 0.0331 | 0.0331 | 0.0361 | 0.0362 | 0.032 | 0.0361 | 0.0335 | 0.0358 | 0.0334 | 0.0314 | 0.0332 | 0.03148 | 0.0342 | < 0.1 | 0.0349 | 0.033118 | 0.03181 | 0.0392 |  |  | 0.034159
| 1 | 5 | 0.0339 | 0.0369 | 0.0345 | 0.0346 | 0.033929 | 0.0362 | 0.0338 | 0.0329 | 0.0361 | 0.0364 | 0.0321 | 0.0355 | 0.034 | 0.0367 | 0.0339 | 0.03 | 0.0334 | 0.03169 | 0.034 | < 0.1 | 0.0341 | 0.03432 | 0.032342 | 0.0375 |  |  | 104
| 1 | 6 | 0.0337 | 0.0343 | 0.0344 | 0.0331 | 0.034221 | 0.0345 | 0.0336 | 0.0333 | 0.0352 | 0.0358 | 0.0328 | 0.0353 | 0.0334 | 0.0363 | 0.0333 | 0.0308 | 0.0329 | 0.03137 | 0.0338 | < 0.1 | 0.0349 | 0.032938 | 0.032368 | 0.0338
| Mean |  | 0.034 | 0.034517 | 0.0346 | 0.033717 | 0.034237 | 0.036117 | 0.033433 | 0.033217 | 0.035317 | 0.03605 | 0.032383 | 0.03555 | 0.033467 | 0.036017 | 0.033517 | 0.0306 | 0.03325 | 0.031473 | 0.033983 |  | 0.034617 | 0.033604 | 0.032268 | 0.036683
| Median |  | 0.0339 | 0.03435 | 0.0346 | 0.03385 | 0.034247 | 0.0361 | 0.0336 | 0.03325 | 0.0353 | 0.036 | 0.03235 | 0.0356 | 0.03345 | 0.03605 | 0.03365 | 0.0306 | 0.0333 | 0.031425 | 0.034 |  | 0.0345 | 0.033591 | 0.032339 | 0.0374
| Std Dev. |  | 0.00035777 | 0.0014525 | 0.00025298 | 0.00074677 | 0.00047349 | 0.0011788 | 0.00044121 | 0.00020412 | 0.00080602 | 0.00056833 | 0.00033116 | 0.00037815 | 0.00032042 | 0.00066758 | 0.00069976 | 0.00050596 | 0.000501 | 0.00036903 | 0.00025626 |  | 0.00062102 | 0.00066079 | 0.00024658 | 0.002116
| Rel.Std.Dev. |  | 1.052% | 4.208% | 0.7312% | 2.215% | 1.383% | 3.264% | 1.32% | 0.6145% | 2.282% | 1.577% | 1.023% | 1.064% | 0.9574% | 1.854% | 2.088% | 1.653% | 1.507% | 1.173% | 0.7541% |  | 1.794% | 1.966% | 0.7642% | 5.768%
| PDM3 |  | -0.4642% | 1.048% | 1.292% | -1.294% | 0.2306% | 5.732% | -2.123% | -2.757% | 3.39% | 5.537% | -5.197% | 4.073% | -2.026% | 5.44% | -1.879% | -10.42% | -2.66% | -7.861% | -0.513% |  | 1.341% | -1.624% | -5.534% | 7.391%
| Z-Score (Absolute) |  | 0.01 | 0.47 | 0.54 | 0.24 | 0.22 | 1.89 | 0.5 | 0.69 | 1.18 | 1.83 | 1.43 | 1.39 | 0.47 | 1.8 | 0.42 | 3.02 | 0.66 | 2.24 | 0.01 | 14.23 | 0.53 | 0.34 | 1.53 | 2.39
| Analytical results for Mo in OREAS 600c (Certified Value 3.63 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 3
| 1 | 1 | 3.61 | 3.72 | 3.5 | 3.7 | 4.0148 | 4.02 | 3.58 | 3.61 | 3.7 | 3.5 | 3.6 | 3.74 | 3.86 | 3.8 | 3.6 | 3.7 | 2.99 | 3.59 | 3.4 | 3.81 | 3.5865 | 3.34 | 3.7952 | < 5 | 3.62 |  | 1
| 1 | 2 | 3.22 | 3.54 | 3.6 | 3.8 | 3.8042 | 3.96 | 3.55 | 3.68 | 3.7 | 3.4 | 3.6 | 3.6 | 3.8 | 3.55 | 3.81 | 3.4 | 3.03 | 3.65 | 3.4 | 4.02 | 4.6117 | 3.19 | 3.78 | < 5 | 3.94 |  | 28
| 1 | 3 | 3.26 | 3.71 | 3.5 | 3.8 | 3.8776 | 3.98 | 3.66 | 3.73 | 3.7 | 3.5 | 3.37 | 3.65 | 3.84 | 3.65 | 3.52 | 3.1 | 3.08 | 3.83 | 3.4 | 3.72 | 3.2836 | 3.07 | 3.349 | < 5 | 3.95 |  | 16
| 1 | 4 | 3.46 | 3.93 | 3.6 | 3.7 | 3.8125 | 4.05 | 3.65 | 3.86 | 3.7 | 3.6 | 3.73 | 3.89 | 3.87 | 3.68 | 3.53 | 3.2 | 3.07 | 3.71 | 3.5 | 3.85 | 3.5586 | 3.45 | 3.5076 | < 5 | 4.13 |  | 3.6283
| 1 | 5 | 3.16 | 3.44 | 3.5 | 3.8 | 3.8364 | 3.88 | 3.62 | 3.86 | 3.7 | 3.6 | 3.49 | 3.63 | 3.93 | 3.77 | 3.65 | 2.5 | 3.18 | 3.65 | 3.4 | 3.66 | 3.4838 | 3.38 | 3.6116 | < 5 | 3.87 |  | 105
| 1 | 6 | 3.27 | 3.56 | 3.6 | 3.8 | 3.8716 | 3.93 | 3.64 | 3.63 | 3.7 | 3.4 | 3.7 | 3.57 | 3.76 | 3.7 | 3.55 | 3.4 | 2.9 | 3.63 | 3.1 | 3.85 | 3.0877 | 3.62 | 3.2806 | < 5 | 3.74
| Mean |  | 3.33 | 3.65 | 3.55 | 3.7667 | 3.8695 | 3.97 | 3.6167 | 3.7283 | 3.7 | 3.5 | 3.5817 | 3.68 | 3.8433 | 3.6917 | 3.61 | 3.2167 | 3.0417 | 3.6767 | 3.3667 | 3.8183 | 3.602 | 3.3417 | 3.554 |  | 3.875
| Median |  | 3.265 | 3.635 | 3.55 | 3.8 | 3.854 | 3.97 | 3.63 | 3.705 | 3.7 | 3.5 | 3.6 | 3.64 | 3.85 | 3.69 | 3.575 | 3.3 | 3.05 | 3.65 | 3.4 | 3.83 | 3.5212 | 3.36 | 3.5596 |  | 3.905
| Std Dev. |  | 0.17018 | 0.1739 | 0.054772 | 0.05164 | 0.077193 | 0.061319 | 0.043205 | 0.11017 | 4.8648e-16 | 0.089443 | 0.13408 | 0.11798 | 0.058878 | 0.089312 | 0.10936 | 0.40702 | 0.09411 | 0.084538 | 0.13663 | 0.12449 | 0.52943 | 0.19364 | 0.2152 |  | 0.17785
| Rel.Std.Dev. |  | 5.11% | 4.764% | 1.543% | 1.371% | 1.995% | 1.545% | 1.195% | 2.955% | 1.315e-14% | 2.556% | 3.743% | 3.206% | 1.532% | 2.419% | 3.029% | 12.65% | 3.094% | 2.299% | 4.058% | 3.26% | 14.7% | 5.795% | 6.055% |  | 4.59%
| PDM3 |  | -8.222% | 0.5973% | -2.159% | 3.813% | 6.648% | 9.417% | -0.3214% | 2.756% | 1.975% | -3.537% | -1.286% | 1.424% | 5.926% | 1.746% | -0.5052% | -11.35% | -16.17% | 1.332% | -7.212% | 5.237% | -0.7261% | -7.901% | -2.049% |  | 6.798%
| Z-Score (Absolute) |  | 1.66 | 0.19 | 0.39 | 0.87 | 1.46 | 2.04 | 0 | 0.65 | 0.48 | 0.67 | 0.2 | 0.37 | 1.31 | 0.43 | 0.04 | 2.31 | 3.32 | 0.35 | 1.44 | 1.17 | 0.08 | 1.59 | 0.36 | 6.45 | 1.49
| Analytical results for Na in OREAS 600c (Certified Value 0.073 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.07 | 0.063 | 0.075 | 0.07 | 0.073264 | 0.074 | 0.07 | 0.07 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.09 | 0.1 | 0.073 | 0.06 | 0.08 | 0.08 | 0.1917 | 0.067 | 0.067021 | 0.067058 |  |  |  | 1
| 1 | 2 | 0.07 | 0.176 | 0.076 | 0.07 | 0.075326 | 0.073 | 0.06 | 0.073 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.08 | 0.11 | 0.076 | 0.07 | 0.07 | 0.08 | 0.1859 | 0.067 | 0.072005 | 0.068529 |  |  |  | 0
| 1 | 3 | 0.07 | 0.062 | 0.074 | 0.07 | 0.072521 | 0.074 | 0.06 | 0.073 | 0.07 | 0.08 | 0.07 | 0.11 | 0.08 | 0.12 | 0.1 | 0.073 | 0.066 | 0.07 | 0.08 | 0.1424 | 0.067 | 0.069226 | 0.067834 |  |  |  | 16
| 1 | 4 | 0.07 | 0.066 | 0.075 | 0.07 | 0.073995 | 0.082 | 0.06 | 0.07 | 0.08 | 0.08 | 0.07 | 0.11 | 0.08 | 0.08 | 0.1 | 0.075 | 0.067 | 0.07 | 0.08 | 0.227 | 0.067 | 0.069451 | 0.06266 |  |  |  | 0.072862
| 1 | 5 | 0.07 | 0.073 | 0.076 | 0.07 | 0.07351 | 0.075 | 0.07 | 0.073 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.09 | 0.1 | 0.083 | 0.064 | 0.08 | 0.08 | 0.1975 | 0.067 | 0.071001 | 0.06658 |  |  |  | 106
| 1 | 6 | 0.07 | 0.174 | 0.074 | 0.07 | 0.072652 | 0.071 | 0.06 | 0.073 | 0.08 | 0.08 | 0.07 | 0.1 | 0.08 | 0.09 | 0.1 | 0.075 | 0.067 | 0.07 | 0.08 | 0.1921 | 0.067 | 0.066778 | 0.069439
| Mean |  | 0.07 | 0.10233 | 0.075 | 0.07 | 0.073545 | 0.074833 | 0.063333 | 0.072 | 0.076667 | 0.08 | 0.07 | 0.088333 | 0.08 | 0.091667 | 0.10167 | 0.075833 | 0.065667 | 0.073333 | 0.08 | 0.18943 | 0.067 | 0.069247 | 0.067016
| Median |  | 0.07 | 0.0695 | 0.075 | 0.07 | 0.073387 | 0.074 | 0.06 | 0.073 | 0.08 | 0.08 | 0.07 | 0.085 | 0.08 | 0.09 | 0.1 | 0.075 | 0.0665 | 0.07 | 0.08 | 0.1919 | 0.067 | 0.069339 | 0.067446
| Std Dev. |  | 0 | 0.056422 | 0.00089443 | 0 | 0.0010298 | 0.0037639 | 0.005164 | 0.0015492 | 0.005164 | 0 | 0 | 0.020412 | 0 | 0.01472 | 0.0040825 | 0.0037103 | 0.0033862 | 0.005164 | 0 | 0.027251 | 0 | 0.0020869 | 0.0023663
| Rel.Std.Dev. |  | 0% | 55.14% | 1.193% | 0% | 1.4% | 5.03% | 8.154% | 2.152% | 6.736% | 0% | 0% | 23.11% | 0% | 16.06% | 4.016% | 4.893% | 5.157% | 7.042% | 0% | 14.39% | 0% | 3.014% | 3.531%
| PDM3 |  | -3.928% | 40.45% | 2.934% | -3.928% | 0.9366% | 2.705% | -13.08% | -1.183% | 5.221% | 9.796% | -3.928% | 21.23% | 9.796% | 25.81% | 39.53% | 4.078% | -9.876% | 0.6465% | 9.796% | 160% | -8.046% | -4.962% | -8.023%
| Z-Score (Absolute) |  | 0.63 | 3.59 | 0.02 | 0.63 | 0.17 | 0 | 1.5 | 0.37 | 0.24 | 0.67 | 0.63 | 1.76 | 0.67 | 2.2 | 3.5 | 0.13 | 1.2 | 0.2 | 0.67 | 14.96 | 1.05 | 0.73 | 1.02
| Analytical results for Nb in OREAS 600c (Certified Value 0.99 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 25g | [6 empty] | 2
| 1 | 1 | 0.57 | 0.9 | 0.91 | 1.0285 | 0.6 | 1.7 | 1.3 | 1.12 | 0.68 | 1.1 | 0.96 | 1 | 0.91 | 1.23 | 0.92 | 1.03 | 10.9943 | 0.02 | 1.261 | 0.64 | [6 empty] | 1
| 1 | 2 | 0.53 | 0.72 | 0.86 | 1.0996 | 0.6 | 1.62 | 1.2 | 0.9 | 0.71 | 0.96 | 1 | 1.04 | 0.94 | 1.25 | 0.81 | 1.14 | 11.1085 | 0.01 | 1.3215 | 0.68 | [6 empty] | 29
| 1 | 3 | 0.54 | 0.9 | 0.87 | 1.0869 | 0.6 | 1.7 | 1.2 | 1.02 | 0.72 | 1 | 0.98 | 0.99 | 0.82 | 1.09 | 0.88 | 1.11 | 13.4792 | 0.01 | 1.3793 | 0.86 | [6 empty] | 16
| 1 | 4 | 0.53 | 1.02 | 0.82 | 1.0966 | 0.6 | 1.66 | 1.1 | 1.05 | 0.71 | 1 | 1.01 | 1.02 | 0.86 | 1.21 | 0.77 | 1.18 | 8.654 | 0.02 | 1.1737 | 1.02 | [6 empty] | 0.98628
| 1 | 5 | 0.51 | 0.87 | 0.91 | 1.0643 | 0.7 | 1.67 | 1.2 | 1.05 | 0.67 | 0.98 | 0.9 | 0.98 | 0.81 | 1.15 | 0.81 | 1.14 | 10.2512 | 0.02 | 1.298 | 0.75 | [6 empty] | 107
| 1 | 6 | 0.53 | 0.75 | 0.78 | 1.0927 | 0.6 | 1.69 | 1.4 | 1.02 | 0.71 | 0.99 | 0.96 | 0.95 | 0.85 | 1.27 | 0.8 | 1.04 | 11.9642 | 0.02 | 1.2072 | 0.85
| Mean |  | 0.535 | 0.86 | 0.85833 | 1.0781 | 0.61667 | 1.6733 | 1.2333 | 1.0267 | 0.7 | 1.005 | 0.96833 | 0.99667 | 0.865 | 1.2 | 0.83167 | 1.1067 | 11.0752 | 0.016667 | 1.2735 | 0.8
| Median |  | 0.53 | 0.885 | 0.865 | 1.0898 | 0.6 | 1.68 | 1.2 | 1.035 | 0.71 | 0.995 | 0.97 | 0.995 | 0.855 | 1.22 | 0.81 | 1.125 | 11.0514 | 0.02 | 1.2795 | 0.8
| Std Dev. |  | 0.019748 | 0.11009 | 0.051153 | 0.027377 | 0.040825 | 0.030768 | 0.10328 | 0.072019 | 0.02 | 0.048888 | 0.0392 | 0.031411 | 0.050892 | 0.067823 | 0.056362 | 0.059889 | 1.6198 | 0.005164 | 0.075644 | 0.13928
| Rel.Std.Dev. |  | 3.691% | 12.8% | 5.96% | 2.539% | 6.62% | 1.839% | 8.374% | 7.015% | 2.857% | 4.864% | 4.048% | 3.152% | 5.883% | 5.652% | 6.777% | 5.412% | 14.63% | 30.98% | 5.94% | 17.41%
| PDM3 |  | -45.76% | -12.8% | -12.97% | 9.309% | -37.48% | 69.66% | 25.05% | 4.095% | -29.03% | 1.898% | -1.819% | 1.054% | -12.3% | 21.67% | -15.68% | 12.21% | 1023% | -98.31% | 29.12% | -18.89%
| Z-Score (Absolute) |  | 1.89 | 0.56 | 0.57 | 0.33 | NA | 2.77 | 0.97 | 0.12 | 1.21 | 0.03 | 0.12 | 0 | 0.54 | 0.83 | 0.67 | 0.45 | 41.19 | 4 | 1.13 | 0.8
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 600c (Certified Value 26.6 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  |  | 1
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  |  | 2
| 1 | 1 | 2.01: 24.04; 2.03: 25.598; 2.04: 24.1327; 2.05: 26; 2.06: 28.933; 2.07: 32.11; 2.08: 21.73; 2.10: 20.0718; 2.11: 16.3; 2.12: 25.86 |  |  |  | 1
| 1 | 2 | 2.01: 25.95; 2.03: 25.126; 2.04: 25.5973; 2.05: 26.8; 2.06: 28.767; 2.07: 31.14; 2.08: 22.16; 2.10: 29.5654; 2.11: 14.2; 2.12: 27.92 |  |  |  | 30
| 1 | 3 | 2.01: 23.24; 2.03: 25.9; 2.04: 25.1224; 2.05: 26.5; 2.06: 28.37; 2.07: 31.42; 2.08: 21.07; 2.10: 31.1109; 2.11: 14.5; 2.12: 27.12 |  |  |  | 16
| 1 | 4 | 2.01: 26.31; 2.03: 25.538; 2.04: 25.2489; 2.05: 26.7; 2.06: 28.933; 2.07: 31.71; 2.08: 21.57; 2.10: 26.2671; 2.11: 17.3; 2.12: 29.12 |  |  |  | 26.5984
| 1 | 5 | 2.01: 23.94; 2.03: 26.216; 2.04: 24.9723; 2.05: 26.8; 2.06: 28.463; 2.07: 32.47; 2.08: 21.42; 2.10: 28.655; 2.11: 16.7; 2.12: 28.49 |  |  |  | 108
| 1 | 6 | 2.01: 25.4; 2.03: 25.606; 2.04: 25.4177; 2.05: 25.7; 2.06: 28.355; 2.07: 30.83; 2.08: 22.33; 2.10: 25.9175; 2.11: 20.3; 2.12: 23.99
| Mean |  | 2.01: 24.8133; 2.03: 25.664; 2.04: 25.0819; 2.05: 26.4167; 2.06: 28.6368; 2.07: 31.6133; 2.08: 21.7133; 2.10: 26.9313; 2.11: 16.55; 2.12: 27.0833
| Median |  | 2.01: 24.72; 2.03: 25.602; 2.04: 25.1856; 2.05: 26.6; 2.06: 28.615; 2.07: 31.565; 2.08: 21.65; 2.10: 27.4611; 2.11: 16.5; 2.12: 27.52
| Std Dev. |  | 2.01: 1.242; 2.03: 0.367; 2.04: 0.51401; 2.05: 0.46224; 2.06: 0.27321; 2.07: 0.61073; 2.08: 0.46916; 2.10: 3.8962; 2.11: 2.2107; 2.12: 1.8917
| Rel.Std.Dev. |  | 2.01: 5.005%; 2.03: 1.43%; 2.04: 2.049%; 2.05: 1.75%; 2.06: 0.9541%; 2.07: 1.932%; 2.08: 2.161%; 2.10: 14.47%; 2.11: 13.36%; 2.12: 6.985%
| PDM3 |  | 2.01: -6.711%; 2.03: -3.513%; 2.04: -5.702%; 2.05: -0.6834%; 2.06: 7.664%; 2.07: 18.85%; 2.08: -18.37%; 2.10: 1.251%; 2.11: -37.78%; 2.12: 1.823%
| Z-Score (Absolute) |  | 2.01: 0.73; 2.03: 0.22; 2.04: 0.57; 2.05: 0.22; 2.06: 1.54; 2.07: 3.31; 2.08: 2.57; 2.10: 0.53; 2.11: 5.64; 2.12: 0.62
| Analytical results for Ni in OREAS 600c (Certified Value 5.59 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 3
| 1 | 1 | 6.2 | 6.4 | 5.7 | 5.7 | 5.5715 | 5.1 | 5.5 | 5.2 | 6 | 12.2 | 5.2 | 5.6 | 5.4 | 5.8 | 6.1 | 5.7 | 5.8 | 5.53 | 4.8 | 6.0418 | 5.79 | 5.4621 | < 5 | 6.7 |  |  | 1
| 1 | 2 | 6.4 | 5.7 | 5.6 | 5.7 | 5.3522 | 5.3 | 5.6 | 5.7 | 6 | 5.6 | 5.4 | 5.8 | 5.4 | 5.7 | 5.9 | 5.7 | 6 | 5.3 | 5 | 7.1882 | 5.59 | 5.109 | < 5 | 7.1 |  |  | 16
| 1 | 3 | 5.5 | 6.4 | 5.5 | 5.7 | 5.5506 | 5.2 | 5.8 | 5.4 | 6 | 5.4 | 5.1 | 5.6 | 5.6 | 6.4 | 5.3 | 5.8 | 5.3 | 5.6 | 4.9 | 8.181 | 5.7 | 5.111 | < 5 | 7.3 |  |  | 16
| 1 | 4 | 6.4 | 7.1 | 5.6 | 5.5 | 5.4833 | 5.4 | 6.1 | 5.3 | 6 | 5.6 | 5.3 | 5.6 | 5.3 | 5.8 | 5.5 | 6.1 | 5.6 | 5.08 | 5 | 7.2575 | 6.02 | 5.2765 | < 5 | 7.3 |  |  | 5.5912
| 1 | 5 | 6.4 | 5.3 | 5.6 | 5.7 | 5.4843 | 5.1 | 5.7 | 5.5 | 6 | 5.8 | 5.4 | 5.6 | 5.6 | 6.3 | 5.6 | 5.9 | 5.7 | 5.09 | 5.1 | 8.9413 | 5.83 | 5.5292 | < 5 | 7.2 |  |  | 109
| 1 | 6 | 6 | 5.2 | 5.5 | 5.7 | 5.7746 | 5.1 | 5.9 | 5.5 | 6 | 5.3 | 5.7 | 5.4 | 5.4 | 5.8 | 5.6 | 5.9 | 5.9 | 5.25 | 5.1 | 7.5526 | 6.09 | 5.2535 | < 5 | 6.5
| Mean |  | 6.15 | 6.0167 | 5.5833 | 5.6667 | 5.5361 | 5.2 | 5.7667 | 5.4333 | 6 | 6.65 | 5.35 | 5.6 | 5.45 | 5.9667 | 5.6667 | 5.85 | 5.7167 | 5.3083 | 4.9833 | 7.5271 | 5.8367 | 5.2902 |  | 7.0167
| Median |  | 6.3 | 6.05 | 5.6 | 5.7 | 5.5175 | 5.15 | 5.75 | 5.45 | 6 | 5.6 | 5.35 | 5.6 | 5.4 | 5.8 | 5.6 | 5.85 | 5.75 | 5.275 | 5 | 7.4051 | 5.81 | 5.265 |  | 7.15
| Std Dev. |  | 0.35637 | 0.7414 | 0.075277 | 0.08165 | 0.13974 | 0.12649 | 0.21602 | 0.17512 | 0 | 2.7245 | 0.20736 | 0.12649 | 0.12247 | 0.30111 | 0.28752 | 0.15166 | 0.24833 | 0.21794 | 0.1169 | 0.98202 | 0.18949 | 0.17503 |  | 0.33714
| Rel.Std.Dev. |  | 5.795% | 12.32% | 1.348% | 1.441% | 2.524% | 2.433% | 3.746% | 3.223% | 0% | 40.97% | 3.876% | 2.259% | 2.247% | 5.047% | 5.074% | 2.592% | 4.344% | 4.106% | 2.346% | 13.05% | 3.247% | 3.309% |  | 4.805%
| PDM3 |  | 9.994% | 7.609% | -0.1413% | 1.349% | -0.9862% | -6.997% | 3.138% | -2.824% | 7.311% | 18.94% | -4.314% | 0.1568% | -2.526% | 6.715% | 1.349% | 4.628% | 2.243% | -5.06% | -10.87% | 34.62% | 4.39% | -5.384% |  | 25.49%
| Z-Score (Absolute) |  | 1.09 | 0.79 | 0.19 | 0 | 0.29 | 1.05 | 0.22 | 0.52 | NA | 2.21 | 0.71 | 0.15 | 0.49 | 0.67 | 0 | 0.41 | 0.11 | 0.81 | 1.54 | 4.18 | 0.38 | 0.85 | 7.12 | 3.03
| Analytical results for P in OREAS 600c (Certified Value 0.03 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 |  |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.03 | 0.029 | 0.029 | 0.03 | 0.030012 | 0.029 | 0.03 | 0.0298 | 0.0303 | 0.0309 | 0.029 | 0.03 | 0.029 | 0.031 | 0.03 | 0.03 | 0.028 | 0.03 | 0.0288 | 0.0236 | 0.028944 | 0.028981 |  |  |  |  | 1
| 1 | 2 | 0.0298 | 0.029 | 0.03 | 0.0305 | 0.030694 | 0.029 | 0.028 | 0.0301 | 0.0304 | 0.0302 | 0.03 | 0.03 | 0.029 | 0.03 | 0.03 | 0.03 | 0.028 | 0.03 | 0.029 | 0.0244 | 0.029295 | 0.028941 |  |  |  |  | 0
| 1 | 3 | 0.0298 | 0.029 | 0.029 | 0.0299 | 0.031278 | 0.028 | 0.029 | 0.0302 | 0.0304 | 0.03 | 0.03 | 0.03 | 0.029 | 0.03 | 0.029 | 0.03 | 0.028 | 0.03 | 0.0292 | 0.0236 | 0.029219 | 0.029049 |  |  |  |  | 16
| 1 | 4 | 0.0303 | 0.03 | 0.029 | 0.0286 | 0.030301 | 0.028 | 0.029 | 0.03 | 0.0304 | 0.0308 | 0.029 | 0.03 | 0.029 | 0.03 | 0.03 | 0.032 | 0.029 | 0.03 | 0.0291 | 0.0236 | 0.029554 | 0.028044 |  |  |  |  | 0.029566
| 1 | 5 | 0.0303 | 0.027 | 0.029 | 0.0309 | 0.030284 | 0.029 | 0.029 | 0.0303 | 0.0307 | 0.0311 | 0.029 | 0.03 | 0.029 | 0.032 | 0.03 | 0.03 | 0.0279 | 0.03 | 0.0288 | 0.024 | 0.029459 | 0.028993 |  |  |  |  | 110
| 1 | 6 | 0.0293 | 0.028 | 0.029 | 0.0289 | 0.030635 | 0.027 | 0.029 | 0.0298 | 0.0306 | 0.0303 | 0.03 | 0.03 | 0.029 | 0.03 | 0.03 | 0.031 | 0.027 | 0.03 | 0.0284 | 0.0231 | 0.02902 | 0.028873
| Mean |  | 0.029917 | 0.028667 | 0.029167 | 0.0298 | 0.030534 | 0.028333 | 0.029 | 0.030033 | 0.030467 | 0.03055 | 0.0295 | 0.03 | 0.029 | 0.0305 | 0.029833 | 0.0305 | 0.027983 | 0.03 | 0.028883 | 0.023717 | 0.029249 | 0.028813
| Median |  | 0.0299 | 0.029 | 0.029 | 0.02995 | 0.030468 | 0.0285 | 0.029 | 0.03005 | 0.0304 | 0.03055 | 0.0295 | 0.03 | 0.029 | 0.03 | 0.03 | 0.03 | 0.028 | 0.03 | 0.0289 | 0.0236 | 0.029257 | 0.028961
| Std Dev. |  | 0.00037639 | 0.0010328 | 0.00040825 | 0.00089443 | 0.00044231 | 0.0008165 | 0.00063246 | 0.00020656 | 0.00015055 | 0.00044159 | 0.00054772 | 0 | 0 | 0.00083666 | 0.00040825 | 0.00083666 | 0.00063377 | 0 | 0.00028577 | 0.00044008 | 0.00023896 | 0.0003815
| Rel.Std.Dev. |  | 1.258% | 3.603% | 1.4% | 3.001% | 1.449% | 2.882% | 2.181% | 0.6878% | 0.4942% | 1.445% | 1.857% | 0% | 0% | 2.743% | 1.368% | 2.743% | 2.265% | 0% | 0.9894% | 1.856% | 0.817% | 1.324%
| PDM3 |  | 1.186% | -3.041% | -1.35% | 0.7919% | 3.274% | -4.169% | -1.914% | 1.581% | 3.047% | 3.329% | -0.2228% | 1.468% | -1.914% | 3.159% | 0.9046% | 3.159% | -5.353% | 1.468% | -2.309% | -19.78% | -1.073% | -2.545%
| Z-Score (Absolute) |  | 0.25 | 0.94 | 0.46 | 0.14 | 0.84 | 1.25 | 0.62 | 0.36 | 0.78 | 0.86 | 0.14 | 0.33 | 0.62 | 0.81 | 0.17 | 0.81 | 1.59 | 0.33 | 0.73 | 5.65 | 0.38 | 0.8
| Analytical results for Pb in OREAS 600c (Certified Value 1098 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 1057 | 1056.4 | 1101.1 | 1108.4 | 1098.9224 | 1030 | 1080 | 1057.4 | 1134 | 1150 | 1118.2 | 1105 | 1130 | 1105 | 1150 | 1120 | 1085 | 1062 | 1510 | 950.5 | 1103 | 951.4999 | 1140 | 1112.5921 | 1057.212 | 941 | 1
| 1 | 2 | 1040 | 1119.72 | 1119.6 | 1102.3 | 1127.7788 | 1020 | 1040 | 1063.57 | 1117 | 1140 | 1083.9 | 1110 | 1125 | 1105 | 1125 | 1125 | 1091 | 1055 | 1500 | 960.3 | 1095 | 959.1689 | 1100 | 1110.3742 | 1083.243 | 969 | 18
| 1 | 3 | 1031 | 1041.6 | 1108.2 | 1128.1 | 1144.7503 | 1030 | 1080 | 1068.32 | 1115 | 1110 | 1091.3 | 1110 | 1125 | 1100 | 1115 | 1075 | 1104 | 1058 | 1540 | 975.5 | 1105 | 969.0489 | 1120 | 1115.1594 | 1078.682 | 957 | 16
| 1 | 4 | 1037 | 1068.48 | 1109.5 | 1119.8 | 1097.1733 | 1030 | 1080 | 1064.205 | 1128 | 1090 | 1092.6 | 1105 | 1135 | 1105 | 1125 | 1105 | 1081 | 1092 | 1540 | 956.4 | 1090 | 961.4838 | 1160 | 1121.8754 | 1008.567 | 986 | 1097.7478
| 1 | 5 | 1033 | 1092.68 | 1109.5 | 1139.1 | 1104.9539 | 1020 | 1060 | 1065.685 | 1133 | 1110 | 1122.6 | 1105 | 1135 | 1115 | 1165 | 1125 | 1080 | 1067 | 1510 | 950.7 | 1096 | 970.1429 | 1110 | 1119.4036 | 1080.404 | 964 | 111
| 1 | 6 | 1027 | 1083.18 | 1108.3 | 1133.7 | 1123.862 | 975 | 1080 | 1061.855 | 1132 | 1110 | 1081.7 | 1130 | 1130 | 1110 | 1150 | 1120 | 1077 | 1076 | 1500 | 955.6 | 1093 | 963.3331 | 1150 | 1109.6408 | 1092.312 | 1051
| Mean |  | 1037.5 | 1077.01 | 1109.3667 | 1121.9 | 1116.2401 | 1017.5 | 1070 | 1063.5058 | 1126.5 | 1118.3333 | 1098.3833 | 1110.8333 | 1130 | 1106.6667 | 1138.3333 | 1111.6667 | 1086.3333 | 1068.3333 | 1516.6667 | 958.1667 | 1097 | 962.4462 | 1130 | 1114.8409 | 1066.7367 | 978
| Median |  | 1035 | 1075.83 | 1108.9 | 1123.95 | 1114.4079 | 1025 | 1080 | 1063.8875 | 1130 | 1110 | 1091.95 | 1107.5 | 1130 | 1105 | 1137.5 | 1120 | 1083 | 1064.5 | 1510 | 956 | 1095.5 | 962.4085 | 1130 | 1113.8758 | 1079.543 | 966.5
| Std Dev. |  | 10.5783 | 27.7819 | 5.9261 | 14.4523 | 18.9467 | 21.3892 | 16.7332 | 3.6993 | 8.4083 | 22.286 | 17.6117 | 9.704 | 4.4721 | 5.164 | 19.4079 | 19.4079 | 9.913 | 13.7502 | 18.619 | 9.2653 | 5.831 | 6.8573 | 23.6643 | 4.9488 | 30.7587 | 38.6782
| Rel.Std.Dev. |  | 1.02% | 2.58% | 0.5342% | 1.288% | 1.697% | 2.102% | 1.564% | 0.3478% | 0.7464% | 1.993% | 1.603% | 0.8736% | 0.3958% | 0.4666% | 1.705% | 1.746% | 0.9125% | 1.287% | 1.228% | 0.967% | 0.5315% | 0.7125% | 2.094% | 0.4439% | 2.883% | 3.955%
| PDM3 |  | -5.488% | -1.889% | 1.058% | 2.2% | 1.685% | -7.31% | -2.528% | -3.119% | 2.619% | 1.875% | 0.0579% | 1.192% | 2.938% | 0.8125% | 3.697% | 1.268% | -1.04% | -2.68% | 38.16% | -12.72% | -0.06812% | -12.33% | 2.938% | 1.557% | -2.825% | -10.91%
| Z-Score (Absolute) |  | 1.65 | 0.65 | 0.17 | 0.49 | 0.35 | 2.16 | 0.83 | 0.99 | 0.61 | 0.4 | 0.11 | 0.21 | 0.7 | 0.11 | 0.91 | 0.23 | 0.41 | 0.87 | 10.54 | 3.67 | 0.14 | 3.57 | 0.7 | 0.31 | 0.91 | 3.17
| Analytical results for Pd in OREAS 600c (Indicative Value 11.9 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.08; 2.04: 2.26 | ppb
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 1
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 01g | 1
| 1 | 1 | 2.01: 24; 2.03: 20; 2.04: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.03: 20; 2.04: < 10 | 5
| 1 | 3 | 2.01: < 10; 2.03: 19; 2.04: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.03: 17; 2.04: < 10 | 11.9167
| 1 | 5 | 2.01: < 10; 2.03: 17; 2.04: < 10 | 11
| 1 | 6 | 2.01: < 10; 2.03: 20; 2.04: < 10
| Mean |  | 2.01: 24; 2.03: 18.8333
| Median |  | 2.01: 24; 2.03: 19.5
| Std Dev. |  | 2.03: 1.472
| Rel.Std.Dev. |  | 2.03: 7.816%
| PDM3 |  | 2.01: 101.4%; 2.03: 58.04%
| Z-Score (Absolute) |  | 2.01: 0; 2.03: 2.27; 2.04: 0.67
| Analytical results for Pr in OREAS 600c (Certified Value 6.7 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g | 3
| 1 | 1 | 2.01: 6.33; 2.03: 6.614; 2.04: 6.4316; 2.05: 6.4; 2.06: 8.7693; 2.07: 8.57; 2.08: 6.58; 2.09: 5.133; 2.10: 4.54; 2.11: 6.65 | 1
| 1 | 2 | 2.01: 7.05; 2.03: 6.599; 2.04: 7.0075; 2.05: 6.6; 2.06: 8.5564; 2.07: 8.16; 2.08: 6.59; 2.09: 7.2312; 2.10: 3.96; 2.11: 7.08 | 33
| 1 | 3 | 2.01: 6.2; 2.03: 6.76; 2.04: 6.9476; 2.05: 6.5; 2.06: 8.5252; 2.07: 8.33; 2.08: 6.26; 2.09: 7.5217; 2.10: 4.06; 2.11: 7.03 | 16
| 1 | 4 | 2.01: 6.9; 2.03: 6.71; 2.04: 6.8856; 2.05: 6.5; 2.06: 8.4365; 2.07: 8.48; 2.08: 6.42; 2.09: 6.578; 2.10: 4.87; 2.11: 7.54 | 6.6982
| 1 | 5 | 2.01: 6.02; 2.03: 6.811; 2.04: 6.8437; 2.05: 6.5; 2.06: 8.5821; 2.07: 8.63; 2.08: 6.37; 2.09: 6.9971; 2.10: 4.64; 2.11: 7.42 | 112
| 1 | 6 | 2.01: 6.52; 2.03: 6.687; 2.04: 6.8616; 2.05: 6.3; 2.06: 8.6669; 2.07: 8.26; 2.08: 6.53; 2.09: 6.1445; 2.10: 5.7; 2.11: 6.03
| Mean |  | 2.01: 6.5033; 2.03: 6.6968; 2.04: 6.8296; 2.05: 6.4667; 2.06: 8.5894; 2.07: 8.405; 2.08: 6.4583; 2.09: 6.6009; 2.10: 4.6283; 2.11: 6.9583
| Median |  | 2.01: 6.425; 2.03: 6.6985; 2.04: 6.8736; 2.05: 6.5; 2.06: 8.5693; 2.07: 8.405; 2.08: 6.475; 2.09: 6.7876; 2.10: 4.59; 2.11: 7.055
| Std Dev. |  | 2.01: 0.40302; 2.03: 0.08213; 2.04: 0.20414; 2.05: 0.10328; 2.06: 0.11577; 2.07: 0.18447; 2.08: 0.13106; 2.09: 0.86783; 2.10: 0.63; 2.11: 0.55286
| Rel.Std.Dev. |  | 2.01: 6.197%; 2.03: 1.226%; 2.04: 2.989%; 2.05: 1.597%; 2.06: 1.348%; 2.07: 2.195%; 2.08: 2.029%; 2.09: 13.15%; 2.10: 13.61%; 2.11: 7.945%
| PDM3 |  | 2.01: -2.909%; 2.03: -0.01996%; 2.04: 1.962%; 2.05: -3.456%; 2.06: 28.24%; 2.07: 25.48%; 2.08: -3.581%; 2.09: -1.452%; 2.10: -30.9%; 2.11: 3.884%
| Z-Score (Absolute) |  | 2.01: 0.53; 2.03: 0.17; 2.04: 0.65; 2.05: 0.66; 2.06: 7.02; 2.07: 6.35; 2.08: 0.69; 2.09: 0.17; 2.10: 7.31; 2.11: 1.12
| Analytical results for Pt in OREAS 600c (Indicative Value 5.5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.03; 2.03: 2.08; 2.04: 2.26 | ppb
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.03: 0.5g; 2.04: 01g | 2
| 1 | 1 | 2.01: 10; 2.03: < 10; 2.04: < 5 | 1
| 1 | 2 | 2.01: 4; 2.03: < 10; 2.04: < 5 | 5
| 1 | 3 | 2.01: 5; 2.03: < 10; 2.04: < 5 | 16
| 1 | 4 | 2.01: 7; 2.03: < 10; 2.04: < 5 | 5.5
| 1 | 5 | 2.01: 3; 2.03: < 10; 2.04: < 5 | 11
| 1 | 6 | 2.01: 4; 2.03: < 10; 2.04: < 5
| Mean |  | 2.01: 5.5
| Median |  | 2.01: 4.5
| Std Dev. |  | 2.01: 2.5884
| Rel.Std.Dev. |  | 2.01: 47.06%
PDM3 | bar: 2.01
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0; 2.04: 3.37
| Analytical results for Rb in OREAS 600c (Certified Value 14.2 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [6 empty] | 1
| 1 | 1 | 15 | 14.7 | 13.91 | 13.9528 | 15.7 | 13.1 | 15.1 | 16.9 | 20.81 | 12.3 | 13.9 | 14.6 | 14.6 | 13.8 | 13.73 | 11.7 | 15.1 | 11.1 | 13.1205 | 15.47 | [6 empty] | 1
| 1 | 2 | 14 | 14.3 | 13.8 | 14.5021 | 16.3 | 12.9 | 14.9 | 16.6 | 19.49 | 12.5 | 13.3 | 14.4 | 14.2 | 14.1 | 14.41 | 11.1 | 16.6 | 10.3 | 14.1754 | 16.34 | [6 empty] | 34
| 1 | 3 | 15.4 | 14.3 | 14.49 | 14.6987 | 15.9 | 13.3 | 14.4 | 16.8 | 20.24 | 12.2 | 13.7 | 14.2 | 15.8 | 13 | 13.81 | 11.9 | 16.3 | 10.5 | 13.9555 | 16.87 | [6 empty] | 16
| 1 | 4 | 14.6 | 16.5 | 14.2 | 14.6981 | 16.3 | 13.1 | 14.5 | 17.4 | 20.64 | 12.1 | 13.5 | 15 | 14.4 | 13.2 | 14.64 | 10.6 | 16.5 | 11.5 | 13.1367 | 17.45 | [6 empty] | 14.1589
| 1 | 5 | 13.8 | 14.4 | 14.16 | 14.1564 | 15.8 | 13.2 | 15.2 | 17.4 | 20.94 | 11.9 | 13.6 | 14.8 | 14.8 | 13.4 | 13.64 | 11.2 | 16 | 11.1 | 14.3166 | 16.78 | [6 empty] | 113
| 1 | 6 | 14.2 | 14.2 | 14.36 | 14.3918 | 15.5 | 13.1 | 14.5 | 17.1 | 19.63 | 12.8 | 13.2 | 14.5 | 14.6 | 13.2 | 14.11 | 11.6 | 15.1 | 12.4 | 13.2273 | 14.41
| Mean |  | 14.5 | 14.7333 | 14.1533 | 14.4 | 15.9167 | 13.1167 | 14.7667 | 17.0333 | 20.2917 | 12.3 | 13.5333 | 14.5833 | 14.7333 | 13.45 | 14.0567 | 11.35 | 15.9333 | 11.15 | 13.6553 | 16.22
| Median |  | 14.4 | 14.35 | 14.18 | 14.447 | 15.85 | 13.1 | 14.7 | 17 | 20.44 | 12.25 | 13.55 | 14.55 | 14.6 | 13.3 | 13.96 | 11.4 | 16.15 | 11.1 | 13.5914 | 16.56
| Std Dev. |  | 0.61644 | 0.88242 | 0.26166 | 0.29927 | 0.32506 | 0.13292 | 0.34448 | 0.3266 | 0.61538 | 0.31623 | 0.2582 | 0.28577 | 0.56095 | 0.41833 | 0.40228 | 0.47645 | 0.67725 | 0.75299 | 0.55427 | 1.1047
| Rel.Std.Dev. |  | 4.251% | 5.989% | 1.849% | 2.078% | 2.042% | 1.013% | 2.333% | 1.917% | 3.033% | 2.571% | 1.908% | 1.96% | 3.807% | 3.11% | 2.862% | 4.198% | 4.251% | 6.753% | 4.059% | 6.811%
| PDM3 |  | 2.409% | 4.057% | -0.03914% | 1.703% | 12.41% | -7.361% | 4.293% | 20.3% | 43.31% | -13.13% | -4.418% | 2.998% | 4.057% | -5.007% | -0.7219% | -19.84% | 12.53% | -21.25% | -3.556% | 14.56%
| Z-Score (Absolute) |  | 0.04 | 0.2 | 0.21 | 0.04 | 1.03 | 0.94 | 0.22 | 1.82 | 4.11 | 1.51 | 0.64 | 0.09 | 0.2 | 0.7 | 0.28 | 2.18 | 1.04 | 2.32 | 0.56 | 1.25
| Analytical results for Re in OREAS 600c (Certified Value < 0.001 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.26 | 2.28 | [10 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 01g | 25g | [10 empty] | 3
| 1 | 1 | 0.02 | 0.002 | < 0.01 | 0.001 | < 0.001 | < 0.05 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | < 0.001 | [10 empty] | 1
| 1 | 2 | 0.031 | < 0.001 | < 0.01 | 0.002 | < 0.001 | < 0.05 | < 0.001 | 0.001 | < 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | 0.002 | [10 empty] | 35
| 1 | 3 | 0.025 | < 0.001 | < 0.01 | 0.002 | 0.001 | < 0.05 | 0.001 | 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | < 0.001 | [10 empty] | 16
| 1 | 4 | 0.044 | 0.006 | < 0.01 | 0.002 | 0.001 | < 0.05 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | < 0.001 | [10 empty] | < 0.001
| 1 | 5 | 0.02 | < 0.001 | < 0.01 | 0.002 | 0.001 | < 0.05 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | < 0.001 | [10 empty] | 114
| 1 | 6 | 0.02 | 0.001 | < 0.01 | 0.001 | < 0.001 | < 0.05 | < 0.001 | 0.001 | 0.001 | 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.002 | < 0.05 | < 0.001
| Mean |  | 2.01: 0.026667; 2.03: 0.003; 2.05: 0.0016667; 2.06: 0.001; 2.08: 0.001; 2.09: 0.001; 2.10: 0.001; 2.11: 0.001; 2.12: 0.001; 2.17: 0.002
| Median |  | 2.01: 0.0225; 2.03: 0.002; 2.05: 0.002; 2.06: 0.001; 2.08: 0.001; 2.09: 0.001; 2.10: 0.001; 2.11: 0.001; 2.12: 0.001; 2.17: 0.002
| Std Dev. |  | 2.01: 0.0095429; 2.03: 0.0026458; 2.05: 0.0005164; 2.06: 0; 2.08: 0; 2.09: 0; 2.10: 0; 2.11: 0; 2.12: 0
| Rel.Std.Dev. |  | 2.01: 35.79%; 2.03: 88.19%; 2.05: 30.98%; 2.06: 0%; 2.08: 0%; 2.09: 0%; 2.10: 0%; 2.11: 0%; 2.12: 0%
PDM3
| Z-Score (Absolute) |  | 41.54 | 1.21 | 6.47 | 1.08 | 0.4 | 38.84 | 0.54 | 0 | 0.27 | 0 | 0.4 | 38.84 | 0.81 | 0 | 38.84 | 0.4
| Analytical results for S in OREAS 600c (Certified Value 0.24 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 3
| 1 | 1 | 0.26 | 0.22 | 0.25 | 0.2345 | 0.2408 | 0.255 | 0.25 | 0.22495 | 0.23 | 0.24 | 0.2407 | 0.23 | 0.25 | 0.24 | 0.26 | 0.26 | 0.251 | 0.236 | 0.24 | 0.12 | 0.24 | 0.2138 | 0.256 | 0.23027 | 0.221 | 0.22 | 1
| 1 | 2 | 0.25 | 0.24 | 0.25 | 0.2356 | 0.24534 | 0.257 | 0.24 | 0.22498 | 0.24 | 0.24 | 0.2356 | 0.24 | 0.25 | 0.24 | 0.26 | 0.26 | 0.2457 | 0.239 | 0.23 | 0.12 | 0.24 | 0.2458 | 0.256 | 0.23527 | 0.217 | 0.24 | 11
| 1 | 3 | 0.25 | 0.22 | 0.24 | 0.2338 | 0.24469 | 0.253 | 0.25 | 0.22696 | 0.23 | 0.24 | 0.2323 | 0.24 | 0.25 | 0.24 | 0.24 | 0.25 | 0.2485 | 0.236 | 0.24 | 0.11 | 0.25 | 0.2333 | 0.25 | 0.23434 | 0.218 | 0.25 | 16
| 1 | 4 | 0.25 | 0.23 | 0.25 | 0.2267 | 0.24463 | 0.253 | 0.24 | 0.22755 | 0.24 | 0.23 | 0.2368 | 0.23 | 0.25 | 0.24 | 0.26 | 0.25 | 0.2524 | 0.244 | 0.24 | 0.12 | 0.24 | 0.2266 | 0.257 | 0.23112 | 0.213 | 0.23 | 0.24046
| 1 | 5 | 0.25 | 0.23 | 0.24 | 0.2392 | 0.24241 | 0.249 | 0.25 | 0.22401 | 0.24 | 0.24 | 0.2392 | 0.23 | 0.25 | 0.24 | 0.27 | 0.26 | 0.2448 | 0.238 | 0.236 | 0.11 | 0.24 | 0.2976 | 0.252 | 0.23467 | 0.214 | 0.23 | 115
| 1 | 6 | 0.25 | 0.23 | 0.24 | 0.2275 | 0.24178 | 0.242 | 0.25 | 0.22124 | 0.23 | 0.24 | 0.2353 | 0.24 | 0.25 | 0.24 | 0.26 | 0.26 | 0.2515 | 0.237 | 0.23 | 0.1 | 0.24 | 0.2292 | 0.255 | 0.23319 | 0.217 | 0.26
| Mean |  | 0.25167 | 0.22833 | 0.245 | 0.23288 | 0.24328 | 0.2515 | 0.24667 | 0.22495 | 0.235 | 0.23833 | 0.23665 | 0.235 | 0.25 | 0.24 | 0.25833 | 0.25667 | 0.24898 | 0.23833 | 0.236 | 0.11333 | 0.24167 | 0.24105 | 0.25433 | 0.23314 | 0.21667 | 0.23833
| Median |  | 0.25 | 0.23 | 0.245 | 0.23415 | 0.24352 | 0.253 | 0.25 | 0.22496 | 0.235 | 0.24 | 0.2362 | 0.235 | 0.25 | 0.24 | 0.26 | 0.26 | 0.24975 | 0.2375 | 0.238 | 0.115 | 0.24 | 0.23125 | 0.2555 | 0.23377 | 0.217 | 0.235
| Std Dev. |  | 0.0040825 | 0.0075277 | 0.0054772 | 0.0048569 | 0.0018541 | 0.0053572 | 0.005164 | 0.0022555 | 0.0054772 | 0.0040825 | 0.0029925 | 0.0054772 | 0 | 0 | 0.0098319 | 0.005164 | 0.0031808 | 0.0030111 | 0.004899 | 0.008165 | 0.0040825 | 0.029571 | 0.0027325 | 0.0020332 | 0.0028752 | 0.01472
| Rel.Std.Dev. |  | 1.622% | 3.297% | 2.236% | 2.086% | 0.7621% | 2.13% | 2.094% | 1.003% | 2.331% | 1.713% | 1.265% | 2.331% | 0% | 0% | 3.806% | 2.012% | 1.278% | 1.263% | 2.076% | 7.204% | 1.689% | 12.27% | 1.074% | 0.8721% | 1.327% | 6.176%
| PDM3 |  | 4.661% | -5.042% | 1.889% | -3.15% | 1.171% | 4.592% | 2.582% | -6.45% | -2.27% | -0.8837% | -1.584% | -2.27% | 3.968% | -0.1906% | 7.434% | 6.741% | 3.545% | -0.8837% | -1.854% | -52.87% | 0.5025% | 0.2461% | 5.77% | -3.042% | -9.894% | -0.8837%
| Z-Score (Absolute) |  | 1.37 | 1.19 | 0.64 | 0.69 | 0.45 | 1.35 | 0.82 | 1.56 | 0.46 | 0.09 | 0.28 | 0.46 | 1.19 | 0.09 | 2.1 | 1.92 | 1.08 | 0.09 | 0.35 | 13.79 | 0.27 | 0.21 | 1.66 | 0.66 | 2.47 | 0.09
| Analytical results for Sb in OREAS 600c (Certified Value 56 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  | 1
| 1 | 1 | 55 | 65.85 | 53.1 | 58.77 | 54.8332 | 52.5 | 56.5 | 52.5 | 48.4 | 47.83 | 54.9 | 61.8 | 57.9 | 60.7 | 61.4 | 55.97 | 58.16 | 58.5 | 58.6 | 52.7 | 59.4643 | 56.5 | 53.1575 | 50.54 | 49.91 |  | 1
| 1 | 2 | 55 | 57.93 | 53.5 | 59.37 | 56.1215 | 55.2 | 56.9 | 54.41 | 48.5 | 44.68 | 55.8 | 59.1 | 57.7 | 58.5 | 60.8 | 50.91 | 57.69 | 59.7 | 58.2 | 56.83 | 53.4182 | 41.1 | 53.44 | 49.56 | 52.07 |  | 36
| 1 | 3 | 55 | 62.73 | 52.9 | 59.35 | 57.0763 | 52.5 | 57.5 | 53.59 | 47.2 | 46.89 | 53.9 | 59.2 | 56.5 | 53.9 | 58.7 | 47.84 | 58 | 59 | 57.8 | 57.48 | 56.7059 | 23.2 | 53.186 | 50.29 | 55.71 |  | 16
| 1 | 4 | 55 | 67.94 | 52.9 | 59.12 | 54.9541 | 52 | 56.8 | 53.8 | 48.4 | 47.53 | 56.7 | 59.5 | 56.5 | 59.5 | 57.8 | 50.04 | 60.29 | 58.9 | 59.3 | 56.57 | 65.1152 | 59.3 | 52.739 | 49.18 | 58.57 |  | 55.6883
| 1 | 5 | 54 | 60.46 | 52.5 | 60.01 | 55.536 | 54.1 | 57.4 | 53.16 | 48.5 | 48.75 | 55 | 60.8 | 55.3 | 59.6 | 59.9 | 38.91 | 57.84 | 57.6 | 58.3 | 54.6 | 63.8105 | 57.7 | 52.8905 | 47.97 | 54.62 |  | 116
| 1 | 6 | 54 | 56.48 | 52.7 | 60.24 | 57.1718 | 52.4 | 57.3 | 54.2 | 48.5 | 46.05 | 56.7 | 59.2 | 57.6 | 60.3 | 59.8 | 53.56 | 57.92 | 57.6 | 59.7 | 53.76 | 60.0813 | 61.3 | 53.5337 | 49.66 | 53.15
| Mean |  | 54.6667 | 61.8983 | 52.9333 | 59.4767 | 55.9488 | 53.1167 | 57.0667 | 53.61 | 48.25 | 46.955 | 55.5 | 59.9333 | 56.9167 | 58.75 | 59.7333 | 49.5383 | 58.3167 | 58.55 | 58.65 | 55.3233 | 59.7659 | 49.85 | 53.1578 | 49.5333 | 54.005
| Median |  | 55 | 61.595 | 52.9 | 59.36 | 55.8288 | 52.5 | 57.1 | 53.695 | 48.45 | 47.21 | 55.4 | 59.35 | 57.05 | 59.55 | 59.85 | 50.475 | 57.96 | 58.7 | 58.45 | 55.585 | 59.7728 | 57.1 | 53.1718 | 49.61 | 53.885
| Std Dev. |  | 0.5164 | 4.4718 | 0.34448 | 0.55157 | 1.0199 | 1.2513 | 0.39328 | 0.70131 | 0.51672 | 1.4365 | 1.1082 | 1.113 | 1.0008 | 2.493 | 1.3231 | 5.9248 | 0.9794 | 0.83126 | 0.71764 | 1.9144 | 4.3548 | 14.92 | 0.30612 | 0.9134 | 3.0123
| Rel.Std.Dev. |  | 0.9446% | 7.224% | 0.6508% | 0.9274% | 1.823% | 2.356% | 0.6892% | 1.308% | 1.071% | 3.059% | 1.997% | 1.857% | 1.758% | 4.243% | 2.215% | 11.96% | 1.679% | 1.42% | 1.224% | 3.46% | 7.286% | 29.93% | 0.5759% | 1.844% | 5.578%
| PDM3 |  | -1.835% | 11.15% | -4.947% | 6.803% | 0.4679% | -4.618% | 2.475% | -3.732% | -13.36% | -15.68% | -0.3381% | 7.623% | 2.206% | 5.498% | 7.264% | -11.04% | 4.72% | 5.139% | 5.318% | -0.6553% | 7.322% | -10.48% | -4.544% | -11.05% | -3.023%
| Z-Score (Absolute) |  | 0.18 | 1.41 | 0.57 | 0.88 | 0.1 | 0.53 | 0.35 | 0.42 | 1.6 | 1.89 | 0 | 0.98 | 0.31 | 0.72 | 0.94 | 1.32 | 0.62 | 0.67 | 0.7 | 0.04 | 0.94 | 1.25 | 0.52 | 1.32 | 0.33
| Analytical results for Sc in OREAS 600c (Certified Value 1.57 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  | 2
| 1 | 1 | 2.1 | 1.6 | 1.4 | 1.4 | 1.531 | 1.6 | 1.6 | 1.9867 | 1.53 | 2 | 1.7 | 1.5 | 1.8 | 1.7 | 1.6 | 1.7 | 2 | 1.3 | 1.3 | 1.6 | 1.3567 | 1.1 | 1.4073 | 1.5 |  |  | 1
| 1 | 2 | 2.2 | 1.5 | 1.4 | 1.4 | 1.5747 | 1.7 | 1.6 | 1.9599 | 1.43 | 2 | 1.8 | 1.5 | 1.7 | 1.7 | 1.6 | 1.7 | 2.7 | 1.4 | 1.4 | 1.6 | 1.7752 | 1 | 1.4242 | 1.9 |  |  | 37
| 1 | 3 | 2 | 1.6 | 1.3 | 1.4 | 1.5323 | 1.7 | 1.6 | 1.6989 | 1.47 | 2 | 1.8 | 1.5 | 1.7 | 1.7 | 1.6 | 1.6 | 2.4 | 1.44 | 1.3 | 1.8 | 1.8213 | 1 | 1.4228 | 1.5 |  |  | 16
| 1 | 4 | 2.1 | 1.7 | 1.4 | 1.4 | 1.5514 | 1.9 | 1.5 | 1.6181 | 1.57 | 2 | 1.8 | 1.4 | 1.7 | 1.7 | 1.6 | 1.6 | 2.4 | 1.39 | 1.3 | 1.7 | 1.6183 | 1.1 | 1.3998 | 1.9 |  |  | 1.5699
| 1 | 5 | 1.9 | 1.5 | 1.4 | 1.5 | 1.5307 | 1.5 | 1.6 | 1.7946 | 1.5 | 2 | 1.8 | 1.4 | 1.7 | 1.7 | 1.6 | 1.7 | 2.3 | 1.3 | 1.3 | 1.6 | 1.6278 | 1.1 | 1.4249 | 1.8 |  |  | 117
| 1 | 6 | 2.1 | 1.5 | 1.3 | 1.5 | 1.5416 | 1.5 | 1.5 | 1.6336 | 1.63 | 2 | 1.8 | 1.6 | 1.6 | 1.7 | 1.6 | 1.6 | 2.3 | 1.46 | 1.3 | 1.6 | 1.5415 | 1.2 | 1.3911 | 1.4
| Mean |  | 2.0667 | 1.5667 | 1.3667 | 1.4333 | 1.5436 | 1.65 | 1.5667 | 1.782 | 1.5217 | 2 | 1.7833 | 1.4833 | 1.7 | 1.7 | 1.6 | 1.65 | 2.35 | 1.3817 | 1.3167 | 1.65 | 1.6235 | 1.0833 | 1.4117 | 1.6667
| Median |  | 2.1 | 1.55 | 1.4 | 1.4 | 1.5369 | 1.65 | 1.6 | 1.7468 | 1.515 | 2 | 1.8 | 1.5 | 1.7 | 1.7 | 1.6 | 1.65 | 2.35 | 1.395 | 1.3 | 1.6 | 1.6231 | 1.1 | 1.4151 | 1.65
| Std Dev. |  | 0.10328 | 0.08165 | 0.05164 | 0.05164 | 0.017225 | 0.15166 | 0.05164 | 0.16094 | 0.071671 | 0 | 0.040825 | 0.075277 | 0.063246 | 0 | 2.4324e-16 | 0.054772 | 0.22583 | 0.06824 | 0.040825 | 0.083666 | 0.16739 | 0.075277 | 0.014423 | 0.22509
| Rel.Std.Dev. |  | 4.997% | 5.212% | 3.779% | 3.603% | 1.116% | 9.191% | 3.296% | 9.031% | 4.71% | 0% | 2.289% | 5.075% | 3.72% | 0% | 1.52e-14% | 3.32% | 9.61% | 4.939% | 3.101% | 5.071% | 10.31% | 6.949% | 1.022% | 13.51%
| PDM3 |  | 31.65% | -0.2041% | -12.94% | -8.697% | -1.673% | 5.104% | -0.2041% | 13.51% | -3.071% | 27.4% | 13.6% | -5.512% | 8.289% | 8.289% | 1.919% | 5.104% | 49.69% | -11.99% | -16.13% | 5.104% | 3.414% | -30.99% | -10.08% | 6.166%
| Z-Score (Absolute) |  | 3.15 | 0.22 | 1.57 | 1.12 | 0.38 | 0.34 | 0.22 | 1.23 | 0.53 | NA | 1.24 | 0.79 | 0.67 | 0.67 | 0 | 0.34 | 5.06 | 1.47 | 1.91 | 0.34 | 0.16 | 3.48 | 1.27 | 0.45
Results from laboratory 2.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 600c (Certified Value 0.33 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 15g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 25g | [6 empty] | 2
| 1 | 1 | < 1 | < 0.1 | < 0.5 | < 1.5 | 0.5 | < 0.2 | < 1 | 0.2 | 0.4 | 0.4 | 0.3 | 0.3 | 0.23 | < 1 | 1.2 | 0.4 | < 0.1 | < 1 | < 10 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 1 | 0.2 | < 0.5 | < 1.5 | 0.5 | 0.2 | < 1 | 0.2 | 0.4 | < 0.2 | 0.3 | 0.3 | 0.19 | < 1 | 1.3 | 0.4 | 0.2 | < 1 | < 10 | < 0.5 | [6 empty] | 38
| 1 | 3 | < 1 | 0.3 | < 0.5 | < 1.5 | 0.6 | < 0.2 | < 1 | 0.3 | 0.5 | 0.3 | 0.2 | 0.3 | 0.18 | < 1 | 0.5 | 0.5 | 0.4 | < 1 | < 10 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 1 | 0.4 | < 0.5 | < 1.5 | 0.5 | 0.2 | < 1 | 0.3 | 0.4 | 0.3 | 0.2 | 0.2 | 0.21 | < 1 | 1.1 | 0.5 | < 0.1 | < 1 | < 10 | < 0.5 | [6 empty] | 0.32606
| 1 | 5 | < 1 | 0.3 | < 0.5 | < 1.5 | 0.5 | 0.2 | < 1 | 0.3 | 0.4 | 0.4 | 0.2 | 0.2 | 0.15 | < 1 | 0.9 | 0.5 | < 0.1 | < 1 | < 10 | < 0.5 | [6 empty] | 118
| 1 | 6 | < 1 | 0.5 | < 0.5 | < 1.5 | 0.5 | 0.2 | < 1 | 0.6 | 0.5 | 0.4 | 0.2 | 0.4 | 0.2 | < 1 | 0.9 | 0.5 | < 0.1 | < 1 | < 10 | < 0.5
| Mean |  | 2.03: 0.34; 2.06: 0.51667; 2.07: 0.2; 2.09: 0.31667; 2.10: 0.43333; 2.11: 0.36; 2.12: 0.23333; 2.13: 0.28333; 2.14: 0.19333; 2.16: 0.98333; 2.17: 0.46667; 2.19: 0.3
| Median |  | 2.03: 0.3; 2.06: 0.5; 2.07: 0.2; 2.09: 0.3; 2.10: 0.4; 2.11: 0.4; 2.12: 0.2; 2.13: 0.3; 2.14: 0.195; 2.16: 1; 2.17: 0.5; 2.19: 0.3
| Std Dev. |  | 2.03: 0.11402; 2.06: 0.040825; 2.07: 0; 2.09: 0.1472; 2.10: 0.05164; 2.11: 0.054772; 2.12: 0.05164; 2.13: 0.075277; 2.14: 0.027325; 2.16: 0.28577; 2.17: 0.05164; 2.19: 0.14142
| Rel.Std.Dev. |  | 2.03: 33.53%; 2.06: 7.902%; 2.07: 0%; 2.09: 46.48%; 2.10: 11.92%; 2.11: 15.21%; 2.12: 22.13%; 2.13: 26.57%; 2.14: 14.13%; 2.16: 29.06%; 2.17: 11.07%; 2.19: 47.14%
| PDM3 |  | 2.03: 4.275%; 2.06: 58.46%; 2.07: -38.66%; 2.09: -2.881%; 2.10: 32.9%; 2.11: 10.41%; 2.12: -28.44%; 2.13: -13.1%; 2.14: -40.71%; 2.16: 201.6%; 2.17: 43.12%; 2.19: -7.993%
| Z-Score (Absolute) |  | 0.67 | 0.45 | 0.67 | 2.02 | 0.76 | 1.12 | 0.67 | 0.31 | 0.31 | 0.31 | 0.76 | 0.49 | 0.98 | 0.67 | 3.28 | 0.49 | 1.3 | 0.67 | 24.95 | 0.67
| Analytical results for Sm in OREAS 600c (Certified Value 4.8 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.09: 2.23; 2.10: 2.25; 2.11: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.09: 0.5g; 2.10: 0.5g; 2.11: 25g |  |  | 2
| 1 | 1 | 2.01: 4.62; 2.03: 4.737; 2.04: 4.3237; 2.05: 4.1; 2.06: 5.5414; 2.07: 6.01; 2.08: 4.08; 2.09: 3.6419; 2.10: 3.26; 2.11: 5.68 |  |  | 1
| 1 | 2 | 2.01: 4.86; 2.03: 4.76; 2.04: 4.7546; 2.05: 4.4; 2.06: 5.9722; 2.07: 5.7; 2.08: 4.09; 2.09: 5.352; 2.10: 2.83; 2.11: 5.85 |  |  | 22
| 1 | 3 | 2.01: 4.47; 2.03: 4.809; 2.04: 4.8107; 2.05: 4.2; 2.06: 5.7531; 2.07: 5.76; 2.08: 3.99; 2.09: 5.6521; 2.10: 2.92; 2.11: 5.85 |  |  | 16
| 1 | 4 | 2.01: 4.93; 2.03: 4.875; 2.04: 4.7119; 2.05: 4.4; 2.06: 5.5795; 2.07: 5.74; 2.08: 3.97; 2.09: 4.726; 2.10: 3.44; 2.11: 6.3 |  |  | 4.8007
| 1 | 5 | 2.01: 4.61; 2.03: 4.827; 2.04: 4.608; 2.05: 4.4; 2.06: 5.7363; 2.07: 5.83; 2.08: 4.07; 2.09: 5.2023; 2.10: 3.34; 2.11: 6.12 |  |  | 119
| 1 | 6 | 2.01: 4.83; 2.03: 4.791; 2.04: 4.7641; 2.05: 4.3; 2.06: 5.1093; 2.07: 5.61; 2.08: 4.21; 2.09: 4.7463; 2.10: 4; 2.11: 5.08
| Mean |  | 2.01: 4.72; 2.03: 4.7998; 2.04: 4.6622; 2.05: 4.3; 2.06: 5.6153; 2.07: 5.775; 2.08: 4.0683; 2.09: 4.8868; 2.10: 3.2983; 2.11: 5.8133
| Median |  | 2.01: 4.725; 2.03: 4.8; 2.04: 4.7332; 2.05: 4.35; 2.06: 5.6579; 2.07: 5.75; 2.08: 4.075; 2.09: 4.9743; 2.10: 3.3; 2.11: 5.85
| Std Dev. |  | 2.01: 0.17911; 2.03: 0.049163; 2.04: 0.17941; 2.05: 0.12649; 2.06: 0.29113; 2.07: 0.13605; 2.08: 0.085421; 2.09: 0.70668; 2.10: 0.41907; 2.11: 0.42169
| Rel.Std.Dev. |  | 2.01: 3.795%; 2.03: 1.024%; 2.04: 3.848%; 2.05: 2.942%; 2.06: 5.185%; 2.07: 2.356%; 2.08: 2.1%; 2.09: 14.46%; 2.10: 12.71%; 2.11: 7.254%
| PDM3 |  | 2.01: -1.68%; 2.03: -0.01751%; 2.04: -2.885%; 2.05: -10.43%; 2.06: 16.97%; 2.07: 20.3%; 2.08: -15.25%; 2.09: 1.793%; 2.10: -31.29%; 2.11: 21.09%
| Z-Score (Absolute) |  | 2.01: 0.05; 2.03: 0.05; 2.04: 0.11; 2.05: 0.54; 2.06: 1; 2.07: 1.19; 2.08: 0.81; 2.09: 0.15; 2.10: 1.71; 2.11: 1.23
| Analytical results for Sn in OREAS 600c (Certified Value 1.2 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 2
| 1 | 1 | 1.2 | 1.5 | 1.14 | 1.1964 | 1.44 | 1 | 1.07 | 1.3 | 1.61 | 1.1 | 1.2 | 1.2 | 1.3 | 1.2 | 1 | 0.6 | < 5 | 1.1 | 0.65 | 1.1199 | < 10 | 1.26 |  |  |  |  | 1
| 1 | 2 | 1.1 | 1.1 | 1.12 | 1.2331 | 1.49 | 1 | 1.23 | 1.3 | 1.53 | 1.1 | 1.2 | 1.2 | 1.2 | 1.2 | 1 | 1 | < 5 | 1.1 | 0.59 | 1.1548 | < 10 | 1.34 |  |  |  |  | 13
| 1 | 3 | 1.2 | 1.5 | 1.17 | 1.2803 | 1.44 | 1 | 1.13 | 1.4 | 1.58 | 1 | 1.2 | 1.1 | 1.3 | 1.1 | 0.9 | 0.6 | < 5 | 1.2 | 0.56 | 1.1479 | < 10 | 1.33 |  |  |  |  | 16
| 1 | 4 | 1 | 1.6 | 1.1 | 1.2521 | 1.45 | 1 | 1.2 | 1.4 | 1.6 | 1 | 1.2 | 1.2 | 1.2 | 1.1 | 1.1 | 0.7 | < 5 | 1.1 | 0.7 | 1.1084 | < 10 | 1.42 |  |  |  |  | 1.2
| 1 | 5 | 1.2 | 1.2 | 1.15 | 1.22 | 1.44 | 1 | 1.2 | 1.4 | 1.65 | 1 | 1.2 | 1.2 | 1.3 | 1.2 | 1 | 0.6 | < 5 | 1.1 | 0.64 | 1.1815 | < 10 | 1.32 |  |  |  |  | 120
| 1 | 6 | 1.3 | 1 | 1.2 | 1.2226 | 1.41 | 1 | 1.23 | 1.4 | 1.55 | 1.1 | 1.1 | 1.2 | 1.3 | 1.2 | 1 | 0.7 | < 5 | 1.1 | 0.75 | 1.1533 | < 10 | 1.14
| Mean |  | 1.1667 | 1.3167 | 1.1467 | 1.2341 | 1.445 | 1 | 1.1767 | 1.3667 | 1.5867 | 1.05 | 1.1833 | 1.1833 | 1.2667 | 1.1667 | 1 | 0.7 |  | 1.1167 | 0.64833 | 1.1443 |  | 1.3017
| Median |  | 1.2 | 1.35 | 1.145 | 1.2278 | 1.44 | 1 | 1.2 | 1.4 | 1.59 | 1.05 | 1.2 | 1.2 | 1.3 | 1.2 | 1 | 0.65 |  | 1.1 | 0.645 | 1.1506 |  | 1.325
| Std Dev. |  | 0.10328 | 0.24833 | 0.03559 | 0.02902 | 0.025884 | 0 | 0.06377 | 0.05164 | 0.043205 | 0.054772 | 0.040825 | 0.040825 | 0.05164 | 0.05164 | 0.063246 | 0.15492 |  | 0.040825 | 0.06969 | 0.026337 |  | 0.094322
| Rel.Std.Dev. |  | 8.853% | 18.86% | 3.104% | 2.352% | 1.791% | 0% | 5.42% | 3.779% | 2.723% | 5.216% | 3.45% | 3.45% | 4.077% | 4.426% | 6.325% | 22.13% |  | 3.656% | 10.75% | 2.302% |  | 7.246%
| PDM3 |  | -2.774% | 9.726% | -4.441% | 2.842% | 20.42% | -16.66% | -1.941% | 13.89% | 32.23% | -12.5% | -1.386% | -1.386% | 5.559% | -2.774% | -16.66% | -41.66% |  | -6.941% | -45.97% | -4.64% |  | 8.476%
| Z-Score (Absolute) |  | 0.09 | 0.76 | 0.21 | 0.29 | 1.49 | NA | 0.04 | 1.04 | 2.3 | 0.76 | 0 | 0 | 0.47 | 0.09 | 1.04 | 2.75 | 7.5 | 0.38 | 3.05 | 0.22 | 21.75 | 0.67
Results from laboratory 2.08 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sr in OREAS 600c (Certified Value 21.6 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 2
| 1 | 1 | 27.2 | 20.6 | 21 | 22.47 | 19.0403 | 20.7 | 20.4 | 20.6 | 22.3 | 22.25 | 20.5 | 22.7 | 22.8 | 22.9 | 23.3 | 22.8 | 32 | 20.8 | 21.2 | 22.0723 | 16 | 20.4175 | 20.57 | 23.8 |  |  | 1
| 1 | 2 | 25.9 | 27.7 | 21 | 22.27 | 20.0546 | 20.8 | 20.4 | 21.3 | 22.7 | 21.29 | 21.3 | 22.5 | 22.5 | 22.2 | 23 | 22.9 | 31 | 20.8 | 22.5 | 22.6824 | 16 | 21.8804 | 20.34 | 25.5 |  |  | 14
| 1 | 3 | 27.1 | 20.4 | 21 | 22.41 | 19.5525 | 19.9 | 20.7 | 20.8 | 21.8 | 21.81 | 20.1 | 22.2 | 22.4 | 20.4 | 21.2 | 22.8 | 32 | 20 | 22.5 | 22.2053 | 16 | 21.6078 | 20.35 | 25.5 |  |  | 16
| 1 | 4 | 26.2 | 21.3 | 21 | 22.33 | 19.6777 | 21.5 | 20.4 | 20.9 | 23.9 | 21.45 | 20.4 | 22.3 | 22.4 | 22.2 | 22.1 | 23.4 | 32 | 20.4 | 22.5 | 24.0616 | 16 | 20.875 | 20.15 | 26.5 |  |  | 21.6444
| 1 | 5 | 26.9 | 23.5 | 21 | 22.46 | 18.9265 | 20.4 | 20.7 | 21.2 | 22.4 | 22.3 | 20 | 22.2 | 22.5 | 22.9 | 22.5 | 23.3 | 31.4 | 20.3 | 21.9 | 21.5518 | 16 | 22.2565 | 20.26 | 25.4 |  |  | 121
| 1 | 6 | 26.8 | 27.5 | 21 | 22.3 | 20.3308 | 19.6 | 20.5 | 20.7 | 22.4 | 21.04 | 22.4 | 22 | 22.5 | 22.4 | 22.2 | 23 | 31 | 20.1 | 21.6 | 21.7918 | 16 | 20.8843 | 20.5 | 23.7
| Mean |  | 26.6833 | 23.5 | 21 | 22.3733 | 19.5971 | 20.4833 | 20.5167 | 20.9167 | 22.5833 | 21.69 | 20.7833 | 22.3167 | 22.5167 | 22.1667 | 22.3833 | 23.0333 | 31.5667 | 20.4 | 22.0333 | 22.3942 | 16 | 21.3202 | 20.3617 | 25.0667
| Median |  | 26.85 | 22.4 | 21 | 22.37 | 19.6151 | 20.55 | 20.45 | 20.85 | 22.4 | 21.63 | 20.45 | 22.25 | 22.5 | 22.3 | 22.35 | 22.95 | 31.7 | 20.35 | 22.2 | 22.1388 | 16 | 21.2461 | 20.345 | 25.45
| Std Dev. |  | 0.51929 | 3.3615 | 0 | 0.08501 | 0.55062 | 0.67946 | 0.1472 | 0.27869 | 0.70828 | 0.51773 | 0.91524 | 0.24833 | 0.1472 | 0.92232 | 0.7414 | 0.2582 | 0.49666 | 0.34059 | 0.55737 | 0.90279 | 0 | 0.70374 | 0.15381 | 1.0967
| Rel.Std.Dev. |  | 1.946% | 14.3% | 0% | 0.38% | 2.81% | 3.317% | 0.7174% | 1.332% | 3.136% | 2.387% | 4.404% | 1.113% | 0.6537% | 4.161% | 3.312% | 1.121% | 1.573% | 1.67% | 2.53% | 4.031% | 0% | 3.301% | 0.7554% | 4.375%
| PDM3 |  | 23.28% | 8.573% | -2.977% | 3.368% | -9.459% | -5.364% | -5.21% | -3.362% | 4.338% | 0.2105% | -3.978% | 3.106% | 4.03% | 2.413% | 3.414% | 6.417% | 45.84% | -5.749% | 1.797% | 3.464% | -26.08% | -1.498% | -5.927% | 15.81%
| Z-Score (Absolute) |  | 2.71 | 0.83 | 0.65 | 0.16 | 1.48 | 0.95 | 0.94 | 0.7 | 0.29 | 0.24 | 0.78 | 0.13 | 0.25 | 0.04 | 0.17 | 0.55 | 5.59 | 1 | 0.04 | 0.17 | 3.6 | 0.46 | 1.03 | 1.75
| Analytical results for Ta in OREAS 600c (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.28 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 25g | [9 empty] | 3
| 1 | 1 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.01 | [9 empty] | 1
| 1 | 2 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.06 | [9 empty] | 25
| 1 | 3 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | 16
| 1 | 4 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | < 0.01
| 1 | 5 | 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.02 | [9 empty] | 122
| 1 | 6 | 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.03
| Mean |  | 2.01: 0.01; 2.06: 0.01; 2.08: 0.01; 2.10: 0.01; 2.11: 0.01; 2.12: 0.01; 2.13: 0.01; 2.14: 0.01; 2.20: 0.026667
| Median |  | 2.01: 0.01; 2.06: 0.01; 2.08: 0.01; 2.10: 0.01; 2.11: 0.01; 2.12: 0.01; 2.13: 0.01; 2.14: 0.01; 2.20: 0.02
| Std Dev. |  | 2.01: 0; 2.08: 0; 2.10: 0; 2.11: 0; 2.12: 0; 2.13: 0; 2.14: 0; 2.20: 0.017512
| Rel.Std.Dev. |  | 2.01: 0%; 2.08: 0%; 2.10: 0%; 2.11: 0%; 2.12: 0%; 2.13: 0%; 2.14: 0%; 2.20: 65.67%
PDM3
| Z-Score (Absolute) |  | 0 | 2.02 | 0.67 | 2.02 | 0.56 | 2.02 | 0 | 0 | 0 | 0 | 0 | 0.11 | 0.67 | 0.67 | 0.67 | 2.02 | 2.25
| Analytical results for Tb in OREAS 600c (Certified Value 0.48 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.09; 2.07: 2.12; 2.08: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 01g; 2.08: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  | 3
| 1 | 1 | 2.01: 0.46; 2.03: 0.45; 2.04: 0.44338; 2.05: 0.5; 2.06: 0.7553; 2.07: 0.55; 2.08: 0.47; 2.10: 0.3488; 2.11: 0.34; 2.12: 0.46 |  |  | 1
| 1 | 2 | 2.01: 0.5; 2.03: 0.435; 2.04: 0.46958; 2.05: 0.5; 2.06: 0.7055; 2.07: 0.53; 2.08: 0.49; 2.10: 0.5202; 2.11: 0.31; 2.12: 0.492 |  |  | 26
| 1 | 3 | 2.01: 0.44; 2.03: 0.471; 2.04: 0.46607; 2.05: 0.5; 2.06: 0.7875; 2.07: 0.55; 2.08: 0.48; 2.10: 0.5401; 2.11: 0.31; 2.12: 0.487 |  |  | 16
| 1 | 4 | 2.01: 0.49; 2.03: 0.483; 2.04: 0.46393; 2.05: 0.5; 2.06: 0.7857; 2.07: 0.54; 2.08: 0.49; 2.10: 0.4481; 2.11: 0.36; 2.12: 0.521 |  |  | 0.47512
| 1 | 5 | 2.01: 0.44; 2.03: 0.463; 2.04: 0.45563; 2.05: 0.5; 2.06: 0.6996; 2.07: 0.55; 2.08: 0.49; 2.10: 0.4948; 2.11: 0.34; 2.12: 0.493 |  |  | 123
| 1 | 6 | 2.01: 0.48; 2.03: 0.46; 2.04: 0.47149; 2.05: 0.4; 2.06: 0.747; 2.07: 0.53; 2.08: 0.5; 2.10: 0.467; 2.11: 0.39; 2.12: 0.425
| Mean |  | 2.01: 0.46833; 2.03: 0.46033; 2.04: 0.46168; 2.05: 0.48333; 2.06: 0.74677; 2.07: 0.54167; 2.08: 0.48667; 2.10: 0.46983; 2.11: 0.34167; 2.12: 0.47967
| Median |  | 2.01: 0.47; 2.03: 0.4615; 2.04: 0.465; 2.05: 0.5; 2.06: 0.75115; 2.07: 0.545; 2.08: 0.49; 2.10: 0.4809; 2.11: 0.34; 2.12: 0.4895
| Std Dev. |  | 2.01: 0.025626; 2.03: 0.016633; 2.04: 0.01053; 2.05: 0.040825; 2.06: 0.037883; 2.07: 0.0098319; 2.08: 0.010328; 2.10: 0.068157; 2.11: 0.030605; 2.12: 0.033068
| Rel.Std.Dev. |  | 2.01: 5.472%; 2.03: 3.613%; 2.04: 2.281%; 2.05: 8.447%; 2.06: 5.073%; 2.07: 1.815%; 2.08: 2.122%; 2.10: 14.51%; 2.11: 8.958%; 2.12: 6.894%
| PDM3 |  | 2.01: -1.428%; 2.03: -3.112%; 2.04: -2.829%; 2.05: 1.729%; 2.06: 57.17%; 2.07: 14.01%; 2.08: 2.43%; 2.10: -1.113%; 2.11: -28.09%; 2.12: 0.957%
| Z-Score (Absolute) |  | 2.01: 0.1; 2.03: 0.65; 2.04: 0.56; 2.05: NA; 2.06: 18.99; 2.07: 4.93; 2.08: 1.15; 2.10: 0; 2.11: 8.79; 2.12: 0.67
Results from laboratory 2.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 600c (Certified Value 0.16 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 25g |  |  |  |  | 2
| 1 | 1 | 0.5 | < 0.02 | 0.2 | 0.2 | < 0.2 | 0.16 | 0.16 | 0.18 | 0.16 | 0.15 | 0.15 | 0.16 | 0.15 | 0.16 | 0.11 | < 0.05 | 0.22 | 0.13 | 0.13 | 0.18526 | < 10 | 0.06 |  |  |  |  | 1
| 1 | 2 | 0.47 | 0.21 | < 0.2 | 0.2 | < 0.2 | 0.15 | 0.19 | 0.15 | 0.14 | 0.14 | 0.14 | 0.17 | 0.14 | 0.17 | 0.11 | < 0.05 | 0.03 | 0.14 | 0.14 | 0.18544 | < 10 | < 0.02 |  |  |  |  | 27
| 1 | 3 | 0.48 | < 0.02 | < 0.2 | 0.2 | < 0.2 | 0.15 | 0.18 | 0.14 | 0.15 | 0.16 | 0.14 | 0.15 | 0.17 | 0.14 | 0.12 | < 0.05 | 0.026 | 0.13 | 0.15 | 0.19164 | < 10 | 0.07 |  |  |  |  | 16
| 1 | 4 | 0.41 | < 0.02 | < 0.2 | 0.2 | < 0.2 | 0.17 | 0.19 | 0.15 | 0.2 | 0.13 | 0.14 | 0.17 | 0.15 | 0.14 | 0.15 | < 0.05 | < 0.01 | 0.13 | 0.15 | 0.18203 | < 10 | 0.11 |  |  |  |  | 0.15765
| 1 | 5 | 0.5 | 0.18 | < 0.2 | 0.2 | < 0.2 | 0.17 | 0.19 | 0.14 | 0.14 | 0.13 | 0.14 | 0.15 | 0.15 | 0.17 | 0.08 | < 0.05 | 0.026 | 0.15 | 0.15 | 0.19055 | < 10 | 0.18 |  |  |  |  | 124
| 1 | 6 | 0.47 | 0.13 | < 0.2 | 0.2 | < 0.2 | 0.15 | 0.18 | 0.15 | 0.17 | 0.13 | 0.14 | 0.17 | 0.15 | 0.17 | 0.12 | < 0.05 | 0.125 | 0.14 | 0.16 | 0.18527 | < 10 | 0.07
| Mean |  | 0.47167 | 0.17333 | 0.2 | 0.2 |  | 0.15833 | 0.18167 | 0.15167 | 0.16 | 0.14 | 0.14167 | 0.16167 | 0.15167 | 0.15833 | 0.115 |  | 0.0854 | 0.13667 | 0.14667 | 0.1867 |  | 0.098
| Median |  | 0.475 | 0.18 | 0.2 | 0.2 |  | 0.155 | 0.185 | 0.15 | 0.155 | 0.135 | 0.14 | 0.165 | 0.15 | 0.165 | 0.115 |  | 0.03 | 0.135 | 0.15 | 0.18536 |  | 0.07
| Std Dev. |  | 0.033116 | 0.040415 |  | 3.0405e-17 |  | 0.0098319 | 0.01169 | 0.01472 | 0.022804 | 0.012649 | 0.0040825 | 0.0098319 | 0.0098319 | 0.01472 | 0.022583 |  | 0.08633 | 0.008165 | 0.010328 | 0.0036548 |  | 0.049699
| Rel.Std.Dev. |  | 7.021% | 23.32% |  | 1.52e-14% |  | 6.21% | 6.435% | 9.705% | 14.25% | 9.035% | 2.882% | 6.082% | 6.483% | 9.297% | 19.64% |  | 101.1% | 5.974% | 7.042% | 1.958% |  | 50.71%
| PDM3 |  | 199.2% | 9.951% | 26.87% | 26.87% |  | 0.4356% | 15.24% | -3.793% | 1.493% | -11.19% | -10.14% | 2.55% | -3.793% | 0.4356% | -27.05% |  | -45.83% | -13.31% | -6.965% | 18.43% |  | -37.84%
| Z-Score (Absolute) |  | 7.31 | 1.24 | 0.67 | NA | 1.05 | 0.26 | 0.79 | 0.11 | 0.3 | 0.15 | 0.11 | 0.34 | 0.11 | 0.26 | 0.71 | 2.73 | 1.68 | 0.22 | 0 | 0.9 | 109.09 | 1.42
Results from laboratory 2.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 600c (Certified Value 10.6 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  |  | 2
| 1 | 1 | 12.3 | 9.7 | 10.2 | 10.24 | 9.1977 | 11 | 10.2 | 13.861 | 11.1 | 11.75 | 10.2 | 11.1 | 10.8 | 10.2 | 11 | 8.47 | < 20 | 11.1 | 12.7 | 7.6925 | 8.45 | 10.3607 | 12.09 |  |  |  | 1
| 1 | 2 | 11.7 | 10 | 10.4 | 10.34 | 9.6901 | 11.2 | 9.7 | 13.527 | 10.8 | 11.05 | 10.2 | 10.8 | 10.8 | 9.9 | 11.3 | 8.43 | < 20 | 11.2 | 13.4 | 10.7931 | 7.76 | 10.9963 | 12.77 |  |  |  | 28
| 1 | 3 | 12 | 9.4 | 10.3 | 10.31 | 9.7761 | 11.3 | 10.2 | 13.587 | 11.4 | 11.29 | 10.2 | 10.4 | 10.8 | 9.8 | 10.2 | 8.18 | < 20 | 11.1 | 13.6 | 11.51 | 7.91 | 10.7907 | 12.57 |  |  |  | 16
| 1 | 4 | 11.6 | 10.7 | 10.4 | 10.55 | 9.4734 | 11.3 | 10.3 | 13.5633 | 11 | 11.34 | 10.1 | 10.2 | 10.6 | 10 | 10.6 | 8.52 | < 20 | 11.1 | 13.4 | 9.6356 | 8.81 | 10.4409 | 13.42 |  |  |  | 10.6206
| 1 | 5 | 11.9 | 9.4 | 10.4 | 10.53 | 9.5688 | 11.3 | 10.1 | 13.705 | 10.8 | 11.72 | 10.1 | 10.2 | 10.8 | 10.8 | 11.2 | 8.47 | < 20 | 11.3 | 13.2 | 10.8247 | 8.45 | 10.9942 | 12.86 |  |  |  | 125
| 1 | 6 | 11.6 | 9.7 | 10.3 | 10.54 | 9.533 | 10.8 | 10.2 | 13.097 | 11.2 | 11.12 | 10.5 | 9.8 | 10.6 | 10.7 | 10.5 | 8.51 | < 20 | 11.1 | 13 | 10.3833 | 9.58 | 10.2228 | 11.57
| Mean |  | 11.85 | 9.8167 | 10.3333 | 10.4183 | 9.5399 | 11.15 | 10.1167 | 13.5567 | 11.05 | 11.3783 | 10.2167 | 10.4167 | 10.7333 | 10.2333 | 10.8 | 8.43 |  | 11.15 | 13.2167 | 10.1399 | 8.4933 | 10.6343 | 12.5467
| Median |  | 11.8 | 9.7 | 10.35 | 10.435 | 9.5509 | 11.25 | 10.2 | 13.5752 | 11.05 | 11.315 | 10.2 | 10.3 | 10.8 | 10.1 | 10.8 | 8.47 |  | 11.1 | 13.3 | 10.5882 | 8.45 | 10.6158 | 12.67
| Std Dev. |  | 0.27386 | 0.48751 | 0.08165 | 0.13732 | 0.20038 | 0.20736 | 0.2137 | 0.25603 | 0.23452 | 0.2962 | 0.1472 | 0.46655 | 0.10328 | 0.42269 | 0.43359 | 0.12665 |  | 0.083666 | 0.32506 | 1.3473 | 0.65771 | 0.33665 | 0.64401
| Rel.Std.Dev. |  | 2.311% | 4.966% | 0.7902% | 1.318% | 2.1% | 1.86% | 2.112% | 1.889% | 2.122% | 2.603% | 1.441% | 4.479% | 0.9622% | 4.131% | 4.015% | 1.502% |  | 0.7504% | 2.46% | 13.29% | 7.744% | 3.166% | 5.133%
| PDM3 |  | 11.58% | -7.57% | -2.705% | -1.904% | -10.18% | 4.985% | -4.745% | 27.65% | 4.043% | 7.135% | -3.803% | -1.92% | 1.061% | -3.646% | 1.689% | -20.63% |  | 4.985% | 24.44% | -4.526% | -20.03% | 0.1287% | 18.14%
| Z-Score (Absolute) |  | 1.6 | 0.67 | 0.1 | 0 | 0.98 | 0.82 | 0.34 | 3.52 | 0.71 | 1.08 | 0.23 | 0 | 0.35 | 0.21 | 0.43 | 2.23 | 0.47 | 0.82 | 3.14 | 0.31 | 2.16 | 0.24 | 2.39
| Analytical results for Ti in OREAS 600c (Certified Value 0.023 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | [5 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | [5 empty] | 3
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | [5 empty] | 3
| 1 | 1 | 0.02 | 0.023 | 0.021 | 0.0227 | 0.022819 | 0.04 | 0.022 | 0.0341 | 0.0303 | 0.021 | 0.024 | 0.023 | 0.026 | 0.023 | 0.025 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0168 | 0.025768 | [5 empty] | 1
| 1 | 2 | 0.02 | 0.023 | 0.021 | 0.023 | 0.024019 | 0.04 | 0.02 | 0.0365 | 0.0284 | 0.021 | 0.023 | 0.024 | 0.025 | 0.024 | 0.025 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0168 | 0.025814 | [5 empty] | 29
| 1 | 3 | 0.02 | 0.023 | 0.02 | 0.0231 | 0.023277 | 0.04 | 0.021 | 0.0347 | 0.029 | 0.022 | 0.024 | 0.023 | 0.029 | 0.022 | 0.023 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0168 | 0.026681 | [5 empty] | 16
| 1 | 4 | 0.02 | 0.027 | 0.022 | 0.0221 | 0.023455 | 0.04 | 0.021 | 0.0365 | 0.0295 | 0.02 | 0.025 | 0.025 | 0.025 | 0.023 | 0.025 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0174 | 0.023538 | [5 empty] | 0.022589
| 1 | 5 | 0.02 | 0.024 | 0.02 | 0.0236 | 0.023023 | 0.04 | 0.022 | 0.0365 | 0.03 | 0.02 | 0.023 | 0.024 | 0.026 | 0.023 | 0.025 | 0.021 | 0.04 | 0.02 | < 0.1 | 0.0162 | 0.024147 | [5 empty] | 126
| 1 | 6 | 0.02 | 0.022 | 0.02 | 0.0223 | 0.023843 | 0.04 | 0.021 | 0.0347 | 0.0289 | 0.021 | 0.024 | 0.023 | 0.026 | 0.023 | 0.026 | 0.02 | 0.04 | 0.02 | < 0.1 | 0.0174 | 0.022927
| Mean |  | 0.02 | 0.023667 | 0.020667 | 0.0228 | 0.023406 | 0.04 | 0.021167 | 0.0355 | 0.02935 | 0.020833 | 0.023833 | 0.023667 | 0.026167 | 0.023 | 0.024833 | 0.020167 | 0.04 | 0.02 |  | 0.0169 | 0.024813
| Median |  | 0.02 | 0.023 | 0.0205 | 0.02285 | 0.023366 | 0.04 | 0.021 | 0.0356 | 0.02925 | 0.021 | 0.024 | 0.0235 | 0.026 | 0.023 | 0.025 | 0.02 | 0.04 | 0.02 |  | 0.0168 | 0.024957
| Std Dev. |  | 0 | 0.0017512 | 0.0008165 | 0.00055136 | 0.00046427 | 0 | 0.00075277 | 0.0011171 | 0.00071764 | 0.00075277 | 0.00075277 | 0.0008165 | 0.001472 | 0.00063246 | 0.00098319 | 0.00040825 | 0 | 0 |  | 0.00045166 | 0.0014854
| Rel.Std.Dev. |  | 0% | 7.399% | 3.951% | 2.418% | 1.984% | 0% | 3.556% | 3.147% | 2.445% | 3.613% | 3.158% | 3.45% | 5.625% | 2.75% | 3.959% | 2.024% | 0% | 0% |  | 2.673% | 5.986%
| PDM3 |  | -11.46% | 4.769% | -8.512% | 0.9321% | 3.615% | 77.07% | -6.298% | 57.15% | 29.93% | -7.774% | 5.506% | 4.769% | 15.84% | 1.817% | 9.933% | -10.73% | 77.07% | -11.46% |  | -25.19% | 9.841%
| Z-Score (Absolute) |  | 0.87 | 0 | 0.71 | 0.21 | 0.06 | 3.89 | 0.59 | 2.82 | 1.35 | 0.67 | 0.04 | 0 | 0.59 | 0.16 | 0.28 | 0.83 | 3.89 | 0.87 | 6.27 | 1.61 | 0.27
| Analytical results for Tl in OREAS 600c (Certified Value 0.47 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.26 | 2.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 01g | 25g |  |  |  |  | 3
| 1 | 1 | 0.39 | 0.44 | 0.5 | 0.46 | 0.45892 | 0.47 | 0.51 | 0.45 | 0.55 | 0.55 | 0.48 | 0.5 | 0.44 | 0.47 | 0.45 | 0.503 | 0.51 | 0.39 | < 5 | 0.31 | 0.45623 | 0.42 |  |  |  |  | 1
| 1 | 2 | 0.42 | 0.49 | 0.5 | 0.47 | 0.46861 | 0.48 | 0.48 | 0.48 | 0.56 | 0.52 | 0.47 | 0.49 | 0.45 | 0.46 | 0.44 | 0.448 | 0.48 | 0.43 | < 5 | 0.3 | 0.45229 | 0.43 |  |  |  |  | 30
| 1 | 3 | 0.43 | 0.43 | 0.5 | 0.48 | 0.4863 | 0.48 | 0.5 | 0.48 | 0.53 | 0.52 | 0.47 | 0.47 | 0.47 | 0.48 | 0.4 | 0.424 | 0.5 | 0.41 | < 5 | 0.35 | 0.47986 | 0.44 |  |  |  |  | 16
| 1 | 4 | 0.37 | 0.49 | 0.5 | 0.49 | 0.46047 | 0.49 | 0.5 | 0.48 | 0.5 | 0.53 | 0.48 | 0.46 | 0.45 | 0.46 | 0.4 | 0.451 | 0.49 | 0.39 | < 5 | 0.33 | 0.47961 | 0.46 |  |  |  |  | 0.46517
| 1 | 5 | 0.38 | 0.41 | 0.5 | 0.48 | 0.45285 | 0.48 | 0.5 | 0.5 | 0.54 | 0.54 | 0.47 | 0.45 | 0.45 | 0.49 | 0.43 | 0.343 | 0.49 | 0.39 | < 5 | 0.32 | 0.46727 | 0.44 |  |  |  |  | 127
| 1 | 6 | 0.37 | 0.47 | 0.5 | 0.49 | 0.47998 | 0.46 | 0.5 | 0.49 | 0.54 | 0.51 | 0.51 | 0.43 | 0.45 | 0.48 | 0.42 | 0.477 | 0.48 | 0.41 | < 5 | 0.29 | 0.45732 | 0.41
| Mean |  | 0.39333 | 0.455 | 0.5 | 0.47833 | 0.46785 | 0.47667 | 0.49833 | 0.48 | 0.53667 | 0.52833 | 0.48 | 0.46667 | 0.45167 | 0.47333 | 0.42333 | 0.441 | 0.49167 | 0.40333 |  | 0.31667 | 0.46543 | 0.43333
| Median |  | 0.385 | 0.455 | 0.5 | 0.48 | 0.46454 | 0.48 | 0.5 | 0.48 | 0.54 | 0.525 | 0.475 | 0.465 | 0.45 | 0.475 | 0.425 | 0.4495 | 0.49 | 0.4 |  | 0.315 | 0.46229 | 0.435
| Std Dev. |  | 0.02582 | 0.033317 | 0 | 0.01169 | 0.013018 | 0.010328 | 0.0098319 | 0.016733 | 0.020656 | 0.01472 | 0.015492 | 0.02582 | 0.0098319 | 0.012111 | 0.020656 | 0.055104 | 0.01169 | 0.01633 |  | 0.021602 | 0.012129 | 0.017512
| Rel.Std.Dev. |  | 6.564% | 7.322% | 0% | 2.444% | 2.782% | 2.167% | 1.973% | 3.486% | 3.849% | 2.786% | 3.227% | 5.533% | 2.177% | 2.559% | 4.879% | 12.5% | 2.378% | 4.049% |  | 6.822% | 2.606% | 4.041%
| PDM3 |  | -15.44% | -2.187% | 7.487% | 2.829% | 0.5765% | 2.471% | 7.129% | 3.187% | 15.37% | 13.58% | 3.187% | 0.3211% | -2.903% | 1.754% | -8.994% | -5.197% | 5.695% | -13.29% |  | -31.92% | 0.05526% | -6.845%
| Z-Score (Absolute) |  | 2.11 | 0.36 | NA | 0.3 | 0 | 0.25 | 0.86 | 0.34 | 1.95 | 1.71 | 0.34 | 0.03 | 0.46 | 0.16 | 1.26 | 0.76 | 0.67 | 1.83 | 57.55 | 4.28 | 0.07 | 0.98
Results from laboratory 2.04 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 600c (Indicative Value 0.048 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.04: 2.06; 2.05: 2.07; 2.06: 2.12; 2.07: 2.23 |  |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*MS |  |  |  |  | 3
|  |  | 2.01: 0.5g; 2.03: 01g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 01g; 2.07: 0.5g |  |  |  |  | 3
| 1 | 1 | 2.01: 0.05; 2.03: 0.047; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.0319 |  |  |  |  | 1
| 1 | 2 | 2.01: 0.05; 2.03: 0.048; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0474 |  |  |  |  | 5
| 1 | 3 | 2.01: 0.04; 2.03: 0.048; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0475 |  |  |  |  | 16
| 1 | 4 | 2.01: 0.05; 2.03: 0.05; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.06; 2.07: 0.0408 |  |  |  |  | 0.048435
| 1 | 5 | 2.01: 0.05; 2.03: 0.048; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.0455 |  |  |  |  | 11
| 1 | 6 | 2.01: 0.04; 2.03: 0.05; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.045
| Mean |  | 2.01: 0.046667; 2.03: 0.0485; 2.06: 0.053333; 2.07: 0.043017
| Median |  | 2.01: 0.05; 2.03: 0.048; 2.06: 0.05; 2.07: 0.04525
| Std Dev. |  | 2.01: 0.005164; 2.03: 0.0012247; 2.06: 0.005164; 2.07: 0.005965
| Rel.Std.Dev. |  | 2.01: 11.07%; 2.03: 2.525%; 2.06: 9.682%; 2.07: 13.87%
| PDM3 |  | 2.01: -3.651%; 2.03: 0.1342%; 2.06: 10.11%; 2.07: -11.19%
| Z-Score (Absolute) |  | 2.01: 0.18; 2.03: 0.18; 2.04: 4.36; 2.05: 0.47; 2.06: 1.11; 2.07: 0.88
| Analytical results for U in OREAS 600c (Certified Value 2.83 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  | 3
| 1 | 1 | 2.77 | 2.76 | 2.8 | 2.85 | 2.7541 | 2.9 | 2.83 | 3.3834 | 2.89 | 3.29 | 3.39 | 2.85 | 2.76 | 2.8 | 3.08 | 2.7 | 3.27 | 2.76 | 2.54 | 2.69 | 1.9123 | 2.07 | 2.8137 | 2.95 |  |  | 1
| 1 | 2 | 2.66 | 2.89 | 2.9 | 2.91 | 2.8452 | 2.9 | 2.69 | 3.8362 | 2.87 | 3.43 | 3.27 | 2.78 | 2.77 | 2.8 | 2.94 | 2.69 | 3 | 2.77 | 2.86 | 2.84 | 2.7629 | 1.94 | 3.076 | 3.06 |  |  | 32
| 1 | 3 | 2.88 | 2.73 | 2.8 | 2.88 | 2.8745 | 3 | 2.87 | 3.6274 | 3.04 | 3.27 | 3.32 | 2.85 | 2.69 | 2.76 | 2.78 | 2.45 | 2.83 | 2.7 | 2.7 | 2.85 | 2.8359 | 2.01 | 2.9055 | 3.09 |  |  | 16
| 1 | 4 | 2.71 | 3.05 | 2.9 | 2.95 | 2.8464 | 2.9 | 2.82 | 3.3851 | 2.93 | 3.2 | 3.28 | 2.73 | 2.65 | 2.78 | 3.03 | 2.51 | 2.99 | 2.78 | 2.63 | 2.84 | 2.3963 | 2.15 | 2.8016 | 3.24 |  |  | 2.8277
| 1 | 5 | 2.66 | 2.66 | 2.9 | 2.98 | 2.7419 | 2.9 | 2.79 | 3.6287 | 3.02 | 3.35 | 3.35 | 2.77 | 2.55 | 2.77 | 3.15 | 2.66 | 2.31 | 2.77 | 2.71 | 2.73 | 2.7282 | 2.07 | 3.1199 | 3.04 |  |  | 128
| 1 | 6 | 2.71 | 2.81 | 2.9 | 2.98 | 2.7746 | 2.8 | 2.78 | 3.7877 | 3.02 | 3.25 | 3.28 | 2.9 | 2.47 | 2.73 | 3.14 | 2.55 | 3.19 | 2.78 | 2.76 | 2.81 | 2.5636 | 2.25 | 2.7617 | 2.75
| Mean |  | 2.7317 | 2.8167 | 2.8667 | 2.925 | 2.8061 | 2.9 | 2.7967 | 3.6081 | 2.9617 | 3.2983 | 3.315 | 2.8133 | 2.6483 | 2.7733 | 3.02 | 2.5933 | 2.9317 | 2.76 | 2.7 | 2.7933 | 2.5332 | 2.0817 | 2.9131 | 3.0217
| Median |  | 2.71 | 2.785 | 2.9 | 2.93 | 2.8099 | 2.9 | 2.805 | 3.628 | 2.975 | 3.28 | 3.3 | 2.815 | 2.67 | 2.775 | 3.055 | 2.605 | 2.995 | 2.77 | 2.705 | 2.825 | 2.6459 | 2.07 | 2.8596 | 3.05
| Std Dev. |  | 0.083287 | 0.13794 | 0.05164 | 0.053944 | 0.05594 | 0.063246 | 0.06121 | 0.19251 | 0.07414 | 0.081097 | 0.047645 | 0.063456 | 0.11873 | 0.026583 | 0.14071 | 0.10443 | 0.34237 | 0.030332 | 0.10936 | 0.067132 | 0.34281 | 0.10815 | 0.15139 | 0.16315
| Rel.Std.Dev. |  | 3.049% | 4.897% | 1.801% | 1.844% | 1.994% | 2.181% | 2.189% | 5.336% | 2.503% | 2.459% | 1.437% | 2.256% | 4.483% | 0.9585% | 4.659% | 4.027% | 11.68% | 1.099% | 4.05% | 2.403% | 13.53% | 5.195% | 5.197% | 5.399%
| PDM3 |  | -3.397% | -0.3906% | 1.378% | 3.441% | -0.7635% | 2.556% | -1.098% | 27.6% | 4.737% | 16.64% | 17.23% | -0.5085% | -6.344% | -1.923% | 6.8% | -8.289% | 3.676% | -2.395% | -4.516% | -1.216% | -10.42% | -26.38% | 3.019% | 6.859%
| Z-Score (Absolute) |  | 0.5 | 0.01 | 0.31 | 0.66 | 0.05 | 0.51 | 0.11 | 4.75 | 0.88 | 2.9 | 3 | 0.01 | 1 | 0.25 | 1.23 | 1.33 | 0.7 | 0.33 | 0.69 | 0.13 | 1.69 | 4.4 | 0.59 | 1.24
| Analytical results for V in OREAS 600c (Certified Value 1.98 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g |  |  |  |  | 2
| 1 | 1 | 2 | 2 | 2 | 1.9504 | 2 | 1.4 | 2.19 | 3 |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 1.9 | 2 | 0.911 | 1 | < 2 | < 10 |  |  |  |  | 1
| 1 | 2 | 2 | 3 | 2 | 1.9875 | 2 | 1.4 | 2.34 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 1.9 | 2 | 1.1321 | 1.2 | < 2 | < 10 |  |  |  |  | 33
| 1 | 3 | 2 | 2 | 2 | 2.0415 | 2 | 1.4 | 2.43 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 1.9 | 2 | 1.2656 | 1.4 | 2.0635 | < 10 |  |  |  |  | 16
| 1 | 4 | 2 | 2 | 2 | 1.933 | 2 | 1.4 | 2.44 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 1.8 | 2 | 1.1966 | 1 | < 2 | < 10 |  |  |  |  | 1.9818
| 1 | 5 | 2 | 2 | 2 | 2.0698 | 2 | 1.5 | 2.19 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2.4 | 1.9 | 2 | 1.2402 | 0.9 | 2.0097 | < 10 |  |  |  |  | 129
| 1 | 6 | 2 | 2 | 2 | 1.962 | 2 | 1.4 | 2.41 | 3 |  | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 1.8 | 2 | 0.9834 | 1 | < 2 | < 10
| Mean |  | 2 | 2.1667 | 2 | 1.9907 | 2 | 1.4167 | 2.3333 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 2.6667 | 2.0667 | 1.8667 | 2 | 1.1215 | 1.0833 | 2.0366
| Median |  | 2 | 2 | 2 | 1.9748 | 2 | 1.4 | 2.375 | 3 | 2 | 2 | 2 | 2 | 2 | 2 | 3 | 2 | 1.9 | 2 | 1.1643 | 1 | 2.0366
| Std Dev. |  | 0 | 0.40825 | 0 | 0.05407 | 0 | 0.040825 | 0.11639 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.5164 | 0.1633 | 0.05164 | 0 | 0.14421 | 0.18348 | 0.038006
| Rel.Std.Dev. |  | 0% | 18.84% | 0% | 2.716% | 0% | 2.882% | 4.988% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 19.36% | 7.902% | 2.766% | 0% | 12.86% | 16.94% | 1.866%
| PDM3 |  | 0.9159% | 9.326% | 0.9159% | 0.4475% | 0.9159% | -28.52% | 17.74% | 51.37% | 0.9159% | 0.9159% | 0.9159% | 0.9159% | 0.9159% | 0.9159% | 34.55% | 4.28% | -5.812% | 0.9159% | -43.41% | -45.34% | 2.763%
| Z-Score (Absolute) |  | 0 | 2.96 | 0 | 0.16 | 0 | 10.36 | 5.92 | 17.76 | 0 | 0 | 0 | 0 | 0 | 0 | 11.84 | 1.18 | 2.37 | 0 | 15.6 | 16.28 | 11.62 | 53.27
| Analytical results for W in OREAS 600c (Certified Value 0.46 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 25g |  |  |  | 3
| 1 | 1 | 0.42 | 0.41 | 0.3 | 0.47 | < 1 | 0.5 | 0.52 | 0.47 | 0.4 | 0.33 | 0.45 | 0.43 | 0.47 | 0.44 | 0.44 | 0.61 | < 0.05 | < 5 | 0.14 | < 0.1 | < 0.1 | 0.44156 | 1 |  |  |  | 1
| 1 | 2 | 0.38 | 0.43 | 0.3 | 0.45 | < 1 | 0.6 | 0.46 | 0.5 | 0.4 | 0.34 | 0.41 | 0.43 | 0.38 | 0.43 | 0.44 | 0.58 | < 0.05 | < 5 | 0.14 | < 0.1 | < 0.1 | 0.51847 | 2 |  |  |  | 34
| 1 | 3 | 0.48 | 0.54 | 0.4 | 0.47 | < 1 | 0.5 | 0.54 | 0.5 | 0.4 | 0.34 | 0.45 | 0.48 | 0.41 | 0.43 | 0.46 | 0.6 | < 0.05 | < 5 | 0.13 | < 0.1 | < 0.1 | 0.43292 | 2 |  |  |  | 16
| 1 | 4 | 0.47 | 0.43 | 0.4 | 0.47 | < 1 | 0.5 | 0.52 | 0.5 | 0.5 | 0.32 | 0.43 | 0.41 | 0.43 | 0.41 | 0.48 | 0.7 | < 0.05 | < 5 | 0.18 | < 0.1 | < 0.1 | 0.42865 | 2 |  |  |  | 0.45696
| 1 | 5 | 0.41 | 0.51 | 0.3 | 0.45 | < 1 | 0.6 | 0.52 | 0.5 | 0.4 | 0.32 | 0.47 | 0.42 | 0.42 | 0.43 | 0.48 | 0.59 | < 0.05 | < 5 | 0.15 | < 0.1 | < 0.1 | 0.51179 | 1 |  |  |  | 130
| 1 | 6 | 0.47 | 0.46 | 0.3 | 0.45 | < 1 | 0.5 | 0.49 | 0.5 | 0.4 | 0.33 | 0.49 | 0.43 | 0.43 | 0.41 | 0.49 | 0.62 | < 0.05 | < 5 | 0.16 | < 0.1 | < 0.1 | 0.42397 | 2
| Mean |  | 0.43833 | 0.46333 | 0.33333 | 0.46 |  | 0.53333 | 0.50833 | 0.495 | 0.41667 | 0.33 | 0.45 | 0.43333 | 0.42333 | 0.425 | 0.465 | 0.61667 |  |  | 0.15 |  |  | 0.45956 | 1.6667
| Median |  | 0.445 | 0.445 | 0.3 | 0.46 |  | 0.5 | 0.52 | 0.5 | 0.4 | 0.33 | 0.45 | 0.43 | 0.425 | 0.43 | 0.47 | 0.605 |  |  | 0.145 |  |  | 0.43724 | 2
| Std Dev. |  | 0.040702 | 0.051251 | 0.05164 | 0.010954 |  | 0.05164 | 0.028577 | 0.012247 | 0.040825 | 0.0089443 | 0.028284 | 0.024221 | 0.029439 | 0.012247 | 0.021679 | 0.043205 |  |  | 0.017889 |  |  | 0.043482 | 0.5164
| Rel.Std.Dev. |  | 9.286% | 11.06% | 15.49% | 2.381% |  | 9.682% | 5.622% | 2.474% | 9.798% | 2.71% | 6.285% | 5.59% | 6.954% | 2.882% | 4.662% | 7.006% |  |  | 11.93% |  |  | 9.462% | 30.98%
| PDM3 |  | -4.076% | 1.395% | -27.05% | 0.6652% |  | 16.71% | 11.24% | 8.325% | -8.818% | -27.78% | -1.523% | -5.17% | -7.359% | -6.994% | 1.759% | 34.95% |  |  | -67.17% |  |  | 0.5693% | 264.7%
| Z-Score (Absolute) |  | 0.17 | 0.2 | NA | 0.15 | 0.75 | NA | 0.87 | 0.67 | NA | 1.8 | 0 | 0.25 | 0.4 | 0.37 | 0.22 | 2.5 | 6.37 | 30.72 | 4.5 | 5.99 | 5.99 | 0.14 | NA
Results from laboratories 2.04, 2.07 and 2.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 600c (Certified Value 8.07 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  |  | 3
| 1 | 1 | 9.25 | 8.85 | 7.86 | 8.0623 | 7.83 | 7.39 | 8.1372 | 7.8 | 9.1 | 8.87 | 7.42 | 8.65 | 8.56 | 8.37 | 8.15 | 6.77 | 7.21 | 7.57 | 5.9147 | 5.94 | 7.8329 | < 10 | 8.43 |  |  |  | 1
| 1 | 2 | 8.98 | 7.69 | 7.96 | 8.3858 | 7.96 | 7.29 | 8.1653 | 8 | 8.61 | 8.65 | 7.52 | 8.31 | 8.28 | 8.3 | 8.21 | 6.99 | 7.78 | 7.71 | 8.1903 | 5.45 | 8.2376 | < 10 | 9.03 |  |  |  | 35
| 1 | 3 | 8.85 | 8.44 | 8.04 | 8.3225 | 7.64 | 7.47 | 8.1842 | 7.9 | 9.17 | 8.86 | 7.31 | 8.38 | 8.25 | 8.1 | 7.49 | 6.67 | 7.6 | 7.6 | 8.4279 | 5.6 | 8.2678 | < 10 | 9.01 |  |  |  | 16
| 1 | 4 | 8.88 | 9.39 | 8.16 | 8.3072 | 8.02 | 7.42 | 8.1412 | 8 | 8.58 | 8.73 | 7.34 | 8.36 | 8.48 | 8.38 | 7.68 | 7.17 | 7.64 | 7.75 | 7.4183 | 6.21 | 7.8814 | < 10 | 9.51 |  |  |  | 8.0702
| 1 | 5 | 8.88 | 7.46 | 8.1 | 8.2716 | 7.93 | 7.48 | 8.1099 | 8.1 | 9.14 | 8.93 | 7.24 | 8.28 | 8.55 | 8.36 | 8.16 | 6.75 | 7.42 | 7.42 | 7.9969 | 5.9 | 8.0723 | < 10 | 9.01 |  |  |  | 131
| 1 | 6 | 8.75 | 7.37 | 8.2 | 8.2385 | 7.59 | 7.48 | 8.1416 | 7.8 | 9.06 | 8.51 | 7.78 | 8.25 | 8.37 | 8.6 | 7.87 | 6.77 | 7.5 | 7.28 | 7.4378 | 6.72 | 7.8923 | < 10 | 7.92
| Mean |  | 8.9317 | 8.2 | 8.0533 | 8.2646 | 7.8283 | 7.4217 | 8.1466 | 7.9333 | 8.9433 | 8.7583 | 7.435 | 8.3717 | 8.415 | 8.3517 | 7.9267 | 6.8533 | 7.525 | 7.555 | 7.5643 | 5.97 | 8.0307 |  | 8.8183
| Median |  | 8.88 | 8.065 | 8.07 | 8.2894 | 7.88 | 7.445 | 8.1414 | 7.95 | 9.08 | 8.795 | 7.38 | 8.335 | 8.425 | 8.365 | 8.01 | 6.77 | 7.55 | 7.585 | 7.7173 | 5.92 | 7.9823 |  | 9.01
| Std Dev. |  | 0.17244 | 0.82375 | 0.12754 | 0.11091 | 0.17702 | 0.07414 | 0.025506 | 0.12111 | 0.27252 | 0.15905 | 0.1945 | 0.1447 | 0.13487 | 0.16055 | 0.29642 | 0.18822 | 0.19736 | 0.17785 | 0.90369 | 0.45458 | 0.19045 |  | 0.55758
| Rel.Std.Dev. |  | 1.931% | 10.05% | 1.584% | 1.342% | 2.261% | 0.999% | 0.3131% | 1.527% | 3.047% | 1.816% | 2.616% | 1.728% | 1.603% | 1.922% | 3.74% | 2.746% | 2.623% | 2.354% | 11.95% | 7.614% | 2.372% |  | 6.323%
| PDM3 |  | 10.67% | 1.608% | -0.2096% | 2.409% | -2.998% | -8.037% | 0.9457% | -1.697% | 10.82% | 8.526% | -7.871% | 3.735% | 4.272% | 3.487% | -1.779% | -15.08% | -6.756% | -6.385% | -6.269% | -26.02% | -0.49% |  | 9.27%
| Z-Score (Absolute) |  | 1.3 | 0.24 | 0.03 | 0.34 | 0.29 | 0.88 | 0.17 | 0.14 | 1.32 | 1.05 | 0.86 | 0.49 | 0.56 | 0.46 | 0.15 | 1.7 | 0.73 | 0.69 | 0.67 | 2.98 | 0 | 4.38 | 1.14
| Analytical results for Yb in OREAS 600c (Certified Value 0.24 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  | 1
| No | No | 2.01: 2.03; 2.03: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.09; 2.08: 2.12; 2.09: 2.19; 2.10: 2.23; 2.11: 2.25; 2.12: 2.28 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.03: AR*MS; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS |  |  |  | 2
|  |  | 2.01: 0.5g; 2.03: 01g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.25g; 2.08: 01g; 2.09: 0.4g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 25g |  |  |  | 3
| 1 | 1 | 2.01: 0.23; 2.03: 0.228; 2.05: 0.23131; 2.06: 0.2; 2.07: 0.3144; 2.08: 0.26; 2.09: 0.3; 2.10: 0.1764; 2.11: 0.18; 2.12: 0.29 |  |  |  | 1
| 1 | 2 | 2.01: 0.25; 2.03: 0.228; 2.05: 0.24189; 2.06: 0.2; 2.07: 0.3485; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2533; 2.11: 0.16; 2.12: 0.33 |  |  |  | 36
| 1 | 3 | 2.01: 0.22; 2.03: 0.238; 2.05: 0.23487; 2.06: 0.2; 2.07: 0.3225; 2.08: 0.25; 2.09: 0.3; 2.10: 0.2718; 2.11: 0.17; 2.12: 0.28 |  |  |  | 16
| 1 | 4 | 2.01: 0.24; 2.03: 0.227; 2.05: 0.22799; 2.06: 0.2; 2.07: 0.326; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2256; 2.11: 0.19; 2.12: 0.29 |  |  |  | 0.23727
| 1 | 5 | 2.01: 0.21; 2.03: 0.221; 2.05: 0.2312; 2.06: 0.2; 2.07: 0.3432; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2475; 2.11: 0.18; 2.12: 0.3 |  |  |  | 132
| 1 | 6 | 2.01: 0.22; 2.03: 0.234; 2.05: 0.23426; 2.06: 0.2; 2.07: 0.3425; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2523; 2.11: 0.2; 2.12: 0.27
| Mean |  | 2.01: 0.22833; 2.03: 0.22933; 2.05: 0.23359; 2.06: 0.2; 2.07: 0.33285; 2.08: 0.245; 2.09: 0.3; 2.10: 0.23782; 2.11: 0.18; 2.12: 0.29333
| Median |  | 2.01: 0.225; 2.03: 0.228; 2.05: 0.23279; 2.06: 0.2; 2.07: 0.33425; 2.08: 0.24; 2.09: 0.3; 2.10: 0.2499; 2.11: 0.18; 2.12: 0.29
| Std Dev. |  | 2.01: 0.01472; 2.03: 0.0059217; 2.05: 0.0047592; 2.06: 3.0405e-17; 2.07: 0.013709; 2.08: 0.0083666; 2.09: 0; 2.10: 0.033524; 2.11: 0.014142; 2.12: 0.020656
| Rel.Std.Dev. |  | 2.01: 6.447%; 2.03: 2.582%; 2.05: 2.037%; 2.06: 1.52e-14%; 2.07: 4.119%; 2.08: 3.415%; 2.09: 0%; 2.10: 14.1%; 2.11: 7.857%; 2.12: 7.042%
| PDM3 |  | 2.01: -3.767%; 2.03: -3.345%; 2.05: -1.553%; 2.06: -15.71%; 2.07: 40.28%; 2.08: 3.258%; 2.09: 26.44%; 2.10: 0.2302%; 2.11: -24.14%; 2.12: 23.63%
| Z-Score (Absolute) |  | 2.01: 0.6; 2.03: 0.52; 2.05: 0.17; 2.06: NA; 2.07: 7.86; 2.08: 0.75; 2.09: NA; 2.10: 0.17; 2.11: 4.51; 2.12: 4.66
Results from laboratories 2.07 and 2.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 600c (Certified Value 2052 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 15g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0
| 1 | 1 | 2214 | 1818.6 | 2013 | 2042 | 1988.6431 | 2080 | 2740 | 1925.12 | 2048 | 2160 | 2024 | 2020 | 2050 | 1990 | 2350 | 2080 | 2197 | 2100 | 2430 | 1971 | 2063 | 1758.3299 | 2030 | 2125.5356 | 1944.967 | 1926 | 1
| 1 | 2 | 2086 | 1624.1 | 2019 | 2038 | 2046.9035 | 2100 | 2080 | 1945.91 | 2058 | 2140 | 1954 | 2010 | 2020 | 2020 | 2310 | 2110 | 2243 | 2094 | 2400 | 1970 | 2085 | 1765.5487 | 1980 | 2087.6401 | 1973.019 | 2076 | 17
| 1 | 3 | 2098 | 1778.3 | 2001 | 2066 | 2036.4008 | 2110 | 2070 | 1953.47 | 2057 | 2070 | 1922 | 2010 | 2040 | 1990 | 2060 | 1995 | 2246 | 2082 | 2470 | 1968 | 2050 | 1789.4648 | 2020 | 2097.5186 | 1972.874 | 2145 | 16
| 1 | 4 | 2117 | 1839 | 2022 | 2044 | 2033.3914 | 2110 | 2040 | 1937.365 | 2056 | 2110 | 1996 | 2020 | 2030 | 1980 | 2310 | 2070 | 2220 | 2155 | 2430 | 1967 | 2022 | 1762.3072 | 2100 | 2106.8093 | 1819.297 | 2179 | 2052.4771
| 1 | 5 | 2034 | 1688.6 | 2010 | 2041 | 1998.9056 | 2070 | 2110 | 1954.415 | 2081 | 2130 | 1994 | 1990 | 2020 | 2030 | 2370 | 2080 | 2217 | 2118 | 2430 | 1982 | 2055 | 1795.9328 | 2020 | 2124.4474 | 1995.782 | 2108 | 133
| 1 | 6 | 2185 | 1601 | 2002 | 2069 | 2011.7946 | 1970 | 2060 | 1964.01 | 2050 | 2150 | 1954 | 2020 | 2040 | 1990 | 2340 | 2080 | 2213 | 2110 | 2410 | 1960 | 2028 | 1769.3393 | 2100 | 2108.0298 | 2030.581 | 1955
| Mean |  | 2122.3333 | 1724.9333 | 2011.1667 | 2050 | 2019.3398 | 2073.3333 | 2183.3333 | 1946.715 | 2058.3333 | 2126.6667 | 1974 | 2011.6667 | 2033.3333 | 2000 | 2290 | 2069.1667 | 2222.6667 | 2109.8333 | 2428.3333 | 1969.6667 | 2050.5 | 1773.4871 | 2041.6667 | 2108.3301 | 1956.0867 | 2064.8333
| Median |  | 2107.5 | 1733.45 | 2011.5 | 2043 | 2022.593 | 2090 | 2075 | 1949.69 | 2056.5 | 2135 | 1974 | 2015 | 2035 | 1990 | 2325 | 2080 | 2218.5 | 2105 | 2430 | 1969 | 2052.5 | 1767.444 | 2025 | 2107.4196 | 1972.9465 | 2092
| Std Dev. |  | 66.443 | 101.4688 | 8.612 | 13.7259 | 23.087 | 53.1664 | 273.6908 | 13.8452 | 11.8096 | 32.6599 | 37.116 | 11.6905 | 12.1106 | 20 | 115.0652 | 38.7836 | 18.7047 | 25.4119 | 24.0139 | 7.174 | 23.1754 | 15.4535 | 48.3391 | 14.8537 | 72.8405 | 102.749
| Rel.Std.Dev. |  | 3.131% | 5.882% | 0.4282% | 0.6696% | 1.143% | 2.564% | 12.54% | 0.7112% | 0.5737% | 1.536% | 1.88% | 0.5811% | 0.5956% | 1% | 5.025% | 1.874% | 0.8415% | 1.204% | 0.9889% | 0.3642% | 1.13% | 0.8714% | 2.368% | 0.7045% | 3.724% | 4.976%
| PDM3 |  | 3.404% | -15.96% | -2.013% | -0.1207% | -1.615% | 1.016% | 6.376% | -5.153% | 0.2853% | 3.615% | -3.824% | -1.988% | -0.9327% | -2.557% | 11.57% | 0.8131% | 8.292% | 2.794% | 18.31% | -4.035% | -0.09633% | -13.59% | -0.5267% | 2.721% | -4.696% | 0.602%
| Z-Score (Absolute) |  | 0.83 | 3.73 | 0.45 | 0 | 0.35 | 0.26 | 1.53 | 1.19 | 0.09 | 0.88 | 0.87 | 0.44 | 0.19 | 0.58 | 2.75 | 0.22 | 1.98 | 0.68 | 4.33 | 0.92 | 0 | 3.17 | 0.1 | 0.67 | 1.08 | 0.17
| Analytical results for Zr in OREAS 600c (Certified Value 49.9 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | 2.28 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 01g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 25g |  |  | 2
| 1 | 1 | 49 | 49.7 | 48.1 | 47.2999 | 12.1 | 46.2 | 57.6425 | 51 | 58.1 | 57.6 | 48.2 | 52.4 | 49.9 | 50.9 | 50.3 | 43.5 | 49.6 | 68.3 | 33.1 | 69.5217 | 29.5 | 41.873 | 53.33 | 38.2 |  |  | 1
| 1 | 2 | 48 | 47.3 | 48.8 | 50.0959 | 13 | 45.3 | 57.786 | 52.3 | 58.4 | 53.7 | 49.9 | 50.2 | 50.7 | 49.8 | 50.3 | 44.2 | 54.9 | 65.1 | 34.1 | 66.0789 | 27.1 | 45.2436 | 52.85 | 35.9 |  |  | 18
| 1 | 3 | 49 | 49.3 | 48.9 | 49.6297 | 9.7 | 46.7 | 58.472 | 51.8 | 56.1 | 57.1 | 48 | 50.7 | 48.6 | 49.9 | 46.9 | 41.9 | 53.2 | 65.7 | 35.8 | 71.4251 | 27.9 | 44.9015 | 53.03 | 42 |  |  | 16
| 1 | 4 | 48 | 52.4 | 48 | 49.2802 | 9.2 | 46.2 | 59.858 | 51.9 | 54.5 | 57.7 | 48.2 | 49.4 | 49.3 | 50.2 | 46.8 | 44.1 | 51.9 | 63.6 | 35 | 73.7217 | 31.2 | 42.0663 | 51.42 | 52.4 |  |  | 49.9236
| 1 | 5 | 48 | 48.1 | 48.9 | 47.5691 | 17.1 | 47.1 | 55.4165 | 52.9 | 56.8 | 58.2 | 49.2 | 49.9 | 49.4 | 50.5 | 48.7 | 44.1 | 51.2 | 64.2 | 34.3 | 76.1658 | 29.9 | 45.1961 | 52.26 | 42 |  |  | 134
| 1 | 6 | 47 | 46.3 | 47.5 | 48.7735 | 12.1 | 46.8 | 56.1095 | 52.8 | 55.2 | 54.7 | 50.9 | 48.6 | 49.6 | 50.9 | 47.6 | 44.3 | 55.1 | 64.2 | 31.1 | 67.9896 | 33.7 | 41.8826 | 52.94 | 49.5
| Mean |  | 48.1667 | 48.85 | 48.3667 | 48.7747 | 12.2 | 46.3833 | 57.5474 | 52.1167 | 56.5167 | 56.5 | 49.0667 | 50.2 | 49.5833 | 50.3667 | 48.4333 | 43.6833 | 52.65 | 65.1833 | 33.9 | 70.8171 | 29.8833 | 43.5272 | 52.6383 | 43.3333
| Median |  | 48 | 48.7 | 48.45 | 49.0268 | 12.1 | 46.45 | 57.7142 | 52.1 | 56.45 | 57.35 | 48.7 | 50.05 | 49.5 | 50.35 | 48.15 | 44.1 | 52.55 | 64.65 | 34.2 | 70.4734 | 29.7 | 43.4839 | 52.895 | 42
| Std Dev. |  | 0.75277 | 2.1445 | 0.58538 | 1.1279 | 2.8256 | 0.63692 | 1.6042 | 0.70828 | 1.5562 | 1.8428 | 1.1587 | 1.2946 | 0.6969 | 0.48028 | 1.5971 | 0.91742 | 2.1603 | 1.6987 | 1.6432 | 3.7293 | 2.3718 | 1.7433 | 0.69211 | 6.4093
| Rel.Std.Dev. |  | 1.563% | 4.39% | 1.21% | 2.312% | 23.16% | 1.373% | 2.788% | 1.359% | 2.753% | 3.262% | 2.362% | 2.579% | 1.406% | 0.9536% | 3.297% | 2.1% | 4.103% | 2.606% | 4.847% | 5.266% | 7.937% | 4.005% | 1.315% | 14.79%
| PDM3 |  | -3.519% | -2.15% | -3.119% | -2.301% | -75.56% | -7.091% | 15.27% | 4.393% | 13.21% | 13.17% | -1.716% | 0.5537% | -0.6815% | 0.8876% | -2.985% | -12.5% | 5.461% | 30.57% | -32.1% | 41.85% | -40.14% | -12.81% | 5.438% | -13.2%
| Z-Score (Absolute) |  | 0.14 | 0.02 | 0.11 | 0.03 | 6.72 | 0.47 | 1.57 | 0.58 | 1.38 | 1.38 | 0.02 | 0.23 | 0.11 | 0.26 | 0.1 | 0.97 | 0.68 | 2.97 | 2.75 | 4 | 3.49 | 0.99 | 0.67 | 1.03
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C856: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C874: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C947: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1131: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ALK
| Analytical results for S-(Sulphate) in OREAS 600c (Indicative Value 0.066 wt.%) | [18 empty] | Indicative
S-(Sulphate) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.25 | 2.26 |  | wt.%
|  |  | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | Na2CO3*IRC | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | Na2CO3*GRAV | I-550*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC | Na2CO3*GRAV |  | 3
3
| 1 | 1 | 0.05 | 0.11 | 0.15687 | 0.04755 | 0.16021 | 0.06 | 0.04 | 0.02 | 0.07 | 0.05 | 0.07 | 0.050065 | 0.17 | 0.18 | 0.15 | 0.05152 |  | 1
| 1 | 2 | 0.05 | 0.1 | 0.16688 | 0.03415 | 0.16021 | 0.06 | 0.05 | 0.04 | 0.07 | 0.03 | 0.07 | 0.060077 | 0.18 | 0.17 | 0.14 | 0.0598 |  | 7
| 1 | 3 | 0.05 | 0.1 | 0.15353 | 0.04627 | 0.16021 | 0.05 | 0.04 | 0.05 | 0.08 | 0.05 | 0.08 | 0.05674 | 0.17 | 0.18 | 0.14 | 0.0552 |  | 16
| 1 | 4 | 0.05 | 0.1 | 0.16354 | 0.0366 | 0.16021 | 0.04 | 0.05 | 0.04 | 0.07 | 0.04 | 0.08 | 0.060077 | 0.16 | 0.16 | 0.15 | 0.05704 |  | 0.0657
| 1 | 5 | 0.05 | 0.1 | 0.15687 | 0.0457 | 0.15687 | 0.06 | 0.04 | 0.03 | 0.08 | 0.08 | 0.09 | 0.063415 | 0.17 | 0.2 | 0.15 | 0.058752 |  | 13
| 1 | 6 | 0.05 | 0.09 | 0.17022 | 0.0407 | 0.16021 | 0.05 | 0.05 | 0.05 | 0.08 | 0.02 | 0.08 | 0.063415 | 0.18 | 0.25 | 0.15 | 0.05336
| Mean |  | 0.05 | 0.1 | 0.16132 | 0.041828 | 0.15965 | 0.053333 | 0.045 | 0.038333 | 0.075 | 0.045 | 0.078333 | 0.058965 | 0.17167 | 0.19 | 0.14667 | 0.055945
| Median |  | 0.05 | 0.1 | 0.16021 | 0.0432 | 0.16021 | 0.055 | 0.045 | 0.04 | 0.075 | 0.045 | 0.08 | 0.060077 | 0.17 | 0.18 | 0.15 | 0.05612
| Std Dev. |  | 7.6012e-18 | 0.0063246 | 0.0065631 | 0.0055681 | 0.0013626 | 0.008165 | 0.0054772 | 0.01169 | 0.0054772 | 0.020736 | 0.0075277 | 0.005025 | 0.0075277 | 0.032249 | 0.005164 | 0.003186
| Rel.Std.Dev. |  | 1.52e-14% | 6.325% | 4.068% | 13.31% | 0.8535% | 15.31% | 12.17% | 30.5% | 7.303% | 46.08% | 9.61% | 8.522% | 4.385% | 16.97% | 3.521% | 5.695%
| PDM3 |  | -23.9% | 52.21% | 145.5% | -36.33% | 143% | -18.82% | -31.51% | -41.65% | 14.15% | -31.51% | 19.23% | -10.25% | 161.3% | 189.2% | 123.2% | -14.85%
| Z-Score (Absolute) |  | 0.49 | 0.94 | 2.7 | 0.72 | 2.65 | 0.39 | 0.63 | 0.82 | 0.23 | 0.63 | 0.32 | 0.23 | 3 | 3.52 | 2.28 | 0.32
| Analytical results for S-(Sulphide) in OREAS 600c (Indicative Value 0.143 wt.%) | [18 empty] | Indicative
S-(Sulphide) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.25 | 2.26 | wt.%
|  |  | IRC-x*Calc | IRC-x*Calc | IRC-x*Calc | IRC | Na2CO3*IRC | IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | HCl*IRC | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | IRC | HCl*IRC | 3
3
| 1 | 1 | 0.26 | 0.13 | 0.09 | 0.21 | 0.19485 | 0.07 | 0.17 | 0.2 | 0.22 | 0.15 | 0.18 | 0.16 | 0.07 | 0.07 | < 0.05 | 0.06 | 0.1604 | 1
| 1 | 2 | 0.25 | 0.14 | 0.08 | 0.21 | 0.1948 | 0.07 | 0.18 | 0.19 | 0.2 | 0.14 | 0.14 | 0.16 | 0.04 | 0.06 | < 0.05 | 0.08 | 0.14829 | 7
| 1 | 3 | 0.27 | 0.13 | 0.08 | 0.22 | 0.19705 | 0.07 | 0.18 | 0.2 | 0.2 | 0.13 | 0.13 | 0.15 | 0.05 | 0.07 | < 0.05 | 0.07 | 0.15 | 16
| 1 | 4 | 0.25 | 0.13 | 0.08 | 0.21 | 0.1909 | 0.07 | 0.19 | 0.18 | 0.2 | 0.15 | 0.14 | 0.15 | 0.04 | 0.07 | < 0.05 | 0.07 | 0.17162 | 0.14305
| 1 | 5 | 0.26 | 0.13 | 0.08 | 0.2 | 0.1857 | 0.07 | 0.18 | 0.19 | 0.2 | 0.14 | 0.11 | 0.15 | 0.03 | 0.06 | < 0.05 | 0.06 | 0.17 | 13
| 1 | 6 | 0.26 | 0.14 | 0.07 | 0.21 | 0.1957 | 0.07 | 0.18 | 0.2 | 0.19 | 0.13 | 0.15 | 0.15 | 0.03 | 0.05 | < 0.05 | 0.06 | 0.15478
| Mean |  | 0.25833 | 0.13333 | 0.08 | 0.21 | 0.19317 | 0.07 | 0.18 | 0.19333 | 0.20167 | 0.14 | 0.14167 | 0.15333 | 0.043333 | 0.063333 |  | 0.066667 | 0.15918
| Median |  | 0.26 | 0.13 | 0.08 | 0.21 | 0.19482 | 0.07 | 0.18 | 0.195 | 0.2 | 0.14 | 0.14 | 0.15 | 0.04 | 0.065 |  | 0.065 | 0.15759
| Std Dev. |  | 0.0075277 | 0.005164 | 0.0063246 | 0.0063246 | 0.0041926 | 0 | 0.0063246 | 0.008165 | 0.0098319 | 0.0089443 | 0.023166 | 0.005164 | 0.015055 | 0.008165 |  | 0.008165 | 0.009956
| Rel.Std.Dev. |  | 2.914% | 3.873% | 7.906% | 3.012% | 2.17% | 0% | 3.514% | 4.223% | 4.875% | 6.389% | 16.35% | 3.368% | 34.74% | 12.89% |  | 12.25% | 6.254%
| PDM3 |  | 80.59% | -6.794% | -44.08% | 46.8% | 35.03% | -51.07% | 25.83% | 35.15% | 40.97% | -2.134% | -0.9689% | 7.187% | -69.71% | -55.73% |  | -53.4% | 11.28%
| Z-Score (Absolute) |  | 1.31 | 0.09 | 0.69 | 0.77 | 0.58 | 0.81 | 0.43 | 0.58 | 0.67 | 0.02 | 0 | 0.13 | 1.11 | 0.88 | 1.31 | 0.84 | 0.2
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Acid Leach
| Analytical results for S-(Sulphate) in OREAS 600c (Indicative Value 0.015 wt.%) |  |  | Indicative
S-(Sulphate) | Round | Replicate | Lab | 1
| No | No | 2.05 | wt.%
|  |  | HCl*OES | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.02 | 7
| 1 | 3 | 0.01 | 16
| 1 | 4 | 0.01 | 0.015
| 1 | 5 | 0.02 | 15
| 1 | 6 | 0.01
| Mean |  | 0.015
| Median |  | 0.015
| Std Dev. |  | 0.0054772
| Rel.Std.Dev. |  | 36.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 600c (Indicative Value 14.85 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.84 | 1
| 1 | 2 | 14.85 | 11
| Mean |  | 14.845 | 16
| Median |  | 14.845 | 14.845
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.04763%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 600c (Indicative Value 60 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 60 | 12
| Mean |  | 60 | 16
| Median |  | 60 | 60
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 600c (Indicative Value 3562 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 3562 | 1
| 1 | 2 | 3562 | 13
| Mean |  | 3562 | 16
| Median |  | 3562 | 3561.635
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.01025%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 600c (Indicative Value 1.73 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.72 | 1
| 1 | 2 | 1.73 | 14
| Mean |  | 1.725 | 16
| Median |  | 1.725 | 1.725
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.4099%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 600c (Indicative Value 25 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 40 | 15
| Mean |  | 25 | 16
| Median |  | 25 | 25
| Std Dev. |  | 21.2132 | 21
| Rel.Std.Dev. |  | 84.85%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 600c (Indicative Value < 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 16
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 22
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 600c (Indicative Value 29.2 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 44 | 1
| 1 | 2 | 15 | 17
| Mean |  | 29.5 | 16
| Median |  | 29.5 | 29.23
| Std Dev. |  | 20.5061 | 23
| Rel.Std.Dev. |  | 69.51%
| PDM3 |  | 0.9237%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 600c (Indicative Value 0.062 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.062 | 1
| 1 | 2 | 0.062 | 18
| Mean |  | 0.062 | 16
| Median |  | 0.062 | 0.062
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 600c (Indicative Value 3.58 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.603 | 1
| 1 | 2 | 3.56 | 55
| Mean |  | 3.5815 | 16
| Median |  | 3.5815 | 3.5814
| Std Dev. |  | 0.030406 | 17
| Rel.Std.Dev. |  | 0.849%
| PDM3 |  | 0.002834%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 600c (Indicative Value 3.78 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.77 | 1
| 1 | 2 | 3.78 | 12
| Mean |  | 3.775 | 16
| Median |  | 3.775 | 3.775
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.1873%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 600c (Indicative Value 0.28 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.27 | 1
| 1 | 2 | 0.29 | 13
| Mean |  | 0.28 | 16
| Median |  | 0.28 | 0.28
| Std Dev. |  | 0.014142 | 19
| Rel.Std.Dev. |  | 5.051%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 600c (Indicative Value 0.048 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.048 | 1
| 1 | 2 | 0.047 | 14
| Mean |  | 0.0475 | 16
| Median |  | 0.0475 | 0.0475
| Std Dev. |  | 0.00070711 | 20
| Rel.Std.Dev. |  | 1.489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 600c (Indicative Value 3.75 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.74 | 1
| 1 | 2 | 3.75 | 15
| Mean |  | 3.745 | 16
| Median |  | 3.745 | 3.745
| Std Dev. |  | 0.0070711 | 21
| Rel.Std.Dev. |  | 0.1888%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 600c (Indicative Value 10 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 16
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 600c (Indicative Value 0.082 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.0825 | 1
| 1 | 2 | 0.0825 | 17
| Mean |  | 0.0825 | 16
| Median |  | 0.0825 | 0.082498
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.002909%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 600c (Indicative Value 1120 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1120 | 1
| 1 | 2 | 1120 | 18
| Mean |  | 1120 | 16
| Median |  | 1120 | 1120
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 600c (Indicative Value 0.239 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.239 | 1
| 1 | 2 | 0.239 | 11
| Mean |  | 0.239 | 16
| Median |  | 0.239 | 0.239
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 600c (Indicative Value 69.03 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69 | 1
| 1 | 2 | 69.06 | 12
| Mean |  | 69.03 | 16
| Median |  | 69.03 | 69.03
| Std Dev. |  | 0.042426 | 18
| Rel.Std.Dev. |  | 0.06146%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 600c (Indicative Value 55 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 50 | 13
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 7.0711 | 19
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 600c (Indicative Value 205 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 200 | 1
| 1 | 2 | 210 | 14
| Mean |  | 205 | 16
| Median |  | 205 | 205
| Std Dev. |  | 7.0711 | 20
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 600c (Indicative Value 0.2 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 15
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 600c (Indicative Value 13.4 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 18 | 1
| 1 | 2 | 18 | 16
| Mean |  | 18 | 16
| Median |  | 18 | 13.389
| Std Dev. |  |  | 22
Rel.Std.Dev.
| PDM3 |  | 34.44%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 600c (Indicative Value 2055 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 2050 | 1
| 1 | 2 | 2060 | 17
| Mean |  | 2055 | 16
| Median |  | 2055 | 2055
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 0.3441%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 600c (Indicative Value 290 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 290 | 1
| 1 | 2 | 290 | 18
| Mean |  | 290 | 16
| Median |  | 290 | 290
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 600c (Indicative Value 1.56 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.58 | 1
| 1 | 2 | 1.54 | 20
| Mean |  | 1.56 | 16
| Median |  | 1.56 | 1.56
| Std Dev. |  | 0.028284 | 26
| Rel.Std.Dev. |  | 1.813%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 600c (Indicative Value 28.6 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27.3 | 1
| 1 | 2 | 29.8 | 22
| Mean |  | 28.55 | 16
| Median |  | 28.55 | 28.55
| Std Dev. |  | 1.7678 | 28
| Rel.Std.Dev. |  | 6.192%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 600c (Indicative Value 40.5 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 39.8 | 1
| 1 | 2 | 41.2 | 12
| Mean |  | 40.5 | 16
| Median |  | 40.5 | 40.5
| Std Dev. |  | 0.98995 | 29
| Rel.Std.Dev. |  | 2.444%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 600c (Indicative Value 3070 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 3020 | 1
| 1 | 2 | 3120 | 24
| Mean |  | 3070 | 16
| Median |  | 3070 | 3070
| Std Dev. |  | 70.7107 | 30
| Rel.Std.Dev. |  | 2.303%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 600c (Indicative Value 3.3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.6 | 1
| 1 | 2 | 3 | 25
| Mean |  | 3.3 | 16
| Median |  | 3.3 | 3.3
| Std Dev. |  | 0.42426 | 31
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 600c (Indicative Value 0.88 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.92 | 26
| Mean |  | 0.88 | 16
| Median |  | 0.88 | 0.88
| Std Dev. |  | 0.056569 | 32
| Rel.Std.Dev. |  | 6.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 600c (Indicative Value 7.5 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.5 | 1
| 1 | 2 | 7.5 | 27
| Mean |  | 7.5 | 16
| Median |  | 7.5 | 7.5
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 600c (Indicative Value 91 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 89.2 | 1
| 1 | 2 | 91.8 | 28
| Mean |  | 90.5 | 16
| Median |  | 90.5 | 90.5
| Std Dev. |  | 1.8385 | 34
| Rel.Std.Dev. |  | 2.031%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 600c (Indicative Value 2.75 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.7 | 16
| Mean |  | 2.75 | 16
| Median |  | 2.75 | 2.75
| Std Dev. |  | 0.070711 | 35
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 600c (Indicative Value 17 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14 | 1
| 1 | 2 | 20 | 30
| Mean |  | 17 | 16
| Median |  | 17 | 17
| Std Dev. |  | 4.2426 | 36
| Rel.Std.Dev. |  | 24.96%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 600c (Indicative Value 6.59 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.39 | 1
| 1 | 2 | 6.79 | 31
| Mean |  | 6.59 | 16
| Median |  | 6.59 | 6.59
| Std Dev. |  | 0.28284 | 37
| Rel.Std.Dev. |  | 4.292%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 600c (Indicative Value 0.059 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0574 | 1
| 1 | 2 | 0.061 | 18
| Mean |  | 0.0592 | 16
| Median |  | 0.0592 | 0.0592
| Std Dev. |  | 0.0025456 | 38
| Rel.Std.Dev. |  | 4.3%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 600c (Indicative Value 3.67 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.56 | 1
| 1 | 2 | 3.78 | 33
| Mean |  | 3.67 | 16
| Median |  | 3.67 | 3.67
| Std Dev. |  | 0.15556 | 39
| Rel.Std.Dev. |  | 4.239%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 600c (Indicative Value 1.08 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.06 | 1
| 1 | 2 | 1.09 | 34
| Mean |  | 1.075 | 16
| Median |  | 1.075 | 1.075
| Std Dev. |  | 0.021213 | 40
| Rel.Std.Dev. |  | 1.973%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 600c (Indicative Value 1.43 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.39 | 1
| 1 | 2 | 1.47 | 35
| Mean |  | 1.43 | 16
| Median |  | 1.43 | 1.43
| Std Dev. |  | 0.056569 | 41
| Rel.Std.Dev. |  | 3.956%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 600c (Indicative Value 21.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.6 | 1
| 1 | 2 | 22 | 36
| Mean |  | 21.8 | 16
| Median |  | 21.8 | 21.8
| Std Dev. |  | 0.28284 | 42
| Rel.Std.Dev. |  | 1.297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 600c (Indicative Value 6.09 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.05 | 1
| 1 | 2 | 6.13 | 37
| Mean |  | 6.09 | 16
| Median |  | 6.09 | 6.09
| Std Dev. |  | 0.056569 | 43
| Rel.Std.Dev. |  | 0.9289%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 600c (Indicative Value 1.33 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.45 | 38
| Mean |  | 1.325 | 16
| Median |  | 1.325 | 1.325
| Std Dev. |  | 0.17678 | 44
| Rel.Std.Dev. |  | 13.34%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 600c (Indicative Value 7.91 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.81 | 1
| 1 | 2 | 8 | 22
| Mean |  | 7.905 | 16
| Median |  | 7.905 | 7.905
| Std Dev. |  | 0.13435 | 28
| Rel.Std.Dev. |  | 1.7%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 600c (Indicative Value 0.54 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.52 | 1
| 1 | 2 | 0.56 | 23
| Mean |  | 0.54 | 16
| Median |  | 0.54 | 0.54
| Std Dev. |  | 0.028284 | 29
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 600c (Indicative Value 0.15 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.15 | 24
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 600c (Indicative Value 47.5 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46.4 | 1
| 1 | 2 | 48.5 | 25
| Mean |  | 47.45 | 16
| Median |  | 47.45 | 47.45
| Std Dev. |  | 1.4849 | 31
| Rel.Std.Dev. |  | 3.129%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 600c (Indicative Value 0.065 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.07 | 26
| Mean |  | 0.065 | 16
| Median |  | 0.065 | 0.065
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 600c (Indicative Value 0.04 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.0402 | 27
| Mean |  | 0.0401 | 16
| Median |  | 0.0401 | 0.0401
| Std Dev. |  | 0.00014142 | 33
| Rel.Std.Dev. |  | 0.3527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 600c (Indicative Value 3.9 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 4 | 28
| Mean |  | 3.9 | 16
| Median |  | 3.9 | 3.9
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 600c (Indicative Value 18.9 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.5 | 1
| 1 | 2 | 19.3 | 29
| Mean |  | 18.9 | 16
| Median |  | 18.9 | 18.9
| Std Dev. |  | 0.56569 | 35
| Rel.Std.Dev. |  | 2.993%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 600c (Indicative Value 41.2 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.5 | 1
| 1 | 2 | 41.8 | 30
| Mean |  | 41.15 | 16
| Median |  | 41.15 | 41.15
| Std Dev. |  | 0.91924 | 36
| Rel.Std.Dev. |  | 2.234%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 600c (Indicative Value 8 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8 | 1
| 1 | 2 | 8 | 16
| Mean |  | 8 | 16
| Median |  | 8 | 8
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 600c (Indicative Value 1195 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1170 | 1
| 1 | 2 | 1220 | 18
| Mean |  | 1195 | 16
| Median |  | 1195 | 1195
| Std Dev. |  | 35.3553 | 38
| Rel.Std.Dev. |  | 2.959%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 600c (Indicative Value 11.2 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.9 | 1
| 1 | 2 | 11.4 | 33
| Mean |  | 11.15 | 16
| Median |  | 11.15 | 11.15
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 3.171%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 600c (Indicative Value 144 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 141 | 1
| 1 | 2 | 147 | 34
| Mean |  | 144 | 16
| Median |  | 144 | 144
| Std Dev. |  | 4.2426 | 40
| Rel.Std.Dev. |  | 2.946%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 600c (Indicative Value 0.008 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | 0.01 | 35
| Mean |  | 0.01 | 16
| Median |  | 0.01 | 0.0075
| Std Dev. |  |  | 41
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 600c (Indicative Value 71 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 69.7 | 1
| 1 | 2 | 72.2 | 36
| Mean |  | 70.95 | 16
| Median |  | 70.95 | 70.95
| Std Dev. |  | 1.7678 | 42
| Rel.Std.Dev. |  | 2.492%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 600c (Indicative Value 4.3 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4.4 | 37
| Mean |  | 4.3 | 16
| Median |  | 4.3 | 4.3
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 3.289%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 600c (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 38
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 600c (Indicative Value 7.74 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.62 | 1
| 1 | 2 | 7.85 | 22
| Mean |  | 7.735 | 16
| Median |  | 7.735 | 7.735
| Std Dev. |  | 0.16263 | 28
| Rel.Std.Dev. |  | 2.103%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 600c (Indicative Value 4.2 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 4.6 | 13
| Mean |  | 4.2 | 16
| Median |  | 4.2 | 4.2
| Std Dev. |  | 0.56569 | 29
| Rel.Std.Dev. |  | 13.47%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 600c (Indicative Value 183 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 179 | 1
| 1 | 2 | 187 | 14
| Mean |  | 183 | 16
| Median |  | 183 | 183
| Std Dev. |  | 5.6569 | 30
| Rel.Std.Dev. |  | 3.091%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 600c (Indicative Value 1.44 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.43 | 1
| 1 | 2 | 1.44 | 25
| Mean |  | 1.435 | 16
| Median |  | 1.435 | 1.435
| Std Dev. |  | 0.0070711 | 31
| Rel.Std.Dev. |  | 0.4928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 600c (Indicative Value 0.81 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.81 | 26
| Mean |  | 0.805 | 16
| Median |  | 0.805 | 0.805
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 0.8784%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 600c (Indicative Value 0.3 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.2 | 27
| Mean |  | 0.3 | 16
| Median |  | 0.3 | 0.3
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 600c (Indicative Value 14.9 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.7 | 1
| 1 | 2 | 15 | 28
| Mean |  | 14.85 | 16
| Median |  | 14.85 | 14.85
| Std Dev. |  | 0.21213 | 34
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 600c (Indicative Value 0.118 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.116 | 1
| 1 | 2 | 0.119 | 29
| Mean |  | 0.1175 | 16
| Median |  | 0.1175 | 0.1175
| Std Dev. |  | 0.0021213 | 35
| Rel.Std.Dev. |  | 1.805%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 600c (Indicative Value 1.1 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1 | 30
| Mean |  | 1.1 | 16
| Median |  | 1.1 | 1.1
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 600c (Indicative Value 0.13 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.12 | 1
| 1 | 2 | 0.13 | 5
| Mean |  | 0.125 | 16
| Median |  | 0.125 | 0.125
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 600c (Indicative Value 5.91 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.81 | 1
| 1 | 2 | 6 | 32
| Mean |  | 5.905 | 16
| Median |  | 5.905 | 5.905
| Std Dev. |  | 0.13435 | 38
| Rel.Std.Dev. |  | 2.275%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 600c (Indicative Value 4.05 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 4.3 | 33
| Mean |  | 4.05 | 16
| Median |  | 4.05 | 4.05
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 8.73%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 600c (Indicative Value 1.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 1.5 | 34
| Mean |  | 1.75 | 16
| Median |  | 1.75 | 1.75
| Std Dev. |  | 0.35355 | 40
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 600c (Indicative Value 15.7 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.2 | 1
| 1 | 2 | 16.2 | 35
| Mean |  | 15.7 | 16
| Median |  | 15.7 | 15.7
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 4.504%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 600c (Indicative Value 0.59 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.56 | 1
| 1 | 2 | 0.62 | 36
| Mean |  | 0.59 | 16
| Median |  | 0.59 | 0.59
| Std Dev. |  | 0.042426 | 42
| Rel.Std.Dev. |  | 7.191%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 600c (Indicative Value 2275 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2230 | 1
| 1 | 2 | 2320 | 17
| Mean |  | 2275 | 16
| Median |  | 2275 | 2275
| Std Dev. |  | 63.6396 | 43
| Rel.Std.Dev. |  | 2.797%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 600c (Indicative Value 288 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 285 | 1
| 1 | 2 | 291 | 18
| Mean |  | 288 | 16
| Median |  | 288 | 288
| Std Dev. |  | 4.2426 | 44
| Rel.Std.Dev. |  | 1.473%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 600c (Indicative Value 24.5 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 24.9 | 1
| 1 | 2 | 24.7 | 22
| 1 | 3 | 24.6 | 16
| 1 | 4 | 24.2 | 24.5333
| 1 | 5 | 24.4 | 46
| 1 | 6 | 24.4
| Mean |  | 24.5333
| Median |  | 24.5
| Std Dev. |  | 0.25033
| Rel.Std.Dev. |  | 1.02%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 600c (Indicative Value 14.25 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 14.4 | 1
| 1 | 2 | 14.6 | 11
| 1 | 3 | 14.1 | 16
| 1 | 4 | 14 | 14.25
| 1 | 5 | 14.2 | 47
| 1 | 6 | 14.2
| Mean |  | 14.25
| Median |  | 14.2
| Std Dev. |  | 0.21679
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 600c (Indicative Value 2908 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 2920 | 1
| 1 | 2 | 2970 | 24
| 1 | 3 | 2900 | 16
| 1 | 4 | 2850 | 2908.3333
| 1 | 5 | 2910 | 48
| 1 | 6 | 2900
| Mean |  | 2908.3333
| Median |  | 2905
| Std Dev. |  | 38.6868
| Rel.Std.Dev. |  | 1.33%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 600c (Indicative Value 3.2 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.3 | 1
| 1 | 2 | 3.2 | 25
| 1 | 3 | 3.2 | 16
| 1 | 4 | 3.2 | 3.2
| 1 | 5 | 3.1 | 49
| 1 | 6 | 3.2
| Mean |  | 3.2
| Median |  | 3.2
| Std Dev. |  | 0.063246
| Rel.Std.Dev. |  | 1.976%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 600c (Indicative Value 1.03 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 0.9 | 26
| 1 | 3 | 1 | 16
| 1 | 4 | 0.8 | 1.0333
| 1 | 5 | 0.9 | 50
| 1 | 6 | 0.8
| Mean |  | 1.0333
| Median |  | 0.9
| Std Dev. |  | 0.38297
| Rel.Std.Dev. |  | 37.06%
| PDM3 |  | 3.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 600c (Indicative Value 1.75 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 1.76 | 1
| 1 | 2 | 1.79 | 14
| 1 | 3 | 1.74 | 16
| 1 | 4 | 1.72 | 1.75
| 1 | 5 | 1.75 | 51
| 1 | 6 | 1.74
| Mean |  | 1.75
| Median |  | 1.745
| Std Dev. |  | 0.023664
| Rel.Std.Dev. |  | 1.352%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 600c (Indicative Value 6.22 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.3 | 1
| 1 | 2 | 6.2 | 27
| 1 | 3 | 6.2 | 16
| 1 | 4 | 6.2 | 6.2167
| 1 | 5 | 6.2 | 52
| 1 | 6 | 6.2
| Mean |  | 6.2167
| Median |  | 6.2
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 0.6567%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 600c (Indicative Value 86 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 87 | 1
| 1 | 2 | 88 | 28
| 1 | 3 | 86 | 16
| 1 | 4 | 84 | 85.5
| 1 | 5 | 84 | 53
| 1 | 6 | 84
| Mean |  | 85.5
| Median |  | 85
| Std Dev. |  | 1.7607
| Rel.Std.Dev. |  | 2.059%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 600c (Indicative Value 2.15 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.22 | 1
| 1 | 2 | 2.16 | 16
| 1 | 3 | 2.21 | 16
| 1 | 4 | 2.15 | 2.1517
| 1 | 5 | 2.07 | 54
| 1 | 6 | 2.1
| Mean |  | 2.1517
| Median |  | 2.155
| Std Dev. |  | 0.059133
| Rel.Std.Dev. |  | 2.748%
| PDM3 |  | -1.443e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 600c (Indicative Value 13.8 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 14 | 1
| 1 | 2 | 14 | 30
| 1 | 3 | 16 | 16
| 1 | 4 | 12 | 13.8333
| 1 | 5 | 14 | 55
| 1 | 6 | 13
| Mean |  | 13.8333
| Median |  | 14
| Std Dev. |  | 1.3292
| Rel.Std.Dev. |  | 9.608%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 600c (Indicative Value 6.4 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.4 | 1
| 1 | 2 | 6.5 | 31
| 1 | 3 | 6.5 | 16
| 1 | 4 | 6.3 | 6.4
| 1 | 5 | 6.4 | 56
| 1 | 6 | 6.3
| Mean |  | 6.4
| Median |  | 6.4
| Std Dev. |  | 0.089443
| Rel.Std.Dev. |  | 1.398%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 600c (Indicative Value 0.057 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*MS | 3
3
| 1 | 1 | 0.0587 | 1
| 1 | 2 | 0.0572 | 18
| 1 | 3 | 0.0571 | 16
| 1 | 4 | 0.0574 | 0.057467
| 1 | 5 | 0.0573 | 57
| 1 | 6 | 0.0571
| Mean |  | 0.057467
| Median |  | 0.05725
| Std Dev. |  | 0.00061536
| Rel.Std.Dev. |  | 1.071%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 600c (Indicative Value 3.65 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.71 | 1
| 1 | 2 | 3.7 | 33
| 1 | 3 | 3.63 | 16
| 1 | 4 | 3.59 | 3.65
| 1 | 5 | 3.69 | 58
| 1 | 6 | 3.58
| Mean |  | 3.65
| Median |  | 3.66
| Std Dev. |  | 0.057619
| Rel.Std.Dev. |  | 1.579%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 600c (Indicative Value 1.01 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.03 | 1
| 1 | 2 | 1.02 | 34
| 1 | 3 | 1.02 | 16
| 1 | 4 | 1.01 | 1.0133
| 1 | 5 | 1 | 59
| 1 | 6 | 1
| Mean |  | 1.0133
| Median |  | 1.015
| Std Dev. |  | 0.012111
| Rel.Std.Dev. |  | 1.195%
| PDM3 |  | 3.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 600c (Indicative Value 1.99 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 2.03 | 1
| 1 | 2 | 2.01 | 35
| 1 | 3 | 1.98 | 16
| 1 | 4 | 1.98 | 1.9933
| 1 | 5 | 2 | 60
| 1 | 6 | 1.96
| Mean |  | 1.9933
| Median |  | 1.99
| Std Dev. |  | 0.025033
| Rel.Std.Dev. |  | 1.256%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 600c (Indicative Value 3.37 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.39 | 1
| 1 | 2 | 3.44 | 55
| 1 | 3 | 3.38 | 16
| 1 | 4 | 3.3 | 3.3667
| 1 | 5 | 3.35 | 61
| 1 | 6 | 3.34
| Mean |  | 3.3667
| Median |  | 3.365
| Std Dev. |  | 0.048028
| Rel.Std.Dev. |  | 1.427%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 600c (Indicative Value 21.6 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 22 | 1
| 1 | 2 | 21.7 | 36
| 1 | 3 | 21.8 | 16
| 1 | 4 | 21.4 | 21.5833
| 1 | 5 | 21 | 62
| 1 | 6 | 21.6
| Mean |  | 21.5833
| Median |  | 21.65
| Std Dev. |  | 0.34881
| Rel.Std.Dev. |  | 1.616%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 600c (Indicative Value 5.92 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.1 | 1
| 1 | 2 | 5.9 | 37
| 1 | 3 | 6 | 16
| 1 | 4 | 5.9 | 5.9167
| 1 | 5 | 5.9 | 46
| 1 | 6 | 5.7
| Mean |  | 5.9167
| Median |  | 5.9
| Std Dev. |  | 0.13292
| Rel.Std.Dev. |  | 2.246%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 600c (Indicative Value 6.73 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.8 | 1
| 1 | 2 | 6.8 | 22
| 1 | 3 | 6.8 | 16
| 1 | 4 | 6.7 | 6.7333
| 1 | 5 | 6.7 | 47
| 1 | 6 | 6.6
| Mean |  | 6.7333
| Median |  | 6.75
| Std Dev. |  | 0.08165
| Rel.Std.Dev. |  | 1.213%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 600c (Indicative Value 0.45 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.46 | 1
| 1 | 2 | 0.45 | 23
| 1 | 3 | 0.46 | 16
| 1 | 4 | 0.44 | 0.45
| 1 | 5 | 0.45 | 48
| 1 | 6 | 0.44
| Mean |  | 0.45
| Median |  | 0.45
| Std Dev. |  | 0.0089443
| Rel.Std.Dev. |  | 1.988%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 600c (Indicative Value 3.64 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.65 | 1
| 1 | 2 | 3.73 | 12
| 1 | 3 | 3.61 | 16
| 1 | 4 | 3.59 | 3.6417
| 1 | 5 | 3.62 | 49
| 1 | 6 | 3.65
| Mean |  | 3.6417
| Median |  | 3.635
| Std Dev. |  | 0.04916
| Rel.Std.Dev. |  | 1.35%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 600c (Indicative Value 46.8 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 48 | 1
| 1 | 2 | 48 | 25
| 1 | 3 | 47 | 16
| 1 | 4 | 45 | 46.8333
| 1 | 5 | 47 | 50
| 1 | 6 | 46
| Mean |  | 46.8333
| Median |  | 47
| Std Dev. |  | 1.169
| Rel.Std.Dev. |  | 2.496%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Li in OREAS 600c (Indicative Value 30.3 ppm) |  |  | Indicative
Li | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 1
1
| 1 | 1 | 31 | 1
| 1 | 2 | 31 | 45
| 1 | 3 | 30 | 16
| 1 | 4 | 30 | 30.3333
| 1 | 5 | 30 | 51
| 1 | 6 | 30
| Mean |  | 30.3333
| Median |  | 30
| Std Dev. |  | 0.5164
| Rel.Std.Dev. |  | 1.702%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 600c (Indicative Value 0.25 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.252 | 1
| 1 | 2 | 0.257 | 13
| 1 | 3 | 0.252 | 16
| 1 | 4 | 0.244 | 0.25017
| 1 | 5 | 0.248 | 52
| 1 | 6 | 0.248
| Mean |  | 0.25017
| Median |  | 0.25
| Std Dev. |  | 0.0044907
| Rel.Std.Dev. |  | 1.795%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 600c (Indicative Value 0.049 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.049 | 1
| 1 | 2 | 0.05 | 14
| 1 | 3 | 0.049 | 16
| 1 | 4 | 0.048 | 0.049
| 1 | 5 | 0.049 | 53
| 1 | 6 | 0.049
| Mean |  | 0.049
| Median |  | 0.049
| Std Dev. |  | 0.00063246
| Rel.Std.Dev. |  | 1.291%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 600c (Indicative Value 4.17 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.36 | 1
| 1 | 2 | 4.28 | 28
| 1 | 3 | 4.37 | 16
| 1 | 4 | 3.91 | 4.1717
| 1 | 5 | 4.1 | 54
| 1 | 6 | 4.01
| Mean |  | 4.1717
| Median |  | 4.19
| Std Dev. |  | 0.19302
| Rel.Std.Dev. |  | 4.627%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 600c (Indicative Value 3.73 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.72 | 1
| 1 | 2 | 3.83 | 15
| 1 | 3 | 3.7 | 16
| 1 | 4 | 3.68 | 3.725
| 1 | 5 | 3.69 | 55
| 1 | 6 | 3.73
| Mean |  | 3.725
| Median |  | 3.71
| Std Dev. |  | 0.054681
| Rel.Std.Dev. |  | 1.468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 600c (Indicative Value 18.6 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 18.8 | 1
| 1 | 2 | 18.8 | 29
| 1 | 3 | 18.9 | 16
| 1 | 4 | 18.2 | 18.6
| 1 | 5 | 18.4 | 56
| 1 | 6 | 18.5
| Mean |  | 18.6
| Median |  | 18.65
| Std Dev. |  | 0.27568
| Rel.Std.Dev. |  | 1.482%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 600c (Indicative Value 37.9 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 38.6 | 1
| 1 | 2 | 38.6 | 30
| 1 | 3 | 38.1 | 16
| 1 | 4 | 37.3 | 37.9333
| 1 | 5 | 37.8 | 57
| 1 | 6 | 37.2
| Mean |  | 37.9333
| Median |  | 37.95
| Std Dev. |  | 0.6121
| Rel.Std.Dev. |  | 1.614%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 600c (Indicative Value 6.43 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 6.3 | 16
| 1 | 3 | 6.9 | 16
| 1 | 4 | 6.3 | 6.4333
| 1 | 5 | 6.2 | 58
| 1 | 6 | 6.3
| Mean |  | 6.4333
| Median |  | 6.3
| Std Dev. |  | 0.26583
| Rel.Std.Dev. |  | 4.132%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 600c (Indicative Value 0.084 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.083 | 1
| 1 | 2 | 0.086 | 17
| 1 | 3 | 0.086 | 16
| 1 | 4 | 0.083 | 0.084167
| 1 | 5 | 0.083 | 59
| 1 | 6 | 0.084
| Mean |  | 0.084167
| Median |  | 0.0835
| Std Dev. |  | 0.001472
| Rel.Std.Dev. |  | 1.749%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 600c (Indicative Value 1100 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 1120 | 1
| 1 | 2 | 1100 | 18
| 1 | 3 | 1100 | 16
| 1 | 4 | 1090 | 1100
| 1 | 5 | 1100 | 60
| 1 | 6 | 1090
| Mean |  | 1100
| Median |  | 1100
| Std Dev. |  | 10.9545
| Rel.Std.Dev. |  | 0.9959%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 600c (Indicative Value 10.4 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 10.6 | 1
| 1 | 2 | 10.5 | 33
| 1 | 3 | 10.4 | 16
| 1 | 4 | 10.3 | 10.3833
| 1 | 5 | 10.4 | 61
| 1 | 6 | 10.1
| Mean |  | 10.3833
| Median |  | 10.4
| Std Dev. |  | 0.17224
| Rel.Std.Dev. |  | 1.659%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 600c (Indicative Value 146 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 148 | 1
| 1 | 2 | 148 | 34
| 1 | 3 | 146 | 16
| 1 | 4 | 145 | 145.6667
| 1 | 5 | 144 | 62
| 1 | 6 | 143
| Mean |  | 145.6667
| Median |  | 145.5
| Std Dev. |  | 2.0656
| Rel.Std.Dev. |  | 1.418%
| PDM3 |  | -2.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 600c (Indicative Value 0.225 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.224 | 1
| 1 | 2 | 0.23 | 11
| 1 | 3 | 0.228 | 16
| 1 | 4 | 0.223 | 0.22483
| 1 | 5 | 0.219 | 46
| 1 | 6 | 0.225
| Mean |  | 0.22483
| Median |  | 0.2245
| Std Dev. |  | 0.0038687
| Rel.Std.Dev. |  | 1.721%
| PDM3 |  | 1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 600c (Indicative Value 4.17 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4.2 | 37
| 1 | 3 | 4.3 | 16
| 1 | 4 | 4.1 | 4.1667
| 1 | 5 | 4.1 | 47
| 1 | 6 | 4.1
| Mean |  | 4.1667
| Median |  | 4.15
| Std Dev. |  | 0.08165
| Rel.Std.Dev. |  | 1.96%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 600c (Indicative Value 7.37 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 7.5 | 1
| 1 | 2 | 7.5 | 22
| 1 | 3 | 7.4 | 16
| 1 | 4 | 7.4 | 7.3667
| 1 | 5 | 7.3 | 48
| 1 | 6 | 7.1
| Mean |  | 7.3667
| Median |  | 7.4
| Std Dev. |  | 0.15055
| Rel.Std.Dev. |  | 2.044%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 600c (Indicative Value 3.93 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 3.95 | 1
| 1 | 2 | 4 | 13
| 1 | 3 | 3.9 | 16
| 1 | 4 | 3.68 | 3.925
| 1 | 5 | 3.75 | 49
| 1 | 6 | 4.27
| Mean |  | 3.925
| Median |  | 3.925
| Std Dev. |  | 0.20811
| Rel.Std.Dev. |  | 5.302%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 600c (Indicative Value 187 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 187 | 1
| 1 | 2 | 192 | 14
| 1 | 3 | 186 | 16
| 1 | 4 | 185 | 187.3333
| 1 | 5 | 186 | 50
| 1 | 6 | 188
| Mean |  | 187.3333
| Median |  | 186.5
| Std Dev. |  | 2.5033
| Rel.Std.Dev. |  | 1.336%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 600c (Indicative Value 1.37 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.36 | 25
| 1 | 3 | 1.37 | 16
| 1 | 4 | 1.35 | 1.3683
| 1 | 5 | 1.37 | 51
| 1 | 6 | 1.36
| Mean |  | 1.3683
| Median |  | 1.365
| Std Dev. |  | 0.017224
| Rel.Std.Dev. |  | 1.259%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 600c (Indicative Value 0.72 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.74 | 1
| 1 | 2 | 0.71 | 26
| 1 | 3 | 0.73 | 16
| 1 | 4 | 0.71 | 0.71667
| 1 | 5 | 0.71 | 52
| 1 | 6 | 0.7
| Mean |  | 0.71667
| Median |  | 0.71
| Std Dev. |  | 0.015055
| Rel.Std.Dev. |  | 2.101%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 600c (Indicative Value 13.5 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 13.9 | 1
| 1 | 2 | 13.7 | 28
| 1 | 3 | 13.4 | 16
| 1 | 4 | 13.4 | 13.5333
| 1 | 5 | 13.4 | 53
| 1 | 6 | 13.4
| Mean |  | 13.5333
| Median |  | 13.4
| Std Dev. |  | 0.21602
| Rel.Std.Dev. |  | 1.596%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 600c (Indicative Value 0.184 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 2.25 | wt.%
|  |  | 3A*OES/MS | 3
3
| 1 | 1 | 0.185 | 1
| 1 | 2 | 0.189 | 15
| 1 | 3 | 0.184 | 16
| 1 | 4 | 0.18 | 0.18417
| 1 | 5 | 0.184 | 54
| 1 | 6 | 0.183
| Mean |  | 0.18417
| Median |  | 0.184
| Std Dev. |  | 0.0029269
| Rel.Std.Dev. |  | 1.589%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 600c (Indicative Value 5.71 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 5.75 | 32
| 1 | 3 | 5.74 | 16
| 1 | 4 | 5.7 | 5.705
| 1 | 5 | 5.6 | 55
| 1 | 6 | 5.64
| Mean |  | 5.705
| Median |  | 5.72
| Std Dev. |  | 0.074229
| Rel.Std.Dev. |  | 1.301%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 600c (Indicative Value 3.8 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3.7 | 33
| 1 | 3 | 3.8 | 16
| 1 | 4 | 3.8 | 3.8
| 1 | 5 | 4 | 56
| 1 | 6 | 4
| Mean |  | 3.8
| Median |  | 3.8
| Std Dev. |  | 0.18974
| Rel.Std.Dev. |  | 4.993%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 600c (Indicative Value 1.8 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.9 | 34
| 1 | 3 | 1.8 | 16
| 1 | 4 | 1.8 | 1.8
| 1 | 5 | 1.7 | 57
| 1 | 6 | 1.8
| Mean |  | 1.8
| Median |  | 1.8
| Std Dev. |  | 0.063246
| Rel.Std.Dev. |  | 3.514%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 600c (Indicative Value 14.9 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 1
1
| 1 | 1 | 15.1 | 1
| 1 | 2 | 15 | 35
| 1 | 3 | 15.1 | 16
| 1 | 4 | 14.8 | 14.9
| 1 | 5 | 14.6 | 58
| 1 | 6 | 14.8
| Mean |  | 14.9
| Median |  | 14.9
| Std Dev. |  | 0.2
| Rel.Std.Dev. |  | 1.342%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 600c (Indicative Value 0.56 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 2
2
| 1 | 1 | 0.57 | 1
| 1 | 2 | 0.57 | 36
| 1 | 3 | 0.57 | 16
| 1 | 4 | 0.56 | 0.56333
| 1 | 5 | 0.56 | 59
| 1 | 6 | 0.55
| Mean |  | 0.56333
| Median |  | 0.565
| Std Dev. |  | 0.008165
| Rel.Std.Dev. |  | 1.449%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 600c (Indicative Value 2082 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*MS | 0
0
| 1 | 1 | 2120 | 1
| 1 | 2 | 2080 | 17
| 1 | 3 | 2080 | 16
| 1 | 4 | 2080 | 2081.6667
| 1 | 5 | 2060 | 60
| 1 | 6 | 2070
| Mean |  | 2081.6667
| Median |  | 2080
| Std Dev. |  | 20.4124
| Rel.Std.Dev. |  | 0.9806%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 600c (Indicative Value 301 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 2.25 | ppm
|  |  | 3A*OES/MS | 0
0
| 1 | 1 | 305 | 1
| 1 | 2 | 308 | 18
| 1 | 3 | 294 | 16
| 1 | 4 | 298 | 300.8333
| 1 | 5 | 300 | 61
| 1 | 6 | 300
| Mean |  | 300.8333
| Median |  | 300
| Std Dev. |  | 4.9967
| Rel.Std.Dev. |  | 1.661%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.