## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 251c
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.50774 | 0.50164 | 0.51385 | 0.50469 | 0.5108
PhotonAssay
Au, Gold (ppm) | 0.50072 | 0.49536 | 0.50609 | 0.49984 | 0.50161
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.4854 | 0.47648 | 0.49432 | 0.4822 | 0.4886
Cyanide Leach
Au, Gold (ppm) | 0.49227 | 0.48371 | 0.50083 | 0.49112 | 0.49342
4-Acid Digestion
Ag, Silver (ppm) | 0.29636 | 0.26054 | 0.33218 | 0.26766 | 0.32506
Al, Aluminium (wt.%) | 6.7694 | 6.5979 | 6.9409 | 6.6753 | 6.8636
As, Arsenic (ppm) | 174.7228 | 168.9221 | 180.5234 | 170.7751 | 178.6704
Ba, Barium (ppm) | 508.9435 | 495.8546 | 522.0324 | 499.4432 | 518.4438
Be, Beryllium (ppm) | 2.3931 | 2.1921 | 2.5942 | 2.3258 | 2.4604
Bi, Bismuth (ppm) | 7.5227 | 6.8967 | 8.1486 | 7.1011 | 7.9443
Ca, Calcium (wt.%) | 1.7149 | 1.6653 | 1.7646 | 1.6786 | 1.7512
Cd, Cadmium (ppm) | 0.051007 | 0.025186 | 0.076828 | 0.041898 | 0.060116
Ce, Cerium (ppm) | 88.3028 | 83.8338 | 92.7717 | 86.2821 | 90.3234
Co, Cobalt (ppm) | 17.1892 | 16.5426 | 17.8357 | 16.7009 | 17.6774
Cr, Chromium (ppm) | 119.6819 | 113.8583 | 125.5055 | 115.9067 | 123.4571
Cs, Caesium (ppm) | 5.9419 | 5.6996 | 6.1842 | 5.7869 | 6.0969
Cu, Copper (ppm) | 27.5167 | 26.0838 | 28.9496 | 26.684 | 28.3493
Dy, Dysprosium (ppm) | 4.2788 | 4.0318 | 4.5257 | 4.14 | 4.4175
Er, Erbium (ppm) | 1.9597 | 1.7814 | 2.138 | 1.8861 | 2.0333
Eu, Europium (ppm) | 1.6785 | 1.5588 | 1.7982 | 1.6193 | 1.7378
Fe, Iron (wt.%) | 4.2771 | 4.1657 | 4.3885 | 4.2196 | 4.3345
Ga, Gallium (ppm) | 18.8914 | 17.9972 | 19.7855 | 18.4845 | 19.2982
Gd, Gadolinium (ppm) | 5.8003 | 5.4852 | 6.1155 | 5.5981 | 6.0026
Hf, Hafnium (ppm) | 4.7008 | 4.4527 | 4.949 | 4.5533 | 4.8483
Ho, Holmium (ppm) | 0.72281 | 0.64023 | 0.80538 | 0.69091 | 0.75471
In, Indium (ppm) | 0.087477 | 0.078773 | 0.096181 | 0.077063 | 0.09789
K, Potassium (wt.%) | 1.8743 | 1.8143 | 1.9343 | 1.8275 | 1.9211
La, Lanthanum (ppm) | 45.3112 | 43.0033 | 47.6191 | 44.2949 | 46.3275
Li, Lithium (ppm) | 30.5942 | 29.4167 | 31.7717 | 30.0426 | 31.1459
Lu, Lutetium (ppm) | 0.26326 | 0.22959 | 0.29693 | 0.23489 | 0.29163
Mg, Magnesium (wt.%) | 1.3663 | 1.3221 | 1.4104 | 1.3477 | 1.3848
Mn, Manganese (wt.%) | 0.04147 | 0.040185 | 0.042755 | 0.040621 | 0.042319
Mo, Molybdenum (ppm) | 2.4831 | 2.319 | 2.6471 | 2.3378 | 2.6283
Na, Sodium (wt.%) | 0.95667 | 0.92978 | 0.98357 | 0.93873 | 0.97462
Nb, Niobium (ppm) | 26.3304 | 24.6685 | 27.9923 | 25.1409 | 27.5199
Nd, Neodymium (ppm) | 38.4355 | 36.4314 | 40.4395 | 37.7375 | 39.1334
Ni, Nickel (ppm) | 72.0797 | 69.837 | 74.3224 | 70.3027 | 73.8567
P, Phosphorus (wt.%) | 0.09159 | 0.088801 | 0.094379 | 0.089818 | 0.093362
Pb, Lead (ppm) | 13.3326 | 12.557 | 14.1083 | 12.7849 | 13.8804
Pr, Praseodymium (ppm) | 10.2291 | 9.6268 | 10.8314 | 10.0536 | 10.4047
Rb, Rubidium (ppm) | 99.0788 | 95.5166 | 102.641 | 96.9949 | 101.1627
S, Sulphur (wt.%) | 0.0097975 | 0.0093442 | 0.010251 | 0.009133 | 0.010462
Sb, Antimony (ppm) | 2.6538 | 2.4758 | 2.8317 | 2.4705 | 2.8371
Sc, Scandium (ppm) | 12.931 | 12.2454 | 13.6166 | 12.6531 | 13.2089
Sm, Samarium (ppm) | 7.3446 | 6.8871 | 7.802 | 7.1997 | 7.4894
Sn, Tin (ppm) | 8.9454 | 8.3274 | 9.5633 | 8.5946 | 9.2961
Sr, Strontium (ppm) | 280.0794 | 272.1105 | 288.0483 | 274.2641 | 285.8947
Ta, Tantalum (ppm) | 1.8329 | 1.7165 | 1.9493 | 1.7447 | 1.9211
Tb, Terbium (ppm) | 0.76538 | 0.71438 | 0.81638 | 0.73451 | 0.79625
Te, Tellurium (ppm) | 0.12144 | 0.090889 | 0.15199 | IND | IND
Th, Thorium (ppm) | 12.2968 | 11.7217 | 12.872 | 12.021 | 12.5727
Ti, Titanium (wt.%) | 0.59256 | 0.57156 | 0.61356 | 0.578 | 0.60712
Tl, Thallium (ppm) | 0.52146 | 0.49147 | 0.55145 | 0.49892 | 0.54399
Tm, Thulium (ppm) | 0.25419 | 0.22293 | 0.28545 | 0.23607 | 0.27231
U, Uranium (ppm) | 2.1802 | 2.0568 | 2.3037 | 2.1202 | 2.2403
V, Vanadium (ppm) | 91.9751 | 88.7369 | 95.2134 | 89.534 | 94.4163
W, Tungsten (ppm) | 31.4592 | 29.0305 | 33.8879 | 29.8467 | 33.0717
Y, Yttrium (ppm) | 18.3027 | 17.2415 | 19.364 | 17.761 | 18.8445
Yb, Ytterbium (ppm) | 1.7247 | 1.577 | 1.8725 | 1.6367 | 1.8128
Zn, Zinc (ppm) | 69.4913 | 67.0884 | 71.8943 | 67.5778 | 71.4049
Zr, Zirconium (ppm) | 182.9031 | 175.6893 | 190.1169 | 176.07 | 189.7361
Aqua Regia Digestion
Ag, Silver (ppm) | 0.26894 | 0.24318 | 0.29469 | 0.24654 | 0.29133
Al, Aluminium (wt.%) | 1.126 | 1.0604 | 1.1917 | 1.0851 | 1.167
As, Arsenic (ppm) | 153.1763 | 148.3924 | 157.9601 | 150.2959 | 156.0566
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 71.9337 | 67.2959 | 76.5716 | 68.3285 | 75.5389
Be, Beryllium (ppm) | 0.93287 | 0.87872 | 0.98701 | 0.89282 | 0.97292
Bi, Bismuth (ppm) | 6.8278 | 6.2538 | 7.4018 | 6.4306 | 7.2249
Ca, Calcium (wt.%) | 0.43656 | 0.41013 | 0.46298 | 0.42161 | 0.4515
Cd, Cadmium (ppm) | 0.025788 | 0.014902 | 0.036674 | IND | IND
Ce, Cerium (ppm) | 52.1856 | 48.514 | 55.8571 | 50.8529 | 53.5183
Co, Cobalt (ppm) | 11.5874 | 10.7759 | 12.3989 | 11.2066 | 11.9681
Cr, Chromium (ppm) | 46.2617 | 42.6865 | 49.837 | 44.7992 | 47.7243
Cs, Caesium (ppm) | 1.4341 | 1.3042 | 1.5639 | 1.3805 | 1.4877
Cu, Copper (ppm) | 22.665 | 21.5667 | 23.7634 | 22.0426 | 23.2875
Eu, Europium (ppm) | 0.9406 | 0.81566 | 1.0655 | 0.90714 | 0.97405
Fe, Iron (wt.%) | 3.0318 | 2.872 | 3.1916 | 2.963 | 3.1006
Ga, Gallium (ppm) | 3.9416 | 3.6267 | 4.2566 | 3.7306 | 4.1527
Ho, Holmium (ppm) | 0.38057 | 0.32804 | 0.43311 | 0.35363 | 0.40751
In, Indium (ppm) | 0.025905 | 0.02276 | 0.02905 | 0.023419 | 0.028391
K, Potassium (wt.%) | 0.18638 | 0.17202 | 0.20073 | 0.17663 | 0.19613
La, Lanthanum (ppm) | 26.6631 | 24.8509 | 28.4753 | 26.006 | 27.3201
Li, Lithium (ppm) | 6.7796 | 6.3045 | 7.2547 | 6.5319 | 7.0272
Mg, Magnesium (wt.%) | 0.53883 | 0.5088 | 0.56887 | 0.52079 | 0.55688
Mn, Manganese (wt.%) | 0.02267 | 0.021062 | 0.024277 | 0.022107 | 0.023232
Mo, Molybdenum (ppm) | 1.9552 | 1.8223 | 2.0882 | 1.8698 | 2.0407
Na, Sodium (wt.%) | 0.19166 | 0.18197 | 0.20134 | 0.18168 | 0.20163
Nb, Niobium (ppm) | 0.61826 | 0.46529 | 0.77122 | 0.55817 | 0.67834
Ni, Nickel (ppm) | 56.639 | 54.3283 | 58.9496 | 55.0544 | 58.2235
P, Phosphorus (wt.%) | 0.064912 | 0.06235 | 0.067473 | 0.063551 | 0.066272
Pb, Lead (ppm) | 8.009 | 7.1967 | 8.8213 | 7.5205 | 8.4974
Pr, Praseodymium (ppm) | 7.4 | 4.9942 | 9.8058 | 7.198 | 7.602
Rb, Rubidium (ppm) | 13.0655 | 12.2651 | 13.8658 | 12.3971 | 13.7338
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.0097259 | 0.0074595 | 0.011992 | 0.0088508 | 0.010601
Sb, Antimony (ppm) | 1.6559 | 1.5158 | 1.7959 | 1.6015 | 1.7102
Sc, Scandium (ppm) | 3.011 | 2.7579 | 3.2642 | 2.8887 | 3.1333
Sn, Tin (ppm) | 1.4556 | 1.3279 | 1.5833 | 1.3418 | 1.5694
Sr, Strontium (ppm) | 39.2312 | 36.5657 | 41.8966 | 37.7749 | 40.6874
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.4887 | 0.36387 | 0.61353 | 0.45865 | 0.51876
Te, Tellurium (ppm) | 0.082578 | 0.064701 | 0.10045 | IND | IND
Th, Thorium (ppm) | 8.6724 | 8.0919 | 9.2529 | 8.4573 | 8.8875
Ti, Titanium (wt.%) | 0.1116 | 0.096409 | 0.1268 | 0.1064 | 0.11681
Tl, Thallium (ppm) | 0.088317 | 0.077213 | 0.099422 | 0.076761 | 0.099873
U, Uranium (ppm) | 1.0538 | 0.98591 | 1.1217 | 1.0134 | 1.0941
V, Vanadium (ppm) | 29.3866 | 27.7263 | 31.0469 | 28.1226 | 30.6506
W, Tungsten (ppm) | 9.7831 | 8.6727 | 10.8935 | 9.2218 | 10.3444
Y, Yttrium (ppm) | 9.1879 | 8.7211 | 9.6546 | 8.9466 | 9.4291
Zn, Zinc (ppm) | 42.1732 | 38.8241 | 45.5223 | 40.8517 | 43.4947
Zr, Zirconium (ppm) | 29.3942 | 24.469 | 34.3195 | 28.2639 | 30.5246
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 251c
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 24.6667 | Hg | ppm | 0.1425 | Re | ppm | 0.0013903
Ge | ppm | 0.18623 | Pt | ppb | 3.4 | Se | ppm | 0.83359
Aqua Regia Digestion
Dy | ppm | 2.2127 | Hg | ppm | 0.026767 | Se | ppm | 0.28687
Er | ppm | 0.86433 | Lu | ppm | 0.077131 | Sm | ppm | 4.5515
Gd | ppm | 3.589 | Nd | ppm | 24.9775 | Tm | ppm | 0.10095
Ge | ppm | 0.125 | Pd | ppb | < 10 | Yb | ppm | 0.58119
Hf | ppm | 0.5266 | Pt | ppb | < 5
Borate Fusion XRF
Al2O3 | wt.% | 13.495 | Fe2O3 | wt.% | 6.4337 | S | wt.% | 0.011
As | ppm | 190 | K2O | wt.% | 2.3 | SiO2 | wt.% | 67.61
BaO | ppm | 574.9975 | MgO | wt.% | 2.375 | Sn | ppm | 65
CaO | wt.% | 2.405 | MnO | wt.% | 0.0575 | Sr | ppm | 200
Cl | ppm | 135 | Na2O | wt.% | 1.315 | TiO2 | wt.% | 1.134
Co | ppm | 20 | Ni | ppm | 95 | V2O5 | ppm | 196.372
Cr2O3 | ppm | 204.61 | P2O5 | wt.% | 0.20739 | Zn | ppm | 70
Cu | ppm | 50 | Pb | ppm | 50 | Zr | ppm | 275
Thermogravimetry
LOI1000 | wt.% | 2.415
Infrared Combustion
C | wt.% | 0.085 | S | wt.% | 0.15
Laser Ablation ICP-MS
Ag | ppm | 0.2 | Hf | ppm | 7.34 | Sm | ppm | 7.755
As | ppm | 167 | Ho | ppm | 1.03 | Sn | ppm | 10.2
Ba | ppm | 505 | In | ppm | 0.075 | Sr | ppm | 276.5
Be | ppm | 2.3 | La | ppm | 46.45 | Ta | ppm | 2.04
Bi | ppm | 8 | Lu | ppm | 0.355 | Tb | ppm | 0.94
Cd | ppm | < 0.1 | Mn | wt.% | 0.0428 | Te | ppm | 0.2
Ce | ppm | 89 | Mo | ppm | 2.7 | Th | ppm | 12.8
Co | ppm | 18.1 | Nb | ppm | 29.95 | Ti | wt.% | 0.6815
Cr | ppm | 146 | Nd | ppm | 40.4 | Tl | ppm | < 0.2
Cs | ppm | 5.57 | Ni | ppm | 75 | Tm | ppm | 0.415
Cu | ppm | 30 | Pb | ppm | 13 | U | ppm | 2.525
Dy | ppm | 5.345 | Pr | ppm | 10.6 | V | ppm | 100
Er | ppm | 2.875 | Rb | ppm | 95.4 | W | ppm | 35
Eu | ppm | 1.795 | Re | ppm | < 0.01 | Y | ppm | 26.75
Ga | ppm | 18.6 | Sb | ppm | 2.8 | Yb | ppm | 2.63
Gd | ppm | 6.2 | Sc | ppm | 13.65 | Zn | ppm | 65
Ge | ppm | 1.375 | Se | ppm | < 5 | Zr | ppm | 276.5
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 251c
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.50774 | 0.014238 | 0.47927 | 0.53622 | 0.46503 | 0.55046 | 2.804% | 5.608% | 8.413% | 0.48236 | 0.53313
PhotonAssay
Au, ppm | 0.50072 | 0.033365 | 0.43399 | 0.56745 | 0.40063 | 0.60082 | 6.663% | 13.33% | 19.99% | 0.47569 | 0.52576
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.4854 | 0.022223 | 0.44096 | 0.52985 | 0.41873 | 0.55207 | 4.578% | 9.157% | 13.73% | 0.46113 | 0.50967
Cyanide Leach
Au, ppm | 0.49227 | 0.020589 | 0.45109 | 0.53345 | 0.4305 | 0.55404 | 4.182% | 8.365% | 12.55% | 0.46766 | 0.51688
4-Acid Digestion
Ag, ppm | 0.29636 | 0.028347 | 0.23967 | 0.35305 | 0.21132 | 0.3814 | 9.565% | 19.13% | 28.7% | 0.28154 | 0.31118
Al, wt.% | 6.7694 | 0.17063 | 6.4281 | 7.1107 | 6.2575 | 7.2813 | 2.521% | 5.041% | 7.562% | 6.4309 | 7.1079
As, ppm | 174.7228 | 5.2392 | 164.2443 | 185.2012 | 159.0051 | 190.4404 | 2.999% | 5.997% | 8.996% | 165.9866 | 183.4589
Ba, ppm | 508.9435 | 11.6749 | 485.5938 | 532.2932 | 473.9189 | 543.9681 | 2.294% | 4.588% | 6.882% | 483.4963 | 534.3907
Be, ppm | 2.3931 | 0.15766 | 2.0778 | 2.7084 | 1.9202 | 2.8661 | 6.588% | 13.18% | 19.76% | 2.2735 | 2.5128
Bi, ppm | 7.5227 | 0.55127 | 6.4201 | 8.6252 | 5.8689 | 9.1765 | 7.328% | 14.66% | 21.98% | 7.1465 | 7.8988
Ca, wt.% | 1.7149 | 0.039887 | 1.6351 | 1.7947 | 1.5953 | 1.8346 | 2.326% | 4.652% | 6.978% | 1.6292 | 1.8007
Cd, ppm | 0.051007 | 0.022614 | 0.0057784 | 0.096235 | 0 | 0.11885 | 44.34% | 88.67% | 133% | 0.048456 | 0.053557
Ce, ppm | 88.3028 | 4.3536 | 79.5955 | 97.0101 | 75.2418 | 101.3637 | 4.93% | 9.861% | 14.79% | 83.8876 | 92.7179
Co, ppm | 17.1892 | 0.58818 | 16.0128 | 18.3655 | 15.4246 | 18.9537 | 3.422% | 6.844% | 10.27% | 16.3297 | 18.0486
Cr, ppm | 119.6819 | 9.1065 | 101.469 | 137.8948 | 92.3625 | 147.0013 | 7.609% | 15.22% | 22.83% | 113.6978 | 125.666
Cs, ppm | 5.9419 | 0.23156 | 5.4788 | 6.405 | 5.2472 | 6.6366 | 3.897% | 7.794% | 11.69% | 5.6448 | 6.239
Cu, ppm | 27.5167 | 1.6508 | 24.2151 | 30.8182 | 22.5644 | 32.469 | 5.999% | 12% | 18% | 26.1408 | 28.8925
Dy, ppm | 4.2788 | 0.24489 | 3.789 | 4.7685 | 3.5441 | 5.0134 | 5.723% | 11.45% | 17.17% | 4.0648 | 4.4927
Er, ppm | 1.9597 | 0.15948 | 1.6407 | 2.2786 | 1.4812 | 2.4381 | 8.138% | 16.28% | 24.41% | 1.8617 | 2.0577
Eu, ppm | 1.6785 | 0.07982 | 1.5189 | 1.8382 | 1.4391 | 1.918 | 4.755% | 9.511% | 14.27% | 1.5946 | 1.7625
Fe, wt.% | 4.2771 | 0.11286 | 4.0514 | 4.5028 | 3.9385 | 4.6157 | 2.639% | 5.278% | 7.916% | 4.0632 | 4.4909
Ga, ppm | 18.8914 | 0.76769 | 17.356 | 20.4268 | 16.5883 | 21.1944 | 4.064% | 8.127% | 12.19% | 17.9468 | 19.8359
Gd, ppm | 5.8003 | 0.23036 | 5.3396 | 6.2611 | 5.1093 | 6.4914 | 3.971% | 7.943% | 11.91% | 5.5103 | 6.0904
Hf, ppm | 4.7008 | 0.20236 | 4.2961 | 5.1056 | 4.0937 | 5.3079 | 4.305% | 8.61% | 12.91% | 4.4658 | 4.9359
Ho, ppm | 0.72281 | 0.070592 | 0.58162 | 0.86399 | 0.51103 | 0.93458 | 9.766% | 19.53% | 29.3% | 0.68667 | 0.75895
In, ppm | 0.087477 | 0.0087491 | 0.069978 | 0.10497 | 0.061229 | 0.11372 | 10% | 20% | 30.01% | 0.083103 | 0.09185
K, wt.% | 1.8743 | 0.072312 | 1.7296 | 2.0189 | 1.6573 | 2.0912 | 3.858% | 7.716% | 11.57% | 1.7806 | 1.968
La, ppm | 45.3112 | 2.9039 | 39.5034 | 51.119 | 36.5996 | 54.0229 | 6.409% | 12.82% | 19.23% | 43.0457 | 47.5768
Li, ppm | 30.5942 | 1.6113 | 27.3717 | 33.8167 | 25.7604 | 35.428 | 5.267% | 10.53% | 15.8% | 29.0645 | 32.1239
Lu, ppm | 0.26326 | 0.021441 | 0.22038 | 0.30614 | 0.19894 | 0.32758 | 8.144% | 16.29% | 24.43% | 0.2501 | 0.27642
Mg, wt.% | 1.3663 | 0.057628 | 1.251 | 1.4815 | 1.1934 | 1.5391 | 4.218% | 8.436% | 12.65% | 1.2979 | 1.4346
Mn, wt.% | 0.04147 | 0.0018284 | 0.037813 | 0.045127 | 0.035985 | 0.046955 | 4.409% | 8.818% | 13.23% | 0.039397 | 0.043544
Mo, ppm | 2.4831 | 0.1829 | 2.1172 | 2.8489 | 1.9343 | 3.0318 | 7.366% | 14.73% | 22.1% | 2.3589 | 2.6072
Na, wt.% | 0.95667 | 0.037872 | 0.88093 | 1.0324 | 0.84306 | 1.0703 | 3.959% | 7.918% | 11.88% | 0.90884 | 1.0045
Nb, ppm | 26.3304 | 1.9519 | 22.4266 | 30.2343 | 20.4746 | 32.1862 | 7.413% | 14.83% | 22.24% | 25.0139 | 27.6469
Nd, ppm | 38.4355 | 1.7048 | 35.0259 | 41.845 | 33.3212 | 43.5498 | 4.435% | 8.871% | 13.31% | 36.5137 | 40.3572
Ni, ppm | 72.0797 | 2.475 | 67.1297 | 77.0297 | 64.6546 | 79.5047 | 3.434% | 6.867% | 10.3% | 68.4757 | 75.6837
P, wt.% | 0.09159 | 0.0025653 | 0.086459 | 0.096721 | 0.083894 | 0.099286 | 2.801% | 5.602% | 8.403% | 0.08701 | 0.096169
Pb, ppm | 13.3326 | 1.1307 | 11.0712 | 15.594 | 9.9405 | 16.7248 | 8.481% | 16.96% | 25.44% | 12.666 | 13.9993
Pr, ppm | 10.2291 | 0.53281 | 9.1635 | 11.2948 | 8.6307 | 11.8276 | 5.209% | 10.42% | 15.63% | 9.7177 | 10.7406
Rb, ppm | 99.0788 | 3.1829 | 92.7131 | 105.4445 | 89.5302 | 108.6274 | 3.212% | 6.425% | 9.637% | 94.1249 | 104.0328
S, wt.% | 0.0097975 | 0.00047819 | 0.0088412 | 0.010754 | 0.008363 | 0.011232 | 4.881% | 9.761% | 14.64% | 0.0093077 | 0.010287
Sb, ppm | 2.6538 | 0.18223 | 2.2893 | 3.0182 | 2.1071 | 3.2005 | 6.867% | 13.73% | 20.6% | 2.5211 | 2.7865
Sc, ppm | 12.931 | 0.61236 | 11.7063 | 14.1557 | 11.0939 | 14.7681 | 4.736% | 9.471% | 14.21% | 12.2844 | 13.5775
Sm, ppm | 7.3446 | 0.43679 | 6.471 | 8.2181 | 6.0342 | 8.6549 | 5.947% | 11.89% | 17.84% | 6.9773 | 7.7118
Sn, ppm | 8.9454 | 0.66616 | 7.613 | 10.2777 | 6.9469 | 10.9438 | 7.447% | 14.89% | 22.34% | 8.4981 | 9.3926
Sr, ppm | 280.0794 | 8.0313 | 264.0168 | 296.142 | 255.9854 | 304.1734 | 2.868% | 5.735% | 8.603% | 266.0754 | 294.0834
Ta, ppm | 1.8329 | 0.14411 | 1.5447 | 2.1211 | 1.4006 | 2.2652 | 7.862% | 15.72% | 23.59% | 1.7412 | 1.9245
Tb, ppm | 0.76538 | 0.030899 | 0.70358 | 0.82718 | 0.67268 | 0.85807 | 4.037% | 8.074% | 12.11% | 0.72711 | 0.80365
Te, ppm | 0.12144 | 0.027891 | 0.065656 | 0.17722 | 0.037765 | 0.20511 | 22.97% | 45.93% | 68.9% | 0.11537 | 0.12751
Th, ppm | 12.2968 | 0.7558 | 10.7852 | 13.8084 | 10.0294 | 14.5642 | 6.146% | 12.29% | 18.44% | 11.682 | 12.9117
Ti, wt.% | 0.59256 | 0.021211 | 0.55014 | 0.63498 | 0.52893 | 0.65619 | 3.58% | 7.159% | 10.74% | 0.56293 | 0.62219
Tl, ppm | 0.52146 | 0.029173 | 0.46311 | 0.5798 | 0.43394 | 0.60897 | 5.594% | 11.19% | 16.78% | 0.49538 | 0.54753
Tm, ppm | 0.25419 | 0.035665 | 0.18286 | 0.32552 | 0.1472 | 0.36119 | 14.03% | 28.06% | 42.09% | 0.24148 | 0.2669
U, ppm | 2.1802 | 0.16557 | 1.8491 | 2.5114 | 1.6835 | 2.677 | 7.594% | 15.19% | 22.78% | 2.0712 | 2.2892
V, ppm | 91.9751 | 2.7959 | 86.3834 | 97.5669 | 83.5876 | 100.3627 | 3.04% | 6.08% | 9.119% | 87.3764 | 96.5739
W, ppm | 31.4592 | 2.43 | 26.5991 | 36.3192 | 24.1691 | 38.7493 | 7.724% | 15.45% | 23.17% | 29.8862 | 33.0322
Y, ppm | 18.3027 | 1.2982 | 15.7063 | 20.8992 | 14.408 | 22.1974 | 7.093% | 14.19% | 21.28% | 17.3876 | 19.2179
Yb, ppm | 1.7247 | 0.14539 | 1.434 | 2.0155 | 1.2886 | 2.1609 | 8.43% | 16.86% | 25.29% | 1.6385 | 1.811
Zn, ppm | 69.4913 | 2.6498 | 64.1917 | 74.791 | 61.5418 | 77.4409 | 3.813% | 7.626% | 11.44% | 66.0168 | 72.9659
Zr, ppm | 182.9031 | 8.2691 | 166.3648 | 199.4414 | 158.0957 | 207.7105 | 4.521% | 9.042% | 13.56% | 173.7579 | 192.0482
Aqua Regia Digestion
Ag, ppm | 0.26894 | 0.028025 | 0.21289 | 0.32499 | 0.18486 | 0.35301 | 10.42% | 20.84% | 31.26% | 0.25549 | 0.28238
Al, wt.% | 1.126 | 0.069748 | 0.98654 | 1.2655 | 0.91679 | 1.3353 | 6.194% | 12.39% | 18.58% | 1.0697 | 1.1823
As, ppm | 153.1763 | 5.32 | 142.5362 | 163.8163 | 137.2162 | 169.1363 | 3.473% | 6.946% | 10.42% | 145.5174 | 160.8351
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 71.9337 | 6.1627 | 59.6082 | 84.2592 | 53.4455 | 90.422 | 8.567% | 17.13% | 25.7% | 68.337 | 75.5304
Be, ppm | 0.93287 | 0.071814 | 0.78924 | 1.0765 | 0.71742 | 1.1483 | 7.698% | 15.4% | 23.09% | 0.88622 | 0.97951
Bi, ppm | 6.8278 | 0.6699 | 5.488 | 8.1676 | 4.8181 | 8.8375 | 9.811% | 19.62% | 29.43% | 6.4864 | 7.1692
Ca, wt.% | 0.43656 | 0.040345 | 0.35586 | 0.51725 | 0.31552 | 0.55759 | 9.242% | 18.48% | 27.73% | 0.41473 | 0.45838
Cd, ppm | 0.025788 | 0.00812 | 0.009548 | 0.042028 | 0.0014279 | 0.050148 | 31.49% | 62.98% | 94.46% | 0.024499 | 0.027077
Ce, ppm | 52.1856 | 5.0735 | 42.0386 | 62.3325 | 36.9652 | 67.406 | 9.722% | 19.44% | 29.17% | 49.5763 | 54.7948
Co, ppm | 11.5874 | 1.0076 | 9.5722 | 13.6025 | 8.5647 | 14.61 | 8.695% | 17.39% | 26.09% | 11.008 | 12.1667
Cr, ppm | 46.2617 | 5.8489 | 34.564 | 57.9594 | 28.7151 | 63.8083 | 12.64% | 25.29% | 37.93% | 43.9486 | 48.5748
Cs, ppm | 1.4341 | 0.15721 | 1.1197 | 1.7485 | 0.96246 | 1.9057 | 10.96% | 21.92% | 32.89% | 1.3624 | 1.5058
Cu, ppm | 22.665 | 1.1293 | 20.4065 | 24.9236 | 19.2772 | 26.0529 | 4.982% | 9.965% | 14.95% | 21.5318 | 23.7983
Eu, ppm | 0.9406 | 0.074451 | 0.79169 | 1.0895 | 0.71724 | 1.1639 | 7.915% | 15.83% | 23.75% | 0.89357 | 0.98763
Fe, wt.% | 3.0318 | 0.27499 | 2.4818 | 3.5818 | 2.2068 | 3.8568 | 9.07% | 18.14% | 27.21% | 2.8802 | 3.1834
Ga, ppm | 3.9416 | 0.37553 | 3.1906 | 4.6927 | 2.815 | 5.0682 | 9.527% | 19.05% | 28.58% | 3.7445 | 4.1387
Ho, ppm | 0.38057 | 0.038925 | 0.30272 | 0.45842 | 0.2638 | 0.49735 | 10.23% | 20.46% | 30.68% | 0.36154 | 0.3996
In, ppm | 0.025905 | 0.002783 | 0.020339 | 0.031471 | 0.017556 | 0.034254 | 10.74% | 21.49% | 32.23% | 0.02461 | 0.0272
K, wt.% | 0.18638 | 0.019425 | 0.14753 | 0.22523 | 0.1281 | 0.24465 | 10.42% | 20.84% | 31.27% | 0.17706 | 0.1957
La, ppm | 26.6631 | 2.9367 | 20.7897 | 32.5365 | 17.853 | 35.4732 | 11.01% | 22.03% | 33.04% | 25.3299 | 27.9962
Li, ppm | 6.7796 | 0.43717 | 5.9053 | 7.6539 | 5.4681 | 8.0911 | 6.448% | 12.9% | 19.34% | 6.4406 | 7.1186
Mg, wt.% | 0.53883 | 0.050792 | 0.43725 | 0.64042 | 0.38646 | 0.69121 | 9.426% | 18.85% | 28.28% | 0.51189 | 0.56578
Mn, wt.% | 0.02267 | 0.0025869 | 0.017496 | 0.027843 | 0.014909 | 0.03043 | 11.41% | 22.82% | 34.23% | 0.021536 | 0.023803
Mo, ppm | 1.9552 | 0.13353 | 1.6882 | 2.2223 | 1.5547 | 2.3558 | 6.829% | 13.66% | 20.49% | 1.8575 | 2.053
Na, wt.% | 0.19166 | 0.0099799 | 0.1717 | 0.21162 | 0.16172 | 0.2216 | 5.207% | 10.41% | 15.62% | 0.18207 | 0.20124
Nb, ppm | 0.61826 | 0.16036 | 0.29753 | 0.93898 | 0.13717 | 1.0993 | 25.94% | 51.88% | 77.81% | 0.58734 | 0.64917
Ni, ppm | 56.639 | 2.9925 | 50.654 | 62.6239 | 47.6616 | 65.6163 | 5.283% | 10.57% | 15.85% | 53.807 | 59.4709
P, wt.% | 0.064912 | 0.0031067 | 0.058698 | 0.071125 | 0.055592 | 0.074232 | 4.786% | 9.572% | 14.36% | 0.061666 | 0.068157
Pb, ppm | 8.009 | 1.2674 | 5.4741 | 10.5438 | 4.2067 | 11.8112 | 15.83% | 31.65% | 47.48% | 7.6085 | 8.4094
Pr, ppm | 7.4 | 1.7516 | 3.8968 | 10.9032 | 2.1452 | 12.6548 | 23.67% | 47.34% | 71.01% | 7.03 | 7.77
Rb, ppm | 13.0655 | 0.7746 | 11.5163 | 14.6146 | 10.7417 | 15.3892 | 5.929% | 11.86% | 17.79% | 12.4122 | 13.7187
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.0097259 | 0.0037797 | 0.0021665 | 0.017285 | 0 | 0.021065 | 38.86% | 77.72% | 116.6% | 0.0092396 | 0.010212
Sb, ppm | 1.6559 | 0.21806 | 1.2197 | 2.092 | 1.0017 | 2.31 | 13.17% | 26.34% | 39.51% | 1.5731 | 1.7387
Sc, ppm | 3.011 | 0.32464 | 2.3617 | 3.6603 | 2.0371 | 3.9849 | 10.78% | 21.56% | 32.34% | 2.8605 | 3.1616
Sn, ppm | 1.4556 | 0.10979 | 1.236 | 1.6752 | 1.1262 | 1.785 | 7.542% | 15.08% | 22.63% | 1.3828 | 1.5284
Sr, ppm | 39.2312 | 4.0374 | 31.1564 | 47.3059 | 27.1191 | 51.3433 | 10.29% | 20.58% | 30.87% | 37.2696 | 41.1927
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.4887 | 0.091287 | 0.30613 | 0.67128 | 0.21484 | 0.76256 | 18.68% | 37.36% | 56.04% | 0.46427 | 0.51314
Te, ppm | 0.082578 | 0.01586 | 0.050857 | 0.1143 | 0.034996 | 0.13016 | 19.21% | 38.41% | 57.62% | 0.078449 | 0.086707
Th, ppm | 8.6724 | 0.89572 | 6.881 | 10.4638 | 5.9852 | 11.3596 | 10.33% | 20.66% | 30.99% | 8.2388 | 9.106
Ti, wt.% | 0.1116 | 0.026738 | 0.058129 | 0.16508 | 0.031391 | 0.19182 | 23.96% | 47.92% | 71.87% | 0.10602 | 0.11718
Tl, ppm | 0.088317 | 0.013854 | 0.060609 | 0.11603 | 0.046754 | 0.12988 | 15.69% | 31.37% | 47.06% | 0.083902 | 0.092733
U, ppm | 1.0538 | 0.072813 | 0.90817 | 1.1994 | 0.83535 | 1.2722 | 6.91% | 13.82% | 20.73% | 1.0011 | 1.1065
V, ppm | 29.3866 | 2.5781 | 24.2304 | 34.5428 | 21.6523 | 37.1209 | 8.773% | 17.55% | 26.32% | 27.9173 | 30.8559
W, ppm | 9.7831 | 1.6753 | 6.4326 | 13.1336 | 4.7574 | 14.8089 | 17.12% | 34.25% | 51.37% | 9.2939 | 10.2723
Y, ppm | 9.1879 | 0.52042 | 8.147 | 10.2287 | 7.6266 | 10.7491 | 5.664% | 11.33% | 16.99% | 8.7285 | 9.6473
Zn, ppm | 42.1732 | 5.642 | 30.8892 | 53.4572 | 25.2472 | 59.0992 | 13.38% | 26.76% | 40.13% | 40.0645 | 44.2818
Zr, ppm | 29.3942 | 6.5147 | 16.3649 | 42.4235 | 9.8503 | 48.9382 | 22.16% | 44.33% | 66.49% | 27.9245 | 30.864
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 251c
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | 350g 300ml raw sample (solid) packed into single use jar with PhotonAssay (X-ray) finish
Text Values:
NA | Unable to report due to QC failure (Lab 1.15)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 251c
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Canning Vale, WA, Australia
ALS, Johannesburg, South Africa
ALS, Kalgoorlie, WA, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST (SÖGÜT), Sögüt, Bilecik Province, Turkey
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Britannia Mining Solutions, Hamilton, Ontario, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Mineral Solutions (BVMS), Al Wadi District, Jeddah, Saudi Arabia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Minerals Limited, Manso Nkwanta, Ashanti Region, Ghana
Intertek Minerals Ltd, Bibiani, Western North Region, Ghana
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Laboratoire LABOMINE SARL, Agadir, Souss-Massa, Morocco
Labwest Minerals Analysis, Perth, WA, Australia
MSA ENVAL Laboratories, Yamoussoukro, Côte d'Ivoire
MSALABS, Bougouni, Bamako, Mali
MSALABS, Prince George, BC, Canada
MSALABS, Val-d'Or, Quebec, Canada
MSALABS Bulyanhulu Gold Mine, Bubada, Shinyanga, United Republic of Tanzania
MSALABS Geita, Geita, Geita, United Republic of Tanzania
MSALABS Ghana Ltd, Obuasi, Ashanti, Ghana
MSALABS Kibali Gold Mines, Doko, Haut-Uélé, Congo, Democratic Republic of the (Zaire)
MSALABS Timmins, Timmins, Ontario, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Ravenswood Gold, Ravenswood, QLD, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 251c (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 5 | 1 | 267799 | 0.086713 | 0.54443 | 0.51668
2 | 1 | ANSLu | 1 | 15 | 3 | 267800 | 0.085684 | 0.53743 | 0.51631
2 | 1 | ANSLu | 1 | 13 | 6 | 267801 | 0.086916 | 0.53928 | 0.51641
2 | 1 | ANSLu | 1 | 18 | 10 | 267802 | 0.083874 | 0.53615 | 0.51624
2 | 1 | ANSLu | 1 | 2 | 12 | 267803 | 0.084572 | 0.50065 | 0.51435
2 | 1 | ANSLu | 1 | 6 | 7 | 267804 | 0.083108 | 0.53497 | 0.51618
2 | 1 | ANSLu | 1 | 16 | 3 | 267805 | 0.085867 | 0.49198 | 0.51389
2 | 1 | ANSLu | 1 | 9 | 4 | 267806 | 0.086463 | 0.49772 | 0.5142
2 | 1 | ANSLu | 1 | 12 | 10 | 267807 | 0.082439 | 0.51363 | 0.51504
2 | 1 | ANSLu | 1 | 3 | 9 | 267808 | 0.083531 | 0.51614 | 0.51518
2 | 1 | ANSLu | 1 | 8 | 11 | 267809 | 0.08457 | 0.53797 | 0.51634
2 | 1 | ANSLu | 1 | 7 | 2 | 267810 | 0.083144 | 0.50076 | 0.51436
2 | 1 | ANSLu | 1 | 19 | 9 | 267811 | 0.086078 | 0.48744 | 0.51365
2 | 1 | ANSLu | 1 | 10 | 4 | 267812 | 0.085894 | 0.50824 | 0.51476
2 | 1 | ANSLu | 1 | 4 | 1 | 267813 | 0.087943 | 0.51612 | 0.51518
2 | 1 | ANSLu | 1 | 11 | 12 | 267814 | 0.087254 | 0.51386 | 0.51506
2 | 1 | ANSLu | 1 | 1 | 5 | 267815 | 0.084376 | 0.52928 | 0.51588
2 | 1 | ANSLu | 1 | 17 | 8 | 267816 | 0.087305 | 0.50075 | 0.51436
2 | 1 | ANSLu | 1 | 20 | 7 | 267817 | 0.083934 | 0.5007 | 0.51435
2 | 1 | ANSLu | 1 | 14 | 6 | 267818 | 0.084257 | 0.49497 | 0.51405
Mean | [6 empty] | 0.085196 | 0.51512 | 0.51512
Median | [7 empty] | 0.51375 | 0.51505
SD | [7 empty] | 0.018343 | 0.00097751
RSD | [7 empty] | 3.561% | 0.1898%
RSD Upscale Factor | [7 empty] | 0.05329
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 251c (Certified Value 0.508 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.28 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 0.085g | 50g | 25g | 30g | 30g | 50g | 50g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 40g | 50g | 30g | 40g | 3
| 1 | 1 | 0.52928 | 0.52 | 0.497 | 0.49482 | 0.466 | 0.50784 | 0.50467 | 0.525 | 0.51 | 0.52 | 0.51 | 0.53 | 0.51 | 0.52 | 0.53207 | 0.5 | 0.52 | 0.501 | 0.462 | 0.516 | 0.51 | 0.507 | 0.495 | 0.47 | 1
| 1 | 2 | 0.50065 | 0.52 | 0.516 | 0.49997 | 0.487 | 0.49536 | 0.5105 | 0.508 | 0.5 | 0.52 | 0.51 | 0.53 | 0.5 | 0.54 | 0.52846 | 0.489 | 0.51 | 0.519 | 0.463 | 0.509 | 0.5 | 0.493 | 0.49 | 0.49 | #N/A
| 1 | 3 | 0.51614 | 0.5 | 0.509 | 0.50139 | 0.486 | 0.47232 | 0.50767 | 0.52 | 0.5 | 0.53 | 0.51 | 0.53 | 0.49 | 0.52 | 0.53804 | 0.497 | 0.51 | 0.531 | 0.488 | 0.497 | 0.51 | 0.49 | 0.489 | 0.47 | 16
| 1 | 4 | 0.51612 | 0.51 | 0.51 | 0.49842 | 0.485 | 0.50784 | 0.517 | 0.51 | 0.5 | 0.53 | 0.52 | 0.52 | 0.49 | 0.54 | 0.53658 | 0.5 | 0.51 | 0.502 | 0.479 | 0.508 | 0.51 | 0.483 | 0.493 | 0.48 | 0.50774
| 1 | 5 | 0.54443 | 0.5 | 0.51 | 0.49552 | 0.466 | 0.51168 | 0.51 | 0.512 | 0.51 | 0.52 | 0.51 | 0.53 | 0.5 | 0.53 | 0.52315 | 0.506 | 0.51 | 0.499 | 0.474 | 0.512 | 0.48 | 0.487 | 0.489 | 0.49 | 7
| 1 | 6 | 0.53497 | 0.49 | 0.506 | 0.49486 | 0.451 | 0.50304 | 0.50567 | 0.511 | 0.5 | 0.52 | 0.51 | 0.54 | 0.5 | 0.52 | 0.52095 | 0.505 | 0.51 | 0.504 | 0.489 | 0.505 | 0.5 | 0.503 | 0.492 | 0.49
|  | 7 | 0.50076
|  | 8 | 0.53797
|  | 9 | 0.49772
|  | 10 | 0.50824
|  | 11 | 0.51386
|  | 12 | 0.51363
|  | 13 | 0.53928
|  | 14 | 0.49497
|  | 15 | 0.53743
|  | 16 | 0.49198
|  | 17 | 0.50075
|  | 18 | 0.53615
|  | 19 | 0.48744
|  | 20 | 0.5007
| Mean |  | 0.51512 | 0.50667 | 0.508 | 0.4975 | 0.4735 | 0.49968 | 0.50925 | 0.51433 | 0.50333 | 0.52333 | 0.51167 | 0.53 | 0.49833 | 0.52833 | 0.52987 | 0.4995 | 0.51167 | 0.50933 | 0.47583 | 0.50783 | 0.50167 | 0.49383 | 0.49133 | 0.48167
| Median |  | 0.51375 | 0.505 | 0.5095 | 0.49697 | 0.4755 | 0.50544 | 0.50883 | 0.5115 | 0.5 | 0.52 | 0.51 | 0.53 | 0.5 | 0.525 | 0.53026 | 0.5 | 0.51 | 0.503 | 0.4765 | 0.5085 | 0.505 | 0.4915 | 0.491 | 0.485
| Std Dev. |  | 0.018343 | 0.012111 | 0.0062929 | 0.0028335 | 0.014761 | 0.014531 | 0.0044394 | 0.0066533 | 0.005164 | 0.005164 | 0.0040825 | 0.0063246 | 0.0075277 | 0.0098319 | 0.0069746 | 0.0061563 | 0.0040825 | 0.012817 | 0.011754 | 0.0064936 | 0.01169 | 0.009347 | 0.0024221 | 0.0098319
| Rel.Std.Dev. |  | 3.561% | 2.39% | 1.239% | 0.5695% | 3.118% | 2.908% | 0.8718% | 1.294% | 1.026% | 0.9867% | 0.7979% | 1.193% | 1.511% | 1.861% | 1.316% | 1.232% | 0.7979% | 2.516% | 2.47% | 1.279% | 2.33% | 1.893% | 0.493% | 2.041%
| PDM3 |  | 1.453% | -0.212% | 0.05058% | -2.018% | -6.744% | -1.588% | 0.2968% | 1.298% | -0.8685% | 3.07% | 0.7727% | 4.383% | -1.853% | 4.055% | 4.359% | -1.623% | 0.7727% | 0.3132% | -6.285% | 0.01776% | -1.197% | -2.74% | -3.232% | -5.136%
| Z-Score (Absolute) |  | NA | 0 | 0.12 | 0.81 | 2.92 | 0.61 | 0.23 | 0.67 | 0.29 | 1.47 | 0.44 | 2.05 | 0.73 | 1.91 | 2.04 | 0.63 | 0.44 | 0.23 | 2.71 | 0.1 | 0.44 | 1.13 | 1.35 | 2.2
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 251c (Certified Value 0.501 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 6.01 | 6.02 | 6.03 | 6.04 | 6.05 | 6.06 | 6.07 | 6.08 | 6.09 | 6.10 | 6.11 | 6.12 | 6.13 | 6.14 | 6.15 | 6.16 | 6.17 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 3
|  |  | 0.085g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 0.52928 | 0.47 | 0.53 | 0.52 | 0.48 | 0.49 | 0.553 | 0.458 | 0.494 | 0.465 | 0.48 | 0.426 | 0.487 | 0.485 | 0.561 | 0.513 | 0.49 | 0.5 | 1
| 1 | 2 | 0.50065 | 0.52 | 0.48 | 0.538 | 0.42 | 0.51 | 0.484 | 0.571 | 0.499 | 0.463 | 0.5 | 0.444 | 0.512 | 0.506 | 0.559 | 0.513 | 0.51 | 0.469 | #N/A
| 1 | 3 | 0.51614 | [16 empty] | 0.545 | 16
| 1 | 4 | 0.51612 | [16 empty] | 0.499 | 0.50072
| 1 | 5 | 0.54443 | 0.51 | 0.47 | 0.512 | 0.44 | 0.48 | 0.484 | 0.494 | 0.493 | 0.461 | 0.57 | 0.464 | 0.496 | 0.511 | 0.503 | 0.513 | 0.52 | 0.533 | 9
| 1 | 6 | 0.53497 | 0.52 | 0.48 | 0.536 | 0.5 | 0.42 | 0.572 | 0.515 | 0.506 | 0.463 | 0.47 | 0.526 | 0.471 | 0.482 | 0.467 | 0.513 | 0.53 | 0.508
| 1 | 7 | 0.50076 | [16 empty] | 0.562
| 1 | 8 | 0.53797 | [16 empty] | 0.566
| 1 | 9 | 0.49772 | 0.44 | 0.54 | 0.505 | 0.49 | 0.49 | 0.464 | 0.511 | 0.562 | 0.538 | 0.51 | 0.547 | 0.48 | 0.482 | 0.491 | 0.467 | 0.47 | 0.478
| 1 | 10 | 0.50824 | 0.59 | 0.51 | 0.462 | 0.41 | 0.52 | 0.487 | 0.499 | 0.482 | 0.515 | 0.46 | 0.533 | 0.501 | 0.493 | 0.487 | 0.467 | 0.48 | 0.512
| 1 | 11 | 0.51386 | [16 empty] | 0.487
| 1 | 12 | 0.51363 | [16 empty] | 0.548
|  | 13 | 0.53928
|  | 14 | 0.49497
|  | 15 | 0.53743
|  | 16 | 0.49198
|  | 17 | 0.50075
|  | 18 | 0.53615
|  | 19 | 0.48744
|  | 20 | 0.5007
| Mean |  | 0.51512 | 0.50833 | 0.50167 | 0.51217 | 0.45667 | 0.485 | 0.50733 | 0.508 | 0.506 | 0.48417 | 0.49833 | 0.49 | 0.49117 | 0.49317 | 0.51133 | 0.49767 | 0.5 | 0.51725
| Median |  | 0.51375 | 0.515 | 0.495 | 0.516 | 0.46 | 0.49 | 0.4855 | 0.505 | 0.4965 | 0.464 | 0.49 | 0.495 | 0.4915 | 0.489 | 0.497 | 0.513 | 0.5 | 0.51
| Std Dev. |  | 0.018343 | 0.051153 | 0.029269 | 0.027802 | 0.038297 | 0.035071 | 0.043926 | 0.036889 | 0.028545 | 0.033612 | 0.039707 | 0.05154 | 0.014851 | 0.01264 | 0.039444 | 0.023754 | 0.023664 | 0.03285
| Rel.Std.Dev. |  | 3.561% | 10.06% | 5.834% | 5.428% | 8.386% | 7.231% | 8.658% | 7.262% | 5.641% | 6.942% | 7.968% | 10.52% | 3.024% | 2.563% | 7.714% | 4.773% | 4.733% | 6.351%
| PDM3 |  | 2.876% | 1.52% | 0.1883% | 2.285% | -8.799% | -3.14% | 1.32% | 1.453% | 1.054% | -3.307% | -0.4774% | -2.142% | -1.909% | -1.509% | 2.119% | -0.6106% | -0.1446% | 3.3%
| Z-Score (Absolute) |  | NA | 0.67 | 0.13 | 0.98 | 3.51 | 1.21 | 0.59 | 0.65 | 0.49 | 1.28 | 0.13 | 0.81 | 0.71 | 0.55 | 0.92 | 0.19 | 0 | 1.4
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 251c (Certified Value 0.485 ppm) | [24 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.28 | 1.29 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*OES | AR*AAS | AR*AAS | AR*MS | AR*AAS | 3
|  |  | 0.085g | 50g | 25g | 30g | 50g | 30g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 10g | 30g | 15g | 10g | 15g | 25g | 30g | 10g | 40g | 25g | 3
| 1 | 1 | 0.52928 | 0.49 | 0.46 | 0.5132 | 0.48 | 0.466 | 0.49 | 0.489 | 0.51 | 0.492 | 0.488 | 0.481 | < 0.5 | 0.473 | 0.49 | 0.4507 | 0.468 | 0.5099 | 0.501 | 0.51 | 0.44 | 0.49 | 0.48 | 1
| 1 | 2 | 0.50065 | 0.5 | 0.47 | 0.49312 | 0.48 | 0.409 | 0.49 | 0.511 | 0.54 | 0.502 | 0.494 | 0.47 | < 0.5 | 0.462 | 0.49 | 0.4683 | 0.461 | 0.4951 | 0.513 | 0.48 | 0.46 | 0.495 | 0.48 | #N/A
| 1 | 3 | 0.51614 | 0.5 | 0.47 | 0.49535 | 0.48 | 0.487 | 0.49 | 0.535 | 0.51 | 0.509 | 0.488 | 0.489 | < 0.5 | 0.475 | 0.49 | 0.4848 | 0.44 | 0.5316 | 0.558 | 0.47 | 0.47 | 0.47 | 0.48 | 16
| 1 | 4 | 0.51612 | 0.5 | 0.46 | 0.52443 | 0.48 | 0.489 | 0.49 | 0.497 | 0.48 | 0.501 | 0.484 | 0.466 | < 0.5 | 0.474 | 0.49 | 0.4464 | 0.454 | 0.5266 | 0.53 | 0.48 | 0.45 | 0.495 | 0.46 | 0.4854
| 1 | 5 | 0.54443 | 0.5 | 0.45 | 0.51519 | 0.52 | 0.437 | 0.49 | 0.521 | 0.51 | 0.509 | 0.479 | 0.468 | < 0.5 | 0.41 | 0.49 | 0.4525 | 0.473 | 0.5073 | 0.514 | 0.5 | 0.47 | 0.471 | 0.47 | 11
| 1 | 6 | 0.53497 | 0.5 | 0.46 | 0.49454 | 0.5 | 0.449 | 0.49 | 0.505 | 0.5 | 0.504 | 0.496 | 0.491 | < 0.5 | 0.461 | 0.49 | 0.4514 | 0.449 | 0.5011 | 0.505 | 0.43 | 0.46 | 0.487 | 0.67
|  | 7 | 0.50076
|  | 8 | 0.53797
|  | 9 | 0.49772
|  | 10 | 0.50824
|  | 11 | 0.51386
|  | 12 | 0.51363
|  | 13 | 0.53928
|  | 14 | 0.49497
|  | 15 | 0.53743
|  | 16 | 0.49198
|  | 17 | 0.50075
|  | 18 | 0.53615
|  | 19 | 0.48744
|  | 20 | 0.5007
| Mean |  | 0.51512 | 0.49833 | 0.46167 | 0.50597 | 0.49 | 0.45617 | 0.49 | 0.50967 | 0.50833 | 0.50283 | 0.48817 | 0.4775 |  | 0.45917 | 0.49 | 0.45902 | 0.4575 | 0.51193 | 0.52017 | 0.47833 | 0.45833 | 0.48467 | 0.50667
| Median |  | 0.51375 | 0.5 | 0.46 | 0.50428 | 0.48 | 0.4575 | 0.49 | 0.508 | 0.51 | 0.503 | 0.488 | 0.4755 |  | 0.4675 | 0.49 | 0.45195 | 0.4575 | 0.5086 | 0.5135 | 0.48 | 0.46 | 0.4885 | 0.48
| Std Dev. |  | 0.018343 | 0.0040825 | 0.0075277 | 0.013316 | 0.016733 | 0.030883 | 6.0809e-17 | 0.016621 | 0.019408 | 0.0063061 | 0.0062743 | 0.011005 |  | 0.024863 | 6.0809e-17 | 0.014695 | 0.012276 | 0.01434 | 0.021047 | 0.027869 | 0.01169 | 0.011396 | 0.080416
| Rel.Std.Dev. |  | 3.561% | 0.8192% | 1.631% | 2.632% | 3.415% | 6.77% | 1.241e-14% | 3.261% | 3.818% | 1.254% | 1.285% | 2.305% |  | 5.415% | 1.241e-14% | 3.202% | 2.683% | 2.801% | 4.046% | 5.826% | 2.551% | 2.351% | 15.87%
| PDM3 |  | 6.123% | 2.664% | -4.89% | 4.238% | 0.9471% | -6.023% | 0.9471% | 4.999% | 4.724% | 3.591% | 0.5694% | -1.628% |  | -5.405% | 0.9471% | -5.436% | -5.748% | 5.466% | 7.162% | -1.456% | -5.577% | -0.1516% | 4.381%
| Z-Score (Absolute) |  | NA | 0.34 | 1 | 0.62 | 0.03 | 1.21 | 0.03 | 0.75 | 0.7 | 0.5 | 0.03 | 0.42 | 8.75 | 1.1 | 0.03 | 1.1 | 1.16 | 0.84 | 1.14 | 0.39 | 1.13 | 0.16 | 0.64
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 251c (Certified Value 0.492 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.22 | 1.23 | 1.27 | 1.29 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | 3
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 30g | 30g | 30g | 30g | 20g | 15g | 05g | 30g | 200g | 50g | 200g | 50g | 3
| 1 | 1 | 0.52928 | 0.5 | 0.49 | 0.4624 | 0.563 | 0.48138 | 0.45 | 0.522 | 0.49 | 0.5 | 0.47 | 0.46 | 0.52 | 0.4886 | 0.48 | 0.5 | 0.49 | 0.522 | 0.465 | 0.48 | 0.48 | 1
| 1 | 2 | 0.50065 | 0.49 | 0.5 | 0.46747 | 0.56 | 0.54545 | 0.5 | 0.522 | 0.49 | 0.5 | 0.45 | 0.54 | 0.53 | 0.49089 | 0.49 | 0.5 | 0.5 | 0.498 | 0.463 | 0.47 | 0.48 | #N/A
| 1 | 3 | 0.51614 | 0.5 | 0.5 | 0.45924 | 0.564 | 0.49338 | 0.55 | 0.517 | 0.48 | 0.51 | 0.45 | 0.22 | 0.53 | 0.49743 | 0.42 | 0.5 | 0.48 | 0.515 | 0.468 | 0.5 | 0.48 | 16
| 1 | 4 | 0.51612 | 0.5 | 0.49 | 0.46451 | 0.574 | 0.49817 | 0.49 | 0.565 | 0.49 | 0.5 | 0.45 | 0.36 | 0.52 | 0.48177 | 0.5 | 0.5 | 0.46 | 0.498 | 0.486 | 0.49 | 0.49 | 0.49227
| 1 | 5 | 0.54443 | 0.5 | 0.5 | 0.45666 | 0.556 | 0.48588 | 0.49 | 0.516 | 0.49 | 0.51 | 0.47 | 0.46 | 0.53 | 0.49003 | 0.41 | 0.5 | 0.45 | 0.491 | 0.466 | 0.48 | 0.5 | 13
| 1 | 6 | 0.53497 | 0.5 | 0.5 | 0.46494 | 0.571 | 0.48323 | 0.51 | 0.523 | 0.49 | 0.51 | 0.49 | 0.4 | 0.53 | 0.47856 | 0.49 | 0.5 | 0.48 | 0.506 | 0.477 | 0.51 | 0.48
|  | 7 | 0.50076
|  | 8 | 0.53797
|  | 9 | 0.49772
|  | 10 | 0.50824
|  | 11 | 0.51386
|  | 12 | 0.51363
|  | 13 | 0.53928
|  | 14 | 0.49497
|  | 15 | 0.53743
|  | 16 | 0.49198
|  | 17 | 0.50075
|  | 18 | 0.53615
|  | 19 | 0.48744
|  | 20 | 0.5007
| Mean |  | 0.51512 | 0.49833 | 0.49667 | 0.46254 | 0.56467 | 0.49792 | 0.49833 | 0.5275 | 0.48833 | 0.505 | 0.46333 | 0.40667 | 0.52667 | 0.48788 | 0.465 | 0.5 | 0.47667 | 0.505 | 0.47083 | 0.48833 | 0.485
| Median |  | 0.51375 | 0.5 | 0.5 | 0.46346 | 0.5635 | 0.48963 | 0.495 | 0.522 | 0.49 | 0.505 | 0.46 | 0.43 | 0.53 | 0.48931 | 0.485 | 0.5 | 0.48 | 0.502 | 0.467 | 0.485 | 0.48
| Std Dev. |  | 0.018343 | 0.0040825 | 0.005164 | 0.0039864 | 0.0067429 | 0.024142 | 0.032506 | 0.018598 | 0.0040825 | 0.0054772 | 0.01633 | 0.11003 | 0.005164 | 0.0067756 | 0.03937 | 0 | 0.018619 | 0.011696 | 0.0088863 | 0.01472 | 0.0083666
| Rel.Std.Dev. |  | 3.561% | 0.8192% | 1.04% | 0.8618% | 1.194% | 4.849% | 6.523% | 3.526% | 0.836% | 1.085% | 3.524% | 27.06% | 0.9805% | 1.389% | 8.467% | 0% | 3.906% | 2.316% | 1.887% | 3.014% | 1.725%
| PDM3 |  | 4.643% | 1.232% | 0.8934% | -6.04% | 14.71% | 1.147% | 1.232% | 7.157% | -0.7994% | 2.586% | -5.878% | -17.39% | 6.988% | -0.8915% | -5.539% | 1.571% | -3.169% | 2.586% | -4.354% | -0.7994% | -1.477%
| Z-Score (Absolute) |  | NA | 0.31 | 0.22 | 1.62 | 3.89 | 0.29 | 0.31 | 1.89 | 0.22 | 0.67 | 1.57 | 4.63 | 1.84 | 0.25 | 1.48 | 0.4 | 0.85 | 0.67 | 1.17 | 0.22 | 0.4
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 251c (Certified Value 0.296 ppm) | [21 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.31 | 0.31604 | 0.5 | 0.5 | 0.28 | 0.31 | 0.28 | 0.31 | 0.28 | 0.36956 | < 0.3 | 0.2 | 0.389 | 0.32 | 0.278 | 0.34 | < 0.5 | 0.2 | 5.326 |  | 1
| 1 | 2 | 0.31 | 0.29002 | 0.4 | 0.5 | 0.28 | 0.31 | 0.29 | 0.26 | 0.25 | 0.28614 | < 0.3 | 0.2 | 0.345 | 0.3 | 0.283 | 0.35 | < 0.5 | 0.2 | 0.722 |  | 23
| 1 | 3 | 0.3 | 0.29327 | < 0.3 | 0.5 | 0.27 | 0.35 | 0.27 | 0.29 | 0.24 | 0.35671 | < 0.3 | 0.2 | 0.405 | 0.31 | 0.272 | 0.32 | < 0.5 | 0.2 | 0.804 |  | 16
| 1 | 4 | 0.31 | 0.28113 | 0.5 | 0.5 | 0.3 | 0.32 | 0.28 | 0.24 | 0.27 | 0.29853 | < 0.3 | 0.2 | 0.405 | 0.31 | 0.303 | 0.3 | < 0.5 | 0.2 | 0.804 |  | 0.29636
| 1 | 5 | 0.32 | 0.31209 | 0.3 | < 0.5 | 0.29 | 0.32 | 0.3 | 0.27 | 0.22 | 0.28635 | < 0.3 | 0.2 | 0.384 | 0.29 | 0.303 | 0.32 | < 0.5 | 0.2 | 0.93 |  | 15
| 1 | 6 | 0.32 | 0.29893 | 0.4 | < 0.5 | 0.26 | 0.33 | 0.3 | 0.27 | 0.24 | 0.29899 | < 0.3 | 0.2 | 0.378 | 0.31 | 0.283 | 0.33 | < 0.5 | 0.2 | 0.881
| Mean |  | 0.31167 | 0.29858 | 0.42 | 0.5 | 0.28 | 0.32333 | 0.28667 | 0.27333 | 0.25 | 0.31605 |  | 0.2 | 0.38433 | 0.30667 | 0.287 | 0.32667 |  | 0.2 | 1.5778
| Median |  | 0.31 | 0.2961 | 0.4 | 0.5 | 0.28 | 0.32 | 0.285 | 0.27 | 0.245 | 0.29876 |  | 0.2 | 0.3865 | 0.31 | 0.283 | 0.325 |  | 0.2 | 0.8425
| Std Dev. |  | 0.0075277 | 0.013367 | 0.083666 | 0 | 0.014142 | 0.015055 | 0.012111 | 0.024221 | 0.021909 | 0.037126 |  | 3.0405e-17 | 0.022196 | 0.010328 | 0.013038 | 0.017512 |  | 3.0405e-17 | 1.8376
| Rel.Std.Dev. |  | 2.415% | 4.477% | 19.92% | 0% | 5.051% | 4.656% | 4.225% | 8.861% | 8.764% | 11.75% |  | 1.52e-14% | 5.775% | 3.368% | 4.543% | 5.361% |  | 1.52e-14% | 116.5%
| PDM3 |  | 5.165% | 0.7494% | 41.72% | 68.71% | -5.52% | 9.102% | -3.271% | -7.77% | -15.64% | 6.643% |  | -32.51% | 29.68% | 3.478% | -3.158% | 10.23% |  | -32.51% | 432.4%
| Z-Score (Absolute) |  | 0.36 | 0 | NA | NA | 0.51 | 0.67 | 0.32 | 0.69 | 1.32 | 0.48 | 4.05 | NA | 2.34 | 0.22 | 0.32 | 0.77 | 1.32 | NA | 34.85
Results from laboratories 1.07, 1.08, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 251c (Certified Value 6.77 wt.%) | [21 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 6.7618 | 6.6192 | 7.0693 | 6.94 | 6.85 | 6.64 | 6.4 | 6.7 | 6.71 | 6.78 | 6.918 | 6.2325 | 6.88 | 7.05 | 7.0631 | 6.46 | 6.58 | 6.67 | 6.76 | 6.6671 | 1
| 1 | 2 | 6.8039 | 6.6021 | 7.1039 | 6.86 | 6.91 | 6.62 | 6.51 | 6.74 | 6.73 | 6.86 | 6.8227 | 6.4219 | 6.87 | 7.01 | 6.9323 | 6.9 | 6.13 | 6.68 | 6.82 | 6.9243 | #N/A
| 1 | 3 | 6.789 | 6.7062 | 7.0799 | 6.86 | 6.8 | 6.57 | 6.6 | 6.68 | 6.79 | 6.78 | 6.7902 | 6.4015 | 6.67 | 7 | 6.9532 | 6.66 | 6.69 | 6.7 | 6.74 | 6.996 | 16
| 1 | 4 | 6.8135 | 6.3512 | 7.0922 | 6.85 | 6.89 | 6.46 | 6.56 | 6.53 | 6.71 | 6.78 | 6.7151 | 6.4446 | 6.62 | 6.88 | 6.919 | 6.72 | 6.46 | 6.64 | 6.85 | 6.9419 | 6.7694
| 1 | 5 | 6.872 | 6.5159 | 7.0786 | 6.84 | 6.73 | 6.44 | 5.91 | 6.63 | 6.72 | 6.72 | 7.011 | 6.1987 | 6.74 | 6.97 | 6.8231 | 6.7 | 6.36 | 6.64 | 6.78 | 6.886 | 16
| 1 | 6 | 6.7854 | 6.6023 | 7.0894 | 6.91 | 6.88 | 6.52 | 6.5 | 6.59 | 6.87 | 6.85 | 7.0095 | 6.1342 | 6.92 | 6.94 | 6.7853 | 6.71 | 6.45 | 6.7 | 6.81 | 6.9304
| Mean |  | 6.8043 | 6.5662 | 7.0855 | 6.8767 | 6.8433 | 6.5417 | 6.4133 | 6.645 | 6.755 | 6.795 | 6.8777 | 6.3056 | 6.7833 | 6.975 | 6.9127 | 6.6917 | 6.445 | 6.6717 | 6.7933 | 6.891
| Median |  | 6.7965 | 6.6022 | 7.0846 | 6.86 | 6.865 | 6.545 | 6.505 | 6.655 | 6.725 | 6.78 | 6.8704 | 6.317 | 6.805 | 6.985 | 6.9257 | 6.705 | 6.455 | 6.675 | 6.795 | 6.9274
| Std Dev. |  | 0.037592 | 0.12146 | 0.012162 | 0.039328 | 0.067429 | 0.082563 | 0.25563 | 0.077136 | 0.063797 | 0.052058 | 0.12158 | 0.13281 | 0.12404 | 0.059582 | 0.098938 | 0.14091 | 0.19233 | 0.027142 | 0.040825 | 0.11527
| Rel.Std.Dev. |  | 0.5525% | 1.85% | 0.1716% | 0.5719% | 0.9853% | 1.262% | 3.986% | 1.161% | 0.9444% | 0.7661% | 1.768% | 2.106% | 1.829% | 0.8542% | 1.431% | 2.106% | 2.984% | 0.4068% | 0.601% | 1.673%
| PDM3 |  | 0.5149% | -3.003% | 4.67% | 1.584% | 1.092% | -3.364% | -5.26% | -1.838% | -0.2128% | 0.3781% | 1.6% | -6.852% | 0.2057% | 3.037% | 2.116% | -1.148% | -4.792% | -1.444% | 0.3534% | 1.796%
| Z-Score (Absolute) |  | 0.1 | 1.37 | 1.83 | 0.54 | 0.34 | 1.52 | 2.31 | 0.88 | 0.2 | 0.04 | 0.55 | 2.97 | 0.03 | 1.15 | 0.76 | 0.59 | 2.11 | 0.72 | 0.03 | 0.63
| Analytical results for As in OREAS 251c (Certified Value 175 ppm) | [21 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 0
0
| 1 | 1 | 183.7 | 175.5909 | 177 | 177 | 171.5 | 165.5 | 183 | 171 | 173 | 177.368 | 147.5 | 172 | 175.5 | 179.1 | 182 | 168 | 103.2 | 180 | 96.419 |  | 1
| 1 | 2 | 181.2 | 174.9919 | 181 | 172 | 173 | 165 | 186 | 167.5 | 173 | 170.9843 | 151.5 | 176 | 172.9 | 176.5 | 183.9 | 167 | 106.88 | 177 | 236.253 |  | 11
| 1 | 3 | 183.8 | 172.7121 | 177 | 171 | 172 | 171.5 | 180.5 | 171.5 | 172.5 | 169.0291 | 152.5 | 173 | 173.9 | 178.3 | 188.5 | 172 | 105.1 | 179 | 229.702 |  | 16
| 1 | 4 | 180.5 | 167.0752 | 180 | 170 | 170 | 168 | 176 | 168.5 | 171.5 | 175.3522 | 153.1 | 174 | 170.6 | 176.6 | 176.9 | 167 | 105.6 | 182 | 231.91 |  | 174.7228
| 1 | 5 | 179.2 | 169.9788 | 177 | 175 | 168.5 | 168.5 | 179.5 | 167 | 171 | 173.9922 | 147.8 | 171 | 173.1 | 176.1 | 180.5 | 167 | 106.4 | 184 | 229.936 |  | 17
| 1 | 6 | 181.1 | 173.5564 | 181 | 178 | 170 | 166.5 | 176.5 | 175.5 | 173.5 | 173.353 | 146.8 | 180 | 171.1 | 175.8 | 182.6 | 166 | 105.8 | 178 | 228.594
| Mean |  | 181.5833 | 172.3175 | 178.8333 | 173.8333 | 170.8333 | 167.5 | 180.25 | 170.1667 | 172.4167 | 173.3465 | 149.8667 | 174.3333 | 172.85 | 177.0667 | 182.4 | 167.8333 | 105.4967 | 180 | 208.8023
| Median |  | 181.15 | 173.1342 | 178.5 | 173.5 | 170.75 | 167.25 | 180 | 169.75 | 172.75 | 173.6726 | 149.65 | 173.5 | 173 | 176.55 | 182.3 | 167 | 105.7 | 179.5 | 229.819
| Std Dev. |  | 1.8236 | 3.2406 | 2.0412 | 3.3116 | 1.633 | 2.3875 | 3.8308 | 3.1885 | 0.9704 | 2.9944 | 2.8048 | 3.266 | 1.8064 | 1.3216 | 3.8346 | 2.137 | 1.2856 | 2.6077 | 55.1229
| Rel.Std.Dev. |  | 1.004% | 1.881% | 1.141% | 1.905% | 0.9559% | 1.425% | 2.125% | 1.874% | 0.5628% | 1.727% | 1.872% | 1.873% | 1.045% | 0.7464% | 2.102% | 1.273% | 1.219% | 1.449% | 26.4%
| PDM3 |  | 3.927% | -1.377% | 2.353% | -0.509% | -2.226% | -4.134% | 3.163% | -2.608% | -1.32% | -0.7877% | -14.23% | -0.2229% | -1.072% | 1.342% | 4.394% | -3.943% | -39.62% | 3.02% | 19.5%
| Z-Score (Absolute) |  | 1.01 | 0.13 | 0.67 | 0.06 | 0.31 | 0.72 | 0.85 | 0.39 | 0.11 | 0 | 2.89 | 0.12 | 0.06 | 0.46 | 1.11 | 0.68 | 8.34 | 0.82 | 4.36
| Analytical results for B in OREAS 251c (Indicative Value 24.7 ppm) | [21 empty] | Indicative
B | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.08 | [19 empty] | ppm
|  |  | 4A*OES | [19 empty] | 1
1
| 1 | 1 | 26 | [19 empty] | 1
| 1 | 2 | 24 | [19 empty] | 1
| 1 | 3 | 25 | [19 empty] | 16
| 1 | 4 | 24 | [19 empty] | 24.6667
| 1 | 5 | 25 | [19 empty] | 7
| 1 | 6 | 24
| Mean |  | 24.6667
| Median |  | 24.5
| Std Dev. |  | 0.8165
| Rel.Std.Dev. |  | 3.31%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 251c (Certified Value 509 ppm) | [21 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 520.2 | 522.9421 | 486.4082 | 527 | 522 | 500 | 510 | 500 | 500 | 510 | 517.8293 | 464 | 499 | 526 | 556.3 | 505 | 594 | 518.1 | 522 | 149.656 | 1
| 1 | 2 | 527.8 | 517.839 | 486.2459 | 516 | 516 | 500 | 510 | 500 | 500 | 510 | 495.7575 | 482 | 513 | 514 | 545.2 | 513 | 605 | 514.3 | 522 | 501.148 | 25
| 1 | 3 | 517.6 | 515.1814 | 490.6989 | 530 | 514 | 490 | 520 | 500 | 510 | 510 | 496.628 | 471 | 500 | 519 | 554.9 | 500 | 484 | 509 | 526 | 502.968 | 16
| 1 | 4 | 515.5 | 498.5348 | 488.5482 | 514 | 513 | 490 | 520 | 490 | 510 | 510 | 515.6983 | 465 | 507 | 511 | 546.5 | 500 | 594 | 515.8 | 524 | 505.233 | 508.9435
| 1 | 5 | 522.5 | 498.2892 | 490.3779 | 532 | 514 | 490 | 510 | 500 | 500 | 500 | 506.6418 | 459 | 495 | 521 | 540.4 | 497 | 588 | 510.6 | 529 | 501.066 | 18
| 1 | 6 | 517.3 | 513.0614 | 490.7738 | 529 | 515 | 490 | 510 | 490 | 510 | 510 | 509.143 | 445 | 524 | 524 | 543.8 | 494 | 525 | 518.8 | 525 | 499.7
| Mean |  | 520.15 | 510.9747 | 488.8422 | 524.6667 | 515.6667 | 493.3333 | 513.3333 | 496.6667 | 505 | 508.3333 | 506.9497 | 464.3333 | 506.3333 | 519.1667 | 547.85 | 501.5 | 565 | 514.4333 | 524.6667 | 443.2952
| Median |  | 518.9 | 514.1214 | 489.4631 | 528 | 514.5 | 490 | 510 | 500 | 505 | 510 | 507.8924 | 464.5 | 503.5 | 520 | 545.85 | 500 | 591 | 515.05 | 524.5 | 501.107
| Std Dev. |  | 4.4787 | 10.2768 | 2.1113 | 7.6855 | 3.266 | 5.164 | 5.164 | 5.164 | 5.4772 | 4.0825 | 9.291 | 12.3234 | 10.7641 | 5.7764 | 6.3538 | 6.7157 | 48.9326 | 3.9652 | 2.6583 | 143.8659
| Rel.Std.Dev. |  | 0.861% | 2.011% | 0.4319% | 1.465% | 0.6334% | 1.047% | 1.006% | 1.04% | 1.085% | 0.8031% | 1.833% | 2.654% | 2.126% | 1.113% | 1.16% | 1.339% | 8.661% | 0.7708% | 0.5067% | 32.45%
| PDM3 |  | 2.202% | 0.3991% | -3.95% | 3.089% | 1.321% | -3.067% | 0.8625% | -2.412% | -0.7748% | -0.1199% | -0.3918% | -8.765% | -0.5129% | 2.009% | 7.645% | -1.463% | 11.01% | 1.079% | 3.089% | -12.9%
| Z-Score (Absolute) |  | 0.71 | 0.09 | 1.4 | 1.01 | 0.41 | 1.1 | 0.25 | 0.88 | 0.31 | 0.09 | 0.18 | 3.05 | 0.22 | 0.64 | 2.57 | 0.55 | 3.73 | 0.32 | 1.01 | 4.47
| Analytical results for Be in OREAS 251c (Certified Value 2.39 ppm) | [21 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 2
3
| 1 | 1 | 2.45 | 2.2653 | 2.3484 | 2.9 | 2.3 | 2.4 | 2.3 | 2.42 | 2.39 | 2.26 | 2.5881 | 2.02 | 2.4 | 2 | 2.48 | 2 | 2.37 | 11.81 | 2.5 |  | 1
| 1 | 2 | 2.77 | 2.2003 | 2.3319 | 2.5 | 2.5 | 2.4 | 2.32 | 2.45 | 2.39 | 2.29 | 2.4647 | 1.96 | 2.4 | 2 | 2.4 | 3 | 2.5 | 12.42 | 2.5 | 1.982 | 26
| 1 | 3 | 2.44 | 2.3132 | 2.2637 | 2.8 | 2.5 | 2.38 | 2.34 | 2.43 | 2.44 | 2.26 | 2.7689 | 2.3 | 2.4 | 2 | 2.41 | 3 | 2.62 | 12.19 | 3 | 1.992 | 16
| 1 | 4 | 2.39 | 2.165 | 2.219 | 2.8 | 2.1 | 2.34 | 2.33 | 2.39 | 2.42 | 2.25 | 2.5888 | 2.31 | 2.3 | 2 | 2.32 | 3 | 2.25 | 12.34 | 3 | 1.955 | 2.3931
| 1 | 5 | 2.66 | 2.1936 | 2.3119 | 2.6 | 2.3 | 2.35 | 2.32 | 2.42 | 2.42 | 2.24 | 2.5788 | 2.26 | 2.3 | 2 | 2.4 | 3 | 2.3 | 12.33 | 3 | 1.911 | 19
| 1 | 6 | 2.56 | 2.3349 | 2.3632 | 2.8 | 2.3 | 2.36 | 2.31 | 2.43 | 2.45 | 2.29 | 2.7319 | 2.25 | 2.4 | 2 | 2.44 | 3 | 2.3 | 12.51 | 3 | 1.952
| Mean |  | 2.545 | 2.2454 | 2.3064 | 2.7333 | 2.3333 | 2.3717 | 2.32 | 2.4233 | 2.4183 | 2.265 | 2.6202 | 2.1833 | 2.3667 | 2 | 2.4083 | 2.8333 | 2.39 | 12.2667 | 2.8333 | 1.9584
| Median |  | 2.505 | 2.2328 | 2.3219 | 2.8 | 2.3 | 2.37 | 2.32 | 2.425 | 2.42 | 2.26 | 2.5885 | 2.255 | 2.4 | 2 | 2.405 | 3 | 2.335 | 12.335 | 3 | 1.955
| Std Dev. |  | 0.14707 | 0.069544 | 0.055001 | 0.15055 | 0.15055 | 0.025626 | 0.014142 | 0.019664 | 0.024833 | 0.020736 | 0.1118 | 0.15267 | 0.05164 | 0 | 0.053072 | 0.40825 | 0.14227 | 0.2476 | 0.2582 | 0.031564
| Rel.Std.Dev. |  | 5.779% | 3.097% | 2.385% | 5.508% | 6.452% | 1.08% | 0.6096% | 0.8114% | 1.027% | 0.9155% | 4.267% | 6.992% | 2.182% | 0% | 2.204% | 14.41% | 5.953% | 2.018% | 9.113% | 1.612%
| PDM3 |  | 6.346% | -6.173% | -3.626% | 14.22% | -2.499% | -0.8968% | -3.056% | 1.262% | 1.053% | -5.354% | 9.488% | -8.767% | -1.106% | -16.43% | 0.6353% | 18.39% | -0.1307% | 412.6% | 18.39% | -18.17%
| Z-Score (Absolute) |  | 1.16 | 0.96 | 0.53 | 2.5 | 0.34 | 0.06 | 0.43 | 0.3 | 0.27 | 0.82 | 1.7 | 1.4 | 0.1 | NA | 0.19 | NA | 0.06 | 70.05 | 3.21 | 2.99
Results from laboratories 1.17 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 251c (Certified Value 7.52 ppm) | [21 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 7.94 | 7.6104 | 7.6 | 8.46 | 7.54 | 7.9 | 7.79 | 7.09 | 8.56 | 6.9586 | 7.53 | 7.73 | 7.36 | 8.18 | 7.47 | 6.48 | < 5 | 9 |  |  | 1
| 1 | 2 | 7.2 | 7.305 | 7.36 | 8.3 | 7.28 | 7.49 | 8.13 | 6.82 | 9.18 | 6.9225 | 8.07 | 7.7 | 6.98 | 8 | 7.39 | 7.03 | < 5 | 9.1 | 0.934 |  | 27
| 1 | 3 | 7.89 | 7.2952 | 7.32 | 7.87 | 7.6 | 8.12 | 7.18 | 7.3 | 8.19 | 7.1648 | 7.34 | 7.48 | 6.74 | 8.51 | 7.48 | 6.54 | < 5 | 9 | 0.846 |  | 16
| 1 | 4 | 7.11 | 6.8658 | 8.04 | 7.85 | 7.63 | 8.07 | 7.82 | 6.38 | 8.23 | 7.503 | 7.41 | 7.64 | 6.32 | 8.09 | 7.28 | 6.63 | < 5 | 8.7 | 1.854 |  | 7.5227
| 1 | 5 | 7.31 | 7.0793 | 9.59 | 8.73 | 7.25 | 8.14 | 7.84 | 6.46 | 8.21 | 7.4435 | 7.27 | 7.78 | 6.68 | 8.62 | 8.1 | 7.68 | < 5 | 8.7 | 0.366 |  | 20
| 1 | 6 | 7.38 | 7.1042 | 7.98 | 8.04 | 6.98 | 9.07 | 7.77 | 7.12 | 7.99 | 7.3801 | 7.5 | 7.55 | 6.63 | 8.42 | 7.21 | 6.64 | < 5 | 8.8 | 0.762
| Mean |  | 7.4717 | 7.21 | 7.9817 | 8.2083 | 7.38 | 8.1317 | 7.755 | 6.8617 | 8.3933 | 7.2287 | 7.52 | 7.6467 | 6.785 | 8.3033 | 7.4883 | 6.8333 |  | 8.8833 | 0.9524
| Median |  | 7.345 | 7.1997 | 7.79 | 8.17 | 7.41 | 8.095 | 7.805 | 6.955 | 8.22 | 7.2725 | 7.455 | 7.67 | 6.71 | 8.3 | 7.43 | 6.635 |  | 8.9 | 0.846
| Std Dev. |  | 0.35594 | 0.25429 | 0.84381 | 0.35074 | 0.25432 | 0.51959 | 0.31118 | 0.37579 | 0.42683 | 0.25104 | 0.28636 | 0.11378 | 0.35269 | 0.24873 | 0.31783 | 0.45719 |  | 0.17224 | 0.54878
| Rel.Std.Dev. |  | 4.764% | 3.527% | 10.57% | 4.273% | 3.446% | 6.39% | 4.013% | 5.477% | 5.085% | 3.473% | 3.808% | 1.488% | 5.198% | 2.996% | 4.244% | 6.691% |  | 1.939% | 57.62%
| PDM3 |  | -0.678% | -4.156% | 6.102% | 9.115% | -1.897% | 8.095% | 3.088% | -8.787% | 11.57% | -3.907% | -0.0355% | 1.648% | -9.806% | 10.38% | -0.4565% | -9.163% |  | 18.09% | -87.34%
| Z-Score (Absolute) |  | 0.02 | 0.3 | 0.53 | 0.77 | 0.12 | 0.69 | 0.29 | 0.67 | 0.97 | 0.28 | 0.03 | 0.17 | 0.76 | 0.88 | 0 | 0.7 | 5.37 | 1.5 | 7.03
| Analytical results for Ca in OREAS 251c (Certified Value 1.71 wt.%) | [21 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.6809 | 1.7177 | 1.7229 | 1.82 | 1.72 | 1.72 | 1.73 | 1.7 | 1.66 | 1.65 | 1.7353 | 1.7489 | 1.76 | 1.76 | 1.7912 | 1.67 | 1.58 | 1.65 | 1.72 | 6.4647 | 1
| 1 | 2 | 1.6922 | 1.7075 | 1.7066 | 1.75 | 1.73 | 1.72 | 1.75 | 1.71 | 1.66 | 1.67 | 1.7204 | 1.8181 | 1.75 | 1.68 | 1.7378 | 1.73 | 1.43 | 1.7 | 1.68 | 1.6873 | #N/A
| 1 | 3 | 1.7013 | 1.6978 | 1.7333 | 1.77 | 1.69 | 1.7 | 1.78 | 1.69 | 1.69 | 1.65 | 1.6735 | 1.8045 | 1.73 | 1.76 | 1.7725 | 1.7 | 1.61 | 1.7 | 1.72 | 1.7344 | 16
| 1 | 4 | 1.6984 | 1.647 | 1.7309 | 1.7 | 1.73 | 1.69 | 1.76 | 1.66 | 1.69 | 1.64 | 1.6699 | 1.8218 | 1.7 | 1.74 | 1.7618 | 1.67 | 1.53 | 1.68 | 1.72 | 1.6965 | 1.7149
| 1 | 5 | 1.7013 | 1.6564 | 1.7301 | 1.81 | 1.71 | 1.69 | 1.72 | 1.69 | 1.67 | 1.63 | 1.7326 | 1.759 | 1.73 | 1.75 | 1.7458 | 1.69 | 1.5 | 1.68 | 1.77 | 1.7104 | 21
| 1 | 6 | 1.6909 | 1.6925 | 1.7195 | 1.78 | 1.71 | 1.69 | 1.75 | 1.68 | 1.69 | 1.68 | 1.724 | 1.7346 | 1.76 | 1.71 | 1.7373 | 1.69 | 1.54 | 1.7 | 1.69 | 1.7252
| Mean |  | 1.6942 | 1.6865 | 1.7239 | 1.7717 | 1.715 | 1.7017 | 1.7483 | 1.6883 | 1.6767 | 1.6533 | 1.7093 | 1.7812 | 1.7383 | 1.7333 | 1.7577 | 1.6917 | 1.5317 | 1.685 | 1.7167 | 2.5031
| Median |  | 1.6953 | 1.6951 | 1.7265 | 1.775 | 1.715 | 1.695 | 1.75 | 1.69 | 1.68 | 1.65 | 1.7222 | 1.7817 | 1.74 | 1.745 | 1.7538 | 1.69 | 1.535 | 1.69 | 1.72 | 1.7178
| Std Dev. |  | 0.007877 | 0.028439 | 0.0099673 | 0.043551 | 0.015166 | 0.01472 | 0.02137 | 0.017224 | 0.015055 | 0.018619 | 0.029632 | 0.038107 | 0.023166 | 0.032042 | 0.021503 | 0.022286 | 0.063061 | 0.019748 | 0.031411 | 1.9409
| Rel.Std.Dev. |  | 0.4649% | 1.686% | 0.5782% | 2.458% | 0.8843% | 0.865% | 1.222% | 1.02% | 0.8979% | 1.126% | 1.734% | 2.139% | 1.333% | 1.849% | 1.223% | 1.317% | 4.117% | 1.172% | 1.83% | 77.54%
| PDM3 |  | -1.21% | -1.659% | 0.522% | 3.309% | 0.004826% | -0.7727% | 1.949% | -1.55% | -2.23% | -3.591% | -0.329% | 3.862% | 1.365% | 1.074% | 2.497% | -1.356% | -10.69% | -1.745% | 0.102% | 45.96%
| Z-Score (Absolute) |  | 0.47 | 0.67 | 0.31 | 1.55 | 0.07 | 0.27 | 0.94 | 0.62 | 0.92 | 1.53 | 0.07 | 1.79 | 0.68 | 0.55 | 1.19 | 0.53 | 4.69 | 0.71 | 0.12 | 20.57
| Analytical results for Cd in OREAS 251c (Certified Value 0.051 ppm) | [21 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 3
3
| 1 | 1 | 0.04 | < 0.2 | < 0.3 | 0.27 | 0.03 | 0.03 | 0.05 | 0.03 | 0.03 | 0.087043 | 0.06 | < 0.05 | 0.08 | 0.05 | 0.03 | 0.05 | 1.43 | < 0.5 | 2.123 |  | 1
| 1 | 2 | 0.06 | < 0.2 | < 0.3 | 0.27 | 0.02 | 0.02 | 0.03 | < 0.02 | 0.02 | 0.059106 | 0.05 | < 0.05 | 0.09 | 0.05 | 0.03 | 0.08 | 1.47 | < 0.5 | 0.747 |  | 28
| 1 | 3 | 0.04 | < 0.2 | < 0.3 | 0.28 | 0.03 | 0.04 | 0.04 | 0.04 | < 0.02 | 0.091073 | 0.06 | < 0.05 | 0.1 | 0.04 | 0.04 | 0.06 | 1.45 | < 0.5 | 0.884 |  | 16
| 1 | 4 | 0.06 | < 0.2 | < 0.3 | 0.27 | 0.03 | 0.04 | 0.02 | 0.02 | 0.06 | 0.057118 | 0.05 | 0.06 | 0.09 | 0.05 | 0.03 | 0.08 | 1.43 | < 0.5 | 0.903 |  | 0.051007
| 1 | 5 | 0.05 | < 0.2 | < 0.3 | 0.3 | 0.03 | 0.05 | < 0.02 | 0.02 | < 0.02 | 0.099577 | 0.06 | < 0.05 | 0.07 | 0.02 | < 0.02 | 0.06 | 1.32 | < 0.5 | 0.927 |  | 22
| 1 | 6 | 0.05 | < 0.2 | < 0.3 | 0.29 | 0.07 | 0.05 | 0.04 | < 0.02 | 0.05 | 0.087605 | 0.06 | 0.08 | 0.1 | 0.05 | 0.02 | 0.09 | 1.32 | < 0.5 | 0.941
| Mean |  | 0.05 |  |  | 0.28 | 0.035 | 0.038333 | 0.036 | 0.0275 | 0.04 | 0.080254 | 0.056667 | 0.07 | 0.088333 | 0.043333 | 0.03 | 0.07 | 1.4033 |  | 1.0875
| Median |  | 0.05 |  |  | 0.275 | 0.03 | 0.04 | 0.04 | 0.025 | 0.04 | 0.087324 | 0.06 | 0.07 | 0.09 | 0.05 | 0.03 | 0.07 | 1.43 |  | 0.915
| Std Dev. |  | 0.0089443 |  |  | 0.012649 | 0.017607 | 0.01169 | 0.011402 | 0.0095743 | 0.018257 | 0.017737 | 0.005164 | 0.014142 | 0.01169 | 0.012111 | 0.0070711 | 0.015492 | 0.066232 |  | 0.51203
| Rel.Std.Dev. |  | 17.89% |  |  | 4.518% | 50.31% | 30.5% | 31.67% | 34.82% | 45.64% | 22.1% | 9.113% | 20.2% | 13.23% | 27.95% | 23.57% | 22.13% | 4.72% |  | 47.08%
| PDM3 |  | -1.974% |  |  | 448.9% | -31.38% | -24.85% | -29.42% | -46.09% | -21.58% | 57.34% | 11.1% | 37.24% | 73.18% | -15.04% | -41.18% | 37.24% | 2651% |  | 2032%
| Z-Score (Absolute) |  | 0.17 | 1.1 | 2.36 | 5.65 | 0.55 | 0.46 | 0.63 | 0.89 | 0.67 | 0.6 | 0 | 0.42 | 0.8 | 0.34 | 0.76 | 0.34 | 34.05 | 4.89 | 26.07
| Analytical results for Ce in OREAS 251c (Certified Value 88 ppm) | [21 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 0
1
| 1 | 1 | 87.51 | 89.6581 | 87.4749 | 89.4 | 93.2 | 85.8 | 91.5 | 84.7 | 91.3 | 75.1 | 79.73 | 94.59 | 82.16 | 88 | 97.6 | 44.26 |  |  |  |  | 1
| 1 | 2 | 88.44 | 87.094 | 87.5658 | 87 | 90.6 | 85.7 | 91.7 | 78.8 | 90.7 | 76.7 | 78.83 | 92.02 | 85.04 | 88 | 95 | 85.008 |  |  |  |  | 29
| 1 | 3 | 85.76 | 88.8741 | 87.6416 | 91.1 | 96.3 | 91.9 | 87.7 | 79.6 | 89.5 | 76.2 | 77.6 | 93.44 | 84.08 | 88 | 95.5 | 87.976 |  |  |  |  | 16
| 1 | 4 | 86.62 | 84.577 | 87.4603 | 88.3 | 92 | 89.4 | 83 | 75.8 | 89.7 | 75.5 | 74.71 | 93 | 85.77 | 87.5 | 97.6 | 83.385 |  |  |  |  | 88.3028
| 1 | 5 | 87.53 | 86.8696 | 87.5629 | 88.7 | 92.9 | 79.6 | 85.4 | 80.1 | 89.1 | 74.6 | 75.68 | 92.39 | 82.69 | 86.9 | 96 | 86.473 |  |  |  |  | 23
| 1 | 6 | 87.21 | 88.791 | 87.3288 | 91.1 | 90.6 | 88.5 | 87.1 | 86 | 93 | 72.8 | 73.99 | 92.36 | 81.03 | 87.3 | 97.3 | 84.986
| Mean |  | 87.1783 | 87.644 | 87.5057 | 89.2667 | 92.6 | 86.8167 | 87.7333 | 80.8333 | 90.55 | 75.15 | 76.7567 | 92.9667 | 83.4617 | 87.6167 | 96.5 | 78.6813
| Median |  | 87.36 | 87.9425 | 87.5189 | 89.05 | 92.45 | 87.15 | 87.4 | 79.85 | 90.2 | 75.3 | 76.64 | 92.695 | 83.385 | 87.75 | 96.65 | 84.997
| Std Dev. |  | 0.91103 | 1.8544 | 0.10922 | 1.6207 | 2.1213 | 4.2358 | 3.411 | 3.8255 | 1.4502 | 1.3751 | 2.3175 | 0.94415 | 1.8113 | 0.46224 | 1.1454 | 16.9342
| Rel.Std.Dev. |  | 1.045% | 2.116% | 0.1248% | 1.816% | 2.291% | 4.879% | 3.888% | 4.733% | 1.602% | 1.83% | 3.019% | 1.016% | 2.17% | 0.5276% | 1.187% | 21.52%
| PDM3 |  | -1.273% | -0.746% | -0.9026% | 1.092% | 4.866% | -1.683% | -0.6449% | -8.459% | 2.545% | -14.9% | -13.08% | 5.282% | -5.482% | -0.777% | 9.283% | -10.9%
| Z-Score (Absolute) |  | 0.07 | 0.02 | 0.01 | 0.32 | 0.96 | 0.14 | 0.03 | 1.28 | 0.57 | 2.36 | 2.06 | 1.03 | 0.78 | 0.01 | 1.7 | 1.69
| Analytical results for Co in OREAS 251c (Certified Value 17.2 ppm) | [21 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 1
2
| 1 | 1 | 18.1 | 17.0272 | 16.8576 | 18 | 17.6 | 18.4 | 16.4 | 17.7 | 16.6 | 17 | 17.2063 | 17 | 17 | 17.8 | 18 | 16.4 | 16.8 | 21.33 | 17 | 39.52 | 1
| 1 | 2 | 18.2 | 17.0025 | 16.7584 | 18 | 16.9 | 17.6 | 16.3 | 17.4 | 15.2 | 17.6 | 17.5314 | 17.3 | 17 | 15.7 | 17.6 | 17.2 | 17.2 | 21.85 | 17 | 17.879 | 15
| 1 | 3 | 18.1 | 16.7899 | 16.2936 | 17 | 17.3 | 18 | 17.5 | 16.8 | 16 | 17 | 17.5337 | 17.4 | 17 | 17.1 | 17.7 | 17 | 17.9 | 21.14 | 17 | 17.899 | 16
| 1 | 4 | 18 | 16.2344 | 18.1128 | 18 | 16.6 | 17.8 | 16.9 | 16.1 | 15 | 17.4 | 17.3022 | 17.1 | 17 | 16.5 | 17.8 | 16.8 | 16.7 | 21.03 | 17 | 18.205 | 17.1892
| 1 | 5 | 18 | 16.3623 | 16.6176 | 17 | 17.7 | 17.5 | 16.6 | 16.2 | 15.5 | 16.5 | 17.2162 | 17 | 17 | 17.3 | 17.5 | 17 | 16.7 | 21.1 | 18 | 17.824 | 24
| 1 | 6 | 18.3 | 16.8972 | 17.0648 | 17 | 17.9 | 17.6 | 16.5 | 16.9 | 16.2 | 17.2 | 17.072 | 16.4 | 17 | 16.9 | 17.4 | 16.8 | 17 | 21.22 | 17 | 18.125
| Mean |  | 18.1167 | 16.7189 | 16.9508 | 17.5 | 17.3333 | 17.8167 | 16.7 | 16.85 | 15.75 | 17.1167 | 17.3103 | 17.0333 | 17 | 16.8833 | 17.6667 | 16.8667 | 17.05 | 21.2783 | 17.1667 | 21.5753
| Median |  | 18.1 | 16.8435 | 16.808 | 17.5 | 17.45 | 17.7 | 16.55 | 16.85 | 15.75 | 17.1 | 17.2592 | 17.05 | 17 | 17 | 17.65 | 16.9 | 16.9 | 21.18 | 17 | 18.012
| Std Dev. |  | 0.1169 | 0.33889 | 0.6248 | 0.54772 | 0.50067 | 0.33714 | 0.44272 | 0.63482 | 0.61887 | 0.38166 | 0.18726 | 0.35024 | 0 | 0.72226 | 0.21602 | 0.27325 | 0.45935 | 0.29849 | 0.40825 | 8.7923
| Rel.Std.Dev. |  | 0.6453% | 2.027% | 3.686% | 3.13% | 2.888% | 1.892% | 2.651% | 3.767% | 3.929% | 2.23% | 1.082% | 2.056% | 0% | 4.278% | 1.223% | 1.62% | 2.694% | 1.403% | 2.378% | 40.75%
| PDM3 |  | 5.396% | -2.736% | -1.387% | 1.808% | 0.8388% | 3.651% | -2.846% | -1.973% | -8.372% | -0.4217% | 0.7047% | -0.9065% | -1.1% | -1.779% | 2.778% | -1.876% | -0.8095% | 23.79% | -0.1308% | 25.52%
| Z-Score (Absolute) |  | 2.54 | 0.79 | 0.24 | NA | 0.67 | 1.82 | 0.83 | 0.48 | 3.09 | 0.16 | 0.62 | 0.04 | NA | 0.4 | 1.47 | 0.44 | 0 | 10.06 | NA | 10.77
| Analytical results for Cr in OREAS 251c (Certified Value 120 ppm) | [21 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 139 | 124.5278 | 81 | 116 | 122 | 121 | 121 | 123 | 120 | 115.7583 | 106.1 | 116 | 106 | 126 | 108 | 105 | 127.8 | 130 | 130.143 |  | 1
| 1 | 2 | 131 | 123.7135 | 80 | 115 | 121 | 122 | 121 | 125 | 123 | 112.5705 | 109.5 | 125 | 97 | 130 | 109 | 99 | 124.2 | 130 | 135.139 |  | 31
| 1 | 3 | 124 | 125.418 | 81 | 114 | 122 | 118 | 122 | 124 | 121 | 111.6958 | 108.8 | 123 | 110 | 132 | 112 | 103 | 121.4 | 140 | 131.432 |  | 16
| 1 | 4 | 130 | 121.3513 | 82 | 113 | 121 | 123 | 118 | 124 | 122 | 117.8139 | 109.5 | 123 | 110 | 130 | 107 | 102 | 124.7 | 130 | 132.053 |  | 119.6819
| 1 | 5 | 130 | 121.9544 | 79 | 114 | 118 | 122 | 121 | 122 | 119 | 118.4967 | 105.3 | 115 | 105 | 132 | 109 | 108 | 124.7 | 130 | 134.758 |  | 25
| 1 | 6 | 132 | 123.1023 | 82 | 112 | 120 | 121 | 121 | 124 | 119 | 112.7583 | 106 | 113 | 110 | 129 | 108 | 105 | 124 | 130 | 132.96
| Mean |  | 131 | 123.3446 | 80.8333 | 114 | 120.6667 | 121.1667 | 120.6667 | 123.6667 | 120.6667 | 114.8489 | 107.5333 | 119.1667 | 106.3333 | 129.8333 | 108.8333 | 103.6667 | 124.4667 | 131.6667 | 132.7475
| Median |  | 130.5 | 123.4079 | 81 | 114 | 121 | 121.5 | 121 | 124 | 120.5 | 114.2583 | 107.45 | 119.5 | 108 | 130 | 108.5 | 104 | 124.45 | 130 | 132.5065
| Std Dev. |  | 4.8166 | 1.5362 | 1.169 | 1.4142 | 1.5055 | 1.7224 | 1.3663 | 1.0328 | 1.633 | 2.913 | 1.9356 | 5.0761 | 5.0859 | 2.2286 | 1.7224 | 3.0768 | 2.0452 | 4.0825 | 1.9392
| Rel.Std.Dev. |  | 3.677% | 1.245% | 1.446% | 1.241% | 1.248% | 1.422% | 1.132% | 0.8351% | 1.353% | 2.536% | 1.8% | 4.26% | 4.783% | 1.717% | 1.583% | 2.968% | 1.643% | 3.101% | 1.461%
| PDM3 |  | 9.457% | 3.06% | -32.46% | -4.748% | 0.8228% | 1.241% | 0.8228% | 3.329% | 0.8228% | -4.038% | -10.15% | -0.4305% | -11.15% | 8.482% | -9.065% | -13.38% | 3.998% | 10.01% | 10.92%
| Z-Score (Absolute) |  | 1.05 | 0.27 | 4.03 | 0.67 | 0 | 0.05 | 0 | 0.3 | 0 | 0.59 | 1.33 | 0.15 | 1.45 | 0.93 | 1.2 | 1.72 | 0.38 | 1.11 | 1.22
| Analytical results for Cs in OREAS 251c (Certified Value 5.94 ppm) | [21 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 5.84 | 6.0645 | 6.14 | 6.12 | 5.66 | 5.82 | 5.49 | 6.25 | 6.1979 | 5.5 | 5.8 | 6.1 | 6.32 | 6 | 6.6 | 6 |  |  |  |  | 1
| 1 | 2 | 5.9 | 5.9881 | 6.02 | 6.03 | 5.71 | 5.89 | 4.94 | 6.21 | 5.9403 | 5.6 | 6 | 5.2 | 6.18 | 6.1 | 7 | 5.9 |  |  |  |  | 32
| 1 | 3 | 5.86 | 5.9525 | 6.1 | 6.17 | 5.99 | 5.53 | 5.31 | 6.33 | 6.1151 | 5.6 | 5.8 | 5.9 | 6.4 | 6.1 | 6.9 | 6 |  |  |  |  | 16
| 1 | 4 | 5.89 | 5.7413 | 6.04 | 6.04 | 5.89 | 5.45 | 5.03 | 6.04 | 6.0546 | 5.5 | 5.9 | 5.9 | 6.28 | 6.1 | 6.9 | 6 |  |  |  |  | 5.9419
| 1 | 5 | 5.87 | 5.8241 | 6.22 | 6.02 | 5.53 | 5.42 | 5.25 | 6.18 | 5.9336 | 5.5 | 5.8 | 6 | 6.25 | 6 | 6.7 | 6 |  |  |  |  | 26
| 1 | 6 | 5.83 | 6.0179 | 6.21 | 5.97 | 5.67 | 5.59 | 5.3 | 6.31 | 6.1283 | 5.4 | 6 | 5.6 | 6.09 | 5.9 | 6.6 | 6
| Mean |  | 5.865 | 5.9314 | 6.1217 | 6.0583 | 5.7417 | 5.6167 | 5.22 | 6.22 | 6.0616 | 5.5167 | 5.8833 | 5.7833 | 6.2533 | 6.0333 | 6.7833 | 5.9833
| Median |  | 5.865 | 5.9703 | 6.12 | 6.035 | 5.69 | 5.56 | 5.275 | 6.23 | 6.0849 | 5.5 | 5.85 | 5.9 | 6.265 | 6.05 | 6.8 | 6
| Std Dev. |  | 0.027386 | 0.1237 | 0.084004 | 0.073052 | 0.1681 | 0.19531 | 0.2014 | 0.10507 | 0.1068 | 0.075277 | 0.098319 | 0.33116 | 0.10838 | 0.08165 | 0.17224 | 0.040825
| Rel.Std.Dev. |  | 0.4669% | 2.085% | 1.372% | 1.206% | 2.928% | 3.477% | 3.858% | 1.689% | 1.762% | 1.365% | 1.671% | 5.726% | 1.733% | 1.353% | 2.539% | 0.6823%
| PDM3 |  | -1.294% | -0.1765% | 3.026% | 1.96% | -3.37% | -5.473% | -12.15% | 4.681% | 2.016% | -7.156% | -0.9854% | -2.668% | 5.242% | 1.539% | 14.16% | 0.6976%
| Z-Score (Absolute) |  | 0.37 | 0.1 | 0.65 | 0.4 | 0.86 | 1.36 | 2.94 | 1.05 | 0.42 | 1.76 | 0.3 | 0.69 | 1.18 | 0.3 | 3.29 | 0.1
| Analytical results for Cu in OREAS 251c (Certified Value 27.5 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 28.5 | 27.8914 | 28.1035 | 29 | 23.6 | 27 | 27.6 | 28.7 | 26.3 | 27.7 | 25.3744 | 20.1 | 28 | 29.1 | 29.5 | 27.6 | 5.2 | 30.77 | 8 | 177.587 | 1
| 1 | 2 | 29.2 | 27.6339 | 27.106 | 29 | 23.4 | 27.2 | 26.9 | 28.7 | 25.5 | 27.2 | 24.6541 | 21.8 | 28 | 29.1 | 29.2 | 28.7 | 4.7 | 29.21 | 8 | 34.26 | 17
| 1 | 3 | 28.9 | 27.2613 | 28.2715 | 29 | 25.4 | 27.4 | 28.8 | 27.5 | 25.7 | 27.6 | 25.4543 | 21.1 | 28 | 28.5 | 29.3 | 28.3 | 5.5 | 29.3 | 8 | 34.797 | 16
| 1 | 4 | 29.3 | 26.399 | 28.177 | 29 | 23.9 | 26.9 | 28.5 | 26.5 | 25.2 | 27.3 | 25.1681 | 21.2 | 26 | 28.2 | 29.6 | 27.9 | 4.9 | 29.34 | 8 | 33.604 | 27.5167
| 1 | 5 | 28.3 | 26.7224 | 28.0405 | 29 | 23.3 | 26.7 | 28.1 | 27.5 | 26 | 27.1 | 24.8229 | 19.1 | 27 | 28.8 | 29.5 | 28.1 | 4.4 | 31.81 | 6 | 35.151 | 27
| 1 | 6 | 29.6 | 27.398 | 29.185 | 29 | 24.4 | 26.1 | 28.2 | 28 | 26.4 | 27.9 | 24.6815 | 19.5 | 28 | 28.6 | 29.3 | 27.7 | 6.5 | 29.76 | 8 | 33.596
| Mean |  | 28.9667 | 27.2177 | 28.1473 | 29 | 24 | 26.8833 | 28.0167 | 27.8167 | 25.85 | 27.4667 | 25.0259 | 20.4667 | 27.5 | 28.7167 | 29.4 | 28.05 | 5.2 | 30.0317 | 7.6667 | 58.1658
| Median |  | 29.05 | 27.3297 | 28.1403 | 29 | 23.75 | 26.95 | 28.15 | 27.75 | 25.85 | 27.45 | 24.9955 | 20.6 | 28 | 28.7 | 29.4 | 28 | 5.05 | 29.55 | 8 | 34.5285
| Std Dev. |  | 0.49666 | 0.56163 | 0.66193 | 0 | 0.79246 | 0.4535 | 0.67946 | 0.84004 | 0.46797 | 0.31411 | 0.35304 | 1.0633 | 0.83666 | 0.35449 | 0.15492 | 0.40866 | 0.74297 | 1.046 | 0.8165 | 58.5075
| Rel.Std.Dev. |  | 1.715% | 2.063% | 2.352% | 0% | 3.302% | 1.687% | 2.425% | 3.02% | 1.81% | 1.144% | 1.411% | 5.195% | 3.042% | 1.234% | 0.5269% | 1.457% | 14.29% | 3.483% | 10.65% | 100.6%
| PDM3 |  | 5.269% | -1.087% | 2.292% | 5.391% | -12.78% | -2.302% | 1.817% | 1.09% | -6.057% | -0.1817% | -9.052% | -25.62% | -0.06061% | 4.361% | 6.844% | 1.938% | -81.1% | 9.14% | -72.14% | 111.4%
| Z-Score (Absolute) |  | 0.67 | 0.35 | 0.19 | NA | 2.24 | 0.55 | 0.12 | 0 | 1.15 | 0.21 | 1.64 | 4.31 | NA | 0.53 | 0.93 | 0.14 | 13.26 | 1.3 | NA | 17.8
Results from laboratories 1.07, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Dy in OREAS 251c (Certified Value 4.28 ppm) | [21 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 4.34 | 4.6561 | 4.5072 | 4.4 | 4 | 4.4 | 4 | 4.18 | 4.75 | 4.149 | [10 empty] | 1
| 1 | 2 | 4.3 | 4.1256 | 4.5942 | 4.2 | 4 | 4.42 | 4 | 4.12 | 4.6 | 4.16 | [10 empty] | 4
| 1 | 3 | 4.16 | 4.1735 | 4.5772 | 4.4 | 3.9 | 4.37 | 4.4 | 4.13 | 4.75 | 4.356 | [10 empty] | 16
| 1 | 4 | 4.2 | 4.0618 | 4.4622 | 4.4 | 3.8 | 4.46 | 4.1 | 4.1 | 4.75 | 4.233 | [10 empty] | 4.2788
| 1 | 5 | 4.41 | 4.3465 | 4.5829 | 4.3 | 3.9 | 4.41 | 4 | 4.25 | 4.75 | 4.344 | [10 empty] | 28
| 1 | 6 | 4.15 | 4.0914 | 4.568 | 4.4 | 3.7 | 4.48 | 4 | 4.24 | 4.6 | 4.393
| Mean |  | 4.26 | 4.2425 | 4.5486 | 4.35 | 3.8833 | 4.4233 | 4.0833 | 4.17 | 4.7 | 4.2725
| Median |  | 4.25 | 4.1496 | 4.5726 | 4.4 | 3.9 | 4.415 | 4 | 4.155 | 4.75 | 4.2885
| Std Dev. |  | 0.10602 | 0.22617 | 0.052206 | 0.083666 | 0.1169 | 0.040332 | 0.16021 | 0.063875 | 0.07746 | 0.1059
| Rel.Std.Dev. |  | 2.489% | 5.331% | 1.148% | 1.923% | 3.01% | 0.9118% | 3.923% | 1.532% | 1.648% | 2.479%
| PDM3 |  | -0.4383% | -0.8478% | 6.307% | 1.665% | -9.241% | 3.379% | -4.567% | -2.542% | 9.845% | -0.1462%
| Z-Score (Absolute) |  | 0.03 | 0.13 | 1.5 | 0.45 | 2.04 | 0.84 | 0.97 | 0.51 | 2.31 | 0.03
| Analytical results for Er in OREAS 251c (Certified Value 1.96 ppm) | [21 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.98 | 2.1773 | 2.2179 | 2 | 2.1 | 1.86 | 1.9 | 1.84 | 2.1 | [11 empty] | 1
| 1 | 2 | 2 | 2.0051 | 2.2041 | 2.1 | 1.7 | 1.87 | 1.8 | 1.88 | 2.15 | [11 empty] | 5
| 1 | 3 | 1.95 | 2.0066 | 2.1501 | 2.2 | 1.8 | 1.86 | 1.7 | 1.91 | 2.15 | [11 empty] | 16
| 1 | 4 | 1.95 | 1.8532 | 2.1616 | 2.1 | 1.7 | 1.88 | 1.8 | 1.78 | 2.1 | [11 empty] | 1.9597
| 1 | 5 | 1.95 | 2.0344 | 2.1749 | 2.1 | 1.7 | 1.93 | 1.8 | 1.83 | 2.15 | [11 empty] | 29
| 1 | 6 | 1.97 | 1.8653 | 2.1917 | 2.2 | 1.8 | 1.93 | 1.7 | 1.82 | 2.1
| Mean |  | 1.9667 | 1.9903 | 2.1834 | 2.1167 | 1.8 | 1.8883 | 1.7833 | 1.8433 | 2.125
| Median |  | 1.96 | 2.0058 | 2.1833 | 2.1 | 1.75 | 1.875 | 1.8 | 1.835 | 2.125
| Std Dev. |  | 0.020656 | 0.11987 | 0.025874 | 0.075277 | 0.15492 | 0.033116 | 0.075277 | 0.045898 | 0.027386
| Rel.Std.Dev. |  | 1.05% | 6.023% | 1.185% | 3.556% | 8.607% | 1.754% | 4.221% | 2.49% | 1.289%
| PDM3 |  | 0.3571% | 1.564% | 11.42% | 8.011% | -8.148% | -3.64% | -8.998% | -5.936% | 8.437%
| Z-Score (Absolute) |  | 0 | 0.11 | 0.97 | 0.67 | 0.75 | 0.35 | 0.82 | 0.55 | 0.71
| Analytical results for Eu in OREAS 251c (Certified Value 1.68 ppm) | [21 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 1.78 | 1.8314 | 2.1707 | 1.85 | 1.7 | 1.75 | 1.6 | 1.64 | 2.1 | 1.356 | [10 empty] | 1
| 1 | 2 | 1.78 | 1.7195 | 2.187 | 1.76 | 1.6 | 1.72 | 1.7 | 1.65 | 2.1 | 1.634 | [10 empty] | 36
| 1 | 3 | 1.7 | 1.7412 | 2.17 | 1.76 | 1.6 | 1.76 | 1.6 | 1.65 | 2.1 | 1.632 | [10 empty] | 16
| 1 | 4 | 1.71 | 1.6718 | 2.1015 | 1.71 | 1.5 | 1.77 | 1.7 | 1.53 | 2.1 | 1.662 | [10 empty] | 1.6785
| 1 | 5 | 1.76 | 1.6601 | 2.1719 | 1.73 | 1.5 | 1.72 | 1.6 | 1.54 | 2.1 | 1.665 | [10 empty] | 30
| 1 | 6 | 1.74 | 1.76 | 2.201 | 1.76 | 1.6 | 1.74 | 1.6 | 1.58 | 2.1 | 1.658
| Mean |  | 1.745 | 1.7307 | 2.167 | 1.7617 | 1.5833 | 1.7433 | 1.6333 | 1.5983 | 2.1 | 1.6012
| Median |  | 1.75 | 1.7303 | 2.1713 | 1.76 | 1.6 | 1.745 | 1.6 | 1.61 | 2.1 | 1.646
| Std Dev. |  | 0.034496 | 0.062751 | 0.034339 | 0.047924 | 0.075277 | 0.020656 | 0.05164 | 0.055648 | 0 | 0.12095
| Rel.Std.Dev. |  | 1.977% | 3.626% | 1.585% | 2.72% | 4.754% | 1.185% | 3.162% | 3.482% | 0% | 7.554%
| PDM3 |  | 3.96% | 3.107% | 29.1% | 4.953% | -5.671% | 3.861% | -2.692% | -4.778% | 25.11% | -4.609%
| Z-Score (Absolute) |  | 0.05 | 0.04 | 2.42 | 0.14 | 0.87 | 0.04 | 0.58 | 0.78 | 2.04 | 0.76
| Analytical results for Fe in OREAS 251c (Certified Value 4.28 wt.%) | [21 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 4.22 | 4.1882 | 4.3268 | 4.34 | 4.34 | 4.18 | 4.22 | 4.37 | 4.18 | 4.14 | 4.3882 | 4.0636 | 4.57 | 4.41 | 4.4048 | 4.17 | 4.46 | 4.26 | 4.2 | 8.0866 | 1
| 1 | 2 | 4.26 | 4.1434 | 4.3388 | 4.29 | 4.34 | 4.2 | 4.24 | 4.44 | 4.18 | 4.17 | 4.4198 | 4.1925 | 4.41 | 4.3 | 4.3118 | 4.31 | 4.4 | 4.39 | 4.11 | 4.6518 | #N/A
| 1 | 3 | 4.24 | 4.2222 | 4.3454 | 4.47 | 4.34 | 4.17 | 4.29 | 4.36 | 4.25 | 4.14 | 4.5168 | 4.1783 | 4.22 | 4.37 | 4.3702 | 4.28 | 4.5 | 4.4 | 4.02 | 4.7272 | 16
| 1 | 4 | 4.24 | 4.0407 | 4.3585 | 4.35 | 4.39 | 4.11 | 4.27 | 4.31 | 4.24 | 4.11 | 4.4435 | 4.2039 | 4.32 | 4.28 | 4.3464 | 4.27 | 4.37 | 4.35 | 4.22 | 4.6797 | 4.2771
| 1 | 5 | 4.24 | 4.0809 | 4.3471 | 4.37 | 4.34 | 4.09 | 4.24 | 4.36 | 4.19 | 4.08 | 4.5208 | 4.0785 | 4.31 | 4.35 | 4.2623 | 4.32 | 4.36 | 4.35 | 4.07 | 4.6575 | 31
| 1 | 6 | 4.2 | 4.1423 | 4.3361 | 4.34 | 4.33 | 4.14 | 4.25 | 4.34 | 4.28 | 4.17 | 4.5058 | 4.0418 | 4.31 | 4.3 | 4.2547 | 4.27 | 4.34 | 4.4 | 4.15 | 4.6865
| Mean |  | 4.2333 | 4.1363 | 4.3421 | 4.36 | 4.3467 | 4.1483 | 4.2517 | 4.3633 | 4.22 | 4.135 | 4.4658 | 4.1264 | 4.3567 | 4.335 | 4.325 | 4.27 | 4.405 | 4.3583 | 4.1283 | 5.2482
| Median |  | 4.24 | 4.1428 | 4.3421 | 4.345 | 4.34 | 4.155 | 4.245 | 4.36 | 4.215 | 4.14 | 4.4746 | 4.1284 | 4.315 | 4.325 | 4.3291 | 4.275 | 4.385 | 4.37 | 4.13 | 4.6831
| Std Dev. |  | 0.020656 | 0.066859 | 0.010822 | 0.06 | 0.021602 | 0.042622 | 0.024833 | 0.043205 | 0.042426 | 0.035071 | 0.056298 | 0.072752 | 0.12061 | 0.0501 | 0.059861 | 0.053292 | 0.06253 | 0.053448 | 0.07679 | 1.3908
| Rel.Std.Dev. |  | 0.4879% | 1.616% | 0.2492% | 1.376% | 0.497% | 1.027% | 0.5841% | 0.9902% | 1.005% | 0.8482% | 1.261% | 1.763% | 2.768% | 1.156% | 1.384% | 1.248% | 1.42% | 1.226% | 1.86% | 26.5%
| PDM3 |  | -1.023% | -3.292% | 1.52% | 1.939% | 1.627% | -3.01% | -0.5943% | 2.016% | -1.335% | -3.322% | 4.412% | -3.522% | 1.861% | 1.354% | 1.121% | -0.1657% | 2.991% | 1.9% | -3.478% | 22.7%
| Z-Score (Absolute) |  | 0.85 | 1.7 | 0.11 | 0.26 | 0.15 | 1.6 | 0.69 | 0.29 | 0.97 | 1.72 | 1.19 | 1.79 | 0.23 | 0.04 | 0.04 | 0.53 | 0.66 | 0.25 | 1.77 | 8.08
| Analytical results for Ga in OREAS 251c (Certified Value 18.9 ppm) | [21 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 19.53 | 19.8039 | 23.565 | 20.4 | 18.8 | 19.7 | 18.35 | 19 | 17.85 | 18.5 | 19.1215 | 19.13 | 19.5 | 18.09 | 21.4 | 17.89 | 18.2 | 19.2 |  |  | 1
| 1 | 2 | 19.78 | 19.4622 | 23.598 | 19.5 | 19.5 | 18.85 | 18.15 | 19.4 | 17 | 18.7 | 18.8373 | 19.34 | 19.3 | 18.6 | 20.49 | 18.27 | 18.9 | 19 |  |  | 37
| 1 | 3 | 19.48 | 19.299 | 23.552 | 20.9 | 18.6 | 19.9 | 19.2 | 18.35 | 17.65 | 18.6 | 19.3441 | 19.48 | 18.7 | 17.79 | 21.2 | 18.34 | 19.2 | 19 |  |  | 16
| 1 | 4 | 19.52 | 18.5537 | 23.529 | 20.9 | 19 | 19.7 | 18.75 | 17.9 | 16.8 | 18.85 | 19.2222 | 19.26 | 18.9 | 17.43 | 21.38 | 18.34 | 18.7 | 19.6 |  |  | 18.8914
| 1 | 5 | 19.49 | 18.9881 | 23.567 | 20.8 | 19 | 19.6 | 18.75 | 17.85 | 17.45 | 18.75 | 18.8676 | 18.9 | 18.6 | 17.5 | 20.96 | 18.56 | 18.6 | 19.4 |  |  | 32
| 1 | 6 | 19.47 | 19.5136 | 23.636 | 19.5 | 19.2 | 18.5 | 18.35 | 18.4 | 18.5 | 19.15 | 19.1163 | 18.32 | 19.7 | 17.38 | 20.59 | 18.05 | 18.5 | 19
| Mean |  | 19.545 | 19.2701 | 23.5745 | 20.3333 | 19.0167 | 19.375 | 18.5917 | 18.4833 | 17.5417 | 18.7583 | 19.0848 | 19.0717 | 19.1167 | 17.7983 | 21.0033 | 18.2417 | 18.6833 | 19.2
| Median |  | 19.505 | 19.3806 | 23.566 | 20.6 | 19 | 19.65 | 18.55 | 18.375 | 17.55 | 18.725 | 19.1189 | 19.195 | 19.1 | 17.645 | 21.08 | 18.305 | 18.65 | 19.1
| Std Dev. |  | 0.11743 | 0.44153 | 0.03757 | 0.67132 | 0.31252 | 0.56192 | 0.38264 | 0.6121 | 0.61271 | 0.22675 | 0.19846 | 0.41763 | 0.44907 | 0.47486 | 0.39348 | 0.23744 | 0.34303 | 0.25298
| Rel.Std.Dev. |  | 0.6008% | 2.291% | 0.1594% | 3.302% | 1.643% | 2.9% | 2.058% | 3.312% | 3.493% | 1.209% | 1.04% | 2.19% | 2.349% | 2.668% | 1.873% | 1.302% | 1.836% | 1.318%
| PDM3 |  | 3.46% | 2.005% | 24.79% | 7.633% | 0.6633% | 2.56% | -1.586% | -2.16% | -7.145% | -0.7042% | 1.024% | 0.9544% | 1.193% | -5.786% | 11.18% | -3.439% | -1.101% | 1.634%
| Z-Score (Absolute) |  | 0.73 | 0.3 | 7.04 | 1.96 | 0.1 | 0.46 | 0.76 | 0.93 | 2.4 | 0.5 | 0.01 | 0.01 | 0.06 | 2 | 3.01 | 1.31 | 0.62 | 0.19
| Analytical results for Gd in OREAS 251c (Certified Value 5.8 ppm) | [21 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [12 empty] | 2
3
| 1 | 1 | 5.81 | 6.0503 | 6.1 | 6.2 | 5.99 | 5.4 | 6.1 | 6 | [12 empty] | 1
| 1 | 2 | 5.87 | 5.7118 | 5.9 | 5.5 | 6.07 | 5.5 | 6.05 | 6.2 | [12 empty] | 6
| 1 | 3 | 5.6 | 5.7706 | 6 | 5.6 | 5.96 | 5.5 | 6.03 | 6 | [12 empty] | 16
| 1 | 4 | 5.56 | 5.6061 | 5.7 | 5.7 | 5.96 | 5.7 | 5.94 | 6 | [12 empty] | 5.8003
| 1 | 5 | 5.71 | 5.7098 | 5.7 | 5.5 | 5.91 | 5.4 | 5.98 | 6 | [12 empty] | 33
| 1 | 6 | 5.66 | 5.8778 | 5.9 | 5.3 | 5.91 | 5.4 | 6.06 | 6
| Mean |  | 5.7017 | 5.7877 | 5.8833 | 5.6333 | 5.9667 | 5.4833 | 6.0267 | 6.0333
| Median |  | 5.685 | 5.7412 | 5.9 | 5.55 | 5.96 | 5.45 | 6.04 | 6
| Std Dev. |  | 0.12024 | 0.15633 | 0.16021 | 0.30768 | 0.059554 | 0.1169 | 0.05785 | 0.08165
| Rel.Std.Dev. |  | 2.109% | 2.701% | 2.723% | 5.462% | 0.9981% | 2.132% | 0.9599% | 1.353%
| PDM3 |  | -1.701% | -0.2175% | 1.431% | -2.879% | 2.868% | -5.465% | 3.902% | 4.017%
| Z-Score (Absolute) |  | 0.56 | 0.2 | 0.2 | 0.84 | 0.54 | 1.46 | 0.79 | 0.82
| Analytical results for Ge in OREAS 251c (Indicative Value 0.19 ppm) | [21 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.7 | 0.26072 | 0.2531 | 0.1 | 0.16 | 0.13 | 0.17 | 0.14 | 0.19 | 0.24214 | NA | 1.3 | 0.16 | 1.1 | 1.51 | [5 empty] | 1
| 1 | 2 | 1.7 | 0.24968 | 0.2504 | 0.1 | 0.14 | 0.14 | 0.18 | 0.14 | 0.21 | 0.23186 | NA | 1.3 | 0.12 | 1.1 | 1.62 | [5 empty] | 2
| 1 | 3 | 1.5 | 0.22255 | 0.2395 | 0.1 | 0.16 | 0.14 | 0.17 | 0.15 | 0.21 | 0.27919 | NA | 1.3 | 0.09 | 1.1 | 1.61 | [5 empty] | 16
| 1 | 4 | 1.6 | 0.2585 | 0.291 | < 0.1 | 0.14 | 0.14 | 0.16 | 0.14 | 0.22 | 0.21074 | NA | 1.4 | 0.15 | 1.1 | 1.55 | [5 empty] | 0.18623
| 1 | 5 | 1.6 | 0.23416 | 0.2741 | 0.1 | 0.15 | 0.17 | 0.17 | 0.15 | 0.19 | 0.24557 | NA | 1.2 | 0.06 | 1.3 | 1.56 | [5 empty] | 8
| 1 | 6 | 1.7 | 0.26145 | 0.2876 | 0.1 | 0.14 | 0.15 | 0.16 | 0.14 | 0.2 | 0.27427 | NA | 1.4 | 0.09 | 1.2 | 1.51
| Mean |  | 1.6333 | 0.24784 | 0.26595 | 0.1 | 0.14833 | 0.145 | 0.16833 | 0.14333 | 0.20333 | 0.2473 |  | 1.3167 | 0.11167 | 1.15 | 1.56
| Median |  | 1.65 | 0.25409 | 0.2636 | 0.1 | 0.145 | 0.14 | 0.17 | 0.14 | 0.205 | 0.24386 |  | 1.3 | 0.105 | 1.1 | 1.555
| Std Dev. |  | 0.08165 | 0.016091 | 0.021303 | 0 | 0.0098319 | 0.013784 | 0.0075277 | 0.005164 | 0.012111 | 0.02588 |  | 0.075277 | 0.038687 | 0.083666 | 0.047329
| Rel.Std.Dev. |  | 4.999% | 6.492% | 8.01% | 0% | 6.628% | 9.506% | 4.472% | 3.603% | 5.956% | 10.47% |  | 5.717% | 34.64% | 7.275% | 3.034%
| PDM3 |  | 777% | 33.08% | 42.81% | -46.3% | -20.35% | -22.14% | -9.611% | -23.04% | 9.183% | 32.79% |  | 607% | -40.04% | 517.5% | 737.7%
| Z-Score (Absolute) |  | NA | 0.8 | 1.04 | NA | 0.49 | 0.53 | 0.23 | 0.55 | 0.23 | 0.8 | NA | NA | 0.96 | NA | 17.83
Results from laboratories 1.02, 1.07, 1.16 and 1.18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 251c (Certified Value 4.7 ppm) | [21 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 4.69 | 4.7038 | 3.3 | 5 | 4.7 | 4.7 | 4.6 | 4.6 | 4.775 | 4.64 | 4.7 | 4.82 | 5.1 | 4.34 | 4.73 | 5.8 |  |  |  |  | 1
| 1 | 2 | 4.66 | 4.4813 | 3 | 4.9 | 4.6 | 4.7 | 4.5 | 4.9 | 4.533 | 4.61 | 4.7 | 4.41 | 4.78 | 4.39 | 4.77 | 6 |  |  |  |  | 8
| 1 | 3 | 4.47 | 4.5908 | 0.8 | 5 | 5 | 4.6 | 4.9 | 4.9 | 5.0262 | 4.6 | 4.6 | 4.57 | 4.87 | 4.37 | 4.89 | 5.8 |  |  |  |  | 16
| 1 | 4 | 4.56 | 4.3728 | 3.8 | 5 | 4.8 | 4.5 | 4.4 | 4.9 | 4.9605 | 4.59 | 4.7 | 4.47 | 5 | 4.36 | 4.8 | 5.8 |  |  |  |  | 4.7008
| 1 | 5 | 4.63 | 4.5998 | 2.2 | 5 | 4.9 | 4.8 | 4.6 | 5 | 5.0726 | 4.53 | 4.5 | 4.5 | 4.86 | 4.41 | 4.86 | 5.8 |  |  |  |  | 34
| 1 | 6 | 4.59 | 4.6939 | 2.5 | 4.9 | 4.8 | 5 | 5 | 4.8 | 4.6437 | 4.5 | 4.6 | 4.36 | 4.76 | 4.21 | 4.81 | 5.8
| Mean |  | 4.6 | 4.5738 | 2.6 | 4.9667 | 4.8 | 4.7167 | 4.6667 | 4.85 | 4.8352 | 4.5783 | 4.6333 | 4.5217 | 4.895 | 4.3467 | 4.81 | 5.8333
| Median |  | 4.61 | 4.5953 | 2.75 | 5 | 4.8 | 4.7 | 4.6 | 4.9 | 4.8677 | 4.595 | 4.65 | 4.485 | 4.865 | 4.365 | 4.805 | 5.8
| Std Dev. |  | 0.078994 | 0.12748 | 1.0488 | 0.05164 | 0.14142 | 0.17224 | 0.23381 | 0.13784 | 0.21916 | 0.052694 | 0.08165 | 0.16315 | 0.13142 | 0.071181 | 0.05831 | 0.08165
| Rel.Std.Dev. |  | 1.717% | 2.787% | 40.34% | 1.04% | 2.946% | 3.652% | 5.01% | 2.842% | 4.533% | 1.151% | 1.762% | 3.608% | 2.685% | 1.638% | 1.212% | 1.4%
| PDM3 |  | -2.145% | -2.703% | -44.69% | 5.655% | 2.11% | 0.337% | -0.7267% | 3.173% | 2.857% | -2.606% | -1.436% | -3.811% | 4.131% | -7.534% | 2.322% | 24.09%
| Z-Score (Absolute) |  | 0.47 | 0.61 | 10.77 | 1.42 | 0.56 | 0.13 | 0.13 | 0.82 | 0.74 | 0.58 | 0.3 | 0.88 | 1.05 | 1.78 | 0.61 | 5.88
| Analytical results for Hg in OREAS 251c (Indicative Value 0.14 ppm) | [21 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.03 | 1.15 | 1.21 | [17 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [17 empty] | 2
2
| 1 | 1 | < 1 | 0.15 | 0.15 | [17 empty] | 1
| 1 | 2 | < 1 | 0.15 | 0.14 | [17 empty] | 1
| 1 | 3 | < 1 | 0.15 | 0.14 | [17 empty] | 16
| 1 | 4 | < 1 | 0.15 | 0.13 | [17 empty] | 0.1425
| 1 | 5 | < 1 | 0.15 | 0.14 | [17 empty] | 7
| 1 | 6 | < 1 | 0.14 | 0.12
| Mean |  |  | 0.14833 | 0.13667
| Median |  |  | 0.15 | 0.14
| Std Dev. |  |  | 0.0040825 | 0.010328
| Rel.Std.Dev. |  |  | 2.752% | 7.557%
| PDM3 |  |  | 4.094% | -4.094%
| Z-Score (Absolute) |  | 20.33 | 0 | 0.67
| Analytical results for Ho in OREAS 251c (Certified Value 0.72 ppm) | [21 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.73 | 0.71076 | 0.84113 | 0.8 | 0.7 | 0.72 | 0.6 | 0.7 | 0.8 | [11 empty] | 1
| 1 | 2 | 0.76 | 0.70661 | 0.87119 | 0.8 | 0.6 | 0.71 | 0.7 | 0.73 | 0.76 | [11 empty] | 24
| 1 | 3 | 0.7 | 0.72497 | 0.8081 | 0.8 | 0.6 | 0.71 | 0.7 | 0.71 | 0.78 | [11 empty] | 16
| 1 | 4 | 0.69 | 0.73106 | 0.80432 | 0.8 | 0.6 | 0.73 | 0.6 | 0.71 | 0.8 | [11 empty] | 0.72281
| 1 | 5 | 0.74 | 0.69844 | 0.82871 | 0.7 | 0.6 | 0.72 | 0.6 | 0.73 | 0.78 | [11 empty] | 35
| 1 | 6 | 0.72 | 0.6991 | 0.8072 | 0.8 | 0.6 | 0.71 | 0.6 | 0.7 | 0.76
| Mean |  | 0.72333 | 0.71182 | 0.82678 | 0.78333 | 0.61667 | 0.71667 | 0.63333 | 0.71333 | 0.78
| Median |  | 0.725 | 0.70868 | 0.81841 | 0.8 | 0.6 | 0.715 | 0.6 | 0.71 | 0.78
| Std Dev. |  | 0.02582 | 0.013508 | 0.026148 | 0.040825 | 0.040825 | 0.008165 | 0.05164 | 0.013663 | 0.017889
| Rel.Std.Dev. |  | 3.57% | 1.898% | 3.163% | 5.212% | 6.62% | 1.139% | 8.154% | 1.915% | 2.293%
| PDM3 |  | 0.07279% | -1.52% | 14.38% | 8.374% | -14.68% | -0.8495% | -12.38% | -1.311% | 7.913%
| Z-Score (Absolute) |  | 0.07 | 0.05 | 1.17 | 0.71 | 1.06 | 0 | 0.89 | 0.04 | 0.67
| Analytical results for In in OREAS 251c (Certified Value 0.087 ppm) | [21 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 0.1 | < 0.1 | 0.1 | 0.093 | 0.082 | 0.088 | 0.077 | 0.08 | 0.081459 | 0.102 | 0.08 | 0.09 | 0.1 | 0.08 | 0.088 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.08 | < 0.1 | 0.09 | 0.089 | 0.079 | 0.086 | 0.068 | 0.089 | 0.087445 | 0.109 | 0.09 | 0.08 | 0.09 | 0.08 | 0.081 | 0.05 |  |  |  |  | 25
| 1 | 3 | 0.08 | < 0.1 | 0.09 | 0.09 | 0.091 | 0.08 | 0.071 | 0.089 | 0.086858 | 0.106 | 0.08 | 0.09 | 0.09 | 0.08 | 0.085 | 0.05 |  |  |  |  | 16
| 1 | 4 | 0.09 | < 0.1 | 0.09 | 0.086 | 0.083 | 0.08 | 0.072 | 0.086 | 0.08468 | 0.105 | 0.08 | 0.1 | 0.09 | 0.08 | 0.092 | 0.05 |  |  |  |  | 0.087477
| 1 | 5 | 0.1 | < 0.1 | 0.1 | 0.092 | 0.08 | 0.086 | 0.077 | 0.083 | 0.080264 | 0.097 | 0.08 | 0.1 | 0.1 | 0.08 | 0.093 | 0.05 |  |  |  |  | 36
| 1 | 6 | 0.1 | < 0.1 | 0.11 | 0.084 | 0.085 | 0.085 | 0.075 | 0.084 | 0.082326 | 0.101 | 0.08 | 0.09 | 0.09 | 0.08 | 0.086 | 0.05
| Mean |  | 0.091667 |  | 0.096667 | 0.089 | 0.083333 | 0.084167 | 0.073333 | 0.085167 | 0.083839 | 0.10333 | 0.081667 | 0.091667 | 0.093333 | 0.08 | 0.0875 | 0.05
| Median |  | 0.095 |  | 0.095 | 0.0895 | 0.0825 | 0.0855 | 0.0735 | 0.085 | 0.083503 | 0.1035 | 0.08 | 0.09 | 0.09 | 0.08 | 0.087 | 0.05
| Std Dev. |  | 0.0098319 |  | 0.008165 | 0.0034641 | 0.0043205 | 0.0033714 | 0.0036148 | 0.0035449 | 0.0029515 | 0.0042269 | 0.0040825 | 0.0075277 | 0.005164 | 0 | 0.0045056 | 7.6012e-18
| Rel.Std.Dev. |  | 10.73% |  | 8.447% | 3.892% | 5.185% | 4.006% | 4.929% | 4.162% | 3.52% | 4.091% | 4.999% | 8.212% | 5.533% | 0% | 5.149% | 1.52e-14%
| PDM3 |  | 4.79% |  | 10.51% | 1.742% | -4.736% | -3.784% | -16.17% | -2.641% | -4.159% | 18.13% | -6.642% | 4.79% | 6.695% | -8.547% | 0.02678% | -42.84%
| Z-Score (Absolute) |  | 0.81 | 4.01 | 1.39 | 0.5 | 0.15 | 0.06 | 1.31 | 0.06 | 0.1 | 2.16 | 0.35 | 0.81 | 1 | 0.54 | 0.33 | 4.01
| Analytical results for K in OREAS 251c (Certified Value 1.87 wt.%) | [21 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.8534 | 1.9476 | 1.7595 | 2.09 | 1.97 | 1.85 | 1.88 | 1.84 | 1.84 | 1.87 | 1.8999 | 1.8698 | 2.08 | 1.86 | 2.0428 | 1.87 | 1.86 | 1.76 | 1.8 | 0.36485 | 1
| 1 | 2 | 1.8573 | 1.8947 | 1.741 | 2.01 | 1.97 | 1.85 | 1.89 | 1.85 | 1.85 | 1.91 | 1.8887 | 1.9217 | 2.03 | 1.31 | 1.9724 | 1.94 | 1.73 | 1.78 | 1.84 | 1.902 | #N/A
| 1 | 3 | 1.8704 | 1.9149 | 1.7585 | 2.06 | 1.95 | 1.84 | 1.92 | 1.84 | 1.86 | 1.87 | 1.8544 | 1.921 | 1.96 | 1.65 | 2.0149 | 1.91 | 1.91 | 1.78 | 1.83 | 1.9113 | 16
| 1 | 4 | 1.8602 | 1.8278 | 1.7474 | 1.98 | 1.96 | 1.81 | 1.91 | 1.79 | 1.84 | 1.88 | 1.8548 | 1.9187 | 1.95 | 1.72 | 2.0138 | 1.9 | 1.82 | 1.77 | 1.88 | 1.912 | 1.8743
| 1 | 5 | 1.8746 | 1.8655 | 1.7646 | 2.1 | 1.94 | 1.81 | 1.87 | 1.82 | 1.86 | 1.86 | 1.8923 | 1.872 | 1.99 | 1.57 | 1.9868 | 1.93 | 1.8 | 1.77 | 1.83 | 1.7186 | 37
| 1 | 6 | 1.857 | 1.8859 | 1.7562 | 2.12 | 1.98 | 1.82 | 1.89 | 1.81 | 1.88 | 1.9 | 1.8936 | 1.8375 | 2.04 | 2.01 | 1.9541 | 1.88 | 1.83 | 1.78 | 1.85 | 1.8859
| Mean |  | 1.8621 | 1.8894 | 1.7545 | 2.06 | 1.9617 | 1.83 | 1.8933 | 1.825 | 1.855 | 1.8817 | 1.8806 | 1.8901 | 2.0083 | 1.6867 | 1.9975 | 1.905 | 1.825 | 1.7733 | 1.8383 | 1.6158
| Median |  | 1.8588 | 1.8903 | 1.7573 | 2.075 | 1.965 | 1.83 | 1.89 | 1.83 | 1.855 | 1.875 | 1.8905 | 1.8954 | 2.01 | 1.685 | 2.0003 | 1.905 | 1.825 | 1.775 | 1.835 | 1.8939
| Std Dev. |  | 0.008408 | 0.04108 | 0.0087139 | 0.054772 | 0.01472 | 0.018974 | 0.018619 | 0.022583 | 0.015166 | 0.019408 | 0.020492 | 0.035434 | 0.050365 | 0.24188 | 0.032399 | 0.027386 | 0.060249 | 0.008165 | 0.026394 | 0.6173
| Rel.Std.Dev. |  | 0.4515% | 2.174% | 0.4967% | 2.659% | 0.7504% | 1.037% | 0.9834% | 1.237% | 0.8176% | 1.031% | 1.09% | 1.875% | 2.508% | 14.34% | 1.622% | 1.438% | 3.301% | 0.4604% | 1.436% | 38.21%
| PDM3 |  | -0.6468% | 0.808% | -6.389% | 9.909% | 4.663% | -2.362% | 1.017% | -2.629% | -1.028% | 0.3945% | 0.3389% | 0.8453% | 7.153% | -10.01% | 6.573% | 1.639% | -2.629% | -5.386% | -1.918% | -13.79%
| Z-Score (Absolute) |  | 0.14 | 0.28 | 1.8 | 2.9 | 1.39 | 0.64 | 0.34 | 0.71 | 0.25 | 0.16 | 0.14 | 0.29 | 2.1 | 2.84 | 1.94 | 0.52 | 0.71 | 1.51 | 0.51 | 3.93
| Analytical results for La in OREAS 251c (Certified Value 45.3 ppm) | [21 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  |  | 1
2
| 1 | 1 | 46.37 | 44.7852 | 51.5524 | 49.3 | 44.5 | 48.5 | 44.1 | 46.4 | 43.5 | 47.3 | 36.7 | 47 | 47.66 | 41.7 | 43.6 | 39.99 | 48.2 | 9.249 |  |  | 1
| 1 | 2 | 46.72 | 44.3144 | 51.5563 | 47.8 | 42.7 | 47.2 | 43.4 | 46.8 | 40.6 | 46.6 | 37.8 | 41.1 | 46.98 | 43.3 | 43.3 | 39.75 | 47.7 | 43.73 |  |  | 26
| 1 | 3 | 45.09 | 45.2835 | 51.5326 | 49.6 | 42.6 | 48.1 | 46.7 | 44 | 40.4 | 46.2 | 37.3 | 47.2 | 47.67 | 42.3 | 43.2 | 39.42 | 47.6 | 44.98 |  |  | 16
| 1 | 4 | 45.41 | 42.3913 | 51.3922 | 48.2 | 42.1 | 46.2 | 45.4 | 43.5 | 39.5 | 47 | 37.8 | 45.7 | 46.91 | 44.8 | 43.2 | 40.83 | 48.1 | 43.563 |  |  | 45.3112
| 1 | 5 | 46.22 | 43.1262 | 51.3952 | 48.4 | 44.3 | 46.6 | 38.6 | 42.9 | 40.8 | 45.1 | 36.8 | 46.9 | 46.85 | 42.7 | 42.6 | 41.58 | 48.2 | 44.519 |  |  | 38
| 1 | 6 | 46.32 | 43.4591 | 51.6436 | 49.5 | 43.5 | 47.5 | 44.3 | 44.4 | 43.8 | 47 | 36.5 | 46.3 | 47.05 | 42.7 | 42.5 | 42.05 | 47.3 | 43.96
| Mean |  | 46.0217 | 43.8933 | 51.5121 | 48.8 | 43.2833 | 47.35 | 43.75 | 44.6667 | 41.4333 | 46.5333 | 37.15 | 45.7 | 47.1867 | 42.9167 | 43.0667 | 40.6033 | 47.85 | 38.3335
| Median |  | 46.27 | 43.8867 | 51.5425 | 48.85 | 43.1 | 47.35 | 44.2 | 44.2 | 40.7 | 46.8 | 37.05 | 46.6 | 47.015 | 42.7 | 43.2 | 40.41 | 47.9 | 43.845
| Std Dev. |  | 0.6292 | 1.0898 | 0.09933 | 0.76158 | 0.97656 | 0.8735 | 2.7747 | 1.5845 | 1.7761 | 0.79917 | 0.56833 | 2.3195 | 0.37655 | 1.0629 | 0.4274 | 1.0586 | 0.37283 | 14.2581
| Rel.Std.Dev. |  | 1.367% | 2.483% | 0.1928% | 1.561% | 2.256% | 1.845% | 6.342% | 3.547% | 4.287% | 1.717% | 1.53% | 5.075% | 0.798% | 2.477% | 0.9924% | 2.607% | 0.7792% | 37.19%
| PDM3 |  | 1.568% | -3.129% | 13.68% | 7.7% | -4.475% | 4.5% | -3.446% | -1.423% | -8.558% | 2.697% | -18.01% | 0.858% | 4.139% | -5.285% | -4.954% | -10.39% | 5.603% | -15.4%
| Z-Score (Absolute) |  | 0.46 | 0.1 | 1.91 | 1.2 | 0.26 | 0.81 | 0.14 | 0.1 | 0.75 | 0.6 | 1.89 | 0.38 | 0.77 | 0.36 | 0.32 | 0.97 | 0.94 | 1.57
| Analytical results for Li in OREAS 251c (Certified Value 30.6 ppm) | [21 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 30.6 | 29.3175 | 22.4493 | 31 | 32 | 31.2 | 31.4 | 32.6 | 31.3 | 32.1 | 30.8708 | 28 | 30.1 | 28.1 | 31.6 | 29.1 | 26.4 | 32.48 | 30 | 19.21 | 1
| 1 | 2 | 30.7 | 29.242 | 22.2524 | 30 | 33 | 31.3 | 31.6 | 32.8 | 31.4 | 33 | 30.4689 | 28 | 31.1 | 28.3 | 30.9 | 30.4 | 26.4 | 33.16 | 30 | 29.173 | #N/A
| 1 | 3 | 30.4 | 28.904 | 22.3012 | 31 | 33 | 30.8 | 31.9 | 32 | 32.1 | 32.3 | 31.1299 | 28 | 31.1 | 26.9 | 31.2 | 30.1 | 28.1 | 33.73 | 30 | 29.23 | 16
| 1 | 4 | 29.4 | 28.1075 | 22.3525 | 31 | 31 | 30.7 | 31.9 | 31.3 | 31.2 | 32.5 | 31.4495 | 28 | 31.1 | 26.8 | 31.3 | 29.8 | 26.7 | 33.43 | 30 | 28.988 | 30.5942
| 1 | 5 | 30.3 | 28.1582 | 22.3941 | 31 | 32 | 30.6 | 31.6 | 31.5 | 31.5 | 32.6 | 30.1376 | 28 | 30.1 | 26.9 | 31.4 | 30.1 | 27.1 | 33.23 | 30 | 28.763 | 39
| 1 | 6 | 30.6 | 29.1215 | 22.2563 | 31 | 32 | 30.6 | 31.6 | 31.6 | 31.8 | 32.7 | 30.3839 | 27 | 31.1 | 26.4 | 30.6 | 29.6 | 25.9 | 33.02 | 30 | 28.94
| Mean |  | 30.3333 | 28.8085 | 22.3343 | 30.8333 | 32.1667 | 30.8667 | 31.6667 | 31.9667 | 31.55 | 32.5333 | 30.7401 | 27.8333 | 30.7667 | 27.2333 | 31.1667 | 29.85 | 26.7667 | 33.175 | 30 | 27.384
| Median |  | 30.5 | 29.0127 | 22.3268 | 31 | 32 | 30.75 | 31.6 | 31.8 | 31.45 | 32.55 | 30.6699 | 28 | 31.1 | 26.9 | 31.25 | 29.95 | 26.55 | 33.195 | 30 | 28.964
| Std Dev. |  | 0.48028 | 0.54191 | 0.078772 | 0.40825 | 0.75277 | 0.30768 | 0.19664 | 0.61536 | 0.33912 | 0.31411 | 0.49714 | 0.40825 | 0.5164 | 0.77374 | 0.36148 | 0.45935 | 0.76333 | 0.42032 | 0 | 4.0079
| Rel.Std.Dev. |  | 1.583% | 1.881% | 0.3527% | 1.324% | 2.34% | 0.9968% | 0.621% | 1.925% | 1.075% | 0.9655% | 1.617% | 1.467% | 1.678% | 2.841% | 1.16% | 1.539% | 2.852% | 1.267% | 0% | 14.64%
| PDM3 |  | -0.8527% | -5.837% | -27% | 0.7816% | 5.14% | 0.8906% | 3.505% | 4.486% | 3.124% | 6.338% | 0.4769% | -9.024% | 0.5637% | -10.99% | 1.871% | -2.433% | -12.51% | 8.436% | -1.942% | -10.49%
| Z-Score (Absolute) |  | 0.27 | 1.23 | 5.34 | 0.05 | 0.9 | 0.07 | 0.58 | 0.77 | 0.51 | 1.13 | 0.01 | 1.85 | 0.01 | 2.23 | 0.26 | 0.57 | 2.53 | 1.54 | 0.48 | 2.14
| Analytical results for Lu in OREAS 251c (Certified Value 0.26 ppm) | [21 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 0.25 | 0.27539 | 0.28232 | 0.3 | 0.2 | 0.27 | 0.2 | 0.23 | 0.3 | [11 empty] | 1
| 1 | 2 | 0.26 | 0.24102 | 0.2897 | 0.3 | 0.2 | 0.28 | 0.2 | 0.24 | 0.28 | [11 empty] | 5
| 1 | 3 | 0.23 | 0.24915 | 0.29528 | 0.3 | 0.2 | 0.27 | 0.2 | 0.25 | 0.28 | [11 empty] | 16
| 1 | 4 | 0.27 | 0.24027 | 0.27008 | 0.3 | 0.2 | 0.27 | 0.2 | 0.23 | 0.3 | [11 empty] | 0.26326
| 1 | 5 | 0.24 | 0.25297 | 0.27278 | 0.3 | 0.2 | 0.29 | 0.2 | 0.23 | 0.28 | [11 empty] | 40
| 1 | 6 | 0.25 | 0.25437 | 0.27404 | 0.3 | 0.2 | 0.27 | 0.2 | 0.23 | 0.28
| Mean |  | 0.25 | 0.2522 | 0.2807 | 0.3 | 0.2 | 0.275 | 0.2 | 0.235 | 0.28667
| Median |  | 0.25 | 0.25106 | 0.27818 | 0.3 | 0.2 | 0.27 | 0.2 | 0.23 | 0.28
| Std Dev. |  | 0.014142 | 0.012804 | 0.010159 | 0 | 3.0405e-17 | 0.0083666 | 3.0405e-17 | 0.0083666 | 0.010328
| Rel.Std.Dev. |  | 5.657% | 5.077% | 3.619% | 0% | 1.52e-14% | 3.042% | 1.52e-14% | 3.56% | 3.603%
| PDM3 |  | -5.037% | -4.203% | 6.624% | 13.96% | -24.03% | 4.459% | -24.03% | -10.73% | 8.891%
| Z-Score (Absolute) |  | 0.6 | 0.5 | 0.75 | NA | NA | 0.5 | NA | 1.26 | 1.01
| Analytical results for Mg in OREAS 251c (Certified Value 1.37 wt.%) | [21 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.3432 | 1.3687 | 1.4179 | 1.56 | 1.39 | 1.33 | 1.34 | 1.38 | 1.28 | 1.27 | 1.4021 | 1.3477 | 1.35 | 1.41 | 1.4247 | 1.32 | 1.31 | 1.42 | 1.51 | 3.6522 | 1
| 1 | 2 | 1.3561 | 1.3312 | 1.4166 | 1.57 | 1.39 | 1.33 | 1.35 | 1.38 | 1.28 | 1.27 | 1.3865 | 1.4001 | 1.39 | 1.4 | 1.3933 | 1.38 | 1.21 | 1.42 | 1.47 | 1.445 | #N/A
| 1 | 3 | 1.3501 | 1.3335 | 1.4222 | 1.56 | 1.37 | 1.32 | 1.37 | 1.38 | 1.3 | 1.27 | 1.3771 | 1.392 | 1.32 | 1.41 | 1.4078 | 1.36 | 1.32 | 1.42 | 1.46 | 1.4643 | 16
| 1 | 4 | 1.3552 | 1.2931 | 1.4163 | 1.55 | 1.38 | 1.3 | 1.36 | 1.34 | 1.29 | 1.26 | 1.3632 | 1.4035 | 1.36 | 1.37 | 1.3992 | 1.35 | 1.28 | 1.42 | 1.52 | 1.4517 | 1.3663
| 1 | 5 | 1.3596 | 1.3255 | 1.4116 | 1.61 | 1.36 | 1.3 | 1.32 | 1.36 | 1.28 | 1.25 | 1.4224 | 1.3563 | 1.38 | 1.39 | 1.3662 | 1.36 | 1.25 | 1.42 | 1.46 | 1.4439 | 41
| 1 | 6 | 1.3504 | 1.3268 | 1.4117 | 1.58 | 1.37 | 1.31 | 1.35 | 1.36 | 1.32 | 1.29 | 1.4018 | 1.3375 | 1.35 | 1.37 | 1.3696 | 1.35 | 1.28 | 1.43 | 1.48 | 1.4556
| Mean |  | 1.3524 | 1.3298 | 1.4161 | 1.5717 | 1.3767 | 1.315 | 1.3483 | 1.3667 | 1.2917 | 1.2683 | 1.3922 | 1.3728 | 1.3583 | 1.3917 | 1.3935 | 1.3533 | 1.275 | 1.4217 | 1.4833 | 1.8188
| Median |  | 1.3528 | 1.329 | 1.4165 | 1.565 | 1.375 | 1.315 | 1.35 | 1.37 | 1.285 | 1.27 | 1.3942 | 1.3741 | 1.355 | 1.395 | 1.3962 | 1.355 | 1.28 | 1.42 | 1.475 | 1.4537
| Std Dev. |  | 0.0057843 | 0.024089 | 0.0040297 | 0.02137 | 0.012111 | 0.013784 | 0.017224 | 0.01633 | 0.016021 | 0.013292 | 0.021 | 0.028999 | 0.024833 | 0.018348 | 0.022481 | 0.019664 | 0.040373 | 0.0040825 | 0.02582 | 0.89822
| Rel.Std.Dev. |  | 0.4277% | 1.811% | 0.2846% | 1.36% | 0.8797% | 1.048% | 1.277% | 1.195% | 1.24% | 1.048% | 1.508% | 2.112% | 1.828% | 1.318% | 1.613% | 1.453% | 3.167% | 0.2872% | 1.741% | 49.39%
| PDM3 |  | -1.012% | -2.668% | 3.645% | 15.03% | 0.7619% | -3.752% | -1.312% | 0.02993% | -5.46% | -7.167% | 1.898% | 0.4825% | -0.58% | 1.86% | 1.991% | -0.946% | -6.679% | 4.056% | 8.569% | 33.12%
| Z-Score (Absolute) |  | 0.37 | 0.85 | 0.98 | 4.28 | 0.15 | 1.16 | 0.45 | 0.07 | 1.65 | 2.15 | 0.48 | 0.07 | 0.24 | 0.46 | 0.5 | 0.35 | 2.01 | 1.1 | 2.41 | 9.51
| Analytical results for Mn in OREAS 251c (Certified Value 0.041 wt.%) | [21 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.0407 | 0.042256 | 0.044012 | 0.0441 | 0.0436 | 0.0398 | 0.0394 | 0.0425 | 0.0402 | 0.0403 | 0.043036 | 0.0377 | 0.0412 | 0.0433 | 0.0443 | 0.0417 | 0.0396 | 0.04076 | 0.0438 | 0.12666 | 1
| 1 | 2 | 0.0408 | 0.041342 | 0.043717 | 0.0434 | 0.0416 | 0.04 | 0.0399 | 0.0428 | 0.04 | 0.0407 | 0.041833 | 0.0385 | 0.0417 | 0.0432 | 0.0431 | 0.0434 | 0.0361 | 0.04189 | 0.0428 | 0.040713 | 28
| 1 | 3 | 0.0411 | 0.041182 | 0.044315 | 0.0438 | 0.0413 | 0.0393 | 0.0404 | 0.0422 | 0.0408 | 0.0403 | 0.042835 | 0.0386 | 0.0409 | 0.0431 | 0.0444 | 0.043 | 0.0405 | 0.04207 | 0.0428 | 0.039789 | 16
| 1 | 4 | 0.0411 | 0.039909 | 0.044266 | 0.0438 | 0.0411 | 0.0388 | 0.0399 | 0.0417 | 0.0403 | 0.0401 | 0.042923 | 0.0385 | 0.0414 | 0.042 | 0.0437 | 0.0424 | 0.0385 | 0.04175 | 0.0442 | 0.040514 | 0.04147
| 1 | 5 | 0.0407 | 0.040781 | 0.044105 | 0.0447 | 0.0428 | 0.0389 | 0.0399 | 0.0421 | 0.0406 | 0.0398 | 0.042118 | 0.0379 | 0.0415 | 0.0429 | 0.0429 | 0.0436 | 0.0378 | 0.04166 | 0.0432 | 0.039289 | 42
| 1 | 6 | 0.0407 | 0.041068 | 0.043661 | 0.0441 | 0.0438 | 0.0389 | 0.0396 | 0.0424 | 0.0414 | 0.0407 | 0.041499 | 0.0371 | 0.042 | 0.0423 | 0.0431 | 0.0428 | 0.0389 | 0.04201 | 0.0436 | 0.040125
| Mean |  | 0.04085 | 0.04109 | 0.044013 | 0.043983 | 0.042367 | 0.039283 | 0.03985 | 0.042283 | 0.04055 | 0.040317 | 0.042374 | 0.03805 | 0.04145 | 0.0428 | 0.043583 | 0.042817 | 0.038567 | 0.04169 | 0.0434 | 0.054515
| Median |  | 0.04075 | 0.041125 | 0.044059 | 0.04395 | 0.0422 | 0.0391 | 0.0399 | 0.0423 | 0.04045 | 0.0403 | 0.042477 | 0.0382 | 0.04145 | 0.043 | 0.0434 | 0.0429 | 0.0387 | 0.04182 | 0.0434 | 0.040319
| Std Dev. |  | 0.00019748 | 0.00076469 | 0.00027373 | 0.00043551 | 0.0011911 | 0.00051153 | 0.00033912 | 0.00037639 | 0.00050498 | 0.00034881 | 0.00064454 | 0.00059245 | 0.00038341 | 0.00052915 | 0.00065243 | 0.00069402 | 0.0015227 | 0.00048087 | 0.00056569 | 0.035347
| Rel.Std.Dev. |  | 0.4834% | 1.861% | 0.6219% | 0.9902% | 2.811% | 1.302% | 0.851% | 0.8902% | 1.245% | 0.8652% | 1.521% | 1.557% | 0.925% | 1.236% | 1.497% | 1.621% | 3.948% | 1.153% | 1.303% | 64.84%
| PDM3 |  | -1.495% | -0.9176% | 6.131% | 6.06% | 2.162% | -5.273% | -3.907% | 1.961% | -2.219% | -2.781% | 2.18% | -8.247% | -0.04853% | 3.207% | 5.096% | 3.247% | -7.001% | 0.5302% | 4.654% | 31.46%
| Z-Score (Absolute) |  | 0.54 | 0.42 | 0.96 | 0.94 | 0.18 | 1.28 | 1.01 | 0.14 | 0.68 | 0.79 | 0.18 | 1.86 | 0.25 | 0.38 | 0.76 | 0.39 | 1.62 | 0.14 | 0.67 | 5.93
| Analytical results for Mo in OREAS 251c (Certified Value 2.48 ppm) | [21 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 2.7 | 2.3014 | < 1 | 1.9 | 2.62 | 2.44 | 2.74 | 2.45 | 2.45 | 2.3975 | 2.4 | 2.5 | 2.14 | 2.6 | 2.38 | 1.8 | < 1 | 2.5 | 8.414 |  | 1
| 1 | 2 | 2.6 | 2.3833 | < 1 | 1.9 | 2.55 | 2.39 | 2.88 | 2.21 | 2.52 | 2.2265 | 2.3 | 2.6 | 2.25 | 2.7 | 2.4 | 1.8 | < 1 | 2.5 | 3.012 |  | 29
| 1 | 3 | 2.6 | 2.3502 | < 1 | 1.9 | 2.6 | 2.58 | 2.57 | 2.34 | 2.52 | 2.3359 | 2.3 | 2.5 | 2.19 | 2.7 | 2.53 | 2.1 | < 1 | 2.5 | 3.181 |  | 16
| 1 | 4 | 2.6 | 2.2564 | < 1 | 2.1 | 2.46 | 2.53 | 2.5 | 2.26 | 2.55 | 2.3716 | 2.3 | 2.5 | 2.17 | 2.6 | 2.51 | 1.8 | < 1 | 2.5 | 2.934 |  | 2.4831
| 1 | 5 | 2.5 | 2.3182 | < 1 | 2.2 | 2.56 | 2.46 | 2.54 | 2.35 | 2.51 | 2.3253 | 2.3 | 2.6 | 2.1 | 2.5 | 2.4 | 1.8 | < 1 | 2.5 | 2.986 |  | 43
| 1 | 6 | 2.5 | 2.3822 | < 1 | 1.9 | 2.47 | 2.46 | 2.69 | 2.4 | 2.6 | 2.2451 | 2.2 | 2.6 | 2.05 | 2.7 | 2.38 | 2 | < 1 | 2.5 | 2.722
| Mean |  | 2.5833 | 2.332 |  | 1.9833 | 2.5433 | 2.4767 | 2.6533 | 2.335 | 2.525 | 2.317 | 2.3 | 2.55 | 2.15 | 2.6333 | 2.4333 | 1.8833 |  | 2.5 | 3.8748
| Median |  | 2.6 | 2.3342 |  | 1.9 | 2.555 | 2.46 | 2.63 | 2.345 | 2.52 | 2.3306 | 2.3 | 2.55 | 2.155 | 2.65 | 2.4 | 1.8 |  | 2.5 | 2.999
| Std Dev. |  | 0.075277 | 0.049634 |  | 0.13292 | 0.065929 | 0.067725 | 0.1439 | 0.088261 | 0.049295 | 0.068178 | 0.063246 | 0.054772 | 0.070143 | 0.08165 | 0.06802 | 0.13292 |  | 0 | 2.2286
| Rel.Std.Dev. |  | 2.914% | 2.128% |  | 6.702% | 2.592% | 2.735% | 5.423% | 3.78% | 1.952% | 2.943% | 2.75% | 2.148% | 3.262% | 3.101% | 2.795% | 7.057% |  | 0% | 57.52%
| PDM3 |  | 4.039% | -6.085% |  | -20.13% | 2.428% | -0.2572% | 6.858% | -5.963% | 1.689% | -6.688% | -7.372% | 2.696% | -13.41% | 6.052% | -2.002% | -24.15% |  | 0.6825% | 56.05%
| Z-Score (Absolute) |  | 0.76 | 0.51 | 9.78 | 2.28 | 0.56 | 0.22 | 1.11 | 0.5 | 0.46 | 0.59 | 0.67 | 0.59 | 1.43 | 1.01 | 0 | 2.78 | 9.78 | 0.34 | 7.29
| Analytical results for Na in OREAS 251c (Certified Value 0.957 wt.%) | [21 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.9369 | 0.94954 | 0.89191 | 1.01 | 1.01 | 0.93 | 0.93 | 0.98 | 0.92 | 1.05 | 0.94065 | 0.8195 | 0.961 | 1.013 | 1.0162 | 0.941 | 0.924 | 0.92 | 0.96 | 1.7392 | 1
| 1 | 2 | 0.9404 | 0.94443 | 0.8918 | 1.02 | 0.99 | 0.93 | 0.94 | 0.98 | 0.92 | 1.08 | 0.93477 | 0.8631 | 0.983 | 1.006 | 0.9986 | 0.981 | 0.856 | 1 | 0.96 | 0.96445 | #N/A
| 1 | 3 | 0.9417 | 0.93929 | 0.89702 | 1.02 | 0.98 | 0.92 | 0.95 | 0.97 | 0.92 | 1.06 | 0.92264 | 0.8547 | 0.989 | 0.999 | 1.0138 | 0.955 | 0.942 | 1.05 | 0.94 | 0.97241 | 16
| 1 | 4 | 0.942 | 0.92707 | 0.89096 | 1.04 | 0.99 | 0.91 | 0.95 | 0.95 | 0.92 | 1.06 | 0.91775 | 0.8595 | 0.96 | 0.985 | 1.0122 | 0.941 | 0.904 | 0.98 | 0.97 | 0.97991 | 0.95667
| 1 | 5 | 0.9452 | 0.92911 | 0.89257 | 1.03 | 0.98 | 0.91 | 0.95 | 0.97 | 0.92 | 1.06 | 0.94495 | 0.8228 | 0.97 | 1.005 | 0.9902 | 0.956 | 0.882 | 0.95 | 0.97 | 0.97207 | 44
| 1 | 6 | 0.9378 | 0.93956 | 0.8934 | 1.1 | 0.99 | 0.92 | 0.94 | 0.96 | 0.94 | 1.07 | 0.94885 | 0.8089 | 0.996 | 1.016 | 0.9884 | 0.946 | 0.904 | 0.94 | 0.96 | 0.97125
| Mean |  | 0.94067 | 0.93817 | 0.89294 | 1.0367 | 0.99 | 0.92 | 0.94333 | 0.96833 | 0.92333 | 1.0633 | 0.93493 | 0.83808 | 0.9765 | 1.004 | 1.0032 | 0.95333 | 0.902 | 0.97333 | 0.96 | 1.0999
| Median |  | 0.94105 | 0.93943 | 0.89224 | 1.025 | 0.99 | 0.92 | 0.945 | 0.97 | 0.92 | 1.06 | 0.93771 | 0.83875 | 0.9765 | 1.0055 | 1.0054 | 0.9505 | 0.904 | 0.965 | 0.96 | 0.97224
| Std Dev. |  | 0.003029 | 0.0086818 | 0.0021591 | 0.03266 | 0.010954 | 0.0089443 | 0.008165 | 0.01169 | 0.008165 | 0.010328 | 0.012435 | 0.023627 | 0.015057 | 0.0111 | 0.012422 | 0.015055 | 0.030358 | 0.047188 | 0.010954 | 0.31326
| Rel.Std.Dev. |  | 0.322% | 0.9254% | 0.2418% | 3.15% | 1.107% | 0.9722% | 0.8655% | 1.207% | 0.8843% | 0.9713% | 1.33% | 2.819% | 1.542% | 1.106% | 1.238% | 1.579% | 3.366% | 4.848% | 1.141% | 28.48%
| PDM3 |  | -1.673% | -1.935% | -6.662% | 8.362% | 3.484% | -3.833% | -1.394% | 1.219% | -3.485% | 11.15% | -2.272% | -12.4% | 2.072% | 4.947% | 4.867% | -0.3492% | -5.715% | 1.741% | 0.3477% | 14.97%
| Z-Score (Absolute) |  | 0.32 | 0.37 | 1.29 | 1.62 | 0.67 | 0.74 | 0.27 | 0.24 | 0.67 | 2.16 | 0.44 | 2.4 | 0.4 | 0.96 | 0.94 | 0.07 | 1.11 | 0.34 | 0.07 | 2.9
| Analytical results for Nb in OREAS 251c (Certified Value 26.3 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 27.85 | 27.5385 | 24.2 | 25 | 29.8 | 26.1 | 29.2 | 26.3 | 25.9 | 23.1851 | NA | 24.7 | 25.29 | 28.45 | 24.76 | 28.6 | 27.5 |  |  |  | 1
| 1 | 2 | 28.13 | 27.1354 | 25 | 23 | 28.6 | 25.4 | 29.3 | 24.4 | 28 | 23.2363 | NA | 25.3 | 24.7 | 27.13 | 26.03 | 29.8 | 28.5 |  |  |  | 30
| 1 | 3 | 26.64 | 26.7632 | 24.5 | 25 | 29.9 | 26.9 | 28.4 | 26.4 | 26.7 | 22.881 | NA | 22.7 | 23.65 | 27.3 | 25.73 | 28.8 | 28.5 |  |  |  | 16
| 1 | 4 | 27.46 | 25.7939 | 24.4 | 24 | 29.2 | 26.7 | 27.7 | 24.6 | 25.8 | 25.8478 | NA | 24.5 | 23.4 | 27.71 | 25.54 | 29.4 | 28.5 |  |  |  | 26.3304
| 1 | 5 | 27.25 | 26.037 | 24.4 | 24 | 29.3 | 26.5 | 28 | 25.7 | 26.6 | 25.6141 | NA | 22.8 | 23.88 | 26.73 | 25.82 | 28.6 | 29 |  |  |  | 45
| 1 | 6 | 27.43 | 27.0544 | 24.9 | 25 | 29 | 26.3 | 28.6 | 25.7 | 26.4 | 22.3133 | NA | 24.2 | 22.79 | 27.28 | 24.97 | 28.7 | 27.5
| Mean |  | 27.46 | 26.7204 | 24.5667 | 24.3333 | 29.3 | 26.3167 | 28.5333 | 25.5167 | 26.5667 | 23.8463 |  | 24.0333 | 23.9517 | 27.4333 | 25.475 | 28.9833 | 28.25
| Median |  | 27.445 | 26.9088 | 24.45 | 24.5 | 29.25 | 26.4 | 28.5 | 25.7 | 26.5 | 23.2107 |  | 24.35 | 23.765 | 27.29 | 25.635 | 28.75 | 28.5
| Std Dev. |  | 0.5135 | 0.67533 | 0.31411 | 0.8165 | 0.4899 | 0.53072 | 0.6377 | 0.84242 | 0.79162 | 1.4982 |  | 1.0577 | 0.90566 | 0.5896 | 0.5025 | 0.49967 | 0.61237
| Rel.Std.Dev. |  | 1.87% | 2.527% | 1.279% | 3.355% | 1.672% | 2.017% | 2.235% | 3.301% | 2.98% | 6.283% |  | 4.401% | 3.781% | 2.149% | 1.973% | 1.724% | 2.168%
| PDM3 |  | 4.29% | 1.481% | -6.699% | -7.585% | 11.28% | -0.05222% | 8.366% | -3.091% | 0.8972% | -9.434% |  | -8.724% | -9.034% | 4.189% | -3.249% | 10.08% | 7.29%
| Z-Score (Absolute) |  | 0.37 | 0.1 | 0.69 | 0.77 | 1.05 | 0.05 | 0.77 | 0.34 | 0.05 | 0.95 | NA | 0.88 | 0.91 | 0.36 | 0.35 | 0.93 | 0.66
| Analytical results for Nd in OREAS 251c (Certified Value 38.4 ppm) | [21 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 1
2
| 1 | 1 | 38.7 | 39.6044 | 46.4226 | 39.4 | 37.4 | 38.59 | 36.1 | 40.8 | 40.5 | [11 empty] | 1
| 1 | 2 | 38.73 | 37.8552 | 46.6111 | 38.4 | 37.1 | 39.27 | 37.1 | 40.9 | 40.2 | [11 empty] | 6
| 1 | 3 | 38.28 | 38.165 | 46.5567 | 39.5 | 34.5 | 38.91 | 37.1 | 39.9 | 40.2 | [11 empty] | 16
| 1 | 4 | 37.76 | 36.5123 | 46.5251 | 38.8 | 35.1 | 39.68 | 38.3 | 40.1 | 40.6 | [11 empty] | 38.4355
| 1 | 5 | 38.63 | 37.4457 | 46.6142 | 39.6 | 35.4 | 38.94 | 37.2 | 39.5 | 40.2 | [11 empty] | 46
| 1 | 6 | 37.66 | 37.9301 | 46.6131 | 39.8 | 33.4 | 40.04 | 36.4 | 38.5 | 40.2
| Mean |  | 38.2933 | 37.9188 | 46.5571 | 39.25 | 35.4833 | 39.2383 | 37.0333 | 39.95 | 40.3167
| Median |  | 38.455 | 37.8927 | 46.5839 | 39.45 | 35.25 | 39.105 | 37.1 | 40 | 40.2
| Std Dev. |  | 0.48065 | 1.0108 | 0.075438 | 0.53572 | 1.5329 | 0.53961 | 0.76333 | 0.88938 | 0.18348
| Rel.Std.Dev. |  | 1.255% | 2.666% | 0.162% | 1.365% | 4.32% | 1.375% | 2.061% | 2.226% | 0.4551%
| PDM3 |  | -0.3698% | -1.344% | 21.13% | 2.119% | -7.681% | 2.089% | -3.648% | 3.94% | 4.894%
| Z-Score (Absolute) |  | 0.59 | 0.83 | 4.58 | 0.01 | 2.35 | 0 | 1.38 | 0.45 | 0.67
| Analytical results for Ni in OREAS 251c (Certified Value 72 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 73.9 | 68.6187 | 75.3913 | 77 | 75 | 72.9 | 71.6 | 74.2 | 73.5 | 73 | 72.3766 | 61.8 | 72 | 71.6 | 74.6 | 71.1 | 70.5 | 69.41 | 88 | 72.561 | 1
| 1 | 2 | 74.4 | 68.0987 | 75.8883 | 75 | 76 | 72.2 | 68.2 | 74.3 | 67.1 | 75.2 | 69.7864 | 61.6 | 72 | 69.1 | 73.2 | 74.4 | 70 | 71.21 | 86 | 72.278 | 15
| 1 | 3 | 75.1 | 68.3034 | 77.3443 | 75 | 73 | 73.4 | 72.2 | 71.8 | 70 | 72.7 | 72.6197 | 64 | 71 | 71.5 | 73.6 | 74.2 | 72.2 | 69.2 | 84 | 71.912 | 16
| 1 | 4 | 73.9 | 65.8789 | 77.2981 | 75 | 75 | 72.8 | 71.2 | 68.8 | 67.5 | 72.9 | 72.0693 | 62.4 | 71 | 70.6 | 73.4 | 72.4 | 69.2 | 71.05 | 88 | 72.64 | 72.0797
| 1 | 5 | 73.7 | 66.1646 | 77.4599 | 76 | 73 | 71.6 | 70.7 | 69.2 | 69.6 | 72.5 | 70.1903 | 61.6 | 71 | 70.9 | 72.4 | 73.1 | 69.6 | 69.73 | 84 | 71.684 | 47
| 1 | 6 | 74 | 68.2461 | 76.1194 | 76 | 75 | 70.5 | 69.8 | 72 | 70.1 | 74.1 | 71.4782 | 59.1 | 71 | 70.2 | 72.4 | 72.6 | 68.8 | 70.49 | 86 | 70.07
| Mean |  | 74.1667 | 67.5517 | 76.5835 | 75.6667 | 74.5 | 72.2333 | 70.6167 | 71.7167 | 69.6333 | 73.4 | 71.4201 | 61.75 | 71.3333 | 70.65 | 73.2667 | 72.9667 | 70.05 | 70.1817 | 86 | 71.8575
| Median |  | 73.95 | 68.1724 | 76.7087 | 75.5 | 75 | 72.5 | 70.95 | 71.9 | 69.8 | 72.95 | 71.7738 | 61.7 | 71 | 70.75 | 73.3 | 72.85 | 69.8 | 70.11 | 86 | 72.095
| Std Dev. |  | 0.51251 | 1.2006 | 0.89191 | 0.8165 | 1.2247 | 1.052 | 1.4372 | 2.3566 | 2.2923 | 1.0431 | 1.1798 | 1.5846 | 0.5164 | 0.92682 | 0.82624 | 1.2275 | 1.2095 | 0.8567 | 1.7889 | 0.94998
| Rel.Std.Dev. |  | 0.691% | 1.777% | 1.165% | 1.079% | 1.644% | 1.456% | 2.035% | 3.286% | 3.292% | 1.421% | 1.652% | 2.566% | 0.7239% | 1.312% | 1.128% | 1.682% | 1.727% | 1.221% | 2.08% | 1.322%
| PDM3 |  | 2.895% | -6.282% | 6.248% | 4.976% | 3.358% | 0.2131% | -2.03% | -0.5036% | -3.394% | 1.832% | -0.9151% | -14.33% | -1.035% | -1.983% | 1.647% | 1.231% | -2.816% | -2.633% | 19.31% | -0.3083%
| Z-Score (Absolute) |  | 1 | 1.77 | 2.01 | 1.63 | 1.14 | 0.19 | 0.49 | 0.03 | 0.9 | 0.68 | 0.15 | 4.21 | 0.19 | 0.48 | 0.62 | 0.49 | 0.73 | 0.67 | 5.96 | 0.03
| Analytical results for P in OREAS 251c (Certified Value 0.092 wt.%) | [21 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.088 | 0.093645 | 0.085 | 0.0943 | 0.089 | 0.09 | 0.091 | 0.091 | 0.093 | 0.094891 | 0.0911 | 0.092 | 0.095 | 0.0923 | 0.093 | 0.0921 | 0.09 | 0.09 | 0.040077 |  | 1
| 1 | 2 | 0.0886 | 0.091113 | 0.082 | 0.0934 | 0.088 | 0.09 | 0.091 | 0.091 | 0.094 | 0.09317 | 0.0938 | 0.093 | 0.091 | 0.0936 | 0.097 | 0.0884 | 0.09 | 0.085 | 0.093603 |  | #N/A
| 1 | 3 | 0.089 | 0.091852 | 0.077 | 0.0932 | 0.088 | 0.091 | 0.091 | 0.092 | 0.093 | 0.091325 | 0.0937 | 0.093 | 0.095 | 0.0902 | 0.096 | 0.0963 | 0.09 | 0.085 | 0.093595 |  | 16
| 1 | 4 | 0.0882 | 0.089413 | 0.087 | 0.0942 | 0.087 | 0.092 | 0.09 | 0.092 | 0.092 | 0.090762 | 0.0936 | 0.092 | 0.093 | 0.0936 | 0.092 | 0.0864 | 0.09 | 0.09 | 0.094497 |  | 0.09159
| 1 | 5 | 0.0894 | 0.089018 | 0.08 | 0.0927 | 0.087 | 0.088 | 0.09 | 0.092 | 0.092 | 0.094255 | 0.091 | 0.093 | 0.094 | 0.0915 | 0.095 | 0.0987 | 0.09 | 0.085 | 0.093988 |  | 48
| 1 | 6 | 0.0888 | 0.092086 | 0.081 | 0.0939 | 0.087 | 0.092 | 0.09 | 0.094 | 0.094 | 0.092482 | 0.0901 | 0.093 | 0.094 | 0.0908 | 0.095 | 0.0914 | 0.09 | 0.09 | 0.092988
| Mean |  | 0.088667 | 0.091188 | 0.082 | 0.093617 | 0.087667 | 0.0905 | 0.0905 | 0.092 | 0.093 | 0.092814 | 0.092217 | 0.092667 | 0.093667 | 0.092 | 0.094667 | 0.092217 | 0.09 | 0.0875 | 0.084791
| Median |  | 0.0887 | 0.091483 | 0.0815 | 0.09365 | 0.0875 | 0.0905 | 0.0905 | 0.092 | 0.093 | 0.092826 | 0.09235 | 0.093 | 0.094 | 0.0919 | 0.095 | 0.09175 | 0.09 | 0.0875 | 0.093599
| Std Dev. |  | 0.0005164 | 0.0017405 | 0.0035777 | 0.00062423 | 0.0008165 | 0.0015166 | 0.00054772 | 0.0010954 | 0.00089443 | 0.0016158 | 0.001663 | 0.0005164 | 0.0015055 | 0.0014241 | 0.0018619 | 0.004643 | 1.5202e-17 | 0.0027386 | 0.021911
| Rel.Std.Dev. |  | 0.5824% | 1.909% | 4.363% | 0.6668% | 0.9314% | 1.676% | 0.6052% | 1.191% | 0.9617% | 1.741% | 1.803% | 0.5573% | 1.607% | 1.548% | 1.967% | 5.035% | 1.689e-14% | 3.13% | 25.84%
| PDM3 |  | -3.192% | -0.4388% | -10.47% | 2.213% | -4.284% | -1.19% | -1.19% | 0.4477% | 1.539% | 1.337% | 0.6842% | 1.176% | 2.267% | 0.4477% | 3.359% | 0.6842% | -1.736% | -4.466% | -7.423%
| Z-Score (Absolute) |  | 1.5 | 0.37 | 4.5 | 0.73 | 1.95 | 0.67 | 0.67 | 0 | 0.45 | 0.37 | 0.1 | 0.3 | 0.75 | 0 | 1.2 | 0.1 | 0.9 | 2.02 | 3.24
| Analytical results for Pb in OREAS 251c (Certified Value 13.3 ppm) | [21 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 14.3 | 13.6463 | 11 | 14.8 | 13.1 | 12.8 | 13.3 | 11.8 | 13.9 | 13.1351 | 12 | 13 | 13.23 | 15.9 | 13.32 | 12.2 | 14.23 | 14 | 21.921 |  | 1
| 1 | 2 | 13.7 | 13.548 | 10 | 14.8 | 12.6 | 13 | 13.4 | 11.4 | 14.7 | 12.8576 | 13 | 13 | 12.19 | 16 | 13.33 | 12.6 | 14.11 | 14 | 16.545 |  | 17
| 1 | 3 | 13.6 | 13.6011 | 10 | 15.4 | 13 | 13.6 | 12.9 | 11.7 | 14.1 | 13.2891 | 12 | 13 | 11.57 | 17.6 | 13.4 | 12.9 | 15.73 | 14 | 11.687 |  | 16
| 1 | 4 | 14.2 | 13.8342 | 11 | 14.1 | 12.8 | 13.3 | 12.7 | 11.4 | 14.5 | 13.3403 | 15 | 13 | 11.69 | 17.1 | 13.51 | 12.3 | 15.17 | 13 | 13.375 |  | 13.3326
| 1 | 5 | 14.1 | 13.4563 | 10 | 16.2 | 12.8 | 13.3 | 13 | 11.8 | 14.3 | 12.9419 | 14 | 13 | 11.49 | 16.4 | 13.43 | 12.8 | 15.7 | 14 | 12.458 |  | 49
| 1 | 6 | 14 | 13.3869 | 11 | 14.7 | 12.7 | 13.1 | 12.8 | 11.6 | 13.9 | 12.9649 | 13 | 13 | 11.28 | 15.7 | 12.88 | 12.2 | 15.09 | 14 | 13.816
| Mean |  | 13.9833 | 13.5788 | 10.5 | 15 | 12.8333 | 13.1833 | 13.0167 | 11.6167 | 14.2333 | 13.0882 | 13.1667 | 13 | 11.9083 | 16.45 | 13.3117 | 12.5 | 15.005 | 13.8333 | 14.967
| Median |  | 14.05 | 13.5746 | 10.5 | 14.8 | 12.8 | 13.2 | 12.95 | 11.65 | 14.2 | 13.05 | 13 | 13 | 11.63 | 16.2 | 13.365 | 12.45 | 15.13 | 14 | 13.5955
| Std Dev. |  | 0.27869 | 0.15688 | 0.54772 | 0.71833 | 0.18619 | 0.27869 | 0.27869 | 0.18348 | 0.3266 | 0.19793 | 1.169 | 0 | 0.71527 | 0.75033 | 0.22266 | 0.30984 | 0.69925 | 0.40825 | 3.7884
| Rel.Std.Dev. |  | 1.993% | 1.155% | 5.216% | 4.789% | 1.451% | 2.114% | 2.141% | 1.579% | 2.295% | 1.512% | 8.879% | 0% | 6.006% | 4.561% | 1.673% | 2.479% | 4.66% | 2.951% | 25.31%
| PDM3 |  | 4.881% | 1.846% | -21.25% | 12.51% | -3.745% | -1.12% | -2.37% | -12.87% | 6.756% | -1.833% | -1.245% | -2.495% | -10.68% | 23.38% | -0.1572% | -6.245% | 12.54% | 3.755% | 12.26%
| Z-Score (Absolute) |  | 0.56 | 0.22 | NA | 1.4 | 0.4 | 0.11 | 0.25 | 1.41 | 0.77 | 0.19 | NA | NA | 1.17 | 2.61 | 0 | 0.67 | 1.41 | NA | 1.38
Results from laboratories 1.07, 1.15, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 251c (Certified Value 10.2 ppm) | [21 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 1
2
| 1 | 1 | 10.32 | 10.4549 | 12.3247 | 10.6 | 9.5 | 11.06 | 9.8 | 10.6 | 10.4 | [11 empty] | 1
| 1 | 2 | 10.32 | 10.2783 | 12.2596 | 10.5 | 9.4 | 10.66 | 10 | 10.6 | 10.3 | [11 empty] | 34
| 1 | 3 | 10.08 | 10.2052 | 12.2642 | 10.9 | 9.2 | 10.85 | 10.2 | 10.6 | 10.3 | [11 empty] | 16
| 1 | 4 | 10.21 | 9.9498 | 12.3665 | 10.8 | 9 | 10.87 | 9.9 | 10.6 | 10.4 | [11 empty] | 10.2291
| 1 | 5 | 10.36 | 10.0497 | 12.31 | 10.4 | 9.1 | 10.77 | 9.5 | 10.5 | 10.2 | [11 empty] | 50
| 1 | 6 | 10.25 | 10.3504 | 12.2972 | 10.8 | 8.9 | 10.86 | 9.3 | 10.6 | 10.2
| Mean |  | 10.2567 | 10.2147 | 12.3037 | 10.6667 | 9.1833 | 10.845 | 9.7833 | 10.5833 | 10.3
| Median |  | 10.285 | 10.2417 | 12.3036 | 10.7 | 9.15 | 10.855 | 9.85 | 10.6 | 10.3
| Std Dev. |  | 0.10211 | 0.1885 | 0.039934 | 0.19664 | 0.23166 | 0.13187 | 0.33116 | 0.040825 | 0.089443
| Rel.Std.Dev. |  | 0.9956% | 1.845% | 0.3246% | 1.843% | 2.523% | 1.216% | 3.385% | 0.3857% | 0.8684%
| PDM3 |  | 0.2692% | -0.1411% | 20.28% | 4.277% | -10.22% | 6.021% | -4.358% | 3.463% | 0.6928%
| Z-Score (Absolute) |  | 0.08 | 0.16 | 3.68 | 0.67 | 2.05 | 1 | 0.95 | 0.52 | 0
| Analytical results for Pt in OREAS 251c (Indicative Value 3.4 ppb) | [21 empty] | Indicative
Pt | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.21 | [19 empty] | ppb
|  |  | 4A*OES/MS | [19 empty] | 2
2
| 1 | 1 | 8 | [19 empty] | 1
| 1 | 2 | < 5 | [19 empty] | 2
| 1 | 3 | 7 | [19 empty] | 16
| 1 | 4 | < 5 | [19 empty] | 3.4
| 1 | 5 | < 5 | [19 empty] | 8
| 1 | 6 | < 5
| Mean |  | 7.5
| Median |  | 7.5
| Std Dev. |  | 0.70711
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 120.6%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 251c (Certified Value 99 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
1
| 1 | 1 | 100.43 | 100.2047 | 117.461 | 94.9 | 101.5 | 99.8 | 104.5 | 88.7 | 96.8 | 101.2052 | 99.6 | 104 | 85.6 | 102.3 | 88.6 | 102 | 96.6 |  |  |  | 1
| 1 | 2 | 100.64 | 98.3475 | 117.4065 | 91.2 | 98.7 | 102 | 100.5 | 85.8 | 97.7 | 97.6095 | 100.5 | 106 | 63.6 | 99.23 | 94.7 | 101 | 97.2 |  |  |  | 35
| 1 | 3 | 100.57 | 99.0275 | 117.395 | 94.5 | 101.5 | 106 | 98.3 | 87.6 | 99.6 | 99.7293 | 101.7 | 106 | 73.8 | 101.15 | 94.2 | 104 | 95.8 |  |  |  | 16
| 1 | 4 | 100.61 | 93.7978 | 117.4255 | 93.7 | 96.8 | 104 | 95.5 | 85.9 | 96.8 | 100.2349 | 99.9 | 106 | 75.2 | 99.49 | 95.9 | 99.8 | 97.4 |  |  |  | 99.0788
| 1 | 5 | 100.29 | 97.319 | 117.398 | 94.8 | 97 | 83 | 94.2 | 86.4 | 99 | 98.7568 | 98.6 | 102 | 67.2 | 98.85 | 92.8 | 100 | 96.6 |  |  |  | 51
| 1 | 6 | 100.73 | 97.7902 | 117.4425 | 93.6 | 101.5 | 101 | 97.6 | 95 | 100.5 | 98.3784 | 96.1 | 105 | 76.4 | 99.05 | 94 | 100 | 96
| Mean |  | 100.545 | 97.7478 | 117.4214 | 93.7833 | 99.5 | 99.3 | 98.4333 | 88.2333 | 98.4 | 99.319 | 99.4 | 104.8333 | 73.6333 | 100.0117 | 93.3667 | 101.1333 | 96.6
| Median |  | 100.59 | 98.0689 | 117.416 | 94.1 | 100.1 | 101.5 | 97.95 | 87 | 98.35 | 99.2431 | 99.75 | 105.5 | 74.5 | 99.36 | 94.1 | 100.5 | 96.6
| Std Dev. |  | 0.15897 | 2.1827 | 0.026445 | 1.379 | 2.2882 | 8.2837 | 3.6963 | 3.4967 | 1.5375 | 1.3182 | 1.9142 | 1.6021 | 7.6863 | 1.3921 | 2.543 | 1.633 | 0.63246
| Rel.Std.Dev. |  | 0.1581% | 2.233% | 0.02252% | 1.47% | 2.3% | 8.342% | 3.755% | 3.963% | 1.563% | 1.327% | 1.926% | 1.528% | 10.44% | 1.392% | 2.724% | 1.615% | 0.6547%
| PDM3 |  | 1.48% | -1.343% | 18.51% | -5.345% | 0.4251% | 0.2232% | -0.6515% | -10.95% | -0.6851% | 0.2424% | 0.3242% | 5.808% | -25.68% | 0.9415% | -5.765% | 2.074% | -2.502%
| Z-Score (Absolute) |  | 0.54 | 0.67 | 7.87 | 2.4 | 0.09 | 0 | 0.38 | 4.81 | 0.39 | 0.01 | 0.04 | 2.4 | 11.15 | 0.31 | 2.58 | 0.8 | 1.17
| Analytical results for Re in OREAS 251c (Indicative Value 0.001 ppm) | [21 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 3
3
| 1 | 1 | 0.012221 | 0.004 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.002 | < 0.05 | 0.009 | < 0.05 | 0.002 | < 0.002 | 0.0015 | < 0.1 | [6 empty] | 1
| 1 | 2 | 0.011466 | 0.003 | < 0.002 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.05 | 0.01 | < 0.05 | 0.003 | < 0.002 | 0.0008 | < 0.1 | [6 empty] | 36
| 1 | 3 | 0.013023 | < 0.001 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.05 | 0.01 | < 0.05 | < 0.002 | < 0.002 | < 0.0005 | < 0.1 | [6 empty] | 16
| 1 | 4 | 0.012607 | < 0.001 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.05 | 0.011 | < 0.05 | 0.003 | < 0.002 | < 0.0005 | < 0.1 | [6 empty] | 0.0013903
| 1 | 5 | 0.01425 | < 0.001 | < 0.002 | < 0.002 | 0.002 | 0.002 | 0.002 | < 0.05 | 0.01 | < 0.05 | < 0.002 | < 0.002 | < 0.0005 | < 0.1 | [6 empty] | 7
| 1 | 6 | 0.01235 | < 0.001 | < 0.002 | < 0.002 | 0.002 | 0.002 | < 0.002 | < 0.05 | 0.012 | < 0.05 | < 0.002 | < 0.002 | 0.001 | < 0.1
| Mean |  | 0.012653 | 0.0035 |  |  | 0.002 | 0.002 | 0.002 |  | 0.010333 |  | 0.0026667 |  | 0.0011
| Median |  | 0.012479 | 0.0035 |  |  | 0.002 | 0.002 | 0.002 |  | 0.01 |  | 0.003 |  | 0.001
| Std Dev. |  | 0.00093542 | 0.00070711 |  |  | 0 | 0 | 0 |  | 0.0010328 |  | 0.00057735 |  | 0.00036056
| Rel.Std.Dev. |  | 7.393% | 20.2% |  |  | 0% | 0% | 0% |  | 9.995% |  | 21.65% |  | 32.78%
| PDM3 |  | 810.1% | 151.7% |  |  | 43.86% | 43.86% | 43.86% |  | 643.3% |  | 91.81% |  | -20.88%
| Z-Score (Absolute) |  | 9.8 | 0.22 | 0.67 | 0.67 | 0.07 | 0.07 | 0.37 | 20.9 | 7.72 | 20.9 | 0.07 | 0.67 | 0.97 | 43.38
| Analytical results for S in OREAS 251c (Certified Value 0.01 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0093 | 0.010563 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0089381 | 0.0081 | 0.015 | < 0.04 | < 0.04 | < 0.005 | 0.01 | 0.015 | 0.32004 |  |  | 1
| 1 | 2 | 0.0092 | 0.010491 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0092913 | 0.009 | 0.015 | < 0.04 | < 0.04 | 0.0066 | 0.01 | 0.015 | 0.019696 |  |  | 21
| 1 | 3 | 0.0091 | 0.010243 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0089722 | 0.0096 | 0.016 | < 0.04 | < 0.04 | < 0.005 | 0.01 | 0.015 | 0.015942 |  |  | 16
| 1 | 4 | 0.0091 | 0.010378 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0087361 | 0.0093 | 0.015 | < 0.04 | < 0.04 | 0.0056 | 0.01 | 0.015 | 0.020356 |  |  | 0.0097975
| 1 | 5 | 0.0095 | 0.010921 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0088847 | 0.0089 | 0.015 | < 0.04 | < 0.04 | 0.0072 | 0.01 | 0.015 | 0.018817 |  |  | 52
| 1 | 6 | 0.0092 | 0.010142 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0091981 | 0.0086 | 0.016 | < 0.04 | < 0.04 | < 0.005 | 0.01 | 0.015 | 0.018421
| Mean |  | 0.0092333 | 0.010456 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.018333 | 0.0090034 | 0.0089167 | 0.015333 |  |  | 0.0064667 | 0.01 | 0.015 | 0.068879
| Median |  | 0.0092 | 0.010435 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.0089552 | 0.00895 | 0.015 |  |  | 0.0066 | 0.01 | 0.015 | 0.019257
| Std Dev. |  | 0.00015055 | 0.00027526 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0040825 | 0.00020569 | 0.00052694 | 0.0005164 |  |  | 0.00080829 | 0 | 0 | 0.12305
| Rel.Std.Dev. |  | 1.631% | 2.632% | 0% | 0% | 0% | 0% | 0% | 0% | 22.27% | 2.285% | 5.91% | 3.368% |  |  | 12.5% | 0% | 0% | 178.7%
| PDM3 |  | -5.759% | 6.723% | 2.066% | 2.066% | 2.066% | 2.066% | 2.066% | 2.066% | 87.12% | -8.105% | -8.991% | 56.5% |  |  | -34% | 2.066% | 53.1% | 603%
| Z-Score (Absolute) |  | 0.5 | 0.3 | 2.16 | 0 | 0 | 0 | 0 | 0 | 5.4 | 0.65 | 0.7 | 3.46 | 6.48 | 6.48 | 3.58 | 0 | 3.24 | 38.18
| Analytical results for Sb in OREAS 251c (Certified Value 2.65 ppm) | [21 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 2.65 | 2.5699 | 2.5 | 1.77 | 2.81 | 2.95 | 2.75 | 2.75 | 2.88 | 2.5021 | 2.6 | 2.5 | 2.62 | 2.71 | 2.28 | 2.7 | < 5 | 2.8 | 2.286 |  | 1
| 1 | 2 | 2.76 | 2.5946 | 2.6 | 1.67 | 2.7 | 2.85 | 2.8 | 2.51 | 2.91 | 2.4254 | 2.54 | 2.6 | 2.36 | 2.71 | 2.33 | 2.8 | < 5 | 3 | 2.185 |  | 37
| 1 | 3 | 2.67 | 2.68 | 2.7 | 1.94 | 2.81 | 3.02 | 2.59 | 2.6 | 2.9 | 2.832 | 2.59 | 2.5 | 2.45 | 2.72 | 2.39 | 2.7 | < 5 | 2.8 | 5.28 |  | 16
| 1 | 4 | 2.53 | 2.6455 | 2.7 | 1.97 | 2.76 | 2.98 | 2.55 | 2.47 | 2.73 | 2.6042 | 2.5 | 2.6 | 2.38 | 2.86 | 2.34 | 2.8 | < 5 | 3 | 3.581 |  | 2.6538
| 1 | 5 | 2.52 | 2.595 | 2.8 | 1.71 | 2.82 | 3.09 | 2.53 | 2.57 | 2.78 | 2.8689 | 2.46 | 2.5 | 2.4 | 2.73 | 2.43 | 2.8 | < 5 | 2.8 | 5.264 |  | 53
| 1 | 6 | 2.64 | 2.5712 | 2.7 | 1.85 | 2.66 | 3.06 | 2.62 | 2.68 | 2.8 | 2.4262 | 2.45 | 2.6 | 2.33 | 2.75 | 2.29 | 2.6 | < 5 | 2.8 | 3.526
| Mean |  | 2.6283 | 2.6094 | 2.6667 | 1.8183 | 2.76 | 2.9917 | 2.64 | 2.5967 | 2.8333 | 2.6098 | 2.5233 | 2.55 | 2.4233 | 2.7467 | 2.3433 | 2.7333 |  | 2.8667 | 3.687
| Median |  | 2.645 | 2.5948 | 2.7 | 1.81 | 2.785 | 3 | 2.605 | 2.585 | 2.84 | 2.5532 | 2.52 | 2.55 | 2.39 | 2.725 | 2.335 | 2.75 |  | 2.8 | 3.5535
| Std Dev. |  | 0.090646 | 0.044137 | 0.10328 | 0.12238 | 0.066633 | 0.08612 | 0.11027 | 0.10463 | 0.073666 | 0.19792 | 0.064083 | 0.054772 | 0.10443 | 0.057504 | 0.05785 | 0.08165 |  | 0.10328 | 1.3624
| Rel.Std.Dev. |  | 3.449% | 1.691% | 3.873% | 6.73% | 2.414% | 2.879% | 4.177% | 4.029% | 2.6% | 7.584% | 2.54% | 2.148% | 4.31% | 2.094% | 2.469% | 2.987% |  | 3.603% | 36.95%
| PDM3 |  | -0.9589% | -1.673% | 0.4855% | -31.48% | 4.003% | 12.73% | -0.5193% | -2.152% | 6.766% | -1.657% | -4.916% | -3.911% | -8.684% | 3.5% | -11.7% | 2.998% |  | 8.022% | 38.93%
| Z-Score (Absolute) |  | 0 | 0.11 | 0.22 | 4.62 | 0.75 | 2.07 | 0.07 | 0.18 | 1.17 | 0.11 | 0.6 | 0.45 | 1.17 | 0.67 | 1.62 | 0.6 | 0.73 | 1.36 | 6.03
| Analytical results for Sc in OREAS 251c (Certified Value 12.9 ppm) | [21 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 13.8 | 13.4547 | 15.369 | 14 | 13.6 | 13.8 | 11.9 | 13.5 | 13.2 | 13 | 13.777 | 11.84 | 13 | 13.2 | 13.7 | 12.1 | 13 | 13 | 35.554 |  | 1
| 1 | 2 | 13.5 | 13.3438 | 15.314 | 14 | 13.3 | 13.6 | 11.9 | 13.6 | 12 | 13 | 13.7374 | 12.12 | 13 | 13.1 | 13 | 12.5 | 11 | 13 | 12.501 |  | 38
| 1 | 3 | 13.5 | 13.2785 | 15.289 | 14 | 13.2 | 14.5 | 12.7 | 13.1 | 12.5 | 12 | 13.2179 | 12.14 | 13 | 13.2 | 13.4 | 12.5 | 13 | 13 | 12.458 |  | 16
| 1 | 4 | 13.5 | 12.7226 | 15.345 | 14 | 13 | 13.6 | 12.3 | 12.6 | 12 | 12.8 | 13.0947 | 12.11 | 13 | 12.9 | 13.1 | 12.4 | 12 | 13 | 12.548 |  | 12.931
| 1 | 5 | 13.5 | 12.916 | 15.22 | 13 | 14.1 | 14 | 12.4 | 12.6 | 12.4 | 12.4 | 13.9227 | 11.91 | 13 | 13.1 | 12.8 | 12.4 | 14 | 13 | 12.531 |  | 54
| 1 | 6 | 13.4 | 13.2425 | 15.329 | 14 | 13.9 | 13.6 | 12 | 13.1 | 12.7 | 13 | 13.7224 | 11.68 | 12 | 13 | 13 | 12.3 | 12 | 13 | 12.423
| Mean |  | 13.5333 | 13.1597 | 15.311 | 13.8333 | 13.5167 | 13.85 | 12.2 | 13.0833 | 12.4667 | 12.7 | 13.5787 | 11.9667 | 12.8333 | 13.0833 | 13.1667 | 12.3667 | 12.5 | 13 | 16.3358
| Median |  | 13.5 | 13.2605 | 15.3215 | 14 | 13.45 | 13.7 | 12.15 | 13.1 | 12.45 | 12.9 | 13.7299 | 12.01 | 13 | 13.1 | 13.05 | 12.4 | 12.5 | 13 | 12.516
| Std Dev. |  | 0.13663 | 0.28011 | 0.052188 | 0.40825 | 0.42622 | 0.35637 | 0.32249 | 0.42622 | 0.45461 | 0.41473 | 0.33702 | 0.18737 | 0.40825 | 0.1169 | 0.3266 | 0.15055 | 1.0488 | 0 | 9.4151
| Rel.Std.Dev. |  | 1.01% | 2.129% | 0.3409% | 2.951% | 3.153% | 2.573% | 2.643% | 3.258% | 3.647% | 3.266% | 2.482% | 1.566% | 3.181% | 0.8935% | 2.48% | 1.217% | 8.39% | 0% | 57.63%
| PDM3 |  | 4.658% | 1.769% | 18.41% | 6.978% | 4.529% | 7.107% | -5.653% | 1.178% | -3.591% | -1.786% | 5.009% | -7.457% | -0.7552% | 1.178% | 1.823% | -4.364% | -3.333% | 0.5337% | 26.33%
| Z-Score (Absolute) |  | 0.55 | 0 | 3.16 | NA | 0.52 | 1.01 | 1.41 | 0.11 | 1.02 | 0.67 | 0.61 | 1.75 | NA | 0.11 | 0.01 | 1.16 | NA | NA | 4.66
Results from laboratories 1.07, 1.16, 1.21 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 251c (Indicative Value 0.83 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  |  | 2
2
| 1 | 1 | < 1.5 | 0.4 | < 1 | 1 | < 1 | < 1 | < 1 | < 1 | < 2 | 0.22 | < 1 | 0.5 | 0.5 | 1.71 | < 5 | 6.445 |  |  |  |  | 1
| 1 | 2 | < 1.5 | 0.4 | < 1 | 1 | < 1 | < 1 | < 1 | 1 | < 2 | 0.26 | < 1 | 1 | 0.6 | 1.95 | < 5 | 4.944 |  |  |  |  | 39
| 1 | 3 | < 1.5 | 0.2 | < 1 | 1 | < 1 | < 1 | < 1 | < 1 | < 2 | 0.22 | < 1 | 0.6 | 0.5 | 1.48 | < 5 | 4.058 |  |  |  |  | 16
| 1 | 4 | < 1.5 | 0.5 | < 1 | 1 | < 1 | < 1 | 1 | < 1 | < 2 | 0.22 | < 1 | 0.7 | 0.3 | 1.77 | < 5 | 2.809 |  |  |  |  | 0.83359
| 1 | 5 | < 1.5 | 0.5 | < 1 | 1 | < 1 | 1 | 1 | 1 | < 2 | 0.21 | < 1 | 1 | 0.4 | 1.91 | < 5 | 3.066 |  |  |  |  | 8
| 1 | 6 | < 1.5 | 0.4 | < 1 | 1 | < 1 | 1 | 1 | 1 | < 2 | 0.2 | < 1 | 0.5 | 0.6 | 1.31 | < 5 | 5.474
| Mean |  |  | 0.4 |  | 1 |  | 1 | 1 | 1 |  | 0.22167 |  | 0.71667 | 0.48333 | 1.6883 |  | 4.466
| Median |  |  | 0.4 |  | 1 |  | 1 | 1 | 1 |  | 0.22 |  | 0.65 | 0.5 | 1.74 |  | 4.501
| Std Dev. |  |  | 0.10954 |  | 0 |  | 0 | 0 | 0 |  | 0.020412 |  | 0.23166 | 0.1169 | 0.24951 |  | 1.4166
| Rel.Std.Dev. |  |  | 27.39% |  | 0% |  | 0% | 0% | 0% |  | 9.209% |  | 32.32% | 24.19% | 14.78% |  | 31.72%
| PDM3 |  |  | -52.01% |  | 19.96% |  | 19.96% | 19.96% | 19.96% |  | -73.41% |  | -14.03% | -42.02% | 102.5% |  | 435.8%
| Z-Score (Absolute) |  | 0.05 | 0.93 | 0.65 | 0.74 | 0.65 | 0.19 | 0.05 | 0.05 | 0.74 | 1.43 | 0.65 | 0.05 | 0.7 | 2.66 | 4.93 | 10.42
| Analytical results for Sm in OREAS 251c (Certified Value 7.34 ppm) | [21 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 7.41 | 7.4164 | 8.1322 | 7.3 | 7.6 | 7.6 | 7 | 7.39 | 7.75 | [11 empty] | 1
| 1 | 2 | 7.47 | 7.1384 | 8.1132 | 7.2 | 6.7 | 7.5 | 7 | 7.3 | 7.75 | [11 empty] | 5
| 1 | 3 | 7.11 | 7.3033 | 8.1203 | 6.8 | 6.7 | 7.69 | 6.9 | 7.37 | 7.75 | [11 empty] | 16
| 1 | 4 | 7.1 | 7.0021 | 8.1985 | 6.2 | 6.3 | 7.61 | 7.1 | 7.28 | 7.75 | [11 empty] | 7.3446
| 1 | 5 | 7.39 | 7.2701 | 8.1015 | 7.3 | 6.8 | 7.38 | 6.9 | 7.37 | 7.7 | [11 empty] | 55
| 1 | 6 | 7.16 | 7.3288 | 8.1816 | 7.6 | 6.5 | 7.67 | 6.8 | 7.41 | 7.65
| Mean |  | 7.2733 | 7.2432 | 8.1412 | 7.0667 | 6.7667 | 7.575 | 6.95 | 7.3533 | 7.725
| Median |  | 7.275 | 7.2867 | 8.1262 | 7.25 | 6.7 | 7.605 | 6.95 | 7.37 | 7.75
| Std Dev. |  | 0.16765 | 0.14885 | 0.039503 | 0.49666 | 0.44572 | 0.1164 | 0.10488 | 0.05164 | 0.041833
| Rel.Std.Dev. |  | 2.305% | 2.055% | 0.4852% | 7.028% | 6.587% | 1.537% | 1.509% | 0.7023% | 0.5415%
| PDM3 |  | -0.9698% | -1.38% | 10.85% | -3.784% | -7.868% | 3.138% | -5.372% | 0.1194% | 5.18%
| Z-Score (Absolute) |  | 0 | 0.07 | 1.94 | 0.46 | 1.13 | 0.67 | 0.72 | 0.18 | 1.01
| Analytical results for Sn in OREAS 251c (Certified Value 8.95 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES |  | 2
3
| 1 | 1 | 9.3 | 9.5374 | 1 | 8.7 | 9.6 | 8.7 | 8.7 | 9.3 | 9.5 | 8.4883 | 6.44 | 8.6 | 9 | 9.6 | 7.9 | 10 | < 5 | 9.6 | 1.038 |  | 1
| 1 | 2 | 9.1 | 9.5069 | 2 | 8.1 | 9.4 | 8.4 | 8.7 | 8.7 | 9.4 | 8.1187 | 7.14 | 8.9 | 8.7 | 9.4 | 8 | 10.4 | < 5 | 9.7 | 9.447 |  | 12
| 1 | 3 | 8.7 | 9.1921 | < 1 | 8.1 | 9.7 | 8.9 | 8.2 | 8.7 | 9.4 | 8.7148 | 5.94 | 8.3 | 8.3 | 9.4 | 8 | 10.2 | < 5 | 9.6 | 9.114 |  | 16
| 1 | 4 | 9 | 9.483 | 2 | 8.3 | 9.5 | 8.9 | 8 | 8.4 | 9 | 8.9083 | 6.93 | 8.3 | 8.2 | 9.4 | 8 | 10.8 | < 5 | 9.7 | 8.725 |  | 8.9454
| 1 | 5 | 8.7 | 10.0044 | < 1 | 8.4 | 9.4 | 8.7 | 8.1 | 8.8 | 9 | 8.9717 | 6.04 | 8.3 | 8.1 | 9.2 | 8 | 10.2 | < 5 | 9.7 | 8.152 |  | 56
| 1 | 6 | 9.4 | 10.0252 | 2 | 8.1 | 9.1 | 8.7 | 8.4 | 9.2 | 9.1 | 8.2523 | 6.74 | 8.5 | 7.9 | 9 | 7.8 | 10.4 | < 5 | 9.6 | 9.021
| Mean |  | 9.0333 | 9.6248 | 1.75 | 8.2833 | 9.45 | 8.7167 | 8.35 | 8.85 | 9.2333 | 8.5757 | 6.5383 | 8.4833 | 8.3667 | 9.3333 | 7.95 | 10.3333 |  | 9.65 | 7.5828
| Median |  | 9.05 | 9.5221 | 2 | 8.2 | 9.45 | 8.7 | 8.3 | 8.75 | 9.25 | 8.6016 | 6.59 | 8.4 | 8.25 | 9.4 | 8 | 10.3 |  | 9.65 | 8.873
| Std Dev. |  | 0.29439 | 0.32661 | 0.5 | 0.24014 | 0.20736 | 0.18348 | 0.30166 | 0.33912 | 0.22509 | 0.34867 | 0.48417 | 0.24014 | 0.40825 | 0.20656 | 0.083666 | 0.27325 |  | 0.054772 | 3.2358
| Rel.Std.Dev. |  | 3.259% | 3.393% | 28.57% | 2.899% | 2.194% | 2.105% | 3.613% | 3.832% | 2.438% | 4.066% | 7.405% | 2.831% | 4.879% | 2.213% | 1.052% | 2.644% |  | 0.5676% | 42.67%
| PDM3 |  | 0.9835% | 7.596% | -80.44% | -7.401% | 5.641% | -2.556% | -6.655% | -1.066% | 3.219% | -4.132% | -26.91% | -5.165% | -6.469% | 4.337% | -11.13% | 15.52% |  | 7.877% | -15.23%
| Z-Score (Absolute) |  | 0.41 | 1.04 | NA | 0.38 | 0.85 | 0.07 | 0.31 | 0.22 | 0.62 | 0.07 | 2.23 | 0.17 | 0.3 | 0.73 | 0.74 | 1.79 | 6.5 | 1.06 | 1.13
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 251c (Certified Value 280 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
0
| 1 | 1 | 282.62 | 279.874 | 282.5606 | 304 | 287 | 274 | 279 | 292 | 283 | 280 | 285.3817 | 251 | 273 | 287 | 292.58 | 265 | 248 | 270.1 | 278 | 173.191 | 1
| 1 | 2 | 289.11 | 277.9063 | 281.7338 | 292 | 271 | 273 | 280 | 292 | 283 | 287 | 283.2307 | 252 | 280 | 290 | 283.43 | 276 | 230 | 279 | 281 | 270.227 | 13
| 1 | 3 | 288.09 | 273.4297 | 284.4474 | 295 | 275 | 269 | 285 | 289 | 289 | 280 | 287.7211 | 255 | 274 | 285 | 287.57 | 274 | 253 | 279.8 | 277 | 267.49 | 16
| 1 | 4 | 281.77 | 263.9446 | 283.1754 | 300 | 266 | 267 | 280 | 280 | 281 | 280 | 286.3978 | 251 | 276 | 281 | 290.31 | 272 | 242 | 276.6 | 279 | 265.557 | 280.0794
| 1 | 5 | 279.55 | 267.3394 | 281.5536 | 307 | 280 | 266 | 272 | 288 | 281 | 284 | 281.0352 | 246 | 275 | 285 | 282.37 | 271 | 237 | 276.6 | 278 | 264.74 | 57
| 1 | 6 | 285.54 | 272.8715 | 280.0272 | 296 | 278 | 270 | 279 | 287 | 286 | 286 | 281.6801 | 241 | 287 | 284 | 284.21 | 265 | 244 | 279.3 | 280 | 265.915
| Mean |  | 284.4467 | 272.5609 | 282.2497 | 299 | 276.1667 | 269.8333 | 279.1667 | 288 | 283.8333 | 282.8333 | 284.2411 | 249.3333 | 277.5 | 285.3333 | 286.745 | 270.5 | 242.3333 | 276.9 | 278.8333 | 251.1867
| Median |  | 284.08 | 273.1506 | 282.1472 | 298 | 276.5 | 269.5 | 279.5 | 288.5 | 283 | 282 | 284.3062 | 251 | 275.5 | 285 | 285.89 | 271.5 | 243 | 277.8 | 278.5 | 265.736
| Std Dev. |  | 3.7602 | 6.0731 | 1.5148 | 5.7271 | 7.3052 | 3.1885 | 4.1673 | 4.4272 | 3.1252 | 3.2506 | 2.6803 | 5.0067 | 5.244 | 3.0111 | 4.0986 | 4.5935 | 8.1158 | 3.6056 | 1.472 | 38.2591
| Rel.Std.Dev. |  | 1.322% | 2.228% | 0.5367% | 1.915% | 2.645% | 1.182% | 1.493% | 1.537% | 1.101% | 1.149% | 0.943% | 2.008% | 1.89% | 1.055% | 1.429% | 1.698% | 3.349% | 1.302% | 0.5279% | 15.23%
| PDM3 |  | 1.559% | -2.684% | 0.7749% | 6.755% | -1.397% | -3.658% | -0.3259% | 2.828% | 1.34% | 0.9833% | 1.486% | -10.98% | -0.921% | 1.876% | 2.38% | -3.42% | -13.48% | -1.135% | -0.4449% | -10.32%
| Z-Score (Absolute) |  | 0.62 | 0.74 | 0.37 | 2.29 | 0.32 | 1.05 | 0.02 | 1.03 | 0.55 | 0.44 | 0.6 | 3.4 | 0.17 | 0.73 | 0.89 | 0.97 | 4.2 | 0.24 | 0.02 | 3.18
| Analytical results for Ta in OREAS 251c (Certified Value 1.83 ppm) | [21 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.04 | 1.7626 | < 0.1 | 1.35 | 1.91 | 1.88 | 1.92 | 1.78 | 1.86 | 1.4403 | 1.77 | 1.62 | 1.9 | 1.95 | 1.7 | 2.11 | 1.9 |  |  |  | 1
| 1 | 2 | 1.97 | 1.7165 | < 0.1 | 1.24 | 1.88 | 1.88 | 1.95 | 1.7 | 1.91 | 1.2046 | 1.77 | 1.6 | 1.7 | 1.9 | 1.7 | 2.2 | 1.9 |  |  |  | 26
| 1 | 3 | 1.82 | 1.8141 | < 0.1 | 1.34 | 1.95 | 1.98 | 1.91 | 1.88 | 1.85 | 1.131 | 1.79 | 1.5 | 1.7 | 1.9 | 1.7 | 2.08 | 2 |  |  |  | 16
| 1 | 4 | 1.94 | 1.6729 | < 0.1 | 1.21 | 1.87 | 1.92 | 1.86 | 1.64 | 1.91 | 1.0951 | 1.76 | 1.62 | 1.7 | 1.95 | 1.7 | 2.15 | 1.9 |  |  |  | 1.8329
| 1 | 5 | 1.75 | 1.7331 | < 0.1 | 1.29 | 1.88 | 1.94 | 1.96 | 1.68 | 1.86 | 1.3765 | 1.73 | 1.51 | 1.7 | 1.87 | 1.7 | 2.09 | 1.9 |  |  |  | 58
| 1 | 6 | 1.82 | 1.7728 | < 0.1 | 1.3 | 1.83 | 1.94 | 1.91 | 1.77 | 1.93 | 1.4115 | 1.73 | 1.59 | 1.6 | 1.89 | 1.6 | 2.13 | 1.8
| Mean |  | 1.89 | 1.7453 |  | 1.2883 | 1.8867 | 1.9233 | 1.9183 | 1.7417 | 1.8867 | 1.2765 | 1.7583 | 1.5733 | 1.7167 | 1.91 | 1.6833 | 2.1267 | 1.9
| Median |  | 1.88 | 1.7478 |  | 1.295 | 1.88 | 1.93 | 1.915 | 1.735 | 1.885 | 1.2906 | 1.765 | 1.595 | 1.7 | 1.9 | 1.7 | 2.12 | 1.9
| Std Dev. |  | 0.11027 | 0.049016 |  | 0.054924 | 0.040332 | 0.038816 | 0.035449 | 0.086352 | 0.033862 | 0.15119 | 0.024014 | 0.054283 | 0.098319 | 0.032863 | 0.040825 | 0.044121 | 0.063246
| Rel.Std.Dev. |  | 5.835% | 2.808% |  | 4.263% | 2.138% | 2.018% | 1.848% | 4.958% | 1.795% | 11.84% | 1.366% | 3.45% | 5.727% | 1.721% | 2.425% | 2.075% | 3.329%
| PDM3 |  | 3.116% | -4.776% |  | -29.71% | 2.934% | 4.935% | 4.662% | -4.977% | 2.934% | -30.35% | -4.067% | -14.16% | -6.341% | 4.207% | -8.159% | 16.03% | 3.662%
| Z-Score (Absolute) |  | 0.63 | 0.06 | 8.13 | 2.24 | 0.61 | 0.79 | 0.76 | 0.08 | 0.61 | 2.29 | 0 | 0.88 | 0.2 | 0.72 | 0.36 | 1.75 | 0.67
| Analytical results for Tb in OREAS 251c (Certified Value 0.77 ppm) | [21 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [10 empty] | 2
3
| 1 | 1 | 0.75 | 0.78457 | 0.9476 | 0.8 | 0.8 | 0.76 | 0.7 | 0.79 | 0.78 | 2.562 | [10 empty] | 1
| 1 | 2 | 0.7 | 0.77946 | 0.92861 | 0.8 | 0.7 | 0.74 | 0.7 | 0.8 | 0.78 | 0.614 | [10 empty] | 27
| 1 | 3 | 0.71 | 0.76456 | 0.9458 | 0.8 | 0.7 | 0.73 | 0.7 | 0.79 | 0.8 | 0.448 | [10 empty] | 16
| 1 | 4 | 0.71 | 0.75852 | 0.92906 | 0.8 | 0.7 | 0.75 | 0.7 | 0.8 | 0.78 | 0.591 | [10 empty] | 0.76538
| 1 | 5 | 0.72 | 0.77052 | 0.97217 | 0.8 | 0.7 | 0.77 | 0.7 | 0.79 | 0.8 | 0.638 | [10 empty] | 59
| 1 | 6 | 0.72 | 0.79375 | 0.95795 | 0.8 | 0.7 | 0.75 | 0.7 | 0.79 | 0.8 | 0.627
| Mean |  | 0.71833 | 0.77523 | 0.94687 | 0.8 | 0.71667 | 0.75 | 0.7 | 0.79333 | 0.79 | 0.91333
| Median |  | 0.715 | 0.77499 | 0.9467 | 0.8 | 0.7 | 0.75 | 0.7 | 0.79 | 0.79 | 0.6205
| Std Dev. |  | 0.017224 | 0.013142 | 0.016824 | 1.2162e-16 | 0.040825 | 0.014142 | 1.2162e-16 | 0.005164 | 0.010954 | 0.81067
| Rel.Std.Dev. |  | 2.398% | 1.695% | 1.777% | 1.52e-14% | 5.696% | 1.886% | 1.737e-14% | 0.6509% | 1.387% | 88.76%
| PDM3 |  | -6.147% | 1.287% | 23.71% | 4.523% | -6.364% | -2.009% | -8.542% | 3.652% | 3.217% | 19.33%
| Z-Score (Absolute) |  | 1.21 | 0.25 | 2.64 | NA | NA | 0.67 | NA | 0.06 | 0 | 2.08
Results from laboratories 1.07, 1.17 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 251c (Certified Value 0.12 ppm) | [21 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | < 0.2 | < 0.1 | 0.52 | 0.13 | 0.11 | 0.1 | 0.17 | 0.1 | 0.13543 | 0.06 | 0.1 | 0.14 | 0.18 | 0.1 | < 0.2 | [5 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.48 | 0.13 | 0.1 | 0.1 | 0.14 | 0.14 | 0.11234 | 0.08 | 0.1 | 0.16 | 0.18 | 0.1 | < 0.2 | [5 empty] | 28
| 1 | 3 | < 0.2 | < 0.1 | 0.51 | 0.1 | 0.11 | 0.09 | 0.15 | 0.12 | < 0.1 | 0.1 | 0.1 | 0.15 | 0.14 | 0.1 | < 0.2 | [5 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.51 | 0.12 | 0.09 | 0.1 | 0.17 | 0.12 | 0.12792 | 0.12 | 0.1 | 0.12 | 0.15 | 0.1 | < 0.2 | [5 empty] | 0.12144
| 1 | 5 | < 0.2 | < 0.1 | 0.5 | 0.11 | 0.13 | 0.12 | 0.16 | 0.11 | 0.1128 | 0.12 | 0.1 | 0.09 | 0.16 | 0.11 | < 0.2 | [5 empty] | 60
| 1 | 6 | < 0.2 | < 0.1 | 0.5 | 0.11 | 0.1 | 0.11 | 0.16 | 0.13 | 0.10839 | 0.14 | 0.1 | 0.06 | 0.18 | 0.1 | < 0.2
| Mean |  |  |  | 0.50333 | 0.11667 | 0.10667 | 0.10333 | 0.15833 | 0.12 | 0.11938 | 0.10333 | 0.1 | 0.12 | 0.165 | 0.10167
| Median |  |  |  | 0.505 | 0.115 | 0.105 | 0.1 | 0.16 | 0.12 | 0.1128 | 0.11 | 0.1 | 0.13 | 0.17 | 0.1
| Std Dev. |  |  |  | 0.013663 | 0.012111 | 0.013663 | 0.010328 | 0.01169 | 0.014142 | 0.011664 | 0.029439 | 1.5202e-17 | 0.038471 | 0.017607 | 0.0040825
| Rel.Std.Dev. |  |  |  | 2.714% | 10.38% | 12.81% | 9.995% | 7.383% | 11.79% | 9.771% | 28.49% | 1.52e-14% | 32.06% | 10.67% | 4.016%
| PDM3 |  |  |  | 314.5% | -3.929% | -12.16% | -14.91% | 30.38% | -1.184% | -1.697% | -14.91% | -17.65% | -1.184% | 35.87% | -16.28%
| Z-Score (Absolute) |  | 0.59 | 4.63 | 32.05 | 0.76 | 0.05 | 0.32 | 4.13 | 1.03 | 0.05 | 0.32 | NA | 1.03 | 4.67 | 0.45 | 0.59
Results from laboratory 1.16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 251c (Certified Value 12.3 ppm) | [21 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 12.66 | 12.7657 | 14.6594 | 15.2 | 13.05 | 11.4 | 13.35 | 12.15 | 11.75 | 13.18 | 12.35 | 12.9 | 11.6 | 13.31 | 12.2 | 11 | 12.9 |  |  |  | 1
| 1 | 2 | 12.43 | 12.3512 | 14.6836 | 14.7 | 12.3 | 11.25 | 13.3 | 11.85 | 12.5 | 13.0167 | 12.38 | 13 | 11.3 | 12.87 | 12.3 | 11.3 | 12.9 |  |  |  | 29
| 1 | 3 | 12.47 | 12.7706 | 14.6478 | 15.2 | 13.05 | 12.1 | 12.8 | 12.55 | 11.9 | 13.1634 | 12.54 | 13.1 | 11 | 13.2 | 12 | 10.9 | 13 |  |  |  | 16
| 1 | 4 | 12.4 | 12.221 | 14.7646 | 14.9 | 12.8 | 11.75 | 12.85 | 10.85 | 12 | 13.1247 | 12.35 | 13 | 10.8 | 13.1 | 12 | 10.9 | 12.9 |  |  |  | 12.2968
| 1 | 5 | 12.52 | 12.3722 | 14.6819 | 14.8 | 12.55 | 10.85 | 12.75 | 11.55 | 11.75 | 12.9927 | 12.2 | 12.9 | 10.8 | 13.07 | 11.9 | 11 | 12.7 |  |  |  | 61
| 1 | 6 | 12.66 | 12.4847 | 14.8392 | 14.9 | 12.65 | 11.65 | 12.7 | 12.2 | 11.9 | 13.127 | 11.97 | 13.3 | 10.7 | 13.21 | 11.5 | 10.8 | 13
| Mean |  | 12.5233 | 12.4942 | 14.7128 | 14.95 | 12.7333 | 11.5 | 12.9583 | 11.8583 | 11.9667 | 13.1007 | 12.2983 | 13.0333 | 11.0333 | 13.1267 | 11.9833 | 10.9833 | 12.9
| Median |  | 12.495 | 12.4285 | 14.6828 | 14.9 | 12.725 | 11.525 | 12.825 | 12 | 11.9 | 13.1258 | 12.35 | 13 | 10.9 | 13.15 | 12 | 10.95 | 12.9
| Std Dev. |  | 0.11325 | 0.2281 | 0.074222 | 0.20736 | 0.29439 | 0.43359 | 0.28882 | 0.59868 | 0.27869 | 0.077728 | 0.19385 | 0.15055 | 0.35024 | 0.15214 | 0.27869 | 0.17224 | 0.10954
| Rel.Std.Dev. |  | 0.9044% | 1.826% | 0.5045% | 1.387% | 2.312% | 3.77% | 2.229% | 5.049% | 2.329% | 0.5933% | 1.576% | 1.155% | 3.174% | 1.159% | 2.326% | 1.568% | 0.8492%
| PDM3 |  | 1.842% | 1.605% | 19.65% | 21.58% | 3.55% | -6.48% | 5.38% | -3.566% | -2.685% | 6.538% | 0.0123% | 5.989% | -10.27% | 6.748% | -2.549% | -10.68% | 4.905%
| Z-Score (Absolute) |  | 0 | 0.04 | 2.65 | 2.94 | 0.25 | 1.24 | 0.53 | 0.81 | 0.67 | 0.7 | 0.27 | 0.62 | 1.8 | 0.73 | 0.65 | 1.87 | 0.46
| Analytical results for Ti in OREAS 251c (Certified Value 0.593 wt.%) | [21 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.5935 | 0.5915 | 0.60198 | 0.629 | 0.59 | 0.596 | 0.599 | 0.623 | 0.592 | 0.6 | 0.53848 | 0.5085 | 0.581 | 0.578 | 0.6518 | 0.59 | 0.597 | 0.58 | 0.64 | 0.56208 | 1
| 1 | 2 | 0.5917 | 0.58848 | 0.60712 | 0.7 | 0.59 | 0.592 | 0.591 | 0.632 | 0.588 | 0.6 | 0.53188 | 0.5176 | 0.57 | 0.597 | 0.6244 | 0.603 | 0.557 | 0.59 | 0.65 | 0.59909 | 30
| 1 | 3 | 0.5596 | 0.58679 | 0.60656 | 0.643 | 0.6 | 0.584 | 0.603 | 0.628 | 0.603 | 0.594 | 0.59162 | 0.5182 | 0.553 | 0.578 | 0.6435 | 0.595 | 0.602 | 0.59 | 0.645 | 0.59415 | 16
| 1 | 4 | 0.5841 | 0.57226 | 0.60559 | 0.648 | 0.58 | 0.576 | 0.603 | 0.622 | 0.609 | 0.59 | 0.59537 | 0.5254 | 0.548 | 0.566 | 0.6451 | 0.587 | 0.586 | 0.6 | 0.635 | 0.59053 | 0.59256
| 1 | 5 | 0.5736 | 0.579 | 0.60516 | 0.601 | 0.59 | 0.574 | 0.595 | 0.628 | 0.597 | 0.587 | 0.58035 | 0.5084 | 0.542 | 0.574 | 0.6194 | 0.601 | 0.577 | 0.6 | 0.655 | 0.60874 | 62
| 1 | 6 | 0.5821 | 0.5864 | 0.6056 | 0.711 | 0.58 | 0.579 | 0.606 | 0.631 | 0.594 | 0.601 | 0.5783 | 0.5089 | 0.558 | 0.564 | 0.6283 | 0.594 | 0.586 | 0.61 | 0.64 | 0.59969
| Mean |  | 0.58077 | 0.58407 | 0.60533 | 0.65533 | 0.58833 | 0.5835 | 0.5995 | 0.62733 | 0.59717 | 0.59533 | 0.56933 | 0.5145 | 0.55867 | 0.57617 | 0.63542 | 0.595 | 0.58417 | 0.595 | 0.64417 | 0.59238
| Median |  | 0.5831 | 0.58659 | 0.60559 | 0.6455 | 0.59 | 0.5815 | 0.601 | 0.628 | 0.5955 | 0.597 | 0.57932 | 0.51325 | 0.5555 | 0.576 | 0.6359 | 0.5945 | 0.586 | 0.595 | 0.6425 | 0.59662
| Std Dev. |  | 0.012601 | 0.0071098 | 0.0017946 | 0.042297 | 0.0075277 | 0.0088938 | 0.005648 | 0.0040825 | 0.007679 | 0.0059217 | 0.027319 | 0.007024 | 0.014501 | 0.011805 | 0.013085 | 0.0061644 | 0.015993 | 0.010488 | 0.0073598 | 0.016067
| Rel.Std.Dev. |  | 2.17% | 1.217% | 0.2965% | 6.454% | 1.28% | 1.524% | 0.9421% | 0.6508% | 1.286% | 0.9947% | 4.798% | 1.365% | 2.596% | 2.049% | 2.059% | 1.036% | 2.738% | 1.763% | 1.143% | 2.712%
| PDM3 |  | -1.99% | -1.433% | 2.155% | 10.59% | -0.7133% | -1.529% | 1.171% | 5.868% | 0.7774% | 0.468% | -3.92% | -13.17% | -5.72% | -2.767% | 7.232% | 0.4117% | -1.417% | 0.4117% | 8.709% | -0.03074%
| Z-Score (Absolute) |  | 0.8 | 0.59 | 0.72 | 3.81 | 0.33 | 0.63 | 0.36 | 2.08 | 0.21 | 0.1 | 1.5 | 4.89 | 2.16 | 1.08 | 2.58 | 0.08 | 0.59 | 0.08 | 3.12 | 0.08
| Analytical results for Tl in OREAS 251c (Certified Value 0.52 ppm) | [21 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 0.53 | 0.5414 | 0.53 | 0.52 | 0.51 | 0.56 | 0.46 | 0.53 | 0.5474 | 0.487 | 0.54 | 0.52 | 0.55 | 0.51 | 0.48 | 0.5 | 0.331 |  |  |  | 1
| 1 | 2 | 0.55 | 0.53271 | 0.51 | 0.5 | 0.51 | 0.58 | 0.44 | 0.56 | 0.53601 | 0.503 | 0.54 | 0.48 | 0.55 | 0.54 | 0.51 | 0.5 |  |  |  |  | 31
| 1 | 3 | 0.56 | 0.54534 | 0.52 | 0.52 | 0.54 | 0.52 | 0.47 | 0.57 | 0.53214 | 0.498 | 0.53 | 0.46 | 0.54 | 0.53 | 0.5 | 0.5 | 0.418 |  |  |  | 16
| 1 | 4 | 0.54 | 0.518 | 0.53 | 0.52 | 0.52 | 0.55 | 0.47 | 0.56 | 0.55876 | 0.497 | 0.54 | 0.5 | 0.55 | 0.53 | 0.49 | 0.5 |  |  |  |  | 0.52146
| 1 | 5 | 0.54 | 0.55168 | 0.52 | 0.51 | 0.53 | 0.55 | 0.47 | 0.56 | 0.52242 | 0.482 | 0.52 | 0.5 | 0.54 | 0.53 | 0.49 | 0.5 | 0.041 |  |  |  | 63
| 1 | 6 | 0.53 | 0.55118 | 0.53 | 0.51 | 0.52 | 0.55 | 0.44 | 0.54 | 0.5451 | 0.462 | 0.54 | 0.49 | 0.53 | 0.51 | 0.5 | 0.5
| Mean |  | 0.54167 | 0.54005 | 0.52333 | 0.51333 | 0.52167 | 0.55167 | 0.45833 | 0.55333 | 0.54031 | 0.48817 | 0.535 | 0.49167 | 0.54333 | 0.525 | 0.495 | 0.5 | 0.26333
| Median |  | 0.54 | 0.54337 | 0.525 | 0.515 | 0.52 | 0.55 | 0.465 | 0.56 | 0.54056 | 0.492 | 0.54 | 0.495 | 0.545 | 0.53 | 0.495 | 0.5 | 0.331
| Std Dev. |  | 0.01169 | 0.012878 | 0.008165 | 0.008165 | 0.01169 | 0.019408 | 0.01472 | 0.015055 | 0.012801 | 0.014959 | 0.0083666 | 0.020412 | 0.008165 | 0.012247 | 0.010488 | 0 | 0.1974
| Rel.Std.Dev. |  | 2.158% | 2.385% | 1.56% | 1.591% | 2.241% | 3.518% | 3.212% | 2.721% | 2.369% | 3.064% | 1.564% | 4.152% | 1.503% | 2.333% | 2.119% | 0% | 74.96%
| PDM3 |  | 3.876% | 3.565% | 0.3598% | -1.558% | 0.0402% | 5.793% | -12.11% | 6.113% | 3.615% | -6.384% | 2.597% | -5.713% | 4.195% | 0.6794% | -5.074% | -4.115% | -49.5%
| Z-Score (Absolute) |  | 0.64 | 0.58 | 0.03 | 0.4 | 0.09 | 1.01 | 2.42 | 1.07 | 0.59 | 1.32 | 0.4 | 1.2 | 0.7 | 0.03 | 1.07 | NA | 9.59
Results from laboratory 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 251c (Certified Value 0.25 ppm) | [21 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.26 | 0.2529 | 0.32049 | 0.3 | 0.2 | 0.3 | 0.2 | 0.25 | 0.25 | [11 empty] | 1
| 1 | 2 | 0.25 | 0.25613 | 0.30348 | 0.3 | 0.2 | 0.28 | 0.2 | 0.25 | 0.25 | [11 empty] | 6
| 1 | 3 | 0.26 | 0.24695 | 0.29205 | 0.3 | 0.2 | 0.28 | 0.2 | 0.24 | 0.25 | [11 empty] | 16
| 1 | 4 | 0.25 | 0.24175 | 0.28971 | 0.3 | 0.2 | 0.28 | 0.2 | 0.25 | 0.25 | [11 empty] | 0.25419
| 1 | 5 | 0.25 | 0.24698 | 0.28746 | 0.3 | 0.2 | 0.28 | 0.2 | 0.25 | 0.25 | [11 empty] | 64
| 1 | 6 | 0.27 | 0.25424 | 0.3042 | 0.3 | 0.2 | 0.28 | 0.2 | 0.25 | 0.25
| Mean |  | 0.25667 | 0.24982 | 0.29957 | 0.3 | 0.2 | 0.28333 | 0.2 | 0.24833 | 0.25
| Median |  | 0.255 | 0.24994 | 0.29777 | 0.3 | 0.2 | 0.28 | 0.2 | 0.25 | 0.25
| Std Dev. |  | 0.008165 | 0.0054811 | 0.012448 | 0 | 3.0405e-17 | 0.008165 | 3.0405e-17 | 0.0040825 | 0
| Rel.Std.Dev. |  | 3.181% | 2.194% | 4.155% | 0% | 1.52e-14% | 2.882% | 1.52e-14% | 1.644% | 0%
| PDM3 |  | 0.9738% | -1.718% | 17.85% | 18.02% | -21.32% | 11.46% | -21.32% | -2.305% | -1.649%
| Z-Score (Absolute) |  | 0.13 | 0 | 1 | 1.01 | 1.01 | 0.67 | 1.01 | 0.03 | 0
| Analytical results for U in OREAS 251c (Certified Value 2.18 ppm) | [21 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.29 | 2.3277 | 2.159 | 2.3 | 2.2 | 2 | 2.4 | 2 | 2.3 | 2.2571 | 1.91 | 2.2 | 2 | 2.43 | 2 | 2.13 | 2.5 |  |  |  | 1
| 1 | 2 | 2.19 | 2.3072 | 2.195 | 2.3 | 2.1 | 2 | 2.3 | 2 | 2.4 | 2.157 | 1.93 | 2.31 | 2 | 2.4 | 2 | 2.17 | 2.5 |  |  |  | 33
| 1 | 3 | 2.19 | 2.3442 | 2.1132 | 2.3 | 2.2 | 2.1 | 2.3 | 2.1 | 2.3 | 2.1528 | 1.9 | 2.25 | 4 | 2.42 | 2 | 2.1 | 2.4 |  |  |  | 16
| 1 | 4 | 2.21 | 2.24 | 2.1787 | 2.3 | 2.1 | 2.1 | 2.2 | 1.9 | 2.3 | 2.2556 | 1.89 | 2.26 | 1.8 | 2.45 | 2.1 | 2.09 | 2.5 |  |  |  | 2.1802
| 1 | 5 | 2.21 | 2.2937 | 2.1383 | 2.3 | 2.1 | 2 | 2.3 | 2 | 2.3 | 2.2463 | 1.88 | 2.22 | 1.9 | 2.43 | 2 | 2.13 | 2.5 |  |  |  | 65
| 1 | 6 | 2.22 | 2.2895 | 2.1347 | 2.3 | 2.2 | 2 | 2.2 | 2.2 | 2.2 | 2.2599 | 1.87 | 2.22 | 1.9 | 2.43 | 2 | 2.07 | 2.4
| Mean |  | 2.2183 | 2.3004 | 2.1532 | 2.3 | 2.15 | 2.0333 | 2.2833 | 2.0333 | 2.3 | 2.2214 | 1.8967 | 2.2433 | 2.2667 | 2.4267 | 2.0167 | 2.115 | 2.4667
| Median |  | 2.21 | 2.3004 | 2.1487 | 2.3 | 2.15 | 2 | 2.3 | 2 | 2.3 | 2.2509 | 1.895 | 2.235 | 1.95 | 2.43 | 2 | 2.115 | 2.5
| Std Dev. |  | 0.037103 | 0.036117 | 0.030314 | 4.8648e-16 | 0.054772 | 0.05164 | 0.075277 | 0.10328 | 0.063246 | 0.051742 | 0.021602 | 0.039328 | 0.85245 | 0.01633 | 0.040825 | 0.035637 | 0.05164
| Rel.Std.Dev. |  | 1.673% | 1.57% | 1.408% | 2.115e-14% | 2.548% | 2.54% | 3.297% | 5.079% | 2.75% | 2.329% | 1.139% | 1.753% | 37.61% | 0.6729% | 2.024% | 1.685% | 2.094%
| PDM3 |  | 1.747% | 5.51% | -1.241% | 5.493% | -1.387% | -6.738% | 4.729% | -6.738% | 5.493% | 1.89% | -13.01% | 2.894% | 3.964% | 11.3% | -7.502% | -2.992% | 13.14%
| Z-Score (Absolute) |  | 0.03 | 0.68 | 0.59 | 0.67 | 0.61 | 1.61 | 0.53 | 1.61 | 0.67 | 0 | 2.79 | 0.19 | 0.39 | 1.76 | 1.76 | 0.91 | 2.1
| Analytical results for V in OREAS 251c (Certified Value 92 ppm) | [21 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 92 | 96.5275 | 91.8015 | 35 | 94 | 89 | 89 | 95 | 89 | 89 | 95.3283 | 89 | 94 | 97 | 96 | 91 | 78 | 89.99 | 90 | 258.128 | 1
| 1 | 2 | 89 | 94.7801 | 90.531 | 37 | 93 | 89 | 89 | 97 | 89 | 89 | 95.0519 | 92 | 94 | 97 | 99 | 93 | 77 | 92 | 90 | 91.368 | 34
| 1 | 3 | 90 | 95.0812 | 95.445 | 30 | 91 | 88 | 91 | 95 | 91 | 88 | 96.5342 | 91 | 87 | 97 | 95 | 95 | 79 | 90.3 | 90 | 91.43 | 16
| 1 | 4 | 90 | 92.0686 | 92.841 | 49 | 92 | 87 | 91 | 93 | 91 | 87 | 97.4788 | 92 | 91 | 95 | 94 | 94 | 75 | 95.1 | 95 | 91.775 | 91.9751
| 1 | 5 | 88 | 93.1489 | 96.478 | 30 | 93 | 87 | 90 | 95 | 91 | 87 | 92.2758 | 89 | 89 | 97 | 95 | 96 | 81 | 89.2 | 90 | 91.203 | 66
| 1 | 6 | 92 | 94.7863 | 91.203 | 36 | 93 | 88 | 91 | 94 | 91 | 90 | 92.6725 | 88 | 90 | 95 | 95 | 93 | 78 | 91.8 | 90 | 90.801
| Mean |  | 90.1667 | 94.3988 | 93.0499 | 36.1667 | 92.6667 | 88 | 90.1667 | 94.8333 | 90.3333 | 88.3333 | 94.8902 | 90.1667 | 90.8333 | 96.3333 | 95.6667 | 93.6667 | 78 | 91.3983 | 90.8333 | 119.1175
| Median |  | 90 | 94.7832 | 92.3213 | 35.5 | 93 | 88 | 90.5 | 95 | 91 | 88.5 | 95.1901 | 90 | 90.5 | 97 | 95 | 93.5 | 78 | 91.05 | 90 | 91.399
| Std Dev. |  | 1.6021 | 1.5675 | 2.4019 | 6.969 | 1.0328 | 0.89443 | 0.98319 | 1.3292 | 1.0328 | 1.2111 | 2.0676 | 1.7224 | 2.7869 | 1.0328 | 1.7512 | 1.7512 | 2 | 2.1094 | 2.0412 | 68.1017
| Rel.Std.Dev. |  | 1.777% | 1.66% | 2.581% | 19.27% | 1.115% | 1.016% | 1.09% | 1.402% | 1.143% | 1.371% | 2.179% | 1.91% | 3.068% | 1.072% | 1.831% | 1.87% | 2.564% | 2.308% | 2.247% | 57.17%
| PDM3 |  | -1.966% | 2.635% | 1.169% | -60.68% | 0.7519% | -4.322% | -1.966% | 3.108% | -1.785% | -3.96% | 3.169% | -1.966% | -1.241% | 4.738% | 4.014% | 1.839% | -15.19% | -0.6271% | -1.241% | 29.51%
| Z-Score (Absolute) |  | 0.24 | 0.83 | 0.49 | 13.89 | 0.39 | 0.79 | 0.24 | 0.94 | 0.2 | 0.7 | 0.95 | 0.24 | 0.07 | 1.32 | 1.15 | 0.65 | 3.32 | 0.07 | 0.07 | 7.08
| Analytical results for W in OREAS 251c (Certified Value 31.5 ppm) | [21 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 32.7 | 33.4617 | 31.2 | 24.7 | 34 | 31.7 | 33.8 | 30.7 | 32.4 | 27.0103 | 26.79 | 30.7 | 32.5 | 33.3 | 30.7 | 36.5 | 18.64 |  |  |  | 1
| 1 | 2 | 32.8 | 32.2151 | 31.7 | 23.5 | 32.8 | 32.1 | 36 | 28.7 | 34.4 | 23.7993 | 25.98 | 30.2 | 29.1 | 31.9 | 30.8 | 35 | 18.84 |  |  |  | 35
| 1 | 3 | 31.8 | 33.3909 | 30.5 | 22.9 | 34 | 32.3 | 34.7 | 32.5 | 33.7 | 23.4028 | 25.87 | 29.2 | 29.8 | 32.4 | 29.4 | 32.7 | 18.3 |  |  |  | 16
| 1 | 4 | 31.5 | 31.6736 | 31.4 | 23.4 | 34 | 32.9 | 31.6 | 28.4 | 32.6 | 28.6739 | 29.31 | 31 | 29.5 | 33.1 | 31.7 | 32.6 | 18.28 |  |  |  | 31.4592
| 1 | 5 | 33.1 | 31.9246 | 31.5 | 24.3 | 36.2 | 32.5 | 35.8 | 30.9 | 33.2 | 28.78 | 28.4 | 29.9 | 32 | 33.8 | 30.4 | 33.1 | 19.13 |  |  |  | 67
| 1 | 6 | 33.1 | 32.8162 | 32.1 | 24.2 | 33.3 | 31.3 | 34.2 | 30.8 | 32.2 | 24.0491 | 27.5 | 30.1 | 30.6 | 33 | 29.6 | 34.1 | 18.07
| Mean |  | 32.5 | 32.5804 | 31.4 | 23.8333 | 34.05 | 32.1333 | 34.35 | 30.3333 | 33.0833 | 25.9526 | 27.3083 | 30.1833 | 30.5833 | 32.9167 | 30.4333 | 34 | 18.5433
| Median |  | 32.75 | 32.5156 | 31.45 | 23.85 | 34 | 32.2 | 34.45 | 30.75 | 32.9 | 25.5297 | 27.145 | 30.15 | 30.2 | 33.05 | 30.55 | 33.6 | 18.47
| Std Dev. |  | 0.68411 | 0.75829 | 0.53666 | 0.67429 | 1.1623 | 0.57155 | 1.6022 | 1.5345 | 0.84951 | 2.5012 | 1.3665 | 0.63061 | 1.3906 | 0.67355 | 0.84538 | 1.531 | 0.39853
| Rel.Std.Dev. |  | 2.105% | 2.327% | 1.709% | 2.829% | 3.414% | 1.779% | 4.664% | 5.059% | 2.568% | 9.637% | 5.004% | 2.089% | 4.547% | 2.046% | 2.778% | 4.503% | 2.149%
| PDM3 |  | 3.308% | 3.564% | -0.1882% | -24.24% | 8.235% | 2.143% | 9.189% | -3.579% | 5.163% | -17.5% | -13.19% | -4.056% | -2.784% | 4.633% | -3.261% | 8.077% | -41.06%
| Z-Score (Absolute) |  | 0.49 | 0.52 | 0 | 3.36 | 1.18 | 0.33 | 1.31 | 0.47 | 0.75 | 2.42 | 1.82 | 0.54 | 0.36 | 0.67 | 0.43 | 1.16 | 5.72
| Analytical results for Y in OREAS 251c (Certified Value 18.3 ppm) | [21 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 19.75 | 18.8404 | 20.7749 | 19.7 | 16.7 | 19.2 | 18 | 20.5 | 25.3 | 17.9 | 19.0545 | 16.26 | 17.8 | 17.3 | 19.43 | 16.7 | 19.3 | 18.5 | 19.061 |  | 1
| 1 | 2 | 18.65 | 18.7915 | 20.6696 | 19.2 | 15.7 | 18.9 | 17.3 | 19 | 17 | 17.9 | 18.4286 | 16.6 | 17.8 | 16.7 | 19.31 | 17.3 | 19.3 | 18.4 | 16.356 |  | 36
| 1 | 3 | 18.32 | 19.2166 | 20.8038 | 19.9 | 15.9 | 19.8 | 18.8 | 17.8 | 17.8 | 18.7 | 19.3213 | 16.5 | 17.2 | 16.9 | 19.2 | 17.2 | 19.5 | 18.5 | 17.345 |  | 16
| 1 | 4 | 18.37 | 17.828 | 20.8956 | 20.1 | 15.4 | 19.5 | 18.7 | 17.4 | 16.6 | 18 | 19.2593 | 16.67 | 17.8 | 16.2 | 19.42 | 17.3 | 19.3 | 18.7 | 16.967 |  | 18.3027
| 1 | 5 | 19.83 | 18.7805 | 20.8173 | 20.1 | 16.7 | 19.2 | 17.5 | 17.7 | 19.4 | 18.4 | 18.5567 | 16.3 | 17 | 16.3 | 20.29 | 16.9 | 19.8 | 18.7 | 17.582 |  | 68
| 1 | 6 | 19.21 | 18.6753 | 20.8027 | 20.1 | 17.1 | 18.8 | 18.1 | 18 | 18.4 | 18.7 | 19.1105 | 16.08 | 17.8 | 16.3 | 19.41 | 18.5 | 19.1 | 18.5 | 17.841
| Mean |  | 19.0217 | 18.6887 | 20.794 | 19.85 | 16.25 | 19.2333 | 18.0667 | 18.4 | 19.0833 | 18.2667 | 18.9552 | 16.4017 | 17.5667 | 16.6167 | 19.51 | 17.3167 | 19.3833 | 18.55 | 17.5253
| Median |  | 18.93 | 18.786 | 20.8032 | 20 | 16.3 | 19.2 | 18.05 | 17.9 | 18.1 | 18.2 | 19.0825 | 16.4 | 17.8 | 16.5 | 19.415 | 17.25 | 19.3 | 18.5 | 17.4635
| Std Dev. |  | 0.67446 | 0.46081 | 0.073318 | 0.35637 | 0.67454 | 0.37238 | 0.60882 | 1.1645 | 3.2053 | 0.38297 | 0.37325 | 0.22578 | 0.36697 | 0.43089 | 0.39217 | 0.62743 | 0.24014 | 0.12247 | 0.91353
| Rel.Std.Dev. |  | 3.546% | 2.466% | 0.3526% | 1.795% | 4.151% | 1.936% | 3.37% | 6.329% | 16.8% | 2.097% | 1.969% | 1.377% | 2.089% | 2.593% | 2.01% | 3.623% | 1.239% | 0.6602% | 5.213%
| PDM3 |  | 3.928% | 2.109% | 13.61% | 8.454% | -11.22% | 5.084% | -1.29% | 0.5314% | 4.265% | -0.1971% | 3.565% | -10.39% | -4.022% | -9.212% | 6.596% | -5.388% | 5.904% | 1.351% | -4.247%
| Z-Score (Absolute) |  | 0.38 | 0.11 | 1.82 | 1.05 | 1.86 | 0.55 | 0.39 | 0.12 | 0.43 | 0.23 | 0.33 | 1.74 | 0.8 | 1.56 | 0.78 | 1 | 0.67 | 0 | 0.83
| Analytical results for Yb in OREAS 251c (Certified Value 1.72 ppm) | [21 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
3
| 1 | 1 | 1.75 | 1.6449 | 1.7488 | 1.9 | 1.6 | 1.82 | 1.6 | 1.79 | 1.9 | [11 empty] | 1
| 1 | 2 | 1.93 | 1.5697 | 1.6647 | 1.8 | 1.5 | 1.75 | 1.6 | 1.58 | 1.95 | [11 empty] | 7
| 1 | 3 | 2.02 | 1.6345 | 1.7015 | 1.9 | 1.5 | 1.78 | 1.5 | 1.66 | 1.95 | [11 empty] | 16
| 1 | 4 | 1.81 | 1.5987 | 1.6911 | 1.9 | 1.5 | 1.81 | 1.6 | 1.62 | 1.9 | [11 empty] | 1.7247
| 1 | 5 | 1.77 | 1.6417 | 1.7165 | 1.9 | 1.6 | 1.75 | 1.6 | 1.67 | 1.9 | [11 empty] | 69
| 1 | 6 | 1.65 | 1.655 | 1.709 | 2 | 1.5 | 1.83 | 1.5 | 1.67 | 1.9
| Mean |  | 1.8217 | 1.6241 | 1.7053 | 1.9 | 1.5333 | 1.79 | 1.5667 | 1.665 | 1.9167
| Median |  | 1.79 | 1.6381 | 1.7053 | 1.9 | 1.5 | 1.795 | 1.6 | 1.665 | 1.9
| Std Dev. |  | 0.13303 | 0.032875 | 0.027902 | 0.063246 | 0.05164 | 0.035214 | 0.05164 | 0.07064 | 0.02582
| Rel.Std.Dev. |  | 7.303% | 2.024% | 1.636% | 3.329% | 3.368% | 1.967% | 3.296% | 4.243% | 1.347%
| PDM3 |  | 5.62% | -5.836% | -1.129% | 10.16% | -11.1% | 3.784% | -9.165% | -3.464% | 11.13%
| Z-Score (Absolute) |  | 0.67 | 0.47 | 0 | 1.13 | 1 | 0.49 | 0.8 | 0.23 | 1.22
| Analytical results for Zn in OREAS 251c (Certified Value 69 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 0
1
| 1 | 1 | 70 | 71.3637 | 70.584 | 71 | 67 | 68 | 68 | 74 | 66 | 64 | 71.4555 | 63 | 70 | 70.8 | 73 | 65.4 | 59.1 | 83.27 | 72 | 110.304 | 1
| 1 | 2 | 72 | 69.7508 | 70.092 | 71 | 73 | 68 | 69 | 74 | 67 | 64 | 69.4208 | 64 | 70 | 68.6 | 70 | 67.1 | 57.5 | 90.48 | 74 | 67.11 | 16
| 1 | 3 | 70 | 69.6094 | 71.268 | 70 | 69 | 67 | 70 | 73 | 68 | 64 | 70.6578 | 65 | 68 | 70.8 | 72 | 67.4 | 59 | 98.65 | 76 | 67.835 | 16
| 1 | 4 | 71 | 66.8552 | 70.464 | 70 | 69 | 67 | 69 | 72 | 67 | 63 | 68.7782 | 65 | 69 | 71 | 71 | 67.4 | 56.8 | 91.2 | 74 | 69.217 | 69.4913
| 1 | 5 | 70 | 68.1592 | 71.088 | 70 | 72 | 66 | 69 | 73 | 67 | 63 | 71.2027 | 63 | 68 | 70.5 | 70 | 66.7 | 57.7 | 86.27 | 76 | 68.477 | 70
| 1 | 6 | 74 | 70.0571 | 70.488 | 71 | 69 | 67 | 69 | 73 | 68 | 65 | 69.636 | 62 | 71 | 68.5 | 70 | 65.7 | 56.9 | 83.74 | 74 | 67.6
| Mean |  | 71.1667 | 69.2992 | 70.664 | 70.5 | 69.8333 | 67.1667 | 69 | 73.1667 | 67.1667 | 63.8333 | 70.1918 | 63.6667 | 69.3333 | 70.0333 | 71 | 66.6167 | 57.8333 | 88.935 | 74.3333 | 75.0905
| Median |  | 70.5 | 69.6801 | 70.536 | 70.5 | 69 | 67 | 69 | 73 | 67 | 64 | 70.1469 | 63.5 | 69.5 | 70.65 | 70.5 | 66.9 | 57.6 | 88.375 | 74 | 68.156
| Std Dev. |  | 1.6021 | 1.5754 | 0.4357 | 0.54772 | 2.2286 | 0.75277 | 0.63246 | 0.75277 | 0.75277 | 0.75277 | 1.0712 | 1.2111 | 1.2111 | 1.1605 | 1.2649 | 0.87044 | 1.0033 | 5.7979 | 1.5055 | 17.2665
| Rel.Std.Dev. |  | 2.251% | 2.273% | 0.6166% | 0.7769% | 3.191% | 1.121% | 0.9166% | 1.029% | 1.121% | 1.179% | 1.526% | 1.902% | 1.747% | 1.657% | 1.782% | 1.307% | 1.735% | 6.519% | 2.025% | 22.99%
| PDM3 |  | 2.411% | -0.2765% | 1.687% | 1.451% | 0.4921% | -3.345% | -0.7071% | 5.289% | -3.345% | -8.142% | 1.008% | -8.382% | -0.2274% | 0.7799% | 2.171% | -4.137% | -16.78% | 27.98% | 6.968% | 8.057%
| Z-Score (Absolute) |  | 0.42 | 0.21 | 0.25 | 0.19 | 0.03 | 0.93 | 0.31 | 1.09 | 0.93 | 2.06 | 0.09 | 2.11 | 0.2 | 0.03 | 0.36 | 1.12 | 4.08 | 6.41 | 1.48 | 1.74
| Analytical results for Zr in OREAS 251c (Certified Value 183 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 187.4 | 180.9976 | 186.0785 | 121 | 190 | 195.5 | 187.5 | 183.5 | 192.5 | 171 | 181.8685 | 163.5 | 191 | 185 | 198.9 | 174.1 | 174 | 161.1 | 185 |  | 1
| 1 | 2 | 191 | 180.5101 | 183.968 | 106 | 185 | 196 | 187.5 | 187 | 177 | 168.5 | 174.1331 | 166.8 | 189 | 184.7 | 192.1 | 178.9 | 181 | 158.4 | 185 |  | 17
| 1 | 3 | 185.9 | 178.8111 | 188.4305 | 48 | 184 | 195.5 | 200 | 183.5 | 187.5 | 188 | 190.3071 | 165.2 | 186 | 185.7 | 194.3 | 179.6 | 179 | 157.2 | 180 |  | 16
| 1 | 4 | 184.3 | 171.6795 | 187.8635 | 144 | 180 | 190 | 193 | 177 | 174.5 | 169.5 | 191.5803 | 167.1 | 190 | 180.1 | 197 | 180.5 | 178 | 159.2 | 175 |  | 182.9031
| 1 | 5 | 183.4 | 174.7616 | 185.9945 | 92 | 189 | 192 | 191 | 180 | 182.5 | 171.5 | 194.8047 | 161.5 | 184 | 184 | 194.7 | 177.8 | 177 | 156.7 | 180 |  | 71
| 1 | 6 | 189.1 | 178.462 | 184.43 | 95 | 186 | 190.5 | 189 | 183.5 | 184 | 175.5 | 177.2345 | 162.8 | 187 | 180.7 | 191.6 | 173.3 | 176 | 168.8 | 185
| Mean |  | 186.85 | 177.537 | 186.1275 | 101 | 185.6667 | 193.25 | 191.3333 | 182.4167 | 183 | 174 | 184.988 | 164.4833 | 187.8333 | 183.3667 | 194.7667 | 177.3667 | 177.5 | 160.2333 | 181.6667
| Median |  | 186.65 | 178.6365 | 186.0365 | 100.5 | 185.5 | 193.75 | 190 | 183.5 | 183.25 | 171.25 | 186.0878 | 164.35 | 188 | 184.35 | 194.5 | 178.35 | 177.5 | 158.8 | 182.5
| Std Dev. |  | 2.8947 | 3.6141 | 1.7823 | 32.249 | 3.6148 | 2.734 | 4.7504 | 3.4557 | 6.6332 | 7.2664 | 8.4356 | 2.2551 | 2.6394 | 2.3695 | 2.8083 | 2.9851 | 2.429 | 4.4778 | 4.0825
| Rel.Std.Dev. |  | 1.549% | 2.036% | 0.9576% | 31.93% | 1.947% | 1.415% | 2.483% | 1.894% | 3.625% | 4.176% | 4.56% | 1.371% | 1.405% | 1.292% | 1.442% | 1.683% | 1.368% | 2.795% | 2.247%
| PDM3 |  | 2.158% | -2.934% | 1.763% | -44.78% | 1.511% | 5.657% | 4.609% | -0.2659% | 0.05299% | -4.868% | 1.14% | -10.07% | 2.696% | 0.2535% | 6.486% | -3.027% | -2.954% | -12.39% | -0.676%
| Z-Score (Absolute) |  | 0.48 | 0.67 | 0.39 | 10.12 | 0.33 | 1.27 | 1.03 | 0.07 | 0 | 1.11 | 0.25 | 2.29 | 0.6 | 0.05 | 1.45 | 0.7 | 0.68 | 2.81 | 0.16
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C471: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C507: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C706: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C724: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C742: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C760: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C833: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C851: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C869: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C888: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C906: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C924: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C943: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 251c (Certified Value 0.269 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.24 | 0.27787 | 0.22384 | 0.304 | < 0.5 | 0.28 | 0.29 | 0.25 | 0.28 | 0.28 | 0.29087 | < 0.3 | 0.2 | 0.243 | 0.262 | 0.28 | < 0.5 | 0.3 | 0.193 |  |  | 1
| 1 | 2 | 0.26 | 0.27337 | 0.091627 | 0.302 | < 0.5 | 0.28 | 0.28 | 0.25 | 0.27 | 0.25 | 0.29183 | < 0.3 | 0.21 | 0.249 | 0.257 | 0.26 | < 0.5 | 0.3 | 0.333 |  |  | 23
| 1 | 3 | 0.23 | 0.25593 | 0.099182 | 0.303 | < 0.5 | 0.29 | 0.3 | 0.26 | 0.27 | 0.26 | 0.30439 | < 0.3 | 0.19 | 0.276 | 0.267 | 0.29 | < 0.5 | 0.3 | 0.113 |  |  | 16
| 1 | 4 | 0.25 | 0.25273 | 0.099392 | 0.311 | < 0.5 | 0.26 | 0.29 | 0.25 | 0.26 | 0.26 | 0.30887 | < 0.3 | 0.2 | 0.253 | 0.266 | 0.29 | < 0.5 | 0.3 | 0.176 |  |  | 0.26894
| 1 | 5 | 0.25 | 0.28089 | 0.33322 | 0.319 | < 0.5 | 0.29 | 0.28 | 0.26 | 0.27 | 0.28 | 0.2976 | < 0.3 | 0.21 | 0.242 | 0.271 | 0.25 | < 0.5 | 0.3 | 0.223 |  |  | 73
| 1 | 6 | 0.24 | 0.28425 | 0.15626 | 0.295 | < 0.5 | 0.27 | 0.29 | 0.26 | 0.26 | 0.27 | 0.30002 | < 0.3 | 0.19 | 0.228 | 0.264 | 0.28 | < 0.5 | 0.25 | 0.079
| Mean |  | 0.245 | 0.27084 | 0.16725 | 0.30567 |  | 0.27833 | 0.28833 | 0.255 | 0.26833 | 0.26667 | 0.29893 |  | 0.2 | 0.2485 | 0.2645 | 0.275 |  | 0.29167 | 0.18617
| Median |  | 0.245 | 0.27562 | 0.12783 | 0.3035 |  | 0.28 | 0.29 | 0.255 | 0.27 | 0.265 | 0.29881 |  | 0.2 | 0.246 | 0.265 | 0.28 |  | 0.3 | 0.1845
| Std Dev. |  | 0.010488 | 0.013319 | 0.095736 | 0.0082865 |  | 0.01169 | 0.0075277 | 0.0054772 | 0.0075277 | 0.012111 | 0.0070258 |  | 0.0089443 | 0.015934 | 0.0047645 | 0.016432 |  | 0.020412 | 0.089336
| Rel.Std.Dev. |  | 4.281% | 4.918% | 57.24% | 2.711% |  | 4.2% | 2.611% | 2.148% | 2.805% | 4.541% | 2.35% |  | 4.472% | 6.412% | 1.801% | 5.975% |  | 6.999% | 47.99%
| PDM3 |  | -8.9% | 0.7085% | -37.81% | 13.66% |  | 3.494% | 7.213% | -5.182% | -0.2241% | -0.8438% | 11.15% |  | -25.63% | -7.599% | -1.649% | 2.255% |  | 8.452% | -30.78%
| Z-Score (Absolute) |  | 0.82 | 0.27 | 4.1 | 1.73 | 0.61 | 0.58 | 1 | 0.4 | 0.16 | 0.09 | 1.45 | 4.83 | 2.72 | 0.67 | 0 | 0.44 | 0.61 | 1.14 | 3.3
| Analytical results for Al in OREAS 251c (Certified Value 1.13 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 3
| 1 | 1 | 1.1409 | 1.2266 | 1.46 | 1.22 | 1.09 | 1.22 | 1.06 | 1.11 | 1.08 | 1.135 | 1.54 | 1.08 | 1.4239 | 0.94 | 1.78 | 1.06 | 1.17 | 1.1528 |  |  |  | 1
| 1 | 2 | 1.1671 | 1.2222 | 1.49 | 1.22 | 1.16 | 1.23 | 1.1 | 1.11 | 1.07 | 1.1154 | 1.62 | 1.17 | 1.3867 | 0.89 | 2.05 | 1.03 | 1.15 | 1.1857 |  |  |  | #N/A
| 1 | 3 | 1.1509 | 1.1925 | 1.37 | 1.13 | 1.13 | 1.23 | 1.07 | 1.1 | 1.06 | 1.1452 | 1.65 | 1.16 | 1.373 | 1 | 1.81 | 1.04 | 1.16 | 1.2198 |  |  |  | 16
| 1 | 4 | 1.1514 | 1.1766 | 1.43 | 1.18 | 1.1 | 1.24 | 1.1 | 1.12 | 1.08 | 1.1974 | 1.56 | 1.13 | 1.3791 | 1 | 2.08 | 1.09 | 1.16 | 1.0624 |  |  |  | 1.126
| 1 | 5 | 1.152 | 1.2081 | 1.4 | 1.22 | 1.12 | 1.25 | 1.1 | 1.08 | 1.08 | 1.1934 | 1.66 | 1.04 | 1.4421 | 0.95 | 1.86 | 1.2 | 1.15 | 1.1579 |  |  |  | 74
| 1 | 6 | 1.1626 | 1.2265 | 1.39 | 1.17 | 1.1 | 1.25 | 1.1 | 1.08 | 1.08 | 1.1737 | 1.63 | 1.04 | 1.4416 | 0.98 | 2.13 | 1.04 | 1.13 | 1.1011
| Mean |  | 1.1541 | 1.2087 | 1.4233 | 1.19 | 1.1167 | 1.2367 | 1.0883 | 1.1 | 1.075 | 1.16 | 1.61 | 1.1033 | 1.4077 | 0.96 | 1.9517 | 1.0767 | 1.1533 | 1.1466
| Median |  | 1.1517 | 1.2151 | 1.415 | 1.2 | 1.11 | 1.235 | 1.1 | 1.105 | 1.08 | 1.1594 | 1.625 | 1.105 | 1.4053 | 0.965 | 1.955 | 1.05 | 1.155 | 1.1553
| Std Dev. |  | 0.0093532 | 0.020554 | 0.045461 | 0.036878 | 0.02582 | 0.012111 | 0.018348 | 0.016733 | 0.0083666 | 0.033263 | 0.04899 | 0.058195 | 0.031806 | 0.042426 | 0.15224 | 0.064083 | 0.013663 | 0.056923
| Rel.Std.Dev. |  | 0.8104% | 1.7% | 3.194% | 3.099% | 2.312% | 0.9793% | 1.686% | 1.521% | 0.7783% | 2.868% | 3.043% | 5.274% | 2.259% | 4.419% | 7.8% | 5.952% | 1.185% | 4.964%
| PDM3 |  | 2.497% | 7.344% | 26.4% | 5.681% | -0.832% | 9.825% | -3.348% | -2.312% | -4.532% | 3.016% | 42.98% | -2.016% | 25.02% | -14.75% | 73.32% | -4.384% | 2.424% | 1.827%
| Z-Score (Absolute) |  | 0 | 0.62 | 3.02 | 0.41 | 0.42 | 0.93 | 0.73 | 0.6 | 0.88 | 0.07 | 5.11 | 0.56 | 2.85 | 2.17 | 8.94 | 0.86 | 0 | 0.08
| Analytical results for As in OREAS 251c (Certified Value 153 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 0
| 1 | 1 | 145 | 153.5792 | 156.558 | 158.845 | 147 | 152 | 157.5 | 159 | 150 | 162 | 153.5 | 152.6212 | 149.2 | 148 | 154.3 | 158 | 148.7 | 156 | 139.4 | 149 | 160.57 | 1
| 1 | 2 | 149 | 155.9065 | 151.257 | 158.988 | 151 | 152 | 162.5 | 159 | 149.5 | 161 | 151 | 155.9743 | 145.7 | 156 | 153.8 | 157 | 144.3 | 158 | 138.5 | 147 | 160.458 | 11
| 1 | 3 | 141 | 154.2475 | 157.932 | 156.183 | 150 | 152 | 159.5 | 160.5 | 151.5 | 157 | 149 | 154.3622 | 143.4 | 155 | 153.5 | 156 | 153.1 | 152 | 117.9 | 153 | 162.998 | 16
| 1 | 4 | 147 | 150.7229 | 145.842 | 160.099 | 147 | 153 | 157.5 | 158 | 150.5 | 162 | 153.5 | 152.7649 | 149.5 | 155 | 151.7 | 156 | 153.4 | 157 | 150.3 | 148 | 161.098 | 153.1763
| 1 | 5 | 152 | 153.5302 | 161.788 | 160.539 | 154 | 153 | 157 | 161.5 | 152 | 161.5 | 154 | 156.1431 | 143.1 | 152 | 154.2 | 156 | 147.1 | 144 | 144.4 | 151 | 161.095 | 75
| 1 | 6 | 146 | 153.8247 | 155.652 | 154.511 | 148 | 153 | 154.5 | 161 | 149 | 159 | 154 | 154.8847 | 143.9 | 153 | 146.6 | 157 | 150.4 | 152 | 136.4 | 149 | 161.034
| Mean |  | 146.6667 | 153.6352 | 154.8382 | 158.1942 | 149.5 | 152.5 | 158.0833 | 159.8333 | 150.4167 | 160.4167 | 152.5 | 154.4584 | 145.8 | 153.1667 | 152.35 | 156.6667 | 149.5 | 153.1667 | 137.8167 | 149.5 | 161.2088
| Median |  | 146.5 | 153.7019 | 156.105 | 158.9165 | 149 | 152.5 | 157.5 | 159.75 | 150.25 | 161.25 | 153.5 | 154.6235 | 144.8 | 154 | 153.65 | 156.5 | 149.55 | 154 | 138.95 | 149 | 161.0645
| Std Dev. |  | 3.7238 | 1.6772 | 5.5722 | 2.3576 | 2.7386 | 0.54772 | 2.691 | 1.3663 | 1.1583 | 2.0104 | 2.0494 | 1.5212 | 2.8955 | 2.9269 | 2.971 | 0.8165 | 3.5333 | 5.1543 | 10.962 | 2.1679 | 0.91965
| Rel.Std.Dev. |  | 2.539% | 1.092% | 3.599% | 1.49% | 1.832% | 0.3592% | 1.702% | 0.8548% | 0.7701% | 1.253% | 1.344% | 0.9849% | 1.986% | 1.911% | 1.95% | 0.5212% | 2.363% | 3.365% | 7.954% | 1.45% | 0.5705%
| PDM3 |  | -4.25% | 0.2996% | 1.085% | 3.276% | -2.4% | -0.4415% | 3.204% | 4.346% | -1.802% | 4.727% | -0.4415% | 0.837% | -4.816% | -0.006261% | -0.5394% | 2.279% | -2.4% | -0.006261% | -10.03% | -2.4% | 5.244%
| Z-Score (Absolute) |  | 1.2 | 0.09 | 0.31 | 0.92 | 0.67 | 0.12 | 0.9 | 1.23 | 0.51 | 1.33 | 0.12 | 0.24 | 1.35 | 0 | 0.15 | 0.64 | 0.67 | 0 | 2.82 | 0.67 | 1.48
| Analytical results for B in OREAS 251c (Certified Value < 10 ppm) | [22 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [9 empty] | 1
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 15g | [9 empty] | 1
| 1 | 1 | 4.8114 | 6 | < 10 | 10 | 10 | 10 | < 10 | < 10 | < 10 | < 20 |  | 4 | [9 empty] | 1
| 1 | 2 | 6.3366 | 6 | < 10 | 10 | 10 | 10 | 10 | < 10 | < 10 | < 20 |  | 4 | [9 empty] | 1
| 1 | 3 | 5.6326 | 6 | < 10 | 10 | 10 | 10 | 10 | < 10 | < 10 | < 20 | 10 | 4 | [9 empty] | 16
| 1 | 4 | 6.4317 | 6 | < 10 | 10 | 10 | 10 | < 10 | < 10 | < 10 | < 20 |  | 4 | [9 empty] | < 10
| 1 | 5 | 5.6185 | 6 | < 10 | 10 | 10 | 10 | < 10 | < 10 | < 10 | < 20 |  | 4 | [9 empty] | 76
| 1 | 6 | 5.2959 | 5 | < 10 | 10 | 10 | 10 | 10 | < 10 | < 10 | < 20 | 11 | 4
| Mean |  | 5.6878 | 5.8333 |  | 10 | 10 | 10 | 10 |  |  |  | 10.5 | 4
| Median |  | 5.6256 | 6 |  | 10 | 10 | 10 | 10 |  |  |  | 10.5 | 4
| Std Dev. |  | 0.61706 | 0.40825 |  | 0 | 0 | 0 | 0 |  |  |  | 0.70711 | 0
| Rel.Std.Dev. |  | 10.85% | 6.999% |  | 0% | 0% | 0% | 0% |  |  |  | 6.734% | 0%
PDM3
| Z-Score (Absolute) |  | 0.3 | 0.26 | 0.52 | 1.04 | 1.04 | 1.04 | 0.26 | 0.52 | 0.52 | 1.04 | 1.19 | 0.83
| Analytical results for Ba in OREAS 251c (Certified Value 72 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 0.5g |  |  |  | 1
| 1 | 1 | 68 | 72.7509 | 68.812 | 86.8 | 73 | 70 | 80 | 70 | 70 | 70 | 70.43 | 92 | 71.2 | 94 | 60.6 | 108 | 66 | 81.007 |  |  |  | 1
| 1 | 2 | 70 | 74.8111 | 67.334 | 87.7 | 73 | 70 | 80 | 70 | 70 | 70 | 70.4977 | 92 | 75.3 | 92 | 59.2 | 129 | 64.8 | 82.81 |  |  |  | 25
| 1 | 3 | 66 | 73.3466 | 67.901 | 85.3 | 72 | 70 | 80 | 70 | 70 | 70 | 71.7389 | 92 | 75.8 | 92 | 65.6 | 98.4 | 63.4 | 82.863 |  |  |  | 16
| 1 | 4 | 68 | 70.4922 | 64.449 | 85.8 | 73 | 70 | 80 | 70 | 70 | 70 | 73.5691 | 90 | 73.9 | 93 | 63.6 | 128 | 68.3 | 76.335 |  |  |  | 71.9337
| 1 | 5 | 70 | 71.8679 | 69.392 | 85.4 | 74 | 70 | 80 | 70 | 70 | 70 | 74.6707 | 92 | 69.1 | 95 | 59.3 | 97.3 | 63.9 | 80.022 |  |  |  | 77
| 1 | 6 | 68 | 73.6615 | 67.265 | 83 | 72 | 70 | 80 | 70 | 70 | 70 | 73.009 | 91 | 69.4 | 95 | 61.3 | 127 | 63.4 | 79.899
| Mean |  | 68.3333 | 72.8217 | 67.5255 | 85.6667 | 72.8333 | 70 | 80 | 70 | 70 | 70 | 72.3193 | 91.5 | 72.45 | 93.5 | 61.6 | 114.6167 | 64.9667 | 80.4893
| Median |  | 68 | 73.0487 | 67.6175 | 85.6 | 73 | 70 | 80 | 70 | 70 | 70 | 72.374 | 92 | 72.55 | 93.5 | 60.95 | 117.5 | 64.35 | 80.5145
| Std Dev. |  | 1.5055 | 1.502 | 1.7233 | 1.5971 | 0.75277 | 0 | 0 | 0 | 0 | 0 | 1.72 | 0.83666 | 2.9508 | 1.3784 | 2.5354 | 15.1391 | 1.9107 | 2.413
| Rel.Std.Dev. |  | 2.203% | 2.063% | 2.552% | 1.864% | 1.034% | 0% | 0% | 0% | 0% | 0% | 2.378% | 0.9144% | 4.073% | 1.474% | 4.116% | 13.21% | 2.941% | 2.998%
| PDM3 |  | -5.005% | 1.234% | -6.128% | 19.09% | 1.251% | -2.688% | 11.21% | -2.688% | -2.688% | -2.688% | 0.536% | 27.2% | 0.7177% | 29.98% | -14.37% | 59.34% | -9.685% | 11.89%
| Z-Score (Absolute) |  | 0.61 | 0.07 | 0.74 | 2.01 | 0.07 | 0.36 | 1.15 | 0.36 | 0.36 | 0.36 | 0.01 | 2.89 | 0.01 | 3.2 | 1.63 | 6.39 | 1.12 | 1.23
| Analytical results for Be in OREAS 251c (Certified Value 0.93 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.91 | 0.94094 | 0.901 | 1 | 0.9 | 0.91 | 0.97 | 0.9 | 0.99 | 0.88 | 1.0729 | 1.11 | 0.95 | < 2 | 0.929 | [6 empty] | 1
| 1 | 2 | 0.97 | 0.89598 | 0.826 | 1 | 0.95 | 0.95 | 0.97 | 0.93 | 0.98 | 0.87 | 1.0544 | 1.12 | 1.06 | < 2 | 0.857 | [6 empty] | 26
| 1 | 3 | 0.81 | 0.88954 | 0.87 | 1 | 0.88 | 0.93 | 0.97 | 0.91 | 0.97 | 0.86 | 1.0433 | 1.14 | 0.92 | < 2 | 0.876 | [6 empty] | 16
| 1 | 4 | 0.94 | 0.86617 | 0.788 | 1 | 0.89 | 0.91 | 0.97 | 0.92 | 0.99 | 0.87 | 1.1177 | 1.1 | 1.02 | < 2 | 0.909 | [6 empty] | 0.93287
| 1 | 5 | 0.86 | 0.89322 | 0.917 | 1 | 0.89 | 0.92 | 0.98 | 0.92 | 0.97 | 0.87 | 1.1043 | 1.12 | 0.78 | < 2 | 0.359 | [6 empty] | 78
| 1 | 6 | 0.94 | 0.8684 | 0.884 | 1 | 0.91 | 0.92 | 0.99 | 0.91 | 0.98 | 0.87 | 1.0722 | 1.12 | 1.06 | < 2 | 0.368
| Mean |  | 0.905 | 0.89237 | 0.86433 | 1 | 0.90333 | 0.92333 | 0.975 | 0.915 | 0.98 | 0.87 | 1.0775 | 1.1183 | 0.965 |  | 0.71633
| Median |  | 0.925 | 0.89138 | 0.877 | 1 | 0.895 | 0.92 | 0.97 | 0.915 | 0.98 | 0.87 | 1.0726 | 1.12 | 0.985 |  | 0.8665
| Std Dev. |  | 0.059582 | 0.026986 | 0.048648 | 0 | 0.025033 | 0.015055 | 0.0083666 | 0.010488 | 0.0089443 | 0.0063246 | 0.028589 | 0.013292 | 0.10728 |  | 0.27446
| Rel.Std.Dev. |  | 6.584% | 3.024% | 5.628% | 0% | 2.771% | 1.631% | 0.8581% | 1.146% | 0.9127% | 0.727% | 2.653% | 1.189% | 11.12% |  | 38.32%
| PDM3 |  | -2.987% | -4.341% | -7.346% | 7.197% | -3.166% | -1.022% | 4.517% | -1.915% | 5.053% | -6.739% | 15.5% | 19.88% | 3.445% |  | -23.21%
| Z-Score (Absolute) |  | 0.18 | 0.35 | 0.71 | NA | 0.21 | 0.05 | 0.72 | 0.05 | 0.79 | 0.64 | 2.05 | 2.58 | 0.59 | 1.05 | 2.63
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 251c (Certified Value 6.83 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 6.84 | 6.4727 | 9.2129 | 7.23 | 7.42 | 6.56 | 7.5 | 6.24 | 6.86 | 6.63 | 7.1706 | 5.97 | 7.51 | 6.27 | 8 | 6.58 | 7.8 | < 5 | 6.7 | 2.19 |  | 1
| 1 | 2 | 6.77 | 6.8026 | 8.4489 | 7.3 | 7.39 | 7.36 | 7.05 | 6.7 | 6.41 | 6.67 | 7.2911 | 4.97 | 7.32 | 6.54 | 7 | 6.38 | 8.68 | < 5 | 6.5 | 0.011 |  | 27
| 1 | 3 | 6.54 | 6.9251 | 8.7227 | 7.36 | 7.02 | 7.29 | 6.87 | 6.73 | 6.53 | 6.33 | 7.3541 | 4.97 | 7.52 | 6.36 | 8 | 6.8 | 8.49 | < 5 | 6.7 | 0.961 |  | 16
| 1 | 4 | 6.48 | 6.287 | 9.7252 | 7.58 | 7.4 | 6.57 | 7.07 | 6.76 | 6.68 | 6.59 | 7.5368 | 5.97 | 7.23 | 6.32 | 8 | 6.73 | 7.73 | < 5 | 6.5 | 0.15 |  | 6.8278
| 1 | 5 | 6.63 | 6.68 | 8.5435 | 7.69 | 7.95 | 7.76 | 7.13 | 6.42 | 6.31 | 6.46 | 7.5184 | 4.97 | 7.35 | 5.72 | 8 | 6.46 | 7.69 | < 5 | 6.8 | 0.798 |  | 79
| 1 | 6 | 6.69 | 6.8344 | 10.3338 | 7.21 | 7 | 6.32 | 8.04 | 6.17 | 6.33 | 6.8 | 7.1397 | 4.97 | 7.42 | 5.93 | 8 | 6.57 | 7.81 | < 5 | 6.5
| Mean |  | 6.6583 | 6.667 | 9.1645 | 7.395 | 7.3633 | 6.9767 | 7.2767 | 6.5033 | 6.52 | 6.58 | 7.3351 | 5.3033 | 7.3917 | 6.19 | 7.8333 | 6.5867 | 8.0333 |  | 6.6167 | 0.822
| Median |  | 6.66 | 6.7413 | 8.9678 | 7.33 | 7.395 | 6.93 | 7.1 | 6.56 | 6.47 | 6.61 | 7.3226 | 4.97 | 7.385 | 6.295 | 8 | 6.575 | 7.805 |  | 6.6 | 0.798
| Std Dev. |  | 0.13644 | 0.24305 | 0.74525 | 0.19644 | 0.34622 | 0.57078 | 0.42744 | 0.26204 | 0.21615 | 0.16492 | 0.16843 | 0.5164 | 0.11339 | 0.30437 | 0.40825 | 0.1582 | 0.4338 |  | 0.13292 | 0.86612
| Rel.Std.Dev. |  | 2.049% | 3.646% | 8.132% | 2.656% | 4.702% | 8.181% | 5.874% | 4.029% | 3.315% | 2.506% | 2.296% | 9.737% | 1.534% | 4.917% | 5.212% | 2.402% | 5.4% |  | 2.009% | 105.4%
| PDM3 |  | -2.482% | -2.355% | 34.22% | 8.308% | 7.844% | 2.181% | 6.574% | -4.752% | -4.508% | -3.629% | 7.431% | -22.33% | 8.259% | -9.341% | 14.73% | -3.531% | 17.66% |  | -3.092% | -87.96%
| Z-Score (Absolute) |  | 0 | 0.01 | 2.73 | 0.8 | 0.77 | 0.35 | 0.67 | 0.17 | 0.15 | 0.09 | 0.74 | 1.48 | 0.8 | 0.51 | NA | 0.08 | 1.5 | 4.53 | 0.05 | 6.36
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 251c (Certified Value 0.437 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 0.42 | 0.44525 | 0.36662 | 0.68 | 0.47 | 0.41 | 0.47 | 0.41 | 0.43 | 0.43 | 0.45379 | 0.3149 | 0.5 | 0.41 | 0.48 | 0.37 | 0.505 | 0.43 | 0.43 | 0.41926 |  | 1
| 1 | 2 | 0.44 | 0.44229 | 0.35972 | 0.66 | 0.47 | 0.44 | 0.47 | 0.42 | 0.43 | 0.42 | 0.44943 | 0.309 | 0.5 | 0.43 | 0.47 | 0.35 | 0.544 | 0.41 | 0.42 | 0.44762 |  | #N/A
| 1 | 3 | 0.4 | 0.43444 | 0.35835 | 0.64 | 0.46 | 0.43 | 0.48 | 0.41 | 0.42 | 0.42 | 0.45504 | 0.3061 | 0.5 | 0.42 | 0.47 | 0.38 | 0.472 | 0.42 | 0.42 | 0.4186 |  | 16
| 1 | 4 | 0.43 | 0.42127 | 0.35433 | 0.67 | 0.46 | 0.41 | 0.47 | 0.42 | 0.43 | 0.42 | 0.46659 | 0.3186 | 0.51 | 0.42 | 0.48 | 0.38 | 0.536 | 0.44 | 0.43 | 0.39266 |  | 0.43656
| 1 | 5 | 0.43 | 0.43096 | 0.36691 | 0.66 | 0.48 | 0.42 | 0.47 | 0.42 | 0.42 | 0.43 | 0.47803 | 0.305 | 0.51 | 0.41 | 0.49 | 0.37 | 0.384 | 0.43 | 0.42 | 0.40201 |  | 80
| 1 | 6 | 0.43 | 0.44606 | 0.34768 | 0.63 | 0.46 | 0.42 | 0.48 | 0.42 | 0.42 | 0.42 | 0.47493 | 0.3055 | 0.5 | 0.39 | 0.48 | 0.38 | 0.511 | 0.41 | 0.41 | 0.42455
| Mean |  | 0.425 | 0.43671 | 0.35893 | 0.65667 | 0.46667 | 0.42167 | 0.47333 | 0.41667 | 0.425 | 0.42333 | 0.46297 | 0.30985 | 0.50333 | 0.41333 | 0.47833 | 0.37167 | 0.492 | 0.42333 | 0.42167 | 0.41745
| Median |  | 0.43 | 0.43836 | 0.35904 | 0.66 | 0.465 | 0.42 | 0.47 | 0.42 | 0.425 | 0.42 | 0.46082 | 0.30755 | 0.5 | 0.415 | 0.48 | 0.375 | 0.508 | 0.425 | 0.42 | 0.41893
| Std Dev. |  | 0.013784 | 0.0096755 | 0.0073698 | 0.018619 | 0.008165 | 0.01169 | 0.005164 | 0.005164 | 0.0054772 | 0.005164 | 0.011944 | 0.0056444 | 0.005164 | 0.013663 | 0.0075277 | 0.01169 | 0.058709 | 0.012111 | 0.0075277 | 0.019071
| Rel.Std.Dev. |  | 3.243% | 2.216% | 2.053% | 2.835% | 1.75% | 2.772% | 1.091% | 1.239% | 1.289% | 1.22% | 2.58% | 1.822% | 1.026% | 3.305% | 1.574% | 3.145% | 11.93% | 2.861% | 1.785% | 4.568%
| PDM3 |  | -2.647% | 0.03548% | -17.78% | 50.42% | 6.897% | -3.41% | 8.425% | -4.556% | -2.647% | -3.029% | 6.05% | -29.02% | 15.3% | -5.319% | 9.57% | -14.86% | 12.7% | -3.029% | -3.41% | -4.376%
| Z-Score (Absolute) |  | 0.02 | 0.33 | 1.71 | 6.11 | 1.12 | 0.07 | 1.29 | 0.2 | 0.02 | 0.02 | 1.02 | 3 | 2.08 | 0.28 | 1.42 | 1.38 | 1.78 | 0.02 | 0.07 | 0.18
| Analytical results for Cd in OREAS 251c (Certified Value 0.026 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  |  | 3
| 1 | 1 | 0.02 | < 0.2 | 0.03441 | 0.04 | 0.07 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 | < 0.05 | < 0.05 | 0.01 | 0.02 | 0.03 | 1 | < 0.1 |  |  |  |  | 1
| 1 | 2 | 0.03 | < 0.2 | 0.020689 | 0.05 | 0.06 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | < 0.05 | < 0.05 | 0.04 | 0.02 | 0.03 | 1.1 | < 0.1 |  |  |  |  | 28
| 1 | 3 | 0.02 | < 0.2 | 0.02877 | 0.02 | 0.07 | 0.02 | 0.02 | 0.02 | 0.03 | 0.03 | < 0.05 | < 0.05 | 0.05 | 0.02 | 0.03 | 1.1 | < 0.1 |  |  |  |  | 16
| 1 | 4 | 0.02 | < 0.2 | 0.023168 | 0.08 | 0.08 | 0.02 | 0.03 | 0.02 | 0.04 | 0.03 | < 0.05 | < 0.05 | 0.05 | 0.02 | 0.03 | 1.3 | < 0.1 |  |  |  |  | 0.025788
| 1 | 5 | 0.03 | < 0.2 | 0.029776 | 0.05 | 0.07 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03 | < 0.05 | < 0.05 | 0.03 | 0.02 | 0.04 | 1.3 | < 0.1 |  |  |  |  | 81
| 1 | 6 | 0.03 | < 0.2 | 0.030468 | 0.04 | 0.07 | 0.02 | 0.02 | 0.02 | 0.03 | 0.01 | < 0.05 | < 0.05 | 0.04 | 0.02 | 0.03 | 1.2 | < 0.1
| Mean |  | 0.025 |  | 0.02788 | 0.046667 | 0.07 | 0.021667 | 0.021667 | 0.02 | 0.03 | 0.023333 |  |  | 0.036667 | 0.02 | 0.031667 | 1.1667
| Median |  | 0.025 |  | 0.029273 | 0.045 | 0.07 | 0.02 | 0.02 | 0.02 | 0.03 | 0.025 |  |  | 0.04 | 0.02 | 0.03 | 1.15
| Std Dev. |  | 0.0054772 |  | 0.0050525 | 0.019664 | 0.0063246 | 0.0040825 | 0.0040825 | 0 | 0.0063246 | 0.008165 |  |  | 0.015055 | 0 | 0.0040825 | 0.12111
| Rel.Std.Dev. |  | 21.91% |  | 18.12% | 42.14% | 9.035% | 18.84% | 18.84% | 0% | 21.08% | 34.99% |  |  | 41.06% | 0% | 12.89% | 10.38%
| PDM3 |  | -3.056% |  | 8.113% | 80.96% | 171.4% | -15.98% | -15.98% | -22.44% | 16.33% | -9.519% |  |  | 42.18% | -22.44% | 22.8% | 4424%
| Z-Score (Absolute) |  | 0.31 | 7.83 | 0 | 2.04 | 4.57 | 0.67 | 0.67 | 0.86 | 0.23 | 0.49 | 0.31 | 0.31 | 0.95 | 0.86 | 0.41 | 123.58 | 2.4
| Analytical results for Ce in OREAS 251c (Certified Value 52 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.18 | 1.21 | 1.28 | 1.30 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 0
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 01g | 0.1g | 40g | 0.5g | [5 empty] | 1
| 1 | 1 | 48.165 | 50.1951 | 74.6212 | 56.7 | 51 | 47.5 | 52.6 | 45.5 | 46.9 | 49.8 | 52.7659 | 61.3 | 55 | 84.8 | 64.6 | 51.386 | [5 empty] | 1
| 1 | 2 | 49.44 | 50.2397 | 74.6248 | 57.6 | 53 | 49.8 | 54.3 | 46.9 | 49.8 | 46.4 | 51.6987 | 63.3 | 54 | 86.3 | 64.8 | 51.541 | [5 empty] | 29
| 1 | 3 | 45.278 | 50.4307 | 74.6272 | 56.2 | 48.69 | 49.5 | 52.8 | 46.9 | 47.2 | 47 | 51.41 | 63.3 | 55 | 80.1 | 58.4 | 53.354 | [5 empty] | 16
| 1 | 4 | 47.689 | 47.2586 | 74.634 | 56.1 | 51 | 47.4 | 53.4 | 46.7 | 46.8 | 47.9 | 53.5574 | 61.2 | 55 | 85.6 | 59.6 | 48.046 | [5 empty] | 52.1856
| 1 | 5 | 49.465 | 48.7016 | 74.7244 | 57.1 | 52 | 48.4 | 53.7 | 46.1 | 47.9 | 47.5 | 54.2177 | 62.9 | 54 | 68.6 | 60.4 | 51.033 | [5 empty] | 82
| 1 | 6 | 48.223 | 50.1067 | 74.5748 | 55.5 | 50 | 47.3 | 53 | 45.4 | 47.7 | 48.1 | 52.2407 | 61.3 | 57 | 82.8 | 60.8 | 49.978
| Mean |  | 48.0433 | 49.4888 | 74.6344 | 56.5333 | 50.9483 | 48.3167 | 53.3 | 46.25 | 47.7167 | 47.7833 | 52.6484 | 62.2167 | 55 | 81.3667 | 61.4333 | 50.8897
| Median |  | 48.194 | 50.1509 | 74.626 | 56.45 | 51 | 47.95 | 53.2 | 46.4 | 47.45 | 47.7 | 52.5033 | 62.1 | 55 | 83.8 | 60.6 | 51.2095
| Std Dev. |  | 1.5357 | 1.259 | 0.04893 | 0.75542 | 1.5047 | 1.1089 | 0.63246 | 0.68629 | 1.1089 | 1.1652 | 1.0867 | 1.0515 | 1.0954 | 6.6437 | 2.6606 | 1.771
| Rel.Std.Dev. |  | 3.196% | 2.544% | 0.06556% | 1.336% | 2.953% | 2.295% | 1.187% | 1.484% | 2.324% | 2.438% | 2.064% | 1.69% | 1.992% | 8.165% | 4.331% | 3.48%
| PDM3 |  | -7.937% | -5.168% | 43.02% | 8.331% | -2.371% | -7.414% | 2.136% | -11.37% | -8.563% | -8.436% | 0.8869% | 19.22% | 5.393% | 55.92% | 17.72% | -2.483%
| Z-Score (Absolute) |  | 0.65 | 0.4 | 3.96 | 0.82 | 0.15 | 0.6 | 0.26 | 0.96 | 0.71 | 0.7 | 0.15 | 1.81 | 0.56 | 5.13 | 1.67 | 0.16
| Analytical results for Co in OREAS 251c (Certified Value 11.6 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 11 | 12.0119 | 11.347 | 14 | 11.5 | 11.4 | 12.7 | 11.1 | 11 | 10.4 | 12.7732 | 13.1 | 10 | 13 | 10.4 | 13.8 | 11.2 | 10.6 | 11.862 |  |  | 1
| 1 | 2 | 11.3 | 11.839 | 10.242 | 14.3 | 11.5 | 12 | 12.4 | 11.4 | 12.2 | 9.9 | 13.0969 | 13.6 | 10.9 | 13 | 9.9 | 13.8 | 11.9 | 11.4 | 11.492 |  |  | 15
| 1 | 3 | 10.9 | 11.9454 | 10.673 | 13.8 | 11 | 11.5 | 12.6 | 11.2 | 11.6 | 10.4 | 12.7379 | 13.4 | 10.4 | 13 | 10.7 | 13.4 | 11.5 | 10.8 | 11.364 |  |  | 16
| 1 | 4 | 11.1 | 11.3958 | 9.636 | 13.9 | 11.6 | 11.3 | 12.5 | 11 | 11.6 | 10.6 | 12.7622 | 13.1 | 11.1 | 13 | 10.7 | 13.8 | 12.2 | 11.4 | 10.776 |  |  | 11.5874
| 1 | 5 | 11.3 | 12.1888 | 10.905 | 13.8 | 11.6 | 11.3 | 12.9 | 11.3 | 11.8 | 10.8 | 12.5419 | 13.3 | 9.1 | 13 | 10.5 | 11.6 | 10.8 | 11.6 | 11.478 |  |  | 83
| 1 | 6 | 11.5 | 11.8647 | 9.965 | 13.7 | 11.5 | 11.5 | 12.7 | 11 | 11.5 | 11 | 13.1176 | 13.3 | 10.1 | 13 | 10.7 | 13.2 | 11.1 | 11.4 | 11.056
| Mean |  | 11.1833 | 11.8743 | 10.4613 | 13.9167 | 11.45 | 11.5 | 12.6333 | 11.1667 | 11.6167 | 10.5167 | 12.8383 | 13.3 | 10.2667 | 13 | 10.4833 | 13.2667 | 11.45 | 11.2 | 11.338
| Median |  | 11.2 | 11.9051 | 10.4575 | 13.85 | 11.5 | 11.45 | 12.65 | 11.15 | 11.6 | 10.5 | 12.7677 | 13.3 | 10.25 | 13 | 10.6 | 13.6 | 11.35 | 11.4 | 11.421
| Std Dev. |  | 0.22286 | 0.26578 | 0.63272 | 0.2137 | 0.22583 | 0.26077 | 0.17512 | 0.1633 | 0.392 | 0.38166 | 0.22487 | 0.18974 | 0.7174 | 0 | 0.31252 | 0.85479 | 0.5244 | 0.4 | 0.37787
| Rel.Std.Dev. |  | 1.993% | 2.238% | 6.048% | 1.536% | 1.972% | 2.268% | 1.386% | 1.462% | 3.374% | 3.629% | 1.752% | 1.427% | 6.988% | 0% | 2.981% | 6.443% | 4.58% | 3.571% | 3.333%
| PDM3 |  | -3.487% | 2.476% | -9.718% | 20.1% | -1.185% | -0.754% | 9.027% | -3.631% | 0.2529% | -9.24% | 10.8% | 14.78% | -11.4% | 12.19% | -9.528% | 14.49% | -1.185% | -3.343% | -2.152%
| Z-Score (Absolute) |  | 0.26 | 0.42 | 0.98 | 2.45 | 0 | 0.05 | 1.18 | 0.28 | 0.17 | 0.93 | 1.38 | 1.84 | 1.18 | NA | 0.96 | 1.8 | 0 | 0.25 | 0.11
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 251c (Certified Value 46.3 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 40g | 0.5g |  |  | 1
| 1 | 1 | 42 | 45.76 | 36.391 | 57.55 | 52 | 44 | 44 | 48 | 43 | 45 | 44 | 49.1453 | 58 | 43 | 54 | 39 | 58 | 40 | 47.398 |  |  | 1
| 1 | 2 | 43 | 46.1679 | 35.269 | 57.28 | 55 | 46 | 46 | 48 | 44 | 45 | 44 | 48.6641 | 57 | 45.1 | 53 | 37.3 | 65 | 40 | 54.277 |  |  | 31
| 1 | 3 | 40 | 45.6217 | 36.032 | 56.28 | 52 | 43 | 45 | 48 | 43 | 43 | 44 | 50.0198 | 57 | 45.3 | 53 | 40.7 | 55 | 40 | 47.65 |  |  | 16
| 1 | 4 | 43 | 44.5477 | 33.862 | 58.72 | 52 | 45 | 44 | 48 | 44 | 45 | 44 | 52.9991 | 56 | 44 | 53 | 41.3 | 62 | 40 | 44.993 |  |  | 46.2617
| 1 | 5 | 43 | 46.0605 | 37.134 | 59.18 | 52 | 45 | 45 | 51 | 44 | 44 | 44 | 51.5702 | 56 | 42.2 | 53 | 38.9 | 64 | 40 | 48.038 |  |  | 84
| 1 | 6 | 43 | 46.1674 | 35.072 | 54.13 | 52 | 43 | 44 | 49 | 44 | 43 | 47 | 53.3383 | 56 | 41.5 | 53 | 40.2 | 61 | 40 | 44.257
| Mean |  | 42.3333 | 45.7209 | 35.6267 | 57.19 | 52.5 | 44.3333 | 44.6667 | 48.6667 | 43.6667 | 44.1667 | 44.5 | 50.9561 | 56.6667 | 43.5167 | 53.1667 | 39.5667 | 60.8333 | 40 | 47.7688
| Median |  | 43 | 45.9103 | 35.6505 | 57.415 | 52 | 44.5 | 44.5 | 48 | 44 | 44.5 | 44 | 50.795 | 56.5 | 43.5 | 53 | 39.6 | 61.5 | 40 | 47.524
| Std Dev. |  | 1.2111 | 0.61682 | 1.1463 | 1.8238 | 1.2247 | 1.2111 | 0.8165 | 1.2111 | 0.5164 | 0.98319 | 1.2247 | 1.9816 | 0.8165 | 1.5484 | 0.40825 | 1.4556 | 3.7639 | 0 | 3.5389
| Rel.Std.Dev. |  | 2.861% | 1.349% | 3.218% | 3.189% | 2.333% | 2.732% | 1.828% | 2.488% | 1.183% | 2.226% | 2.752% | 3.889% | 1.441% | 3.558% | 0.7679% | 3.679% | 6.187% | 0% | 7.408%
| PDM3 |  | -8.492% | -1.169% | -22.99% | 23.62% | 13.48% | -4.168% | -3.448% | 5.199% | -5.609% | -4.529% | -3.808% | 10.15% | 22.49% | -5.934% | 14.93% | -14.47% | 31.5% | -13.54% | 3.258%
| Z-Score (Absolute) |  | 0.39 | 0.18 | 1.52 | 2.11 | 1.32 | 0.06 | 0 | 0.67 | 0.17 | 0.08 | 0.03 | 1.06 | 2.02 | 0.19 | 1.43 | 0.86 | 2.73 | 0.79 | 0.52
| Analytical results for Cs in OREAS 251c (Certified Value 1.43 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 2
| 1 | 1 | 1.65 | 1.6649 | 1.61 | 1.36 | 1.36 | 1.46 | 1.27 | 1.29 | 1.49 | 1.5265 | 2.29 | 2.7 | 1.1 | [8 empty] | 1
| 1 | 2 | 1.7 | 1.6364 | 1.48 | 1.41 | 1.41 | 1.47 | 1.33 | 1.33 | 1.37 | 1.528 | 2.31 | 3.2 | 1.1 | [8 empty] | 32
| 1 | 3 | 1.64 | 1.6873 | 1.54 | 1.28 | 1.43 | 1.47 | 1.28 | 1.28 | 1.43 | 1.534 | 2.32 | 2.6 | 1.2 | [8 empty] | 16
| 1 | 4 | 1.66 | 1.5922 | 1.53 | 1.35 | 1.39 | 1.51 | 1.31 | 1.29 | 1.42 | 1.555 | 2.29 | 3.1 | 1.2 | [8 empty] | 1.4341
| 1 | 5 | 1.66 | 1.6658 | 1.54 | 1.36 | 1.39 | 1.49 | 1.31 | 1.3 | 1.44 | 1.5502 | 2.32 | 2.7 | 1.2 | [8 empty] | 85
| 1 | 6 | 1.68 | 1.6942 | 1.44 | 1.3 | 1.39 | 1.49 | 1.28 | 1.28 | 1.44 | 1.567 | 2.31 | 3.2 | 1.1
| Mean |  | 1.665 | 1.6568 | 1.5233 | 1.3433 | 1.395 | 1.4817 | 1.2967 | 1.295 | 1.4317 | 1.5434 | 2.3067 | 2.9167 | 1.15
| Median |  | 1.66 | 1.6653 | 1.535 | 1.355 | 1.39 | 1.48 | 1.295 | 1.29 | 1.435 | 1.5421 | 2.31 | 2.9 | 1.15
| Std Dev. |  | 0.021679 | 0.037566 | 0.058195 | 0.046762 | 0.023452 | 0.018348 | 0.023381 | 0.018708 | 0.038687 | 0.01643 | 0.013663 | 0.27869 | 0.054772
| Rel.Std.Dev. |  | 1.302% | 2.267% | 3.82% | 3.481% | 1.681% | 1.238% | 1.803% | 1.445% | 2.702% | 1.064% | 0.5923% | 9.555% | 4.763%
| PDM3 |  | 16.1% | 15.53% | 6.224% | -6.328% | -2.725% | 3.318% | -9.582% | -9.698% | -0.1684% | 7.626% | 60.85% | 103.4% | -19.81%
| Z-Score (Absolute) |  | 0.71 | 0.67 | 0.16 | 0.53 | 0.33 | 0 | 0.71 | 0.72 | 0.19 | 0.24 | 3.18 | 5.53 | 1.28
| Analytical results for Cu in OREAS 251c (Certified Value 22.7 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 2
| 1 | 1 | 21.6 | 22.8531 | 21.015 | 23.82 | 27 | 21.5 | 22.2 | 24.3 | 22.4 | 23.7 | 23 | 21.4903 | 20 | 24 | 22.24 | 24 | 20.68 | 18.5 | 32.4 | 7 | 27.552 | 1
| 1 | 2 | 22.1 | 22.7617 | 20.509 | 22.85 | 26.2 | 23 | 23.7 | 24.3 | 22.4 | 23.6 | 21.4 | 23.931 | 20 | 24 | 22.31 | 23 | 20.06 | 18.8 | 30.5 | 9 | 27.776 | 17
| 1 | 3 | 21.1 | 22.8667 | 20.85 | 22.72 | 24.5 | 23.1 | 22.3 | 24.4 | 22.4 | 23.6 | 22 | 23.63 | 19 | 24 | 22.53 | 23 | 21.17 | 17.1 | 31.4 | 7 | 27.27 | 16
| 1 | 4 | 21.8 | 21.7831 | 20.279 | 22.89 | 25.8 | 22.8 | 22.8 | 24.4 | 22.2 | 23.3 | 22.6 | 24.0634 | 20 | 24 | 22.37 | 24 | 21.36 | 18.2 | 31.6 | 7 | 25.419 | 22.665
| 1 | 5 | 22.2 | 23.0641 | 21.012 | 22.86 | 26.3 | 23 | 22.4 | 24.9 | 22.4 | 22.9 | 23.5 | 21.3917 | 19 | 25 | 18.63 | 23 | 20.59 | 16.2 | 30.6 | 7 | 27.395 | 86
| 1 | 6 | 22.8 | 22.9701 | 21.233 | 23.74 | 25.6 | 21.6 | 22.1 | 23.9 | 22.1 | 23 | 23.5 | 23.201 | 19 | 24 | 21.5 | 24 | 20.99 | 16.6 | 31 | 8 | 27.232
| Mean |  | 21.9333 | 22.7165 | 20.8163 | 23.1467 | 25.9 | 22.5 | 22.5833 | 24.3667 | 22.3167 | 23.35 | 22.6667 | 22.9512 | 19.5 | 24.1667 | 21.5967 | 23.5 | 20.8083 | 17.5667 | 31.25 | 7.5 | 27.1073
| Median |  | 21.95 | 22.8599 | 20.931 | 22.875 | 26 | 22.9 | 22.35 | 24.35 | 22.4 | 23.45 | 22.8 | 23.4155 | 19.5 | 24 | 22.275 | 23.5 | 20.835 | 17.65 | 31.2 | 7 | 27.3325
| Std Dev. |  | 0.5785 | 0.46895 | 0.35656 | 0.49468 | 0.83905 | 0.74297 | 0.59805 | 0.32042 | 0.13292 | 0.33912 | 0.84301 | 1.2072 | 0.54772 | 0.40825 | 1.4968 | 0.54772 | 0.46739 | 1.0783 | 0.70922 | 0.83666 | 0.85089
| Rel.Std.Dev. |  | 2.638% | 2.064% | 1.713% | 2.137% | 3.24% | 3.302% | 2.648% | 1.315% | 0.5956% | 1.452% | 3.719% | 5.26% | 2.809% | 1.689% | 6.931% | 2.331% | 2.246% | 6.138% | 2.27% | 11.16% | 3.139%
| PDM3 |  | -3.228% | 0.2269% | -8.157% | 2.125% | 14.27% | -0.7282% | -0.3605% | 7.508% | -1.537% | 3.022% | 0.007151% | 1.263% | -13.96% | 6.625% | -4.714% | 3.684% | -8.192% | -22.49% | 37.88% | -66.91% | 19.6%
| Z-Score (Absolute) |  | 0.46 | 0.03 | 1.17 | 0.3 | 2.04 | 0.11 | 0.05 | 1.07 | 0.22 | 0.43 | 0 | 0.18 | 2 | 0.95 | 0.67 | 0.53 | 1.17 | 3.21 | 5.41 | 9.56 | 2.8
| Analytical results for Dy in OREAS 251c (Indicative Value 2.21 ppm) | [22 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 2
| 1 | 1 | 2.069 | 2.1643 | 3.6662 | 2.5 | 2.04 | 1.124 | [15 empty] | 1
| 1 | 2 | 2.171 | 2.2095 | 3.8774 | 2.6 | 2.06 | 1.411 | [15 empty] | 4
| 1 | 3 | 2.015 | 2.2155 | 3.7302 | 2.4 | 2.17 | 0.878 | [15 empty] | 16
| 1 | 4 | 2.108 | 2.1287 | 3.8165 | 2.4 | 2.11 | 1.358 | [15 empty] | 2.2127
| 1 | 5 | 2.152 | 2.1917 | 3.8304 | 2.4 | 2.14 | 1.468 | [15 empty] | 10
| 1 | 6 | 2.103 | 2.2568 | 3.6946 | 2.4 | 2.1 | 1.367
| Mean |  | 2.103 | 2.1944 | 3.7692 | 2.45 | 2.1033 | 1.2677
| Median |  | 2.1055 | 2.2006 | 3.7734 | 2.4 | 2.105 | 1.3625
| Std Dev. |  | 0.05648 | 0.044272 | 0.084136 | 0.083666 | 0.048442 | 0.22414
| Rel.Std.Dev. |  | 2.686% | 2.017% | 2.232% | 3.415% | 2.303% | 17.68%
| PDM3 |  | -4.957% | -0.8258% | 70.35% | 10.73% | -4.942% | -42.71%
| Z-Score (Absolute) |  | 0.18 | 0.18 | 6.3 | 1.17 | 0.18 | 3.42
| Analytical results for Er in OREAS 251c (Indicative Value 0.86 ppm) | [22 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.8 | 0.82085 | 1.4303 | 0.9 | 0.84 | [16 empty] | 1
| 1 | 2 | 0.862 | 0.83916 | 1.453 | 0.9 | 0.87 | [16 empty] | 5
| 1 | 3 | 0.776 | 0.87097 | 1.4357 | 0.9 | 0.87 | [16 empty] | 16
| 1 | 4 | 0.79 | 0.80631 | 1.4799 | 0.9 | 0.91 | [16 empty] | 0.86433
| 1 | 5 | 0.822 | 0.8169 | 1.4698 | 1 | 0.87 | [16 empty] | 11
| 1 | 6 | 0.814 | 0.88563 | 1.4522 | 1 | 0.88
| Mean |  | 0.81067 | 0.83997 | 1.4535 | 0.93333 | 0.87333
| Median |  | 0.807 | 0.83 | 1.4526 | 0.9 | 0.87
| Std Dev. |  | 0.030058 | 0.031867 | 0.019063 | 0.05164 | 0.022509
| Rel.Std.Dev. |  | 3.708% | 3.794% | 1.312% | 5.533% | 2.577%
| PDM3 |  | -6.208% | -2.818% | 68.16% | 7.984% | 1.042%
| Z-Score (Absolute) |  | 0.7 | 0.37 | 6.52 | 0.67 | 0
| Analytical results for Eu in OREAS 251c (Certified Value 0.94 ppm) | [22 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | 1.30 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [15 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | 0.5g | [15 empty] | 3
| 1 | 1 | 0.911 | 0.92125 | 1.499 | 1.1 | 0.874 | 0.915 | [15 empty] | 1
| 1 | 2 | 0.94 | 0.94597 | 1.6612 | 1.1 | 0.866 | 0.918 | [15 empty] | 36
| 1 | 3 | 0.874 | 0.96518 | 1.5338 | 1.1 | 0.862 | 0.924 | [15 empty] | 16
| 1 | 4 | 0.922 | 0.88875 | 1.5948 | 1 | 0.872 | 0.904 | [15 empty] | 0.9406
| 1 | 5 | 0.94 | 0.91604 | 1.5833 | 1.1 | 0.866 | 0.917 | [15 empty] | 87
| 1 | 6 | 0.898 | 1.0067 | 1.5916 | 1 | 0.88 | 0.891
| Mean |  | 0.91417 | 0.94064 | 1.5773 | 1.0667 | 0.87 | 0.9115
| Median |  | 0.9165 | 0.93361 | 1.5874 | 1.1 | 0.869 | 0.916
| Std Dev. |  | 0.025616 | 0.041625 | 0.055892 | 0.05164 | 0.0065727 | 0.011979
| Rel.Std.Dev. |  | 2.802% | 4.425% | 3.544% | 4.841% | 0.7555% | 1.314%
| PDM3 |  | -2.81% | 0.005079% | 67.69% | 13.4% | -7.505% | -3.093%
| Z-Score (Absolute) |  | 0.24 | 0.24 | 11.96 | 2.56 | 1.06 | 0.29
| Analytical results for Fe in OREAS 251c (Certified Value 3.03 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 2.86 | 3.0341 | 2.655 | 3.5 | 3.06 | 2.85 | 3.04 | 2.81 | 2.96 | 2.74 | 3.206 | 2.2403 | 3.31 | 2.82 | 3.27 | 2.46 | 3.51 | 3.31 | 2.95 | 3.2266 |  | 1
| 1 | 2 | 2.96 | 3.0881 | 2.5965 | 3.6 | 3.07 | 2.96 | 2.97 | 2.88 | 2.95 | 2.75 | 3.0802 | 2.195 | 3.45 | 2.89 | 3.21 | 2.37 | 3.53 | 3.22 | 3.04 | 3.2815 |  | #N/A
| 1 | 3 | 2.93 | 3.0533 | 2.5978 | 3.46 | 3.07 | 2.87 | 3.05 | 2.85 | 2.91 | 2.73 | 3.2225 | 2.1638 | 3.49 | 2.87 | 3.22 | 2.59 | 3.32 | 3.28 | 3.04 | 3.2805 |  | 16
| 1 | 4 | 2.92 | 2.9887 | 2.5591 | 3.5 | 3.07 | 2.81 | 3.06 | 2.89 | 2.97 | 2.74 | 3.2807 | 2.2488 | 3.34 | 2.82 | 3.22 | 2.6 | 3.52 | 3.49 | 3.06 | 2.992 |  | 3.0318
| 1 | 5 | 2.91 | 3.0815 | 2.6651 | 3.53 | 3.05 | 2.87 | 3.08 | 2.87 | 2.91 | 2.74 | 3.321 | 2.125 | 3.48 | 2.77 | 3.2 | 2.48 | 2.96 | 3.22 | 3.02 | 3.1996 |  | 88
| 1 | 6 | 2.94 | 3.0752 | 2.5279 | 3.44 | 3.05 | 2.85 | 3.09 | 2.87 | 2.89 | 2.75 | 3.3813 | 2.1555 | 3.45 | 2.7 | 3.27 | 2.54 | 3.31 | 3.25 | 2.95 | 3.1446
| Mean |  | 2.92 | 3.0535 | 2.6002 | 3.505 | 3.0617 | 2.8683 | 3.0483 | 2.8617 | 2.9317 | 2.7417 | 3.2486 | 2.1881 | 3.42 | 2.8117 | 3.2317 | 2.5067 | 3.3583 | 3.295 | 3.01 | 3.1875
| Median |  | 2.925 | 3.0642 | 2.5972 | 3.5 | 3.065 | 2.86 | 3.055 | 2.87 | 2.93 | 2.74 | 3.2516 | 2.1794 | 3.45 | 2.82 | 3.22 | 2.51 | 3.415 | 3.265 | 3.03 | 3.2131
| Std Dev. |  | 0.034059 | 0.037486 | 0.053199 | 0.05648 | 0.0098319 | 0.049967 | 0.042622 | 0.028577 | 0.032506 | 0.0075277 | 0.10464 | 0.049177 | 0.075895 | 0.069113 | 0.030605 | 0.08756 | 0.21958 | 0.10173 | 0.048166 | 0.10882
| Rel.Std.Dev. |  | 1.166% | 1.228% | 2.046% | 1.611% | 0.3211% | 1.742% | 1.398% | 0.9986% | 1.109% | 0.2746% | 3.221% | 2.248% | 2.219% | 2.458% | 0.947% | 3.493% | 6.538% | 3.088% | 1.6% | 3.414%
| PDM3 |  | -3.687% | 0.7151% | -14.23% | 15.61% | 0.9852% | -5.392% | 0.5454% | -5.612% | -3.303% | -9.57% | 7.151% | -27.83% | 12.8% | -7.261% | 6.592% | -17.32% | 10.77% | 8.681% | -0.719% | 5.135%
| Z-Score (Absolute) |  | 0.35 | 0.08 | 1.38 | 1.53 | 0.1 | 0.52 | 0.06 | 0.54 | 0.31 | 0.92 | 0.7 | 2.7 | 1.25 | 0.7 | 0.65 | 1.68 | 1.06 | 0.85 | 0.06 | 0.51
| Analytical results for Ga in OREAS 251c (Certified Value 3.94 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | [6 empty] | 3
| 1 | 1 | 3.68 | 5.2117 | 4.77 | 4.1 | 3.87 | 4.04 | 3.56 | 3.91 | 3.83 | 4.2301 | 5.3 | 3.9 | 3.2 | 6.76 | 7.9 | [6 empty] | 1
| 1 | 2 | 3.85 | 5.3799 | 4.82 | 4.02 | 4.09 | 4.1 | 3.67 | 3.96 | 3.64 | 4.1841 | 5.1 | 4.3 | 3.1 | 7.53 | 7.8 | [6 empty] | 37
| 1 | 3 | 3.63 | 5.0564 | 4.56 | 3.96 | 4.01 | 4.18 | 3.56 | 3.8 | 3.68 | 4.2001 | 5.2 | 4.1 | 3.4 | 6.12 | 6.8 | [6 empty] | 16
| 1 | 4 | 3.74 | 4.8699 | 4.66 | 3.91 | 3.92 | 4.2 | 3.64 | 3.85 | 3.89 | 4.4342 | 5.1 | 4.2 | 3.3 | 7.11 | 8.8 | [6 empty] | 3.9416
| 1 | 5 | 3.86 | 5.1204 | 4.59 | 4.15 | 3.92 | 4.21 | 3.65 | 3.8 | 3.85 | 4.4158 | 5.1 | 3.5 | 3.2 | 6.49 | 6.4 | [6 empty] | 89
| 1 | 6 | 3.79 | 5.1565 | 4.59 | 4.06 | 3.9 | 4.22 | 3.58 | 3.79 | 3.88 | 4.5093 | 5.3 | 3.8 | 3.4 | 6.87 | 6.9
| Mean |  | 3.7583 | 5.1325 | 4.665 | 4.0333 | 3.9517 | 4.1583 | 3.61 | 3.8517 | 3.795 | 4.3289 | 5.1833 | 3.9667 | 3.2667 | 6.8133 | 7.4333
| Median |  | 3.765 | 5.1385 | 4.625 | 4.04 | 3.92 | 4.19 | 3.61 | 3.825 | 3.84 | 4.3229 | 5.15 | 4 | 3.25 | 6.815 | 7.35
| Std Dev. |  | 0.092394 | 0.16911 | 0.1071 | 0.088919 | 0.082321 | 0.072226 | 0.04899 | 0.06969 | 0.10747 | 0.14036 | 0.098319 | 0.29439 | 0.12111 | 0.48845 | 0.89144
| Rel.Std.Dev. |  | 2.458% | 3.295% | 2.296% | 2.205% | 2.083% | 1.737% | 1.357% | 1.809% | 2.832% | 3.242% | 1.897% | 7.422% | 3.707% | 7.169% | 11.99%
| PDM3 |  | -4.65% | 30.21% | 18.35% | 2.327% | 0.2546% | 5.498% | -8.414% | -2.282% | -3.72% | 9.826% | 31.5% | 0.6352% | -17.12% | 72.86% | 88.59%
| Z-Score (Absolute) |  | 0.63 | 2.51 | 1.44 | 0 | 0.19 | 0.29 | 0.97 | 0.41 | 0.54 | 0.67 | 2.62 | 0.15 | 1.75 | 6.34 | 7.76
| Analytical results for Gd in OREAS 251c (Indicative Value 3.59 ppm) | [22 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 3.262 | 3.5962 | 4.1 | 3.5 | [17 empty] | 1
| 1 | 2 | 3.396 | 3.5384 | 4 | 3.39 | [17 empty] | 6
| 1 | 3 | 3.176 | 3.5976 | 4 | 3.37 | [17 empty] | 16
| 1 | 4 | 3.304 | 3.5327 | 4.2 | 3.52 | [17 empty] | 3.589
| 1 | 5 | 3.45 | 3.4868 | 4.1 | 3.38 | [17 empty] | 12
| 1 | 6 | 3.283 | 3.5841 | 3.9 | 3.47
| Mean |  | 3.3118 | 3.5559 | 4.05 | 3.4383
| Median |  | 3.2935 | 3.5612 | 4.05 | 3.43
| Std Dev. |  | 0.09798 | 0.044208 | 0.10488 | 0.066156
| Rel.Std.Dev. |  | 2.958% | 1.243% | 2.59% | 1.924%
| PDM3 |  | -7.723% | -0.9217% | 12.84% | -4.199%
| Z-Score (Absolute) |  | 1.02 | 0.32 | 3.05 | 0.32
| Analytical results for Ge in OREAS 251c (Indicative Value 0.13 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.21 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | [12 empty] | 2
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.1g | [12 empty] | 2
| 1 | 1 | < 0.2 | < 0.1 | 0.23 | 0.06 | 0.08 | 0.11 | 0.08 | 0.14 | 0.17 | [12 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | 0.23 | 0.07 | 0.09 | 0.11 | 0.08 | 0.15 | 0.2 | [12 empty] | 2
| 1 | 3 | < 0.2 | < 0.1 | 0.22 | 0.07 | 0.08 | 0.1 | 0.08 | 0.14 | 0.15 | [12 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | 0.21 | 0.07 | 0.08 | 0.11 | 0.07 | 0.15 | 0.19 | [12 empty] | 0.125
| 1 | 5 | < 0.2 | < 0.1 | 0.21 | 0.06 | 0.09 | 0.11 | 0.08 | 0.14 | 0.14 | [12 empty] | 13
| 1 | 6 | < 0.2 | < 0.1 | 0.23 | 0.07 | 0.08 | 0.11 | 0.07 | 0.15 | 0.19
| Mean |  |  |  | 0.22167 | 0.066667 | 0.083333 | 0.10833 | 0.076667 | 0.145 | 0.17333
| Median |  |  |  | 0.225 | 0.07 | 0.08 | 0.11 | 0.08 | 0.145 | 0.18
| Std Dev. |  |  |  | 0.0098319 | 0.005164 | 0.005164 | 0.0040825 | 0.005164 | 0.0054772 | 0.024221
| Rel.Std.Dev. |  |  |  | 4.435% | 7.746% | 6.197% | 3.768% | 6.736% | 3.777% | 13.97%
| PDM3 |  |  |  | 77.33% | -46.67% | -33.33% | -13.33% | -38.67% | 16% | 38.67%
| Z-Score (Absolute) |  | 0 | 1.01 | 2.46 | 0.67 | 0.34 | 0.17 | 0.47 | 0.91 | 1.48
| Analytical results for Hf in OREAS 251c (Indicative Value 0.53 ppm) | [22 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 2
| 1 | 1 | 0.1 | 0.55127 | < 0.1 | 0.42 | 0.65 | 0.36 | 0.4 | 0.33 | 0.54 | 0.83949 | 0.73 | 0.84 | [9 empty] | 1
| 1 | 2 | 0.09 | 0.57241 | < 0.1 | 0.41 | 0.64 | 0.37 | 0.42 | 0.31 | 0.46 | 0.83516 | 0.76 | 0.8 | [9 empty] | 8
| 1 | 3 | 0.06 | 0.55067 | < 0.1 | 0.42 | 0.72 | 0.33 | 0.42 | 0.37 | 0.54 | 0.84874 | 0.77 | 0.86 | [9 empty] | 16
| 1 | 4 | 0.09 | 0.54171 | < 0.1 | 0.41 | 0.66 | 0.31 | 0.41 | 0.34 | 0.55 | 0.82848 | 0.78 | 0.84 | [9 empty] | 0.5266
| 1 | 5 | 0.08 | 0.58258 | < 0.1 | 0.41 | 0.62 | 0.38 | 0.43 | 0.32 | 0.56 | 0.87319 | 0.73 | 0.77 | [9 empty] | 14
| 1 | 6 | 0.07 | 0.54079 | < 0.1 | 0.38 | 0.71 | 0.34 | 0.37 | 0.31 | 0.56 | 0.86679 | 0.82 | 0.83
| Mean |  | 0.081667 | 0.55657 |  | 0.40833 | 0.66667 | 0.34833 | 0.40833 | 0.33 | 0.535 | 0.84864 | 0.765 | 0.82333
| Median |  | 0.085 | 0.55097 |  | 0.41 | 0.655 | 0.35 | 0.415 | 0.325 | 0.545 | 0.84411 | 0.765 | 0.835
| Std Dev. |  | 0.01472 | 0.017091 |  | 0.01472 | 0.039833 | 0.026394 | 0.02137 | 0.022804 | 0.037815 | 0.017909 | 0.033912 | 0.03266
| Rel.Std.Dev. |  | 18.02% | 3.071% |  | 3.605% | 5.975% | 7.577% | 5.233% | 6.91% | 7.068% | 2.11% | 4.433% | 3.967%
| PDM3 |  | -84.49% | 5.692% |  | -22.46% | 26.6% | -33.85% | -22.46% | -37.33% | 1.596% | 61.16% | 45.27% | 56.35%
| Z-Score (Absolute) |  | 1.56 | 0.34 | 1.69 | 0.25 | 0.78 | 0.49 | 0.25 | 0.57 | 0.25 | 1.51 | 1.18 | 1.41
| Analytical results for Hg in OREAS 251c (Indicative Value 0.027 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 15g | 0.1g | 01g | 40g | [6 empty] | 3
| 1 | 1 | < 1 | 0.06 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.03 | 0.04 | < 0.05 | 0.07 | 0.033 | 0.016 | 0.05 | < 5 | < 0.5 | [6 empty] | 1
| 1 | 2 | < 1 | 0.03 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | 0.03 | < 0.05 | 0.07 | 0.009 | 0.008 | 0.03 | < 5 | < 0.5 | [6 empty] | 1
| 1 | 3 | < 1 | 0.03 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.03 | 0.03 | < 0.05 | 0.06 | 0.024 | 0.009 | 0.04 | < 5 | < 0.5 | [6 empty] | 16
| 1 | 4 | < 1 | 0.03 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | < 0.05 | 0.06 | 0.021 | 0.013 | 0.02 | < 5 | < 0.5 | [6 empty] | 0.026767
| 1 | 5 | < 1 | 0.05 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.02 | 0.03 | < 0.05 | 0.06 | 0.024 | 0.011 | 0.04 | < 5 | < 0.5 | [6 empty] | 10
| 1 | 6 | < 1 | 0.04 | < 0.01 | 0.01 | 0.01 | 0.02 | 0.04 | 0.03 | < 0.05 | 0.07 | 0.016 | 0.012 | 0.03 | < 5 | < 0.5
| Mean |  |  | 0.04 |  | 0.01 | 0.01 | 0.016667 | 0.026667 | 0.031667 |  | 0.065 | 0.021167 | 0.0115 | 0.035
| Median |  |  | 0.035 |  | 0.01 | 0.01 | 0.02 | 0.025 | 0.03 |  | 0.065 | 0.0225 | 0.0115 | 0.035
| Std Dev. |  |  | 0.012649 |  | 0 | 0 | 0.005164 | 0.008165 | 0.0040825 |  | 0.0054772 | 0.0081343 | 0.002881 | 0.010488
| Rel.Std.Dev. |  |  | 31.62% |  | 0% | 0% | 30.98% | 30.62% | 12.89% |  | 8.427% | 38.43% | 25.05% | 29.97%
| PDM3 |  |  | 49.44% |  | -62.64% | -62.64% | -37.73% | -0.3736% | 18.31% |  | 142.8% | -20.92% | -57.04% | 30.76%
| Z-Score (Absolute) |  | 21.04 | 0.59 | 0.96 | 0.74 | 0.74 | 0.44 | 0 | 0.22 | 0.07 | 1.7 | 0.24 | 0.67 | 0.37 | 109.96 | 9.93
| Analytical results for Ho in OREAS 251c (Certified Value 0.38 ppm) | [22 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.345 | 0.34148 | 0.41872 | 0.4 | 0.38 | [16 empty] | 1
| 1 | 2 | 0.342 | 0.34754 | 0.472 | 0.4 | 0.36 | [16 empty] | 24
| 1 | 3 | 0.355 | 0.35641 | 0.45304 | 0.4 | 0.375 | [16 empty] | 16
| 1 | 4 | 0.331 | 0.33934 | 0.44508 | 0.4 | 0.37 | [16 empty] | 0.38057
| 1 | 5 | 0.351 | 0.35255 | 0.46348 | 0.4 | 0.37 | [16 empty] | 90
| 1 | 6 | 0.346 | 0.35289 | 0.38968 | 0.4 | 0.36
| Mean |  | 0.345 | 0.34837 | 0.44033 | 0.4 | 0.36917
| Median |  | 0.3455 | 0.35004 | 0.44906 | 0.4 | 0.37
| Std Dev. |  | 0.0082704 | 0.0068142 | 0.03084 | 6.0809e-17 | 0.0080104
| Rel.Std.Dev. |  | 2.397% | 1.956% | 7.004% | 1.52e-14% | 2.17%
| PDM3 |  | -9.347% | -8.463% | 15.7% | 5.105% | -2.997%
| Z-Score (Absolute) |  | 0.67 | 0.58 | 1.99 | 0.86 | 0
| Analytical results for In in OREAS 251c (Certified Value 0.026 ppm) | [22 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.03 | 0.087 | 0.025 | 0.026 | 0.023 | 0.026 | 0.025 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.04 | 0.087 | 0.028 | 0.026 | 0.023 | 0.027 | 0.024 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 25
| 1 | 3 | 0.02 | 0.04 | 0.083 | 0.026 | 0.026 | 0.024 | 0.026 | 0.024 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 16
| 1 | 4 | 0.02 | 0.04 | 0.082 | 0.028 | 0.026 | 0.024 | 0.026 | 0.026 | < 0.05 | < 0.05 | 0.04 | < 0.5 | [9 empty] | 0.025905
| 1 | 5 | 0.02 | 0.03 | 0.088 | 0.024 | 0.027 | 0.025 | 0.023 | 0.027 | < 0.05 | < 0.05 | 0.03 | < 0.5 | [9 empty] | 91
| 1 | 6 | 0.03 | 0.04 | 0.085 | 0.026 | 0.028 | 0.026 | 0.025 | 0.025 | < 0.05 | < 0.05 | 0.03 | < 0.5
| Mean |  | 0.023333 | 0.036667 | 0.085333 | 0.026167 | 0.0265 | 0.024167 | 0.0255 | 0.025167 |  |  | 0.033333
| Median |  | 0.02 | 0.04 | 0.086 | 0.026 | 0.026 | 0.024 | 0.026 | 0.025 |  |  | 0.03
| Std Dev. |  | 0.005164 | 0.005164 | 0.0024221 | 0.0016021 | 0.00083666 | 0.001169 | 0.0013784 | 0.001169 |  |  | 0.005164
| Rel.Std.Dev. |  | 22.13% | 14.08% | 2.838% | 6.123% | 3.157% | 4.837% | 5.406% | 4.645% |  |  | 15.49%
| PDM3 |  | -9.926% | 41.54% | 229.4% | 1.011% | 2.298% | -6.71% | -1.563% | -2.849% |  |  | 28.68%
| Z-Score (Absolute) |  | 1.35 | 5.84 | 32.1 | 0.18 | 0.36 | 0.9 | 0.18 | 0.36 | 0.45 | 0.45 | 4.05 | 120.93
| Analytical results for K in OREAS 251c (Certified Value 0.186 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.1868 | 0.20144 | 0.15603 | 0.24 | 0.2 | 0.18 | 0.2 | 0.17 | 0.18 | 0.17 | 0.17676 | 0.276 | 0.18 | 0.2641 | 0.16 | 0.334 | 0.21 | 0.16 | 0.2015 |  |  | 1
| 1 | 2 | 0.1905 | 0.19913 | 0.15471 | 0.24 | 0.19 | 0.18 | 0.2 | 0.18 | 0.18 | 0.17 | 0.18819 | 0.279 | 0.19 | 0.2562 | 0.16 | 0.425 | 0.21 | 0.16 | 0.20395 |  |  | #N/A
| 1 | 3 | 0.1871 | 0.19655 | 0.15507 | 0.23 | 0.18 | 0.18 | 0.2 | 0.17 | 0.18 | 0.17 | 0.17689 | 0.286 | 0.19 | 0.2559 | 0.17 | 0.338 | 0.21 | 0.17 | 0.21052 |  |  | 16
| 1 | 4 | 0.1887 | 0.19093 | 0.15347 | 0.23 | 0.19 | 0.18 | 0.2 | 0.18 | 0.19 | 0.17 | 0.18359 | 0.272 | 0.19 | 0.2546 | 0.17 | 0.416 | 0.22 | 0.17 | 0.19282 |  |  | 0.18638
| 1 | 5 | 0.1874 | 0.19617 | 0.1622 | 0.23 | 0.2 | 0.18 | 0.21 | 0.18 | 0.18 | 0.17 | 0.18326 | 0.288 | 0.17 | 0.2703 | 0.17 | 0.334 | 0.21 | 0.16 | 0.20263 |  |  | 92
| 1 | 6 | 0.1881 | 0.19749 | 0.15059 | 0.22 | 0.19 | 0.18 | 0.21 | 0.18 | 0.18 | 0.17 | 0.18421 | 0.282 | 0.18 | 0.2709 | 0.17 | 0.417 | 0.21 | 0.16 | 0.1939
| Mean |  | 0.1881 | 0.19695 | 0.15534 | 0.23167 | 0.19167 | 0.18 | 0.20333 | 0.17667 | 0.18167 | 0.17 | 0.18215 | 0.2805 | 0.18333 | 0.262 | 0.16667 | 0.37733 | 0.21167 | 0.16333 | 0.20089
| Median |  | 0.18775 | 0.19702 | 0.15489 | 0.23 | 0.19 | 0.18 | 0.2 | 0.18 | 0.18 | 0.17 | 0.18342 | 0.2805 | 0.185 | 0.26015 | 0.17 | 0.377 | 0.21 | 0.16 | 0.20207
| Std Dev. |  | 0.0013638 | 0.0035249 | 0.0038508 | 0.0075277 | 0.0075277 | 3.0405e-17 | 0.005164 | 0.005164 | 0.0040825 | 0 | 0.004489 | 0.0060581 | 0.008165 | 0.0074581 | 0.005164 | 0.046137 | 0.0040825 | 0.005164 | 0.0066265
| Rel.Std.Dev. |  | 0.725% | 1.79% | 2.479% | 3.249% | 3.928% | 1.689e-14% | 2.54% | 2.923% | 2.247% | 0% | 2.465% | 2.16% | 4.454% | 2.847% | 3.098% | 12.23% | 1.929% | 3.162% | 3.299%
| PDM3 |  | 0.9236% | 5.673% | -16.65% | 24.3% | 2.837% | -3.422% | 9.097% | -5.211% | -2.528% | -8.788% | -2.27% | 50.5% | -1.634% | 40.57% | -10.58% | 102.5% | 13.57% | -12.36% | 7.784%
| Z-Score (Absolute) |  | 0 | 0.39 | 1.45 | 1.93 | 0.16 | 0.36 | 0.67 | 0.51 | 0.28 | 0.8 | 0.26 | 4.09 | 0.21 | 3.27 | 0.95 | 8.38 | 1.04 | 1.1 | 0.57
| Analytical results for La in OREAS 251c (Certified Value 26.7 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 1
| 1 | 1 | 25.247 | 26.6023 | 45.8208 | 30.9 | 26 | 25.8 | 27.9 | 24.7 | 25.9 | 25.7 | 27.2474 | 31.9 | 25.2 | 32 | 22.3 | 40.6 | 23.5 | 34.4 | 23.213 |  |  | 1
| 1 | 2 | 26.174 | 26.6997 | 45.7188 | 31.3 | 26 | 26.6 | 28.3 | 25.2 | 27.2 | 24.6 | 26.8972 | 33.3 | 26.8 | 31 | 21.8 | 40.8 | 23.2 | 34.8 | 23.255 |  |  | 26
| 1 | 3 | 24.53 | 26.6791 | 45.7776 | 30.5 | 25 | 25.9 | 28.9 | 24.8 | 26.3 | 25.5 | 26.7911 | 33.6 | 26 | 31 | 24.6 | 37.1 | 23.2 | 36.4 | 23.775 |  |  | 16
| 1 | 4 | 25.309 | 25.4421 | 45.7672 | 30.5 | 26 | 26.3 | 29.5 | 25 | 26.3 | 25.8 | 27.9736 | 32.2 | 27.7 | 32 | 24.1 | 40.4 | 24.6 | 37.8 | 21.707 |  |  | 26.6631
| 1 | 5 | 26.063 | 26.3463 | 45.7048 | 30.8 | 26 | 26.7 | 28.4 | 24.9 | 27 | 25.3 | 27.6265 | 33.2 | 22.3 | 31 | 22.1 | 34.5 | 23.5 | 36 | 22.849 |  |  | 93
| 1 | 6 | 25.472 | 26.4955 | 45.756 | 30 | 26 | 25.7 | 28.6 | 24.5 | 26.7 | 25.7 | 27.4975 | 32.4 | 24.4 | 32 | 23.2 | 39.8 | 23.5 | 35.8 | 22.56
| Mean |  | 25.4658 | 26.3775 | 45.7575 | 30.6667 | 25.8333 | 26.1667 | 28.6 | 24.85 | 26.5667 | 25.4333 | 27.3389 | 32.7667 | 25.4 | 31.5 | 23.0167 | 38.8667 | 23.5833 | 35.8667 | 22.8932
| Median |  | 25.3905 | 26.5489 | 45.7616 | 30.65 | 26 | 26.1 | 28.5 | 24.85 | 26.5 | 25.6 | 27.3724 | 32.8 | 25.6 | 31.5 | 22.75 | 40.1 | 23.5 | 35.9 | 23.031
| Std Dev. |  | 0.60107 | 0.47648 | 0.04191 | 0.44121 | 0.40825 | 0.4274 | 0.55136 | 0.2429 | 0.48854 | 0.44572 | 0.45042 | 0.68896 | 1.911 | 0.54772 | 1.1444 | 2.5359 | 0.51929 | 1.2111 | 0.71145
| Rel.Std.Dev. |  | 2.36% | 1.806% | 0.09159% | 1.439% | 1.58% | 1.633% | 1.928% | 0.9775% | 1.839% | 1.753% | 1.648% | 2.103% | 7.524% | 1.739% | 4.972% | 6.525% | 2.202% | 3.377% | 3.108%
| PDM3 |  | -4.49% | -1.071% | 71.61% | 15.02% | -3.112% | -1.862% | 7.264% | -6.8% | -0.3616% | -4.612% | 2.535% | 22.89% | -4.737% | 18.14% | -13.68% | 45.77% | -11.55% | 34.52% | -14.14%
| Z-Score (Absolute) |  | 0.28 | 0 | 5.88 | 1.3 | 0.17 | 0.06 | 0.67 | 0.46 | 0.06 | 0.29 | 0.29 | 1.94 | 0.3 | 1.55 | 1.02 | 3.79 | 0.85 | 2.88 | 1.06
| Analytical results for Li in OREAS 251c (Certified Value 6.78 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 3
| 1 | 1 | 7 | < 10 | 6.8307 | 9.4 | 6.4 | 6.3 | 7.5 | 6.4 | 6.9 | 6.6 | 6.9732 | 10.1 | 11.3 | < 10 | 6.403 | [6 empty] | 1
| 1 | 2 | 7.1 | < 10 | 6.564 | 9.5 | 6.3 | 6.5 | 7.5 | 6.7 | 7 | 6.5 | 6.9506 | 10.3 | 12.2 | < 10 | 6.862 | [6 empty] | #N/A
| 1 | 3 | 6.7 | < 10 | 6.9595 | 9.1 | 5.8 | 6.5 | 7.6 | 6.7 | 6.8 | 6.5 | 7.3469 | 10 | 10.5 | < 10 | 6.957 | [6 empty] | 16
| 1 | 4 | 7 | < 10 | 6.9242 | 9.2 | 5.7 | 6.2 | 7.6 | 6.7 | 7 | 6.6 | 7.1618 | 10.1 | 13.2 | < 10 | 4.488 | [6 empty] | 6.7796
| 1 | 5 | 7.2 | < 10 | 6.8142 | 9.1 | 6.3 | 6.5 | 7.7 | 6.8 | 6.8 | 6.6 | 7.2481 | 10.1 | 12.7 | < 10 | 6.634 | [6 empty] | 94
| 1 | 6 | 7.2 | < 10 | 7.0361 | 8.8 | 5.9 | 6.4 | 7.6 | 6.7 | 6.8 | 6.5 | 6.9142 | 10 | 12 | < 10 | 6.171
| Mean |  | 7.0333 |  | 6.8548 | 9.1833 | 6.0667 | 6.4 | 7.5833 | 6.6667 | 6.8833 | 6.55 | 7.0991 | 10.1 | 11.9833 |  | 6.2525
| Median |  | 7.05 |  | 6.8774 | 9.15 | 6.1 | 6.45 | 7.6 | 6.7 | 6.85 | 6.55 | 7.0675 | 10.1 | 12.1 |  | 6.5185
| Std Dev. |  | 0.18619 |  | 0.16462 | 0.24833 | 0.30111 | 0.12649 | 0.075277 | 0.13663 | 0.098319 | 0.054772 | 0.17867 | 0.10954 | 0.9704 |  | 0.91171
| Rel.Std.Dev. |  | 2.647% |  | 2.401% | 2.704% | 4.963% | 1.976% | 0.9927% | 2.049% | 1.428% | 0.8362% | 2.517% | 1.085% | 8.098% |  | 14.58%
| PDM3 |  | 3.743% |  | 1.109% | 35.46% | -10.52% | -5.599% | 11.86% | -1.666% | 1.53% | -3.387% | 4.713% | 48.98% | 76.76% |  | -7.775%
| Z-Score (Absolute) |  | 0.2 | 2.08 | 0 | 2.61 | 0.88 | 0.51 | 0.82 | 0.21 | 0.03 | 0.34 | 0.27 | 3.63 | 5.74 | 2.08 | 0.67
| Analytical results for Lu in OREAS 251c (Indicative Value 0.077 ppm) | [22 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 3
| 1 | 1 | 0.072 | 0.071901 | < 0.1 | 0.08 | [17 empty] | 1
| 1 | 2 | 0.076 | 0.079726 | < 0.1 | 0.085 | [17 empty] | 5
| 1 | 3 | 0.068 | 0.073062 | < 0.1 | 0.085 | [17 empty] | 16
| 1 | 4 | 0.074 | 0.072713 | < 0.1 | 0.085 | [17 empty] | 0.077131
| 1 | 5 | 0.075 | 0.072123 | < 0.1 | 0.085 | [17 empty] | 11
| 1 | 6 | 0.074 | 0.079825 | < 0.1 | 0.08
| Mean |  | 0.073167 | 0.074892 |  | 0.083333
| Median |  | 0.074 | 0.072888 |  | 0.085
| Std Dev. |  | 0.0028577 | 0.0038057 |  | 0.002582
| Rel.Std.Dev. |  | 3.906% | 5.082% |  | 3.098%
| PDM3 |  | -5.139% | -2.903% |  | 8.042%
| Z-Score (Absolute) |  | 0.11 | 0.11 | 3.19 | 1.23
| Analytical results for Mg in OREAS 251c (Certified Value 0.539 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.51 | 0.56867 | 0.47439 | 0.72 | 0.56 | 0.5 | 0.55 | 0.48 | 0.52 | 0.47 | 0.55828 | 0.63 | 0.52 | 0.6 | 0.46 | 0.642 | 0.57 | 0.53 | 0.55419 |  |  | 1
| 1 | 2 | 0.53 | 0.56639 | 0.46664 | 0.73 | 0.55 | 0.52 | 0.54 | 0.5 | 0.52 | 0.47 | 0.56034 | 0.64 | 0.54 | 0.59 | 0.44 | 0.668 | 0.55 | 0.55 | 0.56573 |  |  | #N/A
| 1 | 3 | 0.49 | 0.55133 | 0.46572 | 0.69 | 0.56 | 0.51 | 0.55 | 0.49 | 0.51 | 0.46 | 0.57287 | 0.65 | 0.54 | 0.59 | 0.48 | 0.61 | 0.56 | 0.55 | 0.56415 |  |  | 16
| 1 | 4 | 0.52 | 0.5408 | 0.46361 | 0.71 | 0.54 | 0.49 | 0.54 | 0.5 | 0.53 | 0.47 | 0.57446 | 0.64 | 0.53 | 0.59 | 0.48 | 0.712 | 0.58 | 0.55 | 0.5177 |  |  | 0.53883
| 1 | 5 | 0.52 | 0.55107 | 0.47689 | 0.71 | 0.56 | 0.51 | 0.55 | 0.49 | 0.52 | 0.47 | 0.56747 | 0.65 | 0.52 | 0.6 | 0.47 | 0.519 | 0.57 | 0.54 | 0.54482 |  |  | 95
| 1 | 6 | 0.52 | 0.55676 | 0.45134 | 0.69 | 0.55 | 0.5 | 0.55 | 0.49 | 0.51 | 0.47 | 0.56285 | 0.64 | 0.5 | 0.6 | 0.48 | 0.673 | 0.56 | 0.53 | 0.54319
| Mean |  | 0.515 | 0.55584 | 0.46643 | 0.70833 | 0.55333 | 0.505 | 0.54667 | 0.49167 | 0.51833 | 0.46833 | 0.56605 | 0.64167 | 0.525 | 0.595 | 0.46833 | 0.63733 | 0.565 | 0.54167 | 0.5483
| Median |  | 0.52 | 0.55404 | 0.46618 | 0.71 | 0.555 | 0.505 | 0.55 | 0.49 | 0.52 | 0.47 | 0.56516 | 0.64 | 0.525 | 0.595 | 0.475 | 0.655 | 0.565 | 0.545 | 0.5495
| Std Dev. |  | 0.013784 | 0.010452 | 0.0090448 | 0.016021 | 0.008165 | 0.010488 | 0.005164 | 0.0075277 | 0.0075277 | 0.0040825 | 0.0066705 | 0.0075277 | 0.015166 | 0.0054772 | 0.016021 | 0.06717 | 0.010488 | 0.0098319 | 0.017686
| Rel.Std.Dev. |  | 2.677% | 1.88% | 1.939% | 2.262% | 1.476% | 2.077% | 0.9446% | 1.531% | 1.452% | 0.8717% | 1.178% | 1.173% | 2.889% | 0.9205% | 3.421% | 10.54% | 1.856% | 1.815% | 3.226%
| PDM3 |  | -4.423% | 3.156% | -13.44% | 31.46% | 2.691% | -6.279% | 1.454% | -8.754% | -3.805% | -13.08% | 5.05% | 19.08% | -2.567% | 10.42% | -13.08% | 18.28% | 4.856% | 0.5257% | 1.756%
| Z-Score (Absolute) |  | 0.67 | 0.2 | 1.71 | 3.44 | 0.14 | 0.89 | 0 | 1.17 | 0.6 | 1.67 | 0.41 | 2.02 | 0.46 | 1.03 | 1.67 | 1.93 | 0.39 | 0.11 | 0.03
| Analytical results for Mn in OREAS 251c (Certified Value 0.023 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0208 | 0.023021 | 0.018421 | 0.0284 | 0.0219 | 0.0209 | 0.0232 | 0.0203 | 0.0215 | 0.0208 | 0.02398 | 0.0259 | 0.0227 | 0.0256 | 0.0185 | 0.0265 | 0.02482 | 0.02 | 0.023271 |  |  | 1
| 1 | 2 | 0.0214 | 0.023406 | 0.017978 | 0.029 | 0.0229 | 0.0217 | 0.0231 | 0.021 | 0.0214 | 0.0208 | 0.023907 | 0.0264 | 0.0236 | 0.0248 | 0.0177 | 0.0269 | 0.02419 | 0.021 | 0.02369 |  |  | 28
| 1 | 3 | 0.0199 | 0.02324 | 0.017953 | 0.0269 | 0.0211 | 0.0213 | 0.0232 | 0.0207 | 0.021 | 0.0206 | 0.024244 | 0.0268 | 0.0233 | 0.025 | 0.0195 | 0.025 | 0.02479 | 0.021 | 0.02378 |  |  | 16
| 1 | 4 | 0.0211 | 0.022463 | 0.017117 | 0.0278 | 0.0223 | 0.0204 | 0.0232 | 0.0211 | 0.0212 | 0.0206 | 0.02474 | 0.0257 | 0.023 | 0.0245 | 0.0194 | 0.0269 | 0.02477 | 0.021 | 0.02125 |  |  | 0.02267
| 1 | 5 | 0.021 | 0.023115 | 0.018528 | 0.0283 | 0.0222 | 0.0211 | 0.0233 | 0.021 | 0.0209 | 0.0209 | 0.024831 | 0.0267 | 0.0222 | 0.0248 | 0.0185 | 0.0215 | 0.02431 | 0.021 | 0.023068 |  |  | 96
| 1 | 6 | 0.0214 | 0.022925 | 0.017848 | 0.0272 | 0.0213 | 0.0206 | 0.0234 | 0.0209 | 0.0211 | 0.0209 | 0.024533 | 0.0264 | 0.0217 | 0.0254 | 0.0191 | 0.0256 | 0.02437 | 0.02 | 0.022626
| Mean |  | 0.020933 | 0.023028 | 0.017974 | 0.027933 | 0.02195 | 0.021 | 0.023233 | 0.020833 | 0.021183 | 0.020767 | 0.024373 | 0.026317 | 0.02275 | 0.025017 | 0.018783 | 0.0254 | 0.024542 | 0.020667 | 0.022947
| Median |  | 0.02105 | 0.023068 | 0.017966 | 0.02805 | 0.02205 | 0.021 | 0.0232 | 0.02095 | 0.02115 | 0.0208 | 0.024389 | 0.0264 | 0.02285 | 0.0249 | 0.0188 | 0.02605 | 0.02457 | 0.021 | 0.023169
| Std Dev. |  | 0.00055737 | 0.00032408 | 0.00050144 | 0.00078909 | 0.00066858 | 0.00047329 | 0.00010328 | 0.00029439 | 0.00023166 | 0.00013663 | 0.00038941 | 0.00043551 | 0.0007064 | 0.00041191 | 0.0006824 | 0.0020552 | 0.00028216 | 0.0005164 | 0.00093247
| Rel.Std.Dev. |  | 2.663% | 1.407% | 2.79% | 2.825% | 3.046% | 2.254% | 0.4445% | 1.413% | 1.094% | 0.6579% | 1.598% | 1.655% | 3.105% | 1.647% | 3.633% | 8.091% | 1.15% | 2.499% | 4.064%
| PDM3 |  | -7.659% | 1.582% | -20.71% | 23.22% | -3.175% | -7.365% | 2.486% | -8.1% | -6.557% | -8.395% | 7.511% | 16.09% | 0.3543% | 10.35% | -17.14% | 12.04% | 8.258% | -8.836% | 1.225%
| Z-Score (Absolute) |  | 0.67 | 0.1 | 1.77 | 1.92 | 0.3 | 0.65 | 0.18 | 0.71 | 0.58 | 0.74 | 0.6 | 1.32 | 0 | 0.84 | 1.47 | 0.98 | 0.67 | 0.77 | 0.07
| Analytical results for Mo in OREAS 251c (Certified Value 1.96 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 3
| 1 | 1 | 2 | 1.9795 | 2.0271 | 2.1 | 1.82 | 1.96 | 1.94 | 1.82 | 1.89 | 2.18 | 1.5801 | 1.9 | 2.1 | 1.71 | 2 | 1.92 | 1.8 | 6.1 | 2 | 2.538 |  | 1
| 1 | 2 | 2.1 | 2.0869 | 2.0058 | 2.1 | 1.87 | 2.02 | 2 | 1.79 | 1.98 | 2.06 | 1.4471 | 2 | 2.1 | 1.77 | 2 | 1.85 | 1.9 | 6 | 2 | 2.582 |  | 29
| 1 | 3 | 2 | 2.0945 | 1.9596 | 2.05 | 1.82 | 2.02 | 1.98 | 1.85 | 1.87 | 2.07 | 1.4326 | 1.8 | 2.1 | 1.82 | 2 | 1.96 | 1.6 | 5.9 | 2.2 | 2.19 |  | 16
| 1 | 4 | 2 | 1.9788 | 1.9683 | 2.03 | 1.96 | 2 | 2.03 | 1.8 | 1.88 | 2.08 | 1.3621 | 1.9 | 2.1 | 1.84 | 2 | 1.92 | 1.8 | 6 | 2.2 | 2.509 |  | 1.9552
| 1 | 5 | 2 | 1.9744 | 2.069 | 2.18 | 1.83 | 1.94 | 2 | 1.8 | 1.9 | 2.16 | 1.5628 | 1.7 | 2.1 | 1.53 | 2 | 1.89 | 1.6 | 6 | 2.2 | 2.848 |  | 97
| 1 | 6 | 2.1 | 2.1322 | 2.008 | 2.03 | 1.85 | 1.94 | 2.02 | 1.77 | 2 | 2.17 | 1.4385 | 1.9 | 2 | 1.71 | 2 | 1.92 | 1.7 | 6.4 | 2 | 2.793
| Mean |  | 2.0333 | 2.0411 | 2.0063 | 2.0817 | 1.8583 | 1.98 | 1.995 | 1.805 | 1.92 | 2.12 | 1.4705 | 1.8667 | 2.0833 | 1.73 | 2 | 1.91 | 1.7333 | 6.0667 | 2.1 | 2.5767
| Median |  | 2 | 2.0332 | 2.0069 | 2.075 | 1.84 | 1.98 | 2 | 1.8 | 1.895 | 2.12 | 1.4428 | 1.9 | 2.1 | 1.74 | 2 | 1.92 | 1.75 | 6 | 2.1 | 2.56
| Std Dev. |  | 0.05164 | 0.071245 | 0.039987 | 0.057764 | 0.053448 | 0.037947 | 0.032094 | 0.027386 | 0.055498 | 0.055498 | 0.084026 | 0.10328 | 0.040825 | 0.11189 | 0 | 0.036878 | 0.12111 | 0.17512 | 0.10954 | 0.23496
| Rel.Std.Dev. |  | 2.54% | 3.491% | 1.993% | 2.775% | 2.876% | 1.917% | 1.609% | 1.517% | 2.891% | 2.618% | 5.714% | 5.533% | 1.96% | 6.468% | 0% | 1.931% | 6.987% | 2.887% | 5.216% | 9.119%
| PDM3 |  | 3.995% | 4.39% | 2.612% | 6.467% | -4.956% | 1.267% | 2.034% | -7.684% | -1.802% | 8.427% | -24.79% | -4.53% | 6.552% | -11.52% | 2.29% | -2.313% | -11.35% | 210.3% | 7.404% | 31.78%
| Z-Score (Absolute) |  | 0.25 | 0.3 | 0.07 | 0.56 | 0.88 | 0.1 | 0 | 1.22 | 0.48 | 0.8 | 3.37 | 0.82 | 0.57 | 1.7 | NA | 0.55 | 1.68 | 26.15 | 0.67 | 3.74
Results from laboratory 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 251c (Certified Value 0.192 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.2 | 0.20464 | 0.2015 | 0.192 | 0.19 | 0.18 | 0.2 | 0.18 | 0.2 | 0.18 | 0.18333 | 0.2 | 0.189 | 0.21 | 0.181 | 0.176 | 0.13 | 0.19 | 0.18932 |  |  | 1
| 1 | 2 | 0.2 | 0.20538 | 0.20153 | 0.201 | 0.19 | 0.18 | 0.2 | 0.19 | 0.2 | 0.18 | 0.18247 | 0.2 | 0.194 | 0.2 | 0.177 | 0.186 | 0.12 | 0.2 | 0.18908 |  |  | #N/A
| 1 | 3 | 0.19 | 0.19943 | 0.20192 | 0.195 | 0.19 | 0.18 | 0.2 | 0.19 | 0.19 | 0.17 | 0.18516 | 0.21 | 0.195 | 0.2 | 0.186 | 0.17 | 0.12 | 0.2 | 0.1792 |  |  | 16
| 1 | 4 | 0.2 | 0.19584 | 0.20049 | 0.194 | 0.19 | 0.18 | 0.2 | 0.19 | 0.2 | 0.18 | 0.18503 | 0.21 | 0.191 | 0.2 | 0.186 | 0.193 | 0.13 | 0.2 | 0.17297 |  |  | 0.19166
| 1 | 5 | 0.2 | 0.19733 | 0.2034 | 0.195 | 0.19 | 0.18 | 0.2 | 0.19 | 0.19 | 0.18 | 0.19299 | 0.21 | 0.184 | 0.2 | 0.183 | 0.148 | 0.13 | 0.2 | 0.16852 |  |  | 98
| 1 | 6 | 0.2 | 0.20646 | 0.19594 | 0.192 | 0.19 | 0.18 | 0.2 | 0.19 | 0.19 | 0.18 | 0.19234 | 0.22 | 0.184 | 0.21 | 0.186 | 0.186 | 0.12 | 0.19 | 0.18657
| Mean |  | 0.19833 | 0.20151 | 0.2008 | 0.19483 | 0.19 | 0.18 | 0.2 | 0.18833 | 0.195 | 0.17833 | 0.18689 | 0.20833 | 0.1895 | 0.20333 | 0.18317 | 0.1765 | 0.125 | 0.19667 | 0.18094
| Median |  | 0.2 | 0.20203 | 0.20152 | 0.1945 | 0.19 | 0.18 | 0.2 | 0.19 | 0.195 | 0.18 | 0.18509 | 0.21 | 0.19 | 0.2 | 0.1845 | 0.181 | 0.125 | 0.2 | 0.18288
| Std Dev. |  | 0.0040825 | 0.0045441 | 0.002561 | 0.0033116 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 0.0040825 | 0.0054772 | 0.0040825 | 0.0045939 | 0.0075277 | 0.0047645 | 0.005164 | 0.003656 | 0.016171 | 0.0054772 | 0.005164 | 0.0088201
| Rel.Std.Dev. |  | 2.058% | 2.255% | 1.275% | 1.7% | 1.6e-14% | 1.689e-14% | 1.52e-14% | 2.168% | 2.809% | 2.289% | 2.458% | 3.613% | 2.514% | 2.54% | 1.996% | 9.162% | 4.382% | 2.626% | 4.875%
| PDM3 |  | 3.483% | 5.142% | 4.768% | 1.657% | -0.8646% | -6.082% | 4.353% | -1.734% | 1.744% | -6.952% | -2.489% | 8.701% | -1.125% | 6.092% | -4.43% | -7.908% | -34.78% | 2.614% | -5.59%
| Z-Score (Absolute) |  | 0.67 | 0.91 | 0.86 | 0.39 | 0.02 | 0.76 | NA | 0.11 | 0.41 | 0.89 | 0.22 | 1.44 | 0.02 | 1.05 | 0.51 | 1.03 | 5.02 | 0.54 | 0.68
Results from laboratory 1.10 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb in OREAS 251c (Certified Value 0.62 ppm) | [22 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [9 empty] | 2
| 1 | 1 | 0.61 | 0.82355 | 0.5 | 0.32 | 0.63 | 0.74 | 0.52 | 0.53 | 0.78 | 1.1824 | 1.1 | 1.38 | [9 empty] | 1
| 1 | 2 | 0.58 | 0.81865 | 0.5 | 0.33 | 0.77 | 0.75 | 0.56 | 0.53 | 0.7 | 1.1766 | 1 | 1.5 | [9 empty] | 30
| 1 | 3 | 0.52 | 0.76207 | 0.5 | 0.31 | 0.64 | 0.73 | 0.53 | 0.54 | 0.77 | 1.2563 | 1 | 1.46 | [9 empty] | 16
| 1 | 4 | 0.58 | 0.78269 | 0.5 | 0.36 | 0.79 | 0.72 | 0.52 | 0.52 | 0.8 | 1.2895 | 1.1 | 1.59 | [9 empty] | 0.61826
| 1 | 5 | 0.61 | 0.91686 | 0.6 | 0.31 | 0.55 | 0.78 | 0.53 | 0.52 | 0.73 | 1.1544 | 1 | 1.24 | [9 empty] | 99
| 1 | 6 | 0.54 | 0.83444 | 0.5 | 0.33 | 0.59 | 0.76 | 0.49 | 0.48 | 0.79 | 1.2693 | 1.1 | 1.59
| Mean |  | 0.57333 | 0.82304 | 0.51667 | 0.32667 | 0.66167 | 0.74667 | 0.525 | 0.52 | 0.76167 | 1.2214 | 1.05 | 1.46
| Median |  | 0.58 | 0.8211 | 0.5 | 0.325 | 0.635 | 0.745 | 0.525 | 0.525 | 0.775 | 1.2193 | 1.05 | 1.48
| Std Dev. |  | 0.036697 | 0.053474 | 0.040825 | 0.018619 | 0.097245 | 0.021602 | 0.022583 | 0.020976 | 0.038687 | 0.056871 | 0.054772 | 0.13431
| Rel.Std.Dev. |  | 6.401% | 6.497% | 7.902% | 5.7% | 14.7% | 2.893% | 4.302% | 4.034% | 5.079% | 4.656% | 5.216% | 9.2%
| PDM3 |  | -7.266% | 33.12% | -16.43% | -47.16% | 7.022% | 20.77% | -15.08% | -15.89% | 23.2% | 97.56% | 69.83% | 136.1%
| Z-Score (Absolute) |  | 0.57 | 0.52 | NA | 1.64 | 0.18 | 0.18 | 0.78 | 0.8 | 0.25 | 2.25 | NA | 3.29
Results from laboratories 1.07 and 1.16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 251c (Indicative Value 25 ppm) | [22 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 1
| 1 | 1 | 22.358 | 24.1188 | 41.4852 | 26.5 | 26.3 | [16 empty] | 1
| 1 | 2 | 23.386 | 24.2075 | 41.5512 | 27.1 | 26.3 | [16 empty] | 6
| 1 | 3 | 21.71 | 24.5914 | 41.5424 | 25.5 | 26.7 | [16 empty] | 16
| 1 | 4 | 22.481 | 24.4158 | 41.4628 | 26.1 | 27.3 | [16 empty] | 24.9775
| 1 | 5 | 23.566 | 23.719 | 41.484 | 26.7 | 26.3 | [16 empty] | 12
| 1 | 6 | 22.732 | 24.4755 | 41.484 | 25.9 | 27
| Mean |  | 22.7055 | 24.2547 | 41.5016 | 26.3 | 26.65
| Median |  | 22.6065 | 24.3116 | 41.4846 | 26.3 | 26.5
| Std Dev. |  | 0.68796 | 0.31488 | 0.036107 | 0.57966 | 0.42778
| Rel.Std.Dev. |  | 3.03% | 1.298% | 0.087% | 2.204% | 1.605%
| PDM3 |  | -9.096% | -2.894% | 66.16% | 5.295% | 6.696%
| Z-Score (Absolute) |  | 1.19 | 0.67 | 5.01 | 0 | 0.12
| Analytical results for Ni in OREAS 251c (Certified Value 57 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 1
| 1 | 1 | 54.7 | 55.1974 | 51.468 | 59.15 | 59.7 | 57 | 59.2 | 57.6 | 53.3 | 54.6 | 54.4 | 56.0431 | 62 | 56.1 | 59 | 51.2 | 61.8 | 52.91 | 55 | 57.241 |  | 1
| 1 | 2 | 55.2 | 56.088 | 51.773 | 57.85 | 64.2 | 58 | 64.2 | 57.8 | 54.9 | 57.5 | 52.7 | 56.1483 | 62 | 57.2 | 57 | 49.9 | 61.8 | 51.72 | 56 | 57.413 |  | 15
| 1 | 3 | 55.5 | 54.6433 | 51.635 | 57.6 | 61.4 | 56 | 60.9 | 59.3 | 55.9 | 55.3 | 53.8 | 57.2229 | 62 | 56.9 | 57 | 53.8 | 58.3 | 52.25 | 56 | 57.585 |  | 16
| 1 | 4 | 55.7 | 53.7784 | 49.423 | 59.24 | 61 | 59 | 62 | 58.7 | 55.4 | 55.2 | 55.5 | 58.468 | 61 | 56 | 58 | 53.2 | 61.1 | 52.46 | 57 | 53.513 |  | 56.639
| 1 | 5 | 55.6 | 55.7384 | 52.75 | 59.2 | 61.1 | 59 | 60.6 | 60.4 | 55.2 | 55.9 | 55.5 | 57.9737 | 61 | 45.2 | 58 | 51.9 | 53.6 | 52.05 | 56 | 56.001 |  | 100
| 1 | 6 | 56.1 | 55.3585 | 51.62 | 60.2 | 60 | 57 | 58.3 | 59.1 | 54.7 | 56.5 | 55.5 | 57.2454 | 62 | 53.4 | 58 | 52.5 | 58.2 | 52.26 | 55 | 55.671
| Mean |  | 55.4667 | 55.134 | 51.4448 | 58.8733 | 61.2333 | 57.6667 | 60.8667 | 58.8167 | 54.9 | 55.8333 | 54.5667 | 57.1836 | 61.6667 | 54.1333 | 57.8333 | 52.0833 | 59.1333 | 52.275 | 55.8333 | 56.2373
| Median |  | 55.55 | 55.2779 | 51.6275 | 59.175 | 61.05 | 57.5 | 60.75 | 58.9 | 55.05 | 55.6 | 54.95 | 57.2342 | 62 | 56.05 | 58 | 52.2 | 59.7 | 52.255 | 56 | 56.621
| Std Dev. |  | 0.4761 | 0.8256 | 1.0924 | 0.9743 | 1.5983 | 1.2111 | 2.0897 | 1.0342 | 0.88769 | 1.0424 | 1.1587 | 0.96458 | 0.5164 | 4.5772 | 0.75277 | 1.4106 | 3.1709 | 0.39924 | 0.75277 | 1.5495
| Rel.Std.Dev. |  | 0.8583% | 1.497% | 2.123% | 1.655% | 2.61% | 2.1% | 3.433% | 1.758% | 1.617% | 1.867% | 2.124% | 1.687% | 0.8374% | 8.455% | 1.302% | 2.708% | 5.362% | 0.7637% | 1.348% | 2.755%
| PDM3 |  | -2.07% | -2.657% | -9.171% | 3.945% | 8.112% | 1.814% | 7.464% | 3.845% | -3.07% | -1.422% | -3.659% | 0.9616% | 8.877% | -4.424% | 2.109% | -8.043% | 4.404% | -7.705% | -1.422% | -0.7091%
| Z-Score (Absolute) |  | 0.21 | 0.33 | 1.67 | 1.03 | 1.89 | 0.59 | 1.76 | 1.01 | 0.41 | 0.07 | 0.54 | 0.42 | 2.05 | 0.69 | 0.66 | 1.44 | 1.13 | 1.37 | 0.07 | 0.07
| Analytical results for P in OREAS 251c (Certified Value 0.065 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.061 | 0.066212 | 0.067848 | 0.068 | 0.0651 | 0.065 | 0.066 | 0.062 | 0.067 | 0.064 | 0.061163 | 0.0691 | 0.064 | 0.0658 | 0.059 | 0.0743 | 0.07 | 0.062 | 0.070625 |  |  | 1
| 1 | 2 | 0.0618 | 0.067113 | 0.066052 | 0.067 | 0.0646 | 0.067 | 0.066 | 0.063 | 0.067 | 0.064 | 0.060032 | 0.0707 | 0.065 | 0.0625 | 0.057 | 0.0743 | 0.06 | 0.064 | 0.0709 |  |  | #N/A
| 1 | 3 | 0.061 | 0.066886 | 0.067845 | 0.068 | 0.0653 | 0.066 | 0.066 | 0.063 | 0.066 | 0.063 | 0.061716 | 0.0708 | 0.065 | 0.064 | 0.062 | 0.0713 | 0.06 | 0.066 | 0.071408 |  |  | 16
| 1 | 4 | 0.0616 | 0.064069 | 0.062912 | 0.066 | 0.0642 | 0.064 | 0.065 | 0.064 | 0.067 | 0.064 | 0.061574 | 0.0696 | 0.064 | 0.0634 | 0.06 | 0.0801 | 0.07 | 0.064 | 0.068141 |  |  | 0.064912
| 1 | 5 | 0.061 | 0.065467 | 0.070089 | 0.069 | 0.0647 | 0.065 | 0.066 | 0.062 | 0.067 | 0.064 | 0.063525 | 0.0701 | 0.065 | 0.064 | 0.06 | 0.0698 | 0.06 | 0.064 | 0.070335 |  |  | 101
| 1 | 6 | 0.0611 | 0.065508 | 0.065613 | 0.067 | 0.0652 | 0.065 | 0.067 | 0.063 | 0.066 | 0.064 | 0.06315 | 0.0704 | 0.062 | 0.0626 | 0.06 | 0.0718 | 0.06 | 0.064 | 0.070138
| Mean |  | 0.06125 | 0.065876 | 0.066727 | 0.0675 | 0.06485 | 0.065333 | 0.066 | 0.062833 | 0.066667 | 0.063833 | 0.06186 | 0.070117 | 0.064167 | 0.063717 | 0.059667 | 0.0736 | 0.063333 | 0.064 | 0.070258
| Median |  | 0.06105 | 0.06586 | 0.066949 | 0.0675 | 0.0649 | 0.065 | 0.066 | 0.063 | 0.067 | 0.064 | 0.061645 | 0.07025 | 0.0645 | 0.0637 | 0.06 | 0.07305 | 0.06 | 0.064 | 0.07048
| Std Dev. |  | 0.00035637 | 0.0011163 | 0.002452 | 0.0010488 | 0.00042308 | 0.0010328 | 0.00063246 | 0.00075277 | 0.0005164 | 0.00040825 | 0.0012937 | 0.00066156 | 0.001169 | 0.0012106 | 0.001633 | 0.0036387 | 0.005164 | 0.0012649 | 0.0011289
| Rel.Std.Dev. |  | 0.5818% | 1.695% | 3.675% | 1.554% | 0.6524% | 1.581% | 0.9583% | 1.198% | 0.7746% | 0.6396% | 2.091% | 0.9435% | 1.822% | 1.9% | 2.737% | 4.944% | 8.154% | 1.976% | 1.607%
| PDM3 |  | -5.641% | 1.486% | 2.796% | 3.987% | -0.09505% | 0.6495% | 1.677% | -3.202% | 2.704% | -1.661% | -4.701% | 8.019% | -1.148% | -1.841% | -8.08% | 13.38% | -2.432% | -1.405% | 8.236%
| Z-Score (Absolute) |  | 1.34 | 0.38 | 0.7 | 0.98 | 0 | 0.18 | 0.43 | 0.75 | 0.67 | 0.38 | 1.11 | 1.95 | 0.25 | 0.42 | 1.92 | 3.25 | 0.56 | 0.32 | 2.01
| Analytical results for Pb in OREAS 251c (Certified Value 8.01 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  | 2
| 1 | 1 | 7.1 | 9.8058 | 8.0713 | 8.6 | 8.7 | 6.8 | 7.9 | 6.6 | 7 | 7.2 | 6.8599 | 7 | 8.6 | 6.51 | 9 | 6.75 | 10.3 | 10.6 | 7 | 7.538 |  | 1
| 1 | 2 | 7.3 | 8.5963 | 6.4104 | 8.5 | 8.9 | 6.9 | 8.1 | 6.6 | 7.2 | 7.3 | 6.703 | 7 | 8.8 | 7.07 | 8 | 6.43 | 10.9 | 10.4 | 7 | 9.287 |  | 17
| 1 | 3 | 6.8 | 8.6618 | 6.905 | 8.5 | 8.8 | 6.8 | 7.9 | 6.5 | 7 | 7.4 | 13.3898 | 7 | 8.9 | 7.27 | 8 | 7.14 | 10.2 | 11 | 7 | 7.816 |  | 16
| 1 | 4 | 7 | 8.4001 | 6.9392 | 8.5 | 8.1 | 6.7 | 7.8 | 6.7 | 7.1 | 7.5 | 8.5486 | 7 | 8.6 | 7.32 | 8 | 7.11 | 10.9 | 10.3 | 7 | 6.72 |  | 8.009
| 1 | 5 | 7.2 | 8.9935 | 7.8964 | 8.6 | 8.5 | 6.9 | 8 | 7.9 | 6.9 | 8.9 | 9.4685 | 6 | 9.1 | 7.55 | 8 | 6.71 | 8.6 | 10.6 | 7 | 7.845 |  | 102
| 1 | 6 | 7.3 | 9.134 | 9.8721 | 8.2 | 8.2 | 6.8 | 7.6 | 6.5 | 7 | 7.7 | 14.9301 | 156 | 8.6 | 6.56 | 9 | 6.86 | 11 | 10.4 | 7 | 7.972
| Mean |  | 7.1167 | 8.9319 | 7.6824 | 8.4833 | 8.5333 | 6.8167 | 7.8833 | 6.8 | 7.0333 | 7.6667 | 9.9833 | 31.6667 | 8.7667 | 7.0467 | 8.3333 | 6.8333 | 10.3167 | 10.55 | 7 | 7.863
| Median |  | 7.15 | 8.8276 | 7.4178 | 8.5 | 8.6 | 6.8 | 7.9 | 6.6 | 7 | 7.45 | 9.0086 | 7 | 8.7 | 7.17 | 8 | 6.805 | 10.6 | 10.5 | 7 | 7.8305
| Std Dev. |  | 0.19408 | 0.50532 | 1.2463 | 0.1472 | 0.3266 | 0.075277 | 0.17224 | 0.54406 | 0.10328 | 0.62823 | 3.4326 | 60.912 | 0.20656 | 0.42505 | 0.5164 | 0.26688 | 0.90646 | 0.251 | 0 | 0.83118
| Rel.Std.Dev. |  | 2.727% | 5.658% | 16.22% | 1.735% | 3.827% | 1.104% | 2.185% | 8.001% | 1.468% | 8.194% | 34.38% | 192.4% | 2.356% | 6.032% | 6.197% | 3.906% | 8.786% | 2.379% | 0% | 10.57%
| PDM3 |  | -11.14% | 11.52% | -4.077% | 5.923% | 6.547% | -14.89% | -1.569% | -15.1% | -12.18% | -4.274% | 24.65% | 295.4% | 9.461% | -12.02% | 4.05% | -14.68% | 28.81% | 31.73% | -12.6% | -1.822%
| Z-Score (Absolute) |  | 0.61 | 0.87 | 0.15 | 0.5 | 0.54 | 0.85 | 0.02 | 0.86 | 0.67 | 0.16 | 1.72 | NA | 0.73 | 0.66 | NA | 0.84 | 1.99 | 2.18 | NA | 0
Results from laboratories 1.15, 1.18 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 251c (Indicative Value < 10 ppb) | [22 empty] | Indicative
Pd | Round | Replicate | Lab | [20 empty] | 1
| No | No | 1.14 | [20 empty] | ppb
|  |  | AR*OES/MS | [20 empty] | 1
|  |  | 01g | [20 empty] | 1
| 1 | 1 | < 10 | [20 empty] | 1
| 1 | 2 | < 10 | [20 empty] | 7
| 1 | 3 | < 10 | [20 empty] | 16
| 1 | 4 | < 10 | [20 empty] | < 10
| 1 | 5 | < 10 | [20 empty] | 13
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 251c (Certified Value 7.4 ppm) | [22 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 5.968 | 6.1693 | 10.6547 | 7.3 | 6.99 | [16 empty] | 1
| 1 | 2 | 6.106 | 6.2783 | 10.683 | 7.3 | 6.72 | [16 empty] | 34
| 1 | 3 | 5.689 | 6.3728 | 10.6494 | 7.2 | 7.09 | [16 empty] | 16
| 1 | 4 | 5.881 | 6.0191 | 10.7384 | 7.1 | 6.8 | [16 empty] | 7.4
| 1 | 5 | 6.09 | 6.1243 | 10.8016 | 7.2 | 7 | [16 empty] | 103
| 1 | 6 | 5.943 | 6.2479 | 10.7442 | 7.1 | 7.04
| Mean |  | 5.9462 | 6.2019 | 10.7119 | 7.2 | 6.94
| Median |  | 5.9555 | 6.2086 | 10.7107 | 7.2 | 6.995
| Std Dev. |  | 0.15304 | 0.12463 | 0.05965 | 0.089443 | 0.14601
| Rel.Std.Dev. |  | 2.574% | 2.01% | 0.5569% | 1.242% | 2.104%
| PDM3 |  | -19.65% | -16.19% | 44.76% | -2.703% | -6.216%
| Z-Score (Absolute) |  | 0.91 | 0.67 | 3.45 | 0.24 | 0
| Analytical results for Pt in OREAS 251c (Indicative Value < 5 ppb) | [22 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 1.14 | 1.21 | 1.28 | [18 empty] | ppb
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | [18 empty] | 2
|  |  | 01g | 0.1g | 40g | [18 empty] | 2
| 1 | 1 | < 5 | 2 | < 10 | [18 empty] | 1
| 1 | 2 | < 5 | 2 | < 10 | [18 empty] | 2
| 1 | 3 | < 5 | 5 | < 10 | [18 empty] | 16
| 1 | 4 | < 5 | 2 | < 10 | [18 empty] | < 5
| 1 | 5 | < 5 | < 1 | < 10 | [18 empty] | 14
| 1 | 6 | < 5 | 2 | < 10
| Mean |  |  | 2.6
| Median |  |  | 2
| Std Dev. |  |  | 1.3416
| Rel.Std.Dev. |  |  | 51.6%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 6.74
| Analytical results for Rb in OREAS 251c (Certified Value 13.1 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | [8 empty] | 2
| 1 | 1 | 12.99 | 13.832 | 15.1 | 13 | 12.4 | 12.7 | 11.8 | 12.6 | 12.9 | 13.0549 | 20.5 | 23.9 | 15.8 | [8 empty] | 1
| 1 | 2 | 13.25 | 14.4566 | 15 | 13.2 | 13.1 | 12.8 | 12.4 | 12.9 | 12.3 | 12.8949 | 20.9 | 30.3 | 16 | [8 empty] | 35
| 1 | 3 | 13.24 | 13.6532 | 14.6 | 11.6 | 13.1 | 12.7 | 12.2 | 12.5 | 12.6 | 13.0977 | 21.5 | 21 | 16.3 | [8 empty] | 16
| 1 | 4 | 13.1 | 13.5669 | 14.8 | 12.6 | 12.8 | 12.8 | 12.8 | 12.5 | 12.7 | 13.6369 | 20.3 | 28.2 | 16.4 | [8 empty] | 13.0655
| 1 | 5 | 13.26 | 14.0624 | 14.4 | 12.6 | 12.6 | 13 | 12.2 | 12 | 12.6 | 13.5545 | 21 | 24.7 | 16.6 | [8 empty] | 104
| 1 | 6 | 13.01 | 14.0298 | 14.3 | 12.6 | 12.9 | 12.6 | 11.8 | 12.5 | 13.2 | 13.4373 | 20.6 | 27.8 | 15.8
| Mean |  | 13.1417 | 13.9335 | 14.7 | 12.6 | 12.8167 | 12.7667 | 12.2 | 12.5 | 12.7167 | 13.2794 | 20.8 | 25.9833 | 16.15
| Median |  | 13.17 | 13.9309 | 14.7 | 12.6 | 12.85 | 12.75 | 12.2 | 12.5 | 12.65 | 13.2675 | 20.75 | 26.25 | 16.15
| Std Dev. |  | 0.12449 | 0.32336 | 0.32249 | 0.55136 | 0.27869 | 0.13663 | 0.37947 | 0.28983 | 0.30605 | 0.30321 | 0.42895 | 3.3961 | 0.33317
| Rel.Std.Dev. |  | 0.9473% | 2.321% | 2.194% | 4.376% | 2.174% | 1.07% | 3.11% | 2.319% | 2.407% | 2.283% | 2.062% | 13.07% | 2.063%
| PDM3 |  | 0.5833% | 6.644% | 12.51% | -3.562% | -1.904% | -2.287% | -6.624% | -4.328% | -2.67% | 1.637% | 59.2% | 98.87% | 23.61%
| Z-Score (Absolute) |  | 0 | 0.83 | 1.64 | 0.57 | 0.34 | 0.39 | 0.99 | 0.67 | 0.45 | 0.14 | 8.05 | 13.49 | 3.16
| Analytical results for Re in OREAS 251c (Certified Value < 0.001 ppm) | [22 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 36
| 1 | 3 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | < 0.001
| 1 | 5 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001 | [10 empty] | 105
| 1 | 6 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | < 0.05 | < 0.001
| Mean | [6 empty] | 0.001
| Median | [6 empty] | 0.001
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for S in OREAS 251c (Certified Value 0.01 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 3
| 1 | 1 | 0.0071 | 0.0078827 | 0.0043719 | 0.0070324 | 0.007 | 0.01 | 0.34 | 0.01 | 0.01 | 0.03 | 0.01 | 0.012832 | 0.0083 | < 0.01 | < 0.02 | 0.012 | < 0.02 | 0.0058 | 0.01 | 0.015 | 0.0093 | 1
| 1 | 2 | 0.0071 | 0.0077272 | 0.0035608 | 0.0057651 | 0.007 | 0.01 | 0.79 | 0.01 | 0.01 | 0.03 | 0.01 | 0.012602 | 0.008 | < 0.01 | < 0.02 | 0.0113 | < 0.02 | 0.0079 | 0.01 | 0.015 | 0.0149 | 21
| 1 | 3 | 0.0067 | 0.007733 | 0.0024306 | 0.0067573 | 0.008 | 0.01 | 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.01206 | 0.0077 | < 0.01 | < 0.02 | 0.0143 | < 0.02 | < 0.005 | 0.01 | 0.015 | 0.0052 | 16
| 1 | 4 | 0.0071 | 0.007589 | 0.0033278 | 0.0058586 | 0.008 | 0.01 | 0.75 | 0.01 | 0.01 | 0.03 | 0.01 | 0.014901 | 0.0081 | < 0.01 | < 0.02 | 0.0129 | < 0.02 | 0.0068 | 0.02 | 0.015 | 0.008 | 0.0097259
| 1 | 5 | 0.0071 | 0.0074136 | 0.0042029 | 0.0059646 | 0.008 | 0.01 | 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.016135 | 0.0079 | < 0.01 | < 0.02 | 0.0112 | < 0.02 | < 0.005 | 0.02 | 0.015 | 0.0106 | 106
| 1 | 6 | 0.0071 | 0.0077413 | 0.0037292 | 0.0065022 | 0.008 | 0.01 | 0.01 | 0.02 | 0.01 | 0.03 | < 0.01 | 0.015842 | 0.0079 | < 0.01 | < 0.02 | 0.0105 | < 0.02 | < 0.005 | 0.02 | 0.015 | 0.0157
| Mean |  | 0.0070333 | 0.0076811 | 0.0036039 | 0.0063134 | 0.0076667 | 0.01 | 0.31833 | 0.011667 | 0.01 | 0.03 | 0.01 | 0.014062 | 0.0079833 |  |  | 0.012033 |  | 0.0068333 | 0.015 | 0.015 | 0.010617
| Median |  | 0.0071 | 0.0077301 | 0.003645 | 0.0062334 | 0.008 | 0.01 | 0.175 | 0.01 | 0.01 | 0.03 | 0.01 | 0.013866 | 0.00795 |  |  | 0.01165 |  | 0.0068 | 0.015 | 0.015 | 0.00995
| Std Dev. |  | 0.0001633 | 0.00016071 | 0.00069556 | 0.00052516 | 0.0005164 | 0 | 0.37269 | 0.0040825 | 0 | 0 | 0 | 0.001779 | 0.00020412 |  |  | 0.0013765 |  | 0.0010504 | 0.0054772 | 0 | 0.0040524
| Rel.Std.Dev. |  | 2.322% | 2.092% | 19.3% | 8.318% | 6.736% | 0% | 117.1% | 34.99% | 0% | 0% | 0% | 12.65% | 2.557% |  |  | 11.44% |  | 15.37% | 36.51% | 0% | 38.17%
| PDM3 |  | -27.68% | -21.02% | -62.95% | -35.09% | -21.17% | 2.819% | 3173% | 19.96% | 2.819% | 208.5% | 2.819% | 44.58% | -17.92% |  |  | 23.73% |  | -29.74% | 54.23% | 54.23% | 9.159%
| Z-Score (Absolute) |  | 0.86 | 0.67 | 1.85 | 1.07 | 0.67 | 0 | 89.11 | 0.48 | 0 | 5.78 | 0.24 | 1.17 | 0.58 | 1.44 | 0 | 0.59 | 0 | 1.54 | 1.44 | 1.44 | 0.18
| Analytical results for Sb in OREAS 251c (Certified Value 1.66 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 1.91 | 1.3399 | 1.3084 | 1.71 | 1.35 | 1.59 | 1.62 | 1.49 | 1.6 | 1.92 | 1.4866 | 1.93 | 2.12 | 1.61 | 1.6 | 1.55 | < 5 | 1.8 |  |  |  | 1
| 1 | 2 | 1.94 | 1.3944 | 1.3249 | 1.67 | 1.47 | 1.55 | 1.62 | 1.5 | 1.61 | 1.8 | 1.4535 | 1.93 | 2.12 | 1.69 | 1.57 | 1.63 | < 5 | 1.8 |  |  |  | 37
| 1 | 3 | 1.85 | 1.4 | 1.4597 | 1.81 | 1.29 | 1.53 | 1.69 | 1.49 | 1.62 | 1.77 | 1.5379 | 1.89 | 2.14 | 1.78 | 1.62 | 1.65 | < 5 | 1.9 |  |  |  | 16
| 1 | 4 | 1.91 | 1.3247 | 1.5055 | 1.71 | 1.28 | 1.53 | 1.65 | 1.47 | 1.6 | 1.83 | 1.4693 | 1.85 | 2.1 | 1.76 | 1.6 | 1.6 | < 5 | 1.9 | 0.397 |  |  | 1.6559
| 1 | 5 | 1.97 | 1.3401 | 1.2701 | 1.9 | 1.5 | 1.55 | 1.67 | 1.48 | 1.61 | 1.86 | 1.5195 | 1.67 | 2.13 | 1.49 | 1.58 | 1.6 | < 5 | 2 |  |  |  | 107
| 1 | 6 | 1.9 | 1.3428 | 1.4328 | 1.69 | 1.41 | 1.51 | 1.64 | 1.45 | 1.55 | 1.88 | 1.5317 | 1.9 | 2.12 | 1.65 | 1.61 | 1.61 | < 5 | 1.9
| Mean |  | 1.9133 | 1.357 | 1.3836 | 1.7483 | 1.3833 | 1.5433 | 1.6483 | 1.48 | 1.5983 | 1.8433 | 1.4997 | 1.8617 | 2.1217 | 1.6633 | 1.5967 | 1.6067 |  | 1.8833 | 0.397
| Median |  | 1.91 | 1.3415 | 1.3788 | 1.71 | 1.38 | 1.54 | 1.645 | 1.485 | 1.605 | 1.845 | 1.5031 | 1.895 | 2.12 | 1.67 | 1.6 | 1.605 |  | 1.9 | 0.397
| Std Dev. |  | 0.040332 | 0.031831 | 0.094922 | 0.088638 | 0.092014 | 0.027325 | 0.027869 | 0.017889 | 0.024833 | 0.05465 | 0.034923 | 0.098472 | 0.013292 | 0.10652 | 0.018619 | 0.033862 |  | 0.075277
| Rel.Std.Dev. |  | 2.108% | 2.346% | 6.861% | 5.07% | 6.652% | 1.771% | 1.691% | 1.209% | 1.554% | 2.965% | 2.329% | 5.289% | 0.6265% | 6.404% | 1.166% | 2.108% |  | 3.997%
| PDM3 |  | 15.55% | -18.05% | -16.44% | 5.584% | -16.46% | -6.796% | -0.4547% | -10.62% | -3.474% | 11.32% | -9.428% | 12.43% | 28.13% | 0.4512% | -3.575% | -2.971% |  | 13.74% | -76.02%
| Z-Score (Absolute) |  | 0.93 | 0.75 | 0.67 | 0.43 | 0.68 | 0.19 | 0.13 | 0.38 | 0.03 | 0.72 | 0.32 | 0.77 | 1.56 | 0.17 | 0.03 | 0 | 2.7 | 0.84 | 3.66
| Analytical results for Sc in OREAS 251c (Certified Value 3.01 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  |  | 2
| 1 | 1 | 2.7 | 3.0374 | 3.9 | 3.1 | 2.7 | 3.3 | 3.1 | 3 | 3 | 3.2135 | 3 | 2.6 | 4 | 2.4 | 4 | 2.9 | 3 | 2.854 |  |  |  | 1
| 1 | 2 | 2.8 | 3.0984 | 3.8 | 3 | 2.8 | 3.3 | 3.2 | 3.1 | 2.8 | 3.1641 | 4 | 2.9 | 4 | 2.3 | 4 | 2.9 | 3 | 2.893 |  |  |  | 38
| 1 | 3 | 2.6 | 3.0441 | 3.7 | 3.1 | 2.8 | 3.3 | 3.2 | 3.1 | 3 | 3.27 | 4 | 2.8 | 4 | 2.5 | 3 | 2.9 | 3 | 2.966 |  |  |  | 16
| 1 | 4 | 2.8 | 2.9285 | 3.6 | 3.3 | 2.7 | 3.3 | 3.2 | 3.1 | 3 | 3.3758 | 3 | 3 | 4 | 2.5 | 4 | 3 | 3 | 2.598 |  |  |  | 3.011
| 1 | 5 | 2.8 | 3.0841 | 3.7 | 3.4 | 2.9 | 3.4 | 3.2 | 3 | 3 | 3.3742 | 4 | 2.4 | 4 | 2.4 | 3 | 3 | 3 | 2.836 |  |  |  | 108
| 1 | 6 | 2.8 | 3.0543 | 3.7 | 3.2 | 2.8 | 3.3 | 3.1 | 3 | 2.9 | 3.1776 | 3 | 2.7 | 4 | 2.5 | 4 | 2.8 | 3 | 2.722
| Mean |  | 2.75 | 3.0411 | 3.7333 | 3.1833 | 2.7833 | 3.3167 | 3.1667 | 3.05 | 2.95 | 3.2625 | 3.5 | 2.7333 | 4 | 2.4333 | 3.6667 | 2.9167 | 3 | 2.8115
| Median |  | 2.8 | 3.0492 | 3.7 | 3.15 | 2.8 | 3.3 | 3.2 | 3.05 | 3 | 3.2417 | 3.5 | 2.75 | 4 | 2.45 | 4 | 2.9 | 3 | 2.845
| Std Dev. |  | 0.083666 | 0.060019 | 0.10328 | 0.1472 | 0.075277 | 0.040825 | 0.05164 | 0.054772 | 0.083666 | 0.09451 | 0.54772 | 0.21602 | 0 | 0.08165 | 0.5164 | 0.075277 | 0 | 0.13154
| Rel.Std.Dev. |  | 3.042% | 1.974% | 2.766% | 4.624% | 2.705% | 1.231% | 1.631% | 1.796% | 2.836% | 2.897% | 15.65% | 7.903% | 0% | 3.355% | 14.08% | 2.581% | 0% | 4.679%
| PDM3 |  | -8.669% | 0.9991% | 23.99% | 5.723% | -7.562% | 10.15% | 5.169% | 1.294% | -2.027% | 8.353% | 16.24% | -9.222% | 32.85% | -19.19% | 21.77% | -3.134% | -0.3661% | -6.626%
| Z-Score (Absolute) |  | 0.83 | 0.15 | 2.49 | 0.63 | 0.72 | 1.08 | 0.58 | 0.18 | 0.15 | 0.9 | NA | 0.88 | NA | 1.9 | NA | 0.27 | NA | 0.62
Results from laboratories 1.16, 1.18, 1.21 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 251c (Indicative Value 0.29 ppm) | [22 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1.5 | 0.96674 | 0.2 | < 1 | 0.3 | 0.2 | 0.2 | 0.3 | 0.3 | < 1 | 0.16 | < 1 | 0.4 | 0.2 | 1.15 | 0.5 | 0.358 |  |  |  |  | 1
| 1 | 2 | < 1.5 | 0.91755 | 0.3 | < 1 | < 0.2 | < 0.2 | 0.2 | 0.2 | 0.4 | < 1 | 0.25 | < 1 | 0.4 | 0.2 | 0.85 | 0.5 | [5 empty] | 39
| 1 | 3 | < 1.5 | 0.86641 | 0.2 | < 1 | 0.2 | 0.3 | 0.2 | < 0.2 | 0.4 | < 1 | 0.16 | < 1 | 0.7 | 0.2 | 0.9 | 0.5 | [5 empty] | 16
| 1 | 4 | < 1.5 | 0.96242 | 0.3 | < 1 | 0.2 | 0.2 | 0.2 | < 0.2 | 0.3 | < 1 | 0.16 | < 1 | 0.3 | 0.3 | 0.99 | 0.5 | 0.553 |  |  |  |  | 0.28687
| 1 | 5 | < 1.5 | 0.91161 | 0.2 | < 1 | 0.2 | 0.3 | 0.2 | 0.4 | 0.3 | < 1 | 0.13 | < 1 | 0.5 | 0.2 | 0.95 | 0.5 | 0.817 |  |  |  |  | 10
| 1 | 6 | < 1.5 | 0.80237 | 0.2 | < 1 | < 0.2 | 0.2 | 0.2 | 0.2 | 0.4 | < 1 | 0.15 | < 1 | 0.8 | 0.3 | 1.19 | 0.5
| Mean |  |  | 0.90452 | 0.23333 |  | 0.225 | 0.24 | 0.2 | 0.275 | 0.35 |  | 0.16833 |  | 0.51667 | 0.23333 | 1.005 | 0.5 | 0.576
| Median |  |  | 0.91458 | 0.2 |  | 0.2 | 0.2 | 0.2 | 0.25 | 0.35 |  | 0.16 |  | 0.45 | 0.2 | 0.97 | 0.5 | 0.553
| Std Dev. |  |  | 0.062177 | 0.05164 |  | 0.05 | 0.054772 | 3.0405e-17 | 0.095743 | 0.054772 |  | 0.041673 |  | 0.19408 | 0.05164 | 0.13678 | 0 | 0.23036
| Rel.Std.Dev. |  |  | 6.874% | 22.13% |  | 22.22% | 22.82% | 1.52e-14% | 34.82% | 15.65% |  | 24.76% |  | 37.56% | 22.13% | 13.61% | 0% | 39.99%
| PDM3 |  |  | 215.3% | -18.66% |  | -21.57% | -16.34% | -30.28% | -4.137% | 22.01% |  | -41.32% |  | 80.11% | -18.66% | 250.3% | 74.3% | 100.8%
| Z-Score (Absolute) |  | 0.63 | 1.02 | 0.67 | 0 | 0.8 | 0.72 | 0.76 | 0.72 | 0.38 | 0 | 0.84 | 0 | 0.04 | 0.67 | 1.28 | 0 | 0.19
| Analytical results for Sm in OREAS 251c (Indicative Value 4.55 ppm) | [22 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 4.333 | 4.4837 | 7.5246 | 5.5 | 3.9 | [16 empty] | 1
| 1 | 2 | 4.409 | 4.4599 | 7.6048 | 5.8 | 3.91 | [16 empty] | 5
| 1 | 3 | 4.136 | 4.6218 | 7.5446 | 5.9 | 4.1 | [16 empty] | 16
| 1 | 4 | 4.231 | 4.1295 | 7.7312 | 5.3 | 4.19 | [16 empty] | 4.5515
| 1 | 5 | 4.498 | 4.2469 | 7.6478 | 5.2 | 4.06 | [16 empty] | 11
| 1 | 6 | 4.236 | 4.582 | 7.6083 | 4.9 | 4.11
| Mean |  | 4.3072 | 4.4206 | 7.6102 | 5.4333 | 4.045
| Median |  | 4.2845 | 4.4718 | 7.6066 | 5.4 | 4.08
| Std Dev. |  | 0.13238 | 0.19336 | 0.074436 | 0.37771 | 0.1164
| Rel.Std.Dev. |  | 3.073% | 4.374% | 0.9781% | 6.952% | 2.878%
| PDM3 |  | -5.369% | -2.876% | 67.2% | 19.37% | -11.13%
| Z-Score (Absolute) |  | 0.2 | 0 | 5.73 | 1.82 | 0.67
| Analytical results for Sn in OREAS 251c (Certified Value 1.46 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 3
| 1 | 1 | 1.42 | 1.5739 | 1.369 | 1.98 | 1.6 | 1.4 | 1.3 | 1.3 | 1.4 | 1.5 | 1.567 | 2.2 | 2.3 | < 5 | 1 | [6 empty] | 1
| 1 | 2 | 1.45 | 1.5341 | 1.2874 | 1.99 | 1.5 | 1.5 | 1.4 | 1.3 | 1.5 | 1.4 | 1.6237 | 2.2 | 2.8 | < 5 | 1 | [6 empty] | 12
| 1 | 3 | 1.4 | 1.5268 | 1.3387 | 1.97 | 1.5 | 1.6 | 1.4 | 1.3 | 1.4 | 1.4 | 1.6897 | 2.2 | 2.2 | < 5 | 2 | [6 empty] | 16
| 1 | 4 | 1.51 | 1.4923 | 1.2479 | 1.9 | 1.5 | 1.5 | 1.4 | 1.3 | 1.5 | 1.5 | 1.5765 | 2.2 | 2.7 | < 5 | 2 | [6 empty] | 1.4556
| 1 | 5 | 1.4 | 1.4573 | 1.2487 | 1.93 | 1.5 | 1.5 | 1.4 | 1.3 | 1.5 | 1.5 | 1.7085 | 2.2 | 2.4 | < 5 | 2 | [6 empty] | 109
| 1 | 6 | 1.51 | 1.6069 | 1.3685 | 1.89 | 1.5 | 1.5 | 1.4 | 1.3 | 1.5 | 1.8 | 1.669 | 2.3 | 2.8 | < 5 | 1
| Mean |  | 1.4483 | 1.5319 | 1.31 | 1.9433 | 1.5167 | 1.5 | 1.3833 | 1.3 | 1.4667 | 1.5167 | 1.6391 | 2.2167 | 2.5333 |  | 1.5
| Median |  | 1.435 | 1.5305 | 1.3131 | 1.95 | 1.5 | 1.5 | 1.4 | 1.3 | 1.5 | 1.5 | 1.6464 | 2.2 | 2.55 |  | 1.5
| Std Dev. |  | 0.051153 | 0.053952 | 0.056288 | 0.04274 | 0.040825 | 0.063246 | 0.040825 | 0 | 0.05164 | 0.1472 | 0.059358 | 0.040825 | 0.26583 |  | 0.54772
| Rel.Std.Dev. |  | 3.532% | 3.522% | 4.297% | 2.199% | 2.692% | 4.216% | 2.951% | 0% | 3.521% | 9.705% | 3.621% | 1.842% | 10.49% |  | 36.51%
| PDM3 |  | -0.4994% | 5.241% | -9.999% | 33.51% | 4.195% | 3.05% | -4.965% | -10.69% | 0.7601% | 4.195% | 12.61% | 52.29% | 74.04% |  | 3.05%
| Z-Score (Absolute) |  | 0.36 | 0.08 | 1.09 | 2.25 | 0 | 0.09 | 0.7 | 1.14 | 0.26 | 0 | 0.65 | 3.69 | 5.36 | 5.19 | NA
Results from laboratory 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 251c (Certified Value 39.2 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 1
| 1 | 1 | 39.32 | 40.7791 | 34.383 | 61.4 | 44.3 | 41 | 45.5 | 38.1 | 36.4 | 39.9 | 40.918 | 40.5 | 37.9 | 46 | 33 | 32.1 | 40.6 | 36 | 33.293 |  |  | 1
| 1 | 2 | 41.94 | 41.4042 | 33.994 | 61.2 | 46.1 | 42.2 | 45.8 | 39.6 | 37.7 | 38.3 | 40.5955 | 41.7 | 40.9 | 46 | 30.8 | 34.5 | 39.2 | 36 | 34.085 |  |  | 13
| 1 | 3 | 39.06 | 39.932 | 33.948 | 59.7 | 42.3 | 42.4 | 46 | 38.2 | 36.5 | 39.2 | 41.1538 | 41.2 | 41 | 46 | 34.1 | 30 | 39.7 | 38 | 33.717 |  |  | 16
| 1 | 4 | 40.61 | 39.1896 | 33.122 | 59.2 | 45.1 | 40.7 | 46.3 | 39.9 | 36.7 | 39.7 | 42.5144 | 40.5 | 40 | 45 | 34.8 | 36 | 39.2 | 40 | 29.84 |  |  | 39.2312
| 1 | 5 | 41.86 | 40.5602 | 34.538 | 58.5 | 44.7 | 41.5 | 46.3 | 38.9 | 35 | 39.3 | 42.5325 | 41.6 | 36.7 | 45 | 33.5 | 23.3 | 40.9 | 38 | 32.039 |  |  | 110
| 1 | 6 | 41.55 | 41.9648 | 32.828 | 59.2 | 43.1 | 40.8 | 45.5 | 38.6 | 35.2 | 39.9 | 41.9056 | 40.4 | 37.4 | 46 | 34.1 | 34.8 | 39.9 | 39 | 32.11
| Mean |  | 40.7233 | 40.6383 | 33.8022 | 59.8667 | 44.2667 | 41.4333 | 45.9 | 38.8833 | 36.25 | 39.3833 | 41.6033 | 40.9833 | 38.9833 | 45.6667 | 33.3833 | 31.7833 | 39.9167 | 37.8333 | 32.514
| Median |  | 41.08 | 40.6697 | 33.971 | 59.45 | 44.5 | 41.25 | 45.9 | 38.75 | 36.45 | 39.5 | 41.5297 | 40.85 | 38.95 | 46 | 33.8 | 33.3 | 39.8 | 38 | 32.7015
| Std Dev. |  | 1.281 | 0.99697 | 0.68535 | 1.1759 | 1.3765 | 0.72847 | 0.36332 | 0.73598 | 1.0055 | 0.608 | 0.83334 | 0.59133 | 1.8798 | 0.5164 | 1.4049 | 4.6799 | 0.70828 | 1.6021 | 1.5531
| Rel.Std.Dev. |  | 3.146% | 2.453% | 2.028% | 1.964% | 3.109% | 1.758% | 0.7915% | 1.893% | 2.774% | 1.544% | 2.003% | 1.443% | 4.822% | 1.131% | 4.208% | 14.72% | 1.774% | 4.235% | 4.777%
| PDM3 |  | 3.804% | 3.587% | -13.84% | 52.6% | 12.84% | 5.613% | 17% | -0.8866% | -7.599% | 0.3879% | 6.047% | 4.466% | -0.6317% | 16.4% | -14.91% | -18.98% | 1.747% | -3.563% | -17.12%
| Z-Score (Absolute) |  | 0.26 | 0.23 | 1.98 | 6.46 | 1.41 | 0.49 | 1.94 | 0.33 | 1.19 | 0.17 | 0.55 | 0.35 | 0.3 | 1.86 | 2.11 | 2.63 | 0 | 0.67 | 2.4
| Analytical results for Ta in OREAS 251c (Certified Value < 0.01 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | 0.03 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 26
| 1 | 3 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | [10 empty] | 111
| 1 | 6 | < 0.01 | < 0.05 | 0.14 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01
| Mean |  |  |  | 0.14 | [7 empty] | 0.03
| Median |  |  |  | 0.14 | [7 empty] | 0.03
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tb in OREAS 251c (Certified Value 0.49 ppm) | [22 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.28 | [16 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [16 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | 40g | [16 empty] | 2
| 1 | 1 | 0.387 | 0.44975 | 0.664 | 0.5 | 0.445 | [16 empty] | 1
| 1 | 2 | 0.387 | 0.45576 | 0.64128 | 0.5 | 0.445 | [16 empty] | 27
| 1 | 3 | 0.369 | 0.46395 | 0.68032 | 0.5 | 0.46 | [16 empty] | 16
| 1 | 4 | 0.374 | 0.43672 | 0.63636 | 0.5 | 0.46 | [16 empty] | 0.4887
| 1 | 5 | 0.388 | 0.44803 | 0.69172 | 0.5 | 0.455 | [16 empty] | 112
| 1 | 6 | 0.389 | 0.46989 | 0.58828 | 0.5 | 0.46
| Mean |  | 0.38233 | 0.45402 | 0.65033 | 0.5 | 0.45417
| Median |  | 0.387 | 0.45276 | 0.65264 | 0.5 | 0.4575
| Std Dev. |  | 0.0085713 | 0.011889 | 0.03722 | 0 | 0.0073598
| Rel.Std.Dev. |  | 2.242% | 2.619% | 5.723% | 0% | 1.621%
| PDM3 |  | -21.77% | -7.097% | 33.07% | 2.312% | -7.067%
| Z-Score (Absolute) |  | 1.06 | 0 | 2.89 | 0.67 | 0
| Analytical results for Te in OREAS 251c (Certified Value 0.083 ppm) | [22 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g | [5 empty] | 3
| 1 | 1 |  | < 0.2 | 0.13 | 0.39 | 0.07 | 0.08 | 0.08 | 0.07 | 0.09 | 0.08025 | 0.06 | 0.09 | 0.08 | 0.08 | 0.1 | < 0.2 | [5 empty] | 1
| 1 | 2 |  | < 0.2 | 0.08 | 0.38 | 0.09 | 0.07 | 0.09 | 0.05 | 0.09 | 0.086435 | 0.06 | 0.09 | 0.06 | 0.09 | 0.09 | < 0.2 | [5 empty] | 28
| 1 | 3 |  | < 0.2 | 0.1 | 0.34 | 0.09 | 0.08 | 0.08 | 0.07 | 0.09 | 0.11025 | 0.05 | 0.09 | 0.07 | 0.08 | 0.09 | < 0.2 | [5 empty] | 16
| 1 | 4 | 0.1 | < 0.2 | 0.08 | 0.35 | 0.08 | 0.07 | 0.09 | 0.06 | 0.07 | 0.094208 | 0.05 | 0.1 | 0.03 | 0.09 | 0.11 | < 0.2 | [5 empty] | 0.082578
| 1 | 5 | 0.1 | < 0.2 | 0.1 | 0.37 | 0.09 | 0.09 | 0.07 | 0.07 | 0.09 | 0.072169 | 0.07 | 0.08 | 0.05 | 0.09 | 0.09 | < 0.2 | [5 empty] | 113
| 1 | 6 |  | < 0.2 | 0.11 | 0.37 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.11028 | 0.06 | 0.08 | 0.1 | 0.09 | 0.11 | < 0.2
| Mean |  | 0.1 |  | 0.1 | 0.36667 | 0.085 | 0.078333 | 0.081667 | 0.065 | 0.085 | 0.092266 | 0.058333 | 0.088333 | 0.065 | 0.086667 | 0.098333
| Median |  | 0.1 |  | 0.1 | 0.37 | 0.09 | 0.08 | 0.08 | 0.07 | 0.09 | 0.090322 | 0.06 | 0.09 | 0.065 | 0.09 | 0.095
| Std Dev. |  | 0 |  | 0.018974 | 0.018619 | 0.0083666 | 0.0075277 | 0.0075277 | 0.0083666 | 0.0083666 | 0.01571 | 0.0075277 | 0.0075277 | 0.02429 | 0.005164 | 0.0098319
| Rel.Std.Dev. |  | 0% |  | 18.97% | 5.078% | 9.843% | 9.61% | 9.218% | 12.87% | 9.843% | 17.03% | 12.9% | 8.522% | 37.37% | 5.958% | 9.999%
| PDM3 |  | 21.1% |  | 21.1% | 344% | 2.933% | -5.14% | -1.103% | -21.29% | 2.933% | 11.73% | -29.36% | 6.97% | -21.29% | 4.952% | 19.08%
| Z-Score (Absolute) |  | NA | 0.77 | 0.77 | 16.18 | 0.1 | 0.48 | 0.29 | 1.25 | 0.1 | 0.32 | 1.64 | 0.1 | 1.25 | 0 | 0.67 | 0.77
Results from laboratory 1.02 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 251c (Certified Value 8.67 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 2
| 1 | 1 | 7.74 | 8.5078 | 9.534 | 12.7332 | 9.4 | 9 | 8 | 9.2 | 8.1 | 7.9 | 8.1 | 8.7886 | 9.23 | 7.6 | 7.4 | 8.09 | 10.8 | 9.7 |  |  |  | 1
| 1 | 2 | 7.9 | 8.6093 | 9.708 | 12.7088 | 9.5 | 9.5 | 8 | 9.1 | 8.1 | 7.8 | 8.4 | 8.7112 | 9.55 | 7.9 | 7.1 | 8.19 | 10.3 | 10 |  |  |  | 29
| 1 | 3 | 7.4 | 8.4912 | 9.76 | 12.786 | 9.4 | 8.8 | 8 | 9.3 | 7.9 | 7.8 | 8 | 8.7255 | 9.29 | 7.7 | 8 | 7.77 | 9.8 | 10.1 |  |  |  | 16
| 1 | 4 | 7.54 | 8.0465 | 9.145 | 12.786 | 9.3 | 9.2 | 7.9 | 9.2 | 8.1 | 7.9 | 8.5 | 8.8176 | 9.28 | 8.2 | 7.9 | 8.26 | 10.4 | 10.5 |  |  |  | 8.6724
| 1 | 5 | 7.83 | 8.4412 | 9.902 | 12.7424 | 9.6 | 9.3 | 8.5 | 9.4 | 7.9 | 7.6 | 8.5 | 8.8326 | 9.41 | 6.7 | 7.4 | 7.25 | 9.7 | 10.3 |  |  |  | 114
| 1 | 6 | 7.86 | 8.4678 | 9.823 | 12.7375 | 9.2 | 8.9 | 8.2 | 9.1 | 7.7 | 7.8 | 8.2 | 8.8095 | 9.28 | 7.5 | 7.5 | 8.17 | 10 | 10.2
| Mean |  | 7.7117 | 8.4273 | 9.6453 | 12.749 | 9.4 | 9.1167 | 8.1 | 9.2167 | 7.9667 | 7.8 | 8.2833 | 8.7808 | 9.34 | 7.6 | 7.55 | 7.955 | 10.1667 | 10.1333
| Median |  | 7.785 | 8.4795 | 9.734 | 12.7399 | 9.4 | 9.1 | 8 | 9.2 | 8 | 7.8 | 8.3 | 8.7991 | 9.285 | 7.65 | 7.45 | 8.13 | 10.15 | 10.15
| Std Dev. |  | 0.19944 | 0.19519 | 0.27468 | 0.030919 | 0.14142 | 0.26394 | 0.21909 | 0.1169 | 0.1633 | 0.10954 | 0.2137 | 0.050631 | 0.119 | 0.50596 | 0.33912 | 0.38573 | 0.41312 | 0.27325
| Rel.Std.Dev. |  | 2.586% | 2.316% | 2.848% | 0.2425% | 1.504% | 2.895% | 2.705% | 1.268% | 2.05% | 1.404% | 2.58% | 0.5766% | 1.274% | 6.657% | 4.492% | 4.849% | 4.063% | 2.697%
| PDM3 |  | -11.08% | -2.826% | 11.22% | 47.01% | 8.39% | 5.123% | -6.6% | 6.276% | -8.138% | -10.06% | -4.486% | 1.25% | 7.698% | -12.37% | -12.94% | -8.272% | 17.23% | 16.85%
| Z-Score (Absolute) |  | 0.79 | 0.16 | 0.92 | 3.65 | 0.7 | 0.45 | 0.44 | 0.54 | 0.56 | 0.71 | 0.28 | 0.16 | 0.65 | 0.88 | 0.93 | 0.57 | 1.38 | 1.35
| Analytical results for Ti in OREAS 251c (Certified Value 0.112 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 3
| 1 | 1 | 0.0925 | 0.10829 | 0.071108 | 0.15 | 0.105 | 0.112 | 0.131 | 0.096 | 0.107 | 0.104 | 0.14874 | 0.144 | 0.094 | 0.1246 | 0.064 | 0.14 | 0.17 | 0.07 | 0.12123 |  |  | 1
| 1 | 2 | 0.0982 | 0.10765 | 0.067003 | 0.15 | 0.108 | 0.117 | 0.132 | 0.103 | 0.107 | 0.102 | 0.15055 | 0.153 | 0.102 | 0.1265 | 0.061 | 0.151 | 0.16 | 0.08 | 0.12452 |  |  | 30
| 1 | 3 | 0.0881 | 0.10736 | 0.067525 | 0.15 | 0.095 | 0.116 | 0.129 | 0.097 | 0.106 | 0.104 | 0.14911 | 0.152 | 0.101 | 0.1245 | 0.074 | 0.13 | 0.17 | 0.08 | 0.12509 |  |  | 16
| 1 | 4 | 0.0957 | 0.10385 | 0.064307 | 0.14 | 0.103 | 0.112 | 0.131 | 0.103 | 0.106 | 0.103 | 0.15289 | 0.145 | 0.1 | 0.1181 | 0.071 | 0.159 | 0.17 | 0.075 | 0.094982 |  |  | 0.1116
| 1 | 5 | 0.0949 | 0.10957 | 0.070913 | 0.15 | 0.106 | 0.113 | 0.131 | 0.102 | 0.101 | 0.102 | 0.15512 | 0.154 | 0.091 | 0.1218 | 0.065 | 0.154 | 0.16 | 0.08 | 0.11819 |  |  | 115
| 1 | 6 | 0.0999 | 0.11009 | 0.065073 | 0.15 | 0.1 | 0.113 | 0.133 | 0.101 | 0.103 | 0.105 | 0.1563 | 0.15 | 0.092 | 0.1282 | 0.07 | 0.153 | 0.16 | 0.07 | 0.11168
| Mean |  | 0.094883 | 0.1078 | 0.067655 | 0.14833 | 0.10283 | 0.11383 | 0.13117 | 0.10033 | 0.105 | 0.10333 | 0.15212 | 0.14967 | 0.096667 | 0.12395 | 0.0675 | 0.14783 | 0.165 | 0.075833 | 0.11595
| Median |  | 0.0953 | 0.10797 | 0.067264 | 0.15 | 0.104 | 0.113 | 0.131 | 0.1015 | 0.106 | 0.1035 | 0.15172 | 0.151 | 0.097 | 0.12455 | 0.0675 | 0.152 | 0.165 | 0.0775 | 0.11971
| Std Dev. |  | 0.0042078 | 0.0022107 | 0.0028586 | 0.0040825 | 0.0047081 | 0.002137 | 0.0013292 | 0.0030768 | 0.0024495 | 0.0012111 | 0.003161 | 0.0042269 | 0.0048854 | 0.0035815 | 0.0049295 | 0.010759 | 0.0054772 | 0.004916 | 0.011383
| Rel.Std.Dev. |  | 4.435% | 2.051% | 4.225% | 2.752% | 4.578% | 1.877% | 1.013% | 3.067% | 2.333% | 1.172% | 2.078% | 2.824% | 5.054% | 2.889% | 7.303% | 7.278% | 3.32% | 6.483% | 9.817%
| PDM3 |  | -14.98% | -3.41% | -39.38% | 32.91% | -7.859% | 1.997% | 17.53% | -10.1% | -5.918% | -7.411% | 36.3% | 34.1% | -13.38% | 11.06% | -39.52% | 32.46% | 47.84% | -32.05% | 3.892%
| Z-Score (Absolute) |  | 0.54 | 0 | 1.68 | 1.69 | 0.21 | 0.25 | 0.98 | 0.31 | 0.12 | 0.19 | 1.85 | 1.75 | 0.46 | 0.67 | 1.68 | 1.67 | 2.39 | 1.33 | 0.34
| Analytical results for Tl in OREAS 251c (Certified Value 0.088 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.23 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 3
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 01g | 40g |  |  |  | 3
| 1 | 1 | 0.1 | 0.092051 | 0.07716 | 0.12 | 0.07 | 0.08 | 0.1 | 0.08 | 0.08 | 0.08 | 0.098368 | 0.084 | 0.13 | 0.08 | 0.09 | 0.17 | < 5 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.1 | 0.10108 | 0.064954 | 0.11 | 0.08 | 0.09 | 0.1 | 0.08 | 0.08 | 0.09 | 0.093348 | 0.083 | 0.13 | 0.08 | 0.08 | 0.21 | < 5 | < 0.1 |  |  |  | 31
| 1 | 3 | 0.08 | 0.098697 | 0.066422 | 0.11 | 0.06 | 0.09 | 0.1 | 0.08 | 0.08 | 0.09 | 0.10525 | 0.079 | 0.14 | 0.08 | 0.09 | 0.16 | < 5 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.1 | 0.091091 | 0.071211 | 0.12 | 0.06 | 0.09 | 0.1 | 0.08 | 0.08 | 0.09 | 0.11753 | 0.081 | 0.13 | 0.09 | 0.09 | 0.2 | < 5 | < 0.1 |  |  |  | 0.088317
| 1 | 5 | 0.09 | 0.092684 | 0.076341 | 0.11 | 0.07 | 0.09 | 0.11 | 0.08 | 0.07 | 0.09 | 0.10503 | 0.07 | 0.13 | 0.08 | 0.09 | 0.18 | < 5 | < 0.1 |  |  |  | 116
| 1 | 6 | 0.1 | 0.095702 | 0.091175 | 0.11 | 0.06 | 0.08 | 0.1 | 0.08 | 0.07 | 0.09 | 0.12216 | 0.08 | 0.14 | 0.08 | 0.09 | 0.21 | < 5 | < 0.1
| Mean |  | 0.095 | 0.095218 | 0.074544 | 0.11333 | 0.066667 | 0.086667 | 0.10167 | 0.08 | 0.076667 | 0.088333 | 0.10695 | 0.0795 | 0.13333 | 0.081667 | 0.088333 | 0.18833
| Median |  | 0.1 | 0.094193 | 0.073776 | 0.11 | 0.065 | 0.09 | 0.1 | 0.08 | 0.08 | 0.09 | 0.10514 | 0.0805 | 0.13 | 0.08 | 0.09 | 0.19
| Std Dev. |  | 0.0083666 | 0.0040052 | 0.0095468 | 0.005164 | 0.008165 | 0.005164 | 0.0040825 | 0 | 0.005164 | 0.0040825 | 0.011031 | 0.00501 | 0.005164 | 0.0040825 | 0.0040825 | 0.02137
| Rel.Std.Dev. |  | 8.807% | 4.206% | 12.81% | 4.556% | 12.25% | 5.958% | 4.016% | 0% | 6.736% | 4.622% | 10.31% | 6.302% | 3.873% | 4.999% | 4.622% | 11.35%
| PDM3 |  | 7.567% | 7.813% | -15.6% | 28.33% | -24.51% | -1.869% | 15.12% | -9.418% | -13.19% | 0.01803% | 21.1% | -9.984% | 50.97% | -7.531% | 0.01803% | 113.2%
| Z-Score (Absolute) |  | 0.36 | 0.37 | 0.74 | 1.35 | 1.17 | 0.09 | 0.72 | 0.45 | 0.63 | 0 | 1 | 0.48 | 2.43 | 0.36 | 0 | 5.39 | 130.1 | 2.07
| Analytical results for Tm in OREAS 251c (Indicative Value 0.1 ppm) | [22 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.28 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 40g | [17 empty] | 2
| 1 | 1 | 0.1 | 0.099107 | 0.1 | 0.1 | [17 empty] | 1
| 1 | 2 | 0.112 | 0.10178 | 0.1 | 0.095 | [17 empty] | 6
| 1 | 3 | 0.096 | 0.1045 | 0.1 | 0.1 | [17 empty] | 16
| 1 | 4 | 0.099 | 0.097882 | 0.1 | 0.105 | [17 empty] | 0.10095
| 1 | 5 | 0.108 | 0.096039 | 0.1 | 0.1 | [17 empty] | 12
| 1 | 6 | 0.098 | 0.11061 | 0.1 | 0.1
| Mean |  | 0.10217 | 0.10165 | 0.1 | 0.1
| Median |  | 0.0995 | 0.10045 | 0.1 | 0.1
| Std Dev. |  | 0.0063377 | 0.0053021 | 1.5202e-17 | 0.0031623
| Rel.Std.Dev. |  | 6.203% | 5.216% | 1.52e-14% | 3.162%
| PDM3 |  | 1.2% | 0.6917% | -0.9459% | -0.9459%
| Z-Score (Absolute) |  | 1.09 | 0.67 | 0.67 | 0.67
| Analytical results for U in OREAS 251c (Certified Value 1.05 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.28 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 15g | 0.1g | 40g |  |  |  | 3
| 1 | 1 | 0.99 | 0.96797 | 1.1484 | 1.695 | 1 | 1.1 | 0.99 | 1.07 | 0.95 | 1.04 | 0.97 | 1.1208 | 1.17 | 1.15 | 0.97 | 1 | 1.01 | 1.1 |  |  |  | 1
| 1 | 2 | 0.99 | 1.0533 | 1.1152 | 1.8746 | 1.1 | 1.12 | 1.02 | 1.07 | 0.97 | 1.02 | 1.01 | 1.1204 | 1.23 | 1.12 | 0.99 | 0.9 | 1.05 | 1.1 |  |  |  | 33
| 1 | 3 | 0.96 | 1.0493 | 1.2094 | 1.8017 | 1.1 | 1.09 | 1.03 | 1.08 | 0.97 | 1.06 | 0.98 | 1.1478 | 1.13 | 1.15 | 1 | 1 | 1.02 | 1.15 |  |  |  | 16
| 1 | 4 | 0.97 | 0.98072 | 1.1254 | 1.8964 | 1 | 1.08 | 1.03 | 1.05 | 0.97 | 1.08 | 1.07 | 1.1309 | 1.18 | 1.13 | 1.04 | 1 | 1.04 | 1.1 |  |  |  | 1.0538
| 1 | 5 | 1 | 1.038 | 1.1737 | 1.8039 | 1.1 | 1.12 | 1.01 | 1.09 | 0.97 | 0.99 | 1.02 | 1.1565 | 1.02 | 1.14 | 0.85 | 1 | 0.94 | 1.1 |  |  |  | 117
| 1 | 6 | 0.99 | 1.116 | 1.1802 | 1.7639 | 1 | 1.08 | 1 | 1.03 | 0.94 | 1.06 | 1 | 1.163 | 1.13 | 1.13 | 0.9 | 1 | 1.05 | 1.1
| Mean |  | 0.98333 | 1.0342 | 1.1587 | 1.8059 | 1.05 | 1.0983 | 1.0133 | 1.065 | 0.96167 | 1.0417 | 1.0083 | 1.1399 | 1.1433 | 1.1367 | 0.95833 | 0.98333 | 1.0183 | 1.1083
| Median |  | 0.99 | 1.0436 | 1.1611 | 1.8028 | 1.05 | 1.095 | 1.015 | 1.07 | 0.97 | 1.05 | 1.005 | 1.1393 | 1.15 | 1.135 | 0.98 | 1 | 1.03 | 1.1
| Std Dev. |  | 0.015055 | 0.053924 | 0.035698 | 0.07346 | 0.054772 | 0.018348 | 0.01633 | 0.021679 | 0.013292 | 0.032506 | 0.035449 | 0.018408 | 0.070899 | 0.012111 | 0.070261 | 0.040825 | 0.041673 | 0.020412
| Rel.Std.Dev. |  | 1.531% | 5.214% | 3.081% | 4.068% | 5.216% | 1.671% | 1.612% | 2.036% | 1.382% | 3.121% | 3.516% | 1.615% | 6.201% | 1.065% | 7.332% | 4.152% | 4.092% | 1.842%
| PDM3 |  | -6.686% | -1.859% | 9.957% | 71.37% | -0.3599% | 4.227% | -3.839% | 1.064% | -8.742% | -1.151% | -4.314% | 8.17% | 8.497% | 7.864% | -9.059% | -6.686% | -3.365% | 5.176%
| Z-Score (Absolute) |  | 0.67 | 0.13 | 1.22 | 8.2 | 0.04 | 0.57 | 0.35 | 0.21 | 0.91 | 0.04 | 0.4 | 1.01 | 1.05 | 0.98 | 0.94 | 0.67 | 0.3 | 0.67
| Analytical results for V in OREAS 251c (Certified Value 29.4 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g |  |  | 2
| 1 | 1 | 27 | 28.0176 | 26.031 | 34 | 31 | 28 | 30 | 28 | 30 | 27 | 29.1069 | 33 | 29 | 32 | 24 | 34 | 31 | 30 | 28.328 |  |  | 1
| 1 | 2 | 28 | 28.5306 | 25.819 | 33 | 31 | 29 | 30 | 28 | 30 | 27 | 29.2938 | 34 | 30 | 31 | 23 | 37 | 31.1 | 30 | 28.877 |  |  | 34
| 1 | 3 | 26 | 28.1602 | 26.199 | 31 | 31 | 28 | 30 | 28 | 29 | 27 | 29.9598 | 34 | 29 | 31 | 25 | 31 | 30.6 | 30 | 29.142 |  |  | 16
| 1 | 4 | 27 | 27.4917 | 24.139 | 32 | 31 | 28 | 30 | 28 | 30 | 27 | 30.3324 | 33 | 29 | 32 | 25 | 36 | 32.6 | 30 | 26.513 |  |  | 29.3866
| 1 | 5 | 27 | 28.4102 | 26.71 | 32 | 31 | 28 | 30 | 28 | 29 | 27 | 30.5378 | 34 | 28 | 32 | 24 | 27 | 32.5 | 30 | 28.287 |  |  | 118
| 1 | 6 | 27 | 28.5695 | 25.855 | 31 | 31 | 28 | 31 | 28 | 29 | 27 | 30.5499 | 34 | 28 | 32 | 24 | 34 | 30.5 | 30 | 27.928
| Mean |  | 27 | 28.1966 | 25.7922 | 32.1667 | 31 | 28.1667 | 30.1667 | 28 | 29.5 | 27 | 29.9634 | 33.6667 | 28.8333 | 31.6667 | 24.1667 | 33.1667 | 31.3833 | 30 | 28.1792
| Median |  | 27 | 28.2852 | 25.943 | 32 | 31 | 28 | 30 | 28 | 29.5 | 27 | 30.1461 | 34 | 29 | 32 | 24 | 34 | 31.05 | 30 | 28.3075
| Std Dev. |  | 0.63246 | 0.40657 | 0.87208 | 1.169 | 0 | 0.40825 | 0.40825 | 0 | 0.54772 | 0 | 0.63119 | 0.5164 | 0.75277 | 0.5164 | 0.75277 | 3.656 | 0.93256 | 0 | 0.92596
| Rel.Std.Dev. |  | 2.342% | 1.442% | 3.381% | 3.634% | 0% | 1.449% | 1.353% | 0% | 1.857% | 0% | 2.107% | 1.534% | 2.611% | 1.631% | 3.115% | 11.02% | 2.972% | 0% | 3.286%
| PDM3 |  | -8.121% | -4.049% | -12.23% | 9.46% | 5.49% | -4.151% | 2.655% | -4.718% | 0.3859% | -8.121% | 1.963% | 14.56% | -1.883% | 7.759% | -17.76% | 12.86% | 6.795% | 2.087% | -4.109%
| Z-Score (Absolute) |  | 1.12 | 0.59 | 1.67 | 1.2 | 0.67 | 0.6 | 0.3 | 0.67 | 0 | 1.12 | 0.21 | 1.87 | 0.3 | 0.97 | 2.4 | 1.65 | 0.85 | 0.22 | 0.59
| Analytical results for W in OREAS 251c (Certified Value 9.78 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 01g | 15g | 0.1g | 01g |  |  |  |  | 2
| 1 | 1 | 8.43 | 11.8588 | 11.515 | 8.1 | 9.43 | 10.95 | 10.05 | 9.29 | 9.75 | 11.1 | 11.4393 | 10.6 | 8.37 | 7 | 7.2 | 12.9 | < 5 |  |  |  |  | 1
| 1 | 2 | 8.37 | 11.0983 | 12.135 | 7.3 | 9.76 | 10.65 | 10.5 | 9.42 | 9.15 | 10.3 | 11.1701 | 10.9 | 8.35 | 8 | 6.9 | 12.6 | < 5 |  |  |  |  | 35
| 1 | 3 | 8.1 | 11.8647 | 9.897 | 8.3 | 9.41 | 10.75 | 10.5 | 9.49 | 9.6 | 10.3 | 11.3263 | 10.5 | 8.22 | 8 | 7.9 | 13.5 | < 5 |  |  |  |  | 16
| 1 | 4 | 8.53 | 11.1353 | 12.411 | 7.6 | 9.56 | 10.6 | 10.4 | 9.22 | 9.34 | 10.25 | 11.0163 | 10.1 | 8.67 | 8 | 7.6 | 12.3 | < 5 |  |  |  |  | 9.7831
| 1 | 5 | 8.35 | 11.3416 | 13.818 | 9.8 | 9.56 | 10.45 | 11.25 | 9.27 | 8.59 | 10.95 | 11.0345 | 10.2 | 7.79 | 6 | 7.1 | 12 | < 5 |  |  |  |  | 119
| 1 | 6 | 7.82 | 11.1547 | 9.278 | 7.9 | 9.26 | 10.1 | 11.15 | 9.2 | 9.3 | 10.95 | 11.4177 | 10.5 | 8.19 | 7 | 7.3 | 13.6 | < 5
| Mean |  | 8.2667 | 11.4089 | 11.509 | 8.1667 | 9.4967 | 10.5833 | 10.6417 | 9.315 | 9.2883 | 10.6417 | 11.234 | 10.4667 | 8.265 | 7.3333 | 7.3333 | 12.8167
| Median |  | 8.36 | 11.2482 | 11.825 | 8 | 9.495 | 10.625 | 10.5 | 9.28 | 9.32 | 10.625 | 11.2482 | 10.5 | 8.285 | 7.5 | 7.25 | 12.75
| Std Dev. |  | 0.26113 | 0.36074 | 1.6801 | 0.8756 | 0.17049 | 0.28925 | 0.46413 | 0.11537 | 0.40494 | 0.39676 | 0.18746 | 0.28752 | 0.28829 | 0.8165 | 0.36148 | 0.64317
| Rel.Std.Dev. |  | 3.159% | 3.162% | 14.6% | 10.72% | 1.795% | 2.733% | 4.361% | 1.239% | 4.36% | 3.728% | 1.669% | 2.747% | 3.488% | 11.13% | 4.929% | 5.018%
| PDM3 |  | -15.5% | 16.62% | 17.64% | -16.52% | -2.928% | 8.18% | 8.776% | -4.785% | -5.057% | 8.776% | 14.83% | 6.987% | -15.52% | -25.04% | -25.04% | 31.01%
| Z-Score (Absolute) |  | 0.67 | 1.05 | 1.1 | 0.73 | 0 | 0.59 | 0.63 | 0.1 | 0.11 | 0.63 | 0.95 | 0.53 | 0.67 | 1.18 | 1.18 | 1.82 | 3.83
| Analytical results for Y in OREAS 251c (Certified Value 9.19 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 40g | 0.5g | [6 empty] | 3
| 1 | 1 | 8.45 | 9.0191 | 9.476 | 9.5 | 8.52 | 9.03 | 8.93 | 8.75 | 8.7 | 8.9 | 9.4369 | 10.8 | 10.3 | 10 | 9.443 | [6 empty] | 1
| 1 | 2 | 8.7 | 8.9607 | 9.52 | 9.59 | 8.95 | 9.46 | 9.17 | 8.66 | 9.08 | 8.89 | 9.2527 | 11.2 | 10.7 | 10.1 | 9.613 | [6 empty] | 36
| 1 | 3 | 8.25 | 9.0298 | 9.59 | 9.34 | 8.44 | 8.9 | 9.19 | 8.71 | 8.73 | 8.84 | 9.255 | 11.2 | 9.94 | 9.75 | 9.642 | [6 empty] | 16
| 1 | 4 | 8.51 | 8.7212 | 9.544 | 9.38 | 8.52 | 8.97 | 9.16 | 8.82 | 8.79 | 8.84 | 9.7302 | 10.8 | 10.5 | 10.2 | 8.851 | [6 empty] | 9.1879
| 1 | 5 | 8.84 | 8.9815 | 9.572 | 9.35 | 8.63 | 9.26 | 9.16 | 8.65 | 8.66 | 8.87 | 9.6438 | 11.2 | 8.67 | 9.6 | 9.362 | [6 empty] | 120
| 1 | 6 | 8.51 | 9.0831 | 9.332 | 9.28 | 8.23 | 8.98 | 9.36 | 8.69 | 8.74 | 9.03 | 9.3331 | 11 | 10.3 | 9.95 | 9.222
| Mean |  | 8.5433 | 8.9659 | 9.5057 | 9.4067 | 8.5483 | 9.1 | 9.1617 | 8.7133 | 8.7833 | 8.895 | 9.4419 | 11.0333 | 10.0683 | 9.9333 | 9.3555
| Median |  | 8.51 | 9.0003 | 9.532 | 9.365 | 8.52 | 9.005 | 9.165 | 8.7 | 8.735 | 8.88 | 9.385 | 11.1 | 10.3 | 9.975 | 9.4025
| Std Dev. |  | 0.20471 | 0.12711 | 0.09405 | 0.11553 | 0.23777 | 0.21513 | 0.13703 | 0.063456 | 0.15161 | 0.07064 | 0.20321 | 0.19664 | 0.73 | 0.22286 | 0.29272
| Rel.Std.Dev. |  | 2.396% | 1.418% | 0.9894% | 1.228% | 2.782% | 2.364% | 1.496% | 0.7283% | 1.726% | 0.7942% | 2.152% | 1.782% | 7.25% | 2.244% | 3.129%
| PDM3 |  | -7.015% | -2.416% | 3.459% | 2.381% | -6.961% | -0.9563% | -0.2851% | -5.165% | -4.403% | -3.187% | 2.765% | 20.09% | 9.583% | 8.114% | 1.825%
| Z-Score (Absolute) |  | 1.21 | 0.38 | 0.67 | 0.48 | 1.2 | 0.12 | 0 | 0.88 | 0.74 | 0.52 | 0.55 | 3.67 | 1.78 | 1.51 | 0.38
| Analytical results for Yb in OREAS 251c (Indicative Value 0.58 ppm) | [22 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | [17 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | [17 empty] | 2
|  |  | 01g | 0.2g | 0.5g | 0.5g | [17 empty] | 2
| 1 | 1 | 0.544 | 0.56462 | 0.89368 | 0.6 | [17 empty] | 1
| 1 | 2 | 0.572 | 0.57946 | 0.87828 | 0.6 | [17 empty] | 7
| 1 | 3 | 0.556 | 0.59113 | 0.88564 | 0.6 | [17 empty] | 16
| 1 | 4 | 0.568 | 0.54112 | 0.879 | 0.6 | [17 empty] | 0.58119
| 1 | 5 | 0.61 | 0.6012 | 0.8848 | 0.6 | [17 empty] | 13
| 1 | 6 | 0.55 | 0.58385 | 0.847 | 0.6
| Mean |  | 0.56667 | 0.57689 | 0.87807 | 0.6
| Median |  | 0.562 | 0.58165 | 0.8819 | 0.6
| Std Dev. |  | 0.023721 | 0.021348 | 0.016202 | 0
| Rel.Std.Dev. |  | 4.186% | 3.701% | 1.845% | 0%
| PDM3 |  | -2.498% | -0.7386% | 51.08% | 3.237%
| Z-Score (Absolute) |  | 0.88 | 0.47 | 11.72 | 0.47
| Analytical results for Zn in OREAS 251c (Certified Value 42.2 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.28 | 1.30 | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | 1
|  |  | 01g | 0.2g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 01g | 15g | 0.1g | 01g | 40g | 0.5g | 1
| 1 | 1 | 37 | 41.9456 | 37.203 | 47.404 | 47.5 | 41 | 40 | 42 | 37 | 44 | 39 | 42.6908 | 29 | 47 | 36.5 | 46 | 36 | 46.4 | 55.9 | 41 | 45.311 | 1
| 1 | 2 | 38 | 42.5114 | 36.256 | 48.6805 | 49.7 | 43 | 42 | 42 | 38 | 44 | 39 | 43.6361 | 31 | 49 | 39.4 | 46 | 34.6 | 47.6 | 54.3 | 37 | 47.429 | 16
| 1 | 3 | 36 | 41.4899 | 35.425 | 47.5995 | 45.9 | 39 | 41 | 42 | 38 | 44 | 40 | 43.521 | 28 | 51 | 37.9 | 45 | 37.1 | 44.4 | 53.3 | 39 | 45.69 | 16
| 1 | 4 | 38 | 40.558 | 34.975 | 48.324 | 47.3 | 40 | 40 | 42 | 38 | 44 | 39 | 43.3006 | 29 | 49 | 40 | 45 | 36.9 | 47 | 55.2 | 40 | 42.742 | 42.1732
| 1 | 5 | 38 | 42.1778 | 37.713 | 47.4845 | 48.2 | 43 | 41 | 42 | 38 | 43 | 42 | 44.5118 | 26 | 51 | 44.3 | 46 | 36 | 37.5 | 53.7 | 38 | 44.029 | 121
| 1 | 6 | 39 | 42.0425 | 36.128 | 45.081 | 48.3 | 41 | 40 | 43 | 38 | 43 | 40 | 44.7724 | 31 | 51 | 35.7 | 46 | 37.2 | 45.1 | 55.4 | 41 | 45.996
| Mean |  | 37.6667 | 41.7875 | 36.2833 | 47.4289 | 47.8167 | 41.1667 | 40.6667 | 42.1667 | 37.8333 | 43.6667 | 39.8333 | 43.7388 | 29 | 49.6667 | 38.9667 | 45.6667 | 36.3 | 44.6667 | 54.6333 | 39.3333 | 45.1995
| Median |  | 38 | 41.9941 | 36.192 | 47.542 | 47.85 | 41 | 40.5 | 42 | 38 | 44 | 39.5 | 43.5785 | 29 | 50 | 38.65 | 46 | 36.45 | 45.75 | 54.75 | 39.5 | 45.5005
| Std Dev. |  | 1.0328 | 0.68799 | 1.0357 | 1.258 | 1.2624 | 1.6021 | 0.8165 | 0.40825 | 0.40825 | 0.5164 | 1.169 | 0.77641 | 1.8974 | 1.633 | 3.0852 | 0.5164 | 0.98793 | 3.706 | 1.027 | 1.633 | 1.629
| Rel.Std.Dev. |  | 2.742% | 1.646% | 2.854% | 2.652% | 2.64% | 3.892% | 2.008% | 0.9682% | 1.079% | 1.183% | 2.935% | 1.775% | 6.543% | 3.288% | 7.918% | 1.131% | 2.722% | 8.297% | 1.88% | 4.152% | 3.604%
| PDM3 |  | -10.69% | -0.9145% | -13.97% | 12.46% | 13.38% | -2.387% | -3.572% | -0.01546% | -10.29% | 3.541% | -5.548% | 3.712% | -31.24% | 17.77% | -7.603% | 8.284% | -13.93% | 5.912% | 29.55% | -6.734% | 7.176%
| Z-Score (Absolute) |  | 0.81 | 0 | 1.09 | 1.11 | 1.19 | 0.12 | 0.22 | 0.07 | 0.78 | 0.37 | 0.39 | 0.39 | 2.53 | 1.56 | 0.56 | 0.77 | 1.08 | 0.57 | 2.54 | 0.49 | 0.67
| Analytical results for Zr in OREAS 251c (Certified Value 29.4 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.21 | 1.23 | 1.28 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 1
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.1g | 01g | 40g | [6 empty] | 1
| 1 | 1 | 7.9 | 29.0682 | 42.042 | 2.6 | 27.4 | 31.4 | 20.6 | 22.7 | 20.4 | 28.3 | 36.3077 | 31 | 34.8 | 41.6 | 2 | [6 empty] | 1
| 1 | 2 | 8.1 | 28.5662 | 42.336 | 2 | 26.7 | 32.1 | 21.7 | 23.4 | 20.7 | 24.6 | 36.1109 | 30.2 | 37.3 | 40.2 | 2 | [6 empty] | 17
| 1 | 3 | 8 | 28.4049 | 41.7284 | 2.5 | 27 | 33.3 | 18.3 | 23.3 | 22.4 | 27 | 36.2524 | 31.3 | 35.4 | 41.4 | 2 | [6 empty] | 16
| 1 | 4 | 8 | 28.056 | 42.0028 | 2.5 | 27.8 | 31.1 | 18.4 | 22.7 | 21.3 | 27.2 | 36.6019 | 32 | 37.6 | 42.2 | 2 | [6 empty] | 29.3942
| 1 | 5 | 7.8 | 30.0747 | 42.2674 | 3.5 | 27.1 | 30.1 | 21.3 | 24 | 20.4 | 27.4 | 36.5977 | 31.1 | 30.5 | 41.6 | 2 | [6 empty] | 122
| 1 | 6 | 7.7 | 29.5399 | 36.9852 | 2.3 | 26.9 | 31.9 | 20.5 | 21.3 | 19.6 | 29.7 | 36.8397 | 32 | 37.4 | 40 | 2
| Mean |  | 7.9167 | 28.9516 | 41.227 | 2.5667 | 27.15 | 31.65 | 20.1333 | 22.9 | 20.8 | 27.3667 | 36.4517 | 31.2667 | 35.5 | 41.1667 | 2
| Median |  | 7.95 | 28.8172 | 42.0224 | 2.5 | 27.05 | 31.65 | 20.55 | 23 | 20.55 | 27.3 | 36.4527 | 31.2 | 36.35 | 41.5 | 2
| Std Dev. |  | 0.1472 | 0.75771 | 2.0891 | 0.50465 | 0.3937 | 1.0728 | 1.4514 | 0.92304 | 0.95708 | 1.6813 | 0.2724 | 0.6802 | 2.7115 | 0.87101 | 0
| Rel.Std.Dev. |  | 1.859% | 2.617% | 5.067% | 19.66% | 1.45% | 3.39% | 7.209% | 4.031% | 4.601% | 6.143% | 0.7473% | 2.175% | 7.638% | 2.116% | 0%
| PDM3 |  | -73.07% | -1.506% | 40.26% | -91.27% | -7.635% | 7.674% | -31.51% | -22.09% | -29.24% | -6.898% | 24.01% | 6.37% | 20.77% | 40.05% | -93.2%
| Z-Score (Absolute) |  | 1.81 | 0.15 | 1.29 | 2.31 | 0.02 | 0.4 | 0.67 | 0.42 | 0.61 | 0 | 0.85 | 0.36 | 0.76 | 1.29 | 2.36
Results from laboratories 1.02, 1.05, 1.07 and 1.28 were removed due to their high Z score reading.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C708: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C853: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 251c (Indicative Value 13.5 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.5 | 1
| 1 | 2 | 13.49 | 10
| Mean |  | 13.495 | 16
| Median |  | 13.495 | 13.495
| Std Dev. |  | 0.0070711 | 16
| Rel.Std.Dev. |  | 0.0524%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 251c (Indicative Value 190 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 190 | 1
| 1 | 2 | 190 | 11
| Mean |  | 190 | 16
| Median |  | 190 | 190
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 251c (Indicative Value 575 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 569 | 1
| 1 | 2 | 581 | 12
| Mean |  | 575 | 16
| Median |  | 575 | 574.9975
| Std Dev. |  | 8.4853 | 18
| Rel.Std.Dev. |  | 1.476%
| PDM3 |  | 0.0004348%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 251c (Indicative Value 2.41 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.41 | 1
| 1 | 2 | 2.4 | 13
| Mean |  | 2.405 | 16
| Median |  | 2.405 | 2.405
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.294%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 251c (Indicative Value 135 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 130 | 1
| 1 | 2 | 140 | 14
| Mean |  | 135 | 16
| Median |  | 135 | 135
| Std Dev. |  | 7.0711 | 20
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 251c (Indicative Value 20 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 15
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 251c (Indicative Value 205 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 205 | 1
| 1 | 2 | 205 | 16
| Mean |  | 205 | 16
| Median |  | 205 | 204.61
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.1906%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 251c (Indicative Value 50 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 17
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 251c (Indicative Value 6.43 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.434 | 1
| 1 | 2 | 6.434 | 10
| Mean |  | 6.434 | 16
| Median |  | 6.434 | 6.4337
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.00544%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 251c (Indicative Value 2.3 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.3 | 1
| 1 | 2 | 2.3 | 11
| Mean |  | 2.3 | 16
| Median |  | 2.3 | 2.3
| Std Dev. |  | 0 | 17
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 251c (Indicative Value 2.38 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.37 | 1
| 1 | 2 | 2.38 | 12
| Mean |  | 2.375 | 16
| Median |  | 2.375 | 2.375
| Std Dev. |  | 0.0070711 | 18
| Rel.Std.Dev. |  | 0.2977%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 251c (Indicative Value 0.058 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.057 | 1
| 1 | 2 | 0.058 | 13
| Mean |  | 0.0575 | 16
| Median |  | 0.0575 | 0.0575
| Std Dev. |  | 0.00070711 | 19
| Rel.Std.Dev. |  | 1.23%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 251c (Indicative Value 1.32 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.32 | 1
| 1 | 2 | 1.31 | 14
| Mean |  | 1.315 | 16
| Median |  | 1.315 | 1.315
| Std Dev. |  | 0.0070711 | 20
| Rel.Std.Dev. |  | 0.5377%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 251c (Indicative Value 95 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 100 | 15
| Mean |  | 95 | 16
| Median |  | 95 | 95
| Std Dev. |  | 7.0711 | 21
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 251c (Indicative Value 0.207 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.2085 | 1
| 1 | 2 | 0.2062 | 16
| Mean |  | 0.20735 | 16
| Median |  | 0.20735 | 0.20739
| Std Dev. |  | 0.0016263 | 22
| Rel.Std.Dev. |  | 0.7843%
| PDM3 |  | -0.01919%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 251c (Indicative Value 50 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 40 | 1
| 1 | 2 | 60 | 17
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 14.1421 | 23
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 251c (Indicative Value 0.011 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.011 | 1
| 1 | 2 | 0.011 | 21
| Mean |  | 0.011 | 16
| Median |  | 0.011 | 0.011
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 251c (Indicative Value 67.61 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.62 | 1
| 1 | 2 | 67.6 | 11
| Mean |  | 67.61 | 16
| Median |  | 67.61 | 67.61
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.02092%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 251c (Indicative Value 65 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 70 | 12
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 7.0711 | 18
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 251c (Indicative Value 200 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 180 | 1
| 1 | 2 | 220 | 13
| Mean |  | 200 | 16
| Median |  | 200 | 200
| Std Dev. |  | 28.2843 | 19
| Rel.Std.Dev. |  | 14.14%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 251c (Indicative Value 1.13 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.133 | 1
| 1 | 2 | 1.135 | 14
| Mean |  | 1.134 | 16
| Median |  | 1.134 | 1.134
| Std Dev. |  | 0.0014142 | 20
| Rel.Std.Dev. |  | 0.1247%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 251c (Indicative Value 196 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 196 | 1
| 1 | 2 | 196 | 15
| Mean |  | 196 | 16
| Median |  | 196 | 196.372
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.1894%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 251c (Indicative Value 70 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 70 | 16
| Mean |  | 70 | 16
| Median |  | 70 | 70
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 251c (Indicative Value 275 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 270 | 1
| 1 | 2 | 280 | 17
| Mean |  | 275 | 16
| Median |  | 275 | 275
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 251c (Indicative Value 2.42 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.41 | 1
| 1 | 2 | 2.42 | 19
| Mean |  | 2.415 | 16
| Median |  | 2.415 | 2.415
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.2928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 251c (Indicative Value 0.085 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.09 | 21
| Mean |  | 0.085 | 16
| Median |  | 0.085 | 0.085
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 251c (Indicative Value 0.15 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.15 | 1
| 1 | 2 | 0.15 | 21
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 251c (Indicative Value 0.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 23
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 251c (Indicative Value 167 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 165 | 1
| 1 | 2 | 169 | 11
| Mean |  | 167 | 16
| Median |  | 167 | 167
| Std Dev. |  | 2.8284 | 30
| Rel.Std.Dev. |  | 1.694%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 251c (Indicative Value 505 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 504 | 1
| 1 | 2 | 506 | 25
| Mean |  | 505 | 16
| Median |  | 505 | 505
| Std Dev. |  | 1.4142 | 31
| Rel.Std.Dev. |  | 0.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 251c (Indicative Value 2.3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.2 | 26
| Mean |  | 2.3 | 16
| Median |  | 2.3 | 2.3
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 251c (Indicative Value 8 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.8 | 1
| 1 | 2 | 8.2 | 27
| Mean |  | 8 | 16
| Median |  | 8 | 8
| Std Dev. |  | 0.28284 | 33
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 251c (Indicative Value < 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 28
| Mean |  |  | 16
| Median |  |  | < 0.1
| Std Dev. |  |  | 34
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 251c (Indicative Value 89 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 89 | 1
| 1 | 2 | 89 | 29
| Mean |  | 89 | 16
| Median |  | 89 | 89
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 251c (Indicative Value 18.1 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.2 | 1
| 1 | 2 | 18 | 15
| Mean |  | 18.1 | 16
| Median |  | 18.1 | 18.1
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 0.7813%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 251c (Indicative Value 146 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 145 | 1
| 1 | 2 | 147 | 31
| Mean |  | 146 | 16
| Median |  | 146 | 146
| Std Dev. |  | 1.4142 | 37
| Rel.Std.Dev. |  | 0.9686%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 251c (Indicative Value 5.57 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.56 | 1
| 1 | 2 | 5.58 | 32
| Mean |  | 5.57 | 16
| Median |  | 5.57 | 5.57
| Std Dev. |  | 0.014142 | 38
| Rel.Std.Dev. |  | 0.2539%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 251c (Indicative Value 30 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 17
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 251c (Indicative Value 5.35 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.3 | 1
| 1 | 2 | 5.39 | 4
| Mean |  | 5.345 | 16
| Median |  | 5.345 | 5.345
| Std Dev. |  | 0.06364 | 40
| Rel.Std.Dev. |  | 1.191%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 251c (Indicative Value 2.88 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.82 | 1
| 1 | 2 | 2.93 | 5
| Mean |  | 2.875 | 16
| Median |  | 2.875 | 2.875
| Std Dev. |  | 0.077782 | 41
| Rel.Std.Dev. |  | 2.705%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 251c (Indicative Value 1.8 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.81 | 1
| 1 | 2 | 1.78 | 36
| Mean |  | 1.795 | 16
| Median |  | 1.795 | 1.795
| Std Dev. |  | 0.021213 | 42
| Rel.Std.Dev. |  | 1.182%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 251c (Indicative Value 18.6 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.6 | 1
| 1 | 2 | 18.6 | 37
| Mean |  | 18.6 | 16
| Median |  | 18.6 | 18.6
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 251c (Indicative Value 6.2 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.14 | 1
| 1 | 2 | 6.26 | 6
| Mean |  | 6.2 | 16
| Median |  | 6.2 | 6.2
| Std Dev. |  | 0.084853 | 44
| Rel.Std.Dev. |  | 1.369%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 251c (Indicative Value 1.38 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.35 | 1
| 1 | 2 | 1.4 | 2
| Mean |  | 1.375 | 16
| Median |  | 1.375 | 1.375
| Std Dev. |  | 0.035355 | 45
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 251c (Indicative Value 7.34 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.51 | 1
| 1 | 2 | 7.17 | 8
| Mean |  | 7.34 | 16
| Median |  | 7.34 | 7.34
| Std Dev. |  | 0.24042 | 29
| Rel.Std.Dev. |  | 3.275%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 251c (Indicative Value 1.03 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.03 | 1
| 1 | 2 | 1.03 | 24
| Mean |  | 1.03 | 16
| Median |  | 1.03 | 1.03
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 251c (Indicative Value 0.075 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.05 | 25
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.035355 | 31
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 251c (Indicative Value 46.5 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46.5 | 1
| 1 | 2 | 46.4 | 26
| Mean |  | 46.45 | 16
| Median |  | 46.45 | 46.45
| Std Dev. |  | 0.070711 | 32
| Rel.Std.Dev. |  | 0.1522%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 251c (Indicative Value 0.36 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.35 | 5
| Mean |  | 0.355 | 16
| Median |  | 0.355 | 0.355
| Std Dev. |  | 0.0070711 | 33
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 251c (Indicative Value 0.043 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0427 | 1
| 1 | 2 | 0.0429 | 28
| Mean |  | 0.0428 | 16
| Median |  | 0.0428 | 0.0428
| Std Dev. |  | 0.00014142 | 34
| Rel.Std.Dev. |  | 0.3304%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 251c (Indicative Value 2.7 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.6 | 29
| Mean |  | 2.7 | 16
| Median |  | 2.7 | 2.7
| Std Dev. |  | 0.14142 | 35
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 251c (Indicative Value 30 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 29.8 | 1
| 1 | 2 | 30.1 | 30
| Mean |  | 29.95 | 16
| Median |  | 29.95 | 29.95
| Std Dev. |  | 0.21213 | 36
| Rel.Std.Dev. |  | 0.7083%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 251c (Indicative Value 40.4 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.3 | 1
| 1 | 2 | 40.5 | 6
| Mean |  | 40.4 | 16
| Median |  | 40.4 | 40.4
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 0.3501%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 251c (Indicative Value 75 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 76 | 1
| 1 | 2 | 74 | 15
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 1.4142 | 38
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 251c (Indicative Value 13 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13 | 1
| 1 | 2 | 13 | 17
| Mean |  | 13 | 16
| Median |  | 13 | 13
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 251c (Indicative Value 10.6 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.5 | 1
| 1 | 2 | 10.7 | 34
| Mean |  | 10.6 | 16
| Median |  | 10.6 | 10.6
| Std Dev. |  | 0.14142 | 40
| Rel.Std.Dev. |  | 1.334%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 251c (Indicative Value 95 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 94.4 | 1
| 1 | 2 | 96.4 | 35
| Mean |  | 95.4 | 16
| Median |  | 95.4 | 95.4
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 1.482%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 251c (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 36
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 42
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 251c (Indicative Value 2.8 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.9 | 1
| 1 | 2 | 2.7 | 37
| Mean |  | 2.8 | 16
| Median |  | 2.8 | 2.8
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 5.051%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 251c (Indicative Value 13.7 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.7 | 1
| 1 | 2 | 13.6 | 38
| Mean |  | 13.65 | 16
| Median |  | 13.65 | 13.65
| Std Dev. |  | 0.070711 | 44
| Rel.Std.Dev. |  | 0.518%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 251c (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 39
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 45
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 251c (Indicative Value 7.76 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.77 | 1
| 1 | 2 | 7.74 | 5
| Mean |  | 7.755 | 16
| Median |  | 7.755 | 7.755
| Std Dev. |  | 0.021213 | 29
| Rel.Std.Dev. |  | 0.2735%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 251c (Indicative Value 10.2 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.2 | 1
| 1 | 2 | 10.2 | 12
| Mean |  | 10.2 | 16
| Median |  | 10.2 | 10.2
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 251c (Indicative Value 277 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 275 | 1
| 1 | 2 | 278 | 13
| Mean |  | 276.5 | 16
| Median |  | 276.5 | 276.5
| Std Dev. |  | 2.1213 | 31
| Rel.Std.Dev. |  | 0.7672%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 251c (Indicative Value 2.04 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.03 | 1
| 1 | 2 | 2.05 | 26
| Mean |  | 2.04 | 16
| Median |  | 2.04 | 2.04
| Std Dev. |  | 0.014142 | 32
| Rel.Std.Dev. |  | 0.6932%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 251c (Indicative Value 0.94 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.95 | 1
| 1 | 2 | 0.93 | 27
| Mean |  | 0.94 | 16
| Median |  | 0.94 | 0.94
| Std Dev. |  | 0.014142 | 33
| Rel.Std.Dev. |  | 1.504%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 251c (Indicative Value 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 28
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 251c (Indicative Value 12.8 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 13 | 29
| Mean |  | 12.8 | 16
| Median |  | 12.8 | 12.8
| Std Dev. |  | 0.28284 | 35
| Rel.Std.Dev. |  | 2.21%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 251c (Indicative Value 0.682 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.68 | 1
| 1 | 2 | 0.683 | 30
| Mean |  | 0.6815 | 16
| Median |  | 0.6815 | 0.6815
| Std Dev. |  | 0.0021213 | 36
| Rel.Std.Dev. |  | 0.3113%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 251c (Indicative Value < 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 31
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 37
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 251c (Indicative Value 0.42 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.41 | 1
| 1 | 2 | 0.42 | 6
| Mean |  | 0.415 | 16
| Median |  | 0.415 | 0.415
| Std Dev. |  | 0.0070711 | 38
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 251c (Indicative Value 2.53 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.5 | 1
| 1 | 2 | 2.55 | 33
| Mean |  | 2.525 | 16
| Median |  | 2.525 | 2.525
| Std Dev. |  | 0.035355 | 39
| Rel.Std.Dev. |  | 1.4%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 251c (Indicative Value 100 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 34
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 251c (Indicative Value 35 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 34 | 1
| 1 | 2 | 36 | 35
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 251c (Indicative Value 26.8 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.5 | 1
| 1 | 2 | 27 | 36
| Mean |  | 26.75 | 16
| Median |  | 26.75 | 26.75
| Std Dev. |  | 0.35355 | 42
| Rel.Std.Dev. |  | 1.322%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 251c (Indicative Value 2.63 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.64 | 1
| 1 | 2 | 2.62 | 7
| Mean |  | 2.63 | 16
| Median |  | 2.63 | 2.63
| Std Dev. |  | 0.014142 | 43
| Rel.Std.Dev. |  | 0.5377%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 251c (Indicative Value 65 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 65 | 1
| 1 | 2 | 65 | 16
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 251c (Indicative Value 277 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 277 | 1
| 1 | 2 | 276 | 17
| Mean |  | 276.5 | 16
| Median |  | 276.5 | 276.5
| Std Dev. |  | 0.70711 | 45
| Rel.Std.Dev. |  | 0.2557%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.