## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 506b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.35082 | 0.34589 | 0.35576 | 0.34514 | 0.35651
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 0.34831 | 0.34068 | 0.35594 | 0.34213 | 0.3545
4-Acid Digestion
Ag, Silver (ppm) | 1.3353 | 1.248 | 1.4226 | 1.2872 | 1.3834
Al, Aluminium (wt.%) | 7.55 | 7.3189 | 7.7811 | 7.4208 | 7.6791
As, Arsenic (ppm) | 13.5864 | 12.3437 | 14.8291 | 12.861 | 14.3118
Ba, Barium (ppm) | 937.755 | 899.1027 | 976.4073 | 921.5285 | 953.9816
Be, Beryllium (ppm) | 2.246 | 2.0802 | 2.4118 | 2.1662 | 2.3258
Bi, Bismuth (ppm) | 0.93177 | 0.72611 | 1.1374 | 0.8631 | 1.0004
Ca, Calcium (wt.%) | 1.6707 | 1.6296 | 1.7118 | 1.6439 | 1.6976
Cd, Cadmium (ppm) | 0.39658 | 0.34718 | 0.44599 | 0.3644 | 0.42876
Ce, Cerium (ppm) | 63.8192 | 59.0286 | 68.6098 | 61.2001 | 66.4382
Co, Cobalt (ppm) | 7.3353 | 6.9741 | 7.6965 | 7.0741 | 7.5964
Cr, Chromium (ppm) | 39.5348 | 36.9531 | 42.1166 | 37.3777 | 41.692
Cs, Caesium (ppm) | 8.6361 | 8.3069 | 8.9653 | 8.3391 | 8.9331
Cu, Copper (wt.%) | 0.44603 | 0.43548 | 0.45659 | 0.43854 | 0.45352
Dy, Dysprosium (ppm) | 3.2142 | 2.9882 | 3.4402 | 3.0557 | 3.3728
Er, Erbium (ppm) | 1.3315 | 1.1839 | 1.4791 | 1.2288 | 1.4342
Eu, Europium (ppm) | 1.2583 | 1.1803 | 1.3363 | 1.2206 | 1.296
Fe, Iron (wt.%) | 2.7484 | 2.6811 | 2.8157 | 2.7008 | 2.796
Ga, Gallium (ppm) | 19.3091 | 18.5196 | 20.0985 | 18.6126 | 20.0055
Gd, Gadolinium (ppm) | 5.1346 | 4.7498 | 5.5193 | 4.8654 | 5.4038
Ge, Germanium (ppm) | 0.16626 | 0.10742 | 0.2251 | 0.15354 | 0.17899
Hf, Hafnium (ppm) | 2.0909 | 1.9343 | 2.2475 | 1.9906 | 2.1912
Ho, Holmium (ppm) | 0.53363 | 0.48323 | 0.58403 | 0.49415 | 0.57311
In, Indium (ppm) | 0.078844 | 0.070615 | 0.087074 | 0.074366 | 0.083323
K, Potassium (wt.%) | 3.3911 | 3.3138 | 3.4685 | 3.2914 | 3.4909
La, Lanthanum (ppm) | 31.7644 | 29.4514 | 34.0773 | 30.7604 | 32.7683
Li, Lithium (ppm) | 41.8221 | 39.9201 | 43.724 | 40.6191 | 43.0251
Lu, Lutetium (ppm) | 0.17357 | 0.1525 | 0.19464 | IND | IND
Mg, Magnesium (wt.%) | 0.68638 | 0.66341 | 0.70934 | 0.67413 | 0.69862
Mn, Manganese (wt.%) | 0.032867 | 0.031877 | 0.033857 | 0.032166 | 0.033568
Mo, Molybdenum (ppm) | 88.1682 | 84.9318 | 91.4047 | 86.1334 | 90.2031
Na, Sodium (wt.%) | 2.1826 | 2.1323 | 2.2329 | 2.1394 | 2.2258
Nb, Niobium (ppm) | 11.1302 | 10.5898 | 11.6706 | 10.7648 | 11.4957
Nd, Neodymium (ppm) | 28.9279 | 26.3661 | 31.4897 | 27.9806 | 29.8752
Ni, Nickel (ppm) | 15.6991 | 14.8998 | 16.4984 | 15.1549 | 16.2433
P, Phosphorus (wt.%) | 0.084725 | 0.082374 | 0.087076 | 0.082828 | 0.086621
Pb, Lead (ppm) | 55.6601 | 53.6652 | 57.655 | 53.6877 | 57.6325
Pr, Praseodymium (ppm) | 7.4581 | 6.7807 | 8.1354 | 7.1375 | 7.7786
Rb, Rubidium (ppm) | 149.4494 | 141.9456 | 156.9532 | 144.1114 | 154.7873
Re, Rhenium (ppm) | 0.035243 | 0.032054 | 0.038432 | 0.032081 | 0.038405
S, Sulphur (wt.%) | 0.61618 | 0.59553 | 0.63683 | 0.60276 | 0.6296
Sb, Antimony (ppm) | 1.36 | 1.254 | 1.4661 | 1.2555 | 1.4645
Sc, Scandium (ppm) | 7.7035 | 7.3159 | 8.091 | 7.4225 | 7.9845
Se, Selenium (ppm) | 5.0249 | 4.3162 | 5.7336 | 4.6775 | 5.3723
Sm, Samarium (ppm) | 5.9024 | 5.3539 | 6.4509 | 5.5765 | 6.2283
Sn, Tin (ppm) | 4.2747 | 4.0405 | 4.5089 | 4.1211 | 4.4283
Sr, Strontium (ppm) | 276.2316 | 266.9156 | 285.5477 | 269.5991 | 282.8642
Ta, Tantalum (ppm) | 0.92553 | 0.86425 | 0.98682 | 0.89686 | 0.95421
Tb, Terbium (ppm) | 0.66747 | 0.61971 | 0.71522 | 0.62713 | 0.7078
Te, Tellurium (ppm) | 0.21185 | 0.17474 | 0.24896 | 0.17635 | 0.24736
Th, Thorium (ppm) | 11.9562 | 11.0316 | 12.8808 | 11.5813 | 12.331
Ti, Titanium (wt.%) | 0.31876 | 0.30803 | 0.32948 | 0.3108 | 0.32671
Tl, Thallium (ppm) | 0.81087 | 0.76213 | 0.85961 | 0.78522 | 0.83652
Tm, Thulium (ppm) | 0.17588 | 0.14845 | 0.2033 | IND | IND
U, Uranium (ppm) | 3.5015 | 2.995 | 4.0081 | 3.2616 | 3.7415
V, Vanadium (ppm) | 65.4856 | 62.8197 | 68.1515 | 63.4513 | 67.5199
W, Tungsten (ppm) | 10.3101 | 9.4218 | 11.1984 | 9.5343 | 11.0859
Y, Yttrium (ppm) | 14.2861 | 13.6097 | 14.9624 | 13.8214 | 14.7507
Yb, Ytterbium (ppm) | 1.1717 | 1.037 | 1.3064 | 1.1201 | 1.2234
Zn, Zinc (ppm) | 111.906 | 107.3206 | 116.4913 | 109.1617 | 114.6502
Zr, Zirconium (ppm) | 71.0075 | 67.6387 | 74.3762 | 68.9398 | 73.0752
Aqua Regia Digestion
Ag, Silver (ppm) | 1.3139 | 1.2492 | 1.3786 | 1.2726 | 1.3552
Al, Aluminium (wt.%) | 1.6783 | 1.6245 | 1.7322 | 1.6441 | 1.7126
As, Arsenic (ppm) | 12.9541 | 12.2829 | 13.6253 | 12.4625 | 13.4456
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 425.9311 | 409.515 | 442.3471 | 414.4247 | 437.4375
Be, Beryllium (ppm) | 1.2125 | 1.1694 | 1.2555 | 1.1861 | 1.2388
Bi, Bismuth (ppm) | 0.95816 | 0.79693 | 1.1194 | 0.87604 | 1.0403
Ca, Calcium (wt.%) | 0.69545 | 0.67469 | 0.71621 | 0.68203 | 0.70887
Cd, Cadmium (ppm) | 0.31946 | 0.28829 | 0.35064 | 0.29933 | 0.3396
Ce, Cerium (ppm) | 29.4208 | 27.9714 | 30.8703 | 28.4599 | 30.3818
Co, Cobalt (ppm) | 7.0951 | 6.8073 | 7.3828 | 6.8885 | 7.3016
Cr, Chromium (ppm) | 45.4673 | 43.7796 | 47.155 | 44.2206 | 46.714
Cs, Caesium (ppm) | 6.7809 | 6.4046 | 7.1573 | 6.592 | 6.9699
Cu, Copper (wt.%) | 0.44268 | 0.42978 | 0.45558 | 0.43694 | 0.44842
Fe, Iron (wt.%) | 2.6578 | 2.5974 | 2.7182 | 2.6104 | 2.7053
Ga, Gallium (ppm) | 8.027 | 7.7159 | 8.3381 | 7.8146 | 8.2394
Ge, Germanium (ppm) | 0.11133 | 0.077764 | 0.1449 | IND | IND
Hf, Hafnium (ppm) | 0.30977 | 0.28401 | 0.33553 | 0.29672 | 0.32282
In, Indium (ppm) | 0.072341 | 0.064262 | 0.08042 | 0.066596 | 0.078087
K, Potassium (wt.%) | 0.85347 | 0.82478 | 0.88217 | 0.83989 | 0.86705
La, Lanthanum (ppm) | 13.775 | 13.0038 | 14.5461 | 13.257 | 14.2929
Li, Lithium (ppm) | 33.4048 | 32.1944 | 34.6153 | 32.396 | 34.4137
Mg, Magnesium (wt.%) | 0.64469 | 0.62561 | 0.66376 | 0.63294 | 0.65643
Mn, Manganese (wt.%) | 0.028561 | 0.027827 | 0.029295 | 0.028047 | 0.029075
Mo, Molybdenum (ppm) | 83.9642 | 81.7382 | 86.1903 | 82.8741 | 85.0544
Na, Sodium (wt.%) | 0.14197 | 0.13227 | 0.15167 | 0.13513 | 0.14881
Nb, Niobium (ppm) | 0.75864 | 0.62478 | 0.89249 | 0.70725 | 0.81003
Ni, Nickel (ppm) | 15.5003 | 14.8425 | 16.158 | 15.0044 | 15.9961
P, Phosphorus (wt.%) | 0.065458 | 0.063432 | 0.067484 | 0.06397 | 0.066945
Pb, Lead (ppm) | 39.5071 | 37.9239 | 41.0902 | 38.4565 | 40.5576
Rb, Rubidium (ppm) | 79.069 | 75.3681 | 82.7698 | 77.3187 | 80.8193
Re, Rhenium (ppm) | 0.034553 | 0.030784 | 0.038322 | 0.031283 | 0.037823
S, Sulphur (wt.%) | 0.6144 | 0.59755 | 0.63126 | 0.60283 | 0.62598
Sb, Antimony (ppm) | 0.86527 | 0.7799 | 0.95064 | 0.83247 | 0.89808
Sc, Scandium (ppm) | 6.6782 | 6.3417 | 7.0147 | 6.4743 | 6.8821
Se, Selenium (ppm) | 4.8883 | 4.5203 | 5.2563 | 4.6732 | 5.1034
Sn, Tin (ppm) | 2.9487 | 2.8056 | 3.0918 | 2.8209 | 3.0765
Sr, Strontium (ppm) | 94.1065 | 91.518 | 96.695 | 92.1791 | 96.0339
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.45434 | 0.41085 | 0.49784 | 0.43128 | 0.47741
Te, Tellurium (ppm) | 0.19949 | 0.16348 | 0.23549 | 0.16594 | 0.23303
Th, Thorium (ppm) | 5.665 | 5.2569 | 6.0731 | 5.4461 | 5.884
Ti, Titanium (wt.%) | 0.20437 | 0.19573 | 0.213 | 0.19913 | 0.20961
Tl, Thallium (ppm) | 0.51169 | 0.48606 | 0.53733 | 0.494 | 0.52939
U, Uranium (ppm) | 2.8226 | 2.4926 | 3.1527 | 2.6057 | 3.0396
V, Vanadium (ppm) | 54.3956 | 52.7004 | 56.0908 | 52.9422 | 55.8489
W, Tungsten (ppm) | 6.4245 | 5.739 | 7.1101 | 5.9696 | 6.8795
Y, Yttrium (ppm) | 9.6349 | 9.289 | 9.9808 | 9.4057 | 9.8641
Yb, Ytterbium (ppm) | 0.70662 | 0.68089 | 0.73236 | IND | IND
Zn, Zinc (ppm) | 105.7684 | 102.3215 | 109.2152 | 103.6164 | 107.9204
Zr, Zirconium (ppm) | 8.6034 | 8.0172 | 9.1895 | 8.2445 | 8.9622
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 506b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 8.3333 | Pt | ppb | < 10
Sequential Leach
Cu-Res(4A) | wt.% | 0.085733 | Cu-Sol(CN) | wt.% | 0.30773 | Cu-Sol(H2SO4) | wt.% | 0.065319
Gas / Liquid Pycnometry
SG | Unity | 2.6876
Infrared Combustion
C | wt.% | 0.11167 | S | wt.% | 0.5425
Alkaline Leach
S-(Sulphate) | wt.% | 0.43167 | S-(Sulphide) | wt.% | 0.053333
Peroxide Fusion ICP
Ag | ppm | 1.225 | Hg | ppm | < 5 | Si | wt.% | 31.8985
Al | wt.% | 6.2975 | Ho | ppm | 0.63333 | Sm | ppm | 5.7333
As | ppm | 19.2 | K | wt.% | 3.0231 | Sn | ppm | < 5
Ba | ppm | 679.7583 | La | ppm | 30.2583 | Sr | ppm | 211.6661
Be | ppm | < 2 | Li | ppm | 46.3333 | Ta | ppm | 1.1333
Bi | ppm | < 5 | Lu | ppm | < 0.5 | Tb | ppm | 0.73333
Ca | wt.% | 1.6639 | Mg | wt.% | 0.73462 | Th | ppm | 5.6833
Cd | ppm | < 1 | Mn | wt.% | 0.035875 | Ti | wt.% | 0.28417
Ce | ppm | 57.65 | Mo | ppm | 85.6417 | Tl | ppm | < 1
Co | ppm | 10.5083 | Na | wt.% | 0.15167 | Tm | ppm | < 1
Cr | ppm | 44.3333 | Nb | ppm | 12.4833 | U | ppm | 3.8167
Cu | wt.% | 0.43731 | Nd | ppm | 26.3833 | V | ppm | 63.0583
Dy | ppm | 3.6333 | Ni | ppm | 13.8067 | W | ppm | 10
Er | ppm | 1.7167 | P | wt.% | 0.076325 | Y | ppm | 22.05
Eu | ppm | 1.2833 | Pb | ppm | 45.15 | Yb | ppm | 1.4833
Fe | wt.% | 2.8526 | Pr | ppm | 6.8333 | Zn | ppm | 109.8833
Ga | ppm | 20.425 | S | wt.% | 0.55583 | Zr | ppm | 10.1
Gd | ppm | 5.1667 | Sb | ppm | < 5
Hf | ppm | 2.75 | Sc | ppm | 9.05
4-Acid Digestion
B | ppm | < 10 | Hg | ppm | 0.060833 | Pt | ppb | < 5
Aqua Regia Digestion
Dy | ppm | 2.2717 | Ho | ppm | 0.39791 | Pt | ppb | 1.5
Er | ppm | 0.90706 | Lu | ppm | 0.09596 | Sm | ppm | 3.3744
Eu | ppm | 0.42664 | Nd | ppm | 15.0219 | Tm | ppm | 0.1093
Gd | ppm | 3.1291 | Pd | ppb | < 10
Hg | ppm | 0.014568 | Pr | ppm | 3.7527
Borate Fusion XRF
Al2O3 | wt.% | 14.87 | Fe2O3 | wt.% | 3.996 | S | wt.% | 0.619
As | ppm | 10 | K2O | wt.% | 4.085 | SiO2 | wt.% | 67.27
BaO | ppm | 1105.335 | MgO | wt.% | 1.2 | Sn | ppm | < 10
CaO | wt.% | 2.335 | MnO | wt.% | 0.045 | Sr | ppm | 250
Cl | ppm | 85 | Na2O | wt.% | 2.955 | TiO2 | wt.% | 0.565
Co | ppm | 10 | Ni | ppm | 25 | V2O5 | ppm | 133.89
Cr2O3 | ppm | 65.7675 | P2O5 | wt.% | 0.1902 | Zn | ppm | 110
Cu | wt.% | 0.447 | Pb | ppm | 55 | Zr | ppm | 210
Thermogravimetry
LOI1000 | wt.% | 1.51
Laser Ablation ICP-MS
Ag | ppm | 1.25 | Hf | ppm | 6.22 | Sm | ppm | 6.545
As | ppm | 12.6 | Ho | ppm | 1.025 | Sn | ppm | 4.2
Ba | ppm | 917.5 | In | ppm | 0.05 | Sr | ppm | 267
Be | ppm | 2.4 | La | ppm | 33.5 | Ta | ppm | 1.03
Bi | ppm | 1.07 | Lu | ppm | 0.38 | Tb | ppm | 0.935
Cd | ppm | 0.4 | Mn | wt.% | 0.03225 | Te | ppm | 0.15
Ce | ppm | 65.85 | Mo | ppm | 85 | Th | ppm | 12.75
Co | ppm | 7.8 | Nb | ppm | 11.7 | Ti | wt.% | 0.3455
Cr | ppm | 46 | Nd | ppm | 31 | Tl | ppm | 0.6
Cs | ppm | 8.505 | Ni | ppm | 17 | Tm | ppm | 0.425
Cu | wt.% | 0.444 | Pb | ppm | 53 | U | ppm | 4.365
Dy | ppm | 5.17 | Pr | ppm | 8.025 | V | ppm | 70.05
Er | ppm | 2.785 | Rb | ppm | 143.5 | W | ppm | 12
Eu | ppm | 1.27 | Re | ppm | 0.03 | Y | ppm | 26.75
Ga | ppm | 18.55 | Sb | ppm | 1.3 | Yb | ppm | 2.63
Gd | ppm | 6.15 | Sc | ppm | 7.8 | Zn | ppm | 105
Ge | ppm | 1.45 | Se | ppm | < 5 | Zr | ppm | 221
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 506b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.35082 | 0.0080816 | 0.33466 | 0.36699 | 0.32658 | 0.37507 | 2.304% | 4.607% | 6.911% | 0.33328 | 0.36837
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 0.34831 | 0.015218 | 0.31788 | 0.37875 | 0.30266 | 0.39397 | 4.369% | 8.738% | 13.11% | 0.3309 | 0.36573
4-Acid Digestion
Ag, ppm | 1.3353 | 0.092128 | 1.151 | 1.5195 | 1.0589 | 1.6117 | 6.9% | 13.8% | 20.7% | 1.2685 | 1.402
Al, wt.% | 7.55 | 0.2702 | 7.0096 | 8.0904 | 6.7394 | 8.3606 | 3.579% | 7.158% | 10.74% | 7.1725 | 7.9275
As, ppm | 13.5864 | 1.0233 | 11.5399 | 15.6329 | 10.5166 | 16.6562 | 7.532% | 15.06% | 22.59% | 12.9071 | 14.2657
Ba, ppm | 937.755 | 54.116 | 829.523 | 1045.987 | 775.407 | 1100.103 | 5.771% | 11.54% | 17.31% | 890.8673 | 984.6428
Be, ppm | 2.246 | 0.14474 | 1.9565 | 2.5355 | 1.8118 | 2.6802 | 6.444% | 12.89% | 19.33% | 2.1337 | 2.3583
Bi, ppm | 0.93177 | 0.135 | 0.66177 | 1.2018 | 0.52677 | 1.3368 | 14.49% | 28.98% | 43.47% | 0.88518 | 0.97836
Ca, wt.% | 1.6707 | 0.052901 | 1.5649 | 1.7765 | 1.512 | 1.8294 | 3.166% | 6.333% | 9.499% | 1.5872 | 1.7543
Cd, ppm | 0.39658 | 0.051404 | 0.29377 | 0.49939 | 0.24237 | 0.5508 | 12.96% | 25.92% | 38.89% | 0.37675 | 0.41641
Ce, ppm | 63.8192 | 5.1137 | 53.5917 | 74.0466 | 48.478 | 79.1603 | 8.013% | 16.03% | 24.04% | 60.6282 | 67.0101
Co, ppm | 7.3353 | 0.44818 | 6.4389 | 8.2316 | 5.9907 | 8.6798 | 6.11% | 12.22% | 18.33% | 6.9685 | 7.7021
Cr, ppm | 39.5348 | 3.9845 | 31.5658 | 47.5038 | 27.5813 | 51.4883 | 10.08% | 20.16% | 30.24% | 37.5581 | 41.5116
Cs, ppm | 8.6361 | 0.28356 | 8.069 | 9.2032 | 7.7855 | 9.4868 | 3.283% | 6.567% | 9.85% | 8.2043 | 9.0679
Cu, wt.% | 0.44603 | 0.014274 | 0.41748 | 0.47458 | 0.40321 | 0.48885 | 3.2% | 6.4% | 9.601% | 0.42373 | 0.46833
Dy, ppm | 3.2142 | 0.13031 | 2.9536 | 3.4749 | 2.8233 | 3.6052 | 4.054% | 8.108% | 12.16% | 3.0535 | 3.375
Er, ppm | 1.3315 | 0.094784 | 1.142 | 1.5211 | 1.0472 | 1.6159 | 7.118% | 14.24% | 21.36% | 1.265 | 1.3981
Eu, ppm | 1.2583 | 0.032569 | 1.1932 | 1.3234 | 1.1606 | 1.356 | 2.588% | 5.177% | 7.765% | 1.1954 | 1.3212
Fe, wt.% | 2.7484 | 0.076277 | 2.5958 | 2.901 | 2.5196 | 2.9772 | 2.775% | 5.551% | 8.326% | 2.611 | 2.8858
Ga, ppm | 19.3091 | 0.71645 | 17.8762 | 20.7419 | 17.1597 | 21.4584 | 3.71% | 7.421% | 11.13% | 18.3436 | 20.2745
Gd, ppm | 5.1346 | 0.28962 | 4.5553 | 5.7138 | 4.2657 | 6.0034 | 5.641% | 11.28% | 16.92% | 4.8778 | 5.3913
Ge, ppm | 0.16626 | 0.051848 | 0.062566 | 0.26996 | 0.010719 | 0.3218 | 31.18% | 62.37% | 93.55% | 0.15795 | 0.17457
Hf, ppm | 2.0909 | 0.19363 | 1.7036 | 2.4781 | 1.51 | 2.6718 | 9.261% | 18.52% | 27.78% | 1.9863 | 2.1954
Ho, ppm | 0.53363 | 0.031059 | 0.47151 | 0.59575 | 0.44045 | 0.62681 | 5.82% | 11.64% | 17.46% | 0.50695 | 0.56031
In, ppm | 0.078844 | 0.00581 | 0.067224 | 0.090464 | 0.061414 | 0.096274 | 7.369% | 14.74% | 22.11% | 0.074902 | 0.082787
K, wt.% | 3.3911 | 0.085117 | 3.2209 | 3.5614 | 3.1358 | 3.6465 | 2.51% | 5.02% | 7.53% | 3.2216 | 3.5607
La, ppm | 31.7644 | 2.2993 | 27.1657 | 36.363 | 24.8664 | 38.6623 | 7.239% | 14.48% | 21.72% | 30.1761 | 33.3526
Li, ppm | 41.8221 | 2.868 | 36.0861 | 47.558 | 33.2181 | 50.426 | 6.858% | 13.72% | 20.57% | 39.731 | 43.9132
Lu, ppm | 0.17357 | 0.015331 | 0.14291 | 0.20423 | 0.12758 | 0.21956 | 8.833% | 17.67% | 26.5% | 0.16489 | 0.18225
Mg, wt.% | 0.68638 | 0.027912 | 0.63055 | 0.7422 | 0.60264 | 0.77011 | 4.067% | 8.133% | 12.2% | 0.65206 | 0.72069
Mn, wt.% | 0.032867 | 0.0010269 | 0.030813 | 0.034921 | 0.029786 | 0.035948 | 3.124% | 6.249% | 9.373% | 0.031224 | 0.03451
Mo, ppm | 88.1682 | 3.2955 | 81.5773 | 94.7592 | 78.2819 | 98.0546 | 3.738% | 7.475% | 11.21% | 83.7598 | 92.5767
Na, wt.% | 2.1826 | 0.069741 | 2.0431 | 2.3221 | 1.9734 | 2.3918 | 3.195% | 6.391% | 9.586% | 2.0735 | 2.2917
Nb, ppm | 11.1302 | 0.66433 | 9.8016 | 12.4589 | 9.1372 | 13.1232 | 5.969% | 11.94% | 17.91% | 10.5737 | 11.6867
Nd, ppm | 28.9279 | 2.1051 | 24.7177 | 33.1381 | 22.6126 | 35.2432 | 7.277% | 14.55% | 21.83% | 27.4815 | 30.3743
Ni, ppm | 15.6991 | 0.64276 | 14.4136 | 16.9846 | 13.7708 | 17.6274 | 4.094% | 8.188% | 12.28% | 14.9141 | 16.4841
P, wt.% | 0.084725 | 0.0033365 | 0.078052 | 0.091398 | 0.074715 | 0.094734 | 3.938% | 7.876% | 11.81% | 0.080489 | 0.088961
Pb, ppm | 55.6601 | 2.6025 | 50.4551 | 60.8652 | 47.8525 | 63.4677 | 4.676% | 9.352% | 14.03% | 52.8771 | 58.4431
Pr, ppm | 7.4581 | 0.62676 | 6.2045 | 8.7116 | 5.5778 | 9.3383 | 8.404% | 16.81% | 25.21% | 7.0852 | 7.831
Rb, ppm | 149.4494 | 7.3253 | 134.7988 | 164.1 | 127.4735 | 171.4253 | 4.902% | 9.803% | 14.7% | 141.9769 | 156.9218
Re, ppm | 0.035243 | 0.0036124 | 0.028018 | 0.042468 | 0.024406 | 0.04608 | 10.25% | 20.5% | 30.75% | 0.033481 | 0.037005
S, wt.% | 0.61618 | 0.022102 | 0.57198 | 0.66038 | 0.54987 | 0.68249 | 3.587% | 7.174% | 10.76% | 0.58537 | 0.64699
Sb, ppm | 1.36 | 0.11566 | 1.1287 | 1.5913 | 1.013 | 1.707 | 8.504% | 17.01% | 25.51% | 1.292 | 1.428
Sc, ppm | 7.7035 | 0.4683 | 6.7669 | 8.6401 | 6.2986 | 9.1084 | 6.079% | 12.16% | 18.24% | 7.3183 | 8.0887
Se, ppm | 5.0249 | 0.63237 | 3.7602 | 6.2896 | 3.1278 | 6.922 | 12.58% | 25.17% | 37.75% | 4.7736 | 5.2761
Sm, ppm | 5.9024 | 0.41993 | 5.0625 | 6.7422 | 4.6426 | 7.1622 | 7.115% | 14.23% | 21.34% | 5.6073 | 6.1975
Sn, ppm | 4.2747 | 0.23485 | 3.805 | 4.7444 | 3.5701 | 4.9793 | 5.494% | 10.99% | 16.48% | 4.061 | 4.4884
Sr, ppm | 276.2316 | 10.0086 | 256.2144 | 296.2489 | 246.2058 | 306.2575 | 3.623% | 7.247% | 10.87% | 262.4201 | 290.0432
Ta, ppm | 0.92553 | 0.063181 | 0.79917 | 1.0519 | 0.73599 | 1.1151 | 6.826% | 13.65% | 20.48% | 0.87926 | 0.97181
Tb, ppm | 0.66747 | 0.029097 | 0.60927 | 0.72566 | 0.58018 | 0.75476 | 4.359% | 8.719% | 13.08% | 0.63409 | 0.70084
Te, ppm | 0.21185 | 0.031207 | 0.14944 | 0.27427 | 0.11823 | 0.30547 | 14.73% | 29.46% | 44.19% | 0.20126 | 0.22244
Th, ppm | 11.9562 | 0.97789 | 10.0004 | 13.9119 | 9.0225 | 14.8898 | 8.179% | 16.36% | 24.54% | 11.3583 | 12.554
Ti, wt.% | 0.31876 | 0.010618 | 0.29752 | 0.33999 | 0.2869 | 0.35061 | 3.331% | 6.662% | 9.993% | 0.30282 | 0.33469
Tl, ppm | 0.81087 | 0.037169 | 0.73653 | 0.88521 | 0.69936 | 0.92238 | 4.584% | 9.168% | 13.75% | 0.77032 | 0.85141
Tm, ppm | 0.17588 | 0.014944 | 0.14599 | 0.20577 | 0.13105 | 0.22071 | 8.497% | 16.99% | 25.49% | 0.16709 | 0.18467
U, ppm | 3.5015 | 0.33892 | 2.8237 | 4.1794 | 2.4848 | 4.5183 | 9.679% | 19.36% | 29.04% | 3.3265 | 3.6766
V, ppm | 65.4856 | 2.6583 | 60.169 | 70.8021 | 57.5108 | 73.4604 | 4.059% | 8.119% | 12.18% | 62.2113 | 68.7599
W, ppm | 10.3101 | 0.77412 | 8.7619 | 11.8583 | 7.9878 | 12.6325 | 7.508% | 15.02% | 22.52% | 9.7946 | 10.8256
Y, ppm | 14.2861 | 0.72765 | 12.8308 | 15.7414 | 12.1031 | 16.469 | 5.093% | 10.19% | 15.28% | 13.5718 | 15.0004
Yb, ppm | 1.1717 | 0.069426 | 1.0329 | 1.3106 | 0.96345 | 1.38 | 5.925% | 11.85% | 17.78% | 1.1131 | 1.2303
Zn, ppm | 111.906 | 4.8477 | 102.2105 | 121.6014 | 97.3628 | 126.4491 | 4.332% | 8.664% | 13% | 106.3107 | 117.5012
Zr, ppm | 71.0075 | 2.8685 | 65.2705 | 76.7444 | 62.4021 | 79.6129 | 4.04% | 8.079% | 12.12% | 67.4571 | 74.5579
Aqua Regia Digestion
Ag, ppm | 1.3139 | 0.079773 | 1.1544 | 1.4735 | 1.0746 | 1.5532 | 6.071% | 12.14% | 18.21% | 1.2482 | 1.3796
Al, wt.% | 1.6783 | 0.064353 | 1.5496 | 1.807 | 1.4853 | 1.8714 | 3.834% | 7.669% | 11.5% | 1.5944 | 1.7623
As, ppm | 12.9541 | 0.61278 | 11.7285 | 14.1796 | 11.1157 | 14.7924 | 4.73% | 9.461% | 14.19% | 12.3064 | 13.6018
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 425.9311 | 21.7769 | 382.3772 | 469.4849 | 360.6003 | 491.2619 | 5.113% | 10.23% | 15.34% | 404.6345 | 447.2276
Be, ppm | 1.2125 | 0.044959 | 1.1225 | 1.3024 | 1.0776 | 1.3473 | 3.708% | 7.416% | 11.12% | 1.1518 | 1.2731
Bi, ppm | 0.95816 | 0.10011 | 0.75793 | 1.1584 | 0.65782 | 1.2585 | 10.45% | 20.9% | 31.35% | 0.91025 | 1.0061
Ca, wt.% | 0.69545 | 0.028799 | 0.63785 | 0.75305 | 0.60906 | 0.78185 | 4.141% | 8.282% | 12.42% | 0.66068 | 0.73022
Cd, ppm | 0.31946 | 0.036309 | 0.24685 | 0.39208 | 0.21054 | 0.42839 | 11.37% | 22.73% | 34.1% | 0.30349 | 0.33544
Ce, ppm | 29.4208 | 1.7326 | 25.9556 | 32.8861 | 24.2229 | 34.6188 | 5.889% | 11.78% | 17.67% | 27.9498 | 30.8919
Co, ppm | 7.0951 | 0.34184 | 6.4114 | 7.7787 | 6.0695 | 8.1206 | 4.818% | 9.636% | 14.45% | 6.7403 | 7.4498
Cr, ppm | 45.4673 | 2.1336 | 41.2001 | 49.7346 | 39.0664 | 51.8682 | 4.693% | 9.385% | 14.08% | 43.194 | 47.7407
Cs, ppm | 6.7809 | 0.59061 | 5.5997 | 7.9622 | 5.0091 | 8.5528 | 8.71% | 17.42% | 26.13% | 6.4419 | 7.12
Cu, wt.% | 0.44268 | 0.019792 | 0.4031 | 0.48227 | 0.38331 | 0.50206 | 4.471% | 8.942% | 13.41% | 0.42055 | 0.46482
Fe, wt.% | 2.6578 | 0.073215 | 2.5114 | 2.8042 | 2.4382 | 2.8775 | 2.755% | 5.509% | 8.264% | 2.5249 | 2.7907
Ga, ppm | 8.027 | 0.22383 | 7.5793 | 8.4746 | 7.3555 | 8.6985 | 2.788% | 5.577% | 8.365% | 7.6256 | 8.4283
Ge, ppm | 0.11133 | 0.027638 | 0.056058 | 0.16661 | 0.02842 | 0.19425 | 24.82% | 49.65% | 74.47% | 0.10577 | 0.1169
Hf, ppm | 0.30977 | 0.023952 | 0.26186 | 0.35767 | 0.23791 | 0.38162 | 7.732% | 15.46% | 23.2% | 0.29428 | 0.32526
In, ppm | 0.072341 | 0.0038701 | 0.064601 | 0.080081 | 0.060731 | 0.083951 | 5.35% | 10.7% | 16.05% | 0.068724 | 0.075958
K, wt.% | 0.85347 | 0.047637 | 0.7582 | 0.94875 | 0.71056 | 0.99638 | 5.582% | 11.16% | 16.74% | 0.8108 | 0.89615
La, ppm | 13.775 | 0.83235 | 12.1103 | 15.4397 | 11.2779 | 16.272 | 6.042% | 12.08% | 18.13% | 13.0862 | 14.4637
Li, ppm | 33.4048 | 1.291 | 30.8228 | 35.9869 | 29.5318 | 37.2779 | 3.865% | 7.73% | 11.59% | 31.7346 | 35.0751
Mg, wt.% | 0.64469 | 0.019894 | 0.6049 | 0.68447 | 0.585 | 0.70437 | 3.086% | 6.172% | 9.257% | 0.61245 | 0.67692
Mn, wt.% | 0.028561 | 0.0008491 | 0.026863 | 0.030259 | 0.026014 | 0.031108 | 2.973% | 5.946% | 8.919% | 0.027133 | 0.029989
Mo, ppm | 83.9642 | 2.4033 | 79.1576 | 88.7708 | 76.7543 | 91.1741 | 2.862% | 5.725% | 8.587% | 79.766 | 88.1624
Na, wt.% | 0.14197 | 0.01477 | 0.11243 | 0.17151 | 0.09766 | 0.18628 | 10.4% | 20.81% | 31.21% | 0.13487 | 0.14907
Nb, ppm | 0.75864 | 0.15118 | 0.45629 | 1.061 | 0.30511 | 1.2122 | 19.93% | 39.85% | 59.78% | 0.72071 | 0.79657
Ni, ppm | 15.5003 | 0.69051 | 14.1192 | 16.8813 | 13.4287 | 17.5718 | 4.455% | 8.91% | 13.36% | 14.7252 | 16.2753
P, wt.% | 0.065458 | 0.0032761 | 0.058905 | 0.07201 | 0.055629 | 0.075286 | 5.005% | 10.01% | 15.01% | 0.062185 | 0.068731
Pb, ppm | 39.5071 | 2.6144 | 34.2783 | 44.7358 | 31.6639 | 47.3502 | 6.618% | 13.24% | 19.85% | 37.5317 | 41.4824
Rb, ppm | 79.069 | 5.4165 | 68.2361 | 89.9019 | 62.8196 | 95.3184 | 6.85% | 13.7% | 20.55% | 75.1155 | 83.0224
Re, ppm | 0.034553 | 0.0027607 | 0.029032 | 0.040075 | 0.026271 | 0.042836 | 7.99% | 15.98% | 23.97% | 0.032826 | 0.036281
S, wt.% | 0.6144 | 0.024492 | 0.56542 | 0.66339 | 0.54093 | 0.68788 | 3.986% | 7.973% | 11.96% | 0.58368 | 0.64512
Sb, ppm | 0.86527 | 0.13436 | 0.59656 | 1.134 | 0.4622 | 1.2683 | 15.53% | 31.06% | 46.58% | 0.82201 | 0.90854
Sc, ppm | 6.6782 | 0.34575 | 5.9867 | 7.3697 | 5.641 | 7.7154 | 5.177% | 10.35% | 15.53% | 6.3443 | 7.0121
Se, ppm | 4.8883 | 0.32875 | 4.2308 | 5.5458 | 3.902 | 5.8746 | 6.725% | 13.45% | 20.18% | 4.6439 | 5.1327
Sn, ppm | 2.9487 | 0.1078 | 2.7331 | 3.1643 | 2.6253 | 3.2721 | 3.656% | 7.312% | 10.97% | 2.8013 | 3.0961
Sr, ppm | 94.1065 | 3.7528 | 86.6008 | 101.6122 | 82.8479 | 105.365 | 3.988% | 7.976% | 11.96% | 89.4012 | 98.8118
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.45434 | 0.025695 | 0.40295 | 0.50573 | 0.37726 | 0.53143 | 5.655% | 11.31% | 16.97% | 0.43163 | 0.47706
Te, ppm | 0.19949 | 0.020334 | 0.15882 | 0.24015 | 0.13848 | 0.26049 | 10.19% | 20.39% | 30.58% | 0.18951 | 0.20946
Th, ppm | 5.665 | 0.51104 | 4.643 | 6.6871 | 4.1319 | 7.1981 | 9.021% | 18.04% | 27.06% | 5.3818 | 5.9483
Ti, wt.% | 0.20437 | 0.0083466 | 0.18767 | 0.22106 | 0.17933 | 0.22941 | 4.084% | 8.168% | 12.25% | 0.19415 | 0.21459
Tl, ppm | 0.51169 | 0.029304 | 0.45308 | 0.5703 | 0.42378 | 0.5996 | 5.727% | 11.45% | 17.18% | 0.48611 | 0.53728
U, ppm | 2.8226 | 0.19131 | 2.44 | 3.2053 | 2.2487 | 3.3966 | 6.778% | 13.56% | 20.33% | 2.6815 | 2.9638
V, ppm | 54.3956 | 1.6029 | 51.1897 | 57.6014 | 49.5868 | 59.2044 | 2.947% | 5.894% | 8.84% | 51.6758 | 57.1153
W, ppm | 6.4245 | 0.63675 | 5.151 | 7.6981 | 4.5143 | 8.3348 | 9.911% | 19.82% | 29.73% | 6.1033 | 6.7458
Y, ppm | 9.6349 | 0.44487 | 8.7452 | 10.5246 | 8.3003 | 10.9695 | 4.617% | 9.234% | 13.85% | 9.1532 | 10.1166
Yb, ppm | 0.70662 | 0.021647 | 0.66333 | 0.74992 | 0.64168 | 0.77156 | 3.063% | 6.127% | 9.19% | 0.67129 | 0.74196
Zn, ppm | 105.7684 | 4.7334 | 96.3015 | 115.2353 | 91.568 | 119.9687 | 4.475% | 8.951% | 13.43% | 100.48 | 111.0568
Zr, ppm | 8.6034 | 0.80827 | 6.9868 | 10.2199 | 6.1786 | 11.0282 | 9.395% | 18.79% | 28.18% | 8.1732 | 9.0335
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 506b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AL-5S*AAS | 5% sulphuric acid laech with atomic absorption spectroscopy
AL-5S*OES | 5% sulphuric acid leach with inductively coupled plasma optical emission spectroscopy
AL-NaCN*AAS | sodium cyanide leach with atomic absorption spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
GASPYC | gas expansion pycnometry
INAA | instrumental neutron activation analysis
IRC | infrared combustion
IRC-x*Calc | calculation by difference from IRC total value
LOI*TGA | loss on ignition by thermogravimetric analysis
Na2CO3*GRAV | sodium carbonate solution with Gravimetric finish
Na2CO3*IRC | sodium carbonate leach with infrared combustion
NaCN*AAS | sodium cyanide with AAS finish
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
SQL-Tot*AAS | sequential leach analysis for total value method calculation with AAS finish
SQL-Tot*Calc | total value method calculated from sequential leach analysis
Text Values:
- | Not Analysed (Lab 2.04)
-- | Element Not Determined (Lab 2.14)
N.A. | Not Applicable (Lab 2.16)
N.A. | Not Applicable (Lab 2.24)
N.A. | Not Applicable (Lab 2.27)
NA | Not Applicable (Lab 2.02)
NA | Not Applicable (Lab 2.23)
NA | Unable to report due to QC failure (Lab 2.26)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 506b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek Minerals Ltd, Tarkwa, Western Region, Ghana
Labwest Minerals Analysis, Perth, WA, Australia
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS, Randfontein, Gauteng, South Africa
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 506b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
1 | 1 | ACTAn | 1 | 8 | 7 | 270729 | 1.04 | 0.332 | 0.34409
1 | 1 | ACTAn | 1 | 11 | 11 | 270730 | 1.01 | 0.344 | 0.34632
1 | 1 | ACTAn | 1 | 1 | 1 | 270731 | 1.02 | 0.345 | 0.34651
1 | 1 | ACTAn | 1 | 17 | 6 | 270732 | 1.06 | 0.34 | 0.34558
1 | 1 | ACTAn | 1 | 15 | 4 | 270733 | 1.02 | 0.353 | 0.34799
1 | 1 | ACTAn | 1 | 20 | 8 | 270734 | 1.06 | 0.334 | 0.34446
1 | 1 | ACTAn | 1 | 5 | 5 | 270735 | 1 | 0.338 | 0.34521
1 | 1 | ACTAn | 1 | 10 | 9 | 270736 | 1.03 | 0.34 | 0.34558
1 | 1 | ACTAn | 1 | 19 | 8 | 270737 | 1.06 | 0.346 | 0.34669
1 | 1 | ACTAn | 1 | 9 | 9 | 270738 | 1.07 | 0.34 | 0.34558
1 | 1 | ACTAn | 1 | 6 | 5 | 270739 | 1.03 | 0.333 | 0.34428
1 | 1 | ACTAn | 1 | 16 | 4 | 270740 | 1.03 | 0.349 | 0.34725
1 | 1 | ACTAn | 1 | 7 | 7 | 270741 | 1.08 | 0.347 | 0.34688
1 | 1 | ACTAn | 1 | 14 | 2 | 270742 | 1.02 | 0.361 | 0.34948
1 | 1 | ACTAn | 1 | 13 | 2 | 270743 | 1 | 0.356 | 0.34855
1 | 1 | ACTAn | 1 | 12 | 11 | 270744 | 1.02 | 0.349 | 0.34725
1 | 1 | ACTAn | 1 | 2 | 1 | 270745 | 1.02 | 0.367 | 0.35059
1 | 1 | ACTAn | 1 | 3 | 3 | 270746 | 1.06 | 0.357 | 0.34874
1 | 1 | ACTAn | 1 | 18 | 6 | 270747 | 1.07 | 0.351 | 0.34762
1 | 1 | ACTAn | 1 | 4 | 3 | 270748 | 1 | 0.355 | 0.34836
Mean | [6 empty] | 1.035 | 0.34685 | 0.34685
Median | [7 empty] | 0.3465 | 0.34678
SD | [7 empty] | 0.0095492 | 0.0017737
RSD | [7 empty] | 2.753% | 0.5114%
RSD Upscale Factor | [7 empty] | 0.18574
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 506b (Certified Value 0.351 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.23 | 2.24 | 2.26 | 2.27 | ppm
|  |  | INAA | FA*MS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | 3
|  |  | 1.0g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 50g | 30g | 30g | 30g | 3
| 1 | 1 | 0.345 | 0.35 | 0.353 | 0.35304 | 0.372 | 0.321 | 0.362 | 0.35 | 0.336 | 0.345 | 0.353 | 0.348 | 0.36 | 0.323 | 0.351 | 0.34984 | 0.358 | 0.352 | 0.31096 | 0.35 | 0.354 | 0.375 | 0.357 | 0.34 | 1
| 1 | 2 | 0.367 | 0.338 | 0.347 | 0.35563 | 0.358 | 0.34 | 0.361 | 0.351 | 0.34 | 0.345 | 0.356 | 0.344 | 0.354 | 0.346 | 0.349 | 0.34911 | 0.36 | 0.346 | 0.299 | 0.36 | 0.357 | 0.36 | 0.355 | 0.344 | 0
| 1 | 3 | 0.357 | 0.347 | 0.348 | 0.34983 | 0.347 | 0.31467 | 0.359 | 0.348 | 0.347 | 0.341 | 0.363 | 0.346 | 0.357 | 0.324 | 0.351 | 0.34864 | 0.358 | 0.356 | 0.31188 | 0.36 | 0.352 | 0.365 | 0.364 | 0.348 | 16
| 1 | 4 | 0.355 | 0.346 | 0.34 | 0.34608 | 0.338 | 0.322 | 0.359 | 0.344 | 0.35 | 0.34 | 0.349 | 0.343 | 0.349 | 0.328 | 0.345 | 0.34973 | 0.36 | 0.357 | 0.30212 | 0.36 | 0.356 | 0.365 | 0.358 | 0.344 | 0.35082
| 1 | 5 | 0.338 | 0.33 | 0.347 | 0.34761 | 0.356 | 0.321 | 0.364 | 0.34 | 0.336 | 0.344 | 0.347 | 0.349 | 0.355 | 0.344 | 0.353 | 0.34934 | 0.354 | 0.356 | 0.30544 | 0.35 | 0.351 | 0.364 | 0.351 | 0.341 | 7
| 1 | 6 | 0.333 | 0.348 | 0.346 | 0.35491 | 0.349 | 0.34167 | 0.363 | 0.349 | 0.336 | 0.326 | 0.355 | 0.349 | 0.352 | 0.327 | 0.348 | 0.35 | 0.36 | 0.352 | 0.30084 | 0.36 | 0.355 | 0.388 | 0.358 | 0.339
|  | 7 | 00: 0.347
|  | 8 | 00: 0.332
|  | 9 | 00: 0.34
|  | 10 | 00: 0.34
|  | 11 | 00: 0.344
|  | 12 | 00: 0.349
|  | 13 | 00: 0.356
|  | 14 | 00: 0.361
|  | 15 | 00: 0.353
|  | 16 | 00: 0.349
|  | 17 | 00: 0.34
|  | 18 | 00: 0.351
|  | 19 | 00: 0.346
|  | 20 | 00: 0.334
| Mean |  | 0.34685 | 0.34317 | 0.34683 | 0.35118 | 0.35333 | 0.32672 | 0.36133 | 0.347 | 0.34083 | 0.34017 | 0.35383 | 0.3465 | 0.3545 | 0.332 | 0.3495 | 0.34944 | 0.35833 | 0.35317 | 0.30504 | 0.35667 | 0.35417 | 0.3695 | 0.35717 | 0.34267
| Median |  | 0.3465 | 0.3465 | 0.347 | 0.35143 | 0.3525 | 0.3215 | 0.3615 | 0.3485 | 0.338 | 0.3425 | 0.354 | 0.347 | 0.3545 | 0.3275 | 0.35 | 0.34954 | 0.359 | 0.354 | 0.30378 | 0.36 | 0.3545 | 0.365 | 0.3575 | 0.3425
| Std Dev. |  | 0.0095492 | 0.0076529 | 0.0041673 | 0.0039438 | 0.011587 | 0.01125 | 0.0020656 | 0.0041952 | 0.0062102 | 0.0072503 | 0.0056716 | 0.0025884 | 0.0038341 | 0.010257 | 0.0028107 | 0.00051254 | 0.0023381 | 0.0041191 | 0.0053785 | 0.005164 | 0.0023166 | 0.01033 | 0.0042622 | 0.0033267
| Rel.Std.Dev. |  | 2.753% | 2.23% | 1.202% | 1.123% | 3.279% | 3.443% | 0.5717% | 1.209% | 1.822% | 2.131% | 1.603% | 0.747% | 1.082% | 3.089% | 0.8042% | 0.1467% | 0.6525% | 1.166% | 1.763% | 1.448% | 0.6541% | 2.796% | 1.193% | 0.9708%
| PDM3 |  | -1.133% | -2.183% | -1.138% | 0.1016% | 0.7152% | -6.87% | 2.996% | -1.09% | -2.848% | -3.038% | 0.8577% | -1.233% | 1.048% | -5.366% | -0.3774% | -0.3935% | 2.14% | 0.6677% | -13.05% | 1.665% | 0.9528% | 5.323% | 1.808% | -2.325%
| Z-Score (Absolute) |  | NA | 0.67 | 0.28 | 0.18 | 0.41 | 2.43 | 1.26 | 0.27 | 0.92 | 0.99 | 0.46 | 0.32 | 0.53 | 1.86 | 0 | 0.01 | 0.94 | 0.39 | 4.73 | 0.76 | 0.5 | 2.13 | 0.82 | 0.73
| Analytical results for Pd in OREAS 506b (Indicative Value 8.33 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 00: Lab; 2.01: Lab | 1
| No | No | 00: 2.04; 2.01: 2.12 | ppb
|  |  | 00: FA*OES; 2.01: FA*OES | 2
|  |  | 00: 30g; 2.01: 30g | 2
| 1 | 1 | 00: 6; 2.01: 10 | 1
| 1 | 2 | 00: 7; 2.01: 10 | 1
| 1 | 3 | 00: 4; 2.01: 10 | 16
| 1 | 4 | 00: 7; 2.01: 10 | 8.3333
| 1 | 5 | 00: 9; 2.01: 10 | 7
| 1 | 6 | 00: 7; 2.01: 10
| Mean |  | 00: 6.6667; 2.01: 10
| Median |  | 00: 7; 2.01: 10
| Std Dev. |  | 00: 1.633; 2.01: 0
| Rel.Std.Dev. |  | 00: 24.49%; 2.01: 0%
| PDM3 |  | 00: -20%; 2.01: 20%
| Z-Score (Absolute) |  | 00: 0.67; 2.01: 0.67
| Analytical results for Pt in OREAS 506b (Indicative Value < 10 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 00: Lab; 2.01: Lab | 1
| No | No | 00: 2.04; 2.01: 2.12 | ppb
|  |  | 00: FA*OES; 2.01: FA*OES | 1
|  |  | 00: 30g; 2.01: 30g | 1
| 1 | 1 | 00: < 3; 2.01: < 10 | 1
| 1 | 2 | 00: < 3; 2.01: < 10 | 1
| 1 | 3 | 00: < 3; 2.01: < 10 | 16
| 1 | 4 | 00: < 3; 2.01: < 10 | < 10
| 1 | 5 | 00: 3; 2.01: < 10 | 7
| 1 | 6 | 00: < 3; 2.01: < 10
| Mean |  | 00: 3
| Median |  | 00: 3
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 00: 0.67; 2.01: 0.67
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 506b (Certified Value 0.348 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.18 | 2.20 | 2.22 | 2.23 | 2.25 | 2.26 | 2.27 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*OES | AR*AAS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 3
|  |  | 1.0g | 25g | 30g | 25g | 15g | 50g | 10g | 25g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 25g | 50g | 15g | 30g | 15g | 30g | 3
| 1 | 1 | 0.345 | 0.363 | 0.34 | 0.35315 | 0.3564 | 0.34 | 0.36 | 0.369 | 0.339 | 0.352 | 0.337 | 0.396 | 0.32 | 0.343 | 0.3 | 0.356 | 0.342 | 0.356 | 0.328 | 0.34 | 0.3644 | 1
| 1 | 2 | 0.367 | 0.355 | 0.34 | 0.33116 | 0.3473 | 0.32 | 0.36 | 0.382 | 0.337 | 0.347 | 0.345 | 0.368 | 0.322 | 0.341 | 0.307 | 0.332 | 0.338 | 0.351 | 0.327 | 0.35 | 0.3831 | 0
| 1 | 3 | 0.357 | 0.354 | 0.34 | 0.33689 | 0.3769 | 0.32 | 0.35 | 0.358 | 0.343 | 0.347 | 0.347 | 0.358 | 0.323 | 0.349 | 0.31 | 0.363 | 0.341 | 0.358 | 0.337 | 0.36 | 0.3732 | 16
| 1 | 4 | 0.355 | 0.364 | 0.34 | 0.34626 | 0.381 | 0.33 | 0.36 | 0.354 | 0.355 | 0.364 | 0.333 | 0.365 | 0.323 | 0.342 | 0.3 | 0.353 | 0.34 | 0.3596 | 0.355 | 0.35 | 0.3691 | 0.34831
| 1 | 5 | 0.338 | 0.352 | 0.34 | 0.31546 | 0.3745 | 0.32 | 0.36 | 0.339 | 0.357 | 0.335 | 0.331 | 0.364 | 0.317 | 0.35 | 0.3053 | 0.353 | 0.337 | 0.3583 | 0.346 | 0.36 | 0.3663 | 9
| 1 | 6 | 0.333 | 0.363 | 0.34 | 0.32332 | 0.3508 | 0.33 | 0.36 | 0.375 | 0.351 | 0.35 | 0.342 | 0.357 | 0.317 | 0.344 | 0.304 | 0.354 | 0.34 | 0.3614 | 0.343 | 0.35 | 0.3638
|  | 7 | 00: 0.347
|  | 8 | 00: 0.332
|  | 9 | 00: 0.34
|  | 10 | 00: 0.34
|  | 11 | 00: 0.344
|  | 12 | 00: 0.349
|  | 13 | 00: 0.356
|  | 14 | 00: 0.361
|  | 15 | 00: 0.353
|  | 16 | 00: 0.349
|  | 17 | 00: 0.34
|  | 18 | 00: 0.351
|  | 19 | 00: 0.346
|  | 20 | 00: 0.334
| Mean |  | 0.34685 | 0.3585 | 0.34 | 0.33437 | 0.36448 | 0.32667 | 0.35833 | 0.36283 | 0.347 | 0.34917 | 0.33917 | 0.368 | 0.32033 | 0.34483 | 0.30438 | 0.35183 | 0.33967 | 0.35738 | 0.33933 | 0.35167 | 0.36998
| Median |  | 0.3465 | 0.359 | 0.34 | 0.33403 | 0.36545 | 0.325 | 0.36 | 0.3635 | 0.347 | 0.3485 | 0.3395 | 0.3645 | 0.321 | 0.3435 | 0.30465 | 0.3535 | 0.34 | 0.35815 | 0.34 | 0.35 | 0.3677
| Std Dev. |  | 0.0095492 | 0.0053944 | 0 | 0.014073 | 0.014664 | 0.008165 | 0.0040825 | 0.015639 | 0.0084853 | 0.0093684 | 0.0065243 | 0.014353 | 0.0028048 | 0.0037639 | 0.0039448 | 0.01042 | 0.0018619 | 0.0036036 | 0.010857 | 0.0075277 | 0.0072959
| Rel.Std.Dev. |  | 2.753% | 1.505% | 0% | 4.209% | 4.023% | 2.499% | 1.139% | 4.31% | 2.445% | 2.683% | 1.924% | 3.9% | 0.8756% | 1.092% | 1.296% | 2.961% | 0.5482% | 1.008% | 3.199% | 2.141% | 1.972%
| PDM3 |  | -0.4202% | 2.924% | -2.387% | -4.002% | 4.642% | -6.215% | 2.877% | 4.169% | -0.3771% | 0.2449% | -2.626% | 5.652% | -8.033% | -0.9992% | -12.61% | 1.011% | -2.483% | 2.604% | -2.578% | 0.9627% | 6.221%
| Z-Score (Absolute) |  | NA | 0.72 | 0.56 | 0.95 | 1.13 | 1.48 | 0.71 | 1.02 | 0.07 | 0.07 | 0.62 | 1.37 | 1.91 | 0.22 | 3.01 | 0.26 | 0.58 | 0.64 | 0.6 | 0.25 | 1.51
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## SQL
| Analytical results for Cu-Res(4A) in OREAS 506b (Indicative Value 0.086 wt.%) | [8 empty] | Indicative
Cu-Res(4A) | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.14 | 2.16 | 2.22 |  |  | wt.%
|  |  | 4A*AAS | SQL-Tot*Calc | SQL-Tot*AAS | 4A*AAS | 4A*AAS |  |  | 3
3
| 1 | 1 | NA | - | -- | N.A. | 0.0862 |  |  | 1
| 1 | 2 | NA | - | -- | N.A. | 0.087 |  |  | 3
| 1 | 3 | NA | - | -- | N.A. | 0.0848 |  |  | 16
| 1 | 4 | NA | - | -- | N.A. | 0.0862 |  |  | 0.085733
| 1 | 5 | NA | - | -- | N.A. | 0.0858 |  |  | 9
| 1 | 6 | NA | - | -- | N.A. | 0.0844
| Mean | [5 empty] | 0.085733
| Median | [5 empty] | 0.086
| Std Dev. | [5 empty] | 0.00096885
| Rel.Std.Dev. | [5 empty] | 1.13%
| PDM3 | [5 empty] | 2.22e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Cu-Sol(CN) in OREAS 506b (Indicative Value 0.308 wt.%) | [8 empty] | Indicative
Cu-Sol(CN) | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.14 | 2.16 | 2.22 |  |  | wt.%
|  |  | NaCN*AAS | SQL-Tot*Calc | SQL-Tot*AAS | NaCN*AAS | AL-NaCN*AAS |  |  | 3
3
| 1 | 1 | NA | - | -- | N.A. | 0.3064 |  |  | 1
| 1 | 2 | NA | - | -- | N.A. | 0.3088 |  |  | 3
| 1 | 3 | NA | - | -- | N.A. | 0.3066 |  |  | 16
| 1 | 4 | NA | - | -- | N.A. | 0.3138 |  |  | 0.30773
| 1 | 5 | NA | - | -- | N.A. | 0.3072 |  |  | 9
| 1 | 6 | NA | - | -- | N.A. | 0.3036
| Mean | [5 empty] | 0.30773
| Median | [5 empty] | 0.3069
| Std Dev. | [5 empty] | 0.0034168
| Rel.Std.Dev. | [5 empty] | 1.11%
| PDM3 | [5 empty] | 1.11e-13%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Cu-Sol(H2SO4) in OREAS 506b (Indicative Value 0.065 wt.%) | [8 empty] | Indicative
Cu-Sol(H2SO4) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.14 | 2.16 | 2.18 | 2.22 | 2.27 | wt.%
|  |  | AL-5S*AAS | SQL-Tot*Calc | SQL-Tot*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | 3
3
| 1 | 1 | NA | - | -- | N.A. | 0.091515 | 0.0434 | N.A. | 1
| 1 | 2 | NA | - | -- | N.A. | 0.09 | 0.042 | N.A. | 3
| 1 | 3 | NA | - | -- | N.A. | 0.09 | 0.0434 | N.A. | 16
| 1 | 4 | NA | - | -- | N.A. | 0.076518 | 0.0498 | N.A. | 0.065319
| 1 | 5 | NA | - | -- | N.A. | 0.08 | 0.044 | N.A. | 9
| 1 | 6 | NA | - | -- | N.A. | 0.09 | 0.0432 | N.A.
| Mean | [5 empty] | 0.086339 | 0.0443
| Median | [5 empty] | 0.09 | 0.0434
| Std Dev. | [5 empty] | 0.0063818 | 0.0027734
| Rel.Std.Dev. | [5 empty] | 7.392% | 6.261%
| PDM3 | [5 empty] | 32.18% | -32.18%
| Z-Score (Absolute) |  | NA | NA | NA | NA | 0.67 | 0.67 | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Pycnometry
| Analytical results for SG in OREAS 506b (Indicative Value 2.69 Unity) | [9 empty] | Indicative
SG | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | Unity
|  |  | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | 2
2
| 1 | 1 | NA | - | 2.76 | -- | 2.5798 | 2.78 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 2.78 | -- | 2.6155 | 2.68 | N.A. | N.A. | 5
| 1 | 3 | NA | - | 2.77 | -- | 2.6398 | 2.59 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 2.77 | -- | 2.5916 | 2.68 | N.A. | N.A. | 2.6876
| 1 | 5 | NA | - | 2.77 | -- | 2.5707 | 2.68 | N.A. | N.A. | 11
| 1 | 6 | NA | - | 2.78 | -- | 2.6293 | 2.71 | N.A. | N.A.
| Mean |  |  |  | 2.7717 |  | 2.6045 | 2.6867
| Median |  |  |  | 2.77 |  | 2.6036 | 2.68
| Std Dev. |  |  |  | 0.0075277 |  | 0.027905 | 0.06121
| Rel.Std.Dev. |  |  |  | 0.2716% |  | 1.071% | 2.278%
| PDM3 |  |  |  | 3.128% |  | -3.093% | -0.03471%
| Z-Score (Absolute) |  | NA | NA | 0.7 | NA | 0.67 | 0 | NA | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 506b (Indicative Value 0.112 wt.%) | [8 empty] | Indicative
C | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 2.06 | 3.01 | [5 empty] | wt.%
|  |  | IRC | IRC | [5 empty] | 3
3
| 1 | 1 | 0.11 | 0.12 | [5 empty] | 1
| 1 | 2 | 0.12 | 0.1 | [5 empty] | 7
| 1 | 3 | 0.11 | [6 empty] | 16
| 1 | 4 | 0.12 | [6 empty] | 0.11167
| 1 | 5 | 0.11 | [6 empty] | 13
| 1 | 6 | 0.11
| Mean |  | 0.11333 | 0.11
| Median |  | 0.11 | 0.11
| Std Dev. |  | 0.005164 | 0.014142
| Rel.Std.Dev. |  | 4.556% | 12.86%
| PDM3 |  | 1.493% | -1.493%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 506b (Indicative Value 0.543 wt.%) | [8 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.24 | 2.27 | 3.01 | wt.%
|  |  | IRC | IRC-x*Calc | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | NA | - | 0.48 | -- | N.A. | N.A. | 0.61 | 1
| 1 | 2 | NA | - | 0.49 | -- | N.A. | N.A. | 0.59 | 7
| 1 | 3 | NA | - | 0.48 | -- | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 0.49 | -- | N.A. | N.A. |  | 0.5425
| 1 | 5 | NA | - | 0.49 | -- | N.A. | N.A. |  | 13
| 1 | 6 | NA | - | 0.48 | -- | N.A. | N.A.
| Mean |  |  |  | 0.485 |  |  |  | 0.6
| Median |  |  |  | 0.485 |  |  |  | 0.6
| Std Dev. |  |  |  | 0.0054772 |  |  |  | 0.014142
| Rel.Std.Dev. |  |  |  | 1.129% |  |  |  | 2.357%
| PDM3 |  |  |  | -10.6% |  |  |  | 10.6%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | NA | NA | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ALK
| Analytical results for S-(Sulphate) in OREAS 506b (Indicative Value 0.432 wt.%) | [7 empty] | Indicative
S-(Sulphate) | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.24 | 2.27 | wt.%
|  |  | Na2CO3*IRC | IRC-x*Calc | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*GRAV | 3
3
| 1 | 1 | NA | - | 0.44 | -- | N.A. | N.A. | 1
| 1 | 2 | NA | - | 0.43 | -- | N.A. | N.A. | 9
| 1 | 3 | NA | - | 0.43 | -- | N.A. | N.A. | 16
| 1 | 4 | NA | - | 0.42 | -- | N.A. | N.A. | 0.43167
| 1 | 5 | NA | - | 0.44 | -- | N.A. | N.A. | 15
| 1 | 6 | NA | - | 0.43 | -- | N.A. | N.A.
| Mean |  |  |  | 0.43167
| Median |  |  |  | 0.43
| Std Dev. |  |  |  | 0.0075277
| Rel.Std.Dev. |  |  |  | 1.744%
| PDM3 |  |  |  | -6.661e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for S-(Sulphide) in OREAS 506b (Indicative Value 0.053 wt.%) | [7 empty] | Indicative
S-(Sulphide) | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.04 | 2.06 | 2.14 | 2.24 | 2.27 |  | wt.%
|  |  | IRC-x*Calc | Na2CO3*IRC | Na2CO3*IRC | Na2CO3*IRC | IRC-x*Calc |  | 3
3
| 1 | 1 | - | 0.05 | -- | N.A. | N.A. |  | 1
| 1 | 2 | - | 0.06 | -- | N.A. | N.A. |  | 9
| 1 | 3 | - | 0.06 | -- | N.A. | N.A. |  | 16
| 1 | 4 | - | 0.05 | -- | N.A. | N.A. |  | 0.053333
| 1 | 5 | - | 0.05 | -- | N.A. | N.A. |  | 15
| 1 | 6 | - | 0.05 | -- | N.A. | N.A.
| Mean |  |  | 0.053333
| Median |  |  | 0.05
| Std Dev. |  |  | 0.005164
| Rel.Std.Dev. |  |  | 9.682%
| PDM3 |  |  | 4.441e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 506b (Indicative Value 1.23 ppm) | [9 empty] | Indicative
Ag | Round | Replicate | Lab | [7 empty] | 1
| No | No | 2.06 | [7 empty] | ppm
|  |  | PF*OES | [7 empty] | 2
2
| 1 | 1 | 1.22 | [7 empty] | 1
| 1 | 2 | 1.23 | [7 empty] | 11
| 1 | 3 | 1.23 | [7 empty] | 16
| 1 | 4 | 1.22 | [7 empty] | 1.225
| 1 | 5 | 1.22 | [7 empty] | 17
| 1 | 6 | 1.23
| Mean |  | 1.225
| Median |  | 1.225
| Std Dev. |  | 0.0054772
| Rel.Std.Dev. |  | 0.4471%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Al in OREAS 506b (Indicative Value 6.3 wt.%) | [9 empty] | Indicative
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  | 2
2
| 1 | 1 | NA | - | 4.59 | -- | 8.02 | N.A. | N.A. |  | 1
| 1 | 2 | NA | - | 4.56 | -- | 8.01 | N.A. | N.A. |  | 12
| 1 | 3 | NA | - | 4.56 | -- | 8.02 | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 4.58 | -- | 8.01 | N.A. | N.A. |  | 6.2975
| 1 | 5 | NA | - | 4.59 | -- | 8.01 | N.A. | N.A. |  | 18
| 1 | 6 | NA | - | 4.59 | -- | 8.03 | N.A. | N.A.
| Mean |  |  |  | 4.5783 |  | 8.0167
| Median |  |  |  | 4.585 |  | 8.015
| Std Dev. |  |  |  | 0.01472 |  | 0.008165
| Rel.Std.Dev. |  |  |  | 0.3215% |  | 0.1018%
| PDM3 |  |  |  | -27.3% |  | 27.3%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA | NA
| Analytical results for As in OREAS 506b (Indicative Value 19.2 ppm) | [9 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  | 1
1
| 1 | 1 | NA | - | 18.6 | -- | 19 | N.A. | N.A. |  | 1
| 1 | 2 | NA | - | 21.1 | -- | 19 | N.A. | N.A. |  | 13
| 1 | 3 | NA | - | 21.3 | -- | 19 | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 19.5 | -- | 19 | N.A. | N.A. |  | 19.2
| 1 | 5 | NA | - | 18.4 | -- | 19 | N.A. | N.A. |  | 19
| 1 | 6 | NA | - | 17.5 | -- | 19 | N.A. | N.A.
| Mean |  |  |  | 19.4 |  | 19
| Median |  |  |  | 19.05 |  | 19
| Std Dev. |  |  |  | 1.5336 |  | 0
| Rel.Std.Dev. |  |  |  | 7.905% |  | 0%
| PDM3 |  |  |  | 1.042% |  | -1.042%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA | NA
| Analytical results for Ba in OREAS 506b (Indicative Value 680 ppm) | [9 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | NA | - | 414.7 | -- | 947 | N.A. |  |  | 1
| 1 | 2 | NA | - | 409.9 | -- | 946 | N.A. |  |  | 14
| 1 | 3 | NA | - | 419.1 | -- | 946 | N.A. |  |  | 16
| 1 | 4 | NA | - | 410.5 | -- | 947 | N.A. |  |  | 679.7583
| 1 | 5 | NA | - | 408.6 | -- | 949 | N.A. |  |  | 20
| 1 | 6 | NA | - | 410.3 | -- | 949 | N.A.
| Mean |  |  |  | 412.1833 |  | 947.3333
| Median |  |  |  | 410.4 |  | 947
| Std Dev. |  |  |  | 3.9651 |  | 1.3663
| Rel.Std.Dev. |  |  |  | 0.962% |  | 0.1442%
| PDM3 |  |  |  | -39.36% |  | 39.36%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Be in OREAS 506b (Indicative Value < 2 ppm) | [9 empty] | Indicative
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 2
2
| 1 | 1 | NA | - | < 2 | -- | < 5 | N.A. |  |  | 1
| 1 | 2 | NA | - | < 2 | -- | < 5 | N.A. |  |  | 15
| 1 | 3 | NA | - | < 2 | -- | < 5 | N.A. |  |  | 16
| 1 | 4 | NA | - | < 2 | -- | < 5 | N.A. |  |  | < 2
| 1 | 5 | NA | - | < 2 | -- | < 5 | N.A. |  |  | 21
| 1 | 6 | NA | - | < 2 | -- | < 5 | N.A.
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Bi in OREAS 506b (Indicative Value < 5 ppm) | [9 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.06 | 2.14 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | [5 empty] | 2
2
| 1 | 1 | NA | < 5 | -- | [5 empty] | 1
| 1 | 2 | NA | < 5 | -- | [5 empty] | 16
| 1 | 3 | NA | < 5 | -- | [5 empty] | 16
| 1 | 4 | NA | < 5 | -- | [5 empty] | < 5
| 1 | 5 | NA | < 5 | -- | [5 empty] | 22
| 1 | 6 | NA | < 5 | --
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Ca in OREAS 506b (Indicative Value 1.66 wt.%) | [9 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
2
| 1 | 1 | NA | - | 1.48 | -- | 1.8577 | 1.66 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 1.48 | -- | 1.8582 | 1.67 | N.A. | N.A. | 17
| 1 | 3 | NA | - | 1.46 | -- | 1.8557 | 1.67 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 1.47 | -- | 1.8565 | 1.66 | N.A. | N.A. | 1.6639
| 1 | 5 | NA | - | 1.46 | -- | 1.8591 | 1.66 | N.A. | N.A. | 23
| 1 | 6 | NA | - | 1.46 | -- | 1.8528 | 1.68 | N.A. | N.A.
| Mean |  |  |  | 1.4683 |  | 1.8567 | 1.6667
| Median |  |  |  | 1.465 |  | 1.8571 | 1.665
| Std Dev. |  |  |  | 0.0098319 |  | 0.0022517 | 0.008165
| Rel.Std.Dev. |  |  |  | 0.6696% |  | 0.1213% | 0.4899%
| PDM3 |  |  |  | -11.75% |  | 11.59% | 0.1669%
| Z-Score (Absolute) |  | NA | NA | 0.7 | NA | 0.67 | 0 | NA | NA
| Analytical results for Cd in OREAS 506b (Indicative Value < 1 ppm) | [9 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.06 | 2.14 | 2.24 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES |  |  |  |  | 2
2
| 1 | 1 | - | < 1 | -- | N.A. |  |  |  |  | 1
| 1 | 2 | - | < 1 | -- | N.A. |  |  |  |  | 18
| 1 | 3 | - | < 1 | -- | N.A. |  |  |  |  | 16
| 1 | 4 | - | < 1 | -- | N.A. |  |  |  |  | < 1
| 1 | 5 | - | < 1 | -- | N.A. |  |  |  |  | 24
| 1 | 6 | - | < 1 | -- | N.A.
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Ce in OREAS 506b (Indicative Value 58 ppm) | [9 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.14 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | [5 empty] | 0
0
| 1 | 1 | - | 56.7 | -- | [5 empty] | 1
| 1 | 2 | - | 56.3 | -- | [5 empty] | 19
| 1 | 3 | - | 57.9 | -- | [5 empty] | 16
| 1 | 4 | - | 59.8 | -- | [5 empty] | 57.65
| 1 | 5 | - | 58.6 | -- | [5 empty] | 25
| 1 | 6 | - | 56.6 | --
| Mean |  |  | 57.65
| Median |  |  | 57.3
| Std Dev. |  |  | 1.3722
| Rel.Std.Dev. |  |  | 2.38%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Co in OREAS 506b (Indicative Value 10.5 ppm) | [9 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  | 1
1
| 1 | 1 | NA | - | 8.2 | -- | 12 | N.A. | N.A. |  | 1
| 1 | 2 | NA | - | 9.7 | -- | 12 | N.A. | N.A. |  | 20
| 1 | 3 | NA | - | 9.1 | -- | 12 | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 8.8 | -- | 12 | N.A. | N.A. |  | 10.5083
| 1 | 5 | NA | - | 9.5 | -- | 12 | N.A. | N.A. |  | 26
| 1 | 6 | NA | - | 8.8 | -- | 12 | N.A. | N.A.
| Mean |  |  |  | 9.0167 |  | 12
| Median |  |  |  | 8.95 |  | 12
| Std Dev. |  |  |  | 0.54191 |  | 0
| Rel.Std.Dev. |  |  |  | 6.01% |  | 0%
| PDM3 |  |  |  | -14.2% |  | 14.2%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA | NA
| Analytical results for Cr in OREAS 506b (Indicative Value 44.3 ppm) | [9 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.14 | 2.22 | 2.24 | 2.27 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES |  |  | 1
1
| 1 | 1 | NA | - | -- | 44 | N.A. | N.A. |  |  | 1
| 1 | 2 | NA | - | -- | 45 | N.A. | N.A. |  |  | 21
| 1 | 3 | NA | - | -- | 44 | N.A. | N.A. |  |  | 16
| 1 | 4 | NA | - | -- | 44 | N.A. | N.A. |  |  | 44.3333
| 1 | 5 | NA | - | -- | 45 | N.A. | N.A. |  |  | 27
| 1 | 6 | NA | - | -- | 44 | N.A. | N.A.
| Mean |  |  |  |  | 44.3333
| Median |  |  |  |  | 44
| Std Dev. |  |  |  |  | 0.5164
| Rel.Std.Dev. |  |  |  |  | 1.165%
| PDM3 |  |  |  |  | 8.882e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Cu in OREAS 506b (Indicative Value 0.437 wt.%) | [9 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | NA | - | 0.4136 | -- | 0.44448 | 0.4532 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 0.41256 | -- | 0.44035 | 0.4538 | N.A. | N.A. | 22
| 1 | 3 | NA | - | 0.41171 | -- | 0.44567 | 0.4539 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 0.41855 | -- | 0.44559 | 0.4555 | N.A. | N.A. | 0.43731
| 1 | 5 | NA | - | 0.41718 | -- | 0.44477 | 0.4541 | N.A. | N.A. | 28
| 1 | 6 | NA | - | 0.4115 | -- | 0.44114 | 0.4539 | N.A. | N.A.
| Mean |  |  |  | 0.41418 |  | 0.44367 | 0.45407
| Median |  |  |  | 0.41308 |  | 0.44462 | 0.4539
| Std Dev. |  |  |  | 0.0029787 |  | 0.0023217 | 0.00076594
| Rel.Std.Dev. |  |  |  | 0.7192% |  | 0.5233% | 0.1687%
| PDM3 |  |  |  | -5.288% |  | 1.455% | 3.833%
| Z-Score (Absolute) |  | NA | NA | 1.91 | NA | 0 | 0.67 | NA | NA
| Analytical results for Dy in OREAS 506b (Indicative Value 3.63 ppm) | [9 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 3.6 | -- | [6 empty] | 1
| 1 | 2 | 3.4 | -- | [6 empty] | 23
| 1 | 3 | 3.7 | -- | [6 empty] | 16
| 1 | 4 | 3.7 | -- | [6 empty] | 3.6333
| 1 | 5 | 3.8 | -- | [6 empty] | 29
| 1 | 6 | 3.6 | --
| Mean |  | 3.6333
| Median |  | 3.65
| Std Dev. |  | 0.13663
| Rel.Std.Dev. |  | 3.76%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Er in OREAS 506b (Indicative Value 1.72 ppm) | [9 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.5 | -- | [6 empty] | 1
| 1 | 2 | 1.6 | -- | [6 empty] | 24
| 1 | 3 | 1.8 | -- | [6 empty] | 16
| 1 | 4 | 1.7 | -- | [6 empty] | 1.7167
| 1 | 5 | 1.9 | -- | [6 empty] | 30
| 1 | 6 | 1.8 | --
| Mean |  | 1.7167
| Median |  | 1.75
| Std Dev. |  | 0.1472
| Rel.Std.Dev. |  | 8.575%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Eu in OREAS 506b (Indicative Value 1.28 ppm) | [9 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.3 | -- | [6 empty] | 1
| 1 | 2 | 1.2 | -- | [6 empty] | 25
| 1 | 3 | 1.3 | -- | [6 empty] | 16
| 1 | 4 | 1.3 | -- | [6 empty] | 1.2833
| 1 | 5 | 1.3 | -- | [6 empty] | 31
| 1 | 6 | 1.3 | --
| Mean |  | 1.2833
| Median |  | 1.3
| Std Dev. |  | 0.040825
| Rel.Std.Dev. |  | 3.181%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Fe in OREAS 506b (Indicative Value 2.85 wt.%) | [9 empty] | Indicative
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
2
| 1 | 1 | NA | - | 2.69 | -- | 3.0531 | 2.83 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 2.7 | -- | 3.0248 | 2.82 | N.A. | N.A. | 26
| 1 | 3 | NA | - | 2.68 | -- | 3.0109 | 2.81 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 2.69 | -- | 3.0636 | 2.81 | N.A. | N.A. | 2.8526
| 1 | 5 | NA | - | 2.69 | -- | 3.0885 | 2.83 | N.A. | N.A. | 32
| 1 | 6 | NA | - | 2.7 | -- | 3.036 | 2.82 | N.A. | N.A.
| Mean |  |  |  | 2.6917 |  | 3.0461 | 2.82
| Median |  |  |  | 2.69 |  | 3.0445 | 2.82
| Std Dev. |  |  |  | 0.0075277 |  | 0.028104 | 0.0089443
| Rel.Std.Dev. |  |  |  | 0.2797% |  | 0.9226% | 0.3172%
| PDM3 |  |  |  | -5.642% |  | 6.784% | -1.143%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 1.19 | 0 | NA | NA
| Analytical results for Ga in OREAS 506b (Indicative Value 20.4 ppm) | [9 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.06 | 2.14 | 2.22 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES |  |  |  |  | 1
1
| 1 | 1 | NA | 18.2 | -- | 23 |  |  |  |  | 1
| 1 | 2 | NA | 18.1 | -- | 22 |  |  |  |  | 27
| 1 | 3 | NA | 18.3 | -- | 23 |  |  |  |  | 16
| 1 | 4 | NA | 18.2 | -- | 23 |  |  |  |  | 20.425
| 1 | 5 | NA | 18.2 | -- | 23 |  |  |  |  | 33
| 1 | 6 | NA | 18.1 | -- | 22
| Mean |  |  | 18.1833 |  | 22.6667
| Median |  |  | 18.2 |  | 23
| Std Dev. |  |  | 0.075277 |  | 0.5164
| Rel.Std.Dev. |  |  | 0.414% |  | 2.278%
| PDM3 |  |  | -10.98% |  | 10.98%
| Z-Score (Absolute) |  | NA | 0.67 | NA | 0.67
| Analytical results for Gd in OREAS 506b (Indicative Value 5.17 ppm) | [9 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 5.2 | -- | [6 empty] | 1
| 1 | 2 | 5 | -- | [6 empty] | 28
| 1 | 3 | 4.8 | -- | [6 empty] | 16
| 1 | 4 | 5.3 | -- | [6 empty] | 5.1667
| 1 | 5 | 5.4 | -- | [6 empty] | 34
| 1 | 6 | 5.3 | --
| Mean |  | 5.1667
| Median |  | 5.25
| Std Dev. |  | 0.22509
| Rel.Std.Dev. |  | 4.357%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Hf in OREAS 506b (Indicative Value 2.75 ppm) | [9 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 2.5 | -- | [6 empty] | 1
| 1 | 2 | 2.7 | -- | [6 empty] | 29
| 1 | 3 | 2.8 | -- | [6 empty] | 16
| 1 | 4 | 2.6 | -- | [6 empty] | 2.75
| 1 | 5 | 3 | -- | [6 empty] | 35
| 1 | 6 | 2.9 | --
| Mean |  | 2.75
| Median |  | 2.75
| Std Dev. |  | 0.18708
| Rel.Std.Dev. |  | 6.803%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Hg in OREAS 506b (Indicative Value < 5 ppm) | [9 empty] | Indicative
Hg | Round | Replicate | Lab | [7 empty] | 1
| No | No | 2.06 | [7 empty] | ppm
|  |  | PF*OES | [7 empty] | 2
2
| 1 | 1 | < 5 | [7 empty] | 1
| 1 | 2 | < 5 | [7 empty] | 37
| 1 | 3 | < 5 | [7 empty] | 16
| 1 | 4 | < 5 | [7 empty] | < 5
| 1 | 5 | < 5 | [7 empty] | 17
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 506b (Indicative Value 0.63 ppm) | [9 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 0.6 | -- | [6 empty] | 1
| 1 | 2 | 0.6 | -- | [6 empty] | 12
| 1 | 3 | 0.7 | -- | [6 empty] | 16
| 1 | 4 | 0.6 | -- | [6 empty] | 0.63333
| 1 | 5 | 0.7 | -- | [6 empty] | 18
| 1 | 6 | 0.6 | --
| Mean |  | 0.63333
| Median |  | 0.6
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 8.154%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 506b (Indicative Value 3.02 wt.%) | [9 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 2
2
| 1 | 1 | NA | - | 1.79 | -- | 3.7491 | 3.53 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 1.77 | -- | 3.7463 | 3.55 | N.A. | N.A. | 13
| 1 | 3 | NA | - | 1.8 | -- | 3.7498 | 3.53 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 1.77 | -- | 3.7481 | 3.54 | N.A. | N.A. | 3.0231
| 1 | 5 | NA | - | 1.77 | -- | 3.7457 | 3.55 | N.A. | N.A. | 19
| 1 | 6 | NA | - | 1.78 | -- | 3.7465 | 3.55 | N.A. | N.A.
| Mean |  |  |  | 1.78 |  | 3.7476 | 3.5417
| Median |  |  |  | 1.775 |  | 3.7473 | 3.545
| Std Dev. |  |  |  | 0.012649 |  | 0.0016753 | 0.0098319
| Rel.Std.Dev. |  |  |  | 0.7106% |  | 0.0447% | 0.2776%
| PDM3 |  |  |  | -41.12% |  | 23.97% | 17.15%
| Z-Score (Absolute) |  | NA | NA | 5.77 | NA | 0.67 | 0 | NA | NA
| Analytical results for La in OREAS 506b (Indicative Value 30.3 ppm) | [9 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES |  |  |  | 1
1
| 1 | 1 | - | 27.6 | -- | 33 | N.A. |  |  |  | 1
| 1 | 2 | - | 27.7 | -- | 33 | N.A. |  |  |  | 14
| 1 | 3 | - | 26.2 | -- | 33 | N.A. |  |  |  | 16
| 1 | 4 | - | 29.5 | -- | 32 | N.A. |  |  |  | 30.2583
| 1 | 5 | - | 29.2 | -- | 32 | N.A. |  |  |  | 20
| 1 | 6 | - | 27.9 | -- | 32 | N.A.
| Mean |  |  | 28.0167 |  | 32.5
| Median |  |  | 27.8 |  | 32.5
| Std Dev. |  |  | 1.199 |  | 0.54772
| Rel.Std.Dev. |  |  | 4.28% |  | 1.685%
| PDM3 |  |  | -7.408% |  | 7.408%
| Z-Score (Absolute) |  | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Li in OREAS 506b (Indicative Value 46.3 ppm) | [9 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.14 | 2.22 | 2.24 | 2.27 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES |  |  | 1
1
| 1 | 1 | NA | - | -- | 47 | N.A. | N.A. |  |  | 1
| 1 | 2 | NA | - | -- | 46 | N.A. | N.A. |  |  | 15
| 1 | 3 | NA | - | -- | 46 | N.A. | N.A. |  |  | 16
| 1 | 4 | NA | - | -- | 46 | N.A. | N.A. |  |  | 46.3333
| 1 | 5 | NA | - | -- | 46 | N.A. | N.A. |  |  | 21
| 1 | 6 | NA | - | -- | 47 | N.A. | N.A.
| Mean |  |  |  |  | 46.3333
| Median |  |  |  |  | 46
| Std Dev. |  |  |  |  | 0.5164
| Rel.Std.Dev. |  |  |  |  | 1.115%
| PDM3 |  |  |  |  | 6.661e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Lu in OREAS 506b (Indicative Value < 0.5 ppm) | [9 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | < 0.5 | -- | [6 empty] | 1
| 1 | 2 | < 0.5 | -- | [6 empty] | 16
| 1 | 3 | < 0.5 | -- | [6 empty] | 16
| 1 | 4 | < 0.5 | -- | [6 empty] | < 0.5
| 1 | 5 | < 0.5 | -- | [6 empty] | 22
| 1 | 6 | < 0.5 | --
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Mg in OREAS 506b (Indicative Value 0.735 wt.%) | [9 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | NA | - | 0.65 | -- | 0.73439 | 0.74 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 0.66 | -- | 0.73484 | 0.74 | N.A. | N.A. | 17
| 1 | 3 | NA | - | 0.66 | -- | 0.73858 | 0.73 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 0.66 | -- | 0.73282 | 0.73 | N.A. | N.A. | 0.73462
| 1 | 5 | NA | - | 0.65 | -- | 0.7332 | 0.73 | N.A. | N.A. | 23
| 1 | 6 | NA | - | 0.66 | -- | 0.7316 | 0.74 | N.A. | N.A.
| Mean |  |  |  | 0.65667 |  | 0.73424 | 0.735
| Median |  |  |  | 0.66 |  | 0.73379 | 0.735
| Std Dev. |  |  |  | 0.005164 |  | 0.0024214 | 0.0054772
| Rel.Std.Dev. |  |  |  | 0.7864% |  | 0.3298% | 0.7452%
| PDM3 |  |  |  | -10.61% |  | -0.05198% | 0.05198%
| Z-Score (Absolute) |  | NA | NA | 68.49 | NA | 0 | 0.67 | NA | NA
| Analytical results for Mn in OREAS 506b (Indicative Value 0.036 wt.%) | [9 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 | 2.27 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | NA | - | 0.03718 | -- | 0.036748 | 0.0355 | N.A. | N.A. | 1
| 1 | 2 | NA | - | 0.03489 | -- | 0.036271 | 0.0356 | N.A. | N.A. | 18
| 1 | 3 | NA | - | 0.03587 | -- | 0.036775 | 0.0355 | N.A. | N.A. | 16
| 1 | 4 | NA | - | 0.03616 | -- | 0.036777 | 0.0356 | N.A. | N.A. | 0.035875
| 1 | 5 | NA | - | 0.03546 | -- | 0.03616 | 0.0355 | N.A. | N.A. | 24
| 1 | 6 | NA | - | 0.03411 | -- | 0.036041 | 0.0356 | N.A. | N.A.
| Mean |  |  |  | 0.035612 |  | 0.036462 | 0.03555
| Median |  |  |  | 0.035665 |  | 0.036509 | 0.03555
| Std Dev. |  |  |  | 0.0010607 |  | 0.00034154 | 5.4772e-05
| Rel.Std.Dev. |  |  |  | 2.979% |  | 0.9367% | 0.1541%
| PDM3 |  |  |  | -0.7327% |  | 1.637% | -0.9046%
| Z-Score (Absolute) |  | NA | NA | 0 | NA | 9.3 | 0.67 | NA | NA
| Analytical results for Mo in OREAS 506b (Indicative Value 86 ppm) | [9 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.06 | 2.14 | 2.22 | 2.24 |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  |  | 0
0
| 1 | 1 | NA | 82.2 | -- | 89 | N.A. |  |  |  | 1
| 1 | 2 | NA | 82.3 | -- | 89 | N.A. |  |  |  | 19
| 1 | 3 | NA | 81.8 | -- | 89 | N.A. |  |  |  | 16
| 1 | 4 | NA | 82.5 | -- | 89 | N.A. |  |  |  | 85.6417
| 1 | 5 | NA | 82.4 | -- | 89 | N.A. |  |  |  | 25
| 1 | 6 | NA | 82.5 | -- | 89 | N.A.
| Mean |  |  | 82.2833 |  | 89
| Median |  |  | 82.35 |  | 89
| Std Dev. |  |  | 0.26394 |  | 0
| Rel.Std.Dev. |  |  | 0.3208% |  | 0%
| PDM3 |  |  | -3.921% |  | 3.921%
| Z-Score (Absolute) |  | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Na in OREAS 506b (Indicative Value 0.152 wt.%) | [9 empty] | Indicative
Na | Round | Replicate | Lab | [7 empty] | 1
| No | No | 2.06 | [7 empty] | wt.%
|  |  | PF*OES | [7 empty] | 3
3
| 1 | 1 | 0.15 | [7 empty] | 1
| 1 | 2 | 0.16 | [7 empty] | 20
| 1 | 3 | 0.15 | [7 empty] | 16
| 1 | 4 | 0.15 | [7 empty] | 0.15167
| 1 | 5 | 0.15 | [7 empty] | 26
| 1 | 6 | 0.15
| Mean |  | 0.15167
| Median |  | 0.15
| Std Dev. |  | 0.0040825
| Rel.Std.Dev. |  | 2.692%
| PDM3 |  | -2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 506b (Indicative Value 12.5 ppm) | [9 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.14 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | [5 empty] | 1
1
| 1 | 1 | - | 12.2 | -- | [5 empty] | 1
| 1 | 2 | - | 12.3 | -- | [5 empty] | 21
| 1 | 3 | - | 12.5 | -- | [5 empty] | 16
| 1 | 4 | - | 12.5 | -- | [5 empty] | 12.4833
| 1 | 5 | - | 12.8 | -- | [5 empty] | 27
| 1 | 6 | - | 12.6 | --
| Mean |  |  | 12.4833
| Median |  |  | 12.5
| Std Dev. |  |  | 0.2137
| Rel.Std.Dev. |  |  | 1.712%
| PDM3 |  |  | 2.665e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Nd in OREAS 506b (Indicative Value 26.4 ppm) | [9 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 1
1
| 1 | 1 | 26.54 | -- | [6 empty] | 1
| 1 | 2 | 25.91 | -- | [6 empty] | 22
| 1 | 3 | 24.21 | -- | [6 empty] | 16
| 1 | 4 | 27.95 | -- | [6 empty] | 26.3833
| 1 | 5 | 27.5 | -- | [6 empty] | 28
| 1 | 6 | 26.19 | --
| Mean |  | 26.3833
| Median |  | 26.365
| Std Dev. |  | 1.3199
| Rel.Std.Dev. |  | 5.003%
| PDM3 |  | 1.332e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ni in OREAS 506b (Indicative Value 13.8 ppm) | [9 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.24 | 2.27 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 1
1
| 1 | 1 | NA | - | 14.13 | -- | N.A. | N.A. |  |  | 1
| 1 | 2 | NA | - | 13.55 | -- | N.A. | N.A. |  |  | 23
| 1 | 3 | NA | - | 13.52 | -- | N.A. | N.A. |  |  | 16
| 1 | 4 | NA | - | 14.14 | -- | N.A. | N.A. |  |  | 13.8067
| 1 | 5 | NA | - | 13.58 | -- | N.A. | N.A. |  |  | 29
| 1 | 6 | NA | - | 13.92 | -- | N.A. | N.A.
| Mean |  |  |  | 13.8067
| Median |  |  |  | 13.75
| Std Dev. |  |  |  | 0.29255
| Rel.Std.Dev. |  |  |  | 2.119%
| PDM3 |  |  |  | -2.442e-13%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for P in OREAS 506b (Indicative Value 0.076 wt.%) | [9 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 3
3
| 1 | 1 | NA | - | 0.06 | -- | 0.0925 | N.A. |  |  | 1
| 1 | 2 | NA | - | 0.06 | -- | 0.093 | N.A. |  |  | 24
| 1 | 3 | NA | - | 0.06 | -- | 0.0926 | N.A. |  |  | 16
| 1 | 4 | NA | - | 0.06 | -- | 0.0925 | N.A. |  |  | 0.076325
| 1 | 5 | NA | - | 0.06 | -- | 0.0928 | N.A. |  |  | 30
| 1 | 6 | NA | - | 0.06 | -- | 0.0925 | N.A.
| Mean |  |  |  | 0.06 |  | 0.09265
| Median |  |  |  | 0.06 |  | 0.09255
| Std Dev. |  |  |  | 0 |  | 0.00020736
| Rel.Std.Dev. |  |  |  | 0% |  | 0.2238%
| PDM3 |  |  |  | -21.39% |  | 21.39%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Pb in OREAS 506b (Indicative Value 45.2 ppm) | [9 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.24 | 2.27 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 1
1
| 1 | 1 | NA | - | 43.1 | -- | N.A. | N.A. |  |  | 1
| 1 | 2 | NA | - | 44.7 | -- | N.A. | N.A. |  |  | 25
| 1 | 3 | NA | - | 45.7 | -- | N.A. | N.A. |  |  | 16
| 1 | 4 | NA | - | 44.9 | -- | N.A. | N.A. |  |  | 45.15
| 1 | 5 | NA | - | 46.3 | -- | N.A. | N.A. |  |  | 31
| 1 | 6 | NA | - | 46.2 | -- | N.A. | N.A.
| Mean |  |  |  | 45.15
| Median |  |  |  | 45.3
| Std Dev. |  |  |  | 1.1996
| Rel.Std.Dev. |  |  |  | 2.657%
| PDM3 |  |  |  | 0%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Pr in OREAS 506b (Indicative Value 6.83 ppm) | [9 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 6.8 | -- | [6 empty] | 1
| 1 | 2 | 6.8 | -- | [6 empty] | 26
| 1 | 3 | 6.3 | -- | [6 empty] | 16
| 1 | 4 | 7.2 | -- | [6 empty] | 6.8333
| 1 | 5 | 7.1 | -- | [6 empty] | 32
| 1 | 6 | 6.8 | --
| Mean |  | 6.8333
| Median |  | 6.8
| Std Dev. |  | 0.31411
| Rel.Std.Dev. |  | 4.597%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for S in OREAS 506b (Indicative Value 0.556 wt.%) | [9 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.22 | 2.27 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES |  |  |  | 3
3
| 1 | 1 | NA | - | 0.49 | 0.63 | N.A. |  |  |  | 1
| 1 | 2 | NA | - | 0.48 | 0.63 | N.A. |  |  |  | 7
| 1 | 3 | NA | - | 0.48 | 0.63 | N.A. |  |  |  | 16
| 1 | 4 | NA | - | 0.48 | 0.63 | N.A. |  |  |  | 0.55583
| 1 | 5 | NA | - | 0.48 | 0.63 | N.A. |  |  |  | 33
| 1 | 6 | NA | - | 0.48 | 0.63 | N.A.
| Mean |  |  |  | 0.48167 | 0.63
| Median |  |  |  | 0.48 | 0.63
| Std Dev. |  |  |  | 0.0040825 | 0
| Rel.Std.Dev. |  |  |  | 0.8476% | 0%
| PDM3 |  |  |  | -13.34% | 13.34%
| Z-Score (Absolute) |  | NA | NA | 0.67 | 0.67 | NA
| Analytical results for Sb in OREAS 506b (Indicative Value < 5 ppm) | [9 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.06 | 2.14 | 2.24 | [5 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES | [5 empty] | 2
2
| 1 | 1 | < 5 | -- | N.A. | [5 empty] | 1
| 1 | 2 | < 5 | -- | N.A. | [5 empty] | 28
| 1 | 3 | < 5 | -- | N.A. | [5 empty] | 16
| 1 | 4 | < 5 | -- | N.A. | [5 empty] | < 5
| 1 | 5 | < 5 | -- | N.A. | [5 empty] | 34
| 1 | 6 | < 5 | -- | N.A.
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Sc in OREAS 506b (Indicative Value 9.05 ppm) | [9 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.06 | 2.14 | 2.22 | 2.24 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES |  |  |  | 2
2
| 1 | 1 | NA | 8.1 | -- | 10 | N.A. |  |  |  | 1
| 1 | 2 | NA | 8 | -- | 9 | N.A. |  |  |  | 29
| 1 | 3 | NA | 8.6 | -- | 10 | N.A. |  |  |  | 16
| 1 | 4 | NA | 8.5 | -- | 10 | N.A. |  |  |  | 9.05
| 1 | 5 | NA | 8.2 | -- | 10 | N.A. |  |  |  | 35
| 1 | 6 | NA | 8.2 | -- | 10 | N.A.
| Mean |  |  | 8.2667 |  | 9.8333
| Median |  |  | 8.2 |  | 10
| Std Dev. |  |  | 0.23381 |  | 0.40825
| Rel.Std.Dev. |  |  | 2.828% |  | 4.152%
| PDM3 |  |  | -8.656% |  | 8.656%
| Z-Score (Absolute) |  | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Si in OREAS 506b (Indicative Value 31.9 wt.%) | [9 empty] | Indicative
Si | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.18 | 2.22 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES |  |  |  |  | 2
2
| 1 | 1 | - | 31.3773 | 32.46 | N.A. |  |  |  |  | 1
| 1 | 2 | - | 31.3817 | 32.45 | N.A. |  |  |  |  | 11
| 1 | 3 | - | 31.3948 | 32.47 | N.A. |  |  |  |  | 16
| 1 | 4 | - | 31.3601 | 32.42 | N.A. |  |  |  |  | 31.8985
| 1 | 5 | - | 31.3012 | 32.45 | N.A. |  |  |  |  | 17
| 1 | 6 | - | 31.3169 | 32.4 | N.A.
| Mean |  |  | 31.3554 | 32.4417
| Median |  |  | 31.3687 | 32.45
| Std Dev. |  |  | 0.037847 | 0.026394
| Rel.Std.Dev. |  |  | 0.1207% | 0.08136%
| PDM3 |  |  | -1.703% | 1.703%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67 | NA
| Analytical results for Sm in OREAS 506b (Indicative Value 5.73 ppm) | [9 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 5.8 | -- | [6 empty] | 1
| 1 | 2 | 5.7 | -- | [6 empty] | 12
| 1 | 3 | 5.3 | -- | [6 empty] | 16
| 1 | 4 | 5.8 | -- | [6 empty] | 5.7333
| 1 | 5 | 6 | -- | [6 empty] | 18
| 1 | 6 | 5.8 | --
| Mean |  | 5.7333
| Median |  | 5.8
| Std Dev. |  | 0.23381
| Rel.Std.Dev. |  | 4.078%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Sn in OREAS 506b (Indicative Value < 5 ppm) | [9 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.04 | 2.06 | 2.14 | 2.24 | 2.27 |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  |  | 2
2
| 1 | 1 | - | < 5 | -- | N.A. | N.A. |  |  |  | 1
| 1 | 2 | - | < 5 | -- | N.A. | N.A. |  |  |  | 13
| 1 | 3 | - | < 5 | -- | N.A. | N.A. |  |  |  | 16
| 1 | 4 | - | < 5 | -- | N.A. | N.A. |  |  |  | < 5
| 1 | 5 | - | < 5 | -- | N.A. | N.A. |  |  |  | 19
| 1 | 6 | - | < 5 | -- | N.A. | N.A.
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Sr in OREAS 506b (Indicative Value 212 ppm) | [9 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.18 | 2.22 | 2.24 |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  | 0
0
| 1 | 1 | NA | - | 99.8 | -- | 276.94 | 261 | N.A. |  | 1
| 1 | 2 | NA | - | 98.5 | -- | 274.27 | 262 | N.A. |  | 14
| 1 | 3 | NA | - | 100.2 | -- | 270.64 | 261 | N.A. |  | 16
| 1 | 4 | NA | - | 98.1 | -- | 277.56 | 263 | N.A. |  | 211.6661
| 1 | 5 | NA | - | 99.7 | -- | 270.19 | 262 | N.A. |  | 20
| 1 | 6 | NA | - | 99.8 | -- | 271.29 | 264 | N.A.
| Mean |  |  |  | 99.35 |  | 273.4817 | 262.1667
| Median |  |  |  | 99.75 |  | 272.78 | 262
| Std Dev. |  |  |  | 0.84083 |  | 3.2535 | 1.169
| Rel.Std.Dev. |  |  |  | 0.8463% |  | 1.19% | 0.4459%
| PDM3 |  |  |  | -53.06% |  | 29.2% | 23.86%
| Z-Score (Absolute) |  | NA | NA | 9.7 | NA | 0.67 | 0 | NA
| Analytical results for Ta in OREAS 506b (Indicative Value 1.13 ppm) | [9 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.14 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | - | 1.2 | -- | [5 empty] | 1
| 1 | 2 | - | 1.2 | -- | [5 empty] | 15
| 1 | 3 | - | 1.1 | -- | [5 empty] | 16
| 1 | 4 | - | 1.1 | -- | [5 empty] | 1.1333
| 1 | 5 | - | 1.1 | -- | [5 empty] | 21
| 1 | 6 | - | 1.1 | --
| Mean |  |  | 1.1333
| Median |  |  | 1.1
| Std Dev. |  |  | 0.05164
| Rel.Std.Dev. |  |  | 4.556%
| PDM3 |  |  | 2.665e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Tb in OREAS 506b (Indicative Value 0.73 ppm) | [9 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 0.7 | -- | [6 empty] | 1
| 1 | 2 | 0.7 | -- | [6 empty] | 16
| 1 | 3 | 0.7 | -- | [6 empty] | 16
| 1 | 4 | 0.8 | -- | [6 empty] | 0.73333
| 1 | 5 | 0.8 | -- | [6 empty] | 22
| 1 | 6 | 0.7 | --
| Mean |  | 0.73333
| Median |  | 0.7
| Std Dev. |  | 0.05164
| Rel.Std.Dev. |  | 7.042%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 506b (Indicative Value 5.68 ppm) | [9 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.06 | 2.14 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | [5 empty] | 2
2
| 1 | 1 | NA | 5.7 | -- | [5 empty] | 1
| 1 | 2 | NA | 5.7 | -- | [5 empty] | 17
| 1 | 3 | NA | 5.7 | -- | [5 empty] | 16
| 1 | 4 | NA | 5.6 | -- | [5 empty] | 5.6833
| 1 | 5 | NA | 5.7 | -- | [5 empty] | 23
| 1 | 6 | NA | 5.7 | --
| Mean |  |  | 5.6833
| Median |  |  | 5.7
| Std Dev. |  |  | 0.040825
| Rel.Std.Dev. |  |  | 0.7183%
| PDM3 |  |  | 6.661e-14%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Ti in OREAS 506b (Indicative Value 0.284 wt.%) | [9 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | NA | - | 0.23 | -- | 0.35 | N.A. | N.A. |  | 1
| 1 | 2 | NA | - | 0.23 | -- | 0.35 | N.A. | N.A. |  | 18
| 1 | 3 | NA | - | 0.22 | -- | 0.34 | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 0.22 | -- | 0.34 | N.A. | N.A. |  | 0.28417
| 1 | 5 | NA | - | 0.22 | -- | 0.34 | N.A. | N.A. |  | 24
| 1 | 6 | NA | - | 0.22 | -- | 0.35 | N.A. | N.A.
| Mean |  |  |  | 0.22333 |  | 0.345
| Median |  |  |  | 0.22 |  | 0.345
| Std Dev. |  |  |  | 0.005164 |  | 0.0054772
| Rel.Std.Dev. |  |  |  | 2.312% |  | 1.588%
| PDM3 |  |  |  | -21.41% |  | 21.41%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA | NA
| Analytical results for Tl in OREAS 506b (Indicative Value < 1 ppm) | [9 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS |  |  |  |  | 2
2
| 1 | 1 | NA | - | < 1 | -- |  |  |  |  | 1
| 1 | 2 | NA | - | < 1 | -- |  |  |  |  | 19
| 1 | 3 | NA | - | < 1 | -- |  |  |  |  | 16
| 1 | 4 | NA | - | < 1 | -- |  |  |  |  | < 1
| 1 | 5 | NA | - | < 1 | -- |  |  |  |  | 25
| 1 | 6 | NA | - | < 1 | --
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Tm in OREAS 506b (Indicative Value < 1 ppm) | [9 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | < 1 | -- | [6 empty] | 1
| 1 | 2 | < 1 | -- | [6 empty] | 20
| 1 | 3 | < 1 | -- | [6 empty] | 16
| 1 | 4 | < 1 | -- | [6 empty] | < 1
| 1 | 5 | < 1 | -- | [6 empty] | 26
| 1 | 6 | < 1 | --
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U in OREAS 506b (Indicative Value 3.82 ppm) | [9 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.06 | 2.14 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | [5 empty] | 2
2
| 1 | 1 | NA | 3.4 | -- | [5 empty] | 1
| 1 | 2 | NA | 4 | -- | [5 empty] | 21
| 1 | 3 | NA | 4 | -- | [5 empty] | 16
| 1 | 4 | NA | 3.6 | -- | [5 empty] | 3.8167
| 1 | 5 | NA | 3.9 | -- | [5 empty] | 27
| 1 | 6 | NA | 4 | --
| Mean |  |  | 3.8167
| Median |  |  | 3.95
| Std Dev. |  |  | 0.25626
| Rel.Std.Dev. |  |  | 6.714%
| PDM3 |  |  | -8.882e-14%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for V in OREAS 506b (Indicative Value 63 ppm) | [9 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | NA | - | 53.8 | -- | 73 | N.A. |  |  | 1
| 1 | 2 | NA | - | 52.5 | -- | 73 | N.A. |  |  | 22
| 1 | 3 | NA | - | 53.3 | -- | 73 | N.A. |  |  | 16
| 1 | 4 | NA | - | 53 | -- | 73 | N.A. |  |  | 63.0583
| 1 | 5 | NA | - | 52.4 | -- | 73 | N.A. |  |  | 28
| 1 | 6 | NA | - | 53.7 | -- | 73 | N.A.
| Mean |  |  |  | 53.1167 |  | 73
| Median |  |  |  | 53.15 |  | 73
| Std Dev. |  |  |  | 0.59133 |  | 0
| Rel.Std.Dev. |  |  |  | 1.113% |  | 0%
| PDM3 |  |  |  | -15.77% |  | 15.77%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA
| Analytical results for W in OREAS 506b (Indicative Value 10 ppm) | [9 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.06 | 2.14 | 2.24 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*MS | PF*OES |  |  |  |  | 1
1
| 1 | 1 | - | 9.7 | -- | N.A. |  |  |  |  | 1
| 1 | 2 | - | 9.2 | -- | N.A. |  |  |  |  | 23
| 1 | 3 | - | 9.5 | -- | N.A. |  |  |  |  | 16
| 1 | 4 | - | 11.3 | -- | N.A. |  |  |  |  | 10
| 1 | 5 | - | 9.7 | -- | N.A. |  |  |  |  | 29
| 1 | 6 | - | 10.6 | -- | N.A.
| Mean |  |  | 10
| Median |  |  | 9.7
| Std Dev. |  |  | 0.78994
| Rel.Std.Dev. |  |  | 7.899%
| PDM3 |  |  | 2.22e-14%
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Y in OREAS 506b (Indicative Value 22.1 ppm) | [9 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.06 | 2.14 | 2.22 | 2.24 |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES |  |  |  | 1
1
| 1 | 1 | NA | 14.5 | -- | 28 | N.A. |  |  |  | 1
| 1 | 2 | NA | 14.8 | -- | 29 | N.A. |  |  |  | 24
| 1 | 3 | NA | 16.5 | -- | 28 | N.A. |  |  |  | 16
| 1 | 4 | NA | 15.5 | -- | 29 | N.A. |  |  |  | 22.05
| 1 | 5 | NA | 17.1 | -- | 28 | N.A. |  |  |  | 30
| 1 | 6 | NA | 15.2 | -- | 29 | N.A.
| Mean |  |  | 15.6 |  | 28.5
| Median |  |  | 15.35 |  | 28.5
| Std Dev. |  |  | 1.008 |  | 0.54772
| Rel.Std.Dev. |  |  | 6.461% |  | 1.922%
| PDM3 |  |  | -29.25% |  | 29.25%
| Z-Score (Absolute) |  | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Yb in OREAS 506b (Indicative Value 1.48 ppm) | [9 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | [6 empty] | 1
| No | No | 2.06 | 2.14 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.4 | -- | [6 empty] | 1
| 1 | 2 | 1.4 | -- | [6 empty] | 25
| 1 | 3 | 1.6 | -- | [6 empty] | 16
| 1 | 4 | 1.5 | -- | [6 empty] | 1.4833
| 1 | 5 | 1.6 | -- | [6 empty] | 31
| 1 | 6 | 1.4 | --
| Mean |  | 1.4833
| Median |  | 1.45
| Std Dev. |  | 0.098319
| Rel.Std.Dev. |  | 6.628%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Zn in OREAS 506b (Indicative Value 110 ppm) | [9 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.14 | 2.22 | 2.24 | 2.27 |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES |  | 0
0
| 1 | 1 | NA | - | 96.7 | -- | 120 | N.A. | N.A. |  | 1
| 1 | 2 | NA | - | 101.5 | -- | 121 | N.A. | N.A. |  | 26
| 1 | 3 | NA | - | 100.8 | -- | 120 | N.A. | N.A. |  | 16
| 1 | 4 | NA | - | 96.4 | -- | 121 | N.A. | N.A. |  | 109.8833
| 1 | 5 | NA | - | 99.2 | -- | 120 | N.A. | N.A. |  | 32
| 1 | 6 | NA | - | 102 | -- | 120 | N.A. | N.A.
| Mean |  |  |  | 99.4333 |  | 120.3333
| Median |  |  |  | 100 |  | 120
| Std Dev. |  |  |  | 2.4271 |  | 0.5164
| Rel.Std.Dev. |  |  |  | 2.441% |  | 0.4291%
| PDM3 |  |  |  | -9.51% |  | 9.51%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA | NA
| Analytical results for Zr in OREAS 506b (Indicative Value 10.1 ppm) | [9 empty] | Indicative
Zr | Round | Replicate | Lab | [7 empty] | 1
| No | No | 2.06 | [7 empty] | ppm
|  |  | PF*OES | [7 empty] | 1
1
| 1 | 1 | 9.9 | [7 empty] | 1
| 1 | 2 | 10 | [7 empty] | 27
| 1 | 3 | 9.6 | [7 empty] | 16
| 1 | 4 | 10.7 | [7 empty] | 10.1
| 1 | 5 | 9.9 | [7 empty] | 33
| 1 | 6 | 10.5
| Mean |  | 10.1
| Median |  | 9.95
| Std Dev. |  | 0.41473
| Rel.Std.Dev. |  | 4.106%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 506b (Certified Value 1.34 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.44 | 1.3 | 1.4589 | 1.22 | 1.23 | 1.4 | 1.32 | 1.44 | 1.36 | 1.25 | 1.4 | 1.71 | 1.34 | 1.2967 | < 2 | 1.3 | 1.373 | 1.3 | 1.3 | 1.4 | 1.1 | 1.34 | 1.5 | 1 | 1.31 |  | 1
| 1 | 2 | 1.46 | 1.3 | 1.3605 | 1.35 | 1.23 | 1.42 | 1.28 | 1.38 | 1.35 | 1.36 | 1.39 | 1.92 | 1.36 | 1.2873 | < 2 | 1.5 | 1.2995 | 1.3 | 1.4 | 1.4 | 1.13 | 1.37 | 1.5 | 1.1 | 1.244 |  | 11
| 1 | 3 | 1.39 | 1.3 | 1.3964 | 1.314 | 1.22 | 1.41 | 1.28 | 1.36 | 1.27 | 1.35 | 1.42 | 1.57 | 1.3 | 1.3147 | < 2 | 1.4 | 1.334 | 1.2 | 1.3 | 1.5 | 1.2 | 1.37 | 1.5 | 1.2 | 1.271 |  | 16
| 1 | 4 | 1.46 | 1.3 | 1.4526 | 1.239 | 1.22 | 1.39 | 1.33 | 1.4 | 1.36 | 1.24 | 1.42 | 1.66 | 1.36 | 1.37 | < 2 | 1.4 | 1.35 | 1.2 | 1.3 | 1.4 | 1.2 | 1.35 | 1.5 | 1.2 | 1.258 |  | 1.3353
| 1 | 5 | 1.51 | 1.3 | 1.3295 | 1.351 | 1.22 | 1.44 | 1.28 | 1.34 | 1.32 | 1.31 | 1.38 | 1.8 | 1.35 | 1.3016 | < 2 | 1.4 | 1.357 | 1.3 | 1.3 | 1.5 | 1.18 | 1.35 | 1.6 | 1.2 | 1.291 |  | 11
| 1 | 6 | 1.36 | 1.3 | 1.3994 | 1.307 | 1.23 | 1.38 | 1.25 | 1.32 | 1.34 | 1.26 | 1.37 | 1.73 | 1.36 | 1.3609 | < 2 | 1.5 | 1.3455 | 1.3 | 1.4 | 1.5 | 1.08 | 1.36 | 1.5 | 1.2 | 1.306
| Mean |  | 1.4367 | 1.3 | 1.3995 | 1.2968 | 1.225 | 1.4067 | 1.29 | 1.3733 | 1.3333 | 1.295 | 1.3967 | 1.7317 | 1.345 | 1.3219 |  | 1.4167 | 1.3432 | 1.2667 | 1.3333 | 1.45 | 1.1483 | 1.3567 | 1.5167 | 1.15 | 1.28
| Median |  | 1.45 | 1.3 | 1.3979 | 1.3105 | 1.225 | 1.405 | 1.28 | 1.37 | 1.345 | 1.285 | 1.395 | 1.72 | 1.355 | 1.3082 |  | 1.4 | 1.3478 | 1.3 | 1.3 | 1.45 | 1.155 | 1.355 | 1.5 | 1.2 | 1.281
| Std Dev. |  | 0.053914 | 0 | 0.050546 | 0.05551 | 0.0054772 | 0.021602 | 0.029665 | 0.043205 | 0.034448 | 0.05244 | 0.020656 | 0.1199 | 0.023452 | 0.035013 |  | 0.075277 | 0.024997 | 0.05164 | 0.05164 | 0.054772 | 0.052313 | 0.012111 | 0.040825 | 0.083666 | 0.026676
| Rel.Std.Dev. |  | 3.753% | 0% | 3.612% | 4.28% | 0.4471% | 1.536% | 2.3% | 3.146% | 2.584% | 4.049% | 1.479% | 6.924% | 1.744% | 2.649% |  | 5.314% | 1.861% | 4.077% | 3.873% | 3.777% | 4.556% | 0.8927% | 2.692% | 7.275% | 2.084%
| PDM3 |  | 7.593% | -2.642% | 4.813% | -2.879% | -8.259% | 5.346% | -3.391% | 2.85% | -0.1457% | -3.017% | 4.597% | 29.69% | 0.728% | -1.004% |  | 6.095% | 0.5907% | -5.138% | -0.1457% | 8.592% | -14% | 1.602% | 13.58% | -13.88% | -4.14%
| Z-Score (Absolute) |  | 1.1 | 0.35 | 0.7 | 0.39 | 1.15 | 0.78 | 0.46 | 0.43 | 0 | 0.41 | 0.67 | 4.24 | 0.12 | 0.12 | 3.55 | 0.89 | 0.1 | 0.71 | 0 | 1.24 | 1.97 | 0.25 | 1.95 | 1.95 | 0.57
| Analytical results for Al in OREAS 506b (Certified Value 7.55 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 7.675 | 7.53 | 7.7106 | 6.96 | 7.52 | 7.82 | 6.92 | 7.6 | 8.15 | 7.62 | 7.5 | 8.54 | 7.53 | 7.3859 | 7.54 | 7.65 | 7.9467 | 7.84 | 7.53 | 7.27 | 7.56 | 8.21 | 7.37 | 6.9919 | 7.42 |  | 1
| 1 | 2 | 7.7684 | 7.56 | 7.525 | 7.42 | 7.37 | 7.77 | 7.07 | 7.67 | 8.05 | 7.77 | 7.56 | 8.61 | 7.63 | 7.3572 | 7.79 | 7.53 | 7.9455 | 7.58 | 7.48 | 7.06 | 7.5 | 8.25 | 7.9 | 7.0981 | 7.32 |  | 12
| 1 | 3 | 7.6437 | 7.53 | 7.4681 | 7.59 | 7.58 | 7.72 | 6.77 | 7.66 | 8.14 | 8.03 | 7.51 | 8.5 | 7.49 | 7.4237 | 7.72 | 7.6 | 7.94 | 7.52 | 7.5 | 7.22 | 7.54 | 8.29 | 7.65 | 7.1732 | 7.35 |  | 16
| 1 | 4 | 7.6662 | 7.57 | 7.6785 | 6.61 | 7.24 | 7.65 | 7.28 | 7.59 | 8.28 | 7.82 | 7.54 | 8.4 | 7.47 | 7.3412 | 7.6 | 7.56 | 7.9468 | 7.61 | 7.75 | 7.47 | 7.4743 | 8.17 | 7.64 | 7.1641 | 7.36 |  | 7.55
| 1 | 5 | 7.7367 | 7.61 | 7.4978 | 7.4 | 7.36 | 7.81 | 6.89 | 7.58 | 8.13 | 7.75 | 7.5 | 8.17 | 7.38 | 7.1117 | 7.26 | 7.42 | 7.9445 | 7.68 | 7.07 | 7.34 | 7.5692 | 8.19 | 7.91 | 7.1309 | 7.23 |  | 12
| 1 | 6 | 7.7565 | 7.38 | 7.5236 | 7.95 | 7.43 | 7.76 | 7.21 | 7.51 | 8.06 | 7.66 | 7.55 | 8.62 | 7.56 | 7.5795 | 7.91 | 7.49 | 7.9412 | 7.67 | 7.43 | 7.42 | 7.54 | 8.27 | 7.76 | 7.0704 | 7.26
| Mean |  | 7.7077 | 7.53 | 7.5673 | 7.3217 | 7.4167 | 7.755 | 7.0233 | 7.6017 | 8.135 | 7.775 | 7.5267 | 8.4733 | 7.51 | 7.3665 | 7.6367 | 7.5417 | 7.9441 | 7.65 | 7.46 | 7.2967 | 7.5306 | 8.23 | 7.705 | 7.1048 | 7.3233
| Median |  | 7.7058 | 7.545 | 7.5243 | 7.41 | 7.4 | 7.765 | 6.995 | 7.595 | 8.135 | 7.76 | 7.525 | 8.52 | 7.51 | 7.3716 | 7.66 | 7.545 | 7.945 | 7.64 | 7.49 | 7.305 | 7.54 | 8.23 | 7.705 | 7.1145 | 7.335
| Std Dev. |  | 0.052525 | 0.079246 | 0.10126 | 0.47326 | 0.12176 | 0.062849 | 0.19775 | 0.058452 | 0.082644 | 0.14488 | 0.026583 | 0.16896 | 0.085088 | 0.15145 | 0.22704 | 0.081343 | 0.0028771 | 0.11027 | 0.22091 | 0.14814 | 0.036434 | 0.047329 | 0.20127 | 0.067603 | 0.069474
| Rel.Std.Dev. |  | 0.6815% | 1.052% | 1.338% | 6.464% | 1.642% | 0.8104% | 2.816% | 0.7689% | 1.016% | 1.863% | 0.3532% | 1.994% | 1.133% | 2.056% | 2.973% | 1.079% | 0.03622% | 1.441% | 2.961% | 2.03% | 0.4838% | 0.5751% | 2.612% | 0.9515% | 0.9487%
| PDM3 |  | 2.09% | -0.2647% | 0.2291% | -3.024% | -1.766% | 2.715% | -6.976% | 0.6845% | 7.749% | 2.98% | -0.3089% | 12.23% | -0.5296% | -2.43% | 1.148% | -0.1102% | 5.22% | 1.325% | -1.192% | -3.355% | -0.257% | 9.007% | 2.053% | -5.897% | -3.002%
| Z-Score (Absolute) |  | 0.67 | 0.05 | 0.1 | 0.89 | 0.51 | 0.87 | 2.1 | 0.24 | 2.41 | 0.95 | 0.06 | 3.78 | 0.13 | 0.71 | 0.39 | 0 | 1.63 | 0.44 | 0.33 | 0.99 | 0.04 | 2.79 | 0.66 | 1.77 | 0.89
| Analytical results for As in OREAS 506b (Certified Value 13.6 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 13 | 13 | 12.5797 | 13 | 19.34 | 13.9 | 14 | 13.6 | 13.7 | 12.2 | 13 | 2.89 | 15 | 16.0482 | 9 | 13.6 | 13 | 12.1 | 14.2 | NA | 10 | 12.9 | 10.8 | 6.9 |  |  | 1
| 1 | 2 | 13.7 | 13 | 12.933 | 14.4 | 18.22 | 13.9 | 14.6 | 13.9 | 13.6 | 12.6 | 13 | 2.11 | 14 | 15.3162 | 11 | 13.4 | 13 | 13.4 | 14.6 | NA | 10 | 14.8 | 12.5 | 6.7 |  |  | 13
| 1 | 3 | 13.4 | 13 | 13.2341 | 13.8 | 15.96 | 14.4 | 13 | 13.3 | 13.2 | 13.6 | 14 | 2.51 | 15 | 15.6152 | 13 | 14.4 | 12 | 12.4 | 14.4 | NA | 10 | 13.8 | 12.7 | 6.3 |  |  | 16
| 1 | 4 | 13.5 | 13 | 12.7432 | 13.1 | 16.78 | 15.4 | 12.4 | 13.2 | 14.7 | 13.2 | 13 | 8.35 | 15 | 15.4421 | 12 | 13.8 | 13 | 12.9 | 14.5 | NA | 10 | 15.5 | 11.7 | 5.7 |  |  | 13.5864
| 1 | 5 | 13.4 | 13 | 13.0888 | 14.4 | 17.56 | 14.1 | 13 | 13.2 | 14.3 | 11.9 | 13 | < 0.02 | 13 | 14.7303 | 12 | 13.3 | 12 | 12.4 | 14.7 | NA | 10 | 14.4 | 12.5 | 6.8 |  |  | 13
| 1 | 6 | 13.6 | 13 | 12.5627 | 12.9 | 19.72 | 13.7 | 14.2 | 13.4 | 13.7 | 12.2 | 13 | 2.55 | 14 | 15.9636 | 10 | 14.8 | 13 | 12 | 14.6 | NA | 10 | 15 | 12 | 5.9
| Mean |  | 13.4333 | 13 | 12.8569 | 13.6 | 17.93 | 14.2333 | 13.5333 | 13.4333 | 13.8667 | 12.6167 | 13.1667 | 3.682 | 14.3333 | 15.5193 | 11.1667 | 13.8833 | 12.6667 | 12.5333 | 14.5 |  | 10 | 14.4 | 12.0333 | 6.3833
| Median |  | 13.45 | 13 | 12.8381 | 13.45 | 17.89 | 14 | 13.5 | 13.35 | 13.7 | 12.4 | 13 | 2.55 | 14.5 | 15.5287 | 11.5 | 13.7 | 13 | 12.4 | 14.55 |  | 10 | 14.6 | 12.25 | 6.5
| Std Dev. |  | 0.24221 | 0 | 0.27502 | 0.6957 | 1.4572 | 0.6186 | 0.85479 | 0.27325 | 0.53914 | 0.65853 | 0.40825 | 2.6241 | 0.8165 | 0.48086 | 1.472 | 0.5947 | 0.5164 | 0.52789 | 0.17889 |  | 0 | 0.93167 | 0.70899 | 0.49967
| Rel.Std.Dev. |  | 1.803% | 0% | 2.139% | 5.115% | 8.127% | 4.346% | 6.316% | 2.034% | 3.888% | 5.22% | 3.101% | 71.27% | 5.696% | 3.098% | 13.18% | 4.284% | 4.077% | 4.212% | 1.234% |  | 0% | 6.47% | 5.892% | 7.828%
| PDM3 |  | -1.127% | -4.316% | -5.369% | 0.1001% | 31.97% | 4.762% | -0.3906% | -1.127% | 2.063% | -7.138% | -3.089% | -72.9% | 5.498% | 14.23% | -17.81% | 2.186% | -6.769% | -7.751% | 6.724% |  | -26.4% | 5.988% | -11.43% | -53.02%
| Z-Score (Absolute) |  | 0.08 | NA | 0.53 | 0.05 | 3.42 | 0.54 | 0 | 0.08 | 0.26 | 0.71 | NA | 8.14 | NA | 1.55 | NA | 0.27 | NA | 0.78 | 0.75 | NA | NA | 0.67 | 1.17 | 5.56
Results from laboratories 2.02, 2.12, 2.14, 2.16, 2.19 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 506b (Indicative Value < 10 ppm) | [27 empty] | Indicative
B | Round | Replicate | 2.01: Lab |  |  | 1
| No | No | 2.01: 2.22 |  |  | ppm
|  |  | 2.01: 4A*MS |  |  | 1
1
| 1 | 1 | 2.01: < 10 |  |  | 1
| 1 | 2 | 2.01: < 10 |  |  | 31
| 1 | 3 | 2.01: < 10 |  |  | 16
| 1 | 4 | 2.01: < 10 |  |  | < 10
| 1 | 5 | 2.01: < 10 |  |  | 37
| 1 | 6 | 2.01: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Ba in OREAS 506b (Certified Value 938 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 0
0
| 1 | 1 | 968.5 | 939 | 958.7486 | 868 | 965.02 | 960 | 920 | 970 | 1030 | 870 | 923 | 850 | 955 | 916.8054 | 940 | 937 | 970 | 999 | 1076 | 908 | 826 | 890 |  |  |  |  | 1
| 1 | 2 | 993.5 | 952 | 921.0376 | 880 | 1123.5 | 980 | 960 | 960 | 1010 | 880 | 928 | 850 | 971 | 921.7243 | 967 | 915 | 985 | 1025 | 1072 | 907 | 815 | 874 |  |  |  |  | 14
| 1 | 3 | 960 | 943 | 964.9941 | 960 | 1147.83 | 960 | 920 | 960 | 1030 | 910 | 921 | 840 | 922 | 910.6917 | 963 | 910 | 968 | 1012 | 1048 | 909 | 855 | 936 |  |  |  |  | 16
| 1 | 4 | 968.9 | 948 | 975.4938 | 821 | 1136 | 960 | 920 | 960 | 1050 | 880 | 927 | 840 | 941 | 927.0903 | 943 | 914 | 990 | 992 | 1047 | 909 | 867 | 880 |  |  |  |  | 937.755
| 1 | 5 | 976.2 | 933 | 942.6167 | 940 | 1115.5 | 970 | 920 | 950 | 1020 | 880 | 921 | 810 | 946 | 908.8223 | 909 | 924 | 965 | 978 | 1029 | 908 | 894 | 898 |  |  |  |  | 14
| 1 | 6 | 980.4 | 946 | 925.58 | 865 | 1086.73 | 960 | 920 | 950 | 1020 | 880 | 924 | 860 | 950 | 953.1032 | 978 | 921 | 1000 | 969 | 1024 | 907 | 871 | 933
| Mean |  | 974.5833 | 943.5 | 948.0785 | 889 | 1095.7633 | 965 | 926.6667 | 958.3333 | 1026.6667 | 883.3333 | 924 | 841.6667 | 947.5 | 923.0395 | 950 | 920.1667 | 979.6667 | 995.8333 | 1049.3333 | 908 | 854.6667 | 901.8333
| Median |  | 972.55 | 944.5 | 950.6826 | 874 | 1119.5 | 960 | 920 | 960 | 1025 | 880 | 923.5 | 845 | 948 | 919.2648 | 953 | 918 | 977.5 | 995.5 | 1047.5 | 908 | 861 | 894
| Std Dev. |  | 11.6299 | 6.775 | 21.9946 | 51.6991 | 67.3264 | 8.3666 | 16.3299 | 7.5277 | 13.6626 | 13.6626 | 2.9665 | 17.224 | 16.1833 | 16.2199 | 24.7871 | 9.6626 | 14.0949 | 20.8559 | 21.3885 | 0.89443 | 29.5342 | 26.6264
| Rel.Std.Dev. |  | 1.193% | 0.7181% | 2.32% | 5.815% | 6.144% | 0.867% | 1.762% | 0.7855% | 1.331% | 1.547% | 0.321% | 2.046% | 1.708% | 1.757% | 2.609% | 1.05% | 1.439% | 2.094% | 2.038% | 0.09851% | 3.456% | 2.952%
| PDM3 |  | 3.927% | 0.6126% | 1.101% | -5.199% | 16.85% | 2.905% | -1.182% | 2.194% | 9.481% | -5.803% | -1.467% | -10.25% | 1.039% | -1.569% | 1.306% | -1.876% | 4.469% | 6.193% | 11.9% | -3.173% | -8.86% | -3.831%
| Z-Score (Absolute) |  | 0.62 | 0.04 | 0.05 | 1.2 | 3.2 | 0.42 | 0.4 | 0.27 | 1.73 | 1.33 | 0.46 | 2.21 | 0.04 | 0.48 | 0.1 | 0.54 | 0.73 | 1.07 | 2.21 | 0.8 | 1.94 | 0.93
| Analytical results for Be in OREAS 506b (Certified Value 2.25 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 2.31 | 2.2 | 2.2087 | 3 | 1.47 | 2.24 | 2.27 | 2.27 | 2.36 | 2.17 | 2.4 | 2.47 | 2.4 | 2.201 | 1.8 | 2 | 1.99 | 2.24 | 2 | 2 | 2 | 2.4 | 2 |  |  |  | 1
| 1 | 2 | 2.25 | 2.2 | 2.343 | 3 | 1.76 | 2.27 | 2.25 | 2.26 | 2.33 | 2.21 | 2.4 | 2.42 | 2.4 | 2.255 | 2.2 | 2.3 | 2.24 | 2.17 | 2 | 2 | 2 | 2.4 | 2 |  |  |  | 15
| 1 | 3 | 2.23 | 2.2 | 2.2741 | 3 | 1.71 | 2.26 | 2.23 | 2.25 | 2.36 | 2.29 | 2.4 | 2.46 | 2.3 | 2.3014 | 2.3 | 2.2 | 2.03 | 2.25 | 2 | 2.23 | 2 | 2.2 | 2 |  |  |  | 16
| 1 | 4 | 2.29 | 2.2 | 2.343 | 3 | 1.63 | 2.24 | 2.25 | 2.24 | 2.41 | 2.2 | 2.5 | 2.53 | 2.4 | 2.2048 | 1.8 | 2.4 | 2.39 | 2.35 | 2 | 2.1 | 2 | 2.4 | 2 |  |  |  | 2.246
| 1 | 5 | 2.27 | 2.2 | 2.079 | 3 | 1.72 | 2.28 | 2.28 | 2.24 | 2.35 | 2.19 | 2.4 | 2.17 | 2.3 | 2.1607 | 1.7 | 2.1 | 2.31 | 2.27 | 2.1 | 2 | 2 | 2.4 | 2 |  |  |  | 15
| 1 | 6 | 2.35 | 2.1 | 2.1253 | 3 | 1.56 | 2.25 | 2.24 | 2.23 | 2.31 | 2.19 | 2.5 | 2.58 | 2.4 | 2.233 | 2.3 | 2.4 | 2.01 | 2.28 | 1.9 | 2.2 | 2 | 2.2 | 2
| Mean |  | 2.2833 | 2.1833 | 2.2289 | 3 | 1.6417 | 2.2567 | 2.2533 | 2.2483 | 2.3533 | 2.2083 | 2.4333 | 2.4383 | 2.3667 | 2.226 | 2.0167 | 2.2333 | 2.1617 | 2.26 | 2 | 2.0883 | 2 | 2.3333 | 2
| Median |  | 2.28 | 2.2 | 2.2414 | 3 | 1.67 | 2.255 | 2.25 | 2.245 | 2.355 | 2.195 | 2.4 | 2.465 | 2.4 | 2.2189 | 2 | 2.25 | 2.135 | 2.26 | 2 | 2.05 | 2 | 2.4 | 2
| Std Dev. |  | 0.043205 | 0.040825 | 0.11107 | 0 | 0.11053 | 0.01633 | 0.018619 | 0.01472 | 0.033862 | 0.042151 | 0.05164 | 0.14303 | 0.05164 | 0.048843 | 0.27869 | 0.1633 | 0.17325 | 0.058652 | 0.063246 | 0.10591 | 0 | 0.10328 | 0
| Rel.Std.Dev. |  | 1.892% | 1.87% | 4.983% | 0% | 6.733% | 0.7236% | 0.8263% | 0.6547% | 1.439% | 1.909% | 2.122% | 5.866% | 2.182% | 2.194% | 13.82% | 7.312% | 8.015% | 2.595% | 3.162% | 5.071% | 0% | 4.426% | 0%
| PDM3 |  | 1.663% | -2.79% | -0.7628% | 33.57% | -26.91% | 0.4753% | 0.3269% | 0.1042% | 4.779% | -1.677% | 8.341% | 8.564% | 5.373% | -0.8905% | -10.21% | -0.5636% | -3.754% | 0.6237% | -10.95% | -7.02% | -10.95% | 3.889% | -10.95%
| Z-Score (Absolute) |  | 0.42 | 0.57 | 0.12 | NA | 5.91 | 0.16 | 0.12 | 0.07 | 1.11 | 0.32 | 1.9 | 1.95 | 1.24 | 0.15 | 2.21 | 0.07 | 0.78 | 0.19 | 2.38 | 1.5 | NA | 0.91 | NA
Results from laboratories 2.04, 2.24 and 2.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 506b (Certified Value 0.93 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
2
| 1 | 1 | 1.03 | 0.87 | 0.8506 | 0.81 | < 0.5 | 1.14 | 0.8 | 0.98 | 0.85 | 1.15 | 0.82 | 0.75 | 1.15 | 0.97623 | < 5 | 0.83 | 0.59267 | < 5 | 0.88 | 1.36 | 6.37 | 2.48 | 0.74 | 0.98 | 0.15 | 0.8 | 1
| 1 | 2 | 0.99 | 0.9 | 0.87772 | 1.04 | < 0.5 | 0.87 | 0.8 | 0.76 | 1.32 | 0.9 | 0.86 | 0.78 | 1.18 | 0.91097 | < 5 | 0.86 | 0.59493 | < 5 | 0.96 | 1.34 | 6.15 | 2.6 | 1.14 | 0.97 | 0.59 | 0.81 | 16
| 1 | 3 | 0.94 | 0.91 | 0.8431 | 0.96 | < 0.5 | 0.84 | 0.89 | 0.84 | 0.79 | 1.02 | 0.75 | 0.85 | 1.1 | 0.89821 | < 5 | 0.93 | 0.59447 | < 5 | 0.95 | 1.14 | 6.07 | 2.49 | 0.83 | 0.88 | 0.3 | 0.89 | 16
| 1 | 4 | 1.17 | 0.81 | 0.88456 | 1.14 | < 0.5 | 1.1 | 0.85 | 1.11 | 1.1 | 0.98 | 0.84 | 0.88 | 1.16 | 0.94084 | < 5 | 0.94 | 0.52857 | < 5 | 0.97 | 1.29 | 6.23 | 2.6 | 5.5 | 0.94 | 0.33 | 0.8 | 0.93177
| 1 | 5 | 0.85 | 0.9 | 0.83097 | 1.15 | < 0.5 | 0.95 | 0.72 | 0.78 | 0.9 | 0.78 | 0.96 | 1.26 | 1.14 | 0.95182 | < 5 | 0.97 | 0.54222 | < 5 | 0.78 | 1.09 | 6.11 | 2.71 | 0.84 | 1.05 | 0.64 | 0.94 | 16
| 1 | 6 | 0.81 | 0.89 | 0.89816 | 1.05 | < 0.5 | 0.8 | 0.77 | 1.04 | 1.01 | 0.92 | 0.83 | 0.73 | 1.16 | 0.93632 | < 5 | 0.89 | 0.64134 | < 5 | 0.84 | 0.93 | 6.31 | 2.67 | 0.81 | 0.9 | 0.38 | 0.8
| Mean |  | 0.965 | 0.88 | 0.86418 | 1.025 |  | 0.95 | 0.805 | 0.91833 | 0.995 | 0.95833 | 0.84333 | 0.875 | 1.1483 | 0.93573 |  | 0.90333 | 0.58237 |  | 0.89667 | 1.1917 | 6.2067 | 2.5917 | 1.6433 | 0.95333 | 0.39833 | 0.84
| Median |  | 0.965 | 0.895 | 0.86416 | 1.045 |  | 0.91 | 0.8 | 0.91 | 0.955 | 0.95 | 0.835 | 0.815 | 1.155 | 0.93858 |  | 0.91 | 0.59357 |  | 0.915 | 1.215 | 6.19 | 2.6 | 0.835 | 0.955 | 0.355 | 0.805
| Std Dev. |  | 0.13019 | 0.036878 | 0.026401 | 0.12661 |  | 0.14114 | 0.059582 | 0.14538 | 0.1944 | 0.12465 | 0.068313 | 0.19726 | 0.027142 | 0.028095 |  | 0.052789 | 0.040973 |  | 0.076594 | 0.16822 | 0.11759 | 0.092826 | 1.8944 | 0.06121 | 0.18519 | 0.060332
| Rel.Std.Dev. |  | 13.49% | 4.191% | 3.055% | 12.35% |  | 14.86% | 7.401% | 15.83% | 19.54% | 13.01% | 8.1% | 22.54% | 2.364% | 3.002% |  | 5.844% | 7.036% |  | 8.542% | 14.12% | 1.895% | 3.582% | 115.3% | 6.421% | 46.49% | 7.182%
| PDM3 |  | 3.567% | -5.556% | -7.253% | 10.01% |  | 1.957% | -13.61% | -1.442% | 6.786% | 2.851% | -9.491% | -6.092% | 23.24% | 0.4256% |  | -3.052% | -37.5% |  | -3.767% | 27.89% | 566.1% | 178.1% | 76.37% | 2.315% | -57.25% | -9.849%
| Z-Score (Absolute) |  | 0.16 | 0.47 | 0.58 | 0.61 | 5.14 | 0.05 | 1.02 | 0.18 | 0.39 | 0.11 | 0.74 | 0.5 | 1.53 | 0.05 | 11.56 | 0.29 | 2.68 | 11.56 | 0.34 | 1.85 | 39.08 | 12.24 | 5.2 | 0.08 | 4.04 | 0.76
| Analytical results for Ca in OREAS 506b (Certified Value 1.67 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.6679 | 1.66 | 1.6875 | 1.56 | 1.48 | 1.75 | 1.64 | 1.73 | 1.75 | 1.68 | 1.64 | 1.77 | 1.67 | 1.6265 | 1.684 | 1.63 | 1.7207 | 1.7 | 1.64 | 1.65 | 1.672 | 1.79 | 1.66 | 1.6581 | 1.56 |  | 1
| 1 | 2 | 1.6813 | 1.65 | 1.6762 | 1.67 | 1.45 | 1.75 | 1.6 | 1.72 | 1.72 | 1.7 | 1.649 | 1.74 | 1.68 | 1.6556 | 1.731 | 1.61 | 1.7112 | 1.65 | 1.64 | 1.66 | 1.7 | 1.79 | 1.65 | 1.6937 | 1.55 |  | 17
| 1 | 3 | 1.6561 | 1.67 | 1.6586 | 1.7 | 1.49 | 1.74 | 1.57 | 1.72 | 1.74 | 1.78 | 1.635 | 1.75 | 1.65 | 1.6558 | 1.721 | 1.62 | 1.7178 | 1.64 | 1.64 | 1.66 | 1.644 | 1.79 | 1.66 | 1.7015 | 1.55 |  | 16
| 1 | 4 | 1.664 | 1.68 | 1.6859 | 1.52 | 1.43 | 1.7 | 1.62 | 1.71 | 1.78 | 1.72 | 1.646 | 1.75 | 1.65 | 1.6457 | 1.695 | 1.61 | 1.7118 | 1.63 | 1.69 | 1.64 | 1.766 | 1.8 | 1.65 | 1.7079 | 1.56 |  | 1.6707
| 1 | 5 | 1.6781 | 1.64 | 1.6584 | 1.64 | 1.44 | 1.76 | 1.62 | 1.72 | 1.74 | 1.7 | 1.638 | 1.68 | 1.61 | 1.5914 | 1.632 | 1.58 | 1.7247 | 1.65 | 1.53 | 1.66 | 1.7582 | 1.78 | 1.66 | 1.7051 | 1.54 |  | 17
| 1 | 6 | 1.6796 | 1.65 | 1.6532 | 1.72 | 1.48 | 1.75 | 1.67 | 1.68 | 1.7 | 1.68 | 1.632 | 1.75 | 1.67 | 1.6612 | 1.761 | 1.65 | 1.7178 | 1.66 | 1.61 | 1.7 | 1.692 | 1.79 | 1.62 | 1.6846 | 1.54
| Mean |  | 1.6712 | 1.6583 | 1.67 | 1.635 | 1.4617 | 1.7417 | 1.62 | 1.7133 | 1.7383 | 1.71 | 1.64 | 1.74 | 1.655 | 1.6394 | 1.704 | 1.6167 | 1.7173 | 1.655 | 1.625 | 1.6617 | 1.7054 | 1.79 | 1.65 | 1.6918 | 1.55
| Median |  | 1.673 | 1.655 | 1.6674 | 1.655 | 1.465 | 1.75 | 1.62 | 1.72 | 1.74 | 1.7 | 1.639 | 1.75 | 1.66 | 1.6506 | 1.708 | 1.615 | 1.7178 | 1.65 | 1.64 | 1.66 | 1.696 | 1.79 | 1.655 | 1.6976 | 1.55
| Std Dev. |  | 0.010109 | 0.01472 | 0.015139 | 0.079436 | 0.024833 | 0.02137 | 0.034059 | 0.017512 | 0.027142 | 0.037417 | 0.0064807 | 0.030984 | 0.0251 | 0.026535 | 0.044569 | 0.023381 | 0.0051971 | 0.02429 | 0.053198 | 0.020412 | 0.048072 | 0.0063246 | 0.015492 | 0.018554 | 0.0089443
| Rel.Std.Dev. |  | 0.6049% | 0.8876% | 0.9065% | 4.858% | 1.699% | 1.227% | 2.102% | 1.022% | 1.561% | 2.188% | 0.3952% | 1.781% | 1.517% | 1.619% | 2.616% | 1.446% | 0.3026% | 1.468% | 3.274% | 1.228% | 2.819% | 0.3533% | 0.9389% | 1.097% | 0.577%
| PDM3 |  | 0.02657% | -0.7416% | -0.04495% | -2.138% | -12.51% | 4.246% | -3.036% | 2.55% | 4.047% | 2.351% | -1.839% | 4.147% | -0.9411% | -1.878% | 1.992% | -3.235% | 2.791% | -0.9411% | -2.737% | -0.542% | 2.074% | 7.139% | -1.24% | 1.263% | -7.226%
| Z-Score (Absolute) |  | 0.15 | 0.05 | 0.13 | 0.43 | 3.24 | 1.29 | 0.67 | 0.84 | 1.24 | 0.78 | 0.35 | 1.27 | 0.11 | 0.36 | 0.69 | 0.73 | 0.9 | 0.11 | 0.59 | 0 | 0.71 | 2.08 | 0.19 | 0.49 | 1.81
| Analytical results for Cd in OREAS 506b (Certified Value 0.4 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.36 | 0.33 | 0.35598 | 0.39 | < 1 | 0.4 | 0.34 | 0.36 | 0.39 | 0.34 | 0.41 | 0.46 | 0.49 | 0.37281 | < 1 | 0.43 | < 1 | 0.34 | 0.46 | 0.3 | 0.45 | 0.4 | 0.37 | 0.37 |  |  | 1
| 1 | 2 | 0.43 | 0.36 | 0.39674 | 0.5 | < 1 | 0.42 | 0.32 | 0.37 | 0.41 | 0.36 | 0.41 | 0.44 | 0.47 | 0.38543 | < 1 | 0.45 | < 1 | 0.34 | 0.4 | 0.4 | 0.44 | 0.4 | 0.36 | 0.37 |  |  | 18
| 1 | 3 | 0.4 | 0.32 | 0.35545 | 0.47 | < 1 | 0.39 | 0.35 | 0.36 | 0.38 | 0.35 | 0.41 | 0.46 | 0.45 | 0.39866 | < 1 | 0.51 | < 1 | 0.29 | 0.41 | 0.4 | 0.46 | 0.3 | 0.44 | 0.36 |  |  | 16
| 1 | 4 | 0.4 | 0.32 | 0.39039 | 0.38 | < 1 | 0.39 | 0.37 | 0.39 | 0.38 | 0.37 | 0.4 | 0.47 | 0.47 | 0.38594 | < 1 | 0.56 | < 1 | 0.32 | 0.46 | 0.4 | 0.43 | 0.3 | 0.43 | 0.37 |  |  | 0.39658
| 1 | 5 | 0.44 | 0.33 | 0.33659 | 0.48 | < 1 | 0.4 | 0.32 | 0.37 | 0.39 | 0.33 | 0.39 | 0.46 | 0.42 | 0.37665 | < 1 | 0.5 | < 1 | 0.26 | 0.48 | 0.4 | 0.44 | 0.4 | 0.39 | 0.34 |  |  | 18
| 1 | 6 | 0.37 | 0.34 | 0.39964 | 0.42 | < 1 | 0.34 | 0.34 | 0.39 | 0.35 | 0.37 | 0.37 | 0.5 | 0.45 | 0.36616 | < 1 | 0.52 | < 1 | 0.32 | 0.48 | 0.4 | 0.44 | < 0.3 | 0.39 | 0.36
| Mean |  | 0.4 | 0.33333 | 0.37247 | 0.44 |  | 0.39 | 0.34 | 0.37333 | 0.38333 | 0.35333 | 0.39833 | 0.465 | 0.45833 | 0.38094 |  | 0.495 |  | 0.31167 | 0.44833 | 0.38333 | 0.44333 | 0.36 | 0.39667 | 0.36167
| Median |  | 0.4 | 0.33 | 0.37319 | 0.445 |  | 0.395 | 0.34 | 0.37 | 0.385 | 0.355 | 0.405 | 0.46 | 0.46 | 0.38104 |  | 0.505 |  | 0.32 | 0.46 | 0.4 | 0.44 | 0.4 | 0.39 | 0.365
| Std Dev. |  | 0.031623 | 0.015055 | 0.026449 | 0.0502 |  | 0.026833 | 0.018974 | 0.013663 | 0.019664 | 0.01633 | 0.016021 | 0.019748 | 0.024014 | 0.011509 |  | 0.047645 |  | 0.031252 | 0.034881 | 0.040825 | 0.010328 | 0.054772 | 0.032042 | 0.01169
| Rel.Std.Dev. |  | 7.906% | 4.517% | 7.101% | 11.41% |  | 6.88% | 5.58% | 3.66% | 5.13% | 4.622% | 4.022% | 4.247% | 5.239% | 3.021% |  | 9.625% |  | 10.03% | 7.78% | 10.65% | 2.33% | 15.21% | 8.078% | 3.232%
| PDM3 |  | 0.8617% | -15.95% | -6.081% | 10.95% |  | -1.66% | -14.27% | -5.862% | -3.341% | -10.91% | 0.4414% | 17.25% | 15.57% | -3.944% |  | 24.82% |  | -21.41% | 13.05% | -3.341% | 11.79% | -9.224% | 0.02116% | -8.804%
| Z-Score (Absolute) |  | 0.04 | 0.96 | 0.38 | 0.64 | 1.54 | 0.11 | 0.86 | 0.36 | 0.21 | 0.66 | 0.01 | 1.01 | 0.91 | 0.25 | 1.54 | 1.46 | 1.54 | 1.29 | 0.76 | NA | 0.69 | NA | 0.01 | 0.54
Results from laboratory 2.23 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 506b (Certified Value 64 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | 2.27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 0
1
| 1 | 1 | 68.62 | 54 | 63.88 | 66.3 | 60.4 | 64 | 71.9 | 59.3 | 68.43 | 60.4 | 53.55 | 62.4219 | 63.3038 | 68 | 62.6 | 58.8 | 67 | 73 | 62.2 | 59.54 | [6 empty] | 1
| 1 | 2 | 70.97 | 60.6 | 69.29 | 63.7 | 61.1 | 70.5 | 70.7 | 62.5 | 66.07 | 60.5 | 53.35 | 61.9935 | 63.3276 | 66 | 62 | 58.43 | 68 | 70.8 | 67 | 58.88 | [6 empty] | 19
| 1 | 3 | 68.92 | 63.2 | 75.59 | 62 | 60.3 | 69.2 | 67.3 | 67.1 | 68.83 | 58.7 | 51.22 | 64.1979 | 62.9748 | 63 | 54.7 | 59.87 | 69 | 75.5 | 64.5 | 64.48 | [6 empty] | 16
| 1 | 4 | 70.32 | 63.2 | 67.53 | 61.8 | 63.5 | 70.2 | 69.6 | 63.6 | 64.5 | 58.1 | 52.79 | 64.7625 | 63.068 | 63 | 57.2 | 57.35 | 67 | 70.7 | 67.4 | 62.35 | [6 empty] | 63.8192
| 1 | 5 | 70.18 | 58.4 | 58.46 | 64 | 59 | 66.4 | 72.4 | 61.4 | 66.46 | 55.4 | 51.39 | 62.9675 | 63.254 | 66 | 57.2 | 57.88 | 70 | 68.5 | 69.2 | 59.14 | [6 empty] | 19
| 1 | 6 | 68.54 | 64.67 | 68.58 | 61.8 | 63 | 66.6 | 67 | 60.3 | 62.75 | 61.6 | 53.24 | 67.3099 | 63.239 | 63 | 60.6 | 57.49 | 66 | 71.3 | 67.4 | 64.34
| Mean |  | 69.5917 | 60.6783 | 67.2217 | 63.2667 | 61.2167 | 67.8167 | 69.8167 | 62.3667 | 66.1733 | 59.1167 | 52.59 | 63.9422 | 63.1945 | 64.8333 | 59.05 | 58.3033 | 67.8333 | 71.6333 | 66.2833 | 61.455
| Median |  | 69.55 | 61.9 | 68.055 | 62.85 | 60.75 | 67.9 | 70.15 | 61.95 | 66.265 | 59.55 | 53.015 | 63.5827 | 63.2465 | 64.5 | 58.9 | 58.155 | 67.5 | 71.05 | 67.2 | 60.945
| Std Dev. |  | 1.0274 | 3.961 | 5.7303 | 1.7795 | 1.7221 | 2.5631 | 2.2851 | 2.7768 | 2.3135 | 2.2248 | 1.0275 | 1.9564 | 0.14103 | 2.137 | 3.1456 | 0.94517 | 1.472 | 2.3796 | 2.5047 | 2.6051
| Rel.Std.Dev. |  | 1.476% | 6.528% | 8.524% | 2.813% | 2.813% | 3.78% | 3.273% | 4.452% | 3.496% | 3.763% | 1.954% | 3.06% | 0.2232% | 3.296% | 5.327% | 1.621% | 2.17% | 3.322% | 3.779% | 4.239%
| PDM3 |  | 9.045% | -4.921% | 5.331% | -0.8657% | -4.078% | 6.264% | 9.398% | -2.276% | 3.689% | -7.368% | -17.6% | 0.1928% | -0.9788% | 1.589% | -7.473% | -8.643% | 6.29% | 12.24% | 3.861% | -3.704%
| Z-Score (Absolute) |  | 1.23 | 0.6 | 0.75 | 0.07 | 0.49 | 0.87 | 1.28 | 0.26 | 0.53 | 0.92 | 2.27 | 0.07 | 0.08 | 0.25 | 0.94 | 1.09 | 0.87 | 1.65 | 0.55 | 0.44
| Analytical results for Co in OREAS 506b (Certified Value 7.34 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 7.7 | 7 | 7.3679 | 6.8 | 8.85 | 7.4 | 6.8 | 7.3 | 7.8 | 6.5 | 7.8 | 8.2 | 7.7 | 7.4096 | 7 | 7 | 7.3559 | 7 | 6.6 | 7.6 | 7.2 | 7.3 | 7.3 | 7 | 6.1 | 7.4 | 1
| 1 | 2 | 7.9 | 7 | 7.2238 | 7.3 | 9.9 | 7.4 | 6.7 | 7.4 | 7.7 | 6.8 | 7.6 | 8.3 | 7.7 | 7.4184 | 7 | 7.24 | 7.7116 | 7 | 6.6 | 7.4 | 7.5 | 7.5 | 7.4 | 7 | 6.8 | 7.1 | 20
| 1 | 3 | 7.8 | 7 | 7.422 | 7.3 | 10.09 | 7.4 | 6.9 | 7.2 | 7.6 | 7.1 | 7.8 | 8.1 | 7.5 | 7.445 | 7 | 7.35 | 7.4991 | 7 | 6.1 | 7.9 | 7.5 | 7.3 | 7.6 | 7 | 7.1 | 7.3 | 16
| 1 | 4 | 7.7 | 7 | 7.2629 | 6.9 | 9.84 | 7.4 | 7.3 | 7.3 | 7.8 | 6.7 | 7.7 | 8 | 7.7 | 7.3341 | 7 | 7.39 | 7.2265 | 6 | 6.5 | 8.6 | 7.6 | 7.4 | 7.2 | 7 | 6.5 | 7.1 | 7.3353
| 1 | 5 | 7.8 | 7 | 7.2131 | 6.9 | 10.14 | 7.8 | 6.8 | 7 | 7.7 | 6.9 | 7.8 | 7.8 | 7.4 | 7.1715 | 7 | 7.05 | 7.5419 | 7 | 6.1 | 7.7 | 7.3 | 7.2 | 7.4 | 7 | 7.3 | 7.3 | 20
| 1 | 6 | 7.8 | 7 | 7.1384 | 7 | 9.89 | 7.1 | 7 | 7 | 7.7 | 6.7 | 7.5 | 8.3 | 7.8 | 7.5459 | 7 | 7.47 | 7.7188 | 7 | 6.3 | 8.2 | 7.5 | 7.2 | 7.4 | 7 | 6.7 | 7.3
| Mean |  | 7.7833 | 7 | 7.2713 | 7.0333 | 9.785 | 7.4167 | 6.9167 | 7.2 | 7.7167 | 6.7833 | 7.7 | 8.1167 | 7.6333 | 7.3874 | 7 | 7.25 | 7.509 | 6.8333 | 6.3667 | 7.9 | 7.4333 | 7.3167 | 7.3833 | 7 | 6.75 | 7.25
| Median |  | 7.8 | 7 | 7.2433 | 6.95 | 9.895 | 7.4 | 6.85 | 7.25 | 7.7 | 6.75 | 7.75 | 8.15 | 7.7 | 7.414 | 7 | 7.295 | 7.5205 | 7 | 6.4 | 7.8 | 7.5 | 7.3 | 7.4 | 7 | 6.75 | 7.3
| Std Dev. |  | 0.075277 | 0 | 0.10527 | 0.21602 | 0.4734 | 0.22286 | 0.2137 | 0.16733 | 0.075277 | 0.20412 | 0.12649 | 0.19408 | 0.15055 | 0.12598 | 0 | 0.19005 | 0.19467 | 0.40825 | 0.23381 | 0.43818 | 0.15055 | 0.1169 | 0.13292 | 0 | 0.42778 | 0.12247
| Rel.Std.Dev. |  | 0.9672% | 0% | 1.448% | 3.071% | 4.838% | 3.005% | 3.09% | 2.324% | 0.9755% | 3.009% | 1.643% | 2.391% | 1.972% | 1.705% | 0% | 2.621% | 2.592% | 5.974% | 3.672% | 5.547% | 2.025% | 1.598% | 1.8% | 0% | 6.338% | 1.689%
| PDM3 |  | 6.108% | -4.571% | -0.8717% | -4.116% | 33.4% | 1.109% | -5.707% | -1.844% | 5.199% | -7.525% | 4.972% | 10.65% | 4.063% | 0.7104% | -4.571% | -1.163% | 2.368% | -6.843% | -13.2% | 7.699% | 1.337% | -0.2539% | 0.655% | -4.571% | -7.979% | -1.163%
| Z-Score (Absolute) |  | 0.95 | NA | 0.27 | 0.84 | 5.75 | 0.08 | 1.12 | 0.44 | 0.79 | 1.44 | 0.75 | 1.75 | 0.59 | 0 | NA | 0.32 | 0.3 | NA | 2.44 | 1.23 | 0.11 | 0.16 | 0 | NA | 1.52 | 0.32
Results from laboratories 2.02, 2.16, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 506b (Certified Value 39.5 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
1
| 1 | 1 | 43 | 41 | 38.9666 | 35 | 63.98 | 38 | 41 | 39 | 41 | 38 | 31 | 36 | 45 | 39.3183 | 41 | 37.1 | 41 | 36 | 43 | 37.1 | 28 | 43 | 36 |  |  |  | 1
| 1 | 2 | 44 | 41 | 49.7491 | 36 | 76.5 | 41 | 40 | 38 | 40 | 39 | 30 | 37 | 47 | 38.2543 | 40 | 40.1 | 41 | 37 | 44 | 38 | 28 | 46 | 33 |  |  |  | 21
| 1 | 3 | 45 | 42 | 47.4897 | 35 | 71.01 | 40 | 40 | 38 | 39 | 40 | 32 | 34 | 45 | 40.0477 | 40 | 40.2 | 43 | 38 | 44 | 36.4 | 29 | 41 | 35 |  |  |  | 16
| 1 | 4 | 42 | 42 | 45.0757 | 36 | 81.06 | 38 | 40 | 38 | 40 | 41 | 29 | 34 | 45 | 39.2319 | 40 | 40.1 | 40 | 40 | 44 | 37 | 27 | 42 | 33 |  |  |  | 39.5348
| 1 | 5 | 45 | 40 | 46.3611 | 33 | 82.33 | 38 | 40 | 36 | 41 | 40 | 27 | 33 | 45 | 37.6303 | 41 | 40.3 | 41 | 34 | 43 | 36 | 28 | 48 | 35 |  |  |  | 21
| 1 | 6 | 43 | 41 | 37.6099 | 34 | 70.98 | 40 | 42 | 36 | 38 | 39 | 31 | 40 | 46 | 38.7136 | 41 | 42 | 46 | 40 | 42 | 36.4 | 28 | 49 | 35
| Mean |  | 43.6667 | 41.1667 | 44.2087 | 34.8333 | 74.31 | 39.1667 | 40.5 | 37.5 | 39.8333 | 39.5 | 30 | 35.6667 | 45.5 | 38.866 | 40.5 | 39.9667 | 42 | 37.5 | 43.3333 | 36.8167 | 28 | 44.8333 | 34.5
| Median |  | 43.5 | 41 | 45.7184 | 35 | 73.755 | 39 | 40 | 38 | 40 | 39.5 | 30.5 | 35 | 45 | 38.9727 | 40.5 | 40.15 | 41 | 37.5 | 43.5 | 36.7 | 28 | 44.5 | 35
| Std Dev. |  | 1.2111 | 0.75277 | 4.8553 | 1.169 | 6.9766 | 1.3292 | 0.83666 | 1.2247 | 1.169 | 1.0488 | 1.7889 | 2.582 | 0.83666 | 0.85546 | 0.54772 | 1.5845 | 2.1909 | 2.3452 | 0.8165 | 0.7111 | 0.63246 | 3.3116 | 1.2247
| Rel.Std.Dev. |  | 2.773% | 1.829% | 10.98% | 3.356% | 9.389% | 3.394% | 2.066% | 3.266% | 2.935% | 2.655% | 5.963% | 7.239% | 1.839% | 2.201% | 1.352% | 3.965% | 5.216% | 6.254% | 1.884% | 1.931% | 2.259% | 7.386% | 3.55%
| PDM3 |  | 10.45% | 4.128% | 11.82% | -11.89% | 87.96% | -0.9312% | 2.441% | -5.147% | 0.755% | -0.08809% | -24.12% | -9.784% | 15.09% | -1.692% | 2.441% | 1.092% | 6.235% | -5.147% | 9.608% | -6.875% | -29.18% | 13.4% | -12.74%
| Z-Score (Absolute) |  | 0.86 | 0.3 | 0.98 | 1.12 | 7.71 | 0.15 | 0.15 | 0.52 | 0 | 0.07 | 2.2 | 0.93 | 1.27 | 0.22 | 0.15 | 0.03 | 0.48 | 0.52 | 0.78 | 0.67 | 2.65 | 1.12 | 1.19
| Analytical results for Cs in OREAS 506b (Certified Value 8.64 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.24 | 2.25 | 2.27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 8.9 | 8.6 | 8.5771 | 7.9 | 8.49 | 8.53 | 9.04 | 8.51 | 8.13 | 8.95 | 8.85 | 8.62 | 8.2851 | 8.44 | 7.0837 | 8.8 | 8.8 | 9 | 8.41 | 8.3 | [6 empty] | 1
| 1 | 2 | 9.2 | 8.6 | 8.3711 | 8.7 | 8.51 | 8.38 | 9.25 | 8.45 | 8.6 | 8.62 | 8.88 | 8.44 | 8.3445 | 9.16 | 7.0568 | 8.9 | 8.47 | 9 | 8.58 | 8.1 | [6 empty] | 59
| 1 | 3 | 8.96 | 8.5 | 8.6063 | 8.5 | 8.49 | 8.3 | 9.3 | 8.4 | 8.87 | 8.95 | 8.71 | 8.23 | 8.3118 | 8.94 | 6.9959 | 8.6 | 8.8 | 9 | 8.48 | 8.7 | [6 empty] | 16
| 1 | 4 | 9.04 | 8.7 | 8.6756 | 5.8 | 8.33 | 8.89 | 9.17 | 8.71 | 8.5 | 8.83 | 8.58 | 8.59 | 8.6644 | 8.96 | 7.0451 | 9 | 8.95 | 8 | 8.57 | 8.2 | [6 empty] | 8.6361
| 1 | 5 | 8.96 | 8.6 | 8.406 | 9 | 8.68 | 8.38 | 9 | 8.62 | 8.84 | 8.69 | 8.33 | 8.21 | 8.2834 | 8.48 | 7.0837 | 9.1 | 8.43 | 9 | 8.77 | 8.2 | [6 empty] | 22
| 1 | 6 | 9.13 | 8.3 | 8.3633 | 6.5 | 8.17 | 8.51 | 9.13 | 8.35 | 8.52 | 8.66 | 8.89 | 8.51 | 8.4217 | 9.11 | 7.0879 | 8.9 | 8.59 | 9 | 8.66 | 8.6
| Mean |  | 9.0317 | 8.55 | 8.4999 | 7.7333 | 8.445 | 8.4983 | 9.1483 | 8.5067 | 8.5767 | 8.7833 | 8.7067 | 8.4333 | 8.3851 | 8.8483 | 7.0589 | 8.8833 | 8.6733 | 8.8333 | 8.5783 | 8.35
| Median |  | 9 | 8.6 | 8.4916 | 8.2 | 8.49 | 8.445 | 9.15 | 8.48 | 8.56 | 8.76 | 8.78 | 8.475 | 8.3282 | 8.95 | 7.0703 | 8.9 | 8.695 | 9 | 8.575 | 8.25
| Std Dev. |  | 0.11462 | 0.13784 | 0.13579 | 1.2972 | 0.17456 | 0.21066 | 0.11652 | 0.13663 | 0.27001 | 0.1472 | 0.22006 | 0.17694 | 0.14609 | 0.31269 | 0.035299 | 0.17224 | 0.20791 | 0.40825 | 0.12766 | 0.2429
| Rel.Std.Dev. |  | 1.269% | 1.612% | 1.598% | 16.77% | 2.067% | 2.479% | 1.274% | 1.606% | 3.148% | 1.676% | 2.527% | 2.098% | 1.742% | 3.534% | 0.5001% | 1.939% | 2.397% | 4.622% | 1.488% | 2.909%
| PDM3 |  | 4.58% | -0.9973% | -1.577% | -10.45% | -2.213% | -1.596% | 5.931% | -1.499% | -0.6885% | 1.705% | 0.8168% | -2.348% | -2.906% | 2.457% | -18.26% | 2.862% | 0.4308% | 2.283% | -0.6692% | -3.313%
| Z-Score (Absolute) |  | 2.07 | 0 | 0.22 | 3.52 | 0.45 | 0.22 | 2.58 | 0.19 | 0.11 | 1 | 0.67 | 0.5 | 0.71 | 1.28 | 6.42 | 1.43 | 0.53 | NA | 0.12 | 0.86
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 506b (Certified Value 0.446 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.44097 | 0.434 | 0.44048 | 0.48012 | 0.4553 | 0.456 | 0.441 | 0.455 | 0.47 | 0.428 | 0.447 | 0.481 | 0.4567 | 0.4429 | 0.4281 | 0.4251 | 0.46345 | 0.448 | 0.4371 | 0.4164 | 0.4447 | 0.444 | 0.4338 | 0.455 | 0.43342 | 0.43787 | 1
| 1 | 2 | 0.45006 | 0.435 | 0.43322 | 0.47609 | 0.4506 | 0.462 | 0.448 | 0.451 | 0.463 | 0.435 | 0.4517 | 0.482 | 0.4543 | 0.44532 | 0.4376 | 0.4383 | 0.4625 | 0.4382 | 0.4451 | 0.4245 | 0.4449 | 0.446 | 0.4344 | 0.453 | 0.44114 | 0.43148 | 22
| 1 | 3 | 0.4376 | 0.435 | 0.43869 | 0.47647 | 0.4558 | 0.456 | 0.442 | 0.45 | 0.474 | 0.451 | 0.4478 | 0.474 | 0.4504 | 0.44303 | 0.4345 | 0.4337 | 0.46226 | 0.4382 | 0.4259 | 0.4187 | 0.444 | 0.442 | 0.4358 | 0.457 | 0.44243 | 0.4313 | 16
| 1 | 4 | 0.44248 | 0.438 | 0.44113 | 0.44844 | 0.451 | 0.471 | 0.44 | 0.446 | 0.481 | 0.436 | 0.4484 | 0.464 | 0.4416 | 0.44456 | 0.4283 | 0.4449 | 0.46817 | 0.4417 | 0.4362 | 0.4005 | 0.4433 | 0.44809 | 0.4326 | 0.442 | 0.4448 | 0.43292 | 0.44603
| 1 | 5 | 0.44534 | 0.437 | 0.43103 | 0.49486 | 0.4579 | 0.471 | 0.444 | 0.447 | 0.472 | 0.433 | 0.4485 | 0.454 | 0.4519 | 0.43555 | 0.4121 | 0.4362 | 0.46435 | 0.4415 | 0.4317 | 0.3842 | 0.4445 | 0.44352 | 0.4304 | 0.454 | 0.44309 | 0.42962 | 23
| 1 | 6 | 0.44582 | 0.433 | 0.43025 | 0.44734 | 0.459 | 0.452 | 0.441 | 0.444 | 0.471 | 0.433 | 0.4501 | 0.485 | 0.4436 | 0.4502 | 0.4443 | 0.4463 | 0.46925 | 0.4466 | 0.4434 | 0.4635 | 0.4447 | 0.446 | 0.439 | 0.454 | 0.44025 | 0.42418
| Mean |  | 0.44371 | 0.43533 | 0.4358 | 0.47055 | 0.45493 | 0.46133 | 0.44267 | 0.44883 | 0.47183 | 0.436 | 0.44892 | 0.47333 | 0.44975 | 0.44359 | 0.43082 | 0.43742 | 0.465 | 0.44237 | 0.43657 | 0.41797 | 0.44435 | 0.44494 | 0.43433 | 0.4525 | 0.44085 | 0.43123
| Median |  | 0.44391 | 0.435 | 0.43596 | 0.47628 | 0.45555 | 0.459 | 0.4415 | 0.4485 | 0.4715 | 0.434 | 0.44845 | 0.4775 | 0.45115 | 0.4438 | 0.4314 | 0.43725 | 0.4639 | 0.4416 | 0.43665 | 0.41755 | 0.4446 | 0.445 | 0.4341 | 0.454 | 0.44178 | 0.43139
| Std Dev. |  | 0.0043309 | 0.0018619 | 0.0048759 | 0.018846 | 0.0034777 | 0.0081404 | 0.0029439 | 0.0039707 | 0.0058452 | 0.0078486 | 0.0017034 | 0.012094 | 0.0059715 | 0.0047547 | 0.011004 | 0.0077806 | 0.0029883 | 0.004137 | 0.0071749 | 0.026699 | 0.00059917 | 0.0021778 | 0.0029221 | 0.0053198 | 0.003969 | 0.0044566
| Rel.Std.Dev. |  | 0.9761% | 0.4277% | 1.119% | 4.005% | 0.7645% | 1.765% | 0.665% | 0.8847% | 1.239% | 1.8% | 0.3795% | 2.555% | 1.328% | 1.072% | 2.554% | 1.779% | 0.6427% | 0.9352% | 1.643% | 6.388% | 0.1348% | 0.4895% | 0.6728% | 1.176% | 0.9003% | 1.033%
| PDM3 |  | -0.5201% | -2.398% | -2.294% | 5.498% | 1.996% | 3.431% | -0.7544% | 0.6282% | 5.785% | -2.249% | 0.6469% | 6.121% | 0.8337% | -0.5465% | -3.411% | -1.931% | 4.252% | -0.8216% | -2.122% | -6.292% | -0.377% | -0.2458% | -2.623% | 1.45% | -1.161% | -3.319%
| Z-Score (Absolute) |  | 0.01 | 0.72 | 0.68 | 2.34 | 0.98 | 1.54 | 0.09 | 0.45 | 2.45 | 0.67 | 0.46 | 2.58 | 0.53 | 0.01 | 1.12 | 0.54 | 1.86 | 0.11 | 0.62 | 2.23 | 0.06 | 0.11 | 0.81 | 0.77 | 0.24 | 1.08
| Analytical results for Dy in OREAS 506b (Certified Value 3.21 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS | 2
3
| 1 | 1 | 2.01: 3.37; 2.02: 3.2; 2.03: 3.2336; 2.04: 3.1118; 2.06: 3.33; 2.07: 3.09; 2.08: 3.7; 2.09: 3.2 | 1
| 1 | 2 | 2.01: 3.45; 2.02: 3.4; 2.03: 3.2157; 2.04: 3.096; 2.06: 3.21; 2.07: 3.07; 2.08: 3.8; 2.09: 3 | 23
| 1 | 3 | 2.01: 3.42; 2.02: 3.5; 2.03: 3.2846; 2.04: 3.0093; 2.06: 3.18; 2.07: 3.38; 2.08: 3.6; 2.09: 3.3 | 16
| 1 | 4 | 2.01: 3.36; 2.02: 2.9; 2.03: 3.2533; 2.04: 3.0422; 2.06: 3.26; 2.07: 3.25; 2.08: 3.7; 2.09: 3.1 | 3.2142
| 1 | 5 | 2.01: 3.37; 2.02: 3.6; 2.03: 3.1786; 2.04: 3.1321; 2.06: 3.13; 2.07: 3.06; 2.08: 3.8; 2.09: 3.1 | 24
| 1 | 6 | 2.01: 3.44; 2.02: 3.1; 2.03: 3.3797; 2.04: 3.1143; 2.06: 3.22; 2.07: 3.06; 2.08: 3.8; 2.09: 3.3
| Mean |  | 2.01: 3.4017; 2.02: 3.2833; 2.03: 3.2576; 2.04: 3.0843; 2.06: 3.2217; 2.07: 3.1517; 2.08: 3.7333; 2.09: 3.1667
| Median |  | 2.01: 3.395; 2.02: 3.3; 2.03: 3.2435; 2.04: 3.1039; 2.06: 3.215; 2.07: 3.08; 2.08: 3.75; 2.09: 3.15
| Std Dev. |  | 2.01: 0.039707; 2.02: 0.26394; 2.03: 0.069609; 2.04: 0.047929; 2.06: 0.068532; 2.07: 0.13348; 2.08: 0.08165; 2.09: 0.12111
| Rel.Std.Dev. |  | 2.01: 1.167%; 2.02: 8.039%; 2.03: 2.137%; 2.04: 1.554%; 2.06: 2.127%; 2.07: 4.235%; 2.08: 2.187%; 2.09: 3.824%
| PDM3 |  | 2.01: 5.831%; 2.02: 2.15%; 2.03: 1.348%; 2.04: -4.043%; 2.06: 0.231%; 2.07: -1.947%; 2.08: 16.15%; 2.09: -1.48%
| Z-Score (Absolute) |  | 2.01: 1.73; 2.02: NA; 2.03: 0.35; 2.04: 1.32; 2.06: 0; 2.07: 0.67; 2.08: NA; 2.09: NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 506b (Certified Value 1.33 ppm) | [27 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS | 2
3
| 1 | 1 | 2.01: 1.38; 2.02: 1.2; 2.03: 1.3032; 2.04: 1.3282; 2.06: 1.46; 2.07: 1.15; 2.08: 1.6; 2.09: 1.2 | 1
| 1 | 2 | 2.01: 1.37; 2.02: 1.4; 2.03: 1.3102; 2.04: 1.4003; 2.06: 1.38; 2.07: 1.16; 2.08: 1.8; 2.09: 1.2 | 24
| 1 | 3 | 2.01: 1.39; 2.02: 1.4; 2.03: 1.296; 2.04: 1.3041; 2.06: 1.4; 2.07: 1.25; 2.08: 1.5; 2.09: 1.3 | 16
| 1 | 4 | 2.01: 1.43; 2.02: 1.2; 2.03: 1.3864; 2.04: 1.3306; 2.06: 1.52; 2.07: 1.17; 2.08: 1.7; 2.09: 1.2 | 1.3315
| 1 | 5 | 2.01: 1.44; 2.02: 1.4; 2.03: 1.3194; 2.04: 1.3474; 2.06: 1.39; 2.07: 1.15; 2.08: 1.6; 2.09: 1.2 | 25
| 1 | 6 | 2.01: 1.37; 2.02: 1.3; 2.03: 1.3143; 2.04: 1.3658; 2.06: 1.36; 2.07: 1.17; 2.08: 1.5; 2.09: 1.3
| Mean |  | 2.01: 1.3967; 2.02: 1.3167; 2.03: 1.3216; 2.04: 1.3461; 2.06: 1.4183; 2.07: 1.175; 2.08: 1.6167; 2.09: 1.2333
| Median |  | 2.01: 1.385; 2.02: 1.35; 2.03: 1.3123; 2.04: 1.339; 2.06: 1.395; 2.07: 1.165; 2.08: 1.6; 2.09: 1.2
| Std Dev. |  | 2.01: 0.030768; 2.02: 0.098319; 2.03: 0.032789; 2.04: 0.033641; 2.06: 0.060139; 2.07: 0.037815; 2.08: 0.1169; 2.09: 0.05164
| Rel.Std.Dev. |  | 2.01: 2.203%; 2.02: 7.467%; 2.03: 2.481%; 2.04: 2.499%; 2.06: 4.24%; 2.07: 3.218%; 2.08: 7.231%; 2.09: 4.187%
| PDM3 |  | 2.01: 4.892%; 2.02: -1.116%; 2.03: -0.7467%; 2.04: 1.092%; 2.06: 6.519%; 2.07: -11.76%; 2.08: 21.41%; 2.09: -7.375%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: NA; 2.03: 0.33; 2.04: 0; 2.06: 0.96; 2.07: 2.28; 2.08: NA; 2.09: NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 506b (Certified Value 1.26 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS | 2
3
| 1 | 1 | 2.01: 1.27; 2.02: 1.3; 2.03: 1.2325; 2.04: 1.4348; 2.06: 1.26; 2.07: 1.3; 2.08: 1.18; 2.09: 1.1 | 1
| 1 | 2 | 2.01: 1.26; 2.02: 1.4; 2.03: 1.237; 2.04: 1.4394; 2.06: 1.26; 2.07: 1.24; 2.08: 1.27; 2.09: 1 | 25
| 1 | 3 | 2.01: 1.27; 2.02: 1.5; 2.03: 1.2019; 2.04: 1.4417; 2.06: 1.24; 2.07: 1.31; 2.08: 1.24; 2.09: 1.1 | 16
| 1 | 4 | 2.01: 1.25; 2.02: 1.1; 2.03: 1.2842; 2.04: 1.4399; 2.06: 1.29; 2.07: 1.26; 2.08: 1.28; 2.09: 1.1 | 1.2583
| 1 | 5 | 2.01: 1.29; 2.02: 1.5; 2.03: 1.2045; 2.04: 1.387; 2.06: 1.31; 2.07: 1.28; 2.08: 1.22; 2.09: 1.1 | 26
| 1 | 6 | 2.01: 1.31; 2.02: 1.2; 2.03: 1.2492; 2.04: 1.4569; 2.06: 1.25; 2.07: 1.27; 2.08: 1.23; 2.09: 1.1
| Mean |  | 2.01: 1.275; 2.02: 1.3333; 2.03: 1.2349; 2.04: 1.4333; 2.06: 1.2683; 2.07: 1.2767; 2.08: 1.2367; 2.09: 1.0833
| Median |  | 2.01: 1.27; 2.02: 1.35; 2.03: 1.2347; 2.04: 1.4397; 2.06: 1.26; 2.07: 1.275; 2.08: 1.235; 2.09: 1.1
| Std Dev. |  | 2.01: 0.021679; 2.02: 0.1633; 2.03: 0.030537; 2.04: 0.023888; 2.06: 0.026394; 2.07: 0.02582; 2.08: 0.036148; 2.09: 0.040825
| Rel.Std.Dev. |  | 2.01: 1.7%; 2.02: 12.25%; 2.03: 2.473%; 2.04: 1.667%; 2.06: 2.081%; 2.07: 2.022%; 2.08: 2.923%; 2.09: 3.768%
| PDM3 |  | 2.01: 1.326%; 2.02: 5.962%; 2.03: -1.862%; 2.04: 13.91%; 2.06: 0.7966%; 2.07: 1.459%; 2.08: -1.72%; 2.09: -13.91%
| Z-Score (Absolute) |  | 2.01: 0.11; 2.02: NA; 2.03: 1.24; 2.04: 5.45; 2.06: 0.11; 2.07: 0.17; 2.08: 1.18; 2.09: NA
Results from laboratories 2.04 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 506b (Certified Value 2.75 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.77 | 2.81 | 2.8123 | 2.68 | 2.44 | 2.73 | 2.63 | 2.73 | 2.91 | 2.73 | 2.71 | 3.04 | 2.75 | 2.6906 | 2.73 | 2.69 | 2.8283 | 2.81 | 2.49 | 2.2 | 2.81 | 2.808 | 2.83 | 2.69 | 2.6322 | 2.62 | 1
| 1 | 2 | 2.79 | 2.82 | 2.757 | 2.82 | 2.43 | 2.77 | 2.62 | 2.74 | 2.85 | 2.75 | 2.73 | 3.05 | 2.71 | 2.7135 | 2.81 | 2.66 | 2.8259 | 2.74 | 2.52 | 2.59 | 2.88 | 2.788 | 2.82 | 2.79 | 2.6858 | 2.59 | 26
| 1 | 3 | 2.75 | 2.81 | 2.7338 | 2.82 | 2.46 | 2.74 | 2.58 | 2.75 | 2.89 | 2.84 | 2.71 | 3.01 | 2.74 | 2.7171 | 2.8 | 2.67 | 2.8291 | 2.71 | 2.48 | 2.56 | 2.89 | 2.824 | 2.87 | 2.7 | 2.6921 | 2.6 | 16
| 1 | 4 | 2.76 | 2.82 | 2.8131 | 2.63 | 2.43 | 2.71 | 2.7 | 2.72 | 2.95 | 2.77 | 2.73 | 2.98 | 2.71 | 2.6961 | 2.75 | 2.66 | 2.8298 | 2.69 | 2.6 | 2.43 | 2.84 | 3 | 2.8 | 2.78 | 2.7294 | 2.62 | 2.7484
| 1 | 5 | 2.79 | 2.75 | 2.7391 | 2.75 | 2.45 | 2.76 | 2.63 | 2.7 | 2.89 | 2.75 | 2.71 | 2.88 | 2.72 | 2.628 | 2.63 | 2.62 | 2.8284 | 2.73 | 2.28 | 2.31 | 2.78 | 2.8099 | 2.81 | 2.78 | 2.7132 | 2.61 | 27
| 1 | 6 | 2.79 | 2.78 | 2.7138 | 2.72 | 2.46 | 2.74 | 2.65 | 2.68 | 2.87 | 2.73 | 2.72 | 3.07 | 2.7 | 2.7409 | 2.85 | 2.72 | 2.8236 | 2.76 | 2.45 | 2.24 | 2.9 | 2.74 | 2.85 | 2.8 | 2.6761 | 2.57
| Mean |  | 2.775 | 2.7983 | 2.7615 | 2.7367 | 2.445 | 2.7417 | 2.635 | 2.72 | 2.8933 | 2.7617 | 2.7183 | 3.005 | 2.7217 | 2.6977 | 2.7617 | 2.67 | 2.8275 | 2.74 | 2.47 | 2.3883 | 2.85 | 2.8283 | 2.83 | 2.7567 | 2.6881 | 2.6017
| Median |  | 2.78 | 2.81 | 2.7481 | 2.735 | 2.445 | 2.74 | 2.63 | 2.725 | 2.89 | 2.75 | 2.715 | 3.025 | 2.715 | 2.7048 | 2.775 | 2.665 | 2.8283 | 2.735 | 2.485 | 2.37 | 2.86 | 2.8089 | 2.825 | 2.78 | 2.689 | 2.605
| Std Dev. |  | 0.017607 | 0.027869 | 0.041957 | 0.07607 | 0.013784 | 0.02137 | 0.03937 | 0.026077 | 0.034448 | 0.041191 | 0.0098319 | 0.06892 | 0.019408 | 0.038486 | 0.077567 | 0.033466 | 0.0023129 | 0.041952 | 0.10621 | 0.16461 | 0.048166 | 0.089074 | 0.026077 | 0.048442 | 0.033543 | 0.019408
| Rel.Std.Dev. |  | 0.6345% | 0.9959% | 1.519% | 2.78% | 0.5638% | 0.7794% | 1.494% | 0.9587% | 1.191% | 1.492% | 0.3617% | 2.294% | 0.7131% | 1.427% | 2.809% | 1.253% | 0.0818% | 1.531% | 4.3% | 6.892% | 1.69% | 3.149% | 0.9214% | 1.757% | 1.248% | 0.746%
| PDM3 |  | 0.9678% | 1.817% | 0.4774% | -0.427% | -11.04% | -0.2451% | -4.126% | -1.033% | 5.273% | 0.4826% | -1.094% | 9.336% | -0.9728% | -1.844% | 0.4826% | -2.853% | 2.878% | -0.3057% | -10.13% | -13.1% | 3.697% | 2.908% | 2.969% | 0.3007% | -2.193% | -5.339%
| Z-Score (Absolute) |  | 0.42 | 0.7 | 0.25 | 0.05 | 3.62 | 0.01 | 1.3 | 0.25 | 1.87 | 0.25 | 0.28 | 3.23 | 0.23 | 0.53 | 0.25 | 0.87 | 1.06 | 0.01 | 3.31 | 4.31 | 1.34 | 1.07 | 1.09 | 0.19 | 0.64 | 1.7
| Analytical results for Ga in OREAS 506b (Certified Value 19.3 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 18.91 | 19.4 | 19.902 | 18.32 | 17.97 | 18.25 | 19.9 | 19.55 | 20 | 18 | 19.76 | 25.1 | 19.4 | 19.5674 | 19.24 | 19 | 18.7 | 19.8 | 18.3 | 19.7 | 16 | 18.53 |  |  |  |  | 1
| 1 | 2 | 19.57 | 19.5 | 19.8848 | 19.68 | 18.49 | 18.45 | 19.05 | 20.2 | 19.6 | 19.3 | 19.82 | 26 | 19.8 | 19.2979 | 20.56 | 20 | 19.7 | 20.1 | 18.4 | 20.1 | 18 | 18.11 |  |  |  |  | 27
| 1 | 3 | 19.02 | 19.6 | 19.9566 | 19.76 | 18.09 | 18.35 | 19 | 19.65 | 19.2 | 19.9 | 19.66 | 26.7 | 18.6 | 19.6716 | 20.78 | 20 | 19.4 | 20.1 | 19.7 | 20.3 | 19 | 18.08 |  |  |  |  | 16
| 1 | 4 | 19.49 | 19.5 | 20.1261 | 17.95 | 19.02 | 18 | 20.1 | 19.35 | 20 | 18.55 | 19.78 | 26.5 | 19.2 | 19.5528 | 20.92 | 20 | 19.3 | 19.8 | 17.9 | 20.3 | 17 | 18 |  |  |  |  | 19.3091
| 1 | 5 | 19.68 | 19.4 | 18.6253 | 18.8 | 18.89 | 18.7 | 19.1 | 19.75 | 20 | 19.35 | 18.95 | 24.4 | 18.9 | 19.0692 | 19.39 | 20 | 19.3 | 20 | 18.6 | 22.1 | 19 | 18.23 |  |  |  |  | 28
| 1 | 6 | 19.51 | 19.9 | 18.8304 | 18.27 | 18.09 | 17.7 | 19.05 | 19.6 | 19.55 | 18.9 | 18.6 | 27.7 | 19.1 | 20.1287 | 20.77 | 20 | 18.9 | 20.1 | 18.9 | 22.1 | 18 | 18.67
| Mean |  | 19.3633 | 19.55 | 19.5542 | 18.7967 | 18.425 | 18.2417 | 19.3667 | 19.6833 | 19.725 | 19 | 19.4283 | 26.0667 | 19.1667 | 19.5479 | 20.2767 | 19.8333 | 19.2167 | 19.9833 | 18.6333 | 20.7667 | 17.8333 | 18.27
| Median |  | 19.5 | 19.5 | 19.8934 | 18.56 | 18.29 | 18.3 | 19.075 | 19.625 | 19.8 | 19.1 | 19.71 | 26.25 | 19.15 | 19.5601 | 20.665 | 20 | 19.3 | 20.05 | 18.5 | 20.3 | 18 | 18.17
| Std Dev. |  | 0.31747 | 0.18708 | 0.64899 | 0.76547 | 0.44854 | 0.35131 | 0.49565 | 0.28577 | 0.33129 | 0.66858 | 0.52071 | 1.181 | 0.41312 | 0.3591 | 0.75521 | 0.40825 | 0.36009 | 0.1472 | 0.6186 | 1.0558 | 1.169 | 0.26974
| Rel.Std.Dev. |  | 1.64% | 0.9569% | 3.319% | 4.072% | 2.434% | 1.926% | 2.559% | 1.452% | 1.68% | 3.519% | 2.68% | 4.531% | 2.155% | 1.837% | 3.725% | 2.058% | 1.874% | 0.7366% | 3.32% | 5.084% | 6.555% | 1.476%
| PDM3 |  | 0.2811% | 1.248% | 1.27% | -2.654% | -4.578% | -5.528% | 0.2984% | 1.938% | 2.154% | -1.601% | 0.6177% | 35% | -0.7374% | 1.237% | 5.011% | 2.715% | -0.4785% | 3.492% | -3.5% | 7.549% | -7.643% | -5.381%
| Z-Score (Absolute) |  | 0.06 | 0.25 | 0.25 | 0.97 | 1.57 | 1.87 | 0.05 | 0.46 | 0.53 | 0.64 | 0.05 | 10.79 | 0.37 | 0.24 | 1.42 | 0.71 | 0.29 | 0.95 | 1.24 | 2.22 | 2.53 | 1.82
| Analytical results for Gd in OREAS 506b (Certified Value 5.13 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 5.47; 2.02: 4.8; 2.03: 5.1363; 2.04: 5.4814; 2.06: 5.22; 2.07: 4.9; 2.08: 4.8; 2.09: 4.8 |  |  |  |  | 1
| 1 | 2 | 2.01: 5.57; 2.02: 5.4; 2.03: 5.1022; 2.04: 5.5654; 2.06: 5.18; 2.07: 4.79; 2.08: 5.3; 2.09: 4.6 |  |  |  |  | 28
| 1 | 3 | 2.01: 5.45; 2.02: 5.4; 2.03: 5.0157; 2.04: 5.5591; 2.06: 4.58; 2.07: 4.9; 2.08: 4.9; 2.09: 5 |  |  |  |  | 16
| 1 | 4 | 2.01: 5.39; 2.02: 5; 2.03: 5.2658; 2.04: 5.1559; 2.06: 4.86; 2.07: 5.16; 2.08: 5.1; 2.09: 4.9 |  |  |  |  | 5.1346
| 1 | 5 | 2.01: 5.5; 2.02: 5.5; 2.03: 5.1517; 2.04: 5.6452; 2.06: 4.72; 2.07: 4.89; 2.08: 5; 2.09: 4.8 |  |  |  |  | 29
| 1 | 6 | 2.01: 5.41; 2.02: 5.1; 2.03: 5.468; 2.04: 5.5427; 2.06: 5.11; 2.07: 4.77; 2.08: 5.1; 2.09: 5
| Mean |  | 2.01: 5.465; 2.02: 5.2; 2.03: 5.1899; 2.04: 5.4916; 2.06: 4.945; 2.07: 4.9017; 2.08: 5.0333; 2.09: 4.85
| Median |  | 2.01: 5.46; 2.02: 5.25; 2.03: 5.144; 2.04: 5.5509; 2.06: 4.985; 2.07: 4.895; 2.08: 5.05; 2.09: 4.85
| Std Dev. |  | 2.01: 0.065038; 2.02: 0.27568; 2.03: 0.1584; 2.04: 0.17263; 2.06: 0.26425; 2.07: 0.13906; 2.08: 0.17512; 2.09: 0.15166
| Rel.Std.Dev. |  | 2.01: 1.19%; 2.02: 5.302%; 2.03: 3.052%; 2.04: 3.143%; 2.06: 5.344%; 2.07: 2.837%; 2.08: 3.479%; 2.09: 3.127%
| PDM3 |  | 2.01: 6.435%; 2.02: 1.274%; 2.03: 1.078%; 2.04: 6.954%; 2.06: -3.692%; 2.07: -4.536%; 2.08: -1.972%; 2.09: -5.542%
| Z-Score (Absolute) |  | 2.01: 1.27; 2.02: 0.32; 2.03: 0.28; 2.04: 1.36; 2.06: 0.6; 2.07: 0.75; 2.08: 0.28; 2.09: 0.94
| Analytical results for Ge in OREAS 506b (Certified Value 0.17 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.20 | 2.25 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [12 empty] | 2
2
| 1 | 1 | 1.4 | 1.5 | 0.1552 | 0.51 | 0.15 | 0.18 | 0.1 | 0.14 | 0.15 | 0.47 | 0.5 | 0.25963 | 1.54 | 0.2 | [12 empty] | 1
| 1 | 2 | 1.5 | 1.5 | 0.17021 | 0.25 | 0.14 | 0.17 | 0.1 | 0.12 | 0.19 | 0.49 | 0.6 | 0.2737 | 1.6 | 0.2 | [12 empty] | 0
| 1 | 3 | 1.5 | 1.5 | 0.16157 | 0.34 | 0.16 | 0.16 | 0.09 | 0.16 | 0.14 | 0.47 | 0.5 | 0.28411 | 1.58 | 0.2 | [12 empty] | 16
| 1 | 4 | 1.5 | 1.5 | 0.1618 | 0.28 | 0.15 | 0.16 | 0.1 | 0.18 | 0.18 | 0.46 | 0.5 | 0.2817 | 1.53 | 0.2 | [12 empty] | 0.16626
| 1 | 5 | 1.6 | 1.5 | 0.17196 | 0.37 | 0.16 | 0.18 | 0.08 | 0.14 | 0.17 | 0.47 | 0.4 | 0.28953 | 1.54 | 0.2 | [12 empty] | 30
| 1 | 6 | 1.5 | 1.5 | 0.1585 | 0.18 | 0.16 | 0.18 | 0.09 | 0.13 | 0.16 | 0.47 | 0.6 | 0.24507 | 1.53 | 0.2
| Mean |  | 1.5 | 1.5 | 0.16321 | 0.32167 | 0.15333 | 0.17167 | 0.093333 | 0.145 | 0.165 | 0.47167 | 0.51667 | 0.27229 | 1.5533 | 0.2
| Median |  | 1.5 | 1.5 | 0.16169 | 0.31 | 0.155 | 0.175 | 0.095 | 0.14 | 0.165 | 0.47 | 0.5 | 0.2777 | 1.54 | 0.2
| Std Dev. |  | 0.063246 | 0 | 0.006581 | 0.11409 | 0.008165 | 0.0098319 | 0.008165 | 0.021679 | 0.018708 | 0.0098319 | 0.075277 | 0.016903 | 0.029439 | 3.0405e-17
| Rel.Std.Dev. |  | 4.216% | 0% | 4.032% | 35.47% | 5.325% | 5.727% | 8.748% | 14.95% | 11.34% | 2.085% | 14.57% | 6.208% | 1.895% | 1.52e-14%
| PDM3 |  | 802.2% | 802.2% | -1.837% | 93.47% | -7.776% | 3.251% | -43.86% | -12.79% | -0.7588% | 183.7% | 210.8% | 63.77% | 834.3% | 20.29%
| Z-Score (Absolute) |  | NA | NA | 0.07 | 2.1 | 0.21 | 0.05 | 1.03 | 0.32 | 0.05 | 4.16 | NA | 1.43 | 18.99 | NA
Results from laboratories 2.01, 2.02, 2.14 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 506b (Certified Value 2.09 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | 2.27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 2.46 | 1.9 | 2.1248 | 2.01 | 1.35 | 2.1 | 2.1 | 2.1 | 2.4 | 1.9 | 2.05 | 2 | 2.8 | 2.0538 | 2.05 | 2 | 2.09 | 1.87 | 2.5 | 1.78 | [6 empty] | 1
| 1 | 2 | 2.34 | 1.9 | 2.2065 | 2.16 | 1.88 | 2.2 | 2.1 | 2.2 | 2.4 | 2 | 1.99 | 1.9 | 2.91 | 2.0856 | 2.11 | 1.95 | 2.05 | 1.96 | 2.5 | 1.74 | [6 empty] | 29
| 1 | 3 | 2.13 | 2 | 2.2702 | 2.2 | 1.86 | 2.2 | 2.1 | 2.3 | 2.3 | 2.1 | 1.97 | 1.9 | 2.84 | 2.0329 | 2.18 | 2.02 | 2.11 | 2.04 | 2.6 | 1.87 | [6 empty] | 16
| 1 | 4 | 2.17 | 2 | 2.3738 | 2.07 | 1.71 | 2.1 | 2.2 | 2.3 | 2.4 | 2 | 1.96 | 1.9 | 2.8 | 2.1723 | 2.11 | 1.97 | 2.09 | 1.92 | 2.5 | 1.74 | [6 empty] | 2.0909
| 1 | 5 | 2.5 | 2 | 2.4 | 2.27 | 1.84 | 2.3 | 2.1 | 2.1 | 2.3 | 2 | 1.93 | 1.9 | 2.76 | 2.0664 | 2.15 | 1.82 | 2.07 | 1.88 | 2.6 | 1.83 | [6 empty] | 31
| 1 | 6 | 2.15 | 2 | 2.1668 | 2.07 | 1.76 | 2.2 | 2.1 | 2.2 | 2.3 | 1.9 | 1.99 | 2 | 2.93 | 2.0759 | 2.18 | 1.91 | 2.16 | 2 | 2.5 | 1.91
| Mean |  | 2.2917 | 1.9667 | 2.257 | 2.13 | 1.7333 | 2.1833 | 2.1167 | 2.2 | 2.35 | 1.9833 | 1.9817 | 1.9333 | 2.84 | 2.0812 | 2.13 | 1.945 | 2.095 | 1.945 | 2.5333 | 1.8117
| Median |  | 2.255 | 2 | 2.2383 | 2.115 | 1.8 | 2.2 | 2.1 | 2.2 | 2.35 | 2 | 1.98 | 1.9 | 2.82 | 2.0711 | 2.13 | 1.96 | 2.09 | 1.94 | 2.5 | 1.805
| Std Dev. |  | 0.16437 | 0.05164 | 0.11172 | 0.096954 | 0.19856 | 0.075277 | 0.040825 | 0.089443 | 0.054772 | 0.075277 | 0.040208 | 0.05164 | 0.067231 | 0.048273 | 0.0502 | 0.072319 | 0.037815 | 0.067454 | 0.05164 | 0.070261
| Rel.Std.Dev. |  | 7.172% | 2.626% | 4.95% | 4.552% | 11.46% | 3.448% | 1.929% | 4.066% | 2.331% | 3.795% | 2.029% | 2.671% | 2.367% | 2.32% | 2.357% | 3.718% | 1.805% | 3.468% | 2.038% | 3.878%
| PDM3 |  | 9.603% | -5.94% | 7.946% | 1.871% | -17.1% | 4.422% | 1.234% | 5.219% | 12.39% | -5.143% | -5.223% | -7.535% | 35.83% | -0.4645% | 1.871% | -6.977% | 0.1973% | -6.977% | 21.16% | -13.35%
| Z-Score (Absolute) |  | 0.86 | 0.65 | 0.7 | 0.11 | 1.73 | 0.36 | 0.05 | 0.44 | 1.13 | 0.57 | 0.58 | 0.8 | 3.41 | 0.11 | 0.11 | 0.75 | 0.05 | 0.75 | 1.99 | 1.37
| Analytical results for Hg in OREAS 506b (Indicative Value 0.061 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab | [6 empty] | 1
| No | No | 2.01: 2.15; 2.02: 2.19; 2.03: 2.20; 2.04: 2.26 | [6 empty] | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS | [6 empty] | 3
3
| 1 | 1 | 2.01: < 1; 2.02: < 2; 2.03: 0.08; 2.04: 0.05 | [6 empty] | 1
| 1 | 2 | 2.01: < 1; 2.02: < 2; 2.03: 0.06; 2.04: 0.05 | [6 empty] | 37
| 1 | 3 | 2.01: < 1; 2.02: < 2; 2.03: 0.06; 2.04: 0.09 | [6 empty] | 16
| 1 | 4 | 2.01: < 1; 2.02: < 2; 2.03: < 0.05; 2.04: 0.09 | [6 empty] | 0.060833
| 1 | 5 | 2.01: < 1; 2.02: < 2; 2.03: < 0.05; 2.04: 0.07 | [6 empty] | 37
| 1 | 6 | 2.01: < 1; 2.02: < 2; 2.03: 0.07; 2.04: 0.06
| Mean |  | 2.03: 0.0675; 2.04: 0.068333
| Median |  | 2.03: 0.065; 2.04: 0.065
| Std Dev. |  | 2.03: 0.0095743; 2.04: 0.018348
| Rel.Std.Dev. |  | 2.03: 14.18%; 2.04: 26.85%
| PDM3 |  | 2.03: 10.96%; 2.04: 12.33%
| Z-Score (Absolute) |  | 2.01: 0.65; 2.02: 2.16; 2.03: 0.7; 2.04: 0.65
| Analytical results for Ho in OREAS 506b (Certified Value 0.53 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | [6 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 | [6 empty] | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 2.01: 0.55; 2.02: 0.6; 2.03: 0.51815; 2.04: 0.50429; 2.06: 0.58; 2.07: 0.48; 2.08: 0.6; 2.09: 0.5 | [6 empty] | 1
| 1 | 2 | 2.01: 0.6; 2.02: 0.6; 2.03: 0.52785; 2.04: 0.5321; 2.06: 0.56; 2.07: 0.49; 2.08: 0.6; 2.09: 0.5 | [6 empty] | 12
| 1 | 3 | 2.01: 0.55; 2.02: 0.6; 2.03: 0.52292; 2.04: 0.5224; 2.06: 0.54; 2.07: 0.5; 2.08: 0.6; 2.09: 0.5 | [6 empty] | 16
| 1 | 4 | 2.01: 0.6; 2.02: 0.5; 2.03: 0.55276; 2.04: 0.53651; 2.06: 0.55; 2.07: 0.51; 2.08: 0.6; 2.09: 0.5 | [6 empty] | 0.53363
| 1 | 5 | 2.01: 0.56; 2.02: 0.6; 2.03: 0.53422; 2.04: 0.51832; 2.06: 0.54; 2.07: 0.49; 2.08: 0.6; 2.09: 0.5 | [6 empty] | 32
| 1 | 6 | 2.01: 0.54; 2.02: 0.5; 2.03: 0.54892; 2.04: 0.51051; 2.06: 0.56; 2.07: 0.48; 2.08: 0.6; 2.09: 0.5
| Mean |  | 2.01: 0.56667; 2.02: 0.56667; 2.03: 0.53414; 2.04: 0.52069; 2.06: 0.555; 2.07: 0.49167; 2.08: 0.6; 2.09: 0.5
| Median |  | 2.01: 0.555; 2.02: 0.6; 2.03: 0.53103; 2.04: 0.52036; 2.06: 0.555; 2.07: 0.49; 2.08: 0.6; 2.09: 0.5
| Std Dev. |  | 2.01: 0.026583; 2.02: 0.05164; 2.03: 0.014046; 2.04: 0.01234; 2.06: 0.015166; 2.07: 0.01169; 2.08: 0; 2.09: 0
| Rel.Std.Dev. |  | 2.01: 4.691%; 2.02: 9.113%; 2.03: 2.63%; 2.04: 2.37%; 2.06: 2.733%; 2.07: 2.378%; 2.08: 0%; 2.09: 0%
| PDM3 |  | 2.01: 6.191%; 2.02: 6.191%; 2.03: 0.09466%; 2.04: -2.426%; 2.06: 4.004%; 2.07: -7.864%; 2.08: 12.44%; 2.09: -6.302%
| Z-Score (Absolute) |  | 2.01: 1.05; 2.02: NA; 2.03: 0; 2.04: 0.43; 2.06: 0.67; 2.07: 1.37; 2.08: NA; 2.09: NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 506b (Certified Value 0.079 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.08 | 0.07 | 0.083474 | 0.08 | < 1 | 0.077 | 0.078 | 0.074 | 0.096 | 0.075 | 0.081 | 0.096 | 0.08 | < 0.2 | 0.1 | 0.072 | 0.073 | 0.08 | 0.11 | 0.07 | < 0.1 | 0.08 |  |  |  |  | 1
| 1 | 2 | 0.08 | 0.07 | 0.073166 | 0.09 | < 1 | 0.085 | 0.077 | 0.079 | 0.09 | 0.075 | 0.08 | 0.189 | 0.08 | < 0.2 | 0.1 | 0.073 | 0.072 | 0.09 | 0.1 | 0.08 | < 0.1 | 0.07 |  |  |  |  | 52
| 1 | 3 | 0.08 | 0.07 | 0.085095 | 0.09 | < 1 | 0.082 | 0.073 | 0.076 | 0.08 | 0.08 | 0.082 | 0.253 | 0.08 | < 0.2 | 0.1 | 0.079 | 0.073 | 0.08 | 0.09 | 0.08 | < 0.1 | 0.07 |  |  |  |  | 16
| 1 | 4 | 0.08 | 0.07 | 0.075051 | 0.09 | < 1 | 0.086 | 0.079 | 0.074 | 0.085 | 0.074 | 0.08 | 0.159 | 0.08 | < 0.2 | 0.1 | 0.077 | 0.076 | 0.09 | 0.09 | 0.08 | < 0.1 | 0.08 |  |  |  |  | 0.078844
| 1 | 5 | 0.08 | 0.07 | 0.080122 | 0.08 | < 1 | 0.076 | 0.078 | 0.073 | 0.092 | 0.085 | 0.08 | 0.126 | 0.09 | < 0.2 | 0.1 | 0.083 | 0.068 | 0.08 | 0.09 | 0.08 | < 0.1 | 0.08 |  |  |  |  | 33
| 1 | 6 | 0.07 | 0.07 | 0.075942 | 0.09 | < 1 | 0.078 | 0.077 | 0.077 | 0.081 | 0.073 | 0.084 | 0.11 | 0.08 | < 0.2 | 0.1 | 0.076 | 0.077 | 0.08 | 0.1 | 0.08 | < 0.1 | 0.08
| Mean |  | 0.078333 | 0.07 | 0.078809 | 0.086667 |  | 0.080667 | 0.077 | 0.0755 | 0.087333 | 0.077 | 0.081167 | 0.1555 | 0.081667 |  | 0.1 | 0.076667 | 0.073167 | 0.083333 | 0.096667 | 0.078333 |  | 0.076667
| Median |  | 0.08 | 0.07 | 0.078032 | 0.09 |  | 0.08 | 0.0775 | 0.075 | 0.0875 | 0.075 | 0.0805 | 0.1425 | 0.08 |  | 0.1 | 0.0765 | 0.073 | 0.08 | 0.095 | 0.08 |  | 0.08
| Std Dev. |  | 0.0040825 | 0 | 0.0048412 | 0.005164 |  | 0.004274 | 0.0020976 | 0.0022583 | 0.006377 | 0.0046043 | 0.0016021 | 0.058518 | 0.0040825 |  | 1.5202e-17 | 0.0040332 | 0.0031885 | 0.005164 | 0.008165 | 0.0040825 |  | 0.005164
| Rel.Std.Dev. |  | 5.212% | 0% | 6.143% | 5.958% |  | 5.298% | 2.724% | 2.991% | 7.302% | 5.98% | 1.974% | 37.63% | 4.999% |  | 1.52e-14% | 5.261% | 4.358% | 6.197% | 8.447% | 5.212% |  | 6.736%
| PDM3 |  | -0.6481% | -11.22% | -0.04517% | 9.921% |  | 2.311% | -2.339% | -4.242% | 10.77% | -2.339% | 2.946% | 97.22% | 3.58% |  | 26.83% | -2.762% | -7.201% | 5.694% | 22.6% | -0.6481% |  | -2.762%
| Z-Score (Absolute) |  | 0.1 | 1.8 | 0 | 1.6 | 85.84 | 0.38 | 0.37 | 0.67 | 1.74 | 0.37 | 0.48 | 15.63 | 0.58 | 4.32 | NA | 0.44 | 1.15 | 0.92 | 3.64 | 0.1 | 5.87 | 0.44
Results from laboratory 2.17 were removed due to their 0.01 ppm reading resolution.
| Analytical results for K in OREAS 506b (Certified Value 3.39 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.3179 | 3.39 | 3.077 | 1.85 | 3.87 | 3.42 | 3.33 | 3.39 | 3.54 | 3.33 | 3.41 | 2.46 | 3.29 | 3.3196 | 3.33 | 3.39 | 3.5366 | 3.41 | 3.34 | 2.72 | 3.31 | 3.428 | 3.51 | 1.47 | 3.72 |  | 1
| 1 | 2 | 3.342 | 3.32 | 2.9262 | 1.98 | 3.85 | 3.49 | 3.39 | 3.4 | 3.44 | 3.4 | 3.43 | 2.28 | 3.37 | 3.3917 | 3.42 | 3.39 | 3.5384 | 3.35 | 3.33 | 3.08 | 3.36 | 3.412 | 3.55 | 1.42 | 3.63 |  | 13
| 1 | 3 | 3.296 | 3.4 | 2.9032 | 1.9 | 3.88 | 3.43 | 3.34 | 3.4 | 3.54 | 3.5 | 3.1 | 2.53 | 3.26 | 3.3295 | 3.4 | 3.46 | 3.5326 | 3.34 | 3.32 | 2.62 | 3.35 | 3.376 | 3.6 | 1.72 | 3.68 |  | 16
| 1 | 4 | 3.3134 | 3.41 | 3.2358 | 1.97 | 3.85 | 3.4 | 3.34 | 3.35 | 3.57 | 3.4 | 3.43 | 2.5 | 3.28 | 3.323 | 3.35 | 3.42 | 3.5324 | 3.35 | 3.44 | 3.02 | 3.3 | 3.3596 | 3.53 | 1.66 | 3.68 |  | 3.3911
| 1 | 5 | 3.3347 | 3.31 | 3.0983 | 2.01 | 3.85 | 3.55 | 3.36 | 3.36 | 3.53 | 3.38 | 3.41 | 2.67 | 3.25 | 3.2883 | 3.22 | 3.42 | 3.5384 | 3.37 | 3.17 | 2.62 | 3.35 | 3.4033 | 3.52 | 3.65 | 3.68 |  | 34
| 1 | 6 | 3.3537 | 3.37 | 3.1292 | 2.76 | 3.86 | 3.41 | 3.33 | 3.34 | 3.52 | 3.37 | 3.16 | 2.32 | 3.35 | 3.3673 | 3.46 | 3.54 | 3.5314 | 3.39 | 3.31 | 2.62 | 3.4 | 3.344 | 3.52 | 3.72 | 3.61
| Mean |  | 3.3263 | 3.3667 | 3.0616 | 2.0783 | 3.86 | 3.45 | 3.3483 | 3.3733 | 3.5233 | 3.3967 | 3.3233 | 2.46 | 3.3 | 3.3366 | 3.3633 | 3.4367 | 3.535 | 3.3683 | 3.3183 | 2.78 | 3.345 | 3.3872 | 3.5383 | 2.2733 | 3.6667
| Median |  | 3.3263 | 3.38 | 3.0876 | 1.975 | 3.855 | 3.425 | 3.34 | 3.375 | 3.535 | 3.39 | 3.41 | 2.48 | 3.285 | 3.3263 | 3.375 | 3.42 | 3.5346 | 3.36 | 3.325 | 2.67 | 3.35 | 3.3897 | 3.525 | 1.69 | 3.68
| Std Dev. |  | 0.021085 | 0.042269 | 0.12641 | 0.33902 | 0.012649 | 0.05831 | 0.023166 | 0.026583 | 0.044121 | 0.056804 | 0.15122 | 0.14325 | 0.04899 | 0.036984 | 0.084538 | 0.056804 | 0.0031985 | 0.027142 | 0.086583 | 0.21354 | 0.036194 | 0.032534 | 0.033116 | 1.0994 | 0.039833
| Rel.Std.Dev. |  | 0.6339% | 1.256% | 4.129% | 16.31% | 0.3277% | 1.69% | 0.6919% | 0.788% | 1.252% | 1.672% | 4.55% | 5.823% | 1.485% | 1.108% | 2.514% | 1.653% | 0.09048% | 0.8058% | 2.609% | 7.681% | 1.082% | 0.9605% | 0.9359% | 48.36% | 1.086%
| PDM3 |  | -1.913% | -0.7219% | -9.717% | -38.71% | 13.83% | 1.736% | -1.262% | -0.5253% | 3.898% | 0.1628% | -2% | -27.46% | -2.688% | -1.609% | -0.8202% | 1.342% | 4.241% | -0.6727% | -2.147% | -18.02% | -1.361% | -0.1178% | 4.34% | -32.96% | 8.125%
| Z-Score (Absolute) |  | 0.39 | 0.04 | 3.21 | 13.68 | 5.29 | 0.92 | 0.16 | 0.11 | 1.7 | 0.35 | 0.43 | 9.62 | 0.67 | 0.28 | 0 | 0.78 | 1.83 | 0.05 | 0.48 | 6.21 | 0.2 | 0.25 | 1.86 | 11.61 | 3.23
| Analytical results for La in OREAS 506b (Certified Value 31.8 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 32.48 | 29.3 | 32 | 34.5 | 29 | 32.1 | 36 | 29.3 | 33.82 | 28.8 | 22.8 | 30.4092 | 32.2 | 29.8 | 31.6461 | 33 | 30.2 | 28.47 | 30 | 31.9 | 36.9 | 29.4 | 29.3 |  |  |  | 1
| 1 | 2 | 33.59 | 33.2 | 34.84 | 32.8 | 29.2 | 34.6 | 35.5 | 30.6 | 32.7 | 29.1 | 21.4 | 30.8006 | 33.9 | 32.2 | 31.627 | 33 | 30 | 28.54 | 31.7 | 31.1 | 35.3 | 33.1 | 29.2 |  |  |  | 14
| 1 | 3 | 33.01 | 33.4 | 38.04 | 31.9 | 28.7 | 34.3 | 33.3 | 33.8 | 33.5 | 28.5 | 21.2 | 30.2739 | 33 | 32.3 | 31.6507 | 31 | 26.4 | 29.69 | 29.5 | 31.8 | 38.2 | 30 | 31.6 |  |  |  | 16
| 1 | 4 | 33.35 | 30.1 | 33.59 | 32 | 30.4 | 34.8 | 34.4 | 31.8 | 32.33 | 28.2 | 21.9 | 30.955 | 33 | 31.9 | 31.6531 | 31 | 27.2 | 29.39 | 31.7 | 31.9 | 35.4 | 32.4 | 30.8 |  |  |  | 31.7644
| 1 | 5 | 33.77 | 33.5 | 29.55 | 32.6 | 29.3 | 33.7 | 36.2 | 31.7 | 33.55 | 26.7 | 21.6 | 30.6574 | 30 | 31 | 31.6487 | 32 | 26.9 | 27.64 | 28.6 | 31.9 | 34.4 | 33.6 | 29.5 |  |  |  | 35
| 1 | 6 | 32.54 | 32.4 | 34.31 | 31.1 | 31.2 | 34.4 | 33.3 | 30.1 | 31.1 | 29.7 | 21.7 | 32.4518 | 32.8 | 32.5 | 31.6638 | 31 | 28.8 | 27.67 | 31.8 | 31.1 | 35.6 | 32.8 | 31.3
| Mean |  | 33.1233 | 31.9833 | 33.7217 | 32.4833 | 29.6333 | 33.9833 | 34.7833 | 31.2167 | 32.8333 | 28.5 | 21.7667 | 30.9246 | 32.4833 | 31.6167 | 31.6482 | 31.8333 | 28.25 | 28.5667 | 30.55 | 31.6167 | 35.9667 | 31.8833 | 30.2833
| Median |  | 33.18 | 32.8 | 33.95 | 32.3 | 29.25 | 34.35 | 34.95 | 31.15 | 33.1 | 28.65 | 21.65 | 30.729 | 32.9 | 32.05 | 31.6497 | 31.5 | 28 | 28.505 | 30.85 | 31.85 | 35.5 | 32.6 | 30.15
| Std Dev. |  | 0.53932 | 1.828 | 2.8506 | 1.1548 | 0.96056 | 0.99482 | 1.3075 | 1.5842 | 1.0205 | 1.0218 | 0.56095 | 0.78856 | 1.3333 | 1.0342 | 0.012065 | 0.98319 | 1.6441 | 0.85 | 1.3722 | 0.40208 | 1.3574 | 1.7463 | 1.0759
| Rel.Std.Dev. |  | 1.628% | 5.716% | 8.453% | 3.555% | 3.241% | 2.927% | 3.759% | 5.075% | 3.108% | 3.585% | 2.577% | 2.55% | 4.105% | 3.271% | 0.03812% | 3.089% | 5.82% | 2.976% | 4.492% | 1.272% | 3.774% | 5.477% | 3.553%
| PDM3 |  | 4.278% | 0.6894% | 6.162% | 2.263% | -6.709% | 6.986% | 9.504% | -1.724% | 3.365% | -10.28% | -31.47% | -2.644% | 2.263% | -0.4649% | -0.3655% | 0.2172% | -11.06% | -10.07% | -3.823% | -0.4649% | 13.23% | 0.3746% | -4.663%
| Z-Score (Absolute) |  | 0.84 | 0.19 | 1.18 | 0.48 | 1.15 | 1.33 | 1.78 | 0.25 | 0.67 | 1.79 | 5.62 | 0.41 | 0.48 | 0.02 | 0 | 0.11 | 1.93 | 1.75 | 0.62 | 0.02 | 2.46 | 0.13 | 0.78
| Analytical results for Li in OREAS 506b (Certified Value 41.8 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 40.2 | 42.1 | 44.7476 | 41.7 | 41.54 | 43.7 | 41.7 | 44.2 | 43.3 | 41.2 | 45.1 | 40.3 | 39 | 41.4719 | 42 | 42.55 | 52.65 | 40 | 41.4 | 26 | 34.9 | 44.6 | 37 | 38 | 41.8 |  | 1
| 1 | 2 | 40.8 | 42.3 | 44.7114 | 46 | 48.08 | 45 | 43.4 | 44 | 42.7 | 41.6 | 43.1 | 40.7 | 40 | 42.0952 | 43 | 43.83 | 51.27 | 39 | 43.6 | 34 | 36.6 | 44.2 | 36 | 39 | 39.8 |  | 15
| 1 | 3 | 41.4 | 41.9 | 44.9301 | 43.6 | 48.3 | 44.5 | 42.6 | 43.9 | 43.1 | 43.4 | 43.8 | 39.9 | 39 | 42.4892 | 43 | 43.47 | 52.41 | 39 | 43.9 | 31 | 35.5 | 43.8 | 38 | 39 | 40.4 |  | 16
| 1 | 4 | 41.1 | 42.1 | 43.4636 | 36.7 | 49.3 | 44.3 | 41.4 | 43.4 | 44 | 41.9 | 43.4 | 39 | 39 | 41.5324 | 43 | 44.73 | 51.2 | 39 | 44.3 | 30 | 36.6 | 45.1 | 35 | 38 | 39.9 |  | 41.8221
| 1 | 5 | 41.5 | 42.1 | 44.0311 | 45.3 | 48.04 | 45.1 | 42 | 43.6 | 43.4 | 41.7 | 44.8 | 37.8 | 39 | 40.4268 | 40 | 43.67 | 52.37 | 39 | 41.5 | 26 | 36.5 | 44.4 | 36 | 37 | 40.4 |  | 36
| 1 | 6 | 41 | 42.1 | 43.205 | < 0.1 | 46.18 | 44.1 | 41.6 | 43 | 43.9 | 41.9 | 41.9 | 41.6 | 40 | 41.539 | 43 | 44.65 | 52.09 | 39 | 43.6 | 23 | 36 | 43.4 | 37 | 37 | 41.2
| Mean |  | 41 | 42.1 | 44.1815 | 42.66 | 46.9067 | 44.45 | 42.1167 | 43.6833 | 43.4 | 41.95 | 43.6833 | 39.8833 | 39.3333 | 41.5924 | 42.3333 | 43.8167 | 51.9983 | 39.1667 | 43.05 | 28.3333 | 36.0167 | 44.25 | 36.5 | 38 | 40.5833
| Median |  | 41.05 | 42.1 | 44.3713 | 43.6 | 48.06 | 44.4 | 41.85 | 43.75 | 43.35 | 41.8 | 43.6 | 40.1 | 39 | 41.5357 | 43 | 43.75 | 52.23 | 39 | 43.6 | 28 | 36.25 | 44.3 | 36.5 | 38 | 40.4
| Std Dev. |  | 0.46904 | 0.12649 | 0.72845 | 3.7233 | 2.8164 | 0.53572 | 0.75476 | 0.44008 | 0.4899 | 0.75565 | 1.172 | 1.3348 | 0.5164 | 0.69809 | 1.2111 | 0.80931 | 0.61782 | 0.40825 | 1.2661 | 4.0332 | 0.6969 | 0.59917 | 1.0488 | 0.89443 | 0.77567
| Rel.Std.Dev. |  | 1.144% | 0.3005% | 1.649% | 8.728% | 6.004% | 1.205% | 1.792% | 1.007% | 1.129% | 1.801% | 2.683% | 3.347% | 1.313% | 1.678% | 2.861% | 1.847% | 1.188% | 1.042% | 2.941% | 14.23% | 1.935% | 1.354% | 2.873% | 2.354% | 1.911%
| PDM3 |  | -1.966% | 0.6646% | 5.642% | 2.004% | 12.16% | 6.284% | 0.7044% | 4.45% | 3.773% | 0.3059% | 4.45% | -4.636% | -5.951% | -0.5491% | 1.223% | 4.769% | 24.33% | -6.349% | 2.936% | -32.25% | -13.88% | 5.805% | -12.73% | -9.139% | -2.962%
| Z-Score (Absolute) |  | 0.36 | 0 | 0.67 | 2.12 | 1.56 | 0.76 | 0.01 | 0.51 | 0.42 | 0.05 | 0.51 | 0.72 | 0.9 | 0.16 | 0.08 | 0.56 | 3.21 | 0.95 | 0.31 | 4.46 | 1.97 | 0.7 | 1.81 | 1.33 | 0.49
| Analytical results for Lu in OREAS 506b (Certified Value 0.17 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.06; 2.04: 2.12; 2.06: 2.14; 2.07: 2.15; 2.08: 2.20; 2.09: 2.21; 2.10: 2.24; 2.11: 2.25; 2.12: 2.27 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*MS; 2.10: 4A*OES/MS; 2.11: 4A*OES/MS; 2.12: 4A*OES/MS |  | 2
3
| 1 | 1 | 2.01: 0.19; 2.02: 0.2; 2.03: < 1; 2.04: 0.17; 2.06: 0.21; 2.07: 0.16577; 2.08: 0.18; 2.09: 0.16; 2.10: 0.17; 2.11: 0.2; 2.12: 0.2 |  | 1
| 1 | 2 | 2.01: 0.19; 2.02: 0.2; 2.03: < 1; 2.04: 0.16; 2.06: 0.2; 2.07: 0.16615; 2.08: 0.17; 2.09: 0.14; 2.10: 0.17; 2.11: 0.2; 2.12: 0.1 |  | 16
| 1 | 3 | 2.01: 0.19; 2.02: 0.2; 2.03: < 1; 2.04: 0.17; 2.06: 0.18; 2.07: 0.16951; 2.08: 0.16; 2.09: 0.16; 2.10: 0.18; 2.11: 0.2; 2.12: 0.2 |  | 16
| 1 | 4 | 2.01: 0.19; 2.02: 0.1; 2.03: < 1; 2.04: 0.16; 2.06: 0.19; 2.07: 0.17807; 2.08: 0.17; 2.09: 0.16; 2.10: 0.17; 2.11: 0.2; 2.12: 0.2 |  | 0.17357
| 1 | 5 | 2.01: 0.19; 2.02: 0.2; 2.03: < 1; 2.04: 0.16; 2.06: 0.19; 2.07: 0.1821; 2.08: 0.17; 2.09: 0.14; 2.10: 0.17; 2.11: 0.2; 2.12: 0.2 |  | 37
| 1 | 6 | 2.01: 0.15; 2.02: 0.2; 2.03: < 1; 2.04: 0.16; 2.06: 0.2; 2.07: 0.17835; 2.08: 0.17; 2.09: 0.15; 2.10: 0.17; 2.11: 0.2; 2.12: 0.2
| Mean |  | 2.01: 0.18333; 2.02: 0.18333; 2.04: 0.16333; 2.06: 0.195; 2.07: 0.17333; 2.08: 0.17; 2.09: 0.15167; 2.10: 0.17167; 2.11: 0.2; 2.12: 0.18333
| Median |  | 2.01: 0.19; 2.02: 0.2; 2.04: 0.16; 2.06: 0.195; 2.07: 0.17379; 2.08: 0.17; 2.09: 0.155; 2.10: 0.17; 2.11: 0.2; 2.12: 0.2
| Std Dev. |  | 2.01: 0.01633; 2.02: 0.040825; 2.04: 0.005164; 2.06: 0.010488; 2.07: 0.007039; 2.08: 0.0063246; 2.09: 0.0098319; 2.10: 0.0040825; 2.11: 3.0405e-17; 2.12: 0.040825
| Rel.Std.Dev. |  | 2.01: 8.907%; 2.02: 22.27%; 2.04: 3.162%; 2.06: 5.379%; 2.07: 4.061%; 2.08: 3.72%; 2.09: 6.483%; 2.10: 2.378%; 2.11: 1.52e-14%; 2.12: 22.27%
| PDM3 |  | 2.01: 5.625%; 2.02: 5.625%; 2.04: -5.898%; 2.06: 12.35%; 2.07: -0.1412%; 2.08: -2.057%; 2.09: -12.62%; 2.10: -1.097%; 2.11: 15.23%; 2.12: 5.625%
| Z-Score (Absolute) |  | 2.01: 0.73; 2.02: NA; 2.03: 22.08; 2.04: 0.62; 2.06: 1.52; 2.07: 0.06; 2.08: 0.17; 2.09: 1.4; 2.10: 0.06; 2.11: NA; 2.12: NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 506b (Certified Value 0.686 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.6907 | 0.723 | 0.68969 | 0.67 | 0.73 | 0.68 | 0.65 | 0.67 | 0.7 | 0.66 | 0.703 | 0.78 | 0.64 | 0.68498 | 0.717 | 0.69 | 0.69465 | 0.71 | 0.675 | 0.562 | 0.64 | 0.706 | 0.71 | 0.79 | 0.6993 | 0.65 | 1
| 1 | 2 | 0.6958 | 0.705 | 0.68387 | 0.71 | 0.73 | 0.68 | 0.63 | 0.67 | 0.69 | 0.67 | 0.708 | 0.79 | 0.66 | 0.70235 | 0.738 | 0.67 | 0.69127 | 0.7 | 0.678 | 0.66 | 0.65 | 0.704 | 0.72 | 0.78 | 0.7167 | 0.64 | 17
| 1 | 3 | 0.6846 | 0.702 | 0.67295 | 0.72 | 0.74 | 0.67 | 0.63 | 0.67 | 0.69 | 0.7 | 0.703 | 0.78 | 0.64 | 0.69357 | 0.731 | 0.68 | 0.69589 | 0.69 | 0.674 | 0.649 | 0.65 | 0.71 | 0.72 | 0.76 | 0.7166 | 0.64 | 16
| 1 | 4 | 0.6872 | 0.708 | 0.69535 | 0.68 | 0.73 | 0.67 | 0.67 | 0.66 | 0.71 | 0.67 | 0.706 | 0.77 | 0.63 | 0.69373 | 0.722 | 0.68 | 0.69457 | 0.68 | 0.698 | 0.624 | 0.62 | 0.706 | 0.71 | 0.74 | 0.7172 | 0.64 | 0.68638
| 1 | 5 | 0.6957 | 0.709 | 0.65887 | 0.68 | 0.73 | 0.69 | 0.65 | 0.66 | 0.7 | 0.67 | 0.703 | 0.74 | 0.63 | 0.67553 | 0.692 | 0.67 | 0.69593 | 0.69 | 0.63 | 0.588 | 0.63 | 0.708 | 0.72 | 0.79 | 0.7223 | 0.64 | 38
| 1 | 6 | 0.6946 | 0.704 | 0.65838 | 0.7 | 0.74 | 0.68 | 0.65 | 0.66 | 0.69 | 0.67 | 0.703 | 0.79 | 0.67 | 0.69859 | 0.751 | 0.69 | 0.69334 | 0.7 | 0.664 | 0.53 | 0.64 | 0.708 | 0.71 | 0.76 | 0.7109 | 0.64
| Mean |  | 0.69143 | 0.7085 | 0.67652 | 0.69333 | 0.73333 | 0.67833 | 0.64667 | 0.665 | 0.69667 | 0.67333 | 0.70433 | 0.775 | 0.645 | 0.69146 | 0.72517 | 0.68 | 0.69427 | 0.695 | 0.66983 | 0.60217 | 0.63833 | 0.707 | 0.715 | 0.77 | 0.71383 | 0.64167
| Median |  | 0.69265 | 0.7065 | 0.67841 | 0.69 | 0.73 | 0.68 | 0.65 | 0.665 | 0.695 | 0.67 | 0.703 | 0.78 | 0.64 | 0.69365 | 0.7265 | 0.68 | 0.69461 | 0.695 | 0.6745 | 0.606 | 0.64 | 0.707 | 0.715 | 0.77 | 0.71665 | 0.64
| Std Dev. |  | 0.0047424 | 0.0075565 | 0.015717 | 0.019664 | 0.005164 | 0.0075277 | 0.015055 | 0.0054772 | 0.008165 | 0.013663 | 0.0021602 | 0.018708 | 0.016432 | 0.0097449 | 0.020213 | 0.0089443 | 0.0017595 | 0.010488 | 0.022471 | 0.051062 | 0.01169 | 0.0020976 | 0.0054772 | 0.02 | 0.0079839 | 0.0040825
| Rel.Std.Dev. |  | 0.6859% | 1.067% | 2.323% | 2.836% | 0.7042% | 1.11% | 2.328% | 0.8236% | 1.172% | 2.029% | 0.3067% | 2.414% | 2.548% | 1.409% | 2.787% | 1.315% | 0.2534% | 1.509% | 3.355% | 8.48% | 1.831% | 0.2967% | 0.766% | 2.597% | 1.118% | 0.6362%
| PDM3 |  | 0.7368% | 3.223% | -1.436% | 1.014% | 6.841% | -1.172% | -5.785% | -3.114% | 1.499% | -1.9% | 2.616% | 12.91% | -6.028% | 0.7403% | 5.652% | -0.929% | 1.151% | 1.256% | -2.41% | -12.27% | -6.999% | 3.005% | 4.17% | 12.18% | 4% | -6.514%
| Z-Score (Absolute) |  | 0.03 | 0.54 | 0.53 | 0.03 | 1.36 | 0.47 | 1.52 | 0.91 | 0.14 | 0.63 | 0.4 | 2.75 | 1.58 | 0.03 | 1.09 | 0.41 | 0.06 | 0.09 | 0.75 | 3 | 1.8 | 0.49 | 0.75 | 2.58 | 0.71 | 1.69
| Analytical results for Mn in OREAS 506b (Certified Value 0.033 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0321 | 0.0329 | 0.033383 | 0.0342 | 0.043523 | 0.0332 | 0.0321 | 0.0331 | 0.0355 | 0.0325 | 0.0332 | 0.0369 | 0.0333 | 0.032422 | 0.0327 | 0.0315 | 0.034807 | 0.0329 | 0.0308 | 0.0341 | 0.0316 | 0.0325 | 0.0316 | 0.0325 | 0.0331 | 0.0316 | 1
| 1 | 2 | 0.0325 | 0.0326 | 0.033859 | 0.0362 | 0.049675 | 0.0333 | 0.0318 | 0.0328 | 0.0349 | 0.0331 | 0.0334 | 0.0369 | 0.0327 | 0.032942 | 0.0336 | 0.0339 | 0.034137 | 0.0323 | 0.0309 | 0.0388 | 0.0322 | 0.033 | 0.0316 | 0.0324 | 0.0338 | 0.0311 | 18
| 1 | 3 | 0.0318 | 0.0324 | 0.033838 | 0.0362 | 0.049824 | 0.0332 | 0.0316 | 0.0327 | 0.0352 | 0.0346 | 0.0331 | 0.0364 | 0.0326 | 0.032642 | 0.0333 | 0.0336 | 0.034273 | 0.0325 | 0.0306 | 0.0319 | 0.0325 | 0.0321 | 0.0317 | 0.0327 | 0.0335 | 0.0313 | 16
| 1 | 4 | 0.0321 | 0.0329 | 0.031974 | 0.0327 | 0.049763 | 0.0333 | 0.032 | 0.0325 | 0.0362 | 0.0332 | 0.0332 | 0.0357 | 0.0343 | 0.032622 | 0.0329 | 0.0336 | 0.034071 | 0.0324 | 0.0325 | 0.0299 | 0.0323 | 0.0343 | 0.0318 | 0.0329 | 0.0337 | 0.0315 | 0.032867
| 1 | 5 | 0.0323 | 0.032 | 0.033196 | 0.0349 | 0.04901 | 0.0338 | 0.0317 | 0.0327 | 0.0352 | 0.033 | 0.0332 | 0.0347 | 0.0327 | 0.03191 | 0.0315 | 0.0326 | 0.034663 | 0.0328 | 0.0267 | 0.0308 | 0.0315 | 0.0342 | 0.0318 | 0.0337 | 0.0338 | 0.0313 | 39
| 1 | 6 | 0.0324 | 0.032 | 0.032266 | 0.0339 | 0.048908 | 0.033 | 0.0323 | 0.0322 | 0.0347 | 0.0335 | 0.033 | 0.0377 | 0.0337 | 0.033308 | 0.034 | 0.0345 | 0.034631 | 0.0328 | 0.0304 | 0.0351 | 0.0328 | 0.0331 | 0.0319 | 0.0339 | 0.0336 | 0.0309
| Mean |  | 0.0322 | 0.032467 | 0.033086 | 0.034683 | 0.048451 | 0.0333 | 0.031917 | 0.032667 | 0.035283 | 0.033317 | 0.033183 | 0.036383 | 0.033217 | 0.032641 | 0.033 | 0.033283 | 0.03443 | 0.032617 | 0.030317 | 0.033433 | 0.03215 | 0.0332 | 0.031733 | 0.033017 | 0.033583 | 0.031283
| Median |  | 0.0322 | 0.0325 | 0.033289 | 0.03455 | 0.049342 | 0.03325 | 0.0319 | 0.0327 | 0.0352 | 0.03315 | 0.0332 | 0.03665 | 0.033 | 0.032632 | 0.0331 | 0.0336 | 0.034452 | 0.03265 | 0.0307 | 0.033 | 0.03225 | 0.03305 | 0.03175 | 0.0328 | 0.03365 | 0.0313
| Std Dev. |  | 0.00025298 | 0.00040825 | 0.00079649 | 0.0013732 | 0.0024459 | 0.00026833 | 0.00026394 | 0.00030111 | 0.00052694 | 0.00070828 | 0.00013292 | 0.0010553 | 0.0006824 | 0.00047272 | 0.00087178 | 0.0010685 | 0.00030861 | 0.00024833 | 0.001924 | 0.00328 | 0.00050892 | 0.00088994 | 0.00012111 | 0.00063377 | 0.00026394 | 0.00025626
| Rel.Std.Dev. |  | 0.7857% | 1.257% | 2.407% | 3.959% | 5.048% | 0.8058% | 0.827% | 0.9218% | 1.493% | 2.126% | 0.4006% | 2.901% | 2.054% | 1.448% | 2.642% | 3.21% | 0.8963% | 0.7614% | 6.346% | 9.811% | 1.583% | 2.681% | 0.3816% | 1.92% | 0.7859% | 0.8191%
| PDM3 |  | -2.03% | -1.218% | 0.6656% | 5.526% | 47.41% | 1.317% | -2.892% | -0.61% | 7.351% | 1.368% | 0.962% | 10.7% | 1.063% | -0.6886% | 0.4042% | 1.266% | 4.756% | -0.7621% | -7.76% | 1.723% | -2.182% | 1.013% | -3.45% | 0.4549% | 2.179% | -4.819%
| Z-Score (Absolute) |  | 1.25 | 0.89 | 0.06 | 2.06 | 20.42 | 0.22 | 1.62 | 0.62 | 2.86 | 0.24 | 0.06 | 4.33 | 0.11 | 0.66 | 0.18 | 0.2 | 1.73 | 0.69 | 3.76 | 0.4 | 1.31 | 0.09 | 1.87 | 0.16 | 0.6 | 2.47
| Analytical results for Mo in OREAS 506b (Certified Value 88 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
1
| 1 | 1 | 88.4 | 86 | 79.4241 | 84.12 | 83.62 | 89.7 | 93.3 | 89.9 | 91 | 84.3 | 89.44 | 82.6 | 87.52 | 85.7514 | 89 | 83.26 | 88.992 | 88 | 89.5 | 91.01 | 87 | 87.2 | 90.2 | 93 | 72.4 | 87.14 | 1
| 1 | 2 | 92 | 85.2 | 77.9757 | 89.32 | 83.43 | 91.8 | 89.4 | 89.2 | 92.7 | 89.3 | 90.06 | 82.3 | 87.2 | 85.4213 | 90 | 84.69 | 87.4715 | 87 | 90.4 | 91.96 | 89 | 88.6 | 89.9 | 94 | 77.5 | 82.52 | 19
| 1 | 3 | 89 | 86 | 79.6315 | 90.12 | 83.29 | 91.5 | 91.3 | 89.5 | 94.6 | 93.8 | 87.38 | 82.4 | 85.55 | 85.2347 | 90 | 83.61 | 87.842 | 86 | 87.1 | 93.08 | 89 | 87.9 | 93.8 | 96 | 81.8 | 85.29 | 16
| 1 | 4 | 89.9 | 85.8 | 79.2981 | 89.54 | 83.57 | 88.2 | 88.3 | 87.7 | 92.6 | 84.2 | 88.87 | 80.5 | 88.16 | 85.6695 | 90 | 83.78 | 84.7485 | 87 | 88.6 | 90.3 | 87 | 87 | 85.2 | 92 | 76.8 | 82.5 | 88.1682
| 1 | 5 | 89.2 | 86.3 | 76.9972 | 85.84 | 83.77 | 90.3 | 91.5 | 87.8 | 93.8 | 90.3 | 89.17 | 78.3 | 85.85 | 83.7577 | 85 | 85.2 | 88.0605 | 88 | 87.2 | 91.06 | 87 | 88.1 | 93.6 | 98 | 82.7 | 85.46 | 40
| 1 | 6 | 91.7 | 86.8 | 77.1598 | 89.92 | 83.94 | 88.9 | 93.3 | 88 | 91.7 | 91.5 | 85.02 | 83.6 | 87.2 | 85.7272 | 92 | 84.35 | 86.6805 | 87 | 89.1 | 94.31 | 88 | 86.3 | 89.5 | 96 | 77.8 | 85.87
| Mean |  | 90.0333 | 86.0167 | 78.4144 | 88.1433 | 83.6033 | 90.0667 | 91.1833 | 88.6833 | 92.7333 | 88.9 | 88.3233 | 81.6167 | 86.9133 | 85.2603 | 89.3333 | 84.1483 | 87.2992 | 87.1667 | 88.65 | 91.9533 | 87.8333 | 87.5167 | 90.3667 | 94.8333 | 78.1667 | 84.7967
| Median |  | 89.55 | 86 | 78.6369 | 89.43 | 83.595 | 90 | 91.4 | 88.6 | 92.65 | 89.8 | 89.02 | 82.35 | 87.2 | 85.5454 | 90 | 84.065 | 87.6568 | 87 | 88.85 | 91.51 | 87.5 | 87.55 | 90.05 | 95 | 77.65 | 85.375
| Std Dev. |  | 1.4895 | 0.53072 | 1.188 | 2.5255 | 0.23269 | 1.4208 | 2.0282 | 0.96212 | 1.3201 | 3.9023 | 1.8483 | 1.9094 | 1.0076 | 0.76302 | 2.3381 | 0.72865 | 1.46 | 0.75277 | 1.3035 | 1.4996 | 0.98319 | 0.83766 | 3.1602 | 2.2286 | 3.727 | 1.886
| Rel.Std.Dev. |  | 1.654% | 0.617% | 1.515% | 2.865% | 0.2783% | 1.577% | 2.224% | 1.085% | 1.424% | 4.39% | 2.093% | 2.339% | 1.159% | 0.8949% | 2.617% | 0.8659% | 1.672% | 0.8636% | 1.47% | 1.631% | 1.119% | 0.9571% | 3.497% | 2.35% | 4.768% | 2.224%
| PDM3 |  | 2.115% | -2.44% | -11.06% | -0.02825% | -5.177% | 2.153% | 3.42% | 0.5842% | 5.178% | 0.83% | 0.1759% | -7.431% | -1.423% | -3.298% | 1.321% | -4.559% | -0.9857% | -1.136% | 0.5464% | 4.293% | -0.3799% | -0.739% | 2.493% | 7.56% | -11.34% | -3.824%
| Z-Score (Absolute) |  | 0.67 | 0.64 | 3.13 | 0.05 | 1.43 | 0.68 | 1.04 | 0.23 | 1.55 | 0.3 | 0.11 | 2.08 | 0.35 | 0.89 | 0.44 | 1.26 | 0.23 | 0.27 | 0.22 | 1.3 | 0.05 | 0.15 | 0.78 | 2.24 | 3.21 | 1.04
| Analytical results for Na in OREAS 506b (Certified Value 2.18 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.1243 | 2.17 | 2.1518 | 2.181 | 2.29 | 2.2 | 2.12 | 2.17 | 2.38 | 2.14 | 2.164 | 2.37 | 2.29 | 2.1091 | 2.071 | 2.22 | 2.2373 | 2.2 | 2.16 | 1.982 | 2.06 | 2.176 | 2.18 | 2.27 | 2.1062 | 2.275 | 1
| 1 | 2 | 2.1447 | 2.18 | 2.1207 | 2.305 | 2.25 | 2.26 | 2.16 | 2.16 | 2.36 | 2.17 | 2.188 | 2.39 | 2.3 | 2.143 | 2.121 | 2.19 | 2.2395 | 2.17 | 2.26 | 2.095 | 2.07 | 2.252 | 2.18 | 2.25 | 2.1608 | 2.219 | 20
| 1 | 3 | 2.1172 | 2.18 | 2.1127 | 2.331 | 2.3 | 2.25 | 2.12 | 2.17 | 2.4 | 2.25 | 2.165 | 2.34 | 2.27 | 2.1406 | 2.104 | 2.21 | 2.2336 | 2.18 | 2.16 | 2.069 | 2.07 | 2.196 | 2.18 | 2.2 | 2.1667 | 2.203 | 16
| 1 | 4 | 2.125 | 2.18 | 2.1544 | 2.61 | 2.24 | 2.25 | 2.12 | 2.14 | 2.43 | 2.17 | 2.183 | 2.31 | 2.24 | 2.1243 | 2.076 | 2.2 | 2.2383 | 2.17 | 2.15 | 2.197 | 2.05 | 2.3085 | 2.18 | 2.19 | 2.1896 | 2.202 | 2.1826
| 1 | 5 | 2.1332 | 2.13 | 2.0979 | 2.247 | 2.28 | 2.25 | 2.13 | 2.13 | 2.4 | 2.16 | 2.176 | 2.26 | 2.25 | 2.0752 | 1.998 | 2.24 | 2.233 | 2.2 | 2.16 | 2.265 | 2.07 | 2.3126 | 2.19 | 2.22 | 2.1918 | 2.216 | 41
| 1 | 6 | 2.1474 | 2.16 | 2.078 | 2.704 | 2.28 | 2.23 | 2.11 | 2.12 | 2.41 | 2.16 | 2.184 | 2.41 | 2.26 | 2.1605 | 2.155 | 2.25 | 2.236 | 2.2 | 2.23 | 2.096 | 2.12 | 2.128 | 2.2 | 2.23 | 2.154 | 2.194
| Mean |  | 2.132 | 2.1667 | 2.1193 | 2.3963 | 2.2733 | 2.24 | 2.1267 | 2.1483 | 2.3967 | 2.175 | 2.1767 | 2.3467 | 2.2683 | 2.1254 | 2.0875 | 2.2183 | 2.2363 | 2.1867 | 2.1867 | 2.1173 | 2.0733 | 2.2288 | 2.185 | 2.2267 | 2.1615 | 2.2182
| Median |  | 2.1291 | 2.175 | 2.1167 | 2.318 | 2.28 | 2.25 | 2.12 | 2.15 | 2.4 | 2.165 | 2.1795 | 2.355 | 2.265 | 2.1324 | 2.09 | 2.215 | 2.2366 | 2.19 | 2.16 | 2.0955 | 2.07 | 2.224 | 2.18 | 2.225 | 2.1637 | 2.2095
| Std Dev. |  | 0.01206 | 0.019664 | 0.029997 | 0.21052 | 0.023381 | 0.021909 | 0.017512 | 0.02137 | 0.024221 | 0.038341 | 0.010191 | 0.055377 | 0.023166 | 0.030196 | 0.053601 | 0.023166 | 0.002591 | 0.015055 | 0.046332 | 0.099757 | 0.024221 | 0.074749 | 0.0083666 | 0.030111 | 0.031143 | 0.029363
| Rel.Std.Dev. |  | 0.5657% | 0.9076% | 1.415% | 8.785% | 1.028% | 0.9781% | 0.8234% | 0.9947% | 1.011% | 1.763% | 0.4682% | 2.36% | 1.021% | 1.421% | 2.568% | 1.044% | 0.1159% | 0.6885% | 2.119% | 4.711% | 1.168% | 3.354% | 0.3829% | 1.352% | 1.441% | 1.324%
| PDM3 |  | -2.32% | -0.7299% | -2.902% | 9.793% | 4.157% | 2.63% | -2.563% | -1.57% | 9.808% | -0.3481% | -0.2717% | 7.517% | 3.928% | -2.619% | -4.357% | 1.637% | 2.46% | 0.1865% | 0.1865% | -2.99% | -5.006% | 2.119% | 0.1101% | 2.019% | -0.9658% | 1.63%
| Z-Score (Absolute) |  | 0.7 | 0.25 | 0.86 | 2.72 | 1.13 | 0.7 | 0.76 | 0.48 | 2.73 | 0.14 | 0.12 | 2.08 | 1.07 | 0.78 | 1.27 | 0.42 | 0.65 | 0.01 | 0.01 | 0.89 | 1.45 | 0.56 | 0.01 | 0.53 | 0.31 | 0.42
| Analytical results for Nb in OREAS 506b (Certified Value 11.1 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 11.39 | 10.9 | 10.8836 | 11.47 | 4.5 | 10.8 | 11.6 | 12.2 | 11.4 | 9.8 | 11.4 | 12.1 | 10.9 | 11.0954 | 11 | 10.8 | 13.3 | 8 | 11.6 | 10.4 | 10.4 | 10.18 |  |  |  |  | 1
| 1 | 2 | 11.53 | 10.8 | 10.7572 | 12.05 | 4.46 | 10.7 | 11.4 | 12.6 | 11.2 | 10.4 | 12 | 12.1 | 10.9 | 10.8917 | 10 | 11.4 | 12.9 | 8 | 11.7 | 10.5 | 9.6 | 9.66 |  |  |  |  | 21
| 1 | 3 | 11.54 | 10.9 | 10.6887 | 11.96 | 4.66 | 10.7 | 11.6 | 12.2 | 11.1 | 10.9 | 11.7 | 11.9 | 10.5 | 11.0135 | 10 | 11.2 | 13.6 | 8 | 11.5 | 10.9 | 10.7 | 9.92 |  |  |  |  | 16
| 1 | 4 | 11.49 | 10.9 | 10.9579 | 11.58 | 4.31 | 10.9 | 11.8 | 12.1 | 11.5 | 10.3 | 11.4 | 11.7 | 10.6 | 11.1038 | 10 | 11.3 | 13.6 | 8 | 12 | 10.4 | 10.1 | 9.75 |  |  |  |  | 11.1302
| 1 | 5 | 11.54 | 10.7 | 10.4969 | 11.76 | 4.73 | 11.1 | 11.4 | 11.8 | 11.5 | 10.6 | 12 | 11.2 | 10.5 | 10.8675 | 10 | 11.1 | 12.7 | 9 | 11.9 | 10.3 | 12.2 | 10.01 |  |  |  |  | 42
| 1 | 6 | 11.5 | 10.9 | 10.556 | 11.63 | 4.46 | 10.5 | 11.6 | 12 | 11.2 | 10.4 | 11.2 | 12.2 | 10.8 | 11.4012 | 10 | 11.1 | 12.7 | 8 | 11 | 10.5 | 12.7 | 10.07
| Mean |  | 11.4983 | 10.85 | 10.7234 | 11.7417 | 4.52 | 10.7833 | 11.5667 | 12.15 | 11.3167 | 10.4 | 11.6167 | 11.8667 | 10.7 | 11.0622 | 10.1667 | 11.15 | 13.1333 | 8.1667 | 11.6167 | 10.5 | 10.95 | 9.9317
| Median |  | 11.515 | 10.9 | 10.7229 | 11.695 | 4.48 | 10.75 | 11.6 | 12.15 | 11.3 | 10.4 | 11.55 | 12 | 10.7 | 11.0545 | 10 | 11.15 | 13.1 | 8 | 11.65 | 10.45 | 10.55 | 9.965
| Std Dev. |  | 0.057067 | 0.083666 | 0.18016 | 0.22605 | 0.15192 | 0.20412 | 0.15055 | 0.26646 | 0.17224 | 0.36332 | 0.33714 | 0.37238 | 0.18974 | 0.19344 | 0.40825 | 0.20736 | 0.42269 | 0.40825 | 0.35449 | 0.20976 | 1.2276 | 0.19692
| Rel.Std.Dev. |  | 0.4963% | 0.7711% | 1.68% | 1.925% | 3.361% | 1.893% | 1.302% | 2.193% | 1.522% | 3.493% | 2.902% | 3.138% | 1.773% | 1.749% | 4.016% | 1.86% | 3.218% | 4.999% | 3.052% | 1.998% | 11.21% | 1.983%
| PDM3 |  | 3.307% | -2.518% | -3.655% | 5.494% | -59.39% | -3.117% | 3.921% | 9.162% | 1.675% | -6.561% | 4.371% | 6.617% | -3.865% | -0.6112% | -8.657% | 0.1777% | 18% | -26.63% | 4.371% | -5.662% | -1.619% | -10.77%
| Z-Score (Absolute) |  | 0.54 | 0.36 | 0.53 | 0.88 | 9.15 | 0.45 | 0.64 | 1.45 | 0.29 | 0.98 | 0.71 | 1.06 | 0.56 | 0.06 | NA | 0.06 | 2.82 | NA | 0.71 | 0.84 | 0.22 | 1.63
Results from laboratories 2.19 and 2.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 506b (Certified Value 28.9 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.26; 2.10: 2.27 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*MS; 2.10: 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 2.01: 31.7; 2.02: 29.6; 2.03: 27.9151; 2.04: 28.1673; 2.06: 28.4; 2.07: 26.3; 2.08: 27.9; 2.09: 17; 2.10: 28.2 |  |  |  |  | 1
| 1 | 2 | 2.01: 32.52; 2.02: 32.6; 2.03: 28.3805; 2.04: 28.3171; 2.06: 28.5; 2.07: 26.5; 2.08: 30.7; 2.09: 18; 2.10: 26.5 |  |  |  |  | 22
| 1 | 3 | 2.01: 31.98; 2.02: 32.7; 2.03: 27.5083; 2.04: 28.2248; 2.06: 25.3; 2.07: 27.5; 2.08: 28.9; 2.09: 19; 2.10: 29.4 |  |  |  |  | 16
| 1 | 4 | 2.01: 32.15; 2.02: 27.1; 2.03: 29.1061; 2.04: 28.1439; 2.06: 26.9; 2.07: 26.9; 2.08: 30.6; 2.09: 17; 2.10: 28.7 |  |  |  |  | 28.9279
| 1 | 5 | 2.01: 32.53; 2.02: 32.9; 2.03: 28.8295; 2.04: 28.114; 2.06: 26.3; 2.07: 25.7; 2.08: 31.4; 2.09: 19; 2.10: 27.9 |  |  |  |  | 43
| 1 | 6 | 2.01: 32.11; 2.02: 28.1; 2.03: 30.2857; 2.04: 28.1556; 2.06: 27.9; 2.07: 25.6; 2.08: 30; 2.09: 18; 2.10: 29.4
| Mean |  | 2.01: 32.165; 2.02: 30.5; 2.03: 28.6709; 2.04: 28.1871; 2.06: 27.2167; 2.07: 26.4167; 2.08: 29.9167; 2.09: 18; 2.10: 28.35
| Median |  | 2.01: 32.13; 2.02: 31.1; 2.03: 28.605; 2.04: 28.1614; 2.06: 27.4; 2.07: 26.4; 2.08: 30.3; 2.09: 18; 2.10: 28.45
| Std Dev. |  | 2.01: 0.3203; 2.02: 2.5745; 2.03: 0.98278; 2.04: 0.073375; 2.06: 1.275; 2.07: 0.72226; 2.08: 1.2952; 2.09: 0.89443; 2.10: 1.0932
| Rel.Std.Dev. |  | 2.01: 0.9958%; 2.02: 8.441%; 2.03: 3.428%; 2.04: 0.2603%; 2.06: 4.685%; 2.07: 2.734%; 2.08: 4.33%; 2.09: 4.969%; 2.10: 3.856%
| PDM3 |  | 2.01: 11.19%; 2.02: 5.435%; 2.03: -0.8885%; 2.04: -2.561%; 2.06: -5.915%; 2.07: -8.681%; 2.08: 3.418%; 2.09: -37.78%; 2.10: -1.998%
| Z-Score (Absolute) |  | 2.01: 1.64; 2.02: 0.93; 2.03: 0.14; 2.04: 0.07; 2.06: 0.49; 2.07: 0.83; 2.08: 0.67; 2.09: 4.45; 2.10: 0
| Analytical results for Ni in OREAS 506b (Certified Value 15.7 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 16.4 | 16 | 15.2898 | 17.6 | 20.68 | 16.4 | 15.6 | 14.7 | 18.7 | 15.4 | 16 | 16.9 | 16.6 | 15.7083 | 16 | 15.2 | 15 | 15.4 | 15.5 | 15 | 17.7 | 16 | 12.9 | 15.7 |  |  | 1
| 1 | 2 | 16.6 | 16 | 14.8368 | 19.2 | 20.98 | 16 | 15.5 | 15.1 | 16.4 | 15.8 | 16 | 15.7 | 16.2 | 15.7734 | 16 | 15.9 | 15 | 15.8 | 16.7 | 14.6 | 17.7 | 18 | 15 | 15.1 |  |  | 23
| 1 | 3 | 16.3 | 16 | 15.3072 | 19.6 | 21.25 | 16.1 | 15.8 | 14.7 | 15.8 | 16.6 | 16 | 17 | 15.9 | 15.8447 | 16 | 15.7 | 14 | 14.8 | 16.4 | 14.7 | 18.4 | 15 | 15.5 | 15.5 |  |  | 16
| 1 | 4 | 16.4 | 16 | 15.2298 | 17.1 | 21.02 | 16.5 | 15.8 | 14.8 | 16.6 | 15.4 | 16 | 16.5 | 16.5 | 15.7865 | 16 | 16.6 | 15 | 15.6 | 16 | 14.7 | 17.3 | 17 | 14.8 | 15.3 |  |  | 15.6991
| 1 | 5 | 16.5 | 16 | 14.7638 | 19.4 | 21.58 | 16.8 | 15.3 | 15.2 | 16.4 | 16.2 | 16 | 15.2 | 15.7 | 15.3193 | 15 | 15.6 | 15 | 15.5 | 16.1 | 14.5 | 17.8 | 16 | 15.1 | 15.7 |  |  | 44
| 1 | 6 | 16.7 | 16 | 14.9578 | 18.1 | 21.07 | 15.5 | 15.6 | 14.6 | 15.4 | 16.2 | 16 | 16.5 | 16.5 | 16.0965 | 16 | 16 | 15 | 15.8 | 15.9 | 14.8 | 18 | 15 | 14.5 | 15.8
| Mean |  | 16.4833 | 16 | 15.0642 | 18.5 | 21.0967 | 16.2167 | 15.6 | 14.85 | 16.55 | 15.9333 | 16 | 16.3 | 16.2333 | 15.7548 | 15.8333 | 15.8333 | 14.8333 | 15.4833 | 16.1 | 14.7167 | 17.8167 | 16.1667 | 14.6333 | 15.5167
| Median |  | 16.45 | 16 | 15.0938 | 18.65 | 21.045 | 16.25 | 15.6 | 14.75 | 16.4 | 16 | 16 | 16.5 | 16.35 | 15.7799 | 16 | 15.8 | 15 | 15.55 | 16.05 | 14.7 | 17.75 | 16 | 14.9 | 15.6
| Std Dev. |  | 0.1472 | 0 | 0.24107 | 1.0431 | 0.30031 | 0.4535 | 0.18974 | 0.2429 | 1.145 | 0.48442 | 0 | 0.70711 | 0.36697 | 0.25219 | 0.40825 | 0.46762 | 0.40825 | 0.37103 | 0.41473 | 0.17224 | 0.3656 | 1.169 | 0.91141 | 0.27142
| Rel.Std.Dev. |  | 0.893% | 0% | 1.6% | 5.638% | 1.423% | 2.797% | 1.216% | 1.636% | 6.918% | 3.04% | 0% | 4.338% | 2.261% | 1.601% | 2.578% | 2.953% | 2.752% | 2.396% | 2.576% | 1.17% | 2.052% | 7.231% | 6.228% | 1.749%
| PDM3 |  | 4.995% | 1.917% | -4.044% | 17.84% | 34.38% | 3.297% | -0.6313% | -5.409% | 5.42% | 1.492% | 1.917% | 3.828% | 3.403% | 0.3546% | 0.855% | 0.855% | -5.515% | -1.374% | 2.554% | -6.258% | 13.49% | 2.978% | -6.789% | -1.162%
| Z-Score (Absolute) |  | 0.82 | NA | 1.3 | 3.85 | 7.74 | 0.42 | 0.5 | 1.62 | 0.92 | 0 | NA | 0.55 | 0.45 | 0.27 | NA | 0.15 | NA | 0.67 | 0.25 | 1.82 | 2.82 | NA | 1.95 | 0.62
Results from laboratories 2.02, 2.12, 2.16, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 506b (Certified Value 0.085 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.0832 | 0.085 | 0.084405 | 0.083 | 0.277 | 0.085 | 0.083 | 0.087 | 0.089 | 0.083 | 0.086 | 0.095 | 0.0796 | 0.084803 | 0.091 | 0.0843 | 0.0857 | 0.0943 | 0.07 | 0.0819 | 0.0864 | 0.09 | 0.079 | 0.0802 | 0.081 |  | 1
| 1 | 2 | 0.0844 | 0.085 | 0.083909 | 0.087 | 0.313 | 0.087 | 0.08 | 0.086 | 0.089 | 0.084 | 0.087 | 0.094 | 0.0825 | 0.085407 | 0.094 | 0.0836 | 0.0841 | 0.0881 | 0.072 | 0.0861 | 0.0864 | 0.09 | 0.079 | 0.0821 | 0.077 |  | 24
| 1 | 3 | 0.0841 | 0.085 | 0.082287 | 0.089 | 0.316 | 0.087 | 0.081 | 0.086 | 0.089 | 0.087 | 0.085 | 0.091 | 0.0789 | 0.084765 | 0.092 | 0.0835 | 0.0839 | 0.0822 | 0.077 | 0.0842 | 0.0866 | 0.09 | 0.08 | 0.083 | 0.08 |  | 16
| 1 | 4 | 0.0836 | 0.086 | 0.085165 | 0.081 | 0.326 | 0.085 | 0.084 | 0.085 | 0.09 | 0.084 | 0.086 | 0.092 | 0.0807 | 0.08438 | 0.092 | 0.0837 | 0.0844 | 0.0883 | 0.069 | 0.0843 | 0.0878 | 0.09 | 0.08 | 0.0813 | 0.08 |  | 0.084725
| 1 | 5 | 0.0844 | 0.084 | 0.079109 | 0.085 | 0.317 | 0.087 | 0.082 | 0.085 | 0.09 | 0.084 | 0.086 | 0.095 | 0.0777 | 0.083061 | 0.09 | 0.0822 | 0.0849 | 0.0809 | 0.068 | 0.0859 | 0.0866 | 0.09 | 0.08 | 0.0822 | 0.081 |  | 45
| 1 | 6 | 0.0856 | 0.084 | 0.079964 | 0.084 | 0.305 | 0.086 | 0.082 | 0.084 | 0.089 | 0.084 | 0.085 | 0.094 | 0.082 | 0.085714 | 0.096 | 0.0853 | 0.0851 | 0.0864 | 0.078 | 0.0855 | 0.0874 | 0.09 | 0.078 | 0.0822 | 0.08
| Mean |  | 0.084217 | 0.084833 | 0.082473 | 0.084833 | 0.309 | 0.086167 | 0.082 | 0.0855 | 0.089333 | 0.084333 | 0.085833 | 0.0935 | 0.080233 | 0.084688 | 0.0925 | 0.083767 | 0.084683 | 0.0867 | 0.072333 | 0.08465 | 0.086867 | 0.09 | 0.079333 | 0.081833 | 0.079833
| Median |  | 0.08425 | 0.085 | 0.083098 | 0.0845 | 0.3145 | 0.0865 | 0.082 | 0.0855 | 0.089 | 0.084 | 0.086 | 0.094 | 0.08015 | 0.084784 | 0.092 | 0.08365 | 0.08465 | 0.08725 | 0.071 | 0.0849 | 0.0866 | 0.09 | 0.0795 | 0.08215 | 0.08
| Std Dev. |  | 0.00082563 | 0.00075277 | 0.0024775 | 0.0028577 | 0.017076 | 0.00098319 | 0.0014142 | 0.0010488 | 0.0005164 | 0.0013663 | 0.00075277 | 0.0016432 | 0.0018479 | 0.00093081 | 0.0021679 | 0.0010191 | 0.00067651 | 0.004822 | 0.0042269 | 0.0015668 | 0.00058878 | 1.5202e-17 | 0.0008165 | 0.00096471 | 0.001472
| Rel.Std.Dev. |  | 0.9804% | 0.8874% | 3.004% | 3.369% | 5.526% | 1.141% | 1.725% | 1.227% | 0.5781% | 1.62% | 0.877% | 1.757% | 2.303% | 1.099% | 2.344% | 1.217% | 0.7989% | 5.562% | 5.844% | 1.851% | 0.6778% | 1.689e-14% | 1.029% | 1.179% | 1.844%
| PDM3 |  | -0.5997% | 0.1281% | -2.658% | 0.1281% | 264.7% | 1.702% | -3.216% | 0.915% | 5.439% | -0.462% | 1.308% | 10.36% | -5.301% | -0.04304% | 9.177% | -1.131% | -0.0489% | 2.331% | -14.63% | -0.08824% | 2.528% | 6.226% | -6.363% | -3.413% | -5.773%
| Z-Score (Absolute) |  | 0.15 | 0.04 | 0.69 | 0.04 | 69.44 | 0.46 | 0.83 | 0.25 | 1.44 | 0.11 | 0.35 | 2.73 | 1.38 | 0 | 2.42 | 0.29 | 0 | 0.62 | 3.82 | 0.01 | 0.67 | 1.64 | 1.66 | 0.88 | 1.5
| Analytical results for Pb in OREAS 506b (Certified Value 56 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 57 | 56 | 56.2866 | 54.71 | 59.22 | 56.8 | 51 | 56.8 | 62.2 | 54.5 | 57.7 | 100 | 54.9 | 53.5747 | 55 | 53.2 | 54.52 | 50 | 53.6 | 51 | 60 | 56 | 49 | 46 | 54.39 |  | 1
| 1 | 2 | 57.6 | 57 | 55.9581 | 59.45 | 59.18 | 57 | 51.7 | 56.8 | 57.3 | 56.2 | 57.5 | 125 | 56.7 | 57.019 | 60 | 56.3 | 52.69 | 51 | 54 | 52 | 59.8 | 54 | 48 | 47 | 53.66 |  | 25
| 1 | 3 | 57.8 | 56 | 57.1128 | 57.77 | 58.62 | 55.8 | 51.2 | 56.2 | 56.9 | 58.7 | 56.3 | 141 | 53.7 | 54.6214 | 57 | 56.2 | 55.32 | 50 | 53.2 | 53 | 60.7 | 55 | 49 | 46 | 56.14 |  | 16
| 1 | 4 | 57.4 | 56 | 56.6866 | 55.56 | 58.31 | 55.1 | 56.3 | 55.7 | 59.3 | 55.7 | 56.3 | 139 | 54.4 | 57.0341 | 57 | 57.1 | 51 | 50 | 53.1 | 52 | 59.4 | 54 | 49 | 47 | 54.09 |  | 55.6601
| 1 | 5 | 58.1 | 56 | 54.609 | 59.08 | 58.36 | 57.2 | 50.7 | 56.6 | 57.4 | 56.9 | 56.5 | 90.8 | 54.6 | 57.1917 | 54 | 54.2 | 53.91 | 51 | 54.4 | 52 | 59.4 | 54 | 48 | 47 | 54.84 |  | 46
| 1 | 6 | 56.9 | 55 | 54.9522 | 63 | 58.71 | 55.7 | 50.9 | 55.6 | 56.8 | 56.8 | 57.4 | 153 | 54.4 | 57.2182 | 62 | 56.5 | 53.25 | 51 | 54.4 | 53 | 60.4 | 55 | 47 | 46 | 56.31
| Mean |  | 57.4667 | 56 | 55.9342 | 58.2617 | 58.7333 | 56.2667 | 51.9667 | 56.2833 | 58.3167 | 56.4667 | 56.95 | 124.8 | 54.7833 | 56.1098 | 57.5 | 55.5833 | 53.4483 | 50.5 | 53.7833 | 52.1667 | 59.95 | 54.6667 | 48.3333 | 46.5 | 54.905
| Median |  | 57.5 | 56 | 56.1224 | 58.425 | 58.665 | 56.3 | 51.1 | 56.4 | 57.35 | 56.5 | 56.95 | 132 | 54.5 | 57.0266 | 57 | 56.25 | 53.58 | 50.5 | 53.8 | 52 | 59.9 | 54.5 | 48.5 | 46.5 | 54.615
| Std Dev. |  | 0.46332 | 0.63246 | 0.97983 | 2.9896 | 0.392 | 0.84774 | 2.15 | 0.53821 | 2.1085 | 1.401 | 0.65038 | 24.6187 | 1.0187 | 1.5951 | 3.0166 | 1.525 | 1.5148 | 0.54772 | 0.57417 | 0.75277 | 0.5282 | 0.8165 | 0.8165 | 0.54772 | 1.0939
| Rel.Std.Dev. |  | 0.8062% | 1.129% | 1.752% | 5.131% | 0.6674% | 1.507% | 4.137% | 0.9562% | 3.616% | 2.481% | 1.142% | 19.73% | 1.859% | 2.843% | 5.246% | 2.744% | 2.834% | 1.085% | 1.068% | 1.443% | 0.8811% | 1.494% | 1.689% | 1.178% | 1.992%
| PDM3 |  | 3.246% | 0.6106% | 0.4924% | 4.674% | 5.521% | 1.09% | -6.636% | 1.12% | 4.773% | 1.449% | 2.317% | 124.2% | -1.575% | 0.808% | 3.306% | -0.138% | -3.974% | -9.271% | -3.372% | -6.276% | 7.707% | -1.785% | -13.16% | -16.46% | -1.357%
| Z-Score (Absolute) |  | 0.66 | 0 | 0.03 | 1.02 | 1.23 | 0.12 | 1.81 | 0.13 | 1.04 | 0.21 | 0.43 | 30.93 | 0.55 | 0.05 | 0.67 | 0.19 | 1.15 | 2.47 | 1 | 1.72 | 1.78 | 0.6 | 3.45 | 4.27 | 0.49
| Analytical results for Pr in OREAS 506b (Certified Value 7.46 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  | 2
3
| 1 | 1 | 2.01: 8.3; 2.02: 7.7; 2.03: 7.3525; 2.04: 7.4655; 2.06: 7.37; 2.07: 6.61; 2.08: 7.4; 2.09: 7 |  | 1
| 1 | 2 | 2.01: 8.38; 2.02: 8.5; 2.03: 7.4036; 2.04: 7.3987; 2.06: 7.42; 2.07: 6.67; 2.08: 8.1; 2.09: 6.8 |  | 26
| 1 | 3 | 2.01: 8.37; 2.02: 8.6; 2.03: 7.2141; 2.04: 7.4399; 2.06: 6.5; 2.07: 6.92; 2.08: 7.6; 2.09: 7.2 |  | 16
| 1 | 4 | 2.01: 8.51; 2.02: 6.7; 2.03: 7.5261; 2.04: 7.4312; 2.06: 6.9; 2.07: 7.08; 2.08: 8; 2.09: 7.2 |  | 7.4581
| 1 | 5 | 2.01: 8.58; 2.02: 8.5; 2.03: 7.4052; 2.04: 7.4789; 2.06: 6.62; 2.07: 6.37; 2.08: 8.1; 2.09: 6.9 |  | 47
| 1 | 6 | 2.01: 8.3; 2.02: 7; 2.03: 7.8168; 2.04: 7.4309; 2.06: 7.01; 2.07: 6.65; 2.08: 7.9; 2.09: 7.3
| Mean |  | 2.01: 8.4067; 2.02: 7.8333; 2.03: 7.4531; 2.04: 7.4409; 2.06: 6.97; 2.07: 6.7167; 2.08: 7.85; 2.09: 7.0667
| Median |  | 2.01: 8.375; 2.02: 8.1; 2.03: 7.4044; 2.04: 7.4356; 2.06: 6.955; 2.07: 6.66; 2.08: 7.95; 2.09: 7.1
| Std Dev. |  | 2.01: 0.11448; 2.02: 0.83347; 2.03: 0.20474; 2.04: 0.028372; 2.06: 0.37747; 2.07: 0.24961; 2.08: 0.2881; 2.09: 0.19664
| Rel.Std.Dev. |  | 2.01: 1.362%; 2.02: 10.64%; 2.03: 2.747%; 2.04: 0.3813%; 2.06: 5.416%; 2.07: 3.716%; 2.08: 3.67%; 2.09: 2.783%
| PDM3 |  | 2.01: 12.72%; 2.02: 5.032%; 2.03: -0.06712%; 2.04: -0.2306%; 2.06: -6.544%; 2.07: -9.941%; 2.08: 5.255%; 2.09: -5.248%
| Z-Score (Absolute) |  | 2.01: 1.64; 2.02: 0.66; 2.03: 0.01; 2.04: 0.01; 2.06: 0.81; 2.07: 1.25; 2.08: 0.69; 2.09: 0.65
| Analytical results for Pt in OREAS 506b (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab |  | 1
| No | No | 2.01: 2.20 |  | ppb
|  |  | 2.01: 4A*OES/MS |  | 2
2
| 1 | 1 | 2.01: < 5 |  | 1
| 1 | 2 | 2.01: < 5 |  | 1
| 1 | 3 | 2.01: < 5 |  | 16
| 1 | 4 | 2.01: < 5 |  | < 5
| 1 | 5 | 2.01: < 5 |  | 37
| 1 | 6 | 2.01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Rb in OREAS 506b (Certified Value 149 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | 2.27 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
0
| 1 | 1 | 157.62 | 148 | 136.444 | 102 | 185.2 | 152 | 140 | 151 | 155 | 135.5 | 155 | 102 | 149.6 | 156.6161 | 152.4824 | 152 | 137 | 151 | 151 | 104 | 145.3 | [5 empty] | 1
| 1 | 2 | 162.12 | 151 | 131.7482 | 111.2 | 182.92 | 146 | 138.5 | 159.5 | 153.5 | 140 | 152 | 98.6 | 148.5 | 154.8785 | 152.475 | 154 | 133 | 152 | 147 | 106 | 143.6 | [5 empty] | 62
| 1 | 3 | 158.12 | 148 | 137.029 | 109.4 | 183.44 | 148 | 138 | 155.5 | 150 | 144.5 | 154 | 101 | 145 | 153.427 | 152.5118 | 144 | 127 | 151.3 | 153 | 111 | 144 | [5 empty] | 16
| 1 | 4 | 157.31 | 147 | 147.8159 | 74.5 | 186.34 | 144.5 | 148.5 | 156 | 158.5 | 137.5 | 152 | 102 | 147.2 | 152.9392 | 152.4878 | 148 | 143 | 150.2 | 146 | 110 | 139.5 | [5 empty] | 149.4494
| 1 | 5 | 157.84 | 148 | 136.0649 | 120 | 186.29 | 154.5 | 137.5 | 151.5 | 156 | 143 | 152 | 107 | 144.7 | 154.4254 | 152.4354 | 147 | 140 | 153.9 | 149 | 166 | 146.8 | [5 empty] | 48
| 1 | 6 | 161.6 | 149 | 145.7163 | 102.9 | 189.91 | 143 | 145.5 | 157 | 154.5 | 135 | 153 | 97.1 | 147.3 | 157.5518 | 152.4742 | 148 | 142 | 151.8 | 151 | 166 | 148.2
| Mean |  | 159.1017 | 148.5 | 139.1364 | 103.3333 | 185.6833 | 148 | 141.3333 | 155.0833 | 154.5833 | 139.25 | 153 | 101.2833 | 147.05 | 154.973 | 152.4778 | 148.8333 | 137 | 151.7 | 149.5 | 127.1667 | 144.5667
| Median |  | 157.98 | 148 | 136.7365 | 106.15 | 185.745 | 147 | 139.25 | 155.75 | 154.75 | 138.75 | 152.5 | 101.5 | 147.25 | 154.6519 | 152.4787 | 148 | 138.5 | 151.55 | 150 | 110.5 | 144.65
| Std Dev. |  | 2.1593 | 1.3784 | 6.2344 | 15.555 | 2.5126 | 4.4609 | 4.568 | 3.2774 | 2.8181 | 3.9338 | 1.2649 | 3.4202 | 1.9191 | 1.7992 | 0.024883 | 3.6009 | 6.0992 | 1.2522 | 2.6646 | 30.1889 | 3.0217
| Rel.Std.Dev. |  | 1.357% | 0.9282% | 4.481% | 15.05% | 1.353% | 3.014% | 3.232% | 2.113% | 1.823% | 2.825% | 0.8267% | 3.377% | 1.305% | 1.161% | 0.01632% | 2.419% | 4.452% | 0.8254% | 1.782% | 23.74% | 2.09%
| PDM3 |  | 6.459% | -0.6353% | -6.901% | -30.86% | 24.24% | -0.9698% | -5.431% | 3.77% | 3.435% | -6.825% | 2.376% | -32.23% | -1.605% | 3.696% | 2.026% | -0.4122% | -8.33% | 1.506% | 0.03387% | -14.91% | -3.267%
| Z-Score (Absolute) |  | 1.1 | 0 | 0.98 | 4.71 | 3.87 | 0.05 | 0.75 | 0.69 | 0.63 | 0.96 | 0.47 | 4.92 | 0.15 | 0.67 | 0.41 | 0.03 | 1.2 | 0.33 | 0.1 | 2.22 | 0.41
| Analytical results for Re in OREAS 506b (Certified Value 0.035 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.15 | 2.20 | 2.21 | 2.23 | 2.25 | 2.26 | 2.27 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [9 empty] | 3
3
| 1 | 1 | 0.034 | < 0.05 | < 0.05 | 0.305 | 0.043 | 0.032 | 0.037 | 0.036 | 0.031 | 0.033 | 0.070377 | 0.0341 | 0.032 | 0.04 | 0.034 | NA | 0.034 | [9 empty] | 1
| 1 | 2 | 0.037 | < 0.05 | < 0.05 | 0.353 | 0.039 | 0.031 | 0.036 | 0.038 | 0.035 | 0.033 | 0.063748 | 0.0333 | 0.029 | 0.042 | 0.036 | NA | 0.032 | [9 empty] | 63
| 1 | 3 | 0.039 | < 0.05 | < 0.05 | 0.312 | 0.033 | 0.033 | 0.036 | 0.043 | 0.032 | 0.028 | 0.066673 | 0.0338 | 0.034 | 0.043 | 0.034 | NA | 0.035 | [9 empty] | 16
| 1 | 4 | 0.037 | < 0.05 | < 0.05 | 0.311 | 0.036 | 0.033 | 0.038 | 0.038 | 0.031 | 0.033 | 0.0758 | 0.0327 | 0.037 | 0.043 | 0.036 | NA | 0.032 | [9 empty] | 0.035243
| 1 | 5 | 0.035 | < 0.05 | < 0.05 | 0.333 | 0.037 | 0.033 | 0.034 | 0.042 | 0.032 | 0.032 | 0.074199 | 0.0332 | 0.028 | 0.044 | 0.036 | NA | 0.035 | [9 empty] | 49
| 1 | 6 | 0.034 | < 0.05 | < 0.05 | 0.285 | 0.038 | 0.031 | 0.037 | 0.033 | 0.032 | 0.031 | 0.068879 | 0.0354 | 0.036 | 0.043 | 0.036 | NA | 0.034
| Mean |  | 0.036 |  |  | 0.3165 | 0.037667 | 0.032167 | 0.036333 | 0.038333 | 0.032167 | 0.031667 | 0.069946 | 0.03375 | 0.032667 | 0.0425 | 0.035333 |  | 0.033667
| Median |  | 0.036 |  |  | 0.3115 | 0.0375 | 0.0325 | 0.0365 | 0.038 | 0.032 | 0.0325 | 0.069628 | 0.03355 | 0.033 | 0.043 | 0.036 |  | 0.034
| Std Dev. |  | 0.002 |  |  | 0.023578 | 0.0033267 | 0.00098319 | 0.0013663 | 0.0037238 | 0.001472 | 0.0019664 | 0.0045345 | 0.00094393 | 0.0036697 | 0.0013784 | 0.0010328 |  | 0.0013663
| Rel.Std.Dev. |  | 5.556% |  |  | 7.449% | 8.832% | 3.057% | 3.76% | 9.714% | 4.576% | 6.21% | 6.483% | 2.797% | 11.23% | 3.243% | 2.923% |  | 4.058%
| PDM3 |  | 2.148% |  |  | 798% | 6.877% | -8.729% | 3.094% | 8.768% | -8.729% | -10.15% | 98.47% | -4.236% | -7.31% | 20.59% | 0.2562% |  | -4.473%
| Z-Score (Absolute) |  | 0.37 | 2.45 | 2.45 | 72.43 | 0.8 | 0.61 | 0.46 | 0.97 | 0.61 | 0.74 | 9.09 | 0.2 | 0.48 | 2.04 | 0.2 | NA | 0.22
| Analytical results for S in OREAS 506b (Certified Value 0.616 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.6002 | 0.617 | 0.61521 | 0.58 | 0.44 | 0.62 | 0.63 | 0.63 | 0.68 | 0.62 | 0.628 | 0.65 | 0.58 | 0.62398 | 0.638 | 0.62 | 0.58785 | 0.64 | 0.597 | 0.621 | 0.61 | 0.586 | 0.61 | 0.6 | 0.5785 | 0.55 | 1
| 1 | 2 | 0.6022 | 0.621 | 0.60374 | 0.61 | 0.44 | 0.64 | 0.63 | 0.62 | 0.67 | 0.63 | 0.635 | 0.64 | 0.59 | 0.61905 | 0.654 | 0.62 | 0.58805 | 0.63 | 0.583 | 0.774 | 0.6 | 0.586 | 0.61 | 0.61 | 0.5945 | 0.55 | 7
| 1 | 3 | 0.5943 | 0.625 | 0.60286 | 0.62 | 0.43 | 0.64 | 0.63 | 0.62 | 0.69 | 0.64 | 0.625 | 0.63 | 0.58 | 0.61796 | 0.653 | 0.62 | 0.58048 | 0.63 | 0.581 | 0.691 | 0.63 | 0.592 | 0.62 | 0.6 | 0.5998 | 0.55 | 16
| 1 | 4 | 0.5968 | 0.624 | 0.61762 | 0.59 | 0.43 | 0.64 | 0.61 | 0.63 | 0.69 | 0.63 | 0.635 | 0.64 | 0.6 | 0.61927 | 0.638 | 0.62 | 0.58979 | 0.63 | 0.605 | 0.692 | 0.62 | 0.5915 | 0.61 | 0.59 | 0.5962 | 0.55 | 0.61618
| 1 | 5 | 0.599 | 0.613 | 0.60141 | 0.6 | 0.43 | 0.64 | 0.63 | 0.61 | 0.68 | 0.62 | 0.633 | 0.62 | 0.58 | 0.59678 | 0.623 | 0.61 | 0.58424 | 0.64 | 0.559 | 0.752 | 0.62 | 0.587 | 0.61 | 0.63 | 0.5979 | 0.55 | 50
| 1 | 6 | 0.6039 | 0.618 | 0.60557 | 0.6 | 0.43 | 0.63 | 0.63 | 0.6 | 0.67 | 0.62 | 0.629 | 0.64 | 0.61 | 0.63659 | 0.664 | 0.63 | 0.58616 | 0.64 | 0.602 | 0.706 | 0.6 | 0.582 | 0.62 | 0.61 | 0.5884 | 0.54
| Mean |  | 0.5994 | 0.61967 | 0.60773 | 0.6 | 0.43333 | 0.635 | 0.62667 | 0.61833 | 0.68 | 0.62667 | 0.63083 | 0.63667 | 0.59 | 0.61894 | 0.645 | 0.62 | 0.5861 | 0.635 | 0.58783 | 0.706 | 0.61333 | 0.58742 | 0.61333 | 0.60667 | 0.59255 | 0.54833
| Median |  | 0.5996 | 0.6195 | 0.60465 | 0.6 | 0.43 | 0.64 | 0.63 | 0.62 | 0.68 | 0.625 | 0.631 | 0.64 | 0.585 | 0.61916 | 0.6455 | 0.62 | 0.587 | 0.635 | 0.59 | 0.699 | 0.615 | 0.5865 | 0.61 | 0.605 | 0.59535 | 0.55
| Std Dev. |  | 0.0035117 | 0.0045461 | 0.0069024 | 0.014142 | 0.005164 | 0.0083666 | 0.008165 | 0.01169 | 0.0089443 | 0.008165 | 0.0041191 | 0.010328 | 0.012649 | 0.012879 | 0.014751 | 0.0063246 | 0.0033325 | 0.0054772 | 0.017209 | 0.05364 | 0.012111 | 0.0037738 | 0.005164 | 0.013663 | 0.0079112 | 0.0040825
| Rel.Std.Dev. |  | 0.5859% | 0.7336% | 1.136% | 2.357% | 1.192% | 1.318% | 1.303% | 1.891% | 1.315% | 1.303% | 0.653% | 1.622% | 2.144% | 2.081% | 2.287% | 1.02% | 0.5686% | 0.8626% | 2.928% | 7.598% | 1.975% | 0.6424% | 0.842% | 2.252% | 1.335% | 0.7445%
| PDM3 |  | -2.723% | 0.5659% | -1.371% | -2.626% | -29.67% | 3.054% | 1.702% | 0.3495% | 10.36% | 1.702% | 2.378% | 3.325% | -4.249% | 0.448% | 4.677% | 0.62% | -4.882% | 3.054% | -4.6% | 14.58% | -0.462% | -4.668% | -0.462% | -1.544% | -3.835% | -11.01%
| Z-Score (Absolute) |  | 0.62 | 0.15 | 0.31 | 0.6 | 6.91 | 0.73 | 0.41 | 0.09 | 2.43 | 0.41 | 0.57 | 0.79 | 0.98 | 0.12 | 1.1 | 0.16 | 1.13 | 0.73 | 1.06 | 3.42 | 0.09 | 1.08 | 0.09 | 0.35 | 0.88 | 2.56
| Analytical results for Sb in OREAS 506b (Certified Value 1.36 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.45 | 1.3 | 1.531 | 1.25 | < 1 | 1.29 | 1.32 | 1.5 | 1.42 | 1.23 | 1.3 | 2 | 1.32 | 1.4098 | < 5 | 1.19 | < 5 | 1.3 | 1.33 | 2.39 | 1.5 | 1.17 | 1.4 | 1.15 | 1.18 |  | 1
| 1 | 2 | 1.46 | 1.3 | 1.5091 | 1.39 | < 1 | 1.35 | 1.32 | 1.54 | 1.46 | 1.29 | 1.32 | 3.25 | 1.39 | 1.4223 | < 5 | 1.55 | < 5 | 1.4 | 1.3 | 2.38 | 1.5 | 1.18 | 1.4 | 1.22 | 1.17 |  | 28
| 1 | 3 | 1.42 | 1.3 | 1.4829 | 1.33 | < 1 | 1.32 | 1.32 | 1.51 | 1.42 | 1.35 | 1.35 | 4.04 | 1.31 | 1.3628 | < 5 | 1.48 | < 5 | 1.4 | 1.43 | 2.4 | 1.5 | 1.25 | 1.4 | 1.3 | 1.16 |  | 16
| 1 | 4 | 1.52 | 1.3 | 1.5505 | 1.27 | < 1 | 1.24 | 1.4 | 1.56 | 1.47 | 1.36 | 1.4 | 1.36 | 1.38 | 1.447 | < 5 | 1.37 | < 5 | 1.3 | 1.49 | 2.06 | 1.42 | 1.19 | 1.5 | 1.19 | 1.15 |  | 1.36
| 1 | 5 | 1.43 | 1.3 | 1.4743 | 1.38 | < 1 | 1.25 | 1.34 | 2.09 | 1.46 | 1.34 | 1.31 | 1.47 | 1.28 | 1.361 | < 5 | 1.38 | < 5 | 1.3 | 1.39 | 2.24 | 1.6 | 1.22 | 1.5 | 1.25 | 1.18 |  | 51
| 1 | 6 | 1.35 | 1.3 | 1.5562 | 1.33 | < 1 | 1.23 | 1.34 | 1.46 | 1.43 | 1.4 | 1.29 | 3.72 | 1.3 | 1.4548 | < 5 | 1.48 | < 5 | 1.4 | 1.39 | 2.4 | 1.53 | 1.12 | 1.5 | 1.19 | 1.24
| Mean |  | 1.4383 | 1.3 | 1.5173 | 1.325 |  | 1.28 | 1.34 | 1.61 | 1.4433 | 1.3283 | 1.3283 | 2.64 | 1.33 | 1.4096 |  | 1.4083 |  | 1.35 | 1.3883 | 2.3117 | 1.5083 | 1.1883 | 1.45 | 1.2167 | 1.18
| Median |  | 1.44 | 1.3 | 1.5201 | 1.33 |  | 1.27 | 1.33 | 1.525 | 1.445 | 1.345 | 1.315 | 2.625 | 1.315 | 1.4161 |  | 1.43 |  | 1.35 | 1.39 | 2.385 | 1.5 | 1.185 | 1.45 | 1.205 | 1.175
| Std Dev. |  | 0.055648 | 0 | 0.034339 | 0.05648 |  | 0.048166 | 0.030984 | 0.23766 | 0.022509 | 0.059805 | 0.040702 | 1.176 | 0.044721 | 0.040374 |  | 0.12671 |  | 0.054772 | 0.06824 | 0.13776 | 0.058109 | 0.04446 | 0.054772 | 0.052789 | 0.031623
| Rel.Std.Dev. |  | 3.869% | 0% | 2.263% | 4.263% |  | 3.763% | 2.312% | 14.76% | 1.56% | 4.502% | 3.064% | 44.55% | 3.363% | 2.864% |  | 8.998% |  | 4.057% | 4.915% | 5.959% | 3.853% | 3.741% | 3.777% | 4.339% | 2.68%
| PDM3 |  | 5.758% | -4.413% | 11.57% | -2.575% |  | -5.884% | -1.472% | 18.38% | 6.126% | -2.33% | -2.33% | 94.11% | -2.207% | 3.649% |  | 3.552% |  | -0.7367% | 2.082% | 69.97% | 10.91% | -12.62% | 6.616% | -10.54% | -13.24%
| Z-Score (Absolute) |  | 0.24 | NA | 0.7 | 0.43 | 5.31 | 0.7 | 0.34 | 1.25 | 0.27 | 0.41 | 0.41 | 7.33 | 0.4 | 0.07 | 6.51 | 0.06 | 6.51 | NA | 0.06 | 5.39 | 0.65 | 1.24 | NA | 1.07 | 1.29
Results from laboratories 2.02, 2.20 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 506b (Certified Value 7.7 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 8.3 | 8 | 8.6271 | 9.4 | 5.85 | 7.7 | 6.9 | 7.4 | 8.4 | 7.1 | 8.1 | 8 | 7.4 | 7.5312 | 7.9 | 8.2 | 7.8036 | 8 | 8 | 7.5 | 7.5 | 7.52 | 7.6 | 8 | 7.2 |  | 1
| 1 | 2 | 8 | 8 | 8.3273 | 7.4 | 6.14 | 7.4 | 6.7 | 7.8 | 8.3 | 7.4 | 7.8 | 8.1 | 7.6 | 7.5599 | 8.3 | 7.9 | 7.8072 | 8 | 8 | 7.2 | 7.6 | 7.48 | 7.3 | 8 | 6.9 |  | 29
| 1 | 3 | 8.1 | 8 | 8.4798 | 7.4 | 6.11 | 7.4 | 6.7 | 7.7 | 8.2 | 7.7 | 8.2 | 8 | 7.2 | 7.5953 | 8 | 8.4 | 7.8477 | 7 | 8 | 7.3 | 7.5 | 7.4 | 7.8 | 8 | 7 |  | 16
| 1 | 4 | 8.1 | 8 | 8.5868 | 8.2 | 6.01 | 7.4 | 7.5 | 7.6 | 8.5 | 7.3 | 8.1 | 7.9 | 7.3 | 7.5545 | 7.8 | 8.1 | 7.8608 | 7 | 8 | 7.6 | 7.3 | 7.78 | 7.2 | 8 | 6.9 |  | 7.7035
| 1 | 5 | 8.2 | 8 | 8.6603 | 7.3 | 6.12 | 7.9 | 6.6 | 7.4 | 8.5 | 7.6 | 7.8 | 7.6 | 7.2 | 7.4235 | 7.9 | 8.1 | 7.8535 | 8 | 8 | 7.2 | 7.6 | 7.52 | 7.4 | 9 | 6.9 |  | 52
| 1 | 6 | 8.3 | 8 | 8.2983 | 8.7 | 6.02 | 7.3 | 7.4 | 7.5 | 8.2 | 7.3 | 7.6 | 8 | 7.5 | 7.7106 | 8.5 | 8.7 | 7.8913 | 7 | 8 | 7.5 | 7.7 | 7.36 | 7.5 | 8 | 7
| Mean |  | 8.1667 | 8 | 8.4966 | 8.0667 | 6.0417 | 7.5167 | 6.9667 | 7.5667 | 8.35 | 7.4 | 7.9333 | 7.9333 | 7.3667 | 7.5625 | 8.0667 | 8.2333 | 7.844 | 7.5 | 8 | 7.3833 | 7.5333 | 7.51 | 7.4667 | 8.1667 | 6.9833
| Median |  | 8.15 | 8 | 8.5333 | 7.8 | 6.065 | 7.4 | 6.8 | 7.55 | 8.35 | 7.35 | 7.95 | 8 | 7.35 | 7.5572 | 7.95 | 8.15 | 7.8506 | 7.5 | 8 | 7.4 | 7.55 | 7.5 | 7.45 | 8 | 6.95
| Std Dev. |  | 0.12111 | 0 | 0.15508 | 0.85713 | 0.10834 | 0.23166 | 0.38816 | 0.1633 | 0.13784 | 0.21909 | 0.23381 | 0.17512 | 0.1633 | 0.093136 | 0.27325 | 0.28048 | 0.033484 | 0.54772 | 0 | 0.17224 | 0.13663 | 0.14738 | 0.21602 | 0.40825 | 0.1169
| Rel.Std.Dev. |  | 1.483% | 0% | 1.825% | 10.63% | 1.793% | 3.082% | 5.572% | 2.158% | 1.651% | 2.961% | 2.947% | 2.207% | 2.217% | 1.232% | 3.387% | 3.407% | 0.4269% | 7.303% | 0% | 2.333% | 1.814% | 1.962% | 2.893% | 4.999% | 1.674%
| PDM3 |  | 6.013% | 3.849% | 10.3% | 4.714% | -21.57% | -2.425% | -9.565% | -1.776% | 8.392% | -3.94% | 2.984% | 2.984% | -4.372% | -1.83% | 4.714% | 6.878% | 1.824% | -2.642% | 3.849% | -4.156% | -2.209% | -2.512% | -3.074% | 6.013% | -9.348%
| Z-Score (Absolute) |  | 1.1 | NA | 1.7 | 0.92 | 2.77 | 0.08 | 1.08 | 0.01 | 1.43 | 0.3 | 0.67 | 0.67 | 0.36 | 0 | 0.92 | 1.22 | 0.51 | NA | NA | 0.33 | 0.05 | 0.1 | 0.17 | NA | 1.05
Results from laboratories 2.02, 2.19, 2.20 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 506b (Certified Value 5.02 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 5 | 5 | 5.2885 | 5 | 4.02 | 5 | 6 | 5 | 7 | 5 | 5 | 6 | 6 | 6.0224 | 5.2 | < 10 | 3.85 | 4 | 5 | 10 | 4.5 | 3.99 | 5.1 |  |  |  | 1
| 1 | 2 | 4.2 | 5 | 5.2023 | 5.8 | 4.03 | 6 | 6 | 5 | 7 | 5 | 5 | 6 | 5 | 5.6328 | 5.3 | < 10 | 4.35 | 4 | 5 | 9 | 4.6 | 4.36 | 4.8 |  |  |  | 0
| 1 | 3 | 4.7 | 5 | 5.0346 | 4.7 | 4.01 | 5 | 6 | 5 | 7 | 5 | 5 | 5 | 5 | 5.4877 | 5.8 | < 10 | 5.14 | 4 | 6 | 8 | 4.5 | 4.64 | 5 |  |  |  | 16
| 1 | 4 | 5.2 | 5 | 5.1255 | 5.2 | 4.02 | 6 | 7 | 5 | 7 | 5 | 5 | 5 | 6 | 6.1713 | 4.3 | < 10 | 4.06 | 4 | 6 | 10 | 4.7 | 4.5 | 5 |  |  |  | 5.0249
| 1 | 5 | 4.4 | 5 | 4.9923 | 6 | 3.99 | 5 | 6 | 5 | 6 | 5 | 5 | 5 | 6 | 5.6353 | 5.2 | < 10 | 3.59 | 4 | 6 | 9 | 5.5 | 4.6 | 4.8 |  |  |  | 53
| 1 | 6 | 3.8 | 5 | 5.2556 | 5.2 | 4.02 | 5 | 6 | 5 | 5 | 5 | 5 | 5 | 5 | 5.9983 | 5.4 | < 10 | 4.18 | 4 | 6 | 6 | 5.3 | 4.31 | 5.2
| Mean |  | 4.55 | 5 | 5.1498 | 5.3167 | 4.015 | 5.3333 | 6.1667 | 5 | 6.5 | 5 | 5 | 5.3333 | 5.5 | 5.8246 | 5.2 |  | 4.195 | 4 | 5.6667 | 8.6667 | 4.85 | 4.4 | 4.9833
| Median |  | 4.55 | 5 | 5.1639 | 5.2 | 4.02 | 5 | 6 | 5 | 7 | 5 | 5 | 5 | 5.5 | 5.8168 | 5.25 |  | 4.12 | 4 | 6 | 9 | 4.65 | 4.43 | 5
| Std Dev. |  | 0.52058 | 0 | 0.11987 | 0.4916 | 0.013784 | 0.5164 | 0.40825 | 0 | 0.83666 | 0 | 0 | 0.5164 | 0.54772 | 0.27408 | 0.49396 |  | 0.53302 | 0 | 0.5164 | 1.5055 | 0.43704 | 0.23891 | 0.16021
| Rel.Std.Dev. |  | 11.44% | 0% | 2.328% | 9.246% | 0.3433% | 9.682% | 6.62% | 0% | 12.87% | 0% | 0% | 9.682% | 9.959% | 4.705% | 9.499% |  | 12.71% | 0% | 9.113% | 17.37% | 9.011% | 5.43% | 3.215%
| PDM3 |  | -9.451% | -0.4953% | 2.486% | 5.807% | -20.1% | 6.138% | 22.72% | -0.4953% | 29.36% | -0.4953% | -0.4953% | 6.138% | 9.455% | 15.92% | 3.485% |  | -16.52% | -20.4% | 12.77% | 72.47% | -3.48% | -12.44% | -0.827%
| Z-Score (Absolute) |  | 0.91 | 0 | 0.3 | 0.64 | 1.99 | 0.67 | 2.36 | 0 | 3.03 | 0 | 0 | 0.67 | 1.01 | 1.67 | 0.4 | 0 | 1.63 | 2.02 | 1.35 | 7.42 | 0.3 | 1.21 | 0.03
| Analytical results for Sm in OREAS 506b (Certified Value 5.9 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 2.01: 6.11; 2.02: 6; 2.03: 5.816; 2.04: 5.7422; 2.06: 5.91; 2.07: 5.36; 2.08: 6.5; 2.09: 5.5 |  |  |  | 1
| 1 | 2 | 2.01: 6.33; 2.02: 6.8; 2.03: 5.909; 2.04: 5.7847; 2.06: 5.78; 2.07: 5.57; 2.08: 6.5; 2.09: 5.5 |  |  |  | 12
| 1 | 3 | 2.01: 6.16; 2.02: 7; 2.03: 5.8894; 2.04: 5.8065; 2.06: 5.3; 2.07: 5.71; 2.08: 6.3; 2.09: 5.9 |  |  |  | 16
| 1 | 4 | 2.01: 6.04; 2.02: 5.6; 2.03: 6.0913; 2.04: 5.7376; 2.06: 5.42; 2.07: 5.57; 2.08: 6.4; 2.09: 5.7 |  |  |  | 5.9024
| 1 | 5 | 2.01: 6.41; 2.02: 6.9; 2.03: 6.0429; 2.04: 5.8082; 2.06: 5.1; 2.07: 5.41; 2.08: 5.9; 2.09: 5.5 |  |  |  | 54
| 1 | 6 | 2.01: 6.09; 2.02: 5.8; 2.03: 6.3816; 2.04: 5.7552; 2.06: 5.48; 2.07: 5.4; 2.08: 5.7; 2.09: 5.9
| Mean |  | 2.01: 6.19; 2.02: 6.35; 2.03: 6.0217; 2.04: 5.7724; 2.06: 5.4983; 2.07: 5.5033; 2.08: 6.2167; 2.09: 5.6667
| Median |  | 2.01: 6.135; 2.02: 6.4; 2.03: 5.9759; 2.04: 5.7699; 2.06: 5.45; 2.07: 5.49; 2.08: 6.35; 2.09: 5.6
| Std Dev. |  | 2.01: 0.14683; 2.02: 0.61887; 2.03: 0.20363; 2.04: 0.031647; 2.06: 0.30109; 2.07: 0.1353; 2.08: 0.33714; 2.09: 0.19664
| Rel.Std.Dev. |  | 2.01: 2.372%; 2.02: 9.746%; 2.03: 3.382%; 2.04: 0.5483%; 2.06: 5.476%; 2.07: 2.459%; 2.08: 5.423%; 2.09: 3.47%
| PDM3 |  | 2.01: 4.873%; 2.02: 7.584%; 2.03: 2.021%; 2.04: -2.202%; 2.06: -6.846%; 2.07: -6.761%; 2.08: 5.325%; 2.09: -3.994%
| Z-Score (Absolute) |  | 2.01: 0.64; 2.02: 1; 2.03: 0.27; 2.04: 0.27; 2.06: 0.88; 2.07: 0.87; 2.08: 0.7; 2.09: 0.51
| Analytical results for Sn in OREAS 506b (Certified Value 4.27 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 4.4 | 4.3 | 4.5156 | 4.2 | 2.28 | 4.3 | 4.2 | 4.4 | 4.5 | 3.8 | 4.1 | 1 | 4.2 | 4.2734 | < 10 | 4 | < 20 | 4.3 | < 10 | 4.6 | 4.1 | 4 | 3.7 |  |  |  | 1
| 1 | 2 | 4.5 | 4.4 | 4.3347 | 4.4 | 2.32 | 4.3 | 4.1 | 4.5 | 4.5 | 4 | 4 | 5 | 4 | 4.231 | < 10 | 4.3 | < 20 | 4.4 | < 10 | 4.8 | 3.8 | 4 | 3.6 |  |  |  | 13
| 1 | 3 | 4.4 | 4.4 | 4.4961 | 4.4 | 2.26 | 4.2 | 4.1 | 4.4 | 4.4 | 4.2 | 4.1 | 7.4 | 4.1 | 4.2976 | < 10 | 4.1 | < 20 | 4.2 | < 10 | 4.9 | 3.8 | 4 | 3.6 |  |  |  | 16
| 1 | 4 | 4.4 | 4.4 | 4.4817 | 3.9 | 2.28 | 4.4 | 4.3 | 4.4 | 4.4 | 3.9 | 4.1 | 3.6 | 4.2 | 4.303 | < 10 | 4.3 | < 20 | 4.3 | < 10 | 4.9 | 3.9 | 4 | 3.5 |  |  |  | 4.2747
| 1 | 5 | 4.6 | 4.3 | 4.3116 | 4.7 | 2.29 | 4.4 | 4.1 | 4.2 | 4.5 | 4.1 | 4 | 2.4 | 4 | 4.2123 | < 10 | 4.2 | < 20 | 4.3 | < 10 | 4.9 | 4.1 | 5 | 3.7 |  |  |  | 55
| 1 | 6 | 4.2 | 4.4 | 4.3504 | 4.1 | 2.25 | 4.3 | 4.2 | 4.3 | 4.3 | 4.1 | 4 | 1.4 | 4.1 | 4.4659 | < 10 | 4.4 | < 20 | 4.2 | < 10 | 4.7 | 3.7 | 5 | 3.8
| Mean |  | 4.4167 | 4.3667 | 4.415 | 4.2833 | 2.28 | 4.3167 | 4.1667 | 4.3667 | 4.4333 | 4.0167 | 4.05 | 3.4667 | 4.1 | 4.2972 |  | 4.2167 |  | 4.2833 |  | 4.8 | 3.9 | 4.3333 | 3.65
| Median |  | 4.4 | 4.4 | 4.416 | 4.3 | 2.28 | 4.3 | 4.15 | 4.4 | 4.45 | 4.05 | 4.05 | 3 | 4.1 | 4.2855 |  | 4.25 |  | 4.3 |  | 4.85 | 3.85 | 4 | 3.65
| Std Dev. |  | 0.13292 | 0.05164 | 0.092158 | 0.27869 | 0.024495 | 0.075277 | 0.08165 | 0.10328 | 0.08165 | 0.1472 | 0.054772 | 2.4221 | 0.089443 | 0.090179 |  | 0.1472 |  | 0.075277 |  | 0.12649 | 0.16733 | 0.5164 | 0.10488
| Rel.Std.Dev. |  | 3.009% | 1.183% | 2.087% | 6.506% | 1.074% | 1.744% | 1.96% | 2.365% | 1.842% | 3.665% | 1.352% | 69.87% | 2.182% | 2.099% |  | 3.491% |  | 1.757% |  | 2.635% | 4.291% | 11.92% | 2.873%
| PDM3 |  | 3.321% | 2.152% | 3.282% | 0.2021% | -46.66% | 0.9819% | -2.527% | 2.152% | 3.711% | -6.036% | -5.256% | -18.9% | -4.087% | 0.5265% |  | -1.357% |  | 0.2021% |  | 12.29% | -8.765% | 1.372% | -14.61%
| Z-Score (Absolute) |  | 0.51 | 0.31 | 0.5 | 0.03 | 8.13 | 0.11 | 0.5 | 0.31 | 0.58 | 1.11 | 0.97 | 3.33 | 0.77 | 0.03 | 2.87 | 0.3 | 23.1 | 0.03 | 2.87 | 2.06 | 1.58 | NA | 2.59
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 506b (Certified Value 276 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 284.94 | 280 | 297.0059 | 265 | 299.12 | 277 | 264 | 273 | 272 | 267 | 268 | 219 | 284.3 | 268.3325 | 276 | 262.8 | 292.1344 | 274 | 268 | 224 | 280 | 270 | 274 | 269 | 265 |  | 1
| 1 | 2 | 293.57 | 284 | 292.5245 | 275 | 301.52 | 282 | 266 | 273 | 268 | 271 | 270 | 214 | 284.5 | 264.516 | 283 | 276 | 289.6418 | 268 | 266 | 261 | 291 | 279 | 271 | 285 | 259 |  | 14
| 1 | 3 | 288.97 | 286 | 294.8017 | 275 | 293.17 | 280 | 257 | 272 | 273 | 282 | 267 | 215 | 278.2 | 269.3808 | 282 | 274.9 | 290.1465 | 266 | 268 | 256 | 287 | 284 | 275 | 270 | 259 |  | 16
| 1 | 4 | 285.35 | 282 | 294.2139 | 245 | 307.91 | 279 | 266 | 268 | 276 | 272 | 269 | 215 | 288.2 | 265.2967 | 277 | 278.1 | 288.3543 | 269 | 277 | 247 | 285 | 290 | 278 | 274 | 256 |  | 276.2316
| 1 | 5 | 285.32 | 280 | 289.1593 | 281 | 294.4 | 283 | 263 | 269 | 272 | 270 | 269 | 210 | 277.9 | 258.7512 | 265 | 270.2 | 293.6382 | 271 | 250 | 233 | 280 | 277 | 272 | 302 | 258 |  | 56
| 1 | 6 | 292.14 | 287 | 287.8468 | 263 | 305.13 | 279 | 269 | 267 | 278 | 270 | 269 | 215 | 283.6 | 274.1973 | 287 | 281.4 | 293.8442 | 271 | 264 | 239 | 290 | 278 | 273 | 300 | 269
| Mean |  | 288.3817 | 283.1667 | 292.592 | 267.3333 | 300.2083 | 280 | 264.1667 | 270.3333 | 273.1667 | 272 | 268.6667 | 214.6667 | 282.7833 | 266.7458 | 278.3333 | 273.9 | 291.2932 | 269.8333 | 265.5 | 243.3333 | 285.5 | 279.6667 | 273.8333 | 283.3333 | 261
| Median |  | 287.16 | 283 | 293.3692 | 270 | 300.32 | 279.5 | 265 | 270.5 | 272.5 | 270.5 | 269 | 215 | 283.95 | 266.8146 | 279.5 | 275.45 | 291.1404 | 270 | 267 | 243 | 286 | 278.5 | 273.5 | 279.5 | 259
| Std Dev. |  | 3.7895 | 2.9944 | 3.5014 | 12.8634 | 5.8261 | 2.1909 | 4.0702 | 2.6583 | 3.4881 | 5.1769 | 1.0328 | 2.8752 | 4.0027 | 5.2148 | 7.6855 | 6.5757 | 2.2536 | 2.7869 | 8.8034 | 14.038 | 4.7645 | 6.7725 | 2.4833 | 14.8279 | 4.9396
| Rel.Std.Dev. |  | 1.314% | 1.057% | 1.197% | 4.812% | 1.941% | 0.7825% | 1.541% | 0.9833% | 1.277% | 1.903% | 0.3844% | 1.339% | 1.415% | 1.955% | 2.761% | 2.401% | 0.7737% | 1.033% | 3.316% | 5.769% | 1.669% | 2.422% | 0.9069% | 5.233% | 1.893%
| PDM3 |  | 4.398% | 2.511% | 5.923% | -3.221% | 8.68% | 1.364% | -4.368% | -2.135% | -1.11% | -1.532% | -2.739% | -22.29% | 2.372% | -3.434% | 0.7608% | -0.8441% | 5.453% | -2.316% | -3.885% | -11.91% | 3.355% | 1.244% | -0.8682% | 2.571% | -5.514%
| Z-Score (Absolute) |  | 1.18 | 0.76 | 1.52 | 0.53 | 2.13 | 0.5 | 0.78 | 0.28 | 0.05 | 0.15 | 0.42 | 4.79 | 0.72 | 0.57 | 0.36 | 0.01 | 1.41 | 0.32 | 0.67 | 2.47 | 0.94 | 0.47 | 0 | 0.77 | 1.04
| Analytical results for Ta in OREAS 506b (Certified Value 0.93 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.97 | 0.89 | 0.98377 | 1 | 0.93 | 0.9 | 0.93 | 0.93 | 0.96 | 0.84 | 0.76 | 0.64 | 1.44 | 0.85583 | < 10 | 0.96 | 1 | 0.64 | 0.84 | 0.8 | 0.8 | [5 empty] | 1
| 1 | 2 | 1 | 0.89 | 1.0019 | 1.1 | 0.99 | 0.93 | 0.92 | 0.91 | 0.91 | 0.88 | 0.9 | 0.7 | 1.3 | 0.92756 | < 10 | 0.97 | 1.02 | 0.65 | 0.84 | 0.7 | 0.8 | [5 empty] | 15
| 1 | 3 | 0.97 | 0.9 | 1.0182 | 1.1 | 1.02 | 0.93 | 0.93 | 0.9 | 0.9 | 0.87 | 0.82 | 0.59 | 1.41 | 0.90994 | < 10 | 0.96 | 1.06 | 0.7 | 0.87 | 0.8 | 0.8 | [5 empty] | 16
| 1 | 4 | 0.97 | 0.89 | 1.0242 | 0.9 | 0.99 | 0.9 | 0.78 | 0.91 | 0.94 | 0.84 | 0.8 | 0.59 | 1.43 | 0.94614 | < 10 | 0.97 | 0.97 | 0.7 | 0.85 | 0.8 | 0.8 | [5 empty] | 0.92553
| 1 | 5 | 0.99 | 0.88 | 0.9866 | 1.1 | 1.01 | 0.96 | 0.91 | 0.9 | 0.91 | 0.86 | 0.87 | 0.47 | 1.34 | 0.88999 | < 10 | 0.98 | 1.08 | 0.6 | 0.83 | 0.9 | 0.8 | [5 empty] | 57
| 1 | 6 | 0.94 | 0.89 | 1.0015 | 0.9 | 1.24 | 0.91 | 0.92 | 0.92 | 0.9 | 0.87 | 0.81 | 0.78 | 1.33 | 0.90915 | < 10 | 0.95 | 1.05 | 0.67 | 0.84 | 0.9 | 0.8
| Mean |  | 0.97333 | 0.89 | 1.0027 | 1.0167 | 1.03 | 0.92167 | 0.89833 | 0.91167 | 0.92 | 0.86 | 0.82667 | 0.62833 | 1.375 | 0.90644 |  | 0.965 | 1.03 | 0.66 | 0.845 | 0.81667 | 0.8
| Median |  | 0.97 | 0.89 | 1.0017 | 1.05 | 1 | 0.92 | 0.92 | 0.91 | 0.91 | 0.865 | 0.815 | 0.615 | 1.375 | 0.90955 |  | 0.965 | 1.035 | 0.66 | 0.84 | 0.8 | 0.8
| Std Dev. |  | 0.020656 | 0.0063246 | 0.016272 | 0.098319 | 0.10752 | 0.023166 | 0.058452 | 0.01169 | 0.024495 | 0.016733 | 0.050465 | 0.1061 | 0.058907 | 0.031222 |  | 0.010488 | 0.040988 | 0.038471 | 0.013784 | 0.075277 | 1.2162e-16
| Rel.Std.Dev. |  | 2.122% | 0.7106% | 1.623% | 9.671% | 10.44% | 2.513% | 6.507% | 1.282% | 2.662% | 1.946% | 6.105% | 16.89% | 4.284% | 3.444% |  | 1.087% | 3.979% | 5.829% | 1.631% | 9.218% | 1.52e-14%
| PDM3 |  | 5.165% | -3.839% | 8.337% | 9.847% | 11.29% | -0.4178% | -2.939% | -1.498% | -0.5979% | -7.081% | -10.68% | -32.11% | 48.56% | -2.063% |  | 4.264% | 11.29% | -28.69% | -8.701% | -11.76% | -13.56%
| Z-Score (Absolute) |  | 0.6 | 0.27 | 0.91 | NA | 1.2 | 0.06 | 0.18 | 0.04 | 0.04 | 0.59 | 0.94 | 3.02 | 4.83 | 0.1 | 42.92 | 0.52 | 1.2 | 2.69 | 0.74 | NA | NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 506b (Certified Value 0.67 ppm) | [27 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.24 | 2.25 | 2.27 | [13 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [13 empty] | 2
3
| 1 | 1 | 0.69 | 0.7 | 0.61 | 0.68 | 0.67 | 0.65596 | 0.6 | 0.68212 | 0.7 | 0.63 | 0.72 | 0.7 | 0.6 | [13 empty] | 1
| 1 | 2 | 0.71 | 0.9 | 0.7 | 0.68 | 0.69 | 0.6702 | 0.8 | 0.67024 | 0.68 | 0.65 | 0.71 | 0.8 | 0.6 | [13 empty] | 16
| 1 | 3 | 0.69 | 0.8 | 0.73 | 0.69 | 0.65 | 0.63534 | 0.8 | 0.74424 | 0.64 | 0.67 | 0.76 | 0.7 | 0.6 | [13 empty] | 16
| 1 | 4 | 0.72 | 0.7 | 0.67 | 0.65 | 0.68 | 0.66837 | 0.7 | 0.73966 | 0.65 | 0.61 | 0.72 | 0.8 | 0.6 | [13 empty] | 0.66747
| 1 | 5 | 0.7 | 0.8 | 0.65 | 0.67 | 0.64 | 0.65509 | 0.7 | 0.6917 | 0.67 | 0.66 | 0.69 | 0.7 | 0.6 | [13 empty] | 58
| 1 | 6 | 0.7 | 0.7 | 0.67 | 0.64 | 0.65 | 0.67848 | 0.8 | 0.67225 | 0.66 | 0.63 | 0.72 | 0.7 | 0.6
| Mean |  | 0.70167 | 0.76667 | 0.67167 | 0.66833 | 0.66333 | 0.66057 | 0.73333 | 0.70003 | 0.66667 | 0.64167 | 0.72 | 0.73333 | 0.6
| Median |  | 0.7 | 0.75 | 0.67 | 0.675 | 0.66 | 0.66216 | 0.75 | 0.68691 | 0.665 | 0.64 | 0.72 | 0.7 | 0.6
| Std Dev. |  | 0.01169 | 0.08165 | 0.041191 | 0.019408 | 0.019664 | 0.015248 | 0.08165 | 0.03339 | 0.021602 | 0.022286 | 0.022804 | 0.05164 | 0
| Rel.Std.Dev. |  | 1.666% | 10.65% | 6.133% | 2.904% | 2.964% | 2.308% | 11.13% | 4.77% | 3.24% | 3.473% | 3.167% | 7.042% | 0%
| PDM3 |  | 5.124% | 14.86% | 0.6291% | 0.1297% | -0.6194% | -1.033% | 9.868% | 4.879% | -0.12% | -3.865% | 7.87% | 9.868% | -10.11%
| Z-Score (Absolute) |  | 2.9 | NA | 0.29 | 0 | 0.43 | 0.67 | NA | 2.75 | 0.14 | 2.32 | 4.49 | NA | NA
Results from laboratories 2.04, 2.17, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 506b (Certified Value 0.21 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 |  | 0.2 | 0.2172 | 0.19 | < 0.5 | 0.22 | 0.22 | 0.25 | 0.21 | 0.13 | 0.23 | 0.7 | 0.26 | < 0.2 | < 10 | 0.17 | 0.56 | 0.62 | 0.2 | 0.2 | NA | 0.24 |  |  |  |  | 1
| 1 | 2 |  | 0.2 | 0.2208 | 0.19 | < 0.5 | 0.25 | 0.19 | 0.24 | 0.21 | 0.18 | 0.21 | 0.71 | 0.29 | < 0.2 | < 10 | 0.18 | 0.61 | 0.6 | 0.3 | 0.2 | NA | 0.22 |  |  |  |  | 55
| 1 | 3 |  | 0.2 | 0.19561 | 0.16 | < 0.5 | 0.2 | 0.21 | 0.23 | 0.19 | 0.28 | 0.19 | 0.67 | 0.26 | < 0.2 | < 10 | 0.17 | 0.5 | 0.66 | 0.23 | 0.2 | NA | 0.22 |  |  |  |  | 16
| 1 | 4 | 0.2 | 0.2 | 0.22388 | 0.18 | < 0.5 | 0.21 | 0.22 | 0.21 | 0.23 | 0.15 | 0.21 | 0.54 | 0.26 | < 0.2 | < 10 | 0.17 | 0.56 | 0.55 | 0.24 | 0.2 | NA | 0.21 |  |  |  |  | 0.21185
| 1 | 5 | 0.2 | 0.2 | 0.20498 | 0.22 | < 0.5 | 0.24 | 0.18 | 0.2 | 0.26 | 0.17 | 0.2 | 0.44 | 0.27 | < 0.2 | < 10 | 0.16 | 0.56 | 0.63 | 0.21 | 0.2 | NA | 0.23 |  |  |  |  | 59
| 1 | 6 | 0.2 | 0.2 | 0.19286 | 0.2 | < 0.5 | 0.21 | 0.19 | 0.23 | 0.23 | 0.21 | 0.21 | 0.47 | 0.27 | < 0.2 | < 10 | 0.13 | 0.51 | 0.61 | 0.21 | 0.2 | NA | 0.2
| Mean |  | 0.2 | 0.2 | 0.20922 | 0.19 |  | 0.22167 | 0.20167 | 0.22667 | 0.22167 | 0.18667 | 0.20833 | 0.58833 | 0.26833 |  |  | 0.16333 | 0.55 | 0.61167 | 0.23167 | 0.2 |  | 0.22
| Median |  | 0.2 | 0.2 | 0.21109 | 0.19 |  | 0.215 | 0.2 | 0.23 | 0.22 | 0.175 | 0.21 | 0.605 | 0.265 |  |  | 0.17 | 0.56 | 0.615 | 0.22 | 0.2 |  | 0.22
| Std Dev. |  | 3.3993e-17 | 3.0405e-17 | 0.013293 | 0.02 |  | 0.019408 | 0.017224 | 0.018619 | 0.024014 | 0.053166 | 0.013292 | 0.12024 | 0.01169 |  |  | 0.017512 | 0.04 | 0.03656 | 0.03656 | 3.0405e-17 |  | 0.014142
| Rel.Std.Dev. |  | 1.7e-14% | 1.52e-14% | 6.354% | 10.53% |  | 8.755% | 8.541% | 8.214% | 10.83% | 28.48% | 6.38% | 20.44% | 4.357% |  |  | 10.72% | 7.273% | 5.977% | 15.78% | 1.52e-14% |  | 6.428%
| PDM3 |  | -5.594% | -5.594% | -1.241% | -10.31% |  | 4.633% | -4.808% | 6.993% | 4.633% | -11.89% | -1.661% | 177.7% | 26.66% |  |  | -22.9% | 159.6% | 188.7% | 9.353% | -5.594% |  | 3.846%
| Z-Score (Absolute) |  | NA | NA | 0.28 | 0.71 | 0.64 | 0 | 0.45 | 0.11 | 0 | 0.79 | 0.3 | 8.24 | 1.05 | 2.73 | 107.4 | 1.31 | 7.38 | 8.77 | 0.22 | NA | NA | 0.04
Results from laboratories 2.01, 2.02 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 506b (Certified Value 12 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 12.51 | 12.3 | 13.0962 | 11.2 | 9.74 | 12.25 | 12 | 12.6 | 13.15 | 11.95 | 11.71 | 11.7 | 8.2 | 11.9761 | 12.05 | 10.325 | 11 | 11 | 10.36 | 15.1 | 11.8 | 11.9 | 11.5 |  |  |  | 1
| 1 | 2 | 12.81 | 12.3 | 13.2189 | 12.4 | 10.32 | 11.85 | 12.65 | 13.5 | 12.6 | 12.5 | 12.44 | 11.7 | 8.6 | 12.2339 | 13.09 | 10.1924 | 12 | 11.4 | 10.65 | 15.48 | 12.1 | 13.2 | 10.9 |  |  |  | 17
| 1 | 3 | 12.69 | 12.3 | 12.8518 | 12.7 | 11.41 | 11.35 | 12.35 | 13.65 | 12.1 | 13.1 | 12.28 | 11.5 | 8 | 12.0556 | 13.07 | 10.326 | 10 | 9.74 | 10.95 | 15.7 | 12.3 | 12.5 | 11.9 |  |  |  | 16
| 1 | 4 | 12.81 | 12.8 | 13.2758 | 10.9 | 10 | 11.35 | 12.85 | 13.95 | 12.25 | 12.7 | 12.21 | 11.2 | 8 | 12.3266 | 12.96 | 10.4231 | 10 | 9.7 | 11.24 | 15.52 | 11.9 | 13.2 | 11.7 |  |  |  | 11.9562
| 1 | 5 | 13.13 | 12.4 | 12.8494 | 12.5 | 8.28 | 11.75 | 11.95 | 13.4 | 12.45 | 12.15 | 12.57 | 10.7 | 8.1 | 12.2145 | 12.65 | 10.3381 | 12 | 10 | 10.66 | 15.16 | 11.1 | 13.1 | 11.5 |  |  |  | 60
| 1 | 6 | 12.6 | 12.3 | 12.7726 | 11.6 | 10.48 | 11.35 | 12.65 | 13.3 | 11.8 | 12.15 | 12.08 | 12.1 | 8.1 | 12.7655 | 13.08 | 10.1914 | 10 | 10.8 | 10.14 | 15.2 | 12.2 | 12.7 | 12
| Mean |  | 12.7583 | 12.4 | 13.0108 | 11.8833 | 10.0383 | 11.65 | 12.4083 | 13.4 | 12.3917 | 12.425 | 12.215 | 11.4833 | 8.1667 | 12.262 | 12.8167 | 10.2993 | 10.8333 | 10.44 | 10.6667 | 15.36 | 11.9 | 12.7667 | 11.5833
| Median |  | 12.75 | 12.3 | 12.974 | 12 | 10.16 | 11.55 | 12.5 | 13.45 | 12.35 | 12.325 | 12.245 | 11.6 | 8.1 | 12.2242 | 13.015 | 10.3255 | 10.5 | 10.4 | 10.655 | 15.34 | 12 | 12.9 | 11.6
| Std Dev. |  | 0.21674 | 0.2 | 0.21396 | 0.75211 | 1.0333 | 0.36878 | 0.37204 | 0.45277 | 0.46413 | 0.42749 | 0.30138 | 0.48339 | 0.22509 | 0.27753 | 0.41093 | 0.090886 | 0.98319 | 0.72056 | 0.39515 | 0.24033 | 0.43359 | 0.51251 | 0.392
| Rel.Std.Dev. |  | 1.699% | 1.613% | 1.644% | 6.329% | 10.29% | 3.166% | 2.998% | 3.379% | 3.746% | 3.441% | 2.467% | 4.209% | 2.756% | 2.263% | 3.206% | 0.8824% | 9.076% | 6.902% | 3.705% | 1.565% | 3.644% | 4.014% | 3.384%
| PDM3 |  | 6.709% | 3.712% | 8.821% | -0.6091% | -16.04% | -2.561% | 3.782% | 12.08% | 3.643% | 3.921% | 2.165% | -3.955% | -31.69% | 2.558% | 7.197% | -13.86% | -9.391% | -12.68% | -10.79% | 28.47% | -0.4697% | 6.779% | -3.118%
| Z-Score (Absolute) |  | 0.6 | 0.19 | 0.89 | 0.41 | 2.54 | 0.68 | 0.2 | 1.34 | 0.18 | 0.22 | 0.03 | 0.87 | 4.71 | 0.03 | 0.67 | 2.24 | NA | 2.08 | 1.82 | 3.61 | 0.39 | 0.61 | 0.76
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 506b (Certified Value 0.319 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.313 | 0.314 | 0.33028 | 0.319 | 0.54 | 0.313 | 0.325 | 0.325 | 0.346 | 0.322 | 0.311 | 0.358 | 0.3 | 0.31411 | 0.306 | 0.32 | 0.32985 | 0.33 | 0.326 | 0.231 | 0.32 | 0.3312 | 0.31 | 0.315 | 0.323 |  | 1
| 1 | 2 | 0.3154 | 0.311 | 0.32473 | 0.338 | 0.53 | 0.324 | 0.316 | 0.324 | 0.34 | 0.324 | 0.313 | 0.361 | 0.3 | 0.3177 | 0.314 | 0.32 | 0.3234 | 0.32 | 0.329 | 0.278 | 0.32 | 0.3344 | 0.3 | 0.309 | 0.302 |  | 18
| 1 | 3 | 0.3097 | 0.311 | 0.32017 | 0.334 | 0.54 | 0.317 | 0.319 | 0.322 | 0.344 | 0.341 | 0.306 | 0.355 | 0.3 | 0.32156 | 0.311 | 0.32 | 0.32949 | 0.32 | 0.329 | 0.267 | 0.31 | 0.3336 | 0.31 | 0.314 | 0.308 |  | 16
| 1 | 4 | 0.3094 | 0.313 | 0.3309 | 0.323 | 0.53 | 0.313 | 0.314 | 0.324 | 0.354 | 0.327 | 0.311 | 0.349 | 0.3 | 0.31287 | 0.314 | 0.32 | 0.32902 | 0.31 | 0.339 | 0.259 | 0.31 | 0.3383 | 0.31 | 0.311 | 0.303 |  | 0.31876
| 1 | 5 | 0.318 | 0.313 | 0.32185 | 0.326 | 0.53 | 0.321 | 0.324 | 0.32 | 0.344 | 0.324 | 0.304 | 0.338 | 0.29 | 0.30381 | 0.295 | 0.31 | 0.32488 | 0.32 | 0.312 | 0.24 | 0.31 | 0.3305 | 0.31 | 0.307 | 0.313 |  | 61
| 1 | 6 | 0.3017 | 0.313 | 0.32937 | 0.324 | 0.53 | 0.311 | 0.323 | 0.314 | 0.339 | 0.327 | 0.312 | 0.36 | 0.31 | 0.32044 | 0.315 | 0.33 | 0.32725 | 0.32 | 0.322 | 0.22 | 0.32 | 0.3396 | 0.31 | 0.3 | 0.326
| Mean |  | 0.3112 | 0.3125 | 0.32622 | 0.32733 | 0.53333 | 0.3165 | 0.32017 | 0.3215 | 0.3445 | 0.3275 | 0.3095 | 0.3535 | 0.3 | 0.31508 | 0.30917 | 0.32 | 0.32732 | 0.32 | 0.32617 | 0.24917 | 0.315 | 0.3346 | 0.30833 | 0.30933 | 0.3125
| Median |  | 0.31135 | 0.313 | 0.32705 | 0.325 | 0.53 | 0.315 | 0.321 | 0.323 | 0.344 | 0.3255 | 0.311 | 0.3565 | 0.3 | 0.31591 | 0.3125 | 0.32 | 0.32813 | 0.32 | 0.3275 | 0.2495 | 0.315 | 0.334 | 0.31 | 0.31 | 0.3105
| Std Dev. |  | 0.0057072 | 0.0012247 | 0.0046107 | 0.0072019 | 0.005164 | 0.0051284 | 0.004535 | 0.0040866 | 0.0053572 | 0.006892 | 0.0036194 | 0.008735 | 0.0063246 | 0.0064846 | 0.007679 | 0.0063246 | 0.0026562 | 0.0063246 | 0.0089312 | 0.022409 | 0.0054772 | 0.0036905 | 0.0040825 | 0.005465 | 0.010134
| Rel.Std.Dev. |  | 1.834% | 0.3919% | 1.413% | 2.2% | 0.9682% | 1.62% | 1.416% | 1.271% | 1.555% | 2.104% | 1.169% | 2.471% | 2.108% | 2.058% | 2.484% | 1.976% | 0.8115% | 1.976% | 2.738% | 8.994% | 1.739% | 1.103% | 1.324% | 1.767% | 3.243%
| PDM3 |  | -2.371% | -1.963% | 2.34% | 2.691% | 67.32% | -0.708% | 0.4423% | 0.8606% | 8.076% | 2.743% | -2.904% | 10.9% | -5.884% | -1.153% | -3.009% | 0.39% | 2.686% | 0.39% | 2.325% | -21.83% | -1.179% | 4.97% | -3.27% | -2.956% | -1.963%
| Z-Score (Absolute) |  | 0.79 | 0.67 | 0.56 | 0.66 | 19.18 | 0.31 | 0.01 | 0.13 | 2.2 | 0.67 | 0.94 | 3.01 | 1.8 | 0.44 | 0.97 | 0 | 0.66 | 0 | 0.55 | 6.37 | 0.45 | 1.31 | 1.05 | 0.96 | 0.67
| Analytical results for Tl in OREAS 506b (Certified Value 0.81 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.88 | 0.85 | 0.83977 | 0.79 | 0.68 | 0.83 | 0.77 | 0.8 | 0.84 | 0.77 | 0.82 | 0.77 | 0.86 | 0.78854 | 0.76 | < 5 | 0.77 | 0.76 | 0.8 | 0.79 | 0.77 | 0.705 | 0.81 |  |  |  | 1
| 1 | 2 | 0.89 | 0.84 | 0.83564 | 0.84 | 0.78 | 0.81 | 0.78 | 0.81 | 0.83 | 0.79 | 0.83 | 0.81 | 0.87 | 0.82227 | 0.85 | < 5 | 0.8 | 0.77 | 0.8 | 0.77 | 0.79 | 0.753 | 0.79 |  |  |  | 19
| 1 | 3 | 0.87 | 0.84 | 0.86816 | 0.8 | 0.78 | 0.81 | 0.78 | 0.8 | 0.86 | 0.84 | 0.81 | 0.77 | 0.82 | 0.79666 | 0.82 | < 5 | 0.78 | 0.79 | 0.9 | 0.8 | 0.78 | 0.815 | 0.85 |  |  |  | 16
| 1 | 4 | 0.89 | 0.84 | 0.83229 | 0.78 | 0.81 | 0.81 | 0.77 | 0.8 | 0.84 | 0.79 | 0.82 | 0.78 | 0.88 | 0.84221 | 0.9 | < 5 | 0.78 | 0.81 | 0.8 | 0.76 | 0.81 | 0.755 | 0.79 |  |  |  | 0.81087
| 1 | 5 | 0.91 | 0.84 | 0.868 | 0.86 | 0.79 | 0.85 | 0.76 | 0.8 | 0.82 | 0.85 | 0.81 | 0.74 | 0.82 | 0.78779 | 0.85 | < 5 | 0.8 | 0.78 | 0.9 | 0.77 | 0.81 | 0.815 | 0.79 |  |  |  | 62
| 1 | 6 | 0.84 | 0.84 | 0.83903 | 0.81 | 0.79 | 0.77 | 0.76 | 0.78 | 0.84 | 0.79 | 0.83 | 0.82 | 0.87 | 0.80905 | 0.87 | < 5 | 0.78 | 0.81 | 0.8 | 0.76 | 0.85 | 0.767 | 0.83
| Mean |  | 0.88 | 0.84167 | 0.84715 | 0.81333 | 0.77167 | 0.81333 | 0.77 | 0.79833 | 0.83833 | 0.805 | 0.82 | 0.78167 | 0.85333 | 0.80775 | 0.84167 |  | 0.785 | 0.78667 | 0.83333 | 0.775 | 0.80167 | 0.76833 | 0.81
| Median |  | 0.885 | 0.84 | 0.8394 | 0.805 | 0.785 | 0.81 | 0.77 | 0.8 | 0.84 | 0.79 | 0.82 | 0.775 | 0.865 | 0.80285 | 0.85 |  | 0.78 | 0.785 | 0.8 | 0.77 | 0.8 | 0.761 | 0.8
| Std Dev. |  | 0.023664 | 0.0040825 | 0.016431 | 0.030768 | 0.046224 | 0.026583 | 0.0089443 | 0.0098319 | 0.013292 | 0.032094 | 0.0089443 | 0.029269 | 0.026583 | 0.021402 | 0.047924 |  | 0.012247 | 0.020656 | 0.05164 | 0.016432 | 0.028577 | 0.041908 | 0.025298
| Rel.Std.Dev. |  | 2.689% | 0.485% | 1.94% | 3.783% | 5.99% | 3.268% | 1.162% | 1.232% | 1.585% | 3.987% | 1.091% | 3.744% | 3.115% | 2.65% | 5.694% |  | 1.56% | 2.626% | 6.197% | 2.12% | 3.565% | 5.454% | 3.123%
| PDM3 |  | 8.526% | 3.798% | 4.474% | 0.304% | -4.835% | 0.304% | -5.04% | -1.546% | 3.387% | -0.7237% | 1.126% | -3.601% | 5.237% | -0.3843% | 3.798% |  | -3.19% | -2.985% | 2.77% | -4.423% | -1.135% | -5.246% | -0.1071%
| Z-Score (Absolute) |  | 1.69 | 0.78 | 0.91 | 0.11 | 0.89 | 0.11 | 0.93 | 0.25 | 0.7 | 0.09 | 0.26 | 0.65 | 1.06 | 0.03 | 0.78 | 40.25 | 0.57 | 0.53 | NA | 0.81 | 0.17 | 0.96 | 0.03
Results from laboratory 2.23 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Tm in OREAS 506b (Certified Value 0.18 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.15; 2.04: 2.18; 2.06: 2.20; 2.07: 2.21; 2.08: 2.25; 2.09: 2.27 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 2.01: 0.19; 2.02: 0.2; 2.03: 0.17134; 2.04: 0.25444; 2.06: 0.19; 2.07: 0.15; 2.08: 0.2; 2.09: 0.2 |  |  |  | 1
| 1 | 2 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.17789; 2.04: 0.18392; 2.06: 0.2; 2.07: 0.16; 2.08: 0.2; 2.09: 0.2 |  |  |  | 20
| 1 | 3 | 2.01: 0.19; 2.02: 0.2; 2.03: 0.17105; 2.04: 0.1672; 2.06: 0.18; 2.07: 0.17; 2.08: 0.2; 2.09: 0.2 |  |  |  | 16
| 1 | 4 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.18265; 2.04: 0.1583; 2.06: 0.19; 2.07: 0.17; 2.08: 0.2; 2.09: 0.2 |  |  |  | 0.17588
| 1 | 5 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.1818; 2.04: 0.16048; 2.06: 0.17; 2.07: 0.15; 2.08: 0.2; 2.09: 0.2 |  |  |  | 63
| 1 | 6 | 2.01: 0.16; 2.02: 0.2; 2.03: 0.17742; 2.04: 0.16695; 2.06: 0.18; 2.07: 0.16; 2.08: 0.2; 2.09: 0.2
| Mean |  | 2.01: 0.19; 2.02: 0.2; 2.03: 0.17703; 2.04: 0.18188; 2.06: 0.185; 2.07: 0.16; 2.08: 0.2; 2.09: 0.2
| Median |  | 2.01: 0.195; 2.02: 0.2; 2.03: 0.17765; 2.04: 0.16708; 2.06: 0.185; 2.07: 0.16; 2.08: 0.2; 2.09: 0.2
| Std Dev. |  | 2.01: 0.015492; 2.02: 3.0405e-17; 2.03: 0.0049681; 2.04: 0.036663; 2.06: 0.010488; 2.07: 0.0089443; 2.08: 3.0405e-17; 2.09: 3.0405e-17
| Rel.Std.Dev. |  | 2.01: 8.154%; 2.02: 1.52e-14%; 2.03: 2.806%; 2.04: 20.16%; 2.06: 5.669%; 2.07: 5.59%; 2.08: 1.52e-14%; 2.09: 1.52e-14%
| PDM3 |  | 2.01: 8.029%; 2.02: 13.71%; 2.03: 0.6522%; 2.04: 3.412%; 2.06: 5.186%; 2.07: -9.029%; 2.08: 13.71%; 2.09: 13.71%
| Z-Score (Absolute) |  | 2.01: 1.13; 2.02: NA; 2.03: 0.67; 2.04: 0; 2.06: 0.43; 2.07: 3.04; 2.08: NA; 2.09: NA
Results from laboratories 2.04, 2.25 and 2.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 506b (Certified Value 3.5 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.93 | 3.62 | 3.6581 | 3.1 | 3.46 | 3.4 | 3.4 | 3.8 | 4.2 | 3.1 | 3.19 | 3.41 | 3.6 | 3.2962 | 3.86 | 3.1269 | < 10 | 3.3 | 4.16 | 3.7 | 3.71 | 3.7 | 3.07 | 3.7 |  |  | 1
| 1 | 2 | 4.03 | 3.78 | 3.5509 | 3.4 | 3.27 | 3.3 | 2.6 | 3 | 3.1 | 3.6 | 3.53 | 3.21 | 3.5 | 3.4282 | 3.98 | 3.1776 | < 10 | 3.57 | 3.87 | 3.5 | 3.24 | 4.3 | 3.28 | 3.4 |  |  | 21
| 1 | 3 | 3.72 | 3.39 | 3.7653 | 3.2 | 3.38 | 3.7 | 2.9 | 3.2 | 3.1 | 3.6 | 3.18 | 3.96 | 3.6 | 3.6611 | 4.01 | 3.1122 | < 10 | 3.59 | 4.18 | 3.7 | 4.08 | 3.1 | 3.69 | 3.4 |  |  | 16
| 1 | 4 | 3.5 | 3.43 | 3.7617 | 3 | 3.42 | 3.7 | 3.6 | 3.3 | 3.7 | 3.3 | 3.51 | 3.57 | 3.4 | 3.4802 | 4.11 | 2.9619 | < 10 | 3.95 | 3.91 | 3.8 | 3.75 | 4 | 3.17 | 3 |  |  | 3.5015
| 1 | 5 | 3.97 | 3.72 | 3.4805 | 3.4 | 3.36 | 3.4 | 2.8 | 3.7 | 3.4 | 3.8 | 3.54 | 2.96 | 3.6 | 3.5281 | 4.44 | 3.1964 | < 10 | 3.79 | 3.95 | 3.8 | 3.26 | 3.5 | 3.54 | 2.9 |  |  | 64
| 1 | 6 | 3.26 | 3.23 | 3.8316 | 2.9 | 3.42 | 3.7 | 2.7 | 3.3 | 3.1 | 3.4 | 3.76 | 3.26 | 3.7 | 3.6669 | 4.08 | 3.1326 | < 10 | 3.79 | 4.05 | 3.5 | 3.15 | 3.7 | 3.34 | 2.8
| Mean |  | 3.735 | 3.5283 | 3.6747 | 3.1667 | 3.385 | 3.5333 | 3 | 3.3833 | 3.4333 | 3.4667 | 3.4517 | 3.395 | 3.5667 | 3.5101 | 4.08 | 3.1179 |  | 3.665 | 4.02 | 3.6667 | 3.5317 | 3.7167 | 3.3483 | 3.2
| Median |  | 3.825 | 3.525 | 3.7099 | 3.15 | 3.4 | 3.55 | 2.85 | 3.3 | 3.25 | 3.5 | 3.52 | 3.335 | 3.6 | 3.5041 | 4.045 | 3.1298 |  | 3.69 | 4 | 3.7 | 3.485 | 3.7 | 3.31 | 3.2
| Std Dev. |  | 0.30376 | 0.21274 | 0.1369 | 0.20656 | 0.066257 | 0.18619 | 0.40497 | 0.30605 | 0.44572 | 0.25033 | 0.22569 | 0.34402 | 0.10328 | 0.14217 | 0.19688 | 0.082958 |  | 0.22819 | 0.13084 | 0.13663 | 0.37005 | 0.41191 | 0.23129 | 0.35214
| Rel.Std.Dev. |  | 8.133% | 6.029% | 3.726% | 6.523% | 1.957% | 5.27% | 13.5% | 9.046% | 12.98% | 7.221% | 6.539% | 10.13% | 2.896% | 4.05% | 4.825% | 2.661% |  | 6.226% | 3.255% | 3.726% | 10.48% | 11.08% | 6.908% | 11%
| PDM3 |  | 6.668% | 0.7654% | 4.945% | -9.563% | -3.328% | 0.9082% | -14.32% | -3.376% | -1.948% | -0.9958% | -1.424% | -3.042% | 1.86% | 0.245% | 16.52% | -10.95% |  | 4.668% | 14.81% | 4.716% | 0.8606% | 6.144% | -4.375% | -8.611%
| Z-Score (Absolute) |  | 0.99 | 0.04 | 0.71 | 1.62 | 0.62 | 0.06 | 2.39 | 0.63 | 0.4 | 0.24 | 0.31 | 0.57 | 0.22 | 0.04 | 2.58 | 1.85 | 6.81 | 0.67 | 2.3 | 0.68 | 0.06 | 0.91 | 0.79 | 1.47
| Analytical results for V in OREAS 506b (Certified Value 65 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 62 | 66 | 66.3383 | 65 | 58.42 | 66 | 66 | 65 | 68 | 66 | 66 | 72 | 63 | 65.7867 | 63 | 62.5 | 71 | 65 | 66 | 66 | 61 | 66 | 61 | 62 |  |  | 1
| 1 | 2 | 64 | 66 | 66.0553 | 69 | 62.58 | 68 | 65 | 64 | 67 | 67 | 66 | 72 | 65 | 65.7567 | 64 | 67.4 | 69 | 67 | 67 | 65 | 60 | 65 | 63 | 60 |  |  | 22
| 1 | 3 | 64 | 66 | 67.6332 | 69 | 61.4 | 67 | 65 | 64 | 68 | 69 | 66 | 71 | 64 | 66.9303 | 64 | 68 | 67 | 63 | 65 | 67 | 61 | 65 | 63 | 63 |  |  | 16
| 1 | 4 | 63 | 66 | 66.468 | 65 | 60.26 | 67 | 65 | 63 | 70 | 67 | 67 | 70 | 64 | 65.8131 | 63 | 67.1 | 68 | 66 | 65 | 70 | 60 | 64 | 63 | 61 |  |  | 65.4856
| 1 | 5 | 65 | 65 | 64.1419 | 66 | 61.84 | 68 | 66 | 64 | 68 | 66 | 65 | 68 | 63 | 63.889 | 61 | 65.2 | 69 | 64 | 66 | 71 | 59 | 68 | 63 | 61 |  |  | 65
| 1 | 6 | 62 | 64 | 64.6604 | 66 | 59.3 | 67 | 63 | 63 | 67 | 67 | 66 | 72 | 64 | 66.8379 | 65 | 69.7 | 68 | 63 | 67 | 67 | 59 | 68 | 62 | 60
| Mean |  | 63.3333 | 65.5 | 65.8828 | 66.6667 | 60.6333 | 67.1667 | 65 | 63.8333 | 68 | 67 | 66 | 70.8333 | 63.8333 | 65.8356 | 63.3333 | 66.65 | 68.6667 | 64.6667 | 66 | 67.6667 | 60 | 66 | 62.5 | 61.1667
| Median |  | 63.5 | 66 | 66.1968 | 66 | 60.83 | 67 | 65 | 64 | 68 | 67 | 66 | 71.5 | 64 | 65.7999 | 63.5 | 67.25 | 68.5 | 64.5 | 66 | 67 | 60 | 65.5 | 63 | 61
| Std Dev. |  | 1.2111 | 0.83666 | 1.2782 | 1.8619 | 1.5906 | 0.75277 | 1.0954 | 0.75277 | 1.0954 | 1.0954 | 0.63246 | 1.6021 | 0.75277 | 1.0955 | 1.3663 | 2.4986 | 1.3663 | 1.633 | 0.89443 | 2.3381 | 0.89443 | 1.6733 | 0.83666 | 1.169
| Rel.Std.Dev. |  | 1.912% | 1.277% | 1.94% | 2.793% | 2.623% | 1.121% | 1.685% | 1.179% | 1.611% | 1.635% | 0.9583% | 2.262% | 1.179% | 1.664% | 2.157% | 3.749% | 1.99% | 2.525% | 1.355% | 3.455% | 1.491% | 2.535% | 1.339% | 1.911%
| PDM3 |  | -3.287% | 0.02201% | 0.6066% | 1.804% | -7.41% | 2.567% | -0.7415% | -2.523% | 3.84% | 2.313% | 0.7855% | 8.166% | -2.523% | 0.5345% | -3.287% | 1.778% | 4.858% | -1.251% | 0.7855% | 3.331% | -8.377% | 0.7855% | -4.559% | -6.595%
| Z-Score (Absolute) |  | 1.09 | 0.16 | 0.01 | 0.35 | 2.26 | 0.57 | 0.37 | 0.88 | 0.93 | 0.49 | 0.06 | 2.15 | 0.88 | 0.01 | 1.09 | 0.34 | 1.22 | 0.52 | 0.06 | 0.78 | 2.54 | 0.06 | 1.45 | 2.03
| Analytical results for W in OREAS 506b (Certified Value 10.3 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 11.6 | 9.7 | 9.7521 | 9.3 | 9.62 | 10.6 | 10.4 | 11 | 11.6 | 9.1 | 10.1 | 8.93 | 9.6 | 10.2025 | 11 | 10.1 | < 20 | 11.5 | 21.9 | 9.8 | 11 | 10.9 | 10.2 |  |  |  | 1
| 1 | 2 | 10.3 | 10.1 | 10.6531 | 9.8 | 9.23 | 11.3 | 10.5 | 12.4 | 12 | 10.3 | 10.8 | 9.73 | 9.8 | 9.3638 | 11 | 10.2 | < 20 | 13 | 21 | 9.8 | 12 | 10.8 | 10 |  |  |  | 23
| 1 | 3 | 10.5 | 10 | 10.3508 | 9.7 | 9.37 | 12.6 | 10 | 10.2 | 10.7 | 12.6 | 10.3 | 10.4 | 9.2 | 9.9509 | 11 | 11 | < 20 | 12.8 | 20.9 | 10 | 12 | 11.2 | 11.9 |  |  |  | 16
| 1 | 4 | 12.4 | 9.8 | 10.147 | 11.9 | 11.42 | 10.9 | 10.4 | 10.3 | 12.2 | 10.4 | 11 | 9.64 | 9.8 | 9.7108 | 13 | 10.5 | < 20 | 10.6 | 17.9 | 10.5 | 11 | 10.3 | 9.2 |  |  |  | 10.3101
| 1 | 5 | 10.4 | 10.1 | 9.5977 | 11.2 | 9.65 | 12.5 | 9.7 | 10.5 | 12.4 | 9.6 | 11.1 | 9.11 | 9.3 | 10.2894 | 12 | 10.4 | < 20 | 10.5 | 22.3 | 10.4 | 11 | 11 | 10.3 |  |  |  | 66
| 1 | 6 | 9.9 | 9.8 | 10.8034 | 9.5 | 10.36 | 11.1 | 9.8 | 10 | 10.8 | 9.4 | 10.4 | 10.4 | 10 | 10.0845 | 11 | 10.6 | < 20 | 12.5 | 21.3 | 10.3 | 12 | 12.5 | 10.3
| Mean |  | 10.85 | 9.9167 | 10.2173 | 10.2333 | 9.9417 | 11.5 | 10.1333 | 10.7333 | 11.6167 | 10.2333 | 10.6167 | 9.7017 | 9.6167 | 9.9337 | 11.5 | 10.4667 |  | 11.8167 | 20.8833 | 10.1333 | 11.5 | 11.1167 | 10.3167
| Median |  | 10.45 | 9.9 | 10.2489 | 9.75 | 9.635 | 11.2 | 10.2 | 10.4 | 11.8 | 9.95 | 10.6 | 9.685 | 9.7 | 10.0177 | 11 | 10.45 |  | 12 | 21.15 | 10.15 | 11.5 | 10.95 | 10.25
| Std Dev. |  | 0.94816 | 0.17224 | 0.48095 | 1.0577 | 0.82247 | 0.84617 | 0.34448 | 0.88468 | 0.72226 | 1.266 | 0.40702 | 0.62056 | 0.31252 | 0.34535 | 0.83666 | 0.32042 |  | 1.1089 | 1.5575 | 0.30768 | 0.54772 | 0.7414 | 0.87958
| Rel.Std.Dev. |  | 8.739% | 1.737% | 4.707% | 10.34% | 8.273% | 7.358% | 3.399% | 8.242% | 6.217% | 12.37% | 3.834% | 6.396% | 3.25% | 3.477% | 7.275% | 3.061% |  | 9.384% | 7.458% | 3.036% | 4.763% | 6.669% | 8.526%
| PDM3 |  | 5.236% | -3.816% | -0.8998% | -0.7447% | -3.574% | 11.54% | -1.715% | 4.105% | 12.67% | -0.7447% | 2.973% | -5.901% | -6.726% | -3.651% | 11.54% | 1.518% |  | 14.61% | 102.6% | -1.715% | 11.54% | 7.823% | 0.06358%
| Z-Score (Absolute) |  | 1.31 | 0.67 | 0.03 | 0 | 0.62 | 2.7 | 0.21 | 1.06 | 2.95 | 0 | 0.82 | 1.13 | 1.31 | 0.64 | NA | 0.5 | 0.5 | 3.37 | 22.68 | 0.21 | NA | 1.88 | 0.18
Results from laboratories 2.16 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 506b (Certified Value 14.3 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 14.67 | 14.8 | 16.6705 | 12.6 | 11.12 | 14.6 | 14.2 | 14.8 | 14.4 | 12.8 | 13.8 | 16 | 13.8 | 14.2431 | 14 | 13.8 | 14.4847 | 15 | 15.1 | 13.3 | 13.3 | 14.3 | 14.5 | 16.1 | 13.4 |  | 1
| 1 | 2 | 15.12 | 14.7 | 15.3541 | 13.6 | 14.46 | 14.4 | 14.2 | 15.7 | 14.2 | 13.4 | 13.9 | 15.7 | 13.6 | 14.146 | 14.4 | 14.4 | 14.4425 | 14 | 14.8 | 13.5 | 13.6 | 14.3 | 14.1 | 16.7 | 13.1 |  | 24
| 1 | 3 | 14.76 | 14.8 | 15.4172 | 14 | 15.69 | 14.3 | 14 | 17.4 | 14 | 14 | 13.7 | 15.9 | 13.6 | 14.3728 | 14.2 | 14.5 | 14.456 | 14 | 14.1 | 13.7 | 13.4 | 14.7 | 14.8 | 16.2 | 13.5 |  | 16
| 1 | 4 | 14.71 | 14.7 | 16.4019 | 13.2 | 13.39 | 14.4 | 14.9 | 15.5 | 14.4 | 13.4 | 13.7 | 15.5 | 14.2 | 14.1647 | 13.8 | 14.8 | 14.4606 | 14 | 14.2 | 13.6 | 13.6 | 15.1 | 14.2 | 16.6 | 13.4 |  | 14.2861
| 1 | 5 | 14.61 | 14.8 | 14.5784 | 13.3 | 14.14 | 14.7 | 13.7 | 14.9 | 14.4 | 13.9 | 13.8 | 15 | 13.3 | 13.9633 | 13 | 14.4 | 14.493 | 14 | 14.3 | 13.1 | 13.1 | 14.5 | 13.9 | 17.2 | 13.7 |  | 67
| 1 | 6 | 14.76 | 14.8 | 14.6694 | 14.3 | 13.4 | 13.9 | 14.6 | 38.3 | 14 | 13.2 | 13.6 | 16.1 | 14.2 | 14.7428 | 14.1 | 15.8 | 14.4596 | 14 | 14 | 13.6 | 13.5 | 14.9 | 14.5 | 16.9 | 14.1
| Mean |  | 14.7717 | 14.7667 | 15.5153 | 13.5 | 13.7 | 14.3833 | 14.2667 | 19.4333 | 14.2333 | 13.45 | 13.75 | 15.7 | 13.7833 | 14.2721 | 13.9167 | 14.6167 | 14.4661 | 14.1667 | 14.4167 | 13.4667 | 13.4167 | 14.6333 | 14.3333 | 16.6167 | 13.5333
| Median |  | 14.735 | 14.8 | 15.3857 | 13.45 | 13.77 | 14.4 | 14.2 | 15.6 | 14.3 | 13.4 | 13.75 | 15.8 | 13.7 | 14.2039 | 14.05 | 14.45 | 14.4601 | 14 | 14.25 | 13.55 | 13.45 | 14.6 | 14.35 | 16.65 | 13.45
| Std Dev. |  | 0.17994 | 0.05164 | 0.86598 | 0.60663 | 1.5214 | 0.27869 | 0.4274 | 9.2899 | 0.19664 | 0.44609 | 0.10488 | 0.40497 | 0.36009 | 0.26658 | 0.4916 | 0.66458 | 0.018966 | 0.40825 | 0.43551 | 0.22509 | 0.19408 | 0.3266 | 0.3266 | 0.41673 | 0.33862
| Rel.Std.Dev. |  | 1.218% | 0.3497% | 5.581% | 4.494% | 11.11% | 1.938% | 2.996% | 47.8% | 1.382% | 3.317% | 0.7628% | 2.579% | 2.613% | 1.868% | 3.532% | 4.547% | 0.1311% | 2.882% | 3.021% | 1.671% | 1.447% | 2.232% | 2.279% | 2.508% | 2.502%
| PDM3 |  | 3.399% | 3.364% | 8.604% | -5.502% | -4.102% | 0.6809% | -0.1358% | 36.03% | -0.3691% | -5.852% | -3.752% | 9.897% | -3.519% | -0.0975% | -2.586% | 2.314% | 1.26% | -0.8358% | 0.9142% | -5.736% | -6.086% | 2.431% | 0.3309% | 16.31% | -5.269%
| Z-Score (Absolute) |  | 0.64 | 0.63 | 1.65 | 1.1 | 0.82 | 0.11 | 0.05 | 7 | 0.09 | 1.16 | 0.75 | 1.91 | 0.71 | 0.04 | 0.53 | 0.43 | 0.22 | NA | 0.16 | 1.14 | 1.21 | 0.45 | 0.04 | 3.16 | 1.05
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 506b (Certified Value 1.17 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.03: 2.06; 2.04: 2.12; 2.06: 2.14; 2.07: 2.15; 2.08: 2.17; 2.09: 2.18; 2.10: 2.20; 2.11: 2.21; 2.12: 2.24; 2.13: 2.25; 2.14: 2.27 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.06: 4A*OES/MS; 2.07: 4A*MS; 2.08: 4A*OES/MS; 2.09: 4A*MS; 2.10: 4A*OES/MS; 2.11: 4A*MS; 2.12: 4A*OES/MS; 2.13: 4A*OES/MS; 2.14: 4A*OES/MS |  | 2
3
| 1 | 1 | 2.01: 1.22; 2.02: 1.1; 2.03: 1.33; 2.04: 1.1; 2.06: 1.2; 2.07: 1.1007; 2.08: 1.1; 2.09: 1.2241; 2.10: 1.15; 2.11: 1; 2.12: 1.2; 2.13: 1.4; 2.14: 1 |  | 1
| 1 | 2 | 2.01: 1.23; 2.02: 1.3; 2.03: 1.31; 2.04: 1.1; 2.06: 1.2; 2.07: 1.1381; 2.08: 1.3; 2.09: 1.1213; 2.10: 1.13; 2.11: 1; 2.12: 1.2; 2.13: 1.4; 2.14: 1 |  | 25
| 1 | 3 | 2.01: 1.22; 2.02: 1.2; 2.03: 1.34; 2.04: 1.1; 2.06: 1.1; 2.07: 1.1589; 2.08: 1.3; 2.09: 1.1149; 2.10: 1.16; 2.11: 1; 2.12: 1.2; 2.13: 1.4; 2.14: 1.1 |  | 16
| 1 | 4 | 2.01: 1.25; 2.02: 1.1; 2.03: 1.33; 2.04: 1.1; 2.06: 1.2; 2.07: 1.2174; 2.08: 1.3; 2.09: 1.0649; 2.10: 1.21; 2.11: 1; 2.12: 1.2; 2.13: 1.5; 2.14: 1 |  | 1.1717
| 1 | 5 | 2.01: 1.23; 2.02: 1.2; 2.03: 1.35; 2.04: 1.1; 2.06: 1.1; 2.07: 1.1096; 2.08: 1.1; 2.09: 1.3001; 2.10: 1.19; 2.11: 1; 2.12: 1.1; 2.13: 1.4; 2.14: 1 |  | 68
| 1 | 6 | 2.01: 1.24; 2.02: 1.1; 2.03: 1.31; 2.04: 1; 2.06: 1.2; 2.07: 1.1823; 2.08: 1.2; 2.09: 1.2612; 2.10: 1.15; 2.11: 1; 2.12: 1.2; 2.13: 1.4; 2.14: 1
| Mean |  | 2.01: 1.2317; 2.02: 1.1667; 2.03: 1.3283; 2.04: 1.0833; 2.06: 1.1667; 2.07: 1.1512; 2.08: 1.2167; 2.09: 1.1811; 2.10: 1.165; 2.11: 1; 2.12: 1.1833; 2.13: 1.4167; 2.14: 1.0167
| Median |  | 2.01: 1.23; 2.02: 1.15; 2.03: 1.33; 2.04: 1.1; 2.06: 1.2; 2.07: 1.1485; 2.08: 1.25; 2.09: 1.1727; 2.10: 1.155; 2.11: 1; 2.12: 1.2; 2.13: 1.4; 2.14: 1
| Std Dev. |  | 2.01: 0.01169; 2.02: 0.08165; 2.03: 0.016021; 2.04: 0.040825; 2.06: 0.05164; 2.07: 0.044431; 2.08: 0.098319; 2.09: 0.093687; 2.10: 0.029496; 2.11: 0; 2.12: 0.040825; 2.13: 0.040825; 2.14: 0.040825
| Rel.Std.Dev. |  | 2.01: 0.9492%; 2.02: 6.999%; 2.03: 1.206%; 2.04: 3.768%; 2.06: 4.426%; 2.07: 3.86%; 2.08: 8.081%; 2.09: 7.932%; 2.10: 2.532%; 2.11: 0%; 2.12: 3.45%; 2.13: 2.882%; 2.14: 4.016%
| PDM3 |  | 2.01: 5.115%; 2.02: -0.4323%; 2.03: 13.36%; 2.04: -7.544%; 2.06: -0.4323%; 2.07: -1.755%; 2.08: 3.835%; 2.09: 0.798%; 2.10: -0.5745%; 2.11: -14.66%; 2.12: 0.9901%; 2.13: 20.9%; 2.14: -13.23%
| Z-Score (Absolute) |  | 2.01: 1.19; 2.02: 0.15; 2.03: 3.18; 2.04: 1.86; 2.06: 0.15; 2.07: 0.47; 2.08: 0.88; 2.09: 0.15; 2.10: 0.18; 2.11: NA; 2.12: 0.19; 2.13: 5; 2.14: 3.24
Results from laboratory 2.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 506b (Certified Value 112 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 0
0
| 1 | 1 | 115 | 111 | 112.9487 | 107 | 134.27 | 116 | 111 | 115 | 123 | 109 | 109 | 121 | 107 | 114.7705 | 104 | 104.1 | 109.54 | 109 | 109 | 111 | 110 | 94 | 107 | 112 | 115 | 108.4 | 1
| 1 | 2 | 118 | 112 | 110.1287 | 113.3 | 128.76 | 119 | 109 | 114 | 119 | 111 | 110 | 121 | 110 | 113.9891 | 107 | 111.7 | 105.93 | 109 | 110 | 136 | 118 | 94 | 108 | 113 | 117 | 106.3 | 26
| 1 | 3 | 116 | 110 | 112.0627 | 114 | 131.07 | 118 | 110 | 114 | 120 | 116 | 109 | 126 | 107 | 115.6967 | 107 | 109.6 | 106.08 | 109 | 105 | 124 | 115 | 93 | 108 | 114 | 117 | 105.3 | 16
| 1 | 4 | 115 | 111 | 114.3718 | 110 | 133.23 | 116 | 111 | 115 | 123 | 111 | 110 | 117 | 107 | 113.9822 | 105 | 112.7 | 105.46 | 110 | 109 | 105 | 109 | 95 | 108 | 112 | 118 | 102.9 | 111.906
| 1 | 5 | 117 | 108 | 106.3673 | 109 | 127.93 | 119 | 111 | 113 | 120 | 110 | 109 | 116 | 111 | 109.9206 | 100 | 106.5 | 110.85 | 110 | 107 | 120 | 108 | 93 | 108 | 111 | 118 | 106.9 | 69
| 1 | 6 | 117 | 108 | 107.7062 | 112 | 129.49 | 118 | 110 | 112 | 119 | 113 | 109 | 125 | 109 | 116.8226 | 110 | 111.7 | 105.33 | 114 | 109 | 118 | 112 | 92 | 109 | 110 | 116 | 106.3
| Mean |  | 116.3333 | 110 | 110.5976 | 110.8833 | 130.7917 | 117.6667 | 110.3333 | 113.8333 | 120.6667 | 111.6667 | 109.3333 | 121 | 108.5 | 114.197 | 105.5 | 109.3833 | 107.1983 | 110.1667 | 108.1667 | 119 | 112 | 93.5 | 108 | 112 | 116.8333 | 106.0167
| Median |  | 116.5 | 110.5 | 111.0957 | 111 | 130.28 | 118 | 110.5 | 114 | 120 | 111 | 109 | 121 | 108 | 114.3798 | 106 | 110.65 | 106.005 | 109.5 | 109 | 119 | 111 | 93.5 | 108 | 112 | 117 | 106.3
| Std Dev. |  | 1.2111 | 1.6733 | 3.1111 | 2.691 | 2.5351 | 1.3663 | 0.8165 | 1.169 | 1.8619 | 2.5033 | 0.5164 | 4.0497 | 1.7607 | 2.36 | 3.3912 | 3.4061 | 2.3744 | 1.9408 | 1.8348 | 10.7331 | 3.8471 | 1.0488 | 0.63246 | 1.4142 | 1.169 | 1.8357
| Rel.Std.Dev. |  | 1.041% | 1.521% | 2.813% | 2.427% | 1.938% | 1.161% | 0.74% | 1.027% | 1.543% | 2.242% | 0.4723% | 3.347% | 1.623% | 2.067% | 3.214% | 3.114% | 2.215% | 1.762% | 1.696% | 9.019% | 3.435% | 1.122% | 0.5856% | 1.263% | 1.001% | 1.731%
| PDM3 |  | 3.956% | -1.703% | -1.169% | -0.9138% | 16.88% | 5.148% | -1.405% | 1.722% | 7.829% | -0.2138% | -2.299% | 8.127% | -3.044% | 2.047% | -5.724% | -2.254% | -4.207% | -1.554% | -3.341% | 6.339% | 0.08404% | -16.45% | -3.49% | 0.08404% | 4.403% | -5.263%
| Z-Score (Absolute) |  | 1.29 | 0.17 | 0.03 | 0.03 | 4.64 | 1.6 | 0.09 | 0.72 | 2.29 | 0.21 | 0.33 | 2.37 | 0.52 | 0.8 | 1.21 | 0.31 | 0.82 | 0.13 | 0.6 | 1.91 | 0.29 | 3.99 | 0.63 | 0.29 | 1.41 | 1.09
| Analytical results for Zr in OREAS 506b (Certified Value 71 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 74.3 | 68.2 | 71.8181 | 68 | 81.23 | 68.8 | 72.4 | 74.1 | 73.4 | 67.2 | 66.7 | 71.9 | 82 | 70.8028 | 60.1 | 71 | 75.92 | 68 | 69 | 54 | 69.6 | 67.6 | 81 | 63.3 |  |  | 1
| 1 | 2 | 73.7 | 70.1 | 73.3198 | 72.6 | 84.37 | 72.2 | 69.2 | 77 | 73.1 | 70.4 | 67.2 | 71.4 | 81.8 | 70.8554 | 61.8 | 69.7 | 74.3 | 68 | 71 | 55 | 69.4 | 68.1 | 83 | 61.1 |  |  | 27
| 1 | 3 | 71.6 | 70.2 | 72.0743 | 72 | 85.46 | 71 | 71.7 | 75.8 | 71.9 | 73.4 | 65.6 | 72.6 | 83.2 | 70.1303 | 62.2 | 74.5 | 78.74 | 65 | 70 | 54 | 70.6 | 68.9 | 83 | 61.8 |  |  | 16
| 1 | 4 | 72.2 | 69.7 | 73.7996 | 70.2 | 85.78 | 69.2 | 70.3 | 79.6 | 73.9 | 68.1 | 66.4 | 73.3 | 80.2 | 71.5227 | 62.8 | 74.1 | 75.11 | 67 | 68 | 55 | 71 | 68.6 | 84 | 59.2 |  |  | 71.0075
| 1 | 5 | 73.9 | 68.7 | 70.8821 | 72.4 | 85.86 | 72.9 | 71.2 | 72.2 | 72.3 | 71.7 | 66.1 | 68.9 | 81.1 | 69.0435 | 60.2 | 68.8 | 78.59 | 70 | 70 | 57 | 71.7 | 67.2 | 88 | 63.1 |  |  | 70
| 1 | 6 | 68.5 | 69.1 | 70.8669 | 70.2 | 84.31 | 69.1 | 71.4 | 74.1 | 71.6 | 70.1 | 66.4 | 74.9 | 79.8 | 71.5434 | 62.8 | 73.9 | 77.19 | 67 | 70 | 55 | 69.5 | 68.8 | 82 | 63.6
| Mean |  | 72.3667 | 69.3333 | 72.1268 | 70.9 | 84.5017 | 70.5333 | 71.0333 | 75.4667 | 72.7 | 70.15 | 66.4 | 72.1667 | 81.35 | 70.6497 | 61.65 | 72 | 76.6417 | 67.5 | 69.6667 | 55 | 70.3 | 68.2 | 83.5 | 62.0167
| Median |  | 72.95 | 69.4 | 71.9462 | 71.1 | 84.915 | 70.1 | 71.3 | 74.95 | 72.7 | 70.25 | 66.4 | 72.25 | 81.45 | 70.8291 | 62 | 72.45 | 76.555 | 67.5 | 70 | 55 | 70.1 | 68.35 | 83 | 62.45
| Std Dev. |  | 2.1649 | 0.80166 | 1.2212 | 1.7743 | 1.741 | 1.7569 | 1.129 | 2.6059 | 0.90554 | 2.2775 | 0.54037 | 2.0136 | 1.2518 | 0.94614 | 1.2227 | 2.4819 | 1.8357 | 1.6432 | 1.0328 | 1.0954 | 0.94657 | 0.68993 | 2.429 | 1.6822
| Rel.Std.Dev. |  | 2.992% | 1.156% | 1.693% | 2.502% | 2.06% | 2.491% | 1.589% | 3.453% | 1.246% | 3.247% | 0.8138% | 2.79% | 1.539% | 1.339% | 1.983% | 3.447% | 2.395% | 2.434% | 1.482% | 1.992% | 1.346% | 1.012% | 2.909% | 2.712%
| PDM3 |  | 1.914% | -2.358% | 1.576% | -0.1514% | 19% | -0.6678% | 0.0364% | 6.28% | 2.384% | -1.208% | -6.489% | 1.632% | 14.57% | -0.5039% | -13.18% | 1.398% | 7.935% | -4.94% | -1.888% | -22.54% | -0.9964% | -3.954% | 17.59% | -12.66%
| Z-Score (Absolute) |  | 0.61 | 0.55 | 0.52 | 0.05 | 5.26 | 0.09 | 0.1 | 1.8 | 0.74 | 0.24 | 1.68 | 0.53 | 4.06 | 0.05 | 3.5 | 0.47 | 2.25 | 1.26 | 0.42 | 6.05 | 0.18 | 0.99 | 4.88 | 3.36
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C312: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C404: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C478: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C496: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C551: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C569: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C605: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C642: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C660: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C973: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C992: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1010: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1029: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1048: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1066: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1084: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1122: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1159: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 506b (Certified Value 1.31 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 3
| 1 | 1 | 1.42 | 1.23 | 1.3916 | 1.125 | 1.21 | 1.37 | 1.36 | 1.41 | 1.38 | 1.24 | 1.33 | 1.2 | 1.27 | 1.3382 | < 2 | 1.4 | 1.6205 | 1.2 | 1.37 | < 2 | 1.34 | 1.16 | 1.22 | 1.37 | 1.3 | 1.352 | 1
| 1 | 2 | 1.43 | 1.24 | 1.2866 | 1.192 | 1.2 | 1.35 | 1.34 | 1.45 | 1.37 | 1.28 | 1.34 | 1.2 | 1.28 | 1.3924 | < 2 | 1.4 | 1.495 | 1.2 | 1.27 | < 2 | 1.34 | 1.18 | 1.2 | 1.38 | 1.4 | 1.348 | 11
| 1 | 3 | 1.43 | 1.26 | 1.3412 | 1.065 | 1.2 | 1.35 | 1.35 | 1.38 | 1.37 | 1.26 | 1.36 | 1.2 | 1.31 | 1.3296 | < 2 | 1.3 | 1.2765 | 1.2 | 1.33 | < 2 | 1.35 | 1.2 | 1.22 | 1.42 | 1.3 | 1.376 | 16
| 1 | 4 | 1.4 | 1.23 | 1.2943 | 1.073 | 1.19 | 1.39 | 1.34 | 1.42 | 1.35 | 1.31 | 1.33 | 1.2 | 1.31 | 1.4054 | < 2 | 1.4 | 1.47 | 1.2 | 1.34 | < 2 | 1.3 | 1.2 | 1.21 | 1.41 | 1.3 | 1.36 | 1.3139
| 1 | 5 | 1.41 | 1.24 | 1.3287 | 1.211 | 1.2 | 1.31 | 1.32 | 1.4 | 1.34 | 1.25 | 1.3 | 1.1 | 1.33 | 1.3294 | < 2 | 1.3 | 1.64 | 1.2 | 1.34 | < 2 | 1.31 | 1.2 | 1.23 | 1.4 | 1.3 | 1.366 | 72
| 1 | 6 | 1.42 | 1.24 | 1.3824 | 1.206 | 1.21 | 1.34 | 1.4 | 1.36 | 1.35 | 1.29 | 1.31 | 1.1 | 1.3 | 1.3604 | < 2 | 1.4 | 1.4075 | 1.3 | 1.32 | < 2 | 1.35 | 1.2 | 1.19 | 1.42 | 1.3 | 1.334
| Mean |  | 1.4183 | 1.24 | 1.3375 | 1.1453 | 1.2017 | 1.3517 | 1.3517 | 1.4033 | 1.36 | 1.2717 | 1.3283 | 1.1667 | 1.3 | 1.3592 |  | 1.3667 | 1.4849 | 1.2167 | 1.3283 |  | 1.3317 | 1.19 | 1.2117 | 1.4 | 1.3167 | 1.356
| Median |  | 1.42 | 1.24 | 1.3349 | 1.1585 | 1.2 | 1.35 | 1.345 | 1.405 | 1.36 | 1.27 | 1.33 | 1.2 | 1.305 | 1.3493 |  | 1.4 | 1.4825 | 1.2 | 1.335 |  | 1.34 | 1.2 | 1.215 | 1.405 | 1.3 | 1.356
| Std Dev. |  | 0.01169 | 0.010954 | 0.043554 | 0.066737 | 0.0075277 | 0.027142 | 0.027142 | 0.031411 | 0.015492 | 0.026394 | 0.02137 | 0.05164 | 0.021909 | 0.033001 |  | 0.05164 | 0.13579 | 0.040825 | 0.033116 |  | 0.02137 | 0.016733 | 0.01472 | 0.020976 | 0.040825 | 0.014697
| Rel.Std.Dev. |  | 0.8242% | 0.8834% | 3.256% | 5.827% | 0.6264% | 2.008% | 2.008% | 2.238% | 1.139% | 2.076% | 1.609% | 4.426% | 1.685% | 2.428% |  | 3.779% | 9.144% | 3.355% | 2.493% |  | 1.605% | 1.406% | 1.215% | 1.498% | 3.101% | 1.084%
| PDM3 |  | 7.947% | -5.625% | 1.793% | -12.83% | -8.543% | 2.874% | 2.874% | 6.806% | 3.508% | -3.215% | 1.098% | -11.21% | -1.059% | 3.449% |  | 4.015% | 13.02% | -7.401% | 1.098% |  | 1.351% | -9.431% | -7.782% | 6.552% | 0.2098% | 3.203%
| Z-Score (Absolute) |  | 0.95 | 0.93 | 0.1 | 1.92 | 1.33 | 0.25 | 0.25 | 0.79 | 0.33 | 0.6 | 0 | 1.7 | 0.3 | 0.32 | 3.45 | 0.4 | 1.65 | 1.17 | 0 | 3.45 | 0.04 | 1.45 | 1.23 | 0.75 | 0.12 | 0.29
| Analytical results for Al in OREAS 506b (Certified Value 1.68 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 1.6886 | 1.93 | 1.6675 | 1.67 | 4.54 | 1.68 | 1.66 | 1.76 | 1.68 | 1.66 | 1.66 | 1.6 | 1.67 | 1.6935 | 1.788 | 1.65 | 1.72 | 1.77 | 1.592 | 1.72 | 1.7636 | 1.63 | 1.62 | 1.62 |  |  | 1
| 1 | 2 | 1.6908 | 1.94 | 1.5882 | 1.71 | 4.51 | 1.71 | 1.72 | 1.75 | 1.67 | 1.68 | 1.65 | 1.71 | 1.68 | 1.7124 | 1.724 | 1.6 | 1.74 | 1.67 | 1.628 | 1.7 | 1.8233 | 1.59 | 1.64 | 1.68 |  |  | 12
| 1 | 3 | 1.6631 | 1.95 | 1.5688 | 1.69 | 4.51 | 1.7 | 1.7 | 1.78 | 1.64 | 1.66 | 1.67 | 1.63 | 1.67 | 1.6982 | 1.766 | 1.61 | 1.71 | 1.9 | 1.577 | 1.68 | 1.8477 | 1.66 | 1.7 | 1.67 |  |  | 16
| 1 | 4 | 1.6639 | 1.9 | 1.6003 | 1.64 | 4.53 | 1.72 | 1.68 | 1.78 | 1.64 | 1.64 | 1.61 | 1.55 | 1.65 | 1.7131 | 1.771 | 1.63 | 1.71 | 1.83 | 1.536 | 1.7 | 1.8149 | 1.61 | 1.66 | 1.66 |  |  | 1.6783
| 1 | 5 | 1.6812 | 1.95 | 1.6304 | 1.72 | 4.54 | 1.7 | 1.71 | 1.76 | 1.63 | 1.68 | 1.61 | 1.53 | 1.67 | 1.7017 | 1.812 | 1.6 | 1.72 | 1.53 | 1.579 | 1.72 | 1.9031 | 1.68 | 1.6 | 1.63 |  |  | 73
| 1 | 6 | 1.6773 | 1.9 | 1.646 | 1.65 | 4.53 | 1.69 | 1.72 | 1.78 | 1.67 | 1.68 | 1.65 | 1.59 | 1.68 | 1.718 | 1.782 | 1.61 | 1.74 | 1.79 | 1.611 | 1.69 | 1.7874 | 1.57 | 1.64 | 1.62
| Mean |  | 1.6775 | 1.9283 | 1.6169 | 1.68 | 4.5267 | 1.7 | 1.6983 | 1.7683 | 1.655 | 1.6667 | 1.6417 | 1.6017 | 1.67 | 1.7061 | 1.7738 | 1.6167 | 1.7233 | 1.7483 | 1.5872 | 1.7017 | 1.8233 | 1.6233 | 1.6433 | 1.6467
| Median |  | 1.6792 | 1.935 | 1.6154 | 1.68 | 4.53 | 1.7 | 1.705 | 1.77 | 1.655 | 1.67 | 1.65 | 1.595 | 1.67 | 1.707 | 1.7765 | 1.61 | 1.72 | 1.78 | 1.5855 | 1.7 | 1.8191 | 1.62 | 1.64 | 1.645
| Std Dev. |  | 0.011884 | 0.023166 | 0.037418 | 0.032249 | 0.013663 | 0.014142 | 0.024014 | 0.013292 | 0.020736 | 0.01633 | 0.025626 | 0.064005 | 0.010954 | 0.0097038 | 0.029247 | 0.019664 | 0.013663 | 0.13091 | 0.031771 | 0.016021 | 0.048775 | 0.041793 | 0.034448 | 0.026583
| Rel.Std.Dev. |  | 0.7084% | 1.201% | 2.314% | 1.92% | 0.3018% | 0.8319% | 1.414% | 0.7516% | 1.253% | 0.9798% | 1.561% | 3.996% | 0.656% | 0.5688% | 1.649% | 1.216% | 0.7928% | 7.488% | 2.002% | 0.9415% | 2.675% | 2.575% | 2.096% | 1.614%
| PDM3 |  | -0.05122% | 14.9% | -3.663% | 0.09873% | 169.7% | 1.29% | 1.191% | 5.362% | -1.391% | -0.6957% | -2.185% | -4.569% | -0.4971% | 1.657% | 5.69% | -3.675% | 2.681% | 4.17% | -5.433% | 1.39% | 8.639% | -3.278% | -2.086% | -1.887%
| Z-Score (Absolute) |  | 0.02 | 4.12 | 1.02 | 0.02 | 47.03 | 0.35 | 0.32 | 1.48 | 0.39 | 0.2 | 0.61 | 1.27 | 0.14 | 0.45 | 1.57 | 1.03 | 0.74 | 1.15 | 1.51 | 0.38 | 2.39 | 0.91 | 0.58 | 0.53
| Analytical results for As in OREAS 506b (Certified Value 13 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 2
| 1 | 1 | 14 | 13 | 12.0091 | 13.4 | 18.5 | 12.6 | 12.6 | 13.4 | 12.9 | 12.6 | 12 | 13 | 13 | 12.7175 | 19 | 13.4 | 14.45 | 14 | 14.6 | 18 | 13.8 | NA | 11 | 12.2 | 12.4 | 12.4 | 1
| 1 | 2 | 13 | 13 | 11.9478 | 13.7 | 19.3 | 12.5 | 13.1 | 13.3 | 12.7 | 12.4 | 12 | 15 | 14 | 12.8894 | 18 | 13.5 | 12.32 | 14 | 12.1 | 18 | 14.2 | NA | 11 | 12.5 | 12.7 | 12.5 | 13
| 1 | 3 | 13 | 13 | 11.9323 | 13.1 | 19.7 | 12.4 | 13.1 | 13.2 | 13.4 | 12.6 | 12 | 14 | 12 | 13.1455 | 16 | 12.8 | 12.85 | 12 | 13.3 | 18 | 13.8 | NA | 11 | 13.5 | 12.8 | 12.3 | 16
| 1 | 4 | 13 | 13 | 11.8369 | 13.5 | 17.7 | 13.2 | 12.8 | 13.8 | 13.4 | 13.1 | 13 | 14 | 12 | 12.7501 | 16 | 13.6 | 12.44 | 12 | 13.5 | 18 | 14.2 | NA | 12 | 12.3 | 12.9 | 12.5 | 12.9541
| 1 | 5 | 13 | 13 | 12.1405 | 13.6 | 17.6 | 12.6 | 12.4 | 13 | 13.1 | 12.8 | 12 | 13 | 12 | 12.6758 | 18 | 13 | 13.36 | 12 | 13.2 | 20 | 14.1 | NA | 13 | 13.2 | 13.1 | 12.5 | 74
| 1 | 6 | 13 | 13 | 12.0979 | 13.1 | 17.1 | 11.9 | 13.2 | 12.8 | 13.4 | 12.3 | 13 | 12 | 13 | 12.7018 | 16 | 13.2 | 12.9 | 13 | 13.5 | 19 | 14.3 | NA | 11 | 13.5 | 12.5 | 12.2
| Mean |  | 13.1667 | 13 | 11.9941 | 13.4 | 18.3167 | 12.5333 | 12.8667 | 13.25 | 13.15 | 12.6333 | 12.3333 | 13.5 | 12.6667 | 12.8134 | 17.1667 | 13.25 | 13.0533 | 12.8333 | 13.3667 | 18.5 | 14.0667 |  | 11.5 | 12.8667 | 12.7333 | 12.4
| Median |  | 13 | 13 | 11.9785 | 13.45 | 18.1 | 12.55 | 12.95 | 13.25 | 13.25 | 12.6 | 12 | 13.5 | 12.5 | 12.7338 | 17 | 13.3 | 12.875 | 12.5 | 13.4 | 18 | 14.15 |  | 11 | 12.85 | 12.75 | 12.45
| Std Dev. |  | 0.40825 | 0 | 0.11233 | 0.25298 | 1.0284 | 0.41793 | 0.32042 | 0.34496 | 0.30166 | 0.28752 | 0.5164 | 1.0488 | 0.8165 | 0.17926 | 1.3292 | 0.30822 | 0.77763 | 0.98319 | 0.79917 | 0.83666 | 0.21602 |  | 0.83666 | 0.60222 | 0.2582 | 0.12649
| Rel.Std.Dev. |  | 3.101% | 0% | 0.9365% | 1.888% | 5.615% | 3.335% | 2.49% | 2.604% | 2.294% | 2.276% | 4.187% | 7.769% | 6.446% | 1.399% | 7.743% | 2.326% | 5.957% | 7.661% | 5.979% | 4.522% | 1.536% |  | 7.275% | 4.68% | 2.028% | 1.02%
| PDM3 |  | 1.641% | 0.3546% | -7.411% | 3.442% | 41.4% | -3.248% | -0.6747% | 2.284% | 1.513% | -2.476% | -4.792% | 4.214% | -2.219% | -1.086% | 32.52% | 2.284% | 0.7663% | -0.932% | 3.185% | 42.81% | 8.589% |  | -11.22% | -0.6747% | -1.704% | -4.277%
| Z-Score (Absolute) |  | NA | NA | 2.11 | 0.96 | 11.71 | 0.93 | 0.2 | 0.63 | 0.42 | 0.71 | NA | NA | NA | 0.32 | NA | 0.63 | 0.2 | NA | 0.89 | NA | 2.42 | NA | NA | 0.2 | 0.5 | 1.22
Results from laboratories 2.01, 2.02, 2.12, 2.13, 2.14, 2.16, 2.19, 2.21 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 506b (Certified Value < 10 ppm) | [27 empty] | Certified
B | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab | 1
| No | No | 2.01: 2.03; 2.02: 2.04; 2.03: 2.07; 2.04: 2.08; 2.06: 2.09; 2.07: 2.10; 2.08: 2.11; 2.09: 2.14; 2.10: 2.15; 2.11: 2.22; 2.12: 2.23; 2.13: 2.25; 2.14: 2.27 | ppm
|  |  | 2.01: AR*OES/MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*OES/MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*OES/MS; 2.13: AR*OES/MS; 2.14: AR*MS | 1
|  |  | 2.01: 01g; 2.02: 0.5g; 2.03: 0.5g; 2.04: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 0.25g; 2.10: 0.2g; 2.11: 0.5g; 2.12: 0.25g; 2.13: 0.5g; 2.14: 30g | 1
| 1 | 1 | 2.01: < 10; 2.02: < 20; 2.03: < 10; 2.04: < 10; 2.06: < 10; 2.07: < 10; 2.08: < 10; 2.09: < 10; 2.10: < 2; 2.11: < 10; 2.12: < 20; 2.13: 2; 2.14: 1 | 1
| 1 | 2 | 2.01: < 10; 2.02: < 20; 2.03: < 10; 2.04: < 10; 2.06: < 10; 2.07: < 10; 2.08: < 10; 2.09: < 10; 2.10: < 2; 2.11: < 10; 2.12: < 20; 2.13: 2; 2.14: 1 | 31
| 1 | 3 | 2.01: < 10; 2.02: < 20; 2.03: < 10; 2.04: < 10; 2.06: < 10; 2.07: < 10; 2.08: < 10; 2.09: < 10; 2.10: < 2; 2.11: < 10; 2.12: < 20; 2.13: 2; 2.14: 1 | 16
| 1 | 4 | 2.01: < 10; 2.02: < 20; 2.03: < 10; 2.04: < 10; 2.06: < 10; 2.07: < 10; 2.08: < 10; 2.09: < 10; 2.10: < 2; 2.11: < 10; 2.12: < 20; 2.13: 2; 2.14: 1 | < 10
| 1 | 5 | 2.01: < 10; 2.02: < 20; 2.03: < 10; 2.04: < 10; 2.06: < 10; 2.07: < 10; 2.08: < 10; 2.09: < 10; 2.10: < 2; 2.11: < 10; 2.12: < 20; 2.13: 2; 2.14: 1 | 75
| 1 | 6 | 2.01: < 10; 2.02: < 20; 2.03: < 10; 2.04: < 10; 2.06: < 10; 2.07: < 10; 2.08: < 10; 2.09: < 10; 2.10: < 2; 2.11: < 10; 2.12: < 20; 2.13: 2; 2.14: 1
| Mean |  | 2.13: 2; 2.14: 1
| Median |  | 2.13: 2; 2.14: 1
| Std Dev. |  | 2.13: 0; 2.14: 0
Rel.Std.Dev. | bar: 2.13 to 2.14
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.14
| Analytical results for Ba in OREAS 506b (Certified Value 426 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 0
| 1 | 1 | 155 | 456 | 466.428 | 406.6 | 406.3 | 420 | 410 | 450 | 420 | 410 | 418 | 240 | 448 | 419.3537 | 421 | 433 | 411 | 424 | 458 | 224 | 502.55 | 440 | 28.5 | 356.5 |  |  | 1
| 1 | 2 | 229 | 457 | 450.723 | 412.9 | 403.6 | 430 | 410 | 450 | 420 | 420 | 417 | 254 | 439 | 427.2663 | 403 | 425 | 409 | 419 | 474 | 233 | 506 | 435 | 36.4 | 400.7 |  |  | 14
| 1 | 3 | 203 | 450 | 465.7692 | 411.7 | 425.5 | 440 | 410 | 450 | 410 | 420 | 422 | 246 | 443 | 425.7729 | 414 | 428 | 401 | 416 | 459 | 226 | 507.15 | 434 | 31.2 | 393.9 |  |  | 16
| 1 | 4 | 206 | 454 | 450.2284 | 404.6 | 402.9 | 440 | 420 | 440 | 410 | 410 | 407 | 220 | 428 | 436.0928 | 412 | 436 | 402 | 520 | 431 | 240 | 502.55 | 458 | 29.8 | 364.3 |  |  | 425.9311
| 1 | 5 | 163 | 451 | 454.761 | 414 | 400.8 | 440 | 410 | 450 | 410 | 420 | 411 | 258 | 446 | 431.1647 | 418 | 425 | 403 | 429 | 463 | 236 | 509.8 | 452 | 31.1 | 368.5 |  |  | 76
| 1 | 6 | 164 | 455 | 465.0711 | 401 | 403.2 | 430 | 430 | 460 | 400 | 420 | 420 | 240 | 434 | 436.5503 | 411 | 427 | 412 | 423 | 470 | 268 | 509.2 | 431 | 44.7 | 370.1
| Mean |  | 186.6667 | 453.8333 | 458.8301 | 408.4667 | 407.05 | 433.3333 | 415 | 450 | 411.6667 | 416.6667 | 415.8333 | 243 | 439.6667 | 429.3668 | 413.1667 | 429 | 406.3333 | 438.5 | 459.1667 | 237.8333 | 506.2083 | 441.6667 | 33.6167 | 375.6667
| Median |  | 183.5 | 454.5 | 459.9161 | 409.15 | 403.4 | 435 | 410 | 450 | 410 | 420 | 417.5 | 243 | 441 | 429.2155 | 413 | 427.5 | 406 | 423.5 | 461 | 234.5 | 506.575 | 437.5 | 31.15 | 369.3
| Std Dev. |  | 30.0311 | 2.7869 | 7.7602 | 5.1945 | 9.2084 | 8.165 | 8.3666 | 6.3246 | 7.5277 | 5.164 | 5.7067 | 13.4313 | 7.607 | 6.5979 | 6.2423 | 4.5166 | 4.8854 | 40.1734 | 15.1449 | 15.9551 | 3.1479 | 10.8934 | 6.0582 | 17.5387
| Rel.Std.Dev. |  | 16.09% | 0.6141% | 1.691% | 1.272% | 2.262% | 1.884% | 2.016% | 1.405% | 1.829% | 1.239% | 1.372% | 5.527% | 1.73% | 1.537% | 1.511% | 1.053% | 1.202% | 9.162% | 3.298% | 6.709% | 0.6219% | 2.466% | 18.02% | 4.669%
| PDM3 |  | -56.17% | 6.551% | 7.724% | -4.1% | -4.433% | 1.738% | -2.566% | 5.651% | -3.349% | -2.175% | -2.371% | -42.95% | 3.225% | 0.8066% | -2.997% | 0.7205% | -4.601% | 2.951% | 7.803% | -44.16% | 18.85% | 3.694% | -92.11% | -11.8%
| Z-Score (Absolute) |  | 6.78 | 1.11 | 1.26 | 0.23 | 0.27 | 0.5 | 0.04 | 1 | 0.14 | 0.01 | 0.01 | 5.12 | 0.69 | 0.39 | 0.09 | 0.38 | 0.29 | 0.66 | 1.27 | 5.27 | 2.66 | 0.75 | 11.3 | 1.2
| Analytical results for Be in OREAS 506b (Certified Value 1.21 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 1.25 | 1.34 | 1.2008 | < 2 | 1.21 | 1.22 | 1.3 | 1.19 | 1.2 | 1.4 | 0.7 | 1.2 | 1.2388 | 0.8 | 1.2 | 1.34 | 1.16 | 1.04 | 1.4 | 1.2 | [6 empty] | 1
| 1 | 2 | 1.27 | 1.34 | 1.1903 | < 2 | 1.24 | 1.26 | 1.29 | 1.18 | 1.21 | 1.4 | 0.8 | 1.2 | 1.2151 | 0.8 | 1.1 | 1.19 | 1.13 | 1 | 1.3 | 1.3 | [6 empty] | 15
| 1 | 3 | 1.25 | 1.34 | 1.1782 | < 2 | 1.22 | 1.24 | 1.31 | 1.16 | 1.2 | 1.4 | 0.7 | 1.2 | 1.206 | 0.8 | 1.2 | 1.4 | 1.16 | 1.04 | 1.4 | 1.3 | [6 empty] | 16
| 1 | 4 | 1.23 | 1.34 | 1.1849 | < 2 | 1.24 | 1.22 | 1.31 | 1.16 | 1.19 | 1.4 | 0.7 | 1.2 | 1.1956 | 0.9 | 1.2 | 1.37 | 1.16 | 0.98 | 1.3 | 1.3 | [6 empty] | 1.2125
| 1 | 5 | 1.22 | 1.31 | 1.2079 | < 2 | 1.22 | 1.24 | 1.3 | 1.15 | 1.21 | 1.4 | 0.7 | 1.3 | 1.2341 | 0.8 | 1.2 | 1.38 | 1.15 | 0.99 | 1.4 | 1.2 | [6 empty] | 77
| 1 | 6 | 1.23 | 1.32 | 1.1923 | < 2 | 1.22 | 1.26 | 1.31 | 1.17 | 1.16 | 1.4 | 0.7 | 1.3 | 1.1913 | 0.8 | 1.2 | 1.37 | 1.17 | 1.06 | 1.4 | 1.2
| Mean |  | 1.2417 | 1.3317 | 1.1924 |  | 1.225 | 1.24 | 1.3033 | 1.1683 | 1.195 | 1.4 | 0.71667 | 1.2333 | 1.2135 | 0.81667 | 1.1833 | 1.3417 | 1.155 | 1.0183 | 1.3667 | 1.25
| Median |  | 1.24 | 1.34 | 1.1913 |  | 1.22 | 1.24 | 1.305 | 1.165 | 1.2 | 1.4 | 0.7 | 1.2 | 1.2105 | 0.8 | 1.2 | 1.37 | 1.16 | 1.02 | 1.4 | 1.25
| Std Dev. |  | 0.018348 | 0.013292 | 0.010705 |  | 0.012247 | 0.017889 | 0.008165 | 0.01472 | 0.018708 | 2.4324e-16 | 0.040825 | 0.05164 | 0.019691 | 0.040825 | 0.040825 | 0.07679 | 0.013784 | 0.032506 | 0.05164 | 0.054772
| Rel.Std.Dev. |  | 1.478% | 0.9981% | 0.8978% |  | 0.9998% | 1.443% | 0.6265% | 1.26% | 1.566% | 1.737e-14% | 5.696% | 4.187% | 1.623% | 4.999% | 3.45% | 5.723% | 1.193% | 3.192% | 3.779% | 4.382%
| PDM3 |  | 2.41% | 9.832% | -1.654% |  | 1.035% | 2.272% | 7.496% | -3.639% | -1.439% | 15.47% | -40.89% | 1.722% | 0.08326% | -32.64% | -2.402% | 10.66% | -4.739% | -16.01% | 12.72% | 3.097%
| Z-Score (Absolute) |  | 0.69 | 2.34 | 0.22 | 3.76 | 0.38 | 0.66 | 1.82 | 0.66 | 0.17 | NA | NA | NA | 0.17 | NA | 0.38 | 2.53 | 0.91 | 3.42 | NA | NA
Results from laboratories 2.12, 2.13, 2.14, 2.16, 2.24 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 506b (Certified Value 0.96 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  | 2
| 1 | 1 | 1.03 | 0.97 | 1.1011 | 0.94 | < 5 | 0.89 | 1.01 | 1.18 | 0.92 | 0.83 | 0.97 | < 2 | 1.03 | 0.90781 | < 5 | 0.88 | < 5 | 0.87 | 0.76 | NA | 0.92 | 1.03 | 0.89 | 0.99 |  |  | 1
| 1 | 2 | 1 | 0.98 | 0.92915 | 0.92 | < 5 | 0.9 | 1.17 | 0.97 | 0.95 | 0.89 | 1.03 | 7 | 1.01 | 1.0082 | 7 | 0.94 | < 5 | 0.93 | 0.81 | NA | 0.81 | 1.07 | 1.45 | 1.08 |  |  | 16
| 1 | 3 | 1.02 | 0.96 | 0.79669 | 1.24 | < 5 | 0.89 | 0.85 | 0.97 | 1.15 | 0.94 | 0.97 | < 2 | 1.1 | 0.85271 | < 5 | 0.94 | < 5 | 0.83 | 0.83 | NA | 0.97 | 0.94 | 1.16 | 1.01 |  |  | 16
| 1 | 4 | 0.98 | 0.96 | 0.84068 | 0.99 | < 5 | 0.96 | 0.91 | 1.13 | 0.87 | 0.83 | 0.86 | < 2 | 1 | 0.85106 | < 5 | 1.03 | < 5 | 0.92 | 0.78 | NA | 0.72 | 1.03 | 0.95 | 1.03 |  |  | 0.95816
| 1 | 5 | 1.07 | 0.97 | 0.96288 | 0.99 | < 5 | 1.03 | 1.14 | 0.9 | 0.77 | 0.88 | 0.89 | < 2 | 1.04 | 0.9446 | < 5 | 1.02 | < 5 | 0.88 | 0.8 | NA | 6.2 | 1.09 | 1 | 1.06 |  |  | 78
| 1 | 6 | 1.12 | 0.96 | 1.0528 | 0.91 | < 5 | 0.99 | 1.24 | 1.04 | 1 | 0.88 | 0.78 | < 2 | 1.03 | 0.99876 | < 5 | 1 | < 5 | 0.89 | 0.83 | NA | 0.99 | 0.84 | 1.1 | 1.03
| Mean |  | 1.0367 | 0.96667 | 0.94721 | 0.99833 |  | 0.94333 | 1.0533 | 1.0317 | 0.94333 | 0.875 | 0.91667 | 7 | 1.035 | 0.92718 | 7 | 0.96833 |  | 0.88667 | 0.80167 |  | 1.7683 | 1 | 1.0917 | 1.0333
| Median |  | 1.025 | 0.965 | 0.94602 | 0.965 |  | 0.93 | 1.075 | 1.005 | 0.935 | 0.88 | 0.93 | 7 | 1.03 | 0.92621 | 7 | 0.97 |  | 0.885 | 0.805 |  | 0.945 | 1.03 | 1.05 | 1.03
| Std Dev. |  | 0.050859 | 0.008165 | 0.11782 | 0.12319 |  | 0.059217 | 0.15475 | 0.10647 | 0.12801 | 0.041352 | 0.090701 |  | 0.035071 | 0.068845 |  | 0.058109 |  | 0.036148 | 0.027869 |  | 2.1735 | 0.093808 | 0.20114 | 0.03266
| Rel.Std.Dev. |  | 4.906% | 0.8447% | 12.44% | 12.34% |  | 6.277% | 14.69% | 10.32% | 13.57% | 4.726% | 9.895% |  | 3.389% | 7.425% |  | 6.001% |  | 4.077% | 3.476% |  | 122.9% | 9.381% | 18.42% | 3.161%
| PDM3 |  | 8.193% | 0.8877% | -1.143% | 4.193% |  | -1.548% | 9.933% | 7.672% | -1.548% | -8.679% | -4.331% | 630.6% | 8.019% | -3.233% | 630.6% | 1.062% |  | -7.462% | -16.33% |  | 84.55% | 4.367% | 13.93% | 7.845%
| Z-Score (Absolute) |  | 0.47 | 0.39 | 0.63 | 0 | 18.41 | 0.67 | 0.67 | 0.41 | 0.67 | 1.51 | 1 | NA | 0.45 | 0.87 | NA | 0.37 | 18.41 | 1.37 | 2.41 | NA | 9.44 | 0.02 | 1.14 | 0.43
Results from laboratories 2.13 and 2.16 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 506b (Certified Value 0.695 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 0.69 | 0.73 | 0.68441 | 0.65 | 1.27 | 0.71 | 0.69 | 0.74 | 0.73 | 0.71 | 0.72 | 0.75 | 0.69 | 0.68571 | 0.714 | 0.69 | 0.7 | 0.63 | 0.73 | 0.7 | 0.71598 | 0.65 | 0.8 | 0.6411 | 0.69 |  | 1
| 1 | 2 | 0.68 | 0.72 | 0.65188 | 0.65 | 1.26 | 0.71 | 0.7 | 0.73 | 0.72 | 0.71 | 0.74 | 0.797 | 0.69 | 0.69824 | 0.695 | 0.67 | 0.71 | 0.629 | 0.74 | 0.72 | 0.71176 | 0.64 | 0.79 | 0.6544 | 0.69 |  | 17
| 1 | 3 | 0.67 | 0.72 | 0.65216 | 0.66 | 1.2 | 0.71 | 0.7 | 0.75 | 0.71 | 0.71 | 0.73 | 0.747 | 0.68 | 0.69321 | 0.713 | 0.67 | 0.71 | 0.657 | 0.729 | 0.71 | 0.72236 | 0.66 | 0.83 | 0.6543 | 0.7 |  | 16
| 1 | 4 | 0.68 | 0.71 | 0.65339 | 0.65 | 1.27 | 0.72 | 0.69 | 0.74 | 0.71 | 0.7 | 0.71 | 0.72 | 0.67 | 0.69765 | 0.707 | 0.68 | 0.71 | 0.647 | 0.709 | 0.7 | 0.72552 | 0.65 | 0.8 | 0.6514 | 0.69 |  | 0.69545
| 1 | 5 | 0.68 | 0.73 | 0.67021 | 0.66 | 1.28 | 0.71 | 0.69 | 0.73 | 0.7 | 0.71 | 0.72 | 0.719 | 0.68 | 0.69238 | 0.724 | 0.67 | 0.71 | 0.675 | 0.724 | 0.69 | 0.72065 | 0.64 | 0.77 | 0.6622 | 0.69 |  | 79
| 1 | 6 | 0.68 | 0.72 | 0.68269 | 0.65 | 1.27 | 0.71 | 0.71 | 0.73 | 0.72 | 0.69 | 0.71 | 0.73 | 0.68 | 0.69612 | 0.71 | 0.68 | 0.71 | 0.65 | 0.731 | 0.69 | 0.72012 | 0.63 | 0.76 | 0.6513 | 0.68
| Mean |  | 0.68 | 0.72167 | 0.66579 | 0.65333 | 1.2583 | 0.71167 | 0.69667 | 0.73667 | 0.715 | 0.705 | 0.72167 | 0.74383 | 0.68167 | 0.69389 | 0.7105 | 0.67667 | 0.70833 | 0.648 | 0.72717 | 0.70167 | 0.7194 | 0.645 | 0.79167 | 0.65245 | 0.69
| Median |  | 0.68 | 0.72 | 0.6618 | 0.65 | 1.27 | 0.71 | 0.695 | 0.735 | 0.715 | 0.71 | 0.72 | 0.7385 | 0.68 | 0.69467 | 0.7115 | 0.675 | 0.71 | 0.6485 | 0.7295 | 0.7 | 0.72039 | 0.645 | 0.795 | 0.65285 | 0.69
| Std Dev. |  | 0.0063246 | 0.0075277 | 0.015392 | 0.005164 | 0.029269 | 0.0040825 | 0.008165 | 0.008165 | 0.010488 | 0.0083666 | 0.01169 | 0.029158 | 0.0075277 | 0.0046395 | 0.0095237 | 0.008165 | 0.0040825 | 0.017321 | 0.010304 | 0.01169 | 0.0048639 | 0.010488 | 0.024833 | 0.0068343 | 0.0063246
| Rel.Std.Dev. |  | 0.9301% | 1.043% | 2.312% | 0.7904% | 2.326% | 0.5737% | 1.172% | 1.108% | 1.467% | 1.187% | 1.62% | 3.92% | 1.104% | 0.6686% | 1.34% | 1.207% | 0.5764% | 2.673% | 1.417% | 1.666% | 0.6761% | 1.626% | 3.137% | 1.047% | 0.9166%
| PDM3 |  | -2.222% | 3.769% | -4.265% | -6.056% | 80.94% | 2.332% | 0.1747% | 5.926% | 2.811% | 1.373% | 3.769% | 6.957% | -1.982% | -0.2252% | 2.164% | -2.701% | 1.852% | -6.823% | 4.56% | 0.8937% | 3.443% | -7.255% | 13.83% | -6.183% | -0.7839%
| Z-Score (Absolute) |  | 0.76 | 0.51 | 1.19 | 1.57 | 16.83 | 0.2 | 0.25 | 0.96 | 0.3 | 0 | 0.51 | 1.18 | 0.71 | 0.34 | 0.17 | 0.86 | 0.1 | 1.73 | 0.67 | 0.1 | 0.44 | 1.83 | 2.64 | 1.6 | 0.46
| Analytical results for Cd in OREAS 506b (Certified Value 0.32 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 2
| 1 | 1 | 0.29 | 0.32 | 0.295 | 0.34 | < 1 | 0.28 | 0.28 | 0.32 | 0.3 | 0.29 | 0.31 | < 0.5 | 0.34 | < 0.5 | < 1 | 0.4 | < 1 | 0.28 | 0.4 | 0.35 | 0.3 | 0.37 | 0.31 | 0.29 |  |  | 1
| 1 | 2 | 0.28 | 0.29 | 0.33153 | 0.35 | < 1 | 0.3 | 0.28 | 0.34 | 0.33 | 0.3 | 0.32 | < 0.5 | 0.36 | < 0.5 | < 1 | 0.4 | < 1 | 0.29 | 0.33 | 0.35 | 0.3 | 0.37 | 0.31 | 0.3 |  |  | 18
| 1 | 3 | 0.31 | 0.3 | 0.32134 | 0.34 | < 1 | 0.27 | 0.26 | 0.32 | 0.3 | 0.31 | 0.32 | < 0.5 | 0.35 | < 0.5 | < 1 | 0.38 | < 1 | 0.28 | 0.41 | 0.34 | 0.3 | 0.38 | 0.34 | 0.29 |  |  | 16
| 1 | 4 | 0.3 | 0.31 | 0.29669 | 0.32 | < 1 | 0.3 | 0.3 | 0.31 | 0.3 | 0.3 | 0.32 | < 0.5 | 0.34 | < 0.5 | < 1 | 0.4 | < 1 | 0.29 | 0.33 | 0.34 | 0.27 | 0.37 | 0.36 | 0.28 |  |  | 0.31946
| 1 | 5 | 0.3 | 0.31 | 0.31615 | 0.39 | < 1 | 0.29 | 0.25 | 0.32 | 0.3 | 0.28 | 0.32 | < 0.5 | 0.36 | < 0.5 | < 1 | 0.38 | < 1 | 0.27 | 0.31 | 0.36 | 0.3 | 0.38 | 0.36 | 0.29 |  |  | 80
| 1 | 6 | 0.28 | 0.3 | 0.3242 | 0.37 | < 1 | 0.28 | 0.32 | 0.32 | 0.3 | 0.26 | 0.32 | < 0.5 | 0.35 | < 0.5 | < 1 | 0.36 | < 1 | 0.27 | 0.33 | 0.38 | 0.24 | 0.38 | 0.29 | 0.28
| Mean |  | 0.29333 | 0.305 | 0.31415 | 0.35167 |  | 0.28667 | 0.28167 | 0.32167 | 0.305 | 0.29 | 0.31833 |  | 0.35 |  |  | 0.38667 |  | 0.28 | 0.35167 | 0.35333 | 0.285 | 0.375 | 0.32833 | 0.28833
| Median |  | 0.295 | 0.305 | 0.31874 | 0.345 |  | 0.285 | 0.28 | 0.32 | 0.3 | 0.295 | 0.32 |  | 0.35 |  |  | 0.39 |  | 0.28 | 0.33 | 0.35 | 0.3 | 0.375 | 0.325 | 0.29
| Std Dev. |  | 0.012111 | 0.010488 | 0.015036 | 0.024833 |  | 0.012111 | 0.025626 | 0.0098319 | 0.012247 | 0.017889 | 0.0040825 |  | 0.0089443 |  |  | 0.01633 |  | 0.0089443 | 0.042151 | 0.015055 | 0.0251 | 0.0054772 | 0.029269 | 0.0075277
| Rel.Std.Dev. |  | 4.129% | 3.439% | 4.786% | 7.061% |  | 4.225% | 9.098% | 3.057% | 4.016% | 6.168% | 1.282% |  | 2.556% |  |  | 4.223% |  | 3.194% | 11.99% | 4.261% | 8.807% | 1.461% | 8.914% | 2.611%
| PDM3 |  | -8.18% | -4.528% | -1.663% | 10.08% |  | -10.27% | -11.83% | 0.6895% | -4.528% | -9.223% | -0.3539% |  | 9.558% |  |  | 21.04% |  | -12.35% | 10.08% | 10.6% | -10.79% | 17.38% | 2.776% | -9.745%
| Z-Score (Absolute) |  | 0.45 | 0.22 | 0.04 | 0.7 | 3.63 | 0.58 | 0.68 | 0.11 | 0.22 | 0.52 | 0.04 | 1.31 | 0.67 | 1.31 | 3.63 | 1.39 | 3.63 | 0.72 | 0.7 | 0.73 | 0.62 | 1.16 | 0.24 | 0.55
| Analytical results for Ce in OREAS 506b (Certified Value 29.4 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 1
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [6 empty] | 2
| 1 | 1 | 28.443 | 47 | 30.6862 | 30 | 29.2 | 28.4 | 31.7 | 29.6 | 26.7 | 31.01 | 26.57 | 30.4388 | 29.4 | 53.3434 | 25 | 51.6 | 36.18 | 26.51 | 29.5 | 31.3 | [6 empty] | 1
| 1 | 2 | 28.281 | 47.3 | 30.1996 | 29 | 29.9 | 29.6 | 32.5 | 29.7 | 28.2 | 30.08 | 25.78 | 30.8394 | 28.8 | 53.4722 | 25 | 45.9 | 35.52 | 25.92 | 29.8 | 31.3 | [6 empty] | 19
| 1 | 3 | 28.334 | 47.1 | 30.3333 | 32.4 | 29.2 | 28.7 | 31.2 | 29.8 | 28.4 | 30.49 | 26.5 | 30.7186 | 29.1 | 53.6416 | 24 | 58.2 | 35.57 | 26.83 | 30 | 32.4 | [6 empty] | 16
| 1 | 4 | 28.524 | 46.9 | 30.3316 | 28.9 | 29.4 | 29.3 | 31.8 | 30.6 | 28.3 | 29.28 | 26.18 | 31.7611 | 30.1 | 52.962 | 23 | 57.3 | 34.8 | 24.92 | 30.2 | 31.5 | [6 empty] | 29.4208
| 1 | 5 | 28.44 | 47 | 30.7694 | 30.6 | 28.1 | 27.9 | 32.1 | 29.2 | 28.5 | 29.51 | 26.66 | 31.1667 | 28.9 | 53.0754 | 23 | 55.4 | 34.54 | 26.21 | 30.8 | 30 | [6 empty] | 81
| 1 | 6 | 28.907 | 46.6 | 30.9588 | 29 | 29.1 | 29.7 | 32.5 | 27.8 | 25.9 | 29.64 | 26.46 | 31.9639 | 29.4 | 53.6772 | 25 | 53.8 | 36.28 | 26.97 | 28.8 | 31.7
| Mean |  | 28.4882 | 46.9833 | 30.5465 | 29.9833 | 29.15 | 28.9333 | 31.9667 | 29.45 | 27.6667 | 30.0017 | 26.3583 | 31.1481 | 29.2833 | 53.362 | 24.1667 | 53.7 | 35.4817 | 26.2267 | 29.85 | 31.3667
| Median |  | 28.4415 | 47 | 30.5098 | 29.5 | 29.2 | 29 | 31.95 | 29.65 | 28.25 | 29.86 | 26.48 | 31.003 | 29.25 | 53.4078 | 24.5 | 54.6 | 35.545 | 26.36 | 29.9 | 31.4
| Std Dev. |  | 0.2226 | 0.23166 | 0.30038 | 1.3659 | 0.58907 | 0.7174 | 0.50465 | 0.92898 | 1.093 | 0.65646 | 0.32634 | 0.60408 | 0.47081 | 0.29399 | 0.98319 | 4.5024 | 0.70508 | 0.74814 | 0.67454 | 0.78401
| Rel.Std.Dev. |  | 0.7814% | 0.4931% | 0.9833% | 4.556% | 2.021% | 2.48% | 1.579% | 3.154% | 3.951% | 2.188% | 1.238% | 1.939% | 1.608% | 0.5509% | 4.068% | 8.384% | 1.987% | 2.853% | 2.26% | 2.499%
| PDM3 |  | -3.17% | 59.69% | 3.826% | 1.912% | -0.9206% | -1.657% | 8.653% | 0.09909% | -5.962% | 1.974% | -10.41% | 5.871% | -0.4674% | 81.37% | -17.86% | 82.52% | 20.6% | -10.86% | 1.459% | 6.614%
| Z-Score (Absolute) |  | 0.67 | 8 | 0.3 | 0.03 | 0.36 | 0.46 | 0.96 | 0.22 | 1.05 | 0.04 | 1.67 | 0.58 | 0.3 | 10.98 | 2.69 | 11.14 | 2.61 | 1.73 | 0.03 | 0.68
| Analytical results for Co in OREAS 506b (Certified Value 7.1 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 7.4 | 7.2 | 6.8835 | 6.8 | 9.5 | 6.6 | 6.7 | 7.4 | 6.8 | 7 | 6.7 | 7 | 7.6 | 6.8262 | 8 | 7.19 | 8.2701 | 7 | 7.71 | 7.9 | 6.8 | 6.95 | 7.3 | 7.2 | 7.1 |  | 1
| 1 | 2 | 7.4 | 7.1 | 6.5974 | 7.1 | 9.7 | 6.6 | 6.7 | 7.5 | 6.8 | 7.2 | 7 | 7 | 7.4 | 6.9658 | 8 | 7.14 | 8.4122 | 7 | 7.15 | 7.8 | 7.2 | 6.95 | 7 | 7.5 | 7.3 |  | 20
| 1 | 3 | 7.3 | 7.1 | 6.869 | 6.5 | 10.2 | 6.5 | 6.7 | 7.3 | 6.7 | 7 | 6.8 | 6.6 | 7.4 | 7.1733 | 8 | 7 | 8.3798 | 7 | 7.43 | 7.8 | 7.2 | 6.99 | 7.2 | 7.9 | 7.2 |  | 16
| 1 | 4 | 7.4 | 7.1 | 6.718 | 6.5 | 9.5 | 6.7 | 6.7 | 7.6 | 7 | 7.1 | 6.8 | 6 | 7.3 | 7.0923 | 7 | 7.27 | 8.069 | 7 | 7.28 | 7.9 | 7.3 | 6.9 | 7.2 | 7.6 | 7.2 |  | 7.0951
| 1 | 5 | 7.5 | 7.1 | 6.5994 | 7.2 | 9.7 | 6.6 | 6.6 | 7.4 | 6.8 | 7.2 | 6.8 | 6.3 | 7.4 | 6.9357 | 8 | 7.04 | 8.1224 | 7 | 7.36 | 8.3 | 7 | 7.1 | 7.2 | 7.4 | 7.2 |  | 82
| 1 | 6 | 7.6 | 7 | 6.9603 | 7 | 9.5 | 6.4 | 6.8 | 7.3 | 6.6 | 6.9 | 6.8 | 7 | 7.5 | 7.0562 | 8 | 7.13 | 8.4284 | 7 | 7.45 | 8 | 6.9 | 7 | 6.9 | 7.4 | 7.1
| Mean |  | 7.4333 | 7.1 | 6.7713 | 6.85 | 9.6833 | 6.5667 | 6.7 | 7.4167 | 6.7833 | 7.0667 | 6.8167 | 6.65 | 7.4333 | 7.0082 | 7.8333 | 7.1283 | 8.2803 | 7 | 7.3967 | 7.95 | 7.0667 | 6.9817 | 7.1333 | 7.5 | 7.1833
| Median |  | 7.4 | 7.1 | 6.7935 | 6.9 | 9.6 | 6.6 | 6.7 | 7.4 | 6.8 | 7.05 | 6.8 | 6.8 | 7.4 | 7.011 | 8 | 7.135 | 8.325 | 7 | 7.395 | 7.9 | 7.1 | 6.97 | 7.2 | 7.45 | 7.2
| Std Dev. |  | 0.10328 | 0.063246 | 0.15525 | 0.30166 | 0.27142 | 0.10328 | 0.063246 | 0.1169 | 0.13292 | 0.12111 | 0.098319 | 0.42778 | 0.10328 | 0.12397 | 0.40825 | 0.098268 | 0.15421 | 0 | 0.18864 | 0.18708 | 0.19664 | 0.067946 | 0.15055 | 0.23664 | 0.075277
| Rel.Std.Dev. |  | 1.389% | 0.8908% | 2.293% | 4.404% | 2.803% | 1.573% | 0.944% | 1.576% | 1.959% | 1.714% | 1.442% | 6.433% | 1.389% | 1.769% | 5.212% | 1.379% | 1.862% | 0% | 2.55% | 2.353% | 2.783% | 0.9732% | 2.111% | 3.155% | 1.048%
| PDM3 |  | 4.768% | 0.06969% | -4.564% | -3.454% | 36.48% | -7.447% | -5.568% | 4.533% | -4.394% | -0.4001% | -3.924% | -6.273% | 4.768% | -1.224% | 10.41% | 0.469% | 16.71% | -1.34% | 4.251% | 12.05% | -0.4001% | -1.598% | 0.5395% | 5.707% | 1.244%
| Z-Score (Absolute) |  | 0.71 | 0 | 0.7 | 0.53 | 5.5 | 1.14 | 0.85 | 0.67 | 0.67 | 0.07 | 0.6 | 0.96 | 0.71 | 0.2 | NA | 0.06 | 2.51 | NA | 0.63 | 1.81 | 0.07 | 0.25 | 0.07 | 0.85 | 0.18
Results from laboratories 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 506b (Certified Value 45.5 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 2
| 1 | 1 | 45 | 44 | 43.8418 | 44.9 | 46 | 45 | 48 | 45 | 45 | 50 | 41 | 46 | 45.2096 | 45 | 47.7 | 49 | 43 | 47 | 42 | 45.5 | 46 | 41 | 45.7 |  |  |  | 1
| 1 | 2 | 45 | 44 | 42.4865 | 45 | 46 | 46 | 48 | 45 | 46 | 51 | 43 | 47 | 46.875 | 44 | 46 | 49 | 44 | 50 | 44 | 45.5 | 44 | 41 | 46.6 |  |  |  | 21
| 1 | 3 | 44 | 44 | 43.6645 | 45.4 | 46 | 46 | 48 | 44 | 45 | 49 | 40.2 | 46 | 45.8291 | 46 | 46.3 | 49 | 42 | 47 | 42 | 45.8 | 43 | 42 | 45.5 |  |  |  | 16
| 1 | 4 | 44 | 45 | 42.0567 | 44.8 | 47 | 45 | 49 | 44 | 45 | 50 | 39 | 46 | 46.5555 | 45 | 48 | 49 | 46 | 45 | 42 | 46.8 | 43 | 42 | 46.4 |  |  |  | 45.4673
| 1 | 5 | 45 | 45 | 42.8971 | 45.3 | 48 | 46 | 48 | 44 | 46 | 49 | 39.4 | 45 | 45.7753 | 46 | 45.7 | 49 | 44 | 47 | 43 | 45.1 | 43 | 41 | 46 |  |  |  | 83
| 1 | 6 | 44 | 43 | 43.5484 | 44.3 | 46 | 47 | 48 | 44 | 46 | 49 | 39 | 47 | 46.6476 | 45 | 47.3 | 49 | 44 | 48 | 41 | 45.1 | 44 | 41 | 45.6
| Mean |  | 44.5 | 44.1667 | 43.0825 | 44.95 | 46.5 | 45.8333 | 48.1667 | 44.3333 | 45.5 | 49.6667 | 40.2667 | 46.1667 | 46.1487 | 45.1667 | 46.8333 | 49 | 43.8333 | 47.3333 | 42.3333 | 45.6333 | 43.8333 | 41.3333 | 45.9667
| Median |  | 44.5 | 44 | 43.2228 | 44.95 | 46 | 46 | 48 | 44 | 45.5 | 49.5 | 39.8 | 46 | 46.1923 | 45 | 46.8 | 49 | 44 | 47 | 42 | 45.5 | 43.5 | 41 | 45.85
| Std Dev. |  | 0.54772 | 0.75277 | 0.71754 | 0.3937 | 0.83666 | 0.75277 | 0.40825 | 0.5164 | 0.54772 | 0.8165 | 1.5475 | 0.75277 | 0.6427 | 0.75277 | 0.95847 | 0 | 1.3292 | 1.633 | 1.0328 | 0.6314 | 1.169 | 0.5164 | 0.45019
| Rel.Std.Dev. |  | 1.231% | 1.704% | 1.665% | 0.8759% | 1.799% | 1.642% | 0.8476% | 1.165% | 1.204% | 1.644% | 3.843% | 1.631% | 1.393% | 1.667% | 2.047% | 0% | 3.032% | 3.45% | 2.44% | 1.384% | 2.667% | 1.249% | 0.9794%
| PDM3 |  | -2.128% | -2.861% | -5.245% | -1.138% | 2.271% | 0.805% | 5.937% | -2.494% | 0.07186% | 9.236% | -11.44% | 1.538% | 1.499% | -0.6613% | 3.004% | 7.77% | -3.594% | 4.104% | -6.893% | 0.3651% | -3.594% | -9.092% | 1.098%
| Z-Score (Absolute) |  | 0.51 | 0.67 | 1.22 | 0.28 | 0.51 | 0.17 | 1.35 | 0.59 | 0 | 2.11 | 2.65 | 0.34 | 0.33 | 0.17 | 0.67 | 1.77 | 0.84 | 0.93 | 1.6 | 0.07 | 0.84 | 2.11 | 0.24
| Analytical results for Cs in OREAS 506b (Certified Value 6.78 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.5g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [9 empty] | 3
| 1 | 1 | 6.78 | 7.53 | 6.931 | 6.1 | 6.4 | 6.67 | 6.21 | 6 | 7.2 | 6.5655 | 6.92 | 5.7987 | 7.2 | 8.26 | 7.16 | 7.17 | 6.11 | [9 empty] | 1
| 1 | 2 | 6.88 | 7.43 | 6.5578 | 6.13 | 6.56 | 6.89 | 6.24 | 6.23 | 7.28 | 6.6675 | 6.68 | 5.5259 | 7.3 | 8.29 | 7.06 | 7.29 | 6.13 | [9 empty] | 59
| 1 | 3 | 6.82 | 7.57 | 6.9143 | 6.05 | 6.37 | 6.74 | 6.16 | 6.21 | 7.08 | 6.6586 | 6.76 | 5.7984 | 7.4 | 8.2 | 7.24 | 7.27 | 6.34 | [9 empty] | 16
| 1 | 4 | 6.73 | 7.47 | 6.5261 | 6.16 | 6.42 | 6.84 | 6.26 | 6.11 | 7.03 | 6.7955 | 6.97 | 5.6695 | 7.1 | 7.75 | 7.08 | 7.11 | 6.31 | [9 empty] | 6.7809
| 1 | 5 | 6.79 | 7.41 | 6.6731 | 5.92 | 6.31 | 6.76 | 6.14 | 6.27 | 7.23 | 6.8468 | 6.67 | 5.8603 | 7.3 | 8.25 | 7.02 | 7.27 | 6.15 | [9 empty] | 84
| 1 | 6 | 6.87 | 7.32 | 6.8965 | 6.04 | 6.49 | 6.8 | 6 | 6.46 | 7.36 | 6.9374 | 6.79 | 5.775 | 7.2 | 8.47 | 7.28 | 7.14 | 6.1
| Mean |  | 6.8117 | 7.455 | 6.7498 | 6.0667 | 6.425 | 6.7833 | 6.1683 | 6.2133 | 7.1967 | 6.7452 | 6.7983 | 5.738 | 7.25 | 8.2033 | 7.14 | 7.2083 | 6.19
| Median |  | 6.805 | 7.45 | 6.7848 | 6.075 | 6.41 | 6.78 | 6.185 | 6.22 | 7.215 | 6.7315 | 6.775 | 5.7867 | 7.25 | 8.255 | 7.12 | 7.22 | 6.14
| Std Dev. |  | 0.057067 | 0.089387 | 0.18666 | 0.085245 | 0.088713 | 0.077632 | 0.094322 | 0.15526 | 0.1234 | 0.1383 | 0.12352 | 0.12109 | 0.10488 | 0.24064 | 0.10431 | 0.077567 | 0.1064
| Rel.Std.Dev. |  | 0.8378% | 1.199% | 2.765% | 1.405% | 1.381% | 1.144% | 1.529% | 2.499% | 1.715% | 2.05% | 1.817% | 2.11% | 1.447% | 2.933% | 1.461% | 1.076% | 1.719%
| PDM3 |  | 0.4531% | 9.94% | -0.4591% | -10.53% | -5.249% | 0.03521% | -9.034% | -8.371% | 6.131% | -0.5267% | 0.2564% | -15.38% | 6.917% | 20.98% | 5.295% | 6.303% | -8.715%
| Z-Score (Absolute) |  | 0.04 | 1.07 | 0.05 | 1.14 | 0.57 | 0 | 0.98 | 0.9 | 0.66 | 0.06 | 0.02 | 1.66 | 0.74 | 2.25 | 0.57 | 0.67 | 0.94
| Analytical results for Cu in OREAS 506b (Certified Value 0.443 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 3
| 1 | 1 | 0.45727 | 0.423 | 0.42819 | 0.44421 | 0.4045 | 0.442 | 0.45 | 0.459 | 0.452 | 0.429 | 0.4136 | 0.486 | 0.457 | 0.44509 | 0.4538 | 0.4564 | 0.43584 | 0.436 | 0.463 | 0.4755 | 0.4373 | 0.44472 | 0.4098 | 0.437 | 0.4735 | 0.43054 | 1
| 1 | 2 | 0.45723 | 0.424 | 0.41656 | 0.4408 | 0.4034 | 0.447 | 0.455 | 0.459 | 0.443 | 0.43 | 0.4271 | 0.517 | 0.45015 | 0.45011 | 0.4426 | 0.4531 | 0.43068 | 0.4356 | 0.445 | 0.4871 | 0.4385 | 0.44297 | 0.4025 | 0.44 | 0.4817 | 0.43371 | 22
| 1 | 3 | 0.46087 | 0.423 | 0.40823 | 0.44469 | 0.4027 | 0.445 | 0.459 | 0.463 | 0.441 | 0.427 | 0.4222 | 0.469 | 0.45634 | 0.44274 | 0.4541 | 0.4507 | 0.43047 | 0.4378 | 0.462 | 0.4796 | 0.4385 | 0.43725 | 0.4075 | 0.446 | 0.4845 | 0.4369 | 16
| 1 | 4 | 0.45407 | 0.422 | 0.40895 | 0.43717 | 0.4079 | 0.45 | 0.437 | 0.465 | 0.439 | 0.416 | 0.415 | 0.469 | 0.44013 | 0.44744 | 0.4484 | 0.4548 | 0.433 | 0.437 | 0.459 | 0.4621 | 0.4388 | 0.43866 | 0.4084 | 0.441 | 0.4801 | 0.43638 | 0.44268
| 1 | 5 | 0.45799 | 0.421 | 0.41863 | 0.44665 | 0.4056 | 0.445 | 0.442 | 0.455 | 0.435 | 0.429 | 0.4184 | 0.471 | 0.44782 | 0.44531 | 0.4599 | 0.4532 | 0.43543 | 0.4334 | 0.459 | 0.4783 | 0.4382 | 0.44029 | 0.4037 | 0.432 | 0.4921 | 0.4352 | 85
| 1 | 6 | 0.46204 | 0.424 | 0.42847 | 0.4318 | 0.4031 | 0.444 | 0.459 | 0.459 | 0.444 | 0.442 | 0.4153 | 0.462 | 0.44159 | 0.45474 | 0.4478 | 0.4589 | 0.43042 | 0.441 | 0.465 | 0.4817 | 0.4406 | 0.44378 | 0.4026 | 0.436 | 0.4809 | 0.43076
| Mean |  | 0.45825 | 0.42283 | 0.41817 | 0.44089 | 0.40453 | 0.4455 | 0.45033 | 0.46 | 0.44233 | 0.42883 | 0.4186 | 0.479 | 0.44884 | 0.44757 | 0.4511 | 0.45452 | 0.43264 | 0.4368 | 0.45883 | 0.47738 | 0.43865 | 0.44128 | 0.40575 | 0.43867 | 0.48213 | 0.43392
| Median |  | 0.45763 | 0.423 | 0.4176 | 0.4425 | 0.40395 | 0.445 | 0.4525 | 0.459 | 0.442 | 0.429 | 0.41685 | 0.47 | 0.44899 | 0.44637 | 0.4511 | 0.454 | 0.43184 | 0.4365 | 0.4605 | 0.47895 | 0.4385 | 0.44163 | 0.4056 | 0.4385 | 0.4813 | 0.43446
| Std Dev. |  | 0.0028555 | 0.001169 | 0.0088649 | 0.0055703 | 0.0019582 | 0.0027386 | 0.0091579 | 0.0035214 | 0.0057155 | 0.0082805 | 0.0051788 | 0.020229 | 0.0071277 | 0.0042972 | 0.0060689 | 0.0028687 | 0.0025138 | 0.0025424 | 0.0071671 | 0.0084372 | 0.0010858 | 0.0030008 | 0.0031992 | 0.0048028 | 0.0060879 | 0.0027567
| Rel.Std.Dev. |  | 0.6231% | 0.2765% | 2.12% | 1.263% | 0.4841% | 0.6147% | 2.034% | 0.7655% | 1.292% | 1.931% | 1.237% | 4.223% | 1.588% | 0.9601% | 1.345% | 0.6312% | 0.581% | 0.5821% | 1.562% | 1.767% | 0.2475% | 0.68% | 0.7885% | 1.095% | 1.263% | 0.6353%
| PDM3 |  | 3.515% | -4.484% | -5.537% | -0.406% | -8.618% | 0.6364% | 1.728% | 3.912% | -0.07893% | -3.129% | -5.44% | 8.204% | 1.391% | 1.104% | 1.901% | 2.673% | -2.268% | -1.329% | 3.648% | 7.839% | -0.911% | -0.3172% | -8.343% | -0.9072% | 8.912% | -1.98%
| Z-Score (Absolute) |  | 1.01 | 1.16 | 1.45 | 0.06 | 2.28 | 0.23 | 0.52 | 1.12 | 0.03 | 0.8 | 1.42 | 2.28 | 0.43 | 0.35 | 0.57 | 0.78 | 0.56 | 0.31 | 1.04 | 2.18 | 0.19 | 0.03 | 2.21 | 0.19 | 2.47 | 0.48
| Analytical results for Dy in OREAS 506b (Indicative Value 2.27 ppm) | [27 empty] | Indicative
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.04: 2.25 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.04: 0.5g | 2
| 1 | 1 | 2.01: 2.247; 2.02: 2.1292; 2.03: 3.1469; 2.04: 2.3 | 1
| 1 | 2 | 2.01: 2.216; 2.02: 2.295; 2.03: 2.7972; 2.04: 2.2 | 23
| 1 | 3 | 2.01: 2.246; 2.02: 2.1895; 2.03: 3.2295; 2.04: 2.4 | 16
| 1 | 4 | 2.01: 2.257; 2.02: 2.3666; 2.03: 3.0188; 2.04: 2.3 | 2.2717
| 1 | 5 | 2.01: 2.296; 2.02: 2.2196; 2.03: 3.0268; 2.04: 2.2 | 39
| 1 | 6 | 2.01: 2.29; 2.02: 2.3378; 2.03: 2.9241; 2.04: 2.4
| Mean |  | 2.01: 2.2587; 2.02: 2.2563; 2.03: 3.0239; 2.04: 2.3
| Median |  | 2.01: 2.252; 2.02: 2.2573; 2.03: 3.0228; 2.04: 2.3
| Std Dev. |  | 2.01: 0.029984; 2.02: 0.091943; 2.03: 0.15413; 2.04: 0.089443
| Rel.Std.Dev. |  | 2.01: 1.328%; 2.02: 4.075%; 2.03: 5.097%; 2.04: 3.889%
| PDM3 |  | 2.01: -0.5716%; 2.02: -0.6763%; 2.03: 33.11%; 2.04: 1.248%
| Z-Score (Absolute) |  | 2.01: 0.64; 2.02: 0.71; 2.03: 22.97; 2.04: 0.64
| Analytical results for Er in OREAS 506b (Indicative Value 0.91 ppm) | [27 empty] | Indicative
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.04: 2.25 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.04: 0.5g | 2
| 1 | 1 | 2.01: 0.927; 2.02: 0.91066; 2.03: 1.2693; 2.04: 0.9 | 1
| 1 | 2 | 2.01: 0.925; 2.02: 0.86529; 2.03: 1.1365; 2.04: 0.9 | 24
| 1 | 3 | 2.01: 0.898; 2.02: 0.88221; 2.03: 1.2319; 2.04: 0.9 | 16
| 1 | 4 | 2.01: 0.907; 2.02: 0.9153; 2.03: 1.2117; 2.04: 0.9 | 0.90706
| 1 | 5 | 2.01: 0.957; 2.02: 0.90399; 2.03: 1.2242; 2.04: 0.8 | 40
| 1 | 6 | 2.01: 0.939; 2.02: 0.89671; 2.03: 1.2111; 2.04: 0.9
| Mean |  | 2.01: 0.9255; 2.02: 0.89569; 2.03: 1.2141; 2.04: 0.88333
| Median |  | 2.01: 0.926; 2.02: 0.90035; 2.03: 1.218; 2.04: 0.9
| Std Dev. |  | 2.01: 0.021333; 2.02: 0.018893; 2.03: 0.043607; 2.04: 0.040825
| Rel.Std.Dev. |  | 2.01: 2.305%; 2.02: 2.109%; 2.03: 3.592%; 2.04: 4.622%
| PDM3 |  | 2.01: 2.032%; 2.02: -1.254%; 2.03: 33.85%; 2.04: -2.616%
| Z-Score (Absolute) |  | 2.01: 0.48; 2.02: 0.48; 2.03: 9.71; 2.04: 0.87
| Analytical results for Eu in OREAS 506b (Indicative Value 0.43 ppm) | [27 empty] | Indicative
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.04: 2.25 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.04: 0.5g | 2
| 1 | 1 | 2.01: 0.356; 2.02: 0.31977; 2.03: 0.6452; 2.04: 0.4 | 1
| 1 | 2 | 2.01: 0.363; 2.02: 0.34974; 2.03: 0.55482; 2.04: 0.4 | 25
| 1 | 3 | 2.01: 0.361; 2.02: 0.3238; 2.03: 0.65331; 2.04: 0.4 | 16
| 1 | 4 | 2.01: 0.359; 2.02: 0.35131; 2.03: 0.59711; 2.04: 0.4 | 0.42664
| 1 | 5 | 2.01: 0.368; 2.02: 0.34181; 2.03: 0.59787; 2.04: 0.4 | 41
| 1 | 6 | 2.01: 0.368; 2.02: 0.34025; 2.03: 0.58926; 2.04: 0.4
| Mean |  | 2.01: 0.3625; 2.02: 0.33778; 2.03: 0.60626; 2.04: 0.4
| Median |  | 2.01: 0.362; 2.02: 0.34103; 2.03: 0.59749; 2.04: 0.4
| Std Dev. |  | 2.01: 0.0048477; 2.02: 0.013178; 2.03: 0.036931; 2.04: 6.0809e-17
| Rel.Std.Dev. |  | 2.01: 1.337%; 2.02: 3.901%; 2.03: 6.092%; 2.04: 1.52e-14%
| PDM3 |  | 2.01: -15.03%; 2.02: -20.83%; 2.03: 42.1%; 2.04: -6.243%
| Z-Score (Absolute) |  | 2.01: 0.41; 2.02: 0.94; 2.03: 4.88; 2.04: 0.41
| Analytical results for Fe in OREAS 506b (Certified Value 2.66 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 3
| 1 | 1 | 2.69 | 2.73 | 2.6254 | 2.57 | 2.61 | 2.61 | 2.59 | 2.75 | 2.7 | 2.54 | 2.67 | 2.81 | 2.74 | 2.5985 | 2.89 | 2.66 | 2.66 | 2.58 | 2.96 | 2.69 | 2.784 | 2.47 | 2.69 | 2.5832 | 2.55 |  | 1
| 1 | 2 | 2.69 | 2.69 | 2.4967 | 2.58 | 2.63 | 2.65 | 2.68 | 2.72 | 2.69 | 2.55 | 2.8 | 3 | 2.71 | 2.6356 | 2.8 | 2.59 | 2.69 | 2.67 | 3.08 | 2.73 | 2.791 | 2.43 | 2.72 | 2.6428 | 2.57 |  | 26
| 1 | 3 | 2.64 | 2.74 | 2.4903 | 2.62 | 2.64 | 2.65 | 2.64 | 2.74 | 2.64 | 2.54 | 2.7 | 2.81 | 2.73 | 2.6314 | 2.87 | 2.59 | 2.67 | 2.66 | 3 | 2.63 | 2.7472 | 2.49 | 2.78 | 2.6418 | 2.56 |  | 16
| 1 | 4 | 2.64 | 2.67 | 2.5147 | 2.55 | 2.62 | 2.69 | 2.61 | 2.76 | 2.65 | 2.5 | 2.72 | 2.72 | 2.7 | 2.6423 | 2.85 | 2.63 | 2.67 | 2.59 | 2.83 | 2.79 | 2.8138 | 2.43 | 2.75 | 2.6267 | 2.57 |  | 2.6578
| 1 | 5 | 2.68 | 2.78 | 2.6017 | 2.61 | 2.64 | 2.67 | 2.64 | 2.71 | 2.64 | 2.53 | 2.68 | 2.71 | 2.73 | 2.6407 | 2.92 | 2.6 | 2.67 | 2.73 | 3.01 | 2.74 | 2.7849 | 2.48 | 2.68 | 2.6482 | 2.54 |  | 86
| 1 | 6 | 2.68 | 2.65 | 2.6456 | 2.54 | 2.6 | 2.64 | 2.66 | 2.72 | 2.64 | 2.5 | 2.69 | 2.74 | 2.71 | 2.6417 | 2.85 | 2.59 | 2.69 | 2.63 | 3.05 | 2.62 | 2.8029 | 2.41 | 2.69 | 2.6309 | 2.54
| Mean |  | 2.67 | 2.71 | 2.5624 | 2.5783 | 2.6233 | 2.6517 | 2.6367 | 2.7333 | 2.66 | 2.5267 | 2.71 | 2.7983 | 2.72 | 2.6317 | 2.8633 | 2.61 | 2.675 | 2.6433 | 2.9883 | 2.7 | 2.7873 | 2.4517 | 2.7183 | 2.6289 | 2.555
| Median |  | 2.68 | 2.71 | 2.5582 | 2.575 | 2.625 | 2.65 | 2.64 | 2.73 | 2.645 | 2.535 | 2.695 | 2.775 | 2.72 | 2.6382 | 2.86 | 2.595 | 2.67 | 2.645 | 3.005 | 2.71 | 2.7879 | 2.45 | 2.705 | 2.6364 | 2.555
| Std Dev. |  | 0.023664 | 0.04858 | 0.069614 | 0.031885 | 0.01633 | 0.027142 | 0.03266 | 0.019664 | 0.027568 | 0.021602 | 0.047329 | 0.10797 | 0.015492 | 0.016801 | 0.040825 | 0.028983 | 0.012247 | 0.055737 | 0.087958 | 0.066332 | 0.022733 | 0.032506 | 0.039707 | 0.023788 | 0.013784
| Rel.Std.Dev. |  | 0.8863% | 1.793% | 2.717% | 1.237% | 0.6225% | 1.024% | 1.239% | 0.7194% | 1.036% | 0.855% | 1.746% | 3.858% | 0.5696% | 0.6384% | 1.426% | 1.11% | 0.4578% | 2.109% | 2.943% | 2.457% | 0.8156% | 1.326% | 1.461% | 0.9049% | 0.5395%
| PDM3 |  | 0.4583% | 1.963% | -3.59% | -2.991% | -1.297% | -0.2315% | -0.7958% | 2.841% | 0.08208% | -4.935% | 1.963% | 5.287% | 2.34% | -0.9827% | 7.732% | -1.799% | 0.6465% | -0.545% | 12.44% | 1.587% | 4.872% | -7.756% | 2.277% | -1.087% | -3.869%
| Z-Score (Absolute) |  | 0.13 | 0.67 | 1.32 | 1.1 | 0.49 | 0.11 | 0.31 | 0.99 | 0 | 1.8 | 0.67 | 1.87 | 0.81 | 0.38 | 2.74 | 0.67 | 0.2 | 0.22 | 4.43 | 0.54 | 1.72 | 2.81 | 0.79 | 0.42 | 1.42
| Analytical results for Ga in OREAS 506b (Certified Value 8.03 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 8.11 | 8.6 | 8.0258 | 7.4 | 9.3 | 7.71 | 7.95 | 8.17 | 8.15 | 8.14 | 8.1 | < 10 | 8.1 | 7.8608 | 7.95 | 8 | 8.83 | 22.8 | 8.17 | 8.4 | 8.6 | 7.78 | 7.2 |  |  |  | 1
| 1 | 2 | 8.31 | 8.4 | 7.6817 | 7.9 | 8.4 | 7.8 | 8.02 | 8.23 | 8.18 | 8.23 | 8.3 | < 10 | 7.8 | 8.4606 | 8.07 | 9 | 7.57 | 22.5 | 8.09 | 7.9 | 8.7 | 7.87 | 7.4 |  |  |  | 27
| 1 | 3 | 7.92 | 8.3 | 7.9167 | 7.2 | 9.1 | 7.78 | 8.05 | 8.07 | 7.96 | 8.18 | 8.2 | < 10 | 8 | 8.0566 | 7.83 | 8 | 8.95 | 22.2 | 8.06 | 8.2 | 8.7 | 8.14 | 7.5 |  |  |  | 16
| 1 | 4 | 8.06 | 8.5 | 7.802 | 7 | 8.5 | 8.11 | 8.08 | 8.3 | 8.23 | 8.11 | 8.1 | < 10 | 7.8 | 8.428 | 8.02 | 9 | 8.99 | 22.5 | 8.01 | 8 | 8.8 | 7.86 | 7.5 |  |  |  | 8.027
| 1 | 5 | 8.06 | 8.5 | 7.7854 | 8 | 9.5 | 7.66 | 7.9 | 8.17 | 8 | 8.21 | 8.2 | < 10 | 8.1 | 7.9695 | 7.82 | 8 | 8.94 | 23.4 | 7.81 | 7.7 | 8.8 | 7.7 | 7.4 |  |  |  | 87
| 1 | 6 | 8.15 | 8.4 | 8.0174 | 7.8 | 8 | 7.65 | 8 | 8.08 | 7.77 | 7.72 | 8.2 | < 10 | 8 | 8.1023 | 7.75 | 9 | 8.95 | 23.2 | 8.04 | 8.2 | 8.6 | 7.58 | 7.3
| Mean |  | 8.1017 | 8.45 | 7.8715 | 7.55 | 8.8 | 7.785 | 8 | 8.17 | 8.0483 | 8.0983 | 8.1833 |  | 7.9667 | 8.1463 | 7.9067 | 8.5 | 8.705 | 22.7667 | 8.03 | 8.0667 | 8.7 | 7.8217 | 7.3833
| Median |  | 8.085 | 8.45 | 7.8594 | 7.6 | 8.8 | 7.745 | 8.01 | 8.17 | 8.075 | 8.16 | 8.2 |  | 8 | 8.0795 | 7.89 | 8.5 | 8.945 | 22.65 | 8.05 | 8.1 | 8.7 | 7.82 | 7.4
| Std Dev. |  | 0.12828 | 0.10488 | 0.13811 | 0.40866 | 0.58652 | 0.1705 | 0.06603 | 0.087864 | 0.17221 | 0.19052 | 0.075277 |  | 0.13663 | 0.2453 | 0.12596 | 0.54772 | 0.55863 | 0.45898 | 0.12083 | 0.25033 | 0.089443 | 0.18978 | 0.1169
| Rel.Std.Dev. |  | 1.583% | 1.241% | 1.755% | 5.413% | 6.665% | 2.19% | 0.8254% | 1.075% | 2.14% | 2.353% | 0.9199% |  | 1.715% | 3.011% | 1.593% | 6.444% | 6.417% | 2.016% | 1.505% | 3.103% | 1.028% | 2.426% | 1.583%
| PDM3 |  | 0.9303% | 5.27% | -1.937% | -5.942% | 9.63% | -3.015% | -0.3362% | 1.782% | 0.2659% | 0.8888% | 1.948% |  | -0.7515% | 1.486% | -1.499% | 5.893% | 8.447% | 183.6% | 0.03752% | 0.4943% | 8.384% | -2.558% | -8.019%
| Z-Score (Absolute) |  | 0.18 | 1.57 | 0.74 | 2.03 | 2.97 | 1.09 | 0.23 | 0.45 | 0.04 | 0.16 | 0.5 | 12.24 | 0.36 | 0.36 | 0.6 | NA | 2.59 | 58.89 | 0.11 | 0.04 | 2.57 | 0.94 | 2.7
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 506b (Indicative Value 3.13 ppm) | [27 empty] | Indicative
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.04: 2.25 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS |  |  |  | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.04: 0.5g |  |  |  | 2
| 1 | 1 | 2.01: 3.065; 2.02: 3.0936; 2.03: 5.2714; 2.04: 3 |  |  |  | 1
| 1 | 2 | 2.01: 3.126; 2.02: 3.1788; 2.03: 5.1338; 2.04: 3.1 |  |  |  | 28
| 1 | 3 | 2.01: 3.043; 2.02: 3.1427; 2.03: 5.384; 2.04: 3.3 |  |  |  | 16
| 1 | 4 | 2.01: 3.053; 2.02: 3.2659; 2.03: 5.1101; 2.04: 3.1 |  |  |  | 3.1291
| 1 | 5 | 2.01: 3.113; 2.02: 3.2177; 2.03: 5.2088; 2.04: 3.1 |  |  |  | 42
| 1 | 6 | 2.01: 3.059; 2.02: 3.2668; 2.03: 4.977; 2.04: 3.1
| Mean |  | 2.01: 3.0765; 2.02: 3.1942; 2.03: 5.1809; 2.04: 3.1167
| Median |  | 2.01: 3.062; 2.02: 3.1982; 2.03: 5.1713; 2.04: 3.1
| Std Dev. |  | 2.01: 0.034338; 2.02: 0.069233; 2.03: 0.14072; 2.04: 0.098319
| Rel.Std.Dev. |  | 2.01: 1.116%; 2.02: 2.167%; 2.03: 2.716%; 2.04: 3.155%
| PDM3 |  | 2.01: -1.682%; 2.02: 2.081%; 2.03: 65.57%; 2.04: -0.3985%
| Z-Score (Absolute) |  | 2.01: 0.9; 2.02: 0.44; 2.03: 23.2; 2.04: 0.44
| Analytical results for Ge in OREAS 506b (Certified Value 0.11 ppm) | [27 empty] | Certified
Ge | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab |  |  |  | 1
| No | No | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [13 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [13 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [13 empty] | 2
| 1 | 1 | < 5 | 0.09 | 0.11 | 0.11 | 0.09 | 0.1 | 0.1 | 0.26659 | 0.15 | < 0.1 | 0.13 | < 0.1 | < 0.1 | [13 empty] | 1
| 1 | 2 | < 5 | 0.07 | 0.1 | 0.12 | 0.09 | 0.1 | 0.1 | 0.22292 | 0.1 | < 0.1 | 0.14 | < 0.1 | < 0.1 | [13 empty] | 0
| 1 | 3 | < 5 | 0.07 | 0.11 | 0.11 | 0.09 | 0.1 | 0.1 | 0.23903 | 0.17 | < 0.1 | 0.13 | < 0.1 | < 0.1 | [13 empty] | 16
| 1 | 4 | < 5 | 0.08 | 0.09 | 0.13 | 0.09 | 0.08 | 0.1 | 0.23432 | 0.16 | < 0.1 | 0.14 | < 0.1 | < 0.1 | [13 empty] | 0.11133
| 1 | 5 | < 5 | 0.08 | 0.11 | 0.1 | 0.08 | 0.11 | 0.1 | 0.23531 | 0.16 | < 0.1 | 0.14 | < 0.1 | < 0.1 | [13 empty] | 88
| 1 | 6 | < 5 | 0.07 | 0.11 | 0.14 | 0.08 | 0.14 | 0.1 | 0.22843 | 0.15 | < 0.1 | 0.14 | < 0.1 | < 0.1
| Mean |  |  | 0.076667 | 0.105 | 0.11833 | 0.086667 | 0.105 | 0.1 | 0.23777 | 0.14833 |  | 0.13667
| Median |  |  | 0.075 | 0.11 | 0.115 | 0.09 | 0.1 | 0.1 | 0.23482 | 0.155 |  | 0.14
| Std Dev. |  |  | 0.008165 | 0.0083666 | 0.01472 | 0.005164 | 0.019748 | 1.5202e-17 | 0.015217 | 0.024833 |  | 0.005164
| Rel.Std.Dev. |  |  | 10.65% | 7.968% | 12.44% | 5.958% | 18.81% | 1.52e-14% | 6.4% | 16.74% |  | 3.779%
| PDM3 |  |  | -31.14% | -5.689% | 6.287% | -22.16% | -5.689% | -10.18% | 113.6% | 33.23% |  | 22.75%
| Z-Score (Absolute) |  | 43.07 | 0.51 | 0 | 0.24 | 0.33 | 0 | NA | 2.39 | 0.78 | 0.99 | 0.57 | 0.99 | 0.99
Results from laboratory 2.14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 506b (Certified Value 0.31 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.34 | 0.38 | 0.30105 | 0.29 | 0.3 | 0.31 | 0.31 | 0.31 | 0.29 | 0.31791 | 0.36 | 0.44 | 0.44 | 0.29 | 0.32 | 0.3 | [10 empty] | 1
| 1 | 2 | 0.36 | 0.38 | 0.2836 | 0.31 | 0.31 | 0.29 | 0.3 | 0.29 | 0.28 | 0.3241 | 0.36 | 0.48 | 0.4 | 0.29 | 0.32 | 0.3 | [10 empty] | 29
| 1 | 3 | 0.35 | 0.38 | 0.2964 | 0.27 | 0.3 | 0.31 | 0.31 | 0.28 | 0.29 | 0.30833 | 0.33 | 0.5 | 0.4 | 0.32 | 0.31 | 0.3 | [10 empty] | 16
| 1 | 4 | 0.35 | 0.38 | 0.30249 | 0.29 | 0.3 | 0.3 | 0.3 | 0.29 | 0.26 | 0.33239 | 0.35 | 0.48 | 0.44 | 0.32 | 0.32 | 0.3 | [10 empty] | 0.30977
| 1 | 5 | 0.35 | 0.39 | 0.29129 | 0.28 | 0.3 | 0.31 | 0.31 | 0.29 | 0.28 | 0.31879 | 0.34 | 0.48 | 0.39 | 0.31 | 0.32 | 0.3 | [10 empty] | 89
| 1 | 6 | 0.36 | 0.38 | 0.28159 | 0.29 | 0.34 | 0.32 | 0.28 | 0.31 | 0.28 | 0.33331 | 0.36 | 0.46 | 0.43 | 0.32 | 0.32 | 0.4
| Mean |  | 0.35167 | 0.38167 | 0.29274 | 0.28833 | 0.30833 | 0.30667 | 0.30167 | 0.295 | 0.28 | 0.32247 | 0.35 | 0.47333 | 0.41667 | 0.30833 | 0.31833 | 0.31667
| Median |  | 0.35 | 0.38 | 0.29385 | 0.29 | 0.3 | 0.31 | 0.305 | 0.29 | 0.28 | 0.32145 | 0.355 | 0.48 | 0.415 | 0.315 | 0.32 | 0.3
| Std Dev. |  | 0.0075277 | 0.0040825 | 0.0088026 | 0.013292 | 0.016021 | 0.010328 | 0.01169 | 0.012247 | 0.010954 | 0.0095145 | 0.012649 | 0.020656 | 0.022509 | 0.01472 | 0.0040825 | 0.040825
| Rel.Std.Dev. |  | 2.141% | 1.07% | 3.007% | 4.61% | 5.196% | 3.368% | 3.875% | 4.152% | 3.912% | 2.95% | 3.614% | 4.364% | 5.402% | 4.774% | 1.282% | 12.89%
| PDM3 |  | 13.53% | 23.21% | -5.498% | -6.919% | -0.4629% | -1.001% | -2.615% | -4.767% | -9.61% | 4.101% | 12.99% | 52.8% | 34.51% | -0.4629% | 2.765% | 2.227%
| Z-Score (Absolute) |  | 1.87 | 3.17 | 0.67 | 0.86 | 0 | 0.07 | 0.29 | 0.58 | 1.23 | 0.61 | 1.8 | 7.13 | 4.68 | 0 | 0.43 | NA
Results from laboratory 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 506b (Indicative Value 0.015 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  |  |  | 3
| 1 | 1 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.03 | 0.01 | 0.01 | < 0.01 | 0.01 | < 1 | 0.02 | < 1 | < 1 | < 2 | < 0.01 | 0.19 | 0.04 | < 0.01 | < 0.1 | < 0.01 | 0.01 | < 0.005 |  |  |  |  | 1
| 1 | 2 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.03 | 0.02 | 0.01 | < 0.01 | 0.02 | < 1 | 0.02 | < 1 | < 1 | < 2 | 0.02 | 0.12 | 0.04 | < 0.01 | < 0.1 | 0.02 | 0.03 | 0.008 |  |  |  |  | 37
| 1 | 3 | < 0.05 | < 0.005 | < 5 | < 0.01 | 0.03 | 0.01 | 0.02 | < 0.01 | 0.02 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.02 | 0.19 | 0.04 | < 0.01 | < 0.1 | 0.02 | 0.02 | < 0.005 |  |  |  |  | 16
| 1 | 4 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.03 | 0.01 | 0.02 | < 0.01 | 0.02 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.01 | 0.15 | 0.04 | < 0.01 | < 0.1 | 0.02 | 0.02 | 0.005 |  |  |  |  | 0.014568
| 1 | 5 | < 0.05 | < 0.005 | < 5 | < 0.01 | 0.02 | 0.01 | 0.02 | < 0.01 | < 0.01 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.01 | 0.12 | 0.04 | < 0.01 | < 0.1 | 0.01 | 0.02 | 0.007 |  |  |  |  | 43
| 1 | 6 | < 0.05 | < 0.005 | < 5 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.02 | < 1 | 0.01 | < 1 | < 1 | < 2 | 0.01 | 0.13 | 0.04 | < 0.01 | < 0.1 | 0.01 | 0.02 | 0.009
| Mean |  | 2.07: 0.01; 2.08: 0.025; 2.09: 0.011667; 2.10: 0.015; 2.11: 0.01; 2.12: 0.018; 2.14: 0.013333; 2.20: 0.014; 2.21: 0.15; 2.22: 0.04; 2.25: 0.016; 2.26: 0.02; 2.27: 0.00725
| Median |  | 2.07: 0.01; 2.08: 0.03; 2.09: 0.01; 2.10: 0.015; 2.11: 0.01; 2.12: 0.02; 2.14: 0.01; 2.20: 0.01; 2.21: 0.14; 2.22: 0.04; 2.25: 0.02; 2.26: 0.02; 2.27: 0.0075
| Std Dev. |  | 2.07: 0; 2.08: 0.0083666; 2.09: 0.0040825; 2.10: 0.0054772; 2.12: 0.0044721; 2.14: 0.005164; 2.20: 0.0054772; 2.21: 0.032863; 2.22: 0; 2.25: 0.0054772; 2.26: 0.0063246; 2.27: 0.0017078
| Rel.Std.Dev. |  | 2.07: 0%; 2.08: 33.47%; 2.09: 34.99%; 2.10: 36.51%; 2.12: 24.85%; 2.14: 38.73%; 2.20: 39.12%; 2.21: 21.91%; 2.22: 0%; 2.25: 34.23%; 2.26: 31.62%; 2.27: 23.56%
| PDM3 |  | 2.07: -31.36%; 2.08: 71.61%; 2.09: -19.92%; 2.10: 2.964%; 2.11: -31.36%; 2.12: 23.56%; 2.14: -8.476%; 2.20: -3.9%; 2.21: 929.6%; 2.22: 174.6%; 2.25: 9.828%; 2.26: 37.29%; 2.27: -50.23%
| Z-Score (Absolute) |  | 0.39 | 0.85 | 137.56 | 0.53 | 0.39 | 0.35 | 0.16 | 0.67 | 0.12 | 26.72 | 0.25 | 26.72 | 26.72 | 54.43 | 0.3 | 7.32 | 1.22 | 0.72 | 1.78 | 0.21 | 0.12 | 0.68
| Analytical results for Ho in OREAS 506b (Indicative Value 0.4 ppm) | [27 empty] | Indicative
Ho | Round | Replicate | 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab |  |  |  |  | 1
| No | No | 2.03: 2.01; 2.04: 2.15; 2.06: 2.18; 2.07: 2.25 |  |  |  |  | ppm
|  |  | 2.03: AR*MS; 2.04: AR*MS; 2.06: AR*MS; 2.07: AR*OES/MS |  |  |  |  | 2
|  |  | 2.03: 01g; 2.04: 0.2g; 2.06: 0.5g; 2.07: 0.5g |  |  |  |  | 2
| 1 | 1 | 2.03: 0.376; 2.04: 0.3468; 2.06: 0.51366; 2.07: 0.4 |  |  |  |  | 1
| 1 | 2 | 2.03: 0.376; 2.04: 0.35328; 2.06: 0.44923; 2.07: 0.3 |  |  |  |  | 12
| 1 | 3 | 2.03: 0.367; 2.04: 0.32941; 2.06: 0.5234; 2.07: 0.4 |  |  |  |  | 16
| 1 | 4 | 2.03: 0.373; 2.04: 0.35261; 2.06: 0.48846; 2.07: 0.4 |  |  |  |  | 0.39791
| 1 | 5 | 2.03: 0.38; 2.04: 0.3407; 2.06: 0.48409; 2.07: 0.4 |  |  |  |  | 39
| 1 | 6 | 2.03: 0.386; 2.04: 0.34051; 2.06: 0.4696; 2.07: 0.4
| Mean |  | 2.03: 0.37633; 2.04: 0.34389; 2.06: 0.48807; 2.07: 0.38333
| Median |  | 2.03: 0.376; 2.04: 0.34375; 2.06: 0.48628; 2.07: 0.4
| Std Dev. |  | 2.03: 0.0064083; 2.04: 0.0089859; 2.06: 0.027455; 2.07: 0.040825
| Rel.Std.Dev. |  | 2.03: 1.703%; 2.04: 2.613%; 2.06: 5.625%; 2.07: 10.65%
| PDM3 |  | 2.03: -5.422%; 2.04: -13.58%; 2.06: 22.66%; 2.07: -3.663%
| Z-Score (Absolute) |  | 2.03: 0.12; 2.04: 1.23; 2.06: 3.7; 2.07: 0.12
| Analytical results for In in OREAS 506b (Certified Value 0.072 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.07 | 0.07 | 0.076072 | < 5 | 0.069 | 0.074 | 0.077 | 0.072 | 0.071 | 0.07 | < 0.2 | 0.1 | 0.07 | 0.048 | 0.08 | 0.07 | 0.07 | [9 empty] | 1
| 1 | 2 | 0.07 | 0.07 | 0.075609 | < 5 | 0.069 | 0.072 | 0.079 | 0.078 | 0.077 | 0.07 | < 0.2 | 0.1 | 0.07 | 0.04 | 0.065 | 0.08 | 0.07 | [9 empty] | 52
| 1 | 3 | 0.07 | 0.07 | 0.072777 | < 5 | 0.074 | 0.072 | 0.079 | 0.071 | 0.069 | 0.08 | < 0.2 | 0.1 | 0.07 | 0.045 | 0.074 | 0.07 | 0.08 | [9 empty] | 16
| 1 | 4 | 0.08 | 0.07 | 0.066855 | < 5 | 0.071 | 0.074 | 0.077 | 0.073 | 0.075 | 0.07 | < 0.2 | 0.1 | 0.07 | 0.045 | 0.077 | 0.07 | 0.08 | [9 empty] | 0.072341
| 1 | 5 | 0.07 | 0.07 | 0.067583 | < 5 | 0.072 | 0.071 | 0.077 | 0.08 | 0.071 | 0.07 | < 0.2 | 0.1 | 0.07 | 0.033 | 0.073 | 0.08 | 0.09 | [9 empty] | 90
| 1 | 6 | 0.07 | 0.07 | 0.069676 | < 5 | 0.07 | 0.071 | 0.073 | 0.071 | 0.067 | 0.08 | < 0.2 | 0.1 | 0.08 | 0.035 | 0.073 | 0.07 | 0.07
| Mean |  | 0.071667 | 0.07 | 0.071429 |  | 0.070833 | 0.072333 | 0.077 | 0.074167 | 0.071667 | 0.073333 |  | 0.1 | 0.071667 | 0.041 | 0.073667 | 0.073333 | 0.076667
| Median |  | 0.07 | 0.07 | 0.071227 |  | 0.0705 | 0.072 | 0.077 | 0.0725 | 0.071 | 0.07 |  | 0.1 | 0.07 | 0.0425 | 0.0735 | 0.07 | 0.075
| Std Dev. |  | 0.0040825 | 0 | 0.0039907 |  | 0.0019408 | 0.0013663 | 0.0021909 | 0.0038687 | 0.0037238 | 0.005164 |  | 1.5202e-17 | 0.0040825 | 0.0060332 | 0.0050465 | 0.005164 | 0.008165
| Rel.Std.Dev. |  | 5.696% | 0% | 5.587% |  | 2.74% | 1.889% | 2.845% | 5.216% | 5.196% | 7.042% |  | 1.52e-14% | 5.696% | 14.72% | 6.85% | 7.042% | 10.65%
| PDM3 |  | -0.9325% | -3.236% | -1.261% |  | -2.084% | -0.01099% | 6.44% | 2.523% | -0.9325% | 1.371% |  | 38.23% | -0.9325% | -43.32% | 1.832% | 1.371% | 5.979%
| Z-Score (Absolute) |  | 0.57 | 1.4 | 0.69 | 1195.55 | 0.99 | 0.25 | 2.05 | 0.66 | 0.57 | 0.25 | 13.38 | NA | 0.57 | 15.68 | 0.41 | 0.25 | 1.89
Results from laboratory 2.17 were removed due to their 0.01 ppm reading resolution.
| Analytical results for K in OREAS 506b (Certified Value 0.853 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.883 | 0.926 | 0.78797 | 0.8 | 0.89 | 0.85 | 0.85 | 0.9 | 0.85 | 0.81 | 0.82 | 0.99 | 0.9 | 0.83886 | 0.92 | 0.87 | 0.84836 | 0.78 | 0.867 | 0.774 | 0.87 | 0.94204 | 0.88 | 0.78 | 0.86 |  | 1
| 1 | 2 | 0.8767 | 0.939 | 0.74847 | 0.82 | 0.88 | 0.86 | 0.87 | 0.9 | 0.85 | 0.82 | 0.81 | 1.03 | 0.89 | 0.87793 | 0.88 | 0.84 | 0.84464 | 0.79 | 0.835 | 0.793 | 0.88 | 0.93252 | 0.86 | 0.79 | 0.86 |  | 13
| 1 | 3 | 0.8635 | 0.932 | 0.75034 | 0.82 | 0.9 | 0.85 | 0.86 | 0.91 | 0.83 | 0.82 | 0.81 | 0.99 | 0.88 | 0.86277 | 0.91 | 0.85 | 0.84261 | 0.78 | 0.894 | 0.767 | 0.87 | 0.94312 | 0.88 | 0.83 | 0.85 |  | 16
| 1 | 4 | 0.865 | 0.939 | 0.75363 | 0.8 | 0.89 | 0.86 | 0.85 | 0.92 | 0.84 | 0.8 | 0.79 | 0.97 | 0.87 | 0.86776 | 0.9 | 0.86 | 0.84117 | 0.78 | 0.871 | 0.766 | 0.87 | 0.93856 | 0.86 | 0.82 | 0.86 |  | 0.85347
| 1 | 5 | 0.8771 | 0.942 | 0.76718 | 0.82 | 0.89 | 0.86 | 0.87 | 0.9 | 0.83 | 0.82 | 0.82 | 0.962 | 0.89 | 0.8597 | 0.92 | 0.84 | 0.84975 | 0.78 | 0.914 | 0.768 | 0.87 | 0.93744 | 0.88 | 0.78 | 0.85 |  | 91
| 1 | 6 | 0.8728 | 0.933 | 0.79137 | 0.79 | 0.89 | 0.86 | 0.87 | 0.9 | 0.85 | 0.83 | 0.8 | 0.97 | 0.89 | 0.86613 | 0.91 | 0.84 | 0.84661 | 0.79 | 0.884 | 0.78 | 0.87 | 0.94912 | 0.85 | 0.79 | 0.85
| Mean |  | 0.87302 | 0.93517 | 0.76649 | 0.80833 | 0.89 | 0.85667 | 0.86167 | 0.905 | 0.84167 | 0.81667 | 0.80833 | 0.98533 | 0.88667 | 0.86219 | 0.90667 | 0.85 | 0.84552 | 0.78333 | 0.8775 | 0.77467 | 0.87167 | 0.94047 | 0.86833 | 0.79833 | 0.855
| Median |  | 0.87475 | 0.936 | 0.76041 | 0.81 | 0.89 | 0.86 | 0.865 | 0.9 | 0.845 | 0.82 | 0.81 | 0.98 | 0.89 | 0.86445 | 0.91 | 0.845 | 0.84562 | 0.78 | 0.8775 | 0.771 | 0.87 | 0.9403 | 0.87 | 0.79 | 0.855
| Std Dev. |  | 0.007547 | 0.0059133 | 0.019143 | 0.013292 | 0.0063246 | 0.005164 | 0.0098319 | 0.0083666 | 0.0098319 | 0.010328 | 0.01169 | 0.024712 | 0.010328 | 0.013 | 0.015055 | 0.012649 | 0.003326 | 0.005164 | 0.026853 | 0.010424 | 0.0040825 | 0.0056629 | 0.013292 | 0.02137 | 0.0054772
| Rel.Std.Dev. |  | 0.8645% | 0.6323% | 2.497% | 1.644% | 0.7106% | 0.6028% | 1.141% | 0.9245% | 1.168% | 1.265% | 1.446% | 2.508% | 1.165% | 1.508% | 1.661% | 1.488% | 0.3934% | 0.6592% | 3.06% | 1.346% | 0.4684% | 0.6021% | 1.531% | 2.677% | 0.6406%
| PDM3 |  | 2.29% | 9.572% | -10.19% | -5.289% | 4.28% | 0.3742% | 0.96% | 6.037% | -1.383% | -4.313% | -5.289% | 15.45% | 3.889% | 1.022% | 6.233% | -0.407% | -0.9315% | -8.218% | 2.815% | -9.234% | 2.132% | 10.19% | 1.741% | -6.461% | 0.1789%
| Z-Score (Absolute) |  | 0.27 | 1.75 | 2.27 | 1.27 | 0.67 | 0.12 | 0 | 1.03 | 0.48 | 1.07 | 1.27 | 2.94 | 0.59 | 0.01 | 1.07 | 0.28 | 0.38 | 1.86 | 0.38 | 2.07 | 0.24 | 1.88 | 0.16 | 1.51 | 0.16
| Analytical results for La in OREAS 506b (Certified Value 13.8 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 2
| 1 | 1 | 13.482 | 22.7 | 14.2694 | 11.8 | 13.9 | 13.3 | 13.8 | 15 | 14 | 13 | 15.3 | < 10 | 12.9 | 13.4144 | 14 | 13.6 | 25.3453 | 12 | 22.8 | 15.3 | 13.2 | NA | 13.3 | 14.6 | 11.8 |  | 1
| 1 | 2 | 13.629 | 22 | 13.8104 | 13.1 | 13.2 | 13.6 | 14.4 | 15.6 | 14 | 13.7 | 14.8 | 10 | 12.5 | 13.7824 | 13.3 | 13.2 | 25.1844 | 12 | 19.9 | 15.1 | 12.7 | NA | 13.7 | 14.5 | 13.6 |  | 14
| 1 | 3 | 13.425 | 21.9 | 13.9554 | 12.2 | 14.4 | 13.6 | 14 | 14.8 | 14.2 | 13.7 | 15.2 | < 10 | 12.8 | 13.7549 | 13.6 | 13.3 | 25.394 | 12 | 25.6 | 15.2 | 13.2 | NA | 13.9 | 15 | 13 |  | 16
| 1 | 4 | 13.614 | 22 | 13.8828 | 11 | 13.3 | 13.7 | 14 | 14.8 | 14.5 | 13.2 | 14.7 | < 10 | 12.5 | 14.0906 | 14.5 | 13.8 | 25.1929 | 11 | 25.1 | 15.1 | 12.6 | NA | 13.5 | 14.7 | 12.5 |  | 13.775
| 1 | 5 | 13.424 | 22.4 | 14.216 | 13.2 | 13.7 | 12.9 | 13.8 | 15.4 | 13.8 | 13.9 | 14.6 | < 10 | 12.9 | 13.752 | 14.8 | 13.2 | 25.4201 | 12 | 25.2 | 14.8 | 12.9 | NA | 13.3 | 13.9 | 12.3 |  | 92
| 1 | 6 | 13.705 | 22 | 14.2272 | 12.3 | 13.5 | 13.3 | 14.5 | 15.4 | 13.1 | 14 | 14.5 | < 10 | 12.8 | 14.1934 | 14.2 | 13.6 | 25.163 | 12 | 23.6 | 15.4 | 13.4 | NA | 13.3 | 14.9 | 12.3
| Mean |  | 13.5465 | 22.1667 | 14.0602 | 12.2667 | 13.6667 | 13.4 | 14.0833 | 15.1667 | 13.9333 | 13.5833 | 14.85 | 10 | 12.7333 | 13.8313 | 14.0667 | 13.45 | 25.2833 | 11.8333 | 23.7 | 15.15 | 13 |  | 13.5 | 14.6 | 12.5833
| Median |  | 13.548 | 22 | 14.0857 | 12.25 | 13.6 | 13.45 | 14 | 15.2 | 14 | 13.7 | 14.75 | 10 | 12.8 | 13.7687 | 14.1 | 13.45 | 25.2691 | 12 | 24.35 | 15.15 | 13.05 |  | 13.4 | 14.65 | 12.4
| Std Dev. |  | 0.11867 | 0.31411 | 0.20039 | 0.82381 | 0.44121 | 0.29665 | 0.29944 | 0.34448 | 0.47188 | 0.39707 | 0.32711 |  | 0.18619 | 0.27808 | 0.55737 | 0.251 | 0.11597 | 0.40825 | 2.1485 | 0.20736 | 0.31623 |  | 0.25298 | 0.38987 | 0.63061
| Rel.Std.Dev. |  | 0.876% | 1.417% | 1.425% | 6.716% | 3.228% | 2.214% | 2.126% | 2.271% | 3.387% | 2.923% | 2.203% |  | 1.462% | 2.011% | 3.962% | 1.866% | 0.4587% | 3.45% | 9.065% | 1.369% | 2.433% |  | 1.874% | 2.67% | 5.011%
| PDM3 |  | -1.659% | 60.92% | 2.071% | -10.95% | -0.7863% | -2.722% | 2.238% | 10.1% | 1.15% | -1.391% | 7.804% | -27.4% | -7.562% | 0.4088% | 2.117% | -2.359% | 83.54% | -14.1% | 72.05% | 9.982% | -5.626% |  | -1.996% | 5.989% | -8.651%
| Z-Score (Absolute) |  | 0.25 | 7.31 | 0.2 | 1.37 | 0.14 | 0.38 | 0.22 | 1.17 | 0.09 | 0.22 | 0.89 | 7.02 | 0.96 | 0 | 0.21 | 0.33 | 10.05 | NA | 8.66 | 1.16 | 0.73 | NA | 0.29 | 0.67 | 1.09
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 506b (Certified Value 33.4 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 1
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [5 empty] | 2
| 1 | 1 | 34.5 | 33.7 | 35.932 | 32.2 | 34.3 | 33.6 | 35.6 | 34 | 32.3 | 34 | 33 | 31.7652 | 31 | 33.2 | 43.91 | 33 | 35.5 | 38 | 27.6 | 34 | 32.2 | [5 empty] | 1
| 1 | 2 | 34.7 | 32.6 | 34.2178 | 31.4 | 34.7 | 34.6 | 35.2 | 34.1 | 32.7 | 34 | 33 | 33.2589 | 31 | 31.37 | 38.89 | 33 | 32.5 | 39 | 26.2 | 35 | 32.5 | [5 empty] | 15
| 1 | 3 | 33.8 | 33 | 34.8908 | 33.5 | 34.3 | 34.5 | 35.9 | 33.2 | 32.3 | 35 | 33 | 31.9242 | 32 | 32.17 | 40.29 | 33 | 32.3 | 37 | 28 | 34 | 33.7 | [5 empty] | 16
| 1 | 4 | 33.5 | 32.7 | 34.9002 | 31.8 | 34.9 | 33.6 | 36.1 | 33.2 | 31.8 | 35 | 33 | 33.2205 | 31 | 32.87 | 40.39 | 33 | 38.3 | 35 | 26.1 | 34 | 33.1 | [5 empty] | 33.4048
| 1 | 5 | 34 | 33.6 | 34.7194 | 31.7 | 34.8 | 34.4 | 35.7 | 33.2 | 32.4 | 34 | 34 | 32.9587 | 33 | 30.93 | 41.39 | 33 | 34.5 | 38 | 26.3 | 35 | 31.6 | [5 empty] | 93
| 1 | 6 | 34 | 32.7 | 35.8026 | 31.3 | 34.2 | 34.5 | 35.4 | 33.9 | 32 | 34 | 34 | 33.0828 | 31 | 32.03 | 40.1 | 34 | 30.7 | 38 | 28.5 | 34 | 32.5
| Mean |  | 34.0833 | 33.05 | 35.0771 | 31.9833 | 34.5333 | 34.2 | 35.65 | 33.6 | 32.25 | 34.3333 | 33.3333 | 32.7017 | 31.5 | 32.095 | 40.8283 | 33.1667 | 33.9667 | 37.5 | 27.1167 | 34.3333 | 32.6
| Median |  | 34 | 32.85 | 34.8955 | 31.75 | 34.5 | 34.45 | 35.65 | 33.55 | 32.3 | 34 | 33 | 33.0208 | 31 | 32.1 | 40.34 | 33 | 33.5 | 38 | 26.95 | 34 | 32.5
| Std Dev. |  | 0.4446 | 0.48477 | 0.66176 | 0.8085 | 0.30111 | 0.46904 | 0.32711 | 0.44272 | 0.31464 | 0.5164 | 0.5164 | 0.67418 | 0.83666 | 0.8619 | 1.7077 | 0.40825 | 2.7208 | 1.3784 | 1.0458 | 0.5164 | 0.72664
| Rel.Std.Dev. |  | 1.304% | 1.467% | 1.887% | 2.528% | 0.8719% | 1.371% | 0.9176% | 1.318% | 0.9756% | 1.504% | 1.549% | 2.062% | 2.656% | 2.685% | 4.183% | 1.231% | 8.01% | 3.676% | 3.857% | 1.504% | 2.229%
| PDM3 |  | 2.031% | -1.062% | 5.006% | -4.255% | 3.378% | 2.38% | 6.721% | 0.5842% | -3.457% | 2.78% | -0.2141% | -2.105% | -5.702% | -3.921% | 22.22% | -0.713% | 1.682% | 12.26% | -18.82% | 2.78% | -2.409%
| Z-Score (Absolute) |  | 0.35 | 0.4 | 1.07 | 1.17 | 0.67 | 0.43 | 1.48 | 0 | 0.98 | 0.53 | 0.19 | 0.65 | 1.52 | 1.09 | 5.22 | 0.31 | 0.26 | 2.82 | 4.68 | 0.53 | 0.72
| Analytical results for Lu in OREAS 506b (Indicative Value 0.096 ppm) | [27 empty] | Indicative
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | [5 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.06; 2.03: 2.14; 2.06: 2.15; 2.07: 2.24; 2.08: 2.25 | [5 empty] | ppm
|  |  | 2.01: AR*MS; 2.02: AR*OES; 2.03: AR*OES/MS; 2.06: AR*MS; 2.07: AR*OES/MS; 2.08: AR*OES/MS | [5 empty] | 3
|  |  | 2.01: 01g; 2.02: 0.25g; 2.03: 0.25g; 2.06: 0.2g; 2.07: 0.25g; 2.08: 0.5g | [5 empty] | 3
| 1 | 1 | 2.01: 0.104; 2.02: < 5; 2.03: 0.09; 2.06: 0.099888; 2.07: 0.1; 2.08: < 0.1 | [5 empty] | 1
| 1 | 2 | 2.01: 0.105; 2.02: < 5; 2.03: 0.09; 2.06: 0.088816; 2.07: 0.1; 2.08: < 0.1 | [5 empty] | 16
| 1 | 3 | 2.01: 0.1; 2.02: < 5; 2.03: 0.09; 2.06: 0.094575; 2.07: 0.09; 2.08: 0.1 | [5 empty] | 16
| 1 | 4 | 2.01: 0.102; 2.02: < 5; 2.03: 0.09; 2.06: 0.097146; 2.07: 0.1; 2.08: < 0.1 | [5 empty] | 0.09596
| 1 | 5 | 2.01: 0.104; 2.02: < 5; 2.03: 0.08; 2.06: 0.097898; 2.07: 0.1; 2.08: < 0.1 | [5 empty] | 40
| 1 | 6 | 2.01: 0.105; 2.02: < 5; 2.03: 0.08; 2.06: 0.096357; 2.07: 0.09; 2.08: 0.1
| Mean |  | 2.01: 0.10333; 2.03: 0.086667; 2.06: 0.09578; 2.07: 0.096667; 2.08: 0.1
| Median |  | 2.01: 0.104; 2.03: 0.09; 2.06: 0.096751; 2.07: 0.1; 2.08: 0.1
| Std Dev. |  | 2.01: 0.0019664; 2.03: 0.005164; 2.06: 0.0038343; 2.07: 0.005164; 2.08: 0
| Rel.Std.Dev. |  | 2.01: 1.903%; 2.03: 5.958%; 2.06: 4.003%; 2.07: 5.342%; 2.08: 0%
| PDM3 |  | 2.01: 7.684%; 2.03: -9.684%; 2.06: -0.1874%; 2.07: 0.7366%; 2.08: 4.21%
| Z-Score (Absolute) |  | 2.01: 0.58; 2.02: 194.51; 2.03: 0.77; 2.06: 0.04; 2.07: 0.04; 2.08: 2.39
| Analytical results for Mg in OREAS 506b (Certified Value 0.645 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.66 | 0.68 | 0.63967 | 0.64 | 0.66 | 0.63 | 0.63 | 0.68 | 0.65 | 0.63 | 0.71 | 0.64 | 0.66 | 0.65371 | 0.724 | 0.66 | 0.61841 | 0.63 | 0.625 | 0.778 | 0.61 | 0.6711 | 0.65 | 0.66 | 0.63 |  | 1
| 1 | 2 | 0.66 | 0.67 | 0.60716 | 0.64 | 0.64 | 0.64 | 0.65 | 0.68 | 0.65 | 0.64 | 0.73 | 0.685 | 0.67 | 0.65958 | 0.701 | 0.63 | 0.61265 | 0.63 | 0.633 | 0.806 | 0.62 | 0.6675 | 0.64 | 0.66 | 0.64 |  | 17
| 1 | 3 | 0.65 | 0.68 | 0.60456 | 0.65 | 0.66 | 0.64 | 0.64 | 0.69 | 0.63 | 0.63 | 0.73 | 0.643 | 0.66 | 0.65729 | 0.718 | 0.64 | 0.61143 | 0.63 | 0.637 | 0.783 | 0.64 | 0.6776 | 0.65 | 0.68 | 0.63 |  | 16
| 1 | 4 | 0.65 | 0.66 | 0.61289 | 0.63 | 0.63 | 0.64 | 0.64 | 0.69 | 0.63 | 0.63 | 0.72 | 0.62 | 0.65 | 0.66159 | 0.714 | 0.65 | 0.6198 | 0.63 | 0.625 | 0.738 | 0.6 | 0.6749 | 0.64 | 0.67 | 0.63 |  | 0.64469
| 1 | 5 | 0.66 | 0.68 | 0.62289 | 0.65 | 0.64 | 0.64 | 0.65 | 0.68 | 0.63 | 0.64 | 0.71 | 0.618 | 0.66 | 0.66019 | 0.723 | 0.64 | 0.61392 | 0.63 | 0.652 | 0.785 | 0.62 | 0.6809 | 0.65 | 0.66 | 0.63 |  | 94
| 1 | 6 | 0.65 | 0.66 | 0.64177 | 0.63 | 0.65 | 0.63 | 0.65 | 0.68 | 0.64 | 0.61 | 0.72 | 0.63 | 0.66 | 0.66432 | 0.711 | 0.64 | 0.61735 | 0.64 | 0.625 | 0.795 | 0.64 | 0.6318 | 0.63 | 0.65 | 0.63
| Mean |  | 0.655 | 0.67167 | 0.62149 | 0.64 | 0.64667 | 0.63667 | 0.64333 | 0.68333 | 0.63833 | 0.63 | 0.72 | 0.63933 | 0.66 | 0.65944 | 0.71517 | 0.64333 | 0.61559 | 0.63167 | 0.63283 | 0.78083 | 0.62167 | 0.6673 | 0.64333 | 0.66333 | 0.63167
| Median |  | 0.655 | 0.675 | 0.61789 | 0.64 | 0.645 | 0.64 | 0.645 | 0.68 | 0.635 | 0.63 | 0.72 | 0.635 | 0.66 | 0.65988 | 0.716 | 0.64 | 0.61564 | 0.63 | 0.629 | 0.784 | 0.62 | 0.673 | 0.645 | 0.66 | 0.63
| Std Dev. |  | 0.0054772 | 0.0098319 | 0.016184 | 0.0089443 | 0.012111 | 0.005164 | 0.008165 | 0.005164 | 0.0098319 | 0.010954 | 0.0089443 | 0.024557 | 0.0063246 | 0.0036479 | 0.0085654 | 0.010328 | 0.0033915 | 0.0040825 | 0.010666 | 0.023233 | 0.016021 | 0.01802 | 0.008165 | 0.010328 | 0.0040825
| Rel.Std.Dev. |  | 0.8362% | 1.464% | 2.604% | 1.398% | 1.873% | 0.8111% | 1.269% | 0.7557% | 1.54% | 1.739% | 1.242% | 3.841% | 0.9583% | 0.5532% | 1.198% | 1.605% | 0.5509% | 0.6463% | 1.685% | 2.975% | 2.577% | 2.7% | 1.269% | 1.557% | 0.6463%
| PDM3 |  | 1.6% | 4.185% | -3.598% | -0.7269% | 0.3072% | -1.244% | -0.2098% | 5.995% | -0.9854% | -2.278% | 11.68% | -0.8303% | 2.375% | 2.289% | 10.93% | -0.2098% | -4.513% | -2.02% | -1.839% | 21.12% | -3.571% | 3.508% | -0.2098% | 2.892% | -2.02%
| Z-Score (Absolute) |  | 0.59 | 1.43 | 1.1 | 0.17 | 0.17 | 0.34 | 0 | 2.02 | 0.25 | 0.67 | 3.88 | 0.2 | 0.84 | 0.81 | 3.63 | 0 | 1.4 | 0.59 | 0.53 | 6.95 | 1.1 | 1.21 | 0 | 1.01 | 0.59
| Analytical results for Mn in OREAS 506b (Certified Value 0.029 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.0284 | 0.0312 | 0.027874 | 0.029 | 0.03556 | 0.028 | 0.0275 | 0.0301 | 0.0288 | 0.0282 | 0.0275 | 0.0285 | 0.0308 | 0.028447 | 0.0284 | 0.0291 | 0.028863 | 0.0283 | 0.0277 | 0.0326 | 0.0277 | 0.0278 | 0.0275 | 0.0294 | 0.0288 |  | 1
| 1 | 2 | 0.0284 | 0.0314 | 0.026601 | 0.0293 | 0.03339 | 0.0287 | 0.0283 | 0.0301 | 0.0287 | 0.0287 | 0.0285 | 0.0303 | 0.0293 | 0.029133 | 0.0277 | 0.0293 | 0.02847 | 0.0287 | 0.0275 | 0.0336 | 0.0303 | 0.0287 | 0.0274 | 0.0297 | 0.0291 |  | 18
| 1 | 3 | 0.028 | 0.0313 | 0.027543 | 0.0295 | 0.03497 | 0.0285 | 0.0279 | 0.0307 | 0.0281 | 0.0282 | 0.0281 | 0.0286 | 0.0305 | 0.028732 | 0.0281 | 0.0292 | 0.028661 | 0.0284 | 0.0284 | 0.033 | 0.0275 | 0.0284 | 0.0272 | 0.0305 | 0.0289 |  | 16
| 1 | 4 | 0.0281 | 0.0313 | 0.026563 | 0.0289 | 0.03473 | 0.0289 | 0.0278 | 0.0304 | 0.0279 | 0.028 | 0.0276 | 0.0277 | 0.0295 | 0.029451 | 0.028 | 0.0292 | 0.028625 | 0.0286 | 0.0279 | 0.0307 | 0.0289 | 0.0282 | 0.0275 | 0.0298 | 0.0289 |  | 0.028561
| 1 | 5 | 0.0283 | 0.0317 | 0.027423 | 0.0296 | 0.03298 | 0.0286 | 0.028 | 0.03 | 0.028 | 0.0286 | 0.028 | 0.0277 | 0.0301 | 0.028917 | 0.0286 | 0.0291 | 0.028153 | 0.0286 | 0.0293 | 0.0328 | 0.0282 | 0.0279 | 0.0279 | 0.0296 | 0.0286 |  | 95
| 1 | 6 | 0.0282 | 0.0316 | 0.027397 | 0.0287 | 0.03351 | 0.0284 | 0.0285 | 0.0301 | 0.0285 | 0.0286 | 0.0281 | 0.0279 | 0.0306 | 0.029288 | 0.0281 | 0.0291 | 0.028421 | 0.0285 | 0.0282 | 0.0335 | 0.0272 | 0.0282 | 0.0272 | 0.0295 | 0.0287
| Mean |  | 0.028233 | 0.031417 | 0.027233 | 0.029167 | 0.03419 | 0.028517 | 0.028 | 0.030233 | 0.028333 | 0.028383 | 0.027967 | 0.02845 | 0.030133 | 0.028995 | 0.02815 | 0.029167 | 0.028532 | 0.028517 | 0.028167 | 0.0327 | 0.0283 | 0.0282 | 0.02745 | 0.02975 | 0.028833
| Median |  | 0.02825 | 0.03135 | 0.02741 | 0.02915 | 0.03412 | 0.02855 | 0.02795 | 0.0301 | 0.0283 | 0.0284 | 0.02805 | 0.0282 | 0.0303 | 0.029025 | 0.0281 | 0.02915 | 0.028547 | 0.02855 | 0.02805 | 0.0329 | 0.02795 | 0.0282 | 0.02745 | 0.02965 | 0.02885
| Std Dev. |  | 0.0001633 | 0.00019408 | 0.00053249 | 0.0003559 | 0.0010338 | 0.00030605 | 0.00035777 | 0.00026583 | 0.00038297 | 0.00028577 | 0.00036697 | 0.00098742 | 0.00061536 | 0.00037113 | 0.00031464 | 8.165e-05 | 0.00024279 | 0.0001472 | 0.00064395 | 0.0010545 | 0.0011472 | 0.00032863 | 0.00025884 | 0.0003937 | 0.00017512
| Rel.Std.Dev. |  | 0.5784% | 0.6178% | 1.955% | 1.22% | 3.024% | 1.073% | 1.278% | 0.8793% | 1.352% | 1.007% | 1.312% | 3.471% | 2.042% | 1.28% | 1.118% | 0.2799% | 0.8509% | 0.5162% | 2.286% | 3.225% | 4.054% | 1.165% | 0.943% | 1.323% | 0.6073%
| PDM3 |  | -1.147% | 9.999% | -4.648% | 2.121% | 19.71% | -0.155% | -1.964% | 5.856% | -0.7969% | -0.6218% | -2.081% | -0.3884% | 5.505% | 1.519% | -1.439% | 2.121% | -0.1001% | -0.155% | -1.38% | 14.49% | -0.9136% | -1.264% | -3.89% | 4.163% | 0.9538%
| Z-Score (Absolute) |  | 0.4 | 4.09 | 1.81 | 0.92 | 8 | 0 | 0.73 | 2.42 | 0.26 | 0.19 | 0.78 | 0.09 | 2.28 | 0.67 | 0.52 | 0.92 | 0.02 | 0 | 0.49 | 5.9 | 0.31 | 0.45 | 1.5 | 1.74 | 0.45
| Analytical results for Mo in OREAS 506b (Certified Value 84 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 1
| 1 | 1 | 84.6 | 84.6 | 77.7507 | 83.2 | 81.5 | 78.9 | 82.2 | 81.7 | 85.8 | 81.2 | 88.06 | 93 | 85.76 | 84.8433 | 87 | 83.96 | 89 | 84.4 | 83.23 | 86 | 84.5 | 91.3 | 81.1 | 79.4 | 83.88 |  | 1
| 1 | 2 | 85.8 | 83 | 76.4342 | 83.78 | 81.5 | 80.8 | 82.8 | 82.4 | 86.2 | 82.5 | 91.04 | 98.5 | 85.83 | 86.978 | 85 | 82.37 | 88 | 80.1 | 85.27 | 86 | 84.2 | 91.6 | 81.2 | 81.3 | 84.86 |  | 19
| 1 | 3 | 85.2 | 84.3 | 77.309 | 84.17 | 81.6 | 79.7 | 83.9 | 81.9 | 85.9 | 83.4 | 88.95 | 91.4 | 85.15 | 85.6592 | 88 | 82.55 | 89 | 83 | 84.94 | 85 | 84.6 | 90.3 | 84.7 | 81.2 | 84.73 |  | 16
| 1 | 4 | 83.7 | 83.1 | 74.8619 | 82.67 | 81.6 | 82 | 82.4 | 82.2 | 85.5 | 81.4 | 87.46 | 91 | 82.91 | 86.2214 | 87 | 85.32 | 85 | 84.1 | 84.72 | 85 | 83.9 | 91.7 | 83.4 | 80.4 | 84.67 |  | 83.9642
| 1 | 5 | 85.4 | 83.1 | 75.4314 | 84.27 | 81.4 | 78.3 | 82.4 | 82 | 84.8 | 83.2 | 88.54 | 90.3 | 84.26 | 85.7296 | 88 | 83.73 | 86 | 83.9 | 82.34 | 86 | 84.3 | 92.5 | 81.9 | 81.4 | 85.09 |  | 96
| 1 | 6 | 85.1 | 83 | 77.8107 | 82.5 | 81.6 | 79.3 | 83 | 79.5 | 83.1 | 79.9 | 90.44 | 90 | 84.55 | 85.8472 | 87 | 83.54 | 87 | 81.8 | 85.07 | 86 | 84.3 | 88.9 | 81.9 | 80.5 | 84.29
| Mean |  | 84.9667 | 83.5167 | 76.5997 | 83.4317 | 81.5333 | 79.8333 | 82.7833 | 81.6167 | 85.2167 | 81.9333 | 89.0817 | 92.3667 | 84.7433 | 85.8798 | 87 | 83.5783 | 87.3333 | 82.8833 | 84.2617 | 85.6667 | 84.3 | 91.05 | 82.3667 | 80.7 | 84.5867
| Median |  | 85.15 | 83.1 | 76.8716 | 83.49 | 81.55 | 79.5 | 82.6 | 81.95 | 85.65 | 81.95 | 88.745 | 91.2 | 84.85 | 85.7884 | 87 | 83.635 | 87.5 | 83.45 | 84.83 | 86 | 84.3 | 91.45 | 81.9 | 80.85 | 84.7
| Std Dev. |  | 0.73394 | 0.73052 | 1.2415 | 0.75777 | 0.08165 | 1.353 | 0.62102 | 1.0647 | 1.1409 | 1.3441 | 1.3903 | 3.1841 | 1.0965 | 0.70278 | 1.0954 | 1.0697 | 1.633 | 1.6582 | 1.1914 | 0.5164 | 0.24495 | 1.2708 | 1.4081 | 0.7642 | 0.43427
| Rel.Std.Dev. |  | 0.8638% | 0.8747% | 1.621% | 0.9083% | 0.1001% | 1.695% | 0.7502% | 1.305% | 1.339% | 1.641% | 1.561% | 3.447% | 1.294% | 0.8183% | 1.259% | 1.28% | 1.87% | 2.001% | 1.414% | 0.6028% | 0.2906% | 1.396% | 1.71% | 0.947% | 0.5134%
| PDM3 |  | 1.194% | -0.533% | -8.771% | -0.6343% | -2.895% | -4.92% | -1.406% | -2.796% | 1.492% | -2.419% | 6.095% | 10.01% | 0.9279% | 2.281% | 3.616% | -0.4596% | 4.013% | -1.287% | 0.3542% | 2.028% | 0.3999% | 8.439% | -1.903% | -3.888% | 0.7413%
| Z-Score (Absolute) |  | 0.29 | 0.31 | 3.19 | 0.35 | 1.14 | 1.85 | 0.62 | 1.1 | 0.4 | 0.97 | 2.01 | 3.38 | 0.2 | 0.67 | 1.14 | 0.28 | 1.28 | 0.57 | 0 | 0.59 | 0.02 | 2.83 | 0.79 | 1.48 | 0.14
| Analytical results for Na in OREAS 506b (Certified Value 0.142 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  | 3
| 1 | 1 | 0.13 | 0.13 | 0.14123 | 0.137 | 0.13 | 0.13 | 0.12 | 0.17 | 0.16 | 0.13 | 0.15 | 0.26 | 0.12 | 0.13807 | 0.134 | 0.15 | 0.16509 | 0.14 | 0.136 | 0.127 | 0.15 | 0.1583 | 0.14 | 0.163 | 0.139 |  | 1
| 1 | 2 | 0.13 | 0.13 | 0.13358 | 0.144 | 0.13 | 0.13 | 0.13 | 0.17 | 0.16 | 0.14 | 0.15 | 0.264 | 0.12 | 0.14247 | 0.127 | 0.14 | 0.16783 | 0.15 | 0.118 | 0.121 | 0.16 | 0.158 | 0.14 | 0.161 | 0.143 |  | 20
| 1 | 3 | 0.13 | 0.13 | 0.13417 | 0.138 | 0.14 | 0.13 | 0.13 | 0.17 | 0.16 | 0.14 | 0.16 | 0.258 | 0.12 | 0.13619 | 0.129 | 0.14 | 0.16274 | 0.14 | 0.151 | 0.129 | 0.15 | 0.1639 | 0.14 | 0.171 | 0.142 |  | 16
| 1 | 4 | 0.13 | 0.13 | 0.13444 | 0.132 | 0.13 | 0.13 | 0.12 | 0.17 | 0.16 | 0.13 | 0.16 | 0.24 | 0.11 | 0.14009 | 0.134 | 0.14 | 0.16605 | 0.14 | 0.143 | 0.129 | 0.15 | 0.1649 | 0.14 | 0.163 | 0.142 |  | 0.14197
| 1 | 5 | 0.13 | 0.13 | 0.13948 | 0.142 | 0.13 | 0.13 | 0.13 | 0.17 | 0.16 | 0.14 | 0.15 | 0.264 | 0.11 | 0.13965 | 0.136 | 0.14 | 0.16291 | 0.14 | 0.146 | 0.114 | 0.15 | 0.1685 | 0.14 | 0.16 | 0.137 |  | 97
| 1 | 6 | 0.13 | 0.13 | 0.13956 | 0.135 | 0.13 | 0.13 | 0.13 | 0.18 | 0.16 | 0.14 | 0.16 | 0.25 | 0.11 | 0.14203 | 0.137 | 0.14 | 0.16583 | 0.15 | 0.142 | 0.119 | 0.15 | 0.1589 | 0.14 | 0.157 | 0.136
| Mean |  | 0.13 | 0.13 | 0.13708 | 0.138 | 0.13167 | 0.13 | 0.12667 | 0.17167 | 0.16 | 0.13667 | 0.155 | 0.256 | 0.115 | 0.13975 | 0.13283 | 0.14167 | 0.16508 | 0.14333 | 0.13933 | 0.12317 | 0.15167 | 0.16208 | 0.14 | 0.1625 | 0.13983
| Median |  | 0.13 | 0.13 | 0.13696 | 0.1375 | 0.13 | 0.13 | 0.13 | 0.17 | 0.16 | 0.14 | 0.155 | 0.259 | 0.115 | 0.13987 | 0.134 | 0.14 | 0.16546 | 0.14 | 0.1425 | 0.124 | 0.15 | 0.1614 | 0.14 | 0.162 | 0.1405
| Std Dev. |  | 0 | 0 | 0.0033703 | 0.0044272 | 0.0040825 | 0 | 0.005164 | 0.0040825 | 0 | 0.005164 | 0.0054772 | 0.0093808 | 0.0054772 | 0.0023752 | 0.0039707 | 0.0040825 | 0.0019608 | 0.005164 | 0.011553 | 0.0061455 | 0.0040825 | 0.004325 | 0 | 0.0047223 | 0.0029269
| Rel.Std.Dev. |  | 0% | 0% | 2.459% | 3.208% | 3.101% | 0% | 4.077% | 2.378% | 0% | 3.779% | 3.534% | 3.664% | 4.763% | 1.7% | 2.989% | 2.882% | 1.188% | 3.603% | 8.291% | 4.99% | 2.692% | 2.668% | 0% | 2.906% | 2.093%
| PDM3 |  | -8.431% | -8.431% | -3.446% | -2.796% | -7.257% | -8.431% | -10.78% | 20.92% | 12.7% | -3.735% | 9.179% | 80.32% | -19% | -1.564% | -6.435% | -0.2128% | 16.28% | 0.9612% | -1.856% | -13.24% | 6.831% | 14.17% | -1.387% | 14.46% | -1.504%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 0.18 | 0.12 | 0.56 | 0.67 | 0.9 | 2.21 | 1.4 | 0.21 | 1.05 | 8.04 | 1.71 | 0 | 0.48 | 0.13 | 1.75 | 0.25 | 0.03 | 1.15 | 0.82 | 1.54 | 0.02 | 1.57 | 0.01
| Analytical results for Nb in OREAS 506b (Certified Value 0.76 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.68 | 1 | 1.0619 | < 5 | 0.57 | 0.72 | 0.59 | 0.79 | 0.72 | 1.2 | 1.3854 | < 1 | 1.45 | 0.72 | 0.64 | 0.87 | 0.8 | [9 empty] | 1
| 1 | 2 | 0.72 | 1 | 1.1312 | < 5 | 0.61 | 0.64 | 0.58 | 0.77 | 0.67 | 1.21 | 1.446 | < 1 | 1.69 | 0.79 | 0.67 | 1.12 | 0.6 | [9 empty] | 21
| 1 | 3 | 0.71 | 0.9 | 1.1045 | < 5 | 0.66 | 0.9 | 0.68 | 0.82 | 0.86 | 1.19 | 1.4935 | < 1 | 1.76 | 0.77 | 0.62 | 0.92 | 0.8 | [9 empty] | 16
| 1 | 4 | 0.69 | 0.9 | 1.0847 | < 5 | 0.65 | 0.74 | 0.57 | 0.85 | 0.73 | 1.15 | 1.4182 | < 1 | 1.66 | 0.7 | 0.68 | 0.89 | 0.7 | [9 empty] | 0.75864
| 1 | 5 | 0.7 | 0.9 | 1.1346 | < 5 | 0.61 | 0.65 | 0.59 | 1.06 | 0.75 | 1.18 | 1.4631 | < 1 | 1.57 | 0.69 | 0.68 | 0.87 | 0.7 | [9 empty] | 98
| 1 | 6 | 0.7 | 0.9 | 1.0935 | < 5 | 0.62 | 0.71 | 0.58 | 0.84 | 0.71 | 1.18 | 1.4499 | < 1 | 1.61 | 0.69 | 0.64 | 0.97 | 0.6
| Mean |  | 0.7 | 0.93333 | 1.1017 |  | 0.62 | 0.72667 | 0.59833 | 0.855 | 0.74 | 1.185 | 1.4427 |  | 1.6233 | 0.72667 | 0.655 | 0.94 | 0.7
| Median |  | 0.7 | 0.9 | 1.099 |  | 0.615 | 0.715 | 0.585 | 0.83 | 0.725 | 1.185 | 1.4479 |  | 1.635 | 0.71 | 0.655 | 0.905 | 0.7
| Std Dev. |  | 0.014142 | 0.05164 | 0.027937 |  | 0.032249 | 0.093737 | 0.040702 | 0.10483 | 0.064498 | 0.020736 | 0.037244 |  | 0.10727 | 0.043205 | 0.0251 | 0.095917 | 0.089443
| Rel.Std.Dev. |  | 2.02% | 5.533% | 2.536% |  | 5.201% | 12.9% | 6.803% | 12.26% | 8.716% | 1.75% | 2.582% |  | 6.608% | 5.946% | 3.832% | 10.2% | 12.78%
| PDM3 |  | -7.729% | 23.03% | 45.22% |  | -18.27% | -4.214% | -21.13% | 12.7% | -2.457% | 56.2% | 90.17% |  | 114% | -4.214% | -13.66% | 23.91% | -7.729%
| Z-Score (Absolute) |  | 0.19 | NA | 1.72 | 8.38 | 0.57 | 0.06 | 0.67 | 0.55 | 0 | 2.12 | 3.34 | 1.14 | 4.2 | 0.06 | 0.4 | 0.95 | NA
Results from laboratories 2.02 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 506b (Indicative Value 15 ppm) | [27 empty] | Indicative
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.06: Lab | [9 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.06: 2.25 | [9 empty] | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.06: AR*OES/MS | [9 empty] | 1
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.06: 0.5g | [9 empty] | 1
| 1 | 1 | 2.01: 14.199; 2.02: 14.8047; 2.03: 25.1534; 2.06: 15.1 | [9 empty] | 1
| 1 | 2 | 2.01: 14.68; 2.02: 15.4946; 2.03: 25.2472; 2.06: 15 | [9 empty] | 22
| 1 | 3 | 2.01: 14.389; 2.02: 15.1845; 2.03: 25.0156; 2.06: 15.1 | [9 empty] | 16
| 1 | 4 | 2.01: 14.634; 2.02: 16.2887; 2.03: 24.9883; 2.06: 15.3 | [9 empty] | 15.0219
| 1 | 5 | 2.01: 14.435; 2.02: 15.5238; 2.03: 25.0891; 2.06: 14.5 | [9 empty] | 41
| 1 | 6 | 2.01: 14.609; 2.02: 16.1516; 2.03: 25.2119; 2.06: 15
| Mean |  | 2.01: 14.491; 2.02: 15.5746; 2.03: 25.1176; 2.06: 15
| Median |  | 2.01: 14.522; 2.02: 15.5092; 2.03: 25.1213; 2.06: 15.05
| Std Dev. |  | 2.01: 0.18379; 2.02: 0.56514; 2.03: 0.10477; 2.06: 0.26833
| Rel.Std.Dev. |  | 2.01: 1.268%; 2.02: 3.629%; 2.03: 0.4171%; 2.06: 1.789%
| PDM3 |  | 2.01: -3.534%; 2.02: 3.68%; 2.03: 67.21%; 2.06: -0.1457%
| Z-Score (Absolute) |  | 2.01: 0.99; 2.02: 0.36; 2.03: 12.23; 2.06: 0.36
| Analytical results for Ni in OREAS 506b (Certified Value 15.5 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 2
| 1 | 1 | 16.7 | 16 | 14.5015 | 15.6 | 13.93 | 14.6 | 15.9 | 16.2 | 16.4 | 14.8 | 15.2 | 16 | 15.7 | 15.1626 | 17 | 15.6 | 15 | 15.5 | 15.1 | 14.1 | 16.1 | 15.2 | 15.6 |  |  |  | 1
| 1 | 2 | 16.4 | 16 | 14.259 | 15.6 | 13.82 | 14.6 | 16 | 16.3 | 15.9 | 15.3 | 15.6 | 16.7 | 16.3 | 15.7329 | 17 | 15.1 | 16 | 14.8 | 16.5 | 14 | 15.9 | 15.1 | 15.9 |  |  |  | 23
| 1 | 3 | 16.7 | 16 | 14.5825 | 15.3 | 14.14 | 14.3 | 16 | 15.8 | 15.9 | 15 | 14.9 | 16.1 | 15.9 | 15.5236 | 16 | 15.6 | 16 | 15.5 | 14.9 | 14.3 | 16.1 | 15.6 | 15.9 |  |  |  | 16
| 1 | 4 | 16.4 | 16 | 14.2137 | 14.8 | 13.79 | 15 | 15.9 | 16.2 | 16 | 15.1 | 15.3 | 16 | 15.6 | 16.0053 | 16 | 16.1 | 16 | 16.2 | 16 | 14.5 | 15.8 | 15.3 | 15.9 |  |  |  | 15.5003
| 1 | 5 | 16.7 | 16 | 14.351 | 16.1 | 13.63 | 14.4 | 15.6 | 16 | 15.7 | 15.2 | 14.9 | 15.6 | 16.1 | 15.3273 | 17 | 15.1 | 16 | 15.6 | 15.4 | 13.9 | 15.9 | 15.1 | 15.7 |  |  |  | 99
| 1 | 6 | 16.9 | 16 | 14.6236 | 15.6 | 13.75 | 14.2 | 16 | 15.9 | 15.4 | 15.6 | 15.2 | 17 | 16.2 | 15.5857 | 16 | 15.4 | 16 | 15.9 | 14.2 | 14.4 | 15.5 | 14.2 | 15.3
| Mean |  | 16.6333 | 16 | 14.4219 | 15.5 | 13.8433 | 14.5167 | 15.9 | 16.0667 | 15.8833 | 15.1667 | 15.1833 | 16.2333 | 15.9667 | 15.5562 | 16.5 | 15.4833 | 15.8333 | 15.5833 | 15.35 | 14.2 | 15.8833 | 15.0833 | 15.7167
| Median |  | 16.7 | 16 | 14.4263 | 15.6 | 13.805 | 14.5 | 15.95 | 16.1 | 15.9 | 15.15 | 15.2 | 16.05 | 16 | 15.5547 | 16.5 | 15.5 | 16 | 15.55 | 15.25 | 14.2 | 15.9 | 15.15 | 15.8
| Std Dev. |  | 0.19664 | 0 | 0.17187 | 0.42895 | 0.175 | 0.28577 | 0.15492 | 0.19664 | 0.33116 | 0.27325 | 0.26394 | 0.5164 | 0.28048 | 0.29735 | 0.54772 | 0.37639 | 0.40825 | 0.47081 | 0.8167 | 0.23664 | 0.22286 | 0.47081 | 0.24014
| Rel.Std.Dev. |  | 1.182% | 0% | 1.192% | 2.767% | 1.264% | 1.969% | 0.9743% | 1.224% | 2.085% | 1.802% | 1.738% | 3.181% | 1.757% | 1.911% | 3.32% | 2.431% | 2.578% | 3.021% | 5.321% | 1.667% | 1.403% | 3.121% | 1.528%
| PDM3 |  | 7.31% | 3.224% | -6.957% | -0.001622% | -10.69% | -6.346% | 2.579% | 3.654% | 2.471% | -2.152% | -2.045% | 4.729% | 3.009% | 0.3611% | 6.45% | -0.1091% | 2.149% | 0.536% | -0.9693% | -8.389% | 2.471% | -2.69% | 1.396%
| Z-Score (Absolute) |  | 2.03 | NA | 2.03 | 0.05 | 3.1 | 1.86 | 0.68 | 0.99 | 0.65 | 0.66 | 0.63 | 1.3 | 0.81 | 0.05 | NA | 0.08 | NA | 0.1 | 0.33 | 2.44 | 0.65 | 0.82 | 0.35
Results from laboratories 2.02, 2.16 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 506b (Certified Value 0.065 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.0671 | 0.0721 | 0.061202 | 0.062 | 0.06 | 0.066 | 0.066 | 0.069 | 0.065 | 0.065 | 0.064 | 0.069 | 0.0638 | 0.066245 | 0.069 | 0.0658 | 0.0659 | 0.0677 | 0.067 | 0.0642 | 0.0684 | 0.06 | 0.068 | 0.064 |  |  | 1
| 1 | 2 | 0.0668 | 0.071 | 0.058149 | 0.062 | 0.06 | 0.066 | 0.067 | 0.069 | 0.063 | 0.067 | 0.065 | 0.0722 | 0.0666 | 0.066252 | 0.068 | 0.063667 | 0.0663 | 0.0635 | 0.069 | 0.063 | 0.0693 | 0.06 | 0.068 | 0.065 |  |  | 24
| 1 | 3 | 0.0667 | 0.0715 | 0.058345 | 0.063 | 0.06 | 0.065 | 0.067 | 0.069 | 0.063 | 0.066 | 0.064 | 0.0683 | 0.0675 | 0.066347 | 0.07 | 0.064333 | 0.0658 | 0.0631 | 0.065 | 0.0635 | 0.0692 | 0.06 | 0.069 | 0.063 |  |  | 16
| 1 | 4 | 0.0666 | 0.0701 | 0.058385 | 0.063 | 0.06 | 0.067 | 0.068 | 0.07 | 0.063 | 0.064 | 0.066 | 0.066 | 0.0645 | 0.067141 | 0.069 | 0.0653 | 0.0664 | 0.0588 | 0.064 | 0.0626 | 0.0685 | 0.06 | 0.069 | 0.064 |  |  | 0.065458
| 1 | 5 | 0.0664 | 0.0718 | 0.059052 | 0.063 | 0.06 | 0.065 | 0.067 | 0.069 | 0.062 | 0.066 | 0.064 | 0.0665 | 0.0633 | 0.066639 | 0.071 | 0.064367 | 0.066 | 0.0688 | 0.066 | 0.062 | 0.0699 | 0.06 | 0.069 | 0.063 |  |  | 100
| 1 | 6 | 0.0667 | 0.0694 | 0.061371 | 0.061 | 0.06 | 0.065 | 0.069 | 0.068 | 0.063 | 0.064 | 0.065 | 0.065 | 0.0677 | 0.066648 | 0.069 | 0.064367 | 0.0662 | 0.0717 | 0.065 | 0.0645 | 0.0704 | 0.06 | 0.068 | 0.063
| Mean |  | 0.066717 | 0.070983 | 0.059417 | 0.062333 | 0.06 | 0.065667 | 0.067333 | 0.069 | 0.063167 | 0.065333 | 0.064667 | 0.067833 | 0.065567 | 0.066545 | 0.069333 | 0.064639 | 0.0661 | 0.0656 | 0.066 | 0.0633 | 0.069283 | 0.06 | 0.0685 | 0.063667
| Median |  | 0.0667 | 0.07125 | 0.058719 | 0.0625 | 0.06 | 0.0655 | 0.067 | 0.069 | 0.063 | 0.0655 | 0.0645 | 0.0674 | 0.06555 | 0.066493 | 0.069 | 0.064367 | 0.0661 | 0.0656 | 0.0655 | 0.06325 | 0.06925 | 0.06 | 0.0685 | 0.0635
| Std Dev. |  | 0.00023166 | 0.0010458 | 0.0014806 | 0.0008165 | 0 | 0.0008165 | 0.0010328 | 0.00063246 | 0.00098319 | 0.0012111 | 0.0008165 | 0.0026005 | 0.0019367 | 0.00034348 | 0.0010328 | 0.00077087 | 0.00023664 | 0.0046639 | 0.0017889 | 0.00095499 | 0.00077825 | 0 | 0.00054772 | 0.0008165
| Rel.Std.Dev. |  | 0.3472% | 1.473% | 2.492% | 1.31% | 0% | 1.243% | 1.534% | 0.9166% | 1.557% | 1.854% | 1.263% | 3.834% | 2.954% | 0.5162% | 1.49% | 1.193% | 0.358% | 7.11% | 2.71% | 1.509% | 1.123% | 0% | 0.7996% | 1.282%
| PDM3 |  | 1.923% | 8.442% | -9.228% | -4.773% | -8.338% | 0.3192% | 2.865% | 5.412% | -3.5% | -0.19% | -1.208% | 3.629% | 0.1665% | 1.661% | 5.921% | -1.251% | 0.9812% | 0.2174% | 0.8285% | -3.296% | 5.844% | -8.338% | 4.648% | -2.736%
| Z-Score (Absolute) |  | 0.35 | 1.73 | 2.01 | 1.07 | 1.82 | 0.01 | 0.55 | 1.09 | 0.8 | 0.1 | 0.31 | 0.71 | 0.02 | 0.3 | 1.2 | 0.32 | 0.15 | 0.01 | 0.12 | 0.76 | 1.18 | 1.82 | 0.93 | 0.64
| Analytical results for Pb in OREAS 506b (Certified Value 39.5 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 2
| 1 | 1 | 40.2 | 41.1 | 38.1164 | 35.2 | 42.5 | 37.6 | 41.5 | 40.7 | 40.5 | 37.9 | 40 | 45 | 37.3 | 39.8853 | 39 | 40.1 | 36.95 | 38 | 39.4 | 34.7 | 43.8 | 36.4 | 36.9 | 43 | 39.7 |  | 1
| 1 | 2 | 40.4 | 41.5 | 36.8414 | 37.27 | 44.6 | 38.2 | 42.2 | 40.9 | 40.5 | 38.9 | 41.5 | 46 | 37.6 | 40.3044 | 39 | 42.9 | 35.39 | 38 | 39.6 | 34.1 | 44.1 | 36.7 | 37.5 | 43 | 41.93 |  | 25
| 1 | 3 | 39.5 | 41.7 | 38.8853 | 33.23 | 43.9 | 37 | 41.9 | 40.2 | 39.5 | 39 | 40.2 | 43 | 38.2 | 38.9997 | 39 | 39.4 | 37.62 | 38 | 38.2 | 37.5 | 43.2 | 36 | 38.9 | 43 | 40.54 |  | 16
| 1 | 4 | 39.4 | 42.4 | 36.5237 | 32.58 | 44.3 | 38.3 | 41.2 | 40.5 | 40.3 | 38.4 | 39 | 44 | 36.8 | 39.0491 | 39 | 40.6 | 37.05 | 38 | 39.3 | 33.5 | 44.5 | 37.1 | 38 | 43 | 40.34 |  | 39.5071
| 1 | 5 | 40.3 | 41.6 | 36.7519 | 37.55 | 44.5 | 37.6 | 41.1 | 40.2 | 39.5 | 38.6 | 39.5 | 42 | 38.3 | 39.5168 | 40 | 40.3 | 37 | 38 | 40.1 | 33.4 | 43.7 | 37.3 | 42.2 | 44 | 40.39 |  | 101
| 1 | 6 | 40.6 | 41.9 | 38.9974 | 35.77 | 43.6 | 37.3 | 43.3 | 40.3 | 38.2 | 41.6 | 39.2 | 42 | 37.1 | 40.947 | 39 | 39.5 | 36.56 | 39 | 38.9 | 35.8 | 43.9 | 36.3 | 37.5 | 42 | 39.8
| Mean |  | 40.0667 | 41.7 | 37.686 | 35.2667 | 43.9 | 37.6667 | 41.8667 | 40.4667 | 39.75 | 39.0667 | 39.9 | 43.6667 | 37.55 | 39.7837 | 39.1667 | 40.4667 | 36.7617 | 38.1667 | 39.25 | 34.8333 | 43.8667 | 36.6333 | 38.5 | 43 | 40.45
| Median |  | 40.25 | 41.65 | 37.4789 | 35.485 | 44.1 | 37.6 | 41.7 | 40.4 | 39.9 | 38.75 | 39.75 | 43.5 | 37.45 | 39.701 | 39 | 40.2 | 36.975 | 38 | 39.35 | 34.4 | 43.85 | 36.55 | 37.75 | 43 | 40.365
| Std Dev. |  | 0.49666 | 0.43359 | 1.1207 | 2.0421 | 0.7823 | 0.50465 | 0.8165 | 0.28752 | 0.88938 | 1.3018 | 0.90774 | 1.633 | 0.60249 | 0.75631 | 0.40825 | 1.2785 | 0.7529 | 0.40825 | 0.6473 | 1.5769 | 0.43205 | 0.49666 | 1.9318 | 0.63246 | 0.7999
| Rel.Std.Dev. |  | 1.24% | 1.04% | 2.974% | 5.79% | 1.782% | 1.34% | 1.95% | 0.7105% | 2.237% | 3.332% | 2.275% | 3.74% | 1.605% | 1.901% | 1.042% | 3.159% | 2.048% | 1.07% | 1.649% | 4.527% | 0.9849% | 1.356% | 5.018% | 1.471% | 1.978%
| PDM3 |  | 1.416% | 5.551% | -4.609% | -10.73% | 11.12% | -4.658% | 5.973% | 2.429% | 0.6149% | -1.115% | 0.9946% | 10.53% | -4.954% | 0.7003% | -0.8616% | 2.429% | -6.949% | -3.393% | -0.6507% | -11.83% | 11.04% | -7.274% | -2.549% | 8.841% | 2.387%
| Z-Score (Absolute) |  | 0.11 | 0.67 | 0.71 | 1.55 | 1.44 | 0.72 | 0.73 | 0.25 | 0 | 0.24 | 0.05 | 1.35 | 0.76 | 0.01 | 0.2 | 0.25 | 1.03 | 0.55 | 0.17 | 1.7 | 1.42 | 1.08 | 0.43 | 1.12 | 0.24
| Analytical results for Pd in OREAS 506b (Indicative Value < 10 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab |  | 1
| No | No | 2.01: 2.03 |  | ppb
|  |  | 2.01: AR*OES/MS |  | 1
|  |  | 2.01: 01g |  | 1
| 1 | 1 | 2.01: < 10 |  | 1
| 1 | 2 | 2.01: < 10 |  | 1
| 1 | 3 | 2.01: < 10 |  | 16
| 1 | 4 | 2.01: < 10 |  | < 10
| 1 | 5 | 2.01: < 10 |  | 42
| 1 | 6 | 2.01: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Pr in OREAS 506b (Indicative Value 3.75 ppm) | [27 empty] | Indicative
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.04: 2.25 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS |  | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.04: 0.5g |  | 2
| 1 | 1 | 2.01: 3.404; 2.02: 3.6941; 2.03: 6.3948; 2.04: 3.9 |  | 1
| 1 | 2 | 2.01: 3.516; 2.02: 3.8859; 2.03: 6.438; 2.04: 3.9 |  | 26
| 1 | 3 | 2.01: 3.531; 2.02: 3.789; 2.03: 6.3822; 2.04: 3.9 |  | 16
| 1 | 4 | 2.01: 3.513; 2.02: 4.0446; 2.03: 6.4109; 2.04: 3.9 |  | 3.7527
| 1 | 5 | 2.01: 3.481; 2.02: 3.9042; 2.03: 6.4121; 2.04: 3.7 |  | 43
| 1 | 6 | 2.01: 3.535; 2.02: 4.0509; 2.03: 6.1181; 2.04: 3.9
| Mean |  | 2.01: 3.4967; 2.02: 3.8948; 2.03: 6.3594; 2.04: 3.8667
| Median |  | 2.01: 3.5145; 2.02: 3.895; 2.03: 6.4028; 2.04: 3.9
| Std Dev. |  | 2.01: 0.049237; 2.02: 0.14033; 2.03: 0.11966; 2.04: 0.08165
| Rel.Std.Dev. |  | 2.01: 1.408%; 2.02: 3.603%; 2.03: 1.882%; 2.04: 2.112%
| PDM3 |  | 2.01: -6.823%; 2.02: 3.786%; 2.03: 69.46%; 2.04: 3.037%
| Z-Score (Absolute) |  | 2.01: 1.3; 2.02: 0.05; 2.03: 8.4; 2.04: 0.05
| Analytical results for Pt in OREAS 506b (Indicative Value 1.5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.02: Lab |  | 1
| No | No | 2.01: 2.03; 2.02: 2.20 |  | ppb
|  |  | 2.01: AR*OES/MS; 2.02: AR*OES/MS |  | 2
|  |  | 2.01: 01g; 2.02: 0.1g |  | 2
| 1 | 1 | 2.01: < 5; 2.02: 2 |  | 1
| 1 | 2 | 2.01: < 5; 2.02: 1 |  | 1
| 1 | 3 | 2.01: < 5; 2.02: 1 |  | 16
| 1 | 4 | 2.01: < 5; 2.02: 2 |  | 1.5
| 1 | 5 | 2.01: < 5; 2.02: 2 |  | 39
| 1 | 6 | 2.01: < 5; 2.02: 1
| Mean |  | 2.02: 1.5
| Median |  | 2.02: 1.5
| Std Dev. |  | 2.02: 0.54772
| Rel.Std.Dev. |  | 2.02: 36.51%
PDM3 | bar: 2.02
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.67
| Analytical results for Rb in OREAS 506b (Certified Value 79 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 0
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [8 empty] | 1
| 1 | 1 | 85.96 | 97.6 | 78.7378 | 77.6 | 73.9 | 80.7 | 74.8 | 80.5 | 77.6 | 74.66 | 84.6 | 77.307 | 77.1 | 84.2 | 69.73 | 81 | 91.2 | 67.7 | [8 empty] | 1
| 1 | 2 | 85.55 | 95.7 | 78.4154 | 76.8 | 75.8 | 80.6 | 74.3 | 81.8 | 81.2 | 75.72 | 83.6 | 81.9223 | 77.2 | 77.9 | 73.97 | 84 | 91.1 | 67.9 | [8 empty] | 62
| 1 | 3 | 85.14 | 96.9 | 79.7233 | 78.3 | 74.9 | 81.3 | 73.1 | 80 | 79.6 | 75.97 | 84.7 | 78.3376 | 76.9 | 83 | 73.64 | 80 | 92.6 | 70.9 | [8 empty] | 16
| 1 | 4 | 84 | 96.5 | 79.1479 | 77.9 | 77.1 | 81.6 | 74.7 | 82.1 | 80.7 | 74.37 | 82.7 | 80.4895 | 79.7 | 82.7 | 70.16 | 81 | 91.3 | 69.2 | [8 empty] | 79.069
| 1 | 5 | 86.08 | 97.5 | 78.6082 | 78.5 | 75 | 79.9 | 73.1 | 80.4 | 79.5 | 74.79 | 85.7 | 77.2991 | 77.4 | 83.2 | 71.1 | 81 | 93.4 | 69 | [8 empty] | 102
| 1 | 6 | 85.84 | 96 | 79.6161 | 77 | 74.1 | 81.7 | 71.9 | 78.7 | 75.3 | 75.28 | 84.4 | 78.9227 | 79.1 | 82 | 72.43 | 79 | 90 | 68.6
| Mean |  | 85.4283 | 96.7 | 79.0415 | 77.6833 | 75.1333 | 80.9667 | 73.65 | 80.5833 | 78.9833 | 75.1317 | 84.2833 | 79.0464 | 77.9 | 82.1667 | 71.8383 | 81 | 91.6 | 68.8833
| Median |  | 85.695 | 96.7 | 78.9428 | 77.75 | 74.95 | 81 | 73.7 | 80.45 | 79.55 | 75.035 | 84.5 | 78.6302 | 77.3 | 82.85 | 71.765 | 81 | 91.25 | 68.8
| Std Dev. |  | 0.77652 | 0.77717 | 0.54376 | 0.68532 | 1.181 | 0.69186 | 1.1415 | 1.2416 | 2.1904 | 0.63092 | 1.0265 | 1.8412 | 1.1883 | 2.2097 | 1.7859 | 1.6733 | 1.2083 | 1.1514
| Rel.Std.Dev. |  | 0.909% | 0.8037% | 0.6879% | 0.8822% | 1.572% | 0.8545% | 1.55% | 1.541% | 2.773% | 0.8397% | 1.218% | 2.329% | 1.525% | 2.689% | 2.486% | 2.066% | 1.319% | 1.671%
| PDM3 |  | 8.043% | 22.3% | -0.03483% | -1.752% | -4.977% | 2.4% | -6.853% | 1.915% | -0.1083% | -4.98% | 6.595% | -0.02862% | -1.478% | 3.918% | -9.145% | 2.442% | 15.85% | -12.88%
| Z-Score (Absolute) |  | 1.22 | 3.39 | 0 | 0.26 | 0.75 | 0.37 | 1.03 | 0.3 | 0.01 | 0.75 | 1 | 0 | 0.22 | 0.6 | 1.38 | 0.38 | 2.41 | 1.95
| Analytical results for Re in OREAS 506b (Certified Value 0.035 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.20 | 2.22 | 2.24 | 2.25 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [12 empty] | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.1g | 0.5g | 0.25g | 0.5g | [12 empty] | 3
| 1 | 1 | 0.037 | < 0.05 | < 0.05 | 0.031 | 0.035 | 0.034 | 0.033 | 0.031 | 0.035 | < 0.05 | 0.035 | 0.044 | 0.04 | 0.033 | [12 empty] | 1
| 1 | 2 | 0.037 | < 0.05 | < 0.05 | 0.034 | 0.036 | 0.036 | 0.033 | 0.033 | 0.036 | < 0.05 | 0.033 | 0.045 | 0.03 | 0.034 | [12 empty] | 63
| 1 | 3 | 0.033 | < 0.05 | < 0.05 | 0.032 | 0.038 | 0.038 | 0.036 | 0.037 | 0.036 | < 0.05 | 0.031 | 0.044 | 0.03 | 0.035 | [12 empty] | 16
| 1 | 4 | 0.035 | < 0.05 | < 0.05 | 0.032 | 0.041 | 0.036 | 0.035 | 0.037 | 0.033 | < 0.05 | 0.033 | 0.034 | 0.04 | 0.035 | [12 empty] | 0.034553
| 1 | 5 | 0.038 | < 0.05 | < 0.05 | 0.032 | 0.035 | 0.039 | 0.033 | 0.033 | 0.035 | < 0.05 | 0.033 | 0.045 | 0.04 | 0.033 | [12 empty] | 103
| 1 | 6 | 0.037 | < 0.05 | < 0.05 | 0.032 | 0.038 | 0.035 | 0.033 | 0.03 | 0.033 | < 0.05 | 0.031 | 0.039 | 0.03 | 0.033
| Mean |  | 0.036167 |  |  | 0.032167 | 0.037167 | 0.036333 | 0.033833 | 0.0335 | 0.034667 |  | 0.032667 | 0.041833 | 0.035 | 0.033833
| Median |  | 0.037 |  |  | 0.032 | 0.037 | 0.036 | 0.033 | 0.033 | 0.035 |  | 0.033 | 0.044 | 0.035 | 0.0335
| Std Dev. |  | 0.0018348 |  |  | 0.00098319 | 0.0023166 | 0.0018619 | 0.0013292 | 0.0029496 | 0.0013663 |  | 0.0015055 | 0.004446 | 0.0054772 | 0.00098319
| Rel.Std.Dev. |  | 5.073% |  |  | 3.057% | 6.233% | 5.124% | 3.929% | 8.805% | 3.941% |  | 4.609% | 10.63% | 15.65% | 2.906%
| PDM3 |  | 4.669% |  |  | -6.907% | 7.563% | 5.151% | -2.084% | -3.048% | 0.328% |  | -5.46% | 21.07% | 1.293% | -2.084%
| Z-Score (Absolute) |  | 0.79 | 2.98 | 2.98 | 0.56 | 1.12 | 0.84 | 0 | 0.11 | 0.28 | 2.98 | 0.39 | 2.7 | 0.39 | 0
| Analytical results for S in OREAS 506b (Certified Value 0.614 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 3
| 1 | 1 | 0.6177 | 0.6 | 0.59681 | 0.59 | 0.48 | 0.61 | 0.63 | 0.66 | 0.67 | 0.62 | 0.62 | 0.61 | 0.58 | 0.60543 | 0.66 | 0.62 | 0.62532 | 0.61 | 0.59 | 0.67 | 0.6 | 0.58413 | 0.48 | 0.578 | 0.6069 | 0.6 | 1
| 1 | 2 | 0.6176 | 0.59 | 0.56768 | 0.59 | 0.47 | 0.62 | 0.63 | 0.65 | 0.66 | 0.62 | 0.64 | 0.648 | 0.59 | 0.61818 | 0.64 | 0.6 | 0.62149 | 0.61 | 0.603 | 0.7 | 0.58 | 0.606 | 0.47 | 0.584 | 0.6198 | 0.6 | 7
| 1 | 3 | 0.6077 | 0.6 | 0.56579 | 0.6 | 0.48 | 0.61 | 0.63 | 0.66 | 0.66 | 0.62 | 0.63 | 0.601 | 0.58 | 0.61527 | 0.65 | 0.6 | 0.62906 | 0.61 | 0.586 | 0.67 | 0.6 | 0.59462 | 0.48 | 0.593 | 0.6155 | 0.6 | 16
| 1 | 4 | 0.6098 | 0.59 | 0.57313 | 0.59 | 0.47 | 0.62 | 0.63 | 0.66 | 0.66 | 0.61 | 0.63 | 0.6 | 0.59 | 0.61323 | 0.65 | 0.61 | 0.62541 | 0.6 | 0.676 | 0.64 | 0.59 | 0.60775 | 0.47 | 0.598 | 0.6146 | 0.6 | 0.6144
| 1 | 5 | 0.614 | 0.61 | 0.57906 | 0.6 | 0.48 | 0.62 | 0.63 | 0.65 | 0.65 | 0.63 | 0.63 | 0.601 | 0.57 | 0.61507 | 0.66 | 0.61 | 0.62308 | 0.6 | 0.608 | 0.64 | 0.6 | 0.60513 | 0.47 | 0.595 | 0.6262 | 0.6 | 104
| 1 | 6 | 0.6125 | 0.59 | 0.60137 | 0.58 | 0.47 | 0.61 | 0.64 | 0.64 | 0.67 | 0.61 | 0.62 | 0.59 | 0.59 | 0.61376 | 0.65 | 0.61 | 0.62384 | 0.61 | 0.609 | 0.67 | 0.59 | 0.59337 | 0.49 | 0.578 | 0.6098 | 0.6
| Mean |  | 0.61322 | 0.59667 | 0.58064 | 0.59167 | 0.475 | 0.615 | 0.63167 | 0.65333 | 0.66167 | 0.61833 | 0.62833 | 0.60833 | 0.58333 | 0.61349 | 0.65167 | 0.60833 | 0.6247 | 0.60667 | 0.612 | 0.665 | 0.59333 | 0.5985 | 0.47667 | 0.58767 | 0.61547 | 0.6
| Median |  | 0.61325 | 0.595 | 0.57609 | 0.59 | 0.475 | 0.615 | 0.63 | 0.655 | 0.66 | 0.62 | 0.63 | 0.601 | 0.585 | 0.61442 | 0.65 | 0.61 | 0.62458 | 0.61 | 0.6055 | 0.67 | 0.595 | 0.59988 | 0.475 | 0.5885 | 0.61505 | 0.6
| Std Dev. |  | 0.0040632 | 0.008165 | 0.015091 | 0.0075277 | 0.0054772 | 0.0054772 | 0.0040825 | 0.008165 | 0.0075277 | 0.0075277 | 0.0075277 | 0.020442 | 0.008165 | 0.0043078 | 0.0075277 | 0.0075277 | 0.0025912 | 0.005164 | 0.032748 | 0.022583 | 0.008165 | 0.0093126 | 0.008165 | 0.0088242 | 0.0069316 | 0
| Rel.Std.Dev. |  | 0.6626% | 1.368% | 2.599% | 1.272% | 1.153% | 0.8906% | 0.6463% | 1.25% | 1.138% | 1.217% | 1.198% | 3.36% | 1.4% | 0.7022% | 1.155% | 1.237% | 0.4148% | 0.8512% | 5.351% | 3.396% | 1.376% | 1.556% | 1.713% | 1.502% | 1.126% | 0%
| PDM3 |  | -0.1932% | -2.887% | -5.496% | -3.701% | -22.69% | 0.09708% | 2.81% | 6.336% | 7.693% | 0.6396% | 2.267% | -0.988% | -5.057% | -0.1485% | 6.065% | -0.988% | 1.676% | -1.259% | -0.3912% | 8.235% | -3.429% | -2.588% | -22.42% | -4.352% | 0.173% | -2.344%
| Z-Score (Absolute) |  | 0.05 | 0.63 | 1.29 | 0.84 | 5.66 | 0.12 | 0.81 | 1.71 | 2.05 | 0.26 | 0.67 | 0.15 | 1.18 | 0.06 | 1.64 | 0.15 | 0.52 | 0.22 | 0 | 2.19 | 0.77 | 0.56 | 5.59 | 1 | 0.14 | NA
Results from laboratory 2.27 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Sb in OREAS 506b (Certified Value 0.87 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  | 2
| 1 | 1 | 1.11 | 0.7 | 0.80496 | 0.9 | < 5 | 0.93 | 0.99 | 1 | 0.86 | 0.92 | 0.6 | < 2 | 0.76 | 0.87307 | < 5 | 0.84 | < 5 | 0.93 | 0.71 | 0.47 | 1 | 0.66 | 0.99 | 0.91 | 0.9 |  | 1
| 1 | 2 | 1.09 | 0.71 | 0.78585 | 0.96 | < 5 | 0.93 | 0.98 | 1.02 | 0.84 | 0.97 | 0.59 | < 2 | 0.73 | 0.89722 | < 5 | 0.88 | < 5 | 0.93 | 0.69 | 0.43 | 1 | 0.75 | 0.95 | 1.06 | 0.88 |  | 28
| 1 | 3 | 1.13 | 0.73 | 0.78535 | 0.88 | < 5 | 0.95 | 0.97 | 1.01 | 0.88 | 1 | 0.58 | < 2 | 0.77 | 0.93933 | < 5 | 0.77 | < 5 | 0.92 | 0.75 | 0.5 | 1 | 0.71 | 1.01 | 0.88 | 0.89 |  | 16
| 1 | 4 | 1.07 | 0.75 | 0.74159 | 0.83 | < 5 | 0.99 | 0.98 | 1.03 | 0.92 | 0.95 | 0.55 | < 2 | 0.72 | 0.9381 | < 5 | 0.89 | < 5 | 0.91 | 0.68 | 0.5 | 1 | 0.68 | 1.02 | 0.89 | 0.88 |  | 0.86527
| 1 | 5 | 1.11 | 0.71 | 0.79413 | 0.94 | < 5 | 0.92 | 0.95 | 1.07 | 0.9 | 0.97 | 0.55 | < 2 | 0.74 | 0.94758 | < 5 | 0.8 | < 5 | 0.91 | 0.69 | 0.48 | 1 | 0.64 | 1.01 | 0.88 | 0.9 |  | 105
| 1 | 6 | 1.14 | 0.73 | 0.72952 | 0.94 | < 5 | 0.9 | 0.97 | 1.03 | 0.84 | 0.97 | 0.57 | < 2 | 0.72 | 0.88053 | < 5 | 0.82 | < 5 | 0.91 | 0.69 | 0.48 | 1.1 | 0.78 | 1 | 0.87 | 0.89
| Mean |  | 1.1083 | 0.72167 | 0.77357 | 0.90833 |  | 0.93667 | 0.97333 | 1.0267 | 0.87333 | 0.96333 | 0.57333 |  | 0.74 | 0.91264 |  | 0.83333 |  | 0.91833 | 0.70167 | 0.47667 | 1.0167 | 0.70333 | 0.99667 | 0.915 | 0.89
| Median |  | 1.11 | 0.72 | 0.7856 | 0.92 |  | 0.93 | 0.975 | 1.025 | 0.87 | 0.97 | 0.575 |  | 0.735 | 0.91766 |  | 0.83 |  | 0.915 | 0.69 | 0.48 | 1 | 0.695 | 1.005 | 0.885 | 0.89
| Std Dev. |  | 0.025626 | 0.018348 | 0.030532 | 0.048339 |  | 0.030768 | 0.013663 | 0.024221 | 0.03266 | 0.026583 | 0.020656 |  | 0.020976 | 0.032912 |  | 0.046332 |  | 0.0098319 | 0.025626 | 0.02582 | 0.040825 | 0.053914 | 0.025033 | 0.072319 | 0.0089443
| Rel.Std.Dev. |  | 2.312% | 2.543% | 3.947% | 5.322% |  | 3.285% | 1.404% | 2.359% | 3.74% | 2.76% | 3.603% |  | 2.835% | 3.606% |  | 5.56% |  | 1.071% | 3.652% | 5.417% | 4.016% | 7.665% | 2.512% | 7.904% | 1.005%
| PDM3 |  | 28.09% | -16.6% | -10.6% | 4.976% |  | 8.251% | 12.49% | 18.65% | 0.9314% | 11.33% | -33.74% |  | -14.48% | 5.474% |  | -3.691% |  | 6.132% | -18.91% | -44.91% | 17.5% | -18.72% | 15.19% | 5.747% | 2.858%
| Z-Score (Absolute) |  | 1.32 | 1.3 | 0.95 | 0.04 | 10.75 | 0.15 | 0.4 | 0.76 | 0.27 | 0.34 | 2.31 | 0.58 | 1.18 | 0.01 | 10.75 | 0.55 | 10.75 | 0.03 | 1.44 | 2.96 | NA | 1.43 | 0.56 | 0.01 | 0.16
Results from laboratory 2.23 were removed due to their 0.001 ppm reading resolution.
| Analytical results for Sc in OREAS 506b (Certified Value 6.68 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 6.4 | 7 | 6.4809 | 6.1 | < 1 | 6.4 | 6.9 | 7 | 6.6 | 6.6 | 6.6 | 7 | 7.2 | 6.5263 | 6.6 | 6.8 | 7 | 7 | 6.4 | 7.23 | 4.9 | 6.1 | 6.3 |  |  |  | 1
| 1 | 2 | 6.4 | 7 | 6.7041 | 6.4 | < 1 | 6.4 | 7.1 | 7 | 6.7 | 6.8 | 6.9 | 7.4 | 7.2 | 6.8516 | 5.8 | 6.6 | 7 | 7 | 7 | 7.41 | 4.9 | 6.5 | 6.6 |  |  |  | 29
| 1 | 3 | 6.3 | 7 | 6.8395 | 5.9 | < 1 | 6.2 | 7 | 6.9 | 6.6 | 6.7 | 6.7 | 7 | 7.1 | 6.7459 | 6.4 | 6.7 | 7 | 7 | 6.4 | 7.15 | 5.1 | 6.6 | 6.4 |  |  |  | 16
| 1 | 4 | 6.3 | 7 | 6.5356 | 5.8 | < 1 | 6.5 | 6.9 | 7.1 | 6.8 | 6.7 | 6.6 | 7 | 7.1 | 6.797 | 6 | 7.1 | 7 | 9 | 6.9 | 7.15 | 4.8 | 6.6 | 6.4 |  |  |  | 6.6782
| 1 | 5 | 6.4 | 7 | 6.6722 | 6.4 | < 1 | 6.3 | 6.8 | 7 | 6.6 | 6.7 | 6.7 | 6.6 | 7.3 | 6.5816 | 6.6 | 6.7 | 7 | 7 | 6.5 | 7.41 | 4.9 | 6.6 | 6.3 |  |  |  | 106
| 1 | 6 | 6.3 | 7 | 6.8414 | 6.2 | < 1 | 6.3 | 7 | 7 | 6.4 | 6.2 | 6.7 | 7 | 7.2 | 6.8295 | 6.3 | 6.9 | 7 | 7 | 6.3 | 7.15 | 4.9 | 6.4 | 6.2
| Mean |  | 6.35 | 7 | 6.6789 | 6.1333 |  | 6.35 | 6.95 | 7 | 6.6167 | 6.6167 | 6.7 | 7 | 7.1833 | 6.722 | 6.2833 | 6.8 | 7 | 7.3333 | 6.5833 | 7.25 | 4.9167 | 6.4667 | 6.3667
| Median |  | 6.35 | 7 | 6.6881 | 6.15 |  | 6.35 | 6.95 | 7 | 6.6 | 6.7 | 6.7 | 7 | 7.2 | 6.7715 | 6.35 | 6.75 | 7 | 7 | 6.45 | 7.19 | 4.9 | 6.55 | 6.35
| Std Dev. |  | 0.054772 | 0 | 0.15007 | 0.25033 |  | 0.10488 | 0.10488 | 0.063246 | 0.13292 | 0.2137 | 0.10954 | 0.25298 | 0.075277 | 0.13605 | 0.32506 | 0.17889 | 0 | 0.8165 | 0.29269 | 0.12775 | 0.098319 | 0.19664 | 0.13663
| Rel.Std.Dev. |  | 0.8626% | 0% | 2.247% | 4.082% |  | 1.652% | 1.509% | 0.9035% | 2.009% | 3.23% | 1.635% | 3.614% | 1.048% | 2.024% | 5.173% | 2.631% | 0% | 11.13% | 4.446% | 1.762% | 2% | 3.041% | 2.146%
| PDM3 |  | -4.914% | 4.819% | 0.01107% | -8.159% |  | -4.914% | 4.07% | 4.819% | -0.9214% | -0.9214% | 0.3264% | 4.819% | 7.564% | 0.6557% | -5.913% | 1.824% | 4.819% | 9.81% | -1.421% | 8.562% | -26.38% | -3.168% | -4.665%
| Z-Score (Absolute) |  | 0.67 | NA | 0.16 | 1.22 | 15.47 | 0.67 | 0.84 | 0.97 | 0 | 0 | 0.21 | 0.97 | 1.43 | 0.27 | 0.84 | 0.46 | NA | NA | 0.08 | 1.6 | 4.3 | 0.38 | 0.63
Results from laboratories 2.02, 2.19 and 2.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 506b (Certified Value 4.89 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  |  |  | 3
| 1 | 1 | 5 | 5 | 4.409 | 4.7 | < 5 | 5 | 4.5 | 5.5 | 4.9 | 4.9 | 5 | 4.6367 | 4.9 | < 10 | 5.12 | 4 | < 5 | 4.6 | 4 | 4.7 | 5.14 | 4.9 |  |  |  |  | 1
| 1 | 2 | 5 | 5 | 4.4121 | 4.4 | < 5 | 4.9 | 5.1 | 5.4 | 4.9 | 4.6 | 6 | 4.4773 | 4.4 | < 10 | 4.9 | 4 | < 5 | 5 | 3 | 4.5 | 5.36 | 4.8 |  |  |  |  | 0
| 1 | 3 | 5 | 5 | 4.3771 | 3.8 | < 5 | 4.6 | 4.7 | 5.4 | 5.3 | 4.6 | 5 | 4.4358 | 4.9 | < 10 | 4.9 | 4 | < 5 | 4.8 | 3 | 4.7 | 5.3 | 4.9 |  |  |  |  | 16
| 1 | 4 | 5 | 5 | 4.5949 | 3.9 | < 5 | 4.8 | 4.6 | 5.4 | 5.5 | 5.1 | 5 | 4.3742 | 5.1 | < 10 | 5.05 | 5 | < 5 | 5 | 4 | 4.7 | 5.29 | 4.8 |  |  |  |  | 4.8883
| 1 | 5 | 5 | 5 | 4.4274 | 4 | < 5 | 5 | 4.3 | 5.4 | 4.9 | 5.1 | 5 | 4.3178 | 4.9 | < 10 | 5.22 | 4 | < 5 | 5 | 4 | 5 | 5.44 | 4.7 |  |  |  |  | 107
| 1 | 6 | 4 | 5 | 4.5722 | 4.1 | < 5 | 5.1 | 4.8 | 5.4 | 5.1 | 4.6 | 5 | 4.2934 | 4.9 | < 10 | 5.34 | 4 | < 5 | 5.1 | 3 | 4.6 | 5.21 | 4.8
| Mean |  | 4.8333 | 5 | 4.4654 | 4.15 |  | 4.9 | 4.6667 | 5.4167 | 5.1 | 4.8167 | 5.1667 | 4.4225 | 4.85 |  | 5.0883 | 4.1667 |  | 4.9167 | 3.5 | 4.7 | 5.29 | 4.8167
| Median |  | 5 | 5 | 4.4198 | 4.05 |  | 4.95 | 4.65 | 5.4 | 5 | 4.75 | 5 | 4.405 | 4.9 |  | 5.085 | 4 |  | 5 | 3.5 | 4.7 | 5.295 | 4.8
| Std Dev. |  | 0.40825 | 0 | 0.093221 | 0.33912 |  | 0.17889 | 0.27325 | 0.040825 | 0.25298 | 0.24833 | 0.40825 | 0.1257 | 0.23452 |  | 0.17555 | 0.40825 |  | 0.18348 | 0.54772 | 0.16733 | 0.10621 | 0.075277
| Rel.Std.Dev. |  | 8.447% | 0% | 2.088% | 8.171% |  | 3.651% | 5.855% | 0.7537% | 4.96% | 5.156% | 7.902% | 2.842% | 4.835% |  | 3.45% | 9.798% |  | 3.732% | 15.65% | 3.56% | 2.008% | 1.563%
| PDM3 |  | -1.125% | 2.285% | -8.651% | -15.1% |  | 0.2392% | -4.534% | 10.81% | 4.331% | -1.466% | 5.694% | -9.528% | -0.7837% |  | 4.092% | -14.76% |  | 0.5801% | -28.4% | -3.852% | 8.217% | -1.466%
| Z-Score (Absolute) |  | 0 | NA | 1.49 | 2.76 | 9.44 | 0.27 | 0.67 | 2.36 | 1.08 | 0.07 | NA | 1.66 | 0.07 | 0.67 | 1.03 | NA | NA | 0.34 | NA | 0.54 | 1.85 | 0.07
Results from laboratories 2.02, 2.14, 2.21, 2.22 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 506b (Indicative Value 3.37 ppm) | [27 empty] | Indicative
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab |  |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.18; 2.04: 2.25 |  |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*OES/MS |  |  |  |  | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g; 2.04: 0.5g |  |  |  |  | 2
| 1 | 1 | 2.01: 3.222; 2.02: 3.1724; 2.03: 5.5919; 2.04: 3.5 |  |  |  |  | 1
| 1 | 2 | 2.01: 3.258; 2.02: 3.4238; 2.03: 5.8825; 2.04: 3.8 |  |  |  |  | 12
| 1 | 3 | 2.01: 3.233; 2.02: 3.266; 2.03: 5.6297; 2.04: 3.7 |  |  |  |  | 16
| 1 | 4 | 2.01: 3.219; 2.02: 3.2957; 2.03: 5.3731; 2.04: 3.4 |  |  |  |  | 3.3744
| 1 | 5 | 2.01: 3.23; 2.02: 3.37; 2.03: 5.4755; 2.04: 3.4 |  |  |  |  | 40
| 1 | 6 | 2.01: 3.191; 2.02: 3.4578; 2.03: 5.0967; 2.04: 3.6
| Mean |  | 2.01: 3.2255; 2.02: 3.331; 2.03: 5.5082; 2.04: 3.5667
| Median |  | 2.01: 3.226; 2.02: 3.3329; 2.03: 5.5337; 2.04: 3.55
| Std Dev. |  | 2.01: 0.021806; 2.02: 0.1066; 2.03: 0.26477; 2.04: 0.1633
| Rel.Std.Dev. |  | 2.01: 0.676%; 2.02: 3.2%; 2.03: 4.807%; 2.04: 4.578%
| PDM3 |  | 2.01: -4.412%; 2.02: -1.287%; 2.03: 63.24%; 2.04: 5.699%
| Z-Score (Absolute) |  | 2.01: 0.88; 2.02: 0.47; 2.03: 8.14; 2.04: 0.47
| Analytical results for Sn in OREAS 506b (Certified Value 2.95 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 3.17 | 3.1 | 2.9389 | < 5 | 2.9 | 2.9 | 3.2 | 3.1 | 2.8 | 3 | 3 | 2.8469 | < 10 | 2.9 | < 20 | 2.9 | 2.9 | 1.9 | 3 | 2.8 | [6 empty] | 1
| 1 | 2 | 3.14 | 3.2 | 2.7569 | < 5 | 2.9 | 2.9 | 3.1 | 3.1 | 3 | 3 | 3 | 2.8132 | < 10 | 2.9 | < 20 | 2.9 | 3 | 1.6 | 3.1 | 2.8 | [6 empty] | 13
| 1 | 3 | 3.1 | 3.1 | 2.9587 | < 5 | 2.8 | 2.9 | 3.2 | 3.1 | 2.9 | 3 | 3 | 2.9177 | < 10 | 3 | < 20 | 3 | 3 | 1.8 | 3 | 2.9 | [6 empty] | 16
| 1 | 4 | 3.03 | 3.2 | 2.819 | < 5 | 3 | 2.9 | 3.3 | 3 | 2.9 | 2.9 | 2.8 | 2.891 | < 10 | 3.1 | < 20 | 2.9 | 2.8 | 1.8 | 3 | 2.85 | [6 empty] | 2.9487
| 1 | 5 | 3.11 | 3.1 | 2.8734 | < 5 | 2.8 | 2.9 | 3.2 | 3 | 2.9 | 2.9 | 2.9 | 2.9153 | < 10 | 3 | < 20 | 2.9 | 3 | 1.8 | 3 | 2.92 | [6 empty] | 108
| 1 | 6 | 3.08 | 3.2 | 2.8096 | < 5 | 2.8 | 2.9 | 3.1 | 2.9 | 2.7 | 2.9 | 2.9 | 2.9322 | < 10 | 2.9 | < 20 | 2.9 | 3 | 1.7 | 2.9 | 2.81
| Mean |  | 3.105 | 3.15 | 2.8594 |  | 2.8667 | 2.9 | 3.1833 | 3.0333 | 2.8667 | 2.95 | 2.9333 | 2.8861 |  | 2.9667 |  | 2.9167 | 2.95 | 1.7667 | 3 | 2.8467
| Median |  | 3.105 | 3.15 | 2.8462 |  | 2.85 | 2.9 | 3.2 | 3.05 | 2.9 | 2.95 | 2.95 | 2.9032 |  | 2.95 |  | 2.9 | 3 | 1.8 | 3 | 2.83
| Std Dev. |  | 0.048477 | 0.054772 | 0.078733 |  | 0.08165 | 0 | 0.075277 | 0.08165 | 0.10328 | 0.054772 | 0.08165 | 0.04657 |  | 0.08165 |  | 0.040825 | 0.083666 | 0.10328 | 0.063246 | 0.052789
| Rel.Std.Dev. |  | 1.561% | 1.739% | 2.753% |  | 2.848% | 0% | 2.365% | 2.692% | 3.603% | 1.857% | 2.784% | 1.614% |  | 2.752% |  | 1.4% | 2.836% | 5.846% | 2.108% | 1.854%
| PDM3 |  | 5.301% | 6.827% | -3.028% |  | -2.782% | -1.652% | 7.957% | 2.87% | -2.782% | 0.04413% | -0.5211% | -2.124% |  | 0.6093% |  | -1.086% | 0.04413% | -40.09% | 1.74% | -3.46%
| Z-Score (Absolute) |  | 1.4 | 1.79 | 0.71 | 3.79 | 0.64 | 0.36 | 2.07 | 0.79 | 0.64 | 0.07 | 0.07 | 0.48 | 17.65 | 0.21 | 60.54 | 0.21 | 0.07 | 10.08 | 0.5 | 0.81
| Analytical results for Sr in OREAS 506b (Certified Value 94 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 1
| 1 | 1 | 98.04 | 97.5 | 86.8694 | 90.7 | 99.4 | 96.4 | 92.4 | 95.6 | 97.7 | 90.4 | 91.4 | 109 | 98.7 | 92.6555 | 91.2 | 94 | 70.5 | 96.1 | 92 | 96 | 88.9 | 94.8 | 91 |  |  |  | 1
| 1 | 2 | 98.68 | 97.2 | 88.0802 | 94.2 | 98.2 | 96.5 | 92.1 | 94.1 | 97.5 | 92.3 | 93 | 115 | 96 | 97.4736 | 88.7 | 94 | 69.5 | 99.7 | 92 | 99 | 89.4 | 95.2 | 91.5 |  |  |  | 14
| 1 | 3 | 98.36 | 98.8 | 88.3903 | 95.2 | 99.7 | 96.3 | 91.6 | 94.1 | 96.2 | 91.1 | 94 | 108 | 96.2 | 93.8354 | 90.3 | 94 | 71.7 | 97.6 | 89 | 99 | 88 | 96.3 | 90.5 |  |  |  | 16
| 1 | 4 | 97.9 | 98.8 | 86.2123 | 93.7 | 97.5 | 99.5 | 91.4 | 95.2 | 97.8 | 91.9 | 90.5 | 103 | 94.4 | 94.7082 | 90.2 | 94 | 72.3 | 90.8 | 91 | 101 | 87 | 96.4 | 91.2 |  |  |  | 94.1065
| 1 | 5 | 96.16 | 99.4 | 86.1823 | 97 | 99.3 | 95.7 | 90.9 | 93.1 | 95.3 | 91.3 | 90.3 | 105 | 97.1 | 93.4218 | 92.2 | 93 | 74.1 | 97.3 | 91 | 103 | 87.4 | 97.3 | 89.6 |  |  |  | 109
| 1 | 6 | 96.92 | 97.2 | 87.9241 | 92.4 | 99.4 | 96 | 93.1 | 92.7 | 94.5 | 90.7 | 91.3 | 105 | 97.5 | 94.4117 | 90.5 | 95 | 71 | 98.7 | 90 | 104 | 88.7 | 95.3 | 88.9
| Mean |  | 97.6767 | 98.15 | 87.2764 | 93.8667 | 98.9167 | 96.7333 | 91.9167 | 94.1333 | 96.5 | 91.2833 | 91.75 | 107.5 | 96.65 | 94.4177 | 90.5167 | 94 | 71.5167 | 96.7 | 90.8333 | 100.3333 | 88.2333 | 95.8833 | 90.45
| Median |  | 97.97 | 98.15 | 87.3968 | 93.95 | 99.35 | 96.35 | 91.85 | 94.1 | 96.85 | 91.2 | 91.35 | 106.5 | 96.65 | 94.1236 | 90.4 | 94 | 71.35 | 97.45 | 91 | 100 | 88.35 | 95.8 | 90.75
| Std Dev. |  | 0.95167 | 0.96281 | 0.97994 | 2.1869 | 0.86583 | 1.3866 | 0.78337 | 1.1325 | 1.39 | 0.71671 | 1.4571 | 4.2778 | 1.4707 | 1.6649 | 1.1618 | 0.63246 | 1.593 | 3.1413 | 1.169 | 2.9439 | 0.92664 | 0.9411 | 1.0095
| Rel.Std.Dev. |  | 0.9743% | 0.981% | 1.123% | 2.33% | 0.8753% | 1.433% | 0.8523% | 1.203% | 1.44% | 0.7851% | 1.588% | 3.979% | 1.522% | 1.763% | 1.283% | 0.6728% | 2.227% | 3.249% | 1.287% | 2.934% | 1.05% | 0.9815% | 1.116%
| PDM3 |  | 3.794% | 4.297% | -7.258% | -0.2548% | 5.111% | 2.791% | -2.327% | 0.02852% | 2.543% | -3% | -2.504% | 14.23% | 2.703% | 0.3307% | -3.815% | -0.1132% | -24% | 2.756% | -3.478% | 6.617% | -6.241% | 1.888% | -3.885%
| Z-Score (Absolute) |  | 0.84 | 0.95 | 1.62 | 0.06 | 1.13 | 0.62 | 0.52 | 0 | 0.56 | 0.67 | 0.56 | 3.16 | 0.6 | 0.07 | 0.86 | 0.03 | 5.35 | 0.61 | 0.78 | 1.47 | 1.4 | 0.41 | 0.87
| Analytical results for Ta in OREAS 506b (Certified Value < 0.01 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.19 | 2.20 | 2.22 | 2.24 | 2.25 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [11 empty] | 3
|  |  | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | [11 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | [11 empty] | 15
| 1 | 3 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | [11 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | [11 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | [11 empty] | 110
| 1 | 6 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 10 | < 0.01 | < 0.01 | < 0.05 | < 0.05
| Mean |  | 2.07: 0.01
| Median |  | 2.07: 0.01
| Std Dev. |  | 2.07: 0
Rel.Std.Dev. | bar: 2.07
PDM3
| Z-Score (Absolute) |  | 7.42 | 7.42 | 1008.77 | 0.67 | 0.67 | 0 | 0.67 | 0.67 | 7.42 | 0.67 | 2020.23 | 0.67 | 0.67 | 7.42 | 7.42
| Analytical results for Tb in OREAS 506b (Certified Value 0.45 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.06; 2.03: 2.14; 2.04: 2.15; 2.06: 2.17; 2.07: 2.18; 2.08: 2.21; 2.09: 2.24; 2.10: 2.25 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*OES; 2.03: AR*OES/MS; 2.04: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*OES/MS; 2.10: AR*OES/MS |  |  |  | 2
|  |  | 2.01: 01g; 2.02: 0.25g; 2.03: 0.25g; 2.04: 0.2g; 2.06: 0.2g; 2.07: 0.5g; 2.08: 0.25g; 2.09: 0.25g; 2.10: 0.5g |  |  |  | 3
| 1 | 1 | 2.01: 0.425; 2.02: < 5; 2.03: 0.43; 2.04: 0.4563; 2.06: 0.4; 2.07: 0.6964; 2.08: 0.5; 2.09: 0.45; 2.10: 0.4 |  |  |  | 1
| 1 | 2 | 2.01: 0.441; 2.02: < 5; 2.03: 0.43; 2.04: 0.45849; 2.06: 0.4; 2.07: 0.6035; 2.08: 0.48; 2.09: 0.44; 2.10: 0.4 |  |  |  | 16
| 1 | 3 | 2.01: 0.439; 2.02: < 5; 2.03: 0.44; 2.04: 0.44858; 2.06: 0.4; 2.07: 0.71165; 2.08: 0.5; 2.09: 0.45; 2.10: 0.4 |  |  |  | 16
| 1 | 4 | 2.01: 0.435; 2.02: < 5; 2.03: 0.42; 2.04: 0.4784; 2.06: 0.4; 2.07: 0.65722; 2.08: 0.47; 2.09: 0.44; 2.10: 0.4 |  |  |  | 0.45434
| 1 | 5 | 2.01: 0.443; 2.02: < 5; 2.03: 0.44; 2.04: 0.46573; 2.06: 0.4; 2.07: 0.66179; 2.08: 0.5; 2.09: 0.44; 2.10: 0.4 |  |  |  | 111
| 1 | 6 | 2.01: 0.447; 2.02: < 5; 2.03: 0.42; 2.04: 0.49284; 2.06: 0.5; 2.07: 0.62903; 2.08: 0.51; 2.09: 0.44; 2.10: 0.4
| Mean |  | 2.01: 0.43833; 2.03: 0.43; 2.04: 0.46672; 2.06: 0.41667; 2.07: 0.65993; 2.08: 0.49333; 2.09: 0.44333; 2.10: 0.4
| Median |  | 2.01: 0.44; 2.03: 0.43; 2.04: 0.46211; 2.06: 0.4; 2.07: 0.65951; 2.08: 0.5; 2.09: 0.44; 2.10: 0.4
| Std Dev. |  | 2.01: 0.0076594; 2.03: 0.0089443; 2.04: 0.016278; 2.06: 0.040825; 2.07: 0.040385; 2.08: 0.015055; 2.09: 0.005164; 2.10: 6.0809e-17
| Rel.Std.Dev. |  | 2.01: 1.747%; 2.03: 2.08%; 2.04: 3.488%; 2.06: 9.798%; 2.07: 6.12%; 2.08: 3.052%; 2.09: 1.165%; 2.10: 1.52e-14%
| PDM3 |  | 2.01: -3.524%; 2.03: -5.358%; 2.04: 2.725%; 2.06: -8.293%; 2.07: 45.25%; 2.08: 8.581%; 2.09: -2.424%; 2.10: -11.96%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 48.29; 2.03: 0.87; 2.04: 0; 2.06: NA; 2.07: 4.59; 2.08: 0.63; 2.09: 0.56; 2.10: NA
Results from laboratories 2.17 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 506b (Certified Value 0.2 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.15 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | [5 empty] | 2
| 1 | 1 | 0.2 | 0.2 | 0.18381 | 0.16 | < 5 | 0.18 | 0.21 | 0.21 | 0.21 | 0.15 | 0.25 | < 0.2 | < 10 | 0.22 | 0.28 | < 5 | 0.62 | 0.18 | 0.17 | 0.2 | 0.2 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.2 | 0.19159 | 0.15 | < 5 | 0.19 | 0.22 | 0.15 | 0.22 | 0.2 | 0.22 | < 0.2 | < 10 | 0.2 | 0.23 | < 5 | 0.54 | 0.18 | 0.21 | 0.2 | 0.21 | [5 empty] | 55
| 1 | 3 | 0.2 | 0.2 | 0.19098 | 0.14 | < 5 | 0.19 | 0.21 | 0.2 | 0.19 | 0.2 | 0.21 | < 0.2 | < 10 | 0.19 | 0.24 | < 5 | 0.6 | 0.19 | 0.21 | 0.19 | 0.21 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.2 | 0.167 | 0.15 | < 5 | 0.22 | 0.18 | 0.21 | 0.18 | 0.19 | 0.18 | < 0.2 | < 10 | 0.19 | 0.38 | < 5 | 0.56 | 0.17 | 0.2 | 0.23 | 0.22 | [5 empty] | 0.19949
| 1 | 5 | 0.2 | 0.2 | 0.17115 | 0.17 | < 5 | 0.17 | 0.17 | 0.19 | 0.2 | 0.21 | 0.24 | < 0.2 | < 10 | 0.2 | 0.22 | 5 | 0.61 | 0.19 | 0.24 | 0.18 | 0.22 | [5 empty] | 112
| 1 | 6 | 0.3 | 0.2 | 0.1903 | 0.15 | < 5 | 0.19 | 0.2 | 0.19 | 0.19 | 0.16 | 0.22 | < 0.2 | < 10 | 0.2 | 0.22 | < 5 | 0.63 | 0.2 | 0.19 | 0.21 | 0.2
| Mean |  | 0.21667 | 0.2 | 0.18247 | 0.15333 |  | 0.19 | 0.19833 | 0.19167 | 0.19833 | 0.185 | 0.22 |  |  | 0.2 | 0.26167 | 5 | 0.59333 | 0.185 | 0.20333 | 0.20167 | 0.21
| Median |  | 0.2 | 0.2 | 0.18706 | 0.15 |  | 0.19 | 0.205 | 0.195 | 0.195 | 0.195 | 0.22 |  |  | 0.2 | 0.235 | 5 | 0.605 | 0.185 | 0.205 | 0.2 | 0.21
| Std Dev. |  | 0.040825 | 3.0405e-17 | 0.010826 | 0.010328 |  | 0.016733 | 0.019408 | 0.022286 | 0.01472 | 0.02429 | 0.024495 |  |  | 0.010954 | 0.062102 |  | 0.03559 | 0.010488 | 0.023381 | 0.017224 | 0.0089443
| Rel.Std.Dev. |  | 18.84% | 1.52e-14% | 5.933% | 6.736% |  | 8.807% | 9.785% | 11.63% | 7.422% | 13.13% | 11.13% |  |  | 5.477% | 23.73% |  | 5.998% | 5.669% | 11.5% | 8.541% | 4.259%
| PDM3 |  | 8.613% | 0.2581% | -8.528% | -23.14% |  | -4.755% | -0.5774% | -3.919% | -0.5774% | -7.261% | 10.28% |  |  | 0.2581% | 31.17% | 2406% | 197.4% | -7.261% | 1.929% | 1.094% | 5.271%
| Z-Score (Absolute) |  | NA | NA | 0.79 | 2.18 | 109.52 | 0.44 | 0.04 | 0.36 | 0.04 | 0.67 | 0.99 | 4.72 | 228.51 | 0.04 | 2.97 | NA | 18.76 | 0.67 | 0.2 | 0.12 | 0.52
Results from laboratory 2.01 and 2.02 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.22 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 506b (Certified Value 5.67 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 3
| 1 | 1 | 5.58 | 9.7 | 6.5348 | 5.8 | 5.2 | 5.2 | 5 | 6.6 | 5.9 | 5.2 | 6.74 | < 20 | 5.5 | 5.9055 | 10.442 | < 10 | 11 | 7.77 | 4.8 | 6 | 5.5 | 5.7 | 4.6 |  |  |  | 1
| 1 | 2 | 5.6 | 9.88 | 6.2708 | 6.1 | 5.2 | 5.4 | 5.2 | 6.8 | 6 | 5.5 | 6.47 | < 20 | 5.4 | 6.2113 | 9.9552 | < 10 | 10 | 7.71 | 4.7 | 5.8 | 5.7 | 5.8 | 5.3 |  |  |  | 17
| 1 | 3 | 5.58 | 9.96 | 6.4025 | 5.6 | 6.1 | 5.3 | 5.2 | 6.2 | 6 | 5.5 | 6.43 | < 20 | 5.5 | 5.8687 | 10.5412 | < 10 | 11.6 | 7.72 | 5.1 | 5.8 | 5.8 | 5.8 | 5.1 |  |  |  | 16
| 1 | 4 | 5.66 | 9.69 | 6.3636 | 5 | 5.4 | 5.3 | 5 | 6.3 | 6.2 | 5.4 | 6.12 | < 20 | 5.5 | 6.0482 | 10.3473 | < 10 | 11.8 | 7.68 | 4.7 | 6 | 5.7 | 5.7 | 4.8 |  |  |  | 5.665
| 1 | 5 | 5.58 | 9.74 | 6.372 | 6.1 | 5.5 | 5.1 | 4.9 | 6.5 | 5.9 | 5.5 | 6.32 | < 20 | 5.6 | 6.1403 | 10.4056 | < 10 | 11.7 | 7.45 | 4.9 | 5.9 | 5.5 | 5.5 | 4.8 |  |  |  | 113
| 1 | 6 | 5.63 | 9.75 | 6.4998 | 5.7 | 5.4 | 5.2 | 5.4 | 6.7 | 5.6 | 5.4 | 6.12 | < 20 | 5.5 | 6.1455 | 9.9945 | < 10 | 11.1 | 7.94 | 5.3 | 5.9 | 5.5 | 6 | 4.8
| Mean |  | 5.605 | 9.7867 | 6.4072 | 5.7167 | 5.4667 | 5.25 | 5.1167 | 6.5167 | 5.9333 | 5.4167 | 6.3667 |  | 5.5 | 6.0532 | 10.281 |  | 11.2 | 7.7117 | 4.9167 | 5.9 | 5.6167 | 5.75 | 4.9
| Median |  | 5.59 | 9.745 | 6.3873 | 5.75 | 5.4 | 5.25 | 5.1 | 6.55 | 5.95 | 5.45 | 6.375 |  | 5.5 | 6.0942 | 10.3765 |  | 11.35 | 7.715 | 4.85 | 5.9 | 5.6 | 5.75 | 4.8
| Std Dev. |  | 0.033317 | 0.10875 | 0.096574 | 0.40702 | 0.33267 | 0.10488 | 0.18348 | 0.23166 | 0.19664 | 0.1169 | 0.23577 |  | 0.063246 | 0.13926 | 0.24569 |  | 0.67231 | 0.15817 | 0.24014 | 0.089443 | 0.13292 | 0.16432 | 0.25298
| Rel.Std.Dev. |  | 0.5944% | 1.111% | 1.507% | 7.12% | 6.085% | 1.998% | 3.586% | 3.555% | 3.314% | 2.158% | 3.703% |  | 1.15% | 2.301% | 2.39% |  | 6.003% | 2.051% | 4.884% | 1.516% | 2.366% | 2.858% | 5.163%
| PDM3 |  | -1.06% | 72.76% | 13.1% | 0.9115% | -3.502% | -7.326% | -9.68% | 15.03% | 4.736% | -4.384% | 12.39% |  | -2.913% | 6.853% | 81.48% |  | 97.7% | 36.13% | -13.21% | 4.148% | -0.8537% | 1.5% | -13.5%
| Z-Score (Absolute) |  | 0.16 | 4.41 | 0.72 | 0.04 | 0.31 | 0.55 | 0.69 | 0.84 | 0.2 | 0.36 | 0.67 | 4.65 | 0.27 | 0.33 | 4.95 | 0.82 | 5.96 | 2.15 | 0.91 | 0.16 | 0.15 | 0 | 0.93
| Analytical results for Ti in OREAS 506b (Certified Value 0.204 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.2016 | 0.237 | 0.22513 | 0.198 | 0.21 | 0.21 | 0.198 | 0.213 | 0.21 | 0.21 | 0.23 | 0.25 | 0.21 | 0.19912 | 0.238 | 0.21 | 0.2 | 0.187 | 0.187 | 0.201 | 0.2241 | 0.2 | 0.21 | 0.19 |  |  | 1
| 1 | 2 | 0.2024 | 0.234 | 0.21616 | 0.198 | 0.2 | 0.211 | 0.202 | 0.212 | 0.207 | 0.211 | 0.24 | 0.263 | 0.21 | 0.20723 | 0.23 | 0.21 | 0.2 | 0.184 | 0.194 | 0.202 | 0.2328 | 0.21 | 0.2 | 0.203 |  |  | 18
| 1 | 3 | 0.1975 | 0.238 | 0.21347 | 0.198 | 0.21 | 0.211 | 0.203 | 0.212 | 0.205 | 0.212 | 0.23 | 0.254 | 0.21 | 0.20332 | 0.236 | 0.21 | 0.2 | 0.195 | 0.187 | 0.196 | 0.2377 | 0.21 | 0.21 | 0.198 |  |  | 16
| 1 | 4 | 0.1991 | 0.232 | 0.21672 | 0.196 | 0.2 | 0.214 | 0.196 | 0.214 | 0.206 | 0.204 | 0.23 | 0.24 | 0.2 | 0.20622 | 0.234 | 0.21 | 0.2 | 0.19 | 0.175 | 0.213 | 0.2334 | 0.21 | 0.21 | 0.192 |  |  | 0.20437
| 1 | 5 | 0.2003 | 0.241 | 0.22048 | 0.201 | 0.21 | 0.212 | 0.202 | 0.216 | 0.205 | 0.207 | 0.23 | 0.24 | 0.21 | 0.20332 | 0.24 | 0.21 | 0.2 | 0.199 | 0.189 | 0.209 | 0.2397 | 0.21 | 0.21 | 0.193 |  |  | 114
| 1 | 6 | 0.1999 | 0.239 | 0.22686 | 0.194 | 0.2 | 0.211 | 0.205 | 0.215 | 0.208 | 0.201 | 0.23 | 0.25 | 0.21 | 0.2058 | 0.235 | 0.21 | 0.2 | 0.19 | 0.193 | 0.197 | 0.2305 | 0.21 | 0.2 | 0.19
| Mean |  | 0.20013 | 0.23683 | 0.2198 | 0.1975 | 0.205 | 0.2115 | 0.201 | 0.21367 | 0.20683 | 0.2075 | 0.23167 | 0.2495 | 0.20833 | 0.20417 | 0.2355 | 0.21 | 0.2 | 0.19083 | 0.1875 | 0.203 | 0.23303 | 0.20833 | 0.20667 | 0.19433
| Median |  | 0.2001 | 0.2375 | 0.2186 | 0.198 | 0.205 | 0.211 | 0.202 | 0.2135 | 0.2065 | 0.2085 | 0.23 | 0.25 | 0.21 | 0.20456 | 0.2355 | 0.21 | 0.2 | 0.19 | 0.188 | 0.2015 | 0.2331 | 0.21 | 0.21 | 0.1925
| Std Dev. |  | 0.0017535 | 0.0033116 | 0.0053204 | 0.0023452 | 0.0054772 | 0.0013784 | 0.0033466 | 0.001633 | 0.0019408 | 0.0043243 | 0.0040825 | 0.0087579 | 0.0040825 | 0.0029387 | 0.0034496 | 0 | 3.0405e-17 | 0.0054191 | 0.0068044 | 0.0067231 | 0.0055251 | 0.0040825 | 0.005164 | 0.005164
| Rel.Std.Dev. |  | 0.8762% | 1.398% | 2.421% | 1.187% | 2.672% | 0.6517% | 1.665% | 0.7643% | 0.9383% | 2.084% | 1.762% | 3.51% | 1.96% | 1.439% | 1.465% | 0% | 1.52e-14% | 2.84% | 3.629% | 3.312% | 2.371% | 1.96% | 2.499% | 2.657%
| PDM3 |  | -2.072% | 15.89% | 7.553% | -3.36% | 0.3098% | 3.49% | -1.647% | 4.551% | 1.207% | 1.533% | 13.36% | 22.08% | 1.941% | -0.09786% | 15.23% | 2.756% | -2.137% | -6.622% | -8.253% | -0.6688% | 14.03% | 1.941% | 1.125% | -4.91%
| Z-Score (Absolute) |  | 0.76 | 3.04 | 1.28 | 1.04 | 0.26 | 0.41 | 0.67 | 0.64 | 0.07 | 0 | 2.51 | 4.36 | 0.09 | 0.35 | 2.9 | 0.26 | NA | 1.73 | 2.07 | 0.47 | 2.65 | 0.09 | 0.09 | 1.37
Results from laboratory 2.19 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Tl in OREAS 506b (Certified Value 0.51 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  | 3
| 1 | 1 | 0.55 | 0.56 | 0.52804 | 0.46 | < 5 | 0.5 | 0.51 | 0.54 | 0.46 | 0.48 | 0.54 | < 10 | 0.5 | 0.51379 | 0.52 | < 5 | 0.53 | 0.76 | 0.45 | 0.4 | 0.48 | 0.49 | 0.514 | 0.53 |  |  | 1
| 1 | 2 | 0.54 | 0.56 | 0.51426 | 0.49 | < 5 | 0.5 | 0.54 | 0.53 | 0.47 | 0.5 | 0.54 | < 10 | 0.49 | 0.51247 | 0.53 | < 5 | 0.52 | 0.73 | 0.46 | 0.4 | 0.49 | 0.51 | 0.535 | 0.56 |  |  | 19
| 1 | 3 | 0.53 | 0.57 | 0.53091 | 0.46 | < 5 | 0.5 | 0.53 | 0.53 | 0.46 | 0.51 | 0.53 | < 10 | 0.51 | 0.49531 | 0.53 | < 5 | 0.51 | 0.76 | 0.47 | 0.4 | 0.49 | 0.52 | 0.523 | 0.55 |  |  | 16
| 1 | 4 | 0.53 | 0.57 | 0.51051 | 0.44 | < 5 | 0.51 | 0.51 | 0.53 | 0.48 | 0.51 | 0.52 | < 10 | 0.49 | 0.51514 | 0.54 | < 5 | 0.53 | 0.82 | 0.44 | 0.4 | 0.49 | 0.5 | 0.52 | 0.55 |  |  | 0.51169
| 1 | 5 | 0.54 | 0.56 | 0.52023 | 0.5 | < 5 | 0.5 | 0.51 | 0.52 | 0.47 | 0.5 | 0.53 | < 10 | 0.52 | 0.5235 | 0.52 | < 5 | 0.53 | 0.67 | 0.44 | 0.4 | 0.49 | 0.5 | 0.518 | 0.53 |  |  | 115
| 1 | 6 | 0.55 | 0.57 | 0.52596 | 0.5 | < 5 | 0.49 | 0.55 | 0.52 | 0.44 | 0.54 | 0.51 | < 10 | 0.49 | 0.5248 | 0.52 | < 5 | 0.51 | 0.72 | 0.48 | 0.4 | 0.48 | 0.49 | 0.515 | 0.54
| Mean |  | 0.54 | 0.565 | 0.52165 | 0.475 |  | 0.5 | 0.525 | 0.52833 | 0.46333 | 0.50667 | 0.52833 |  | 0.5 | 0.51417 | 0.52667 |  | 0.52167 | 0.74333 | 0.45667 | 0.4 | 0.48667 | 0.50167 | 0.52083 | 0.54333
| Median |  | 0.54 | 0.565 | 0.5231 | 0.475 |  | 0.5 | 0.52 | 0.53 | 0.465 | 0.505 | 0.53 |  | 0.495 | 0.51446 | 0.525 |  | 0.525 | 0.745 | 0.455 | 0.4 | 0.49 | 0.5 | 0.519 | 0.545
| Std Dev. |  | 0.0089443 | 0.0054772 | 0.0080726 | 0.0251 |  | 0.0063246 | 0.017607 | 0.0075277 | 0.013663 | 0.019664 | 0.01169 |  | 0.012649 | 0.010582 | 0.008165 |  | 0.0098319 | 0.050067 | 0.01633 | 6.0809e-17 | 0.005164 | 0.01169 | 0.007679 | 0.012111
| Rel.Std.Dev. |  | 1.656% | 0.9694% | 1.548% | 5.284% |  | 1.265% | 3.354% | 1.425% | 2.949% | 3.881% | 2.213% |  | 2.53% | 2.058% | 1.55% |  | 1.885% | 6.735% | 3.576% | 1.52e-14% | 1.061% | 2.33% | 1.474% | 2.229%
| PDM3 |  | 5.532% | 10.42% | 1.947% | -7.171% |  | -2.285% | 2.601% | 3.252% | -9.451% | -0.9822% | 3.252% |  | -2.285% | 0.4837% | 2.926% |  | 1.949% | 45.27% | -10.75% | -21.83% | -4.891% | -1.959% | 1.786% | 6.184%
| Z-Score (Absolute) |  | 0.57 | 1.35 | 0 | 1.45 | 61.57 | 0.67 | 0.1 | 0.21 | 1.82 | 0.47 | 0.21 | 139.37 | 0.67 | 0.23 | 0.16 | 61.57 | 0 | 6.9 | 2.02 | NA | 1.09 | 0.62 | 0.03 | 0.67
Results from laboratory 2.23 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Tm in OREAS 506b (Indicative Value 0.11 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.15; 2.03: 2.25 |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  |  | 2
|  |  | 2.01: 01g; 2.02: 0.2g; 2.03: 0.5g |  |  | 2
| 1 | 1 | 2.01: 0.114; 2.02: 0.11279; 2.03: 0.1 |  |  | 1
| 1 | 2 | 2.01: 0.12; 2.02: 0.10995; 2.03: 0.1 |  |  | 20
| 1 | 3 | 2.01: 0.118; 2.02: 0.10716; 2.03: 0.1 |  |  | 16
| 1 | 4 | 2.01: 0.116; 2.02: 0.11343; 2.03: 0.1 |  |  | 0.1093
| 1 | 5 | 2.01: 0.116; 2.02: 0.10977; 2.03: 0.1 |  |  | 41
| 1 | 6 | 2.01: 0.117; 2.02: 0.11331; 2.03: 0.1
| Mean |  | 2.01: 0.11683; 2.02: 0.11107; 2.03: 0.1
| Median |  | 2.01: 0.1165; 2.02: 0.11137; 2.03: 0.1
| Std Dev. |  | 2.01: 0.0020412; 2.02: 0.0025217; 2.03: 1.5202e-17
| Rel.Std.Dev. |  | 2.01: 1.747%; 2.02: 2.27%; 2.03: 1.52e-14%
| PDM3 |  | 2.01: 6.893%; 2.02: 1.616%; 2.03: -8.508%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0; 2.03: 1.29
| Analytical results for U in OREAS 506b (Certified Value 2.82 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.2g | 0.5g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g |  |  | 3
| 1 | 1 | 2.85 | 3.27 | 2.9819 | 2.75 | < 5 | 2.55 | 2.88 | 2.64 | 2.69 | 2.48 | 3.59 | < 10 | 2.75 | 2.9603 | 2.64 | 3.4177 | < 10 | 3.13 | 2.11 | 2.8 | 2.78 | 2.8 | 2.8 | 2.8 |  |  | 1
| 1 | 2 | 2.76 | 3.38 | 3.1578 | 2.69 | < 5 | 3 | 2.92 | 3.12 | 2.93 | 2.88 | 2.98 | < 10 | 2.98 | 2.8688 | 2.7 | 3.446 | < 10 | 2.71 | 2.3 | 2.9 | 2.63 | 2.8 | 3 | 3 |  |  | 21
| 1 | 3 | 2.89 | 3.34 | 3.1315 | 2.76 | < 5 | 2.57 | 2.77 | 3.01 | 2.49 | 2.81 | 3.52 | < 10 | 2.81 | 2.7271 | 2.96 | 3.7768 | < 10 | 3.07 | 2.37 | 2.6 | 2.52 | 3.1 | 2.75 | 2.9 |  |  | 16
| 1 | 4 | 2.72 | 3.26 | 2.9995 | 2.58 | < 5 | 2.86 | 2.99 | 3.03 | 2.63 | 2.67 | 2.85 | < 10 | 2.89 | 2.5418 | 2.94 | 3.5784 | < 10 | 4.03 | 2.32 | 2.6 | 2.41 | 2.7 | 2.92 | 2.8 |  |  | 2.8226
| 1 | 5 | 2.95 | 3.35 | 3.1069 | 2.55 | < 5 | 2.46 | 2.72 | 2.99 | 3.01 | 3.06 | 3.61 | < 10 | 2.88 | 2.8744 | 2.79 | 3.3518 | < 10 | 2.86 | 2.22 | 2.8 | 3.63 | 2.4 | 3.05 | 2.8 |  |  | 116
| 1 | 6 | 3.09 | 3.29 | 3.1859 | 2.96 | < 5 | 2.41 | 3.14 | 2.51 | 2.71 | 2.85 | 3.31 | < 10 | 2.67 | 2.7621 | 2.94 | 3.3122 | < 10 | 3.14 | 2.43 | 2.6 | 2.75 | 3 | 3.01 | 2.8
| Mean |  | 2.8767 | 3.315 | 3.0939 | 2.715 |  | 2.6417 | 2.9033 | 2.8833 | 2.7433 | 2.7917 | 3.31 |  | 2.83 | 2.7891 | 2.8283 | 3.4805 |  | 3.1567 | 2.2917 | 2.7167 | 2.7867 | 2.8 | 2.9217 | 2.85
| Median |  | 2.87 | 3.315 | 3.1192 | 2.72 |  | 2.56 | 2.9 | 3 | 2.7 | 2.83 | 3.415 |  | 2.845 | 2.8154 | 2.865 | 3.4319 |  | 3.1 | 2.31 | 2.7 | 2.69 | 2.8 | 2.96 | 2.8
| Std Dev. |  | 0.13411 | 0.048477 | 0.084361 | 0.14789 |  | 0.23507 | 0.15214 | 0.24639 | 0.19336 | 0.19773 | 0.3265 |  | 0.11045 | 0.14743 | 0.13834 | 0.17168 |  | 0.46016 | 0.11339 | 0.13292 | 0.43592 | 0.24495 | 0.12222 | 0.083666
| Rel.Std.Dev. |  | 4.662% | 1.462% | 2.727% | 5.447% |  | 8.898% | 5.24% | 8.545% | 7.048% | 7.083% | 9.864% |  | 3.903% | 5.286% | 4.891% | 4.933% |  | 14.58% | 4.948% | 4.893% | 15.64% | 8.748% | 4.183% | 2.936%
| PDM3 |  | 1.915% | 17.44% | 9.611% | -3.813% |  | -6.411% | 2.859% | 2.151% | -2.809% | -1.097% | 17.27% |  | 0.2612% | -1.189% | 0.2022% | 23.31% |  | 11.83% | -18.81% | -3.754% | -1.274% | -0.8016% | 3.509% | 0.9698%
| Z-Score (Absolute) |  | 0.22 | 2.91 | 1.56 | 0.77 | 2.08 | 1.22 | 0.39 | 0.27 | 0.59 | 0.3 | 2.88 | 13.24 | 0.06 | 0.31 | 0.07 | 3.93 | 13.24 | 1.94 | 3.36 | 0.76 | 0.33 | 0.25 | 0.5 | 0.06
| Analytical results for V in OREAS 506b (Certified Value 54 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  |  | 1
| 1 | 1 | 53 | 62 | 55.8129 | 53 | 53.7 | 54 | 54 | 57 | 54 | 53 | 56 | 57 | 53 | 53.9321 | 56 | 57 | 54 | 49 | 52 | 54.4 | 56 | 52 | 53 |  |  |  | 1
| 1 | 2 | 53 | 61 | 53.1682 | 54 | 52.5 | 55 | 56 | 57 | 54 | 54 | 57 | 60.7 | 54 | 55.7773 | 55 | 56 | 54 | 49 | 57 | 54.6 | 53 | 52 | 54 |  |  |  | 22
| 1 | 3 | 53 | 61 | 55.6172 | 54 | 53.2 | 55 | 55 | 58 | 53 | 54 | 54 | 57.8 | 53 | 55.1797 | 56 | 56.5 | 53 | 49 | 51 | 56 | 53 | 54 | 54 |  |  |  | 16
| 1 | 4 | 53 | 62 | 54.062 | 53 | 52.9 | 56 | 55 | 58 | 53 | 53 | 54 | 55 | 53 | 56.1494 | 55 | 57.5 | 54 | 64 | 55 | 54.6 | 54 | 53 | 54 |  |  |  | 54.3956
| 1 | 5 | 53 | 60 | 54.0201 | 54 | 52.3 | 55 | 55 | 57 | 52 | 54 | 56 | 55.3 | 53 | 55.0265 | 57 | 56 | 53 | 51 | 54 | 53.4 | 56 | 51 | 53 |  |  |  | 117
| 1 | 6 | 53 | 62 | 55.4564 | 52 | 53.6 | 55 | 56 | 57 | 53 | 53 | 55 | 56 | 54 | 56.0193 | 56 | 56.5 | 54 | 49 | 51 | 54.6 | 53 | 51 | 53
| Mean |  | 53 | 61.3333 | 54.6895 | 53.3333 | 53.0333 | 55 | 55.1667 | 57.3333 | 53.1667 | 53.5 | 55.3333 | 56.9667 | 53.3333 | 55.3474 | 55.8333 | 56.5833 | 53.6667 | 51.8333 | 53.3333 | 54.6 | 54.1667 | 52.1667 | 53.5
| Median |  | 53 | 61.5 | 54.7592 | 53.5 | 53.05 | 55 | 55 | 57 | 53 | 53.5 | 55.5 | 56.5 | 53 | 55.4785 | 56 | 56.5 | 54 | 49 | 53 | 54.6 | 53.5 | 52 | 53.5
| Std Dev. |  | 0 | 0.8165 | 1.0833 | 0.8165 | 0.57155 | 0.63246 | 0.75277 | 0.5164 | 0.75277 | 0.54772 | 1.2111 | 2.1078 | 0.5164 | 0.82625 | 0.75277 | 0.58452 | 0.5164 | 6.0139 | 2.4221 | 0.82946 | 1.472 | 1.169 | 0.54772
| Rel.Std.Dev. |  | 0% | 1.331% | 1.981% | 1.531% | 1.078% | 1.15% | 1.365% | 0.9007% | 1.416% | 1.024% | 2.189% | 3.7% | 0.9682% | 1.493% | 1.348% | 1.033% | 0.9622% | 11.6% | 4.541% | 1.519% | 2.717% | 2.241% | 1.024%
| PDM3 |  | -2.566% | 12.75% | 0.5403% | -1.953% | -2.504% | 1.111% | 1.418% | 5.401% | -2.259% | -1.646% | 1.724% | 4.727% | -1.953% | 1.75% | 2.643% | 4.022% | -1.34% | -4.71% | -1.953% | 0.3758% | -0.4208% | -4.098% | -1.646%
| Z-Score (Absolute) |  | 0.79 | 4.83 | 0.35 | 0.56 | 0.76 | 0.56 | 0.67 | 2.14 | 0.67 | 0.45 | 0.79 | 1.89 | 0.56 | 0.8 | 1.12 | 1.63 | 0.34 | 1.57 | 0.56 | 0.29 | 0 | 1.35 | 0.45
| Analytical results for W in OREAS 506b (Certified Value 6.42 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 30g |  |  | 3
| 1 | 1 | 5.11 | 7.2 | 7.2427 | 4.61 | 9.4 | 6.55 | 6.49 | 6.98 | 5.91 | 5.97 | 6.25 | < 10 | 7 | 7.99 | < 10 | 6.5 | < 20 | 6.83 | 5.51 | 6.38 | 4.86 | 6.4 | 6.6 | 4 |  |  | 1
| 1 | 2 | 5.17 | 7.1 | 6.7131 | 4.86 | 9.1 | 6.44 | 6.66 | 6.57 | 5.85 | 6.68 | 6.83 | < 10 | 6.4 | 8.4693 | < 10 | 6 | < 20 | 6.25 | 5.05 | 6.52 | 5.4 | 7 | 6.8 | 5 |  |  | 23
| 1 | 3 | 5.16 | 7 | 7.3599 | 4.52 | 9.7 | 6.12 | 6.22 | 6.13 | 5.83 | 5.72 | 6.82 | < 10 | 6.8 | 8.8506 | < 10 | 6.7 | < 20 | 6.36 | 4.79 | 6.51 | 4.98 | 6.3 | 7.3 | 4.7 |  |  | 16
| 1 | 4 | 5.3 | 6.8 | 6.9815 | 4.54 | 11.1 | 6.84 | 6.47 | 7.04 | 6.06 | 5.86 | 6.74 | < 10 | 6.2 | 7.9462 | 10 | 7.1 | < 20 | 6.91 | 4.95 | 6.7 | 5 | 6 | 6.9 | 4.3 |  |  | 6.4245
| 1 | 5 | 5.31 | 7.2 | 6.6088 | 4.53 | 9.6 | 6.5 | 6.49 | 6.64 | 6.12 | 5.96 | 6.33 | < 10 | 6.8 | 8.1901 | 11 | 6.7 | < 20 | 7.35 | 4.52 | 6.61 | 5.4 | 7.1 | 6.7 | 4.3 |  |  | 118
| 1 | 6 | 5.09 | 6.7 | 7.3403 | 4.78 | 10.3 | 6.3 | 6.75 | 7.65 | 5.6 | 6.54 | 6.69 | < 10 | 6.8 | 8.5774 | < 10 | 6.9 | < 20 | 7.31 | 4.61 | 6.34 | 4.86 | 6.2 | 7 | 4.5
| Mean |  | 5.19 | 7 | 7.041 | 4.64 | 9.8667 | 6.4583 | 6.5133 | 6.835 | 5.895 | 6.1217 | 6.61 |  | 6.6667 | 8.3373 | 10.5 | 6.65 |  | 6.835 | 4.905 | 6.51 | 5.0833 | 6.5 | 6.8833 | 4.4667
| Median |  | 5.165 | 7.05 | 7.1121 | 4.575 | 9.65 | 6.47 | 6.49 | 6.81 | 5.88 | 5.965 | 6.715 |  | 6.8 | 8.3297 | 10.5 | 6.7 |  | 6.87 | 4.87 | 6.515 | 4.99 | 6.35 | 6.85 | 4.4
| Std Dev. |  | 0.094021 | 0.20976 | 0.32552 | 0.14519 | 0.72296 | 0.24318 | 0.1825 | 0.51625 | 0.18512 | 0.39133 | 0.25448 |  | 0.30111 | 0.35604 | 0.70711 | 0.37815 |  | 0.46138 | 0.35709 | 0.13565 | 0.25216 | 0.44721 | 0.24833 | 0.35024
| Rel.Std.Dev. |  | 1.812% | 2.997% | 4.623% | 3.129% | 7.327% | 3.765% | 2.802% | 7.553% | 3.14% | 6.392% | 3.85% |  | 4.517% | 4.27% | 6.734% | 5.687% |  | 6.75% | 7.28% | 2.084% | 4.961% | 6.88% | 3.608% | 7.841%
| PDM3 |  | -19.22% | 8.957% | 9.596% | -27.78% | 53.58% | 0.5259% | 1.382% | 6.389% | -8.243% | -4.714% | 2.887% |  | 3.769% | 29.77% | 63.44% | 3.509% |  | 6.389% | -23.65% | 1.33% | -20.88% | 1.174% | 7.141% | -30.47%
| Z-Score (Absolute) |  | 1.83 | 0.67 | 0.73 | 2.6 | 4.65 | 0.08 | 0 | 0.45 | 0.86 | 0.54 | 0.13 | 2.1 | 0.21 | 2.53 | NA | 0.19 | 4.83 | 0.45 | 2.23 | 0 | 1.98 | 0.02 | 0.51 | 2.84
Results from laboratory 2.16 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 506b (Certified Value 9.63 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [5 empty] | 3
| 1 | 1 | 9.48 | 11.4 | 9.7534 | 8.8 | 9.98 | 9.71 | 9.72 | 9.62 | 8.94 | 10.2 | 9.89 | 9.739 | 10.7 | 9.4 | 13.7231 | 10 | 11.2 | 10.28 | 8.88 | 9.65 | 9.12 | [5 empty] | 1
| 1 | 2 | 9.56 | 11.2 | 9.4852 | 8.7 | 10.15 | 9.85 | 9.78 | 10.05 | 9.3 | 10.28 | 9.67 | 10.085 | 9.9 | 9.6 | 12.0611 | 10 | 10 | 10.16 | 9.31 | 9.56 | 9.11 | [5 empty] | 24
| 1 | 3 | 9.48 | 11.4 | 9.6244 | 9.3 | 9.67 | 9.67 | 9.69 | 9.91 | 9.08 | 10.25 | 9.8 | 9.7329 | 10.4 | 9.6 | 13.6811 | 9 | 11.5 | 10.12 | 8.77 | 9.48 | 9.27 | [5 empty] | 16
| 1 | 4 | 9.43 | 11.4 | 9.2678 | 8.7 | 10.25 | 9.81 | 9.86 | 10.15 | 9.14 | 9.97 | 9.56 | 10.0074 | 10.2 | 9.8 | 12.9242 | 9 | 11.3 | 9.22 | 9.03 | 9.57 | 9.24 | [5 empty] | 9.6349
| 1 | 5 | 9.33 | 11.6 | 9.2672 | 8.9 | 9.49 | 9.59 | 9.73 | 9.86 | 9.23 | 10.14 | 9.76 | 9.7275 | 10.7 | 9.6 | 13.259 | 9 | 11.4 | 10.31 | 9.01 | 9.64 | 9.07 | [5 empty] | 119
| 1 | 6 | 9.48 | 11.3 | 9.6073 | 8.8 | 9.71 | 9.87 | 9.68 | 9.5 | 9 | 10.19 | 9.87 | 9.9108 | 10.2 | 9.6 | 12.637 | 10 | 11.4 | 10.29 | 8.97 | 9.16 | 9.19
| Mean |  | 9.46 | 11.3833 | 9.5009 | 8.8667 | 9.875 | 9.75 | 9.7433 | 9.8483 | 9.115 | 10.1717 | 9.7583 | 9.8671 | 10.35 | 9.6 | 13.0476 | 9.5 | 11.1333 | 10.0633 | 8.995 | 9.51 | 9.1667
| Median |  | 9.48 | 11.4 | 9.5462 | 8.8 | 9.845 | 9.76 | 9.725 | 9.885 | 9.11 | 10.195 | 9.78 | 9.8249 | 10.3 | 9.6 | 13.0916 | 9.5 | 11.35 | 10.22 | 8.99 | 9.565 | 9.155
| Std Dev. |  | 0.076158 | 0.13292 | 0.19978 | 0.22509 | 0.29824 | 0.111 | 0.067132 | 0.24863 | 0.13649 | 0.11017 | 0.12545 | 0.15685 | 0.31464 | 0.12649 | 0.64175 | 0.54772 | 0.56451 | 0.42022 | 0.18174 | 0.18221 | 0.079162
| Rel.Std.Dev. |  | 0.8051% | 1.168% | 2.103% | 2.539% | 3.02% | 1.138% | 0.689% | 2.525% | 1.497% | 1.083% | 1.286% | 1.59% | 3.04% | 1.318% | 4.919% | 5.766% | 5.07% | 4.176% | 2.02% | 1.916% | 0.8636%
| PDM3 |  | -1.815% | 18.15% | -1.391% | -7.973% | 2.492% | 1.195% | 1.125% | 2.215% | -5.396% | 5.571% | 1.281% | 2.41% | 7.422% | -0.3622% | 35.42% | -1.4% | 15.55% | 4.447% | -6.641% | -1.296% | -4.86%
| Z-Score (Absolute) |  | 0.66 | 3.64 | 0.57 | 1.98 | 0.27 | 0.01 | 0.02 | 0.21 | 1.43 | 0.93 | 0.01 | 0.25 | 1.33 | 0.34 | 7.36 | NA | 3.08 | 0.69 | 1.7 | 0.55 | 1.31
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 506b (Certified Value 0.71 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab | [5 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.06; 2.03: 2.14; 2.06: 2.15; 2.07: 2.17; 2.08: 2.18; 2.09: 2.21; 2.10: 2.24; 2.11: 2.25 | [5 empty] | ppm
|  |  | 2.01: AR*MS; 2.02: AR*OES; 2.03: AR*OES/MS; 2.06: AR*MS; 2.07: AR*OES/MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*OES/MS; 2.11: AR*OES/MS | [5 empty] | 2
|  |  | 2.01: 01g; 2.02: 0.25g; 2.03: 0.25g; 2.06: 0.2g; 2.07: 0.2g; 2.08: 0.5g; 2.09: 0.25g; 2.10: 0.25g; 2.11: 0.5g | [5 empty] | 3
| 1 | 1 | 2.01: 0.736; 2.02: < 5; 2.03: 0.6; 2.06: 0.70332; 2.07: 0.7; 2.08: 0.97868; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7 | [5 empty] | 1
| 1 | 2 | 2.01: 0.773; 2.02: < 5; 2.03: 0.7; 2.06: 0.69282; 2.07: 0.7; 2.08: 0.86331; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7 | [5 empty] | 25
| 1 | 3 | 2.01: 0.758; 2.02: < 5; 2.03: 0.6; 2.06: 0.67379; 2.07: 0.7; 2.08: 0.93622; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7 | [5 empty] | 16
| 1 | 4 | 2.01: 0.74; 2.02: < 5; 2.03: 0.7; 2.06: 0.70402; 2.07: 0.7; 2.08: 0.93383; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7 | [5 empty] | 0.70662
| 1 | 5 | 2.01: 0.762; 2.02: < 5; 2.03: 0.7; 2.06: 0.69489; 2.07: 0.7; 2.08: 0.94924; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7 | [5 empty] | 120
| 1 | 6 | 2.01: 0.76; 2.02: < 5; 2.03: 0.7; 2.06: 0.68038; 2.07: 0.7; 2.08: 0.90304; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7
| Mean |  | 2.01: 0.75483; 2.03: 0.66667; 2.06: 0.69154; 2.07: 0.7; 2.08: 0.92739; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7
| Median |  | 2.01: 0.759; 2.03: 0.7; 2.06: 0.69386; 2.07: 0.7; 2.08: 0.93502; 2.09: 0.7; 2.10: 0.7; 2.11: 0.7
| Std Dev. |  | 2.01: 0.014091; 2.03: 0.05164; 2.06: 0.012223; 2.07: 1.2162e-16; 2.08: 0.039817; 2.09: 1.2162e-16; 2.10: 1.2162e-16; 2.11: 1.2162e-16
| Rel.Std.Dev. |  | 2.01: 1.867%; 2.03: 7.746%; 2.06: 1.768%; 2.07: 1.737e-14%; 2.08: 4.293%; 2.09: 1.737e-14%; 2.10: 1.737e-14%; 2.11: 1.737e-14%
| PDM3 |  | 2.01: 6.822%; 2.03: -5.655%; 2.06: -2.135%; 2.07: -0.9374%; 2.08: 31.24%; 2.09: -0.9374%; 2.10: -0.9374%; 2.11: -0.9374%
| Z-Score (Absolute) |  | 2.01: 4.37; 2.02: 143.39; 2.03: 2.66; 2.06: 0.67; 2.07: 0; 2.08: 18.11; 2.09: 0; 2.10: 0; 2.11: 0
| Analytical results for Zn in OREAS 506b (Certified Value 106 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | 2.26 | 2.27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 01g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.2g | 0.25g | 0.2g | 0.5g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 30g | 0
| 1 | 1 | 99 | 110 | 107.8391 | 101 | 96.2 | 104 | 100 | 109 | 111 | 104 | 109 | 140 | 111 | 105.5412 | 113 | 109.3 | 98.19 | 105 | 111 | 119 | 105 | 90 | 93 | 102 | 111 | 106.5 | 1
| 1 | 2 | 97 | 108 | 101.0326 | 99 | 93.9 | 106 | 103 | 108 | 110 | 104 | 113 | 152 | 108 | 108.9457 | 111 | 107.6 | 97.04 | 106 | 104 | 122 | 109 | 87 | 92 | 105 | 112 | 106.2 | 26
| 1 | 3 | 99 | 109 | 105.3633 | 101.6 | 99.8 | 106 | 103 | 110 | 108 | 103 | 111 | 141 | 110 | 108.6045 | 112 | 105.9 | 98.83 | 105 | 107 | 122 | 103 | 89 | 95 | 110 | 112 | 108.7 | 16
| 1 | 4 | 99 | 110 | 101.7816 | 100 | 97.2 | 107 | 102 | 109 | 109 | 101 | 109 | 141 | 106 | 109.625 | 114 | 109.2 | 96.99 | 105 | 107 | 115 | 108 | 91 | 93 | 107 | 112 | 108.8 | 105.7684
| 1 | 5 | 98 | 108 | 102.8227 | 102.1 | 94.1 | 106 | 103 | 107 | 108 | 103 | 110 | 142 | 111 | 108.8795 | 117 | 107.3 | 97.86 | 105 | 107 | 116 | 105 | 92 | 95 | 103 | 111 | 107.7 | 121
| 1 | 6 | 99 | 107 | 108.2312 | 99.3 | 94.4 | 106 | 104 | 107 | 109 | 101 | 111 | 139 | 108 | 110.2577 | 114 | 108.5 | 97.57 | 107 | 106 | 118 | 105 | 91 | 92 | 104 | 111 | 106
| Mean |  | 98.5 | 108.6667 | 104.5118 | 100.5 | 95.9333 | 105.8333 | 102.5 | 108.3333 | 109.1667 | 102.6667 | 110.5 | 142.5 | 109 | 108.6423 | 113.5 | 107.9667 | 97.7467 | 105.5 | 107 | 118.6667 | 105.8333 | 90 | 93.3333 | 105.1667 | 111.5 | 107.3167
| Median |  | 99 | 108.5 | 104.093 | 100.5 | 95.3 | 106 | 103 | 108.5 | 109 | 103 | 110.5 | 141 | 109 | 108.9126 | 113.5 | 108.05 | 97.715 | 105 | 107 | 118.5 | 105 | 90.5 | 93 | 104.5 | 111.5 | 107.1
| Std Dev. |  | 0.83666 | 1.2111 | 3.0997 | 1.2617 | 2.301 | 0.98319 | 1.3784 | 1.2111 | 1.169 | 1.3663 | 1.5166 | 4.7645 | 2 | 1.6335 | 2.0736 | 1.2987 | 0.70514 | 0.83666 | 2.2804 | 2.9439 | 2.2286 | 1.7889 | 1.3663 | 2.9269 | 0.54772 | 1.2576
| Rel.Std.Dev. |  | 0.8494% | 1.114% | 2.966% | 1.255% | 2.399% | 0.929% | 1.345% | 1.118% | 1.071% | 1.331% | 1.372% | 3.343% | 1.835% | 1.504% | 1.827% | 1.203% | 0.7214% | 0.793% | 2.131% | 2.481% | 2.106% | 1.988% | 1.464% | 2.783% | 0.4912% | 1.172%
| PDM3 |  | -6.872% | 2.74% | -1.188% | -4.981% | -9.299% | 0.06142% | -3.09% | 2.425% | 3.213% | -2.933% | 4.474% | 34.73% | 3.055% | 2.717% | 7.31% | 2.078% | -7.584% | -0.2537% | 1.164% | 12.19% | 0.06142% | -14.91% | -11.76% | -0.5689% | 5.419% | 1.464%
| Z-Score (Absolute) |  | 1.64 | 0.47 | 0.4 | 1.23 | 2.18 | 0.12 | 0.81 | 0.4 | 0.57 | 0.78 | 0.85 | 7.49 | 0.54 | 0.46 | 1.47 | 0.32 | 1.8 | 0.19 | 0.12 | 2.54 | 0.12 | 3.41 | 2.71 | 0.26 | 1.05 | 0.19
| Analytical results for Zr in OREAS 506b (Certified Value 8.6 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [7 empty] | 2
|  |  | 01g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.2g | 0.25g | 0.2g | 0.1g | 0.25g | 0.5g | 0.25g | 0.5g | [7 empty] | 3
| 1 | 1 | 9.6 | 10.7 | 8.2838 | 7.7 | 8.3 | 8 | 8.9 | 8.3 | 8.2 | 8.9 | 7.6 | 8.6737 | 7.4 | 11 | 10.8 | 9.7 | 11.3 | 8.2 | 11.3 | [7 empty] | 1
| 1 | 2 | 9.6 | 10.4 | 8.1093 | 8.8 | 8.3 | 8.2 | 8.9 | 8.3 | 8.8 | 8.8 | 7.4 | 8.9736 | 7.5 | 11 | 8.5 | 9.8 | 11.2 | 7.9 | 11.1 | [7 empty] | 27
| 1 | 3 | 9.5 | 10.7 | 8.3831 | 9.2 | 8 | 8 | 8.8 | 8.2 | 8.5 | 8.6 | 7.7 | 8.746 | 7.2 | 10 | 12.6 | 9.5 | 11 | 8.5 | 11.2 | [7 empty] | 16
| 1 | 4 | 9.3 | 10.3 | 7.9982 | 9.3 | 8.5 | 8.1 | 9 | 8.5 | 8.4 | 8.7 | 7.3 | 8.902 | 7.7 | 10 | 11.8 | 9.2 | 11.1 | 8.2 | 11.1 | [7 empty] | 8.6034
| 1 | 5 | 9.5 | 10.7 | 8.2272 | 8.6 | 8.1 | 7.8 | 9 | 8.2 | 8.6 | 8.9 | 7.5 | 8.6245 | 7.4 | 10 | 11.4 | 9.3 | 11.3 | 8.6 | 11.4 | [7 empty] | 122
| 1 | 6 | 9.5 | 10.3 | 8.0826 | 8.8 | 8.1 | 8.2 | 9.3 | 8 | 7.8 | 8.3 | 7.5 | 8.7981 | 7.9 | 10 | 11.4 | 9.7 | 10.9 | 7.8 | 11.6
| Mean |  | 9.5 | 10.5167 | 8.1807 | 8.7333 | 8.2167 | 8.05 | 8.9833 | 8.25 | 8.3833 | 8.7 | 7.5 | 8.7863 | 7.5167 | 10.3333 | 11.0833 | 9.5333 | 11.1333 | 8.2 | 11.2833
| Median |  | 9.5 | 10.55 | 8.1682 | 8.8 | 8.2 | 8.05 | 8.95 | 8.25 | 8.45 | 8.75 | 7.5 | 8.772 | 7.45 | 10 | 11.4 | 9.6 | 11.15 | 8.2 | 11.25
| Std Dev. |  | 0.10954 | 0.20412 | 0.14267 | 0.57155 | 0.18348 | 0.15166 | 0.17224 | 0.16432 | 0.34881 | 0.22804 | 0.14142 | 0.13348 | 0.24833 | 0.5164 | 1.3977 | 0.24221 | 0.1633 | 0.31623 | 0.19408
| Rel.Std.Dev. |  | 1.153% | 1.941% | 1.744% | 6.544% | 2.233% | 1.884% | 1.917% | 1.992% | 4.161% | 2.621% | 1.886% | 1.519% | 3.304% | 4.997% | 12.61% | 2.541% | 1.467% | 3.856% | 1.72%
| PDM3 |  | 10.42% | 22.24% | -4.913% | 1.511% | -4.495% | -6.432% | 4.417% | -4.107% | -2.557% | 1.123% | -12.82% | 2.127% | -12.63% | 20.11% | 28.83% | 10.81% | 29.41% | -4.688% | 31.15%
| Z-Score (Absolute) |  | 0.88 | 2.02 | 0.6 | 0.02 | 0.56 | 0.75 | 0.3 | 0.52 | 0.37 | 0.02 | 1.36 | 0.08 | 1.35 | NA | 2.65 | 0.92 | 2.71 | 0.58 | 2.88
Results from laboratory 2.19 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1154: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1172: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 506b (Indicative Value 14.87 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.86 | 1
| 1 | 2 | 14.88 | 39
| Mean |  | 14.87 | 16
| Median |  | 14.87 | 14.87
| Std Dev. |  | 0.014142 | 45
| Rel.Std.Dev. |  | 0.09511%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 506b (Indicative Value 10 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 13
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 506b (Indicative Value 1105 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1105 | 1
| 1 | 2 | 1105 | 41
| Mean |  | 1105 | 16
| Median |  | 1105 | 1105.335
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.03031%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 506b (Indicative Value 2.34 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.34 | 1
| 1 | 2 | 2.33 | 42
| Mean |  | 2.335 | 16
| Median |  | 2.335 | 2.335
| Std Dev. |  | 0.0070711 | 48
| Rel.Std.Dev. |  | 0.3028%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 506b (Indicative Value 85 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 90 | 1
| 1 | 2 | 80 | 43
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 7.0711 | 49
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 506b (Indicative Value 10 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 20
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 50
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 506b (Indicative Value 66 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 58 | 1
| 1 | 2 | 73 | 45
| Mean |  | 65.5 | 16
| Median |  | 65.5 | 65.7675
| Std Dev. |  | 10.6066 | 51
| Rel.Std.Dev. |  | 16.19%
| PDM3 |  | -0.4067%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 506b (Indicative Value 0.447 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.447 | 1
| 1 | 2 | 0.447 | 22
| Mean |  | 0.447 | 16
| Median |  | 0.447 | 0.447
| Std Dev. |  | 0 | 52
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 506b (Indicative Value 4 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.003 | 1
| 1 | 2 | 3.989 | 39
| Mean |  | 3.996 | 16
| Median |  | 3.996 | 3.996
| Std Dev. |  | 0.0098995 | 45
| Rel.Std.Dev. |  | 0.2477%
| PDM3 |  | -0.0002878%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 506b (Indicative Value 4.09 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.09 | 1
| 1 | 2 | 4.08 | 40
| Mean |  | 4.085 | 16
| Median |  | 4.085 | 4.085
| Std Dev. |  | 0.0070711 | 46
| Rel.Std.Dev. |  | 0.1731%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 506b (Indicative Value 1.2 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 41
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 506b (Indicative Value 0.045 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.05 | 42
| Mean |  | 0.045 | 16
| Median |  | 0.045 | 0.045
| Std Dev. |  | 0.0070711 | 48
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 506b (Indicative Value 2.96 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.96 | 1
| 1 | 2 | 2.95 | 43
| Mean |  | 2.955 | 16
| Median |  | 2.955 | 2.955
| Std Dev. |  | 0.0070711 | 49
| Rel.Std.Dev. |  | 0.2393%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 506b (Indicative Value 25 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 30 | 23
| Mean |  | 25 | 16
| Median |  | 25 | 25
| Std Dev. |  | 7.0711 | 50
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 506b (Indicative Value 0.19 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.1902 | 1
| 1 | 2 | 0.1902 | 45
| Mean |  | 0.1902 | 16
| Median |  | 0.1902 | 0.1902
| Std Dev. |  | 0 | 51
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.001472%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 506b (Indicative Value 55 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 50 | 25
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 7.0711 | 52
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 506b (Indicative Value 0.619 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.621 | 1
| 1 | 2 | 0.617 | 7
| Mean |  | 0.619 | 16
| Median |  | 0.619 | 0.619
| Std Dev. |  | 0.0028284 | 45
| Rel.Std.Dev. |  | 0.4569%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 506b (Indicative Value 67.27 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.29 | 1
| 1 | 2 | 67.25 | 40
| Mean |  | 67.27 | 16
| Median |  | 67.27 | 67.27
| Std Dev. |  | 0.028284 | 46
| Rel.Std.Dev. |  | 0.04205%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 506b (Indicative Value < 10 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 13
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 506b (Indicative Value 250 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 250 | 1
| 1 | 2 | 250 | 14
| Mean |  | 250 | 16
| Median |  | 250 | 250
| Std Dev. |  | 0 | 48
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 506b (Indicative Value 0.565 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.57 | 1
| 1 | 2 | 0.56 | 43
| Mean |  | 0.565 | 16
| Median |  | 0.565 | 0.565
| Std Dev. |  | 0.0070711 | 49
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 506b (Indicative Value 134 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 143 | 1
| 1 | 2 | 125 | 44
| Mean |  | 134 | 16
| Median |  | 134 | 133.89
| Std Dev. |  | 12.7279 | 50
| Rel.Std.Dev. |  | 9.498%
| PDM3 |  | 0.08216%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 506b (Indicative Value 110 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 110 | 1
| 1 | 2 | 110 | 26
| Mean |  | 110 | 16
| Median |  | 110 | 110
| Std Dev. |  | 0 | 51
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 506b (Indicative Value 210 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 210 | 1
| 1 | 2 | 210 | 27
| Mean |  | 210 | 16
| Median |  | 210 | 210
| Std Dev. |  | 0 | 52
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 506b (Indicative Value 1.51 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.49 | 48
| Mean |  | 1.51 | 16
| Median |  | 1.51 | 1.51
| Std Dev. |  | 0.028284 | 54
| Rel.Std.Dev. |  | 1.873%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 506b (Indicative Value 1.25 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.3 | 11
| Mean |  | 1.25 | 16
| Median |  | 1.25 | 1.25
| Std Dev. |  | 0.070711 | 56
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 506b (Indicative Value 12.6 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 12.6 | 13
| Mean |  | 12.6 | 16
| Median |  | 12.6 | 12.6
| Std Dev. |  | 0 | 57
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 506b (Indicative Value 918 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 916 | 1
| 1 | 2 | 919 | 14
| Mean |  | 917.5 | 16
| Median |  | 917.5 | 917.5
| Std Dev. |  | 2.1213 | 58
| Rel.Std.Dev. |  | 0.2312%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 506b (Indicative Value 2.4 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.4 | 15
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0 | 59
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 506b (Indicative Value 1.07 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.96 | 1
| 1 | 2 | 1.18 | 16
| Mean |  | 1.07 | 16
| Median |  | 1.07 | 1.07
| Std Dev. |  | 0.15556 | 60
| Rel.Std.Dev. |  | 14.54%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 506b (Indicative Value 0.4 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 18
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 61
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 506b (Indicative Value 66 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 64.6 | 1
| 1 | 2 | 67.1 | 19
| Mean |  | 65.85 | 16
| Median |  | 65.85 | 65.85
| Std Dev. |  | 1.7678 | 62
| Rel.Std.Dev. |  | 2.685%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 506b (Indicative Value 7.8 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.7 | 1
| 1 | 2 | 7.9 | 20
| Mean |  | 7.8 | 16
| Median |  | 7.8 | 7.8
| Std Dev. |  | 0.14142 | 63
| Rel.Std.Dev. |  | 1.813%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 506b (Indicative Value 46 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 46 | 21
| Mean |  | 46 | 16
| Median |  | 46 | 46
| Std Dev. |  | 0 | 64
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 506b (Indicative Value 8.51 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.46 | 1
| 1 | 2 | 8.55 | 59
| Mean |  | 8.505 | 16
| Median |  | 8.505 | 8.505
| Std Dev. |  | 0.06364 | 65
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 506b (Indicative Value 0.444 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.438 | 1
| 1 | 2 | 0.45 | 22
| Mean |  | 0.444 | 16
| Median |  | 0.444 | 0.444
| Std Dev. |  | 0.0084853 | 66
| Rel.Std.Dev. |  | 1.911%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 506b (Indicative Value 5.17 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.11 | 1
| 1 | 2 | 5.23 | 23
| Mean |  | 5.17 | 16
| Median |  | 5.17 | 5.17
| Std Dev. |  | 0.084853 | 67
| Rel.Std.Dev. |  | 1.641%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 506b (Indicative Value 2.79 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.77 | 24
| Mean |  | 2.785 | 16
| Median |  | 2.785 | 2.785
| Std Dev. |  | 0.021213 | 68
| Rel.Std.Dev. |  | 0.7617%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 506b (Indicative Value 1.27 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.23 | 1
| 1 | 2 | 1.31 | 25
| Mean |  | 1.27 | 16
| Median |  | 1.27 | 1.27
| Std Dev. |  | 0.056569 | 69
| Rel.Std.Dev. |  | 4.454%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 506b (Indicative Value 18.6 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.1 | 1
| 1 | 2 | 19 | 27
| Mean |  | 18.55 | 16
| Median |  | 18.55 | 18.55
| Std Dev. |  | 0.6364 | 70
| Rel.Std.Dev. |  | 3.431%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 506b (Indicative Value 6.15 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.09 | 1
| 1 | 2 | 6.21 | 28
| Mean |  | 6.15 | 16
| Median |  | 6.15 | 6.15
| Std Dev. |  | 0.084853 | 71
| Rel.Std.Dev. |  | 1.38%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 506b (Indicative Value 1.45 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.45 | 1
| 1 | 2 | 1.45 | 0
| Mean |  | 1.45 | 16
| Median |  | 1.45 | 1.45
| Std Dev. |  | 0 | 72
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 506b (Indicative Value 6.22 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.15 | 1
| 1 | 2 | 6.29 | 29
| Mean |  | 6.22 | 16
| Median |  | 6.22 | 6.22
| Std Dev. |  | 0.098995 | 56
| Rel.Std.Dev. |  | 1.592%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 506b (Indicative Value 1.03 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.03 | 1
| 1 | 2 | 1.02 | 12
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.0070711 | 57
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 506b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 52
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 58
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 506b (Indicative Value 33.5 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.2 | 1
| 1 | 2 | 33.8 | 14
| Mean |  | 33.5 | 16
| Median |  | 33.5 | 33.5
| Std Dev. |  | 0.42426 | 59
| Rel.Std.Dev. |  | 1.266%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 506b (Indicative Value 0.38 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.38 | 16
| Mean |  | 0.38 | 16
| Median |  | 0.38 | 0.38
| Std Dev. |  | 0 | 60
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 506b (Indicative Value 0.032 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0318 | 1
| 1 | 2 | 0.0327 | 18
| Mean |  | 0.03225 | 16
| Median |  | 0.03225 | 0.03225
| Std Dev. |  | 0.0006364 | 61
| Rel.Std.Dev. |  | 1.973%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 506b (Indicative Value 85 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84.4 | 1
| 1 | 2 | 85.6 | 19
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 0.84853 | 62
| Rel.Std.Dev. |  | 0.9983%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 506b (Indicative Value 11.7 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.6 | 1
| 1 | 2 | 11.8 | 21
| Mean |  | 11.7 | 16
| Median |  | 11.7 | 11.7
| Std Dev. |  | 0.14142 | 63
| Rel.Std.Dev. |  | 1.209%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 506b (Indicative Value 31 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.6 | 1
| 1 | 2 | 31.4 | 22
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 0.56569 | 64
| Rel.Std.Dev. |  | 1.825%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 506b (Indicative Value 17 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 16 | 23
| Mean |  | 17 | 16
| Median |  | 17 | 17
| Std Dev. |  | 1.4142 | 65
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 506b (Indicative Value 53 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 53 | 1
| 1 | 2 | 53 | 25
| Mean |  | 53 | 16
| Median |  | 53 | 53
| Std Dev. |  | 0 | 66
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 506b (Indicative Value 8.03 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.96 | 1
| 1 | 2 | 8.09 | 26
| Mean |  | 8.025 | 16
| Median |  | 8.025 | 8.025
| Std Dev. |  | 0.091924 | 67
| Rel.Std.Dev. |  | 1.145%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 506b (Indicative Value 144 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 142 | 1
| 1 | 2 | 145 | 62
| Mean |  | 143.5 | 16
| Median |  | 143.5 | 143.5
| Std Dev. |  | 2.1213 | 68
| Rel.Std.Dev. |  | 1.478%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 506b (Indicative Value 0.03 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.02 | 63
| Mean |  | 0.03 | 16
| Median |  | 0.03 | 0.03
| Std Dev. |  | 0.014142 | 69
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 506b (Indicative Value 1.3 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.3 | 28
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0 | 70
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 506b (Indicative Value 7.8 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.6 | 1
| 1 | 2 | 8 | 29
| Mean |  | 7.8 | 16
| Median |  | 7.8 | 7.8
| Std Dev. |  | 0.28284 | 71
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 506b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 0
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 72
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 506b (Indicative Value 6.55 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.45 | 1
| 1 | 2 | 6.64 | 12
| Mean |  | 6.545 | 16
| Median |  | 6.545 | 6.545
| Std Dev. |  | 0.13435 | 56
| Rel.Std.Dev. |  | 2.053%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 506b (Indicative Value 4.2 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4.2 | 13
| Mean |  | 4.2 | 16
| Median |  | 4.2 | 4.2
| Std Dev. |  | 0 | 57
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 506b (Indicative Value 267 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 265 | 1
| 1 | 2 | 269 | 14
| Mean |  | 267 | 16
| Median |  | 267 | 267
| Std Dev. |  | 2.8284 | 58
| Rel.Std.Dev. |  | 1.059%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 506b (Indicative Value 1.03 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 1.04 | 15
| Mean |  | 1.03 | 16
| Median |  | 1.03 | 1.03
| Std Dev. |  | 0.014142 | 59
| Rel.Std.Dev. |  | 1.373%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 506b (Indicative Value 0.94 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.93 | 1
| 1 | 2 | 0.94 | 16
| Mean |  | 0.935 | 16
| Median |  | 0.935 | 0.935
| Std Dev. |  | 0.0070711 | 60
| Rel.Std.Dev. |  | 0.7563%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 506b (Indicative Value 0.15 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | < 0.2 | 55
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.15
| Std Dev. |  |  | 61
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 506b (Indicative Value 12.8 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.6 | 1
| 1 | 2 | 12.9 | 17
| Mean |  | 12.75 | 16
| Median |  | 12.75 | 12.75
| Std Dev. |  | 0.21213 | 62
| Rel.Std.Dev. |  | 1.664%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 506b (Indicative Value 0.346 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.344 | 1
| 1 | 2 | 0.347 | 18
| Mean |  | 0.3455 | 16
| Median |  | 0.3455 | 0.3455
| Std Dev. |  | 0.0021213 | 63
| Rel.Std.Dev. |  | 0.614%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 506b (Indicative Value 0.6 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 19
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 64
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 506b (Indicative Value 0.43 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.43 | 1
| 1 | 2 | 0.42 | 20
| Mean |  | 0.425 | 16
| Median |  | 0.425 | 0.425
| Std Dev. |  | 0.0070711 | 65
| Rel.Std.Dev. |  | 1.664%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 506b (Indicative Value 4.37 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.12 | 1
| 1 | 2 | 4.61 | 21
| Mean |  | 4.365 | 16
| Median |  | 4.365 | 4.365
| Std Dev. |  | 0.34648 | 66
| Rel.Std.Dev. |  | 7.938%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 506b (Indicative Value 70 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 69.4 | 1
| 1 | 2 | 70.7 | 22
| Mean |  | 70.05 | 16
| Median |  | 70.05 | 70.05
| Std Dev. |  | 0.91924 | 67
| Rel.Std.Dev. |  | 1.312%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 506b (Indicative Value 12 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.5 | 1
| 1 | 2 | 12.5 | 23
| Mean |  | 12 | 16
| Median |  | 12 | 12
| Std Dev. |  | 0.70711 | 68
| Rel.Std.Dev. |  | 5.893%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 506b (Indicative Value 26.8 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.7 | 1
| 1 | 2 | 26.8 | 24
| Mean |  | 26.75 | 16
| Median |  | 26.75 | 26.75
| Std Dev. |  | 0.070711 | 69
| Rel.Std.Dev. |  | 0.2643%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 506b (Indicative Value 2.63 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.59 | 1
| 1 | 2 | 2.67 | 25
| Mean |  | 2.63 | 16
| Median |  | 2.63 | 2.63
| Std Dev. |  | 0.056569 | 70
| Rel.Std.Dev. |  | 2.151%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 506b (Indicative Value 105 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 105 | 1
| 1 | 2 | 105 | 26
| Mean |  | 105 | 16
| Median |  | 105 | 105
| Std Dev. |  | 0 | 71
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 506b (Indicative Value 221 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 219 | 1
| 1 | 2 | 223 | 27
| Mean |  | 221 | 16
| Median |  | 221 | 221
| Std Dev. |  | 2.8284 | 72
| Rel.Std.Dev. |  | 1.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.