## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 753b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 0.27288 | 0.23969 | 0.30607 | 0.25064 | 0.29511
Al, Aluminium (wt.%) | 6.0109 | 5.7755 | 6.2462 | 5.8619 | 6.1598
As, Arsenic (ppm) | 4.7453 | 4.4589 | 5.0316 | 4.367 | 5.1235
Ba, Barium (ppm) | 130.4267 | 125.6068 | 135.2466 | 126.7056 | 134.1478
Be, Beryllium (ppm) | 11.6959 | 11.3396 | 12.0522 | 11.4324 | 11.9593
Bi, Bismuth (ppm) | 0.72074 | 0.67056 | 0.77091 | 0.68027 | 0.76121
Ca, Calcium (wt.%) | 0.22763 | 0.21835 | 0.2369 | 0.22052 | 0.23474
Cd, Cadmium (ppm) | 0.13611 | 0.10598 | 0.16624 | IND | IND
Ce, Cerium (ppm) | 7.9229 | 7.0822 | 8.7636 | 7.4364 | 8.4094
Co, Cobalt (ppm) | 1.3821 | 1.2975 | 1.4667 | 1.3373 | 1.4269
Cr, Chromium (ppm) | 16.9538 | 14.7863 | 19.1214 | 16.0097 | 17.8979
Cs, Caesium (ppm) | 169.5518 | 162.8692 | 176.2345 | 165.9473 | 173.1564
Cu, Copper (ppm) | 18.0328 | 17.2962 | 18.7693 | 17.3871 | 18.6784
Dy, Dysprosium (ppm) | 0.45965 | 0.40181 | 0.51748 | 0.43613 | 0.48316
Er, Erbium (ppm) | 0.20707 | 0.17093 | 0.24322 | 0.18986 | 0.22429
Eu, Europium (ppm) | 0.15253 | 0.13332 | 0.17174 | IND | IND
Fe, Iron (wt.%) | 0.61476 | 0.59634 | 0.63319 | 0.5988 | 0.63072
Ga, Gallium (ppm) | 46.5152 | 44.6372 | 48.3931 | 45.3483 | 47.682
Gd, Gadolinium (ppm) | 0.71187 | 0.64758 | 0.77616 | 0.64812 | 0.77561
Ge, Germanium (ppm) | 0.076944 | 0.041526 | 0.11236 | IND | IND
Hf, Hafnium (ppm) | 1.8617 | 1.7652 | 1.9582 | 1.7874 | 1.936
Ho, Holmium (ppm) | 0.076111 | 0.05638 | 0.095843 | IND | IND
K, Potassium (wt.%) | 0.93837 | 0.91283 | 0.96391 | 0.91995 | 0.95678
La, Lanthanum (ppm) | 3.8022 | 3.3206 | 4.2839 | 3.4923 | 4.1122
Li, Lithium (wt.%) | 1.0345 | 1.0133 | 1.0558 | 1.0234 | 1.0456
Li2O, Lithium oxide (wt.%) | 2.2271 | 2.1812 | 2.2729 | 2.2032 | 2.251
Lu, Lutetium (ppm) | 0.025 | 0.015215 | 0.034785 | IND | IND
Mg, Magnesium (wt.%) | 0.12043 | 0.11359 | 0.12728 | 0.11793 | 0.12294
Mn, Manganese (wt.%) | 0.13262 | 0.12859 | 0.13665 | 0.12972 | 0.13553
Mo, Molybdenum (ppm) | 3.3625 | 3.2053 | 3.5197 | 3.2402 | 3.4848
Na, Sodium (wt.%) | 0.41235 | 0.40096 | 0.42373 | 0.40262 | 0.42208
Nb, Niobium (ppm) | 102.0064 | 98.1644 | 105.8483 | 98.6705 | 105.3423
Nd, Neodymium (ppm) | 3.9442 | 3.5721 | 4.3162 | 3.6465 | 4.2418
Ni, Nickel (ppm) | 8.3177 | 7.9101 | 8.7253 | 8.0373 | 8.5981
P, Phosphorus (wt.%) | 0.049133 | 0.047429 | 0.050838 | 0.048038 | 0.050229
Pb, Lead (ppm) | 4.6367 | 4.4114 | 4.862 | 4.4166 | 4.8568
Pr, Praseodymium (ppm) | 1.0748 | 0.96801 | 1.1817 | 0.95922 | 1.1905
Rb, Rubidium (ppm) | 1215.2833 | 1171.1762 | 1259.3904 | 1191.3847 | 1239.1819
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.025183 | 0.019442 | 0.030924 | 0.023869 | 0.026497
Sb, Antimony (ppm) | 0.64158 | 0.60403 | 0.67912 | 0.6101 | 0.67305
Sc, Scandium (ppm) | 1.0272 | 0.90993 | 1.1445 | 0.92727 | 1.1272
Sm, Samarium (ppm) | 0.81696 | 0.75066 | 0.88326 | 0.71238 | 0.92153
Sn, Tin (ppm) | 109.4955 | 104.2281 | 114.763 | 106.0399 | 112.9512
Sr, Strontium (ppm) | 38.2516 | 36.6451 | 39.858 | 37.1101 | 39.393
Ta, Tantalum (ppm) | 282.846 | 267.4755 | 298.2165 | 274.8753 | 290.8166
Tb, Terbium (ppm) | 0.10039 | 0.086728 | 0.11405 | IND | IND
Te, Tellurium (ppm) | 0.071203 | 0.041028 | 0.10138 | IND | IND
Th, Thorium (ppm) | 1.9228 | 1.7248 | 2.1208 | 1.8021 | 2.0434
Ti, Titanium (wt.%) | 0.052081 | 0.05041 | 0.053751 | 0.050637 | 0.053525
Tl, Thallium (ppm) | 8.7265 | 8.3703 | 9.0826 | 8.5153 | 8.9376
U, Uranium (ppm) | 1.8629 | 1.6591 | 2.0668 | 1.777 | 1.9488
V, Vanadium (ppm) | 8.1126 | 7.6504 | 8.5748 | 7.7298 | 8.4954
W, Tungsten (ppm) | 3.3661 | 3.0276 | 3.7046 | 3.0864 | 3.6458
Y, Yttrium (ppm) | 1.9698 | 1.8116 | 2.1279 | 1.843 | 2.0966
Yb, Ytterbium (ppm) | 0.178 | 0.15594 | 0.20006 | IND | IND
Zn, Zinc (ppm) | 169.4825 | 163.9037 | 175.0614 | 165.9812 | 172.9839
Zr, Zirconium (ppm) | 16.2335 | 14.8643 | 17.6028 | 15.4058 | 17.0613
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.2977 | 6.17 | 6.4254 | 6.1932 | 6.4021
B, Boron (ppm) | 80.2083 | 65.2773 | 95.1394 | 65.0913 | 95.3254
Ba, Barium (ppm) | 131.4135 | 125.3572 | 137.4698 | 126.7097 | 136.1173
Be, Beryllium (ppm) | 11.9803 | 10.9365 | 13.0242 | IND | IND
Bi, Bismuth (ppm) | 0.74583 | 0.55986 | 0.9318 | IND | IND
Ca, Calcium (wt.%) | 0.22363 | 0.19915 | 0.24811 | 0.20709 | 0.24017
Ce, Cerium (ppm) | 8.8877 | 8.1732 | 9.6022 | 8.4878 | 9.2876
Co, Cobalt (ppm) | 1.58 | 0.96199 | 2.198 | IND | IND
Cs, Caesium (ppm) | 167.5068 | 161.7336 | 173.2801 | 164.9483 | 170.0654
Cu, Copper (ppm) | 20.8927 | 18.2978 | 23.4877 | IND | IND
Dy, Dysprosium (ppm) | 0.68717 | 0.57101 | 0.80333 | 0.6113 | 0.76305
Er, Erbium (ppm) | 0.3663 | 0.30209 | 0.43051 | 0.30583 | 0.42677
Eu, Europium (ppm) | 0.16377 | 0.13929 | 0.18825 | IND | IND
Fe, Iron (wt.%) | 0.62378 | 0.60006 | 0.6475 | 0.60677 | 0.64079
Ga, Gallium (ppm) | 48.4981 | 45.571 | 51.4253 | 46.8692 | 50.1271
Gd, Gadolinium (ppm) | 0.81533 | 0.6346 | 0.99605 | 0.7246 | 0.90605
Ge, Germanium (ppm) | 2.6646 | 2.0987 | 3.2305 | IND | IND
Ho, Holmium (ppm) | 0.13081 | 0.099852 | 0.16177 | IND | IND
In, Indium (ppm) | < 0.2 | IND | IND | IND | IND
K, Potassium (wt.%) | 0.96599 | 0.92719 | 1.0048 | 0.93279 | 0.9992
La, Lanthanum (ppm) | 4.5454 | 4.1794 | 4.9114 | 4.2066 | 4.8841
Li, Lithium (wt.%) | 1.0426 | 1.0165 | 1.0687 | 1.0212 | 1.0639
Li2O, Lithium oxide (wt.%) | 2.2444 | 2.1881 | 2.3006 | 2.1984 | 2.2904
Lu, Lutetium (ppm) | 0.066035 | 0.034083 | 0.097987 | IND | IND
Mg, Magnesium (wt.%) | 0.12778 | 0.11969 | 0.13586 | 0.12435 | 0.13121
Mn, Manganese (wt.%) | 0.13484 | 0.12933 | 0.14035 | 0.13108 | 0.13859
Mo, Molybdenum (ppm) | 3.4976 | 2.7959 | 4.1993 | IND | IND
Nb, Niobium (ppm) | 101.5547 | 97.957 | 105.1524 | 99.2569 | 103.8524
Nd, Neodymium (ppm) | 4.0208 | 3.5634 | 4.4782 | 3.6889 | 4.3527
P, Phosphorus (wt.%) | 0.047133 | 0.04474 | 0.049526 | 0.045051 | 0.049215
Pr, Praseodymium (ppm) | 1.0836 | 0.96268 | 1.2045 | 0.98152 | 1.1857
Rb, Rubidium (ppm) | 1226.4281 | 1164.6762 | 1288.1801 | 1205.3771 | 1247.4792
Re, Rhenium (ppm) | < 0.1 | IND | IND | IND | IND
Sb, Antimony (ppm) | 0.66778 | 0.50746 | 0.82809 | IND | IND
Si, Silicon (wt.%) | 38.0785 | 36.9964 | 39.1605 | 37.495 | 38.6619
Sm, Samarium (ppm) | 0.84026 | 0.68765 | 0.99288 | 0.75751 | 0.92301
Sn, Tin (ppm) | 161.5085 | 154.741 | 168.276 | 156.0377 | 166.9793
Sr, Strontium (ppm) | 47.3626 | 42.9193 | 51.8059 | 44.9007 | 49.8245
Ta, Tantalum (ppm) | 295.1226 | 280.4368 | 309.8084 | 287.8256 | 302.4196
Tb, Terbium (ppm) | 0.12428 | 0.10452 | 0.14404 | IND | IND
Th, Thorium (ppm) | 2.0797 | 1.8907 | 2.2687 | 1.7944 | 2.365
Ti, Titanium (wt.%) | 0.051604 | 0.047685 | 0.055523 | IND | IND
Tl, Thallium (ppm) | 8.8554 | 8.3608 | 9.3501 | 8.4634 | 9.2474
U, Uranium (ppm) | 2.0742 | 1.8549 | 2.2936 | 1.8569 | 2.2915
V, Vanadium (ppm) | 10.5417 | 8.5849 | 12.4984 | IND | IND
W, Tungsten (ppm) | 3.8143 | 2.7884 | 4.8402 | IND | IND
Y, Yttrium (ppm) | 3.4603 | 2.9752 | 3.9453 | 3.0403 | 3.8803
Yb, Ytterbium (ppm) | 0.33744 | 0.22509 | 0.44978 | IND | IND
Zn, Zinc (ppm) | 164.2492 | 151.6092 | 176.8891 | 159.0632 | 169.4351
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 12.0365 | 11.8914 | 12.1815 | 11.9527 | 12.1202
BaO, Barium oxide (ppm) | 150.4346 | 105.3835 | 195.4856 | IND | IND
CaO, Calcium oxide (wt.%) | 0.32299 | 0.31562 | 0.33037 | 0.31455 | 0.33144
Cr2O3, Chromium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 0.89891 | 0.88381 | 0.91402 | 0.88319 | 0.91463
K2O, Potassium oxide (wt.%) | 1.1346 | 1.1162 | 1.1529 | 1.1162 | 1.1529
MgO, Magnesium oxide (wt.%) | 0.21363 | 0.20235 | 0.2249 | 0.20574 | 0.22152
MnO, Manganese oxide (wt.%) | 0.17766 | 0.17109 | 0.18424 | 0.17226 | 0.18306
Na2O, Sodium oxide (wt.%) | 0.55906 | 0.53837 | 0.57976 | 0.54646 | 0.57166
Nb, Niobium (ppm) | 93.2649 | 70.552 | 115.9779 | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.10955 | 0.10392 | 0.11518 | 0.10591 | 0.11318
Rb, Rubidium (ppm) | 1200.0351 | 1121.9381 | 1278.1321 | 1129.5876 | 1270.4826
SiO2, Silicon dioxide (wt.%) | 81.4827 | 80.952 | 82.0134 | 81.0976 | 81.8677
Sn, Tin (ppm) | 150.9245 | 102.8423 | 199.0068 | IND | IND
SO3, Sulphur trioxide (wt.%) | 0.056915 | 0.046454 | 0.067377 | 0.050228 | 0.063603
TiO2, Titanium dioxide (wt.%) | 0.086777 | 0.078869 | 0.094684 | 0.083965 | 0.089588
V2O5, Vanadium(V) oxide (ppm) | < 100 | IND | IND | IND | IND
Zn, Zinc (ppm) | 169.7826 | 145.9621 | 193.6031 | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 0.87179 | 0.78607 | 0.9575 | 0.82966 | 0.91391
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 753b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 4 | In | ppm | 0.023235 | Tm | ppm | 0.031181
Hg | ppm | < 1 | Se | ppm | 0.7519
Borate / Peroxide Fusion ICP
Ag | ppm | < 10 | Hg | ppm | < 5 | Sc | ppm | < 5
As | ppm | 4.0285 | Na | wt.% | 0.44833 | Se | ppm | < 10
Cd | ppm | < 10 | Ni | ppm | 17.2705 | Te | ppm | < 1
Cr | ppm | 39.2489 | Pb | ppm | 5.8873 | Tm | ppm | 0.047408
Hf | ppm | 2.6109 | S | wt.% | 0.028141 | Zr | ppm | 38.8176
Borate Fusion XRF
Ag | ppm | 0.2005 | Hf | ppm | < 80 | Sm | ppm | 0.86667
As | ppm | < 100 | Ho | ppm | 0.11833 | SrO | ppm | 85.4507
Be | ppm | 12.1667 | In | ppm | < 0.1 | Ta | ppm | 262.8492
Bi | ppm | < 100 | La | ppm | 34.4875 | Tb | ppm | 0.13
Ce | ppm | 8.5787 | Lu | ppm | 0.05 | Te | ppm | < 0.1
Co | ppm | < 100 | Mo | ppm | < 50 | Th | ppm | 21.1328
Cs | ppm | 183.2385 | Nd | ppm | 3.8167 | Tl | ppm | 8.23
CuO | ppm | 39.4125 | NiO | ppm | 26.1137 | Tm | ppm | 0.051667
Dy | ppm | 0.64333 | Pb | ppm | 11.4 | U | ppm | 7.0546
Er | ppm | 0.35833 | Pr | ppm | 1.0633 | W | ppm | < 10
Eu | ppm | 0.16833 | Re | ppm | < 0.1 | Y | ppm | 41.9596
Ga | ppm | 45.5833 | Sb | ppm | < 50 | Yb | ppm | 0.35
Gd | ppm | 0.81333 | Sc | ppm | 2.6667 | Zr | ppm | 63.1301
H2O- | wt.% | 0.18006 | Se | ppm | 0.91667
Aqua Regia Digestion
Cs | ppm | 170.4217
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 753b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 0.27288 | 0.042446 | 0.18798 | 0.35777 | 0.14554 | 0.40022 | 15.56% | 31.11% | 46.67% | 0.25923 | 0.28652
Al, wt.% | 6.0109 | 0.35969 | 5.2915 | 6.7302 | 4.9318 | 7.0899 | 5.984% | 11.97% | 17.95% | 5.7103 | 6.3114
As, ppm | 4.7453 | 0.35934 | 4.0266 | 5.4639 | 3.6672 | 5.8233 | 7.573% | 15.15% | 22.72% | 4.508 | 4.9825
Ba, ppm | 130.4267 | 5.8305 | 118.7657 | 142.0877 | 112.9352 | 147.9182 | 4.47% | 8.941% | 13.41% | 123.9054 | 136.948
Be, ppm | 11.6959 | 0.6003 | 10.4953 | 12.8965 | 9.895 | 13.4968 | 5.133% | 10.27% | 15.4% | 11.1111 | 12.2807
Bi, ppm | 0.72074 | 0.048009 | 0.62472 | 0.81676 | 0.57671 | 0.86477 | 6.661% | 13.32% | 19.98% | 0.6847 | 0.75677
Ca, wt.% | 0.22763 | 0.0090645 | 0.2095 | 0.24576 | 0.20043 | 0.25482 | 3.982% | 7.964% | 11.95% | 0.21625 | 0.23901
Cd, ppm | 0.13611 | 0.023257 | 0.089595 | 0.18262 | 0.066337 | 0.20588 | 17.09% | 34.17% | 51.26% | 0.1293 | 0.14291
Ce, ppm | 7.9229 | 1.4779 | 4.9671 | 10.8788 | 3.4891 | 12.3567 | 18.65% | 37.31% | 55.96% | 7.5268 | 8.3191
Co, ppm | 1.3821 | 0.072751 | 1.2366 | 1.5276 | 1.1638 | 1.6003 | 5.264% | 10.53% | 15.79% | 1.313 | 1.4512
Cr, ppm | 16.9538 | 3.7116 | 9.5307 | 24.377 | 5.8191 | 28.0886 | 21.89% | 43.78% | 65.68% | 16.1061 | 17.8015
Cs, ppm | 169.5518 | 9.3695 | 150.8129 | 188.2908 | 141.4434 | 197.6603 | 5.526% | 11.05% | 16.58% | 161.0742 | 178.0294
Cu, ppm | 18.0328 | 0.86599 | 16.3008 | 19.7647 | 15.4348 | 20.6307 | 4.802% | 9.605% | 14.41% | 17.1311 | 18.9344
Dy, ppm | 0.45965 | 0.048511 | 0.36263 | 0.55667 | 0.31411 | 0.60518 | 10.55% | 21.11% | 31.66% | 0.43667 | 0.48263
Er, ppm | 0.20707 | 0.034754 | 0.13757 | 0.27658 | 0.10281 | 0.31134 | 16.78% | 33.57% | 50.35% | 0.19672 | 0.21743
Eu, ppm | 0.15253 | 0.014623 | 0.12329 | 0.18178 | 0.10866 | 0.1964 | 9.587% | 19.17% | 28.76% | 0.14491 | 0.16016
Fe, wt.% | 0.61476 | 0.02657 | 0.56162 | 0.6679 | 0.53505 | 0.69447 | 4.322% | 8.644% | 12.97% | 0.58402 | 0.6455
Ga, ppm | 46.5152 | 2.9415 | 40.6321 | 52.3982 | 37.6906 | 55.3397 | 6.324% | 12.65% | 18.97% | 44.1894 | 48.8409
Gd, ppm | 0.71187 | 0.037293 | 0.63728 | 0.78646 | 0.59999 | 0.82375 | 5.239% | 10.48% | 15.72% | 0.67627 | 0.74746
Ge, ppm | 0.076944 | 0.019982 | 0.036981 | 0.11691 | 0.017 | 0.13689 | 25.97% | 51.94% | 77.91% | 0.073097 | 0.080792
Hf, ppm | 1.8617 | 0.1008 | 1.6601 | 2.0633 | 1.5593 | 2.1641 | 5.415% | 10.83% | 16.24% | 1.7686 | 1.9548
Ho, ppm | 0.076111 | 0.011198 | 0.053714 | 0.098508 | 0.042516 | 0.10971 | 14.71% | 29.43% | 44.14% | 0.072306 | 0.079917
K, wt.% | 0.93837 | 0.040492 | 0.85738 | 1.0194 | 0.81689 | 1.0598 | 4.315% | 8.63% | 12.95% | 0.89145 | 0.98529
La, ppm | 3.8022 | 0.71396 | 2.3743 | 5.2301 | 1.6604 | 5.9441 | 18.78% | 37.55% | 56.33% | 3.6121 | 3.9923
Li, wt.% | 1.0345 | 0.025612 | 0.98332 | 1.0858 | 0.95771 | 1.1114 | 2.476% | 4.951% | 7.427% | 0.98282 | 1.0863
Li2O, wt.% | 2.2271 | 0.055135 | 2.1168 | 2.3373 | 2.0617 | 2.3925 | 2.476% | 4.951% | 7.427% | 2.1157 | 2.3384
Lu, ppm | 0.025 | 0.005644 | 0.013712 | 0.036288 | 0.008068 | 0.041932 | 22.58% | 45.15% | 67.73% | 0.02375 | 0.02625
Mg, wt.% | 0.12043 | 0.0072281 | 0.10598 | 0.13489 | 0.09875 | 0.14212 | 6.002% | 12% | 18.01% | 0.11441 | 0.12646
Mn, wt.% | 0.13262 | 0.0045419 | 0.12354 | 0.14171 | 0.119 | 0.14625 | 3.425% | 6.849% | 10.27% | 0.12599 | 0.13925
Mo, ppm | 3.3625 | 0.18367 | 2.9951 | 3.7298 | 2.8115 | 3.9135 | 5.462% | 10.92% | 16.39% | 3.1943 | 3.5306
Na, wt.% | 0.41235 | 0.014543 | 0.38326 | 0.44143 | 0.36872 | 0.45597 | 3.527% | 7.054% | 10.58% | 0.39173 | 0.43296
Nb, ppm | 102.0064 | 5.2526 | 91.5011 | 112.5116 | 86.2485 | 117.7642 | 5.149% | 10.3% | 15.45% | 96.9061 | 107.1067
Nd, ppm | 3.9442 | 0.24584 | 3.4525 | 4.4358 | 3.2066 | 4.6817 | 6.233% | 12.47% | 18.7% | 3.7469 | 4.1414
Ni, ppm | 8.3177 | 0.43774 | 7.4422 | 9.1931 | 7.0045 | 9.6309 | 5.263% | 10.53% | 15.79% | 7.9018 | 8.7336
P, wt.% | 0.049133 | 0.0025835 | 0.043966 | 0.0543 | 0.041383 | 0.056884 | 5.258% | 10.52% | 15.77% | 0.046677 | 0.05159
Pb, ppm | 4.6367 | 0.25456 | 4.1276 | 5.1458 | 3.873 | 5.4004 | 5.49% | 10.98% | 16.47% | 4.4049 | 4.8685
Pr, ppm | 1.0748 | 0.073161 | 0.92853 | 1.2212 | 0.85537 | 1.2943 | 6.807% | 13.61% | 20.42% | 1.0211 | 1.1286
Rb, ppm | 1215.2833 | 73.4703 | 1068.3426 | 1362.224 | 994.8723 | 1435.6943 | 6.046% | 12.09% | 18.14% | 1154.5191 | 1276.0475
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.025183 | 0.004094 | 0.016995 | 0.033371 | 0.012901 | 0.037465 | 16.26% | 32.51% | 48.77% | 0.023924 | 0.026442
Sb, ppm | 0.64158 | 0.042374 | 0.55683 | 0.72632 | 0.51446 | 0.7687 | 6.605% | 13.21% | 19.81% | 0.6095 | 0.67366
Sc, ppm | 1.0272 | 0.13535 | 0.75653 | 1.2979 | 0.62117 | 1.4333 | 13.18% | 26.35% | 39.53% | 0.97587 | 1.0786
Sm, ppm | 0.81696 | 0.056604 | 0.70375 | 0.93017 | 0.64714 | 0.98677 | 6.929% | 13.86% | 20.79% | 0.77611 | 0.85781
Sn, ppm | 109.4955 | 6.4482 | 96.5991 | 122.392 | 90.1508 | 128.8403 | 5.889% | 11.78% | 17.67% | 104.0208 | 114.9703
Sr, ppm | 38.2516 | 2.3344 | 33.5827 | 42.9204 | 31.2483 | 45.2548 | 6.103% | 12.21% | 18.31% | 36.339 | 40.1641
Ta, ppm | 282.846 | 27.4063 | 228.0335 | 337.6585 | 200.6272 | 365.0647 | 9.689% | 19.38% | 29.07% | 268.7037 | 296.9883
Tb, ppm | 0.10039 | 0.014469 | 0.071452 | 0.12933 | 0.056983 | 0.1438 | 14.41% | 28.83% | 43.24% | 0.09537 | 0.10541
Te, ppm | 0.071203 | 0.014793 | 0.041617 | 0.10079 | 0.026824 | 0.11558 | 20.78% | 41.55% | 62.33% | 0.067643 | 0.074763
Th, ppm | 1.9228 | 0.29731 | 1.3282 | 2.5174 | 1.0309 | 2.8147 | 15.46% | 30.92% | 46.39% | 1.8266 | 2.0189
Ti, wt.% | 0.052081 | 0.0027904 | 0.0465 | 0.057662 | 0.04371 | 0.060452 | 5.358% | 10.72% | 16.07% | 0.049477 | 0.054685
Tl, ppm | 8.7265 | 0.39528 | 7.9359 | 9.517 | 7.5406 | 9.9123 | 4.53% | 9.059% | 13.59% | 8.2901 | 9.1628
U, ppm | 1.8629 | 0.2342 | 1.3945 | 2.3313 | 1.1603 | 2.5655 | 12.57% | 25.14% | 37.72% | 1.7698 | 1.9561
V, ppm | 8.1126 | 0.4266 | 7.2594 | 8.9658 | 6.8328 | 9.3924 | 5.258% | 10.52% | 15.78% | 7.707 | 8.5182
W, ppm | 3.3661 | 0.36106 | 2.644 | 4.0882 | 2.2829 | 4.4493 | 10.73% | 21.45% | 32.18% | 3.1978 | 3.5344
Y, ppm | 1.9698 | 0.23438 | 1.501 | 2.4385 | 1.2666 | 2.6729 | 11.9% | 23.8% | 35.7% | 1.8713 | 2.0683
Yb, ppm | 0.178 | 0.032615 | 0.11277 | 0.24323 | 0.080155 | 0.27584 | 18.32% | 36.65% | 54.97% | 0.1691 | 0.1869
Zn, ppm | 169.4825 | 9.035 | 151.4126 | 187.5525 | 142.3776 | 196.5874 | 5.331% | 10.66% | 15.99% | 161.0084 | 177.9567
Zr, ppm | 16.2335 | 1.5889 | 13.0558 | 19.4113 | 11.467 | 21.0001 | 9.788% | 19.58% | 29.36% | 15.4219 | 17.0452
Borate / Peroxide Fusion ICP
Al, wt.% | 6.2977 | 0.11239 | 6.0729 | 6.5225 | 5.9605 | 6.6349 | 1.785% | 3.569% | 5.354% | 5.9828 | 6.6126
B, ppm | 80.2083 | 13.75 | 52.7084 | 107.7083 | 38.9584 | 121.4583 | 17.14% | 34.29% | 51.43% | 76.1979 | 84.2188
Ba, ppm | 131.4135 | 7.0297 | 117.3541 | 145.4729 | 110.3244 | 152.5026 | 5.349% | 10.7% | 16.05% | 124.8428 | 137.9842
Be, ppm | 11.9803 | 1.041 | 9.8983 | 14.0624 | 8.8572 | 15.1034 | 8.69% | 17.38% | 26.07% | 11.3813 | 12.5793
Bi, ppm | 0.74583 | 0.10711 | 0.53162 | 0.96004 | 0.42452 | 1.0672 | 14.36% | 28.72% | 43.08% | 0.70854 | 0.78312
Ca, wt.% | 0.22363 | 0.025683 | 0.17226 | 0.275 | 0.14658 | 0.30068 | 11.48% | 22.97% | 34.45% | 0.21245 | 0.23481
Ce, ppm | 8.8877 | 0.93902 | 7.0097 | 10.7658 | 6.0707 | 11.7048 | 10.57% | 21.13% | 31.7% | 8.4433 | 9.3321
Co, ppm | 1.58 | 0.38409 | 0.81184 | 2.3482 | 0.42774 | 2.7323 | 24.31% | 48.62% | 72.93% | 1.501 | 1.659
Cs, ppm | 167.5068 | 5.3267 | 156.8534 | 178.1603 | 151.5266 | 183.4871 | 3.18% | 6.36% | 9.54% | 159.1315 | 175.8822
Cu, ppm | 20.8927 | 2.6068 | 15.6792 | 26.1062 | 13.0725 | 28.713 | 12.48% | 24.95% | 37.43% | 19.8481 | 21.9374
Dy, ppm | 0.68717 | 0.071284 | 0.54461 | 0.82974 | 0.47332 | 0.90102 | 10.37% | 20.75% | 31.12% | 0.65281 | 0.72153
Er, ppm | 0.3663 | 0.050845 | 0.26461 | 0.46799 | 0.21377 | 0.51883 | 13.88% | 27.76% | 41.64% | 0.34799 | 0.38462
Eu, ppm | 0.16377 | 0.022879 | 0.11801 | 0.20953 | 0.095131 | 0.2324 | 13.97% | 27.94% | 41.91% | 0.15558 | 0.17196
Fe, wt.% | 0.62378 | 0.022355 | 0.57907 | 0.66849 | 0.55672 | 0.69084 | 3.584% | 7.167% | 10.75% | 0.59259 | 0.65497
Ga, ppm | 48.4981 | 2.1312 | 44.2357 | 52.7606 | 42.1045 | 54.8918 | 4.394% | 8.789% | 13.18% | 46.0732 | 50.923
Gd, ppm | 0.81533 | 0.097524 | 0.62028 | 1.0104 | 0.52276 | 1.1079 | 11.96% | 23.92% | 35.88% | 0.77456 | 0.8561
Ge, ppm | 2.6646 | 0.45363 | 1.7573 | 3.5719 | 1.3037 | 4.0255 | 17.02% | 34.05% | 51.07% | 2.5314 | 2.7978
Ho, ppm | 0.13081 | 0.020891 | 0.089028 | 0.17259 | 0.068137 | 0.19349 | 15.97% | 31.94% | 47.91% | 0.12427 | 0.13735
In, ppm | < 0.2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K, wt.% | 0.96599 | 0.052725 | 0.86054 | 1.0714 | 0.80782 | 1.1242 | 5.458% | 10.92% | 16.37% | 0.91769 | 1.0143
La, ppm | 4.5454 | 0.43998 | 3.6654 | 5.4253 | 3.2254 | 5.8653 | 9.68% | 19.36% | 29.04% | 4.3181 | 4.7726
Li, wt.% | 1.0426 | 0.033645 | 0.97529 | 1.1099 | 0.94165 | 1.1435 | 3.227% | 6.454% | 9.681% | 0.99045 | 1.0947
Li2O, wt.% | 2.2444 | 0.072427 | 2.0995 | 2.3892 | 2.0271 | 2.4616 | 3.227% | 6.454% | 9.681% | 2.1321 | 2.3566
Lu, ppm | 0.066035 | 0.017435 | 0.031166 | 0.1009 | 0.013731 | 0.11834 | 26.4% | 52.8% | 79.21% | 0.062733 | 0.069337
Mg, wt.% | 0.12778 | 0.0068988 | 0.11398 | 0.14158 | 0.10708 | 0.14847 | 5.399% | 10.8% | 16.2% | 0.12139 | 0.13417
Mn, wt.% | 0.13484 | 0.0051176 | 0.1246 | 0.14507 | 0.11949 | 0.15019 | 3.795% | 7.591% | 11.39% | 0.1281 | 0.14158
Mo, ppm | 3.4976 | 0.68627 | 2.1251 | 4.8702 | 1.4388 | 5.5564 | 19.62% | 39.24% | 58.86% | 3.3227 | 3.6725
Nb, ppm | 101.5547 | 3.3689 | 94.8169 | 108.2925 | 91.448 | 111.6613 | 3.317% | 6.635% | 9.952% | 96.4769 | 106.6324
Nd, ppm | 4.0208 | 0.49309 | 3.0346 | 5.007 | 2.5415 | 5.5001 | 12.26% | 24.53% | 36.79% | 3.8198 | 4.2218
P, wt.% | 0.047133 | 0.0036594 | 0.039814 | 0.054452 | 0.036155 | 0.058111 | 7.764% | 15.53% | 23.29% | 0.044776 | 0.04949
Pr, ppm | 1.0836 | 0.13166 | 0.82027 | 1.3469 | 0.68861 | 1.4786 | 12.15% | 24.3% | 36.45% | 1.0294 | 1.1378
Rb, ppm | 1226.4281 | 46.3665 | 1133.6952 | 1319.1611 | 1087.3287 | 1365.5276 | 3.781% | 7.561% | 11.34% | 1165.1067 | 1287.7495
Re, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sb, ppm | 0.66778 | 0.096512 | 0.47475 | 0.8608 | 0.37824 | 0.95731 | 14.45% | 28.91% | 43.36% | 0.63439 | 0.70117
Si, wt.% | 38.0785 | 1.2759 | 35.5267 | 40.6303 | 34.2508 | 41.9062 | 3.351% | 6.701% | 10.05% | 36.1745 | 39.9824
Sm, ppm | 0.84026 | 0.11313 | 0.614 | 1.0665 | 0.50087 | 1.1797 | 13.46% | 26.93% | 40.39% | 0.79825 | 0.88227
Sn, ppm | 161.5085 | 11.0428 | 139.423 | 183.594 | 128.3802 | 194.6368 | 6.837% | 13.67% | 20.51% | 153.4331 | 169.5839
Sr, ppm | 47.3626 | 6.1984 | 34.9657 | 59.7594 | 28.7673 | 65.9579 | 13.09% | 26.17% | 39.26% | 44.9945 | 49.7307
Ta, ppm | 295.1226 | 22.192 | 250.7386 | 339.5066 | 228.5467 | 361.6986 | 7.52% | 15.04% | 22.56% | 280.3665 | 309.8787
Tb, ppm | 0.12428 | 0.016243 | 0.091795 | 0.15677 | 0.075553 | 0.17301 | 13.07% | 26.14% | 39.21% | 0.11807 | 0.13049
Th, ppm | 2.0797 | 0.20862 | 1.6624 | 2.4969 | 1.4538 | 2.7055 | 10.03% | 20.06% | 30.09% | 1.9757 | 2.1836
Ti, wt.% | 0.051604 | 0.0031475 | 0.045309 | 0.057899 | 0.042161 | 0.061046 | 6.099% | 12.2% | 18.3% | 0.049024 | 0.054184
Tl, ppm | 8.8554 | 0.44823 | 7.959 | 9.7519 | 7.5107 | 10.2001 | 5.062% | 10.12% | 15.18% | 8.4126 | 9.2982
U, ppm | 2.0742 | 0.21322 | 1.6478 | 2.5007 | 1.4346 | 2.7139 | 10.28% | 20.56% | 30.84% | 1.9705 | 2.1779
V, ppm | 10.5417 | 2.168 | 6.2057 | 14.8777 | 4.0377 | 17.0457 | 20.57% | 41.13% | 61.7% | 10.0146 | 11.0687
W, ppm | 3.8143 | 0.64054 | 2.5333 | 5.0954 | 1.8927 | 5.736 | 16.79% | 33.59% | 50.38% | 3.6236 | 4.0051
Y, ppm | 3.4603 | 0.45594 | 2.5484 | 4.3721 | 2.0925 | 4.8281 | 13.18% | 26.35% | 39.53% | 3.2873 | 3.6333
Yb, ppm | 0.33744 | 0.060161 | 0.21712 | 0.45776 | 0.15696 | 0.51792 | 17.83% | 35.66% | 53.49% | 0.32057 | 0.35431
Zn, ppm | 164.2492 | 13.7569 | 136.7353 | 191.763 | 122.9784 | 205.52 | 8.376% | 16.75% | 25.13% | 156.0367 | 172.4616
Borate Fusion XRF
Al2O3, wt.% | 12.0365 | 0.20268 | 11.6311 | 12.4418 | 11.4284 | 12.6445 | 1.684% | 3.368% | 5.052% | 11.4346 | 12.6383
BaO, ppm | 150.4346 | 43.001 | 64.4326 | 236.4366 | 21.4316 | 279.4376 | 28.58% | 57.17% | 85.75% | 142.9128 | 157.9563
CaO, wt.% | 0.32299 | 0.0074419 | 0.30811 | 0.33788 | 0.30067 | 0.34532 | 2.304% | 4.608% | 6.912% | 0.30684 | 0.33914
Cr2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Fe2O3, wt.% | 0.89891 | 0.022253 | 0.85441 | 0.94342 | 0.83215 | 0.96567 | 2.476% | 4.951% | 7.427% | 0.85397 | 0.94386
K2O, wt.% | 1.1346 | 0.021997 | 1.0906 | 1.1786 | 1.0686 | 1.2005 | 1.939% | 3.878% | 5.816% | 1.0778 | 1.1913
MgO, wt.% | 0.21363 | 0.018171 | 0.17728 | 0.24997 | 0.15911 | 0.26814 | 8.506% | 17.01% | 25.52% | 0.20295 | 0.22431
MnO, wt.% | 0.17766 | 0.0060911 | 0.16548 | 0.18985 | 0.15939 | 0.19594 | 3.428% | 6.857% | 10.29% | 0.16878 | 0.18655
Na2O, wt.% | 0.55906 | 0.02877 | 0.50152 | 0.6166 | 0.47275 | 0.64537 | 5.146% | 10.29% | 15.44% | 0.53111 | 0.58702
Nb, ppm | 93.2649 | 23.3938 | 46.4773 | 140.0526 | 23.0834 | 163.4464 | 25.08% | 50.17% | 75.25% | 88.6017 | 97.9282
P2O5, wt.% | 0.10955 | 0.0057593 | 0.098029 | 0.12107 | 0.09227 | 0.12683 | 5.257% | 10.51% | 15.77% | 0.10407 | 0.11502
Rb, ppm | 1200.0351 | 52.6435 | 1094.7482 | 1305.322 | 1042.1047 | 1357.9655 | 4.387% | 8.774% | 13.16% | 1140.0334 | 1260.0369
SiO2, wt.% | 81.4827 | 0.90332 | 79.6761 | 83.2893 | 78.7727 | 84.1927 | 1.109% | 2.217% | 3.326% | 77.4086 | 85.5568
Sn, ppm | 150.9245 | 37.1866 | 76.5514 | 225.2977 | 39.3649 | 262.4842 | 24.64% | 49.28% | 73.92% | 143.3783 | 158.4708
SO3, wt.% | 0.056915 | 0.013501 | 0.029914 | 0.083916 | 0.016413 | 0.097417 | 23.72% | 47.44% | 71.16% | 0.054069 | 0.059761
TiO2, wt.% | 0.086777 | 0.0064525 | 0.073872 | 0.099682 | 0.067419 | 0.10613 | 7.436% | 14.87% | 22.31% | 0.082438 | 0.091115
V2O5, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Zn, ppm | 169.7826 | 17.0335 | 135.7156 | 203.8496 | 118.6821 | 220.883 | 10.03% | 20.07% | 30.1% | 161.2935 | 178.2717
Thermogravimetry
LOI1000, wt.% | 0.87179 | 0.083242 | 0.7053 | 1.0383 | 0.62206 | 1.1215 | 9.548% | 19.1% | 28.65% | 0.8282 | 0.91538
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 753b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 753b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
African Natural Resources & Mines Ltd, Suleja, Niger State, Nigeria
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
CERTIMIN, Lima, Peru
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Labwest Minerals Analysis, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 753b (Certified Value 0.273 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.26 | 0.2433 | 0.27 | 0.3 | 0.31 | 0.2 | 0.3 | 0.277 | < 0.3 | 0.26 | 0.37 | 0.17 | 0.15 | 0.22 | 0.26 | 0.25 | 0.22 | < 0.5 | 0.3 | 0.5 | < 0.5 | 1.33 | 0.34 |  |  |  | 1
| 1 | 2 | 0.28 | 0.26118 | 0.3 | 0.26 | 0.3 | 0.3 | 0.3 | 0.228 | < 0.3 | 0.27 | 0.36 | 0.17 | 0.16 | 0.21 | 0.27 | 0.2 | 0.25 | < 0.5 | 0.3 | 0.6 | < 0.5 | 1.36 | 0.35 |  |  |  | 4
| 1 | 3 | 0.27 | 0.24998 | 0.28 | 0.28 | 0.3 | 0.3 | 0.3 | 0.252 | < 0.3 | 0.27 | 0.33 | 0.16 | 0.27 | 0.19 | 0.24 | 0.26 | 0.23 | < 0.5 | 0.3 | 0.7 | < 0.5 | 1.32 | 0.34 |  |  |  | 16
| 1 | 4 | 0.28 | 0.18966 | 0.28 | 0.32 | 0.26 | 0.3 | 0.3 | 0.256 | < 0.3 | 0.26 | 0.34 | 0.15 | 0.27 | 0.24 | 0.31 | 0.28 | 0.24 | < 0.5 | 0.3 | 0.7 | < 0.5 | 1.38 | 0.33 |  |  |  | 0.27288
| 1 | 5 | 0.29 | 0.25901 | 0.28 | 0.32 | 0.32 | 0.3 | 0.3 | 0.261 | < 0.3 | 0.26 | 0.31 | 0.16 | 0.28 | 0.24 | 0.3 | 0.28 | 0.23 | < 0.5 | 0.3 | 0.8 | < 0.5 | 1.39 | 0.31 |  |  |  | 7
| 1 | 6 | 0.27 | 0.25292 | 0.29 | 0.31 | 0.29 | 0.2 | 0.3 | 0.238 | < 0.3 | 0.27 | 0.33 | 0.16 | 0.23 | 0.21 | 0.36 | 0.25 | 0.25 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 1.34 | 0.32
| Mean |  | 0.275 | 0.24268 | 0.28333 | 0.29833 | 0.29667 | 0.26667 | 0.3 | 0.252 |  | 0.265 | 0.34 | 0.16167 | 0.22667 | 0.21833 | 0.29 | 0.25333 | 0.23667 |  | 0.3 | 0.66 |  | 1.3533 | 0.33167
| Median |  | 0.275 | 0.25145 | 0.28 | 0.305 | 0.3 | 0.3 | 0.3 | 0.254 |  | 0.265 | 0.335 | 0.16 | 0.25 | 0.215 | 0.285 | 0.255 | 0.235 |  | 0.3 | 0.7 |  | 1.35 | 0.335
| Std Dev. |  | 0.010488 | 0.026751 | 0.010328 | 0.024014 | 0.020656 | 0.05164 | 0 | 0.017286 |  | 0.0054772 | 0.021909 | 0.0075277 | 0.058195 | 0.019408 | 0.042895 | 0.029439 | 0.012111 |  | 0 | 0.11402 |  | 0.028048 | 0.01472
| Rel.Std.Dev. |  | 3.814% | 11.02% | 3.645% | 8.049% | 6.963% | 19.36% | 0% | 6.859% |  | 2.067% | 6.444% | 4.656% | 25.67% | 8.889% | 14.79% | 11.62% | 5.117% |  | 0% | 17.28% |  | 2.072% | 4.438%
| PDM3 |  | 0.778% | -11.07% | 3.832% | 9.329% | 8.718% | -2.276% | 9.94% | -7.651% |  | -2.887% | 24.6% | -40.75% | -16.93% | -19.99% | 6.275% | -7.162% | -13.27% |  | 9.94% | 141.9% |  | 395.9% | 21.54%
| Z-Score (Absolute) |  | 0.34 | 0.35 | 0.51 | 0.83 | 0.8 | NA | NA | 0.15 | 2.32 | 0.12 | 1.72 | 2.08 | 0.69 | 0.87 | 0.66 | 0.12 | 0.48 | 0.2 | NA | 7.08 | 0.2 | 23.3 | 1.54
Results from laboratories 6, 7 and 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 753b (Certified Value 6.01 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 6.0543 | 5.8863 | 6.22 | 5.96 | 6.4577 | 5.9 | 6.18 | 5.44 | 6.01 | 6.7 | 6.15 | 5.74 | 5.47 | 5.63 | 5.51 | 5.92 | 5.09 | 6.18 | 6.2751 | 6.11 | 5.74 | 6.16 | 6.192 | 6.84 | 5.29 |  | 1
| 1 | 2 | 6.2138 | 5.8875 | 6.16 | 5.93 | 6.5188 | 5.73 | 6.253 | 4.74 | 6.14 | 6.67 | 6.65 | 5.67 | 5.58 | 5.49 | 5.61 | 6.05 | 4.89 | 6.2 | 6.2856 | 6.01 | 5.86 | 5.83 | 6.245 | 6.38 | 5.27 |  | 0
| 1 | 3 | 6.312 | 5.8339 | 6.31 | 5.96 | 6.4687 | 5.63 | 6.074 | 5.87 | 5.74 | 6.49 | 7.01 | 5.76 | 5.79 | 5.41 | 5.66 | 5.88 | 5.03 | 6.18 | 6.2503 | 6 | 6.05 | 6.04 | 6.192 | 6.39 | 5.51 |  | 16
| 1 | 4 | 6.2413 | 5.6973 | 6.13 | 5.96 | 6.5223 | 5.82 | 6.201 | 5.42 | 6.06 | 6.52 | 6.58 | 5.73 | 4.71 | 5.68 | 5.6 | 5.94 | 4.92 | 6.12 | 6.2728 | 6.18 | 5.84 | 5.94 | 6.192 | 6.56 | 5.64 |  | 6.0109
| 1 | 5 | 6.1641 | 5.8013 | 5.93 | 5.96 | 6.4836 | 5.84 | 6.129 | 5.67 | 6.06 | 6.47 | 6.62 | 5.64 | 5.05 | 5.52 | 5.5 | 5.72 | 5.11 | 6.14 | 6.2569 | 6.05 | 4.43 | 5.38 | 6.192 | 6.3 | 5.39 |  | 8
| 1 | 6 | 6.479 | 5.7479 | 6 | 6.04 | 6.4987 | 5.8 | 6.069 | 5.66 | 6.06 | 6.78 | 6.9 | 5.78 | 5.25 | 5.56 | 5.81 | 5.92 | 4.92 | 6.2 | 6.2671 | 6.04 | 4.6 | 5.93 | 6.192 | 6.43 | 6.07
| Mean |  | 6.2441 | 5.8091 | 6.125 | 5.9683 | 6.4916 | 5.7867 | 6.151 | 5.4667 | 6.0117 | 6.605 | 6.6517 | 5.72 | 5.3083 | 5.5483 | 5.615 | 5.905 | 4.9933 | 6.17 | 6.268 | 6.065 | 5.42 | 5.88 | 6.2008 | 6.4833 | 5.5283
| Median |  | 6.2275 | 5.8176 | 6.145 | 5.96 | 6.4912 | 5.81 | 6.1545 | 5.55 | 6.06 | 6.595 | 6.635 | 5.735 | 5.36 | 5.54 | 5.605 | 5.92 | 4.975 | 6.18 | 6.2699 | 6.045 | 5.79 | 5.935 | 6.192 | 6.41 | 5.45
| Std Dev. |  | 0.14362 | 0.076185 | 0.14011 | 0.037103 | 0.026345 | 0.094587 | 0.073324 | 0.39302 | 0.13949 | 0.12849 | 0.29929 | 0.054037 | 0.38989 | 0.097451 | 0.11362 | 0.10728 | 0.095638 | 0.032863 | 0.012798 | 0.068337 | 0.71018 | 0.2693 | 0.021637 | 0.19439 | 0.29963
| Rel.Std.Dev. |  | 2.3% | 1.311% | 2.287% | 0.6217% | 0.4058% | 1.635% | 1.192% | 7.189% | 2.32% | 1.945% | 4.5% | 0.9447% | 7.345% | 1.756% | 2.024% | 1.817% | 1.915% | 0.5326% | 0.2042% | 1.127% | 13.1% | 4.58% | 0.3489% | 2.998% | 5.42%
| PDM3 |  | 3.88% | -3.357% | 1.899% | -0.7074% | 7.999% | -3.73% | 2.332% | -9.053% | 0.01352% | 9.885% | 10.66% | -4.839% | -11.69% | -7.695% | -6.586% | -1.761% | -16.93% | 2.648% | 4.277% | 0.9008% | -9.83% | -2.177% | 3.161% | 7.86% | -8.027%
| Z-Score (Absolute) |  | 0.67 | 0.39 | 0.38 | 0 | 1.28 | 0.44 | 0.45 | 1.23 | 0.11 | 1.56 | 1.67 | 0.61 | 1.61 | 1.03 | 0.86 | 0.15 | 2.38 | 0.49 | 0.73 | 0.24 | 1.34 | 0.22 | 0.57 | 1.26 | 1.08
| Analytical results for As in OREAS 753b (Certified Value 4.75 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.28 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 4.7 | 5.9166 | 4.7 | 5 | 4.4 | 5 | 5 | 5.5 | 5.3 | 4.7 | 4.5 | 4.8 | 5.3 | 4.4 | 4.7 | 4.9 | 4 | 3.6 | 4.7 | 4 | < 5 | 3.32 |  |  |  |  | 1
| 1 | 2 | 4.9 | 6.1581 | 4.8 | 5 | 4.6 | 5 | 5 | 4.4 | 4.7 | 4.9 | 4.5 | 4.8 | 4.8 | 4.2 | 4.5 | 5.9 | 5 | 3.3 | 4.9 | 5 | < 5 | 3.44 |  |  |  |  | 5
| 1 | 3 | 4.9 | 6.1804 | 4.7 | 4 | 4.5 | 5 | 5 | 4.9 | 3.9 | 5 | 4.4 | 4.6 | 4.6 | 4.9 | 4.7 | 3.5 | 4 | 3.1 | 4.7 | 5 | < 5 | 3.37 |  |  |  |  | 16
| 1 | 4 | 4.6 | 5.814 | 4.4 | 4 | 4.6 | 5 | 5 | 5 | 5.2 | 5 | 4.5 | 4.5 | 4.8 | 4.7 | 5.2 | 5.8 | 5 | 3.6 | 4.8 | 4 | 5 | 3.53 |  |  |  |  | 4.7453
| 1 | 5 | 4.8 | 5.9252 | 4.6 | 5 | 4.7 | 5 | 5 | 4.9 | 5.6 | 4.9 | 4.6 | 4.9 | 4.8 | 4.1 | 5 | 2.4 | 4 | 3.3 | 5 | 4 | < 5 | 3.71 |  |  |  |  | 9
| 1 | 6 | 4.9 | 5.6398 | 4.5 | 5 | 4.7 | 5 | 5 | 4.6 | 4.7 | 4.7 | 4.4 | 4.6 | 4.9 | 4.5 | 4.5 | 6.6 | 5 | 3.4 | 4.8 | 4 | 5 | 3.69
| Mean |  | 4.8 | 5.939 | 4.6167 | 4.6667 | 4.5833 | 5 | 5 | 4.8833 | 4.9 | 4.8667 | 4.4833 | 4.7 | 4.8667 | 4.4667 | 4.7667 | 4.85 | 4.5 | 3.3833 | 4.8167 | 4.3333 | 5 | 3.51
| Median |  | 4.85 | 5.9209 | 4.65 | 5 | 4.6 | 5 | 5 | 4.9 | 4.95 | 4.9 | 4.5 | 4.7 | 4.8 | 4.45 | 4.7 | 5.35 | 4.5 | 3.35 | 4.8 | 4 | 5 | 3.485
| Std Dev. |  | 0.12649 | 0.20596 | 0.1472 | 0.5164 | 0.1169 | 0 | 0 | 0.37639 | 0.60332 | 0.13663 | 0.075277 | 0.15492 | 0.23381 | 0.30111 | 0.28048 | 1.6059 | 0.54772 | 0.19408 | 0.1169 | 0.5164 | 0 | 0.16334
| Rel.Std.Dev. |  | 2.635% | 3.468% | 3.188% | 11.07% | 2.551% | 0% | 0% | 7.708% | 12.31% | 2.807% | 1.679% | 3.296% | 4.804% | 6.741% | 5.884% | 33.11% | 12.17% | 5.736% | 2.427% | 11.92% | 0% | 4.654%
| PDM3 |  | 1.154% | 25.16% | -2.71% | -1.656% | -3.412% | 5.368% | 5.368% | 2.91% | 3.261% | 2.559% | -5.52% | -0.9537% | 2.559% | -5.871% | 0.4512% | 2.207% | -5.168% | -28.7% | 1.505% | -8.681% | 5.368% | -26.03%
| Z-Score (Absolute) |  | 0.1 | 7.19 | 1.04 | NA | 1.24 | NA | NA | 0.62 | 0.73 | 0.52 | 1.87 | 0.52 | 0.52 | 1.97 | 0.1 | 0.41 | NA | 8.71 | 0.21 | NA | NA | 7.93
| Analytical results for B in OREAS 753b (Indicative Value 4 ppm) | [27 empty] | Indicative
B | Round | Replicate | 1.01: Lab |  |  |  |  | 1
| No | No | 1.01: 1.19 |  |  |  |  | ppm
|  |  | 1.01: 4A*OES |  |  |  |  | 2
2
| 1 | 1 | 1.01: 4 |  |  |  |  | 1
| 1 | 2 | 1.01: 4 |  |  |  |  | 1
| 1 | 3 | 1.01: 4 |  |  |  |  | 16
| 1 | 4 | 1.01: 4 |  |  |  |  | 4
| 1 | 5 | 1.01: 4 |  |  |  |  | 7
| 1 | 6 | 1.01: 4
| Mean |  | 1.01: 4
| Median |  | 1.01: 4
| Std Dev. |  | 1.01: 0
Rel.Std.Dev. | bar: 1.01
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ba in OREAS 753b (Certified Value 130 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 0
0
| 1 | 1 | 130.5 | 127.2964 | 148 | 133 | 140.3 | 129 | 133 | 137 | 128 | 148 | 127.6648 | 144 | 140 | 130 | 130 | 130 | 140 | 119 | 133 | 123.948 | 133 | 130 | 131 | 130 | 119.59 |  | 1
| 1 | 2 | 127.2 | 130.8773 | 143 | 132 | 139.2 | 129 | 135 | 129 | 127 | 148 | 127.4572 | 144 | 130 | 130 | 120 | 130 | 140 | 126 | 128 | 125.059 | 132 | 134 | 129 | 133 | 117.35 |  | 0
| 1 | 3 | 129.1 | 128.9167 | 140 | 132 | 138.4 | 126 | 136 | 141 | 125 | 144 | 126.9501 | 144 | 130 | 130 | 120 | 130 | 140 | 121 | 124 | 126.951 | 128 | 132 | 128 | 128 | 118.72 |  | 16
| 1 | 4 | 129.8 | 126.1209 | 137 | 133 | 137.4 | 133 | 134 | 136 | 127 | 146 | 129.6021 | 141 | 130 | 120 | 130 | 130 | 140 | 127 | 147 | 125.928 | 137 | 129 | 129 | 129 | 121.3 |  | 130.4267
| 1 | 5 | 129.3 | 131.6173 | 136 | 133 | 136.7 | 131 | 132 | 138 | 132 | 146 | 127.4504 | 136 | 130 | 120 | 130 | 130 | 130 | 120 | 129 | 124.96 | 132 | 115 | 112 | 131 | 119.96 |  | 10
| 1 | 6 | 132.5 | 128.2792 | 137 | 134 | 139 | 132 | 133 | 139 | 129 | 153 | 130.5658 | 146 | 130 | 130 | 130 | 130 | 130 | 123 | 133 | 125.422 | 133 | 126 | 114 | 130 | 124.9
| Mean |  | 129.7333 | 128.8513 | 140.1667 | 132.8333 | 138.5 | 130 | 133.8333 | 136.6667 | 128 | 147.5 | 128.2817 | 142.5 | 131.6667 | 126.6667 | 126.6667 | 130 | 136.6667 | 122.6667 | 132.3333 | 125.378 | 132.5 | 127.6667 | 123.8333 | 130.1667 | 120.3033
| Median |  | 129.55 | 128.598 | 138.5 | 133 | 138.7 | 130 | 133.5 | 137.5 | 127.5 | 147 | 127.561 | 144 | 130 | 130 | 130 | 130 | 140 | 122 | 131 | 125.2405 | 132.5 | 129.5 | 128.5 | 130 | 119.775
| Std Dev. |  | 1.7466 | 2.0957 | 4.6224 | 0.75277 | 1.2992 | 2.5298 | 1.472 | 4.1312 | 2.3664 | 3.0822 | 1.4481 | 3.5637 | 4.0825 | 5.164 | 5.164 | 0 | 5.164 | 3.266 | 7.9415 | 1.01 | 2.881 | 6.7725 | 8.4715 | 1.7224 | 2.6069
| Rel.Std.Dev. |  | 1.346% | 1.626% | 3.298% | 0.5667% | 0.9381% | 1.946% | 1.1% | 3.023% | 1.849% | 2.09% | 1.129% | 2.501% | 3.101% | 4.077% | 4.077% | 0% | 3.779% | 2.662% | 6.001% | 0.8055% | 2.174% | 5.305% | 6.841% | 1.323% | 2.167%
| PDM3 |  | -0.5316% | -1.208% | 7.468% | 1.845% | 6.19% | -0.3272% | 2.612% | 4.784% | -1.861% | 13.09% | -1.645% | 9.257% | 0.9507% | -2.883% | -2.883% | -0.3272% | 4.784% | -5.95% | 1.462% | -3.871% | 1.59% | -2.116% | -5.055% | -0.1994% | -7.762%
| Z-Score (Absolute) |  | 0.05 | 0.23 | 2.06 | 0.57 | 1.72 | 0 | 0.78 | 1.35 | 0.4 | 3.54 | 0.35 | 2.53 | 0.34 | 0.67 | 0.67 | 0 | 1.35 | 1.48 | 0.47 | 0.93 | 0.51 | 0.47 | 1.25 | 0.03 | 1.96
| Analytical results for Be in OREAS 753b (Certified Value 11.7 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 12.14 | 12.0728 | 12.88 | 11.68 | 11.43 | 11.5 | 12.5 | 10 | 12.1 | 11.7 | 11.4284 | 9.96 | 11.6 | 11.5 | 10.9 | 11.45 | 11.95 | 11.2 | 12 | 12.5964 | 11.6 | 12.4 | 12 | 13.5 | 10.38 | 10.03 | 1
| 1 | 2 | 11.88 | 12.2647 | 11.81 | 12.39 | 11.31 | 11.6 | 12.6 | 9 | 12 | 11.4 | 11.7102 | 9.51 | 11.5 | 11.7 | 10.65 | 11.45 | 12.1 | 10.7 | 12 | 12.4524 | 11.6 | 12.7 | 12 | 13.5 | 9.97 | 9.94 | 12
| 1 | 3 | 11.95 | 11.8892 | 11.88 | 11.71 | 11.8 | 11.5 | 12.7 | 10 | 11.9 | 10.7 | 11.6947 | 9.1 | 11.5 | 11.4 | 10.65 | 11.65 | 11.8 | 11.3 | 12 | 12.466 | 11.4 | 12.5 | 12 | 14.2 | 10.1 | 10.04 | 16
| 1 | 4 | 12.04 | 11.7465 | 12.4 | 11.81 | 11.16 | 11.3 | 12.6 | 10 | 11.8 | 11 | 11.6112 | 9.49 | 11.55 | 10.85 | 11 | 11.55 | 11.95 | 10.8 | 12 | 12.4316 | 11.9 | 12.6 | 12 | 13.9 | 10.28 | 10.88 | 11.6959
| 1 | 5 | 11.85 | 11.8758 | 11.92 | 11.59 | 11.31 | 11.6 | 12.6 | 10 | 12.5 | 11.6 | 11.8887 | 9.28 | 11.45 | 10.95 | 10.9 | 11.3 | 11.55 | 11.1 | 12 | 12.506 | 11.4 | 12.4 | 13 | 13.8 | 10.6 | 10.74 | 11
| 1 | 6 | 12.22 | 11.8555 | 12.52 | 12.06 | 11.38 | 11.6 | 12.8 | 9 | 12 | 12.3 | 11.8272 | 9.01 | 11.65 | 11.1 | 10.9 | 11.95 | 11.95 | 10.9 | 12 | 12.47 | 11.6 | 12.4 | 12 | 13.7 | 10.64 | 9.92
| Mean |  | 12.0133 | 11.9507 | 12.235 | 11.8733 | 11.3983 | 11.5167 | 12.6333 | 9.6667 | 12.05 | 11.45 | 11.6934 | 9.3917 | 11.5417 | 11.25 | 10.8333 | 11.5583 | 11.8833 | 11 | 12 | 12.4871 | 11.5833 | 12.5 | 12.1667 | 13.7667 | 10.3283 | 10.2583
| Median |  | 11.995 | 11.8825 | 12.16 | 11.76 | 11.345 | 11.55 | 12.6 | 10 | 12 | 11.5 | 11.7025 | 9.385 | 11.525 | 11.25 | 10.9 | 11.5 | 11.95 | 11 | 12 | 12.468 | 11.6 | 12.45 | 12 | 13.75 | 10.33 | 10.035
| Std Dev. |  | 0.14692 | 0.18636 | 0.43136 | 0.30004 | 0.21683 | 0.1169 | 0.10328 | 0.5164 | 0.2429 | 0.56125 | 0.16333 | 0.34348 | 0.073598 | 0.33466 | 0.1472 | 0.22454 | 0.18886 | 0.23664 | 0 | 0.058865 | 0.18348 | 0.12649 | 0.40825 | 0.26583 | 0.26701 | 0.43222
| Rel.Std.Dev. |  | 1.223% | 1.559% | 3.526% | 2.527% | 1.902% | 1.015% | 0.8175% | 5.342% | 2.016% | 4.902% | 1.397% | 3.657% | 0.6377% | 2.975% | 1.359% | 1.943% | 1.589% | 2.151% | 0% | 0.4714% | 1.584% | 1.012% | 3.355% | 1.931% | 2.585% | 4.213%
| PDM3 |  | 2.714% | 2.179% | 4.61% | 1.517% | -2.544% | -1.532% | 8.015% | -17.35% | 3.028% | -2.102% | -0.02101% | -19.7% | -1.318% | -3.812% | -7.375% | -1.176% | 1.603% | -5.95% | 2.6% | 6.765% | -0.9621% | 6.875% | 4.025% | 17.71% | -11.69% | -12.29%
| Z-Score (Absolute) |  | 0.66 | 0.55 | 1.05 | 0.42 | 0.42 | 0.22 | 1.76 | NA | 0.73 | 0.33 | 0.1 | 3.97 | 0.17 | 0.69 | 1.42 | 0.14 | 0.43 | 1.13 | 0.64 | 1.5 | 0.1 | 1.52 | NA | 3.76 | 2.31 | 2.44
| Analytical results for Bi in OREAS 753b (Certified Value 0.72 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.7 | 0.71013 | 0.71 | 0.76 | 0.68 | 0.76 | 0.71 | 0.64 | 0.71 | 0.72 | 0.81047 | 0.68 | 0.73 | 0.66 | 0.71 | 0.67 | 0.74 | 0.68 | 0.44 | 0.74 | 0.73 | 0.8 | 0.8 | 0.76 |  |  | 1
| 1 | 2 | 0.77 | 0.74996 | 0.72 | 0.75 | 0.72 | 0.75 | 0.68 | 0.64 | 0.72 | 0.77 | 0.68324 | 0.74 | 0.76 | 0.7 | 0.63 | 0.68 | 0.71 | 0.62 | 0.4 | 0.76 | 0.61 | 0.7 | 0.8 | 0.73 |  |  | 13
| 1 | 3 | 0.75 | 0.74938 | 0.73 | 0.78 | 0.72 | 0.74 | 0.68 | 0.68 | 0.76 | 0.76 | 0.69813 | 0.69 | 0.73 | 0.7 | 0.68 | 0.7 | 0.71 | 0.67 | 0.42 | 0.73 | 0.78 | 0.7 | 0.9 | 0.72 |  |  | 16
| 1 | 4 | 0.77 | 0.76867 | 0.74 | 0.76 | 0.76 | 0.76 | 0.7 | 0.66 | 0.71 | 0.79 | 0.71241 | 0.69 | 0.7 | 0.82 | 0.73 | 0.67 | 0.78 | 0.63 | 0.41 | 0.71 | 0.6 | 0.7 | 0.8 | 0.75 |  |  | 0.72074
| 1 | 5 | 0.73 | 0.73453 | 0.73 | 0.76 | 0.72 | 0.77 | 0.68 | 0.67 | 0.7 | 0.76 | 0.74132 | 0.62 | 0.75 | 0.68 | 0.65 | 0.72 | 0.78 | 0.68 | 0.44 | 0.77 | 0.64 | 0.7 | 0.8 | 0.81 |  |  | 12
| 1 | 6 | 0.78 | 0.72246 | 0.73 | 0.8 | 0.69 | 0.76 | 0.71 | 0.68 | 0.77 | 0.76 | 0.84828 | 0.67 | 0.78 | 1.06 | 0.63 | 0.77 | 0.81 | 0.65 | 0.43 | 0.74 | 0.68 | 0.7 | 0.9 | 0.79
| Mean |  | 0.75 | 0.73919 | 0.72667 | 0.76833 | 0.715 | 0.75667 | 0.69333 | 0.66167 | 0.72833 | 0.76 | 0.74897 | 0.68167 | 0.74167 | 0.77 | 0.67167 | 0.70167 | 0.755 | 0.655 | 0.42333 | 0.74167 | 0.67333 | 0.71667 | 0.83333 | 0.76
| Median |  | 0.76 | 0.74196 | 0.73 | 0.76 | 0.72 | 0.76 | 0.69 | 0.665 | 0.715 | 0.76 | 0.72686 | 0.685 | 0.74 | 0.7 | 0.665 | 0.69 | 0.76 | 0.66 | 0.425 | 0.74 | 0.66 | 0.7 | 0.8 | 0.755
| Std Dev. |  | 0.030332 | 0.021142 | 0.010328 | 0.018348 | 0.028107 | 0.010328 | 0.015055 | 0.018348 | 0.029269 | 0.022804 | 0.06625 | 0.038687 | 0.027869 | 0.15271 | 0.042151 | 0.038687 | 0.041352 | 0.025884 | 0.01633 | 0.02137 | 0.070899 | 0.040825 | 0.05164 | 0.034641
| Rel.Std.Dev. |  | 4.044% | 2.86% | 1.421% | 2.388% | 3.931% | 1.365% | 2.171% | 2.773% | 4.019% | 3% | 8.845% | 5.675% | 3.758% | 19.83% | 6.276% | 5.514% | 5.477% | 3.952% | 3.857% | 2.881% | 10.53% | 5.696% | 6.197% | 4.558%
| PDM3 |  | 4.06% | 2.56% | 0.8226% | 6.604% | -0.7961% | 4.985% | -3.802% | -8.196% | 1.054% | 5.447% | 3.918% | -5.421% | 2.904% | 6.835% | -6.808% | -2.646% | 4.754% | -9.121% | -41.26% | 2.904% | -6.577% | -0.5649% | 15.62% | 5.447%
| Z-Score (Absolute) |  | 0.42 | 0.14 | 0.18 | 0.89 | 0.48 | 0.59 | 1.04 | 1.85 | 0.14 | 0.67 | 0.39 | 1.34 | 0.2 | 0.93 | 1.6 | 0.82 | 0.55 | 2.02 | 7.98 | 0.2 | 1.55 | NA | NA | 0.67
Results from laboratories 23 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 753b (Certified Value 0.228 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.2289 | 0.23011 | 0.24 | 0.225 | 0.234 | 0.229 | 0.231 | 0.23 | 0.25 | 0.22 | 0.153 | 0.23 | 0.22 | 0.23 | 0.23 | 0.24 | 0.21 | 0.22 | 0.22462 | 0.22 | 0.236 | 0.23 | 0.214 | 0.311 | 0.23 |  | 1
| 1 | 2 | 0.2305 | 0.23026 | 0.24 | 0.226 | 0.2332 | 0.215 | 0.233 | 0.23 | 0.24 | 0.22 | 0.162 | 0.23 | 0.22 | 0.21 | 0.23 | 0.24 | 0.19 | 0.22 | 0.22464 | 0.22 | 0.241 | 0.22 | 0.214 | 0.294 | 0.23 |  | 0
| 1 | 3 | 0.2343 | 0.22835 | 0.24 | 0.226 | 0.2336 | 0.213 | 0.229 | 0.24 | 0.23 | 0.21 | 0.178 | 0.23 | 0.23 | 0.22 | 0.23 | 0.24 | 0.2 | 0.22 | 0.2239 | 0.22 | 0.242 | 0.23 | 0.214 | 0.31 | 0.23 |  | 16
| 1 | 4 | 0.2349 | 0.22487 | 0.24 | 0.227 | 0.2322 | 0.23 | 0.232 | 0.23 | 0.24 | 0.22 | 0.164 | 0.23 | 0.19 | 0.22 | 0.23 | 0.24 | 0.21 | 0.22 | 0.22528 | 0.22 | 0.236 | 0.22 | 0.214 | 0.357 | 0.23 |  | 0.22763
| 1 | 5 | 0.235 | 0.23423 | 0.23 | 0.227 | 0.2372 | 0.219 | 0.229 | 0.24 | 0.24 | 0.21 | 0.162 | 0.23 | 0.2 | 0.22 | 0.23 | 0.23 | 0.19 | 0.22 | 0.22614 | 0.22 | 0.195 | 0.21 | 0.214 | 0.305 | 0.23 |  | 13
| 1 | 6 | 0.2359 | 0.22222 | 0.23 | 0.229 | 0.2408 | 0.219 | 0.228 | 0.24 | 0.25 | 0.22 | 0.176 | 0.23 | 0.21 | 0.22 | 0.24 | 0.24 | 0.2 | 0.23 | 0.22663 | 0.23 | 0.208 | 0.23 | 0.214 | 0.283 | 0.22
| Mean |  | 0.23325 | 0.22834 | 0.23667 | 0.22667 | 0.23517 | 0.22083 | 0.23033 | 0.235 | 0.24167 | 0.21667 | 0.16583 | 0.23 | 0.21167 | 0.22 | 0.23167 | 0.23833 | 0.2 | 0.22167 | 0.2252 | 0.22167 | 0.22633 | 0.22333 | 0.214 | 0.31 | 0.22833
| Median |  | 0.2346 | 0.22923 | 0.24 | 0.2265 | 0.2338 | 0.219 | 0.23 | 0.235 | 0.24 | 0.22 | 0.163 | 0.23 | 0.215 | 0.22 | 0.23 | 0.24 | 0.2 | 0.22 | 0.22496 | 0.22 | 0.236 | 0.225 | 0.214 | 0.3075 | 0.23
| Std Dev. |  | 0.0028424 | 0.0042671 | 0.005164 | 0.0013663 | 0.0032358 | 0.007111 | 0.0019664 | 0.0054772 | 0.0075277 | 0.005164 | 0.0094745 | 0 | 0.01472 | 0.0063246 | 0.0040825 | 0.0040825 | 0.0089443 | 0.0040825 | 0.0010265 | 0.0040825 | 0.019826 | 0.008165 | 0 | 0.025377 | 0.0040825
| Rel.Std.Dev. |  | 1.219% | 1.869% | 2.182% | 0.6028% | 1.376% | 3.22% | 0.8537% | 2.331% | 3.115% | 2.383% | 5.713% | 0% | 6.954% | 2.875% | 1.762% | 1.713% | 4.472% | 1.842% | 0.4558% | 1.842% | 8.76% | 3.656% | 0% | 8.186% | 1.788%
| PDM3 |  | 2.47% | 0.3133% | 3.971% | -0.422% | 3.312% | -2.985% | 1.189% | 3.239% | 6.168% | -4.815% | -27.15% | 1.042% | -7.012% | -3.351% | 1.775% | 4.703% | -12.14% | -2.619% | -1.066% | -2.619% | -0.5685% | -1.886% | -5.987% | 36.19% | 0.3102%
| Z-Score (Absolute) |  | 0.67 | 0.17 | 1.02 | 0 | 0.87 | 0.6 | 0.38 | 0.85 | 1.54 | 1.02 | 6.23 | 0.34 | 1.54 | 0.68 | 0.51 | 1.19 | 2.73 | 0.51 | 0.15 | 0.51 | 0.03 | 0.34 | 1.26 | 8.54 | 0.17
| Analytical results for Cd in OREAS 753b (Certified Value 0.14 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | 0.14 | < 0.2 | 0.15 | 0.03 | 0.11 | 0.16 | 0.13 | 0.1 | < 0.3 | 0.15 | 0.12606 | 0.11 | < 0.02 | 0.17 | < 0.02 | 0.09 | < 0.02 | 0.35 | 0.15 | < 1 | < 0.5 | 0.3 | < 0.02 | < 0.2 |  |  | 1
| 1 | 2 | 0.15 | < 0.2 | 0.14 | 0.03 | 0.11 | 0.16 | 0.12 | 0.11 | < 0.3 | 0.15 | 0.12732 | 0.09 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.1 | 0.33 | 0.15 | < 1 | < 0.5 | 0.3 | < 0.02 | < 0.2 |  |  | 6
| 1 | 3 | 0.16 | < 0.2 | 0.14 | 0.11 | 0.14 | 0.15 | 0.13 | 0.17 | < 0.3 | 0.15 | 0.12947 | 0.11 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.13 | 0.36 | 0.17 | < 1 | < 0.5 | 0.3 | < 0.02 | < 0.2 |  |  | 16
| 1 | 4 | 0.16 | < 0.2 | 0.17 | 0.09 | 0.12 | 0.15 | 0.13 | 0.1 | < 0.3 | 0.15 | 0.11536 | 0.11 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.05 | 0.32 | 0.16 | < 1 | < 0.5 | 0.3 | < 0.02 | < 0.2 |  |  | 0.13611
| 1 | 5 | 0.15 | < 0.2 | 0.15 | 0.07 | 0.14 | 0.15 | 0.16 | 0.14 | < 0.3 | 0.13 | 0.12606 | 0.11 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.09 | 0.36 | 0.14 | < 1 | < 0.5 | 0.3 | < 0.02 | < 0.2 |  |  | 14
| 1 | 6 | 0.11 | < 0.2 | 0.14 | 0.06 | 0.15 | 0.16 | 0.16 | 0.13 | < 0.3 | 0.14 | 0.1282 | 0.1 | < 0.02 | < 0.02 | < 0.02 | < 0.02 | 0.2 | 0.33 | 0.16 | < 1 | < 0.5 | 0.3 | < 0.02 | < 0.2
| Mean |  | 0.145 |  | 0.14833 | 0.065 | 0.12833 | 0.155 | 0.13833 | 0.125 |  | 0.145 | 0.12541 | 0.105 |  | 0.17 |  | 0.09 | 0.114 | 0.34167 | 0.155 |  |  | 0.3
| Median |  | 0.15 |  | 0.145 | 0.065 | 0.13 | 0.155 | 0.13 | 0.12 |  | 0.15 | 0.12669 | 0.11 |  | 0.17 |  | 0.09 | 0.1 | 0.34 | 0.155 |  |  | 0.3
| Std Dev. |  | 1.01: 0.018708; 1.03: 0.01169; 1.04: 0.032094; 1.05: 0.017224; 1.06: 0.0054772; 1.07: 0.017224; 1.08: 0.027386; 1.10: 0.0083666; 1.11: 0.0050944; 1.12: 0.0083666; 1.17: 0.055946; 1.18: 0.017224; 1.19: 0.010488; 1.24: 0
| Rel.Std.Dev. |  | 1.01: 12.9%; 1.03: 7.881%; 1.04: 49.37%; 1.05: 13.42%; 1.06: 3.534%; 1.07: 12.45%; 1.08: 21.91%; 1.10: 5.77%; 1.11: 4.062%; 1.12: 7.968%; 1.17: 49.08%; 1.18: 5.041%; 1.19: 6.767%; 1.24: 0%
| PDM3 |  | 6.532% |  | 8.981% | -52.24% | -5.713% | 13.88% | 1.634% | -8.162% |  | 6.532% | -7.86% | -22.86% |  | 24.9% |  | -33.88% | -16.24% | 151% | 13.88% |  |  | 120.4%
| Z-Score (Absolute) |  | 0.46 | 0.6 | 0.54 | 1.42 | 0.07 | 0.69 | 0.3 | 0.01 | 0.58 | 0.46 | 0 | 0.48 | 2.71 | 2.08 | 2.71 | 2.39 | 0.67 | 5.07 | 0.69 | 8.79 | 2.92 | NA | 2.71 | 0.6
Results from laboratory 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 753b (Certified Value 7.92 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 9.74 | 6.8998 | 7.98 | 9.77 | 9.39 | 6.72 | 10.2 | 9.35 | 9.1738 | 8.2 | 7.72 | 8.15 | 7.37 | 5.91 | 7.39 | 4.51 | 9.6324 | 7.6 | 5.39 | 9 | 9 | 8.86 | 8.63 |  |  |  | 1
| 1 | 2 | 8.75 | 7.2442 | 8.28 | 9.31 | 9.06 | 5.93 | 8.6 | 10.8 | 8.2223 | 8.14 | 6.88 | 8.08 | 7.19 | 6.14 | 8.45 | 4.64 | 9.6851 | 7.7 | 4.5 | 8 | 9 | 8.88 | 8.22 |  |  |  | 14
| 1 | 3 | 8.78 | 7.2664 | 7.7 | 9.35 | 9.06 | 8.22 | 8.5 | 9.14 | 8.3496 | 7.88 | 7.42 | 7.72 | 6.94 | 6.43 | 7.35 | 4.57 | 9.6376 | 7.6 | 4.86 | 9 | 8 | 9.46 | 8.08 |  |  |  | 16
| 1 | 4 | 8.94 | 7.1223 | 8.05 | 9.68 | 9.48 | 6.79 | 8.5 | 10.2 | 8.3463 | 8.89 | 8.03 | 6.15 | 8.04 | 6.26 | 8 | 4.71 | 9.6126 | 7.7 | 5.38 | 9 | 8 | 9.22 | 7.19 |  |  |  | 7.9229
| 1 | 5 | 9.36 | 7.0048 | 8.52 | 10.29 | 8.94 | 7.03 | 9.7 | 10 | 8.8103 | 7.8 | 7.05 | 6.47 | 7.08 | 6.73 | 7.38 | 4.52 | 9.6985 | 8.1 | 7.01 | 7 | 9 | 9.02 | 8.15 |  |  |  | 15
| 1 | 6 | 9.41 | 7.4463 | 8.13 | 9.26 | 9.4 | 6.82 | 9.3 | 11.7 | 8.9279 | 8.59 | 7.61 | 6.39 | 6.97 | 6.57 | 7.52 | 4.75 | 9.6324 | 7.9 | 5.79 | 7 | 8 | 9.19 | 7.84
| Mean |  | 9.1633 | 7.164 | 8.11 | 9.61 | 9.2217 | 6.9183 | 9.1333 | 10.1983 | 8.6384 | 8.25 | 7.4517 | 7.16 | 7.265 | 6.34 | 7.6817 | 4.6167 | 9.6498 | 7.7667 | 5.4883 | 8.1667 | 8.5 | 9.105 | 8.0183
| Median |  | 9.15 | 7.1832 | 8.09 | 9.515 | 9.225 | 6.805 | 8.95 | 10.1 | 8.5799 | 8.17 | 7.515 | 7.095 | 7.135 | 6.345 | 7.455 | 4.605 | 9.635 | 7.7 | 5.385 | 8.5 | 8.5 | 9.105 | 8.115
| Std Dev. |  | 0.39993 | 0.19669 | 0.27785 | 0.39319 | 0.22737 | 0.7415 | 0.7174 | 0.94827 | 0.38521 | 0.41933 | 0.42911 | 0.9197 | 0.41089 | 0.29813 | 0.44808 | 0.099933 | 0.033929 | 0.19664 | 0.87172 | 0.98319 | 0.54772 | 0.22993 | 0.48039
| Rel.Std.Dev. |  | 4.364% | 2.746% | 3.426% | 4.091% | 2.466% | 10.72% | 7.855% | 9.298% | 4.459% | 5.083% | 5.759% | 12.84% | 5.656% | 4.702% | 5.833% | 2.165% | 0.3516% | 2.532% | 15.88% | 12.04% | 6.444% | 2.525% | 5.991%
| PDM3 |  | 15.66% | -9.579% | 2.361% | 21.29% | 16.39% | -12.68% | 15.28% | 28.72% | 9.03% | 4.128% | -5.948% | -9.629% | -8.304% | -19.98% | -3.045% | -41.73% | 21.8% | -1.972% | -30.73% | 3.076% | 7.284% | 14.92% | 1.204%
| Z-Score (Absolute) |  | 0.71 | 0.53 | 0.06 | 0.99 | 0.75 | 0.68 | 0.69 | 1.35 | 0.38 | 0.14 | 0.35 | 0.53 | 0.47 | 1.04 | 0.21 | 2.11 | 1.01 | 0.16 | 1.57 | NA | NA | 0.67 | 0
Results from laboratories 23 and 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 753b (Certified Value 1.38 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 1.4 | 1.2715 | 1.4 | 1.4 | 1.4 | 1 | 1.3 | 1.3 | 1 | 1.48 | 1.3948 | 1.4 | 1.4 | 1.3 | 1.4 | 1.4 | 1.4 | 1.4 | 1.3 | 1.9766 | 1.3 | 1 | 1.48 | 2.3 | 1.39 |  | 1
| 1 | 2 | 1.5 | 1.3225 | 1.4 | 1.4 | 1.4 | 1 | 1.3 | 1.3 | 1 | 1.46 | 1.3747 | 1.4 | 1.5 | 1.3 | 1.3 | 1.4 | 1.5 | 1.3 | 1.2 | 1.9681 | 1.3 | 1 | 1.48 | 1.7 | 1.3 |  | 15
| 1 | 3 | 1.4 | 1.361 | 1.4 | 1.4 | 1.5 | 1 | 1.3 | 1.4 | 1 | 1.44 | 1.3971 | 1.4 | 1.4 | 1.4 | 1.3 | 1.4 | 1.4 | 1.4 | 1.3 | 1.9533 | 1.3 | 1 | 1.53 | 1.6 | 1.34 |  | 16
| 1 | 4 | 1.4 | 1.3016 | 1.4 | 1.4 | 1.4 | 1 | 1.3 | 1.4 | 1 | 1.46 | 1.4205 | 1.4 | 1.4 | 1.4 | 1.3 | 1.5 | 1.4 | 1.3 | 1.4 | 1.9649 | 1.3 | 1 | 1.45 | 1.8 | 1.46 |  | 1.3821
| 1 | 5 | 1.4 | 1.3384 | 1.4 | 1.3 | 1.4 | 1 | 1.3 | 1.4 | 1 | 1.44 | 1.3858 | 1.4 | 1.4 | 1.3 | 1.3 | 1.4 | 1.5 | 1.4 | 1.2 | 1.9605 | 1.3 | 1 | 1.53 | 1.7 | 1.46 |  | 16
| 1 | 6 | 1.5 | 1.2843 | 1.5 | 1.4 | 1.4 | 1 | 1.3 | 1.3 | 1 | 1.51 | 1.4389 | 1.4 | 1.4 | 1.3 | 1.3 | 1.6 | 1.4 | 1.3 | 1.3 | 1.949 | 1.4 | 1 | 1.46 | 1.7 | 1.29
| Mean |  | 1.4333 | 1.3132 | 1.4167 | 1.3833 | 1.4167 | 1 | 1.3 | 1.35 | 1 | 1.465 | 1.4019 | 1.4 | 1.4167 | 1.3333 | 1.3167 | 1.45 | 1.4333 | 1.35 | 1.2833 | 1.9621 | 1.3167 | 1 | 1.4883 | 1.8 | 1.3733
| Median |  | 1.4 | 1.3121 | 1.4 | 1.4 | 1.4 | 1 | 1.3 | 1.35 | 1 | 1.46 | 1.3959 | 1.4 | 1.4 | 1.3 | 1.3 | 1.4 | 1.4 | 1.35 | 1.3 | 1.9627 | 1.3 | 1 | 1.48 | 1.7 | 1.365
| Std Dev. |  | 0.05164 | 0.033819 | 0.040825 | 0.040825 | 0.040825 | 0 | 0 | 0.054772 | 0 | 0.026646 | 0.023595 | 2.4324e-16 | 0.040825 | 0.05164 | 0.040825 | 0.083666 | 0.05164 | 0.054772 | 0.075277 | 0.01006 | 0.040825 | 0 | 0.034303 | 0.25298 | 0.075807
| Rel.Std.Dev. |  | 3.603% | 2.575% | 2.882% | 2.951% | 2.882% | 0% | 0% | 4.057% | 0% | 1.819% | 1.683% | 1.737e-14% | 2.882% | 3.873% | 3.101% | 5.77% | 3.603% | 4.057% | 5.866% | 0.5127% | 3.101% | 0% | 2.305% | 14.05% | 5.52%
| PDM3 |  | 3.708% | -4.984% | 2.502% | 0.08988% | 2.502% | -27.65% | -5.94% | -2.322% | -27.65% | 5.999% | 1.437% | 1.296% | 2.502% | -3.528% | -4.734% | 4.913% | 3.708% | -2.322% | -7.146% | 41.96% | -4.734% | -27.65% | 7.687% | 30.24% | -0.6337%
| Z-Score (Absolute) |  | 0.43 | 1.16 | 0.21 | 0.23 | 0.21 | NA | 1.34 | 0.67 | NA | 0.85 | 0.01 | 0.01 | 0.21 | 0.89 | 1.12 | 0.65 | 0.43 | 0.67 | 1.56 | 7.42 | 1.12 | NA | 1.16 | 5.28 | 0.37
Results from laboratories 6, 9 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 753b (Certified Value 17 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 21 | 17.4634 | 16 | 9 | 18 | 18 | 12 | 12 | 18 | 19.1 | 23.4284 | 17 | 17 | 18 | 21 | 20 | 20 | 20 | 12 | 12 | 11 | 12 | 18 | 35.22 |  |  | 1
| 1 | 2 | 21 | 17.7249 | 16 | 6 | 17 | 18 | 13 | 12 | 17 | 18.1 | 23.4803 | 21 | 16 | 19 | 20 | 20 | 19 | 21 | 14 | 11 | 11 | 11 | 8 | 36.03 |  |  | 7
| 1 | 3 | 17 | 16.9703 | 16 | 9 | 16 | 20 | 14 | 13 | 23 | 18.3 | 23.3814 | 19 | 16 | 19 | 19 | 19 | 20 | 20 | 12 | 10 | 11 | 17 | 9 | 34.7 |  |  | 16
| 1 | 4 | 18 | 16.3694 | 15 | 7 | 15 | 21 | 11 | 13 | 24 | 17.6 | 24.4427 | 20 | 16 | 19 | 23 | 19 | 19 | 21 | 14 | < 10 | 10 | 11 | 6 | 33.54 |  |  | 16.9538
| 1 | 5 | 19 | 16.286 | 15 | 6 | 16 | 18 | 11 | 13 | 16 | 18.2 | 23.6421 | 18 | 16 | 18 | 22 | 20 | 18 | 20 | 13 | < 10 | 9 | 12 | 6 | 32.93 |  |  | 17
| 1 | 6 | 22 | 16.4525 | 16 | 8 | 18 | 19 | 15 | 12 | 19 | 19.2 | 23.2478 | 20 | 17 | 19 | 24 | 22 | 19 | 21 | 14 | < 10 | 10 | 13 | 4 | 36.08
| Mean |  | 19.6667 | 16.8778 | 15.6667 | 7.5 | 16.6667 | 19 | 12.6667 | 12.5 | 19.5 | 18.4167 | 23.6038 | 19.1667 | 16.3333 | 18.6667 | 21.5 | 20 | 19.1667 | 20.5 | 13.1667 | 11 | 10.3333 | 12.6667 | 8.5 | 34.75
| Median |  | 20 | 16.7114 | 16 | 7.5 | 16.5 | 18.5 | 12.5 | 12.5 | 18.5 | 18.25 | 23.4544 | 19.5 | 16 | 19 | 21.5 | 20 | 19 | 20.5 | 13.5 | 11 | 10.5 | 12 | 7 | 34.96
| Std Dev. |  | 1.9664 | 0.6097 | 0.5164 | 1.3784 | 1.2111 | 1.2649 | 1.633 | 0.54772 | 3.2711 | 0.61779 | 0.43065 | 1.472 | 0.5164 | 0.5164 | 1.8708 | 1.0954 | 0.75277 | 0.54772 | 0.98319 | 1 | 0.8165 | 2.2509 | 4.9699 | 1.2967
| Rel.Std.Dev. |  | 9.999% | 3.612% | 3.296% | 18.38% | 7.266% | 6.657% | 12.89% | 4.382% | 16.77% | 3.355% | 1.825% | 7.68% | 3.162% | 2.766% | 8.702% | 5.477% | 3.928% | 2.672% | 7.467% | 9.091% | 7.902% | 17.77% | 58.47% | 3.732%
| PDM3 |  | 16% | -0.4487% | -7.592% | -55.76% | -1.694% | 12.07% | -25.29% | -26.27% | 15.02% | 8.628% | 39.22% | 13.05% | -3.66% | 10.1% | 26.82% | 17.97% | 13.05% | 20.92% | -22.34% | -35.12% | -39.05% | -25.29% | -49.86% | 105%
| Z-Score (Absolute) |  | 0.52 | 0.2 | 0.51 | 2.63 | 0.25 | 0.35 | 1.29 | 1.33 | 0.48 | 0.2 | 1.54 | 0.39 | 0.34 | 0.26 | 1 | 0.61 | 0.39 | 0.74 | 1.16 | 2.5 | 1.89 | 1.29 | 2.37 | 4.43
| Analytical results for Cs in OREAS 753b (Certified Value 170 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 174.15 | 175.48 | 165 | 162.25 | 173 | 170.8 | 150.6 | > 100 | 185 | 166.2076 | 187 | 175 | 170.5 | 162.5 | 171 | 163 | 82.7 | 171.023 | 172.6 | 163 | 174 | 188 | 162 |  |  |  | 1
| 1 | 2 | 170.52 | 164.25 | 173 | 161.68 | 173 | 172.1 | 145.9 | > 100 | 176 | 167.3676 | 181 | 174.5 | 176.5 | 157 | 166.5 | 167 | 85.5 | 171.088 | 170.3 | 193 | 173 | 193 | 160 |  |  |  | 16
| 1 | 3 | 172.74 | 166.66 | 164 | 163.19 | 172 | 172.2 | 164.8 | > 100 | 176 | 167.8804 | 186 | 174.5 | 161 | 157.5 | 170.5 | 164 | 83.4 | 170.288 | 164.3 | 193 | 173 | 193 | 159 |  |  |  | 16
| 1 | 4 | 173.74 | 171.65 | 164 | 162.9 | 174 | 171 | 158.4 | > 100 | 181 | 167.7099 | 186 | 178.5 | 149.5 | 160.5 | 172.5 | 159.5 | 85.7 | 170.308 | 176.3 | 194 | 171 | 188 | 160 |  |  |  | 169.5518
| 1 | 5 | 171.6 | 163.5 | 168 | 163.7 | 171 | 170.9 | 165.1 | > 100 | 181 | 168.9082 | 180 | 175 | 155 | 162 | 170.5 | 155 | 83.9 | 171.074 | 172.7 | 205 | 151 | 188 | 162 |  |  |  | 18
| 1 | 6 | 175.35 | 163.98 | 173 | 162.74 | 173 | 170.2 | 149.6 | > 100 | 174 | 168.7511 | 187 | 174 | 152 | 163 | 178 | 159.5 | 86.1 | 171.035 | 171.6 | 201 | 154 | 192 | 165
| Mean |  | 173.0167 | 167.5867 | 167.8333 | 162.7433 | 172.6667 | 171.2 | 155.7333 |  | 178.8333 | 167.8041 | 184.5 | 175.25 | 160.75 | 160.4167 | 171.5 | 161.3333 | 84.55 | 170.8027 | 171.3 | 191.5 | 166 | 190.3333 | 161.3333
| Median |  | 173.24 | 165.455 | 166.5 | 162.82 | 173 | 170.95 | 154.5 |  | 178.5 | 167.7952 | 186 | 174.75 | 158 | 161.25 | 170.75 | 161.25 | 84.7 | 171.029 | 172.1 | 193.5 | 172 | 190 | 161
| Std Dev. |  | 1.7646 | 4.9095 | 4.2622 | 0.70899 | 1.0328 | 0.7874 | 8.2184 |  | 4.1673 | 0.9869 | 3.1464 | 1.6355 | 10.7645 | 2.5965 | 3.755 | 4.2151 | 1.3996 | 0.3917 | 3.9684 | 14.802 | 10.5451 | 2.582 | 2.1602
| Rel.Std.Dev. |  | 1.02% | 2.93% | 2.54% | 0.4356% | 0.5981% | 0.4599% | 5.277% |  | 2.33% | 0.5881% | 1.705% | 0.9333% | 6.696% | 1.619% | 2.19% | 2.613% | 1.655% | 0.2293% | 2.317% | 7.73% | 6.352% | 1.357% | 1.339%
| PDM3 |  | 2.044% | -1.159% | -1.014% | -4.016% | 1.837% | 0.9721% | -8.15% |  | 5.474% | -1.031% | 8.816% | 3.361% | -5.191% | -5.388% | 1.149% | -4.847% | -50.13% | 0.7377% | 1.031% | 12.94% | -2.095% | 12.26% | -4.847%
| Z-Score (Absolute) |  | 0.4 | 0.19 | 0.16 | 0.71 | 0.36 | 0.2 | 1.46 | NA | 1.03 | 0.16 | 1.64 | 0.64 | 0.92 | 0.96 | 0.24 | 0.86 | 9.14 | 0.16 | 0.21 | 2.39 | 0.36 | 2.27 | 0.86
| Analytical results for Cu in OREAS 753b (Certified Value 18 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 18.2 | 17.2062 | 16.7 | 17.3 | 17.9 | 16 | 16 | 16.9 | 18 | 18.4 | 17.5837 | 18.2 | 18.3 | 17 | 18 | 17.6 | 19.1 | 14.8 | 19 | 17.4 | 17.3 | 18.1 | 19.1 | 20.5 | 19.28 |  | 1
| 1 | 2 | 17.4 | 17.4529 | 16.1 | 17.4 | 17.6 | 17 | 17 | 16.6 | 18 | 18.6 | 17.128 | 18.5 | 18.4 | 17.9 | 17 | 17.5 | 19.2 | 15.7 | 18.2 | 17.5 | 19.3 | 18 | 19.9 | 22 | 19.18 |  | 0
| 1 | 3 | 17.7 | 17.5029 | 16.7 | 17.6 | 17.4 | 16 | 16 | 17.5 | 18 | 18.3 | 17.4555 | 18.2 | 17.6 | 17.4 | 17 | 18 | 18.4 | 14.8 | 18.8 | 16.7 | 19.9 | 18.4 | 19.1 | 21 | 18.49 |  | 16
| 1 | 4 | 18 | 17.3087 | 16.6 | 17.4 | 17.9 | 17 | 16 | 17.7 | 18 | 18.1 | 17.6752 | 18.7 | 18.7 | 17.4 | 17.6 | 17.9 | 18 | 15.4 | 19.3 | 17.8 | 19.6 | 19.6 | 19.3 | 23.1 | 19.04 |  | 18.0328
| 1 | 5 | 17.7 | 17.6901 | 17 | 18.3 | 17.4 | 16 | 16 | 18.1 | 18 | 18.1 | 17.9864 | 18.9 | 18.6 | 16.9 | 17 | 17.8 | 20.4 | 14.8 | 17.8 | 17.3 | 19.1 | 17.7 | 23.3 | 20.7 | 19.55 |  | 19
| 1 | 6 | 17.7 | 17.2522 | 16.6 | 17.1 | 17.9 | 17 | 17 | 16.8 | 19 | 18.9 | 17.9483 | 18.1 | 18.2 | 17 | 17.2 | 18.6 | 19.1 | 15.7 | 20 | 17.6 | 19.2 | 21.3 | 18.9 | 20.7 | 18.5
| Mean |  | 17.7833 | 17.4022 | 16.6167 | 17.5167 | 17.6833 | 16.5 | 16.3333 | 17.2667 | 18.1667 | 18.4 | 17.6295 | 18.4333 | 18.3 | 17.2667 | 17.3 | 17.9 | 19.0333 | 15.2 | 18.85 | 17.3833 | 19.0667 | 18.85 | 19.9333 | 21.3333 | 19.0067
| Median |  | 17.7 | 17.3808 | 16.65 | 17.4 | 17.75 | 16.5 | 16 | 17.2 | 18 | 18.35 | 17.6294 | 18.35 | 18.35 | 17.2 | 17.1 | 17.85 | 19.1 | 15.1 | 18.9 | 17.45 | 19.25 | 18.25 | 19.2 | 20.85 | 19.11
| Std Dev. |  | 0.27869 | 0.1818 | 0.29269 | 0.41673 | 0.24833 | 0.54772 | 0.5164 | 0.58878 | 0.40825 | 0.30984 | 0.32092 | 0.32042 | 0.38987 | 0.37771 | 0.41473 | 0.38987 | 0.82138 | 0.45166 | 0.78422 | 0.37639 | 0.9136 | 1.3693 | 1.6848 | 1.0172 | 0.43005
| Rel.Std.Dev. |  | 1.567% | 1.045% | 1.761% | 2.379% | 1.404% | 3.32% | 3.162% | 3.41% | 2.247% | 1.684% | 1.82% | 1.738% | 2.13% | 2.188% | 2.397% | 2.178% | 4.315% | 2.971% | 4.16% | 2.165% | 4.792% | 7.264% | 8.452% | 4.768% | 2.263%
| PDM3 |  | -1.383% | -3.497% | -7.853% | -2.862% | -1.938% | -8.5% | -9.424% | -4.248% | 0.7426% | 2.037% | -2.236% | 2.221% | 1.482% | -4.248% | -4.063% | -0.7362% | 5.549% | -15.71% | 4.532% | -3.601% | 5.734% | 4.532% | 10.54% | 18.3% | 5.401%
| Z-Score (Absolute) |  | 0.07 | 0.52 | 1.44 | 0.38 | 0.19 | NA | NA | 0.67 | NA | 0.65 | 0.25 | 0.69 | 0.54 | 0.67 | 0.64 | 0.07 | 1.4 | 3.1 | 1.18 | 0.54 | 1.44 | 1.18 | 2.45 | 4.09 | 1.37
Results from laboratories 6, 7 and 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Dy in OREAS 753b (Certified Value 0.46 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.20; 1.08: 1.22; 1.09: 1.24; 1.10: 1.25 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.49; 1.02: 0.49; 1.03: 0.5; 1.04: 0.5; 1.05: 0.63238; 1.06: 0.45; 1.07: 0.45036; 1.08: 0.38; 1.09: 0.66; 1.10: 0.57 |  | 1
| 1 | 2 | 1.01: 0.5; 1.02: 0.49; 1.03: 0.4; 1.04: 0.5; 1.05: 0.6783; 1.06: 0.45; 1.07: 0.47826; 1.08: 0.33; 1.09: 0.63; 1.10: 0.5 |  | 18
| 1 | 3 | 1.01: 0.48; 1.02: 0.46; 1.03: 0.5; 1.04: 0.4; 1.05: 0.61734; 1.06: 0.42; 1.07: 0.51057; 1.08: 0.34; 1.09: 0.61; 1.10: 0.51 |  | 16
| 1 | 4 | 1.01: 0.49; 1.02: 0.48; 1.03: 0.4; 1.04: 0.5; 1.05: 0.61889; 1.06: 0.45; 1.07: 0.49734; 1.08: 0.38; 1.09: 0.6; 1.10: 0.5 |  | 0.45965
| 1 | 5 | 1.01: 0.5; 1.02: 0.47; 1.03: 0.4; 1.04: 0.5; 1.05: 0.65484; 1.06: 0.44; 1.07: 0.41373; 1.08: 0.42; 1.09: 0.6; 1.10: 0.5 |  | 20
| 1 | 6 | 1.01: 0.49; 1.02: 0.48; 1.03: 0.4; 1.04: 0.5; 1.05: 0.66577; 1.06: 0.47; 1.07: 0.44307; 1.08: 0.38; 1.09: 0.62; 1.10: 0.51
| Mean |  | 1.01: 0.49167; 1.02: 0.47833; 1.03: 0.43333; 1.04: 0.48333; 1.05: 0.64459; 1.06: 0.44667; 1.07: 0.46555; 1.08: 0.37167; 1.09: 0.62; 1.10: 0.515
| Median |  | 1.01: 0.49; 1.02: 0.48; 1.03: 0.4; 1.04: 0.5; 1.05: 0.64361; 1.06: 0.45; 1.07: 0.46431; 1.08: 0.38; 1.09: 0.615; 1.10: 0.505
| Std Dev. |  | 1.01: 0.0075277; 1.02: 0.01169; 1.03: 0.05164; 1.04: 0.040825; 1.05: 0.025464; 1.06: 0.01633; 1.07: 0.036388; 1.08: 0.032506; 1.09: 0.022804; 1.10: 0.027386
| Rel.Std.Dev. |  | 1.01: 1.531%; 1.02: 2.444%; 1.03: 11.92%; 1.04: 8.447%; 1.05: 3.95%; 1.06: 3.656%; 1.07: 7.816%; 1.08: 8.746%; 1.09: 3.678%; 1.10: 5.318%
| PDM3 |  | 1.01: 6.966%; 1.02: 4.065%; 1.03: -5.725%; 1.04: 5.153%; 1.05: 40.23%; 1.06: -2.824%; 1.07: 1.285%; 1.08: -19.14%; 1.09: 34.89%; 1.10: 12.04%
| Z-Score (Absolute) |  | 1.01: 0.13; 1.02: 0.13; 1.03: NA; 1.04: NA; 1.05: 3.15; 1.06: 0.76; 1.07: 0.38; 1.08: 2.24; 1.09: 2.66; 1.10: 0.59
Results from laboratories 8 and 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 753b (Certified Value 0.21 ppm) | [27 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.20; 1.08: 1.22; 1.09: 1.24; 1.10: 1.25 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.19; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2; 1.05: 0.3364; 1.06: 0.17; 1.07: 0.20826; 1.08: 0.15; 1.09: 0.27; 1.10: 0.24 |  | 1
| 1 | 2 | 1.01: 0.19; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2; 1.05: 0.37508; 1.06: 0.17; 1.07: 0.21591; 1.08: 0.15; 1.09: 0.26; 1.10: 0.22 |  | 19
| 1 | 3 | 1.01: 0.18; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2; 1.05: 0.32147; 1.06: 0.17; 1.07: 0.26199; 1.08: 0.17; 1.09: 0.26; 1.10: 0.21 |  | 16
| 1 | 4 | 1.01: 0.19; 1.02: 0.22; 1.03: 0.2; 1.04: 0.2; 1.05: 0.3353; 1.06: 0.16; 1.07: 0.2809; 1.08: 0.16; 1.09: 0.26; 1.10: 0.2 |  | 0.20707
| 1 | 5 | 1.01: 0.2; 1.02: 0.2; 1.03: 0.2; 1.04: 0.2; 1.05: 0.32792; 1.06: 0.18; 1.07: 0.20277; 1.08: 0.19; 1.09: 0.25; 1.10: 0.22 |  | 21
| 1 | 6 | 1.01: 0.21; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2; 1.05: 0.33611; 1.06: 0.17; 1.07: 0.19728; 1.08: 0.19; 1.09: 0.27; 1.10: 0.22
| Mean |  | 1.01: 0.19333; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2; 1.05: 0.33871; 1.06: 0.17; 1.07: 0.22785; 1.08: 0.16833; 1.09: 0.26167; 1.10: 0.21833
| Median |  | 1.01: 0.19; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2; 1.05: 0.3357; 1.06: 0.17; 1.07: 0.21208; 1.08: 0.165; 1.09: 0.26; 1.10: 0.22
| Std Dev. |  | 1.01: 0.010328; 1.02: 0.0063246; 1.03: 3.0405e-17; 1.04: 3.0405e-17; 1.05: 0.018762; 1.06: 0.0063246; 1.07: 0.034842; 1.08: 0.018348; 1.09: 0.0075277; 1.10: 0.013292
| Rel.Std.Dev. |  | 1.01: 5.342%; 1.02: 3.012%; 1.03: 1.52e-14%; 1.04: 1.52e-14%; 1.05: 5.539%; 1.06: 3.72%; 1.07: 15.29%; 1.08: 10.9%; 1.09: 2.877%; 1.10: 6.088%
| PDM3 |  | 1.01: -6.636%; 1.02: 1.413%; 1.03: -3.416%; 1.04: -3.416%; 1.05: 63.57%; 1.06: -17.9%; 1.07: 10.03%; 1.08: -18.71%; 1.09: 26.36%; 1.10: 5.437%
| Z-Score (Absolute) |  | 1.01: 0.43; 1.02: 0.09; 1.03: NA; 1.04: NA; 1.05: 2.58; 1.06: 0.92; 1.07: 0.28; 1.08: 0.95; 1.09: 0.99; 1.10: 0.09
Results from laboratories 8 and 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 753b (Certified Value 0.15 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.20; 1.08: 1.22; 1.09: 1.24; 1.10: 1.25 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 0.17; 1.02: 0.17; 1.03: 0.1; 1.04: 0.17; 1.05: 0.15915; 1.06: 0.14; 1.07: 0.2204; 1.08: 0.12; 1.09: 0.21; 1.10: 0.14 |  | 1
| 1 | 2 | 1.01: 0.17; 1.02: 0.18; 1.03: 0.3; 1.04: 0.15; 1.05: 0.15633; 1.06: 0.18; 1.07: 0.25163; 1.08: 0.13; 1.09: 0.22; 1.10: 0.14 |  | 20
| 1 | 3 | 1.01: 0.15; 1.02: 0.18; 1.03: 0.1; 1.04: 0.15; 1.05: 0.15154; 1.06: 0.15; 1.07: 0.23921; 1.08: 0.14; 1.09: 0.21; 1.10: 0.14 |  | 16
| 1 | 4 | 1.01: 0.14; 1.02: 0.17; 1.03: 0.1; 1.04: 0.15; 1.05: 0.1474; 1.06: 0.17; 1.07: 0.24524; 1.08: 0.13; 1.09: 0.21; 1.10: 0.14 |  | 0.15253
| 1 | 5 | 1.01: 0.16; 1.02: 0.15; 1.03: 0.1; 1.04: 0.15; 1.05: 0.1501; 1.06: 0.16; 1.07: 0.20933; 1.08: 0.15; 1.09: 0.21; 1.10: 0.13 |  | 22
| 1 | 6 | 1.01: 0.16; 1.02: 0.15; 1.03: 0.1; 1.04: 0.16; 1.05: 0.15185; 1.06: 0.17; 1.07: 0.21905; 1.08: 0.14; 1.09: 0.21; 1.10: 0.14
| Mean |  | 1.01: 0.15833; 1.02: 0.16667; 1.03: 0.13333; 1.04: 0.155; 1.05: 0.15273; 1.06: 0.16167; 1.07: 0.23081; 1.08: 0.135; 1.09: 0.21167; 1.10: 0.13833
| Median |  | 1.01: 0.16; 1.02: 0.17; 1.03: 0.1; 1.04: 0.15; 1.05: 0.15169; 1.06: 0.165; 1.07: 0.22981; 1.08: 0.135; 1.09: 0.21; 1.10: 0.14
| Std Dev. |  | 1.01: 0.01169; 1.02: 0.013663; 1.03: 0.08165; 1.04: 0.0083666; 1.05: 0.0042824; 1.06: 0.01472; 1.07: 0.016854; 1.08: 0.010488; 1.09: 0.0040825; 1.10: 0.0040825
| Rel.Std.Dev. |  | 1.01: 7.383%; 1.02: 8.198%; 1.03: 61.24%; 1.04: 5.398%; 1.05: 2.804%; 1.06: 9.105%; 1.07: 7.302%; 1.08: 7.769%; 1.09: 1.929%; 1.10: 2.951%
| PDM3 |  | 1.01: 3.803%; 1.02: 9.266%; 1.03: -12.59%; 1.04: 1.618%; 1.05: 0.1283%; 1.06: 5.988%; 1.07: 51.32%; 1.08: -11.49%; 1.09: 38.77%; 1.10: -9.309%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: NA; 1.04: 0.27; 1.05: 0.45; 1.06: 0.27; 1.07: 5.86; 1.08: 1.89; 1.09: 4.32; 1.10: 1.62
| Analytical results for Fe in OREAS 753b (Certified Value 0.615 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.61 | 0.59332 | 0.63 | 0.63 | 0.6081 | 0.64 | 0.6 | 0.64 | 0.61 | 0.6 | 0.68009 | 0.6 | 0.6 | 0.58 | 0.59 | 0.63 | 0.56 | 0.6 | 0.93287 | 0.62 | 0.64 | 0.64 | 0.609 | 0.68 | 0.61 |  | 1
| 1 | 2 | 0.61 | 0.60149 | 0.63 | 0.63 | 0.6034 | 0.63 | 0.61 | 0.61 | 0.61 | 0.59 | 0.67568 | 0.59 | 0.6 | 0.57 | 0.59 | 0.64 | 0.58 | 0.61 | 0.94093 | 0.61 | 0.65 | 0.63 | 0.63 | 0.63 | 0.67 |  | 0
| 1 | 3 | 0.62 | 0.59465 | 0.63 | 0.63 | 0.6091 | 0.62 | 0.6 | 0.66 | 0.6 | 0.59 | 0.65502 | 0.59 | 0.61 | 0.57 | 0.59 | 0.63 | 0.55 | 0.61 | 0.94089 | 0.6 | 0.64 | 0.63 | 0.609 | 0.63 | 0.75 |  | 16
| 1 | 4 | 0.62 | 0.58209 | 0.62 | 0.63 | 0.6093 | 0.63 | 0.61 | 0.65 | 0.61 | 0.58 | 0.67776 | 0.6 | 0.56 | 0.59 | 0.59 | 0.63 | 0.56 | 0.61 | 0.95455 | 0.62 | 0.63 | 0.62 | 0.616 | 0.64 | 0.66 |  | 0.61476
| 1 | 5 | 0.62 | 0.5977 | 0.62 | 0.63 | 0.6099 | 0.62 | 0.61 | 0.66 | 0.63 | 0.57 | 0.6801 | 0.59 | 0.57 | 0.58 | 0.57 | 0.59 | 0.58 | 0.61 | 0.93361 | 0.61 | 0.62 | 0.61 | 0.616 | 0.61 | 0.72 |  | 23
| 1 | 6 | 0.64 | 0.57866 | 0.62 | 0.64 | 0.609 | 0.64 | 0.6 | 0.67 | 0.63 | 0.6 | 0.67225 | 0.6 | 0.58 | 0.59 | 0.61 | 0.61 | 0.57 | 0.62 | 0.95389 | 0.63 | 0.62 | 0.64 | 0.616 | 0.61 | 0.59
| Mean |  | 0.62 | 0.59132 | 0.625 | 0.63167 | 0.60813 | 0.63 | 0.605 | 0.64833 | 0.615 | 0.58833 | 0.67348 | 0.595 | 0.58667 | 0.58 | 0.59 | 0.62167 | 0.56667 | 0.61 | 0.94279 | 0.615 | 0.63333 | 0.62833 | 0.616 | 0.63333 | 0.66667
| Median |  | 0.62 | 0.59399 | 0.625 | 0.63 | 0.60905 | 0.63 | 0.605 | 0.655 | 0.61 | 0.59 | 0.67672 | 0.595 | 0.59 | 0.58 | 0.59 | 0.63 | 0.565 | 0.61 | 0.94091 | 0.615 | 0.635 | 0.63 | 0.616 | 0.63 | 0.665
| Std Dev. |  | 0.010954 | 0.0089948 | 0.0054772 | 0.0040825 | 0.0023905 | 0.0089443 | 0.0054772 | 0.02137 | 0.012247 | 0.01169 | 0.0095169 | 0.0054772 | 0.019664 | 0.0089443 | 0.012649 | 0.018348 | 0.012111 | 0.0063246 | 0.0094998 | 0.010488 | 0.012111 | 0.01169 | 0.0076681 | 0.02582 | 0.061536
| Rel.Std.Dev. |  | 1.767% | 1.521% | 0.8764% | 0.6463% | 0.3931% | 1.42% | 0.9053% | 3.296% | 1.991% | 1.987% | 1.413% | 0.9205% | 3.352% | 1.542% | 2.144% | 2.951% | 2.137% | 1.037% | 1.008% | 1.705% | 1.912% | 1.861% | 1.245% | 4.077% | 9.23%
| PDM3 |  | 0.8519% | -3.814% | 1.665% | 2.75% | -1.078% | 2.479% | -1.588% | 5.461% | 0.03857% | -4.299% | 9.552% | -3.215% | -4.57% | -5.655% | -4.028% | 1.123% | -7.824% | -0.7747% | 53.36% | 0.03857% | 3.021% | 2.207% | 0.2012% | 3.021% | 8.443%
| Z-Score (Absolute) |  | 0.17 | 0.92 | 0.36 | 0.61 | 0.28 | 0.55 | 0.4 | 1.24 | 0.02 | 1.03 | 2.19 | 0.78 | 1.09 | 1.34 | 0.97 | 0.23 | 1.85 | 0.21 | 12.38 | 0.02 | 0.67 | 0.49 | 0 | 0.67 | 1.94
| Analytical results for Ga in OREAS 753b (Certified Value 46.5 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 48.25 | 43.5644 | 44.93 | 45.2 | 48.39 | 48.5 | 49.2 | 49.32 | 48.5 | 49 | 46.3093 | 44.6 | 50.7 | 46.3 | 46.5 | 47.4 | 52 | 42.3 | 43.2 | 49.5312 | 44.61 | 43.33 | 44.3 | 20.07 | 42.26 |  | 1
| 1 | 2 | 49.17 | 44.7862 | 44.74 | 47.41 | 49.26 | 47.2 | 50.2 | 47.12 | 48.2 | 48.7 | 46.8713 | 41.8 | 50.4 | 48.6 | 46.2 | 44.7 | 49.8 | 41.6 | 41.2 | 49.5029 | 44.46 | 41.73 | 44.9 | 17.95 | 40.48 |  | 21
| 1 | 3 | 49.62 | 44.9667 | 43.84 | 45.59 | 49.48 | 47.9 | 50.6 | 49.68 | 47.4 | 47.4 | 47.0094 | 43 | 49.2 | 52.4 | 45.4 | 45.4 | 49.5 | 42.5 | 42.1 | 48.5967 | 43.85 | 41.57 | 45.4 | 16.23 | 40.85 |  | 16
| 1 | 4 | 49.61 | 43.7831 | 45.86 | 45.69 | 48.74 | 48.6 | 48.6 | 49.28 | 48.3 | 48.3 | 46.7491 | 43.3 | 51.4 | 49.1 | 46.1 | 47.2 | 47.9 | 42.4 | 42.8 | 49.5322 | 45.17 | 40.94 | 44.6 | 15.31 | 43.46 |  | 46.5152
| 1 | 5 | 49.19 | 44.931 | 45.63 | 47.5 | 48.82 | 48.1 | 49.3 | 49.08 | 49.2 | 48.4 | 47.4944 | 42.6 | 51.2 | 48.7 | 44.9 | 47.9 | 48.6 | 41.6 | 42.9 | 49.5538 | 44.36 | 42.53 | 44 | 15.14 | 43.13 |  | 24
| 1 | 6 | 50.37 | 42.2421 | 45.32 | 47.74 | 49.21 | 47.6 | 49.6 | 49.68 | 48.8 | 50.8 | 47.2967 | 43.1 | 49.1 | 50.7 | 45.7 | 49.8 | 49 | 42 | 43.2 | 49.5669 | 43.68 | 41.36 | 43.4 | 14.73 | 39.92
| Mean |  | 49.3683 | 44.0456 | 45.0533 | 46.5217 | 48.9833 | 47.9833 | 49.5833 | 49.0267 | 48.4 | 48.7667 | 46.955 | 43.0667 | 50.3333 | 49.3 | 45.8 | 47.0667 | 49.4667 | 42.0667 | 42.5667 | 49.3806 | 44.355 | 41.91 | 44.4333 | 16.5717 | 41.6833
| Median |  | 49.4 | 44.2847 | 45.125 | 46.55 | 49.015 | 48 | 49.45 | 49.3 | 48.4 | 48.55 | 46.9404 | 43.05 | 50.55 | 48.9 | 45.9 | 47.3 | 49.25 | 42.15 | 42.85 | 49.5317 | 44.41 | 41.65 | 44.45 | 15.77 | 41.555
| Std Dev. |  | 0.69955 | 1.0709 | 0.72663 | 1.1434 | 0.40312 | 0.53448 | 0.72226 | 0.96355 | 0.60992 | 1.1325 | 0.41881 | 0.92014 | 0.98319 | 2.0717 | 0.58652 | 1.8261 | 1.4109 | 0.39833 | 0.78145 | 0.38466 | 0.53869 | 0.87088 | 0.70048 | 2.0634 | 1.472
| Rel.Std.Dev. |  | 1.417% | 2.431% | 1.613% | 2.458% | 0.823% | 1.114% | 1.457% | 1.965% | 1.26% | 2.322% | 0.8919% | 2.137% | 1.953% | 4.202% | 1.281% | 3.88% | 2.852% | 0.9469% | 1.836% | 0.779% | 1.215% | 2.078% | 1.576% | 12.45% | 3.531%
| PDM3 |  | 6.134% | -5.309% | -3.143% | 0.01395% | 5.306% | 3.156% | 6.596% | 5.399% | 4.052% | 4.84% | 0.9456% | -7.414% | 8.208% | 5.987% | -1.538% | 1.186% | 6.345% | -9.564% | -8.489% | 6.16% | -4.644% | -9.9% | -4.476% | -64.37% | -10.39%
| Z-Score (Absolute) |  | 0.67 | 0.81 | 0.53 | 0.12 | 0.56 | 0.29 | 0.73 | 0.58 | 0.4 | 0.5 | 0 | 1.08 | 0.94 | 0.65 | 0.32 | 0.03 | 0.7 | 1.36 | 1.22 | 0.67 | 0.72 | 1.4 | 0.7 | 8.45 | 1.47
| Analytical results for Gd in OREAS 753b (Certified Value 0.71 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.22; 1.08: 1.24; 1.09: 1.25 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 0.7; 1.02: 0.73; 1.03: 0.6; 1.04: 0.8; 1.05: 0.79293; 1.06: 0.71; 1.07: 0.5; 1.08: 0.8; 1.09: 0.77 |  | 1
| 1 | 2 | 1.01: 0.69; 1.02: 0.72; 1.03: 0.6; 1.04: 0.7; 1.05: 0.726; 1.06: 0.69; 1.07: 0.5; 1.08: 0.7; 1.09: 0.68 |  | 22
| 1 | 3 | 1.01: 0.64; 1.02: 0.73; 1.03: 0.7; 1.04: 0.7; 1.05: 0.70512; 1.06: 0.7; 1.07: 0.5; 1.08: 0.7; 1.09: 0.71 |  | 16
| 1 | 4 | 1.01: 0.62; 1.02: 0.74; 1.03: 0.6; 1.04: 0.6; 1.05: 0.72368; 1.06: 0.9; 1.07: 0.5; 1.08: 0.7; 1.09: 0.7 |  | 0.71187
| 1 | 5 | 1.01: 0.71; 1.02: 0.73; 1.03: 0.6; 1.04: 0.7; 1.05: 0.78338; 1.06: 0.66; 1.07: 0.7; 1.08: 0.7; 1.09: 0.71 |  | 25
| 1 | 6 | 1.01: 0.71; 1.02: 0.7; 1.03: 0.6; 1.04: 0.7; 1.05: 0.76294; 1.06: 0.7; 1.07: 0.5; 1.08: 0.7; 1.09: 0.72
| Mean |  | 1.01: 0.67833; 1.02: 0.725; 1.03: 0.61667; 1.04: 0.7; 1.05: 0.74901; 1.06: 0.72667; 1.07: 0.53333; 1.08: 0.71667; 1.09: 0.715
| Median |  | 1.01: 0.695; 1.02: 0.73; 1.03: 0.6; 1.04: 0.7; 1.05: 0.74447; 1.06: 0.7; 1.07: 0.5; 1.08: 0.7; 1.09: 0.71
| Std Dev. |  | 1.01: 0.038687; 1.02: 0.013784; 1.03: 0.040825; 1.04: 0.063246; 1.05: 0.035782; 1.06: 0.086641; 1.07: 0.08165; 1.08: 0.040825; 1.09: 0.030166
| Rel.Std.Dev. |  | 1.01: 5.703%; 1.02: 1.901%; 1.03: 6.62%; 1.04: 9.035%; 1.05: 4.777%; 1.06: 11.92%; 1.07: 15.31%; 1.08: 5.696%; 1.09: 4.219%
| PDM3 |  | 1.01: -4.711%; 1.02: 1.845%; 1.03: -13.37%; 1.04: -1.667%; 1.05: 5.217%; 1.06: 2.079%; 1.07: -25.08%; 1.08: 0.674%; 1.09: 0.4399%
| Z-Score (Absolute) |  | 1.01: 3.15; 1.02: 0; 1.03: NA; 1.04: NA; 1.05: 1.62; 1.06: 0.11; 1.07: NA; 1.08: NA; 1.09: 0.67
| Analytical results for Ge in OREAS 753b (Certified Value 0.077 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | [9 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [9 empty] | 3
3
| 1 | 1 | 2.8 | < 0.2 | < 0.05 | 2.7 | 2.7 | 1.3 | 0.28 | 0.3 | < 0.05 | 2.21 | 0.09 | 0.09 | 0.05 | 0.05 | 0.09 | 0.7 | < 0.2 | [9 empty] | 1
| 1 | 2 | 2.8 | < 0.2 | < 0.05 | 2.8 | 2.6 | 1.3 | 0.27 | 0.2 | < 0.05 | 2.12 | 0.08 | 0.1 | 0.05 | 0.05 | 0.05 | 0.8 | < 0.2 | [9 empty] | 0
| 1 | 3 | 2.7 | < 0.2 | < 0.05 | 2.9 | 2.6 | 1.3 | 0.13 | 0.2 | 0.09 | 2.17 | 0.09 | 0.07 | 0.06 | 0.05 | 0.07 | 0.7 | < 0.2 | [9 empty] | 16
| 1 | 4 | 2.8 | < 0.2 | < 0.05 | 2.7 | 2.7 | 1.4 | 0.29 | 0.2 | 0.11 | 2.16 | 0.11 | < 0.05 | 0.07 | 0.06 | 0.08 | 0.8 | < 0.2 | [9 empty] | 0.076944
| 1 | 5 | 2.7 | < 0.2 | < 0.05 | 2.7 | 2.7 | 1.4 | 0.22 | 0.2 | < 0.05 | 2.12 | 0.09 | < 0.05 | 0.06 | 0.06 | 0.06 | 0.7 | < 0.2 | [9 empty] | 26
| 1 | 6 | 2.9 | < 0.2 | < 0.05 | 2.8 | 2.6 | 1.4 | 0.21 | 0.3 | < 0.05 | 2.16 | 0.1 | < 0.05 | 0.05 | 0.08 | 0.05 | 0.8 | < 0.2
| Mean |  | 2.7833 |  |  | 2.7667 | 2.65 | 1.35 | 0.23333 | 0.23333 | 0.1 | 2.1567 | 0.093333 | 0.086667 | 0.056667 | 0.058333 | 0.066667 | 0.75
| Median |  | 2.8 |  |  | 2.75 | 2.65 | 1.35 | 0.245 | 0.2 | 0.1 | 2.16 | 0.09 | 0.09 | 0.055 | 0.055 | 0.065 | 0.75
| Std Dev. |  | 0.075277 |  |  | 0.08165 | 0.054772 | 0.054772 | 0.060222 | 0.05164 | 0.014142 | 0.033862 | 0.010328 | 0.015275 | 0.008165 | 0.01169 | 0.01633 | 0.054772
| Rel.Std.Dev. |  | 2.705% |  |  | 2.951% | 2.067% | 4.057% | 25.81% | 22.13% | 14.14% | 1.57% | 11.07% | 17.63% | 14.41% | 20.04% | 24.49% | 7.303%
| PDM3 |  | 3517% |  |  | 3496% | 3344% | 1655% | 203.2% | 203.2% | 29.96% | 2703% | 21.3% | 12.64% | -26.35% | -24.19% | -13.36% | 874.7%
| Z-Score (Absolute) |  | NA | 0.57 | 1.56 | NA | NA | NA | 4.35 | 4.35 | 0.85 | 58.96 | 0.38 | 0.69 | 0.66 | 0.62 | 0.38 | NA | 0.57
| Analytical results for Hf in OREAS 753b (Certified Value 1.86 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.82 | 1.9 | 1.9 | 1.88 | 2 | 1.4 | 1.83 | 1.9 | 1.9 | 2.458 | 1.8 | 1.9 | 1.8 | 1.7 | 2 | 1.9 | 1.92 | 1.89 | 1.81 | 1.9 | 2.1 | 2.39 | 1.97 |  |  |  | 1
| 1 | 2 | 1.88 | 1.8 | 1.82 | 1.78 | 2 | 1.4 | 1.66 | 1.8 | 1.9 | 2.453 | 1.8 | 1.9 | 1.9 | 1.7 | 1.9 | 2.1 | 1.85 | 1.87 | 1.72 | 1.8 | 2.2 | 2.02 | 1.84 |  |  |  | 8
| 1 | 3 | 1.86 | 1.9 | 1.86 | 1.8 | 2 | 1.6 | 2 | 1.7 | 1.9 | 2.4215 | 1.82 | 1.9 | 1.8 | 1.6 | 2 | 1.9 | 1.96 | 1.85 | 1.81 | 1.8 | 2.2 | 1.93 | 2.02 |  |  |  | 16
| 1 | 4 | 1.83 | 1.9 | 1.84 | 1.78 | 2 | 1.4 | 1.88 | 1.8 | 1.9 | 2.5578 | 1.74 | 1.9 | 1.7 | 1.7 | 2 | 2 | 1.91 | 1.93 | 1.75 | 1.9 | 2.2 | 1.9 | 1.99 |  |  |  | 1.8617
| 1 | 5 | 2.09 | 1.9 | 1.95 | 1.88 | 2 | 1.6 | 1.83 | 1.7 | 1.9 | 2.4984 | 1.82 | 1.9 | 1.7 | 1.7 | 1.9 | 1.8 | 1.86 | 1.85 | 1.77 | 1.7 | 2.2 | 1.88 | 1.96 |  |  |  | 27
| 1 | 6 | 1.96 | 1.9 | 1.88 | 1.8 | 2.1 | 1.6 | 1.76 | 1.7 | 2 | 2.4943 | 1.81 | 2 | 1.7 | 1.7 | 2 | 1.9 | 1.94 | 1.83 | 1.75 | 1.6 | 2.3 | 1.89 | 1.8
| Mean |  | 1.9067 | 1.8833 | 1.875 | 1.82 | 2.0167 | 1.5 | 1.8267 | 1.7667 | 1.9167 | 2.4805 | 1.7983 | 1.9167 | 1.7667 | 1.6833 | 1.9667 | 1.9333 | 1.9067 | 1.87 | 1.7683 | 1.7833 | 2.2 | 2.0017 | 1.93
| Median |  | 1.87 | 1.9 | 1.87 | 1.8 | 2 | 1.5 | 1.83 | 1.75 | 1.9 | 2.4761 | 1.805 | 1.9 | 1.75 | 1.7 | 2 | 1.9 | 1.915 | 1.86 | 1.76 | 1.8 | 2.2 | 1.915 | 1.965
| Std Dev. |  | 0.1027 | 0.040825 | 0.046368 | 0.047329 | 0.040825 | 0.10954 | 0.11413 | 0.08165 | 0.040825 | 0.047401 | 0.029944 | 0.040825 | 0.08165 | 0.040825 | 0.05164 | 0.10328 | 0.043665 | 0.035777 | 0.036009 | 0.1169 | 0.063246 | 0.19692 | 0.088544
| Rel.Std.Dev. |  | 5.386% | 2.168% | 2.473% | 2.6% | 2.024% | 7.303% | 6.248% | 4.622% | 2.13% | 1.911% | 1.665% | 2.13% | 4.622% | 2.425% | 2.626% | 5.342% | 2.29% | 1.913% | 2.036% | 6.555% | 2.875% | 9.838% | 4.588%
| PDM3 |  | 2.415% | 1.162% | 0.7144% | -2.24% | 8.324% | -19.43% | -1.882% | -5.105% | 2.952% | 33.24% | -3.404% | 2.952% | -5.105% | -9.581% | 5.638% | 3.848% | 2.415% | 0.4458% | -5.015% | -4.209% | 18.17% | 7.518% | 3.669%
| Z-Score (Absolute) |  | 0.19 | 0 | 0.07 | 0.51 | 1.08 | 3.1 | 0.46 | 0.94 | 0.27 | 4.83 | 0.69 | 0.27 | 0.94 | 1.62 | 0.67 | 0.4 | 0.19 | 0.11 | 0.93 | 0.81 | 2.56 | 0.96 | 0.38
| Analytical results for Hg in OREAS 753b (Indicative Value < 1 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab |  |  |  | 1
| No | No | 1.01: 1.02 |  |  |  | ppm
|  |  | 1.01: 4A*MS |  |  |  | 2
2
| 1 | 1 | 1.01: < 1 |  |  |  | 1
| 1 | 2 | 1.01: < 1 |  |  |  | 2
| 1 | 3 | 1.01: < 1 |  |  |  | 16
| 1 | 4 | 1.01: < 1 |  |  |  | < 1
| 1 | 5 | 1.01: < 1 |  |  |  | 8
| 1 | 6 | 1.01: < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ho in OREAS 753b (Certified Value 0.076 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.05; 1.04: 1.08; 1.05: 1.09; 1.06: 1.11; 1.07: 1.12; 1.08: 1.20; 1.09: 1.22; 1.10: 1.24; 1.11: 1.25 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*MS |  |  |  | 3
3
| 1 | 1 | 1.01: 0.05; 1.03: 0.09; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.11877; 1.07: 0.07; 1.08: < 0.1; 1.09: 0.06; 1.10: 0.09; 1.11: 0.1 |  |  |  | 1
| 1 | 2 | 1.01: 0.08; 1.03: 0.08; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.12947; 1.07: 0.07; 1.08: < 0.1; 1.09: < 0.05; 1.10: 0.09; 1.11: 0.08 |  |  |  | 11
| 1 | 3 | 1.01: 0.07; 1.03: 0.08; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.10984; 1.07: 0.07; 1.08: < 0.1; 1.09: < 0.05; 1.10: 0.09; 1.11: 0.08 |  |  |  | 16
| 1 | 4 | 1.01: 0.07; 1.03: 0.08; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.10904; 1.07: 0.07; 1.08: < 0.1; 1.09: < 0.05; 1.10: 0.09; 1.11: 0.08 |  |  |  | 0.076111
| 1 | 5 | 1.01: 0.08; 1.03: 0.08; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.11647; 1.07: 0.06; 1.08: < 0.1; 1.09: 0.06; 1.10: 0.09; 1.11: 0.08 |  |  |  | 28
| 1 | 6 | 1.01: 0.08; 1.03: 0.08; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.12057; 1.07: 0.07; 1.08: < 0.1; 1.09: < 0.05; 1.10: 0.09; 1.11: 0.09
| Mean |  | 1.01: 0.071667; 1.03: 0.081667; 1.06: 0.11736; 1.07: 0.068333; 1.09: 0.06; 1.10: 0.09; 1.11: 0.085
| Median |  | 1.01: 0.075; 1.03: 0.08; 1.06: 0.11762; 1.07: 0.07; 1.09: 0.06; 1.10: 0.09; 1.11: 0.08
| Std Dev. |  | 1.01: 0.01169; 1.03: 0.0040825; 1.06: 0.0075566; 1.07: 0.0040825; 1.09: 0; 1.10: 1.5202e-17; 1.11: 0.0083666
| Rel.Std.Dev. |  | 1.01: 16.31%; 1.03: 4.999%; 1.06: 6.439%; 1.07: 5.974%; 1.09: 0%; 1.10: 1.689e-14%; 1.11: 9.843%
| PDM3 |  | 1.01: -5.839%; 1.03: 7.299%; 1.06: 54.2%; 1.07: -10.22%; 1.09: -21.17%; 1.10: 18.25%; 1.11: 11.68%
| Z-Score (Absolute) |  | 1.01: 0.06; 1.03: 0.39; 1.04: 0.67; 1.05: 0.67; 1.06: 1.6; 1.07: 0.06; 1.08: 0.67; 1.09: 1.12; 1.10: 0.67; 1.11: 0.51
| Analytical results for In in OREAS 753b (Indicative Value 0.023 ppm) | [27 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.25 | 1.28 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | [5 empty] | 3
3
| 1 | 1 | 0.02 | < 0.005 | < 0.005 | 0.01 | < 0.05 | < 0.05 | 0.29 | < 0.1 | < 0.005 |  | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.013 | 0.02 | 0.2 | < 0.1 | 0.058 | < 0.005 | < 1 | [5 empty] | 1
| 1 | 2 |  | < 0.005 | 0.007 | 0.03 | < 0.05 | < 0.05 | 0.28 | < 0.1 | < 0.005 | 0.012 | < 0.005 | 0.007 | < 0.005 | < 0.005 | 0.008 | 0.03 | 0.2 | < 0.1 | 0.069 | < 0.005 | < 1 | [5 empty] | 1
| 1 | 3 | 0.04 | < 0.005 | 0.019 | 0.04 | < 0.05 | < 0.05 | 0.31 | < 0.1 | < 0.005 | 0.026 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.014 | 0.02 | 0.19 | < 0.1 | 0.074 | < 0.005 | < 1 | [5 empty] | 16
| 1 | 4 |  | < 0.005 | 0.034 | 0.02 | < 0.05 | < 0.05 | 0.31 | < 0.1 | < 0.005 | 0.006 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.03 | 0.2 | < 0.1 | 0.062 | < 0.005 | < 1 | [5 empty] | 0.023235
| 1 | 5 | 0.15 | < 0.005 | 0.019 | 0.02 | < 0.05 | < 0.05 | 0.29 | < 0.1 | < 0.005 | 0.031 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | 0.015 | 0.02 | 0.21 | < 0.1 | 0.066 | < 0.005 | < 1 | [5 empty] | 7
| 1 | 6 |  | < 0.005 | 0.019 |  | < 0.05 | < 0.05 | 0.29 | < 0.1 | < 0.005 |  | < 0.005 | < 0.005 | 0.005 | < 0.005 | 0.023 | 0.03 | 0.21 | < 0.1 | 0.062 | < 0.005 | < 1
| Mean |  | 1.01: 0.07; 1.04: 0.0196; 1.05: 0.024; 1.08: 0.295; 1.12: 0.01875; 1.14: 0.007; 1.15: 0.005; 1.17: 0.0146; 1.18: 0.025; 1.19: 0.20167; 1.22: 0.065167
| Median |  | 1.01: 0.04; 1.04: 0.019; 1.05: 0.02; 1.08: 0.29; 1.12: 0.019; 1.14: 0.007; 1.15: 0.005; 1.17: 0.014; 1.18: 0.025; 1.19: 0.2; 1.22: 0.064
| Std Dev. |  | 1.01: 0.07; 1.04: 0.0095812; 1.05: 0.011402; 1.08: 0.012247; 1.12: 0.011701; 1.17: 0.0054129; 1.18: 0.0054772; 1.19: 0.0075277; 1.22: 0.0057417
| Rel.Std.Dev. |  | 1.01: 100%; 1.04: 48.88%; 1.05: 47.51%; 1.08: 4.152%; 1.12: 62.41%; 1.17: 37.07%; 1.18: 21.91%; 1.19: 3.733%; 1.22: 8.811%
| PDM3 |  | 1.01: 201.3%; 1.04: -15.65%; 1.05: 3.292%; 1.08: 1170%; 1.12: -19.3%; 1.14: -69.87%; 1.15: -78.48%; 1.17: -37.16%; 1.18: 7.595%; 1.19: 767.9%; 1.22: 180.5%
| Z-Score (Absolute) |  | 1.44 | 0.67 | 0.23 | 0 | 0.03 | 0.03 | 8.5 | 0.82 | 0.67 | 0.16 | 0.67 | 0.65 | 0.66 | 0.67 | 0.36 | 0.03 | 5.57 | 0.82 | 1.29 | 0.67 | 14.93
| Analytical results for K in OREAS 753b (Certified Value 0.938 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.9567 | 0.91734 | 0.98 | 0.93 | 0.9744 | 0.931 | 0.919 | 0.95 | 1 | 0.94 | 0.917 | 0.97 | 0.91 | 0.88 | 0.9 | 0.94 | 0.91 | 0.98 | 0.92777 | 0.95 | 0.99 | 0.91 | 0.921 | 1.19 | 0.79 |  | 1
| 1 | 2 | 0.9583 | 0.94072 | 0.98 | 0.93 | 0.9806 | 0.943 | 0.921 | 0.92 | 1 | 0.95 | 0.996 | 0.95 | 0.92 | 0.86 | 0.91 | 0.97 | 0.88 | 0.99 | 0.92644 | 0.94 | 1.01 | 0.91 | 0.946 | 1.12 | 0.85 |  | 0
| 1 | 3 | 0.9786 | 0.91388 | 0.98 | 0.94 | 0.9807 | 0.928 | 0.907 | 0.96 | 1.01 | 0.91 | 0.996 | 0.94 | 0.92 | 0.86 | 0.91 | 0.93 | 0.92 | 1 | 0.9268 | 0.93 | 0.99 | 0.91 | 0.913 | 1.09 | 0.87 |  | 16
| 1 | 4 | 0.9845 | 0.90077 | 0.96 | 0.94 | 0.9947 | 0.924 | 0.924 | 0.96 | 1.01 | 0.92 | 0.98 | 0.94 | 0.88 | 0.89 | 0.91 | 0.94 | 0.89 | 1 | 0.92552 | 0.95 | 1 | 0.9 | 0.93 | 1.1 | 0.89 |  | 0.93837
| 1 | 5 | 0.984 | 0.93181 | 0.97 | 0.94 | 0.9859 | 0.882 | 0.912 | 0.96 | 1.02 | 0.92 | 0.954 | 0.94 | 0.89 | 0.88 | 0.89 | 0.9 | 0.92 | 1 | 0.92489 | 0.94 | 0.98 | 0.89 | 0.921 | 1.05 | 0.88 |  | 29
| 1 | 6 | 0.979 | 0.89009 | 0.97 | 0.95 | 0.9834 | 0.876 | 0.909 | 0.96 | 1.03 | 0.96 | 1.02 | 0.96 | 0.89 | 0.88 | 0.94 | 0.95 | 0.9 | 1 | 0.92548 | 0.96 | 0.98 | 0.9 | 0.921 | 1.06 | 0.82
| Mean |  | 0.97352 | 0.91577 | 0.97333 | 0.93833 | 0.98328 | 0.914 | 0.91533 | 0.95167 | 1.0117 | 0.93333 | 0.97717 | 0.95 | 0.90167 | 0.875 | 0.91 | 0.93833 | 0.90333 | 0.995 | 0.92615 | 0.945 | 0.99167 | 0.90333 | 0.92533 | 1.1017 | 0.85
| Median |  | 0.9788 | 0.91561 | 0.975 | 0.94 | 0.98205 | 0.926 | 0.9155 | 0.96 | 1.01 | 0.93 | 0.988 | 0.945 | 0.9 | 0.88 | 0.91 | 0.94 | 0.905 | 1 | 0.92598 | 0.945 | 0.99 | 0.905 | 0.921 | 1.095 | 0.86
| Std Dev. |  | 0.012655 | 0.018816 | 0.008165 | 0.0075277 | 0.0067822 | 0.027907 | 0.0069474 | 0.016021 | 0.01169 | 0.019664 | 0.03663 | 0.012649 | 0.017224 | 0.012247 | 0.016733 | 0.023166 | 0.01633 | 0.0083666 | 0.0010538 | 0.010488 | 0.01169 | 0.008165 | 0.011466 | 0.050365 | 0.038471
| Rel.Std.Dev. |  | 1.3% | 2.055% | 0.8389% | 0.8022% | 0.6897% | 3.053% | 0.759% | 1.683% | 1.156% | 2.107% | 3.749% | 1.331% | 1.91% | 1.4% | 1.839% | 2.469% | 1.808% | 0.8409% | 0.1138% | 1.11% | 1.179% | 0.9039% | 1.239% | 4.572% | 4.526%
| PDM3 |  | 3.746% | -2.408% | 3.726% | -0.00372% | 4.787% | -2.597% | -2.455% | 1.417% | 7.811% | -0.5366% | 4.135% | 1.24% | -3.911% | -6.753% | -3.023% | -0.00372% | -3.734% | 6.035% | -1.302% | 0.7067% | 5.68% | -3.734% | -1.389% | 17.4% | -9.417%
| Z-Score (Absolute) |  | 0.68 | 0.43 | 0.67 | 0 | 0.87 | 0.47 | 0.44 | 0.26 | 1.41 | 0.1 | 0.75 | 0.22 | 0.71 | 1.22 | 0.55 | 0 | 0.67 | 1.09 | 0.23 | 0.13 | 1.03 | 0.67 | 0.24 | 3.15 | 1.7
| Analytical results for La in OREAS 753b (Certified Value 3.8 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | 4.74 | 2.9859 | 3.2 | 4.76 | 4.52 | 3.2 | 5.1 | 4.7 | 4.4512 | 3.78 | 3.5 | 3.8 | 3.4 | 2.5 | 3.5 | 2.1 | 4 | 4.435 | 3.5 | 2.76 | 4.3 | 5 | 4.41 | 2.72 |  |  | 1
| 1 | 2 | 4.32 | 3.1691 | 3.2 | 4.57 | 4.43 | 2.6 | 4.3 | 5.4 | 3.9319 | 3.74 | 3 | 3.8 | 3.3 | 2.6 | 4.1 | 1.8 | 3.7 | 4.6061 | 3.6 | 2.13 | 3.7 | 4 | 4.42 | 3.41 |  |  | 13
| 1 | 3 | 4.32 | 3.124 | 3.3 | 4.57 | 4.4 | 3.4 | 4.1 | 4.6 | 4.0248 | 3.72 | 3.3 | 3.4 | 3.1 | 2.7 | 3.4 | 2 | 3.8 | 4.47 | 3.5 | 2.27 | 4.2 | 4 | 4.75 | 3.1 |  |  | 16
| 1 | 4 | 4.11 | 3.2068 | 3.1 | 4.72 | 4.55 | 3.2 | 4.1 | 5.2 | 3.9448 | 4.1 | 3.6 | 2.7 | 3.7 | 2.7 | 3.8 | 1.8 | 4.4 | 4.5689 | 3.7 | 2.5 | 4 | 4 | 4.59 | 4.18 |  |  | 3.8022
| 1 | 5 | 4.42 | 3.0956 | 2.8 | 4.56 | 4.13 | 3.4 | 4.8 | 5 | 4.2355 | 3.62 | 3 | 2.9 | 3.3 | 2.9 | 3.3 | 1.9 | 4.1 | 4.4103 | 3.7 | 3.39 | 3.1 | 4 | 4.49 | 4.31 |  |  | 30
| 1 | 6 | 4.64 | 3.182 | 3 | 4.69 | 4.05 | 3.2 | 4.6 | 5.9 | 4.3108 | 3.94 | 3.4 | 2.8 | 3.2 | 2.8 | 3.4 | 2.1 | 4.2 | 4.6247 | 3.6 | 2.76 | 3.4 | 4 | 4.57 | 3.29
| Mean |  | 4.425 | 3.1272 | 3.1 | 4.645 | 4.3467 | 3.1667 | 4.5 | 5.1333 | 4.1498 | 3.8167 | 3.3 | 3.2333 | 3.3333 | 2.7 | 3.5833 | 1.95 | 4.0333 | 4.5191 | 3.6 | 2.635 | 3.7833 | 4.1667 | 4.5383 | 3.5017
| Median |  | 4.37 | 3.1465 | 3.15 | 4.63 | 4.415 | 3.2 | 4.45 | 5.1 | 4.1302 | 3.76 | 3.35 | 3.15 | 3.3 | 2.7 | 3.45 | 1.95 | 4.05 | 4.5194 | 3.6 | 2.63 | 3.85 | 4 | 4.53 | 3.35
| Std Dev. |  | 0.23098 | 0.080067 | 0.17889 | 0.088713 | 0.20791 | 0.29439 | 0.40497 | 0.48028 | 0.21411 | 0.17363 | 0.25298 | 0.50067 | 0.20656 | 0.14142 | 0.30605 | 0.13784 | 0.2582 | 0.092213 | 0.089443 | 0.44894 | 0.47081 | 0.40825 | 0.1275 | 0.62275
| Rel.Std.Dev. |  | 5.22% | 2.56% | 5.77% | 1.91% | 4.783% | 9.297% | 8.999% | 9.356% | 5.16% | 4.549% | 7.666% | 15.48% | 6.197% | 5.238% | 8.541% | 7.069% | 6.402% | 2.04% | 2.485% | 17.04% | 12.44% | 9.798% | 2.809% | 17.78%
| PDM3 |  | 16.38% | -17.75% | -18.47% | 22.17% | 14.32% | -16.72% | 18.35% | 35.01% | 9.142% | 0.3798% | -13.21% | -14.96% | -12.33% | -28.99% | -5.757% | -48.71% | 6.078% | 18.86% | -5.319% | -30.7% | -0.4969% | 9.585% | 19.36% | -7.905%
| Z-Score (Absolute) |  | 0.91 | 0.7 | 0.73 | 1.18 | 0.81 | 0.65 | 1 | 1.78 | 0.57 | 0.15 | 0.48 | 0.57 | 0.44 | 1.23 | 0.13 | 2.16 | 0.42 | 1.02 | 0.11 | 1.31 | 0.11 | 0.59 | 1.05 | 0.24
| Analytical results for Li in OREAS 753b (Certified Value 1.03 wt.%) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 1.0424 | 1.038 | 1.039 | 0.9926 | 1.0544 | 1.09 | 1.0215 | 1.059 | > 1 | 1.1 | 1.0251 | 1.07 | 1.03 | 0.971 | > 1 | > 1 | 1.025 | 0.9665 | 0.9502 | 1.064 | > 1 |  | 1.0221 | 1.04 | > 1 | 0.9982 | 1
| 1 | 2 | 1.0603 | 1.054 | 1.039 | > 1 | 1.0446 | 1.07 | 1.0244 | 1.036 | > 1 | 1.07 | 1.0346 | 1 | 1.03 | 0.972 | 1 | > 1 | 1.055 | 0.9702 | 0.9526 | 1.0643 | > 1 |  | 1.0212 | 1.06 | > 1 | 0.9986 | 0
| 1 | 3 | 1.0541 | 1.0424 | 1.021 | 0.9935 | 1.0409 | 1.09 | 1.036 | 1.04 | > 1 | 1.03 | 1.0406 | 0.986 | 1.03 | 1 | 0.999 | > 1 | 1.045 | 0.9655 | 0.9719 | 1.0639 | > 1 |  | 1.0311 | 1.03 | > 1 | 0.9981 | 16
| 1 | 4 | 1.0634 | 0.99748 | 1.046 | 0.995 | 1.0367 | 1.07 | 1.0304 | 1.069 | > 1 | 1.09 | 1.0342 | 1.01 | 1.035 | 0.966 | > 1 | > 1 | 1.05 | 0.9636 | 0.9773 | 1.0747 | > 1 |  | 1.0242 | 1.05 | > 1 | 0.9896 | 1.0345
| 1 | 5 | 1.0443 | 1.0271 | 1.032 | > 1 | 1.0454 | 1.09 | 1.0203 | 1.036 | > 1 | 1.06 | 1.0525 | 0.998 | 1.03 | 0.966 | > 1 | > 1 | 1.045 | 0.9739 | 0.9817 | 1.07 | > 1 |  | 1.031 | 1.04 | > 1 | 0.9868 | 31
| 1 | 6 | 1.0139 | 1.0256 | 1.031 | > 1 | 1.0454 | 1.08 | 1.0304 | 1.028 | > 1 | 1.01 | 1.0475 | 1.01 | 1.035 | 0.984 | > 1 | > 1 | 1.045 | 0.9656 | 0.9828 | 1.0652 | > 1 |  | 1.0203 | 1.05 | > 1 | 0.9956
| Mean |  | 1.0464 | 1.0308 | 1.0347 | 0.9937 | 1.0446 | 1.0817 | 1.0272 | 1.0447 |  | 1.06 | 1.0391 | 1.0123 | 1.0317 | 0.9765 | 0.9995 |  | 1.0442 | 0.96755 | 0.96942 | 1.067 |  |  | 1.025 | 1.045 |  | 0.99448
| Median |  | 1.0492 | 1.0325 | 1.0355 | 0.9935 | 1.045 | 1.085 | 1.0274 | 1.038 |  | 1.065 | 1.0376 | 1.005 | 1.03 | 0.9715 | 0.9995 |  | 1.045 | 0.96605 | 0.9746 | 1.0647 |  |  | 1.0232 | 1.045 |  | 0.99685
| Std Dev. |  | 0.01799 | 0.019373 | 0.008641 | 0.0012124 | 0.0058892 | 0.0098319 | 0.0060968 | 0.015795 |  | 0.034641 | 0.0099543 | 0.029622 | 0.002582 | 0.013263 | 0.00070711 |  | 0.010206 | 0.0037962 | 0.014494 | 0.0044141 |  |  | 0.0048866 | 0.010488 |  | 0.005059
| Rel.Std.Dev. |  | 1.719% | 1.879% | 0.8351% | 0.122% | 0.5638% | 0.909% | 0.5936% | 1.512% |  | 3.268% | 0.958% | 2.926% | 0.2503% | 1.358% | 0.07075% |  | 0.9775% | 0.3923% | 1.495% | 0.4137% |  |  | 0.4767% | 1.004% |  | 0.5087%
| PDM3 |  | 1.146% | -0.3651% | 0.0116% | -3.948% | 0.9685% | 4.555% | -0.7134% | 0.9782% |  | 2.46% | 0.4384% | -2.147% | -0.2784% | -5.611% | -3.388% |  | 0.9299% | -6.476% | -6.296% | 3.138% |  |  | -0.9233% | 1.01% |  | -3.873%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 0.14 | 1.74 | 0.59 | 2.29 | 0.21 | 0.59 | NA | 1.3 | 0.34 | 0.88 | 0 | 2.52 | 1.47 | NA | 0.57 | 2.93 | 2.85 | 1.62 | NA | NA | 0.31 | 0.61 | NA | 1.7
| Analytical results for Li2O in OREAS 753b (Certified Value 2.23 wt.%) | [27 empty] | Certified
Li2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 2.244 | 2.2345 | 2.2367 | 2.1368 | 2.2698 | 2.3464 | 2.199 | 2.2797 | > 2.2 | 2.368 | 2.2067 | 2.3034 | 2.2173 | 2.0903 | > 2.2 | > 2.2 | 2.2065 | 2.0806 | 2.0455 | 2.2905 | > 2.2 |  | 2.2002 | 2.2388 | > 2.2 | 2.1488 | 1
| 1 | 2 | 2.2825 | 2.269 | 2.2367 | > 2.2 | 2.2487 | 2.3034 | 2.2052 | 2.2302 | > 2.2 | 2.3034 | 2.2272 | 2.1527 | 2.2173 | 2.0924 | 2.1527 | > 2.2 | 2.2711 | 2.0886 | 2.0507 | 2.291 | > 2.2 |  | 2.1984 | 2.2819 | > 2.2 | 2.1497 | 0
| 1 | 3 | 2.2692 | 2.2439 | 2.1979 | 2.1387 | 2.2407 | 2.3464 | 2.2302 | 2.2388 | > 2.2 | 2.2173 | 2.2401 | 2.1226 | 2.2173 | 2.1527 | 2.1505 | > 2.2 | 2.2496 | 2.0784 | 2.0922 | 2.2903 | > 2.2 |  | 2.2197 | 2.2173 | > 2.2 | 2.1486 | 16
| 1 | 4 | 2.2892 | 2.1473 | 2.2517 | 2.1419 | 2.2317 | 2.3034 | 2.2181 | 2.3012 | > 2.2 | 2.3464 | 2.2262 | 2.1742 | 2.228 | 2.0795 | > 2.2 | > 2.2 | 2.2603 | 2.0743 | 2.1038 | 2.3135 | > 2.2 |  | 2.2049 | 2.2603 | > 2.2 | 2.1303 | 2.2271
| 1 | 5 | 2.2481 | 2.211 | 2.2216 | > 2.2 | 2.2504 | 2.3464 | 2.1964 | 2.2302 | > 2.2 | 2.2819 | 2.2658 | 2.1484 | 2.2173 | 2.0795 | > 2.2 | > 2.2 | 2.2496 | 2.0965 | 2.1133 | 2.3033 | > 2.2 |  | 2.2195 | 2.2388 | > 2.2 | 2.1243 | 32
| 1 | 6 | 2.1826 | 2.2079 | 2.2194 | > 2.2 | 2.2504 | 2.3249 | 2.2181 | 2.213 | > 2.2 | 2.1742 | 2.255 | 2.1742 | 2.228 | 2.1183 | > 2.2 | > 2.2 | 2.2496 | 2.0787 | 2.1157 | 2.293 | > 2.2 |  | 2.1964 | 2.2603 | > 2.2 | 2.1432
| Mean |  | 2.2526 | 2.2189 | 2.2273 | 2.1391 | 2.2486 | 2.3285 | 2.2112 | 2.2489 |  | 2.2819 | 2.2368 | 2.1793 | 2.2209 | 2.1021 | 2.1516 |  | 2.2478 | 2.0828 | 2.0869 | 2.297 |  |  | 2.2065 | 2.2496 |  | 2.1408
| Median |  | 2.2586 | 2.2228 | 2.2292 | 2.1387 | 2.2496 | 2.3357 | 2.2117 | 2.2345 |  | 2.2926 | 2.2336 | 2.1635 | 2.2173 | 2.0913 | 2.1516 |  | 2.2496 | 2.0796 | 2.098 | 2.292 |  |  | 2.2025 | 2.2496 |  | 2.1459
| Std Dev. |  | 0.038728 | 0.041703 | 0.018608 | 0.0026098 | 0.012679 | 0.021165 | 0.013125 | 0.033983 |  | 0.074572 | 0.021429 | 0.063767 | 0.0055255 | 0.028551 | 0.0015203 |  | 0.021975 | 0.0081711 | 0.031201 | 0.0095022 |  |  | 0.010519 | 0.022578 |  | 0.010891
| Rel.Std.Dev. |  | 1.719% | 1.879% | 0.8354% | 0.122% | 0.5639% | 0.909% | 0.5936% | 1.511% |  | 3.268% | 0.958% | 2.926% | 0.2488% | 1.358% | 0.07066% |  | 0.9776% | 0.3923% | 1.495% | 0.4137% |  |  | 0.4767% | 1.004% |  | 0.5087%
| PDM3 |  | 1.146% | -0.3651% | 0.01188% | -3.948% | 0.9685% | 4.555% | -0.7134% | 0.978% |  | 2.46% | 0.4384% | -2.147% | -0.2785% | -5.611% | -3.388% |  | 0.9301% | -6.476% | -6.296% | 3.138% |  |  | -0.9233% | 1.011% |  | -3.873%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 0.14 | 1.74 | 0.59 | 2.29 | 0.21 | 0.59 | NA | 1.3 | 0.34 | 0.88 | 0 | 2.52 | 1.47 | NA | 0.57 | 2.93 | 2.85 | 1.62 | NA | NA | 0.31 | 0.61 | NA | 1.7
| Analytical results for Lu in OREAS 753b (Certified Value 0.025 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.04; 1.03: 1.05; 1.04: 1.08; 1.05: 1.09; 1.06: 1.11; 1.07: 1.12; 1.08: 1.18; 1.09: 1.22; 1.10: 1.25; 1.11: 1.28 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*MS | 3
3
| 1 | 1 | 1.01: 0.02; 1.02: 0.03; 1.03: 0.03; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.04359; 1.08: 0.01; 1.09: 0.02; 1.10: 0.04; 1.11: < 1 | 1
| 1 | 2 | 1.01: 0.03; 1.02: 0.03; 1.03: 0.03; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.05421; 1.08: 0.02; 1.09: 0.02; 1.10: 0.03; 1.11: < 1 | 14
| 1 | 3 | 1.01: 0.02; 1.02: 0.02; 1.03: 0.03; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.04551; 1.08: 0.01; 1.09: 0.02; 1.10: 0.03; 1.11: < 1 | 16
| 1 | 4 | 1.01: 0.02; 1.02: 0.03; 1.03: 0.03; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.04498; 1.08: 0.02; 1.09: 0.02; 1.10: 0.03; 1.11: < 1 | 0.025
| 1 | 5 | 1.01: 0.02; 1.02: 0.02; 1.03: 0.03; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.04217; 1.07: 0.02; 1.08: 0.01; 1.09: 0.02; 1.10: 0.03; 1.11: < 1 | 33
| 1 | 6 | 1.01: 0.03; 1.02: 0.02; 1.03: 0.03; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.04351; 1.07: 0.02; 1.08: 0.02; 1.09: 0.02; 1.10: 0.03; 1.11: < 1
| Mean |  | 1.01: 0.023333; 1.02: 0.025; 1.03: 0.03; 1.06: 0.045662; 1.07: 0.02; 1.08: 0.015; 1.09: 0.02; 1.10: 0.031667
| Median |  | 1.01: 0.02; 1.02: 0.025; 1.03: 0.03; 1.06: 0.044285; 1.07: 0.02; 1.08: 0.015; 1.09: 0.02; 1.10: 0.03
| Std Dev. |  | 1.01: 0.005164; 1.02: 0.0054772; 1.03: 0; 1.06: 0.0043514; 1.07: 0; 1.08: 0.0054772; 1.09: 0; 1.10: 0.0040825
| Rel.Std.Dev. |  | 1.01: 22.13%; 1.02: 21.91%; 1.03: 0%; 1.06: 9.53%; 1.07: 0%; 1.08: 36.51%; 1.09: 0%; 1.10: 12.89%
| PDM3 |  | 1.01: -6.667%; 1.02: -2.22e-14%; 1.03: 20%; 1.06: 82.65%; 1.07: -20%; 1.08: -40%; 1.09: -20%; 1.10: 26.67%
| Z-Score (Absolute) |  | 1.01: 0.45; 1.02: 0.34; 1.03: 0; 1.04: 1.35; 1.05: 1.35; 1.06: 1.06; 1.07: 0.67; 1.08: 1.01; 1.09: 0.67; 1.10: 0.11; 1.11: 31.69
| Analytical results for Mg in OREAS 753b (Certified Value 0.12 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.113 | 0.11724 | 0.13 | 0.124 | 0.1257 | 0.127 | 0.125 | 0.11 | 0.12 | 0.13 | 0.0772 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.12 | 0.12303 | 0.12 | 0.128 | 0.12 | 0.1339 | 0.132 | 0.11 |  | 1
| 1 | 2 | 0.1201 | 0.11977 | 0.13 | 0.124 | 0.1241 | 0.126 | 0.125 | 0.1 | 0.12 | 0.13 | 0.0847 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.11 | 0.12 | 0.12373 | 0.12 | 0.132 | 0.12 | 0.1357 | 0.122 | 0.11 |  | 0
| 1 | 3 | 0.1215 | 0.11691 | 0.13 | 0.124 | 0.1255 | 0.125 | 0.123 | 0.12 | 0.12 | 0.13 | 0.0912 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.11 | 0.12 | 0.12525 | 0.12 | 0.131 | 0.12 | 0.1327 | 0.12 | 0.12 |  | 16
| 1 | 4 | 0.1182 | 0.11561 | 0.13 | 0.125 | 0.1259 | 0.126 | 0.125 | 0.12 | 0.12 | 0.13 | 0.0844 | 0.11 | 0.09 | 0.12 | 0.11 | 0.12 | 0.12 | 0.12 | 0.12532 | 0.12 | 0.126 | 0.12 | 0.1339 | 0.123 | 0.13 |  | 0.12043
| 1 | 5 | 0.1131 | 0.11979 | 0.12 | 0.124 | 0.1256 | 0.119 | 0.124 | 0.12 | 0.13 | 0.13 | 0.0821 | 0.11 | 0.1 | 0.11 | 0.11 | 0.11 | 0.11 | 0.12 | 0.1235 | 0.12 | 0.108 | 0.11 | 0.1351 | 0.118 | 0.13 |  | 34
| 1 | 6 | 0.1282 | 0.11535 | 0.12 | 0.125 | 0.1252 | 0.12 | 0.123 | 0.12 | 0.12 | 0.13 | 0.0911 | 0.11 | 0.1 | 0.11 | 0.11 | 0.12 | 0.11 | 0.12 | 0.12368 | 0.12 | 0.12 | 0.11 | 0.1339 | 0.118 | 0.11
| Mean |  | 0.11902 | 0.11745 | 0.12667 | 0.12433 | 0.12533 | 0.12383 | 0.12417 | 0.115 | 0.12167 | 0.13 | 0.085117 | 0.11 | 0.10333 | 0.11167 | 0.11 | 0.11667 | 0.11333 | 0.12 | 0.12409 | 0.12 | 0.12417 | 0.11667 | 0.1342 | 0.12217 | 0.11833
| Median |  | 0.11915 | 0.11708 | 0.13 | 0.124 | 0.12555 | 0.1255 | 0.1245 | 0.12 | 0.12 | 0.13 | 0.08455 | 0.11 | 0.105 | 0.11 | 0.11 | 0.12 | 0.11 | 0.12 | 0.1237 | 0.12 | 0.127 | 0.12 | 0.1339 | 0.121 | 0.115
| Std Dev. |  | 0.0057192 | 0.0019479 | 0.005164 | 0.0005164 | 0.00064704 | 0.0034303 | 0.00098319 | 0.0083666 | 0.0040825 | 0 | 0.0053901 | 0 | 0.008165 | 0.0040825 | 0 | 0.005164 | 0.005164 | 0 | 0.00096206 | 0 | 0.0089981 | 0.005164 | 0.0010564 | 0.0052313 | 0.0098319
| Rel.Std.Dev. |  | 4.805% | 1.659% | 4.077% | 0.4153% | 0.5163% | 2.77% | 0.7918% | 7.275% | 3.355% | 0% | 6.333% | 0% | 7.902% | 3.656% | 0% | 4.426% | 4.556% | 0% | 0.7753% | 0% | 7.247% | 4.426% | 0.7872% | 4.282% | 8.309%
| PDM3 |  | -1.177% | -2.481% | 5.175% | 3.237% | 4.068% | 2.822% | 3.099% | -4.513% | 1.023% | 7.942% | -29.33% | -8.664% | -14.2% | -7.28% | -8.664% | -3.129% | -5.896% | -0.3609% | 3.031% | -0.3609% | 3.099% | -3.129% | 11.43% | 1.438% | -1.745%
| Z-Score (Absolute) |  | 0.16 | 0.41 | 1.08 | 0.7 | 0.86 | 0.62 | 0.67 | 0.81 | 0.27 | 1.62 | 5.65 | 1.62 | 2.7 | 1.35 | 1.62 | 0.54 | 1.08 | 0 | 0.66 | 0 | 0.67 | 0.54 | 2.3 | 0.35 | 0.27
| Analytical results for Mn in OREAS 753b (Certified Value 0.133 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.136 | 0.13455 | 0.1365 | 0.134 | 0.1316 | 0.13 | 0.1302 | 0.1428 | 0.131 | 0.131 | 0.13 | 0.131 | 0.1265 | 0.1285 | 0.127 | 0.142 | 0.1272 | 0.1303 | 0.13524 | 0.1339 | 0.1384 | 0.1267 | 0.1293 | 0.1549 | 0.13884 |  | 1
| 1 | 2 | 0.1363 | 0.13633 | 0.1374 | 0.133 | 0.1323 | 0.126 | 0.1301 | 0.136 | 0.136 | 0.13 | 0.129 | 0.132 | 0.1295 | 0.126 | 0.1285 | 0.142 | 0.1296 | 0.1284 | 0.13583 | 0.1349 | 0.1426 | 0.1271 | 0.1317 | 0.1408 | 0.13146 |  | 0
| 1 | 3 | 0.1389 | 0.13401 | 0.1373 | 0.1357 | 0.1321 | 0.134 | 0.1299 | 0.145 | 0.127 | 0.129 | 0.126 | 0.1295 | 0.1325 | 0.124 | 0.129 | 0.139 | 0.1309 | 0.1288 | 0.13411 | 0.1324 | 0.1367 | 0.1287 | 0.1293 | 0.1407 | 0.13479 |  | 16
| 1 | 4 | 0.1398 | 0.13106 | 0.1352 | 0.1346 | 0.1325 | 0.134 | 0.1322 | 0.1447 | 0.14 | 0.127 | 0.129 | 0.1325 | 0.123 | 0.1295 | 0.1295 | 0.1395 | 0.1272 | 0.1313 | 0.1344 | 0.1374 | 0.1402 | 0.1269 | 0.1301 | 0.1438 | 0.14415 |  | 0.13262
| 1 | 5 | 0.1402 | 0.13507 | 0.1346 | 0.1348 | 0.1326 | 0.129 | 0.1296 | 0.1455 | 0.132 | 0.126 | 0.129 | 0.131 | 0.1255 | 0.127 | 0.126 | 0.1355 | 0.1308 | 0.1303 | 0.13235 | 0.1346 | 0.1334 | 0.1287 | 0.1301 | 0.135 | 0.14445 |  | 35
| 1 | 6 | 0.1389 | 0.13121 | 0.135 | 0.1377 | 0.1312 | 0.128 | 0.132 | 0.1465 | 0.134 | 0.132 | 0.124 | 0.131 | 0.1275 | 0.128 | 0.134 | 0.144 | 0.1276 | 0.1316 | 0.13296 | 0.1336 | 0.135 | 0.1319 | 0.1301 | 0.1362 | 0.13294
| Mean |  | 0.13835 | 0.1337 | 0.136 | 0.13497 | 0.13205 | 0.13017 | 0.13067 | 0.14342 | 0.13333 | 0.12917 | 0.12783 | 0.13117 | 0.12742 | 0.12717 | 0.129 | 0.14033 | 0.12888 | 0.13012 | 0.13415 | 0.13447 | 0.13772 | 0.12833 | 0.1301 | 0.1419 | 0.13777
| Median |  | 0.1389 | 0.13428 | 0.13585 | 0.1347 | 0.1322 | 0.1295 | 0.13015 | 0.14485 | 0.133 | 0.1295 | 0.129 | 0.131 | 0.127 | 0.1275 | 0.12875 | 0.14075 | 0.1286 | 0.1303 | 0.13425 | 0.13425 | 0.13755 | 0.1279 | 0.1301 | 0.14075 | 0.13681
| Std Dev. |  | 0.0017807 | 0.0021341 | 0.0012247 | 0.0016108 | 0.00054681 | 0.0032506 | 0.0011308 | 0.0038311 | 0.0044572 | 0.0023166 | 0.0023166 | 0.0010328 | 0.0032927 | 0.0019664 | 0.0027749 | 0.0029944 | 0.0017646 | 0.0012922 | 0.0013187 | 0.0016825 | 0.003392 | 0.0019613 | 0.00087636 | 0.0071465 | 0.0056292
| Rel.Std.Dev. |  | 1.287% | 1.596% | 0.9005% | 1.193% | 0.4141% | 2.497% | 0.8654% | 2.671% | 3.343% | 1.794% | 1.812% | 0.7874% | 2.584% | 1.546% | 2.151% | 2.134% | 1.369% | 0.9931% | 0.983% | 1.251% | 2.463% | 1.528% | 0.6736% | 5.036% | 4.086%
| PDM3 |  | 4.319% | 0.8151% | 2.547% | 1.768% | -0.4314% | -1.851% | -1.474% | 8.139% | 0.5363% | -2.605% | -3.611% | -1.097% | -3.925% | -4.114% | -2.731% | 5.814% | -2.819% | -1.889% | 1.151% | 1.391% | 3.841% | -3.234% | -1.902% | 6.996% | 3.882%
| Z-Score (Absolute) |  | 1.39 | 0.37 | 0.87 | 0.64 | 0 | 0.42 | 0.31 | 2.51 | 0.28 | 0.64 | 0.93 | 0.2 | 1.02 | 1.08 | 0.67 | 1.83 | 0.7 | 0.43 | 0.46 | 0.53 | 1.25 | 0.82 | 0.43 | 2.18 | 1.26
| Analytical results for Mo in OREAS 753b (Certified Value 3.36 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 3.6 | 3.3051 | 3.39 | 3.41 | 3.4 | 3.3 | 2.8 | 3.16 | 5 | 3.48 | 3.2989 | 3.2 | 3.64 | 3.43 | 3.44 | 3.42 | 3.44 | 3.7 | 2.7 | 3 | 3.1 | 3.57 | 3.43 |  |  |  | 1
| 1 | 2 | 3.4 | 3.4952 | 3.28 | 3.56 | 3.5 | 3.3 | 2.8 | 3.02 | 4 | 3.56 | 3.3287 | 3.3 | 3.43 | 3.6 | 3.42 | 3.44 | 3.85 | 3.64 | 2.6 | 3.2 | 3.2 | 3.54 | 3.19 |  |  |  | 15
| 1 | 3 | 3.4 | 3.3493 | 3.32 | 3.28 | 3.5 | 3.3 | 2.9 | 3.24 | 5 | 3.27 | 3.4047 | 3.1 | 3.39 | 3.46 | 3.3 | 3.8 | 3.34 | 3.69 | 2.7 | 3.2 | 3.1 | 3.96 | 3.27 |  |  |  | 16
| 1 | 4 | 3.4 | 3.2608 | 3.25 | 3.26 | 3.4 | 3.2 | 2.8 | 2.9 | 5 | 3.51 | 3.348 | 3.2 | 3.24 | 3.32 | 3.42 | 3.65 | 3.35 | 3.61 | 2.5 | 3.2 | 3 | 3.46 | 3.16 |  |  |  | 3.3625
| 1 | 5 | 3.5 | 3.2853 | 3.38 | 3.25 | 3.3 | 3.3 | 3 | 3.31 | 7 | 3.43 | 3.2727 | 3.3 | 3.37 | 3.31 | 3.35 | 3.6 | 3.47 | 3.7 | 2.8 | 3.1 | 3.1 | 3.47 | 3.47 |  |  |  | 36
| 1 | 6 | 3.6 | 3.1073 | 3.35 | 3.42 | 3.5 | 3.3 | 3 | 2.88 | 5 | 3.86 | 3.3462 | 3.1 | 3.29 | 3.36 | 3.43 | 3.74 | 3.36 | 3.65 | 2.7 | 3.1 | 3.1 | 3.56 | 3.18
| Mean |  | 3.4833 | 3.3005 | 3.3283 | 3.3633 | 3.4333 | 3.2833 | 2.8833 | 3.085 | 5.1667 | 3.5183 | 3.3332 | 3.2 | 3.3933 | 3.4133 | 3.3933 | 3.6083 | 3.4683 | 3.665 | 2.6667 | 3.1333 | 3.1 | 3.5933 | 3.2833
| Median |  | 3.45 | 3.2952 | 3.335 | 3.345 | 3.45 | 3.3 | 2.85 | 3.09 | 5 | 3.495 | 3.3374 | 3.2 | 3.38 | 3.395 | 3.42 | 3.625 | 3.4 | 3.67 | 2.7 | 3.15 | 3.1 | 3.55 | 3.23
| Std Dev. |  | 0.098319 | 0.12607 | 0.055648 | 0.12209 | 0.08165 | 0.040825 | 0.098319 | 0.1793 | 0.98319 | 0.19467 | 0.045494 | 0.089443 | 0.13924 | 0.10912 | 0.055737 | 0.15471 | 0.19426 | 0.037283 | 0.10328 | 0.08165 | 0.063246 | 0.18544 | 0.13501
| Rel.Std.Dev. |  | 2.823% | 3.82% | 1.672% | 3.63% | 2.378% | 1.243% | 3.41% | 5.812% | 19.03% | 5.533% | 1.365% | 2.795% | 4.103% | 3.197% | 1.643% | 4.288% | 5.601% | 1.017% | 3.873% | 2.606% | 2.04% | 5.161% | 4.112%
| PDM3 |  | 3.595% | -1.843% | -1.015% | 0.02577% | 2.108% | -2.353% | -14.25% | -8.252% | 53.66% | 4.635% | -0.8713% | -4.832% | 0.918% | 1.513% | 0.918% | 7.312% | 3.148% | 8.997% | -20.69% | -6.814% | -7.806% | 6.866% | -2.353%
| Z-Score (Absolute) |  | 0.71 | 0.25 | 0.11 | 0.08 | 0.45 | 0.34 | 2.46 | 1.39 | NA | 0.9 | 0.08 | 0.78 | 0.24 | 0.34 | 0.24 | 1.37 | 0.63 | 1.67 | 3.6 | 1.14 | 1.31 | 1.3 | 0.34
Results from laboratory 9 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 753b (Certified Value 0.412 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.4071 | 0.41284 | 0.44 | 0.401 | 0.4231 | 0.4 | 0.404 | 0.439 | 0.46 | 0.41 | 0.39 | 0.42 | 0.4 | 0.39 | 0.41 | 0.42 | 0.52 | 0.41 | 0.40429 | 0.41 | 0.374 | 0.44 | 0.393 | 0.654 | 0.35 |  | 1
| 1 | 2 | 0.4066 | 0.43003 | 0.44 | 0.402 | 0.4251 | 0.394 | 0.408 | 0.426 | 0.44 | 0.41 | 0.415 | 0.41 | 0.41 | 0.38 | 0.41 | 0.42 | 0.49 | 0.42 | 0.40283 | 0.4 | 0.378 | 0.44 | 0.401 | 0.623 | 0.33 |  | 5
| 1 | 3 | 0.4152 | 0.4166 | 0.44 | 0.404 | 0.4238 | 0.397 | 0.404 | 0.442 | 0.46 | 0.41 | 0.42 | 0.41 | 0.42 | 0.38 | 0.41 | 0.41 | 0.51 | 0.42 | 0.40516 | 0.4 | 0.373 | 0.44 | 0.386 | 0.569 | 0.37 |  | 16
| 1 | 4 | 0.419 | 0.40771 | 0.43 | 0.406 | 0.4277 | 0.41 | 0.407 | 0.442 | 0.45 | 0.4 | 0.421 | 0.41 | 0.4 | 0.39 | 0.41 | 0.41 | 0.48 | 0.42 | 0.40341 | 0.41 | 0.378 | 0.44 | 0.401 | 0.672 | 0.39 |  | 0.41235
| 1 | 5 | 0.4193 | 0.41882 | 0.43 | 0.404 | 0.422 | 0.4 | 0.402 | 0.44 | 0.45 | 0.4 | 0.396 | 0.41 | 0.4 | 0.39 | 0.4 | 0.4 | 0.52 | 0.42 | 0.40241 | 0.4 | 0.375 | 0.43 | 0.393 | 0.709 | 0.38 |  | 37
| 1 | 6 | 0.4136 | 0.40766 | 0.43 | 0.409 | 0.4223 | 0.401 | 0.402 | 0.442 | 0.46 | 0.41 | 0.422 | 0.42 | 0.41 | 0.39 | 0.42 | 0.42 | 0.5 | 0.43 | 0.40357 | 0.41 | 0.375 | 0.44 | 0.393 | 0.639 | 0.34
| Mean |  | 0.41347 | 0.41561 | 0.435 | 0.40433 | 0.424 | 0.40033 | 0.4045 | 0.4385 | 0.45333 | 0.40667 | 0.41067 | 0.41333 | 0.40667 | 0.38667 | 0.41 | 0.41333 | 0.50333 | 0.42 | 0.40361 | 0.405 | 0.3755 | 0.43833 | 0.3945 | 0.64433 | 0.36
| Median |  | 0.4144 | 0.41472 | 0.435 | 0.404 | 0.42345 | 0.4 | 0.404 | 0.441 | 0.455 | 0.41 | 0.4175 | 0.41 | 0.405 | 0.39 | 0.41 | 0.415 | 0.505 | 0.42 | 0.40349 | 0.405 | 0.375 | 0.44 | 0.393 | 0.6465 | 0.36
| Std Dev. |  | 0.0055741 | 0.0083967 | 0.0054772 | 0.0028752 | 0.0021279 | 0.0053914 | 0.00251 | 0.006253 | 0.008165 | 0.005164 | 0.014024 | 0.005164 | 0.008165 | 0.005164 | 0.0063246 | 0.008165 | 0.01633 | 0.0063246 | 0.00099301 | 0.0054772 | 0.0020736 | 0.0040825 | 0.0057184 | 0.047327 | 0.023664
| Rel.Std.Dev. |  | 1.348% | 2.02% | 1.259% | 0.7111% | 0.5019% | 1.347% | 0.6205% | 1.426% | 1.801% | 1.27% | 3.415% | 1.249% | 2.008% | 1.336% | 1.543% | 1.975% | 3.244% | 1.506% | 0.246% | 1.352% | 0.5522% | 0.9314% | 1.45% | 7.345% | 6.573%
| PDM3 |  | 0.2716% | 0.7917% | 5.494% | -1.943% | 2.826% | -2.913% | -1.903% | 6.343% | 9.94% | -1.377% | -0.4074% | 0.2393% | -1.377% | -6.228% | -0.5691% | 0.2393% | 22.07% | 1.856% | -2.118% | -1.782% | -8.936% | 6.302% | -4.328% | 56.26% | -12.69%
| Z-Score (Absolute) |  | 0.2 | 0.36 | 1.76 | 0.46 | 0.96 | 0.75 | 0.45 | 2.01 | 3.08 | 0.29 | 0 | 0.19 | 0.29 | 1.73 | 0.05 | 0.19 | 6.69 | 0.67 | 0.51 | 0.41 | 2.54 | 2 | 1.17 | 16.88 | 3.66
| Analytical results for Nb in OREAS 753b (Certified Value 102 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 108.96 | 98.5368 | 109.4 | 76.4 | 101.63 | 101 | 98.7999 | 98.38 | 49.3 | 66.7 | 101.6945 | 112.5 | 101.5 | 104 | 105.5 | 105.5 | 101.5 | 90 | 107.8055 | 92.9 | 86.7 | 109 | 101 | 101.04 |  |  | 1
| 1 | 2 | 103.38 | 101.3784 | 102.3 | 79.3 | 101.74 | 100 | 99.5568 | 92.85 | 36.3 | 78.6 | 102.222 | 111.5 | 103.5 | 99.9 | 102.5 | 105 | 103.1 | 93 | 109.5945 | 78.5 | 82.5 | 108 | 101 | 100.56 |  |  | 0
| 1 | 3 | 102.67 | 98.7952 | 102.3 | 76.5 | 102.75 | 101 | 102.5395 | 80.56 | 35.8 | 96.8 | 102.6132 | 107 | 108 | 100.5 | 106.5 | 104 | 100.2 | 91 | 105.3915 | 78.8 | 83.1 | 106 | 101 | 99.61 |  |  | 16
| 1 | 4 | 104.4 | 96.1278 | 105.5 | 72.1 | 102.4 | 102 | 109.6256 | 81.72 | 33.2 | 75.2 | 101.7542 | 109.5 | 103.5 | 103 | 107 | 101.5 | 103.2 | 88 | 104.346 | 75.8 | 87.7 | 106 | 90.1 | 103.06 |  |  | 102.0064
| 1 | 5 | 105.08 | 98.6843 | 102.7 | 72.6 | 101.95 | 101 | 100.2987 | 80.76 | 35.8 | 70.7 | 103.0475 | 112.5 | 101.5 | 101 | 110 | 104 | 100 | 87 | 105.5615 | 78.7 | 72 | 105 | 89.8 | 103.36 |  |  | 38
| 1 | 6 | 104.22 | 96.1599 | 103.2 | 81 | 101.83 | 102 | 110.7765 | 95.16 | 37.5 | 72 | 102.256 | 111 | 101 | 102.5 | 111.5 | 110.5 | 102.7 | 86 | 101.8403 | 75.2 | 69.1 | 106 | 91.3 | 100.9
| Mean |  | 104.785 | 98.2804 | 104.2333 | 76.3167 | 102.05 | 101.1667 | 103.5995 | 88.2383 | 37.9833 | 76.6667 | 102.2646 | 110.6667 | 103.1667 | 101.8167 | 107.1667 | 105.0833 | 101.7833 | 89.1667 | 105.7566 | 79.9833 | 80.1833 | 106.6667 | 95.7 | 101.4217
| Median |  | 104.31 | 98.6106 | 102.95 | 76.45 | 101.89 | 101 | 101.4191 | 87.285 | 36.05 | 73.6 | 102.239 | 111.25 | 102.5 | 101.75 | 106.75 | 104.5 | 102.1 | 89 | 105.4765 | 78.6 | 82.8 | 106 | 96.15 | 100.97
| Std Dev. |  | 2.2105 | 1.9607 | 2.7998 | 3.5346 | 0.43437 | 0.75277 | 5.2768 | 8.1167 | 5.7192 | 10.6579 | 0.51417 | 2.1134 | 2.6013 | 1.5943 | 3.2197 | 2.9903 | 1.4386 | 2.6394 | 2.6991 | 6.5187 | 7.7801 | 1.5055 | 5.8275 | 1.4754
| Rel.Std.Dev. |  | 2.11% | 1.995% | 2.686% | 4.632% | 0.4256% | 0.7441% | 5.093% | 9.199% | 15.06% | 13.9% | 0.5028% | 1.91% | 2.521% | 1.566% | 3.004% | 2.846% | 1.413% | 2.96% | 2.552% | 8.15% | 9.703% | 1.411% | 6.089% | 1.455%
| PDM3 |  | 2.724% | -3.653% | 2.183% | -25.18% | 0.04277% | -0.8232% | 1.562% | -13.5% | -62.76% | -24.84% | 0.2531% | 8.49% | 1.137% | -0.186% | 5.059% | 3.016% | -0.2186% | -12.59% | 3.676% | -21.59% | -21.39% | 4.569% | -6.182% | -0.5732%
| Z-Score (Absolute) |  | 0.54 | 0.63 | 0.44 | 4.6 | 0.05 | 0.11 | 0.32 | 2.45 | 11.51 | 4.53 | 0.08 | 1.6 | 0.25 | 0 | 0.97 | 0.59 | 0 | 2.28 | 0.71 | 3.94 | 3.9 | 0.88 | 1.1 | 0.07
| Analytical results for Nd in OREAS 753b (Certified Value 3.94 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.20; 1.08: 1.22; 1.09: 1.24; 1.10: 1.25 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 4.35; 1.02: 4.26; 1.03: 3.5; 1.04: 4.3; 1.05: 4.2341; 1.06: 3.78; 1.07: 4.4412; 1.08: 2.8; 1.09: 5.2; 1.10: 4.4 |  |  | 1
| 1 | 2 | 1.01: 3.92; 1.02: 4.13; 1.03: 3.1; 1.04: 3.8; 1.05: 3.8131; 1.06: 3.86; 1.07: 4.4967; 1.08: 2.3; 1.09: 4.5; 1.10: 3.7 |  |  | 16
| 1 | 3 | 1.01: 4.28; 1.02: 4.15; 1.03: 4; 1.04: 3.6; 1.05: 3.8033; 1.06: 3.67; 1.07: 4.5616; 1.08: 2.5; 1.09: 4.2; 1.10: 4 |  |  | 16
| 1 | 4 | 1.01: 3.86; 1.02: 3.86; 1.03: 3.6; 1.04: 3.7; 1.05: 3.8448; 1.06: 4.23; 1.07: 4.4968; 1.08: 2.8; 1.09: 4.2; 1.10: 4 |  |  | 3.9442
| 1 | 5 | 1.01: 4.15; 1.02: 3.97; 1.03: 3.6; 1.04: 4.3; 1.05: 3.9869; 1.06: 3.7; 1.07: 4.4358; 1.08: 3.4; 1.09: 5; 1.10: 3.9 |  |  | 39
| 1 | 6 | 1.01: 4.19; 1.02: 3.84; 1.03: 3.5; 1.04: 4.1; 1.05: 3.9523; 1.06: 3.98; 1.07: 4.4791; 1.08: 2.8; 1.09: 4.2; 1.10: 4.2
| Mean |  | 1.01: 4.125; 1.02: 4.035; 1.03: 3.55; 1.04: 3.9667; 1.05: 3.9391; 1.06: 3.87; 1.07: 4.4852; 1.08: 2.7667; 1.09: 4.55; 1.10: 4.0333
| Median |  | 1.01: 4.17; 1.02: 4.05; 1.03: 3.55; 1.04: 3.95; 1.05: 3.8986; 1.06: 3.82; 1.07: 4.4879; 1.08: 2.8; 1.09: 4.35; 1.10: 4
| Std Dev. |  | 1.01: 0.19583; 1.02: 0.17073; 1.03: 0.2881; 1.04: 0.30768; 1.05: 0.16288; 1.06: 0.20919; 1.07: 0.045861; 1.08: 0.37238; 1.09: 0.44609; 1.10: 0.24221
| Rel.Std.Dev. |  | 1.01: 4.747%; 1.02: 4.231%; 1.03: 8.115%; 1.04: 7.757%; 1.05: 4.135%; 1.06: 5.405%; 1.07: 1.022%; 1.08: 13.46%; 1.09: 9.804%; 1.10: 6.005%
| PDM3 |  | 1.01: 4.585%; 1.02: 2.303%; 1.03: -9.993%; 1.04: 0.5708%; 1.05: -0.1285%; 1.06: -1.88%; 1.07: 13.72%; 1.08: -29.85%; 1.09: 15.36%; 1.10: 2.261%
| Z-Score (Absolute) |  | 1.01: 0.66; 1.02: 0.19; 1.03: 2.38; 1.04: 0.18; 1.05: 0.32; 1.06: 0.69; 1.07: 2.57; 1.08: 6.52; 1.09: 2.91; 1.10: 0.18
| Analytical results for Ni in OREAS 753b (Certified Value 8.32 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 9.1 | 7.665 | 8.6 | 10 | 8.2 | 8 | 8 | 8 | 10 | 9.5 | 7.9461 | 7.7 | 8.5 | 7.9 | 8.1 | 8.2 | 8.7 | 8.3 | 8.7 | 8.2 | 10 | 8.3 | 9 | 8.34 |  |  | 1
| 1 | 2 | 8.1 | 7.7201 | 8.6 | 10 | 8.4 | 8 | 8 | 7.9 | 10 | 9.4 | 8.4477 | 8.1 | 8.7 | 8.1 | 7.8 | 8.3 | 8.6 | 7.9 | 8.5 | 7.9 | 10 | 8.6 | 9.2 | 8.26 |  |  | 6
| 1 | 3 | 8 | 7.7012 | 8.7 | 10 | 8.3 | 8 | 8 | 8 | 10 | 9.9 | 8.5598 | 7.6 | 8.2 | 8 | 7.8 | 8.2 | 8.4 | 8.5 | 8.2 | 7.8 | 10 | 8.3 | 8.8 | 8.47 |  |  | 16
| 1 | 4 | 8.5 | 7.642 | 8.7 | 9 | 8.4 | 8 | 8 | 8.7 | 9 | 8.9 | 8.1049 | 7.6 | 8.3 | 7.9 | 7.9 | 8.8 | 8 | 8 | 8.5 | 8.3 | 9 | 8.3 | 9 | 9.34 |  |  | 8.3177
| 1 | 5 | 8.5 | 7.716 | 8.8 | 10 | 8.4 | 8 | 8 | 7.8 | 10 | 8.9 | 8.0078 | 8.1 | 8.6 | 7.7 | 8 | 8.5 | 8.1 | 8.5 | 8.4 | 8.1 | 10 | 7.9 | 9.8 | 8.63 |  |  | 40
| 1 | 6 | 8.5 | 7.4026 | 8.7 | 10 | 8.6 | 8 | 8 | 7.8 | 11 | 9.5 | 8.3754 | 7.7 | 8.4 | 7.7 | 8.1 | 10.1 | 8.5 | 7.9 | 8.7 | 8.2 | 11 | 7.7 | 8.7 | 9.12
| Mean |  | 8.45 | 7.6412 | 8.6833 | 9.8333 | 8.3833 | 8 | 8 | 8.0333 | 10 | 9.35 | 8.2403 | 7.8 | 8.45 | 7.8833 | 7.95 | 8.6833 | 8.3833 | 8.1833 | 8.5 | 8.0833 | 10 | 8.1833 | 9.0833 | 8.6933
| Median |  | 8.5 | 7.6831 | 8.7 | 10 | 8.4 | 8 | 8 | 7.95 | 10 | 9.45 | 8.2402 | 7.7 | 8.45 | 7.9 | 7.95 | 8.4 | 8.45 | 8.15 | 8.5 | 8.15 | 10 | 8.3 | 9 | 8.55
| Std Dev. |  | 0.38859 | 0.12076 | 0.075277 | 0.40825 | 0.13292 | 0 | 0 | 0.33862 | 0.63246 | 0.38859 | 0.25388 | 0.23664 | 0.18708 | 0.16021 | 0.13784 | 0.73052 | 0.27869 | 0.28577 | 0.18974 | 0.19408 | 0.63246 | 0.32506 | 0.392 | 0.43971
| Rel.Std.Dev. |  | 4.599% | 1.58% | 0.8669% | 4.152% | 1.585% | 0% | 0% | 4.215% | 6.325% | 4.156% | 3.081% | 3.034% | 2.214% | 2.032% | 1.734% | 8.413% | 3.324% | 3.492% | 2.232% | 2.401% | 6.325% | 3.972% | 4.316% | 5.058%
| PDM3 |  | 1.591% | -8.133% | 4.396% | 18.22% | 0.7894% | -3.819% | -3.819% | -3.419% | 20.23% | 12.41% | -0.9304% | -6.224% | 1.591% | -5.222% | -4.42% | 4.396% | 0.7894% | -1.615% | 2.192% | -2.817% | 20.23% | -1.615% | 9.205% | 4.516%
| Z-Score (Absolute) |  | 0.15 | 1.67 | 0.67 | NA | 0 | NA | NA | 0.79 | NA | 2.17 | 0.32 | 1.31 | 0.15 | 1.12 | 0.97 | 0.67 | 0 | 0.45 | 0.26 | 0.67 | 3.63 | 0.45 | 1.57 | 0.7
| Analytical results for P in OREAS 753b (Certified Value 0.049 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.0511 | 0.049489 | 0.0539 | 0.049 | 0.0467 | 0.05 | 0.046 | 0.052 | 0.049 | 0.0467 | 0.0532 | 0.05 | 0.048 | 0.046 | 0.048 | 0.051 | 0.0491 | 0.0512 | 0.044326 | 0.0496 | 0.054 | 0.04 | 0.0511 | 0.047 | 0.12 |  | 1
| 1 | 2 | 0.0516 | 0.050042 | 0.053 | 0.049 | 0.0457 | 0.049 | 0.047 | 0.05 | 0.049 | 0.047 | 0.0529 | 0.05 | 0.048 | 0.045 | 0.048 | 0.051 | 0.0505 | 0.0514 | 0.044235 | 0.0483 | 0.054 | 0.04 | 0.0515 | 0.043 | 0.12 |  | 0
| 1 | 3 | 0.0523 | 0.049654 | 0.0531 | 0.048 | 0.0469 | 0.049 | 0.046 | 0.053 | 0.043 | 0.0473 | 0.0552 | 0.049 | 0.049 | 0.046 | 0.048 | 0.05 | 0.0506 | 0.0514 | 0.044257 | 0.0483 | 0.054 | 0.04 | 0.0502 | 0.045 | 0.13 |  | 16
| 1 | 4 | 0.0514 | 0.048322 | 0.0517 | 0.048 | 0.047 | 0.05 | 0.046 | 0.053 | 0.047 | 0.0462 | 0.0523 | 0.049 | 0.046 | 0.047 | 0.048 | 0.05 | 0.0489 | 0.0502 | 0.044855 | 0.0492 | 0.052 | 0.04 | 0.0506 | 0.046 | 0.14 |  | 0.049133
| 1 | 5 | 0.0515 | 0.049549 | 0.0518 | 0.048 | 0.0483 | 0.049 | 0.045 | 0.053 | 0.048 | 0.0452 | 0.0527 | 0.049 | 0.046 | 0.046 | 0.047 | 0.049 | 0.0506 | 0.0519 | 0.044687 | 0.0487 | 0.052 | 0.04 | 0.0489 | 0.045 | 0.14 |  | 41
| 1 | 6 | 0.0524 | 0.047788 | 0.0522 | 0.049 | 0.0488 | 0.05 | 0.047 | 0.053 | 0.049 | 0.0478 | 0.0528 | 0.05 | 0.047 | 0.047 | 0.05 | 0.05 | 0.0487 | 0.0518 | 0.044335 | 0.049 | 0.054 | 0.04 | 0.0497 | 0.045 | 0.12
| Mean |  | 0.051717 | 0.04914 | 0.052617 | 0.0485 | 0.047233 | 0.0495 | 0.046167 | 0.052333 | 0.0475 | 0.0467 | 0.053183 | 0.0495 | 0.047333 | 0.046167 | 0.048167 | 0.050167 | 0.049733 | 0.051317 | 0.044449 | 0.04885 | 0.053333 | 0.04 | 0.050333 | 0.045167 | 0.12833
| Median |  | 0.05155 | 0.049519 | 0.0526 | 0.0485 | 0.04695 | 0.0495 | 0.046 | 0.053 | 0.0485 | 0.04685 | 0.05285 | 0.0495 | 0.0475 | 0.046 | 0.048 | 0.05 | 0.0498 | 0.0514 | 0.044331 | 0.04885 | 0.054 | 0.04 | 0.0504 | 0.045 | 0.125
| Std Dev. |  | 0.00051929 | 0.00087911 | 0.0008612 | 0.00054772 | 0.0011308 | 0.00054772 | 0.00075277 | 0.0012111 | 0.0023452 | 0.00091214 | 0.0010304 | 0.00054772 | 0.0012111 | 0.00075277 | 0.00098319 | 0.00075277 | 0.00092232 | 0.000608 | 0.00025778 | 0.00051672 | 0.0010328 | 0 | 0.00094798 | 0.0013292 | 0.0098319
| Rel.Std.Dev. |  | 1.004% | 1.789% | 1.637% | 1.129% | 2.394% | 1.107% | 1.631% | 2.314% | 4.937% | 1.953% | 1.937% | 1.107% | 2.559% | 1.631% | 2.041% | 1.501% | 1.855% | 1.185% | 0.58% | 1.058% | 1.936% | 0% | 1.883% | 2.943% | 7.661%
| PDM3 |  | 5.258% | 0.01463% | 7.09% | -1.289% | -3.867% | 0.7464% | -6.038% | 6.513% | -3.324% | -4.952% | 8.243% | 0.7464% | -3.663% | -6.038% | -1.967% | 2.103% | 1.221% | 4.444% | -9.533% | -0.5765% | 8.548% | -18.59% | 2.442% | -8.073% | 161.2%
| Z-Score (Absolute) |  | 0.8 | 0 | 1.08 | 0.2 | 0.59 | 0.11 | 0.92 | 0.99 | 0.51 | 0.76 | 1.25 | 0.11 | 0.56 | 0.92 | 0.3 | 0.32 | 0.18 | 0.67 | 1.45 | 0.09 | 1.3 | 2.83 | 0.37 | 1.23 | 24.54
| Analytical results for Pb in OREAS 753b (Certified Value 4.64 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 4.6 | 4.4822 | 4.9 | 4.9 | 4.9 | 5 | 4 | 3.4 | < 3 | 4.4 | 4.5493 | 4.4 | 4.9 | 4.3 | 4.4 | 4.4 | 4.6 | 3.9 | 5.4 | 4.5 | 4.3 | 4.9 | 4.88 | 10.9 |  |  | 1
| 1 | 2 | 4.9 | 4.8496 | 4.8 | 4.9 | 4.8 | 5 | 4 | 3.33 | < 3 | 4.6 | 4.4401 | 4.7 | 4.6 | 4.4 | 4.3 | 4.3 | 5 | 3.4 | 4.8 | 4.7 | 3.5 | 4.8 | 4.97 | 11.8 |  |  | 7
| 1 | 3 | 4.9 | 4.3961 | 4.7 | 4.8 | 4.9 | 5 | 4 | 3.41 | < 3 | 4.3 | 4.4334 | 4.4 | 4.5 | 4.4 | 4.3 | 4.4 | 4.8 | 3.7 | 5.1 | 4.5 | 3.6 | 4.8 | 5.03 | 12.8 |  |  | 16
| 1 | 4 | 4.9 | 4.4796 | 4.9 | 4.6 | 4.7 | 5 | 4 | 3.58 | < 3 | 4.4 | 4.4656 | 4.5 | 4.5 | 4.1 | 4.4 | 4.5 | 4.8 | 3.4 | 5.1 | 4.7 | 3.6 | 4.7 | 4.81 | 12.7 |  |  | 4.6367
| 1 | 5 | 4.9 | 4.5597 | 4.8 | 4.6 | 4.4 | 5 | 4 | 3.6 | < 3 | 4.4 | 4.4185 | 4.3 | 4.7 | 4.2 | 4.3 | 4.3 | 4.4 | 3.9 | 5.2 | 4.6 | 3.5 | 4.4 | 4.75 | 17 |  |  | 42
| 1 | 6 | 5 | 4.6386 | 4.8 | 4.4 | 4.6 | 4 | 4 | 3.24 | < 3 | 4.5 | 4.5998 | 4.5 | 4.7 | 4.2 | 4.3 | 4.6 | 4.8 | 3.6 | 5.1 | 4.7 | 3.5 | 4.5 | 4.97 | 11
| Mean |  | 4.8667 | 4.5676 | 4.8167 | 4.7 | 4.7167 | 4.8333 | 4 | 3.4267 |  | 4.4333 | 4.4844 | 4.4667 | 4.65 | 4.2667 | 4.3333 | 4.4167 | 4.7333 | 3.65 | 5.1167 | 4.6167 | 3.6667 | 4.6833 | 4.9017 | 12.7
| Median |  | 4.9 | 4.521 | 4.8 | 4.7 | 4.75 | 5 | 4 | 3.405 |  | 4.4 | 4.4528 | 4.45 | 4.65 | 4.25 | 4.3 | 4.4 | 4.8 | 3.65 | 5.1 | 4.65 | 3.55 | 4.75 | 4.925 | 12.25
| Std Dev. |  | 0.13663 | 0.16068 | 0.075277 | 0.2 | 0.19408 | 0.40825 | 0 | 0.14052 |  | 0.10328 | 0.073183 | 0.13663 | 0.15166 | 0.12111 | 0.05164 | 0.1169 | 0.20656 | 0.22583 | 0.19408 | 0.098319 | 0.31411 | 0.19408 | 0.10741 | 2.2557
| Rel.Std.Dev. |  | 2.807% | 3.518% | 1.563% | 4.255% | 4.115% | 8.447% | 0% | 4.101% |  | 2.33% | 1.632% | 3.059% | 3.261% | 2.838% | 1.192% | 2.647% | 4.364% | 6.187% | 3.793% | 2.13% | 8.567% | 4.144% | 2.191% | 17.76%
| PDM3 |  | 4.96% | -1.49% | 3.882% | 1.365% | 1.725% | 4.241% | -13.73% | -26.1% |  | -4.386% | -3.284% | -3.667% | 0.2871% | -7.98% | -6.543% | -4.745% | 2.084% | -21.28% | 10.35% | -0.4318% | -20.92% | 1.006% | 5.715% | 173.9%
| Z-Score (Absolute) |  | 0.93 | 0.08 | 0.76 | 0.36 | 0.42 | NA | NA | 3.93 | 10.43 | 0.54 | 0.36 | 0.42 | 0.2 | 1.1 | 0.87 | 0.59 | 0.48 | 3.18 | 1.77 | 0.08 | 3.12 | 0.31 | 1.04 | 27.34
Results from laboratories 6 and 7 were removed due to their 1 ppm reading resolution.
| Analytical results for Pr in OREAS 753b (Certified Value 1.07 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.20; 1.08: 1.22; 1.09: 1.24; 1.10: 1.25 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 1.14; 1.02: 1.12; 1.03: 0.8; 1.04: 1.2; 1.05: 1.1024; 1.06: 0.97; 1.07: 1.1851; 1.08: 0.7; 1.09: 1.4; 1.10: 1.25 |  |  | 1
| 1 | 2 | 1.01: 1.03; 1.02: 1.1; 1.03: 0.7; 1.04: 1; 1.05: 1.025; 1.06: 0.96; 1.07: 1.1523; 1.08: 0.56; 1.09: 1.3; 1.10: 1.05 |  |  | 19
| 1 | 3 | 1.01: 1.07; 1.02: 1.08; 1.03: 1; 1.04: 1; 1.05: 1.0204; 1.06: 0.94; 1.07: 1.1832; 1.08: 0.59; 1.09: 1.2; 1.10: 1.14 |  |  | 16
| 1 | 4 | 1.01: 1.05; 1.02: 0.99; 1.03: 0.9; 1.04: 1; 1.05: 0.9922; 1.06: 1.06; 1.07: 1.1524; 1.08: 0.68; 1.09: 1.1; 1.10: 1.12 |  |  | 1.0748
| 1 | 5 | 1.01: 1.1; 1.02: 1.03; 1.03: 0.9; 1.04: 1.1; 1.05: 1.0434; 1.06: 0.94; 1.07: 1.0977; 1.08: 0.84; 1.09: 1.3; 1.10: 1.12 |  |  | 43
| 1 | 6 | 1.01: 1.13; 1.02: 1; 1.03: 0.9; 1.04: 1.1; 1.05: 1.0691; 1.06: 1.03; 1.07: 1.1014; 1.08: 0.7; 1.09: 1.2; 1.10: 1.15
| Mean |  | 1.01: 1.0867; 1.02: 1.0533; 1.03: 0.86667; 1.04: 1.0667; 1.05: 1.0421; 1.06: 0.98333; 1.07: 1.1454; 1.08: 0.67833; 1.09: 1.25; 1.10: 1.1383
| Median |  | 1.01: 1.085; 1.02: 1.055; 1.03: 0.9; 1.04: 1.05; 1.05: 1.0342; 1.06: 0.965; 1.07: 1.1523; 1.08: 0.69; 1.09: 1.25; 1.10: 1.13
| Std Dev. |  | 1.01: 0.044121; 1.02: 0.054283; 1.03: 0.10328; 1.04: 0.08165; 1.05: 0.039038; 1.06: 0.050067; 1.07: 0.038243; 1.08: 0.098877; 1.09: 0.10488; 1.10: 0.064936
| Rel.Std.Dev. |  | 1.01: 4.06%; 1.02: 5.153%; 1.03: 11.92%; 1.04: 7.655%; 1.05: 3.746%; 1.06: 5.092%; 1.07: 3.339%; 1.08: 14.58%; 1.09: 8.39%; 1.10: 5.704%
| PDM3 |  | 1.01: 1.099%; 1.02: -2.002%; 1.03: -19.37%; 1.04: -0.7613%; 1.05: -3.049%; 1.06: -8.514%; 1.07: 6.56%; 1.08: -36.89%; 1.09: 16.3%; 1.10: 5.906%
| Z-Score (Absolute) |  | 1.01: 0.32; 1.02: 0; 1.03: NA; 1.04: NA; 1.05: 0.11; 1.06: 0.67; 1.07: 0.89; 1.08: 3.61; 1.09: NA; 1.10: 0.82
Results from laboratories 8, 9 and 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Rb in OREAS 753b (Certified Value 1215 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 1334.21 | 1123.6467 | 1212.1 | 1147 | 1200.89 | 1410 | 1240.8 | 1167.8 | 1200 | 1330 | 1175.997 | 1210 | 1225 | 1265 | 1165 | 1150 | 1275 | 1103 | 1456.63 | 1220.8 | 1561 | 1210.9 | 1300 | 1378 | 1399.15 |  | 1
| 1 | 2 | 1293.41 | 1158.3454 | 1186.9 | 1175 | 1198.49 | 1420 | 1256.6 | 1092.1 | 1210 | 1330 | 1204.17 | 1210 | 1210 | 1315 | 1155 | 1140 | 1255 | 1089 | 1451.14 | 1213.3 | 1604 | 1175.9 | 1310 | 1368 | 1397.44 |  | 0
| 1 | 3 | 1286.83 | 1125.341 | 1200.7 | 1146 | 1204.45 | 1410 | 1249.6 | 1244.9 | 1210 | 1310 | 1201.343 | 1200 | 1220 | 1175 | 1115 | 1185 | 1240 | 1067 | 1456.92 | 1213 | 1567 | 1196.5 | 1310 | 1346 | 1400.02 |  | 16
| 1 | 4 | 1278.76 | 1110.1665 | 1239.3 | 1183 | 1212.02 | 1420 | 1252.1 | 1215.8 | 1200 | 1330 | 1231.763 | 1190 | 1240 | 1065 | 1180 | 1195 | 1195 | 1105 | 1459.06 | 1239.8 | 1496 | 1202.9 | 1280 | 1382 | 1400.42 |  | 1215.2833
| 1 | 5 | 1351.56 | 1142.3964 | 1212.4 | 1180 | 1213.93 | 1420 | 1239.5 | 1245.4 | 1320 | 1320 | 1205.906 | 1200 | 1215 | 1095 | 1150 | 1185 | 1130 | 1072 | 1456.73 | 1225.5 | 1411 | 1060.7 | 1300 | 1386 | 1401.62 |  | 44
| 1 | 6 | 1361.87 | 1095.8451 | 1245.3 | 1260 | 1209.47 | 1470 | 1245.4 | 1133.5 | 1250 | 1380 | 1202.907 | 1200 | 1220 | 1110 | 1125 | 1225 | 1195 | 1114 | 1459.25 | 1216.1 | 1499 | 1167.2 | 1280 | 1409 | 1396.84
| Mean |  | 1317.7733 | 1125.9569 | 1216.1167 | 1181.8333 | 1206.5417 | 1425 | 1247.3333 | 1183.25 | 1231.6667 | 1333.3333 | 1203.681 | 1201.6667 | 1221.6667 | 1170.8333 | 1148.3333 | 1180 | 1215 | 1091.6667 | 1456.6217 | 1221.4167 | 1523 | 1169.0167 | 1296.6667 | 1378.1667 | 1399.2483
| Median |  | 1313.81 | 1124.4938 | 1212.25 | 1177.5 | 1206.96 | 1420 | 1247.5 | 1191.8 | 1210 | 1330 | 1203.5385 | 1200 | 1220 | 1142.5 | 1152.5 | 1185 | 1217.5 | 1096 | 1456.825 | 1218.45 | 1530 | 1186.2 | 1300 | 1380 | 1399.585
| Std Dev. |  | 35.8589 | 22.2757 | 22.407 | 41.6145 | 6.2341 | 22.5832 | 6.656 | 62.841 | 47.0815 | 24.2212 | 17.6989 | 7.5277 | 10.328 | 100.3203 | 24.4268 | 30.9839 | 52.6308 | 19.0123 | 2.9327 | 10.1979 | 68.9783 | 55.5712 | 13.6626 | 20.8079 | 1.8261
| Rel.Std.Dev. |  | 2.721% | 1.978% | 1.843% | 3.521% | 0.5167% | 1.585% | 0.5336% | 5.311% | 3.823% | 1.817% | 1.47% | 0.6264% | 0.8454% | 8.568% | 2.127% | 2.626% | 4.332% | 1.742% | 0.2013% | 0.8349% | 4.529% | 4.754% | 1.054% | 1.51% | 0.1305%
| PDM3 |  | 8.433% | -7.35% | 0.06857% | -2.752% | -0.7193% | 17.26% | 2.637% | -2.636% | 1.348% | 9.714% | -0.9547% | -1.12% | 0.5253% | -3.658% | -5.509% | -2.903% | -0.02331% | -10.17% | 19.86% | 0.5047% | 25.32% | -3.807% | 6.697% | 13.4% | 15.14%
| Z-Score (Absolute) |  | 1.51 | 1.34 | 0 | 0.51 | 0.14 | 3.11 | 0.46 | 0.49 | 0.23 | 1.75 | 0.19 | 0.22 | 0.08 | 0.67 | 1.01 | 0.54 | 0.02 | 1.85 | 3.58 | 0.08 | 4.57 | 0.7 | 1.2 | 2.41 | 2.73
| Analytical results for Re in OREAS 753b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.22 | 1.25 | [11 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [11 empty] | 3
3
| 1 | 1 | < 0.01 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | < 0.001 | < 0.002 |  | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.1 | 0.064 | [11 empty] | 1
| 1 | 2 | < 0.01 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | < 0.001 | 0.002 | 0.0006 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.1 | 0.065 | [11 empty] | 20
| 1 | 3 | < 0.01 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.001 | < 0.002 | 0.0005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.1 | 0.043 | [11 empty] | 16
| 1 | 4 | < 0.01 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.001 | < 0.002 | 0.0006 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.1 | 0.041 | [11 empty] | < 0.002
| 1 | 5 | < 0.01 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | 0.002 | < 0.002 | 0.0008 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.1 | 0.041 | [11 empty] | 45
| 1 | 6 | < 0.01 | < 0.002 | < 0.05 | < 0.05 | < 0.002 | < 0.001 | < 0.002 | 0.0007 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.1 | 0.033
| Mean |  | 1.09: 0.0013333; 1.10: 0.002; 1.12: 0.00064; 1.16: 0.002; 1.17: 0.002; 1.25: 0.047833
| Median |  | 1.09: 0.001; 1.10: 0.002; 1.12: 0.0006; 1.16: 0.002; 1.17: 0.002; 1.25: 0.042
| Std Dev. |  | 1.09: 0.00057735; 1.12: 0.00011402; 1.17: 0; 1.25: 0.013363
| Rel.Std.Dev. |  | 1.09: 43.3%; 1.12: 17.82%; 1.17: 0%; 1.25: 27.94%
PDM3
| Z-Score (Absolute) |  | 15.51 | 0.67 | 96.43 | 96.43 | 0.67 | 1.01 | 0 | 2.13 | 0.67 | 0.67 | 0.67 | 0 | 0.67 | 197.57 | 188.81
| Analytical results for S in OREAS 753b (Certified Value 0.025 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.0231 | 0.02601 | 0.03 | 0.027 |  | 0.019 | 0.024 | < 0.04 | 0.03 | 0.03 | 0.0228 | 0.03 | 0.02 | 0.02 | 0.03 | 0.04 | 0.02 | 0.03 | 0.02448 | 0.025 | 0.025 | < 0.05 | 0.0279 | 0.028 | 0.02 |  | 1
| 1 | 2 | 0.0229 | 0.030314 | 0.03 | 0.026 |  | 0.02 | 0.023 | < 0.04 | 0.02 | 0.02 | 0.022 | 0.03 | 0.03 | 0.02 | 0.03 | 0.04 | 0.03 | 0.03 | 0.024437 | 0.026 | 0.025 | < 0.05 | 0.0289 | 0.026 | 0.02 |  | 8
| 1 | 3 | 0.0237 | 0.026857 | 0.03 | 0.027 |  | 0.019 | 0.023 | < 0.04 | 0.02 | 0.02 | 0.0184 | 0.03 | 0.03 | 0.02 | 0.03 | 0.04 | 0.02 | 0.03 | 0.024127 | 0.025 | 0.025 | < 0.05 | 0.027 | 0.027 | 0.02 |  | 16
| 1 | 4 | 0.0241 | 0.024416 | 0.03 | 0.027 |  | 0.02 | 0.024 | < 0.04 | 0.02 | 0.02 | 0.0202 | 0.03 | 0.02 | 0.03 | 0.03 | 0.04 | 0.02 | 0.03 | 0.024757 | 0.026 | 0.025 | < 0.05 | 0.0276 | 0.03 | 0.02 |  | 0.025183
| 1 | 5 | 0.0237 | 0.028214 | 0.03 | 0.027 | 0.06 | 0.019 | 0.023 | < 0.04 | 0.03 | 0.02 | 0.0186 | 0.02 | 0.03 | 0.02 | 0.03 | 0.04 | 0.03 | 0.03 | 0.024948 | 0.025 | 0.019 | < 0.05 | 0.0267 | 0.029 | 0.02 |  | 46
| 1 | 6 | 0.0234 | 0.029621 | 0.03 | 0.026 |  | 0.019 | 0.023 | < 0.04 | 0.02 | 0.03 | 0.0229 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.02 | 0.03 | 0.024283 | 0.026 | 0.016 | < 0.05 | 0.0249 | 0.03 | 0.01
| Mean |  | 0.023483 | 0.027572 | 0.03 | 0.026667 | 0.06 | 0.019333 | 0.023333 |  | 0.023333 | 0.023333 | 0.020817 | 0.028333 | 0.026667 | 0.023333 | 0.03 | 0.04 | 0.023333 | 0.03 | 0.024505 | 0.0255 | 0.0225 |  | 0.027167 | 0.028333 | 0.018333
| Median |  | 0.02355 | 0.027535 | 0.03 | 0.027 | 0.06 | 0.019 | 0.023 |  | 0.02 | 0.02 | 0.0211 | 0.03 | 0.03 | 0.02 | 0.03 | 0.04 | 0.02 | 0.03 | 0.024458 | 0.0255 | 0.025 |  | 0.0273 | 0.0285 | 0.02
| Std Dev. |  | 0.00044008 | 0.0022379 | 0 | 0.0005164 |  | 0.0005164 | 0.0005164 |  | 0.005164 | 0.005164 | 0.00204 | 0.0040825 | 0.005164 | 0.005164 | 0 | 0 | 0.005164 | 0 | 0.00030233 | 0.00054772 | 0.0039875 |  | 0.0013501 | 0.001633 | 0.0040825
| Rel.Std.Dev. |  | 1.874% | 8.117% | 0% | 1.936% |  | 2.671% | 2.213% |  | 22.13% | 22.13% | 9.8% | 14.41% | 19.36% | 22.13% | 0% | 0% | 22.13% | 0% | 1.234% | 2.148% | 17.72% |  | 4.97% | 5.764% | 22.27%
| PDM3 |  | -6.749% | 9.487% | 19.13% | 5.891% | 138.3% | -23.23% | -7.345% |  | -7.345% | -7.345% | -17.34% | 12.51% | 5.891% | -7.345% | 19.13% | 58.84% | -7.345% | 19.13% | -2.691% | 1.259% | -10.65% |  | 7.877% | 12.51% | -27.2%
| Z-Score (Absolute) |  | 0.41 | 0.69 | 1.35 | 0.45 | 9.44 | 1.53 | 0.45 | 1.35 | 0.45 | 0.45 | 1.13 | 0.9 | 0.45 | 0.45 | 1.35 | 4.05 | 0.45 | 1.35 | 0.13 | 0.13 | 0.67 | 0 | 0.58 | 0.9 | 1.8
| Analytical results for Sb in OREAS 753b (Certified Value 0.64 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 0.62 | 0.59862 | 0.7 | 0.61 | 0.6 | 0.6 | 0.7 | 0.62 | 0.6 | 0.63 | 0.62901 | 0.7 | 0.75 | 0.63 | 0.63 | 0.71 | 0.71 | 0.66 | 5.28 | 0.64 | 0.6 | 0.91 | < 1 |  |  |  | 1
| 1 | 2 | 0.62 | 0.60929 | 0.6 | 0.63 | 0.61 | 0.6 | 0.7 | 0.6 | 0.5 | 0.67 | 0.61854 | 0.6 | 0.69 | 0.67 | 0.62 | 0.68 | 0.73 | 0.63 | 5.45 | 0.53 | 0.6 | 0.9 | < 1 |  |  |  | 21
| 1 | 3 | 0.66 | 0.60635 | 0.6 | 0.57 | 0.61 | 0.7 | 0.7 | 0.7 | 0.4 | 0.64 | 0.63695 | 0.7 | 0.68 | 0.61 | 0.6 | 0.68 | 0.74 | 0.67 | 5.31 | 0.54 | 0.6 | 0.8 | < 1 |  |  |  | 16
| 1 | 4 | 0.64 | 0.58195 | 0.7 | 0.61 | 0.63 | 0.6 | 0.7 | 0.69 | 0.4 | 0.66 | 0.61198 | 0.7 | 0.65 | 0.58 | 0.62 | 0.72 | 0.67 | 0.65 | 5.27 | 0.5 | 0.6 | 0.83 | < 1 |  |  |  | 0.64158
| 1 | 5 | 0.66 | 0.59214 | 0.6 | 0.62 | 0.6 | 0.7 | 0.7 | 0.63 | 0.4 | 0.61 | 0.60464 | 0.6 | 0.7 | 0.6 | 0.67 | 0.69 | 0.72 | 0.63 | 5.6 | 0.53 | 0.5 | 0.8 | < 1 |  |  |  | 47
| 1 | 6 | 0.64 | 0.60604 | 0.7 | 0.64 | 0.61 | 0.6 | 0.7 | 0.6 | 0.5 | 0.65 | 0.64098 | 0.6 | 0.69 | 0.59 | 0.64 | 0.71 | 0.73 | 0.67 | 5.31 | 0.53 | 0.6 | 0.74 | < 1
| Mean |  | 0.64 | 0.59907 | 0.65 | 0.61333 | 0.61 | 0.63333 | 0.7 | 0.64 | 0.46667 | 0.64333 | 0.62368 | 0.65 | 0.69333 | 0.61333 | 0.63 | 0.69833 | 0.71667 | 0.65167 | 5.37 | 0.545 | 0.58333 | 0.83
| Median |  | 0.64 | 0.60233 | 0.65 | 0.615 | 0.61 | 0.6 | 0.7 | 0.625 | 0.45 | 0.645 | 0.62377 | 0.65 | 0.69 | 0.605 | 0.625 | 0.7 | 0.725 | 0.655 | 5.31 | 0.53 | 0.6 | 0.815
| Std Dev. |  | 0.017889 | 0.010466 | 0.054772 | 0.024221 | 0.010954 | 0.05164 | 1.2162e-16 | 0.044272 | 0.08165 | 0.021602 | 0.01435 | 0.054772 | 0.03266 | 0.03266 | 0.023664 | 0.017224 | 0.025033 | 0.018348 | 0.13008 | 0.048477 | 0.040825 | 0.065115
| Rel.Std.Dev. |  | 2.795% | 1.747% | 8.427% | 3.949% | 1.796% | 8.154% | 1.737e-14% | 6.917% | 17.5% | 3.358% | 2.301% | 8.427% | 4.711% | 5.325% | 3.756% | 2.466% | 3.493% | 2.816% | 2.422% | 8.895% | 6.999% | 7.845%
| PDM3 |  | -0.2458% | -6.626% | 1.313% | -4.402% | -4.922% | -1.285% | 9.106% | -0.2458% | -27.26% | 0.2737% | -2.789% | 1.313% | 8.067% | -4.402% | -1.805% | 8.846% | 11.7% | 1.573% | 737% | -15.05% | -9.078% | 29.37%
| Z-Score (Absolute) |  | 0.11 | 0.8 | NA | 0.48 | 0.55 | NA | NA | 0.11 | NA | 0.18 | 0.25 | NA | 1.29 | 0.48 | 0.11 | 1.4 | 1.81 | 0.37 | 104.8 | 1.99 | 1.14 | 4.32 | 2.99
Results from laboratories 3, 6, 7, 9 and 12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 753b (Certified Value 1.03 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
2
| 1 | 1 | 0.9 | 1.0103 | 1.1 | 1.1 | 1.1 | 1 | 1 | 0.9 | 1 | 1.3 | 1.0749 | 2 | 1 | 1.1 | 1.1 | 1 | 1.3 | 0.8 | 0.9 | 0.84398 | 1.2 | 1.1 | < 1 | 1.2 | 2.5 | 0.93 | 1
| 1 | 2 | 1.1 | 1 | 1.1 | 1.1 | 1 | < 1 | 1 | 0.8 | 1 | 1.2 | 1.0659 | 2 | 1 | 1.1 | 1 | 1 | 1.3 | 0.7 | 0.9 | 0.85281 | 1.1 | 1.5 | < 1 | 1.2 | 2.3 | 0.82 | 22
| 1 | 3 | 1.1 | 0.9779 | 1.1 | 1.1 | 1 | < 1 | 1 | 1 | 1 | 1.2 | 1.0564 | 2 | 1 | 1 | 1 | 1 | 1.2 | 0.8 | 0.9 | 0.89117 | 1 | 1.4 | < 1 | 1.2 | 1.9 | 0.85 | 16
| 1 | 4 | 1 | 0.97535 | 1.1 | 1.1 | 1.1 | 1 | 1 | 0.9 | 1 | 1.2 | 1.0425 | 2 | 1 | 0.8 | 1 | 1 | 1.1 | 0.7 | 0.9 | 0.85066 | 1.1 | 1.1 | < 1 | 1.2 | 1.6 | 1.11 | 1.0272
| 1 | 5 | 1.1 | 0.9627 | 1 | 1.1 | 1 | 1 | 1 | 1 | 1 | 1.1 | 1.126 | 2 | 1 | 0.8 | 1 | 1 | 1.2 | 0.7 | 0.9 | 0.88616 | 1.1 | 1 | < 1 | 1.2 | 1.5 | 1.18 | 48
| 1 | 6 | 1.2 | 0.94006 | 1.1 | 1.2 | 1 | 1 | 1 | 0.8 | 1 | 1.3 | 1.1051 | 2 | 1 | 0.9 | 1 | 1.1 | 1.2 | 0.8 | 0.9 | 0.80953 | 1.1 | 1 | < 1 | 1.2 | 1.3 | 0.95
| Mean |  | 1.0667 | 0.97772 | 1.0833 | 1.1167 | 1.0333 | 1 | 1 | 0.9 | 1 | 1.2167 | 1.0785 | 2 | 1 | 0.95 | 1.0167 | 1.0167 | 1.2167 | 0.75 | 0.9 | 0.85572 | 1.1 | 1.1833 |  | 1.2 | 1.85 | 0.97333
| Median |  | 1.1 | 0.97662 | 1.1 | 1.1 | 1 | 1 | 1 | 0.9 | 1 | 1.2 | 1.0704 | 2 | 1 | 0.95 | 1 | 1 | 1.2 | 0.75 | 0.9 | 0.85174 | 1.1 | 1.1 |  | 1.2 | 1.75 | 0.94
| Std Dev. |  | 0.10328 | 0.025325 | 0.040825 | 0.040825 | 0.05164 | 0 | 0 | 0.089443 | 0 | 0.075277 | 0.031375 | 0 | 0 | 0.13784 | 0.040825 | 0.040825 | 0.075277 | 0.054772 | 0 | 0.029962 | 0.063246 | 0.2137 |  | 0 | 0.47223 | 0.1432
| Rel.Std.Dev. |  | 9.682% | 2.59% | 3.768% | 3.656% | 4.997% | 0% | 0% | 9.938% | 0% | 6.187% | 2.909% | 0% | 0% | 14.51% | 4.016% | 4.016% | 6.187% | 7.303% | 0% | 3.501% | 5.75% | 18.06% |  | 0% | 25.53% | 14.71%
| PDM3 |  | 3.839% | -4.82% | 5.461% | 8.706% | 0.5937% | -2.651% | -2.651% | -12.39% | -2.651% | 18.44% | 4.99% | 94.7% | -2.651% | -7.519% | -1.029% | -1.029% | 18.44% | -26.99% | -12.39% | -16.7% | 7.084% | 15.2% |  | 16.82% | 80.1% | -5.247%
| Z-Score (Absolute) |  | 0.34 | 0.26 | 0.45 | 0.67 | 0.11 | 1.24 | 0.11 | 0.79 | NA | 1.35 | 0.42 | NA | NA | 0.45 | 0 | 0 | 1.35 | 1.8 | 0.79 | 1.09 | 0.56 | 1.12 | 3.48 | 1.24 | 5.62 | 0.29
Results from laboratories 9, 12 and 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 753b (Indicative Value 0.75 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.28 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | < 1.5 | < 1 | < 1 | < 1 | < 2 | 0.4 | 0.7 | < 0.5 | 0.39 | 1 | 2 | < 1 | < 1 | 1 | < 2 | < 5 | < 1 | < 5 | 1 | 1.37 | [6 empty] | 1
| 1 | 2 | < 1.5 | < 1 | < 1 | < 1 | < 2 | < 0.3 | 0.9 | < 0.5 | 0.21 | < 1 | 1 | < 1 | < 1 | 1 | < 2 | < 5 | < 1 | < 5 | < 1 | 1.36 | [6 empty] | 2
| 1 | 3 | < 1.5 | < 1 | < 1 | < 1 | < 2 | 0.7 | 0.9 | < 0.5 | 0.5 | < 1 | 2 | < 1 | < 1 | 1 | < 2 | < 5 | < 1 | < 5 | < 1 | 1.38 | [6 empty] | 16
| 1 | 4 | < 1.5 | < 1 | < 1 | < 1 | < 2 | 0.5 | 0.8 | < 0.5 | 0.34 | < 1 | 2 | < 1 | < 1 | 1 | < 2 | < 5 | < 1 | < 5 | < 1 | 1.42 | [6 empty] | 0.7519
| 1 | 5 | < 1.5 | < 1 | < 1 | < 1 | < 2 | 0.9 | 1 | < 0.5 | 0.43 | < 1 | 2 | < 1 | < 1 | 1 | < 2 | < 5 | < 1 | < 5 | < 1 | 1.39 | [6 empty] | 8
| 1 | 6 | < 1.5 | < 1 | < 1 | < 1 | < 2 | < 0.3 | 1 | < 0.5 | 0.34 | < 1 | 2 | < 1 | < 1 | 1 | < 2 | < 5 | < 1 | < 5 | < 1 | 1.35
| Mean |  | 1.08: 0.625; 1.09: 0.88333; 1.12: 0.36833; 1.13: 1; 1.14: 1.8333; 1.17: 1; 1.25: 1; 1.28: 1.3783
| Median |  | 1.08: 0.6; 1.09: 0.9; 1.12: 0.365; 1.13: 1; 1.14: 2; 1.17: 1; 1.25: 1; 1.28: 1.375
| Std Dev. |  | 1.08: 0.22174; 1.09: 0.1169; 1.12: 0.098268; 1.14: 0.40825; 1.17: 0; 1.28: 0.024833
| Rel.Std.Dev. |  | 1.08: 35.48%; 1.09: 13.23%; 1.12: 26.68%; 1.14: 22.27%; 1.17: 0%; 1.28: 1.802%
| PDM3 |  | 1.08: -16.88%; 1.09: 17.48%; 1.12: -51.01%; 1.13: 33%; 1.14: 143.8%; 1.17: 33%; 1.25: 33%; 1.28: 83.31%
| Z-Score (Absolute) |  | 0.59 | 0.29 | 0.29 | 0.29 | 1.47 | 0.41 | 1.06 | 1.18 | 0.76 | 0 | 4.42 | 0.29 | 0.29 | 1.47 | 1.47 | 6.77 | 0.29 | 6.77 | 0 | 2.81
| Analytical results for Sm in OREAS 753b (Certified Value 0.82 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab | [6 empty] | 1
| No | No | 1.02: 1.01; 1.03: 1.05; 1.04: 1.08; 1.06: 1.09; 1.07: 1.11; 1.08: 1.12; 1.09: 1.20; 1.10: 1.22; 1.12: 1.24; 1.13: 1.25 | [6 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.12: 4A*MS; 1.13: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.02: 0.92; 1.03: 0.9; 1.04: 0.7; 1.06: 1.1; 1.07: 0.86861; 1.08: 0.75; 1.09: 0.78578; 1.10: 0.6; 1.12: 1; 1.13: 0.89 | [6 empty] | 1
| 1 | 2 | 1.02: 0.85; 1.03: 0.85; 1.04: 0.6; 1.06: 0.6; 1.07: 0.77774; 1.08: 0.76; 1.09: 0.71745; 1.10: 0.5; 1.12: 0.9; 1.13: 0.87 | [6 empty] | 10
| 1 | 3 | 1.02: 0.82; 1.03: 0.9; 1.04: 0.8; 1.06: 0.7; 1.07: 0.76318; 1.08: 0.78; 1.09: 0.78875; 1.10: 0.5; 1.12: 0.9; 1.13: 0.83 | [6 empty] | 16
| 1 | 4 | 1.02: 0.78; 1.03: 0.84; 1.04: 0.7; 1.06: 0.7; 1.07: 0.78486; 1.08: 0.86; 1.09: 0.70575; 1.10: 0.6; 1.12: 0.9; 1.13: 0.82 | [6 empty] | 0.81696
| 1 | 5 | 1.02: 0.84; 1.03: 0.87; 1.04: 0.7; 1.06: 0.6; 1.07: 0.82946; 1.08: 0.74; 1.09: 0.76238; 1.10: 0.7; 1.12: 0.9; 1.13: 0.85 | [6 empty] | 49
| 1 | 6 | 1.02: 0.89; 1.03: 0.82; 1.04: 0.7; 1.06: 0.8; 1.07: 0.84022; 1.08: 0.81; 1.09: 0.7063; 1.10: 0.6; 1.12: 0.9; 1.13: 0.84
| Mean |  | 1.02: 0.85; 1.03: 0.86333; 1.04: 0.7; 1.06: 0.75; 1.07: 0.81068; 1.08: 0.78333; 1.09: 0.7444; 1.10: 0.58333; 1.12: 0.91667; 1.13: 0.85
| Median |  | 1.02: 0.845; 1.03: 0.86; 1.04: 0.7; 1.06: 0.7; 1.07: 0.80716; 1.08: 0.77; 1.09: 0.73991; 1.10: 0.6; 1.12: 0.9; 1.13: 0.845
| Std Dev. |  | 1.02: 0.0498; 1.03: 0.03266; 1.04: 0.063246; 1.06: 0.18708; 1.07: 0.041447; 1.08: 0.045019; 1.09: 0.039177; 1.10: 0.075277; 1.12: 0.040825; 1.13: 0.026077
| Rel.Std.Dev. |  | 1.02: 5.859%; 1.03: 3.783%; 1.04: 9.035%; 1.06: 24.94%; 1.07: 5.113%; 1.08: 5.747%; 1.09: 5.263%; 1.10: 12.9%; 1.12: 4.454%; 1.13: 3.068%
| PDM3 |  | 1.02: 4.045%; 1.03: 5.677%; 1.04: -14.32%; 1.06: -8.196%; 1.07: -0.7686%; 1.08: -4.116%; 1.09: -8.881%; 1.10: -28.6%; 1.12: 12.2%; 1.13: 4.045%
| Z-Score (Absolute) |  | 1.02: 0.5; 1.03: 0.85; 1.04: NA; 1.06: NA; 1.07: 0.5; 1.08: 1.2; 1.09: 2.2; 1.10: NA; 1.12: NA; 1.13: 0.5
| Analytical results for Sn in OREAS 753b (Certified Value 109 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 117.2 | 109.9096 | 105.7 | 106 | 108.1 | 125 | 100.5 | 110.8 | 95 | 109 | 113.1108 | 115 | 120.5 | 105.5 | 104.5 | 118 | 115 | 104.4 | 108 | 107.8 | 101 | 98.1 | 118 | 124 | 92.14 |  | 1
| 1 | 2 | 116.1 | 112.744 | 105 | 103 | 104.2 | 120 | 104.2 | 107 | 87 | 112 | 110.2469 | 106 | 122 | 108 | 99.8 | 116 | 114 | 107.9 | 106 | 106 | 91.9 | 107.4 | 121 | 103 | 89.96 |  | 0
| 1 | 3 | 112.2 | 110.2598 | 106.8 | 96.7 | 109.6 | 119 | 105.3 | 109.5 | 83 | 111 | 112.8926 | 109 | 118 | 108.5 | 100.5 | 120 | 114.5 | 104.9 | 108 | 101.7 | 91.5 | 98.5 | 119 | 107 | 91.95 |  | 16
| 1 | 4 | 116.1 | 107.4782 | 110.5 | 100 | 109.9 | 125 | 100.2 | 111.8 | 84 | 112 | 111.3685 | 115 | 118.5 | 108 | 105 | 124.5 | 112.5 | 107 | 108 | 108.7 | 90 | 97.4 | 118 | 113 | 92.14 |  | 109.4955
| 1 | 5 | 126.5 | 109.9545 | 109.5 | 101 | 107.8 | 124 | 97.3 | 110.8 | 87 | 113 | 112.7323 | 110 | 120 | 105.5 | 102 | 120 | 103.5 | 104.9 | 109 | 102.8 | 94.2 | 87.9 | 117 | 113 | 90.98 |  | 50
| 1 | 6 | 119.8 | 106.8102 | 111.8 | 104 | 108.6 | 130 | 99.5 | 112 | 90 | 112 | 111.1115 | 107 | 120.5 | 107 | 103.5 | 126 | 99.8 | 106 | 110 | 102.6 | 93 | 98.3 | 120 | 110 | 90.54
| Mean |  | 117.9833 | 109.526 | 108.2167 | 101.7833 | 108.0333 | 123.8333 | 101.1667 | 110.3167 | 87.6667 | 111.5 | 111.9104 | 110.3333 | 119.9167 | 107.0833 | 102.55 | 120.75 | 109.8833 | 105.85 | 108.1667 | 104.9333 | 93.6 | 97.9333 | 118.8333 | 111.6667 | 91.285
| Median |  | 116.65 | 109.932 | 108.15 | 102 | 108.35 | 124.5 | 100.35 | 110.8 | 87 | 112 | 112.0504 | 109.5 | 120.25 | 107.5 | 102.75 | 120 | 113.25 | 105.45 | 108 | 104.4 | 92.45 | 98.2 | 118.5 | 111.5 | 91.465
| Std Dev. |  | 4.8371 | 2.1351 | 2.7709 | 3.2805 | 2.05 | 3.9707 | 3.0131 | 1.853 | 4.3665 | 1.3784 | 1.1645 | 3.8816 | 1.4634 | 1.3197 | 2.1342 | 3.8177 | 6.5377 | 1.3751 | 1.3292 | 2.9663 | 3.892 | 6.1815 | 1.472 | 7.1461 | 0.92822
| Rel.Std.Dev. |  | 4.1% | 1.949% | 2.56% | 3.223% | 1.898% | 3.207% | 2.978% | 1.68% | 4.981% | 1.236% | 1.041% | 3.518% | 1.22% | 1.232% | 2.081% | 3.162% | 5.95% | 1.299% | 1.229% | 2.827% | 4.158% | 6.312% | 1.239% | 6.399% | 1.017%
| PDM3 |  | 7.752% | 0.02785% | -1.168% | -7.043% | -1.335% | 13.09% | -7.607% | 0.7499% | -19.94% | 1.831% | 2.205% | 0.7651% | 9.517% | -2.203% | -6.343% | 10.28% | 0.3542% | -3.329% | -1.214% | -4.167% | -14.52% | -10.56% | 8.528% | 1.983% | -16.63%
| Z-Score (Absolute) |  | 1.78 | 0.24 | 0 | 1.17 | 0.03 | 2.85 | 1.29 | 0.38 | 3.75 | 0.6 | 0.67 | 0.39 | 2.14 | 0.21 | 1.03 | 2.29 | 0.3 | 0.43 | 0.01 | 0.6 | 2.67 | 1.88 | 1.94 | 0.63 | 3.09
| Analytical results for Sr in OREAS 753b (Certified Value 38.3 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 1
2
| 1 | 1 | 41.42 | 36.8328 | 36.5 | 49.8 | 37.87 | 38.2 | 39.8 | 51 | 37.5 | 42.1 | 36.1174 | 40 | 39.4 | 36.7 | 39 | 38 | 39.6 | 28.3 | 38.7 | 50.529 | 38 | 23.2 | 35 | 52 | 37.13 |  | 1
| 1 | 2 | 40.32 | 37.6325 | 35.3 | 49.9 | 37.92 | 37.7 | 40.3 | 45 | 36 | 42 | 36.0836 | 44 | 38.6 | 37.6 | 37.6 | 38.3 | 39.7 | 26.8 | 35 | 51.537 | 37.6 | 25.4 | 34 | 52 | 39.43 |  | 0
| 1 | 3 | 40.06 | 37.8207 | 36.6 | 49.9 | 38.16 | 37.7 | 40.3 | 55 | 36.3 | 40.5 | 36.0511 | 44.3 | 40 | 38.5 | 37.2 | 39 | 38.4 | 28.2 | 33.6 | 52.881 | 37 | 26.8 | 34 | 52 | 38.42 |  | 16
| 1 | 4 | 39.77 | 36.6636 | 37.2 | 50.3 | 37.9 | 37.6 | 39.7 | 51 | 35.2 | 41.6 | 36.3055 | 43.7 | 38.7 | 32.2 | 38 | 39.4 | 37.8 | 26.7 | 40.3 | 52.755 | 38.7 | 25.7 | 35 | 52 | 40.25 |  | 38.2516
| 1 | 5 | 41.19 | 37.5811 | 37 | 49.8 | 37.78 | 38.4 | 39.6 | 53 | 37.6 | 41.2 | 36.468 | 43.3 | 38.3 | 34.6 | 37 | 38.6 | 37.8 | 28.4 | 35.5 | 51.18 | 37.9 | 31 | 31 | 52 | 40.24 |  | 51
| 1 | 6 | 40.89 | 36.9278 | 37 | 50.6 | 37.51 | 38.5 | 39.8 | 52 | 36.7 | 43.2 | 36.3431 | 45.9 | 39.2 | 35.6 | 37.4 | 40.5 | 39.7 | 27.9 | 38.1 | 51.1905 | 38.2 | 30.3 | 33 | 52 | 38.98
| Mean |  | 40.6083 | 37.2431 | 36.6 | 50.05 | 37.8567 | 38.0167 | 39.9167 | 51.1667 | 36.55 | 41.7667 | 36.2281 | 43.5333 | 39.0333 | 35.8667 | 37.7 | 38.9667 | 38.8333 | 27.7167 | 36.8667 | 51.6788 | 37.9 | 27.0667 | 33.6667 | 52 | 39.075
| Median |  | 40.605 | 37.2545 | 36.8 | 49.9 | 37.885 | 37.95 | 39.8 | 51.5 | 36.5 | 41.8 | 36.2114 | 43.85 | 38.95 | 36.15 | 37.5 | 38.8 | 39 | 28.05 | 36.8 | 51.3638 | 37.95 | 26.25 | 34 | 52 | 39.205
| Std Dev. |  | 0.65774 | 0.49052 | 0.68993 | 0.32711 | 0.21172 | 0.39707 | 0.30605 | 3.3714 | 0.91815 | 0.9136 | 0.16808 | 1.948 | 0.62183 | 2.2695 | 0.72388 | 0.90037 | 0.9395 | 0.7679 | 2.5571 | 0.94158 | 0.57271 | 3.0197 | 1.5055 | 0 | 1.1904
| Rel.Std.Dev. |  | 1.62% | 1.317% | 1.885% | 0.6536% | 0.5593% | 1.044% | 0.7667% | 6.589% | 2.512% | 2.187% | 0.4639% | 4.475% | 1.593% | 6.328% | 1.92% | 2.311% | 2.419% | 2.771% | 6.936% | 1.822% | 1.511% | 11.16% | 4.472% | 0% | 3.046%
| PDM3 |  | 6.161% | -2.636% | -4.318% | 30.84% | -1.032% | -0.6141% | 4.353% | 33.76% | -4.448% | 9.189% | -5.29% | 13.81% | 2.044% | -6.235% | -1.442% | 1.87% | 1.521% | -27.54% | -3.62% | 35.1% | -0.9191% | -29.24% | -11.99% | 35.94% | 2.153%
| Z-Score (Absolute) |  | 0.98 | 0.29 | 0.53 | 4.54 | 0.06 | 0 | 0.72 | 4.96 | 0.55 | 1.41 | 0.67 | 2.08 | 0.38 | 0.81 | 0.12 | 0.36 | 0.31 | 3.88 | 0.43 | 5.15 | 0.04 | 4.13 | 1.64 | 5.27 | 0.4
| Analytical results for Ta in OREAS 753b (Certified Value 283 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 297.08 | 281.6483 | > 100 | 205 | 280.73 | 287 | 241.6 | 263.1 | 86.2 | 144 | 290.4721 | 353.17 | 302 | 280 | 287 | 300 | 310 | 485.8 | 240 | 268.6 | 225 | 161.4 | 281 | 288 | 300.56 |  | 1
| 1 | 2 | 294.31 | 287.1416 | > 100 | 191 | 277.45 | 290 | 231.07 | 247.6 | 56 | 183 | 290.075 | 350.19 | 307 | 295 | 278 | 302 | 314 | 476 | 242 | 270.6 | 244 | 170.8 | 287 | 286 | 292.33 |  | 12
| 1 | 3 | 299.43 | 282.3869 | > 100 | 191 | 277.23 | 288 | 237.55 | 267.6 | 49.9 | 240 | 292.4419 | 351.58 | 298 | 287 | 275 | 301 | 300 | 490.3 | 246 | 260.5 | 219 | 140 | 280 | 287 | 289.43 |  | 16
| 1 | 4 | 298.2 | 272.7136 | > 100 | 207 | 273.61 | 296 | 235.49 | 264.4 | 59.3 | 141 | 296.7939 | 354.69 | 302 | 277 | 287 | 303 | 310 | 478.4 | 248 | 273.1 | 214 | 152.1 | 271 | 294 | 297.61 |  | 282.846
| 1 | 5 | 296.38 | 281.5546 | > 100 | 199 | 275.8 | 292 | 241.11 | 265.7 | 65.4 | 145 | 291.3413 | 351.12 | 307 | 277 | 284 | 294 | 290 | 492 | 244 | 275.1 | 255 | 121.2 | 278 | 294 | 299.5 |  | 52
| 1 | 6 | 299.18 | 272.6563 | > 100 | 210 | 283.41 | 296 | 233.19 | 249.3 | 54.1 | 161 | 291.3559 | 353.17 | 306 | 281 | 286 | 310 | 298 | 475.7 | 242 | 261.6 | 248 | 86.5 | 276 | 295 | 299.49
| Mean |  | 297.43 | 279.6835 |  | 200.5 | 278.0383 | 291.5 | 236.6683 | 259.6167 | 61.8167 | 169 | 292.08 | 352.32 | 303.6667 | 282.8333 | 282.8333 | 301.6667 | 303.6667 | 483.0333 | 243.6667 | 268.25 | 234.1667 | 138.6667 | 278.8333 | 290.6667 | 296.4867
| Median |  | 297.64 | 281.6014 |  | 202 | 277.34 | 291 | 236.52 | 263.75 | 57.65 | 153 | 291.3486 | 352.375 | 304 | 280.5 | 285 | 301.5 | 305 | 482.1 | 243 | 269.6 | 234.5 | 146.05 | 279 | 291 | 298.55
| Std Dev. |  | 1.9295 | 5.8009 |  | 8.1915 | 3.5134 | 3.8859 | 4.2345 | 8.793 | 13.0316 | 38.1733 | 2.4506 | 1.6484 | 3.6148 | 6.9976 | 5.1153 | 5.164 | 9.1579 | 7.288 | 2.9439 | 6.0062 | 16.9873 | 30.8481 | 5.3448 | 4.0825 | 4.5394
| Rel.Std.Dev. |  | 0.6487% | 2.074% |  | 4.086% | 1.264% | 1.333% | 1.789% | 3.387% | 21.08% | 22.59% | 0.839% | 0.4679% | 1.19% | 2.474% | 1.809% | 1.712% | 3.016% | 1.509% | 1.208% | 2.239% | 7.254% | 22.25% | 1.917% | 1.405% | 1.531%
| PDM3 |  | 5.156% | -1.118% |  | -29.11% | -1.7% | 3.06% | -16.33% | -8.213% | -78.14% | -40.25% | 3.265% | 24.56% | 7.361% | -0.00447% | -0.00447% | 6.654% | 7.361% | 70.78% | -13.85% | -5.16% | -17.21% | -50.97% | -1.419% | 2.765% | 4.823%
| Z-Score (Absolute) |  | 0.52 | 0.05 | NA | 2.59 | 0.1 | 0.33 | 1.43 | 0.69 | 7.04 | 3.6 | 0.35 | 2.28 | 0.72 | 0.05 | 0.05 | 0.65 | 0.72 | 6.47 | 1.21 | 0.42 | 1.51 | 4.57 | 0.08 | 0.3 | 0.49
| Analytical results for Tb in OREAS 753b (Certified Value 0.1 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.04; 1.03: 1.05; 1.04: 1.08; 1.05: 1.09; 1.06: 1.11; 1.07: 1.12; 1.08: 1.18; 1.09: 1.20; 1.10: 1.21; 1.11: 1.22; 1.12: 1.24; 1.13: 1.25; 1.14: 1.28 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*MS; 1.14: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.09; 1.02: 0.11; 1.03: 0.11; 1.04: < 0.1; 1.05: 0.1; 1.06: 0.11503; 1.07: 0.09; 1.08: 0.07; 1.09: < 0.1; 1.10: 0.1; 1.11: 0.08; 1.12: 0.12; 1.13: 0.11; 1.14: < 0.5 |  | 1
| 1 | 2 | 1.01: 0.09; 1.02: 0.1; 1.03: 0.1; 1.04: < 0.1; 1.05: 0.1; 1.06: 0.11701; 1.07: 0.09; 1.08: 0.06; 1.09: 0.11826; 1.10: 0.1; 1.11: 0.05; 1.12: 0.12; 1.13: 0.1; 1.14: < 0.5 |  | 13
| 1 | 3 | 1.01: 0.1; 1.02: 0.1; 1.03: 0.1; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.11454; 1.07: 0.09; 1.08: 0.07; 1.09: 0.11574; 1.10: 0.1; 1.11: 0.06; 1.12: 0.11; 1.13: 0.11; 1.14: < 0.5 |  | 16
| 1 | 4 | 1.01: 0.09; 1.02: 0.1; 1.03: 0.1; 1.04: < 0.1; 1.05: < 0.1; 1.06: 0.10511; 1.07: 0.1; 1.08: 0.07; 1.09: 0.1152; 1.10: 0.11; 1.11: 0.06; 1.12: 0.11; 1.13: 0.1; 1.14: < 0.5 |  | 0.10039
| 1 | 5 | 1.01: 0.09; 1.02: 0.09; 1.03: 0.1; 1.04: < 0.1; 1.05: 0.1; 1.06: 0.11632; 1.07: 0.09; 1.08: 0.06; 1.09: 0.10656; 1.10: 0.11; 1.11: 0.07; 1.12: 0.11; 1.13: 0.1; 1.14: < 0.5 |  | 53
| 1 | 6 | 1.01: 0.09; 1.02: 0.1; 1.03: 0.1; 1.04: < 0.1; 1.05: 0.1; 1.06: 0.118; 1.07: 0.1; 1.08: 0.07; 1.09: 0.10782; 1.10: 0.1; 1.11: 0.06; 1.12: 0.11; 1.13: 0.11; 1.14: < 0.5
| Mean |  | 1.01: 0.091667; 1.02: 0.1; 1.03: 0.10167; 1.05: 0.1; 1.06: 0.11434; 1.07: 0.093333; 1.08: 0.066667; 1.09: 0.11272; 1.10: 0.10333; 1.11: 0.063333; 1.12: 0.11333; 1.13: 0.105
| Median |  | 1.01: 0.09; 1.02: 0.1; 1.03: 0.1; 1.05: 0.1; 1.06: 0.11568; 1.07: 0.09; 1.08: 0.07; 1.09: 0.1152; 1.10: 0.1; 1.11: 0.06; 1.12: 0.11; 1.13: 0.105
| Std Dev. |  | 1.01: 0.0040825; 1.02: 0.0063246; 1.03: 0.0040825; 1.05: 0; 1.06: 0.0046937; 1.07: 0.005164; 1.08: 0.005164; 1.09: 0.0051942; 1.10: 0.005164; 1.11: 0.010328; 1.12: 0.005164; 1.13: 0.0054772
| Rel.Std.Dev. |  | 1.01: 4.454%; 1.02: 6.325%; 1.03: 4.016%; 1.05: 0%; 1.06: 4.105%; 1.07: 5.533%; 1.08: 7.746%; 1.09: 4.608%; 1.10: 4.997%; 1.11: 16.31%; 1.12: 4.556%; 1.13: 5.216%
| PDM3 |  | 1.01: -8.689%; 1.02: -0.3881%; 1.03: 1.272%; 1.05: -0.3881%; 1.06: 13.89%; 1.07: -7.029%; 1.08: -33.59%; 1.09: 12.28%; 1.10: 2.932%; 1.11: -36.91%; 1.12: 12.89%; 1.13: 4.593%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.11; 1.03: 0; 1.04: 3.48; 1.05: NA; 1.06: 0.85; 1.07: 0.56; 1.08: 2.36; 1.09: 0.04; 1.10: 0.11; 1.11: 2.58; 1.12: 0.79; 1.13: 0.22; 1.14: 10
Results from laboratory 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 753b (Certified Value 0.071 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.25 | 1.28 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | < 0.2 | 0.2 | 0.07 | 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.11 | 0.07007 | 0.05 | 0.08 | 0.09 | 0.07 | 0.08 | 0.05 | 0.99 | < 0.05 | 0.54 | < 0.5 | [7 empty] | 1
| 1 | 2 | < 0.2 | 0.18 | 0.08 | 0.1 | < 0.1 | 0.07 | < 0.1 | 0.07 | 0.0679 | 0.09 | 0.06 | 0.07 | 0.11 | 0.1 | < 0.05 | 0.96 | < 0.05 | 0.35 | < 0.5 | [7 empty] | 6
| 1 | 3 | < 0.2 | 0.17 | 0.06 | 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.1 | 0.06043 | 0.06 | 0.07 | 0.07 | 0.14 | 0.08 | 0.05 | 1.1 | < 0.05 | 0.27 | < 0.5 | [7 empty] | 16
| 1 | 4 | < 0.2 | 0.16 | 0.06 | 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.11 | 0.06982 | 0.09 | 0.09 | 0.1 | 0.13 | 0.08 | < 0.05 | 1.11 | < 0.05 | 0.2 | < 0.5 | [7 empty] | 0.071203
| 1 | 5 | < 0.2 | 0.15 | 0.08 | 0.1 | < 0.1 | < 0.05 | < 0.1 | 0.07 | 0.06456 | 0.08 | 0.07 | 0.07 | 0.11 | 0.07 | < 0.05 | 1.1 | < 0.05 | 0.16 | < 0.5 | [7 empty] | 54
| 1 | 6 | < 0.2 | 0.18 | 0.08 | 0.1 | < 0.1 | 0.07 | < 0.1 | 0.06 | 0.05817 | 0.05 | 0.07 | 0.07 | 0.11 | 0.07 | 0.05 | 1.05 | < 0.05 | 0.15 | < 0.5
| Mean |  |  | 0.17333 | 0.071667 | 0.1 |  | 0.07 |  | 0.086667 | 0.065158 | 0.07 | 0.073333 | 0.078333 | 0.11167 | 0.08 | 0.05 | 1.0517 |  | 0.27833
| Median |  |  | 0.175 | 0.075 | 0.1 |  | 0.07 |  | 0.085 | 0.06623 | 0.07 | 0.07 | 0.07 | 0.11 | 0.08 | 0.05 | 1.075 |  | 0.235
| Std Dev. |  |  | 0.017512 | 0.0098319 | 1.5202e-17 |  | 0 |  | 0.022509 | 0.0049985 | 0.018974 | 0.010328 | 0.013292 | 0.024014 | 0.010954 | 8.4984e-18 | 0.063692 |  | 0.14851
| Rel.Std.Dev. |  |  | 10.1% | 13.72% | 1.52e-14% |  | 0% |  | 25.97% | 7.671% | 27.11% | 14.08% | 16.97% | 21.5% | 13.69% | 1.7e-14% | 6.056% |  | 53.36%
| PDM3 |  |  | 143.4% | 0.6515% | 40.44% |  | -1.689% |  | 21.72% | -8.489% | -1.689% | 2.992% | 10.01% | 56.83% | 12.36% | -29.78% | 1377% |  | 290.9%
| Z-Score (Absolute) |  | 0.63 | 2.54 | 0.11 | NA | 0.67 | 0.94 | 0.67 | 0.28 | 0.28 | 0.15 | 0.07 | 0.07 | 0.94 | 0.11 | 1 | 25.47 | 1.33 | 5.29 | 4.55
Results from laboratory 6 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 753b (Certified Value 1.92 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.1 | 1.7027 | 1.9 | 2.27 | 2.06 | 1.98 | 2.17 | 1.8 | 2.3 | 2 | 2.1775 | 2.03 | 1.79 | 1.93 | 1.78 | 1.57 | 1.88 | 1.4 | 1.42 | 2.1 | 2.77 | 2.51 | 1.27 |  |  |  | 1
| 1 | 2 | 2.08 | 1.8369 | 1.9 | 2.07 | 2.03 | 2.04 | 2.25 | 1.6 | 2 | 2.3 | 1.9063 | 1.97 | 1.65 | 1.99 | 1.76 | 1.62 | 2.17 | 1.3 | 1.25 | 1.9 | 2.5 | 2.1 | 1.44 |  |  |  | 15
| 1 | 3 | 2.1 | 1.8038 | 1.9 | 2.14 | 2.07 | 2.06 | 2.16 | 2.1 | 1.9 | 2 | 1.9651 | 1.9 | 1.74 | 1.8 | 1.8 | 1.7 | 1.8 | 1.4 | 1.37 | 2.1 | 2.37 | 2.33 | 1.78 |  |  |  | 16
| 1 | 4 | 2.06 | 1.7692 | 1.7 | 2.15 | 2.15 | 2.08 | 2.25 | 1.8 | 1.9 | 2.2 | 1.9754 | 2.04 | 1.86 | 1.36 | 1.95 | 1.69 | 2.05 | 1.3 | 1.47 | 2.1 | 2.31 | 2.27 | 1.97 |  |  |  | 1.9228
| 1 | 5 | 2.08 | 1.6962 | 1.7 | 2.19 | 2.06 | 1.99 | 2.17 | 1.9 | 2.2 | 2.2 | 2.1395 | 1.76 | 1.79 | 1.42 | 1.8 | 1.68 | 1.81 | 1.4 | 1.77 | 1.7 | 2.5 | 2.28 | 2.05 |  |  |  | 55
| 1 | 6 | 2.11 | 1.8291 | 1.7 | 2.05 | 2.2 | 2.02 | 2.19 | 1.8 | 2.1 | 2.4 | 2.1319 | 2.13 | 1.89 | 1.47 | 1.72 | 1.7 | 1.71 | 1.3 | 1.52 | 1.7 | 2.53 | 2.31 | 1.35
| Mean |  | 2.0883 | 1.773 | 1.8 | 2.145 | 2.095 | 2.0283 | 2.1983 | 1.8333 | 2.0667 | 2.1833 | 2.0493 | 1.9717 | 1.7867 | 1.6617 | 1.8017 | 1.66 | 1.9033 | 1.35 | 1.4667 | 1.9333 | 2.4967 | 2.3 | 1.6433
| Median |  | 2.09 | 1.7865 | 1.8 | 2.145 | 2.065 | 2.03 | 2.18 | 1.8 | 2.05 | 2.2 | 2.0536 | 2 | 1.79 | 1.635 | 1.79 | 1.685 | 1.845 | 1.35 | 1.445 | 2 | 2.5 | 2.295 | 1.61
| Std Dev. |  | 0.018348 | 0.061706 | 0.10954 | 0.080436 | 0.065345 | 0.0392 | 0.041191 | 0.1633 | 0.1633 | 0.16021 | 0.11348 | 0.12891 | 0.085946 | 0.27752 | 0.078592 | 0.053292 | 0.17317 | 0.054772 | 0.17512 | 0.19664 | 0.15921 | 0.13145 | 0.33393
| Rel.Std.Dev. |  | 0.8786% | 3.48% | 6.086% | 3.75% | 3.119% | 1.933% | 1.874% | 8.907% | 7.902% | 7.338% | 5.537% | 6.538% | 4.81% | 16.7% | 4.362% | 3.21% | 9.098% | 4.057% | 11.94% | 10.17% | 6.377% | 5.715% | 20.32%
| PDM3 |  | 8.61% | -7.791% | -6.385% | 11.56% | 8.957% | 5.49% | 14.33% | -4.652% | 7.483% | 13.55% | 6.579% | 2.543% | -7.079% | -13.58% | -6.299% | -13.67% | -1.011% | -29.79% | -23.72% | 0.5489% | 29.85% | 19.62% | -14.53%
| Z-Score (Absolute) |  | 0.65 | 0.67 | 0.56 | 0.89 | 0.68 | 0.4 | 1.11 | 0.42 | 0.56 | 1.05 | 0.49 | 0.16 | 0.62 | 1.14 | 0.55 | 1.15 | 0.13 | 2.45 | 1.96 | 0 | 2.37 | 1.54 | 1.22
| Analytical results for Ti in OREAS 753b (Certified Value 0.052 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.0523 | 0.051029 | 0.06 | 0.05 | 0.0545 | 0.0503 | 0.0489 | 0.057 | 0.0552 | 0.05 | 0.051748 | 0.0461 | 0.051 | 0.051 | 0.049 | 0.052 | 0.056 | 0.05 | 0.05 | 0.050583 | 0.052 | 0.055 | 0.049 | 0.0528 | 0.052 | 0.06 | 1
| 1 | 2 | 0.0528 | 0.051902 | 0.06 | 0.05 | 0.0554 | 0.0525 | 0.0484 | 0.056 | 0.0556 | 0.05 | 0.05124 | 0.0486 | 0.052 | 0.051 | 0.047 | 0.052 | 0.057 | 0.04 | 0.05 | 0.050247 | 0.052 | 0.056 | 0.05 | 0.0552 | 0.05 | 0.05 | 0
| 1 | 3 | 0.0531 | 0.050672 | 0.06 | 0.05 | 0.0552 | 0.0503 | 0.0471 | 0.058 | 0.0532 | 0.05 | 0.050285 | 0.0521 | 0.05 | 0.053 | 0.049 | 0.053 | 0.056 | 0.05 | 0.05 | 0.050913 | 0.052 | 0.056 | 0.05 | 0.0534 | 0.05 | 0.05 | 16
| 1 | 4 | 0.0539 | 0.049631 | 0.05 | 0.05 | 0.0548 | 0.051 | 0.0514 | 0.057 | 0.0536 | 0.05 | 0.052537 | 0.0488 | 0.051 | 0.051 | 0.049 | 0.053 | 0.056 | 0.04 | 0.05 | 0.050985 | 0.053 | 0.056 | 0.048 | 0.054 | 0.05 | 0.06 | 0.052081
| 1 | 5 | 0.0536 | 0.051589 | 0.07 | 0.049 | 0.054 | 0.052 | 0.0545 | 0.057 | 0.053 | 0.05 | 0.051906 | 0.0487 | 0.051 | 0.051 | 0.049 | 0.051 | 0.054 | 0.05 | 0.05 | 0.050534 | 0.052 | 0.057 | 0.049 | 0.054 | 0.049 | 0.06 | 56
| 1 | 6 | 0.0537 | 0.04994 | 0.06 | 0.051 | 0.0545 | 0.0528 | 0.0539 | 0.058 | 0.0554 | 0.05 | 0.053802 | 0.0518 | 0.051 | 0.052 | 0.049 | 0.053 | 0.057 | 0.04 | 0.05 | 0.050371 | 0.053 | 0.056 | 0.047 | 0.0546 | 0.049 | 0.05
| Mean |  | 0.053233 | 0.050794 | 0.06 | 0.05 | 0.054733 | 0.051483 | 0.0507 | 0.057167 | 0.054333 | 0.05 | 0.05192 | 0.04935 | 0.051 | 0.0515 | 0.048667 | 0.052333 | 0.056 | 0.045 | 0.05 | 0.050605 | 0.052333 | 0.056 | 0.048833 | 0.054 | 0.05 | 0.055
| Median |  | 0.05335 | 0.050851 | 0.06 | 0.05 | 0.05465 | 0.0515 | 0.05015 | 0.057 | 0.0544 | 0.05 | 0.051827 | 0.04875 | 0.051 | 0.051 | 0.049 | 0.0525 | 0.056 | 0.045 | 0.05 | 0.050558 | 0.052 | 0.056 | 0.049 | 0.054 | 0.05 | 0.055
| Std Dev. |  | 0.0006121 | 0.0008957 | 0.0063246 | 0.00063246 | 0.00051251 | 0.0011017 | 0.0030548 | 0.00075277 | 0.0011911 | 7.6012e-18 | 0.0011907 | 0.0022546 | 0.00063246 | 0.00083666 | 0.0008165 | 0.0008165 | 0.0010954 | 0.0054772 | 7.6012e-18 | 0.00029246 | 0.0005164 | 0.00063246 | 0.001169 | 0.00084853 | 0.0010954 | 0.0054772
| Rel.Std.Dev. |  | 1.15% | 1.763% | 10.54% | 1.265% | 0.9364% | 2.14% | 6.025% | 1.317% | 2.192% | 1.52e-14% | 2.293% | 4.568% | 1.24% | 1.625% | 1.678% | 1.56% | 1.956% | 12.17% | 1.52e-14% | 0.5779% | 0.9867% | 1.129% | 2.394% | 1.571% | 2.191% | 9.959%
| PDM3 |  | 2.213% | -2.471% | 15.21% | -3.995% | 5.093% | -1.147% | -2.651% | 9.765% | 4.325% | -3.995% | -0.3093% | -5.243% | -2.075% | -1.115% | -6.555% | 0.4849% | 7.525% | -13.6% | -3.995% | -2.833% | 0.4849% | 7.525% | -6.235% | 3.685% | -3.995% | 5.605%
| Z-Score (Absolute) |  | 0.73 | 0.29 | 3.55 | 0.62 | 1.35 | 0 | 0.33 | 2.37 | 1.19 | 0.62 | 0.18 | 0.89 | 0.21 | 0 | 1.18 | 0.35 | 1.88 | 2.71 | 0.62 | 0.37 | 0.35 | 1.88 | 1.11 | 1.03 | 0.62 | 1.46
| Analytical results for Tl in OREAS 753b (Certified Value 8.73 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.25 | 1.28 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 8.87 | 8.5529 | 8.76 | 8.64 | 9.11 | 8.45 | 8.62 | 8.5 | 8.71 | 9.3 | 8.7224 | 8.74 | 8.76 | 8.23 | 8.13 | 8.61 | 9.4 | 8.21 | 9.0851 | 9.27 | 7.52 | 8.23 | 8.45 |  |  |  | 1
| 1 | 2 | 8.99 | 8.6441 | 8.08 | 9.07 | 8.81 | 8.44 | 8.56 | 8.18 | 8.69 | 9.46 | 8.7463 | 8.51 | 8.96 | 8.66 | 8.02 | 8.39 | 10 | 8.03 | 9.0701 | 8.98 | 6.81 | 8.22 | 7.95 |  |  |  | 16
| 1 | 3 | 9.18 | 8.4924 | 8.24 | 8.63 | 8.94 | 8.41 | 8.55 | 8.93 | 8.63 | 9 | 8.7886 | 8.5 | 8.57 | 9.37 | 8.08 | 8.51 | 9.75 | 8.23 | 9.0864 | 8.85 | 6.94 | 8.2 | 9.05 |  |  |  | 16
| 1 | 4 | 9.13 | 8.25 | 8.53 | 8.58 | 8.86 | 8.7 | 8.63 | 8.67 | 8.55 | 9.14 | 8.9049 | 8.48 | 8.8 | 8.9 | 8.18 | 8.82 | 9.75 | 8.11 | 9.0049 | 9.16 | 6.88 | 8.38 | 9.52 |  |  |  | 8.7265
| 1 | 5 | 9.09 | 8.3593 | 8.42 | 9.12 | 8.84 | 8.64 | 8.65 | 8.88 | 8.79 | 9.14 | 8.9205 | 8.42 | 9.19 | 8.79 | 8.23 | 8.15 | 9.22 | 8.04 | 9.0279 | 9.01 | 6.86 | 8.5 | 9.31 |  |  |  | 57
| 1 | 6 | 9.08 | 8.2155 | 8.3 | 8.93 | 8.99 | 8.58 | 8.63 | 8.34 | 8.7 | 9.65 | 8.9358 | 8.69 | 9.34 | 8.91 | 8.09 | 8.85 | 9.36 | 8.17 | 9 | 9.09 | 6.83 | 8.58 | 8.38
| Mean |  | 9.0567 | 8.419 | 8.3883 | 8.8283 | 8.925 | 8.5367 | 8.6067 | 8.5833 | 8.6783 | 9.2817 | 8.8364 | 8.5567 | 8.9367 | 8.81 | 8.1217 | 8.555 | 9.58 | 8.1317 | 9.0457 | 9.06 | 6.9733 | 8.3517 | 8.7767
| Median |  | 9.085 | 8.4258 | 8.36 | 8.785 | 8.9 | 8.515 | 8.625 | 8.585 | 8.695 | 9.22 | 8.8467 | 8.505 | 8.88 | 8.845 | 8.11 | 8.56 | 9.575 | 8.14 | 9.049 | 9.05 | 6.87 | 8.305 | 8.75
| Std Dev. |  | 0.11094 | 0.17175 | 0.23836 | 0.24095 | 0.11256 | 0.12011 | 0.041312 | 0.29817 | 0.081097 | 0.2397 | 0.094913 | 0.12754 | 0.28696 | 0.37175 | 0.075211 | 0.26591 | 0.29819 | 0.085421 | 0.039684 | 0.14697 | 0.27156 | 0.1613 | 0.60978
| Rel.Std.Dev. |  | 1.225% | 2.04% | 2.842% | 2.729% | 1.261% | 1.407% | 0.48% | 3.474% | 0.9345% | 2.583% | 1.074% | 1.491% | 3.211% | 4.22% | 0.9261% | 3.108% | 3.113% | 1.05% | 0.4387% | 1.622% | 3.894% | 1.931% | 6.948%
| PDM3 |  | 3.784% | -3.523% | -3.875% | 1.167% | 2.275% | -2.175% | -1.373% | -1.64% | -0.5515% | 6.362% | 1.26% | -1.946% | 2.409% | 0.9573% | -6.931% | -1.965% | 9.781% | -6.816% | 3.659% | 3.822% | -20.09% | -4.295% | 0.5753%
| Z-Score (Absolute) |  | 0.99 | 0.68 | 0.76 | 0.39 | 0.64 | 0.37 | 0.19 | 0.25 | 0 | 1.57 | 0.41 | 0.32 | 0.67 | 0.34 | 1.45 | 0.32 | 2.35 | 1.43 | 0.96 | 1 | 4.45 | 0.85 | 0.26
| Analytical results for Tm in OREAS 753b (Indicative Value 0.031 ppm) | [27 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.08; 1.04: 1.09; 1.05: 1.11; 1.06: 1.12; 1.07: 1.22; 1.08: 1.25 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*OES/MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*MS |  |  |  | 3
3
| 1 | 1 | 1.01: 0.02; 1.02: 0.03; 1.03: < 0.1; 1.04: < 0.1; 1.05: 0.04396; 1.06: 0.02; 1.07: < 0.03; 1.08: 0.03 |  |  |  | 1
| 1 | 2 | 1.01: 0.03; 1.02: 0.03; 1.03: < 0.1; 1.04: < 0.1; 1.05: 0.05604; 1.06: 0.03; 1.07: < 0.03; 1.08: 0.03 |  |  |  | 1
| 1 | 3 | 1.01: 0.02; 1.02: 0.03; 1.03: < 0.1; 1.04: < 0.1; 1.05: 0.04547; 1.06: 0.02; 1.07: < 0.03; 1.08: 0.03 |  |  |  | 16
| 1 | 4 | 1.01: 0.02; 1.02: 0.03; 1.03: < 0.1; 1.04: < 0.1; 1.05: 0.04251; 1.06: 0.02; 1.07: < 0.03; 1.08: 0.03 |  |  |  | 0.031181
| 1 | 5 | 1.01: 0.03; 1.02: 0.03; 1.03: < 0.1; 1.04: < 0.1; 1.05: 0.04848; 1.06: 0.03; 1.07: < 0.03; 1.08: 0.03 |  |  |  | 7
| 1 | 6 | 1.01: 0.03; 1.02: 0.03; 1.03: < 0.1; 1.04: < 0.1; 1.05: 0.04897; 1.06: 0.02; 1.07: < 0.03; 1.08: 0.03
| Mean |  | 1.01: 0.025; 1.02: 0.03; 1.05: 0.047572; 1.06: 0.023333; 1.08: 0.03
| Median |  | 1.01: 0.025; 1.02: 0.03; 1.05: 0.046975; 1.06: 0.02; 1.08: 0.03
| Std Dev. |  | 1.01: 0.0054772; 1.02: 0; 1.05: 0.0048492; 1.06: 0.005164; 1.08: 0
| Rel.Std.Dev. |  | 1.01: 21.91%; 1.02: 0%; 1.05: 10.19%; 1.06: 22.13%; 1.08: 0%
| PDM3 |  | 1.01: -19.82%; 1.02: -3.788%; 1.05: 52.57%; 1.06: -25.17%; 1.08: -3.788%
| Z-Score (Absolute) |  | 1.01: 0.31; 1.02: 0; 1.03: 1.24; 1.04: 1.24; 1.05: 1.09; 1.06: 0.41; 1.07: 0.93; 1.08: 0
| Analytical results for U in OREAS 753b (Certified Value 1.86 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | 2.04 | 1.7694 | 1.9 | 2.16 | 1.94 | 1.88 | 2.22 | 1.8 | 1.9 | 1.97 | 1.9817 | 2.2 | 1.8 | 1.9 | 1.8 | 1.6 | 1.8 | 1.9 | 1.1262 | 1.79 | 1.57 | 1.9 | 2.34 | 2.38 | 1.73 |  | 1
| 1 | 2 | 2.04 | 1.6261 | 1.7 | 2.18 | 1.98 | 1.83 | 2 | 1.6 | 1.8 | 2.17 | 1.9758 | 2.25 | 1.7 | 1.9 | 1.5 | 1.3 | 2.1 | 1.7 | 1.1813 | 1.79 | 1.58 | 1.7 | 2.29 | 2.09 | 1.55 |  | 18
| 1 | 3 | 1.95 | 1.6196 | 1.7 | 2.18 | 2.1 | 1.9 | 2.07 | 1.9 | 1.7 | 1.86 | 1.9729 | 2.14 | 1.6 | 1.5 | 1.5 | 1.6 | 1.8 | 1.8 | 1.1979 | 1.78 | 1.82 | 1.9 | 2.16 | 2.25 | 1.87 |  | 16
| 1 | 4 | 1.98 | 1.6936 | 1.7 | 2.05 | 2 | 1.83 | 1.98 | 1.8 | 1.8 | 1.96 | 2.1267 | 2.09 | 1.9 | 1.4 | 1.7 | 1.4 | 1.9 | 1.9 | 1.1576 | 1.78 | 1.57 | 1.8 | 2.07 | 2.14 | 1.96 |  | 1.8629
| 1 | 5 | 2.22 | 1.5855 | 1.6 | 2.08 | 1.92 | 1.83 | 2.18 | 1.8 | 1.9 | 2.06 | 1.979 | 2.08 | 1.6 | 1.7 | 1.6 | 1.5 | 1.7 | 1.7 | 1.1774 | 1.79 | 1.7 | 1.4 | 2.08 | 2.24 | 1.67 |  | 58
| 1 | 6 | 2.21 | 1.665 | 1.7 | 2.02 | 2.06 | 1.86 | 2.03 | 1.7 | 1.9 | 2.09 | 2.0467 | 2.35 | 1.6 | 1.5 | 1.6 | 1.5 | 1.7 | 1.8 | 1.1428 | 1.88 | 1.77 | 1.5 | 2.11 | 2.52 | 1.64
| Mean |  | 2.0733 | 1.6599 | 1.7167 | 2.1117 | 2 | 1.855 | 2.08 | 1.7667 | 1.8333 | 2.0183 | 2.0138 | 2.185 | 1.7 | 1.65 | 1.6167 | 1.4833 | 1.8333 | 1.8 | 1.1639 | 1.8017 | 1.6683 | 1.7 | 2.175 | 2.27 | 1.7367
| Median |  | 2.04 | 1.6456 | 1.7 | 2.12 | 1.99 | 1.845 | 2.05 | 1.8 | 1.85 | 2.015 | 1.9803 | 2.17 | 1.65 | 1.6 | 1.6 | 1.5 | 1.8 | 1.8 | 1.1675 | 1.79 | 1.64 | 1.75 | 2.135 | 2.245 | 1.7
| Std Dev. |  | 0.11518 | 0.065514 | 0.098319 | 0.070545 | 0.069282 | 0.030166 | 0.09859 | 0.10328 | 0.08165 | 0.11017 | 0.061965 | 0.10368 | 0.12649 | 0.21679 | 0.1169 | 0.1169 | 0.15055 | 0.089443 | 0.026641 | 0.038687 | 0.11089 | 0.20976 | 0.11397 | 0.15824 | 0.15253
| Rel.Std.Dev. |  | 5.555% | 3.947% | 5.727% | 3.341% | 3.464% | 1.626% | 4.74% | 5.846% | 4.454% | 5.458% | 3.077% | 4.745% | 7.441% | 13.14% | 7.231% | 7.881% | 8.212% | 4.969% | 2.289% | 2.147% | 6.647% | 12.34% | 5.24% | 6.971% | 8.783%
| PDM3 |  | 11.29% | -10.9% | -7.851% | 13.35% | 7.358% | -0.4259% | 11.65% | -5.168% | -1.589% | 8.342% | 8.098% | 17.29% | -8.746% | -11.43% | -13.22% | -20.38% | -1.589% | -3.378% | -37.53% | -3.289% | -10.45% | -8.746% | 16.75% | 21.85% | -6.778%
| Z-Score (Absolute) |  | 1.21 | 0.63 | 0.38 | 1.38 | 0.88 | 0.24 | 1.24 | 0.16 | 0.14 | 0.96 | 0.94 | 1.7 | 0.45 | 0.67 | 0.82 | 1.42 | 0.14 | 0.01 | 2.84 | 0 | 0.59 | 0.45 | 1.66 | 2.08 | 0.29
| Analytical results for V in OREAS 753b (Certified Value 8.11 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.28 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 8 | 8.0644 | 8 | 8 | 9 | 8 | 8 | 9 | 9 | 8.9 | 8.3262 | 7 | 8 | 8 | 8 | 8 | 8 | 7 | 8 | 7.4445 | 8.23 | 9 | < 10 | 8 | 10.02 |  | 1
| 1 | 2 | 8 | 8.2143 | 8 | 8 | 9 | 8 | 8 | 9 | 8 | 8.9 | 8.2175 | 7 | 8 | 8 | 8 | 8 | 9 | 6 | 8 | 7.2555 | 8.08 | 9 | < 10 | 8.1 | 9.49 |  | 0
| 1 | 3 | 7 | 7.9811 | 8 | 8 | 9 | 8 | 8 | 9 | 9 | 8.6 | 8.4126 | 7 | 8 | 8 | 8 | 8 | 8 | 7 | 8 | 7.497 | 7.9 | 9 | < 10 | 7.6 | 9.76 |  | 16
| 1 | 4 | 7 | 7.7362 | 8 | 8 | 9 | 8 | 8 | 9 | 8 | 8.6 | 8.5689 | 7 | 8 | 8 | 8 | 8 | 8 | 6 | 9 | 7.539 | 8.26 | 9 | < 10 | 8.5 | 10.03 |  | 8.1126
| 1 | 5 | 7 | 7.966 | 8 | 8 | 9 | 8 | 8 | 9 | 8 | 8.8 | 8.5905 | 7 | 8 | 8 | 8 | 8 | 8 | 7 | 8 | 7.4445 | 8.16 | 9 | < 10 | 8.1 | 10.28 |  | 59
| 1 | 6 | 8 | 7.804 | 8 | 8 | 9 | 8 | 8 | 9 | 9 | 8.9 | 8.6432 | 7 | 8 | 8 | 8 | 8 | 8 | 6 | 9 | 7.3605 | 8.04 | 8 | < 10 | 8.1 | 9.17
| Mean |  | 7.5 | 7.961 | 8 | 8 | 9 | 8 | 8 | 9 | 8.5 | 8.7833 | 8.4598 | 7 | 8 | 8 | 8 | 8 | 8.1667 | 6.5 | 8.3333 | 7.4235 | 8.1117 | 8.8333 |  | 8.0667 | 9.7917
| Median |  | 7.5 | 7.9735 | 8 | 8 | 9 | 8 | 8 | 9 | 8.5 | 8.85 | 8.4908 | 7 | 8 | 8 | 8 | 8 | 8 | 6.5 | 8 | 7.4445 | 8.12 | 9 |  | 8.1 | 9.89
| Std Dev. |  | 0.54772 | 0.17348 | 0 | 0 | 0 | 0 | 0 | 0 | 0.54772 | 0.1472 | 0.16819 | 0 | 0 | 0 | 0 | 0 | 0.40825 | 0.54772 | 0.5164 | 0.1018 | 0.13363 | 0.40825 |  | 0.28752 | 0.40642
| Rel.Std.Dev. |  | 7.303% | 2.179% | 0% | 0% | 0% | 0% | 0% | 0% | 6.444% | 1.676% | 1.988% | 0% | 0% | 0% | 0% | 0% | 4.999% | 8.427% | 6.197% | 1.371% | 1.647% | 4.622% |  | 3.564% | 4.151%
| PDM3 |  | -7.551% | -1.869% | -1.388% | -1.388% | 10.94% | -1.388% | -1.388% | 10.94% | 4.775% | 8.268% | 4.28% | -13.71% | -1.388% | -1.388% | -1.388% | -1.388% | 0.6665% | -19.88% | 2.721% | -8.494% | -0.01145% | 8.884% |  | -0.5661% | 20.7%
| Z-Score (Absolute) |  | 1.01 | 0.08 | 0 | 0 | 2.02 | 0 | 0 | 2.02 | 1.01 | 1.58 | 0.93 | 2.02 | 0 | 0 | 0 | 0 | 0.34 | 3.03 | 0.67 | 1.17 | 0.23 | 1.69 | 6.07 | 0.13 | 3.62
| Analytical results for W in OREAS 753b (Certified Value 3.37 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | 3.7 | 3.5369 | 3.5 | 2.6 | 3.5 | 3.3 | 2.1 | 3.3 | 1.7 | 2.6 | 3.4236 | 3.5 | 3.6 | 3.5 | 4.1 | 3.5 | 3.4 | 3.2 | 3.1 | 3.48 | 3.4 | 3.3 | 4.5 | 4.5 | 4.4 |  | 1
| 1 | 2 | 3.3 | 3.5217 | 3.7 | 3.1 | 3.6 | 3.2 | 2.2 | 3.5 | 1.5 | 3 | 3.133 | 3.4 | 3.5 | 3.7 | 3.3 | 3.6 | 3.7 | 3 | 3.3 | 3.53 | 2.8 | 2.9 | 5 | 3.9 | 4.33 |  | 19
| 1 | 3 | 3.6 | 3.4683 | 3.6 | 2.7 | 3.5 | 3.3 | 2.5 | 3.5 | 1.1 | 3.1 | 3.3147 | 3.7 | 3.2 | 3.9 | 4.1 | 3.6 | 4.6 | 3.2 | 2.9 | 3.28 | 3.2 | 3.5 | 4.7 | 3.8 | 4.94 |  | 16
| 1 | 4 | 3.5 | 3.3457 | 3.7 | 2.6 | 3.4 | 3.4 | 2.1 | 3.4 | 1.4 | 2.8 | 3.389 | 3.5 | 3.9 | 3.7 | 3.4 | 4.3 | 3.7 | 3.1 | 3.3 | 3.52 | 3 | 3.1 | 5 | 3.2 | 4.5 |  | 3.3661
| 1 | 5 | 3.5 | 3.3176 | 3.9 | 2.9 | 3.9 | 3.2 | 2.2 | 3.5 | 1.4 | 2.6 | 3.1937 | 3.7 | 3.7 | 3.1 | 3.3 | 3.3 | 4.1 | 3 | 3 | 3.48 | 3.1 | 2.9 | 4.9 | 3 | 3.88 |  | 60
| 1 | 6 | 3.4 | 3.257 | 3.6 | 2.7 | 3.6 | 3.2 | 2.5 | 3.3 | 1.4 | 2.8 | 3.3204 | 3.5 | 3.7 | 3.7 | 3.1 | 3.6 | 4.1 | 3.1 | 3 | 3.3 | 2.7 | 2.3 | 4.9 | 3.1 | 4.63
| Mean |  | 3.5 | 3.4079 | 3.6667 | 2.7667 | 3.5833 | 3.2667 | 2.2667 | 3.4167 | 1.4167 | 2.8167 | 3.2957 | 3.55 | 3.6 | 3.6 | 3.55 | 3.65 | 3.9333 | 3.1 | 3.1 | 3.4317 | 3.0333 | 3 | 4.8333 | 3.5833 | 4.4467
| Median |  | 3.5 | 3.407 | 3.65 | 2.7 | 3.55 | 3.25 | 2.2 | 3.45 | 1.4 | 2.8 | 3.3176 | 3.5 | 3.65 | 3.7 | 3.35 | 3.6 | 3.9 | 3.1 | 3.05 | 3.48 | 3.05 | 3 | 4.9 | 3.5 | 4.45
| Std Dev. |  | 0.14142 | 0.11665 | 0.13663 | 0.19664 | 0.17224 | 0.08165 | 0.18619 | 0.098319 | 0.19408 | 0.20412 | 0.11216 | 0.12247 | 0.23664 | 0.27568 | 0.43704 | 0.33912 | 0.42269 | 0.089443 | 0.16733 | 0.11179 | 0.2582 | 0.41473 | 0.19664 | 0.58452 | 0.35121
| Rel.Std.Dev. |  | 4.041% | 3.423% | 3.726% | 7.107% | 4.807% | 2.499% | 8.214% | 2.878% | 13.7% | 7.247% | 3.403% | 3.45% | 6.573% | 7.658% | 12.31% | 9.291% | 10.75% | 2.885% | 5.398% | 3.258% | 8.512% | 13.82% | 4.068% | 16.31% | 7.898%
| PDM3 |  | 3.977% | 1.24% | 8.928% | -17.81% | 6.453% | -2.955% | -32.66% | 1.502% | -57.91% | -16.32% | -2.091% | 5.463% | 6.948% | 6.948% | 5.463% | 8.433% | 16.85% | -7.906% | -7.906% | 1.947% | -9.886% | -10.88% | 43.59% | 6.453% | 32.1%
| Z-Score (Absolute) |  | 0.21 | 0.07 | 0.73 | 2.05 | 0.47 | 0.51 | 3.6 | 0.05 | 6.22 | 1.9 | 0.42 | 0.37 | 0.52 | 0.52 | 0.37 | 0.67 | 1.55 | 1.02 | 1.02 | 0 | 1.23 | 1.33 | 4.33 | 0.47 | 3.13
| Analytical results for Y in OREAS 753b (Certified Value 1.97 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
2
| 1 | 1 | 2.23 | 1.883 | 2 | 2 | 2.22 | 2.1 | 2.2 | 1.7 | 2.5 | 2.2 | 3.2985 | 1.98 | 1.9 | 2 | 1.9 | 1.6 | 2 | 1.2 | 1.6 | 2.1671 | 1.98 | 1.61 | 2 | 2.8 | 2.4 | 1.16 | 1
| 1 | 2 | 2.12 | 1.8542 | 1.8 | 2 | 2.09 | 2 | 2.2 | 1.4 | 2.4 | 2.5 | 3.8146 | 1.95 | 1.8 | 2 | 1.8 | 1.6 | 2.2 | 1.1 | 1.5 | 2.1167 | 1.98 | 1.49 | 1.9 | 2.8 | 2.1 | 1.15 | 20
| 1 | 3 | 2.11 | 1.8663 | 2 | 1.9 | 2.13 | 2.1 | 2.3 | 2.1 | 2.3 | 2.1 | 3.3853 | 2 | 1.9 | 1.8 | 1.8 | 1.7 | 2 | 1.2 | 1.5 | 2.1731 | 1.97 | 1.58 | 2 | 2.6 | 2.1 | 1.62 | 16
| 1 | 4 | 1.99 | 1.9417 | 1.9 | 2 | 2.09 | 2 | 2.3 | 1.7 | 2.3 | 2.2 | 3.1978 | 2.01 | 2 | 1.4 | 2 | 1.6 | 2 | 1 | 1.7 | 2.1683 | 2.02 | 1.65 | 2 | 2.7 | 2.1 | 2.29 | 1.9698
| 1 | 5 | 2.27 | 1.9034 | 1.8 | 2 | 2.17 | 2 | 2.2 | 1.8 | 2.6 | 2.2 | 3.3868 | 2.04 | 1.8 | 1.6 | 1.9 | 1.7 | 1.8 | 1.1 | 1.7 | 2.1283 | 2.08 | 2 | 1.6 | 2.6 | 2.1 | 2.24 | 61
| 1 | 6 | 2.27 | 1.9218 | 1.8 | 2.1 | 2.05 | 2 | 2.3 | 1.7 | 2.5 | 2.3 | 3.5256 | 2.02 | 1.8 | 1.6 | 1.8 | 1.7 | 1.9 | 1.2 | 1.5 | 2.1234 | 2.06 | 1.92 | 1.8 | 2.6 | 2.1 | 1.07
| Mean |  | 2.165 | 1.895 | 1.8833 | 2 | 2.125 | 2.0333 | 2.25 | 1.7333 | 2.4333 | 2.25 | 3.4348 | 2 | 1.8667 | 1.7333 | 1.8667 | 1.65 | 1.9833 | 1.1333 | 1.5833 | 2.1461 | 2.015 | 1.7083 | 1.8833 | 2.6833 | 2.15 | 1.5883
| Median |  | 2.175 | 1.8932 | 1.85 | 2 | 2.11 | 2 | 2.25 | 1.7 | 2.45 | 2.2 | 3.386 | 2.005 | 1.85 | 1.7 | 1.85 | 1.65 | 2 | 1.15 | 1.55 | 2.1477 | 2 | 1.63 | 1.95 | 2.65 | 2.1 | 1.39
| Std Dev. |  | 0.11131 | 0.033462 | 0.098319 | 0.063246 | 0.061887 | 0.05164 | 0.054772 | 0.22509 | 0.12111 | 0.13784 | 0.2154 | 0.031623 | 0.08165 | 0.24221 | 0.08165 | 0.054772 | 0.13292 | 0.08165 | 0.098319 | 0.025931 | 0.046368 | 0.20351 | 0.16021 | 0.098319 | 0.12247 | 0.55898
| Rel.Std.Dev. |  | 5.141% | 1.766% | 5.22% | 3.162% | 2.912% | 2.54% | 2.434% | 12.99% | 4.977% | 6.126% | 6.271% | 1.581% | 4.374% | 13.97% | 4.374% | 3.32% | 6.702% | 7.204% | 6.21% | 1.208% | 2.301% | 11.91% | 8.507% | 3.664% | 5.696% | 35.19%
| PDM3 |  | 9.912% | -3.793% | -4.388% | 1.535% | 7.881% | 3.227% | 14.23% | -12% | 23.53% | 14.23% | 74.38% | 1.535% | -5.234% | -12% | -5.234% | -16.23% | 0.689% | -42.46% | -19.62% | 8.954% | 2.297% | -13.27% | -4.388% | 36.23% | 9.15% | -19.36%
| Z-Score (Absolute) |  | 0.7 | 0.39 | 0.44 | 0.03 | 0.54 | 0.17 | 1.05 | 1.05 | 1.8 | 1.05 | 5.87 | 0.03 | 0.51 | 1.05 | 0.51 | 1.39 | 0.03 | 3.49 | 1.66 | 0.63 | 0.09 | 1.15 | 0.44 | 2.81 | 0.64 | 1.64
| Analytical results for Yb in OREAS 753b (Certified Value 0.18 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.04; 1.03: 1.05; 1.04: 1.08; 1.05: 1.09; 1.06: 1.11; 1.07: 1.12; 1.08: 1.18; 1.09: 1.21; 1.10: 1.22; 1.11: 1.24; 1.12: 1.25; 1.13: 1.28 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*MS; 1.13: 4A*MS | 2
2
| 1 | 1 | 1.01: 0.18; 1.02: 0.2; 1.03: 0.19; 1.04: 0.2; 1.05: 0.2; 1.06: 0.2976; 1.07: 0.18; 1.08: 0.2; 1.09: 0.18; 1.10: 0.1; 1.11: 0.19; 1.12: 0.2; 1.13: < 0.5 | 1
| 1 | 2 | 1.01: 0.17; 1.02: 0.2; 1.03: 0.19; 1.04: 0.1; 1.05: 0.2; 1.06: 0.3537; 1.07: 0.18; 1.08: 0.1; 1.09: 0.17; 1.10: 0.1; 1.11: 0.2; 1.12: 0.2; 1.13: < 0.5 | 21
| 1 | 3 | 1.01: 0.18; 1.02: 0.2; 1.03: 0.17; 1.04: 0.2; 1.05: 0.2; 1.06: 0.30075; 1.07: 0.15; 1.08: 0.2; 1.09: 0.17; 1.10: 0.1; 1.11: 0.21; 1.12: 0.2; 1.13: < 0.5 | 16
| 1 | 4 | 1.01: 0.18; 1.02: 0.2; 1.03: 0.17; 1.04: 0.1; 1.05: 0.2; 1.06: 0.306; 1.07: 0.15; 1.08: 0.1; 1.09: 0.18; 1.10: 0.1; 1.11: 0.21; 1.12: 0.2; 1.13: < 0.5 | 0.178
| 1 | 5 | 1.01: 0.18; 1.02: 0.2; 1.03: 0.19; 1.04: 0.1; 1.05: 0.2; 1.06: 0.30039; 1.07: 0.14; 1.08: 0.2; 1.09: 0.17; 1.10: 0.1; 1.11: 0.2; 1.12: 0.2; 1.13: < 0.5 | 62
| 1 | 6 | 1.01: 0.19; 1.02: 0.2; 1.03: 0.18; 1.04: 0.1; 1.05: 0.2; 1.06: 0.33634; 1.07: 0.15; 1.08: 0.1; 1.09: 0.17; 1.10: 0.1; 1.11: 0.21; 1.12: 0.2; 1.13: < 0.5
| Mean |  | 1.01: 0.18; 1.02: 0.2; 1.03: 0.18167; 1.04: 0.13333; 1.05: 0.2; 1.06: 0.3158; 1.07: 0.15833; 1.08: 0.15; 1.09: 0.17333; 1.10: 0.1; 1.11: 0.20333; 1.12: 0.2
| Median |  | 1.01: 0.18; 1.02: 0.2; 1.03: 0.185; 1.04: 0.1; 1.05: 0.2; 1.06: 0.30338; 1.07: 0.15; 1.08: 0.15; 1.09: 0.17; 1.10: 0.1; 1.11: 0.205; 1.12: 0.2
| Std Dev. |  | 1.01: 0.0063246; 1.02: 3.0405e-17; 1.03: 0.0098319; 1.04: 0.05164; 1.05: 3.0405e-17; 1.06: 0.02345; 1.07: 0.017224; 1.08: 0.054772; 1.09: 0.005164; 1.10: 1.5202e-17; 1.11: 0.008165; 1.12: 3.0405e-17
| Rel.Std.Dev. |  | 1.01: 3.514%; 1.02: 1.52e-14%; 1.03: 5.412%; 1.04: 38.73%; 1.05: 1.52e-14%; 1.06: 7.426%; 1.07: 10.88%; 1.08: 36.51%; 1.09: 2.979%; 1.10: 1.52e-14%; 1.11: 4.016%; 1.12: 1.52e-14%
| PDM3 |  | 1.01: 1.124%; 1.02: 12.36%; 1.03: 2.06%; 1.04: -25.09%; 1.05: 12.36%; 1.06: 77.41%; 1.07: -11.05%; 1.08: -15.73%; 1.09: -2.622%; 1.10: -43.82%; 1.11: 14.23%; 1.12: 12.36%
| Z-Score (Absolute) |  | 1.01: 0.05; 1.02: 0.57; 1.03: 0; 1.04: 1.5; 1.05: 0.57; 1.06: 4.17; 1.07: 0.73; 1.08: 0.99; 1.09: 0.26; 1.10: 2.54; 1.11: 0.67; 1.12: 0.57; 1.13: 2.13
| Analytical results for Zn in OREAS 753b (Certified Value 169 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 171 | 165.467 | 178 | 168 | 167 | 160 | 167 | 176.6 | 186 | 175 | 158.5043 | 126 | 173 | 158 | 161 | 168 | 181 | 145 | 174 | 154.915 | 166.08 | 181 | 161 | 171 | 182 | 173.43 | 1
| 1 | 2 | 169 | 169.3298 | 179 | 169 | 168 | 159 | 168 | 172.8 | 185 | 173 | 159.8824 | 126 | 171 | 160 | 157 | 169 | 182 | 152 | 178 | 157.335 | 167.24 | 183 | 161 | 170 | 179 | 172.04 | 0
| 1 | 3 | 170 | 165.494 | 180 | 169 | 167 | 155 | 168 | 178.7 | 182 | 168 | 161.2687 | 124 | 170 | 172 | 158 | 170 | 178 | 147 | 176 | 152.241 | 162 | 183 | 160 | 173 | 175 | 170.26 | 16
| 1 | 4 | 172 | 161.0839 | 176 | 169 | 168 | 169 | 167 | 182.2 | 184 | 173 | 163.0973 | 130 | 170 | 166 | 162 | 169 | 180 | 149 | 174 | 158.391 | 167.93 | 180 | 157 | 170 | 181 | 184.92 | 169.4825
| 1 | 5 | 171 | 164.5505 | 177 | 170 | 164 | 162 | 167 | 180.9 | 183 | 172 | 162.8672 | 129 | 170 | 166 | 160 | 167 | 175 | 152 | 179 | 155.234 | 164.8 | 182 | 178 | 168 | 176 | 192.67 | 63
| 1 | 6 | 171 | 159.8452 | 177 | 171 | 164 | 161 | 167 | 183.5 | 190 | 179 | 162.3865 | 124 | 172 | 168 | 161 | 175 | 182 | 146 | 177 | 159.414 | 164.47 | 179 | 180 | 168 | 180 | 172.06
| Mean |  | 170.6667 | 164.2951 | 177.8333 | 169.3333 | 166.3333 | 161 | 167.3333 | 179.1167 | 185 | 173.3333 | 161.3344 | 126.5 | 171 | 165 | 159.8333 | 169.6667 | 179.6667 | 148.5 | 176.3333 | 156.255 | 165.42 | 181.3333 | 166.1667 | 170 | 178.8333 | 177.5633
| Median |  | 171 | 165.0088 | 177.5 | 169 | 167 | 160.5 | 167 | 179.8 | 184.5 | 173 | 161.8276 | 126 | 170.5 | 166 | 160.5 | 169 | 180.5 | 148 | 176.5 | 156.2845 | 165.44 | 181.5 | 161 | 170 | 179.5 | 172.745
| Std Dev. |  | 1.0328 | 3.4158 | 1.472 | 1.0328 | 1.8619 | 4.6043 | 0.5164 | 3.9575 | 2.8284 | 3.6148 | 1.8268 | 2.51 | 1.2649 | 5.1769 | 1.9408 | 2.8048 | 2.7325 | 3.0166 | 2.0656 | 2.6328 | 2.1462 | 1.633 | 10.0681 | 1.8974 | 2.7869 | 9.0945
| Rel.Std.Dev. |  | 0.6052% | 2.079% | 0.8277% | 0.6099% | 1.119% | 2.86% | 0.3086% | 2.209% | 1.529% | 2.085% | 1.132% | 1.984% | 0.7397% | 3.137% | 1.214% | 1.653% | 1.521% | 2.031% | 1.171% | 1.685% | 1.297% | 0.9005% | 6.059% | 1.116% | 1.558% | 5.122%
| PDM3 |  | 0.6987% | -3.061% | 4.927% | -0.08803% | -1.858% | -5.005% | -1.268% | 5.684% | 9.156% | 2.272% | -4.808% | -25.36% | 0.8954% | -2.645% | -5.693% | 0.1086% | 6.009% | -12.38% | 4.042% | -7.805% | -2.397% | 6.992% | -1.956% | 0.3053% | 5.517% | 4.768%
| Z-Score (Absolute) |  | 0.11 | 0.47 | 0.75 | 0.02 | 0.29 | 0.77 | 0.2 | 0.87 | 1.4 | 0.35 | 0.74 | 3.89 | 0.14 | 0.41 | 0.88 | 0.02 | 0.92 | 1.9 | 0.62 | 1.2 | 0.37 | 1.07 | 0.3 | 0.05 | 0.84 | 0.73
| Analytical results for Zr in OREAS 753b (Certified Value 16.2 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.28 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 16.4 | 16.2507 | 17.9 | 15 | 15.1 | 16.4 | 13.6 | 14.3 | 20 | 16.3 | 41.3759 | 15 | 16.2 | 15.2 | 17.6 | 16.5 | 16.2 | 28.1 | 11.4 | 14.4165 | 16.72 | 14.6 | 15.4 | 16 | 28.9 | 20.78 | 1
| 1 | 2 | 16.9 | 16.713 | 17 | 15.5 | 15.3 | 16.7 | 14.4 | 13.5 | 19 | 16.9 | 39.5025 | 15 | 16.6 | 17 | 16.2 | 16 | 17.8 | 26.8 | 11.6 | 14.4585 | 16.67 | 15.4 | 16 | 17 | 23 | 17.97 | 8
| 1 | 3 | 15.9 | 16.3731 | 16.8 | 14.6 | 16 | 16.8 | 15.4 | 15.9 | 19 | 16.2 | 40.3283 | 14 | 16.2 | 16.8 | 16.5 | 16.2 | 15.6 | 28.5 | 11.4 | 14.7945 | 15.9 | 15.5 | 15.5 | 18 | 29.1 | 22.17 | 16
| 1 | 4 | 16 | 16.3398 | 16.9 | 15.1 | 14.7 | 16.9 | 13 | 14.9 | 21 | 15.7 | 42.4787 | 14 | 17.3 | 15 | 16.6 | 16.9 | 15.7 | 27.9 | 12.3 | 14.7 | 17.18 | 15.5 | 16.4 | 18 | 26.4 | 20.86 | 16.2335
| 1 | 5 | 17.2 | 16.3836 | 17 | 15.3 | 15.6 | 16.7 | 13.9 | 14.2 | 19 | 16.6 | 41.0481 | 15 | 16.4 | 15.5 | 17.2 | 16.5 | 14.5 | 26.7 | 11.4 | 14.0595 | 16.78 | 14.6 | 14.1 | 18 | 30.4 | 19.52 | 64
| 1 | 6 | 16.6 | 15.7484 | 17.8 | 15.9 | 15.2 | 16.8 | 14.9 | 13.9 | 21 | 16.8 | 39.9006 | 14 | 16.5 | 15.6 | 18.8 | 16.9 | 15.6 | 28.4 | 12.4 | 14.616 | 17.18 | 14.6 | 14.5 | 18 | 26.2 | 20.02
| Mean |  | 16.5 | 16.3014 | 17.2333 | 15.2333 | 15.3167 | 16.7167 | 14.2 | 14.45 | 19.8333 | 16.4167 | 40.7723 | 14.5 | 16.5333 | 15.85 | 17.15 | 16.5 | 15.9 | 27.7333 | 11.75 | 14.5075 | 16.7383 | 15.0333 | 15.3167 | 17.5 | 27.3333 | 20.22
| Median |  | 16.5 | 16.3564 | 17 | 15.2 | 15.25 | 16.75 | 14.15 | 14.25 | 19.5 | 16.45 | 40.6882 | 14.5 | 16.45 | 15.55 | 16.9 | 16.5 | 15.65 | 28 | 11.5 | 14.5373 | 16.75 | 15 | 15.45 | 18 | 27.65 | 20.4
| Std Dev. |  | 0.50596 | 0.31343 | 0.48442 | 0.44572 | 0.4446 | 0.17224 | 0.87864 | 0.84794 | 0.98319 | 0.4446 | 1.0883 | 0.54772 | 0.40825 | 0.84321 | 0.95446 | 0.36332 | 1.0844 | 0.79162 | 0.47223 | 0.26176 | 0.46897 | 0.4761 | 0.87502 | 0.83666 | 2.68 | 1.4225
| Rel.Std.Dev. |  | 3.066% | 1.923% | 2.811% | 2.926% | 2.903% | 1.03% | 6.188% | 5.868% | 4.957% | 2.708% | 2.669% | 3.777% | 2.469% | 5.32% | 5.565% | 2.202% | 6.82% | 2.854% | 4.019% | 1.804% | 2.802% | 3.167% | 5.713% | 4.781% | 9.805% | 7.035%
| PDM3 |  | 1.641% | 0.4182% | 6.159% | -6.161% | -5.648% | 2.976% | -12.53% | -10.99% | 22.18% | 1.128% | 151.2% | -10.68% | 1.847% | -2.363% | 5.645% | 1.641% | -2.055% | 70.84% | -27.62% | -10.63% | 3.11% | -7.393% | -5.648% | 7.802% | 68.38% | 24.56%
| Z-Score (Absolute) |  | 0.02 | 0.09 | 0.46 | 0.72 | 0.67 | 0.15 | 1.33 | 1.19 | 1.99 | 0.02 | 14.36 | 1.16 | 0.04 | 0.36 | 0.41 | 0.02 | 0.33 | 6.66 | 2.78 | 1.15 | 0.17 | 0.84 | 0.67 | 0.62 | 6.42 | 2.22
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C294: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C894: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C912: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C930: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 753b (Indicative Value < 10 ppm) | [25 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.03 | 1.04 | 1.12 | 1.22 | 1.28 | [19 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | [19 empty] | 1
1
| 1 | 1 | < 10 | < 1 | 0.6 | < 10 | < 0.5 | [19 empty] | 1
| 1 | 2 | < 10 | < 1 | 0.7 | < 10 | < 0.5 | [19 empty] | 4
| 1 | 3 | < 10 | < 1 | 0.4 | < 10 | < 0.5 | [19 empty] | 16
| 1 | 4 | < 10 | < 1 | 1 | < 10 | < 0.5 | [19 empty] | < 10
| 1 | 5 | < 10 | < 1 | 1.2 | < 10 | < 0.5 | [19 empty] | 10
| 1 | 6 | < 10 | < 1 | 0.5 | < 10 | < 0.5
| Mean |  |  |  | 0.73333
| Median |  |  |  | 0.65
| Std Dev. |  |  |  | 0.30768
| Rel.Std.Dev. |  |  |  | 41.96%
PDM3
| Z-Score (Absolute) |  | 5.95 | 0.33 | 0 | 5.95 | 0.67
| Analytical results for Al in OREAS 753b (Certified Value 6.3 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 2
3
| 1 | 1 | 6.32 | 6.2951 | 6.4 | 6.47 | 6.39 | 6.14 | 6.21 | 6.48 | 6.24 | 6.32 | 5.2924 | 6.166 | 5.98 | 6.219 | 6.219 | 6.139 | 6.05 | 6.46 | 6.3379 | 5.8 | 6.13 | 5.6 | 6.2724 | 5.66 | 1
| 1 | 2 | 6.22 | 6.3519 | 6.32 | 6.4 | 6.37 | 6.15 | 6.24 | 6.46 | 6.31 | 6.32 | 3.1119 | 6.086 | 6.298 | 6.272 | 6.298 | 6.192 | 6.2 | 6.51 | 6.3484 | 6.43 | 6.17 | 5.18 | 6.2468 | 5.64 | 0
| 1 | 3 | 6.48 | 6.2089 | 6.36 | 6.38 | 6.26 | 6.21 | 6.31 | 6.51 | 6.27 | 6.37 | 3.9428 | 6.219 | 6.219 | 6.192 | 6.298 | 6.139 | 6.06 | 6.41 | 6.3128 | 6.36 | 6.06 | 5.49 | 6.34 | 5.55 | 16
| 1 | 4 | 6.33 | 6.3639 | 6.39 | 6.44 | 6.51 | 6.11 | 6.17 | 6.34 | 6.39 | 6.37 | 4.6944 | 6.272 | 6.272 | 6.324 | 6.219 | 6.219 | 6.22 | 6.52 | 6.3355 | 6.56 | 6.14 | 5.26 | 6.2928 | 5.09 | 6.2977
| 1 | 5 | 6.31 | 6.1168 | 6.26 | 6.32 | 6.43 | 6.13 | 6.23 | 6.17 | 6.32 | 6.36 | 7.3564 | 6.166 | 6.298 | 6.192 | 6.245 | 6.272 | 6.08 | 6.47 | 6.3195 | 6.24 | 6.19 | 5.37 | 6.3347 | 5.73 | 66
| 1 | 6 | 6.33 | 6.3044 | 6.28 | 6.31 | 6.39 | 6.16 | 6.31 | 6.57 | 6.25 | 6.41 | 5.7687 | 6.351 | 6.298 | 6.351 | 6.324 | 6.324 | 6.1 | 6.47 | 6.3298 | 6.38 | 6.33 | 5.17 | 6.3045 | 6.12
| Mean |  | 6.3317 | 6.2735 | 6.335 | 6.3867 | 6.3917 | 6.15 | 6.245 | 6.4217 | 6.2967 | 6.3583 | 5.0278 | 6.21 | 6.2275 | 6.2583 | 6.2672 | 6.2142 | 6.1183 | 6.4733 | 6.3306 | 6.295 | 6.17 | 5.345 | 6.2985 | 5.6317
| Median |  | 6.325 | 6.2998 | 6.34 | 6.39 | 6.39 | 6.145 | 6.235 | 6.47 | 6.29 | 6.365 | 4.9934 | 6.1925 | 6.285 | 6.2455 | 6.2715 | 6.2055 | 6.09 | 6.47 | 6.3326 | 6.37 | 6.155 | 5.315 | 6.2987 | 5.65
| Std Dev. |  | 0.083766 | 0.094316 | 0.057879 | 0.06377 | 0.081588 | 0.034059 | 0.055767 | 0.1447 | 0.055737 | 0.034303 | 1.4833 | 0.092794 | 0.12508 | 0.068459 | 0.045305 | 0.073787 | 0.073326 | 0.039328 | 0.012926 | 0.2638 | 0.090111 | 0.17421 | 0.035946 | 0.33139
| Rel.Std.Dev. |  | 1.323% | 1.503% | 0.9136% | 0.9985% | 1.276% | 0.5538% | 0.893% | 2.253% | 0.8852% | 0.5395% | 29.5% | 1.494% | 2.008% | 1.094% | 0.7229% | 1.187% | 1.198% | 0.6075% | 0.2042% | 4.191% | 1.46% | 3.259% | 0.5707% | 5.884%
| PDM3 |  | 0.5396% | -0.3842% | 0.5925% | 1.413% | 1.492% | -2.345% | -0.8366% | 1.969% | -0.01619% | 0.963% | -20.16% | -1.392% | -1.114% | -0.6249% | -0.4846% | -1.326% | -2.848% | 2.789% | 0.5233% | -0.04265% | -2.028% | -15.13% | 0.01352% | -10.58%
| Z-Score (Absolute) |  | 0.66 | 0.03 | 0.69 | 1.24 | 1.3 | 1.29 | 0.27 | 1.62 | 0.28 | 0.94 | 13.29 | 0.65 | 0.46 | 0.13 | 0.03 | 0.6 | 1.63 | 2.17 | 0.65 | 0.26 | 1.07 | 9.9 | 0.3 | 6.83
| Analytical results for As in OREAS 753b (Indicative Value 4.03 ppm) | [25 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | [5 empty] | 2
2
| 1 | 1 | < 10 | < 5 | < 20 | < 20 | < 100 | < 5 | 5 | 4 | < 100 | < 100 | < 100 | < 100 | 100 | < 30 | 7 | < 10 | < 100 | 5.9054 | 3.4 | [5 empty] | 1
| 1 | 2 | < 10 | 5 | < 20 | < 20 | < 100 | < 5 | < 5 | 4 | < 100 | 100 | < 100 | < 100 | 100 | < 30 | 8 | < 10 | < 100 | < 5 | 3.5 | [5 empty] | 5
| 1 | 3 | < 10 | < 5 | < 20 | < 20 | < 100 | 5 | < 5 | 4 | < 100 | < 100 | < 100 | < 100 | 100 | < 30 | 8 | < 10 | < 100 | < 5 | 3.4 | [5 empty] | 16
| 1 | 4 | < 10 | < 5 | < 20 | < 20 | < 100 | 6 | < 5 | 5 | < 100 | < 100 | < 100 | < 100 | 100 | < 30 | 8 | < 10 | < 100 | < 5 | 3.6 | [5 empty] | 4.0285
| 1 | 5 | < 10 | < 5 | < 20 | < 20 | < 100 | 7 | < 5 | 5 | < 100 | 100 | 100 | < 100 | < 100 | < 30 | 8 | < 10 | < 100 | < 5 | 3.8 | [5 empty] | 11
| 1 | 6 | < 10 | < 5 | < 20 | < 20 | < 100 | < 5 | < 5 | 4 | 100 | 200 | 100 | < 100 | 100 | < 30 | 9 | < 10 | < 100 | 5.2485 | 3.7
| Mean |  | 1.04: 5; 1.09: 6; 1.10: 5; 1.12: 4.3333; 1.13: 100; 1.14: 133.3333; 1.15: 100; 1.17: 100; 1.19: 8; 1.27: 5.577; 1.28: 3.5667
| Median |  | 1.04: 5; 1.09: 6; 1.10: 5; 1.12: 4; 1.13: 100; 1.14: 100; 1.15: 100; 1.17: 100; 1.19: 8; 1.27: 5.577; 1.28: 3.55
| Std Dev. |  | 1.09: 1; 1.12: 0.5164; 1.14: 57.735; 1.15: 0; 1.17: 0; 1.19: 0.63246; 1.27: 0.46446; 1.28: 0.1633
| Rel.Std.Dev. |  | 1.09: 16.67%; 1.12: 11.92%; 1.14: 43.3%; 1.15: 0%; 1.17: 0%; 1.19: 7.906%; 1.27: 8.328%; 1.28: 4.578%
| PDM3 |  | 1.04: 24.11%; 1.09: 48.94%; 1.10: 24.11%; 1.12: 7.566%; 1.13: 2382%; 1.14: 3210%; 1.15: 2382%; 1.17: 2382%; 1.19: 98.58%; 1.27: 38.44%; 1.28: -11.46%
| Z-Score (Absolute) |  | 0.52 | 0.74 | 0 | 0 | 4.17 | 0.6 | 0.74 | 0.59 | 5.03 | 8.51 | 5.9 | 4.17 | 8.51 | 0.52 | 0.21 | 0.52 | 4.17 | 0.67 | 0.67
| Analytical results for B in OREAS 753b (Certified Value 80 ppm) | [25 empty] | Certified
B | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab | [5 empty] | 1
| No | No | 1.03: 1.01; 1.04: 1.05; 1.06: 1.07; 1.07: 1.09; 1.08: 1.10; 1.09: 1.12; 1.10: 1.18; 1.12: 1.19; 1.13: 1.22 | [5 empty] | ppm
|  |  | 1.03: PF*OES; 1.04: PF*OES; 1.06: PF*OES; 1.07: PF*OES/MS; 1.08: PF*OES/MS; 1.09: PF*OES; 1.10: PF*OES; 1.12: PF*OES/MS; 1.13: PF*OES | [5 empty] | 0
0
| 1 | 1 | 1.03: 90; 1.04: 78; 1.06: 72; 1.07: 90; 1.08: 88; 1.09: 1930; 1.10: 60; 1.12: 78; 1.13: 66 | [5 empty] | 1
| 1 | 2 | 1.03: 91; 1.04: 80; 1.06: 98; 1.07: 100; 1.08: 96; 1.09: 1890; 1.10: 50; 1.12: 76; 1.13: 70 | [5 empty] | 1
| 1 | 3 | 1.03: 85; 1.04: 81; 1.06: 78; 1.07: 100; 1.08: 94; 1.09: 1540; 1.10: 60; 1.12: 75; 1.13: 69 | [5 empty] | 16
| 1 | 4 | 1.03: 89; 1.04: 81; 1.06: 97; 1.07: 90; 1.08: 86; 1.09: 1420; 1.10: 50; 1.12: 76; 1.13: 66 | [5 empty] | 80.2083
| 1 | 5 | 1.03: 89; 1.04: 76; 1.06: 99; 1.07: 90; 1.08: 93; 1.09: 1290; 1.10: 50; 1.12: 76; 1.13: 64 | [5 empty] | 67
| 1 | 6 | 1.03: 87; 1.04: 75; 1.06: 84; 1.07: 100; 1.08: 95; 1.09: 552; 1.10: 60; 1.12: 77; 1.13: 75
| Mean |  | 1.03: 88.5; 1.04: 78.5; 1.06: 88; 1.07: 95; 1.08: 92; 1.09: 1437; 1.10: 55; 1.12: 76.3333; 1.13: 68.3333
| Median |  | 1.03: 89; 1.04: 79; 1.06: 90.5; 1.07: 95; 1.08: 93.5; 1.09: 1480; 1.10: 55; 1.12: 76; 1.13: 67.5
| Std Dev. |  | 1.03: 2.1679; 1.04: 2.5884; 1.06: 11.6103; 1.07: 5.4772; 1.08: 4.0497; 1.09: 502.7902; 1.10: 5.4772; 1.12: 1.0328; 1.13: 3.9328
| Rel.Std.Dev. |  | 1.03: 2.45%; 1.04: 3.297%; 1.06: 13.19%; 1.07: 5.766%; 1.08: 4.402%; 1.09: 34.99%; 1.10: 9.959%; 1.12: 1.353%; 1.13: 5.755%
| PDM3 |  | 1.03: 10.34%; 1.04: -2.13%; 1.06: 9.714%; 1.07: 18.44%; 1.08: 14.7%; 1.09: 1692%; 1.10: -31.43%; 1.12: -4.831%; 1.13: -14.81%
| Z-Score (Absolute) |  | 1.03: 0.04; 1.04: 0.67; 1.06: 0; 1.07: 0.5; 1.08: 0.28; 1.09: 95.75; 1.10: 2.34; 1.12: 0.83; 1.13: 1.4
| Analytical results for Ba in OREAS 753b (Certified Value 131 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | 130 | 133.0943 | 137 | 133 | 135 | 134 | 130 | 200 | 132 | 138 | 126 | 130 | 117 | 143 | 124 | 147.53 | 125.9 | [7 empty] | 1
| 1 | 2 | 133 | 132.2903 | 136 | 135 | 134 | 133 | 138 | 200 | 131 | 139 | 80 | 120 | 122 | 141 | 129 | 135.42 | 119.7 | [7 empty] | 0
| 1 | 3 | 129 | 131.7955 | 137 | 134 | 132 | 132 | 136 | 200 | 137 | 141 | 90 | 130 | 120 | 91 | 114 | 135.79 | 121.5 | [7 empty] | 16
| 1 | 4 | 130 | 132.1773 | 135 | 136 | 139 | 129 | 127 | 200 | 139 | 139 | 113 | 120 | 125 | 139 | 110 | 135.94 | 124.3 | [7 empty] | 131.4135
| 1 | 5 | 132 | 132.7543 | 138 | 133 | 137 | 134 | 129 | 200 | 133 | 140 | 129 | 130 | 121 | 139 | 110 | 138.35 | 122.5 | [7 empty] | 68
| 1 | 6 | 128 | 140.889 | 140 | 134 | 134 | 127 | 131 | 200 | 136 | 142 | 106 | 130 | 122 | 134 | 109 | 135.9 | 126.5
| Mean |  | 130.3333 | 133.8335 | 137.1667 | 134.1667 | 135.1667 | 131.5 | 131.8333 | 200 | 134.6667 | 139.8333 | 107.3333 | 126.6667 | 121.1667 | 131.1667 | 116 | 138.155 | 123.4
| Median |  | 130 | 132.5223 | 137 | 134 | 134.5 | 132.5 | 130.5 | 200 | 134.5 | 139.5 | 109.5 | 130 | 121.5 | 139 | 112 | 135.92 | 123.4
| Std Dev. |  | 1.8619 | 3.4862 | 1.7224 | 1.169 | 2.4833 | 2.881 | 4.2622 | 0 | 3.1411 | 1.472 | 19.4902 | 5.164 | 2.6394 | 19.9039 | 8.4617 | 4.7115 | 2.6374
| Rel.Std.Dev. |  | 1.429% | 2.605% | 1.256% | 0.8713% | 1.837% | 2.191% | 3.233% | 0% | 2.333% | 1.053% | 18.16% | 4.077% | 2.178% | 15.17% | 7.295% | 3.41% | 2.137%
| PDM3 |  | -0.822% | 1.841% | 4.378% | 2.095% | 2.856% | 0.06583% | 0.3195% | 52.19% | 2.476% | 6.407% | -18.32% | -3.612% | -7.797% | -0.1878% | -11.73% | 5.13% | -6.098%
| Z-Score (Absolute) |  | 0.2 | 0.26 | 0.7 | 0.3 | 0.44 | 0.04 | 0 | 8.9 | 0.37 | 1.04 | 3.2 | 0.67 | 1.39 | 0.09 | 2.07 | 0.83 | 1.1
| Analytical results for Be in OREAS 753b (Certified Value 12 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | [8 empty] | 1
2
| 1 | 1 | 12 | 13 | 13 | 13 | 12 | 11 | 13 | 11 | 14 | 11 | 12 | 12.8231 | 12.2 | 13 | < 10 | 10.1 | [8 empty] | 1
| 1 | 2 | 12 | 13 | 13 | 13 | 11 | 13 | 13 | 11 | 13 | 10 | 12 | 12.6765 | 12.1 | 12 | < 10 | 10 | [8 empty] | 12
| 1 | 3 | 12 | 13 | 13 | 13 | 12 | 13 | 13 | 11 | 9 | 9 | 12 | 12.6904 | 7.7 | 12 | < 10 | 10.1 | [8 empty] | 16
| 1 | 4 | 12 | 13 | 13 | 13 | 12 | 12 | 13 | 11 | 14 | 10 | 12 | 12.6554 | 11.4 | 12 | < 10 | 10.9 | [8 empty] | 11.9803
| 1 | 5 | 12 | 12 | 13 | 13 | 12 | 12 | 12 | 11 | 13 | 9 | 12 | 12.7311 | 11.5 | 13 | < 10 | 10.8 | [8 empty] | 69
| 1 | 6 | 12 | 13 | 13 | 13 | 12 | 13 | 12 | 11 | 12 | 11 | 12 | 12.6945 | 11.1 | 12 | 10.78 | 10.1
| Mean |  | 12 | 12.8333 | 13 | 13 | 11.8333 | 12.3333 | 12.6667 | 11 | 12.5 | 10 | 12 | 12.7118 | 11 | 12.3333 | 10.78 | 10.3333
| Median |  | 12 | 13 | 13 | 13 | 12 | 12.5 | 13 | 11 | 13 | 10 | 12 | 12.6924 | 11.45 | 12 | 10.78 | 10.1
| Std Dev. |  | 0 | 0.40825 | 0 | 0 | 0.40825 | 0.8165 | 0.5164 | 0 | 1.8708 | 0.89443 | 0 | 0.059925 | 1.6709 | 0.5164 |  | 0.40332
| Rel.Std.Dev. |  | 0% | 3.181% | 0% | 0% | 3.45% | 6.62% | 4.077% | 0% | 14.97% | 8.944% | 0% | 0.4714% | 15.19% | 4.187% |  | 3.903%
| PDM3 |  | 0.1642% | 7.12% | 8.511% | 8.511% | -1.227% | 2.947% | 5.729% | -8.183% | 4.338% | -16.53% | 0.1642% | 6.106% | -8.183% | 2.947% | -10.02% | -13.75%
| Z-Score (Absolute) |  | 0.19 | 0.74 | 0.93 | 0.93 | 0.37 | 0.19 | 0.56 | 1.3 | 0.37 | 2.41 | 0.19 | 0.61 | 1.3 | 0.19 | 6.9 | 2.04
| Analytical results for Bi in OREAS 753b (Certified Value 0.75 ppm) | [25 empty] | Certified
Bi | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab | [5 empty] | 1
| No | No | 1.03: 1.01; 1.04: 1.03; 1.06: 1.04; 1.07: 1.05; 1.08: 1.06; 1.09: 1.07; 1.10: 1.09; 1.12: 1.10; 1.13: 1.12; 1.14: 1.22; 1.15: 1.28 | [5 empty] | ppm
|  |  | 1.03: PF*MS; 1.04: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*OES/MS; 1.09: PF*MS; 1.10: PF*OES/MS; 1.12: PF*OES/MS; 1.13: PF*OES; 1.14: PF*MS; 1.15: PF*OES | [5 empty] | 2
2
| 1 | 1 | 1.03: 0.6; 1.04: 0.8; 1.06: 1; 1.07: 0.7; 1.08: 0.7; 1.09: 0.3; 1.10: < 2; 1.12: 0.8; 1.13: 0.7; 1.14: 0.9; 1.15: < 5 | [5 empty] | 1
| 1 | 2 | 1.03: 0.6; 1.04: 0.8; 1.06: 0.8; 1.07: 0.9; 1.08: 0.6; 1.09: 0.3; 1.10: < 2; 1.12: 0.7; 1.13: 0.7; 1.14: 1; 1.15: < 5 | [5 empty] | 13
| 1 | 3 | 1.03: 0.7; 1.04: 0.7; 1.06: 0.9; 1.07: 0.8; 1.08: 0.7; 1.09: 0.4; 1.10: < 2; 1.12: 0.7; 1.13: 0.5; 1.14: 0.5; 1.15: < 5 | [5 empty] | 16
| 1 | 4 | 1.03: 0.7; 1.04: 0.7; 1.06: 0.8; 1.07: 0.7; 1.08: 0.7; 1.09: 0.4; 1.10: < 2; 1.12: 0.7; 1.13: 0.7; 1.14: 0.8; 1.15: < 5 | [5 empty] | 0.74583
| 1 | 5 | 1.03: 0.8; 1.04: 0.8; 1.06: 0.8; 1.07: 0.8; 1.08: 0.7; 1.09: 0.3; 1.10: < 2; 1.12: 0.7; 1.13: 0.8; 1.14: 0.8; 1.15: < 5 | [5 empty] | 70
| 1 | 6 | 1.03: 0.6; 1.04: 0.7; 1.06: 0.9; 1.07: 0.7; 1.08: 0.7; 1.09: 0.3; 1.10: < 2; 1.12: 0.7; 1.13: 0.9; 1.14: 0.8; 1.15: < 5
| Mean |  | 1.03: 0.66667; 1.04: 0.75; 1.06: 0.86667; 1.07: 0.76667; 1.08: 0.68333; 1.09: 0.33333; 1.12: 0.71667; 1.13: 0.71667; 1.14: 0.8
| Median |  | 1.03: 0.65; 1.04: 0.75; 1.06: 0.85; 1.07: 0.75; 1.08: 0.7; 1.09: 0.3; 1.12: 0.7; 1.13: 0.7; 1.14: 0.8
| Std Dev. |  | 1.03: 0.08165; 1.04: 0.054772; 1.06: 0.08165; 1.07: 0.08165; 1.08: 0.040825; 1.09: 0.05164; 1.12: 0.040825; 1.13: 0.13292; 1.14: 0.16733
| Rel.Std.Dev. |  | 1.03: 12.25%; 1.04: 7.303%; 1.06: 9.421%; 1.07: 10.65%; 1.08: 5.974%; 1.09: 15.49%; 1.12: 5.696%; 1.13: 18.55%; 1.14: 20.92%
| PDM3 |  | 1.03: -10.61%; 1.04: 0.5587%; 1.06: 16.2%; 1.07: 2.793%; 1.08: -8.38%; 1.09: -55.31%; 1.12: -3.911%; 1.13: -3.911%; 1.14: 7.263%
| Z-Score (Absolute) |  | 1.03: 0.84; 1.04: 0; 1.06: 1.18; 1.07: 0.17; 1.08: 0.67; 1.09: 4.21; 1.10: 2.53; 1.12: 0.34; 1.13: 0.34; 1.14: 0.51; 1.15: 17.7
| Analytical results for Ca in OREAS 753b (Certified Value 0.224 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.48 | 0.19 | 0.26 | 0.4045 | 0.179 | 0.207 | 0.229 | 0.243 | 0.222 | 0.23 | 0.2 | 0.22911 | 0.2 | 0.22 | 0.3 | 0.23779 | 0.24 |  | 1
| 1 | 2 | 0.2 | 0.3 | 0.3 | 0.2 | 0.2 | 0.2 | 0.44 | 0.17 | 0.25 | 0.1815 | 0.15 | 0.222 | 0.207 | 0.243 | 0.243 | 0.22 | 0.22 | 0.22913 | 0.2 | 0.22 | 0.3 | 0.21066 | 0.25 |  | 0
| 1 | 3 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.46 | 0.19 | 0.25 | 0.2666 | 0.2 | 0.179 | 0.25 | 0.243 | 0.243 | 0.24 | 0.22 | 0.22838 | 0.2 | 0.22 | 0.3 | 0.21557 | 0.24 |  | 16
| 1 | 4 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.4 | 0.19 | 0.25 | 0.3466 | 0.193 | 0.2 | 0.222 | 0.257 | 0.243 | 0.22 | 0.23 | 0.22979 | 0.2 | 0.23 | 0.2 | 0.21728 | 0.24 |  | 0.22363
| 1 | 5 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.42 | 0.17 | 0.26 | 0.4545 | 0.179 | 0.207 | 0.236 | 0.272 | 0.229 | 0.23 | 0.22 | 0.23066 | 0.2 | 0.21 | 0.3 | 0.20089 | 0.24 |  | 71
| 1 | 6 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.52 | 0.16 | 0.24 | 0.2051 | 0.243 | 0.207 | 0.229 | 0.293 | 0.257 | 0.24 | 0.22 | 0.16301 | 0.2 | 0.23 | 0.3 | 0.18912 | 0.23
| Mean |  | 0.2 | 0.3 | 0.21667 | 0.2 | 0.2 | 0.2 | 0.45333 | 0.17833 | 0.25167 | 0.3098 | 0.19067 | 0.20367 | 0.22883 | 0.2585 | 0.2395 | 0.23 | 0.21833 | 0.21835 | 0.2 | 0.22167 | 0.28333 | 0.21189 | 0.24
| Median |  | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.45 | 0.18 | 0.25 | 0.3066 | 0.186 | 0.207 | 0.229 | 0.25 | 0.243 | 0.23 | 0.22 | 0.22912 | 0.2 | 0.22 | 0.3 | 0.21311 | 0.24
| Std Dev. |  | 3.0405e-17 | 0 | 0.040825 | 3.0405e-17 | 3.0405e-17 | 3.0405e-17 | 0.043205 | 0.013292 | 0.0075277 | 0.11005 | 0.030846 | 0.014081 | 0.014303 | 0.020472 | 0.012325 | 0.0089443 | 0.0098319 | 0.027122 | 3.0405e-17 | 0.0075277 | 0.040825 | 0.016458 | 0.0063246
| Rel.Std.Dev. |  | 1.52e-14% | 0% | 18.84% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 9.531% | 7.453% | 2.991% | 35.52% | 16.18% | 6.914% | 6.25% | 7.92% | 5.146% | 3.889% | 4.503% | 12.42% | 1.52e-14% | 3.396% | 14.41% | 7.767% | 2.635%
| PDM3 |  | -10.57% | 34.15% | -3.114% | -10.57% | -10.57% | -10.57% | 102.7% | -20.25% | 12.54% | 38.53% | -14.74% | -8.927% | 2.327% | 15.59% | 7.097% | 2.849% | -2.368% | -2.362% | -10.57% | -0.8777% | 26.7% | -5.251% | 7.32%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | 9.01 | 2.02 | 0.92 | 3.25 | 1.53 | 1 | 0 | 1.19 | 0.43 | 0.05 | 0.42 | 0.42 | NA | 0.28 | NA | 0.67 | 0.45
Results from laboratories 1, 3, 4, 5, 6, 7, 22 and 25 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 753b (Indicative Value < 10 ppm) | [25 empty] | Indicative
Cd | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  | 1
| No | No | 1.01: 1.03; 1.03: 1.04; 1.04: 1.07; 1.05: 1.08; 1.06: 1.09; 1.07: 1.10; 1.08: 1.12; 1.09: 1.18; 1.10: 1.22; 1.12: 1.25; 1.13: 1.27; 1.14: 1.28 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*OES/MS; 1.08: PF*OES; 1.09: PF*OES; 1.10: PF*MS; 1.12: PF*MS; 1.13: PF*OES; 1.14: PF*OES |  | 1
1
| 1 | 1 | 1.01: 3; 1.03: 0.2; 1.04: < 10; 1.05: < 50; 1.06: < 2; 1.07: < 5; 1.08: 0.8; 1.09: < 10; 1.10: < 0.5; 1.12: 0.2; 1.13: < 10; 1.14: < 1 |  | 1
| 1 | 2 | 1.01: < 1; 1.03: < 0.2; 1.04: < 10; 1.05: < 50; 1.06: < 2; 1.07: < 5; 1.08: 0.8; 1.09: < 10; 1.10: < 0.5; 1.12: 0.2; 1.13: < 10; 1.14: < 1 |  | 6
| 1 | 3 | 1.01: < 1; 1.03: < 0.2; 1.04: < 10; 1.05: < 50; 1.06: < 2; 1.07: < 5; 1.08: < 0.1; 1.09: < 10; 1.10: 0.6; 1.12: 0.2; 1.13: < 10; 1.14: < 1 |  | 16
| 1 | 4 | 1.01: < 1; 1.03: < 0.2; 1.04: < 10; 1.05: < 50; 1.06: < 2; 1.07: < 5; 1.08: 0.7; 1.09: < 10; 1.10: 0.5; 1.12: 0.2; 1.13: < 10; 1.14: < 1 |  | < 10
| 1 | 5 | 1.01: 3; 1.03: < 0.2; 1.04: < 10; 1.05: < 50; 1.06: < 2; 1.07: < 5; 1.08: 0.9; 1.09: < 10; 1.10: 0.6; 1.12: < 0.2; 1.13: < 10; 1.14: < 1 |  | 12
| 1 | 6 | 1.01: < 1; 1.03: 0.4; 1.04: < 10; 1.05: < 50; 1.06: < 2; 1.07: < 5; 1.08: 0.9; 1.09: < 10; 1.10: < 0.5; 1.12: < 0.2; 1.13: < 10; 1.14: < 1
| Mean |  | 1.01: 3; 1.03: 0.3; 1.08: 0.82; 1.10: 0.56667; 1.12: 0.2
| Median |  | 1.01: 3; 1.03: 0.3; 1.08: 0.8; 1.10: 0.6; 1.12: 0.2
| Std Dev. |  | 1.01: 0; 1.03: 0.14142; 1.08: 0.083666; 1.10: 0.057735; 1.12: 0
| Rel.Std.Dev. |  | 1.01: 0%; 1.03: 47.14%; 1.08: 10.2%; 1.10: 10.19%; 1.12: 0%
PDM3
| Z-Score (Absolute) |  | 1.01: 0.11; 1.03: 0.67; 1.04: 2.58; 1.05: 16.07; 1.06: 0.11; 1.07: 0.9; 1.08: 0.32; 1.09: 2.58; 1.10: 0.51; 1.12: 0.67; 1.13: 2.58; 1.14: 0.45
| Analytical results for Ce in OREAS 753b (Certified Value 8.89 ppm) | [25 empty] | Certified
Ce | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  | 1
| No | No | 1.01: 1.01; 1.03: 1.02; 1.04: 1.03; 1.05: 1.04; 1.06: 1.09; 1.07: 1.10; 1.08: 1.12; 1.09: 1.20; 1.10: 1.22; 1.12: 1.25; 1.13: 1.27; 1.14: 1.28 |  | ppm
|  |  | 1.01: PF*MS; 1.03: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*OES/MS; 1.08: PF*OES; 1.09: PF*MS; 1.10: PF*MS; 1.12: PF*MS; 1.13: PF*MS; 1.14: PF*MS |  | 2
2
| 1 | 1 | 1.01: 9.5; 1.03: 10.2075; 1.04: 9.9; 1.05: 10.5; 1.06: 9.1; 1.07: 9; 1.08: 6.5; 1.09: 9.7769; 1.10: 8.2; 1.12: 8.4; 1.13: 7.4424; 1.14: 8.6 |  | 1
| 1 | 2 | 1.01: 9.2; 1.03: 9.8372; 1.04: 9.6; 1.05: 9.9; 1.06: 9.2; 1.07: 9; 1.08: 8.8; 1.09: 9.8304; 1.10: 8.4; 1.12: 8.6; 1.13: 6.7314; 1.14: 8.3 |  | 14
| 1 | 3 | 1.01: 9.8; 1.03: 10.292; 1.04: 9.9; 1.05: 10.1; 1.06: 8.9; 1.07: 8.8; 1.08: 7.9; 1.09: 9.7822; 1.10: 8.7; 1.12: 7.7; 1.13: 7.4783; 1.14: 7.9 |  | 16
| 1 | 4 | 1.01: 9.5; 1.03: 9.2467; 1.04: 10.1; 1.05: 9.1; 1.06: 9.2; 1.07: 8; 1.08: 9.5; 1.09: 9.7568; 1.10: 8.6; 1.12: 7.9; 1.13: 7.186; 1.14: 7.3 |  | 8.8877
| 1 | 5 | 1.01: 9.3; 1.03: 9.1666; 1.04: 9.4; 1.05: 9.8; 1.06: 9.4; 1.07: 7.9; 1.08: 12; 1.09: 9.844; 1.10: 7.8; 1.12: 8; 1.13: 8.1276; 1.14: 8.1 |  | 72
| 1 | 6 | 1.01: 8.8; 1.03: 8.9991; 1.04: 9.6; 1.05: 9; 1.06: 8.4; 1.07: 8.9; 1.08: 10.4; 1.09: 9.7769; 1.10: 8.8; 1.12: 8; 1.13: 6.9543; 1.14: 8.9
| Mean |  | 1.01: 9.35; 1.03: 9.6248; 1.04: 9.75; 1.05: 9.7333; 1.06: 9.0333; 1.07: 8.6; 1.08: 9.1833; 1.09: 9.7945; 1.10: 8.4167; 1.12: 8.1; 1.13: 7.32; 1.14: 8.1833
| Median |  | 1.01: 9.4; 1.03: 9.5419; 1.04: 9.75; 1.05: 9.85; 1.06: 9.15; 1.07: 8.85; 1.08: 9.15; 1.09: 9.7795; 1.10: 8.5; 1.12: 8; 1.13: 7.3142; 1.14: 8.2
| Std Dev. |  | 1.01: 0.33912; 1.03: 0.56112; 1.04: 0.25884; 1.05: 0.58195; 1.06: 0.35024; 1.07: 0.5099; 1.08: 1.924; 1.09: 0.034438; 1.10: 0.37103; 1.12: 0.33466; 1.13: 0.48792; 1.14: 0.56006
| Rel.Std.Dev. |  | 1.01: 3.627%; 1.03: 5.83%; 1.04: 2.655%; 1.05: 5.979%; 1.06: 3.877%; 1.07: 5.929%; 1.08: 20.95%; 1.09: 0.3516%; 1.10: 4.408%; 1.12: 4.132%; 1.13: 6.665%; 1.14: 6.844%
| PDM3 |  | 1.01: 5.201%; 1.03: 8.294%; 1.04: 9.702%; 1.05: 9.514%; 1.06: 1.638%; 1.07: -3.237%; 1.08: 3.326%; 1.09: 10.2%; 1.10: -5.3%; 1.12: -8.863%; 1.13: -17.64%; 1.14: -7.925%
| Z-Score (Absolute) |  | 1.01: 0.26; 1.03: 0.55; 1.04: 0.68; 1.05: 0.67; 1.06: 0.08; 1.07: 0.54; 1.08: 0.08; 1.09: 0.73; 1.10: 0.74; 1.12: 1.07; 1.13: 1.9; 1.14: 0.98
| Analytical results for Co in OREAS 753b (Certified Value 1.58 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*OES |  |  | 2
2
| 1 | 1 | 1.3802 | 1 | 1.5 | 2 | 1 | 1 | < 10 | 1.7 | 2 | 1 | < 20 | < 20 | < 20 | < 20 | 20 | 10 | < 1 | 2 | < 20 | 2.9 | < 10 | 1.4 |  |  | 1
| 1 | 2 | 1.3811 | < 1 | 1.6 | 2 | 2 | 2 | < 10 | 1.6 | 2 | 1 | < 20 | < 20 | < 20 | < 20 | < 20 | 10 | < 1 | 2 | < 20 | 2.7 | < 10 | 1.3 |  |  | 15
| 1 | 3 | 1.4932 | < 1 | 1.9 | 1 | 1 | 1 | < 10 | 1.6 | 2 | < 1 | < 20 | < 20 | < 20 | < 20 | 30 | 10 | < 1 | 1 | < 20 | 2.5 | < 10 | 1.4 |  |  | 16
| 1 | 4 | 1.3427 | 1 | 2 | 1 | 2 | 1 | < 10 | 1.6 | 1 | 1 | < 20 | < 20 | < 20 | < 20 | < 20 | 10 | < 1 | 2 | < 20 | 2.4 | < 10 | 1.5 |  |  | 1.58
| 1 | 5 | 1.489 | 1 | 1.4 | 2 | 2 | 1 | < 10 | 1.5 | 2 | 1 | < 20 | < 20 | < 20 | < 20 | 20 | 10 | < 1 | 2 | < 20 | 2.2 | < 10 | 1.5 |  |  | 73
| 1 | 6 | 1.5348 | 1 | 1.5 | 2 | 2 | 2 | < 10 | 1.4 | 1 | 1 | < 20 | < 20 | < 20 | < 20 | 20 | 10 | < 1 | 2 | < 20 | 2.3 | < 10 | 1.3
| Mean |  | 1.4368 | 1 | 1.65 | 1.6667 | 1.6667 | 1.3333 |  | 1.5667 | 1.6667 | 1 |  |  |  |  | 22.5 | 10 |  | 1.8333 |  | 2.5 |  | 1.4
| Median |  | 1.4351 | 1 | 1.55 | 2 | 2 | 1 |  | 1.6 | 2 | 1 |  |  |  |  | 20 | 10 |  | 2 |  | 2.45 |  | 1.4
| Std Dev. |  | 0.078321 | 0 | 0.2429 | 0.5164 | 0.5164 | 0.5164 |  | 0.10328 | 0.5164 | 0 |  |  |  |  | 5 | 0 |  | 0.40825 |  | 0.26077 |  | 0.089443
| Rel.Std.Dev. |  | 5.451% | 0% | 14.72% | 30.98% | 30.98% | 38.73% |  | 6.592% | 30.98% | 0% |  |  |  |  | 22.22% | 0% |  | 22.27% |  | 10.43% |  | 6.389%
| PDM3 |  | -9.063% | -36.71% | 4.429% | 5.484% | 5.484% | -15.61% |  | -0.845% | 5.484% | -36.71% |  |  |  |  | 1324% | 532.9% |  | 16.03% |  | 58.23% |  | -11.39%
| Z-Score (Absolute) |  | 0.24 | 0.71 | 0.08 | 0.06 | 0.06 | 0.32 | 2.5 | 0.14 | 0.06 | 0.64 | 6.36 | 6.36 | 6.36 | 6.36 | 12.78 | 6.36 | 0.96 | 0.06 | 6.36 | 0.58 | 2.5 | 0.27
| Analytical results for Cr in OREAS 753b (Indicative Value 39.2 ppm) | [25 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | [5 empty] | 1
1
| 1 | 1 | 28.138 | 18 |  | < 50 | < 50 | < 100 | < 30 | 20 | 38 | 70 | 140 | < 70 | 70 | < 70 | 30 | 26 | 35 | < 100 | < 50 | [5 empty] | 1
| 1 | 2 | 28.6141 | 19 | 63 | < 50 | < 50 | < 100 | < 30 | 20 | 11 | 70 | 140 | < 70 | < 70 | < 70 | 40 | 24 | 24 | < 100 | 51.49 | [5 empty] | 7
| 1 | 3 | 27.1144 | 20 |  | < 50 | < 50 | < 100 | < 30 | 20 | 16 | 70 | 140 | 70 | < 70 | < 70 | 30 | 24 | 26 | < 100 | < 50 | [5 empty] | 16
| 1 | 4 | 28.3063 | 28 |  | < 50 | < 50 | < 100 | < 30 | 20 | 22 | 70 | 70 | 70 | 140 | < 70 | 30 | 24 | 27 | < 100 | 58.13 | [5 empty] | 39.2489
| 1 | 5 | 26.0552 | 20 |  | < 50 | < 50 | < 100 | < 30 | 30 | 21 | 70 | 140 | < 70 | 70 | < 70 | 40 | 24 | 35 | < 100 | < 50 | [5 empty] | 13
| 1 | 6 | 27.2654 | 20 |  | < 50 | < 50 | < 100 | < 30 | 20 | < 1 | 70 | 140 | 70 | 140 | < 70 | 30 | 26 | 26 | < 100 | < 50
| Mean |  | 27.5822 | 20.8333 | 63 |  |  |  |  | 21.6667 | 21.6 | 70 | 128.3333 | 70 | 105 |  | 33.3333 | 24.6667 | 28.8333 |  | 54.81
| Median |  | 27.7017 | 20 | 63 |  |  |  |  | 20 | 21 | 70 | 140 | 70 | 105 |  | 30 | 24 | 26.5 |  | 54.81
| Std Dev. |  | 0.95384 | 3.6009 | [5 empty] | 4.0825 | 10.1637 | 0 | 28.5774 | 0 | 40.4145 |  | 5.164 | 1.0328 | 4.8751 |  | 4.6952
| Rel.Std.Dev. |  | 3.458% | 17.28% | [5 empty] | 18.84% | 47.05% | 0% | 22.27% | 0% | 38.49% |  | 15.49% | 4.187% | 16.91% |  | 8.566%
| PDM3 |  | -29.72% | -46.92% | 60.51% |  |  |  |  | -44.8% | -44.97% | 78.35% | 227% | 78.35% | 167.5% |  | -15.07% | -37.15% | -26.54% |  | 39.65%
| Z-Score (Absolute) |  | 0.31 | 0.67 | 1.6 | 0.45 | 0.45 | 0.9 | 0.99 | 0.63 | 0.82 | 1.89 | 4.97 | 0.97 | 2.51 | 0.05 | 0 | 0.47 | 0.24 | 0.9 | 0.09
| Analytical results for Cs in OREAS 753b (Certified Value 168 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.20 | 1.22 | 1.25 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [11 empty] | 0
0
| 1 | 1 | 160.9 | 159.2567 | 168.6 | 171 | 173.3 | 169 | 171.1 | 161 | 161 | 163 | 172.1537 | 172 | 170 | [11 empty] | 1
| 1 | 2 | 169 | 159.4661 | 162.8 | 165 | 176.3 | 167 | 173 | 160 | 159 | 167 | 172.2206 | 174 | 173 | [11 empty] | 16
| 1 | 3 | 166.5 | 157.4875 | 167.6 | 169 | 172.3 | 171 | 174.9 | 162 | 160 | 160 | 172.3966 | 111 | 167 | [11 empty] | 16
| 1 | 4 | 162.8 | 158.0385 | 164.5 | 175 | 176 | 171 | 169.2 | 164 | 163 | 166 | 172.4172 | 171 | 171 | [11 empty] | 167.5068
| 1 | 5 | 165.8 | 161.2398 | 163.9 | 171 | 171.2 | 170 | 176.7 | 163 | 163 | 165 | 172.2062 | 170 | 168 | [11 empty] | 74
| 1 | 6 | 160.2 | 158.1844 | 169.2 | 177 | 172.5 | 169 | 177 | 164 | 161 | 163 | 172.1661 | 168 | 168
| Mean |  | 164.2 | 158.9455 | 166.1 | 171.3333 | 173.6 | 169.5 | 173.65 | 162.3333 | 161.1667 | 164 | 172.2601 | 161 | 169.5
| Median |  | 164.3 | 158.7206 | 166.05 | 171 | 172.9 | 169.5 | 173.95 | 162.5 | 161 | 164 | 172.2134 | 170.5 | 169
| Std Dev. |  | 3.4566 | 1.3535 | 2.6981 | 4.274 | 2.0881 | 1.5166 | 3.1246 | 1.633 | 1.6021 | 2.5298 | 0.11659 | 24.5764 | 2.2583
| Rel.Std.Dev. |  | 2.105% | 0.8515% | 1.624% | 2.495% | 1.203% | 0.8947% | 1.799% | 1.006% | 0.9941% | 1.543% | 0.06768% | 15.26% | 1.332%
| PDM3 |  | -1.974% | -5.111% | -0.8399% | 2.284% | 3.638% | 1.19% | 3.667% | -3.089% | -3.785% | -2.094% | 2.838% | -3.885% | 1.19%
| Z-Score (Absolute) |  | 0.26 | 0.98 | 0 | 0.72 | 1.03 | 0.46 | 1.03 | 0.51 | 0.67 | 0.29 | 0.84 | 0.7 | 0.46
| Analytical results for Cu in OREAS 753b (Certified Value 20.9 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 1
1
| 1 | 1 | < 50 | 19 | < 20 | < 20 | < 50 | 18 | 21 | 19 | 20 | < 20 | 20 | 40 | < 20 | 20 | 19 | 28 | < 50 | 25 | 38.32 | 19.95 |  |  |  |  | 1
| 1 | 2 | < 50 | 19 | < 20 | < 20 | < 50 | 20 | 21 | 21 | 20 | < 20 | 20 | < 20 | < 20 | 20 | 19 | 24 | < 50 | 30 | < 30 | 20.47 |  |  |  |  | 0
| 1 | 3 | < 50 | 18 | < 20 | < 20 | < 50 | 20 | 21 | 19 | 20 | < 20 | 40 | 20 | < 20 | 20 | 20 | 22 | 70 | 25 | < 30 | 19.22 |  |  |  |  | 16
| 1 | 4 | < 50 | 20 | < 20 | < 20 | < 50 | 20 | 20 | 22 | < 20 | < 20 | 20 | 40 | 990 | 20 | 19 | 22 | < 50 | 30 | < 30 | 21.31 |  |  |  |  | 20.8927
| 1 | 5 | < 50 | 22 | < 20 | < 20 | < 50 | 21 | 21 | 23 | < 20 | < 20 | < 20 | 30 | < 20 | 20 | 19 | 22 | < 50 | 25 | < 30 | 19.74 |  |  |  |  | 75
| 1 | 6 | < 50 | 19 | < 20 | < 20 | < 50 | 17 | 21 | 21 | < 20 | < 20 | < 20 | 30 | 20 | 20 | 20 | 48 | < 50 | 25 | < 30 | 20.63
| Mean |  |  | 19.5 |  |  |  | 19.3333 | 20.8333 | 20.8333 | 20 |  | 25 | 32 | 505 | 20 | 19.3333 | 27.6667 | 70 | 26.6667 | 38.32 | 20.22
| Median |  |  | 19 |  |  |  | 20 | 21 | 21 | 20 |  | 20 | 30 | 505 | 20 | 19 | 23 | 70 | 25 | 38.32 | 20.21
| Std Dev. |  |  | 1.3784 |  |  |  | 1.5055 | 0.40825 | 1.6021 | 0 |  | 10 | 8.3666 | 685.8936 | 0 | 0.5164 | 10.2307 |  | 2.582 |  | 0.73783
| Rel.Std.Dev. |  |  | 7.069% |  |  |  | 7.787% | 1.96% | 7.69% | 0% |  | 40% | 26.15% | 135.8% | 0% | 2.671% | 36.98% |  | 9.682% |  | 3.649%
| PDM3 |  |  | -6.666% |  |  |  | -7.464% | -0.2843% | -0.2843% | -4.273% |  | 19.66% | 53.16% | 2317% | -4.273% | -7.464% | 32.42% | 235% | 27.64% | 83.41% | -3.22%
| Z-Score (Absolute) |  | 0.67 | 0.08 | 1.39 | 1.39 | 0.67 | 0.11 | 0.1 | 0.1 | 0.7 | 1.39 | 0.02 | 1.13 | 21.36 | 0.02 | 0.11 | 1.04 | 1.71 | 0.9 | 0.17 | 0.02
| Analytical results for Dy in OREAS 753b (Certified Value 0.69 ppm) | [25 empty] | Certified
Dy | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  |  | 1
| No | No | 1.03: 1.01; 1.04: 1.02; 1.06: 1.03; 1.07: 1.04; 1.08: 1.09; 1.09: 1.10; 1.10: 1.12; 1.12: 1.22; 1.13: 1.25; 1.14: 1.27; 1.15: 1.28 |  |  |  |  | ppm
|  |  | 1.03: PF*MS; 1.04: PF*OES/MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES/MS; 1.09: PF*OES/MS; 1.10: PF*OES; 1.12: PF*MS; 1.13: PF*MS; 1.14: PF*MS; 1.15: PF*MS |  |  |  |  | 2
2
| 1 | 1 | 1.03: 0.7; 1.04: 0.789; 1.06: 0.77; 1.07: 0.74; 1.08: 0.8; 1.09: 0.64; 1.10: 0.58; 1.12: 0.7; 1.13: 0.67; 1.14: 0.58502; 1.15: 0.5 |  |  |  |  | 1
| 1 | 2 | 1.03: 0.6; 1.04: 0.73362; 1.06: 0.76; 1.07: 0.73; 1.08: 0.6; 1.09: 0.61; 1.10: 0.7; 1.12: 0.74; 1.13: 0.63; 1.14: 0.49229; 1.15: 0.5 |  |  |  |  | 18
| 1 | 3 | 1.03: 0.7; 1.04: 0.78992; 1.06: 0.76; 1.07: 0.65; 1.08: 0.9; 1.09: 0.66; 1.10: 0.6; 1.12: 0.54; 1.13: 0.64; 1.14: 0.60216; 1.15: < 0.5 |  |  |  |  | 16
| 1 | 4 | 1.03: 0.8; 1.04: 0.71732; 1.06: 0.78; 1.07: 0.74; 1.08: 0.8; 1.09: 0.66; 1.10: 0.79; 1.12: 0.66; 1.13: 0.71; 1.14: 0.63823; 1.15: < 0.5 |  |  |  |  | 0.68717
| 1 | 5 | 1.03: 0.6; 1.04: 0.71222; 1.06: 0.7; 1.07: 0.76; 1.08: 0.6; 1.09: 0.62; 1.10: 0.79; 1.12: 0.64; 1.13: 0.63; 1.14: 0.59986; 1.15: < 0.5 |  |  |  |  | 76
| 1 | 6 | 1.03: 0.7; 1.04: 0.7134; 1.06: 0.74; 1.07: 0.94; 1.08: 0.7; 1.09: 0.74; 1.10: 0.67; 1.12: 0.7; 1.13: 0.69; 1.14: 0.6853; 1.15: 0.6
| Mean |  | 1.03: 0.68333; 1.04: 0.74258; 1.06: 0.75167; 1.07: 0.76; 1.08: 0.73333; 1.09: 0.655; 1.10: 0.68833; 1.12: 0.66333; 1.13: 0.66167; 1.14: 0.60048; 1.15: 0.53333
| Median |  | 1.03: 0.7; 1.04: 0.72547; 1.06: 0.76; 1.07: 0.74; 1.08: 0.75; 1.09: 0.65; 1.10: 0.685; 1.12: 0.68; 1.13: 0.655; 1.14: 0.60101; 1.15: 0.5
| Std Dev. |  | 1.03: 0.075277; 1.04: 0.037117; 1.06: 0.028577; 1.07: 0.096125; 1.08: 0.12111; 1.09: 0.046368; 1.10: 0.090203; 1.12: 0.069761; 1.13: 0.033714; 1.14: 0.064136; 1.15: 0.057735
| Rel.Std.Dev. |  | 1.03: 11.02%; 1.04: 4.998%; 1.06: 3.802%; 1.07: 12.65%; 1.08: 16.51%; 1.09: 7.079%; 1.10: 13.1%; 1.12: 10.52%; 1.13: 5.095%; 1.14: 10.68%; 1.15: 10.83%
| PDM3 |  | 1.03: -0.5589%; 1.04: 8.063%; 1.06: 9.385%; 1.07: 10.6%; 1.08: 6.717%; 1.09: -4.682%; 1.10: 0.1688%; 1.12: -3.469%; 1.13: -3.712%; 1.14: -12.62%; 1.15: -22.39%
| Z-Score (Absolute) |  | 1.03: NA; 1.04: 1.03; 1.06: 1.17; 1.07: 1.3; 1.08: NA; 1.09: 0.32; 1.10: 0.19; 1.12: 0.19; 1.13: 0.22; 1.14: 1.16; 1.15: NA
Results from laboratories 1, 9 and 28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 753b (Certified Value 0.37 ppm) | [25 empty] | Certified
Er | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  |  | 1
| No | No | 1.03: 1.01; 1.04: 1.02; 1.06: 1.03; 1.07: 1.04; 1.08: 1.09; 1.09: 1.10; 1.10: 1.12; 1.12: 1.22; 1.13: 1.25; 1.14: 1.27; 1.15: 1.28 |  |  |  |  | ppm
|  |  | 1.03: PF*MS; 1.04: PF*OES/MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES/MS; 1.09: PF*OES/MS; 1.10: PF*OES; 1.12: PF*MS; 1.13: PF*MS; 1.14: PF*MS; 1.15: PF*MS |  |  |  |  | 2
2
| 1 | 1 | 1.03: 0.4; 1.04: 0.46684; 1.06: 0.46; 1.07: 0.44; 1.08: 0.3; 1.09: 0.38; 1.10: 0.37; 1.12: 0.39; 1.13: 0.41; 1.14: 0.31521; 1.15: < 0.5 |  |  |  |  | 1
| 1 | 2 | 1.03: 0.4; 1.04: 0.35074; 1.06: 0.42; 1.07: 0.37; 1.08: 0.4; 1.09: 0.32; 1.10: 0.37; 1.12: 0.34; 1.13: 0.35; 1.14: < 0.3; 1.15: < 0.5 |  |  |  |  | 19
| 1 | 3 | 1.03: 0.3; 1.04: 0.41037; 1.06: 0.44; 1.07: 0.32; 1.08: 0.4; 1.09: 0.32; 1.10: 0.31; 1.12: 0.23; 1.13: 0.4; 1.14: 0.33688; 1.15: < 0.5 |  |  |  |  | 16
| 1 | 4 | 1.03: 0.4; 1.04: 0.38123; 1.06: 0.44; 1.07: 0.39; 1.08: 0.3; 1.09: 0.35; 1.10: 0.27; 1.12: 0.34; 1.13: 0.32; 1.14: 0.44832; 1.15: < 0.5 |  |  |  |  | 0.3663
| 1 | 5 | 1.03: 0.4; 1.04: 0.44082; 1.06: 0.4; 1.07: 0.4; 1.08: 0.4; 1.09: 0.31; 1.10: 0.33; 1.12: 0.39; 1.13: 0.34; 1.14: 0.31155; 1.15: < 0.5 |  |  |  |  | 77
| 1 | 6 | 1.03: 0.4; 1.04: 0.34446; 1.06: 0.42; 1.07: 0.46; 1.08: 0.5; 1.09: 0.38; 1.10: 0.27; 1.12: 0.34; 1.13: 0.31; 1.14: 0.34165; 1.15: < 0.5
| Mean |  | 1.03: 0.38333; 1.04: 0.39908; 1.06: 0.43; 1.07: 0.39667; 1.08: 0.38333; 1.09: 0.34333; 1.10: 0.32; 1.12: 0.33833; 1.13: 0.355; 1.14: 0.35072
| Median |  | 1.03: 0.4; 1.04: 0.3958; 1.06: 0.43; 1.07: 0.395; 1.08: 0.4; 1.09: 0.335; 1.10: 0.32; 1.12: 0.34; 1.13: 0.345; 1.14: 0.33688
| Std Dev. |  | 1.03: 0.040825; 1.04: 0.049193; 1.06: 0.020976; 1.07: 0.050067; 1.08: 0.075277; 1.09: 0.031411; 1.10: 0.045166; 1.12: 0.058452; 1.13: 0.041352; 1.14: 0.056111
| Rel.Std.Dev. |  | 1.03: 10.65%; 1.04: 12.33%; 1.06: 4.878%; 1.07: 12.62%; 1.08: 19.64%; 1.09: 9.149%; 1.10: 14.11%; 1.12: 17.28%; 1.13: 11.65%; 1.14: 16%
| PDM3 |  | 1.03: 4.65%; 1.04: 8.948%; 1.06: 17.39%; 1.07: 8.29%; 1.08: 4.65%; 1.09: -6.27%; 1.10: -12.64%; 1.12: -7.635%; 1.13: -3.085%; 1.14: -4.253%
| Z-Score (Absolute) |  | 1.03: NA; 1.04: 1.44; 1.06: 2.24; 1.07: 1.38; 1.08: NA; 1.09: 0; 1.10: 0.6; 1.12: 0.13; 1.13: 0.3; 1.14: 0.67; 1.15: 2.41
Results from laboratories 1 and 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 753b (Certified Value 0.16 ppm) | [25 empty] | Certified
Eu | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  |  | 1
| No | No | 1.03: 1.01; 1.04: 1.02; 1.06: 1.03; 1.07: 1.04; 1.08: 1.09; 1.09: 1.10; 1.10: 1.12; 1.12: 1.22; 1.13: 1.25; 1.14: 1.27; 1.15: 1.28 |  |  |  |  | ppm
|  |  | 1.03: PF*MS; 1.04: PF*OES/MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES/MS; 1.09: PF*OES/MS; 1.10: PF*OES; 1.12: PF*MS; 1.13: PF*MS; 1.14: PF*MS; 1.15: PF*MS |  |  |  |  | 2
2
| 1 | 1 | 1.03: 0.2; 1.04: 0.18412; 1.06: 0.17; 1.07: 0.16; 1.08: 0.2; 1.09: 0.15; 1.10: 0.1; 1.12: 0.17; 1.13: 0.14; 1.14: < 0.2; 1.15: < 0.5 |  |  |  |  | 1
| 1 | 2 | 1.03: 0.2; 1.04: 0.17331; 1.06: 0.18; 1.07: 0.18; 1.08: 0.2; 1.09: 0.18; 1.10: 0.18; 1.12: 0.17; 1.13: 0.13; 1.14: < 0.2; 1.15: < 0.5 |  |  |  |  | 20
| 1 | 3 | 1.03: 0.2; 1.04: 0.18776; 1.06: 0.18; 1.07: 0.16; 1.08: 0.1; 1.09: 0.15; 1.10: 0.15; 1.12: 0.14; 1.13: 0.16; 1.14: < 0.2; 1.15: < 0.5 |  |  |  |  | 16
| 1 | 4 | 1.03: 0.2; 1.04: 0.20581; 1.06: 0.17; 1.07: 0.18; 1.08: 0.2; 1.09: 0.17; 1.10: 0.13; 1.12: 0.16; 1.13: 0.15; 1.14: < 0.2; 1.15: < 0.5 |  |  |  |  | 0.16377
| 1 | 5 | 1.03: 0.1; 1.04: 0.20018; 1.06: 0.19; 1.07: 0.17; 1.08: 0.1; 1.09: 0.16; 1.10: 0.12; 1.12: 0.16; 1.13: 0.15; 1.14: < 0.2; 1.15: < 0.5 |  |  |  |  | 78
| 1 | 6 | 1.03: 0.2; 1.04: 0.20706; 1.06: 0.17; 1.07: 0.17; 1.08: 0.1; 1.09: 0.16; 1.10: 0.13; 1.12: 0.19; 1.13: 0.14; 1.14: < 0.2; 1.15: < 0.5
| Mean |  | 1.03: 0.18333; 1.04: 0.19304; 1.06: 0.17667; 1.07: 0.17; 1.08: 0.15; 1.09: 0.16167; 1.10: 0.135; 1.12: 0.165; 1.13: 0.145
| Median |  | 1.03: 0.2; 1.04: 0.19397; 1.06: 0.175; 1.07: 0.17; 1.08: 0.15; 1.09: 0.16; 1.10: 0.13; 1.12: 0.165; 1.13: 0.145
| Std Dev. |  | 1.03: 0.040825; 1.04: 0.013472; 1.06: 0.008165; 1.07: 0.0089443; 1.08: 0.054772; 1.09: 0.01169; 1.10: 0.027386; 1.12: 0.016432; 1.13: 0.010488
| Rel.Std.Dev. |  | 1.03: 22.27%; 1.04: 6.979%; 1.06: 4.622%; 1.07: 5.261%; 1.08: 36.51%; 1.09: 7.231%; 1.10: 20.29%; 1.12: 9.959%; 1.13: 7.233%
| PDM3 |  | 1.03: 11.95%; 1.04: 17.87%; 1.06: 7.876%; 1.07: 3.806%; 1.08: -8.407%; 1.09: -1.283%; 1.10: -17.57%; 1.12: 0.7525%; 1.13: -11.46%
| Z-Score (Absolute) |  | 1.03: NA; 1.04: 0.95; 1.06: 0.39; 1.07: 0.17; 1.08: NA; 1.09: 0.11; 1.10: 1.01; 1.12: 0; 1.13: 0.67; 1.14: 2.19; 1.15: 2.87
Results from laboratories 1 and 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 753b (Certified Value 0.624 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.63271 | 0.62 | 0.63 | 0.63 | 0.6 | 0.62 | 0.61 | 0.62 | 0.64 | 0.61 | 0.616 | 0.602 | 0.595 | 0.65 | 0.588 | 0.61 | 0.61 | 0.9422 | 0.57 | 0.63 | 0.64 | 0.65659 | 0.89 |  | 1
| 1 | 2 | 0.64425 | 0.62 | 0.63 | 0.64 | 0.61 | 0.6 | 0.62 | 0.63 | 0.64 | 0.6 | 0.602 | 0.643 | 0.602 | 0.63 | 0.609 | 0.59 | 0.61 | 0.95034 | 0.62 | 0.62 | 0.61 | 0.64499 | 0.89 |  | 0
| 1 | 3 | 0.63026 | 0.63 | 0.62 | 0.62 | 0.63 | 0.65 | 0.62 | 0.63 | 0.65 | 0.61 | 0.63 | 0.692 | 0.595 | 0.609 | 0.609 | 0.58 | 0.6 | 0.95029 | 0.6 | 0.62 | 0.62 | 0.63746 | 0.86 |  | 16
| 1 | 4 | 0.63278 | 0.63 | 0.65 | 0.64 | 0.62 | 0.61 | 0.6 | 0.63 | 0.64 | 0.6 | 0.643 | 0.692 | 0.609 | 0.734 | 0.623 | 0.61 | 0.62 | 0.96409 | 0.63 | 0.63 | 0.6 | 0.64637 | 0.88 |  | 0.62378
| 1 | 5 | 0.62994 | 0.64 | 0.62 | 0.63 | 0.62 | 0.65 | 0.59 | 0.62 | 0.65 | 0.61 | 0.595 | 0.713 | 0.595 | 0.636 | 0.867 | 0.59 | 0.62 | 0.94294 | 0.6 | 0.62 | 0.62 | 0.63877 | 0.88 |  | 79
| 1 | 6 | 0.64138 | 0.62 | 0.62 | 0.63 | 0.61 | 0.65 | 0.63 | 0.62 | 0.65 | 0.6 | 0.664 | 0.699 | 0.63 | 0.671 | 0.602 | 0.6 | 0.62 | 0.96343 | 0.6 | 0.64 | 0.57 | 0.64736 | 0.89
| Mean |  | 0.63522 | 0.62667 | 0.62833 | 0.63167 | 0.615 | 0.63 | 0.61167 | 0.625 | 0.645 | 0.605 | 0.625 | 0.6735 | 0.60433 | 0.655 | 0.64967 | 0.59667 | 0.61333 | 0.95222 | 0.60333 | 0.62667 | 0.61 | 0.64526 | 0.88167
| Median |  | 0.63275 | 0.625 | 0.625 | 0.63 | 0.615 | 0.635 | 0.615 | 0.625 | 0.645 | 0.605 | 0.623 | 0.692 | 0.5985 | 0.643 | 0.609 | 0.595 | 0.615 | 0.95032 | 0.6 | 0.625 | 0.615 | 0.64568 | 0.885
| Std Dev. |  | 0.0060708 | 0.008165 | 0.01169 | 0.0075277 | 0.010488 | 0.022804 | 0.01472 | 0.0054772 | 0.0054772 | 0.0054772 | 0.026 | 0.042279 | 0.013765 | 0.043873 | 0.10708 | 0.012111 | 0.008165 | 0.0095948 | 0.020656 | 0.008165 | 0.023664 | 0.0068851 | 0.01169
| Rel.Std.Dev. |  | 0.9557% | 1.303% | 1.861% | 1.192% | 1.705% | 3.62% | 2.406% | 0.8764% | 0.8492% | 0.9053% | 4.16% | 6.277% | 2.278% | 6.698% | 16.48% | 2.03% | 1.331% | 1.008% | 3.424% | 1.303% | 3.879% | 1.067% | 1.326%
| PDM3 |  | 1.834% | 0.463% | 0.7301% | 1.265% | -1.407% | 0.9973% | -1.942% | 0.1958% | 3.402% | -3.01% | 0.1958% | 7.971% | -3.117% | 5.005% | 4.15% | -4.346% | -1.675% | 52.65% | -3.278% | 0.463% | -2.209% | 3.443% | 41.34%
| Z-Score (Absolute) |  | 0.35 | 0 | 0.07 | 0.2 | 0.47 | 0.13 | 0.61 | 0.07 | 0.74 | 0.88 | 0.07 | 1.91 | 0.92 | 1.15 | 0.92 | 1.21 | 0.54 | 13.17 | 0.94 | 0 | 0.67 | 0.75 | 10.32
| Analytical results for Ga in OREAS 753b (Certified Value 48.5 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.19 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | [8 empty] | 1
2
| 1 | 1 | 46 | 44.4914 | 48 | 49 | 49 | 49 | 48 | 45.6 | 44.4 | 47 | 36 | 49.6303 | 49 | 41 | < 100 | 50.8 | [8 empty] | 1
| 1 | 2 | 47 | 44.921 | 50 | 53 | 51 | 49 | 47 | 51 | 43.5 | 47 | 36 | 49.6019 | 49 | 40 | < 100 | 49.9 | [8 empty] | 21
| 1 | 3 | 46 | 44.9197 | 50 | 49 | 50 | 49 | 50 | 47.6 | 43.7 | 44 | 36 | 48.6939 | 32 | 39 | < 100 | 48.2 | [8 empty] | 16
| 1 | 4 | 46 | 44.2779 | 50 | 51 | 47 | 50 | 46 | 50.2 | 41.2 | 50 | 36 | 49.6313 | 48 | 39 | < 100 | 49.3 | [8 empty] | 48.4981
| 1 | 5 | 46 | 43.9416 | 49 | 48 | 52 | 49 | 50 | 50.8 | 43.2 | 46 | 36 | 49.6529 | 48 | 38 | < 100 | 49.5 | [8 empty] | 80
| 1 | 6 | 45 | 45.5572 | 51 | 51 | 50 | 48 | 51 | 50.3 | 43.1 | 45 | 36 | 49.666 | 47 | 38 | < 100 | 51.1
| Mean |  | 46 | 44.6848 | 49.6667 | 50.1667 | 49.8333 | 49 | 48.6667 | 49.25 | 43.1833 | 46.5 | 36 | 49.4794 | 45.5 | 39.1667 |  | 49.8
| Median |  | 46 | 44.7055 | 50 | 50 | 50 | 49 | 49 | 50.25 | 43.35 | 46.5 | 36 | 49.6308 | 48 | 39 |  | 49.7
| Std Dev. |  | 0.63246 | 0.57049 | 1.0328 | 1.8348 | 1.7224 | 0.63246 | 1.9664 | 2.1686 | 1.0759 | 2.0736 | 0 | 0.38543 | 6.6558 | 1.169 |  | 1.0583
| Rel.Std.Dev. |  | 1.375% | 1.277% | 2.079% | 3.658% | 3.456% | 1.291% | 4.041% | 4.403% | 2.492% | 4.459% | 0% | 0.779% | 14.63% | 2.985% |  | 2.125%
| PDM3 |  | -5.151% | -7.863% | 2.409% | 3.44% | 2.753% | 1.035% | 0.3475% | 1.55% | -10.96% | -4.12% | -25.77% | 2.023% | -6.182% | -19.24% |  | 2.684%
| Z-Score (Absolute) |  | 1.53 | 2.24 | 0.45 | 0.72 | 0.54 | 0.09 | 0.09 | 0.22 | 3.05 | 1.26 | 6.92 | 0.35 | 1.8 | 5.21 | 0.63 | 0.52
| Analytical results for Gd in OREAS 753b (Certified Value 0.82 ppm) | [25 empty] | Certified
Gd | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.02: 1.01; 1.03: 1.02; 1.04: 1.03; 1.05: 1.04; 1.06: 1.09; 1.07: 1.10; 1.08: 1.12; 1.09: 1.22; 1.10: 1.25; 1.12: 1.27; 1.13: 1.28 |  | ppm
|  |  | 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*OES/MS; 1.08: PF*OES; 1.09: PF*MS; 1.10: PF*MS; 1.12: PF*MS; 1.13: PF*MS |  | 2
2
| 1 | 1 | 1.02: 0.7; 1.03: 0.9112; 1.04: < 3; 1.05: 0.97; 1.06: 1; 1.07: 0.83; 1.08: 0.66; 1.09: 0.7; 1.10: 0.81; 1.12: 0.70032; 1.13: 0.7 |  | 1
| 1 | 2 | 1.02: 0.8; 1.03: 0.87464; 1.04: < 3; 1.05: 0.97; 1.06: 0.8; 1.07: 0.79; 1.08: 0.89; 1.09: 0.8; 1.10: 0.8; 1.12: 0.63663; 1.13: 0.7 |  | 22
| 1 | 3 | 1.02: 0.7; 1.03: 0.90182; 1.04: < 3; 1.05: 0.9; 1.06: 0.7; 1.07: 0.83; 1.08: 0.68; 1.09: 0.67; 1.10: 0.76; 1.12: 0.7299; 1.13: 0.6 |  | 16
| 1 | 4 | 1.02: 0.8; 1.03: 0.93156; 1.04: < 3; 1.05: 0.9; 1.06: 0.9; 1.07: 0.84; 1.08: 0.85; 1.09: 0.75; 1.10: 0.71; 1.12: 0.81327; 1.13: 0.6 |  | 0.81533
| 1 | 5 | 1.02: 0.7; 1.03: 0.89195; 1.04: < 3; 1.05: 0.89; 1.06: 0.7; 1.07: 0.83; 1.08: 0.97; 1.09: 0.71; 1.10: 0.94; 1.12: 0.62986; 1.13: 0.7 |  | 81
| 1 | 6 | 1.02: 0.8; 1.03: 0.83632; 1.04: < 3; 1.05: 0.93; 1.06: 1.1; 1.07: 0.86; 1.08: 0.72; 1.09: 0.85; 1.10: 0.87; 1.12: 0.70635; 1.13: 0.7
| Mean |  | 1.02: 0.75; 1.03: 0.89125; 1.05: 0.92667; 1.06: 0.86667; 1.07: 0.83; 1.08: 0.795; 1.09: 0.74667; 1.10: 0.815; 1.12: 0.70272; 1.13: 0.66667
| Median |  | 1.02: 0.75; 1.03: 0.89688; 1.05: 0.915; 1.06: 0.85; 1.07: 0.83; 1.08: 0.785; 1.09: 0.73; 1.10: 0.805; 1.12: 0.70334; 1.13: 0.7
| Std Dev. |  | 1.02: 0.054772; 1.03: 0.032951; 1.05: 0.036148; 1.06: 0.1633; 1.07: 0.022804; 1.08: 0.12629; 1.09: 0.067725; 1.10: 0.081179; 1.12: 0.067313; 1.13: 0.05164
| Rel.Std.Dev. |  | 1.02: 7.303%; 1.03: 3.697%; 1.05: 3.901%; 1.06: 18.84%; 1.07: 2.747%; 1.08: 15.89%; 1.09: 9.07%; 1.10: 9.961%; 1.12: 9.579%; 1.13: 7.746%
| PDM3 |  | 1.02: -8.013%; 1.03: 9.311%; 1.05: 13.66%; 1.06: 6.297%; 1.07: 1.799%; 1.08: -2.493%; 1.09: -8.421%; 1.10: -0.04035%; 1.12: -13.81%; 1.13: -18.23%
| Z-Score (Absolute) |  | 1.02: NA; 1.03: 0.64; 1.04: 6.32; 1.05: 0.97; 1.06: NA; 1.07: 0.07; 1.08: 0.26; 1.09: 0.71; 1.10: 0.07; 1.12: 1.12; 1.13: NA
Results from laboratories 1, 9 and 28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ge in OREAS 753b (Certified Value 2.66 ppm) | [25 empty] | Certified
Ge | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.02: 1.02; 1.03: 1.03; 1.04: 1.04; 1.05: 1.05; 1.06: 1.06; 1.07: 1.07; 1.08: 1.09; 1.09: 1.10; 1.10: 1.12; 1.12: 1.22; 1.13: 1.25 |  | ppm
|  |  | 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES/MS; 1.09: PF*OES/MS; 1.10: PF*OES; 1.12: PF*MS; 1.13: PF*MS |  | 2
2
| 1 | 1 | 1.02: 2.2841; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 2.1; 1.09: 3; 1.10: < 5; 1.12: 3; 1.13: 2 |  | 1
| 1 | 2 | 1.02: 2.5162; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 2.3; 1.09: 3; 1.10: < 5; 1.12: 3; 1.13: 2 |  | 0
| 1 | 3 | 1.02: 2.2596; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 2.5; 1.09: 3; 1.10: < 5; 1.12: 3; 1.13: 2 |  | 16
| 1 | 4 | 1.02: 2.733; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 1.9; 1.09: 3; 1.10: < 5; 1.12: 3; 1.13: 2 |  | 2.6646
| 1 | 5 | 1.02: 2.8984; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 3.4; 1.09: 3; 1.10: < 5; 1.12: 3; 1.13: 2 |  | 82
| 1 | 6 | 1.02: 2.1971; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 2.8; 1.09: 3; 1.10: < 5; 1.12: 3; 1.13: 2
| Mean |  | 1.02: 2.4814; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 2.5; 1.09: 3; 1.12: 3; 1.13: 2
| Median |  | 1.02: 2.4001; 1.03: 2; 1.04: 3; 1.05: 3; 1.06: 3; 1.07: 1; 1.08: 2.4; 1.09: 3; 1.12: 3; 1.13: 2
| Std Dev. |  | 1.02: 0.28543; 1.03: 0; 1.04: 0; 1.05: 0; 1.06: 0; 1.07: 0; 1.08: 0.54037; 1.09: 0; 1.12: 0; 1.13: 0
| Rel.Std.Dev. |  | 1.02: 11.5%; 1.03: 0%; 1.04: 0%; 1.05: 0%; 1.06: 0%; 1.07: 0%; 1.08: 21.61%; 1.09: 0%; 1.12: 0%; 1.13: 0%
| PDM3 |  | 1.02: -6.876%; 1.03: -24.94%; 1.04: 12.59%; 1.05: 12.59%; 1.06: 12.59%; 1.07: -62.47%; 1.08: -6.177%; 1.09: 12.59%; 1.12: 12.59%; 1.13: -24.94%
| Z-Score (Absolute) |  | 1.02: 0.03; 1.03: 0.67; 1.04: 0.67; 1.05: 0.67; 1.06: 0.67; 1.07: 2.02; 1.08: 0; 1.09: 0.67; 1.10: 0; 1.12: 0.67; 1.13: 0.67
| Analytical results for Hf in OREAS 753b (Indicative Value 2.61 ppm) | [25 empty] | Indicative
Hf | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  | 1
| No | No | 1.02: 1.01; 1.03: 1.02; 1.04: 1.04; 1.05: 1.09; 1.06: 1.12; 1.07: 1.28 |  | ppm
|  |  | 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*OES; 1.07: PF*MS |  | 2
2
| 1 | 1 | 1.02: 2.7; 1.03: 3.188; 1.04: 3; 1.05: < 10; 1.06: < 1; 1.07: 1.9 |  | 1
| 1 | 2 | 1.02: 2.7; 1.03: 3.1887; 1.04: 3; 1.05: < 10; 1.06: 1; 1.07: 1.9 |  | 8
| 1 | 3 | 1.02: 2.5; 1.03: 2.7853; 1.04: 3; 1.05: 50; 1.06: < 1; 1.07: 1.9 |  | 16
| 1 | 4 | 1.02: 2.4; 1.03: 2.9012; 1.04: 3; 1.05: < 10; 1.06: < 1; 1.07: 1.9 |  | 2.6109
| 1 | 5 | 1.02: 2.6; 1.03: 3.0786; 1.04: 3; 1.05: < 10; 1.06: 3; 1.07: 2.3 |  | 14
| 1 | 6 | 1.02: 2; 1.03: 2.821; 1.04: 3; 1.05: < 10; 1.06: < 1; 1.07: 1.9
| Mean |  | 1.02: 2.4833; 1.03: 2.9938; 1.04: 3; 1.05: 50; 1.06: 2; 1.07: 1.9667
| Median |  | 1.02: 2.55; 1.03: 2.9899; 1.04: 3; 1.05: 50; 1.06: 2; 1.07: 1.9
| Std Dev. |  | 1.02: 0.26394; 1.03: 0.18155; 1.04: 0; 1.06: 1.4142; 1.07: 0.1633
| Rel.Std.Dev. |  | 1.02: 10.63%; 1.03: 6.064%; 1.04: 0%; 1.06: 70.71%; 1.07: 8.303%
| PDM3 |  | 1.02: -4.888%; 1.03: 14.66%; 1.04: 14.9%; 1.05: 1815%; 1.06: -23.4%; 1.07: -24.68%
| Z-Score (Absolute) |  | 1.02: 0.33; 1.03: 0.33; 1.04: 0.34; 1.05: 12.74; 1.06: 2.27; 1.07: 1.01
| Analytical results for Hg in OREAS 753b (Indicative Value < 5 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | 1.02: Lab |  | 1
| No | No | 1.02: 1.28 |  | ppm
|  |  | 1.02: PF*OES |  | 2
2
| 1 | 1 | 1.02: < 5 |  | 1
| 1 | 2 | 1.02: < 5 |  | 2
| 1 | 3 | 1.02: < 5 |  | 16
| 1 | 4 | 1.02: < 5 |  | < 5
| 1 | 5 | 1.02: < 5 |  | 10
| 1 | 6 | 1.02: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.02
| Analytical results for Ho in OREAS 753b (Certified Value 0.13 ppm) | [25 empty] | Certified
Ho | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.02: 1.01; 1.03: 1.02; 1.04: 1.03; 1.05: 1.04; 1.06: 1.09; 1.07: 1.10; 1.08: 1.12; 1.09: 1.22; 1.10: 1.25; 1.12: 1.27; 1.13: 1.28 |  | ppm
|  |  | 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*OES/MS; 1.08: PF*OES; 1.09: PF*MS; 1.10: PF*MS; 1.12: PF*MS; 1.13: PF*MS |  | 2
2
| 1 | 1 | 1.02: 0.1; 1.03: 0.16885; 1.04: 0.15; 1.05: 0.16; 1.06: < 0.2; 1.07: 0.12; 1.08: 0.11; 1.09: 0.12; 1.10: 0.15; 1.12: < 0.2; 1.13: < 0.5 |  | 1
| 1 | 2 | 1.02: 0.1; 1.03: 0.14217; 1.04: 0.13; 1.05: 0.13; 1.06: < 0.2; 1.07: 0.11; 1.08: 0.11; 1.09: 0.12; 1.10: 0.1; 1.12: < 0.2; 1.13: < 0.5 |  | 11
| 1 | 3 | 1.02: 0.1; 1.03: 0.14312; 1.04: 0.15; 1.05: 0.12; 1.06: < 0.2; 1.07: 0.12; 1.08: 0.09; 1.09: 0.08; 1.10: 0.1; 1.12: < 0.2; 1.13: < 0.5 |  | 16
| 1 | 4 | 1.02: 0.1; 1.03: 0.12982; 1.04: 0.13; 1.05: 0.14; 1.06: < 0.2; 1.07: 0.12; 1.08: 0.13; 1.09: 0.13; 1.10: 0.11; 1.12: 0.21145; 1.13: < 0.5 |  | 0.13081
| 1 | 5 | 1.02: 0.1; 1.03: 0.17843; 1.04: 0.15; 1.05: 0.16; 1.06: < 0.2; 1.07: 0.11; 1.08: 0.13; 1.09: 0.14; 1.10: 0.1; 1.12: < 0.2; 1.13: < 0.5 |  | 83
| 1 | 6 | 1.02: 0.1; 1.03: 0.12569; 1.04: 0.16; 1.05: 0.17; 1.06: < 0.2; 1.07: 0.13; 1.08: 0.12; 1.09: 0.12; 1.10: 0.14; 1.12: < 0.2; 1.13: < 0.5
| Mean |  | 1.02: 0.1; 1.03: 0.14801; 1.04: 0.145; 1.05: 0.14667; 1.07: 0.11833; 1.08: 0.115; 1.09: 0.11833; 1.10: 0.11667; 1.12: 0.21145
| Median |  | 1.02: 0.1; 1.03: 0.14264; 1.04: 0.15; 1.05: 0.15; 1.07: 0.12; 1.08: 0.115; 1.09: 0.12; 1.10: 0.105; 1.12: 0.21145
| Std Dev. |  | 1.02: 1.5202e-17; 1.03: 0.021199; 1.04: 0.012247; 1.05: 0.019664; 1.07: 0.0075277; 1.08: 0.015166; 1.09: 0.020412; 1.10: 0.022509
| Rel.Std.Dev. |  | 1.02: 1.52e-14%; 1.03: 14.32%; 1.04: 8.447%; 1.05: 13.41%; 1.07: 6.361%; 1.08: 13.19%; 1.09: 17.25%; 1.10: 19.29%
| PDM3 |  | 1.02: -23.55%; 1.03: 13.15%; 1.04: 10.85%; 1.05: 12.12%; 1.07: -9.539%; 1.08: -12.09%; 1.09: -9.539%; 1.10: -10.81%; 1.12: 61.64%
| Z-Score (Absolute) |  | 1.02: NA; 1.03: 1.82; 1.04: 1.63; 1.05: 1.74; 1.06: 1.14; 1.07: 0.01; 1.08: 0.21; 1.09: 0.01; 1.10: 0.11; 1.12: 0.01; 1.13: 8.1
Results from laboratory 1 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 753b (Certified Value < 0.2 ppm) | [25 empty] | Certified
In | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab |  | 1
| No | No | 1.02: 1.01; 1.03: 1.03; 1.04: 1.04; 1.05: 1.06; 1.06: 1.07; 1.07: 1.09; 1.08: 1.10; 1.09: 1.12; 1.10: 1.22; 1.12: 1.25 |  | ppm
|  |  | 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES/MS; 1.08: PF*OES/MS; 1.09: PF*OES; 1.10: PF*MS; 1.12: PF*MS |  | 2
2
| 1 | 1 | 1.03: < 0.3; 1.04: < 0.2; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.2; 1.08: < 0.2; 1.09: < 0.05; 1.10: < 0.2; 1.12: < 0.2 |  | 1
| 1 | 2 | 1.03: < 0.3; 1.04: < 0.2; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.2; 1.08: < 0.2; 1.09: < 0.05; 1.10: < 0.2; 1.12: < 0.2 |  | 1
| 1 | 3 | 1.03: < 0.3; 1.04: < 0.2; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.2; 1.08: < 0.2; 1.09: < 0.05; 1.10: < 0.2; 1.12: < 0.2 |  | 16
| 1 | 4 | 1.03: < 0.3; 1.04: < 0.2; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.2; 1.08: < 0.2; 1.09: 0.07; 1.10: < 0.2; 1.12: < 0.2 |  | < 0.2
| 1 | 5 | 1.03: < 0.3; 1.04: < 0.2; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.2; 1.08: < 0.2; 1.09: < 0.05; 1.10: < 0.2; 1.12: < 0.2 |  | 84
| 1 | 6 | 1.02: 0.1; 1.03: < 0.3; 1.04: < 0.2; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.2; 1.08: < 0.2; 1.09: < 0.05; 1.10: < 0.2; 1.12: < 0.2
| Mean |  | 1.02: 0.1; 1.09: 0.07
| Median |  | 1.02: 0.1; 1.09: 0.07
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.02 to 1.12
| Analytical results for K in OREAS 753b (Certified Value 0.966 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.96 | 1.0704 | 0.96 | 1 | 0.93 | 0.96 | 0.96 | 1.11 | 1 | 0.92 | 0.78 | 0.913 | 0.888 | 0.996 | 1.046 | 0.996 | 0.9 | 0.98 | 0.8 | 0.98 | 1 | 0.91385 | 0.89 |  | 1
| 1 | 2 | 0.95 | 1.0964 | 0.95 | 1 | 0.93 | 0.92 | 0.94 | 1.07 | 1 | 0.91 | 0.35 | 0.888 | 0.938 | 0.963 | 0.98 | 0.996 | 0.91 | 0.98 | 1 | 0.97 | 0.9 | 0.94658 | 0.99 |  | 0
| 1 | 3 | 0.99 | 1.055 | 0.96 | 1 | 0.91 | 0.93 | 1 | 1.09 | 1 | 0.94 | 0.51 | 0.921 | 0.963 | 0.98 | 0.955 | 0.996 | 0.94 | 0.98 | 1 | 0.97 | 0.9 | 0.92528 | 0.91 |  | 16
| 1 | 4 | 0.97 | 1.0805 | 0.94 | 1 | 0.95 | 0.95 | 0.93 | 1.02 | 1 | 0.92 | 0.64 | 0.921 | 0.946 | 0.996 | 1.021 | 0.996 | 0.92 | 0.98 | 1 | 0.97 | 0.9 | 0.91138 | 1.12 |  | 0.96599
| 1 | 5 | 0.97 | 1.0631 | 0.98 | 1 | 0.94 | 0.93 | 0.95 | 1.01 | 1 | 0.94 | 1.09 | 0.913 | 0.98 | 0.83 | 0.946 | 1.029 | 0.9 | 0.99 | 0.9 | 0.95 | 0.9 | 0.91526 | 1.15 |  | 85
| 1 | 6 | 0.99 | 1.0355 | 0.98 | 1 | 0.93 | 0.93 | 1 | 1.17 | 1 | 0.94 | 0.99 | 0.955 | 0.996 | 0.872 | 1.029 | 1.046 | 0.93 | 0.98 | 0.9 | 0.99 | 0.9 | 0.90767 | 0.96
| Mean |  | 0.97167 | 1.0668 | 0.96167 | 1 | 0.93167 | 0.93667 | 0.96333 | 1.0783 | 1 | 0.92833 | 0.72667 | 0.9185 | 0.95183 | 0.9395 | 0.99617 | 1.0098 | 0.91667 | 0.98167 | 0.93333 | 0.97167 | 0.91667 | 0.92 | 1.0033
| Median |  | 0.97 | 1.0667 | 0.96 | 1 | 0.93 | 0.93 | 0.955 | 1.08 | 1 | 0.93 | 0.71 | 0.917 | 0.9545 | 0.9715 | 1.0005 | 0.996 | 0.915 | 0.98 | 0.95 | 0.97 | 0.9 | 0.91456 | 0.975
| Std Dev. |  | 0.016021 | 0.021021 | 0.016021 | 0 | 0.013292 | 0.015055 | 0.030111 | 0.05947 | 0 | 0.013292 | 0.28303 | 0.021612 | 0.037865 | 0.070885 | 0.041595 | 0.022094 | 0.01633 | 0.0040825 | 0.08165 | 0.013292 | 0.040825 | 0.01429 | 0.10838
| Rel.Std.Dev. |  | 1.649% | 1.971% | 1.666% | 0% | 1.427% | 1.607% | 3.126% | 5.515% | 0% | 1.432% | 38.95% | 2.353% | 3.978% | 7.545% | 4.176% | 2.188% | 1.781% | 0.4159% | 8.748% | 1.368% | 4.454% | 1.553% | 10.8%
| PDM3 |  | 0.5873% | 10.44% | -0.4479% | 3.52% | -3.554% | -3.036% | -0.2754% | 11.63% | 3.52% | -3.899% | -24.78% | -4.917% | -1.466% | -2.743% | 3.124% | 4.538% | -5.106% | 1.622% | -3.381% | 0.5873% | -5.106% | -4.761% | 3.865%
| Z-Score (Absolute) |  | 0.2 | 2.13 | 0 | NA | 0.61 | 0.51 | 0.03 | 2.36 | NA | 0.67 | 4.75 | 0.87 | 0.2 | 0.45 | 0.7 | 0.98 | 0.91 | 0.4 | NA | 0.2 | NA | 0.84 | 0.84
Results from laboratories 4, 9, 22 and 25 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for La in OREAS 753b (Certified Value 4.55 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | [10 empty] | 2
3
| 1 | 1 | 4.7 | 4.9176 | 4.88 | 5.3 | 4.7 | 4.4 | 3.3 | < 10 | < 5 | 4.4882 | < 10 | 4.4 | 3.9134 | 4.5 | [10 empty] | 1
| 1 | 2 | 4.6 | 5.0881 | 4.54 | 5.1 | 4.9 | 4.4 | 4.3 | 10 | < 5 | 4.6613 | < 10 | 4.6 | 3.5648 | 4.5 | [10 empty] | 13
| 1 | 3 | 4.7 | 5.2613 | 4.86 | 5.1 | 4.7 | 4.4 | 3.8 | < 10 | < 5 | 4.5236 | < 10 | 4.3 | 3.9963 | 3.9 | [10 empty] | 16
| 1 | 4 | 4.6 | 4.793 | 4.97 | 4.6 | 5.1 | 3.9 | 5 | < 10 | 5 | 4.6237 | < 10 | 4.4 | 3.721 | 3.9 | [10 empty] | 4.5454
| 1 | 5 | 4.6 | 4.4769 | 4.58 | 4.7 | 5.1 | 3.8 | 5.5 | < 10 | < 5 | 4.4632 | < 10 | 4.3 | 4.29 | 4.3 | [10 empty] | 86
| 1 | 6 | 4.3 | 4.7931 | 4.82 | 4.5 | 4.6 | 4.4 | 5.3 | 10 | 5 | 4.6802 | < 10 | 4.3 | 3.7486 | 4.7
| Mean |  | 1.01: 4.5833; 1.02: 4.8883; 1.03: 4.775; 1.04: 4.8833; 1.05: 4.85; 1.06: 4.2167; 1.07: 4.5333; 1.08: 10; 1.09: 5; 1.10: 4.5734; 1.13: 4.3833; 1.14: 3.8723; 1.15: 4.3
| Median |  | 1.01: 4.6; 1.02: 4.8554; 1.03: 4.84; 1.04: 4.9; 1.05: 4.8; 1.06: 4.4; 1.07: 4.65; 1.08: 10; 1.09: 5; 1.10: 4.5737; 1.13: 4.35; 1.14: 3.831; 1.15: 4.4
| Std Dev. |  | 1.01: 0.1472; 1.02: 0.27104; 1.03: 0.1741; 1.04: 0.32506; 1.05: 0.21679; 1.06: 0.28577; 1.07: 0.87788; 1.08: 0; 1.09: 0; 1.10: 0.09332; 1.13: 0.1169; 1.14: 0.25467; 1.15: 0.33466
| Rel.Std.Dev. |  | 1.01: 3.212%; 1.02: 5.545%; 1.03: 3.646%; 1.04: 6.657%; 1.05: 4.47%; 1.06: 6.777%; 1.07: 19.36%; 1.08: 0%; 1.09: 0%; 1.10: 2.04%; 1.13: 2.667%; 1.14: 6.577%; 1.15: 7.783%
| PDM3 |  | 1.01: 0.8353%; 1.02: 7.545%; 1.03: 5.052%; 1.04: 7.435%; 1.05: 6.702%; 1.06: -7.232%; 1.07: -0.2648%; 1.08: 120%; 1.09: 10%; 1.10: 0.6163%; 1.13: -3.565%; 1.14: -14.81%; 1.15: -5.398%
| Z-Score (Absolute) |  | 0.01 | 0.76 | 0.48 | 0.75 | 0.67 | 0.89 | 0.11 | NA | NA | 0.01 | 1.03 | 0.48 | 1.73 | 0.68
Results from laboratory 19 were removed due to their 1 ppm reading resolution. Results from laboratory 18 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 753b (Certified Value 1.04 wt.%) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | [6 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [6 empty] | 2
3
| 1 | 1 | 1.0852 | 1.0197 | 1.0576 | 1.04 | 1.05 | 1.0189 | 1.02 | 1.07 | 0.87 | 1.05 | 1.045 | 1.022 | 0.9882 | 1.0747 | 1.0706 | 1.061 | 0.997 | 1.0266 | [6 empty] | 1
| 1 | 2 | 1.0994 | 1.019 | 1.0526 | 1.03 | 1.07 | 1.0082 | 1.01 | 1.07 | 0.538 | 1.06 | 1.054 | 1.035 | 0.9868 | 1.0702 | 1.1148 | 1.063 | 1.015 | 1.0216 | [6 empty] | 0
| 1 | 3 | 1.0927 | 1.0308 | 1.0446 | 1.01 | 1.06 | 1.0673 | 1.03 | 1.09 | 0.658 | 1.065 | 1.045 | 1.041 | 0.9849 | 1.0744 | 1.1909 | 1.059 | 0.968 | 1.0407 | [6 empty] | 16
| 1 | 4 | 1.0819 | 1.0201 | 1.0688 | 1.06 | 1.06 | 1.0423 | 1.02 | 1.08 | 0.768 | 1.05 | 1.059 | 1.021 | 0.9836 | 1.0855 | 1.1661 | 1.057 | 0.955 | 1.0331 | [6 empty] | 1.0426
| 1 | 5 | 1.0971 | 1.0344 | 1.0478 | 1.04 | 1.08 | 1.0498 | 1.02 | 1.06 | 0.787 | 1.06 | 1.073 | 1.038 | 0.9821 | 1.0807 | 1.0464 | 1.057 | 0.968 | 1.0291 | [6 empty] | 87
| 1 | 6 | 1.0941 | 1.0268 | 1.0503 | 1.04 | 1.07 | 1.044 | 1 | 1.1 | 0.709 | 1.065 | 1.087 | 1.018 | 0.9804 | 1.0758 | 1.226 | 1.074 | 0.946 | 1.0269
| Mean |  | 1.0917 | 1.0251 | 1.0536 | 1.0367 | 1.065 | 1.0384 | 1.0167 | 1.0783 | 0.72167 | 1.0583 | 1.0605 | 1.0292 | 0.98433 | 1.0769 | 1.1358 | 1.0618 | 0.97483 | 1.0297
| Median |  | 1.0934 | 1.0234 | 1.0514 | 1.04 | 1.065 | 1.0432 | 1.02 | 1.075 | 0.7385 | 1.06 | 1.0565 | 1.0285 | 0.98425 | 1.0753 | 1.1404 | 1.06 | 0.968 | 1.028
| Std Dev. |  | 0.0068336 | 0.0065304 | 0.0086437 | 0.01633 | 0.010488 | 0.021466 | 0.010328 | 0.01472 | 0.11521 | 0.0068313 | 0.016634 | 0.0099482 | 0.0029104 | 0.0053783 | 0.070391 | 0.0064005 | 0.026164 | 0.0065639
| Rel.Std.Dev. |  | 0.6259% | 0.637% | 0.8204% | 1.575% | 0.9848% | 2.067% | 1.016% | 1.365% | 15.96% | 0.6455% | 1.569% | 0.9666% | 0.2957% | 0.4994% | 6.197% | 0.6028% | 2.684% | 0.6375%
| PDM3 |  | 4.715% | -1.673% | 1.059% | -0.5673% | 2.15% | -0.3994% | -2.486% | 3.429% | -30.78% | 1.511% | 1.719% | -1.287% | -5.587% | 3.289% | 8.941% | 1.847% | -6.498% | -1.238%
| Z-Score (Absolute) |  | 1.55 | 0.71 | 0.26 | 0.32 | 0.64 | 0.26 | 0.99 | 1.09 | 10.97 | 0.42 | 0.5 | 0.57 | 2.09 | 1.04 | 3.04 | 0.54 | 2.41 | 0.55
| Analytical results for Li2O in OREAS 753b (Certified Value 2.24 wt.%) | [25 empty] | Certified
Li2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | [6 empty] | wt.%
|  |  | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [6 empty] | 2
3
| 1 | 1 | 2.3361 | 2.1951 | 2.2767 | 2.239 | 2.2603 | 2.1934 | 2.196 | 2.3034 | 1.8728 | 2.2603 | 2.25 | 2.2001 | 2.1273 | 2.3134 | 2.3047 | 2.284 | 2.1462 | 2.21 | [6 empty] | 1
| 1 | 2 | 2.3667 | 2.1936 | 2.2659 | 2.217 | 2.3034 | 2.1704 | 2.174 | 2.3034 | 1.1582 | 2.2819 | 2.27 | 2.228 | 2.1243 | 2.3039 | 2.3998 | 2.2883 | 2.185 | 2.1992 | [6 empty] | 0
| 1 | 3 | 2.3523 | 2.219 | 2.2487 | 2.174 | 2.2819 | 2.2976 | 2.217 | 2.3464 | 1.4165 | 2.2926 | 2.25 | 2.241 | 2.1202 | 2.3128 | 2.5636 | 2.2797 | 2.0838 | 2.2403 | [6 empty] | 16
| 1 | 4 | 2.329 | 2.196 | 2.3008 | 2.282 | 2.2819 | 2.2438 | 2.196 | 2.3249 | 1.6533 | 2.2603 | 2.28 | 2.1979 | 2.1174 | 2.3367 | 2.5103 | 2.2754 | 2.0558 | 2.224 | [6 empty] | 2.2444
| 1 | 5 | 2.3617 | 2.2268 | 2.2556 | 2.239 | 2.3249 | 2.2599 | 2.196 | 2.2819 | 1.6942 | 2.2819 | 2.31 | 2.2345 | 2.1142 | 2.3264 | 2.2526 | 2.2754 | 2.0838 | 2.2153 | [6 empty] | 88
| 1 | 6 | 2.3553 | 2.2104 | 2.261 | 2.239 | 2.3034 | 2.2474 | 2.153 | 2.368 | 1.5263 | 2.2926 | 2.34 | 2.1914 | 2.1105 | 2.316 | 2.6392 | 2.312 | 2.0364 | 2.2107
| Mean |  | 2.3502 | 2.2068 | 2.2681 | 2.2317 | 2.2926 | 2.2354 | 2.1887 | 2.3213 | 1.5535 | 2.2783 | 2.2833 | 2.2155 | 2.119 | 2.3182 | 2.445 | 2.2858 | 2.0985 | 2.2166
| Median |  | 2.3538 | 2.2032 | 2.2635 | 2.239 | 2.2926 | 2.2456 | 2.196 | 2.3142 | 1.5898 | 2.2819 | 2.275 | 2.214 | 2.1188 | 2.3147 | 2.455 | 2.2818 | 2.0838 | 2.213
| Std Dev. |  | 0.014711 | 0.014056 | 0.018609 | 0.035314 | 0.022578 | 0.046211 | 0.022142 | 0.031687 | 0.24801 | 0.014717 | 0.03559 | 0.021425 | 0.0062634 | 0.011578 | 0.15153 | 0.01378 | 0.056324 | 0.01413
| Rel.Std.Dev. |  | 0.6259% | 0.637% | 0.8205% | 1.582% | 0.9848% | 2.067% | 1.012% | 1.365% | 15.96% | 0.646% | 1.559% | 0.9671% | 0.2956% | 0.4994% | 6.197% | 0.6029% | 2.684% | 0.6375%
| PDM3 |  | 4.715% | -1.674% | 1.059% | -0.5657% | 2.15% | -0.3994% | -2.482% | 3.429% | -30.78% | 1.511% | 1.736% | -1.287% | -5.587% | 3.289% | 8.941% | 1.846% | -6.498% | -1.238%
| Z-Score (Absolute) |  | 1.55 | 0.71 | 0.26 | 0.32 | 0.64 | 0.26 | 0.99 | 1.09 | 10.97 | 0.42 | 0.5 | 0.57 | 2.09 | 1.04 | 3.04 | 0.54 | 2.41 | 0.55
| Analytical results for Lu in OREAS 753b (Certified Value 0.066 ppm) | [25 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.10; 1.05: 1.12; 1.06: 1.22; 1.07: 1.25; 1.08: 1.27; 1.09: 1.28 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*OES; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS; 1.09: PF*MS |  | 3
3
| 1 | 1 | 1.01: 0.059216; 1.02: 0.09; 1.03: 0.05; 1.04: < 0.05; 1.05: 0.08; 1.06: < 0.1; 1.07: 0.09; 1.08: < 0.2; 1.09: < 0.5 |  | 1
| 1 | 2 | 1.01: 0.053076; 1.02: 0.08; 1.03: 0.06; 1.04: < 0.05; 1.05: 0.37; 1.06: < 0.1; 1.07: 0.09; 1.08: < 0.2; 1.09: < 0.5 |  | 14
| 1 | 3 | 1.01: 0.054682; 1.02: 0.12; 1.03: 0.05; 1.04: < 0.05; 1.05: 0.22; 1.06: < 0.1; 1.07: 0.07; 1.08: < 0.2; 1.09: < 0.5 |  | 16
| 1 | 4 | 1.01: 0.048592; 1.02: 0.08; 1.03: 0.05; 1.04: < 0.05; 1.05: 0.21; 1.06: < 0.1; 1.07: 0.06; 1.08: < 0.2; 1.09: < 0.5 |  | 0.066035
| 1 | 5 | 1.01: 0.06686; 1.02: 0.08; 1.03: < 0.05; 1.04: < 0.05; 1.05: 0.26; 1.06: < 0.1; 1.07: 0.05; 1.08: < 0.2; 1.09: < 0.5 |  | 89
| 1 | 6 | 1.01: 0.05342; 1.02: 0.09; 1.03: 0.07; 1.04: < 0.05; 1.05: 0.2; 1.06: < 0.1; 1.07: 0.08; 1.08: < 0.2; 1.09: < 0.5
| Mean |  | 1.01: 0.055974; 1.02: 0.09; 1.03: 0.056; 1.05: 0.22333; 1.07: 0.073333
| Median |  | 1.01: 0.054051; 1.02: 0.085; 1.03: 0.05; 1.05: 0.215; 1.07: 0.075
| Std Dev. |  | 1.01: 0.0063262; 1.02: 0.015492; 1.03: 0.0089443; 1.05: 0.09395; 1.07: 0.01633
| Rel.Std.Dev. |  | 1.01: 11.3%; 1.02: 17.21%; 1.03: 15.97%; 1.05: 42.07%; 1.07: 22.27%
| PDM3 |  | 1.01: -15.24%; 1.02: 36.29%; 1.03: -15.2%; 1.05: 238.2%; 1.07: 11.05%
| Z-Score (Absolute) |  | 1.01: 0.5; 1.02: 0.48; 1.03: 0.65; 1.04: 1.4; 1.05: 4.33; 1.06: 0.67; 1.07: 0; 1.08: 0.77; 1.09: 5.11
| Analytical results for Mg in OREAS 753b (Certified Value 0.128 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | 3
3
| 1 | 1 | 0.14 | 0.13455 | 0.14 | 0.13 | 0.13 | 0.13 | 0.12 | 0.13 | 0.12 | 0.12 | 0.2798 | 0.121 | 0.127 | 0.127 | 0.127 | 0.121 | 0.12 | 0.19 | 0.12549 | 0.1 | 0.12 | 0.11 | 0.13324 | 0.12 | 1
| 1 | 2 | 0.13 | 0.13635 | 0.13 | 0.13 | 0.13 | 0.13 | 0.12 | 0.13 | 0.13 | 0.12 | 0.1315 | 0.121 | 0.127 | 0.127 | 0.133 | 0.121 | 0.13 | 0.19 | 0.1262 | 0.11 | 0.12 | 0.12 | 0.13318 | 0.11 | 0
| 1 | 3 | 0.14 | 0.13701 | 0.13 | 0.13 | 0.13 | 0.14 | 0.13 | 0.13 | 0.13 | 0.12 | 0.1538 | 0.121 | 0.133 | 0.127 | 0.133 | 0.121 | 0.12 | 0.19 | 0.12776 | 0.1 | 0.11 | 0.1 | 0.13162 | 0.12 | 16
| 1 | 4 | 0.14 | 0.13217 | 0.13 | 0.13 | 0.13 | 0.14 | 0.12 | 0.13 | 0.13 | 0.12 | 0.2219 | 0.127 | 0.133 | 0.133 | 0.127 | 0.133 | 0.12 | 0.19 | 0.12783 | 0.11 | 0.12 | 0.1 | 0.13134 | 0.13 | 0.12778
| 1 | 5 | 0.14 | 0.1355 | 0.13 | 0.13 | 0.13 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.2099 | 0.121 | 0.133 | 0.127 | 0.133 | 0.133 | 0.13 | 0.19 | 0.12597 | 0.1 | 0.12 | 0.1 | 0.13231 | 0.13 | 90
| 1 | 6 | 0.14 | 0.13303 | 0.13 | 0.13 | 0.13 | 0.14 | 0.12 | 0.14 | 0.12 | 0.12 | 0.1423 | 0.127 | 0.133 | 0.139 | 0.133 | 0.133 | 0.12 | 0.2 | 0.12615 | 0.11 | 0.11 | 0.1 | 0.1322 | 0.11
| Mean |  | 0.13833 | 0.13477 | 0.13167 | 0.13 | 0.13 | 0.135 | 0.12167 | 0.13 | 0.125 | 0.12 | 0.18987 | 0.123 | 0.131 | 0.13 | 0.131 | 0.127 | 0.12333 | 0.19167 | 0.12657 | 0.105 | 0.11667 | 0.105 | 0.13231 | 0.12
| Median |  | 0.14 | 0.13502 | 0.13 | 0.13 | 0.13 | 0.135 | 0.12 | 0.13 | 0.125 | 0.12 | 0.18185 | 0.121 | 0.133 | 0.127 | 0.133 | 0.127 | 0.12 | 0.19 | 0.12618 | 0.105 | 0.12 | 0.1 | 0.13225 | 0.12
| Std Dev. |  | 0.0040825 | 0.0018924 | 0.0040825 | 0 | 0 | 0.0054772 | 0.0040825 | 0.0063246 | 0.0054772 | 0 | 0.05742 | 0.0030984 | 0.0030984 | 0.00502 | 0.0030984 | 0.0065727 | 0.005164 | 0.0040825 | 0.0009813 | 0.0054772 | 0.005164 | 0.0083666 | 0.00078128 | 0.0089443
| Rel.Std.Dev. |  | 2.951% | 1.404% | 3.101% | 0% | 0% | 4.057% | 3.355% | 4.865% | 4.382% | 0% | 30.24% | 2.519% | 2.365% | 3.862% | 2.365% | 5.175% | 4.187% | 2.13% | 0.7753% | 5.216% | 4.426% | 7.968% | 0.5905% | 7.454%
| PDM3 |  | 8.26% | 5.469% | 3.043% | 1.739% | 1.739% | 5.652% | -4.783% | 1.739% | -2.174% | -6.087% | 48.59% | -3.74% | 2.521% | 1.739% | 2.521% | -0.6091% | -3.479% | 50% | -0.9482% | -17.83% | -8.696% | -17.83% | 3.55% | -6.087%
| Z-Score (Absolute) |  | 1.13 | 0.66 | 0.25 | 0.02 | 0.02 | 0.69 | 1.09 | 0.02 | 0.64 | 1.31 | 8.01 | 0.96 | 0.11 | 0.02 | 0.11 | 0.43 | 0.87 | 8.25 | 0.44 | 3.31 | 1.76 | 3.31 | 0.33 | 1.31
| Analytical results for Mn in OREAS 753b (Certified Value 0.135 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 3
3
| 1 | 1 | 0.1362 | 0.1394 | 0.1383 | 0.131 | 0.1317 | 0.14 | 0.138 | 0.137 | 0.1216 | 0.132 | 0.124 | 0.132 | 0.139 | 0.132 | 0.129 | 0.1341 | 0.13727 | 0.1175 | 0.14 | 0.1205 | 0.13885 | 0.14266 |  |  | 1
| 1 | 2 | 0.1335 | 0.1367 | 0.1387 | 0.131 | 0.136 | 0.14 | 0.136 | 0.137 | 0.072 | 0.124 | 0.132 | 0.132 | 0.132 | 0.132 | 0.131 | 0.1342 | 0.13786 | 0.127 | 0.14 | 0.1277 | 0.13794 | 0.13942 |  |  | 0
| 1 | 3 | 0.1339 | 0.1379 | 0.1358 | 0.132 | 0.1373 | 0.14 | 0.14 | 0.139 | 0.0898 | 0.132 | 0.132 | 0.124 | 0.124 | 0.132 | 0.128 | 0.1333 | 0.13612 | 0.1233 | 0.13 | 0.1183 | 0.13925 | 0.14215 |  |  | 16
| 1 | 4 | 0.134 | 0.1404 | 0.1409 | 0.13 | 0.1298 | 0.14 | 0.137 | 0.138 | 0.1069 | 0.132 | 0.139 | 0.132 | 0.132 | 0.132 | 0.131 | 0.1344 | 0.13641 | 0.1277 | 0.14 | 0.1119 | 0.13893 | 0.14861 |  |  | 0.13484
| 1 | 5 | 0.1329 | 0.1353 | 0.1394 | 0.13 | 0.1338 | 0.14 | 0.136 | 0.137 | 0.1371 | 0.124 | 0.139 | 0.132 | 0.139 | 0.139 | 0.129 | 0.1346 | 0.13434 | 0.1231 | 0.13 | 0.1165 | 0.13578 | 0.14957 |  |  | 91
| 1 | 6 | 0.1335 | 0.1366 | 0.1369 | 0.131 | 0.1377 | 0.14 | 0.134 | 0.138 | 0.1185 | 0.139 | 0.132 | 0.132 | 0.139 | 0.139 | 0.13 | 0.1341 | 0.13496 | 0.1258 | 0.14 | 0.1095 | 0.13642 | 0.14761
| Mean |  | 0.134 | 0.13772 | 0.13833 | 0.13083 | 0.13438 | 0.14 | 0.13683 | 0.13767 | 0.10765 | 0.1305 | 0.133 | 0.13067 | 0.13417 | 0.13433 | 0.12967 | 0.13412 | 0.13616 | 0.12407 | 0.13667 | 0.1174 | 0.13786 | 0.145
| Median |  | 0.1337 | 0.1373 | 0.1385 | 0.131 | 0.1349 | 0.14 | 0.1365 | 0.1375 | 0.1127 | 0.132 | 0.132 | 0.132 | 0.1355 | 0.132 | 0.1295 | 0.13415 | 0.13627 | 0.12455 | 0.14 | 0.1174 | 0.13839 | 0.14513
| Std Dev. |  | 0.0011454 | 0.0019073 | 0.0018074 | 0.00075277 | 0.0031833 | 0 | 0.0020412 | 0.0008165 | 0.023536 | 0.0057184 | 0.0055857 | 0.003266 | 0.006047 | 0.0036148 | 0.0012111 | 0.0004446 | 0.0013385 | 0.003727 | 0.005164 | 0.0064804 | 0.0014462 | 0.0041343
| Rel.Std.Dev. |  | 0.8548% | 1.385% | 1.307% | 0.5754% | 2.369% | 0% | 1.492% | 0.5931% | 21.86% | 4.382% | 4.2% | 2.499% | 4.507% | 2.691% | 0.934% | 0.3315% | 0.983% | 3.004% | 3.779% | 5.52% | 1.049% | 2.851%
| PDM3 |  | -0.6223% | 2.134% | 2.591% | -2.971% | -0.3381% | 3.827% | 1.479% | 2.097% | -20.16% | -3.218% | -1.364% | -3.094% | -0.4987% | -0.3751% | -3.836% | -0.5358% | 0.9799% | -7.989% | 1.355% | -12.93% | 2.242% | 7.538%
| Z-Score (Absolute) |  | 0.05 | 0.68 | 0.8 | 0.67 | 0.03 | 1.13 | 0.51 | 0.67 | 5.19 | 0.76 | 0.25 | 0.76 | 0 | 0 | 0.89 | 0.02 | 0.38 | 1.99 | 0.47 | 3.29 | 0.71 | 2.1
| Analytical results for Mo in OREAS 753b (Certified Value 3.5 ppm) | [25 empty] | Certified
Mo | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.04; 1.05: 1.05; 1.06: 1.06; 1.07: 1.07; 1.08: 1.09; 1.09: 1.10; 1.10: 1.12; 1.12: 1.18; 1.13: 1.19; 1.14: 1.22; 1.15: 1.27; 1.16: 1.28 |  |  | ppm
|  |  | 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES/MS; 1.09: PF*OES/MS; 1.10: PF*OES; 1.12: PF*OES; 1.13: PF*OES/MS; 1.14: PF*MS; 1.15: PF*OES; 1.16: PF*OES |  |  | 2
2
| 1 | 1 | 1.03: < 10; 1.04: 4; 1.06: 4; 1.07: 4; 1.08: 4; 1.09: < 5; 1.10: 2; 1.12: < 10; 1.13: 3; 1.14: 5; 1.15: < 30; 1.16: 3.5 |  |  | 1
| 1 | 2 | 1.03: < 10; 1.04: 4; 1.05: 1; 1.06: 3; 1.07: 4; 1.08: 3; 1.09: < 5; 1.10: 4; 1.12: < 10; 1.13: 3; 1.14: 5; 1.15: < 30; 1.16: 3.2 |  |  | 15
| 1 | 3 | 1.03: < 10; 1.04: 4; 1.06: 3; 1.07: 3; 1.08: 4; 1.09: < 5; 1.10: 2; 1.12: < 10; 1.13: 3; 1.14: 5; 1.15: < 30; 1.16: 3.3 |  |  | 16
| 1 | 4 | 1.03: < 10; 1.04: 4; 1.05: 2; 1.06: 3; 1.07: 4; 1.08: 4; 1.09: < 5; 1.10: 2; 1.12: < 10; 1.13: 4; 1.14: 5; 1.15: < 30; 1.16: 3.2 |  |  | 3.4976
| 1 | 5 | 1.03: < 10; 1.04: 4; 1.05: 1; 1.06: 3; 1.07: 3; 1.08: 5; 1.09: < 5; 1.10: 5; 1.12: < 10; 1.13: 4; 1.14: 4; 1.15: < 30; 1.16: 3.5 |  |  | 92
| 1 | 6 | 1.03: < 10; 1.04: 4; 1.05: 1; 1.06: 3; 1.07: 4; 1.08: 4; 1.09: < 5; 1.10: 3; 1.12: < 10; 1.13: 3; 1.14: 5; 1.15: < 30; 1.16: 3.2
| Mean |  | 1.04: 4; 1.05: 1.25; 1.06: 3.1667; 1.07: 3.6667; 1.08: 4; 1.10: 3; 1.13: 3.3333; 1.14: 4.8333; 1.16: 3.3167
| Median |  | 1.04: 4; 1.05: 1; 1.06: 3; 1.07: 4; 1.08: 4; 1.10: 2.5; 1.13: 3; 1.14: 5; 1.16: 3.25
| Std Dev. |  | 1.04: 0; 1.05: 0.5; 1.06: 0.40825; 1.07: 0.5164; 1.08: 0.63246; 1.10: 1.2649; 1.13: 0.5164; 1.14: 0.40825; 1.16: 0.1472
| Rel.Std.Dev. |  | 1.04: 0%; 1.05: 40%; 1.06: 12.89%; 1.07: 14.08%; 1.08: 15.81%; 1.10: 42.16%; 1.13: 15.49%; 1.14: 8.447%; 1.16: 4.438%
| PDM3 |  | 1.04: 14.36%; 1.05: -64.26%; 1.06: -9.462%; 1.07: 4.833%; 1.08: 14.36%; 1.10: -14.23%; 1.13: -4.697%; 1.14: 38.19%; 1.16: -5.174%
| Z-Score (Absolute) |  | 1.03: 1.35; 1.04: 0.34; 1.05: 2.44; 1.06: 0.51; 1.07: 0; 1.08: 0.34; 1.09: 1.18; 1.10: 0.67; 1.12: 1.35; 1.13: 0.34; 1.14: 1.18; 1.15: 11.46; 1.16: 0.35
| Analytical results for Na in OREAS 753b (Indicative Value 0.448 wt.%) | [25 empty] | Indicative
Na | Round | Replicate | 1.03: Lab |  |  | 1
| No | No | 1.03: 1.28 |  |  | wt.%
|  |  | 1.03: PF*OES |  |  | 3
3
| 1 | 1 | 1.03: 0.45 |  |  | 1
| 1 | 2 | 1.03: 0.43 |  |  | 5
| 1 | 3 | 1.03: 0.46 |  |  | 16
| 1 | 4 | 1.03: 0.48 |  |  | 0.44833
| 1 | 5 | 1.03: 0.45 |  |  | 11
| 1 | 6 | 1.03: 0.42
| Mean |  | 1.03: 0.44833
| Median |  | 1.03: 0.45
| Std Dev. |  | 1.03: 0.02137
| Rel.Std.Dev. |  | 1.03: 4.766%
| PDM3 |  | 1.03: 6.661e-14%
Z-Score (Absolute) | bar: 1.03
| Analytical results for Nb in OREAS 753b (Certified Value 102 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | [10 empty] | 0
0
| 1 | 1 | 102 | 105.7285 | 89 | 96 | 74 | 101.2 | 105 | 106.768 | 99.4 | 98 | 102 | 106 | 98.1305 | 101.2 | [10 empty] | 1
| 1 | 2 | 1.03: 103; 1.04: 106.6778; 1.05: 86; 1.06: 103; 1.07: 77; 1.08: 101.9; 1.09: 98; 1.10: 104.639; 1.13: 98; 1.14: 101; 1.15: 105; 1.16: 94.0664; 1.17: 100.8 |  |  | 0
| 1 | 3 | 103 | 103.4415 | 88 | 106 | 77 | 105 | 101 | 108.289 | 101 | 98 | 79 | 104 | 95.654 | 106.4 | [10 empty] | 16
| 1 | 4 | 1.03: 102; 1.04: 105.2657; 1.05: 94; 1.06: 99; 1.07: 75; 1.08: 101.1; 1.09: 97; 1.10: 103.2325; 1.13: 99; 1.14: 101; 1.15: 104; 1.16: 98.7478; 1.17: 103.8 |  |  | 101.5547
| 1 | 5 | 105 | 104.9334 | 90 | 108 | 77 | 102.3 | 102 | 102.4313 | 101 | 99 | 98 | 100 | 96.3003 | 102.5 | [10 empty] | 93
| 1 | 6 | 99 | 106.1758 | 86 | 99 | 74 | 101.5 | 101 | 106.1766 | 111 | 98 | 99 | 103 | 93.4774 | 102.5
| Mean |  | 102.3333 | 105.3704 | 88.8333 | 101.8333 | 75.6667 | 102.1667 | 100.6667 | 105.2561 | 103.1 | 98.3333 | 96.6667 | 103.6667 | 96.0627 | 102.8667
| Median |  | 102.5 | 105.4971 | 88.5 | 101 | 76 | 101.7 | 101 | 105.4078 | 101 | 98 | 100 | 104 | 95.9771 | 102.5
| Std Dev. |  | 1.9664 | 1.1321 | 2.9944 | 4.6224 | 1.5055 | 1.4583 | 2.8752 | 2.2265 | 5.3204 | 0.5164 | 8.7788 | 2.0656 | 2.1156 | 2.0334
| Rel.Std.Dev. |  | 1.922% | 1.074% | 3.371% | 4.539% | 1.99% | 1.427% | 2.856% | 2.115% | 5.16% | 0.5252% | 9.081% | 1.993% | 2.202% | 1.977%
| PDM3 |  | 0.7668% | 3.757% | -12.53% | 0.2744% | -25.49% | 0.6026% | -0.8744% | 3.645% | 1.522% | -3.172% | -4.813% | 2.08% | -5.408% | 1.292%
| Z-Score (Absolute) |  | 0.09 | 0.92 | 3.61 | 0.05 | 7.21 | 0.05 | 0.37 | 0.89 | 0.3 | 1 | 1.46 | 0.46 | 1.63 | 0.24
| Analytical results for Nd in OREAS 753b (Certified Value 4.02 ppm) | [25 empty] | Certified
Nd | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  | 1
| No | No | 1.03: 1.01; 1.04: 1.02; 1.05: 1.03; 1.06: 1.04; 1.07: 1.09; 1.08: 1.10; 1.09: 1.12; 1.10: 1.20; 1.12: 1.22; 1.13: 1.25; 1.14: 1.27; 1.15: 1.28 |  |  | ppm
|  |  | 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES/MS; 1.08: PF*OES/MS; 1.09: PF*OES; 1.10: PF*MS; 1.12: PF*MS; 1.13: PF*MS; 1.14: PF*MS; 1.15: PF*MS |  |  | 2
2
| 1 | 1 | 1.03: 3.9; 1.04: 3.8841; 1.05: 4.4; 1.06: 4.9; 1.07: 4.9; 1.08: 4; 1.09: 2.8; 1.10: 4.53; 1.12: 3.5; 1.13: 4; 1.14: 3.6483; 1.15: 3.92 |  |  | 1
| 1 | 2 | 1.03: 3.9; 1.04: 4.5441; 1.05: 4.1; 1.06: 4.7; 1.07: 4.5; 1.08: 4; 1.09: 3.3; 1.10: 4.5866; 1.12: 3.7; 1.13: 3.9; 1.14: 2.9939; 1.15: 3.82 |  |  | 16
| 1 | 3 | 1.03: 4.4; 1.04: 4.31; 1.05: 4.4; 1.06: 4.5; 1.07: 4.4; 1.08: 4; 1.09: 3.2; 1.10: 4.6528; 1.12: 3.6; 1.13: 3.9; 1.14: 3.4764; 1.15: 3.32 |  |  | 16
| 1 | 4 | 1.03: 4.4; 1.04: 3.8532; 1.05: 4.6; 1.06: 4.2; 1.07: 4.7; 1.08: 4; 1.09: 3.7; 1.10: 4.5867; 1.12: 3.7; 1.13: 3.3; 1.14: 3.3682; 1.15: 2.99 |  |  | 4.0208
| 1 | 5 | 1.03: 4; 1.04: 3.9185; 1.05: 4.3; 1.06: 4.4; 1.07: 4.8; 1.08: 3; 1.09: 4.4; 1.10: 4.5245; 1.12: 3.5; 1.13: 3.9; 1.14: 4.0989; 1.15: 3.6 |  |  | 94
| 1 | 6 | 1.03: 3.9; 1.04: 3.9521; 1.05: 4.3; 1.06: 4.1; 1.07: 4.2; 1.08: 4; 1.09: 4; 1.10: 4.5687; 1.12: 4; 1.13: 3.7; 1.14: 3.2468; 1.15: 3.96
| Mean |  | 1.03: 4.0833; 1.04: 4.077; 1.05: 4.35; 1.06: 4.4667; 1.07: 4.5833; 1.08: 3.8333; 1.09: 3.5667; 1.10: 4.5749; 1.12: 3.6667; 1.13: 3.7833; 1.14: 3.4721; 1.15: 3.6017
| Median |  | 1.03: 3.95; 1.04: 3.9353; 1.05: 4.35; 1.06: 4.45; 1.07: 4.6; 1.08: 4; 1.09: 3.5; 1.10: 4.5777; 1.12: 3.65; 1.13: 3.9; 1.14: 3.4223; 1.15: 3.71
| Std Dev. |  | 1.03: 0.24833; 1.04: 0.28302; 1.05: 0.16432; 1.06: 0.30111; 1.07: 0.26394; 1.08: 0.40825; 1.09: 0.58195; 1.10: 0.046778; 1.12: 0.18619; 1.13: 0.25626; 1.14: 0.37792; 1.15: 0.38233
| Rel.Std.Dev. |  | 1.03: 6.081%; 1.04: 6.942%; 1.05: 3.777%; 1.06: 6.741%; 1.07: 5.759%; 1.08: 10.65%; 1.09: 16.32%; 1.10: 1.022%; 1.12: 5.078%; 1.13: 6.773%; 1.14: 10.88%; 1.15: 10.62%
| PDM3 |  | 1.03: 1.555%; 1.04: 1.398%; 1.05: 8.187%; 1.06: 11.09%; 1.07: 13.99%; 1.08: -4.663%; 1.09: -11.29%; 1.10: 13.78%; 1.12: -8.808%; 1.13: -5.906%; 1.14: -13.65%; 1.15: -10.42%
| Z-Score (Absolute) |  | 1.03: 0.01; 1.04: 0; 1.05: 0.45; 1.06: 0.64; 1.07: 0.83; 1.08: NA; 1.09: 0.84; 1.10: 0.82; 1.12: 0.67; 1.13: 0.48; 1.14: 0.99; 1.15: 0.78
Results from laboratory 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 753b (Indicative Value 17.3 ppm) | [25 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.27 | 1.28 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 1
1
| 1 | 1 | < 50 | 11 |  | < 20 | < 20 | < 50 | 20 | < 10 | < 5 | < 20 | < 20 | 20 | 30 | < 20 | 10 | 11 | < 50 | < 20 | 9.16 | [5 empty] | 1
| 1 | 2 | < 50 | 13 | 37 | < 20 | < 20 | < 50 | 20 | < 10 | 7 | 20 | 20 | 20 | 100 | < 20 | 10 | 10 | < 50 | 31.43 | 9.14 | [5 empty] | 6
| 1 | 3 | < 50 | 11 | 20 | < 20 | < 20 | < 50 | 10 | < 10 | 14 | < 20 | < 20 | 30 | < 20 | < 20 | 10 | 16 | < 50 | 30.56 | 9.18 | [5 empty] | 16
| 1 | 4 | < 50 | 16 | 22 | < 20 | < 20 | < 50 | 10 | < 10 | 6 | < 20 | < 20 | < 20 | 70 | < 20 | 10 | 13 | < 50 | 39.01 | 9.29 | [5 empty] | 17.2705
| 1 | 5 | < 50 | 11 | 22 | < 20 | < 20 | < 50 | 20 | < 10 | 9 | < 20 | 20 | < 20 | 40 | < 20 | 10 | 19 | < 50 | 25.01 | 9.21 | [5 empty] | 12
| 1 | 6 | < 50 | 11 | 55 | < 20 | < 20 | < 50 | 10 | < 10 | 7 | 20 | 20 | < 20 | 40 | < 20 | 10 | 32 | < 50 | 32.5 | 9.26
| Mean |  |  | 12.1667 | 31.2 |  |  |  | 15 |  | 8.6 | 20 | 20 | 23.3333 | 56 |  | 10 | 16.8333 |  | 31.702 | 9.2067
| Median |  |  | 11 | 22 |  |  |  | 15 |  | 7 | 20 | 20 | 20 | 40 |  | 10 | 14.5 |  | 31.43 | 9.195
| Std Dev. |  |  | 2.0412 | 14.9566 |  |  |  | 5.4772 |  | 3.2094 | 0 | 0 | 5.7735 | 28.8097 |  | 0 | 8.1343 |  | 5.0051 | 0.058538
| Rel.Std.Dev. |  |  | 16.78% | 47.94% |  |  |  | 36.51% |  | 37.32% | 0% | 0% | 24.74% | 51.45% |  | 0% | 48.32% |  | 15.79% | 0.6358%
| PDM3 |  |  | -29.55% | 80.66% |  |  |  | -13.15% |  | -50.2% | 15.8% | 15.8% | 35.11% | 224.3% |  | -42.1% | -2.531% |  | 83.56% | -46.69%
| Z-Score (Absolute) |  | 1.35 | 0.38 | 2.18 | 0.67 | 0.67 | 1.35 | 0 | 1.35 | 1 | 0.22 | 0 | 0.22 | 4.5 | 0.67 | 0.67 | 0.25 | 1.35 | 1.76 | 0.78
| Analytical results for P in OREAS 753b (Certified Value 0.047 wt.%) | [25 empty] | Certified
P | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab | [5 empty] | 1
| No | No | 1.03: 1.01; 1.04: 1.02; 1.05: 1.04; 1.06: 1.06; 1.07: 1.10; 1.08: 1.12; 1.09: 1.18; 1.10: 1.19; 1.12: 1.20; 1.13: 1.22; 1.14: 1.28 | [5 empty] | wt.%
|  |  | 1.03: PF*OES; 1.04: PF*OES/MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*OES/MS; 1.08: PF*OES; 1.09: PF*OES; 1.10: PF*OES/MS; 1.12: PF*OES; 1.13: PF*OES; 1.14: PF*OES | [5 empty] | 3
3
| 1 | 1 | 1.03: 0.05; 1.04: 0.047087; 1.05: 0.05; 1.06: 0.05; 1.07: 0.04; 1.08: 0.0375; 1.09: 0.05; 1.10: 0.045; 1.12: 0.045035; 1.13: 0.04; 1.14: 0.11 | [5 empty] | 1
| 1 | 2 | 1.03: 0.05; 1.04: 0.049574; 1.05: 0.05; 1.06: 0.05; 1.07: 0.04; 1.08: 0.0319; 1.09: 0.05; 1.10: 0.045; 1.12: 0.044943; 1.13: 0.04; 1.14: 0.13 | [5 empty] | 0
| 1 | 3 | 1.03: 0.05; 1.04: 0.044993; 1.05: 0.05; 1.06: 0.05; 1.07: 0.05; 1.08: 0.0327; 1.09: 0.05; 1.10: 0.0447; 1.12: 0.044965; 1.13: 0.04; 1.14: 0.12 | [5 empty] | 16
| 1 | 4 | 1.03: 0.05; 1.04: 0.046334; 1.05: 0.05; 1.06: 0.05; 1.07: 0.05; 1.08: 0.0214; 1.09: 0.05; 1.10: 0.045; 1.12: 0.045573; 1.13: 0.04; 1.14: 0.13 | [5 empty] | 0.047133
| 1 | 5 | 1.03: 0.05; 1.04: 0.047799; 1.05: 0.05; 1.06: 0.05; 1.07: 0.04; 1.08: 0.0537; 1.09: 0.05; 1.10: 0.0447; 1.12: 0.045402; 1.13: 0.04; 1.14: 0.13 | [5 empty] | 95
| 1 | 6 | 1.03: 0.05; 1.04: 0.049438; 1.05: 0.05; 1.06: 0.05; 1.07: 0.05; 1.08: 0.0375; 1.09: 0.05; 1.10: 0.0446; 1.12: 0.045044; 1.13: 0.05; 1.14: 0.11
| Mean |  | 1.03: 0.05; 1.04: 0.047537; 1.05: 0.05; 1.06: 0.05; 1.07: 0.045; 1.08: 0.035783; 1.09: 0.05; 1.10: 0.044833; 1.12: 0.04516; 1.13: 0.041667; 1.14: 0.12167
| Median |  | 1.03: 0.05; 1.04: 0.047443; 1.05: 0.05; 1.06: 0.05; 1.07: 0.045; 1.08: 0.0351; 1.09: 0.05; 1.10: 0.04485; 1.12: 0.04504; 1.13: 0.04; 1.14: 0.125
| Std Dev. |  | 1.03: 7.6012e-18; 1.04: 0.0017863; 1.05: 7.6012e-18; 1.06: 7.6012e-18; 1.07: 0.0054772; 1.08: 0.010568; 1.09: 7.6012e-18; 1.10: 0.00018619; 1.12: 0.00026191; 1.13: 0.0040825; 1.14: 0.0098319
| Rel.Std.Dev. |  | 1.03: 1.52e-14%; 1.04: 3.758%; 1.05: 1.52e-14%; 1.06: 1.52e-14%; 1.07: 12.17%; 1.08: 29.53%; 1.09: 1.52e-14%; 1.10: 0.4153%; 1.12: 0.58%; 1.13: 9.798%; 1.14: 8.081%
| PDM3 |  | 1.03: 6.083%; 1.04: 0.8579%; 1.05: 6.083%; 1.06: 6.083%; 1.07: -4.526%; 1.08: -24.08%; 1.09: 6.083%; 1.10: -4.879%; 1.12: -4.185%; 1.13: -11.6%; 1.14: 158.1%
| Z-Score (Absolute) |  | 1.03: 0.67; 1.04: 0; 1.05: 0.67; 1.06: 0.67; 1.07: 0.69; 1.08: 3.22; 1.09: 0.67; 1.10: 0.74; 1.12: 0.65; 1.13: 1.61; 1.14: 20.3
| Analytical results for Pb in OREAS 753b (Indicative Value 5.89 ppm) | [25 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  |  | 2
2
| 1 | 1 | < 100 | 5 |  | < 20 | < 20 | < 100 | 4.2 | 5 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | < 5 | < 20 | < 300 | < 100 | 10.8 |  |  |  |  | 1
| 1 | 2 | < 100 | 6 | 24 | < 20 | < 20 | < 100 | 5.5 | 4 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | < 5 | < 20 | < 300 | < 100 | 12 |  |  |  |  | 7
| 1 | 3 | < 100 | 5 |  | < 20 | < 20 | < 100 | 4.5 | 5 | 2 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | < 5 | < 20 | < 300 | < 100 | 13.2 |  |  |  |  | 16
| 1 | 4 | < 100 | 5 |  | < 20 | < 20 | < 100 | 4.8 | 4 | 4 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | < 5 | < 20 | < 300 | < 100 | 13.1 |  |  |  |  | 5.8873
| 1 | 5 | < 100 | 5 |  | < 20 | < 20 | < 100 | 3.6 | 4 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | < 5 | < 20 | < 300 | < 100 | 17.4 |  |  |  |  | 13
| 1 | 6 | < 100 | 6 | 30 | < 20 | < 20 | < 100 | 5.3 | 4 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | < 20 | < 5 | < 20 | < 300 | < 100 | 11.5
| Mean |  | 1.04: 5.3333; 1.05: 27; 1.09: 4.65; 1.10: 4.3333; 1.12: 3; 1.17: 100; 1.28: 13
| Median |  | 1.04: 5; 1.05: 27; 1.09: 4.65; 1.10: 4; 1.12: 3; 1.17: 100; 1.28: 12.55
| Std Dev. |  | 1.04: 0.5164; 1.05: 4.2426; 1.09: 0.7064; 1.10: 0.5164; 1.12: 0.63246; 1.17: 0; 1.28: 2.3452
| Rel.Std.Dev. |  | 1.04: 9.682%; 1.05: 15.71%; 1.09: 15.19%; 1.10: 11.92%; 1.12: 21.08%; 1.17: 0%; 1.28: 18.04%
| PDM3 |  | 1.04: -9.41%; 1.05: 358.6%; 1.09: -21.02%; 1.10: -26.4%; 1.12: -49.04%; 1.17: 1599%; 1.28: 120.8%
| Z-Score (Absolute) |  | 1.17 | 0.57 | 0.27 | 0.39 | 0.39 | 1.17 | 0.6 | 0.61 | 0.66 | 1.17 | 1.17 | 1.17 | 1.17 | 3.13 | 0.39 | 0.68 | 0.39 | 5.08 | 1.17 | 0.27
| Analytical results for Pr in OREAS 753b (Certified Value 1.08 ppm) | [25 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | 1.1 | 1.1245 | 1.22 | 1.26 | 1.1 | 0.96 | 0.86 | 1.2444 | 0.97 | 1.16 | 0.95033 | 1.3 | [12 empty] | 1
| 1 | 2 | 1.1 | 1.178 | 1.25 | 1.19 | 1.1 | 0.99 | 1.03 | 1.2099 | 0.97 | 1.16 | 0.78881 | < 1 | [12 empty] | 19
| 1 | 3 | 1.1 | 1.0984 | 1.2 | 1.21 | 1 | 0.95 | 0.9 | 1.2424 | 0.96 | 1.08 | 0.88119 | < 1 | [12 empty] | 16
| 1 | 4 | 1.1 | 1.0806 | 1.22 | 1.1 | 1.2 | 0.85 | 1.15 | 1.21 | 0.98 | 1.01 | 0.80719 | < 1 | [12 empty] | 1.0836
| 1 | 5 | 1.1 | 1.1084 | 1.21 | 1 | 1.2 | 0.88 | 1.39 | 1.1526 | 0.89 | 1.05 | 0.94923 | < 1 | [12 empty] | 96
| 1 | 6 | 1 | 1.0464 | 1.23 | 1.1 | 1 | 0.97 | 1.22 | 1.1565 | 1.03 | 1.08 | 0.80959 | 1.1
| Mean |  | 1.0833 | 1.1061 | 1.2217 | 1.1433 | 1.1 | 0.93333 | 1.0917 | 1.2026 | 0.96667 | 1.09 | 0.86439 | 1.2
| Median |  | 1.1 | 1.1034 | 1.22 | 1.145 | 1.1 | 0.955 | 1.09 | 1.2099 | 0.97 | 1.08 | 0.84539 | 1.2
| Std Dev. |  | 0.040825 | 0.044248 | 0.017224 | 0.094375 | 0.089443 | 0.055377 | 0.20154 | 0.040155 | 0.045019 | 0.06 | 0.073285 | 0.14142
| Rel.Std.Dev. |  | 3.768% | 4.001% | 1.41% | 8.254% | 8.131% | 5.933% | 18.46% | 3.339% | 4.657% | 5.505% | 8.478% | 11.79%
| PDM3 |  | -0.02359% | 2.073% | 12.74% | 5.514% | 1.515% | -13.87% | 0.7455% | 10.99% | -10.79% | 0.5917% | -20.23% | 10.74%
| Z-Score (Absolute) |  | 0.06 | 0.12 | 1.07 | 0.43 | 0.08 | 1.29 | 0.01 | 0.92 | 1.02 | 0.01 | 1.86 | 2.93
| Analytical results for Rb in OREAS 753b (Certified Value 1226 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.20 | 1.22 | 1.25 | [11 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [11 empty] | 0
0
| 1 | 1 | 1239.9 | 1241.1571 | 1251 | 1213 | 1275.8 | 1390 | 1244.5 | 1230 | 1190 | 1240 | 1436.9286 | 1168 | 1400 | [11 empty] | 1
| 1 | 2 | 1243.1 | 1241.0563 | 1265 | 1174 | 1289.5 | 1420 | 1224 | 1240 | 1240 | 1240 | 1446.6207 | 1172 | 1399 | [11 empty] | 0
| 1 | 3 | 1224.3 | 1226.6424 | 1285 | 1205 | 1289.4 | 1390 | 1316.1 | 1250 | 1190 | 1150 | 1448.6498 | 761 | 1385 | [11 empty] | 16
| 1 | 4 | 1217.3 | 1222.5428 | 1259 | 1251 | 1285 | 1410 | 1205.1 | 1230 | 1210 | 1130 | 1441.5763 | 1141 | 1395 | [11 empty] | 1226.4281
| 1 | 5 | 1220.4 | 1224.5449 | 1270 | 1210 | 1282.3 | 1420 | 1272.1 | 1250 | 1210 | 1200 | 1438.7353 | 1142 | 1389 | [11 empty] | 97
| 1 | 6 | 1199.2 | 1240.7447 | 1289 | 1276 | 1288.5 | 1410 | 1306.9 | 1220 | 1180 | 1120 | 1442.97 | 1125 | 1389
| Mean |  | 1224.0333 | 1232.7814 | 1269.8333 | 1221.5 | 1285.0833 | 1406.6667 | 1261.45 | 1236.6667 | 1203.3333 | 1180 | 1442.5801 | 1084.8333 | 1392.8333
| Median |  | 1222.35 | 1233.6936 | 1267.5 | 1211.5 | 1286.75 | 1410 | 1258.3 | 1235 | 1200 | 1175 | 1442.2731 | 1141.5 | 1392
| Std Dev. |  | 16.0661 | 9.0818 | 14.784 | 36.2588 | 5.3619 | 13.6626 | 44.7903 | 12.1106 | 21.6025 | 54.037 | 4.4955 | 159.6364 | 6.08
| Rel.Std.Dev. |  | 1.313% | 0.7367% | 1.164% | 2.968% | 0.4172% | 0.9713% | 3.551% | 0.9793% | 1.795% | 4.579% | 0.3116% | 14.72% | 0.4365%
| PDM3 |  | -0.1953% | 0.518% | 3.539% | -0.4018% | 4.783% | 14.7% | 2.856% | 0.8348% | -1.883% | -3.786% | 17.62% | -11.55% | 13.57%
| Z-Score (Absolute) |  | 0.26 | 0.08 | 0.67 | 0.31 | 0.98 | 3.44 | 0.5 | 0 | 0.67 | 1.15 | 4.17 | 3.07 | 3.16
| Analytical results for Re in OREAS 753b (Certified Value < 0.1 ppm) | [25 empty] | Certified
Re | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.06; 1.05: 1.07; 1.06: 1.12; 1.07: 1.22; 1.08: 1.25 |  |  |  |  | ppm
|  |  | 1.03: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES; 1.07: PF*MS; 1.08: PF*MS |  |  |  |  | 2
2
| 1 | 1 | 1.03: < 0.01; 1.04: < 0.1; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.05; 1.08: < 0.05 |  |  |  |  | 1
| 1 | 2 | 1.03: < 0.01; 1.04: < 0.1; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.05; 1.08: < 0.05 |  |  |  |  | 20
| 1 | 3 | 1.03: < 0.01; 1.04: < 0.1; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.05; 1.08: < 0.05 |  |  |  |  | 16
| 1 | 4 | 1.03: < 0.01; 1.04: < 0.1; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.05; 1.08: < 0.05 |  |  |  |  | < 0.1
| 1 | 5 | 1.03: < 0.01; 1.04: < 0.1; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.05; 1.08: < 0.05 |  |  |  |  | 98
| 1 | 6 | 1.03: < 0.01; 1.04: < 0.1; 1.05: < 0.1; 1.06: < 0.1; 1.07: < 0.05; 1.08: < 0.05
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 1.03: 1.75; 1.04: 0.67; 1.05: 0.67; 1.06: 0.67; 1.07: 0.67; 1.08: 0.67
| Analytical results for S in OREAS 753b (Indicative Value 0.028 wt.%) | [25 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.03 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | [6 empty] | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [6 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | 0.03 | 0.03 | 0.02 | 0.01 | 0.02 | 0.02 | < 0.01 | 0.03 | 0.03 | 0.02 | 0.01 | 0.02 | < 0.1 | 0.03 | [6 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.05 | 0.03 | 0.02 | < 0.01 | 0.05 | 0.01 | 0.03 | 0.03 | < 0.01 | 0.04 | 0.03 | 0.03 | < 0.01 | 0.02 | < 0.1 | 0.02 | [6 empty] | 8
| 1 | 3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | 0.02 | 0.02 | 0.03 | 0.04 | 0.03 | 0.04 | < 0.01 | 0.04 | 0.03 | 0.02 | < 0.01 | 0.02 | < 0.1 | 0.02 | [6 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.05 | 0.04 | 0.02 | 0.02 | 0.03 | 0.06 | 0.03 | 0.04 | < 0.01 | 0.03 | 0.03 | 0.01 | 0.03 | 0.01 | < 0.1 | 0.02 | [6 empty] | 0.028141
| 1 | 5 | < 0.05 | < 0.05 | < 0.05 | 0.04 | 0.02 | 0.05 | 0.04 | 0.04 | 0.03 | 0.02 | < 0.01 | 0.04 | 0.03 | 0.03 | 0.02 | 0.01 | < 0.1 | 0.02 | [6 empty] | 14
| 1 | 6 | < 0.05 | < 0.05 | < 0.05 | 0.02 | 0.02 | 0.04 | 0.03 | 0.05 | 0.03 | 0.05 | < 0.01 | 0.03 | 0.03 | 0.02 | 0.03 | 0.01 | < 0.1 | 0.01
| Mean |  |  |  |  | 0.038 | 0.021667 | 0.032 | 0.033333 | 0.035 | 0.028333 | 0.033333 |  | 0.035 | 0.03 | 0.021667 | 0.0225 | 0.015 |  | 0.02
| Median |  |  |  |  | 0.04 | 0.02 | 0.03 | 0.03 | 0.04 | 0.03 | 0.035 |  | 0.035 | 0.03 | 0.02 | 0.025 | 0.015 |  | 0.02
| Std Dev. |  |  |  |  | 0.014832 | 0.0040825 | 0.013038 | 0.010328 | 0.020736 | 0.0040825 | 0.012111 |  | 0.0054772 | 0 | 0.0075277 | 0.0095743 | 0.0054772 |  | 0.0063246
| Rel.Std.Dev. |  |  |  |  | 39.03% | 18.84% | 40.75% | 30.98% | 59.25% | 14.41% | 36.33% |  | 15.65% | 0% | 34.74% | 42.55% | 36.51% |  | 31.62%
| PDM3 |  |  |  |  | 35.03% | -23.01% | 13.71% | 18.45% | 24.37% | 0.6834% | 18.45% |  | 24.37% | 6.606% | -23.01% | -20.05% | -46.7% |  | -28.93%
| Z-Score (Absolute) |  | 0.13 | 0.13 | 0.13 | 0.67 | 0.49 | 0.13 | 0.76 | 0.94 | 0.22 | 0.76 | 2.29 | 0.94 | 0.4 | 0.49 | 1.03 | 1.21 | 2.56 | 0.67
| Analytical results for Sb in OREAS 753b (Certified Value 0.67 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.18 | 1.22 | 1.25 | 1.28 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | [12 empty] | 2
2
| 1 | 1 | 1.03: 0.6; 1.04: < 4; 1.05: 0.7; 1.07: 0.6; 1.08: < 2; 1.09: < 1; 1.10: 0.6; 1.12: < 50; 1.13: 0.6; 1.14: < 1; 1.15: < 5 |  |  |  |  | 1
| 1 | 2 | 0.8 | < 4 | 0.7 | 0.7 | 0.7 | < 2 | < 1 | 0.6 | < 50 | 0.6 | < 1 | < 5 | [12 empty] | 21
| 1 | 3 | 0.6 | < 4 | 0.7 | 0.6 | 0.7 | < 2 | < 1 | 0.4 | < 50 | 1.3 | < 1 | < 5 | [12 empty] | 16
| 1 | 4 | 0.8 | < 4 | 0.7 | 0.7 | 0.7 | < 2 | < 1 | 0.5 | < 50 | 0.8 | < 1 | < 5 | [12 empty] | 0.66778
| 1 | 5 | 1.03: 0.8; 1.04: < 4; 1.05: 0.7; 1.07: 0.7; 1.08: < 2; 1.09: < 1; 1.10: 0.7; 1.12: < 50; 1.13: 0.6; 1.14: < 1; 1.15: < 5 |  |  |  |  | 99
| 1 | 6 | 1.03: 0.9; 1.04: < 4; 1.05: 0.7; 1.07: 0.7; 1.08: < 2; 1.09: < 1; 1.10: 0.6; 1.12: < 50; 1.13: 0.6; 1.14: < 1; 1.15: < 5
| Mean |  | 1.03: 0.75; 1.05: 0.7; 1.06: 0.66667; 1.07: 0.68333; 1.10: 0.56667; 1.13: 0.75
| Median |  | 1.03: 0.8; 1.05: 0.7; 1.06: 0.7; 1.07: 0.7; 1.10: 0.6; 1.13: 0.6
| Std Dev. |  | 1.03: 0.12247; 1.05: 1.2162e-16; 1.06: 0.057735; 1.07: 0.040825; 1.10: 0.10328; 1.13: 0.28107
| Rel.Std.Dev. |  | 1.03: 16.33%; 1.05: 1.737e-14%; 1.06: 8.66%; 1.07: 5.974%; 1.10: 18.23%; 1.13: 37.48%
| PDM3 |  | 1.03: 12.31%; 1.05: 4.825%; 1.06: -0.1664%; 1.07: 2.329%; 1.10: -15.14%; 1.13: 12.31%
| Z-Score (Absolute) |  | 0.09 | 4.49 | 0.09 | 0.21 | 0.15 | 0.97 | 0.79 | 0.56 | 85.4 | 0.09 | 0.79 | 6.24
| Analytical results for Sc in OREAS 753b (Indicative Value < 5 ppm) | [25 empty] | Indicative
Sc | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab |  |  |  |  | 1
| No | No | 1.03: 1.04; 1.04: 1.06; 1.05: 1.07; 1.06: 1.10; 1.07: 1.12; 1.08: 1.18; 1.09: 1.19; 1.10: 1.22; 1.12: 1.27; 1.13: 1.28 |  |  |  |  | ppm
|  |  | 1.03: PF*OES; 1.04: PF*OES/MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*OES; 1.08: PF*OES; 1.09: PF*OES/MS; 1.10: PF*OES; 1.12: PF*OES; 1.13: PF*OES |  |  |  |  | 2
2
| 1 | 1 | 1.03: < 5; 1.04: < 20; 1.05: < 20; 1.06: < 10; 1.07: 6; 1.08: < 5; 1.09: < 5; 1.10: < 5; 1.12: < 30; 1.13: 1.3 |  |  |  |  | 1
| 1 | 2 | 1.03: < 5; 1.04: < 20; 1.05: < 20; 1.06: < 10; 1.07: 7; 1.08: < 5; 1.09: < 5; 1.10: < 5; 1.12: < 30; 1.13: 1.1 |  |  |  |  | 22
| 1 | 3 | 1.03: < 5; 1.04: < 20; 1.05: < 20; 1.06: < 10; 1.07: < 5; 1.08: < 5; 1.09: 5; 1.10: < 5; 1.12: < 30; 1.13: 1.3 |  |  |  |  | 16
| 1 | 4 | 1.03: < 5; 1.04: < 20; 1.05: < 20; 1.06: < 10; 1.07: 7; 1.08: < 5; 1.09: 5; 1.10: < 5; 1.12: < 30; 1.13: 1.1 |  |  |  |  | < 5
| 1 | 5 | 1.03: < 5; 1.04: < 20; 1.05: < 20; 1.06: < 10; 1.07: 12; 1.08: < 5; 1.09: 5; 1.10: < 5; 1.12: < 30; 1.13: 1.4 |  |  |  |  | 10
| 1 | 6 | 1.03: < 5; 1.04: < 20; 1.05: < 20; 1.06: < 10; 1.07: 11; 1.08: < 5; 1.09: 5; 1.10: < 5; 1.12: < 30; 1.13: 1.2
| Mean |  | 1.07: 8.6; 1.09: 5; 1.13: 1.2333
| Median |  | 1.07: 7; 1.09: 5; 1.13: 1.25
| Std Dev. |  | 1.07: 2.7019; 1.09: 0; 1.13: 0.12111
| Rel.Std.Dev. |  | 1.07: 31.42%; 1.09: 0%; 1.13: 9.819%
PDM3
| Z-Score (Absolute) |  | 1.03: 0.55; 1.04: 1.44; 1.05: 1.44; 1.06: 0.11; 1.07: 0.8; 1.08: 0.55; 1.09: 0.11; 1.10: 0.55; 1.12: 2.76; 1.13: 0.89
| Analytical results for Se in OREAS 753b (Indicative Value < 10 ppm) | [25 empty] | Indicative
Se | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  |  | 1
| No | No | 1.03: 1.03; 1.04: 1.07; 1.05: 1.09; 1.06: 1.10; 1.07: 1.12; 1.08: 1.22 |  |  |  |  | ppm
|  |  | 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*OES/MS; 1.07: PF*OES; 1.08: PF*MS |  |  |  |  | 1
1
| 1 | 1 | 1.03: < 10; 1.04: < 20; 1.05: < 8; 1.06: < 5; 1.07: 1; 1.08: < 50 |  |  |  |  | 1
| 1 | 2 | 1.03: < 10; 1.04: < 20; 1.05: < 8; 1.06: < 5; 1.07: 1; 1.08: < 50 |  |  |  |  | 2
| 1 | 3 | 1.03: < 10; 1.04: < 20; 1.05: < 8; 1.06: < 5; 1.07: 2; 1.08: < 50 |  |  |  |  | 16
| 1 | 4 | 1.03: < 10; 1.04: < 20; 1.05: < 8; 1.06: < 5; 1.07: < 1; 1.08: < 50 |  |  |  |  | < 10
| 1 | 5 | 1.03: < 10; 1.04: < 20; 1.05: < 8; 1.06: < 5; 1.07: 3; 1.08: < 50 |  |  |  |  | 11
| 1 | 6 | 1.03: < 10; 1.04: < 20; 1.05: < 8; 1.06: < 5; 1.07: 3; 1.08: < 50
| Mean |  | 1.07: 2
| Median |  | 1.07: 2
Std Dev. | bar: 1.07
| Rel.Std.Dev. |  | 1.07: 50%
PDM3
| Z-Score (Absolute) |  | 1.03: 0.14; 1.04: 1.56; 1.05: 0.14; 1.06: 0.57; 1.07: 0.78; 1.08: 5.82
| Analytical results for Si in OREAS 753b (Certified Value 38.08 wt.%) | [25 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES |  |  |  |  | 2
3
| 1 | 1 | 38.9 | 38.2745 | 37.7 | > 30 | 40.2 | 30.6 | 35.73 | 37.8 | 37.9 | 33.9847 | 37.4 | 38 | 37.58 | 39.59 | 38.71 | > 30 | 38.35 | > 30 | > 30 | 36.639 |  |  |  |  | 1
| 1 | 2 | 38.3 | 38.4296 | 37.1 | > 30 | 40.4 | 31.5 | 36.38 | 38.3 | 37.8 | 20.4282 | 37.12 | 40.11 | 37.58 | 39.69 | 39.08 | > 30 | 38.17 | > 30 | > 30 | 37.7441 |  |  |  |  | 0
| 1 | 3 | 39.9 | 36.9032 | 36 | > 30 | 39.9 | 31.8 | 36.58 | 37.3 | 38.2 | 25.1963 | 37.82 | 37.4 | 37.21 | 39.92 | 38.61 | > 30 | 38.09 | > 30 | > 30 | 37.196 |  |  |  |  | 16
| 1 | 4 | 39.1 | 37.7498 | 37 | > 30 | 41 | 30.2 | 36.04 | 38.1 | 38.1 | 30.1515 | 38.19 | 37.91 | 38.29 | 39.59 | 38.89 | > 30 | 38.39 | > 30 | > 30 | 36.891 |  |  |  |  | 38.0785
| 1 | 5 | 38.9 | 37.7646 | 37.7 | > 30 | 41 | 30.9 | 34.89 | 37.7 | 38 | 41.7913 | 37.49 | 38.1 | 38.19 | 39.78 | 40.11 | > 30 | 38.08 | > 30 | > 30 | 36.2246 |  |  |  |  | 100
| 1 | 6 | 38.8 | 37.8177 | 36.4 | > 30 | 40.1 | 30.7 | 35.07 | 37.7 | 38.5 | 32.255 | 37.02 | 37.68 | 38.38 | 40.01 | 40.2 | > 30 | 38.16 | > 30 | > 30 | 35.8522
| Mean |  | 38.9833 | 37.8232 | 36.9833 |  | 40.4333 | 30.95 | 35.7817 | 37.8167 | 38.0833 | 30.6345 | 37.5067 | 38.2 | 37.8717 | 39.7633 | 39.2667 |  | 38.2067 |  |  | 36.7578
| Median |  | 38.9 | 37.7911 | 37.05 |  | 40.3 | 30.8 | 35.885 | 37.75 | 38.05 | 31.2033 | 37.445 | 37.955 | 37.885 | 39.735 | 38.985 |  | 38.165 |  |  | 36.765
| Std Dev. |  | 0.52313 | 0.53415 | 0.68532 |  | 0.46762 | 0.59582 | 0.68794 | 0.34881 | 0.24833 | 7.3797 | 0.43889 | 0.96877 | 0.47806 | 0.17386 | 0.70713 |  | 0.13216 |  |  | 0.6784
| Rel.Std.Dev. |  | 1.342% | 1.412% | 1.853% |  | 1.157% | 1.925% | 1.923% | 0.9224% | 0.6521% | 24.09% | 1.17% | 2.536% | 1.262% | 0.4372% | 1.801% |  | 0.3459% |  |  | 1.846%
| PDM3 |  | 2.376% | -0.6703% | -2.876% |  | 6.184% | -18.72% | -6.032% | -0.6875% | 0.01278% | -19.55% | -1.502% | 0.3192% | -0.5431% | 4.425% | 3.12% |  | 0.3367% |  |  | -3.468%
| Z-Score (Absolute) |  | 0.78 | 0.02 | 0.6 | NA | 1.78 | 4.76 | 1.43 | 0.02 | 0.16 | 4.98 | 0.24 | 0.24 | 0.02 | 1.32 | 0.98 | NA | 0.25 | NA | NA | 0.75
| Analytical results for Sm in OREAS 753b (Certified Value 0.84 ppm) | [25 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | 0.9 | 0.89654 | 0.96 | 1 | 0.9 | 0.8 | 0.7 | 0.8565 | 0.7 | 0.8 | 0.77714 | < 1 | [12 empty] | 1
| 1 | 2 | 0.8 | 0.91674 | 0.95 | 1 | 1 | 0.8 | 0.7 | 0.78202 | 0.9 | 0.8 | 0.67533 | < 1 | [12 empty] | 10
| 1 | 3 | 0.9 | 0.95225 | 0.94 | 1 | 0.5 | 0.9 | 0.7 | 0.85974 | 0.8 | 0.7 | 0.54316 | < 1 | [12 empty] | 16
| 1 | 4 | 0.9 | 0.85776 | 0.98 | 0.8 | 0.7 | 0.7 | 0.9 | 0.76927 | 0.9 | 0.8 | 0.9101 | < 1 | [12 empty] | 0.84026
| 1 | 5 | 0.9 | 0.88693 | 0.94 | 0.9 | 1.4 | 0.8 | 1 | 0.83099 | 0.8 | 0.8 | 0.79374 | < 1 | [12 empty] | 101
| 1 | 6 | 0.7 | 0.92919 | 0.95 | 0.9 | 1 | 0.9 | 1 | 0.76987 | 0.8 | 0.8 | 0.60997 | < 1
| Mean |  | 1.01: 0.85; 1.02: 0.90657; 1.03: 0.95333; 1.04: 0.93333; 1.05: 0.91667; 1.07: 0.81667; 1.08: 0.83333; 1.09: 0.8114; 1.10: 0.81667; 1.12: 0.78333; 1.13: 0.71824
| Median |  | 1.01: 0.9; 1.02: 0.90664; 1.03: 0.95; 1.04: 0.95; 1.05: 0.95; 1.07: 0.8; 1.08: 0.8; 1.09: 0.80651; 1.10: 0.8; 1.12: 0.8; 1.13: 0.72624
| Std Dev. |  | 1.01: 0.083666; 1.02: 0.033377; 1.03: 0.015055; 1.04: 0.08165; 1.05: 0.30605; 1.07: 0.075277; 1.08: 0.15055; 1.09: 0.042703; 1.10: 0.075277; 1.12: 0.040825; 1.13: 0.13431
| Rel.Std.Dev. |  | 1.01: 9.843%; 1.02: 3.682%; 1.03: 1.579%; 1.04: 8.748%; 1.05: 33.39%; 1.07: 9.218%; 1.08: 18.07%; 1.09: 5.263%; 1.10: 9.218%; 1.12: 5.212%; 1.13: 18.7%
| PDM3 |  | 1.01: 1.159%; 1.02: 7.891%; 1.03: 13.46%; 1.04: 11.08%; 1.05: 9.093%; 1.07: -2.808%; 1.08: -0.8245%; 1.09: -3.435%; 1.10: -2.808%; 1.12: -6.775%; 1.13: -14.52%
| Z-Score (Absolute) |  | 0.27 | 0.89 | 1.4 | 1.19 | 1 | 0.09 | 0.09 | 0.15 | 0.09 | 0.46 | 1.17 | 3.56
| Analytical results for Sn in OREAS 753b (Certified Value 162 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | [7 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | [7 empty] | 0
0
| 1 | 1 | 163 | 159.8776 | 166 | 163 | 200 | 145 | 280 | 148 | 170 | 169 | 100 | 118 | 171 | 140 | 158 | 173.6169 | 162.1 | [7 empty] | 1
| 1 | 2 | 169 | 161.6086 | 160 | 172 | 200 | 143 | 200 | 147 | 169 | 171 | 110 | 120 | 157 | 200 | 158 | 153.1073 | 161.4 | [7 empty] | 0
| 1 | 3 | 170 | 160.3908 | 165 | 169 | 200 | 153 | 210 | 150 | 168 | 115 | 100 | 119 | 164 | 200 | 156 | 155.934 | 160.8 | [7 empty] | 16
| 1 | 4 | 167 | 159.5597 | 160 | 167 | 200 | 142 | 180 | 153 | 163 | 173 | 120 | 122 | 168 | < 50 | 158 | 160.2306 | 161.9 | [7 empty] | 161.5085
| 1 | 5 | 171 | 159.3792 | 167 | 161 | 200 | 147 | 180 | 148 | 167 | 170 | 100 | 120 | 142 | 140 | 159 | 159.0015 | 161.5 | [7 empty] | 102
| 1 | 6 | 161 | 161.5152 | 161 | 167 | 200 | 151 | 230 | 148 | 167 | 166 | 110 | 121 | 187 | 220 | 158 | 153.5022 | 161.5
| Mean |  | 166.8333 | 160.3885 | 163.1667 | 166.5 | 200 | 146.8333 | 213.3333 | 149 | 167.3333 | 160.6667 | 106.6667 | 120 | 164.8333 | 180 | 157.8333 | 159.2321 | 161.5333
| Median |  | 168 | 160.1342 | 163 | 167 | 200 | 146 | 205 | 148 | 167.5 | 169.5 | 105 | 120 | 166 | 200 | 158 | 157.4677 | 161.5
| Std Dev. |  | 4.0208 | 0.97206 | 3.1885 | 3.9875 | 0 | 4.4008 | 37.7712 | 2.1909 | 2.4221 | 22.4915 | 8.165 | 1.4142 | 14.9855 | 37.4166 | 0.98319 | 7.6052 | 0.45019
| Rel.Std.Dev. |  | 2.41% | 0.6061% | 1.954% | 2.395% | 0% | 2.997% | 17.71% | 1.47% | 1.447% | 14% | 7.655% | 1.179% | 9.091% | 20.79% | 0.6229% | 4.776% | 0.2787%
| PDM3 |  | 3.297% | -0.6934% | 1.027% | 3.091% | 23.83% | -9.086% | 32.09% | -7.745% | 3.607% | -0.5212% | -33.96% | -25.7% | 2.059% | 11.45% | -2.276% | -1.409% | 0.01539%
| Z-Score (Absolute) |  | 0.67 | 0.03 | 0.27 | 0.64 | 4.3 | 1.51 | 5.76 | 1.28 | 0.73 | 0 | 5.9 | 4.45 | 0.46 | 0.71 | 0.31 | 0.16 | 0.09
| Analytical results for Sr in OREAS 753b (Certified Value 47.4 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [6 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | [6 empty] | 1
1
| 1 | 1 | 47 | 48.0867 | 53 | 52 | 39 | 40 | 36 | 60 | 49 | 52 | 34 | 50 | 47 | 51.2364 | 43.3 | 50 | 52.54 | 37.2 | [6 empty] | 1
| 1 | 2 | 48 | 52.2314 | 53 | 54 | 41 | 40 | 37 | 60 | 48 | 52 | 40 | 40 | 47 | 52.2585 | 43.2 | 50 | 49.24 | 39.6 | [6 empty] | 0
| 1 | 3 | 48 | 53.9153 | 59 | 52 | 42 | 40 | 39 | 60 | 48 | 52 | 41 | 50 | 47 | 53.6213 | 32.9 | 50 | 49.89 | 38.6 | [6 empty] | 16
| 1 | 4 | 47 | 45.4425 | 52 | 53 | 40 | 40 | 39 | 60 | 49 | 52 | 55 | 50 | 47 | 53.4936 | 43.3 | 45 | 50.09 | 40.3 | [6 empty] | 47.3626
| 1 | 5 | 47 | 53.7329 | 55 | 52 | 40 | 40 | 37 | 50 | 46 | 53 | 56 | 40 | 47 | 51.8965 | 42.8 | 45 | 49.04 | 40.3 | [6 empty] | 103
| 1 | 6 | 45 | 47.3769 | 56 | 52 | 41 | 40 | 39 | 60 | 47 | 52 | 51 | 50 | 47 | 51.9072 | 42.3 | 45 | 48.08 | 38.2
| Mean |  | 47 | 50.131 | 54.6667 | 52.5 | 40.5 | 40 | 37.8333 | 58.3333 | 47.8333 | 52.1667 | 46.1667 | 46.6667 | 47 | 52.4023 | 41.3 | 47.5 | 49.8133 | 39.0333
| Median |  | 47 | 50.1591 | 54 | 52 | 40.5 | 40 | 38 | 60 | 48 | 52 | 46 | 50 | 47 | 52.0828 | 43 | 47.5 | 49.565 | 39.1
| Std Dev. |  | 1.0954 | 3.6181 | 2.582 | 0.83666 | 1.0488 | 0 | 1.3292 | 4.0825 | 1.169 | 0.40825 | 9.0646 | 5.164 | 0 | 0.95476 | 4.1333 | 2.7386 | 1.5132 | 1.2469
| Rel.Std.Dev. |  | 2.331% | 7.217% | 4.723% | 1.594% | 2.59% | 0% | 3.513% | 6.999% | 2.444% | 0.7826% | 19.63% | 11.07% | 0% | 1.822% | 10.01% | 5.766% | 3.038% | 3.194%
| PDM3 |  | -0.7656% | 5.845% | 15.42% | 10.85% | -14.49% | -15.55% | -20.12% | 23.16% | 0.9939% | 10.14% | -2.525% | -1.469% | -0.7656% | 10.64% | -12.8% | 0.2901% | 5.174% | -17.59%
| Z-Score (Absolute) |  | 0.03 | 0.39 | 0.99 | 0.7 | 0.9 | 0.97 | 1.26 | 1.48 | 0.08 | 0.66 | 0.15 | 0.08 | 0.03 | 0.69 | 0.8 | 0.03 | 0.34 | 1.1
| Analytical results for Ta in OREAS 753b (Certified Value 295 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.09 | 1.10 | 1.12 | 1.14 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | [9 empty] | 0
0
| 1 | 1 | 297.1 | 292.1885 | 283 | 314 | 286.6 | 257.0798 | 309 | 320 | 328 | 281 | 146 | 283 | 305 | 278.7018 | 304.3 | [9 empty] | 1
| 1 | 2 | 294.4 | 288.9872 | 287 | 346 | 294.9 | 262.3567 | 313 | 315 | 336 |  | 146 | 269 | 302 | 258.6185 | 302.3 | [9 empty] | 12
| 1 | 3 | 294.8 | 297.131 | 279 | 295 | 294.6 | 256.7393 | 313 | 335 | 237 | 293 | 147 | 402 | 299 | 264.7232 | 302.8 | [9 empty] | 16
| 1 | 4 | 292.4 | 293.2078 | 278 | 342 | 289.2 | 258.7372 | 304 | 315 | 344 |  | 147 | 283 | 299 | 276.0595 | 301.5 | [9 empty] | 295.1226
| 1 | 5 | 291.7 | 292.4525 | 300 | 297 | 293.9 | 261.6552 | 307 | 311 | 287 | 290 | 147 | 246 | 299 | 267.8577 | 303.2 | [9 empty] | 104
| 1 | 6 | 288.6 | 286.9709 | 288 | 359 | 290.3 | 259.0954 | 311 | 314 | 336 | 300 | 146 | 290 | 297 | 247.523 | 304.8
| Mean |  | 293.1667 | 291.823 | 285.8333 | 325.5 | 291.5833 | 259.2773 | 309.5 | 318.3333 | 311.3333 | 291 | 146.5 | 295.5 | 300.1667 | 265.5806 | 303.15
| Median |  | 293.4 | 292.3205 | 285 | 328 | 292.1 | 258.9163 | 310 | 315 | 332 | 291.5 | 146.5 | 283 | 299 | 266.2905 | 303
| Std Dev. |  | 2.9412 | 3.5273 | 8.0353 | 27.1643 | 3.3949 | 2.3122 | 3.5637 | 8.6641 | 41.6637 | 7.874 | 0.54772 | 54.4748 | 2.8577 | 11.5104 | 1.2341
| Rel.Std.Dev. |  | 1.003% | 1.209% | 2.811% | 8.345% | 1.164% | 0.8918% | 1.151% | 2.722% | 13.38% | 2.706% | 0.3739% | 18.43% | 0.9521% | 4.334% | 0.4071%
| PDM3 |  | -0.6628% | -1.118% | -3.148% | 10.29% | -1.199% | -12.15% | 4.872% | 7.865% | 5.493% | -1.397% | -50.36% | 0.1279% | 1.709% | -10.01% | 2.72%
| Z-Score (Absolute) |  | 0 | 0.09 | 0.5 | 2.18 | 0.11 | 2.29 | 1.1 | 1.7 | 1.22 | 0.15 | 9.91 | 0.16 | 0.47 | 1.86 | 0.67
| Analytical results for Tb in OREAS 753b (Certified Value 0.12 ppm) | [25 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.09 | 1.10 | 1.12 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [12 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | 0.1 | 0.13536 | 0.14 | 0.15 | 0.1 | 0.13 | 0.13 | 0.11979 | 0.13 | 0.1 | < 0.2 | < 0.5 | [12 empty] | 1
| 1 | 2 | 0.1 | 0.1231 | 0.15 | 0.14 | 0.2 | 0.11 | 0.12 | 0.14428 | 0.12 | 0.12 | < 0.2 | < 0.5 | [12 empty] | 13
| 1 | 3 | 0.1 | 0.14685 | 0.14 | 0.13 | 0.2 | 0.12 | 0.13 | 0.1412 | 0.1 | 0.1 | < 0.2 | < 0.5 | [12 empty] | 16
| 1 | 4 |  | 0.13285 | 0.13 | 0.13 | 0.2 | 0.1 | 0.09 | 0.14054 | 0.12 | 0.1 | < 0.2 | < 0.5 | [12 empty] | 0.12428
| 1 | 5 | 0.1 | 0.14524 | 0.13 | 0.12 | 0.2 | 0.1 | 0.11 | 0.13 | 0.11 | 0.11 | < 0.2 | < 0.5 | [12 empty] | 105
| 1 | 6 |  | 0.13469 | 0.14 | 0.15 | 0.1 | 0.12 | 0.1 | 0.13154 | 0.12 | 0.1 | < 0.2 | < 0.5
| Mean |  | 1.01: 0.1; 1.02: 0.13635; 1.03: 0.13833; 1.04: 0.13667; 1.05: 0.16667; 1.06: 0.11333; 1.07: 0.11333; 1.08: 0.13456; 1.09: 0.11667; 1.10: 0.105
| Median |  | 1.01: 0.1; 1.02: 0.13502; 1.03: 0.14; 1.04: 0.135; 1.05: 0.2; 1.06: 0.115; 1.07: 0.115; 1.08: 0.13604; 1.09: 0.12; 1.10: 0.1
| Std Dev. |  | 1.01: 0; 1.02: 0.0087274; 1.03: 0.0075277; 1.04: 0.012111; 1.05: 0.05164; 1.06: 0.012111; 1.07: 0.01633; 1.08: 0.0091907; 1.09: 0.010328; 1.10: 0.0083666
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 6.401%; 1.03: 5.442%; 1.04: 8.861%; 1.05: 30.98%; 1.06: 10.69%; 1.07: 14.41%; 1.08: 6.83%; 1.09: 8.853%; 1.10: 7.968%
| PDM3 |  | 1.01: -19.54%; 1.02: 9.711%; 1.03: 11.31%; 1.04: 9.966%; 1.05: 34.11%; 1.06: -8.808%; 1.07: -8.808%; 1.08: 8.271%; 1.09: -6.126%; 1.10: -15.51%
| Z-Score (Absolute) |  | NA | 0.59 | 0.7 | 0.61 | NA | 0.67 | 0.67 | 0.49 | 0.49 | 1.13 | 1.41 | 6.83
Results from laboratories 1 and 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 753b (Indicative Value < 1 ppm) | [25 empty] | Indicative
Te | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.01: 1.03; 1.02: 1.07; 1.03: 1.09; 1.04: 1.10; 1.05: 1.12; 1.06: 1.22; 1.07: 1.25 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*OES; 1.06: PF*MS; 1.07: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: < 5; 1.02: < 1; 1.03: < 6; 1.04: < 5; 1.05: 0.3; 1.06: < 1; 1.07: 2 | [6 empty] | 1
| 1 | 2 | 1.01: < 5; 1.02: < 1; 1.03: < 6; 1.04: < 5; 1.05: 0.3; 1.06: < 1; 1.07: 2 | [6 empty] | 6
| 1 | 3 | 1.01: < 5; 1.02: < 1; 1.03: < 6; 1.04: < 5; 1.05: 0.1; 1.06: < 1; 1.07: < 1 | [6 empty] | 16
| 1 | 4 | 1.01: < 5; 1.02: < 1; 1.03: < 6; 1.04: < 5; 1.05: 0.3; 1.06: < 1; 1.07: < 1 | [6 empty] | < 1
| 1 | 5 | 1.01: < 5; 1.02: < 1; 1.03: < 6; 1.04: < 5; 1.05: 0.5; 1.06: < 1; 1.07: < 1 | [6 empty] | 12
| 1 | 6 | 1.01: < 5; 1.02: < 1; 1.03: < 6; 1.04: < 5; 1.05: < 0.1; 1.06: < 1; 1.07: < 1
| Mean |  | 1.05: 0.3; 1.07: 2
| Median |  | 1.05: 0.3; 1.07: 2
| Std Dev. |  | 1.05: 0.14142; 1.07: 0
| Rel.Std.Dev. |  | 1.05: 47.14%; 1.07: 0%
PDM3
| Z-Score (Absolute) |  | 1.01: 1.36; 1.02: 0.45; 1.03: 1.82; 1.04: 1.36; 1.05: 0.67; 1.06: 0.45; 1.07: 0
| Analytical results for Th in OREAS 753b (Certified Value 2.08 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.22 | 1.25 | 1.27 | 1.28 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*MS | [10 empty] | 2
2
| 1 | 1 | 2.1 | 2.3504 | 2 | 2.5 | 2.3 | 1.9 | 2.1 | 2.2 | 2 | 1.9 | 1.9 | 2.4 | 1.9085 | 2 | [10 empty] | 1
| 1 | 2 | 2.1 | 2.2243 | 2 | 2.4 | 2.5 | 1.9 | 2.1 | 2.2 | 2 | 1.9 | 2.2 | 2.4 | 1.7083 | 2 | [10 empty] | 15
| 1 | 3 | 2.3 | 2.1479 | 2 | 2.3 | 2 | 1.9 | 2.3 | 2.2 | 2 | 1.3 | 1.9 | 2.2 | 1.8468 | 1.8 | [10 empty] | 16
| 1 | 4 | 2.2 | 2.0956 | 2 | 2.2 | 2.1 | 1.8 | 2.1 | 2.2 | 1.7 | 2.3 | 1.9 | 2.1 | 1.7692 | 1.8 | [10 empty] | 2.0797
| 1 | 5 | 2.3 | 2.3979 | 2 | 2 | 2 | 1.9 | 2.3 | 2.2 | 1.7 | 2.4 | 1.7 | 2.3 | 2.114 | 1.9 | [10 empty] | 106
| 1 | 6 | 2.1 | 2.2012 | 2 | 2.2 | 2.3 | 1.8 | 2.2 | 2 | 1.9 | 2.3 | 1.9 | 2.2 | 1.7299 | 2
| Mean |  | 2.1833 | 2.2362 | 2 | 2.2667 | 2.2 | 1.8667 | 2.1833 | 2.1667 | 1.8833 | 2.0167 | 1.9167 | 2.2667 | 1.8461 | 1.9167
| Median |  | 2.15 | 2.2127 | 2 | 2.25 | 2.2 | 1.9 | 2.15 | 2.2 | 1.95 | 2.1 | 1.9 | 2.25 | 1.808 | 1.95
| Std Dev. |  | 0.098319 | 0.11672 | 0 | 0.17512 | 0.2 | 0.05164 | 0.098319 | 0.08165 | 0.1472 | 0.41191 | 0.16021 | 0.12111 | 0.15105 | 0.098319
| Rel.Std.Dev. |  | 4.503% | 5.219% | 0% | 7.726% | 9.091% | 2.766% | 4.503% | 3.768% | 7.816% | 20.43% | 8.359% | 5.343% | 8.182% | 5.13%
| PDM3 |  | 4.985% | 7.527% | -3.831% | 8.992% | 5.786% | -10.24% | 4.985% | 4.183% | -9.441% | -3.029% | -7.838% | 8.992% | -11.23% | -7.838%
| Z-Score (Absolute) |  | 0.11 | 0.47 | NA | 0.67 | 0.22 | 2.02 | 0.11 | 0 | 1.91 | 1.01 | 1.69 | 0.67 | 2.16 | 1.69
Results from laboratory 3 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 753b (Certified Value 0.052 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.05 | 0.054326 | 0.049 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.03 | 0.078 | 0.048 | 0.054 | 0.054 | 0.012 | 0.04 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05412 | 0.14 |  | 1
| 1 | 2 | 0.05 | 0.055627 | 0.049 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.01 | 0.072 | 0.054 | 0.054 | 0.048 | 0.012 | 0.05 | 0.06 | 0.05 | 0.05 | 0.06 | 0.05513 | 0.13 |  | 0
| 1 | 3 | 0.05 | 0.052916 | 0.055 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.02 | 0.084 | 0.048 | 0.054 | 0.048 | 0.012 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.056552 | 0.13 |  | 16
| 1 | 4 | 0.05 | 0.055869 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.03 | 0.054 | 0.048 | 0.054 | 0.054 | 0.03 | 0.04 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05548 | 0.14 |  | 0.051604
| 1 | 5 | 0.06 | 0.054271 | 0.052 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.07 | 0.072 | 0.06 | 0.054 | 0.054 | 0.018 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.051723 | 0.14 |  | 107
| 1 | 6 | 0.06 | 0.055535 | 0.049 | 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.04 | 0.054 | 0.06 | 0.054 | 0.054 | 0.03 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.055957 | 0.13
| Mean |  | 0.053333 | 0.054757 | 0.050667 | 0.051667 | 0.05 | 0.05 | 0.051667 | 0.051667 | 0.05 | 0.05 | 0.033333 | 0.069 | 0.053 | 0.054 | 0.052 | 0.019 | 0.046667 | 0.06 | 0.05 | 0.05 | 0.051667 | 0.054827 | 0.135
| Median |  | 0.05 | 0.05493 | 0.0495 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.03 | 0.072 | 0.051 | 0.054 | 0.054 | 0.015 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.055305 | 0.135
| Std Dev. |  | 0.005164 | 0.0011321 | 0.0024221 | 0.0040825 | 7.6012e-18 | 7.6012e-18 | 0.0040825 | 0.0040825 | 7.6012e-18 | 7.6012e-18 | 0.020656 | 0.012442 | 0.0058992 | 0 | 0.0030984 | 0.0088318 | 0.005164 | 0 | 7.6012e-18 | 7.6012e-18 | 0.0040825 | 0.0017266 | 0.0054772
| Rel.Std.Dev. |  | 9.682% | 2.067% | 4.781% | 7.902% | 1.52e-14% | 1.52e-14% | 7.902% | 7.902% | 1.52e-14% | 1.52e-14% | 61.97% | 18.03% | 11.13% | 0% | 5.958% | 46.48% | 11.07% | 0% | 1.52e-14% | 1.52e-14% | 7.902% | 3.149% | 4.057%
| PDM3 |  | 3.352% | 6.111% | -1.816% | 0.1218% | -3.108% | -3.108% | 0.1218% | 0.1218% | -3.108% | -3.108% | -35.41% | 33.71% | 2.706% | 4.643% | 0.7678% | -63.18% | -9.567% | 16.27% | -3.108% | -3.108% | 0.1218% | 6.246% | 161.6%
| Z-Score (Absolute) |  | 0.67 | 1.25 | 0.4 | 0 | 0.67 | 0.67 | 0 | 0 | 0.67 | 0.67 | 7.42 | 6.99 | 0.52 | 0.93 | 0.12 | 13.22 | 2.02 | 3.37 | 0.67 | 0.67 | 0 | 1.28 | 33.72
| Analytical results for Tl in OREAS 753b (Certified Value 8.86 ppm) | [25 empty] | Certified
Tl | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.03; 1.03: 1.04; 1.04: 1.05; 1.05: 1.06; 1.06: 1.07; 1.07: 1.09; 1.08: 1.10; 1.09: 1.12; 1.10: 1.22; 1.12: 1.25; 1.13: 1.28 |  | ppm
|  |  | 1.01: PF*MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES/MS; 1.08: PF*OES/MS; 1.09: PF*OES; 1.10: PF*MS; 1.12: PF*MS; 1.13: PF*MS |  | 2
3
| 1 | 1 | 1.01: 9.4; 1.02: 8.71; 1.03: 9.3; 1.04: 9.3; 1.05: 9; 1.06: 8.4; 1.07: 8; 1.08: 8.9; 1.09: 8.9; 1.10: 8.8; 1.12: 8.7; 1.13: 9 |  | 1
| 1 | 2 | 1.01: 9.3; 1.02: 8.89; 1.03: 9.6; 1.04: 9.4; 1.05: 9; 1.06: 8.9; 1.07: 8.7; 1.08: 8.9; 1.09: 9.1; 1.10: 9.2; 1.12: 8.1; 1.13: 8.9 |  | 16
| 1 | 3 | 1.01: 9.3; 1.02: 8.81; 1.03: 9.2; 1.04: 9.6; 1.05: 9; 1.06: 8.9; 1.07: 8.4; 1.08: 8.7; 1.09: 8.3; 1.10: 6.1; 1.12: 8.1; 1.13: 8.5 |  | 16
| 1 | 4 | 1.01: 9.4; 1.02: 9.15; 1.03: 9.5; 1.04: 9.1; 1.05: 9.2; 1.06: 8.6; 1.07: 8.6; 1.08: 8.2; 1.09: 9.2; 1.10: 8.4; 1.12: 8.5; 1.13: 8.8 |  | 8.8554
| 1 | 5 | 1.01: 9; 1.02: 9.05; 1.03: 9.2; 1.04: 9.4; 1.05: 9.2; 1.06: 9.1; 1.07: 8.4; 1.08: 8.2; 1.09: 7.7; 1.10: 8.7; 1.12: 8.1; 1.13: 8.7 |  | 108
| 1 | 6 | 1.01: 9.2; 1.02: 8.96; 1.03: 9.8; 1.04: 9.2; 1.05: 9.1; 1.06: 9.2; 1.07: 8.6; 1.08: 8.5; 1.09: 8.3; 1.10: 8.5; 1.12: 7.8; 1.13: 9.1
| Mean |  | 1.01: 9.2667; 1.02: 8.9283; 1.03: 9.4333; 1.04: 9.3333; 1.05: 9.0833; 1.06: 8.85; 1.07: 8.45; 1.08: 8.5667; 1.09: 8.5833; 1.10: 8.2833; 1.12: 8.2167; 1.13: 8.8333
| Median |  | 1.01: 9.3; 1.02: 8.925; 1.03: 9.4; 1.04: 9.35; 1.05: 9.05; 1.06: 8.9; 1.07: 8.5; 1.08: 8.6; 1.09: 8.6; 1.10: 8.6; 1.12: 8.1; 1.13: 8.85
| Std Dev. |  | 1.01: 0.15055; 1.02: 0.16005; 1.03: 0.24221; 1.04: 0.17512; 1.05: 0.098319; 1.06: 0.30166; 1.07: 0.251; 1.08: 0.32042; 1.09: 0.58109; 1.10: 1.1053; 1.12: 0.32506; 1.13: 0.21602
| Rel.Std.Dev. |  | 1.01: 1.625%; 1.02: 1.793%; 1.03: 2.568%; 1.04: 1.876%; 1.05: 1.082%; 1.06: 3.409%; 1.07: 2.97%; 1.08: 3.74%; 1.09: 6.77%; 1.10: 13.34%; 1.12: 3.956%; 1.13: 2.446%
| PDM3 |  | 1.01: 4.644%; 1.02: 0.8234%; 1.03: 6.526%; 1.04: 5.397%; 1.05: 2.574%; 1.06: -0.06117%; 1.07: -4.578%; 1.08: -3.261%; 1.09: -3.073%; 1.10: -6.46%; 1.12: -7.213%; 1.13: -0.2494%
| Z-Score (Absolute) |  | 1.01: 0.86; 1.02: 0.18; 1.03: 1.2; 1.04: 0.99; 1.05: 0.49; 1.06: 0.02; 1.07: 0.79; 1.08: 0.56; 1.09: 0.52; 1.10: 1.13; 1.12: 1.26; 1.13: 0.02
| Analytical results for Tm in OREAS 753b (Indicative Value 0.047 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.04; 1.04: 1.09; 1.05: 1.10; 1.06: 1.12; 1.07: 1.22; 1.08: 1.25; 1.09: 1.27; 1.10: 1.28 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*OES/MS; 1.06: PF*OES; 1.07: PF*MS; 1.08: PF*MS; 1.09: PF*MS; 1.10: PF*MS |  | 3
3
| 1 | 1 | 1.01: 0.069288; 1.02: 0.06; 1.03: 0.06; 1.04: < 0.1; 1.05: < 0.05; 1.06: 0.05; 1.07: 0.06; 1.08: 0.06; 1.09: < 0.2; 1.10: < 1 |  | 1
| 1 | 2 | 1.01: 0.059642; 1.02: 0.06; 1.03: < 0.05; 1.04: < 0.1; 1.05: < 0.05; 1.06: 0.05; 1.07: 0.05; 1.08: < 0.05; 1.09: < 0.2; 1.10: < 1 |  | 1
| 1 | 3 | 1.01: 0.055914; 1.02: 0.05; 1.03: < 0.05; 1.04: < 0.1; 1.05: 0.05; 1.06: 0.07; 1.07: < 0.05; 1.08: 0.05; 1.09: < 0.2; 1.10: < 1 |  | 16
| 1 | 4 | 1.01: 0.05734; 1.02: 0.05; 1.03: 0.06; 1.04: < 0.1; 1.05: < 0.05; 1.06: < 0.05; 1.07: 0.06; 1.08: 0.05; 1.09: < 0.2; 1.10: < 1 |  | 0.047408
| 1 | 5 | 1.01: 0.085592; 1.02: 0.06; 1.03: < 0.05; 1.04: < 0.1; 1.05: < 0.05; 1.06: 0.09; 1.07: 0.06; 1.08: < 0.05; 1.09: < 0.2; 1.10: < 1 |  | 13
| 1 | 6 | 1.01: 0.062929; 1.02: 0.05; 1.03: 0.06; 1.04: < 0.1; 1.05: 0.06; 1.06: 0.06; 1.07: < 0.05; 1.08: < 0.05; 1.09: < 0.2; 1.10: < 1
| Mean |  | 1.01: 0.065117; 1.02: 0.055; 1.03: 0.06; 1.05: 0.055; 1.06: 0.064; 1.07: 0.0575; 1.08: 0.053333
| Median |  | 1.01: 0.061285; 1.02: 0.055; 1.03: 0.06; 1.05: 0.055; 1.06: 0.06; 1.07: 0.06; 1.08: 0.05
| Std Dev. |  | 1.01: 0.011105; 1.02: 0.0054772; 1.03: 0; 1.05: 0.0070711; 1.06: 0.016733; 1.07: 0.005; 1.08: 0.0057735
| Rel.Std.Dev. |  | 1.01: 17.05%; 1.02: 9.959%; 1.03: 0%; 1.05: 12.86%; 1.06: 26.15%; 1.07: 8.696%; 1.08: 10.83%
| PDM3 |  | 1.01: 37.36%; 1.02: 16.01%; 1.03: 26.56%; 1.05: 16.01%; 1.06: 35%; 1.07: 21.29%; 1.08: 12.5%
| Z-Score (Absolute) |  | 1.01: 0.75; 1.02: 0.15; 1.03: 0.6; 1.04: 0.15; 1.05: 1.04; 1.06: 0.3; 1.07: 0.35; 1.08: 0.8; 1.09: 2.83; 1.10: 26.68
| Analytical results for U in OREAS 753b (Certified Value 2.07 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.20 | 1.22 | 1.25 | 1.27 | 1.28 | [9 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 2.1 | 3.2014 | 4 | 2.38 | 2.3 | 2.1 | 2 | 2.2 | 2.1 | 1.57 | 1.2275 | 2 | 3 | 3.9277 | 2 | [9 empty] | 1
| 1 | 2 | 1.9 | 3.0707 | 3 | 2.3 | 2.2 | 2 | 2 | 2.2 | 2.1 | 1.57 | 1.2876 | 2 | 2 | 3.7786 | 1.8 | [9 empty] | 18
| 1 | 3 | 2.1 | 3.0189 | 4 | 2.68 | 2 | 2 | 2.2 | 2.2 | 2.2 | 0.39 | 1.3057 | < 2 | 2 | 3.6188 | 1.9 | [9 empty] | 16
| 1 | 4 | 2.1 | 2.7995 | 3 | 2.41 | 2 | 2.1 | 2 | 2.1 | 2 | 1.64 | 1.2618 | 2 | 2 | 4.4313 | 1.8 | [9 empty] | 2.0742
| 1 | 5 | 2.1 | 3.0165 | 4 | 2.15 | 2.1 | 2 | 2.1 | 2 | 2 | 2.36 | 1.2834 | < 2 | 2 | 2.395 | 1.8 | [9 empty] | 109
| 1 | 6 | 2.1 | 2.9733 | 3 | 2.58 | 2.4 | 2 | 2.2 | 2.2 | 2 | 1.95 | 1.2457 | 2 | 2 | 3.1632 | 1.9
| Mean |  | 2.0667 | 3.0134 | 3.5 | 2.4167 | 2.1667 | 2.0333 | 2.0833 | 2.15 | 2.0667 | 1.58 | 1.2686 | 2 | 2.1667 | 3.5524 | 1.8667
| Median |  | 2.1 | 3.0177 | 3.5 | 2.395 | 2.15 | 2 | 2.05 | 2.2 | 2.05 | 1.605 | 1.2726 | 2 | 2 | 3.6987 | 1.85
| Std Dev. |  | 0.08165 | 0.13118 | 0.54772 | 0.19086 | 0.1633 | 0.05164 | 0.098319 | 0.083666 | 0.08165 | 0.65806 | 0.029039 | 0 | 0.40825 | 0.70143 | 0.08165
| Rel.Std.Dev. |  | 3.951% | 4.353% | 15.65% | 7.898% | 7.537% | 2.54% | 4.719% | 3.891% | 3.951% | 41.65% | 2.289% | 0% | 18.84% | 19.75% | 4.374%
| PDM3 |  | -0.3643% | 45.28% | 68.74% | 16.51% | 4.457% | -1.971% | 0.4393% | 3.653% | -0.3643% | -23.83% | -38.84% | -3.578% | 4.457% | 71.27% | -10.01%
| Z-Score (Absolute) |  | 0.04 | 4.22 | NA | 1.54 | 0.41 | 0.19 | 0.04 | 0.34 | 0.04 | 2.23 | 3.63 | NA | NA | 6.64 | 0.94
Results from laboratories 3, 22 and 25 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 753b (Certified Value 10.5 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 1
1
| 1 | 1 | 12 | 12 | < 20 | < 20 | < 50 | 11 | < 10 | 14 | 10 | 8 | 11 | < 10 | < 50 | 11.3 | [10 empty] | 1
| 1 | 2 | 14 | 8 | < 20 | < 20 | < 50 | 11 | < 10 | 11 | 10 | 8 | 13 | < 10 | < 50 | 10.3 | [10 empty] | 0
| 1 | 3 | 21 | 9 | < 20 | < 20 | < 50 | 10 | < 10 | 22 | 10 | 7 | < 10 | < 10 | < 50 | 10.4 | [10 empty] | 16
| 1 | 4 | 12 | 9 | < 20 | < 20 | < 50 | 10 | < 10 | 7 | 10 | 7 | < 10 | < 10 | < 50 | 10.9 | [10 empty] | 10.5417
| 1 | 5 | 12 | 9 | < 20 | < 20 | < 50 | 9 | < 10 | 14 | 10 | 7 | 17 | < 10 | < 50 | 11.1 | [10 empty] | 110
| 1 | 6 | 12 | 8 | < 20 | < 20 | < 50 | 9 | < 10 | < 5 | 10 | 8 | < 10 | < 10 | < 50 | 10
| Mean |  | 1.01: 13.8333; 1.02: 9.1667; 1.06: 10; 1.08: 13.6; 1.09: 10; 1.10: 7.5; 1.12: 13.6667; 1.15: 10.6667
| Median |  | 1.01: 12; 1.02: 9; 1.06: 10; 1.08: 14; 1.09: 10; 1.10: 7.5; 1.12: 13; 1.15: 10.65
| Std Dev. |  | 1.01: 3.6009; 1.02: 1.472; 1.06: 0.89443; 1.08: 5.5045; 1.09: 0; 1.10: 0.54772; 1.12: 3.0551; 1.15: 0.50859
| Rel.Std.Dev. |  | 1.01: 26.03%; 1.02: 16.06%; 1.06: 8.944%; 1.08: 40.47%; 1.09: 0%; 1.10: 7.303%; 1.12: 22.35%; 1.15: 4.768%
| PDM3 |  | 1.01: 31.23%; 1.02: -13.04%; 1.06: -5.138%; 1.08: 29.01%; 1.09: -5.138%; 1.10: -28.85%; 1.12: 29.64%; 1.15: 1.186%
| Z-Score (Absolute) |  | 2 | 0.44 | 0 | 0 | 7.83 | 0 | 2.61 | 0.91 | 0 | 1.31 | 0.35 | 2.61 | 7.83 | 0.35
| Analytical results for W in OREAS 753b (Certified Value 3.81 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.18 | 1.22 | 1.27 | 1.28 | [10 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*MS | PF*OES | [10 empty] | 2
2
| 1 | 1 | 5 | < 5 | < 5 | 3 | 5 | 3 | < 100 | 3.5 | < 5 | 3 | < 50 | 4 | 3.5035 | < 5 | [10 empty] | 1
| 1 | 2 | 4 | < 5 | < 5 | 4 | 4 | 4 | < 100 | 4 | < 5 | 4 | < 50 | 4 | 4.3942 | < 5 | [10 empty] | 19
| 1 | 3 | 5 | < 5 | < 5 | 3 | 4 | 3 | < 100 | 3.9 | < 5 | 2 | < 50 | 4 | 4.0812 | < 5 | [10 empty] | 16
| 1 | 4 | 4 | < 5 | < 5 | 4 | 4 | 3 | < 100 | 3.5 | < 5 | 4 | < 50 | 3 | 4.5084 | < 5 | [10 empty] | 3.8143
| 1 | 5 | 4 | < 5 | < 5 | 3 | 3 | 3 | < 100 | 3.9 | < 5 | 4 | < 50 | 4 | 3.5283 | < 5 | [10 empty] | 111
| 1 | 6 | 5 | < 5 | < 5 | 4 | 5 | 4 | < 100 | 4 | < 5 | 4 | < 50 | 4 | 3.2728 | < 5
| Mean |  | 1.01: 4.5; 1.04: 3.5; 1.05: 4.1667; 1.06: 3.3333; 1.08: 3.8; 1.10: 3.5; 1.13: 3.8333; 1.14: 3.8814
| Median |  | 1.01: 4.5; 1.04: 3.5; 1.05: 4; 1.06: 3; 1.08: 3.9; 1.10: 4; 1.13: 4; 1.14: 3.8047
| Std Dev. |  | 1.01: 0.54772; 1.04: 0.54772; 1.05: 0.75277; 1.06: 0.5164; 1.08: 0.23664; 1.10: 0.83666; 1.13: 0.40825; 1.14: 0.51652
| Rel.Std.Dev. |  | 1.01: 12.17%; 1.04: 15.65%; 1.05: 18.07%; 1.06: 15.49%; 1.08: 6.227%; 1.10: 23.9%; 1.13: 10.65%; 1.14: 13.31%
| PDM3 |  | 1.01: 17.98%; 1.04: -8.241%; 1.05: 9.237%; 1.06: -12.61%; 1.08: -0.376%; 1.10: -8.241%; 1.13: 0.4979%; 1.14: 1.758%
| Z-Score (Absolute) |  | 0.84 | 1.13 | 1.13 | 0.15 | 0.51 | 0.31 | 45.74 | 0.15 | 1.13 | 0.15 | 21.07 | 0.18 | 0.23 | 1.13
| Analytical results for Y in OREAS 753b (Certified Value 3.46 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | [8 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 3.7 | 5.1757 | 6 | 4.1 | 4 | 2.8 | 3.4 | 4.3 | 3.2 | 2.1 | < 5 | < 5 | 3.7 | 4.2 | 2.7546 | 2.4 | [8 empty] | 1
| 1 | 2 | 3.5 | 3.895 | 4 | 3.8 | 4 | 2.6 | 3.3 | 3.6 | 3.2 | 3.4 | < 5 | < 5 | 3.3 | 3.5 | 2.9898 | 2.2 | [8 empty] | 20
| 1 | 3 | 3.5 | 3.9807 | 5 | 3.3 | 3.3 | 2.8 | 3.5 | 4.4 | 3.4 | 0.4 | < 5 | < 5 | 2.7 | 3.6 | 3.4685 | 1.9 | [8 empty] | 16
| 1 | 4 | 3.7 | 3.9441 | 6 | 3.9 | 3.5 | 2.7 | 3.6 | 3.7 | 3.4 | 2.8 | < 5 | < 5 | 3.6 | 3.4 | 2.9481 | 1.6 | [8 empty] | 3.4603
| 1 | 5 | 3.7 | 3.9478 | 6 | 3.5 | 3.8 | 2.6 | 3.6 | 3.6 | 3.3 | 14.2 | < 5 | < 5 | 3.4 | 3.6 | 2.8982 | 2.1 | [8 empty] | 112
| 1 | 6 | 3.5 | 3.5621 | 6 | 4.5 | 4 | 2.9 | 3.5 | 3.3 | 3.6 | 6.6 | < 5 | < 5 | 3.5 | 4 | 3.3852 | 2.5
| Mean |  | 3.6 | 4.0842 | 5.5 | 3.85 | 3.7667 | 2.7333 | 3.4833 | 3.8167 | 3.35 | 4.9167 |  |  | 3.3667 | 3.7167 | 3.074 | 2.1167
| Median |  | 3.6 | 3.9459 | 6 | 3.85 | 3.9 | 2.75 | 3.5 | 3.65 | 3.35 | 3.1 |  |  | 3.45 | 3.6 | 2.9689 | 2.15
| Std Dev. |  | 0.10954 | 0.55657 | 0.83666 | 0.42778 | 0.30111 | 0.12111 | 0.1169 | 0.43551 | 0.15166 | 4.9825 |  |  | 0.3559 | 0.31252 | 0.28577 | 0.33116
| Rel.Std.Dev. |  | 3.043% | 13.63% | 15.21% | 11.11% | 7.994% | 4.431% | 3.356% | 11.41% | 4.527% | 101.3% |  |  | 10.57% | 8.409% | 9.296% | 15.65%
| PDM3 |  | 4.038% | 18.03% | 58.95% | 11.26% | 8.855% | -21.01% | 0.6664% | 10.3% | -3.187% | 42.09% |  |  | -2.705% | 7.41% | -11.16% | -38.83%
| Z-Score (Absolute) |  | 0.21 | 1.11 | NA | 0.67 | 0.52 | 1.38 | 0 | 0.61 | 0.25 | 2.64 | 1.81 | 1.81 | 0.21 | 0.43 | 0.75 | 2.51
Results from laboratory 3 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 753b (Certified Value 0.34 ppm) | [25 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.12: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.04; 1.05: 1.09; 1.06: 1.10; 1.07: 1.12; 1.08: 1.22; 1.09: 1.25; 1.10: 1.27; 1.12: 1.28 |  | ppm
|  |  | 1.01: PF*MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*OES/MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*MS; 1.10: PF*MS; 1.12: PF*MS |  | 2
2
| 1 | 1 | 1.01: 0.3; 1.02: 0.38845; 1.03: 0.37; 1.04: 0.4; 1.05: 0.3; 1.06: 0.3; 1.07: 0.3; 1.08: 0.4; 1.09: 0.3; 1.10: 0.26094; 1.12: < 0.5 |  | 1
| 1 | 2 | 1.01: 0.5; 1.02: 0.36167; 1.03: 0.34; 1.04: 0.4; 1.05: 0.4; 1.06: 0.3; 1.07: 0.4; 1.08: 0.4; 1.09: 0.3; 1.10: 0.28423; 1.12: < 0.5 |  | 21
| 1 | 3 | 1.01: 0.4; 1.02: 0.42623; 1.03: 0.34; 1.04: 0.3; 1.05: 0.4; 1.06: 0.3; 1.07: 0.4; 1.08: 0.3; 1.09: 0.2; 1.10: 0.32842; 1.12: < 0.5 |  | 16
| 1 | 4 | 1.01: 0.4; 1.02: 0.31727; 1.03: 0.34; 1.04: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.2; 1.08: 0.3; 1.09: 0.3; 1.10: 0.2824; 1.12: < 0.5 |  | 0.33744
| 1 | 5 | 1.01: 0.3; 1.02: 0.43653; 1.03: 0.35; 1.04: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.3; 1.08: 0.4; 1.09: 0.4; 1.10: 0.30185; 1.12: < 0.5 |  | 113
| 1 | 6 | 1.01: 0.4; 1.02: 0.35878; 1.03: 0.35; 1.04: 0.4; 1.05: 0.4; 1.06: 0.4; 1.07: 0.3; 1.08: 0.3; 1.09: 0.2; 1.10: 0.3096; 1.12: < 0.5
| Mean |  | 1.01: 0.38333; 1.02: 0.38149; 1.03: 0.34833; 1.04: 0.35; 1.05: 0.35; 1.06: 0.31667; 1.07: 0.31667; 1.08: 0.35; 1.09: 0.28333; 1.10: 0.29457
| Median |  | 1.01: 0.4; 1.02: 0.37506; 1.03: 0.345; 1.04: 0.35; 1.05: 0.35; 1.06: 0.3; 1.07: 0.3; 1.08: 0.35; 1.09: 0.3; 1.10: 0.29304
| Std Dev. |  | 1.01: 0.075277; 1.02: 0.044973; 1.03: 0.01169; 1.04: 0.054772; 1.05: 0.054772; 1.06: 0.040825; 1.07: 0.075277; 1.08: 0.054772; 1.09: 0.075277; 1.10: 0.023711
| Rel.Std.Dev. |  | 1.01: 19.64%; 1.02: 11.79%; 1.03: 3.356%; 1.04: 15.65%; 1.05: 15.65%; 1.06: 12.89%; 1.07: 23.77%; 1.08: 15.65%; 1.09: 26.57%; 1.10: 8.049%
| PDM3 |  | 1.01: 13.6%; 1.02: 13.05%; 1.03: 3.228%; 1.04: 3.722%; 1.05: 3.722%; 1.06: -6.156%; 1.07: -6.156%; 1.08: 3.722%; 1.09: -16.03%; 1.10: -12.7%
| Z-Score (Absolute) |  | 1.01: 0.75; 1.02: 0.71; 1.03: 0; 1.04: 0.04; 1.05: 0.04; 1.06: 0.67; 1.07: 0.67; 1.08: 0.04; 1.09: 1.38; 1.10: 1.14; 1.12: 2.09
| Analytical results for Zn in OREAS 753b (Certified Value 164 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 0
0
| 1 | 1 | 157 | 174 | 175 | 160 | 137 | 160 | 150 | 172 | 131 | 160 | 170 | 160 | 130 | < 20 | 160 | 154 | 158.0133 | 161 | 100 | 119 | 171.23 | 174.3 |  |  | 1
| 1 | 2 | 134 | 172 | 175 | 160 | 121 | 160 | 160 | 172 | 158 | 160 | 180 | 160 | 180 | 20 | 150 | 156 | 160.4817 | 165 | 200 | 104 | 170.73 | 169.3 |  |  | 0
| 1 | 3 | 148 | 172 | 172 | 170 | 111 | 170 | 180 | 174 | 136 | 160 | 150 | 170 | 190 | 40 | 160 | 156 | 155.2858 | 105 | 200 | 144 | 165.91 | 171.4 |  |  | 16
| 1 | 4 | 147 | 172 | 179 | 160 | 121 | 160 | 160 | 174 | 155 | 160 | 150 | 140 | 200 | 40 | 150 | 156 | 161.5588 | 160 | 200 | 130 | 166.49 | 185.6 |  |  | 164.2492
| 1 | 5 | 145 | 173 | 173 | 160 | 134 | 160 | 160 | 172 | 187 | 160 | 150 | 150 | 210 | < 20 | 150 | 156 | 158.3387 | 171 | 200 | 144 | 173.08 | 193.4 |  |  | 114
| 1 | 6 | 150 | 172 | 173 | 160 | 118 | 170 | 160 | 174 | 147 | 160 | 140 | 140 | 210 | 50 | 160 | 157 | 162.6023 | 157 | 200 | 114 | 171.19 | 173.2
| Mean |  | 146.8333 | 172.5 | 174.5 | 161.6667 | 123.6667 | 163.3333 | 161.6667 | 173 | 152.3333 | 160 | 156.6667 | 153.3333 | 186.6667 | 37.5 | 155 | 155.8333 | 159.3801 | 153.1667 | 183.3333 | 125.8333 | 169.7717 | 177.8667
| Median |  | 147.5 | 172 | 174 | 160 | 121 | 160 | 160 | 173 | 151 | 160 | 150 | 155 | 195 | 40 | 155 | 156 | 159.4102 | 160.5 | 200 | 124.5 | 170.96 | 173.75
| Std Dev. |  | 7.5211 | 0.83666 | 2.51 | 4.0825 | 9.913 | 5.164 | 9.8319 | 1.0954 | 19.9566 | 0 | 15.0555 | 12.1106 | 30.1109 | 12.5831 | 5.4772 | 0.98319 | 2.6855 | 24.0866 | 40.8248 | 16.3758 | 2.8873 | 9.4952
| Rel.Std.Dev. |  | 5.122% | 0.485% | 1.438% | 2.525% | 8.016% | 3.162% | 6.082% | 0.6332% | 13.1% | 0% | 9.61% | 7.898% | 16.13% | 33.55% | 3.534% | 0.6309% | 1.685% | 15.73% | 22.27% | 13.01% | 1.701% | 5.338%
| PDM3 |  | -10.6% | 5.023% | 6.241% | -1.572% | -24.71% | -0.5576% | -1.572% | 5.328% | -7.255% | -2.587% | -4.616% | -6.646% | 13.65% | -77.17% | -5.631% | -5.124% | -2.964% | -6.747% | 11.62% | -23.39% | 3.362% | 8.291%
| Z-Score (Absolute) |  | 0.99 | 0.99 | 1.15 | 0.15 | 2.79 | 0.28 | 0.15 | 1.03 | 0.57 | 0.02 | 0.23 | 0.49 | 2.09 | 10.16 | 0.36 | 0.3 | 0.02 | 0.5 | 1.83 | 2.62 | 0.78 | 1.41
| Analytical results for Zr in OREAS 753b (Indicative Value 38.8 ppm) | [25 empty] | Indicative
Zr | Round | Replicate | 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  | 1
| No | No | 1.03: 1.01; 1.04: 1.02; 1.05: 1.04; 1.06: 1.28 |  |  | ppm
|  |  | 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES |  |  | 1
1
| 1 | 1 | 1.03: 43; 1.04: 49.5242; 1.05: 39.2; 1.06: 21.6 |  |  | 1
| 1 | 2 | 1.03: 40; 1.04: 49.4575; 1.05: 41.4; 1.06: 19.2 |  |  | 8
| 1 | 3 | 1.03: 45; 1.04: 47.1055; 1.05: 41.2; 1.06: 23.6 |  |  | 16
| 1 | 4 | 1.03: 45; 1.04: 48.2258; 1.05: 45.3; 1.06: 22.2 |  |  | 38.8176
| 1 | 5 | 1.03: 42; 1.04: 46.216; 1.05: 45.6; 1.06: 20.8 |  |  | 14
| 1 | 6 | 1.03: 44; 1.04: 48.1944; 1.05: 43.1; 1.06: 20.7
| Mean |  | 1.03: 43.1667; 1.04: 48.1206; 1.05: 42.6333; 1.06: 21.35
| Median |  | 1.03: 43.5; 1.04: 48.2101; 1.05: 42.25; 1.06: 21.2
| Std Dev. |  | 1.03: 1.9408; 1.04: 1.2986; 1.05: 2.5097; 1.06: 1.4963
| Rel.Std.Dev. |  | 1.03: 4.496%; 1.04: 2.699%; 1.05: 5.887%; 1.06: 7.009%
| PDM3 |  | 1.03: 11.2%; 1.04: 23.97%; 1.05: 9.83%; 1.06: -45%
| Z-Score (Absolute) |  | 1.03: 0.07; 1.04: 1.28; 1.05: 0.07; 1.06: 5.3
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Ag in OREAS 753b (Indicative Value 0.201 ppm) | [21 empty] | Indicative
Ag | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 3
3
| 1 | 1 | 0.215 | [19 empty] | 1
| 1 | 2 | 0.243 | [19 empty] | 4
| 1 | 3 | 0.197 | [19 empty] | 16
| 1 | 4 | 0.199 | [19 empty] | 0.2005
| 1 | 5 | 0.19 | [19 empty] | 16
| 1 | 6 | 0.159
| Mean |  | 0.2005
| Median |  | 0.198
| Std Dev. |  | 0.027797
| Rel.Std.Dev. |  | 13.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 753b (Certified Value 12.04 wt.%) | [21 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 12.09 | 12.0674 | 12.41 | 12.02 | 12.34 | 11.99 | 11.912 | 11.87 | 12.1229 | 12.2 | 12.1 | 11.9 | 11.75 | 11.94 | 10.6679 | 12.0078 | 12.12 | 11.8 | 12.0002 | 12.29 | 1
| 1 | 2 | 12.05 | 12.0863 | 12.43 | 12.09 | 12.34 | 11.98 | 11.905 | 11.84 | 12.1105 | 12.66 | 12.05 | 12 | 11.65 | 11.76 | 10.6766 | 11.9605 | 12.12 | 11.89 | 12.0305 | 12.1 | 0
| 1 | 3 | 12.12 | 11.9957 | 12.39 | 12.02 | 12.32 | 11.96 | 11.839 | 11.89 | 12.0995 | 11.44 | 12.1 | 11.85 | 11.7 | 11.88 | 10.6878 | 12.2062 | 12.11 | 11.98 | 12.0587 | 12.17 | 16
| 1 | 4 | 12.11 | 11.9509 | 12.36 | 11.97 | 12.35 | 12.02 | 11.829 | 11.88 | 12.0712 | 12.39 | 11.95 | 11.85 | 11.55 | 12.05 | 10.6414 | 12.0871 | 12.09 | 11.97 | 11.974 | 12.78 | 12.0365
| 1 | 5 | 12.12 | 12.0614 | 12.4 | 12.11 | 12.35 | 12.03 | 11.736 | 11.8 | 12.1205 | 12.27 | 12.05 | 11.65 | 11.85 | 11.92 | 10.6846 | 12.2761 | 12.15 | 11.92 | 11.966 | 12.02 | 116
| 1 | 6 | 12.03 | 12.0415 | 12.41 | 12.06 | 12.31 | 11.97 | 12.108 | 11.88 | 12.1027 | 12.78 | 12.1 | 11.8 | 11.5 | 11.81 | 10.6661 | 12.089 | 12.17 | 11.88 | 12.0305 | 12.16
| Mean |  | 12.0867 | 12.0339 | 12.4 | 12.045 | 12.335 | 11.9917 | 11.8882 | 11.86 | 12.1045 | 12.29 | 12.0583 | 11.8417 | 11.6667 | 11.8933 | 10.6707 | 12.1044 | 12.1267 | 11.9067 | 12.01 | 12.2533
| Median |  | 12.1 | 12.0514 | 12.405 | 12.04 | 12.34 | 11.985 | 11.872 | 11.875 | 12.1066 | 12.33 | 12.075 | 11.85 | 11.675 | 11.9 | 10.6722 | 12.0881 | 12.12 | 11.905 | 12.0154 | 12.165
| Std Dev. |  | 0.038297 | 0.051018 | 0.023664 | 0.051672 | 0.016432 | 0.027869 | 0.1251 | 0.034059 | 0.018798 | 0.47286 | 0.058452 | 0.11583 | 0.1291 | 0.10231 | 0.016806 | 0.11866 | 0.028752 | 0.066232 | 0.036163 | 0.27289
| Rel.Std.Dev. |  | 0.3169% | 0.424% | 0.1908% | 0.429% | 0.1332% | 0.2324% | 1.052% | 0.2872% | 0.1553% | 3.848% | 0.4847% | 0.9782% | 1.107% | 0.8602% | 0.1575% | 0.9803% | 0.2371% | 0.5563% | 0.3011% | 2.227%
| PDM3 |  | 0.4172% | -0.02156% | 3.02% | 0.07098% | 2.48% | -0.3721% | -1.232% | -1.466% | 0.5657% | 2.106% | 0.1818% | -1.618% | -3.072% | -1.189% | -11.35% | 0.5649% | 0.7495% | -1.078% | -0.2199% | 1.802%
| Z-Score (Absolute) |  | 0.23 | 0.03 | 1.74 | 0.03 | 1.43 | 0.23 | 0.73 | 0.87 | 0.31 | 1.21 | 0.09 | 0.96 | 1.8 | 0.71 | 6.62 | 0.31 | 0.42 | 0.64 | 0.14 | 1.03
| Analytical results for As in OREAS 753b (Indicative Value < 100 ppm) | [21 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 5 | 200 | < 100 | < 100 | [16 empty] | 1
| 1 | 2 | < 5 | < 100 | < 100 | < 100 | 8 | [15 empty] | 5
| 1 | 3 | < 5 | < 100 | 100 | < 100 | [16 empty] | 16
| 1 | 4 | < 5 | < 100 | < 100 | < 100 | [16 empty] | < 100
| 1 | 5 | < 5 | < 100 | < 100 | 100 | 8 | [15 empty] | 17
| 1 | 6 | < 5 | 100 | < 100 | < 100
| Mean |  | 5 | 150 | 100 | 100 | 8
| Median |  | 5 | 150 | 100 | 100 | 8
| Std Dev. |  |  | 70.7107 |  |  | 0
| Rel.Std.Dev. |  |  | 47.14% |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 1.49 | 0.67 | 0 | 0 | 1.37
| Analytical results for BaO in OREAS 753b (Certified Value 150 ppm) | [21 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 190 | 200 | 200 | 100 | 140.7229 | 150.7 | 200 | 100 | 100 | 100 | < 100 | 220 | < 100 | 130 | 100 | [5 empty] | 1
| 1 | 2 | 180 | 200 | 200 | 100 | 172.5179 | 147.4 | 200 | 200 | 100 | 100 | 134 | 220 | < 100 | 120 | 100 | [5 empty] | 0
| 1 | 3 | 160 | 200 | 200 | 200 | 160.5261 | 148.5 | 200 | 200 | < 100 | 100 | 123 | 220 | < 100 | 150 | 100 | [5 empty] | 16
| 1 | 4 | 160 | 200 | 200 | 100 | 149.6448 | 149.6 | 200 | 200 | 100 | 100 | 167 | 110 | < 100 | 110 | 100 | [5 empty] | 150.4346
| 1 | 5 | 180 | 100 | 200 | 200 | 146.6627 | 149.6 | 200 | 100 | < 100 | 100 | 167 | 220 | < 100 | 130 | 100 | [5 empty] | 117
| 1 | 6 | 145 | 200 | 200 | 100 | 167.4619 | 147.4 | 200 | 100 | < 100 | 100 | 145 | 110 | < 100 | 150 | 100
| Mean |  | 169.1667 | 183.3333 | 200 | 133.3333 | 156.256 | 148.8667 | 200 | 150 | 100 | 100 | 147.2 | 183.3333 |  | 131.6667 | 100
| Median |  | 170 | 200 | 200 | 100 | 155.0854 | 149.05 | 200 | 150 | 100 | 100 | 145 | 220 |  | 130 | 100
| Std Dev. |  | 16.8572 | 40.8248 | 0 | 51.6398 | 12.5322 | 1.3322 | 0 | 54.7723 | 0 | 0 | 19.6774 | 56.8038 |  | 16.0208 | 0
| Rel.Std.Dev. |  | 9.965% | 22.27% | 0% | 38.73% | 8.02% | 0.8949% | 0% | 36.51% | 0% | 0% | 13.37% | 30.98% |  | 12.17% | 0%
| PDM3 |  | 12.45% | 21.87% | 32.95% | -11.37% | 3.87% | -1.042% | 32.95% | -0.2889% | -33.53% | -33.53% | -2.15% | 21.87% |  | -12.48% | -33.53%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 1 | 0.3 | 0.14 | 0 | 1 | 0.02 | 1.45 | 0.96 | 0.35 | 0.73 | 1.93 | 0.34 | 0.96
| Analytical results for Be in OREAS 753b (Indicative Value 12.2 ppm) | [21 empty] | Indicative
Be | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 1
1
| 1 | 1 | 12 | [19 empty] | 1
| 1 | 2 | 12 | [19 empty] | 12
| 1 | 3 | 12 | [19 empty] | 16
| 1 | 4 | 13 | [19 empty] | 12.1667
| 1 | 5 | 13 | [19 empty] | 18
| 1 | 6 | 11
| Mean |  | 12.1667
| Median |  | 12
| Std Dev. |  | 0.75277
| Rel.Std.Dev. |  | 6.187%
| PDM3 |  | -2.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 753b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | [16 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | 0.5 | < 100 | < 100 | < 100 | [16 empty] | 1
| 1 | 2 | 0.5 | < 100 | < 100 | < 100 | [16 empty] | 13
| 1 | 3 | 0.5 | < 100 | < 100 | < 100 | [16 empty] | 16
| 1 | 4 | 0.7 | < 100 | < 100 | < 100 | [16 empty] | < 100
| 1 | 5 | 0.6 | < 100 | < 100 | < 100 | [16 empty] | 19
| 1 | 6 | 0.5 | < 100 | < 100 | < 100
| Mean |  | 0.55
| Median |  | 0.5
| Std Dev. |  | 0.083666
| Rel.Std.Dev. |  | 15.21%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for CaO in OREAS 753b (Certified Value 0.323 wt.%) | [21 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.32 | 0.293 | 0.33 | 0.31 | 0.33 | 0.33 | 0.329 | 0.34 | 0.32426 | 0.29 | 0.32 | 0.31 | 0.32 | 0.31 | 0.31429 | 0.31762 | 0.33 | 0.33 | 0.31693 | 0.57 | 1
| 1 | 2 | 0.32 | 0.293 | 0.32 | 0.32 | 0.33 | 0.33 | 0.33 | 0.33 | 0.32501 | 0.28 | 0.32 | 0.31 | 0.33 | 0.32 | 0.31851 | 0.31202 | 0.33 | 0.33 | 0.3159 | 0.56 | 0
| 1 | 3 | 0.32 | 0.292 | 0.33 | 0.32 | 0.33 | 0.34 | 0.326 | 0.33 | 0.32412 | 0.26 | 0.32 | 0.32 | 0.32 | 0.32 | 0.32728 | 0.31902 | 0.33 | 0.33 | 0.3159 | 0.55 | 16
| 1 | 4 | 0.32 | 0.295 | 0.33 | 0.32 | 0.32 | 0.33 | 0.324 | 0.34 | 0.32595 | 0.28 | 0.32 | 0.32 | 0.32 | 0.31 | 0.32221 | 0.31622 | 0.33 | 0.34 | 0.3159 | 0.61 | 0.32299
| 1 | 5 | 0.32 | 0.297 | 0.33 | 0.32 | 0.33 | 0.33 | 0.337 | 0.32 | 0.32553 | 0.28 | 0.32 | 0.31 | 0.33 | 0.31 | 0.31641 | 0.32322 | 0.33 | 0.33 | 0.31282 | 0.53 | 118
| 1 | 6 | 0.32 | 0.2925 | 0.32 | 0.32 | 0.32 | 0.33 | 0.331 | 0.31 | 0.32363 | 0.3 | 0.32 | 0.31 | 0.32 | 0.31 | 0.3171 | 0.32042 | 0.33 | 0.33 | 0.31796 | 0.59
| Mean |  | 0.32 | 0.29375 | 0.32667 | 0.31833 | 0.32667 | 0.33167 | 0.3295 | 0.32833 | 0.32475 | 0.28167 | 0.32 | 0.31333 | 0.32333 | 0.31333 | 0.3193 | 0.31809 | 0.33 | 0.33167 | 0.3159 | 0.56833
| Median |  | 0.32 | 0.293 | 0.33 | 0.32 | 0.33 | 0.33 | 0.3295 | 0.33 | 0.32463 | 0.28 | 0.32 | 0.31 | 0.32 | 0.31 | 0.31781 | 0.31832 | 0.33 | 0.33 | 0.3159 | 0.565
| Std Dev. |  | 0 | 0.0018908 | 0.005164 | 0.0040825 | 0.005164 | 0.0040825 | 0.0045056 | 0.01169 | 0.00089526 | 0.013292 | 0 | 0.005164 | 0.005164 | 0.005164 | 0.0047112 | 0.0038255 | 0 | 0.0040825 | 0.0017218 | 0.028577
| Rel.Std.Dev. |  | 0% | 0.6437% | 1.581% | 1.282% | 1.581% | 1.231% | 1.367% | 3.561% | 0.2757% | 4.719% | 0% | 1.648% | 1.597% | 1.648% | 1.475% | 1.203% | 0% | 1.231% | 0.5451% | 5.028%
| PDM3 |  | -0.9265% | -9.054% | 1.138% | -1.442% | 1.138% | 2.686% | 2.015% | 1.654% | 0.544% | -12.79% | -0.9265% | -2.99% | 0.1056% | -2.99% | -1.143% | -1.519% | 2.17% | 2.686% | -2.195% | 75.96%
| Z-Score (Absolute) |  | 0.18 | 3.03 | 0.54 | 0.36 | 0.54 | 1.08 | 0.85 | 0.72 | 0.33 | 4.34 | 0.18 | 0.9 | 0.18 | 0.9 | 0.26 | 0.39 | 0.9 | 1.08 | 0.63 | 26.76
| Analytical results for Ce in OREAS 753b (Indicative Value 8.58 ppm) | [21 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.22 | 1.26 | 1.28 | [13 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [13 empty] | 2
2
| 1 | 1 | 8.5 | < 80 | < 80 | < 80 | 90 |  | < 10 | [13 empty] | 1
| 1 | 2 | 9.2 | < 80 | < 80 | < 80 | 140 |  | < 10 | [13 empty] | 14
| 1 | 3 | 9.6 | < 80 | < 80 | < 80 | 100 |  | < 10 | [13 empty] | 16
| 1 | 4 | 8.9 | < 80 | < 80 | < 80 | 240 |  | < 10 | [13 empty] | 8.5787
| 1 | 5 | 8.4 | < 80 | < 80 | < 80 | 230 | 8 | < 10 | [13 empty] | 20
| 1 | 6 | 9.5 | < 80 | < 80 | < 80 | 160 |  | < 10
| Mean |  | 9.0167 |  |  |  | 160 | 8
| Median |  | 9.05 |  |  |  | 150 | 8
| Std Dev. |  | 0.50365 |  |  |  | 63.561
| Rel.Std.Dev. |  | 5.586% |  |  |  | 39.73%
| PDM3 |  | 5.106% |  |  |  | 1765% | -6.745%
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 0 | 2.54 | 0.69 | 0.76
| Analytical results for Co in OREAS 753b (Indicative Value < 100 ppm) | [21 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.12 | 1.13 | 1.14 | 1.16 | 1.22 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | < 37 | 1.3 | < 100 | < 100 | < 100 | < 70 | [14 empty] | 1
| 1 | 2 | < 37 | 1.4 | < 100 | < 100 | < 100 | < 70 | [14 empty] | 15
| 1 | 3 | < 37 | 1.3 | < 100 | < 100 | < 100 | < 70 | [14 empty] | 16
| 1 | 4 | < 37 | 1.3 | < 100 | < 100 | < 100 | < 70 | [14 empty] | < 100
| 1 | 5 | < 37 | 1.4 | < 100 | < 100 | < 100 | < 70 | [14 empty] | 21
| 1 | 6 | < 37 | 1.3 | < 100 | < 100 | < 100 | < 70
| Mean |  |  | 1.3333
| Median |  |  | 1.3
| Std Dev. |  |  | 0.05164
| Rel.Std.Dev. |  |  | 3.873%
PDM3
| Z-Score (Absolute) |  | 2.54 | 4.26 | 0.67 | 0.67 | 0.67 | 0.67
| Analytical results for Cr2O3 in OREAS 753b (Certified Value < 100 ppm) | [21 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 50 | < 100 | < 100 | < 100 | < 50 | < 100 | < 100 | 30.7 | 150 | < 150 | < 150 | < 100 | < 100 | < 100 | < 100 | 20 | < 100 |  |  |  | 1
| 1 | 2 | 60 | < 100 | < 100 | < 100 | < 50 | < 100 | < 100 | 35.1 | < 150 | < 150 | < 150 | < 100 | < 100 | < 100 | < 100 | 20 | < 100 |  |  |  | 0
| 1 | 3 | < 50 | < 100 | < 100 | < 100 | < 50 | < 100 | < 100 | 17.5 | < 150 | < 150 | < 150 | < 100 | < 100 | < 100 | < 100 | 20 | < 100 |  |  |  | 16
| 1 | 4 | < 50 | < 100 | < 100 | < 100 | < 50 | < 100 | < 100 | 16.1 | < 150 | < 150 | < 150 | < 100 | < 100 | < 100 | < 100 | 20 | < 100 |  |  |  | < 100
| 1 | 5 | 60 | < 100 | < 100 | < 100 | < 50 | < 100 | < 100 | 27.8 | < 150 | < 150 | < 150 | < 100 | < 100 | < 100 | < 100 | 40 | < 100 |  |  |  | 119
| 1 | 6 | 50 | < 100 | < 100 | < 100 | < 50 | < 100 | < 100 | 36.5 | < 150 | < 150 | < 150 | < 100 | < 100 | < 100 | < 100 | 30 | < 100
| Mean |  | 55 | [6 empty] | 27.2833 | 150 | [6 empty] | 25
| Median |  | 55 | [6 empty] | 29.25 | 150 | [6 empty] | 20
| Std Dev. |  | 5.7735 | [6 empty] | 8.7037 | [7 empty] | 8.3666
| Rel.Std.Dev. |  | 10.5% | [6 empty] | 31.9% | [7 empty] | 33.47%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cs in OREAS 753b (Indicative Value 183 ppm) | [21 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.22 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 0
0
| 1 | 1 | 176 | < 90 | [18 empty] | 1
| 1 | 2 | 179 | 190 | [18 empty] | 16
| 1 | 3 | 178 | 190 | [18 empty] | 16
| 1 | 4 | 178 | 190 | [18 empty] | 183.2385
| 1 | 5 | 178 | 190 | [18 empty] | 22
| 1 | 6 | 178 | 190
| Mean |  | 177.8333 | 190
| Median |  | 178 | 190
| Std Dev. |  | 0.98319 | 0
| Rel.Std.Dev. |  | 0.5529% | 0%
| PDM3 |  | -2.95% | 3.69%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for CuO in OREAS 753b (Indicative Value 39.4 ppm) | [21 empty] | Indicative
CuO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 50 | 26.3 | < 63 | 63 | 138 | 50 | [14 empty] | 1
| 1 | 2 | < 50 | 33.8 | < 63 | < 63 | < 63 | 50 | [14 empty] | 17
| 1 | 3 | < 50 | 31.3 | < 63 | 75 | 63 | 60 | [14 empty] | 16
| 1 | 4 | < 50 | 25 | < 63 | < 63 | < 63 | 60 | [14 empty] | 39.4125
| 1 | 5 | < 50 | 28.8 | < 63 | < 63 | < 63 | 50 | [14 empty] | 23
| 1 | 6 | < 50 | 30 | < 63 | < 63 | < 63 | 50
| Mean |  |  | 29.2 |  | 69 | 100.5 | 53.3333
| Median |  |  | 29.4 |  | 69 | 100.5 | 50
| Std Dev. |  |  | 3.2379 |  | 8.4853 | 53.033 | 5.164
| Rel.Std.Dev. |  |  | 11.09% |  | 12.3% | 52.77% | 9.682%
| PDM3 |  |  | -25.91% |  | 75.07% | 155% | 35.32%
| Z-Score (Absolute) |  | 0.81 | 0.54 | 0.4 | 0.4 | 1.08 | 1.02
| Analytical results for Dy in OREAS 753b (Indicative Value 0.64 ppm) | [21 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.6 | < 10 | [18 empty] | 1
| 1 | 2 | 0.63 | < 10 | [18 empty] | 18
| 1 | 3 | 0.63 | < 10 | [18 empty] | 16
| 1 | 4 | 0.65 | < 10 | [18 empty] | 0.64333
| 1 | 5 | 0.65 | < 10 | [18 empty] | 24
| 1 | 6 | 0.7 | < 10
| Mean |  | 0.64333
| Median |  | 0.64
| Std Dev. |  | 0.033267
| Rel.Std.Dev. |  | 5.171%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Er in OREAS 753b (Indicative Value 0.36 ppm) | [21 empty] | Indicative
Er | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.36 | < 10 | [18 empty] | 1
| 1 | 2 | 0.37 | < 10 | [18 empty] | 19
| 1 | 3 | 0.32 | < 10 | [18 empty] | 16
| 1 | 4 | 0.36 | < 10 | [18 empty] | 0.35833
| 1 | 5 | 0.33 | < 10 | [18 empty] | 25
| 1 | 6 | 0.41 | < 10
| Mean |  | 0.35833
| Median |  | 0.36
| Std Dev. |  | 0.031885
| Rel.Std.Dev. |  | 8.898%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Eu in OREAS 753b (Indicative Value 0.17 ppm) | [21 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.16 | < 10 | [18 empty] | 1
| 1 | 2 | 0.17 | < 10 | [18 empty] | 20
| 1 | 3 | 0.14 | < 10 | [18 empty] | 16
| 1 | 4 | 0.18 | < 10 | [18 empty] | 0.16833
| 1 | 5 | 0.18 | < 10 | [18 empty] | 26
| 1 | 6 | 0.18 | < 10
| Mean |  | 0.16833
| Median |  | 0.175
| Std Dev. |  | 0.016021
| Rel.Std.Dev. |  | 9.517%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Fe2O3 in OREAS 753b (Certified Value 0.899 wt.%) | [21 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.901 | 0.886 | 0.91 | 0.91 | 0.92 | 0.97 | 0.921 | 1.06 | 0.91708 | 0.88 | 0.915 | 0.901 | 0.872 | 0.93 | 1.3023 | 0.9 | 0.89 | 0.9 | 0.89782 | 0.89 | 1
| 1 | 2 | 0.901 | 0.888 | 0.9 | 0.91 | 0.97 | 0.98 | 0.919 | 1.07 | 0.92019 | 0.92 | 0.901 | 0.872 | 0.872 | 0.93 | 1.3135 | 0.9 | 0.9 | 0.91 | 0.89146 | 0.88 | 0
| 1 | 3 | 0.901 | 0.875 | 0.89 | 0.91 | 0.95 | 0.85 | 0.915 | 1.07 | 0.91143 | 0.85 | 0.886 | 0.886 | 0.872 | 0.93 | 1.333 | 0.91 | 0.89 | 0.9 | 0.88555 | 0.89 | 16
| 1 | 4 | 0.901 | 0.877 | 0.9 | 0.91 | 0.9 | 0.87 | 0.916 | 1.07 | 0.92136 | 0.86 | 0.886 | 0.886 | 0.872 | 0.93 | 1.3326 | 0.91 | 0.89 | 0.91 | 0.88828 | 0.88 | 0.89891
| 1 | 5 | 0.901 | 0.882 | 0.9 | 0.92 | 0.93 | 0.86 | 0.937 | 1.05 | 0.91704 | 0.89 | 0.886 | 0.872 | 0.886 | 0.92 | 1.3033 | 0.92 | 0.88 | 0.9 | 0.84376 | 0.91 | 120
| 1 | 6 | 0.901 | 0.88 | 0.89 | 0.93 | 0.9 | 0.87 | 0.928 | 1.05 | 0.91837 | 0.88 | 0.886 | 0.872 | 0.858 | 0.93 | 1.3316 | 0.91 | 0.89 | 0.9 | 0.89252 | 0.88
| Mean |  | 0.901 | 0.88133 | 0.89833 | 0.915 | 0.92833 | 0.9 | 0.92267 | 1.0617 | 0.91758 | 0.88 | 0.89333 | 0.8815 | 0.872 | 0.92833 | 1.3194 | 0.90833 | 0.89 | 0.90333 | 0.88323 | 0.88833
| Median |  | 0.901 | 0.881 | 0.9 | 0.91 | 0.925 | 0.87 | 0.92 | 1.065 | 0.91772 | 0.88 | 0.886 | 0.879 | 0.872 | 0.93 | 1.3226 | 0.91 | 0.89 | 0.9 | 0.88987 | 0.885
| Std Dev. |  | 1.2162e-16 | 0.0050465 | 0.0075277 | 0.0083666 | 0.027869 | 0.058652 | 0.0084063 | 0.0098319 | 0.0034658 | 0.024495 | 0.012193 | 0.01176 | 0.0088544 | 0.0040825 | 0.014787 | 0.0075277 | 0.0063246 | 0.005164 | 0.019777 | 0.01169
| Rel.Std.Dev. |  | 1.35e-14% | 0.5726% | 0.838% | 0.9144% | 3.002% | 6.517% | 0.9111% | 0.9261% | 0.3777% | 2.784% | 1.365% | 1.334% | 1.015% | 0.4398% | 1.121% | 0.8287% | 0.7106% | 0.5717% | 2.239% | 1.316%
| PDM3 |  | 0.232% | -1.956% | -0.06461% | 1.789% | 3.273% | 0.1208% | 2.642% | 18.11% | 2.076% | -2.104% | -0.6208% | -1.937% | -2.994% | 3.273% | 46.78% | 1.048% | -0.9917% | 0.4916% | -1.745% | -1.177%
| Z-Score (Absolute) |  | 0.01 | 0.75 | 0.08 | 0.58 | 1.1 | 0.01 | 0.88 | 6.33 | 0.68 | 0.8 | 0.27 | 0.73 | 1.11 | 1.1 | 16.45 | 0.31 | 0.41 | 0.12 | 0.67 | 0.47
| Analytical results for Ga in OREAS 753b (Indicative Value 45.6 ppm) | [21 empty] | Indicative
Ga | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 1
1
| 1 | 1 | 45.5 | [19 empty] | 1
| 1 | 2 | 46.4 | [19 empty] | 21
| 1 | 3 | 45.1 | [19 empty] | 16
| 1 | 4 | 45.1 | [19 empty] | 45.5833
| 1 | 5 | 45.5 | [19 empty] | 27
| 1 | 6 | 45.9
| Mean |  | 45.5833
| Median |  | 45.5
| Std Dev. |  | 0.49967
| Rel.Std.Dev. |  | 1.096%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 753b (Indicative Value 0.81 ppm) | [21 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.83 | < 10 | [18 empty] | 1
| 1 | 2 | 0.84 | < 10 | [18 empty] | 22
| 1 | 3 | 0.89 | < 10 | [18 empty] | 16
| 1 | 4 | 0.74 | < 10 | [18 empty] | 0.81333
| 1 | 5 | 0.82 | < 10 | [18 empty] | 28
| 1 | 6 | 0.76 | < 10
| Mean |  | 0.81333
| Median |  | 0.825
| Std Dev. |  | 0.055015
| Rel.Std.Dev. |  | 6.764%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for H2O- in OREAS 753b (Indicative Value 0.18 wt.%) | [21 empty] | Indicative
H2O- | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.11 | [19 empty] | wt.%
|  |  | BF*XRF | [19 empty] | 3
3
| 1 | 1 | 0.1798 | [19 empty] | 1
| 1 | 2 | 0.0722 | [19 empty] | 23
| 1 | 3 | 0.1852 | [19 empty] | 16
| 1 | 4 | 0.1841 | [19 empty] | 0.18006
| 1 | 5 | 0.1641 | [19 empty] | 29
| 1 | 6 | 0.1871
| Mean |  | 0.16208
| Median |  | 0.18195
| Std Dev. |  | 0.044815
| Rel.Std.Dev. |  | 27.65%
| PDM3 |  | -9.984%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 753b (Indicative Value < 80 ppm) | [21 empty] | Indicative
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | 1.28 | [14 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 3.1 | < 80 | < 80 | < 80 | 8 | < 10 | [14 empty] | 1
| 1 | 2 | 2.8 | < 80 | < 80 | < 80 |  | < 10 | [14 empty] | 8
| 1 | 3 | 2.8 | < 80 | < 80 | < 80 |  | < 10 | [14 empty] | 16
| 1 | 4 | 2.8 | < 80 | < 80 | < 80 | 8 | < 10 | [14 empty] | < 80
| 1 | 5 | 2.9 | < 80 | < 80 | < 80 | 8 | < 10 | [14 empty] | 16
| 1 | 6 | 3.1 | < 80 | < 80 | < 80 |  | < 10
| Mean |  | 2.9167 |  |  |  | 8
| Median |  | 2.85 |  |  |  | 8
| Std Dev. |  | 0.1472 |  |  |  | 0
| Rel.Std.Dev. |  | 5.047% |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.9 | 0.67 | 0.67 | 0.67 | 0.67 | 0.81
| Analytical results for Ho in OREAS 753b (Indicative Value 0.12 ppm) | [21 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.1 | < 10 | [18 empty] | 1
| 1 | 2 | 0.13 | < 10 | [18 empty] | 11
| 1 | 3 | 0.12 | < 10 | [18 empty] | 16
| 1 | 4 | 0.13 | < 10 | [18 empty] | 0.11833
| 1 | 5 | 0.11 | < 10 | [18 empty] | 17
| 1 | 6 | 0.12 | < 10
| Mean |  | 0.11833
| Median |  | 0.12
| Std Dev. |  | 0.01169
| Rel.Std.Dev. |  | 9.879%
| PDM3 |  | 2.887e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for In in OREAS 753b (Indicative Value < 0.1 ppm) | [21 empty] | Indicative
In | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | < 0.1 | [19 empty] | 1
| 1 | 2 | < 0.1 | [19 empty] | 1
| 1 | 3 | < 0.1 | [19 empty] | 16
| 1 | 4 | < 0.1 | [19 empty] | < 0.1
| 1 | 5 | < 0.1 | [19 empty] | 18
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 753b (Certified Value 1.13 wt.%) | [21 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.15 | 1.14 | 1.2 | 1.14 | 1.13 | 1.13 | 1.14 | 1.16 | 1.115 | 1.18 | 1.11 | 1.12 | 1.11 | 1.17 | 1.0935 | 1.14 | 1.12 | 1.15 | 1.1278 | 1.55 | 1
| 1 | 2 | 1.14 | 1.145 | 1.22 | 1.14 | 1.13 | 1.12 | 1.14 | 1.17 | 1.1125 | 1.18 | 1.12 | 1.12 | 1.1 | 1.17 | 1.1039 | 1.12 | 1.12 | 1.16 | 1.1297 | 1.56 | 0
| 1 | 3 | 1.14 | 1.124 | 1.22 | 1.13 | 1.13 | 1.12 | 1.149 | 1.17 | 1.1192 | 1.1 | 1.12 | 1.12 | 1.1 | 1.18 | 1.1285 | 1.15 | 1.12 | 1.17 | 1.1406 | 1.1 | 16
| 1 | 4 | 1.14 | 1.128 | 1.2 | 1.13 | 1.12 | 1.13 | 1.142 | 1.17 | 1.1114 | 1.24 | 1.12 | 1.12 | 1.08 | 1.17 | 1.1269 | 1.13 | 1.11 | 1.16 | 1.1317 | 1.44 | 1.1346
| 1 | 5 | 1.15 | 1.139 | 1.23 | 1.15 | 1.13 | 1.14 | 1.17 | 1.16 | 1.1151 | 1.16 | 1.13 | 1.1 | 1.11 | 1.15 | 1.1093 | 1.14 | 1.14 | 1.14 | 1.1218 | 1.63 | 121
| 1 | 6 | 1.14 | 1.135 | 1.21 | 1.14 | 1.12 | 1.13 | 1.157 | 1.18 | 1.1119 | 1.38 | 1.13 | 1.11 | 1.08 | 1.18 | 1.1185 | 1.13 | 1.13 | 1.16 | 1.1357 | 1.56
| Mean |  | 1.1433 | 1.1352 | 1.2133 | 1.1383 | 1.1267 | 1.1283 | 1.1497 | 1.1683 | 1.1142 | 1.2067 | 1.1217 | 1.115 | 1.0967 | 1.17 | 1.1134 | 1.135 | 1.1233 | 1.1567 | 1.1312 | 1.4733
| Median |  | 1.14 | 1.137 | 1.215 | 1.14 | 1.13 | 1.13 | 1.1455 | 1.17 | 1.1138 | 1.18 | 1.12 | 1.12 | 1.1 | 1.17 | 1.1139 | 1.135 | 1.12 | 1.16 | 1.1307 | 1.555
| Std Dev. |  | 0.005164 | 0.0078846 | 0.012111 | 0.0075277 | 0.005164 | 0.0075277 | 0.011944 | 0.0075277 | 0.0029141 | 0.096056 | 0.0075277 | 0.0083666 | 0.013663 | 0.010954 | 0.013693 | 0.010488 | 0.010328 | 0.010328 | 0.006506 | 0.19284
| Rel.Std.Dev. |  | 0.4517% | 0.6946% | 0.9981% | 0.6613% | 0.4583% | 0.6672% | 1.039% | 0.6443% | 0.2615% | 7.96% | 0.6711% | 0.7504% | 1.246% | 0.9363% | 1.23% | 0.9241% | 0.9194% | 0.8929% | 0.5751% | 13.09%
| PDM3 |  | 0.7736% | 0.05374% | 6.943% | 0.3328% | -0.6955% | -0.5486% | 1.332% | 2.977% | -1.794% | 6.356% | -1.136% | -1.724% | -3.34% | 3.124% | -1.862% | 0.03905% | -0.9893% | 1.949% | -0.2935% | 29.86%
| Z-Score (Absolute) |  | 0.32 | 0 | 3.04 | 0.13 | 0.33 | 0.26 | 0.57 | 1.29 | 0.81 | 2.78 | 0.52 | 0.78 | 1.49 | 1.36 | 0.84 | 0 | 0.46 | 0.84 | 0.15 | 13.16
| Analytical results for La in OREAS 753b (Indicative Value 34.5 ppm) | [21 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | 1.28 | [14 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | 4.1 | < 90 | < 90 | < 90 | 25.5798 | < 10 | [14 empty] | 1
| 1 | 2 | 4.5 | < 90 | 90 | < 90 | 17.0532 | < 10 | [14 empty] | 13
| 1 | 3 | 4.7 | < 90 | < 90 | < 90 |  | < 10 | [14 empty] | 16
| 1 | 4 | 4.3 | < 90 | < 90 | < 90 |  | < 10 | [14 empty] | 34.4875
| 1 | 5 | 4.1 | 90 | < 90 | < 90 | 51.1596 | < 10 | [14 empty] | 19
| 1 | 6 | 4.5 | < 90 | < 90 | < 90 | 42.633 | < 10
| Mean |  | 4.3667 | 90 | 90 |  | 34.1064
| Median |  | 4.4 | 90 | 90 |  | 34.1064
| Std Dev. |  | 0.24221 |  |  |  | 15.5674
| Rel.Std.Dev. |  | 5.547% |  |  |  | 45.64%
| PDM3 |  | -87.34% | 161% | 161% |  | -1.105%
| Z-Score (Absolute) |  | 2.02 | 0.67 | 0.67 | 0.25 | 0.25 | 1.98
| Analytical results for Lu in OREAS 753b (Indicative Value 0.05 ppm) | [21 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 3
3
| 1 | 1 | 0.05 | < 10 | [18 empty] | 1
| 1 | 2 | 0.05 | < 10 | [18 empty] | 14
| 1 | 3 | 0.05 | < 10 | [18 empty] | 16
| 1 | 4 | 0.05 | < 10 | [18 empty] | 0.05
| 1 | 5 | 0.04 | < 10 | [18 empty] | 20
| 1 | 6 | 0.06 | < 10
| Mean |  | 0.05
| Median |  | 0.05
| Std Dev. |  | 0.0063246
| Rel.Std.Dev. |  | 12.65%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 753b (Certified Value 0.214 wt.%) | [21 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.21 | 0.14935 | 0.24 | 0.24 | 0.21 | 0.24 | 0.223 | 0.24 | 0.2091 | 0.19 | 0.22 | 0.22 | 0.22 | 0.16 | 0.20236 | 0.21 | 0.2 | 0.22 | 0.2033 | 0.24 | 1
| 1 | 2 | 0.2 | 0.14729 | 0.24 | 0.23 | 0.19 | 0.24 | 0.216 | 0.23 | 0.2097 | 0.19 | 0.21 | 0.23 | 0.21 | 0.16 | 0.20518 | 0.21 | 0.2 | 0.21 | 0.20758 | 0.21 | 0
| 1 | 3 | 0.21 | 0.14729 | 0.23 | 0.23 | 0.2 | 0.24 | 0.213 | 0.24 | 0.21314 | 0.17 | 0.21 | 0.23 | 0.23 | 0.18 | 0.2077 | 0.21 | 0.2 | 0.22 | 0.20758 | 0.22 | 16
| 1 | 4 | 0.2 | 0.14523 | 0.23 | 0.24 | 0.19 | 0.25 | 0.214 | 0.23 | 0.20921 | 0.19 | 0.21 | 0.21 | 0.23 | 0.17 | 0.20782 | 0.21 | 0.2 | 0.21 | 0.20651 | 0.23 | 0.21363
| 1 | 5 | 0.21 | 0.1442 | 0.23 | 0.23 | 0.21 | 0.24 | 0.215 | 0.23 | 0.2084 | 0.2 | 0.21 | 0.21 | 0.22 | 0.18 | 0.2048 | 0.22 | 0.2 | 0.22 | 0.20865 | 0.24 | 122
| 1 | 6 | 0.2 | 0.1478 | 0.23 | 0.24 | 0.2 | 0.24 | 0.225 | 0.24 | 0.21156 | 0.18 | 0.22 | 0.24 | 0.22 | 0.19 | 0.2051 | 0.21 | 0.19 | 0.22 | 0.20972 | 0.2
| Mean |  | 0.205 | 0.14686 | 0.23333 | 0.235 | 0.2 | 0.24167 | 0.21767 | 0.235 | 0.21018 | 0.18667 | 0.21333 | 0.22333 | 0.22167 | 0.17333 | 0.20549 | 0.21167 | 0.19833 | 0.21667 | 0.20722 | 0.22333
| Median |  | 0.205 | 0.14729 | 0.23 | 0.235 | 0.2 | 0.24 | 0.2155 | 0.235 | 0.20945 | 0.19 | 0.21 | 0.225 | 0.22 | 0.175 | 0.20514 | 0.21 | 0.2 | 0.22 | 0.20758 | 0.225
| Std Dev. |  | 0.0054772 | 0.0018545 | 0.005164 | 0.0054772 | 0.0089443 | 0.0040825 | 0.0050465 | 0.0054772 | 0.0017993 | 0.010328 | 0.005164 | 0.012111 | 0.0075277 | 0.012111 | 0.0020406 | 0.0040825 | 0.0040825 | 0.005164 | 0.0022102 | 0.01633
| Rel.Std.Dev. |  | 2.672% | 1.263% | 2.213% | 2.331% | 4.472% | 1.689% | 2.318% | 2.331% | 0.8561% | 5.533% | 2.421% | 5.423% | 3.396% | 6.987% | 0.993% | 1.929% | 2.058% | 2.383% | 1.067% | 7.312%
| PDM3 |  | -4.038% | -31.25% | 9.225% | 10.01% | -6.379% | 13.13% | 1.891% | 10.01% | -1.611% | -12.62% | -0.1372% | 4.544% | 3.764% | -18.86% | -3.807% | -0.9174% | -7.159% | 1.423% | -2.997% | 4.544%
| Z-Score (Absolute) |  | 0.47 | 4.09 | 1.3 | 1.4 | 0.78 | 1.82 | 0.32 | 1.4 | 0.14 | 1.61 | 0.05 | 0.67 | 0.57 | 2.44 | 0.44 | 0.05 | 0.88 | 0.26 | 0.33 | 0.67
| Analytical results for MnO in OREAS 753b (Certified Value 0.178 wt.%) | [21 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.18 | 0.1872 | 0.19 | 0.177 | 0.18 | 0.18 | 0.18 | 0.17 | 0.16599 | 0.177 | 0.168 | 0.181 | 0.168 | 0.2 | 0.16687 | 0.18 | 0.17 | 0.18 | 0.1702 | 0.23 | 1
| 1 | 2 | 0.18 | 0.1872 | 0.19 | 0.177 | 0.18 | 0.18 | 0.179 | 0.18 | 0.17103 | 0.184 | 0.168 | 0.181 | 0.181 | 0.2 | 0.16763 | 0.18 | 0.17 | 0.18 | 0.17671 | 0.19 | 0
| 1 | 3 | 0.18 | 0.1872 | 0.19 | 0.177 | 0.18 | 0.17 | 0.181 | 0.18 | 0.17024 | 0.174 | 0.168 | 0.181 | 0.181 | 0.2 | 0.17316 | 0.18 | 0.17 | 0.17 | 0.17671 | 0.2 | 16
| 1 | 4 | 0.18 | 0.18512 | 0.19 | 0.186 | 0.18 | 0.18 | 0.176 | 0.18 | 0.16798 | 0.178 | 0.181 | 0.181 | 0.168 | 0.2 | 0.17353 | 0.18 | 0.17 | 0.17 | 0.17485 | 0.21 | 0.17766
| 1 | 5 | 0.18 | 0.1872 | 0.19 | 0.186 | 0.18 | 0.18 | 0.18 | 0.18 | 0.17015 | 0.177 | 0.181 | 0.181 | 0.181 | 0.19 | 0.1709 | 0.19 | 0.17 | 0.18 | 0.1722 | 0.24 | 123
| 1 | 6 | 0.18 | 0.18564 | 0.18 | 0.177 | 0.18 | 0.18 | 0.18 | 0.18 | 0.17182 | 0.181 | 0.168 | 0.181 | 0.168 | 0.2 | 0.17168 | 0.18 | 0.17 | 0.18 | 0.17299 | 0.24
| Mean |  | 0.18 | 0.18659 | 0.18833 | 0.18 | 0.18 | 0.17833 | 0.17933 | 0.17833 | 0.16954 | 0.1785 | 0.17233 | 0.181 | 0.1745 | 0.19833 | 0.17063 | 0.18167 | 0.17 | 0.17667 | 0.17395 | 0.21833
| Median |  | 0.18 | 0.1872 | 0.19 | 0.177 | 0.18 | 0.18 | 0.18 | 0.18 | 0.17019 | 0.1775 | 0.168 | 0.181 | 0.1745 | 0.2 | 0.17129 | 0.18 | 0.17 | 0.18 | 0.17392 | 0.22
| Std Dev. |  | 3.0405e-17 | 0.00095412 | 0.0040825 | 0.0046476 | 3.0405e-17 | 0.0040825 | 0.0017512 | 0.0040825 | 0.0021587 | 0.0035071 | 0.0067132 | 3.0405e-17 | 0.0071204 | 0.0040825 | 0.0027975 | 0.0040825 | 0 | 0.005164 | 0.0026124 | 0.02137
| Rel.Std.Dev. |  | 1.689e-14% | 0.5113% | 2.168% | 2.582% | 1.689e-14% | 2.289% | 0.9765% | 2.289% | 1.273% | 1.965% | 3.895% | 1.68e-14% | 4.08% | 2.058% | 1.64% | 2.247% | 0% | 2.923% | 1.502% | 9.788%
| PDM3 |  | 1.315% | 5.026% | 6.006% | 1.315% | 1.315% | 0.3772% | 0.9401% | 0.3772% | -4.575% | 0.471% | -3% | 1.878% | -1.78% | 11.63% | -3.959% | 2.253% | -4.313% | -0.5609% | -2.092% | 22.89%
| Z-Score (Absolute) |  | 0.2 | 1.41 | 1.73 | 0.16 | 0.2 | 0.11 | 0.08 | 0.11 | 1.72 | 0.08 | 1.24 | 0.34 | 0.84 | 3.56 | 1.52 | 0.5 | 1.64 | 0.41 | 0.91 | 7.23
| Analytical results for Mo in OREAS 753b (Indicative Value < 50 ppm) | [21 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | [16 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | 3 | < 50 | < 50 | < 50 | [16 empty] | 1
| 1 | 2 | 3.2 | < 50 | < 50 | < 50 | [16 empty] | 15
| 1 | 3 | 3.4 | < 50 | < 50 | < 50 | [16 empty] | 16
| 1 | 4 | 2.9 | < 50 | < 50 | < 50 | [16 empty] | < 50
| 1 | 5 | 4.7 | < 50 | < 50 | < 50 | [16 empty] | 21
| 1 | 6 | 3.6 | < 50 | < 50 | < 50
| Mean |  | 3.4667
| Median |  | 3.3
| Std Dev. |  | 0.65625
| Rel.Std.Dev. |  | 18.93%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Na2O in OREAS 753b (Certified Value 0.559 wt.%) | [21 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.52 | 0.52894 | 0.62 | 0.57 | 0.58 | 0.57 | 0.578 | 0.47 | 0.57746 | 0.52 | 0.58 | 0.53553 | 0.54998 | 0.56 | 0.55 | 0.527 | 0.62 |  |  |  | 1
| 1 | 2 | 0.52 | 0.52894 | 0.57 | 0.57 | 0.58 | 0.59 | 0.582 | 0.48 | 0.5699 | 0.5 | 0.56 | 0.54301 | 0.54998 | 0.57 | 0.54 | 0.514 | 0.59 |  |  |  | 0
| 1 | 3 | 0.54 | 0.51622 | 0.62 | 0.56 | 0.57 | 0.57 | 0.577 | 0.47 | 0.57581 | 0.47 | 0.6 | 0.55018 | 0.55133 | 0.55 | 0.53 | 0.529 | 0.64 |  |  |  | 16
| 1 | 4 | 0.52 | 0.51834 | 0.56 | 0.55 | 0.57 | 0.59 | 0.584 | 0.46 | 0.57876 | 0.53 | 0.58 | 0.55457 | 0.54864 | 0.58 | 0.54 | 0.526 | 0.66 |  |  |  | 0.55906
| 1 | 5 | 0.54 | 0.53106 | 0.59 | 0.58 | 0.59 | 0.56 | 0.577 | 0.46 | 0.57714 | 0.5 | 0.57 | 0.54244 | 0.54998 | 0.57 | 0.54 | 0.529 | 0.61 |  |  |  | 124
| 1 | 6 | 0.52 | 0.53371 | 0.58 | 0.56 | 0.58 | 0.57 | 0.575 | 0.47 | 0.58413 | 0.56 | 0.58 | 0.544 | 0.54864 | 0.56 | 0.54 | 0.531 | 0.58
| Mean |  | 0.52667 | 0.5262 | 0.59 | 0.565 | 0.57833 | 0.575 | 0.57883 | 0.46833 | 0.5772 | 0.51333 | 0.57833 | 0.54495 | 0.54976 | 0.565 | 0.54 | 0.526 | 0.61667
| Median |  | 0.52 | 0.52894 | 0.585 | 0.565 | 0.58 | 0.57 | 0.5775 | 0.47 | 0.5773 | 0.51 | 0.58 | 0.5435 | 0.54998 | 0.565 | 0.54 | 0.528 | 0.615
| Std Dev. |  | 0.010328 | 0.0071612 | 0.025298 | 0.010488 | 0.0075277 | 0.012247 | 0.0034303 | 0.0075277 | 0.0045989 | 0.030768 | 0.013292 | 0.0066269 | 0.0010118 | 0.010488 | 0.0063246 | 0.0061319 | 0.030111
| Rel.Std.Dev. |  | 1.961% | 1.361% | 4.288% | 1.856% | 1.302% | 2.13% | 0.5926% | 1.607% | 0.7968% | 5.994% | 2.298% | 1.216% | 0.184% | 1.856% | 1.171% | 1.166% | 4.883%
| PDM3 |  | -5.795% | -5.878% | 5.534% | 1.062% | 3.447% | 2.851% | 3.536% | -16.23% | 3.244% | -8.18% | 3.447% | -2.524% | -1.664% | 1.062% | -3.41% | -5.914% | 10.3%
| Z-Score (Absolute) |  | 1.29 | 1.31 | 0.84 | 0 | 0.45 | 0.34 | 0.47 | 3.25 | 0.41 | 1.74 | 0.45 | 0.67 | 0.51 | 0 | 0.84 | 1.31 | 1.74
| Analytical results for Nb in OREAS 753b (Certified Value 93 ppm) | [21 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.12 | 1.13 | 1.14 | 1.16 | 1.22 | 1.26 | 1.28 | [12 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [12 empty] | 0
0
| 1 | 1 | 70 | 102 | 70 | < 50 | 100 | 110 | 97.8679 | 106.61 | [12 empty] | 1
| 1 | 2 | 70 | 103 | 70 | 50 | 120 | 170 | 90.8773 | 103.41 | [12 empty] | 0
| 1 | 3 | 70 | 103 | 80 | < 50 | 120 | 110 | 97.8679 | 106.42 | [12 empty] | 16
| 1 | 4 | 70 | 104 | 90 | 60 | 120 | 80 | 76.8962 | 107.38 | [12 empty] | 93.2649
| 1 | 5 | 70 | 103 | 100 | < 50 | 130 | 210 | 76.8962 | 104.75 | [12 empty] | 125
| 1 | 6 | 70 | 103 | 90 | < 50 | 110 | 150 | 90.8773 | 105.43
| Mean |  | 70 | 103 | 83.3333 | 55 | 116.6667 | 138.3333 | 88.5471 | 105.6667
| Median |  | 70 | 103 | 85 | 55 | 120 | 130 | 90.8773 | 105.925
| Std Dev. |  | 0 | 0.63246 | 12.1106 | 7.0711 | 10.328 | 47.5044 | 9.5509 | 1.441
| Rel.Std.Dev. |  | 0% | 0.614% | 14.53% | 12.86% | 8.853% | 34.34% | 10.79% | 1.364%
| PDM3 |  | -24.94% | 10.44% | -10.65% | -41.03% | 25.09% | 48.32% | -5.058% | 13.3%
| Z-Score (Absolute) |  | 1.05 | 0.29 | 0.5 | 2.46 | 0.85 | 1.72 | 0.29 | 0.4
| Analytical results for Nd in OREAS 753b (Indicative Value 3.82 ppm) | [21 empty] | Indicative
Nd | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 3.4 | < 10 | [18 empty] | 1
| 1 | 2 | 3.9 | < 10 | [18 empty] | 16
| 1 | 3 | 4.2 | < 10 | [18 empty] | 16
| 1 | 4 | 3.8 | < 10 | [18 empty] | 3.8167
| 1 | 5 | 3.6 | < 10 | [18 empty] | 22
| 1 | 6 | 4 | < 10
| Mean |  | 3.8167
| Median |  | 3.85
| Std Dev. |  | 0.28577
| Rel.Std.Dev. |  | 7.488%
| PDM3 |  | -7.772e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for NiO in OREAS 753b (Indicative Value 26.1 ppm) | [21 empty] | Indicative
NiO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.09 | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 30 | 10.2 | < 64 | < 64 | < 64 | 50 | [14 empty] | 1
| 1 | 2 | < 30 | 11.5 | 64 | < 64 | < 64 | 30 | [14 empty] | 17
| 1 | 3 | < 30 | 7.6 | < 64 | < 64 | < 64 | 40 | [14 empty] | 16
| 1 | 4 | < 30 | < 6.4 | < 64 | < 64 | < 64 | 30 | [14 empty] | 26.1137
| 1 | 5 | < 30 | 12.7 | < 64 | < 64 | < 64 | 30 | [14 empty] | 23
| 1 | 6 | < 30 | 14 | < 64 | < 64 | < 64 | 40
| Mean |  |  | 11.2 | 64 |  |  | 36.6667
| Median |  |  | 11.5 | 64 |  |  | 35
| Std Dev. |  |  | 2.4566 |  |  |  | 8.165
| Rel.Std.Dev. |  |  | 21.93% |  |  |  | 22.27%
| PDM3 |  |  | -57.11% | 145.1% |  |  | 40.41%
| Z-Score (Absolute) |  | 2.23 | 2.91 | 0.7 | 0 | 0 | 0.64
| Analytical results for P2O5 in OREAS 753b (Certified Value 0.11 wt.%) | [21 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.113 | 0.107 | 0.1 | 0.11 | 0.112 | 0.116 | 0.115 | 0.1 | 0.10564 | 0.11 | 0.1 | 0.1 | 0.09 | 0.11 | 0.10157 | 0.11 | 0.11 | 0.1192 | 0.1122 | 0.21 | 1
| 1 | 2 | 0.112 | 0.107 | 0.09 | 0.11 | 0.112 | 0.119 | 0.115 | 0.1 | 0.1052 | 0.11 | 0.1 | 0.11 | 0.09 | 0.11 | 0.10365 | 0.11 | 0.11 | 0.1169 | 0.11118 | 0.22 | 0
| 1 | 3 | 0.111 | 0.107 | 0.1 | 0.11 | 0.111 | 0.117 | 0.113 | 0.1 | 0.10613 | 0.1 | 0.1 | 0.11 | 0.09 | 0.11 | 0.10599 | 0.11 | 0.11 | 0.1215 | 0.1122 | 0.24 | 16
| 1 | 4 | 0.112 | 0.107 | 0.11 | 0.11 | 0.109 | 0.121 | 0.118 | 0.1 | 0.10624 | 0.08 | 0.1 | 0.1 | 0.1 | 0.11 | 0.10507 | 0.11 | 0.11 | 0.1169 | 0.11016 | 0.26 | 0.10955
| 1 | 5 | 0.109 | 0.10593 | 0.1 | 0.1 | 0.111 | 0.122 | 0.116 | 0.1 | 0.10514 | 0.14 | 0.1 | 0.1 | 0.1 | 0.11 | 0.10239 | 0.11 | 0.11 | 0.1123 | 0.10914 | 0.27 | 126
| 1 | 6 | 0.113 | 0.10647 | 0.1 | 0.11 | 0.111 | 0.118 | 0.114 | 0.1 | 0.10563 | 0.12 | 0.1 | 0.1 | 0.09 | 0.11 | 0.10159 | 0.12 | 0.11 | 0.1123 | 0.10812 | 0.24
| Mean |  | 0.11167 | 0.10673 | 0.1 | 0.10833 | 0.111 | 0.11883 | 0.11517 | 0.1 | 0.10566 | 0.11 | 0.1 | 0.10333 | 0.093333 | 0.11 | 0.10338 | 0.11167 | 0.11 | 0.11652 | 0.1105 | 0.24
| Median |  | 0.112 | 0.107 | 0.1 | 0.11 | 0.111 | 0.1185 | 0.115 | 0.1 | 0.10564 | 0.11 | 0.1 | 0.1 | 0.09 | 0.11 | 0.10302 | 0.11 | 0.11 | 0.1169 | 0.11067 | 0.24
| Std Dev. |  | 0.0015055 | 0.00044761 | 0.0063246 | 0.0040825 | 0.0010954 | 0.0023166 | 0.0017224 | 1.5202e-17 | 0.00045522 | 0.02 | 1.5202e-17 | 0.005164 | 0.005164 | 0 | 0.0018559 | 0.0040825 | 0 | 0.0036848 | 0.0016657 | 0.022804
| Rel.Std.Dev. |  | 1.348% | 0.4194% | 6.325% | 3.768% | 0.9869% | 1.949% | 1.496% | 1.52e-14% | 0.4308% | 18.18% | 1.52e-14% | 4.997% | 5.533% | 0% | 1.795% | 3.656% | 0% | 3.162% | 1.507% | 9.501%
| PDM3 |  | 1.934% | -2.57% | -8.715% | -1.108% | 1.326% | 8.476% | 5.129% | -8.715% | -3.547% | 0.413% | -8.715% | -5.673% | -14.8% | 0.413% | -5.633% | 1.934% | 0.413% | 6.362% | 0.8694% | 119.1%
| Z-Score (Absolute) |  | 0.3 | 0.58 | 1.77 | 0.3 | 0.18 | 1.57 | 0.92 | NA | 0.77 | 0 | NA | 1.18 | 2.96 | 0 | 1.17 | 0.3 | 0 | 1.15 | 0.09 | 23.05
| Analytical results for Pb in OREAS 753b (Indicative Value 11.4 ppm) | [21 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 1
1
| 1 | 1 | 4.2 | < 50 | < 50 | < 50 | 19 | [15 empty] | 1
| 1 | 2 | 4.1 | < 50 | < 50 | < 50 | [16 empty] | 7
| 1 | 3 | 4.5 | < 50 | < 50 | < 50 | [16 empty] | 16
| 1 | 4 | 4.2 | < 50 | < 50 | < 50 | 28 | [15 empty] | 11.4
| 1 | 5 | 4.2 | < 50 | < 50 | < 50 | [16 empty] | 24
| 1 | 6 | 4.2 | < 50 | < 50 | < 50 | 9
| Mean |  | 4.2333 |  |  |  | 18.6667
| Median |  | 4.2 |  |  |  | 19
| Std Dev. |  | 0.13663 |  |  |  | 9.5044
| Rel.Std.Dev. |  | 3.227% |  |  |  | 50.92%
| PDM3 |  | -62.87% |  |  |  | 63.74%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Pr in OREAS 753b (Indicative Value 1.06 ppm) | [21 empty] | Indicative
Pr | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 1.05 | < 10 | [18 empty] | 1
| 1 | 2 | 1.05 | < 10 | [18 empty] | 19
| 1 | 3 | 1.15 | < 10 | [18 empty] | 16
| 1 | 4 | 1.03 | < 10 | [18 empty] | 1.0633
| 1 | 5 | 0.98 | < 10 | [18 empty] | 25
| 1 | 6 | 1.12 | < 10
| Mean |  | 1.0633
| Median |  | 1.05
| Std Dev. |  | 0.06186
| Rel.Std.Dev. |  | 5.818%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Rb in OREAS 753b (Certified Value 1200 ppm) | [21 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.22 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 1220 | 1120 | 1180 | 1170 | 1280 | [15 empty] | 1
| 1 | 2 | 1260 | 1110 | 1180 | 1180 | 1280 | [15 empty] | 0
| 1 | 3 | 1210 | 1150 | 1200 | 1170 | 1280 | [15 empty] | 16
| 1 | 4 | 1230 | 1180 | 1160 | 1120 | 1280 | [15 empty] | 1200.0351
| 1 | 5 | 1210 | 1190 | 1160 | 1180 | 1280 | [15 empty] | 127
| 1 | 6 | 1230 | 1180 | 1200 | 1130 | 1280
| Mean |  | 1226.6667 | 1155 | 1180 | 1158.3333 | 1280
| Median |  | 1225 | 1165 | 1180 | 1170 | 1280
| Std Dev. |  | 18.619 | 33.9116 | 17.8885 | 26.3944 | 0
| Rel.Std.Dev. |  | 1.518% | 2.936% | 1.516% | 2.279% | 0%
| PDM3 |  | 2.219% | -3.753% | -1.67% | -3.475% | 6.664%
| Z-Score (Absolute) |  | 1.26 | 0.67 | 0 | 0.58 | 2.7
| Analytical results for Re in OREAS 753b (Indicative Value < 0.1 ppm) | [21 empty] | Indicative
Re | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | < 0.1 | [19 empty] | 1
| 1 | 2 | < 0.1 | [19 empty] | 20
| 1 | 3 | < 0.1 | [19 empty] | 16
| 1 | 4 | < 0.1 | [19 empty] | < 0.1
| 1 | 5 | < 0.1 | [19 empty] | 26
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 753b (Indicative Value < 50 ppm) | [21 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | [16 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 0
0
| 1 | 1 | 1 | < 50 | < 50 | < 50 | [16 empty] | 1
| 1 | 2 | < 1 | < 50 | < 50 | < 50 | [16 empty] | 21
| 1 | 3 | 1 | < 50 | < 50 | < 50 | [16 empty] | 16
| 1 | 4 | < 1 | 80 | < 50 | < 50 | [16 empty] | < 50
| 1 | 5 | < 1 | < 50 | < 50 | < 50 | [16 empty] | 27
| 1 | 6 | < 1 | 90 | < 50 | < 50
| Mean |  | 1 | 85
| Median |  | 1 | 85
| Std Dev. |  | 0 | 7.0711
| Rel.Std.Dev. |  | 0% | 8.319%
PDM3
| Z-Score (Absolute) |  | 1.64 | 1.35 | 0 | 0
| Analytical results for Sc in OREAS 753b (Indicative Value 2.67 ppm) | [21 empty] | Indicative
Sc | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | 2 | [19 empty] | 1
| 1 | 2 | 3 | [19 empty] | 22
| 1 | 3 | 3 | [19 empty] | 16
| 1 | 4 | 3 | [19 empty] | 2.6667
| 1 | 5 | 2 | [19 empty] | 28
| 1 | 6 | 3
| Mean |  | 2.6667
| Median |  | 3
| Std Dev. |  | 0.5164
| Rel.Std.Dev. |  | 19.36%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 753b (Indicative Value 0.92 ppm) | [21 empty] | Indicative
Se | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | 1 | [19 empty] | 1
| 1 | 2 | < 1 | [19 empty] | 2
| 1 | 3 | < 1 | [19 empty] | 16
| 1 | 4 | < 1 | [19 empty] | 0.91667
| 1 | 5 | 1 | [19 empty] | 29
| 1 | 6 | 2
| Mean |  | 1.3333
| Median |  | 1
| Std Dev. |  | 0.57735
| Rel.Std.Dev. |  | 43.3%
| PDM3 |  | 45.45%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 753b (Certified Value 81.48 wt.%) | [21 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 81.26 | 80.926 | 82.66 | 81.61 | 81.93 | 81.01 | 81.14 | 80.43 | 81.6604 | 84.47 | 81.9 | 80.2 | 78.6 | 82.89 | 81.9825 | 80.75 | 81.24 | 81.61 | 81.1216 | 82.25 | 1
| 1 | 2 | 81.04 | 81.15 | 82.56 | 82.2 | 81.88 | 81 | 81.017 | 80.61 | 81.7248 | 80.4 | 81.6 | 80.3 | 77.4 | 82.77 | 81.9495 | 80.65 | 81.36 | 81.64 | 80.9125 | 82.54 | 0
| 1 | 3 | 81.15 | 80.488 | 82.56 | 81.76 | 81.98 | 81.12 | 81.094 | 80.5 | 81.7137 | 85.5 | 81.8 | 80.5 | 77.7 | 82.6 | 81.9213 | 82.2 | 81.3 | 81.75 | 81.0163 | 82.54 | 16
| 1 | 4 | 80.95 | 80.271 | 82.61 | 81.78 | 81.65 | 81.07 | 80.676 | 80.67 | 81.4873 | 78.41 | 81.1 | 79.6 | 76.8 | 83.84 | 82.0507 | 81.49 | 81.25 | 81.84 | 80.9063 | 82.27 | 81.4827
| 1 | 5 | 81.06 | 81.06 | 82.44 | 82.08 | 81.72 | 81.11 | 82.133 | 79.97 | 81.6893 | 85.33 | 81.6 | 78.7 | 78.7 | 83.41 | 82.015 | 83.13 | 81.51 | 81.44 | 80.9259 | 82.44 | 128
| 1 | 6 | 81.21 | 80.833 | 82.63 | 81.53 | 81.62 | 81.11 | 81.887 | 80.66 | 81.6957 | 85.79 | 81.5 | 79.8 | 76.5 | 82.63 | 82.0579 | 81.62 | 81.38 | 81.43 | 80.9914 | 82.47
| Mean |  | 81.1117 | 80.788 | 82.5767 | 81.8267 | 81.7967 | 81.07 | 81.3245 | 80.4733 | 81.6619 | 83.3167 | 81.5833 | 79.85 | 77.6167 | 83.0233 | 81.9961 | 81.64 | 81.34 | 81.6183 | 80.979 | 82.4183
| Median |  | 81.105 | 80.8795 | 82.585 | 81.77 | 81.8 | 81.09 | 81.117 | 80.555 | 81.6925 | 84.9 | 81.6 | 80 | 77.55 | 82.83 | 81.9987 | 81.555 | 81.33 | 81.625 | 80.9587 | 82.455
| Std Dev. |  | 0.11583 | 0.34158 | 0.077632 | 0.26273 | 0.15293 | 0.053292 | 0.56078 | 0.26387 | 0.088331 | 3.1257 | 0.27869 | 0.65345 | 0.90646 | 0.49645 | 0.05497 | 0.93046 | 0.1006 | 0.16388 | 0.08294 | 0.12891
| Rel.Std.Dev. |  | 0.1428% | 0.4228% | 0.09401% | 0.3211% | 0.187% | 0.06574% | 0.6896% | 0.3279% | 0.1082% | 3.752% | 0.3416% | 0.8183% | 1.168% | 0.598% | 0.06704% | 1.14% | 0.1237% | 0.2008% | 0.1024% | 0.1564%
| PDM3 |  | -0.4553% | -0.8526% | 1.343% | 0.4221% | 0.3853% | -0.5065% | -0.1941% | -1.239% | 0.2199% | 2.251% | 0.1235% | -2.004% | -4.745% | 1.891% | 0.6301% | 0.1931% | -0.1751% | 0.1665% | -0.6182% | 1.148%
| Z-Score (Absolute) |  | 0.65 | 1.07 | 1.29 | 0.3 | 0.26 | 0.7 | 0.37 | 1.49 | 0.08 | 2.27 | 0.02 | 2.31 | 5.27 | 1.88 | 0.52 | 0.05 | 0.34 | 0.02 | 0.82 | 1.08
| Analytical results for Sm in OREAS 753b (Indicative Value 0.87 ppm) | [21 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.7 | < 10 | [18 empty] | 1
| 1 | 2 | 0.9 | < 10 | [18 empty] | 10
| 1 | 3 | 1 | < 10 | [18 empty] | 16
| 1 | 4 | 0.9 | < 10 | [18 empty] | 0.86667
| 1 | 5 | 0.7 | < 10 | [18 empty] | 16
| 1 | 6 | 1 | < 10
| Mean |  | 0.86667
| Median |  | 0.9
| Std Dev. |  | 0.13663
| Rel.Std.Dev. |  | 15.76%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 753b (Certified Value 151 ppm) | [21 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.12 | 1.13 | 1.14 | 1.16 | 1.22 | 1.26 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 130 | 165 | 70 | 140 | 90 | 160 | 165.4064 | [13 empty] | 1
| 1 | 2 | 150 | 168 | 70 | 140 | 130 | 240 | 157.5299 | [13 empty] | 0
| 1 | 3 | 170 | 162 | 80 | 130 | 90 | 240 | 165.4064 | [13 empty] | 16
| 1 | 4 | 160 | 165 | 70 | 140 | 100 | 160 | 165.4064 | [13 empty] | 150.9245
| 1 | 5 | 150 | 166 | 70 | 80 | 100 | 160 | 157.5299 | [13 empty] | 129
| 1 | 6 | 160 | 165 | 80 | 110 | 70 | 240 | 157.5299
| Mean |  | 153.3333 | 165.1667 | 73.3333 | 123.3333 | 96.6667 | 200 | 161.4682
| Median |  | 155 | 165 | 70 | 135 | 95 | 200 | 161.4682
| Std Dev. |  | 13.6626 | 1.9408 | 5.164 | 24.2212 | 19.6638 | 43.8178 | 4.3141
| Rel.Std.Dev. |  | 8.91% | 1.175% | 7.042% | 19.64% | 20.34% | 21.91% | 2.672%
| PDM3 |  | 1.596% | 9.437% | -51.41% | -18.28% | -35.95% | 32.52% | 6.986%
| Z-Score (Absolute) |  | 0 | 0.27 | 1.8 | 0.67 | 1.27 | 0.98 | 0.18
| Analytical results for SO3 in OREAS 753b (Certified Value 0.057 wt.%) | [21 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.06 | 1.07 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.26 | 1.28 | [7 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 3
3
| 1 | 1 | 0.075 | 0.041 | 0.04 | 0.068 | 0.05 | 0.075 | 0.05 | < 0.01 | 0.058631 | 0.0749 | 0.034 | 0.057 | 0.37 | [7 empty] | 1
| 1 | 2 | 0.05 | 0.04 | 0.0549 | 0.075 | 0.05 | 0.075 | 0.05 | < 0.01 | 0.066018 | 0.0749 | 0.031 | 0.059 | 0.42 | [7 empty] | 0
| 1 | 3 | 0.075 | 0.041 | 0.0574 | 0.075 | 0.05 | 0.05 | 0.05 | < 0.01 | 0.06774 | 0.0699 | 0.032 | 0.058143 | 0.38 | [7 empty] | 16
| 1 | 4 | 0.05 | 0.04 | 0.0549 | 0.07 | 0.05 | 0.05 | 0.05 | < 0.01 | 0.059324 | 0.0799 | 0.033 | 0.055 | 0.41 | [7 empty] | 0.056915
| 1 | 5 | 0.05 | 0.042 | 0.0549 | 0.077 | 0.05 | 0.075 | 0.05 | < 0.01 | 0.062298 | 0.0799 | 0.04 | 0.06 | 0.3 | [7 empty] | 130
| 1 | 6 | 0.075 | 0.04 | 0.0574 | 0.079 | 0.05 | 0.05 | 0.05 | < 0.01 | 0.063146 | 0.0724 | 0.039 | 0.057 | 0.38
| Mean |  | 0.0625 | 0.040667 | 0.05325 | 0.074 | 0.05 | 0.0625 | 0.05 |  | 0.062859 | 0.075317 | 0.034833 | 0.05769 | 0.37667
| Median |  | 0.0625 | 0.0405 | 0.0549 | 0.075 | 0.05 | 0.0625 | 0.05 |  | 0.062722 | 0.0749 | 0.0335 | 0.057571 | 0.38
| Std Dev. |  | 0.013693 | 0.0008165 | 0.0066057 | 0.0041952 | 7.6012e-18 | 0.013693 | 7.6012e-18 |  | 0.0035936 | 0.0040052 | 0.0037639 | 0.0017576 | 0.042269
| Rel.Std.Dev. |  | 21.91% | 2.008% | 12.41% | 5.669% | 1.52e-14% | 21.91% | 1.52e-14% |  | 5.717% | 5.318% | 10.81% | 3.047% | 11.22%
| PDM3 |  | 9.813% | -28.55% | -6.44% | 30.02% | -12.15% | 9.813% | -12.15% |  | 10.44% | 32.33% | -38.8% | 1.362% | 561.8%
| Z-Score (Absolute) |  | 0.41 | 1.48 | 0.38 | 1.42 | 0.67 | 0.41 | 0.67 | 4.59 | 0.45 | 1.54 | 1.99 | 0 | 27.77
| Analytical results for SrO in OREAS 753b (Indicative Value 85 ppm) | [21 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.23 | 1.26 | 1.28 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 50 | < 100 | < 100 | 58.9615 | 47.3 | < 120 | 120 | < 120 | 100 | 95 | 120 | < 100 | [8 empty] | 1
| 1 | 2 | < 50 | < 100 | 200 | 57.5945 | 48.5 | < 120 | 120 | < 120 | 100 | 95 | 130 | < 100 | [8 empty] | 11
| 1 | 3 | < 50 | < 100 | 200 | 56.846 | 48.5 | < 120 | 120 | < 120 | 100 | 47 | 140 | < 100 | [8 empty] | 16
| 1 | 4 | < 50 | < 100 | < 100 | 57.4039 | 47.3 | < 120 | 120 | < 120 | 100 | 71 | 97.1429 | < 100 | [8 empty] | 85.4507
| 1 | 5 | < 50 | < 100 | < 100 | 56.5566 | 48.5 | < 120 | < 120 | < 120 | 100 | 59 | 100 | < 100 | [8 empty] | 17
| 1 | 6 | < 50 | < 100 | < 100 | 56.2247 | 48.5 | < 120 | 120 | < 120 | 100 | 47 | 105.3061 | < 100
| Mean |  |  |  | 200 | 57.2645 | 48.1 |  | 120 |  | 100 | 69 | 115.4082
| Median |  |  |  | 200 | 57.1249 | 48.5 |  | 120 |  | 100 | 65 | 112.6531
| Std Dev. |  |  |  | 0 | 0.97627 | 0.61968 |  | 0 |  | 0 | 22.0182 | 17.3888
| Rel.Std.Dev. |  |  |  | 0% | 1.705% | 1.288% |  | 0% |  | 0% | 31.91% | 15.07%
| PDM3 |  |  |  | 134.1% | -32.99% | -43.71% |  | 40.43% |  | 17.03% | -19.25% | 35.06%
| Z-Score (Absolute) |  | 2.2 | 0.59 | 2.64 | 0.12 | 0.71 | 0 | 3.18 | 0 | 2.64 | 0.64 | 3.63 | 0.59
| Analytical results for Ta in OREAS 753b (Indicative Value 263 ppm) | [21 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.01 | 1.12 | 1.22 | 1.26 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 330 | 265 | 270 | 262 | [16 empty] | 1
| 1 | 2 | 250 | 261 | 250 | 270 | [16 empty] | 12
| 1 | 3 | 250 | 265 | 300 | 287 | [16 empty] | 16
| 1 | 4 | 250 | 264 | 260 | 262 | [16 empty] | 262.8492
| 1 | 5 | 250 | 263 | 280 | 262 | [16 empty] | 18
| 1 | 6 | 250 | 263 | 280 | 270
| Mean |  | 263.3333 | 263.5 | 273.3333 | 268.8333
| Median |  | 250 | 263.5 | 275 | 266
| Std Dev. |  | 32.6599 | 1.5166 | 17.5119 | 9.7245
| Rel.Std.Dev. |  | 12.4% | 0.5756% | 6.407% | 3.617%
| PDM3 |  | 0.1842% | 0.2476% | 3.989% | 2.277%
| Z-Score (Absolute) |  | 0.97 | 0.38 | 1.01 | 0.38
| Analytical results for Tb in OREAS 753b (Indicative Value 0.13 ppm) | [21 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.12 | < 10 | [18 empty] | 1
| 1 | 2 | 0.14 | < 10 | [18 empty] | 13
| 1 | 3 | 0.15 | < 10 | [18 empty] | 16
| 1 | 4 | 0.13 | < 10 | [18 empty] | 0.13
| 1 | 5 | 0.13 | < 10 | [18 empty] | 19
| 1 | 6 | 0.11 | < 10
| Mean |  | 0.13
| Median |  | 0.13
| Std Dev. |  | 0.014142
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 753b (Indicative Value < 0.1 ppm) | [21 empty] | Indicative
Te | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | < 0.1 | [19 empty] | 1
| 1 | 2 | < 0.1 | [19 empty] | 6
| 1 | 3 | < 0.1 | [19 empty] | 16
| 1 | 4 | < 0.1 | [19 empty] | < 0.1
| 1 | 5 | < 0.1 | [19 empty] | 20
| 1 | 6 | < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 753b (Indicative Value 21.1 ppm) | [21 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.12 | 1.22 | 1.26 | 1.28 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [16 empty] | 1
1
| 1 | 1 | 2.1 | 53 | 43.9406 | < 10 | [16 empty] | 1
| 1 | 2 | 2.3 | 44 | 41.4297 | < 10 | [16 empty] | 15
| 1 | 3 | 2.6 | < 44 | 26.3644 | < 10 | [16 empty] | 16
| 1 | 4 | 2.2 | < 44 | 26.3644 | < 10 | [16 empty] | 21.1328
| 1 | 5 | 2.2 | < 44 | 35.1525 | < 10 | [16 empty] | 21
| 1 | 6 | 2.4 | < 44 | 35.1525 | < 10
| Mean |  | 2.3 | 48.5 | 34.734
| Median |  | 2.25 | 48.5 | 35.1525
| Std Dev. |  | 0.17889 | 6.364 | 7.3491
| Rel.Std.Dev. |  | 7.778% | 13.12% | 21.16%
| PDM3 |  | -89.12% | 129.5% | 64.36%
| Z-Score (Absolute) |  | 0.74 | 0.61 | 0.8 | 0.61
| Analytical results for TiO2 in OREAS 753b (Certified Value 0.087 wt.%) | [21 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.09 | 0.07777 | 0.09 | 0.08 | 0.1 | 0.08 | 0.093 | 0.09 | 0.10911 | 0.088 | 0.08 | 0.08 | 0.1 | 0.09 | 0.082706 | 0.09 | 0.09 | 0.08 | 0.0816 | 0.25 | 1
| 1 | 2 | 0.09 | 0.07777 | 0.08 | 0.09 | 0.09 | 0.08 | 0.095 | 0.09 | 0.10797 | 0.086 | 0.08 | 0.08 | 0.09 | 0.09 | 0.083813 | 0.09 | 0.09 | 0.08 | 0.08262 | 0.23 | 0
| 1 | 3 | 0.09 | 0.07979 | 0.08 | 0.09 | 0.08 | 0.08 | 0.094 | 0.09 | 0.1086 | 0.084 | 0.09 | 0.09 | 0.1 | 0.1 | 0.086592 | 0.09 | 0.09 | 0.08 | 0.07956 | 0.23 | 16
| 1 | 4 | 0.09 | 0.0808 | 0.09 | 0.09 | 0.08 | 0.07 | 0.09 | 0.09 | 0.10733 | 0.085 | 0.09 | 0.08 | 0.08 | 0.09 | 0.085044 | 0.09 | 0.09 | 0.08 | 0.086 | 0.25 | 0.086777
| 1 | 5 | 0.09 | 0.07878 | 0.09 | 0.09 | 0.09 | 0.07 | 0.097 | 0.09 | 0.10762 | 0.086 | 0.09 | 0.08 | 0.09 | 0.09 | 0.084292 | 0.09 | 0.11 | 0.08 | 0.083 | 0.25 | 131
| 1 | 6 | 0.09 | 0.086355 | 0.08 | 0.09 | 0.09 | 0.07 | 0.095 | 0.09 | 0.1084 | 0.085 | 0.09 | 0.09 | 0.1 | 0.1 | 0.084021 | 0.09 | 0.1 | 0.08 | 0.08874 | 0.23
| Mean |  | 0.09 | 0.080211 | 0.085 | 0.088333 | 0.088333 | 0.075 | 0.094 | 0.09 | 0.10817 | 0.085667 | 0.086667 | 0.083333 | 0.093333 | 0.093333 | 0.084411 | 0.09 | 0.095 | 0.08 | 0.083587 | 0.24
| Median |  | 0.09 | 0.079285 | 0.085 | 0.09 | 0.09 | 0.075 | 0.0945 | 0.09 | 0.10819 | 0.0855 | 0.09 | 0.08 | 0.095 | 0.09 | 0.084156 | 0.09 | 0.09 | 0.08 | 0.08281 | 0.24
| Std Dev. |  | 1.5202e-17 | 0.0032323 | 0.0054772 | 0.0040825 | 0.0075277 | 0.0054772 | 0.0023664 | 1.5202e-17 | 0.0006599 | 0.0013663 | 0.005164 | 0.005164 | 0.008165 | 0.005164 | 0.0013106 | 1.5202e-17 | 0.0083666 | 0 | 0.0032809 | 0.010954
| Rel.Std.Dev. |  | 1.689e-14% | 4.03% | 6.444% | 4.622% | 8.522% | 7.303% | 2.517% | 1.689e-14% | 0.61% | 1.595% | 5.958% | 6.197% | 8.748% | 5.533% | 1.553% | 1.689e-14% | 8.807% | 0% | 3.925% | 4.564%
| PDM3 |  | 3.715% | -7.566% | -2.047% | 1.794% | 1.794% | -13.57% | 8.324% | 3.715% | 24.66% | -1.279% | -0.1268% | -3.968% | 7.556% | 7.556% | -2.726% | 3.715% | 9.476% | -7.809% | -3.676% | 176.6%
| Z-Score (Absolute) |  | 0.23 | 1.12 | 0.46 | 0 | 0 | 1.84 | 0.78 | 0.23 | 2.75 | 0.37 | 0.23 | 0.69 | 0.69 | 0.69 | 0.54 | 0.23 | 0.92 | 1.15 | 0.66 | 20.99
| Analytical results for Tl in OREAS 753b (Indicative Value 8.23 ppm) | [21 empty] | Indicative
Tl | Round | Replicate | Lab | [19 empty] | 1
| No | No | 1.12 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 2
2
| 1 | 1 | 8.27 | [19 empty] | 1
| 1 | 2 | 8.33 | [19 empty] | 16
| 1 | 3 | 8.26 | [19 empty] | 16
| 1 | 4 | 8.41 | [19 empty] | 8.23
| 1 | 5 | 8.43 | [19 empty] | 22
| 1 | 6 | 7.68
| Mean |  | 8.23
| Median |  | 8.3
| Std Dev. |  | 0.27835
| Rel.Std.Dev. |  | 3.382%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 753b (Indicative Value 0.052 ppm) | [21 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 3
3
| 1 | 1 | 0.04 | < 10 | [18 empty] | 1
| 1 | 2 | 0.05 | < 10 | [18 empty] | 1
| 1 | 3 | 0.05 | < 10 | [18 empty] | 16
| 1 | 4 | 0.05 | < 10 | [18 empty] | 0.051667
| 1 | 5 | 0.06 | < 10 | [18 empty] | 23
| 1 | 6 | 0.06 | < 10
| Mean |  | 0.051667
| Median |  | 0.05
| Std Dev. |  | 0.0075277
| Rel.Std.Dev. |  | 14.57%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for U in OREAS 753b (Indicative Value 7.05 ppm) | [21 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | [17 empty] | 1
| No | No | 1.12 | 1.22 | 1.26 | [17 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [17 empty] | 2
2
| 1 | 1 | 2.27 | < 42 | 8 | [17 empty] | 1
| 1 | 2 | 2.24 | < 42 | 8 | [17 empty] | 18
| 1 | 3 | 2.45 | < 42 | 8 | [17 empty] | 16
| 1 | 4 | 2.18 | < 42 | [18 empty] | 7.0546
| 1 | 5 | 2.16 | < 42 | 17 | [17 empty] | 24
| 1 | 6 | 2.12 | < 42 | 17
| Mean |  | 2.2367 |  | 11.6
| Median |  | 2.21 |  | 8
| Std Dev. |  | 0.11776 |  | 4.9295
| Rel.Std.Dev. |  | 5.265% |  | 42.5%
| PDM3 |  | -68.29% |  | 64.43%
| Z-Score (Absolute) |  | 0.71 | 0.67 | 0
| Analytical results for V2O5 in OREAS 753b (Certified Value < 100 ppm) | [21 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.04 | 1.07 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.16 | 1.22 | 1.25 | 1.26 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | 60 | < 30 | < 100 | 14.3 | < 180 | < 180 | < 180 | < 100 | < 100 | 30 | [9 empty] | 1
| 1 | 2 | < 100 | 70 | 30 | < 100 | 14.3 | < 180 | < 180 | < 180 | < 100 | < 100 | [10 empty] | 0
| 1 | 3 | < 100 | 60 | < 30 | < 100 | 16.1 | < 180 | < 180 | < 180 | < 100 | < 100 | 20 | [9 empty] | 16
| 1 | 4 | < 100 | 60 | < 30 | < 100 | 16.1 | < 180 | < 180 | < 180 | < 100 | < 100 | 10 | [9 empty] | < 100
| 1 | 5 | < 100 | 50 | < 30 | < 100 | 16.1 | < 180 | < 180 | < 180 | < 100 | < 100 | [10 empty] | 132
| 1 | 6 | < 100 | 60 | 50 | < 100 | 14.3 | < 180 | < 180 | < 180 | < 100 | < 100 | 10
| Mean |  |  | 60 | 40 |  | 15.2 | [5 empty] | 17.5
| Median |  |  | 60 | 40 |  | 15.2 | [5 empty] | 15
| Std Dev. |  |  | 6.3246 | 14.1421 |  | 0.9859 | [5 empty] | 9.5743
| Rel.Std.Dev. |  |  | 10.54% | 35.36% |  | 6.486% | [5 empty] | 54.71%
PDM3
| Z-Score (Absolute) |  | 0 | 0.25 | 0.67 | 0 | 0.88 | 0.99 | 0.99 | 0.99 | 0 | 0 | 0.82
| Analytical results for W in OREAS 753b (Indicative Value < 10 ppm) | [21 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | [16 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [16 empty] | 1
1
| 1 | 1 | 3.8 | < 10 | < 10 | < 10 | [16 empty] | 1
| 1 | 2 | 3.3 | < 10 | < 10 | < 10 | [16 empty] | 19
| 1 | 3 | 3.6 | < 10 | < 10 | < 10 | [16 empty] | 16
| 1 | 4 | 3.3 | < 10 | < 10 | < 10 | [16 empty] | < 10
| 1 | 5 | 4 | < 10 | < 10 | < 10 | [16 empty] | 25
| 1 | 6 | 3.2 | < 10 | < 10 | < 10
| Mean |  | 3.5333
| Median |  | 3.45
| Std Dev. |  | 0.32042
| Rel.Std.Dev. |  | 9.068%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Y in OREAS 753b (Indicative Value 42 ppm) | [21 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | 1.28 | [14 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | 3.7 | 47 | 47 | < 39 | 102.3703 | < 10 | [14 empty] | 1
| 1 | 2 | 3.5 | 39 | < 39 | < 39 | 94.4956 | < 10 | [14 empty] | 20
| 1 | 3 | 4 | 47 | < 39 | < 39 | 102.3703 | < 10 | [14 empty] | 16
| 1 | 4 | 3.7 | 55 | < 39 | < 39 | 78.7464 | < 10 | [14 empty] | 41.9596
| 1 | 5 | 3.7 | 63 | < 39 | < 39 | 86.621 | < 10 | [14 empty] | 26
| 1 | 6 | 4.1 | 55 | < 39 | < 39 | 94.4956 | < 10
| Mean |  | 3.7833 | 51 | 47 |  | 93.1832
| Median |  | 3.7 | 51 | 47 |  | 94.4956
| Std Dev. |  | 0.22286 | 8.3905 |  |  | 9.2058
| Rel.Std.Dev. |  | 5.891% | 16.45% |  |  | 9.879%
| PDM3 |  | -90.98% | 21.55% | 12.01% |  | 122.1%
| Z-Score (Absolute) |  | 0.7 | 1.12 | 0.09 | 0.09 | 2.73 | 0.65
| Analytical results for Yb in OREAS 753b (Indicative Value 0.35 ppm) | [21 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 1.12 | 1.28 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 2
2
| 1 | 1 | 0.4 | < 10 | [18 empty] | 1
| 1 | 2 | 0.3 | < 10 | [18 empty] | 21
| 1 | 3 | 0.4 | < 10 | [18 empty] | 16
| 1 | 4 | 0.4 | < 10 | [18 empty] | 0.35
| 1 | 5 | 0.3 | < 10 | [18 empty] | 27
| 1 | 6 | 0.3 | < 10
| Mean |  | 0.35
| Median |  | 0.35
| Std Dev. |  | 0.054772
| Rel.Std.Dev. |  | 15.65%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 753b (Certified Value 170 ppm) | [21 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | [15 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [15 empty] | 0
0
| 1 | 1 | 152 | 200 | 180 | 140 | 169 | [15 empty] | 1
| 1 | 2 | 161 | 200 | 170 | 150 | 177 | [15 empty] | 0
| 1 | 3 | 160 | 200 | 170 | 160 | 169 | [15 empty] | 16
| 1 | 4 | 156 | 190 | 170 | 150 | 177 | [15 empty] | 169.7826
| 1 | 5 | 160 | 200 | 160 | 170 | 161 | [15 empty] | 133
| 1 | 6 | 159 | 190 | 170 | 140 | 177
| Mean |  | 158 | 196.6667 | 170 | 151.6667 | 171.6667
| Median |  | 159.5 | 200 | 170 | 150 | 173
| Std Dev. |  | 3.4059 | 5.164 | 6.3246 | 11.6905 | 6.532
| Rel.Std.Dev. |  | 2.156% | 2.626% | 3.72% | 7.708% | 3.805%
| PDM3 |  | -6.94% | 15.83% | 0.1281% | -10.67% | 1.11%
| Z-Score (Absolute) |  | 0.67 | 1.5 | 0 | 1.03 | 0.08
| Analytical results for Zr in OREAS 753b (Indicative Value 63 ppm) | [21 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.07 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.26 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 70 | 42.8907 | 49 | < 100 | < 100 | 100 | 52 | [13 empty] | 1
| 1 | 2 | 70 | 45.162 | 46 | < 100 | < 100 | 100 | 59 | [13 empty] | 8
| 1 | 3 | 70 | 42.916 | 47 | < 100 | < 100 | 100 | 67 | [13 empty] | 16
| 1 | 4 | 70 | 41.3158 | 43 | < 100 | < 100 | 100 | 59 | [13 empty] | 63.1301
| 1 | 5 | 70 | 44.4546 | 46 | < 100 | < 100 | 100 | 52 | [13 empty] | 28
| 1 | 6 | 70 | 42.2024 | 48 | < 100 | 100 | 100 | 52
| Mean |  | 70 | 43.1569 | 46.5 |  | 100 | 100 | 56.8333
| Median |  | 70 | 42.9034 | 46.5 |  | 100 | 100 | 55.5
| Std Dev. |  | 0 | 1.4237 | 2.0736 |  |  | 0 | 6.047
| Rel.Std.Dev. |  | 0% | 3.299% | 4.459% |  |  | 0% | 10.64%
| PDM3 |  | 10.88% | -31.64% | -26.34% |  | 58.4% | 58.4% | -9.974%
| Z-Score (Absolute) |  | 1.14 | 0.89 | 0.67 | 0.44 | 0.1 | 2.84 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 753b (Certified Value 0.872 wt.%) | [22 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.28 | wt.%
|  |  | LOI*TGA | LOI*TGA | BF*XRF | BF*XRF | LOI*TGA | LOI*Furnace | BF*XRF | BF*XRF | BF*XRF | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*Furnace | BF*XRF | LOI*TGA | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | 3
3
| 1 | 1 | 0.76 | 0.83514 | 0.95 | 1.2701 | 0.73835 | 0.91 | 1.02 | 1.01 | 0.8994 | 0.788 | 0.74 | 0.91 | 0.77 | 0.85 | 0.91643 | 0.9 | 0.9 | 0.8 | 0.94012 | 0.74425 | 0.76 | 1
| 1 | 2 | 0.86 | 0.82373 | 0.91 | 1.2899 | 0.73192 | 0.87 | 1.02 | 0.95 | 0.8518 | 1.591 | 0.84 | 0.76 | 0.83 | 0.84 | 0.87667 | 0.9 | 0.93 | 0.9 | 1.03 | 0.77369 | 0.74 | 0
| 1 | 3 | 0.76 | 0.8414 | 0.94 | 1.1998 | 0.73665 | 0.93 | 1 | 1.01 | 0.9098 | 1.042 | 0.77 | 0.93 | 0.87 | 0.84 | 0.87843 | 0.89 | 0.91 | 0.8 | 0.89629 | 0.71649 | 0.73 | 16
| 1 | 4 | 0.86 | 0.82363 | 0.92 | 1.1594 | 0.74487 | 0.88 | 1.01 | 0.96 | 0.8975 | 0.913 | 0.79 | 1.23 | 0.9 | 0.83 | 0.83825 | 0.89 | 0.92 | 0.9 | 0.95337 | 0.82528 | 0.75 | 0.87179
| 1 | 5 | 0.85 | 0.80634 | 0.94 | 1.2194 | 0.74154 | 0.9 | 0.98 | 0.97 | 0.9036 | 0.942 | 0.89 | 0.46 | 0.78 | 0.85 | 0.87656 | 0.93 | 0.92 | 0.9 | 1.0215 | 0.87074 | 0.74 | 135
| 1 | 6 | 0.83 | 0.845 | 0.92 | 1.1598 | 0.70359 | 0.93 | 0.92 | 0.98 | 0.9101 | 1.03 | 0.73 | 0.91 | 0.85 | 0.86 | 0.85888 | 0.97 | 0.93 | 0.9 | 0.89867 | 0.73609 | 0.7
| Mean |  | 0.82 | 0.82921 | 0.93 | 1.2164 | 0.73282 | 0.90333 | 0.99167 | 0.98 | 0.89537 | 1.051 | 0.79333 | 0.86667 | 0.83333 | 0.845 | 0.8742 | 0.91333 | 0.91833 | 0.86667 | 0.95666 | 0.77775 | 0.73667
| Median |  | 0.84 | 0.82943 | 0.93 | 1.2096 | 0.7375 | 0.905 | 1.005 | 0.975 | 0.9015 | 0.986 | 0.78 | 0.91 | 0.84 | 0.845 | 0.87661 | 0.9 | 0.92 | 0.9 | 0.94674 | 0.75897 | 0.74
| Std Dev. |  | 0.047749 | 0.014266 | 0.015492 | 0.054817 | 0.014978 | 0.025033 | 0.038166 | 0.025298 | 0.021964 | 0.28008 | 0.061536 | 0.25161 | 0.050859 | 0.010488 | 0.025845 | 0.031411 | 0.01169 | 0.05164 | 0.058097 | 0.059216 | 0.020656
| Rel.Std.Dev. |  | 5.823% | 1.72% | 1.666% | 4.506% | 2.044% | 2.771% | 3.849% | 2.581% | 2.453% | 26.65% | 7.757% | 29.03% | 6.103% | 1.241% | 2.956% | 3.439% | 1.273% | 5.958% | 6.073% | 7.614% | 2.804%
| PDM3 |  | -5.94% | -4.884% | 6.677% | 39.53% | -15.94% | 3.619% | 13.75% | 12.41% | 2.705% | 20.56% | -8.999% | -0.5873% | -4.411% | -3.073% | 0.2771% | 4.766% | 5.339% | -0.5873% | 9.736% | -10.79% | -15.5%
| Z-Score (Absolute) |  | 0.67 | 0.56 | 0.69 | 4.26 | 1.76 | 0.36 | 1.46 | 1.32 | 0.26 | 2.2 | 1.01 | 0.09 | 0.51 | 0.36 | 0 | 0.49 | 0.55 | 0.09 | 1.03 | 1.2 | 1.71
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Cs in OREAS 753b (Indicative Value 170 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 1.28 | ppm
|  |  | AR*MS | 0
|  |  | 01g | 0
| 1 | 1 | 170.64 | 1
| 1 | 2 | 170.13 | 16
| 1 | 3 | 170.52 | 16
| 1 | 4 | 170.34 | 170.4217
| 1 | 5 | 170.47 | 31
| 1 | 6 | 170.43
| Mean |  | 170.4217
| Median |  | 170.45
| Std Dev. |  | 0.17406
| Rel.Std.Dev. |  | 0.1021%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.