## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 922b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 1.1725 | 1.0967 | 1.2482 | 1.1231 | 1.2219
Al, Aluminium (wt.%) | 6.3566 | 6.1978 | 6.5155 | 6.2291 | 6.4842
As, Arsenic (ppm) | 54.1355 | 51.9013 | 56.3697 | 52.1957 | 56.0753
Ba, Barium (wt.%) | 0.19098 | 0.18355 | 0.19842 | 0.18553 | 0.19644
Be, Beryllium (ppm) | 2.3037 | 2.2 | 2.4075 | 2.22 | 2.3874
Bi, Bismuth (ppm) | 12.2198 | 11.7539 | 12.6858 | 11.8841 | 12.5556
Ca, Calcium (wt.%) | 0.039966 | 0.038961 | 0.040972 | 0.03908 | 0.040853
Cd, Cadmium (ppm) | 0.078209 | 0.050784 | 0.10563 | IND | IND
Ce, Cerium (ppm) | 79.0115 | 75.3057 | 82.7174 | 76.4981 | 81.525
Co, Cobalt (ppm) | 3.7014 | 3.4665 | 3.9364 | 3.5489 | 3.854
Cr, Chromium (ppm) | 79.3998 | 75.2694 | 83.5302 | 76.5397 | 82.2599
Cs, Caesium (ppm) | 4.5919 | 4.3572 | 4.8266 | 4.4403 | 4.7436
Cu, Copper (wt.%) | 0.21938 | 0.21414 | 0.22462 | 0.21517 | 0.22359
Dy, Dysprosium (ppm) | 3.2818 | 3.0549 | 3.5087 | 3.1515 | 3.4121
Er, Erbium (ppm) | 1.7221 | 1.5641 | 1.88 | 1.5815 | 1.8627
Eu, Europium (ppm) | 1.2077 | 1.108 | 1.3075 | 1.1393 | 1.2762
Fe, Iron (wt.%) | 3.0965 | 3.0333 | 3.1597 | 3.025 | 3.168
Ga, Gallium (ppm) | 18.7734 | 17.9547 | 19.5921 | 18.1838 | 19.363
Gd, Gadolinium (ppm) | 4.8695 | 4.4307 | 5.3083 | 4.5269 | 5.2121
Ge, Germanium (ppm) | 0.14193 | 0.10124 | 0.18263 | IND | IND
Hf, Hafnium (ppm) | 2.5163 | 2.3897 | 2.643 | 2.3882 | 2.6445
Ho, Holmium (ppm) | 0.59956 | 0.51314 | 0.68598 | 0.58455 | 0.61457
In, Indium (ppm) | 0.24803 | 0.23529 | 0.26078 | 0.23451 | 0.26156
K, Potassium (wt.%) | 2.5897 | 2.5261 | 2.6533 | 2.5236 | 2.6558
La, Lanthanum (ppm) | 42.5453 | 40.7033 | 44.3874 | 41.4447 | 43.6459
Li, Lithium (ppm) | 23.854 | 23.0598 | 24.6482 | 23.1219 | 24.5861
Lu, Lutetium (ppm) | 0.27385 | 0.25433 | 0.29336 | 0.26097 | 0.28672
Mg, Magnesium (wt.%) | 0.40712 | 0.397 | 0.41725 | 0.39762 | 0.41663
Mn, Manganese (wt.%) | 0.011145 | 0.010913 | 0.011378 | 0.010878 | 0.011412
Mo, Molybdenum (ppm) | 5.011 | 4.7805 | 5.2415 | 4.8429 | 5.1791
Na, Sodium (wt.%) | 0.096132 | 0.087519 | 0.10475 | 0.09384 | 0.098425
Nb, Niobium (ppm) | 5.7452 | 5.0205 | 6.4699 | 5.3718 | 6.1186
Nd, Neodymium (ppm) | 34.9836 | 33.623 | 36.3442 | 34.1617 | 35.8055
Ni, Nickel (ppm) | 12.2469 | 11.3593 | 13.1345 | 11.6254 | 12.8684
P, Phosphorus (wt.%) | 0.033391 | 0.032249 | 0.034534 | 0.032458 | 0.034324
Pb, Lead (ppm) | 43.5078 | 41.0674 | 45.9482 | 41.6349 | 45.3807
Pr, Praseodymium (ppm) | 9.4587 | 9.1262 | 9.7911 | 9.0914 | 9.8259
Rb, Rubidium (ppm) | 152.2434 | 146.7837 | 157.7032 | 147.3392 | 157.1477
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.27391 | 0.26447 | 0.28336 | 0.26581 | 0.28202
Sb, Antimony (ppm) | 3.7549 | 3.4446 | 4.0651 | 3.6059 | 3.9039
Sc, Scandium (ppm) | 11.5431 | 10.756 | 12.3301 | 11.1363 | 11.9498
Se, Selenium (ppm) | 4.9004 | 4.297 | 5.5039 | 4.5513 | 5.2495
Sm, Samarium (ppm) | 6.4812 | 6.0605 | 6.9018 | 6.124 | 6.8383
Sn, Tin (ppm) | 5.7562 | 5.5471 | 5.9653 | 5.5554 | 5.957
Sr, Strontium (ppm) | 43.0725 | 41.5483 | 44.5967 | 42.0487 | 44.0963
Ta, Tantalum (ppm) | 0.45637 | 0.37895 | 0.53378 | 0.42 | 0.49273
Tb, Terbium (ppm) | 0.6339 | 0.55536 | 0.71244 | 0.58574 | 0.68206
Te, Tellurium (ppm) | 0.083654 | 0.063053 | 0.10426 | IND | IND
Th, Thorium (ppm) | 14.2597 | 13.4815 | 15.0379 | 13.8579 | 14.6616
Ti, Titanium (wt.%) | 0.22422 | 0.2113 | 0.23714 | 0.21735 | 0.23109
Tl, Thallium (ppm) | 0.94234 | 0.89877 | 0.98591 | 0.9125 | 0.97218
Tm, Thulium (ppm) | 0.2275 | 0.1789 | 0.27611 | IND | IND
U, Uranium (ppm) | 4.2872 | 4.1059 | 4.4686 | 4.1711 | 4.4033
V, Vanadium (ppm) | 183.6626 | 176.7547 | 190.5704 | 180.1409 | 187.1842
W, Tungsten (ppm) | 2.8047 | 2.5897 | 3.0197 | 2.6781 | 2.9313
Y, Yttrium (ppm) | 16.4057 | 15.7438 | 17.0675 | 15.7604 | 17.0509
Yb, Ytterbium (ppm) | 1.6735 | 1.5346 | 1.8123 | 1.5902 | 1.7568
Zn, Zinc (ppm) | 74.7683 | 71.9707 | 77.5659 | 71.8572 | 77.6794
Zr, Zirconium (ppm) | 85.0321 | 80.4205 | 89.6437 | 82.0809 | 87.9833
Aqua Regia Digestion
Ag, Silver (ppm) | 0.97259 | 0.92378 | 1.0214 | 0.93145 | 1.0137
Al, Aluminium (wt.%) | 0.79305 | 0.75778 | 0.82831 | 0.76989 | 0.81621
As, Arsenic (ppm) | 51.6941 | 49.8509 | 53.5373 | 50.2908 | 53.0973
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 187.6625 | 175.3809 | 199.9441 | 182.3425 | 192.9824
Be, Beryllium (ppm) | 0.46974 | 0.42486 | 0.51463 | 0.44882 | 0.49067
Bi, Bismuth (ppm) | 12.0429 | 11.5261 | 12.5597 | 11.7183 | 12.3674
Ca, Calcium (wt.%) | 0.036287 | 0.033578 | 0.038995 | 0.035046 | 0.037528
Cd, Cadmium (ppm) | 0.072215 | 0.061387 | 0.083043 | IND | IND
Ce, Cerium (ppm) | 60.1333 | 57.1061 | 63.1604 | 58.8459 | 61.4206
Co, Cobalt (ppm) | 3.4127 | 3.2538 | 3.5717 | 3.2966 | 3.5288
Cr, Chromium (ppm) | 19.7065 | 17.5624 | 21.8506 | 18.8189 | 20.5941
Cs, Caesium (ppm) | 1.1112 | 0.99714 | 1.2253 | 1.0747 | 1.1477
Cu, Copper (wt.%) | 0.22066 | 0.21447 | 0.22685 | 0.2163 | 0.22501
Fe, Iron (wt.%) | 2.8617 | 2.7757 | 2.9477 | 2.8102 | 2.9132
Ga, Gallium (ppm) | 2.4167 | 2.2079 | 2.6255 | 2.3313 | 2.5021
Ge, Germanium (ppm) | 0.10272 | 0.075267 | 0.13018 | IND | IND
Hf, Hafnium (ppm) | 0.31807 | 0.28402 | 0.35211 | 0.30168 | 0.33446
Hg, Mercury (ppm) | 0.022682 | 0.014355 | 0.031008 | IND | IND
In, Indium (ppm) | 0.2002 | 0.18906 | 0.21135 | 0.1927 | 0.20771
K, Potassium (wt.%) | 0.28104 | 0.26792 | 0.29417 | 0.272 | 0.29009
La, Lanthanum (ppm) | 28.5128 | 27.3977 | 29.6279 | 28.003 | 29.0226
Li, Lithium (ppm) | 3.8097 | 3.4245 | 4.1949 | 3.6208 | 3.9986
Lu, Lutetium (ppm) | 0.097565 | 0.085525 | 0.10961 | IND | IND
Mg, Magnesium (wt.%) | 0.12973 | 0.12233 | 0.13712 | 0.12563 | 0.13383
Mn, Manganese (wt.%) | 0.0087381 | 0.0084244 | 0.0090518 | 0.0085404 | 0.0089358
Mo, Molybdenum (ppm) | 4.7338 | 4.5397 | 4.9278 | 4.5828 | 4.8847
Na, Sodium (wt.%) | 0.009753 | 0.008707 | 0.010799 | IND | IND
Nb, Niobium (ppm) | 0.06881 | 0.054522 | 0.083098 | IND | IND
Ni, Nickel (ppm) | 9.3093 | 8.5749 | 10.0437 | 8.8575 | 9.7611
P, Phosphorus (wt.%) | 0.02879 | 0.027433 | 0.030146 | 0.02805 | 0.029529
Pb, Lead (ppm) | 29.6025 | 27.6658 | 31.5392 | 28.8027 | 30.4022
Rb, Rubidium (ppm) | 16.7091 | 15.6328 | 17.7854 | 16.0386 | 17.3796
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.27082 | 0.2629 | 0.27875 | 0.2636 | 0.27804
Sb, Antimony (ppm) | 2.2749 | 2.0223 | 2.5275 | 2.1811 | 2.3686
Sc, Scandium (ppm) | 2.0295 | 1.8716 | 2.1874 | 1.8971 | 2.1618
Se, Selenium (ppm) | 4.455 | 4.0227 | 4.8873 | 4.1888 | 4.7212
Sn, Tin (ppm) | 2.4202 | 2.2957 | 2.5448 | 2.3401 | 2.5003
Sr, Strontium (ppm) | 9.0363 | 8.1346 | 9.9381 | 8.6446 | 9.4281
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.41632 | 0.36151 | 0.47112 | 0.38468 | 0.44796
Te, Tellurium (ppm) | 0.058656 | 0.039988 | 0.077323 | IND | IND
Th, Thorium (ppm) | 9.9094 | 9.4749 | 10.344 | 9.6597 | 10.1592
Ti, Titanium (wt.%) | 0.011405 | 0.0098267 | 0.012984 | 0.010781 | 0.012029
Tl, Thallium (ppm) | 0.13549 | 0.11867 | 0.15231 | 0.12188 | 0.1491
U, Uranium (ppm) | 2.2736 | 2.1727 | 2.3745 | 2.1976 | 2.3496
V, Vanadium (ppm) | 22.2827 | 20.7898 | 23.7756 | 21.181 | 23.3844
W, Tungsten (ppm) | 0.74785 | 0.66528 | 0.83042 | 0.70415 | 0.79154
Y, Yttrium (ppm) | 7.2863 | 6.9213 | 7.6514 | 7.103 | 7.4696
Yb, Ytterbium (ppm) | 0.70815 | 0.60723 | 0.80906 | IND | IND
Zn, Zinc (ppm) | 65.8353 | 63.093 | 68.5776 | 64.049 | 67.6216
Zr, Zirconium (ppm) | 11.2207 | 10.4634 | 11.9781 | 10.8569 | 11.5846
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 12.3955 | 12.225 | 12.5659 | 12.2875 | 12.5034
BaO, Barium oxide (wt.%) | 0.22129 | 0.20648 | 0.23611 | 0.21342 | 0.22917
Bi, Bismuth (ppm) | < 100 | IND | IND | IND | IND
CaO, Calcium oxide (wt.%) | 0.054497 | 0.044817 | 0.064178 | IND | IND
Co, Cobalt (ppm) | < 100 | IND | IND | IND | IND
Cr2O3, Chromium(III) oxide (ppm) | 141.3358 | 129.0363 | 153.6354 | IND | IND
Cu, Copper (wt.%) | 0.21955 | 0.20931 | 0.22978 | 0.21389 | 0.2252
Fe2O3, Iron(III) oxide (wt.%) | 4.4937 | 4.4456 | 4.5418 | 4.4498 | 4.5376
Hf, Hafnium (ppm) | < 80 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 3.1918 | 3.1615 | 3.2221 | 3.1568 | 3.2269
MgO, Magnesium oxide (wt.%) | 0.70402 | 0.67365 | 0.7344 | 0.68099 | 0.72706
Mo, Molybdenum (ppm) | < 50 | IND | IND | IND | IND
Na2O, Sodium oxide (wt.%) | 0.12378 | 0.09672 | 0.15083 | IND | IND
Ni, Nickel (ppm) | < 50 | IND | IND | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.075809 | 0.067423 | 0.084196 | 0.072656 | 0.078963
Rb, Rubidium (ppm) | 124.6667 | 81.7071 | 167.6262 | IND | IND
Sb, Antimony (ppm) | < 50 | IND | IND | IND | IND
SiO2, Silicon dioxide (wt.%) | 74.4396 | 73.6121 | 75.2671 | 73.9483 | 74.9309
Sn, Tin (ppm) | < 50 | IND | IND | IND | IND
SO3, Sulphur trioxide (wt.%) | 0.66734 | 0.6355 | 0.69917 | 0.64949 | 0.68519
SrO, Strontium oxide (ppm) | 62.5794 | 40.4004 | 84.7584 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.52814 | 0.51332 | 0.54296 | 0.51127 | 0.54502
V, Vanadium (ppm) | 161.1741 | 125.4995 | 196.8487 | 156.427 | 165.9212
Zn, Zinc (ppm) | 74.5238 | 55.9952 | 93.0525 | IND | IND
Zr, Zirconium (ppm) | 111.7222 | 92.6257 | 130.8188 | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000 °C (wt.%) | 3.5917 | 3.4544 | 3.729 | 3.5122 | 3.6712
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.4632 | 6.3709 | 6.5556 | 6.3962 | 6.5302
As, Arsenic (ppm) | 54.8555 | 50.8532 | 58.8578 | 52.3426 | 57.3685
Ba, Barium (wt.%) | 0.19505 | 0.18846 | 0.20165 | 0.19235 | 0.19776
Be, Beryllium (ppm) | 2.5087 | 1.9724 | 3.045 | 2.2481 | 2.7693
Bi, Bismuth (ppm) | 12.8155 | 12.1131 | 13.5178 | 12.0994 | 13.5315
Ca, Calcium (wt.%) | 0.053466 | 0.038742 | 0.068191 | IND | IND
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (ppm) | 80.5077 | 77.8716 | 83.1437 | 77.4017 | 83.6136
Co, Cobalt (ppm) | 3.7023 | 3.3447 | 4.0599 | 3.4866 | 3.9181
Cr2O3, Chromium(III) oxide (ppm) | 130.3621 | 115.2881 | 145.4361 | 126.5092 | 134.215
Cs, Caesium (ppm) | 4.7758 | 4.4357 | 5.116 | 4.5818 | 4.9699
Cu, Copper (wt.%) | 0.2197 | 0.21305 | 0.22634 | 0.21661 | 0.22278
Dy, Dysprosium (ppm) | 4.788 | 4.5491 | 5.0269 | 4.5964 | 4.9795
Er, Erbium (ppm) | 2.7486 | 2.519 | 2.9781 | 2.5508 | 2.9463
Eu, Europium (ppm) | 1.2989 | 1.1975 | 1.4003 | 1.2107 | 1.387
Fe, Iron (wt.%) | 3.1447 | 3.0781 | 3.2112 | 3.1029 | 3.1864
Ga, Gallium (ppm) | 19.1347 | 18.0476 | 20.2218 | 18.0892 | 20.1802
Gd, Gadolinium (ppm) | 5.511 | 5.2911 | 5.7309 | 5.3216 | 5.7004
Ge, Germanium (ppm) | 2.0527 | 1.7987 | 2.3066 | IND | IND
Hf, Hafnium (ppm) | 3.6265 | 2.9841 | 4.269 | 3.439 | 3.8141
Ho, Holmium (ppm) | 0.96641 | 0.91967 | 1.0131 | 0.92419 | 1.0086
In, Indium (ppm) | 0.26713 | 0.14602 | 0.38824 | IND | IND
K, Potassium (wt.%) | 2.6236 | 2.5613 | 2.686 | 2.5884 | 2.6588
La, Lanthanum (ppm) | 43.025 | 41.4638 | 44.5862 | 41.8153 | 44.2347
Li, Lithium (ppm) | 25.0743 | 22.793 | 27.3557 | 23.2091 | 26.9395
Lu, Lutetium (ppm) | 0.41575 | 0.3586 | 0.4729 | 0.38861 | 0.44289
Mg, Magnesium (wt.%) | 0.41755 | 0.40616 | 0.42894 | 0.40853 | 0.42657
Mn, Manganese (wt.%) | 0.010488 | 0.0097418 | 0.011235 | 0.010275 | 0.010701
Mo, Molybdenum (ppm) | 5.1865 | 4.1368 | 6.2363 | IND | IND
Nb, Niobium (ppm) | 10.9675 | 9.8406 | 12.0943 | 10.4295 | 11.5055
Nd, Neodymium (ppm) | 36.0147 | 34.4816 | 37.5479 | 34.6057 | 37.4238
P, Phosphorus (wt.%) | 0.031513 | 0.02606 | 0.036965 | IND | IND
Pb, Lead (ppm) | 46.3563 | 40.26 | 52.4525 | 43.6816 | 49.0309
Pr, Praseodymium (ppm) | 9.8868 | 9.493 | 10.2805 | 9.575 | 10.1985
Rb, Rubidium (ppm) | 154.4527 | 150.6044 | 158.3011 | 150.419 | 158.4864
S, Sulphur (wt.%) | 0.26664 | 0.25278 | 0.2805 | 0.25338 | 0.2799
Sb, Antimony (ppm) | 4.0329 | 3.2413 | 4.8245 | 3.5351 | 4.5307
Sc, Scandium (ppm) | 10.8063 | 9.8965 | 11.716 | IND | IND
Si, Silicon (wt.%) | 35.2128 | 34.2494 | 36.1763 | 34.8882 | 35.5375
Sm, Samarium (ppm) | 6.838 | 6.4164 | 7.2595 | 6.5649 | 7.111
Sn, Tin (ppm) | 6.2725 | 4.9952 | 7.5498 | IND | IND
Sr, Strontium (ppm) | 47.6273 | 43.9098 | 51.3449 | 46.1586 | 49.096
Ta, Tantalum (ppm) | 0.96677 | 0.82227 | 1.1113 | IND | IND
Tb, Terbium (ppm) | 0.89995 | 0.72037 | 1.0795 | 0.84843 | 0.95147
Th, Thorium (ppm) | 13.9161 | 13.2012 | 14.6311 | 13.363 | 14.4693
Ti, Titanium (wt.%) | 0.30958 | 0.30083 | 0.31834 | 0.30097 | 0.3182
Tl, Thallium (ppm) | 0.9917 | 0.86953 | 1.1139 | IND | IND
Tm, Thulium (ppm) | 0.40111 | 0.37456 | 0.42766 | 0.36709 | 0.43513
U, Uranium (ppm) | 4.5355 | 4.2927 | 4.7782 | 4.3793 | 4.6916
V, Vanadium (ppm) | 192.6098 | 184.4299 | 200.7897 | 189.7148 | 195.5048
W, Tungsten (ppm) | 4.0707 | 2.8402 | 5.3011 | IND | IND
Y, Yttrium (ppm) | 26.5616 | 25.1611 | 27.9621 | 25.5807 | 27.5425
Yb, Ytterbium (ppm) | 2.7204 | 2.5574 | 2.8835 | 2.5361 | 2.9048
Zn, Zinc (ppm) | 75.5121 | 62.3041 | 88.72 | 73.5882 | 77.436
Zr, Zirconium (ppm) | 130.4863 | 121.011 | 139.9616 | 125.4331 | 135.5395
Infrared Combustion
C, Carbon (wt.%) | 0.76584 | 0.75022 | 0.78146 | 0.75095 | 0.78074
S, Sulphur (wt.%) | 0.26444 | 0.25305 | 0.27584 | 0.25487 | 0.27401
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 922b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | < 10 | Hg | ppm | 0.04
Aqua Regia Digestion
Dy | ppm | 1.996 | Ho | ppm | 0.29881 | Pt | ppb | 1.8
Er | ppm | 0.8186 | Nd | ppm | 25.6786 | Sm | ppm | 4.778
Eu | ppm | 0.88265 | Pd | ppb | 3.3333 | Tm | ppm | 0.09772
Gd | ppm | 4.0669 | Pr | ppm | 6.9059
Borate Fusion XRF
Ag | ppm | < 10 | Ge | ppm | < 10 | Se | ppm | < 10
As | ppm | 82.5 | Hg | ppm | < 100 | Sm | ppm | 28.5
Au | ppm | < 10 | Ho | ppm | < 10 | Ta | ppm | 14.6667
Cd | ppm | < 10 | In | ppm | < 100 | Tb | ppm | < 10
Ce | ppm | 78.5309 | La | ppm | 54.5608 | Te | ppm | < 100
Cl | ppm | 19.3333 | Lu | ppm | < 10 | Th | ppm | < 10
Cs | ppm | < 10 | MnO | wt.% | 0.012879 | Tl | ppm | < 10
Dy | ppm | < 10 | Nb | ppm | 48.8333 | Tm | ppm | < 10
Er | ppm | < 10 | Nd | ppm | 61 | U | ppm | < 10
Eu | ppm | < 10 | Pb | ppm | 43.2 | W | ppm | 10.125
Ga | ppm | 13 | Pr | ppm | < 10 | Y | ppm | 36.3043
Gd | ppm | < 10 | Sc | ppm | 15 | Yb | ppm | < 10
Borate / Peroxide Fusion ICP
Ag | ppm | 0.98 | Na | wt.% | 0.083505 | Se | ppm | < 10
B | ppm | 117.0333 | Ni | ppm | 20.9137 | Te | ppm | < 1
Hg | ppm | < 5 | Re | ppm | < 0.1
Laser Ablation ICP-MS
Ag | ppm | 1.45 | Hf | ppm | 3.775 | Sm | ppm | 6.98
As | ppm | 64.4 | Ho | ppm | 0.985 | Sn | ppm | 6.6
Ba | wt.% | 0.203 | In | ppm | 0.25 | Sr | ppm | 43.5
Be | ppm | 2.4 | La | ppm | 42.7 | Ta | ppm | 0.95
Bi | ppm | 13.85 | Lu | ppm | 0.4 | Tb | ppm | 0.805
Cd | ppm | < 0.1 | Mn | wt.% | 0.0116 | Te | ppm | < 0.2
Ce | ppm | 75.45 | Mo | ppm | 4.8 | Th | ppm | 14.05
Co | ppm | 4.2 | Nb | ppm | 11.6 | Ti | wt.% | 0.323
Cr | ppm | 84.5 | Nd | ppm | 36.4 | Tl | ppm | 0.4
Cs | ppm | 4.95 | Ni | ppm | 16 | Tm | ppm | 0.395
Cu | wt.% | 0.236 | Pb | ppm | 46.5 | U | ppm | 4.525
Dy | ppm | 4.73 | Pr | ppm | 9.89 | V | ppm | 196.5
Er | ppm | 2.875 | Rb | ppm | 157.5 | W | ppm | 3.5
Eu | ppm | 1.26 | Re | ppm | < 0.01 | Y | ppm | 26.85
Ga | ppm | 18.15 | Sb | ppm | 4.85 | Yb | ppm | 2.8
Gd | ppm | 5.625 | Sc | ppm | 12.1 | Zn | ppm | 70
Ge | ppm | 1.875 | Se | ppm | < 5 | Zr | ppm | 131
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 922b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 1.1725 | 0.073833 | 1.0248 | 1.3202 | 0.95099 | 1.394 | 6.297% | 12.59% | 18.89% | 1.1139 | 1.2311
Al, wt.% | 6.3566 | 0.21001 | 5.9366 | 6.7767 | 5.7266 | 6.9867 | 3.304% | 6.608% | 9.912% | 6.0388 | 6.6745
As, ppm | 54.1355 | 2.7786 | 48.5783 | 59.6927 | 45.7997 | 62.4714 | 5.133% | 10.27% | 15.4% | 51.4287 | 56.8423
Ba, wt.% | 0.19098 | 0.010278 | 0.17043 | 0.21154 | 0.16015 | 0.22182 | 5.382% | 10.76% | 16.15% | 0.18144 | 0.20053
Be, ppm | 2.3037 | 0.11667 | 2.0704 | 2.5371 | 1.9537 | 2.6538 | 5.065% | 10.13% | 15.19% | 2.1885 | 2.4189
Bi, ppm | 12.2198 | 0.58354 | 11.0528 | 13.3869 | 10.4692 | 13.9705 | 4.775% | 9.551% | 14.33% | 11.6089 | 12.8308
Ca, wt.% | 0.039966 | 0.00086388 | 0.038239 | 0.041694 | 0.037375 | 0.042558 | 2.162% | 4.323% | 6.485% | 0.037968 | 0.041965
Cd, ppm | 0.078209 | 0.020613 | 0.036983 | 0.11943 | 0.01637 | 0.14005 | 26.36% | 52.71% | 79.07% | 0.074298 | 0.082119
Ce, ppm | 79.0115 | 4.0219 | 70.9677 | 87.0553 | 66.9458 | 91.0772 | 5.09% | 10.18% | 15.27% | 75.0609 | 82.9621
Co, ppm | 3.7014 | 0.25408 | 3.1933 | 4.2096 | 2.9392 | 4.4637 | 6.864% | 13.73% | 20.59% | 3.5164 | 3.8865
Cr, ppm | 79.3998 | 7.7085 | 63.9828 | 94.8168 | 56.2743 | 102.5254 | 9.708% | 19.42% | 29.13% | 75.4298 | 83.3698
Cs, ppm | 4.5919 | 0.1817 | 4.2285 | 4.9553 | 4.0468 | 5.137 | 3.957% | 7.914% | 11.87% | 4.3623 | 4.8215
Cu, wt.% | 0.21938 | 0.0066907 | 0.206 | 0.23276 | 0.19931 | 0.23945 | 3.05% | 6.1% | 9.149% | 0.20841 | 0.23035
Dy, ppm | 3.2818 | 0.19455 | 2.8927 | 3.6709 | 2.6981 | 3.8654 | 5.928% | 11.86% | 17.78% | 3.1177 | 3.4459
Er, ppm | 1.7221 | 0.080354 | 1.5614 | 1.8828 | 1.481 | 1.9631 | 4.666% | 9.332% | 14% | 1.636 | 1.8082
Eu, ppm | 1.2077 | 0.071755 | 1.0642 | 1.3513 | 0.99248 | 1.423 | 5.941% | 11.88% | 17.82% | 1.1474 | 1.2681
Fe, wt.% | 3.0965 | 0.076316 | 2.9439 | 3.2491 | 2.8676 | 3.3255 | 2.465% | 4.929% | 7.394% | 2.9417 | 3.2513
Ga, ppm | 18.7734 | 0.95874 | 16.8559 | 20.6909 | 15.8972 | 21.6496 | 5.107% | 10.21% | 15.32% | 17.8347 | 19.7121
Gd, ppm | 4.8695 | 0.26931 | 4.3309 | 5.4081 | 4.0616 | 5.6774 | 5.531% | 11.06% | 16.59% | 4.626 | 5.113
Ge, ppm | 0.14193 | 0.021455 | 0.099025 | 0.18484 | 0.07757 | 0.2063 | 15.12% | 30.23% | 45.35% | 0.13484 | 0.14903
Hf, ppm | 2.5163 | 0.15375 | 2.2088 | 2.8238 | 2.0551 | 2.9776 | 6.11% | 12.22% | 18.33% | 2.3905 | 2.6421
Ho, ppm | 0.59956 | 0.059073 | 0.48142 | 0.71771 | 0.42234 | 0.77678 | 9.853% | 19.71% | 29.56% | 0.56959 | 0.62954
In, ppm | 0.24803 | 0.01329 | 0.22145 | 0.27461 | 0.20816 | 0.2879 | 5.358% | 10.72% | 16.07% | 0.23563 | 0.26043
K, wt.% | 2.5897 | 0.101 | 2.3877 | 2.7917 | 2.2867 | 2.8927 | 3.9% | 7.8% | 11.7% | 2.4602 | 2.7192
La, ppm | 42.5453 | 2.3462 | 37.8529 | 47.2377 | 35.5068 | 49.5839 | 5.515% | 11.03% | 16.54% | 40.4181 | 44.6726
Li, ppm | 23.854 | 0.82889 | 22.1962 | 25.5117 | 21.3673 | 26.3406 | 3.475% | 6.95% | 10.42% | 22.6613 | 25.0467
Lu, ppm | 0.27385 | 0.01045 | 0.25295 | 0.29475 | 0.2425 | 0.3052 | 3.816% | 7.632% | 11.45% | 0.26016 | 0.28754
Mg, wt.% | 0.40712 | 0.011666 | 0.38379 | 0.43045 | 0.37213 | 0.44212 | 2.865% | 5.731% | 8.596% | 0.38677 | 0.42748
Mn, wt.% | 0.011145 | 0.00028729 | 0.010571 | 0.01172 | 0.010283 | 0.012007 | 2.578% | 5.155% | 7.733% | 0.010588 | 0.011703
Mo, ppm | 5.011 | 0.27381 | 4.4634 | 5.5586 | 4.1896 | 5.8324 | 5.464% | 10.93% | 16.39% | 4.7605 | 5.2616
Na, wt.% | 0.096132 | 0.012579 | 0.070975 | 0.12129 | 0.058397 | 0.13387 | 13.08% | 26.17% | 39.25% | 0.091326 | 0.10094
Nb, ppm | 5.7452 | 1.0114 | 3.7224 | 7.7679 | 2.711 | 8.7793 | 17.6% | 35.21% | 52.81% | 5.4579 | 6.0324
Nd, ppm | 34.9836 | 0.75836 | 33.4669 | 36.5004 | 32.7085 | 37.2587 | 2.168% | 4.336% | 6.503% | 33.2344 | 36.7328
Ni, ppm | 12.2469 | 0.74038 | 10.7661 | 13.7277 | 10.0258 | 14.4681 | 6.045% | 12.09% | 18.14% | 11.6346 | 12.8593
P, wt.% | 0.033391 | 0.0018283 | 0.029735 | 0.037048 | 0.027907 | 0.038876 | 5.475% | 10.95% | 16.43% | 0.031722 | 0.035061
Pb, ppm | 43.5078 | 3.7643 | 35.9792 | 51.0363 | 32.215 | 54.8006 | 8.652% | 17.3% | 25.96% | 41.3324 | 45.6832
Pr, ppm | 9.4587 | 0.17808 | 9.1025 | 9.8148 | 8.9244 | 9.9929 | 1.883% | 3.766% | 5.648% | 8.9857 | 9.9316
Rb, ppm | 152.2434 | 9.2572 | 133.7291 | 170.7578 | 124.4719 | 180.015 | 6.081% | 12.16% | 18.24% | 144.6313 | 159.8556
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.27391 | 0.012938 | 0.24804 | 0.29979 | 0.2351 | 0.31273 | 4.723% | 9.447% | 14.17% | 0.26022 | 0.28761
Sb, ppm | 3.7549 | 0.2169 | 3.3211 | 4.1887 | 3.1042 | 4.4056 | 5.776% | 11.55% | 17.33% | 3.5671 | 3.9426
Sc, ppm | 11.5431 | 0.66111 | 10.2208 | 12.8653 | 9.5597 | 13.5264 | 5.727% | 11.45% | 17.18% | 10.9659 | 12.1202
Se, ppm | 4.9004 | 0.51424 | 3.8719 | 5.9289 | 3.3577 | 6.4431 | 10.49% | 20.99% | 31.48% | 4.6554 | 5.1454
Sm, ppm | 6.4812 | 0.21819 | 6.0448 | 6.9175 | 5.8266 | 7.1357 | 3.367% | 6.733% | 10.1% | 6.1571 | 6.8052
Sn, ppm | 5.7562 | 0.23691 | 5.2824 | 6.23 | 5.0455 | 6.4669 | 4.116% | 8.232% | 12.35% | 5.4684 | 6.044
Sr, ppm | 43.0725 | 1.6158 | 39.8409 | 46.3041 | 38.2251 | 47.9199 | 3.751% | 7.503% | 11.25% | 40.9189 | 45.2261
Ta, ppm | 0.45637 | 0.11436 | 0.22766 | 0.68508 | 0.1133 | 0.79943 | 25.06% | 50.12% | 75.17% | 0.43355 | 0.47918
Tb, ppm | 0.6339 | 0.083908 | 0.46608 | 0.80171 | 0.38217 | 0.88562 | 13.24% | 26.47% | 39.71% | 0.6022 | 0.66559
Te, ppm | 0.083654 | 0.018348 | 0.046958 | 0.12035 | 0.02861 | 0.1387 | 21.93% | 43.87% | 65.8% | 0.079471 | 0.087837
Th, ppm | 14.2597 | 0.7843 | 12.6911 | 15.8283 | 11.9068 | 16.6126 | 5.5% | 11% | 16.5% | 13.5467 | 14.9727
Ti, wt.% | 0.22422 | 0.020398 | 0.18342 | 0.26501 | 0.16302 | 0.28541 | 9.097% | 18.19% | 27.29% | 0.21301 | 0.23543
Tl, ppm | 0.94234 | 0.045584 | 0.85117 | 1.0335 | 0.80559 | 1.0791 | 4.837% | 9.675% | 14.51% | 0.89522 | 0.98946
Tm, ppm | 0.2275 | 0.034719 | 0.15807 | 0.29694 | 0.12335 | 0.33166 | 15.26% | 30.52% | 45.78% | 0.21613 | 0.23888
U, ppm | 4.2872 | 0.23463 | 3.818 | 4.7565 | 3.5833 | 4.9911 | 5.473% | 10.95% | 16.42% | 4.0729 | 4.5016
V, ppm | 183.6626 | 9.4583 | 164.746 | 202.5791 | 155.2877 | 212.0374 | 5.15% | 10.3% | 15.45% | 174.4794 | 192.8457
W, ppm | 2.8047 | 0.25931 | 2.2861 | 3.3233 | 2.0268 | 3.5826 | 9.246% | 18.49% | 27.74% | 2.6645 | 2.9449
Y, ppm | 16.4057 | 0.78125 | 14.8432 | 17.9682 | 14.0619 | 18.7494 | 4.762% | 9.524% | 14.29% | 15.5854 | 17.2259
Yb, ppm | 1.6735 | 0.10186 | 1.4698 | 1.8772 | 1.3679 | 1.9791 | 6.087% | 12.17% | 18.26% | 1.5898 | 1.7572
Zn, ppm | 74.7683 | 3.8281 | 67.112 | 82.4246 | 63.2839 | 86.2527 | 5.12% | 10.24% | 15.36% | 71.0299 | 78.5067
Zr, ppm | 85.0321 | 4.0144 | 77.0032 | 93.061 | 72.9888 | 97.0754 | 4.721% | 9.442% | 14.16% | 80.7805 | 89.2837
Aqua Regia Digestion
Ag, ppm | 0.97259 | 0.057605 | 0.85738 | 1.0878 | 0.79977 | 1.1454 | 5.923% | 11.85% | 17.77% | 0.92396 | 1.0212
Al, wt.% | 0.79305 | 0.059814 | 0.67342 | 0.91267 | 0.6136 | 0.97249 | 7.542% | 15.08% | 22.63% | 0.75339 | 0.8327
As, ppm | 51.6941 | 1.996 | 47.7021 | 55.6861 | 45.7061 | 57.6821 | 3.861% | 7.722% | 11.58% | 49.1094 | 54.2788
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 187.6625 | 17.7748 | 152.113 | 223.212 | 134.3382 | 240.9868 | 9.472% | 18.94% | 28.41% | 178.2794 | 197.0456
Be, ppm | 0.46974 | 0.042051 | 0.38564 | 0.55385 | 0.34359 | 0.5959 | 8.952% | 17.9% | 26.86% | 0.44626 | 0.49323
Bi, ppm | 12.0429 | 0.90943 | 10.224 | 13.8617 | 9.3146 | 14.7712 | 7.552% | 15.1% | 22.65% | 11.4407 | 12.645
Ca, wt.% | 0.036287 | 0.0058422 | 0.024603 | 0.047971 | 0.01876 | 0.053813 | 16.1% | 32.2% | 48.3% | 0.034473 | 0.038101
Cd, ppm | 0.072215 | 0.011322 | 0.049571 | 0.094859 | 0.038249 | 0.10618 | 15.68% | 31.36% | 47.03% | 0.068604 | 0.075826
Ce, ppm | 60.1333 | 3.3996 | 53.3341 | 66.9325 | 49.9345 | 70.3321 | 5.653% | 11.31% | 16.96% | 57.1266 | 63.1399
Co, ppm | 3.4127 | 0.19225 | 3.0282 | 3.7972 | 2.836 | 3.9895 | 5.633% | 11.27% | 16.9% | 3.2421 | 3.5834
Cr, ppm | 19.7065 | 1.5167 | 16.6731 | 22.7399 | 15.1564 | 24.2566 | 7.696% | 15.39% | 23.09% | 18.7212 | 20.6918
Cs, ppm | 1.1112 | 0.17314 | 0.76493 | 1.4575 | 0.59179 | 1.6306 | 15.58% | 31.16% | 46.74% | 1.0556 | 1.1668
Cu, wt.% | 0.22066 | 0.005272 | 0.21012 | 0.2312 | 0.20484 | 0.23648 | 2.389% | 4.778% | 7.168% | 0.20963 | 0.23169
Fe, wt.% | 2.8617 | 0.10314 | 2.6554 | 3.0679 | 2.5522 | 3.1711 | 3.604% | 7.208% | 10.81% | 2.7186 | 3.0047
Ga, ppm | 2.4167 | 0.30993 | 1.7969 | 3.0366 | 1.4869 | 3.3465 | 12.82% | 25.65% | 38.47% | 2.2959 | 2.5376
Ge, ppm | 0.10272 | 0.018325 | 0.066073 | 0.13937 | 0.047748 | 0.1577 | 17.84% | 35.68% | 53.52% | 0.097586 | 0.10786
Hf, ppm | 0.31807 | 0.043295 | 0.23148 | 0.40466 | 0.18818 | 0.44795 | 13.61% | 27.22% | 40.84% | 0.30216 | 0.33397
Hg, ppm | 0.022682 | 0.0060107 | 0.010661 | 0.034703 | 0.0046499 | 0.040714 | 26.5% | 53% | 79.5% | 0.021548 | 0.023816
In, ppm | 0.2002 | 0.0087419 | 0.18272 | 0.21769 | 0.17398 | 0.22643 | 4.366% | 8.733% | 13.1% | 0.19019 | 0.21021
K, wt.% | 0.28104 | 0.021394 | 0.23826 | 0.32383 | 0.21686 | 0.34523 | 7.612% | 15.22% | 22.84% | 0.26699 | 0.2951
La, ppm | 28.5128 | 1.305 | 25.9028 | 31.1228 | 24.5978 | 32.4279 | 4.577% | 9.154% | 13.73% | 27.0872 | 29.9384
Li, ppm | 3.8097 | 0.46321 | 2.8833 | 4.7361 | 2.4201 | 5.1993 | 12.16% | 24.32% | 36.48% | 3.6192 | 4.0002
Lu, ppm | 0.097565 | 0.0082978 | 0.08097 | 0.11416 | 0.072672 | 0.12246 | 8.505% | 17.01% | 25.51% | 0.092687 | 0.10244
Mg, wt.% | 0.12973 | 0.01012 | 0.10949 | 0.14997 | 0.099367 | 0.16009 | 7.801% | 15.6% | 23.4% | 0.12324 | 0.13621
Mn, wt.% | 0.0087381 | 0.00037311 | 0.0079919 | 0.0094843 | 0.0076188 | 0.0098574 | 4.27% | 8.54% | 12.81% | 0.0083012 | 0.009175
Mo, ppm | 4.7338 | 0.30515 | 4.1235 | 5.3441 | 3.8183 | 5.6492 | 6.446% | 12.89% | 19.34% | 4.4971 | 4.9705
Na, wt.% | 0.009753 | 0.0016609 | 0.0064312 | 0.013075 | 0.0047702 | 0.014736 | 17.03% | 34.06% | 51.09% | 0.0092654 | 0.010241
Nb, ppm | 0.06881 | 0.0095936 | 0.049622 | 0.087997 | 0.040029 | 0.09759 | 13.94% | 27.88% | 41.83% | 0.065369 | 0.07225
Ni, ppm | 9.3093 | 0.44425 | 8.4208 | 10.1978 | 7.9765 | 10.6421 | 4.772% | 9.544% | 14.32% | 8.8438 | 9.7748
P, wt.% | 0.02879 | 0.0017672 | 0.025255 | 0.032324 | 0.023488 | 0.034091 | 6.138% | 12.28% | 18.42% | 0.02735 | 0.030229
Pb, ppm | 29.6025 | 3.6995 | 22.2035 | 37.0014 | 18.504 | 40.7009 | 12.5% | 24.99% | 37.49% | 28.1223 | 31.0826
Rb, ppm | 16.7091 | 1.5557 | 13.5977 | 19.8206 | 12.0419 | 21.3763 | 9.311% | 18.62% | 27.93% | 15.8737 | 17.5446
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.27082 | 0.011418 | 0.24799 | 0.29366 | 0.23657 | 0.30507 | 4.216% | 8.432% | 12.65% | 0.25728 | 0.28436
Sb, ppm | 2.2749 | 0.49602 | 1.2828 | 3.2669 | 0.7868 | 3.7629 | 21.8% | 43.61% | 65.41% | 2.1611 | 2.3886
Sc, ppm | 2.0295 | 0.19896 | 1.6315 | 2.4274 | 1.4326 | 2.6263 | 9.804% | 19.61% | 29.41% | 1.928 | 2.1309
Se, ppm | 4.455 | 0.35441 | 3.7462 | 5.1638 | 3.3918 | 5.5182 | 7.955% | 15.91% | 23.87% | 4.2322 | 4.6777
Sn, ppm | 2.4202 | 0.16004 | 2.1001 | 2.7403 | 1.9401 | 2.9003 | 6.613% | 13.23% | 19.84% | 2.2992 | 2.5412
Sr, ppm | 9.0363 | 1.3683 | 6.2997 | 11.773 | 4.9313 | 13.1414 | 15.14% | 30.29% | 45.43% | 8.5845 | 9.4882
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.41632 | 0.039289 | 0.33774 | 0.4949 | 0.29845 | 0.53418 | 9.437% | 18.87% | 28.31% | 0.3955 | 0.43713
Te, ppm | 0.058656 | 0.01101 | 0.036636 | 0.080675 | 0.025626 | 0.091685 | 18.77% | 37.54% | 56.31% | 0.055723 | 0.061588
Th, ppm | 9.9094 | 0.66701 | 8.5754 | 11.2435 | 7.9084 | 11.9105 | 6.731% | 13.46% | 20.19% | 9.414 | 10.4049
Ti, wt.% | 0.011405 | 0.002278 | 0.0068495 | 0.015961 | 0.0045715 | 0.018239 | 19.97% | 39.95% | 59.92% | 0.010835 | 0.011976
Tl, ppm | 0.13549 | 0.018818 | 0.097855 | 0.17313 | 0.079037 | 0.19194 | 13.89% | 27.78% | 41.67% | 0.12872 | 0.14226
U, ppm | 2.2736 | 0.15396 | 1.9657 | 2.5815 | 1.8117 | 2.7354 | 6.772% | 13.54% | 20.31% | 2.1599 | 2.3872
V, ppm | 22.2827 | 2.2785 | 17.7257 | 26.8396 | 15.4473 | 29.1181 | 10.23% | 20.45% | 30.68% | 21.1685 | 23.3968
W, ppm | 0.74785 | 0.1068 | 0.53424 | 0.96145 | 0.42744 | 1.0683 | 14.28% | 28.56% | 42.84% | 0.71045 | 0.78524
Y, ppm | 7.2863 | 0.52719 | 6.232 | 8.3407 | 5.7048 | 8.8679 | 7.235% | 14.47% | 21.71% | 6.922 | 7.6506
Yb, ppm | 0.70815 | 0.06787 | 0.57241 | 0.84389 | 0.50453 | 0.91176 | 9.584% | 19.17% | 28.75% | 0.67274 | 0.74355
Zn, ppm | 65.8353 | 3.1916 | 59.452 | 72.2186 | 56.2604 | 75.4103 | 4.848% | 9.696% | 14.54% | 62.5436 | 69.1271
Zr, ppm | 11.2207 | 1.273 | 8.6748 | 13.7667 | 7.4018 | 15.0396 | 11.34% | 22.69% | 34.03% | 10.6597 | 11.7818
Borate Fusion XRF
Al2O3, wt.% | 12.3955 | 0.11526 | 12.1649 | 12.626 | 12.0497 | 12.7412 | 0.9299% | 1.86% | 2.79% | 11.7757 | 13.0152
BaO, wt.% | 0.22129 | 0.013418 | 0.19446 | 0.24813 | 0.18104 | 0.26155 | 6.063% | 12.13% | 18.19% | 0.21023 | 0.23236
Bi, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
CaO, wt.% | 0.054497 | 0.0066577 | 0.041182 | 0.067813 | 0.034524 | 0.074471 | 12.22% | 24.43% | 36.65% | 0.051772 | 0.057222
Co, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cr2O3, ppm | 141.3358 | 7.6242 | 126.0873 | 156.5843 | 118.4631 | 164.2085 | 5.394% | 10.79% | 16.18% | 134.269 | 148.4026
Cu, wt.% | 0.21955 | 0.0069096 | 0.20573 | 0.23336 | 0.19882 | 0.24027 | 3.147% | 6.294% | 9.442% | 0.20857 | 0.23052
Fe2O3, wt.% | 4.4937 | 0.060442 | 4.3728 | 4.6146 | 4.3124 | 4.675 | 1.345% | 2.69% | 4.035% | 4.269 | 4.7184
Hf, ppm | < 80 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 3.1918 | 0.038774 | 3.1143 | 3.2694 | 3.0755 | 3.3082 | 1.215% | 2.43% | 3.644% | 3.0322 | 3.3514
MgO, wt.% | 0.70402 | 0.039588 | 0.62485 | 0.7832 | 0.58526 | 0.82279 | 5.623% | 11.25% | 16.87% | 0.66882 | 0.73922
Mo, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Na2O, wt.% | 0.12378 | 0.029708 | 0.06436 | 0.18319 | 0.034651 | 0.2129 | 24% | 48% | 72.01% | 0.11759 | 0.12997
Ni, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
P2O5, wt.% | 0.075809 | 0.0063499 | 0.063109 | 0.088509 | 0.05676 | 0.094859 | 8.376% | 16.75% | 25.13% | 0.072019 | 0.0796
Rb, ppm | 124.6667 | 28.6587 | 67.3493 | 181.984 | 38.6906 | 210.6427 | 22.99% | 45.98% | 68.96% | 118.4333 | 130.9
Sb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SiO2, wt.% | 74.4396 | 0.53912 | 73.3614 | 75.5178 | 72.8223 | 76.057 | 0.7242% | 1.448% | 2.173% | 70.7176 | 78.1616
Sn, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SO3, wt.% | 0.66734 | 0.026393 | 0.61455 | 0.72012 | 0.58816 | 0.74652 | 3.955% | 7.91% | 11.87% | 0.63397 | 0.7007
SrO, ppm | 62.5794 | 23.3443 | 15.8908 | 109.268 | 0 | 132.6123 | 37.3% | 74.61% | 111.9% | 59.4504 | 65.7084
TiO2, wt.% | 0.52814 | 0.018032 | 0.49208 | 0.5642 | 0.47404 | 0.58224 | 3.414% | 6.829% | 10.24% | 0.50173 | 0.55455
V, ppm | 161.1741 | 46.3771 | 68.4198 | 253.9283 | 22.0427 | 300.3054 | 28.77% | 57.55% | 86.32% | 153.1154 | 169.2328
Zn, ppm | 74.5238 | 17.095 | 40.3339 | 108.7137 | 23.2389 | 125.8087 | 22.94% | 45.88% | 68.82% | 70.7976 | 78.25
Zr, ppm | 111.7222 | 18.0928 | 75.5366 | 147.9078 | 57.4438 | 166.0006 | 16.19% | 32.39% | 48.58% | 106.1361 | 117.3083
Thermogravimetry
LOI1000, wt.% | 3.5917 | 0.20291 | 3.1859 | 3.9975 | 2.983 | 4.2004 | 5.649% | 11.3% | 16.95% | 3.4121 | 3.7713
Borate / Peroxide Fusion ICP
Al, wt.% | 6.4632 | 0.14307 | 6.1771 | 6.7494 | 6.034 | 6.8924 | 2.214% | 4.427% | 6.641% | 6.1401 | 6.7864
As, ppm | 54.8555 | 3.9345 | 46.9866 | 62.7244 | 43.0522 | 66.6589 | 7.172% | 14.34% | 21.52% | 52.1127 | 57.5983
Ba, wt.% | 0.19505 | 0.0074728 | 0.18011 | 0.21 | 0.17263 | 0.21747 | 3.831% | 7.662% | 11.49% | 0.1853 | 0.20481
Be, ppm | 2.5087 | 0.42706 | 1.6546 | 3.3629 | 1.2276 | 3.7899 | 17.02% | 34.05% | 51.07% | 2.3833 | 2.6342
Bi, ppm | 12.8155 | 0.47319 | 11.8691 | 13.7618 | 11.3959 | 14.235 | 3.692% | 7.385% | 11.08% | 12.1747 | 13.4562
Ca, wt.% | 0.053466 | 0.01126 | 0.030946 | 0.075986 | 0.019686 | 0.087247 | 21.06% | 42.12% | 63.18% | 0.050793 | 0.056139
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 80.5077 | 2.4977 | 75.5123 | 85.5031 | 73.0146 | 88.0008 | 3.102% | 6.205% | 9.307% | 76.4823 | 84.5331
Co, ppm | 3.7023 | 0.30947 | 3.0834 | 4.3213 | 2.7739 | 4.6308 | 8.359% | 16.72% | 25.08% | 3.5172 | 3.8875
Cr2O3, ppm | 130.3621 | 12.6632 | 105.0357 | 155.6886 | 92.3725 | 168.3518 | 9.714% | 19.43% | 29.14% | 123.844 | 136.8802
Cs, ppm | 4.7758 | 0.40711 | 3.9616 | 5.5901 | 3.5545 | 5.9972 | 8.524% | 17.05% | 25.57% | 4.537 | 5.0146
Cu, wt.% | 0.2197 | 0.0071607 | 0.20537 | 0.23402 | 0.19821 | 0.24118 | 3.259% | 6.519% | 9.778% | 0.20871 | 0.23068
Dy, ppm | 4.788 | 0.20741 | 4.3732 | 5.2028 | 4.1657 | 5.4102 | 4.332% | 8.664% | 13% | 4.5486 | 5.0274
Er, ppm | 2.7486 | 0.14026 | 2.468 | 3.0291 | 2.3278 | 3.1693 | 5.103% | 10.21% | 15.31% | 2.6111 | 2.886
Eu, ppm | 1.2989 | 0.068018 | 1.1629 | 1.4349 | 1.0948 | 1.5029 | 5.237% | 10.47% | 15.71% | 1.2339 | 1.3638
Fe, wt.% | 3.1447 | 0.067632 | 3.0094 | 3.2799 | 2.9418 | 3.3476 | 2.151% | 4.301% | 6.452% | 2.9874 | 3.3019
Ga, ppm | 19.1347 | 1.3633 | 16.4081 | 21.8613 | 15.0448 | 23.2246 | 7.125% | 14.25% | 21.37% | 18.178 | 20.0914
Gd, ppm | 5.511 | 0.27584 | 4.9593 | 6.0627 | 4.6835 | 6.3385 | 5.005% | 10.01% | 15.02% | 5.2354 | 5.7865
Ge, ppm | 2.0527 | 0.24107 | 1.5705 | 2.5348 | 1.3294 | 2.7759 | 11.74% | 23.49% | 35.23% | 1.95 | 2.1553
Hf, ppm | 3.6265 | 0.46952 | 2.6875 | 4.5656 | 2.218 | 5.0351 | 12.95% | 25.89% | 38.84% | 3.4452 | 3.8079
Ho, ppm | 0.96641 | 0.034167 | 0.89807 | 1.0347 | 0.86391 | 1.0689 | 3.535% | 7.071% | 10.61% | 0.91809 | 1.0147
In, ppm | 0.26713 | 0.061262 | 0.14461 | 0.38965 | 0.083346 | 0.45092 | 22.93% | 45.87% | 68.8% | 0.25377 | 0.28049
K, wt.% | 2.6236 | 0.097824 | 2.428 | 2.8193 | 2.3301 | 2.9171 | 3.729% | 7.457% | 11.19% | 2.4924 | 2.7548
La, ppm | 43.025 | 1.3002 | 40.4247 | 45.6253 | 39.1245 | 46.9255 | 3.022% | 6.044% | 9.066% | 40.8738 | 45.1763
Li, ppm | 25.0743 | 2.0352 | 21.0039 | 29.1447 | 18.9687 | 31.18 | 8.117% | 16.23% | 24.35% | 23.8206 | 26.328
Lu, ppm | 0.41575 | 0.04161 | 0.33253 | 0.49897 | 0.29092 | 0.54058 | 10.01% | 20.02% | 30.03% | 0.39496 | 0.43654
Mg, wt.% | 0.41755 | 0.01205 | 0.39345 | 0.44165 | 0.3814 | 0.4537 | 2.886% | 5.772% | 8.658% | 0.39667 | 0.43843
Mn, wt.% | 0.010488 | 0.0011228 | 0.0082428 | 0.012734 | 0.00712 | 0.013857 | 10.7% | 21.41% | 32.11% | 0.0099639 | 0.011013
Mo, ppm | 5.1865 | 0.55363 | 4.0793 | 6.2938 | 3.5257 | 6.8474 | 10.67% | 21.35% | 32.02% | 4.9272 | 5.4459
Nb, ppm | 10.9675 | 0.95122 | 9.0651 | 12.8699 | 8.1138 | 13.8211 | 8.673% | 17.35% | 26.02% | 10.4191 | 11.5159
Nd, ppm | 36.0147 | 1.2508 | 33.5131 | 38.5164 | 32.2623 | 39.7672 | 3.473% | 6.946% | 10.42% | 34.214 | 37.8155
P, wt.% | 0.031513 | 0.0037128 | 0.024087 | 0.038938 | 0.020375 | 0.042651 | 11.78% | 23.56% | 35.35% | 0.029937 | 0.033088
Pb, ppm | 46.3563 | 6.3232 | 33.7098 | 59.0027 | 27.3865 | 65.326 | 13.64% | 27.28% | 40.92% | 44.0384 | 48.6741
Pr, ppm | 9.8868 | 0.29276 | 9.3012 | 10.4723 | 9.0085 | 10.765 | 2.961% | 5.922% | 8.883% | 9.3924 | 10.3811
Rb, ppm | 154.4527 | 2.9258 | 148.6011 | 160.3044 | 145.6752 | 163.2302 | 1.894% | 3.789% | 5.683% | 146.7301 | 162.1754
S, wt.% | 0.26664 | 0.015809 | 0.23502 | 0.29826 | 0.21921 | 0.31407 | 5.929% | 11.86% | 17.79% | 0.25331 | 0.27997
Sb, ppm | 4.0329 | 0.54886 | 2.9352 | 5.1306 | 2.3863 | 5.6795 | 13.61% | 27.22% | 40.83% | 3.8312 | 4.2345
Sc, ppm | 10.8063 | 0.6466 | 9.513 | 12.0995 | 8.8664 | 12.7461 | 5.984% | 11.97% | 17.95% | 10.2659 | 11.3466
Si, wt.% | 35.2128 | 1.216 | 32.7809 | 37.6447 | 31.565 | 38.8607 | 3.453% | 6.906% | 10.36% | 33.4522 | 36.9735
Sm, ppm | 6.838 | 0.33993 | 6.1581 | 7.5178 | 5.8182 | 7.8577 | 4.971% | 9.942% | 14.91% | 6.4961 | 7.1799
Sn, ppm | 6.2725 | 0.6181 | 5.0363 | 7.5087 | 4.4182 | 8.1268 | 9.854% | 19.71% | 29.56% | 5.9589 | 6.5861
Sr, ppm | 47.6273 | 4.9304 | 37.7665 | 57.4882 | 32.836 | 62.4186 | 10.35% | 20.7% | 31.06% | 45.246 | 50.0087
Ta, ppm | 0.96677 | 0.15679 | 0.65319 | 1.2803 | 0.49641 | 1.4371 | 16.22% | 32.44% | 48.65% | 0.91843 | 1.0151
Tb, ppm | 0.89995 | 0.12983 | 0.6403 | 1.1596 | 0.51047 | 1.2894 | 14.43% | 28.85% | 43.28% | 0.85496 | 0.94495
Th, ppm | 13.9161 | 0.76186 | 12.3924 | 15.4398 | 11.6306 | 16.2017 | 5.475% | 10.95% | 16.42% | 13.2203 | 14.6119
Ti, wt.% | 0.30958 | 0.0085148 | 0.29255 | 0.32661 | 0.28404 | 0.33513 | 2.75% | 5.501% | 8.251% | 0.2941 | 0.32506
Tl, ppm | 0.9917 | 0.07909 | 0.83352 | 1.1499 | 0.75443 | 1.229 | 7.975% | 15.95% | 23.93% | 0.94211 | 1.0413
Tm, ppm | 0.40111 | 0.014095 | 0.37292 | 0.4293 | 0.35883 | 0.44339 | 3.514% | 7.028% | 10.54% | 0.38106 | 0.42117
U, ppm | 4.5355 | 0.21599 | 4.1035 | 4.9674 | 3.8875 | 5.1834 | 4.762% | 9.525% | 14.29% | 4.3087 | 4.7622
V, ppm | 192.6098 | 9.314 | 173.9818 | 211.2378 | 164.6677 | 220.5518 | 4.836% | 9.671% | 14.51% | 182.9793 | 202.2403
W, ppm | 4.0707 | 1.1553 | 1.76 | 6.3813 | 0.6047 | 7.5366 | 28.38% | 56.76% | 85.14% | 3.8671 | 4.2742
Y, ppm | 26.5616 | 1.4912 | 23.5792 | 29.544 | 22.088 | 31.0353 | 5.614% | 11.23% | 16.84% | 25.2335 | 27.8897
Yb, ppm | 2.7204 | 0.17461 | 2.3712 | 3.0697 | 2.1966 | 3.2443 | 6.419% | 12.84% | 19.26% | 2.5844 | 2.8565
Zn, ppm | 75.5121 | 8.2305 | 59.0511 | 91.973 | 50.8207 | 100.2035 | 10.9% | 21.8% | 32.7% | 71.7365 | 79.2877
Zr, ppm | 130.4863 | 7.0629 | 116.3605 | 144.612 | 109.2976 | 151.6749 | 5.413% | 10.83% | 16.24% | 123.962 | 137.0106
Infrared Combustion
C, wt.% | 0.76584 | 0.018119 | 0.7296 | 0.80208 | 0.71148 | 0.8202 | 2.366% | 4.732% | 7.098% | 0.72755 | 0.80413
S, wt.% | 0.26444 | 0.014071 | 0.2363 | 0.29259 | 0.22223 | 0.30666 | 5.321% | 10.64% | 15.96% | 0.25122 | 0.27767
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 922b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PSF*XRF | pyro-sulfate fusion with X-ray fluorescence spectroscopy
Text Values:
NA | Not Applicable (Lab 2.03)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 922b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Genalysis, Adelaide, SA, Australia
MSALABS, Vancouver, BC, Canada
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 922b (Certified Value 1.17 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 1.1 | 1.23 | 1.1 | 1.28 | 1.1 | 1.1299 | 1.17 | 1.12 | 1.16 | 1.11 | 1.22 | 1.1 | 1.18 | 1.0798 | 1.17 | 1.18 | 1.24 | 1 | 1.2 | 1.1 | 1.01 | 1.3 | 1.26 | 1.21 | 1.1 | 1
| 1 | 2 | 1.1 | 1.14 | 1.1 | 1.29 | 1.07 | 1.0838 | 1.24 | 1.14 | 1.16 | 1.14 | 1.12 | 1.2 | 1.19 | 1.0944 | 1.19 | 1.15 | 1.197 | 0.9 | 1.2 | 1.1 | 1.02 | 1.3 | 1.26 | 1.22 | 1.5 | 8
| 1 | 3 | 1.1 | 1.19 | 1.1 | 1.27 | 1.06 | 1.2357 | 1.16 | 1.11 | 1.14 | 1.14 | 1.22 | 1.2 | 1.2 | 1.1221 | 1.21 | 1.18 | 1.234 | 1 | 1.2 | 1.1 | 1 | 1.2 | 1.26 | 1.23 | 1.1 | 16
| 1 | 4 | 1.1 | 1.08 | 1.1 | 1.28 | 1.1 | 1.1751 | 1.16 | 1.11 | 1.2 | 1.14 | 1.2 | 1.2 | 1.17 | 1.1378 | 0.99 | 1.22 | 1.285 | 1 | 1.2 | 1.2 | 1.1 | 1.2 | 1.25 | 1.2 | 1.1 | 1.1725
| 1 | 5 | 1.1 | 1.24 | 1.1 | 1.26 | 1.15 | 1.1929 | 1.22 | 1.1 | 1.14 | 1.07 | 1.15 | 1.2 | 1.21 | 1.112 | 0.81 | 1.17 | 1.332 | 1 | 1.2 | 1.1 | 1.1 | 1.2 | 1.24 | 1.23 | 1.1 | 7
| 1 | 6 | 1.1 | 1.34 | 1.1 | 1.31 | 1.08 | 1.1679 | 1.2 | 1.13 | 1.13 | 1.14 | 1.08 | 1.2 | 1.2 | 1.1449 | 1.19 | 1.21 | 1.3 | 0.9 | 1.1 | 1.1 | 1.04 | 1.3 | 1.28 | 1.27 | 1.2
| Mean |  | 1.1 | 1.2033 | 1.1 | 1.2817 | 1.0933 | 1.1642 | 1.1917 | 1.1183 | 1.155 | 1.1233 | 1.165 | 1.1833 | 1.1917 | 1.1152 | 1.0933 | 1.185 | 1.2647 | 0.96667 | 1.1833 | 1.1167 | 1.045 | 1.25 | 1.2583 | 1.2267 | 1.1833
| Median |  | 1.1 | 1.21 | 1.1 | 1.28 | 1.09 | 1.1715 | 1.185 | 1.115 | 1.15 | 1.14 | 1.175 | 1.2 | 1.195 | 1.1171 | 1.18 | 1.18 | 1.2625 | 1 | 1.2 | 1.1 | 1.03 | 1.25 | 1.26 | 1.225 | 1.1
| Std Dev. |  | 2.4324e-16 | 0.089592 | 2.4324e-16 | 0.017224 | 0.032042 | 0.052363 | 0.033714 | 0.01472 | 0.0251 | 0.028752 | 0.057879 | 0.040825 | 0.01472 | 0.025033 | 0.16071 | 0.025884 | 0.049645 | 0.05164 | 0.040825 | 0.040825 | 0.044609 | 0.054772 | 0.013292 | 0.024221 | 0.16021
| Rel.Std.Dev. |  | 2.211e-14% | 7.445% | 2.211e-14% | 1.344% | 2.931% | 4.498% | 2.829% | 1.316% | 2.173% | 2.56% | 4.968% | 3.45% | 1.235% | 2.245% | 14.7% | 2.184% | 3.926% | 5.342% | 3.45% | 3.656% | 4.269% | 4.382% | 1.056% | 1.975% | 13.54%
| PDM3 |  | -6.183% | 2.63% | -6.183% | 9.311% | -6.751% | -0.7055% | 1.635% | -4.619% | -1.492% | -4.193% | -0.6391% | 0.9245% | 1.635% | -4.889% | -6.751% | 1.067% | 7.861% | -17.55% | 0.9245% | -4.761% | -10.87% | 6.61% | 7.321% | 4.62% | 0.9245%
| Z-Score (Absolute) |  | NA | 0.55 | NA | 1.69 | 1.04 | 0.01 | 0.39 | 0.67 | 0.14 | 0.6 | 0 | NA | 0.39 | 0.72 | 1.04 | 0.29 | 1.44 | NA | NA | NA | 1.73 | NA | 1.35 | 0.89 | NA
Results from laboratories 2.01, 2.03, 2.12, 2.18, 2.19, 2.20, 2.22 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 922b (Certified Value 6.36 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 6.29 | 6.2117 | 6.36 | 5.7 | 6.2949 | 6.4044 | 6.7 | 6.36 | 6.59 | 6.47 | 6.29 | 6.23 | 6.2 | 5.8044 | 6.57 | 6.5564 | 6.39 | 6.1279 | 6.6 | 6.39 | 6.02 | 6.24 | 6.5432 | 6.42 | 6.68 | 1
| 1 | 2 | 6.27 | 6.2164 | 6.35 | 5.77 | 6.0346 | 6.3568 | 6.69 | 6.34 | 6.59 | 6.55 | 6.48 | 6.34 | 6.38 | 5.8598 | 6.58 | 6.4943 | 5.64 | 6.1492 | 6.62 | 6.16 | 6.02 | 6.45 | 6.5065 | 6.21 | 6.75 | 0
| 1 | 3 | 6.2 | 6.2106 | 6.38 | 5.82 | 6.21 | 6.4739 | 6.44 | 6.49 | 6.42 | 6.52 | 6.28 | 6.32 | 6.2 | 6.4457 | 6.67 | 6.441 | 6.01 | 5.9747 | 6.49 | 6.38 | 6.04 | 6.3 | 6.5243 | 6.49 | 6.76 | 16
| 1 | 4 | 6.33 | 6.2164 | 6.3 | 5.73 | 6.1516 | 6.277 | 6.71 | 6.36 | 6.65 | 6.61 | 6.46 | 6.3 | 6.29 | 6.1518 | 5.32 | 6.4384 | 5.75 | 6.0888 | 6.65 | 6.62 | 6.06 | 6.37 | 6.3943 | 6.33 | 6.74 | 6.3566
| 1 | 5 | 6.27 | 6.2126 | 6.38 | 5.83 | 6.1299 | 6.3813 | 6.74 | 6.44 | 6.62 | 6.49 | 6.36 | 6.24 | 6.43 | 5.9871 | 4.46 | 6.4369 | 6.19 | 5.9547 | 6.52 | 6.31 | 6.03 | 6.2 | 6.4717 | 6.36 | 6.74 | 8
| 1 | 6 | 6.35 | 6.2096 | 6.4 | 5.77 | 6.0018 | 6.3895 | 6.55 | 6.5 | 6.53 | 6.56 | 6.38 | 6.18 | 6.42 | 6.4637 | 6.36 | 6.4532 | 5.86 | 6.0754 | 6.45 | 6.36 | 6.04 | 6.28 | 6.5374 | 6.3 | 6.62
| Mean |  | 6.285 | 6.2129 | 6.3617 | 5.77 | 6.1371 | 6.3805 | 6.6383 | 6.415 | 6.5667 | 6.5333 | 6.375 | 6.2683 | 6.32 | 6.1187 | 5.9933 | 6.47 | 5.9733 | 6.0618 | 6.555 | 6.37 | 6.035 | 6.3067 | 6.4962 | 6.3517 | 6.715
| Median |  | 6.28 | 6.2122 | 6.37 | 5.77 | 6.1407 | 6.3854 | 6.695 | 6.4 | 6.59 | 6.535 | 6.37 | 6.27 | 6.335 | 6.0694 | 6.465 | 6.4471 | 5.935 | 6.0821 | 6.56 | 6.37 | 6.035 | 6.29 | 6.5154 | 6.345 | 6.74
| Std Dev. |  | 0.05282 | 0.0029055 | 0.034881 | 0.0502 | 0.10889 | 0.064274 | 0.11754 | 0.070922 | 0.082138 | 0.050859 | 0.083367 | 0.061455 | 0.10526 | 0.28642 | 0.90261 | 0.047482 | 0.28133 | 0.079973 | 0.079687 | 0.14886 | 0.015166 | 0.090701 | 0.056189 | 0.09704 | 0.054314
| Rel.Std.Dev. |  | 0.8404% | 0.04676% | 0.5483% | 0.87% | 1.774% | 1.007% | 1.771% | 1.106% | 1.251% | 0.7785% | 1.308% | 0.9804% | 1.666% | 4.681% | 15.06% | 0.7339% | 4.71% | 1.319% | 1.216% | 2.337% | 0.2513% | 1.438% | 0.8649% | 1.528% | 0.8088%
| PDM3 |  | -1.127% | -2.261% | 0.07936% | -9.229% | -3.453% | 0.3753% | 4.432% | 0.9184% | 3.304% | 2.78% | 0.2891% | -1.389% | -0.5761% | -3.743% | -5.715% | 1.784% | -6.03% | -4.638% | 3.121% | 0.2105% | -5.06% | -0.7859% | 2.196% | -0.07796% | 5.638%
| Z-Score (Absolute) |  | 0.31 | 0.65 | 0.05 | 2.71 | 1 | 0.13 | 1.34 | 0.3 | 1 | 0.85 | 0.11 | 0.39 | 0.15 | 1.09 | 1.67 | 0.55 | 1.76 | 1.35 | 0.95 | 0.09 | 1.48 | 0.21 | 0.67 | 0 | 1.69
| Analytical results for As in OREAS 922b (Certified Value 54 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 53 | 26.08 | 53.6 | 53.1188 | 55.1 | 49.8 | 54.3 | 53.5 | 58.5 | 51.7 | 56 | 48.5702 | 34.8 | 54.9 | 54.4 | 53.4 | 58 | 57 | 52 | 56.8 | 57.7 | 48.5 | 57.5 |  |  | 1
| 1 | 2 | 54 | 30.68 | 51.6 | 53.6165 | 55.9 | 51.2 | 55 | 58 | 55.1 | 52.6 | 57 | 48.1041 | 51.6 | 56.7 | 50.6 | 54.1 | 57 | 58 | 53 | 56.5 | 58 | 49.3 | 59.3 |  |  | 9
| 1 | 3 | 52 | 29.17 | 53.5 | 52.1127 | 48.3 | 53.1 | 52.6 | 54.9 | 58.7 | 51 | 54 | 53.5551 | 41.1 | 56.1 | 54 | 54.4 | 57 | 56 | 53 | 54.8 | 57.6 | 50.6 | 57.9 |  |  | 16
| 1 | 4 | 54 | 30.73 | 53.2 | 53.7362 | 54.9 | 48.9 | 55.8 | 55.2 | 58.3 | 52.2 | 56 | 50.9548 | 45 | 57.3 | 51.1 | 54.5 | 58 | 58 | 53 | 52.8 | 55.9 | 48 | 58.4 |  |  | 54.1355
| 1 | 5 | 52 | 28.25 | 53.4 | 53.5439 | 51.7 | 51.4 | 54.1 | 53.7 | 59.2 | 55.4 | 56 | 51.1311 | 41.9 | 57.4 | 53.9 | 53.8 | 57 | 59 | 53 | 55.2 | 56.7 | 47.7 | 57.8 |  |  | 9
| 1 | 6 | 54 | 31.08 | 51.3 | 54.2643 | 48.9 | 52.6 | 52.8 | 53.6 | 54.4 | 52.2 | 54 | 52.3458 | 43.5 | 57.4 | 51.2 | 53 | 56 | 55 | 53 | 55.8 | 58.3 | 50.4 | 57.4
| Mean |  | 53.1667 | 29.3317 | 52.7667 | 53.3988 | 52.4667 | 51.1667 | 54.1 | 54.8167 | 57.3667 | 52.5167 | 55.5 | 50.7769 | 42.9833 | 56.6333 | 52.5333 | 53.8667 | 57.1667 | 57.1667 | 52.8333 | 55.3167 | 57.3667 | 49.0833 | 58.05
| Median |  | 53.5 | 29.925 | 53.3 | 53.5802 | 53.3 | 51.3 | 54.2 | 54.3 | 58.4 | 52.2 | 56 | 51.0429 | 42.7 | 57 | 52.55 | 53.95 | 57 | 57.5 | 53 | 55.5 | 57.65 | 48.9 | 57.85
| Std Dev. |  | 0.98319 | 1.9287 | 1.0328 | 0.7299 | 3.3261 | 1.6033 | 1.2394 | 1.7175 | 2.0607 | 1.5158 | 1.2247 | 2.1142 | 5.4821 | 0.9913 | 1.7363 | 0.58538 | 0.75277 | 1.472 | 0.40825 | 1.4456 | 0.89815 | 1.2254 | 0.7064
| Rel.Std.Dev. |  | 1.849% | 6.575% | 1.957% | 1.367% | 6.339% | 3.134% | 2.291% | 3.133% | 3.592% | 2.886% | 2.207% | 4.164% | 12.75% | 1.75% | 3.305% | 1.087% | 1.317% | 2.575% | 0.7727% | 2.613% | 1.566% | 2.497% | 1.217%
| PDM3 |  | -1.79% | -45.82% | -2.529% | -1.361% | -3.083% | -5.484% | -0.06559% | 1.258% | 5.969% | -2.99% | 2.521% | -6.204% | -20.6% | 4.614% | -2.96% | -0.4966% | 5.599% | 5.599% | -2.405% | 2.182% | 5.969% | -9.332% | 7.231%
| Z-Score (Absolute) |  | 0.07 | 7.72 | 0.2 | 0 | 0.3 | 0.72 | 0.23 | 0.46 | 1.27 | 0.28 | 0.67 | 0.84 | 3.34 | 1.04 | 0.28 | 0.15 | 1.21 | 1.21 | 0.18 | 0.62 | 1.27 | 1.38 | 1.49
| Analytical results for B in OREAS 922b (Indicative Value < 10 ppm) | [26 empty] | Indicative
B | Round | Replicate | 2.01: Lab |  |  | 1
| No | No | 2.01: 2.20 |  |  | ppm
|  |  | 2.01: 4A*MS |  |  | 1
1
| 1 | 1 | 2.01: < 10 |  |  | 1
| 1 | 2 | 2.01: < 10 |  |  | 1
| 1 | 3 | 2.01: < 10 |  |  | 16
| 1 | 4 | 2.01: < 10 |  |  | < 10
| 1 | 5 | 2.01: < 10 |  |  | 7
| 1 | 6 | 2.01: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Ba in OREAS 922b (Certified Value 0.191 wt.%) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.192 | 0.1932 | 0.18568 | 0.19636 | 0.19537 | 0.206 | 0.196 | 0.2 | 0.198 | 0.183 | 0.145 | 0.1781 | 0.17727 | 0.195 | 0.19672 | 0.1321 | 0.1551 | 0.2062 | 0.1901 | 0.1847 | > 0.1 |  |  |  |  | 1
| 1 | 2 | 0.197 | 0.191 | 0.18715 | 0.18964 | 0.19863 | 0.206 | 0.194 | 0.202 | 0.199 | 0.188 | 0.154 | 0.1826 | 0.17872 | 0.195 | 0.19839 | 0.1126 | 0.1719 | 0.2027 | 0.1906 | 0.1792 | > 0.1 |  |  |  |  | 27
| 1 | 3 | 0.192 | 0.192 | 0.17843 | 0.19771 | 0.19613 | 0.196 | 0.2 | 0.196 | 0.198 | 0.183 | 0.141 | 0.1775 | 0.16855 | 0.2 | 0.19536 | 0.1308 | 0.1654 | 0.2046 | 0.192 | 0.1878 | > 0.1 |  |  |  |  | 16
| 1 | 4 | 0.194 | 0.1952 | 0.18157 | 0.19295 | 0.19271 | 0.206 | 0.194 | 0.203 | 0.202 | 0.187 | 0.165 | 0.1798 | 0.16996 | 0.159 | 0.20128 | 0.1134 | 0.0965 | 0.2049 | 0.1913 | 0.1853 | > 0.1 |  |  |  |  | 0.19098
| 1 | 5 | 0.191 | 0.1964 | 0.18854 | 0.19371 | 0.19725 | 0.204 | 0.198 | 0.201 | 0.197 | 0.184 | 0.142 | 0.1846 | 0.17317 | 0.128 | 0.20114 | 0.1277 | 0.1702 | 0.2092 | 0.1919 | 0.1835 | > 0.1 |  |  |  |  | 10
| 1 | 6 | 0.199 | 0.1935 | 0.18555 | 0.19289 | 0.19645 | 0.2 | 0.2 | 0.2 | 0.199 | 0.186 | 0.144 | 0.183 | 0.16802 | 0.191 | 0.20291 | 0.0991 | 0.1718 | 0.2063 | 0.1914 | 0.1816 | > 0.1
| Mean |  | 0.19417 | 0.19355 | 0.18449 | 0.19388 | 0.19609 | 0.203 | 0.197 | 0.20033 | 0.19883 | 0.18517 | 0.1485 | 0.18093 | 0.17262 | 0.178 | 0.1993 | 0.11928 | 0.15515 | 0.20565 | 0.19122 | 0.18368
| Median |  | 0.193 | 0.19335 | 0.18561 | 0.19333 | 0.19629 | 0.205 | 0.197 | 0.2005 | 0.1985 | 0.185 | 0.1445 | 0.1812 | 0.17157 | 0.193 | 0.19976 | 0.12055 | 0.1678 | 0.20555 | 0.19135 | 0.1841
| Std Dev. |  | 0.0031885 | 0.0019937 | 0.0037772 | 0.002852 | 0.0019961 | 0.0041473 | 0.0027568 | 0.0024221 | 0.0017224 | 0.002137 | 0.0093113 | 0.0028842 | 0.0045589 | 0.028608 | 0.0029444 | 0.013072 | 0.029425 | 0.0021769 | 0.0007414 | 0.0030023
| Rel.Std.Dev. |  | 1.642% | 1.03% | 2.047% | 1.471% | 1.018% | 2.043% | 1.399% | 1.209% | 0.8663% | 1.154% | 6.27% | 1.594% | 2.641% | 16.07% | 1.477% | 10.96% | 18.97% | 1.059% | 0.3877% | 1.634%
| PDM3 |  | 1.666% | 1.343% | -3.403% | 1.514% | 2.673% | 6.291% | 3.15% | 4.895% | 4.109% | -3.046% | -22.25% | -5.263% | -9.618% | -6.799% | 4.354% | -37.54% | -18.76% | 7.679% | 0.1214% | -3.823%
| Z-Score (Absolute) |  | 0.15 | 0.1 | 0.67 | 0.13 | 0.32 | 0.9 | 0.39 | 0.68 | 0.55 | 0.61 | 3.73 | 0.97 | 1.68 | 1.22 | 0.59 | 6.22 | 3.17 | 1.13 | 0.1 | 0.74 | NA
| Analytical results for Be in OREAS 922b (Certified Value 2.3 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.3 | 2.4952 | 2 | 2.2 | 2.25 | 2.2434 | 2.4 | 2.27 | 2.31 | 2.28 | 2.18 | 2.3 | 2.49 | 1.9718 | 2.12 | 2.49 | 2 | 3 | 2.3 | 2 | 2.5 | 2.3 | 2.25 | 2 |  | 1
| 1 | 2 | 2.2 | 2.5649 | 2 | 2.22 | 2.36 | 2.2912 | 2.38 | 2.26 | 2.35 | 2.29 | 2.25 | 2.5 | 2.41 | 2.0697 | 2.3 | 2.44 | 2 | 3 | 2.1 | 2 | 2.32 | 2.19 | 2.17 | 2 |  | 28
| 1 | 3 | 2.3 | 2.47 | 2 | 2.2 | 2.49 | 2.1798 | 2.27 | 2.29 | 2.28 | 2.3 | 2.17 | 2.5 | 2.37 | 2.3839 | 1.79 | 2.44 | 2 | 3 | 2.1 | 2 | 2.36 | 2.27 | 2.31 | 2 |  | 16
| 1 | 4 | 2.2 | 2.5604 | 2 | 2.11 | 2.39 | 2.2304 | 2.4 | 2.26 | 2.37 | 2.33 | 2.27 | 2.4 | 2.26 | 2.2195 | 2.21 | 2.41 | 2 | 3 | 2.3 | 2 | 2.38 | 2.21 | 2.25 | 2 |  | 2.3037
| 1 | 5 | 2.3 | 2.5679 | 2 | 2.22 | 2.38 | 2.1889 | 2.36 | 2.29 | 2.34 | 2.28 | 2.2 | 2.7 | 2.47 | 2.3026 | 1.47 | 2.45 | 2 | 3 | 2.2 | 2 | 2.43 | 2.15 | 2.23 | 2 |  | 11
| 1 | 6 | 2.3 | 2.5589 | 2 | 2.37 | 2.3 | 2.2587 | 2.31 | 2.3 | 2.32 | 2.31 | 2.21 | 2.3 | 2.31 | 2.1993 | 2.11 | 2.49 | 2 | 3 | 2.2 | 2 | 2.51 | 2.29 | 2.19 | 2
| Mean |  | 2.2667 | 2.5362 | 2 | 2.22 | 2.3617 | 2.2321 | 2.3533 | 2.2783 | 2.3283 | 2.2983 | 2.2133 | 2.45 | 2.385 | 2.1911 | 2 | 2.4533 | 2 | 3 | 2.2 | 2 | 2.4167 | 2.235 | 2.2333 | 2
| Median |  | 2.3 | 2.5597 | 2 | 2.21 | 2.37 | 2.2369 | 2.37 | 2.28 | 2.33 | 2.295 | 2.205 | 2.45 | 2.39 | 2.2094 | 2.115 | 2.445 | 2 | 3 | 2.2 | 2 | 2.405 | 2.24 | 2.24 | 2
| Std Dev. |  | 0.05164 | 0.042427 | 0 | 0.084143 | 0.082321 | 0.04228 | 0.052789 | 0.017224 | 0.031885 | 0.019408 | 0.039328 | 0.15166 | 0.089833 | 0.15053 | 0.31164 | 0.031411 | 0 | 0 | 0.089443 | 0 | 0.077115 | 0.060581 | 0.049666 | 0
| Rel.Std.Dev. |  | 2.278% | 1.673% | 0% | 3.79% | 3.486% | 1.894% | 2.243% | 0.756% | 1.369% | 0.8444% | 1.777% | 6.19% | 3.767% | 6.87% | 15.58% | 1.28% | 0% | 0% | 4.066% | 0% | 3.191% | 2.711% | 2.224% | 0%
| PDM3 |  | -1.609% | 10.09% | -13.18% | -3.635% | 2.515% | -3.111% | 2.153% | -1.103% | 1.068% | -0.2345% | -3.924% | 6.349% | 3.527% | -4.888% | -13.18% | 6.494% | -13.18% | 30.22% | -4.503% | -13.18% | 4.902% | -2.984% | -3.056% | -13.18%
| Z-Score (Absolute) |  | 0.1 | 2.32 | NA | 0.52 | 0.75 | 0.42 | 0.67 | 0 | 0.45 | 0.18 | 0.58 | 1.54 | 0.96 | 0.78 | 2.5 | 1.57 | NA | NA | 0.7 | NA | 1.24 | 0.39 | 0.4 | NA
Results from laboratories 2.03, 2.17, 2.19, 2.21 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 922b (Certified Value 12.2 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 11.9 | 12.788 | 11.4 | 12.16 | 12.25 | 12.207 | 11.9 | 11.8 | 13.1 | 11.7 | 11.45 | 12.4 | 13.01 | 10.9291 | 11 | 13.27 | 11.42 | 7.87 | 12.6 | 12.41 | 11.1 | 12.13 | 12.37 | 12.53 | 12.3 | 1
| 1 | 2 | 12.4 | 13.589 | 11.8 | 12.23 | 12.29 | 12.1474 | 12 | 11.75 | 12.85 | 12.25 | 11.15 | 12.7 | 12.9 | 10.7068 | 11 | 13.19 | 10.78 | 7.95 | 12.8 | 13.43 | 11.4 | 11.91 | 12.61 | 12.54 | 12.2 | 29
| 1 | 3 | 12.1 | 12.802 | 11.5 | 12.31 | 12.34 | 11.9873 | 12.6 | 12.25 | 12.85 | 12.1 | 11.85 | 12.1 | 12.3 | 12.304 | 11.3 | 13.2 | 11.59 | 8.03 | 12.5 | 12.11 | 11.6 | 11.67 | 12.18 | 12.7 | 12.2 | 16
| 1 | 4 | 12.4 | 12.407 | 11.7 | 12.07 | 11.92 | 11.6956 | 11.9 | 11.6 | 13.2 | 11.95 | 11.15 | 12.5 | 12.57 | 11.5953 | 9.15 | 13.45 | 11.1 | 7.89 | 13.2 | 12.7 | 11.6 | 11.6 | 12.23 | 12.26 | 12.3 | 12.2198
| 1 | 5 | 12.4 | 12.644 | 11.4 | 12.39 | 12.32 | 12.0561 | 12.8 | 11.65 | 13 | 11.9 | 12 | 12.8 | 13.08 | 11.8354 | 7.47 | 13.27 | 11.88 | 7.92 | 12.8 | 14.14 | 11.5 | 11.66 | 12.27 | 12.66 | 12.1 | 12
| 1 | 6 | 12.4 | 12.527 | 11.6 | 12.22 | 12.45 | 12.3077 | 12.25 | 12.05 | 12.95 | 11.75 | 10.85 | 12.4 | 12.73 | 11.8675 | 10.9 | 13.63 | 11.64 | 7.75 | 12.4 | 12.79 | 11.5 | 11.89 | 12.32 | 13.21 | 12.2
| Mean |  | 12.2667 | 12.7928 | 11.5667 | 12.23 | 12.2617 | 12.0669 | 12.2417 | 11.85 | 12.9917 | 11.9417 | 11.4083 | 12.4833 | 12.765 | 11.5397 | 10.1367 | 13.335 | 11.4017 | 7.9017 | 12.7167 | 12.93 | 11.45 | 11.81 | 12.33 | 12.65 | 12.2167
| Median |  | 12.4 | 12.716 | 11.55 | 12.225 | 12.305 | 12.1018 | 12.125 | 11.775 | 12.975 | 11.925 | 11.3 | 12.45 | 12.815 | 11.7153 | 10.95 | 13.27 | 11.505 | 7.905 | 12.7 | 12.745 | 11.5 | 11.78 | 12.295 | 12.6 | 12.2
| Std Dev. |  | 0.21602 | 0.41849 | 0.1633 | 0.11189 | 0.18038 | 0.21375 | 0.38264 | 0.251 | 0.13934 | 0.20837 | 0.44544 | 0.24833 | 0.29426 | 0.60805 | 1.5173 | 0.17202 | 0.39972 | 0.093041 | 0.28577 | 0.73884 | 0.18708 | 0.20248 | 0.15245 | 0.31458 | 0.075277
| Rel.Std.Dev. |  | 1.761% | 3.271% | 1.412% | 0.9149% | 1.471% | 1.771% | 3.126% | 2.118% | 1.073% | 1.745% | 3.905% | 1.989% | 2.305% | 5.269% | 14.97% | 1.29% | 3.506% | 1.177% | 2.247% | 5.714% | 1.634% | 1.715% | 1.236% | 2.487% | 0.6162%
| PDM3 |  | 0.3831% | 4.689% | -5.345% | 0.08309% | 0.3422% | -1.252% | 0.1786% | -3.027% | 6.316% | -2.276% | -6.641% | 2.156% | 4.461% | -5.566% | -17.05% | 9.126% | -6.696% | -35.34% | 4.066% | 5.811% | -6.3% | -3.354% | 0.9014% | 3.52% | -0.02603%
| Z-Score (Absolute) |  | 0.05 | 0.78 | 0.92 | 0 | 0.04 | 0.23 | 0.02 | 0.53 | 1.06 | 0.4 | 1.14 | 0.35 | 0.74 | 0.96 | 2.9 | 1.53 | 1.15 | 6 | 0.67 | 0.97 | 1.08 | 0.58 | 0.14 | 0.58 | 0.02
| Analytical results for Ca in OREAS 922b (Certified Value 0.04 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.037 | 0.036088 | 0.0612 | 0.04 | 0.0401 | 0.040968 | 0.05 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.039 | 0.048012 | 0.06 | 0.0418 | 0.04 | 0.0609 | 0.04 | 0.04 | 0.04 | 0.04 | 0.0419 | 0.04 | 0.04 | 1
| 1 | 2 | 0.038 | 0.035254 | 0.0612 | 0.04 | 0.0358 | 0.040488 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.04 | 0.047304 | 0.07 | 0.0411 | 0.04 | 0.0605 | 0.04 | 0.04 | 0.04 | 0.041 | 0.0416 | 0.04 | 0.04 | 0
| 1 | 3 | 0.038 | 0.03617 | 0.061 | 0.05 | 0.0376 | 0.040811 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.052178 | 0.09 | 0.0407 | 0.04 | 0.0603 | 0.04 | 0.04 | 0.04 | 0.04 | 0.0396 | 0.04 | 0.04 | 16
| 1 | 4 | 0.038 | 0.037187 | 0.0611 | 0.04 | 0.0391 | 0.040816 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.046428 | 0.04 | 0.0409 | 0.04 | 0.0611 | 0.04 | 0.04 | 0.04 | 0.041 | 0.041 | 0.04 | 0.04 | 0.039966
| 1 | 5 | 0.038 | 0.037685 | 0.061 | 0.05 | 0.039 | 0.04284 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.05 | 0.041 | 0.048299 | 0.06 | 0.0408 | 0.04 | 0.0605 | 0.04 | 0.04 | 0.04 | 0.039 | 0.0409 | 0.04 | 0.04 | 13
| 1 | 6 | 0.041 | 0.036311 | 0.061 | 0.04 | 0.0362 | 0.04065 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.053747 | 0.08 | 0.0408 | 0.04 | 0.0592 | 0.04 | 0.04 | 0.04 | 0.04 | 0.0426 | 0.04 | 0.04
| Mean |  | 0.038333 | 0.036449 | 0.061083 | 0.043333 | 0.037967 | 0.041095 | 0.041667 | 0.04 | 0.04 | 0.04 | 0.04 | 0.043333 | 0.04 | 0.049328 | 0.066667 | 0.041017 | 0.04 | 0.060417 | 0.04 | 0.04 | 0.04 | 0.040167 | 0.041267 | 0.04 | 0.04
| Median |  | 0.038 | 0.036241 | 0.06105 | 0.04 | 0.0383 | 0.040813 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.04 | 0.048155 | 0.065 | 0.04085 | 0.04 | 0.0605 | 0.04 | 0.04 | 0.04 | 0.04 | 0.0413 | 0.04 | 0.04
| Std Dev. |  | 0.0013663 | 0.00086361 | 9.8319e-05 | 0.005164 | 0.0017236 | 0.0008704 | 0.0040825 | 0 | 0 | 0 | 0 | 0.005164 | 0.00063246 | 0.0029307 | 0.017512 | 0.00040702 | 0 | 0.00066458 | 0 | 0 | 0 | 0.00075277 | 0.001027 | 0 | 0
| Rel.Std.Dev. |  | 3.564% | 2.369% | 0.161% | 11.92% | 4.54% | 2.118% | 9.798% | 0% | 0% | 0% | 0% | 11.92% | 1.581% | 5.941% | 26.27% | 0.9923% | 0% | 1.1% | 0% | 0% | 0% | 1.874% | 2.489% | 0% | 0%
| PDM3 |  | -4.086% | -8.8% | 52.84% | 8.424% | -5.004% | 2.825% | 4.254% | 0.08389% | 0.08389% | 0.08389% | 0.08389% | 8.424% | 0.08389% | 23.42% | 66.81% | 2.628% | 0.08389% | 51.17% | 0.08389% | 0.08389% | 0.08389% | 0.5009% | 3.253% | 0.08389% | 0.08389%
| Z-Score (Absolute) |  | 1.11 | 2.35 | 13.98 | 2.21 | 1.35 | 0.73 | 1.11 | 0 | 0 | 0 | 0 | 2.21 | 0 | 6.19 | 17.69 | 0.67 | 0 | 13.54 | 0 | 0 | 0 | 0.11 | 0.84 | 0 | 0
| Analytical results for Cd in OREAS 922b (Certified Value 0.078 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.07 | 0.1 | < 0.2 | 0.1 | < 0.2 | 0.1 | 0.04 | 0.08 | 0.06 | 0.08 | < 0.3 | 0.08 | 0.09937 | 0.74 | 0.07 | 0.04 | 0.09 | < 0.1 | 0.1 | 0.07 | 0.14 | 0.07 | < 0.1 |  |  | 1
| 1 | 2 | 0.09 | 0.1 | < 0.2 | 0.09 | < 0.2 | 0.06 | 0.07 | 0.07 | 0.06 | 0.07 | < 0.3 | 0.08 | 0.095044 | 0.84 | 0.08 | 0.06 | 0.09 | 0.1 | 0.11 | 0.06 | 0.07 | 0.08 | < 0.1 |  |  | 11
| 1 | 3 | 0.09 | 0.1 | < 0.2 | 0.07 | < 0.2 | 0.11 | 0.04 | 0.1 | 0.06 | 0.1 | < 0.3 | 0.08 | 0.089942 | 1.01 | 0.07 | < 0.02 | 0.14 | 0.2 | 0.1 | 0.09 | 0.09 | 0.07 | < 0.1 |  |  | 16
| 1 | 4 | 0.09 | 0.1 | < 0.2 | 0.07 | < 0.2 | 0.07 | 0.1 | 0.07 | 0.06 | 0.09 | < 0.3 | 0.09 | 0.11051 | 0.51 | 0.09 | 0.02 | 0.17 | 0.2 | 0.1 | 0.07 | 0.1 | 0.08 | < 0.1 |  |  | 0.078209
| 1 | 5 | 0.07 | 0.1 | < 0.2 | 0.09 | < 0.2 | 0.11 | 0.09 | 0.07 | 0.04 | 0.07 | < 0.3 | 0.07 | 0.085631 | 0.64 | 0.07 | < 0.02 | 0.2 | < 0.1 | 0.1 | 0.08 | 0.07 | 0.08 | < 0.1 |  |  | 14
| 1 | 6 | 0.06 | 0.1 | < 0.2 | 0.07 | < 0.2 | 0.1 | 0.04 | 0.04 | 0.05 | 0.05 | < 0.3 | 0.09 | 0.11328 | 0.83 | 0.08 | 0.07 | 0.09 | 0.1 | 0.11 | 0.07 | 0.11 | 0.08 | < 0.1
| Mean |  | 0.078333 | 0.1 |  | 0.081667 |  | 0.091667 | 0.063333 | 0.071667 | 0.055 | 0.076667 |  | 0.081667 | 0.098963 | 0.76167 | 0.076667 | 0.0475 | 0.13 | 0.15 | 0.10333 | 0.073333 | 0.096667 | 0.076667
| Median |  | 0.08 | 0.1 |  | 0.08 |  | 0.1 | 0.055 | 0.07 | 0.06 | 0.075 |  | 0.08 | 0.097207 | 0.785 | 0.075 | 0.05 | 0.115 | 0.15 | 0.1 | 0.07 | 0.095 | 0.08
| Std Dev. |  | 0.013292 | 1.5202e-17 |  | 0.013292 |  | 0.02137 | 0.027325 | 0.019408 | 0.0083666 | 0.017512 |  | 0.0075277 | 0.011072 | 0.17383 | 0.008165 | 0.022174 | 0.047749 | 0.057735 | 0.005164 | 0.010328 | 0.026583 | 0.005164
| Rel.Std.Dev. |  | 16.97% | 1.52e-14% |  | 16.28% |  | 23.31% | 43.15% | 27.08% | 15.21% | 22.84% |  | 9.218% | 11.19% | 22.82% | 10.65% | 46.68% | 36.73% | 38.49% | 4.997% | 14.08% | 27.5% | 6.736%
| PDM3 |  | 0.1594% | 27.86% |  | 4.422% |  | 17.21% | -19.02% | -8.365% | -29.68% | -1.972% |  | 4.422% | 26.54% | 873.9% | -1.972% | -39.27% | 66.22% | 91.79% | 32.13% | -6.234% | 23.6% | -1.972%
| Z-Score (Absolute) |  | 0.13 | NA | 0.71 | 0 | 0.71 | 0.39 | 0.71 | 0.39 | 1.04 | 0.19 | 2.66 | 0 | 0.67 | 26.51 | 0.19 | 1.82 | 1.88 | NA | 0.84 | 0.32 | 0.58 | 0.19 | 1.23
Results from laboratories 2.03 and 2.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 922b (Certified Value 79 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 0
1
| 1 | 1 | 85.4082 | 80 | 80.86 | 77.47 | 77.5961 | 79.8 | 75.9 | 79.4 | 78.4 | 84.9 | 70.3 | 79.05 | 74.4 | 77.76 | 76.66 | 86 | 82 | 81.6 | 76.4 | 79.6 | 76.32 | 62 |  |  |  | 1
| 1 | 2 | 85.7732 | 80 | 82.62 | 74.9 | 80.3741 | 78.5 | 74.1 | 80.7 | 80.4 | 79.9 | 71.3 | 76.11 | 75.3 | 77.7 | 70.51 | 84 | 82 | 83.6 | 76.1 | 80.66 | 76.29 | 67 |  |  |  | 12
| 1 | 3 | 85.6959 | 80 | 83.01 | 76.2 | 78.8151 | 83.4 | 78.1 | 77.4 | 81.6 | 87.1 | 69.5 | 71.27 | 79.6 | 75.25 | 77.9 | 83 | 77 | 86.9 | 77.3 | 80.18 | 78.43 | 65 |  |  |  | 16
| 1 | 4 | 85.6894 | 80 | 78.61 | 74.15 | 77.9648 | 77.9 | 73.6 | 82.6 | 79.2 | 82.3 | 70.9 | 74.18 | 61.6 | 77.98 | 73.19 | 86 | 83 | 85.2 | 77.3 | 78.12 | 74.01 | 63 |  |  |  | 79.0115
| 1 | 5 | 85.4313 | 80 | 81.68 | 74.54 | 80.165 | 87.2 | 75.1 | 78.8 | 77.1 | 85.9 | 72 | 78.86 | 49.9 | 77.07 | 78.13 | 84 | 80 | 84.2 | 78 | 78.39 | 74.51 | 62 |  |  |  | 15
| 1 | 6 | 85.4615 | 80 | 81.59 | 75.18 | 79.5277 | 81.3 | 79.4 | 79 | 77.1 | 78.5 | 70.2 | 75.12 | 75 | 79.4 | 78.47 | 81 | 82 | 84.3 | 76.4 | 78.46 | 79.81 | 63
| Mean |  | 85.5766 | 80 | 81.395 | 75.4067 | 79.0738 | 81.35 | 76.0333 | 79.65 | 78.9667 | 83.1 | 70.7 | 75.765 | 69.3 | 77.5267 | 75.81 | 84 | 81 | 84.3 | 76.9167 | 79.235 | 76.5617 | 63.6667
| Median |  | 85.5755 | 80 | 81.635 | 75.04 | 79.1714 | 80.55 | 75.5 | 79.2 | 78.8 | 83.6 | 70.6 | 75.615 | 74.7 | 77.73 | 77.28 | 84 | 82 | 84.25 | 76.85 | 79.03 | 76.305 | 63
| Std Dev. |  | 0.1602 | 0 | 1.5666 | 1.2268 | 1.1461 | 3.4887 | 2.2871 | 1.793 | 1.8074 | 3.4386 | 0.88769 | 2.9537 | 11.2804 | 1.3558 | 3.2405 | 1.8974 | 2.1909 | 1.7527 | 0.73052 | 1.0597 | 2.2314 | 1.9664
| Rel.Std.Dev. |  | 0.1872% | 0% | 1.925% | 1.627% | 1.449% | 4.289% | 3.008% | 2.251% | 2.289% | 4.138% | 1.256% | 3.899% | 16.28% | 1.749% | 4.274% | 2.259% | 2.705% | 2.079% | 0.9498% | 1.337% | 2.914% | 3.089%
| PDM3 |  | 8.309% | 1.251% | 3.017% | -4.562% | 0.07881% | 2.96% | -3.769% | 0.8081% | -0.05678% | 5.175% | -10.52% | -4.109% | -12.29% | -1.879% | -4.052% | 6.314% | 2.517% | 6.693% | -2.651% | 0.2828% | -3.101% | -19.42%
| Z-Score (Absolute) |  | 1.62 | 0.24 | 0.59 | 0.89 | 0.01 | 0.58 | 0.74 | 0.16 | 0.01 | 1.01 | 2.06 | 0.81 | 2.41 | 0.37 | 0.8 | 1.23 | 0.49 | 1.31 | 0.52 | 0.05 | 0.61 | 3.8
| Analytical results for Co in OREAS 922b (Certified Value 3.7 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 4 | 4.0479 | 4.2 | 3.79 | 3.9 | 2.8091 | 3.7 | 3.4 | 3.4 | 3.4 | 3.4 | 4 | 3.7 | 3.2588 | 3.9 | 3.7 | 3.5 | 4.1 | 3.8 | 3.9 | 3.5 | 3.76 | 3.8 | 3.7 | 3.6 | 1
| 1 | 2 | 4 | 4.0602 | 4.2 | 3.72 | 3.6 | 2.7453 | 3.5 | 3.5 | 3.4 | 3.6 | 3.4 | 4 | 3.6 | 3.2614 | 3.8 | 3.8 | 3.4 | 4 | 3.9 | 4.1 | 3.8 | 3.82 | 3.8 | 3.8 | 3.6 | 32
| 1 | 3 | 4 | 4.0694 | 4.2 | 3.27 | 3.7 | 2.9181 | 3.2 | 3.5 | 3.3 | 3.5 | 3.6 | 4 | 3.5 | 3.7409 | 3.9 | 3.7 | 3.8 | 3.9 | 3.9 | 3.8 | 3.7 | 3.84 | 3.8 | 3.9 | 3.6 | 16
| 1 | 4 | 4 | 4.0545 | 4.2 | 3.65 | 3.7 | 2.9559 | 3.5 | 3.4 | 3.5 | 3.5 | 3.4 | 4 | 3.6 | 3.5841 | 3.2 | 3.8 | 3.5 | 4 | 3.9 | 4 | 3.7 | 3.75 | 3.7 | 3.8 | 3.6 | 3.7014
| 1 | 5 | 4 | 4.0418 | 4.2 | 3.68 | 3.8 | 2.766 | 3.4 | 3.4 | 3.3 | 3.4 | 3.6 | 4 | 3.9 | 3.594 | 2.6 | 3.7 | 3.8 | 4 | 3.8 | 4 | 3.7 | 3.84 | 3.9 | 4 | 3.6 | 16
| 1 | 6 | 4 | 4.044 | 4.2 | 3.09 | 3.6 | 2.8844 | 3.3 | 3.6 | 3.4 | 3.3 | 3.2 | 4 | 3.5 | 3.634 | 3.8 | 3.8 | 3.7 | 4.2 | 3.7 | 3.7 | 3.5 | 3.87 | 3.9 | 4 | 3.6
| Mean |  | 4 | 4.0529 | 4.2 | 3.5333 | 3.7167 | 2.8465 | 3.4333 | 3.4667 | 3.3833 | 3.45 | 3.4333 | 4 | 3.6333 | 3.5122 | 3.5333 | 3.75 | 3.6167 | 4.0333 | 3.8333 | 3.9167 | 3.65 | 3.8133 | 3.8167 | 3.8667 | 3.6
| Median |  | 4 | 4.0512 | 4.2 | 3.665 | 3.7 | 2.8468 | 3.45 | 3.45 | 3.4 | 3.45 | 3.4 | 4 | 3.6 | 3.5891 | 3.8 | 3.75 | 3.6 | 4 | 3.85 | 3.95 | 3.7 | 3.83 | 3.8 | 3.85 | 3.6
| Std Dev. |  | 0 | 0.010535 | 0 | 0.28345 | 0.1169 | 0.085583 | 0.17512 | 0.08165 | 0.075277 | 0.10488 | 0.15055 | 0 | 0.15055 | 0.20303 | 0.52789 | 0.054772 | 0.17224 | 0.10328 | 0.08165 | 0.1472 | 0.12247 | 0.048028 | 0.075277 | 0.12111 | 0
| Rel.Std.Dev. |  | 0% | 0.2599% | 0% | 8.022% | 3.145% | 3.007% | 5.101% | 2.355% | 2.225% | 3.04% | 4.385% | 0% | 4.144% | 5.781% | 14.94% | 1.461% | 4.762% | 2.561% | 2.13% | 3.758% | 3.355% | 1.259% | 1.972% | 3.132% | 0%
| PDM3 |  | 8.066% | 9.496% | 13.47% | -4.542% | 0.4112% | -23.1% | -7.243% | -6.343% | -8.594% | -6.793% | -7.243% | 8.066% | -1.84% | -5.112% | -4.542% | 1.312% | -2.29% | 8.966% | 3.563% | 5.814% | -1.39% | 3.023% | 3.113% | 4.464% | -2.741%
| Z-Score (Absolute) |  | NA | 1.54 | 2.08 | 0.37 | 0.31 | 2.89 | 0.74 | 0.61 | 0.92 | 0.67 | 0.74 | NA | 0 | 0.45 | 0.37 | 0.43 | 0.06 | 1.47 | 0.74 | 1.04 | 0.06 | 0.66 | 0.67 | 0.86 | 0.12
Results from laboratories 2.01 and 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 922b (Certified Value 79 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 87 | 95.02 | 87 | 84.22 | 91 | 75.4238 | 80 | 76 | 74 | 74 | 73 | 89 | 78 | 75.0871 | 71 | 91 | 91 | 62 | 80 | 85 | 69.8 | 84 | 82 | 80 |  | 1
| 1 | 2 | 81 | 94.657 | 85 | 75.39 | 83 | 73.676 | 81 | 72 | 72 | 73 | 76 | 95 | 83 | 78.8024 | 74 | 90 | 86 | 65 | 80 | 84 | 69.1 | 82 | 75 | 83 |  | 33
| 1 | 3 | 84 | 94.514 | 86 | 75.13 | 92 | 73.1722 | 76 | 75 | 73 | 73 | 72 | 89 | 83 | 76.4188 | 74 | 86 | 90 | 64 | 78 | 86 | 68.8 | 83 | 80 | 81 |  | 16
| 1 | 4 | 86 | 94.239 | 86 | 76.46 | 85 | 75.4831 | 80 | 72 | 72 | 72 | 75 | 69 | 80 | 81.1915 | 58 | 87 | 85 | 71 | 85 | 86 | 68.9 | 83 | 79 | 83 |  | 79.3998
| 1 | 5 | 87 | 94.683 | 86 | 76.08 | 86 | 77.5002 | 77 | 73 | 76 | 74 | 73 | 74 | 86 | 73.0981 | 49 | 84 | 94 | 63 | 82 | 84 | 65.1 | 85 | 81 | 82 |  | 17
| 1 | 6 | 87 | 94.404 | 86 | 75.91 | 83 | 72.6601 | 76 | 76 | 76 | 73 | 76 | 67 | 76 | 78.4181 | 69 | 88 | 87 | 64 | 83 | 86 | 69.2 | 78 | 78 | 83
| Mean |  | 85.3333 | 94.5862 | 86 | 77.1983 | 86.6667 | 74.6526 | 78.3333 | 74 | 73.8333 | 73.1667 | 74.1667 | 80.5 | 81 | 77.1693 | 65.8333 | 87.6667 | 88.8333 | 64.8333 | 81.3333 | 85.1667 | 68.4833 | 82.5 | 79.1667 | 82
| Median |  | 86.5 | 94.5855 | 86 | 75.995 | 85.5 | 74.5499 | 78.5 | 74 | 73.5 | 73 | 74 | 81.5 | 81.5 | 77.4185 | 70 | 87.5 | 88.5 | 64 | 81 | 85.5 | 69 | 83 | 79.5 | 82.5
| Std Dev. |  | 2.4221 | 0.2689 | 0.63246 | 3.473 | 3.9328 | 1.8171 | 2.2509 | 1.8974 | 1.8348 | 0.75277 | 1.7224 | 11.929 | 3.6878 | 2.8941 | 10.1472 | 2.582 | 3.4303 | 3.1885 | 2.5033 | 0.98319 | 1.694 | 2.429 | 2.4833 | 1.2649
| Rel.Std.Dev. |  | 2.838% | 0.2843% | 0.7354% | 4.499% | 4.538% | 2.434% | 2.874% | 2.564% | 2.485% | 1.029% | 2.322% | 14.82% | 4.553% | 3.75% | 15.41% | 2.945% | 3.861% | 4.918% | 3.078% | 1.154% | 2.474% | 2.944% | 3.137% | 1.543%
| PDM3 |  | 7.473% | 19.13% | 8.313% | -2.773% | 9.152% | -5.979% | -1.343% | -6.801% | -7.011% | -7.85% | -6.591% | 1.386% | 2.015% | -2.809% | -17.09% | 10.41% | 11.88% | -18.35% | 2.435% | 7.263% | -13.75% | 3.905% | -0.2936% | 3.275%
| Z-Score (Absolute) |  | 0.66 | 1.78 | 0.74 | 0.32 | 0.83 | 0.63 | 0.18 | 0.7 | 0.72 | 0.81 | 0.68 | 0.08 | 0.14 | 0.32 | 1.69 | 0.95 | 1.09 | 1.81 | 0.18 | 0.64 | 1.37 | 0.32 | 0.08 | 0.26
| Analytical results for Cs in OREAS 922b (Certified Value 4.59 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 4.6 | 6.2114 | 4.42 | 4.76 | 4.5009 | 4.64 | 4.7 | 4.8 | 4.52 | 4.46 | 4.7 | 4.76 | 4.2556 | 4.64 | 4.67 | 4.5 | 4 | 4.59 | 4.95 | 4.35 | 5 |  |  |  |  | 1
| 1 | 2 | 4.7 | 6.276 | 4.34 | 4.63 | 4.6986 | 4.59 | 4.43 | 4.99 | 4.67 | 4.29 | 4.7 | 4.56 | 4.2515 | 4.57 | 4.63 | 4.2 | 5 | 4.49 | 4.93 | 4.38 | 5.1 |  |  |  |  | 14
| 1 | 3 | 4.6 | 6.2856 | 4.24 | 4.75 | 4.5641 | 4.28 | 4.72 | 4.79 | 4.58 | 4.67 | 4.62 | 4.53 | 4.7105 | 4.6 | 4.62 | 4.6 | 5 | 4.3 | 5 | 4.48 | 5.1 |  |  |  |  | 16
| 1 | 4 | 4.7 | 6.2457 | 4.54 | 4.78 | 4.5173 | 4.51 | 4.39 | 5.05 | 4.64 | 4.46 | 4.66 | 4.53 | 4.5569 | 3.79 | 4.72 | 4.4 | 5 | 4.42 | 4.82 | 4.26 | 4.9 |  |  |  |  | 4.5919
| 1 | 5 | 4.6 | 6.2069 | 4.64 | 4.63 | 4.597 | 4.6 | 4.65 | 4.85 | 4.5 | 4.59 | 4.8 | 4.86 | 4.4231 | 3.05 | 4.79 | 4.8 | 4 | 4.45 | 4.83 | 4.29 | 5.2 |  |  |  |  | 18
| 1 | 6 | 4.7 | 6.2577 | 4.32 | 4.59 | 4.5226 | 4.37 | 4.75 | 4.72 | 4.48 | 4.24 | 4.68 | 4.58 | 4.7261 | 4.48 | 4.7 | 4.6 | 4 | 4.47 | 4.91 | 4.57 | 4.9
| Mean |  | 4.65 | 6.2472 | 4.4167 | 4.69 | 4.5668 | 4.4983 | 4.6067 | 4.8667 | 4.565 | 4.4517 | 4.6933 | 4.6367 | 4.4873 | 4.1883 | 4.6883 | 4.5167 | 4.5 | 4.4533 | 4.9067 | 4.3883 | 5.0333
| Median |  | 4.65 | 6.2517 | 4.38 | 4.69 | 4.5434 | 4.55 | 4.675 | 4.825 | 4.55 | 4.46 | 4.69 | 4.57 | 4.49 | 4.525 | 4.685 | 4.55 | 4.5 | 4.46 | 4.92 | 4.365 | 5.05
| Std Dev. |  | 0.054772 | 0.03263 | 0.14935 | 0.082219 | 0.073529 | 0.14359 | 0.15629 | 0.12723 | 0.07791 | 0.16606 | 0.060222 | 0.13924 | 0.21218 | 0.64167 | 0.063061 | 0.20412 | 0.54772 | 0.094798 | 0.070048 | 0.11754 | 0.12111
| Rel.Std.Dev. |  | 1.178% | 0.5223% | 3.382% | 1.753% | 1.61% | 3.192% | 3.393% | 2.614% | 1.707% | 3.73% | 1.283% | 3.003% | 4.728% | 15.32% | 1.345% | 4.519% | 12.17% | 2.129% | 1.428% | 2.679% | 2.406%
| PDM3 |  | 1.264% | 36.05% | -3.817% | 2.136% | -0.5485% | -2.038% | 0.3208% | 5.983% | -0.5866% | -3.055% | 2.208% | 0.9741% | -2.279% | -8.789% | 2.099% | -1.639% | -2.002% | -3.018% | 6.854% | -4.434% | 9.612%
| Z-Score (Absolute) |  | 0.41 | 10.67 | 1.09 | 0.66 | 0.13 | 0.57 | 0.13 | 1.8 | 0.14 | 0.87 | 0.69 | 0.32 | 0.64 | 2.56 | 0.65 | 0.45 | NA | 0.86 | 2.06 | 1.27 | 2.87
Results from laboratory 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 922b (Certified Value 0.219 wt.%) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.221 | 0.20903 | 0.224 | 0.21816 | 0.22334 | 0.21586 | 0.225 | 0.212 | 0.217 | 0.22 | 0.209 | 0.226 | 0.2169 | 0.20469 | 0.216 | 0.23093 | 0.22794 | 0.21951 | 0.22955 | 0.2377 | 0.2011 | 0.2269 | 0.22682 | 0.21874 | 0.21326 | 1
| 1 | 2 | 0.222 | 0.21142 | 0.22 | 0.21953 | 0.20577 | 0.21364 | 0.223 | 0.211 | 0.22 | 0.219 | 0.214 | 0.23 | 0.2233 | 0.20796 | 0.213 | 0.23296 | 0.20852 | 0.21959 | 0.2293 | 0.2348 | 0.2011 | 0.2234 | 0.22506 | 0.21557 | 0.21462 | 35
| 1 | 3 | 0.22 | 0.21063 | 0.222 | 0.21758 | 0.21338 | 0.21595 | 0.21 | 0.216 | 0.215 | 0.221 | 0.206 | 0.229 | 0.2167 | 0.22163 | 0.221 | 0.22964 | 0.21955 | 0.2137 | 0.22635 | 0.2351 | 0.201 | 0.2246 | 0.22455 | 0.22325 | 0.21611 | 16
| 1 | 4 | 0.223 | 0.20782 | 0.222 | 0.21772 | 0.21033 | 0.21493 | 0.223 | 0.211 | 0.223 | 0.22 | 0.214 | 0.23 | 0.2192 | 0.22092 | 0.227 | 0.23299 | 0.21544 | 0.21712 | 0.23188 | 0.2368 | 0.2018 | 0.2202 | 0.21916 | 0.22299 | 0.2202 | 0.21938
| 1 | 5 | 0.221 | 0.21124 | 0.224 | 0.21935 | 0.21442 | 0.21539 | 0.217 | 0.211 | 0.219 | 0.217 | 0.21 | 0.224 | 0.2244 | 0.21723 | 0.218 | 0.23299 | 0.23302 | 0.21434 | 0.22889 | 0.2345 | 0.2013 | 0.2238 | 0.2225 | 0.22279 | 0.21862 | 19
| 1 | 6 | 0.225 | 0.20707 | 0.224 | 0.21124 | 0.20888 | 0.21303 | 0.21 | 0.216 | 0.219 | 0.217 | 0.213 | 0.224 | 0.2236 | 0.21426 | 0.232 | 0.23804 | 0.22184 | 0.21893 | 0.22445 | 0.2352 | 0.2017 | 0.2219 | 0.22382 | 0.21456 | 0.22101
| Mean |  | 0.222 | 0.20953 | 0.22267 | 0.21726 | 0.21269 | 0.2148 | 0.218 | 0.21283 | 0.21883 | 0.219 | 0.211 | 0.22717 | 0.22068 | 0.21445 | 0.22117 | 0.23292 | 0.22105 | 0.2172 | 0.2284 | 0.23568 | 0.20133 | 0.22347 | 0.22365 | 0.21965 | 0.2173
| Median |  | 0.2215 | 0.20983 | 0.223 | 0.21794 | 0.21185 | 0.21516 | 0.22 | 0.2115 | 0.219 | 0.2195 | 0.2115 | 0.2275 | 0.22125 | 0.21574 | 0.2195 | 0.23297 | 0.2207 | 0.21802 | 0.22909 | 0.23515 | 0.2012 | 0.2236 | 0.22418 | 0.22076 | 0.21737
| Std Dev. |  | 0.0017889 | 0.0018402 | 0.001633 | 0.0030632 | 0.0060805 | 0.0012089 | 0.0067528 | 0.0024833 | 0.0027142 | 0.0016733 | 0.0032249 | 0.0028577 | 0.0035085 | 0.0069074 | 0.0071391 | 0.0028616 | 0.0087419 | 0.0026259 | 0.0026187 | 0.0012703 | 0.00033862 | 0.0022941 | 0.0026214 | 0.0039334 | 0.0031253
| Rel.Std.Dev. |  | 0.8058% | 0.8782% | 0.7334% | 1.41% | 2.859% | 0.5628% | 3.098% | 1.167% | 1.24% | 0.7641% | 1.528% | 1.258% | 1.59% | 3.221% | 3.228% | 1.229% | 3.955% | 1.209% | 1.147% | 0.539% | 0.1682% | 1.027% | 1.172% | 1.791% | 1.438%
| PDM3 |  | 1.194% | -4.488% | 1.498% | -0.9643% | -3.051% | -2.087% | -0.6289% | -2.984% | -0.2491% | -0.1731% | -3.82% | 3.55% | 0.5942% | -2.248% | 0.8145% | 6.174% | 0.7621% | -0.9943% | 4.113% | 7.432% | -8.226% | 1.863% | 1.947% | 0.1232% | -0.9465%
| Z-Score (Absolute) |  | 0.48 | 1.52 | 0.59 | 0.28 | 1.01 | 0.67 | 0.16 | 0.99 | 0.03 | 0 | 1.28 | 1.31 | 0.27 | 0.73 | 0.35 | 2.24 | 0.33 | 0.29 | 1.51 | 2.68 | 2.84 | 0.72 | 0.75 | 0.1 | 0.27
| Analytical results for Dy in OREAS 922b (Certified Value 3.28 ppm) | [26 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 3.5497; 2.02: 3.5; 2.03: 3.4739; 2.04: 4; 2.05: 3.43; 2.06: 3.2; 2.07: 2.93 | 1
| 1 | 2 | 2.01: 3.3205; 2.02: 3.5; 2.03: 3.3257; 2.04: 4.2; 2.05: 3.46; 2.06: 3; 2.07: 2.97 | 3
| 1 | 3 | 2.01: 3.3291; 2.02: 3.5; 2.03: 3.225; 2.04: 4.1; 2.05: 3.45; 2.06: 3.2; 2.07: 2.98 | 16
| 1 | 4 | 2.01: 3.2635; 2.02: 3.5; 2.03: 3.2517; 2.04: 4.4; 2.05: 3.52; 2.06: 3.1; 2.07: 3.01 | 3.2818
| 1 | 5 | 2.01: 3.4176; 2.02: 3.5; 2.03: 3.247; 2.04: 4.1; 2.05: 3.31; 2.06: 3.4; 2.07: 2.94 | 20
| 1 | 6 | 2.01: 3.344; 2.02: 3.5; 2.03: 3.263; 2.04: 4.3; 2.05: 3.49; 2.06: 3.2; 2.07: 2.97
| Mean |  | 2.01: 3.3707; 2.02: 3.5; 2.03: 3.2977; 2.04: 4.1833; 2.05: 3.4433; 2.06: 3.1833; 2.07: 2.9667
| Median |  | 2.01: 3.3365; 2.02: 3.5; 2.03: 3.2573; 2.04: 4.15; 2.05: 3.455; 2.06: 3.2; 2.07: 2.97
| Std Dev. |  | 2.01: 0.10069; 2.02: 0; 2.03: 0.092746; 2.04: 0.1472; 2.05: 0.072572; 2.06: 0.13292; 2.07: 0.028752
| Rel.Std.Dev. |  | 2.01: 2.987%; 2.02: 0%; 2.03: 2.812%; 2.04: 3.519%; 2.05: 2.108%; 2.06: 4.175%; 2.07: 0.9692%
| PDM3 |  | 2.01: 2.71%; 2.02: 6.649%; 2.03: 0.485%; 2.04: 27.47%; 2.05: 4.922%; 2.06: -3%; 2.07: -9.602%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.03: 0.38; 2.04: 4.24; 2.05: 0.38; 2.06: 0.98; 2.07: 2.11
| Analytical results for Er in OREAS 922b (Certified Value 1.72 ppm) | [26 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 1.7054; 2.02: 1.71; 2.03: 1.6431; 2.04: 2.1; 2.05: 1.77; 2.06: 1.6; 2.07: 1.59 | 1
| 1 | 2 | 2.01: 1.7475; 2.02: 1.71; 2.03: 1.7165; 2.04: 2.2; 2.05: 1.85; 2.06: 1.6; 2.07: 1.5 | 4
| 1 | 3 | 2.01: 1.7163; 2.02: 1.71; 2.03: 1.6958; 2.04: 2.2; 2.05: 1.79; 2.06: 1.8; 2.07: 1.47 | 16
| 1 | 4 | 2.01: 1.6848; 2.02: 1.71; 2.03: 1.6261; 2.04: 2.2; 2.05: 1.84; 2.06: 1.6; 2.07: 1.55 | 1.7221
| 1 | 5 | 2.01: 1.7486; 2.02: 1.71; 2.03: 1.6521; 2.04: 2.1; 2.05: 1.94; 2.06: 1.8; 2.07: 1.53 | 21
| 1 | 6 | 2.01: 1.6529; 2.02: 1.71; 2.03: 1.6826; 2.04: 2.1; 2.05: 1.84; 2.06: 1.7; 2.07: 1.58
| Mean |  | 2.01: 1.7092; 2.02: 1.71; 2.03: 1.6694; 2.04: 2.15; 2.05: 1.8383; 2.06: 1.6833; 2.07: 1.5367
| Median |  | 2.01: 1.7108; 2.02: 1.71; 2.03: 1.6674; 2.04: 2.15; 2.05: 1.84; 2.06: 1.65; 2.07: 1.54
| Std Dev. |  | 2.01: 0.037011; 2.02: 2.4324e-16; 2.03: 0.034515; 2.04: 0.054772; 2.05: 0.059133; 2.06: 0.098319; 2.07: 0.046332
| Rel.Std.Dev. |  | 2.01: 2.165%; 2.02: 1.422e-14%; 2.03: 2.068%; 2.04: 2.548%; 2.05: 3.217%; 2.06: 5.841%; 2.07: 3.015%
| PDM3 |  | 2.01: -0.7438%; 2.02: -0.7003%; 2.03: -3.059%; 2.04: 24.85%; 2.05: 6.752%; 2.06: -2.249%; 2.07: -10.77%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.01; 2.03: 0.67; 2.04: 7.45; 2.05: 2.18; 2.06: 0.44; 2.07: 2.92
| Analytical results for Eu in OREAS 922b (Certified Value 1.21 ppm) | [26 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 1.3522; 2.02: 0.52; 2.03: 1.1981; 2.04: 1.21; 2.05: 1.22; 2.06: 1.2; 2.07: 1.18 | 1
| 1 | 2 | 2.01: 1.3215; 2.02: 0.52; 2.03: 1.2563; 2.04: 1.27; 2.05: 1.19; 2.06: 1.1; 2.07: 1.11 | 5
| 1 | 3 | 2.01: 1.3096; 2.02: 0.52; 2.03: 1.2242; 2.04: 1.17; 2.05: 1.18; 2.06: 1.2; 2.07: 1.08 | 16
| 1 | 4 | 2.01: 1.3227; 2.02: 0.52; 2.03: 1.1648; 2.04: 1.21; 2.05: 1.23; 2.06: 1.1; 2.07: 1.16 | 1.2077
| 1 | 5 | 2.01: 1.3091; 2.02: 0.52; 2.03: 1.2176; 2.04: 1.21; 2.05: 1.22; 2.06: 1.2; 2.07: 1.08 | 22
| 1 | 6 | 2.01: 1.3412; 2.02: 0.52; 2.03: 1.2115; 2.04: 1.18; 2.05: 1.25; 2.06: 1.2; 2.07: 1.1
| Mean |  | 2.01: 1.326; 2.02: 0.52; 2.03: 1.2121; 2.04: 1.2083; 2.05: 1.215; 2.06: 1.1667; 2.07: 1.1183
| Median |  | 2.01: 1.3221; 2.02: 0.52; 2.03: 1.2145; 2.04: 1.21; 2.05: 1.22; 2.06: 1.2; 2.07: 1.105
| Std Dev. |  | 2.01: 0.017339; 2.02: 0; 2.03: 0.030193; 2.04: 0.034881; 2.05: 0.025884; 2.06: 0.05164; 2.07: 0.042151
| Rel.Std.Dev. |  | 2.01: 1.308%; 2.02: 0%; 2.03: 2.491%; 2.04: 2.887%; 2.05: 2.13%; 2.06: 4.426%; 2.07: 3.769%
| PDM3 |  | 2.01: 9.796%; 2.02: -56.94%; 2.03: 0.3589%; 2.04: 0.04883%; 2.05: 0.6008%; 2.06: -3.401%; 2.07: -7.403%
| Z-Score (Absolute) |  | 2.01: 1.91; 2.02: 11.14; 2.03: 0.06; 2.04: 0; 2.05: 0.11; 2.06: 0.67; 2.07: 1.46
| Analytical results for Fe in OREAS 922b (Certified Value 3.1 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.13 | 3.1484 | 3.16 | 3.09 | 3.12 | 3.1228 | 3.15 | 3.01 | 3.09 | 3.03 | 3 | 3.11 | 3.06 | 2.6759 | 3.13 | 3.19 | 3.19 | 3.0306 | 3.19 | 3.22 | 2.98 | 2.99 | 3.03 | 3.08 | 3.04 | 1
| 1 | 2 | 3.19 | 3.1428 | 3.14 | 3.13 | 2.99 | 3.135 | 3.17 | 2.98 | 3.12 | 3.03 | 3.08 | 3.29 | 3.14 | 2.7199 | 3.15 | 3.15 | 2.93 | 3.0341 | 3.16 | 3.23 | 2.97 | 3.09 | 3.1 | 3 | 3.1 | 0
| 1 | 3 | 3.1 | 3.1251 | 3.1 | 3.06 | 3.1 | 3.1377 | 3.05 | 3.05 | 3.02 | 3.03 | 2.99 | 3.14 | 3.06 | 3.0056 | 3.2 | 3.11 | 3.09 | 2.947 | 3.13 | 3.14 | 2.98 | 3.02 | 3.07 | 3.13 | 3.1 | 16
| 1 | 4 | 3.17 | 3.1758 | 3.22 | 3.08 | 3.04 | 3.1384 | 3.17 | 2.96 | 3.14 | 3.09 | 3.08 | 3.2 | 3.09 | 2.6008 | 2.57 | 3.13 | 2.94 | 3.0091 | 3.23 | 3.27 | 2.99 | 3.07 | 3 | 3.08 | 3.13 | 3.0965
| 1 | 5 | 3.14 | 3.1334 | 3.2 | 3.14 | 3.09 | 3.1378 | 3.17 | 3.01 | 3.1 | 3.01 | 3.04 | 3.32 | 3.17 | 2.8948 | 2.14 | 3.14 | 3.17 | 2.9626 | 3.17 | 3.18 | 2.98 | 3 | 3.07 | 3.05 | 3.12 | 23
| 1 | 6 | 3.19 | 3.1324 | 3.08 | 3.03 | 2.97 | 3.128 | 3.11 | 3.06 | 3.09 | 3.03 | 3.05 | 3.14 | 3.16 | 3.1452 | 3.07 | 3.13 | 3.03 | 2.9765 | 3.11 | 3.22 | 2.98 | 3.05 | 3.11 | 3.02 | 3.12
| Mean |  | 3.1533 | 3.143 | 3.15 | 3.0883 | 3.0517 | 3.1333 | 3.1367 | 3.0117 | 3.0933 | 3.0367 | 3.04 | 3.2 | 3.1133 | 2.8404 | 2.8767 | 3.1417 | 3.0583 | 2.9933 | 3.165 | 3.21 | 2.98 | 3.0367 | 3.0633 | 3.06 | 3.1017
| Median |  | 3.155 | 3.1381 | 3.15 | 3.085 | 3.065 | 3.1363 | 3.16 | 3.01 | 3.095 | 3.03 | 3.045 | 3.17 | 3.115 | 2.8073 | 3.1 | 3.135 | 3.06 | 2.9928 | 3.165 | 3.22 | 2.98 | 3.035 | 3.07 | 3.065 | 3.11
| Std Dev. |  | 0.036148 | 0.018055 | 0.054772 | 0.041673 | 0.061779 | 0.0064047 | 0.048442 | 0.038687 | 0.040825 | 0.027325 | 0.038471 | 0.086948 | 0.049666 | 0.21079 | 0.42838 | 0.027142 | 0.11143 | 0.036535 | 0.042778 | 0.044721 | 0.0063246 | 0.039833 | 0.041793 | 0.046904 | 0.032506
| Rel.Std.Dev. |  | 1.146% | 0.5745% | 1.739% | 1.349% | 2.024% | 0.2044% | 1.544% | 1.285% | 1.32% | 0.8998% | 1.265% | 2.717% | 1.595% | 7.421% | 14.89% | 0.8639% | 3.643% | 1.221% | 1.352% | 1.393% | 0.2122% | 1.312% | 1.364% | 1.533% | 1.048%
| PDM3 |  | 1.835% | 1.501% | 1.728% | -0.2638% | -1.448% | 1.187% | 1.297% | -2.74% | -0.1024% | -1.932% | -1.825% | 3.342% | 0.5435% | -8.272% | -7.1% | 1.459% | -1.233% | -3.332% | 2.212% | 3.665% | -3.762% | -1.932% | -1.071% | -1.179% | 0.1668%
| Z-Score (Absolute) |  | 0.85 | 0.71 | 0.8 | 0 | 0.48 | 0.59 | 0.63 | 1 | 0.07 | 0.67 | 0.63 | 1.46 | 0.33 | 3.24 | 2.76 | 0.7 | 0.39 | 1.24 | 1 | 1.59 | 1.41 | 0.67 | 0.33 | 0.37 | 0.17
| Analytical results for Ga in OREAS 922b (Certified Value 18.8 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 19.8 | 22.4488 | 20.4 | 18.26 | 19.68 | 18.0832 | 19.85 | 17.65 | 18.4 | 17.7 | 18.8 | 14.4 | 18.67 | 19.0173 | 40 | 18.19 | 18.29 | 18 | 19.1 | 19.4 | 19.75 | 19.42 | 16.52 | 21 |  | 1
| 1 | 2 | 19.7 | 22.3287 | 20.6 | 18.4 | 18.62 | 18.9822 | 19.9 | 17.8 | 18.85 | 19.1 | 18.2 | 15.5 | 19.06 | 18.8776 | 40 | 18.29 | 17.66 | 18 | 19.5 | 20.5 | 18.95 | 19.51 | 16.67 | 21 |  | 18
| 1 | 3 | 19.3 | 22.4838 | 19.68 | 18.35 | 19.12 | 18.0477 | 16.9 | 18.55 | 18.35 | 17.95 | 19.45 | 15.4 | 17.8 | 20.5815 | 40 | 18.08 | 19.08 | 18 | 19.2 | 20.5 | 19.03 | 19.74 | 17.04 | 21 |  | 16
| 1 | 4 | 19.8 | 22.575 | 20.4 | 18.34 | 18.95 | 17.8852 | 19.45 | 17.8 | 19.05 | 17.9 | 18.6 | 16.4 | 18.59 | 19.3282 | 40 | 18.5 | 18.3 | 18 | 19.9 | 20.2 | 18.59 | 18.96 | 16.15 | 20 |  | 18.7734
| 1 | 5 | 19.4 | 22.2087 | 19.79 | 18.29 | 19.28 | 18.3366 | 18.25 | 18.3 | 18.2 | 17.15 | 19.45 | 17.8 | 19.17 | 20.2424 | 40 | 18.64 | 19.25 | 18 | 19.3 | 19.9 | 19.31 | 18.76 | 16.3 | 20 |  | 24
| 1 | 6 | 19.7 | 22.5225 | 19.88 | 18.21 | 18.52 | 18.2501 | 17.25 | 18.45 | 18.65 | 17.15 | 17.15 | 15.7 | 18.42 | 19.7869 | 40 | 19.02 | 18.61 | 18 | 19.1 | 20.1 | 18.95 | 19.54 | 17.38 | 20
| Mean |  | 19.6167 | 22.4279 | 20.125 | 18.3083 | 19.0283 | 18.2642 | 18.6 | 18.0917 | 18.5833 | 17.825 | 18.6083 | 15.8667 | 18.6183 | 19.639 | 40 | 18.4533 | 18.5317 | 18 | 19.35 | 20.1 | 19.0967 | 19.3217 | 16.6767 | 20.5
| Median |  | 19.7 | 22.4662 | 20.14 | 18.315 | 19.035 | 18.1667 | 18.85 | 18.05 | 18.525 | 17.8 | 18.7 | 15.6 | 18.63 | 19.5576 | 40 | 18.395 | 18.455 | 18 | 19.25 | 20.15 | 18.99 | 19.465 | 16.595 | 20.5
| Std Dev. |  | 0.2137 | 0.13565 | 0.38656 | 0.068532 | 0.43074 | 0.38582 | 1.3274 | 0.38654 | 0.32506 | 0.71746 | 0.86569 | 1.1448 | 0.49232 | 0.6837 | 0 | 0.34454 | 0.58218 | 0 | 0.30822 | 0.41473 | 0.39409 | 0.37791 | 0.46272 | 0.54772
| Rel.Std.Dev. |  | 1.089% | 0.6048% | 1.921% | 0.3743% | 2.264% | 2.112% | 7.137% | 2.137% | 1.749% | 4.025% | 4.652% | 7.215% | 2.644% | 3.481% | 0% | 1.867% | 3.142% | 0% | 1.593% | 2.063% | 2.064% | 1.956% | 2.775% | 2.672%
| PDM3 |  | 4.492% | 19.47% | 7.2% | -2.477% | 1.358% | -2.712% | -0.9236% | -3.631% | -1.012% | -5.052% | -0.8792% | -15.48% | -0.826% | 4.611% | 113.1% | -1.705% | -1.288% | -4.12% | 3.071% | 7.066% | 1.722% | 2.92% | -11.17% | 9.197%
| Z-Score (Absolute) |  | 1.1 | 4.18 | 1.66 | 0.33 | 0.46 | 0.38 | 0.01 | 0.57 | 0.03 | 0.86 | 0.01 | 3.01 | 0.01 | 1.12 | 23.45 | 0.18 | 0.09 | NA | 0.81 | 1.63 | 0.53 | 0.78 | 2.12 | NA
Results from laboratories 2.18 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 922b (Certified Value 4.87 ppm) | [26 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 6.3944; 2.02: 4.8; 2.03: 4.6196; 2.04: 5.2; 2.05: 4.82; 2.06: 4.8; 2.07: 4.68 |  | 1
| 1 | 2 | 2.01: 6.5324; 2.02: 4.8; 2.03: 4.844; 2.04: 5.4; 2.05: 4.89; 2.06: 4.4; 2.07: 4.85 |  | 6
| 1 | 3 | 2.01: 6.1997; 2.02: 4.8; 2.03: 4.7142; 2.04: 4.9; 2.05: 5.12; 2.06: 4.9; 2.07: 4.45 |  | 16
| 1 | 4 | 2.01: 6.6906; 2.02: 4.8; 2.03: 4.5667; 2.04: 5.6; 2.05: 5.04; 2.06: 4.5; 2.07: 4.69 |  | 4.8695
| 1 | 5 | 2.01: 6.6189; 2.02: 4.8; 2.03: 4.7746; 2.04: 5.3; 2.05: 4.95; 2.06: 5; 2.07: 4.6 |  | 25
| 1 | 6 | 2.01: 6.2405; 2.02: 4.8; 2.03: 4.7228; 2.04: 5.4; 2.05: 5; 2.06: 5.1; 2.07: 4.67
| Mean |  | 2.01: 6.4461; 2.02: 4.8; 2.03: 4.707; 2.04: 5.3; 2.05: 4.97; 2.06: 4.7833; 2.07: 4.6567
| Median |  | 2.01: 6.4634; 2.02: 4.8; 2.03: 4.7185; 2.04: 5.35; 2.05: 4.975; 2.06: 4.85; 2.07: 4.675
| Std Dev. |  | 2.01: 0.20137; 2.02: 0; 2.03: 0.10098; 2.04: 0.23664; 2.05: 0.10733; 2.06: 0.27869; 2.07: 0.13049
| Rel.Std.Dev. |  | 2.01: 3.124%; 2.02: 0%; 2.03: 2.145%; 2.04: 4.465%; 2.05: 2.16%; 2.06: 5.826%; 2.07: 2.802%
| PDM3 |  | 2.01: 32.38%; 2.02: -1.427%; 2.03: -3.338%; 2.04: 8.841%; 2.05: 2.064%; 2.06: -1.769%; 2.07: -4.371%
| Z-Score (Absolute) |  | 2.01: 7.74; 2.02: 0; 2.03: 0.44; 2.04: 2.35; 2.05: 0.8; 2.06: 0.08; 2.07: 0.67
| Analytical results for Ge in OREAS 922b (Certified Value 0.14 ppm) | [26 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.23 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 1.8 | 2.03 | 1.9 | < 0.2 | 0.15 | 0.14 | 0.13 | 0.15 | 0.17 | 0.2 | 0.12466 | 0.34 | 2.1 | 0.11 | 2 | < 0.05 | 4 | [8 empty] | 1
| 1 | 2 | 1.8 | 2.0185 | 1.7 | < 0.2 | 0.17 | 0.18 | 0.15 | 0.15 | 0.14 | 0.1 | 0.13794 | 0.39 | 2.2 | 0.13 | 1.6 | < 0.05 | 4.2 | [8 empty] | 41
| 1 | 3 | 1.8 | 2.023 | 1.7 | < 0.2 | 0.13 | 0.16 | 0.13 | 0.15 | 0.16 | < 0.1 | 0.13826 | 0.36 | 2.4 | 0.12 | 1.8 | 0.05 | 4.1 | [8 empty] | 16
| 1 | 4 | 1.9 | 2.005 | 1.8 | < 0.2 | 0.15 | 0.15 | 0.15 | 0.15 | 0.13 | < 0.1 | 0.11346 | 0.29 | 2.1 | 0.1 | 1.8 | < 0.05 | 4.3 | [8 empty] | 0.14193
| 1 | 5 | 1.8 | 2.0324 | 1.7 | < 0.2 | 0.15 | 0.2 | 0.14 | 0.14 | 0.17 | 0.1 | 0.11952 | 0.25 | 2.2 | 0.16 | 1.8 | 0.05 | 4.1 | [8 empty] | 26
| 1 | 6 | 1.9 | 2.02 | 1.8 | < 0.2 | 0.11 | 0.15 | 0.15 | 0.14 | 0.1 | < 0.1 | 0.10739 | 0.35 | 2.3 | 0.16 | 1.6 | < 0.05 | 4.1
| Mean |  | 1.8333 | 2.0215 | 1.7667 |  | 0.14333 | 0.16333 | 0.14167 | 0.14667 | 0.145 | 0.13333 | 0.12354 | 0.33 | 2.2167 | 0.13 | 1.7667 | 0.05 | 4.1333
| Median |  | 1.8 | 2.0215 | 1.75 |  | 0.15 | 0.155 | 0.145 | 0.15 | 0.15 | 0.1 | 0.12209 | 0.345 | 2.2 | 0.125 | 1.8 | 0.05 | 4.1
| Std Dev. |  | 0.05164 | 0.0097653 | 0.08165 |  | 0.020656 | 0.022509 | 0.0098319 | 0.005164 | 0.027386 | 0.057735 | 0.012681 | 0.05099 | 0.1169 | 0.025298 | 0.15055 | 0 | 0.10328
| Rel.Std.Dev. |  | 2.817% | 0.4831% | 4.622% |  | 14.41% | 13.78% | 6.94% | 3.521% | 18.89% | 43.3% | 10.26% | 15.45% | 5.274% | 19.46% | 8.522% | 0% | 2.499%
| PDM3 |  | 1192% | 1324% | 1145% |  | 0.9858% | 15.08% | -0.1885% | 3.334% | 2.16% | -6.06% | -12.96% | 132.5% | 1462% | -8.408% | 1145% | -64.77% | 2812%
| Z-Score (Absolute) |  | NA | 63.98 | NA | 1.48 | 0 | 0.68 | 0.06 | 0.11 | 0.06 | NA | 0.67 | 6.36 | NA | 0.45 | NA | 3.75 | NA
Results from laboratories 2.01, 2.05, 2.12, 2.16, 2.23 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 922b (Certified Value 2.52 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 2.4 | 2.8 | < 5 | 2.59 | 2.447 | 2.7 | 2.4 | 2.2 | 2.3 | 2.5 | 2.7 | 2.45 | 2.4373 | 2.4 | 2.5 | 2.49 | 2.4 | 2.49 | 2.77 | 2.54 | 2.6 | 2.6 |  |  |  | 1
| 1 | 2 | 2.4 | 2.8 | < 5 | 2.39 | 2.4584 | 2.7 | 2.5 | 2.3 | 2.4 | 2.5 | 2.7 | 2.45 | 2.4821 | 2.3 | 2.56 | 2.45 | 2.4 | 2.6 | 2.54 | 2.72 | 2.6 | 2.5 |  |  |  | 25
| 1 | 3 | 2.6 | 2.8 | < 5 | 2.42 | 2.3663 | 2.7 | 2.5 | 2.2 | 2.4 | 2.6 | 2.6 | 2.39 | 2.8502 | 2.4 | 2.47 | 2.67 | 2.3 | 2.65 | 2.51 | 2.55 | 2.6 | 2.6 |  |  |  | 16
| 1 | 4 | 2.5 | 2.8 | < 5 | 2.52 | 2.3732 | 2.7 | 2.4 | 2.4 | 2.4 | 2.5 | 2.6 | 2.52 | 2.8545 | 1.9 | 2.61 | 2.56 | 2.5 | 2.51 | 2.38 | 2.45 | 2.6 | 2.4 |  |  |  | 2.5163
| 1 | 5 | 2.4 | 2.8 | < 5 | 2.39 | 2.4949 | 2.8 | 2.5 | 2.1 | 2.3 | 2.5 | 2.7 | 2.51 | 2.6887 | 1.5 | 2.93 | 2.93 | 2.4 | 2.71 | 2.53 | 2.5 | 2.7 | 2.5 |  |  |  | 27
| 1 | 6 | 2.4 | 2.8 | < 5 | 2.34 | 2.4281 | 2.6 | 2.5 | 2.2 | 2.4 | 2.4 | 2.7 | 2.41 | 2.7392 | 2.4 | 2.57 | 2.75 | 2.3 | 2.61 | 2.55 | 2.48 | 2.6 | 2.4
| Mean |  | 2.45 | 2.8 |  | 2.4417 | 2.428 | 2.7 | 2.4667 | 2.2333 | 2.3667 | 2.5 | 2.6667 | 2.455 | 2.6753 | 2.15 | 2.6067 | 2.6417 | 2.3833 | 2.595 | 2.5467 | 2.54 | 2.6167 | 2.5
| Median |  | 2.4 | 2.8 |  | 2.405 | 2.4375 | 2.7 | 2.5 | 2.2 | 2.4 | 2.5 | 2.7 | 2.45 | 2.714 | 2.35 | 2.565 | 2.615 | 2.4 | 2.605 | 2.535 | 2.52 | 2.6 | 2.5
| Std Dev. |  | 0.083666 | 4.8648e-16 |  | 0.09411 | 0.050135 | 0.063246 | 0.05164 | 0.10328 | 0.05164 | 0.063246 | 0.05164 | 0.052058 | 0.17939 | 0.37283 | 0.16621 | 0.18005 | 0.075277 | 0.083367 | 0.12596 | 0.095708 | 0.040825 | 0.089443
| Rel.Std.Dev. |  | 3.415% | 1.737e-14% |  | 3.854% | 2.065% | 2.342% | 2.094% | 4.624% | 2.182% | 2.53% | 1.936% | 2.12% | 6.705% | 17.34% | 6.376% | 6.816% | 3.158% | 3.213% | 4.946% | 3.768% | 1.56% | 3.578%
| PDM3 |  | -2.636% | 11.27% |  | -2.967% | -3.511% | 7.299% | -1.974% | -11.25% | -5.948% | -0.6491% | 5.974% | -2.437% | 6.319% | -14.56% | 3.59% | 4.981% | -5.285% | 3.126% | 1.205% | 0.9405% | 3.987% | -0.6491%
| Z-Score (Absolute) |  | 0.33 | 2.01 | 0 | 0.39 | 0.48 | 1.34 | 0.22 | 1.78 | 0.89 | 0 | 1.11 | 0.3 | 1.17 | 2.34 | 0.71 | 0.95 | 0.78 | 0.64 | 0.31 | 0.27 | 0.78 | 0
| Analytical results for Hg in OREAS 922b (Indicative Value 0.04 ppm) | [26 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.03: Lab |  |  |  | 1
| No | No | 2.01: 2.06; 2.03: 2.18 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS |  |  |  | 3
3
| 1 | 1 | 2.01: < 1; 2.03: 0.04 |  |  |  | 1
| 1 | 2 | 2.01: < 1; 2.03: 0.04 |  |  |  | 23
| 1 | 3 | 2.01: < 1; 2.03: 0.04 |  |  |  | 16
| 1 | 4 | 2.01: < 1; 2.03: 0.04 |  |  |  | 0.04
| 1 | 5 | 2.01: < 1; 2.03: 0.04 |  |  |  | 7
| 1 | 6 | 2.01: < 1; 2.03: 0.04
| Mean |  | 2.03: 0.04
| Median |  | 2.03: 0.04
| Std Dev. |  | 2.03: 0
Rel.Std.Dev. | bar: 2.03
PDM3 | bar: 2.03
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.67
| Analytical results for Ho in OREAS 922b (Certified Value 0.6 ppm) | [26 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab |  |  |  | 1
| No | No | 2.01: 2.02; 2.03: 2.03; 2.04: 2.06; 2.05: 2.12; 2.06: 2.16; 2.07: 2.17; 2.08: 2.20 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.01: 0.65806; 2.03: 0.63; 2.04: 0.60846; 2.05: 0.7; 2.06: 0.62; 2.07: 0.6; 2.08: 0.51 |  |  |  | 1
| 1 | 2 | 2.01: 0.67106; 2.03: 0.623; 2.04: 0.62056; 2.05: 0.7; 2.06: 0.59; 2.07: 0.5; 2.08: 0.51 |  |  |  | 3
| 1 | 3 | 2.01: 0.63973; 2.03: 0.625; 2.04: 0.59748; 2.05: 0.8; 2.06: 0.63; 2.07: 0.6; 2.08: 0.46 |  |  |  | 16
| 1 | 4 | 2.01: 0.64867; 2.03: 0.622; 2.04: 0.57839; 2.05: 0.7; 2.06: 0.62; 2.07: 0.5; 2.08: 0.49 |  |  |  | 0.59956
| 1 | 5 | 2.01: 0.63934; 2.03: 0.628; 2.04: 0.61943; 2.05: 0.8; 2.06: 0.62; 2.07: 0.6; 2.08: 0.49 |  |  |  | 28
| 1 | 6 | 2.01: 0.66053; 2.03: 0.62; 2.04: 0.60318; 2.05: 0.7; 2.06: 0.63; 2.07: 0.6; 2.08: 0.49
| Mean |  | 2.01: 0.6529; 2.03: 0.62467; 2.04: 0.60458; 2.05: 0.73333; 2.06: 0.61833; 2.07: 0.56667; 2.08: 0.49167
| Median |  | 2.01: 0.65336; 2.03: 0.624; 2.04: 0.60582; 2.05: 0.7; 2.06: 0.62; 2.07: 0.6; 2.08: 0.49
| Std Dev. |  | 2.01: 0.012569; 2.03: 0.0037771; 2.04: 0.015679; 2.05: 0.05164; 2.06: 0.01472; 2.07: 0.05164; 2.08: 0.018348
| Rel.Std.Dev. |  | 2.01: 1.925%; 2.03: 0.6047%; 2.04: 2.593%; 2.05: 7.042%; 2.06: 2.381%; 2.07: 9.113%; 2.08: 3.732%
| PDM3 |  | 2.01: 8.896%; 2.03: 4.187%; 2.04: 0.8375%; 2.05: 22.31%; 2.06: 3.131%; 2.07: -5.487%; 2.08: -18%
| Z-Score (Absolute) |  | 2.01: 1.7; 2.03: 0.31; 2.04: 0.67; 2.05: NA; 2.06: 0; 2.07: NA; 2.08: 6.21
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 922b (Certified Value 0.25 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.26 | 0.3206 | 0.3 | < 1 | 0.25 | 0.2375 | 0.248 | 0.24 | 0.239 | 0.23 | 0.265 | 0.2 | 0.246 | 0.19094 | 0.239 | 0.23 | 0.24 | 0.25 | 0.29 | 0.24 | 0.28 | 0.26 | 0.237 | 0.27 |  | 1
| 1 | 2 | 0.26 | 0.31794 | 0.3 | < 1 | 0.26 | 0.24162 | 0.252 | 0.246 | 0.237 | 0.243 | 0.259 | 0.3 | 0.244 | 0.21706 | 0.255 | 0.26 | 0.23 | 0.25 | 0.3 | 0.25 | 0.32 | 0.27 | 0.242 | 0.28 |  | 11
| 1 | 3 | 0.25 | 0.30282 | 0.3 | < 1 | 0.24 | 0.23003 | 0.25 | 0.257 | 0.229 | 0.236 | 0.273 | 0.2 | 0.242 | 0.23416 | 0.244 | 0.26 | 0.24 | 0.25 | 0.27 | 0.26 | 0.24 | 0.25 | 0.246 | 0.26 |  | 16
| 1 | 4 | 0.26 | 0.34524 | 0.3 | < 1 | 0.25 | 0.23941 | 0.255 | 0.244 | 0.243 | 0.235 | 0.249 | 0.2 | 0.238 | 0.22969 | 0.213 | 0.27 | 0.23 | 0.26 | 0.28 | 0.25 | 0.25 | 0.26 | 0.236 | 0.26 |  | 0.24803
| 1 | 5 | 0.26 | 0.30632 | 0.3 | < 1 | 0.25 | 0.24363 | 0.259 | 0.25 | 0.228 | 0.228 | 0.275 | 0.2 | 0.24 | 0.23378 | 0.176 | 0.25 | 0.26 | 0.25 | 0.3 | 0.25 | 0.25 | 0.24 | 0.234 | 0.27 |  | 29
| 1 | 6 | 0.26 | 0.27076 | 0.3 | < 1 | 0.25 | 0.24291 | 0.261 | 0.249 | 0.225 | 0.229 | 0.254 | 0.2 | 0.241 | 0.22086 | 0.259 | 0.26 | 0.25 | 0.24 | 0.29 | 0.25 | 0.28 | 0.24 | 0.252 | 0.26
| Mean |  | 0.25833 | 0.31061 | 0.3 |  | 0.25 | 0.23918 | 0.25417 | 0.24767 | 0.2335 | 0.2335 | 0.2625 | 0.21667 | 0.24183 | 0.22108 | 0.231 | 0.255 | 0.24167 | 0.25 | 0.28833 | 0.25 | 0.27 | 0.25333 | 0.24117 | 0.26667
| Median |  | 0.26 | 0.31213 | 0.3 |  | 0.25 | 0.24051 | 0.2535 | 0.2475 | 0.233 | 0.2325 | 0.262 | 0.2 | 0.2415 | 0.22528 | 0.2415 | 0.26 | 0.24 | 0.25 | 0.29 | 0.25 | 0.265 | 0.255 | 0.2395 | 0.265
| Std Dev. |  | 0.0040825 | 0.024576 | 0 |  | 0.0063246 | 0.0050255 | 0.0051153 | 0.0058195 | 0.0071484 | 0.0056833 | 0.010387 | 0.040825 | 0.0028577 | 0.016316 | 0.031439 | 0.013784 | 0.01169 | 0.0063246 | 0.01169 | 0.0063246 | 0.029665 | 0.012111 | 0.0068823 | 0.008165
| Rel.Std.Dev. |  | 1.58% | 7.912% | 0% |  | 2.53% | 2.101% | 2.013% | 2.35% | 3.061% | 2.434% | 3.957% | 18.84% | 1.182% | 7.38% | 13.61% | 5.406% | 4.837% | 2.53% | 4.054% | 2.53% | 10.99% | 4.781% | 2.854% | 3.062%
| PDM3 |  | 4.153% | 25.23% | 20.95% |  | 0.7931% | -3.569% | 2.473% | -0.1477% | -5.859% | -5.859% | 5.833% | -12.65% | -2.499% | -10.87% | -6.867% | 2.809% | -2.567% | 0.7931% | 16.25% | 0.7931% | 8.857% | 2.137% | -2.768% | 7.513%
| Z-Score (Absolute) |  | 0.57 | 4.16 | NA | 17.16 | 0 | 0.74 | 0.29 | 0.16 | 1.13 | 1.13 | 0.86 | NA | 0.56 | 1.98 | 1.3 | 0.34 | 0.57 | 0 | 2.63 | 0 | 1.37 | 0.23 | 0.61 | 1.14
Results from laboratory 2.03 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for K in OREAS 922b (Certified Value 2.59 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.58 | 2.6438 | 2.62 | 2.34 | 2.572 | 2.5275 | 2.7 | 2.56 | 2.66 | 2.59 | 2.54 | 2.32 | 2.5 | 2.0379 | 2.74 | 2.6255 | 2.74 | 2.82 | 2.7 | 2.45 | 2.57 | 2.5193 | 2.58 | 2.63 |  | 1
| 1 | 2 | 2.6 | 2.6595 | 2.58 | 2.3 | 2.4373 | 2.5274 | 2.7 | 2.54 | 2.68 | 2.6 | 2.61 | 2.38 | 2.54 | 1.7925 | 2.74 | 2.5843 | 2.18 | 2.81 | 2.7 | 2.46 | 2.65 | 2.5456 | 2.51 | 2.58 |  | 0
| 1 | 3 | 2.55 | 2.631 | 2.54 | 2.39 | 2.498 | 2.5483 | 2.63 | 2.61 | 2.6 | 2.6 | 2.53 | 2.14 | 2.49 | 2.1794 | 2.79 | 2.5626 | 2.13 | 2.77 | 2.68 | 2.45 | 2.58 | 2.5271 | 2.62 | 2.61 |  | 16
| 1 | 4 | 2.61 | 2.6477 | 2.64 | 2.38 | 2.4918 | 2.5094 | 2.7 | 2.56 | 2.71 | 2.63 | 2.62 | 2.09 | 2.53 | 1.9874 | 2.22 | 2.5797 | 2.52 | 2.87 | 2.77 | 2.45 | 2.61 | 2.4728 | 2.58 | 2.56 |  | 2.5897
| 1 | 5 | 2.57 | 2.6629 | 2.66 | 2.37 | 2.5045 | 2.5273 | 2.74 | 2.55 | 2.65 | 2.58 | 2.57 | 2.3 | 2.59 | 1.6916 | 1.82 | 2.5911 | 2.07 | 2.81 | 2.69 | 2.45 | 2.56 | 2.4768 | 2.57 | 2.61 |  | 30
| 1 | 6 | 2.61 | 2.6214 | 2.56 | 2.39 | 2.4451 | 2.5281 | 2.64 | 2.62 | 2.7 | 2.6 | 2.6 | 2.05 | 2.46 | 1.7284 | 2.68 | 2.5781 | 2.31 | 2.76 | 2.68 | 2.45 | 2.58 | 2.5186 | 2.54 | 2.58
| Mean |  | 2.5867 | 2.6444 | 2.6 | 2.3617 | 2.4914 | 2.528 | 2.685 | 2.5733 | 2.6667 | 2.6 | 2.5783 | 2.2133 | 2.5183 | 1.9029 | 2.4983 | 2.5869 | 2.325 | 2.8067 | 2.7033 | 2.4517 | 2.5917 | 2.51 | 2.5667 | 2.595
| Median |  | 2.59 | 2.6458 | 2.6 | 2.375 | 2.4949 | 2.5274 | 2.7 | 2.56 | 2.67 | 2.6 | 2.585 | 2.22 | 2.515 | 1.89 | 2.71 | 2.582 | 2.245 | 2.81 | 2.695 | 2.45 | 2.58 | 2.5189 | 2.575 | 2.595
| Std Dev. |  | 0.024221 | 0.016063 | 0.047329 | 0.035449 | 0.048543 | 0.012303 | 0.041833 | 0.033267 | 0.039833 | 0.016733 | 0.037639 | 0.13706 | 0.04535 | 0.19449 | 0.39306 | 0.021138 | 0.25867 | 0.039328 | 0.033862 | 0.0040825 | 0.033116 | 0.029004 | 0.037771 | 0.025884
| Rel.Std.Dev. |  | 0.9364% | 0.6074% | 1.82% | 1.501% | 1.948% | 0.4867% | 1.558% | 1.293% | 1.494% | 0.6436% | 1.46% | 6.193% | 1.801% | 10.22% | 15.73% | 0.8171% | 11.13% | 1.401% | 1.253% | 0.1665% | 1.278% | 1.156% | 1.472% | 0.9975%
| PDM3 |  | -0.1163% | 2.112% | 0.3985% | -8.805% | -3.793% | -2.381% | 3.681% | -0.6312% | 2.973% | 0.3985% | -0.4381% | -14.53% | -2.755% | -26.52% | -3.527% | -0.108% | -10.22% | 8.379% | 4.389% | -5.329% | 0.07676% | -3.076% | -0.8886% | 0.2055%
| Z-Score (Absolute) |  | 0.11 | 0.69 | 0.24 | 2.15 | 0.85 | 0.48 | 1.1 | 0.03 | 0.91 | 0.24 | 0.03 | 3.64 | 0.58 | 6.76 | 0.78 | 0.11 | 2.52 | 2.32 | 1.28 | 1.25 | 0.16 | 0.66 | 0.09 | 0.19
| Analytical results for La in OREAS 922b (Certified Value 42.5 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 39.5 | 40.4 | 43.07 | 41.36 | 41.5794 | 43.9 | 41.5 | 42.4 | 41.8 | 44.2 | 45.3 | 43.04 | 44.1 | 41.67 | 40.2 | 46.8 | 46.5 | 43.2 | 40.1 | 41.55 | 41.3 | 46 |  |  |  | 1
| 1 | 2 | 40.83 | 40.5 | 44.24 | 40.23 | 41.9928 | 43.8 | 39.9 | 43.4 | 42.9 | 41.5 | 45.7 | 39.41 | 44.5 | 41.22 | 37.3 | 46.2 | 46.5 | 43.3 | 40.6 | 42.33 | 41.5 | 47 |  |  |  | 12
| 1 | 3 | 39.6 | 40.5 | 44.31 | 40.8 | 41.8356 | 43.2 | 40.8 | 42.1 | 43.3 | 45.7 | 45.3 | 38.66 | 45.1 | 40.58 | 41.2 | 45.2 | 44.3 | 45.1 | 41.2 | 42.02 | 42.2 | 46 |  |  |  | 16
| 1 | 4 | 39.6 | 40.5 | 42.25 | 39.75 | 40.4353 | 43.1 | 39.6 | 44.2 | 43.2 | 44.6 | 46 | 40.08 | 36.2 | 41.94 | 38 | 46.6 | 47.7 | 44.8 | 40.4 | 40.73 | 39.8 | 45 |  |  |  | 42.5453
| 1 | 5 | 40.64 | 40.5 | 44.11 | 40.46 | 41.7261 | 46.4 | 40.6 | 42.7 | 41.5 | 44.8 | 47.2 | 41.32 | 29.5 | 41.2 | 40.9 | 47 | 44.7 | 45.2 | 40.3 | 41.13 | 40.5 | 47 |  |  |  | 31
| 1 | 6 | 39.92 | 40.6 | 43.72 | 40.42 | 41.8032 | 45.4 | 42.3 | 43 | 41.5 | 40.9 | 45.5 | 39.98 | 43.4 | 42.44 | 40.5 | 44.5 | 46.9 | 44.6 | 40.8 | 41.33 | 42.4 | 44
| Mean |  | 40.015 | 40.5 | 43.6167 | 40.5033 | 41.5621 | 44.3 | 40.7833 | 42.9667 | 42.3667 | 43.6167 | 45.8333 | 40.415 | 40.4667 | 41.5083 | 39.6833 | 46.05 | 46.1 | 44.3667 | 40.5667 | 41.515 | 41.2833 | 45.8333
| Median |  | 39.76 | 40.5 | 43.915 | 40.44 | 41.7646 | 43.85 | 40.7 | 42.85 | 42.35 | 44.4 | 45.6 | 40.03 | 43.75 | 41.445 | 40.35 | 46.4 | 46.5 | 44.7 | 40.5 | 41.44 | 41.4 | 46
| Std Dev. |  | 0.57854 | 0.063246 | 0.81116 | 0.54276 | 0.56837 | 1.3176 | 1.0028 | 0.75542 | 0.85713 | 1.9447 | 0.72019 | 1.5557 | 6.2931 | 0.65123 | 1.6266 | 0.99146 | 1.3206 | 0.89144 | 0.39328 | 0.58637 | 0.99482 | 1.169
| Rel.Std.Dev. |  | 1.446% | 0.1562% | 1.86% | 1.34% | 1.368% | 2.974% | 2.459% | 1.758% | 2.023% | 4.459% | 1.571% | 3.849% | 15.55% | 1.569% | 4.099% | 2.153% | 2.865% | 2.009% | 0.9695% | 1.412% | 2.41% | 2.551%
| PDM3 |  | -5.947% | -4.807% | 2.518% | -4.8% | -2.311% | 4.124% | -4.141% | 0.9903% | -0.42% | 2.518% | 7.728% | -5.007% | -4.886% | -2.437% | -6.727% | 8.237% | 8.355% | 4.281% | -4.651% | -2.422% | -2.966% | 7.728%
| Z-Score (Absolute) |  | 0.81 | 0.55 | 1.1 | 0.55 | 0.01 | 1.46 | 0.4 | 0.75 | 0.44 | 1.1 | 2.27 | 0.59 | 0.57 | 0.02 | 0.98 | 2.38 | 2.41 | 1.49 | 0.51 | 0.01 | 0.13 | 2.27
| Analytical results for Li in OREAS 922b (Certified Value 23.9 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 24.1 | 30.91 | 24.2 | 23.02 | 23.4 | 22.9586 | 24.8 | 23.1 | 24.3 | 24.3 | 23.5 | 24 | 23 | 22.0869 | 24.8 | 24.4 | 24.7 | 25.5 | 20.8 | 23 | 25.5 | 22.8 | 24 | 22.3 |  | 1
| 1 | 2 | 24.3 | 30.71 | 24.3 | 23.13 | 24.4 | 23.0124 | 24.9 | 23.1 | 24.3 | 24.3 | 24.1 | 25 | 23.8 | 21.575 | 23.7 | 23.7 | 23.7 | 25.6 | 21.2 | 25 | 24.1 | 22.9 | 24.4 | 23.4 |  | 0
| 1 | 3 | 24 | 31.08 | 24.3 | 22.97 | 24.7 | 22.9007 | 23.8 | 23.6 | 23.6 | 24.3 | 23.2 | 24 | 23.1 | 22.8973 | 23.4 | 23.7 | 25.7 | 25 | 20.7 | 25 | 24 | 22.9 | 24.8 | 22.6 |  | 16
| 1 | 4 | 23.7 | 30.32 | 24.6 | 22.7 | 24.3 | 22.582 | 24.8 | 23.1 | 24.5 | 24.6 | 24.1 | 24 | 23.4 | 22.7164 | 18.9 | 24.4 | 24.3 | 25.4 | 22 | 25 | 23.6 | 22.7 | 23.5 | 22.4 |  | 23.854
| 1 | 5 | 24.1 | 31.69 | 24.2 | 23.83 | 23.6 | 23.1996 | 24.6 | 23.4 | 24.3 | 23.9 | 23.8 | 24 | 23.8 | 21.896 | 16.6 | 24 | 25.6 | 25.6 | 21.1 | 24 | 24.3 | 23.3 | 24 | 23.2 |  | 32
| 1 | 6 | 24 | 30.42 | 24.2 | 23.45 | 23.5 | 23.5632 | 24.1 | 23.6 | 24.2 | 24.3 | 23.8 | 24 | 23.7 | 22.8307 | 22.6 | 24.8 | 24.7 | 24.7 | 21 | 22 | 23.9 | 23 | 25.2 | 23.5
| Mean |  | 24.0333 | 30.855 | 24.3 | 23.1833 | 23.9833 | 23.0361 | 24.5 | 23.3167 | 24.2 | 24.2833 | 23.75 | 24.1667 | 23.4667 | 22.3337 | 21.6667 | 24.1667 | 24.7833 | 25.3 | 21.1333 | 24 | 24.2333 | 22.9333 | 24.3167 | 22.9
| Median |  | 24.05 | 30.81 | 24.25 | 23.075 | 23.95 | 22.9855 | 24.7 | 23.25 | 24.3 | 24.3 | 23.8 | 24 | 23.55 | 22.4017 | 23 | 24.2 | 24.7 | 25.45 | 21.05 | 24.5 | 24.05 | 22.9 | 24.2 | 22.9
| Std Dev. |  | 0.19664 | 0.49947 | 0.15492 | 0.39948 | 0.54924 | 0.3273 | 0.44721 | 0.24833 | 0.30984 | 0.22286 | 0.35071 | 0.40825 | 0.3559 | 0.55484 | 3.1985 | 0.44121 | 0.76529 | 0.36878 | 0.46332 | 1.2649 | 0.66232 | 0.20656 | 0.61455 | 0.52915
| Rel.Std.Dev. |  | 0.8182% | 1.619% | 0.6375% | 1.723% | 2.29% | 1.421% | 1.825% | 1.065% | 1.28% | 0.9177% | 1.477% | 1.689% | 1.517% | 2.484% | 14.76% | 1.826% | 3.088% | 1.458% | 2.192% | 5.27% | 2.733% | 0.9007% | 2.527% | 2.311%
| PDM3 |  | 0.752% | 29.35% | 1.87% | -2.811% | 0.5424% | -3.429% | 2.708% | -2.252% | 1.451% | 1.8% | -0.4358% | 1.311% | -1.624% | -6.373% | -9.17% | 1.311% | 3.896% | 6.062% | -11.41% | 0.6122% | 1.59% | -3.859% | 1.94% | -3.999%
| Z-Score (Absolute) |  | 0.02 | 8.92 | 0.37 | 1.09 | 0.04 | 1.28 | 0.63 | 0.91 | 0.24 | 0.35 | 0.35 | 0.2 | 0.72 | 2.2 | 3.07 | 0.2 | 1 | 1.67 | 3.76 | 0.02 | 0.28 | 1.41 | 0.39 | 1.46
| Analytical results for Lu in OREAS 922b (Certified Value 0.27 ppm) | [26 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.16; 2.08: 2.17; 2.09: 2.20; 2.10: 2.21 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*OES/MS |  | 2
3
| 1 | 1 | 2.01: 0.1929; 2.02: 0.262; 2.03: < 1; 2.04: 0.28702; 2.05: 0.3; 2.06: 0.27; 2.07: 0.27; 2.08: 0.2; 2.09: 0.22; 2.10: 0.27 |  | 1
| 1 | 2 | 2.01: 0.1973; 2.02: 0.256; 2.03: < 1; 2.04: 0.29234; 2.05: 0.3; 2.06: 0.28; 2.07: 0.27; 2.08: 0.2; 2.09: 0.22; 2.10: 0.28 |  | 13
| 1 | 3 | 2.01: 0.1885; 2.02: 0.252; 2.03: < 1; 2.04: 0.28595; 2.05: 0.3; 2.06: 0.26; 2.07: 0.28; 2.08: 0.3; 2.09: 0.21; 2.10: 0.28 |  | 16
| 1 | 4 | 2.01: 0.2038; 2.02: 0.258; 2.03: < 1; 2.04: 0.27293; 2.05: 0.3; 2.06: 0.27; 2.07: 0.28; 2.08: 0.2; 2.09: 0.23; 2.10: 0.27 |  | 0.27385
| 1 | 5 | 2.01: 0.1866; 2.02: 0.264; 2.03: < 1; 2.04: 0.28374; 2.05: 0.3; 2.06: 0.28; 2.07: 0.27; 2.08: 0.3; 2.09: 0.2; 2.10: 0.28 |  | 33
| 1 | 6 | 2.01: 0.185; 2.02: 0.26; 2.03: < 1; 2.04: 0.28714; 2.05: 0.4; 2.06: 0.28; 2.07: 0.28; 2.08: 0.3; 2.09: 0.2; 2.10: 0.27
| Mean |  | 2.01: 0.19235; 2.02: 0.25867; 2.04: 0.28486; 2.05: 0.31667; 2.06: 0.27333; 2.07: 0.275; 2.08: 0.25; 2.09: 0.21333; 2.10: 0.275
| Median |  | 2.01: 0.1907; 2.02: 0.259; 2.04: 0.28649; 2.05: 0.3; 2.06: 0.275; 2.07: 0.275; 2.08: 0.25; 2.09: 0.215; 2.10: 0.275
| Std Dev. |  | 2.01: 0.0071807; 2.02: 0.0043205; 2.04: 0.0064899; 2.05: 0.040825; 2.06: 0.008165; 2.07: 0.0054772; 2.08: 0.054772; 2.09: 0.012111; 2.10: 0.0054772
| Rel.Std.Dev. |  | 2.01: 3.733%; 2.02: 1.67%; 2.04: 2.278%; 2.05: 12.89%; 2.06: 2.987%; 2.07: 1.992%; 2.08: 21.91%; 2.09: 5.677%; 2.10: 1.992%
| PDM3 |  | 2.01: -29.76%; 2.02: -5.544%; 2.04: 4.02%; 2.05: 15.64%; 2.06: -0.188%; 2.07: 0.4207%; 2.08: -8.708%; 2.09: -22.1%; 2.10: 0.4207%
| Z-Score (Absolute) |  | 2.01: 4.21; 2.02: 0.8; 2.03: 11.63; 2.04: 0.55; 2.05: NA; 2.06: 0.04; 2.07: 0.04; 2.08: NA; 2.09: 3.13; 2.10: 0.04
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 922b (Certified Value 0.407 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.413 | 0.41435 | 0.412 | 0.41 | 0.4153 | 0.42282 | 0.41 | 0.38 | 0.4 | 0.39 | 0.4 | 0.34 | 0.395 | 0.35655 | 0.42 | 0.4291 | 0.41 | 0.4081 | 0.38 | 0.42 | 0.4 | 0.4 | 0.4178 | 0.42 | 0.41 | 1
| 1 | 2 | 0.412 | 0.41276 | 0.415 | 0.41 | 0.3939 | 0.4222 | 0.41 | 0.39 | 0.4 | 0.39 | 0.41 | 0.35 | 0.405 | 0.36211 | 0.42 | 0.4217 | 0.38 | 0.4088 | 0.39 | 0.41 | 0.4 | 0.41 | 0.4112 | 0.41 | 0.41 | 0
| 1 | 3 | 0.407 | 0.41262 | 0.415 | 0.42 | 0.4001 | 0.41981 | 0.4 | 0.39 | 0.39 | 0.39 | 0.4 | 0.33 | 0.393 | 0.40319 | 0.43 | 0.4197 | 0.41 | 0.3972 | 0.39 | 0.41 | 0.4 | 0.41 | 0.4147 | 0.43 | 0.41 | 16
| 1 | 4 | 0.42 | 0.41233 | 0.415 | 0.41 | 0.4042 | 0.42265 | 0.41 | 0.39 | 0.4 | 0.4 | 0.41 | 0.34 | 0.398 | 0.34817 | 0.34 | 0.4228 | 0.37 | 0.4047 | 0.4 | 0.43 | 0.4 | 0.41 | 0.4048 | 0.42 | 0.41 | 0.40712
| 1 | 5 | 0.411 | 0.41222 | 0.41 | 0.42 | 0.4022 | 0.42487 | 0.42 | 0.39 | 0.4 | 0.39 | 0.4 | 0.36 | 0.408 | 0.37339 | 0.28 | 0.4229 | 0.4 | 0.3978 | 0.39 | 0.41 | 0.4 | 0.4 | 0.4101 | 0.42 | 0.42 | 34
| 1 | 6 | 0.42 | 0.41385 | 0.41 | 0.41 | 0.3909 | 0.42339 | 0.41 | 0.39 | 0.4 | 0.39 | 0.41 | 0.33 | 0.407 | 0.39532 | 0.41 | 0.4202 | 0.39 | 0.4061 | 0.39 | 0.41 | 0.39 | 0.41 | 0.4188 | 0.41 | 0.41
| Mean |  | 0.41383 | 0.41302 | 0.41283 | 0.41333 | 0.4011 | 0.42262 | 0.41 | 0.38833 | 0.39833 | 0.39167 | 0.405 | 0.34167 | 0.401 | 0.37312 | 0.38333 | 0.42273 | 0.39333 | 0.40378 | 0.39 | 0.415 | 0.39833 | 0.40667 | 0.4129 | 0.41833 | 0.41167
| Median |  | 0.4125 | 0.41269 | 0.4135 | 0.41 | 0.40115 | 0.42274 | 0.41 | 0.39 | 0.4 | 0.39 | 0.405 | 0.34 | 0.4015 | 0.36775 | 0.415 | 0.42225 | 0.395 | 0.4054 | 0.39 | 0.41 | 0.4 | 0.41 | 0.41295 | 0.42 | 0.41
| Std Dev. |  | 0.0051929 | 0.00087403 | 0.0024833 | 0.005164 | 0.0085946 | 0.0016574 | 0.0063246 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0054772 | 0.01169 | 0.0064807 | 0.021977 | 0.060222 | 0.0033839 | 0.01633 | 0.0050823 | 0.0063246 | 0.0083666 | 0.0040825 | 0.005164 | 0.0052612 | 0.0075277 | 0.0040825
| Rel.Std.Dev. |  | 1.255% | 0.2116% | 0.6015% | 1.249% | 2.143% | 0.3922% | 1.543% | 1.051% | 1.025% | 1.042% | 1.352% | 3.422% | 1.616% | 5.89% | 15.71% | 0.8005% | 4.152% | 1.259% | 1.622% | 2.016% | 1.025% | 1.27% | 1.274% | 1.799% | 0.9917%
| PDM3 |  | 1.648% | 1.449% | 1.403% | 1.525% | -1.479% | 3.807% | 0.7067% | -4.615% | -2.159% | -3.796% | -0.5214% | -16.08% | -1.504% | -8.352% | -5.843% | 3.834% | -3.387% | -0.8203% | -4.206% | 1.935% | -2.159% | -0.1121% | 1.419% | 2.754% | 1.116%
| Z-Score (Absolute) |  | 0.71 | 0.65 | 0.63 | 0.67 | 0.32 | 1.43 | 0.4 | 1.35 | 0.54 | 1.08 | 0 | 5.12 | 0.32 | 2.58 | 1.75 | 1.43 | 0.94 | 0.1 | 1.21 | 0.81 | 0.54 | 0.13 | 0.64 | 1.08 | 0.54
| Analytical results for Mn in OREAS 922b (Certified Value 0.011 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0109 | 0.011313 | 0.0117 | 0.01093 | 0.0111 | 0.01075 | 0.0114 | 0.011 | 0.0112 | 0.0112 | 0.0111 | 0.0117 | 0.0111 | 0.011275 | 0.0123 | 0.0111 | 0.0115 | 0.0118 | 0.0121 | 0.0113 | 0.0107 | 0.0115 | 0.0116 | 0.011 | 0.0109 | 1
| 1 | 2 | 0.0112 | 0.011161 | 0.0117 | 0.01081 | 0.0106 | 0.010985 | 0.0113 | 0.0109 | 0.0112 | 0.0112 | 0.0114 | 0.0122 | 0.0114 | 0.011185 | 0.0119 | 0.0112 | 0.0106 | 0.012 | 0.012 | 0.0112 | 0.0106 | 0.011 | 0.0118 | 0.0107 | 0.0111 | 30
| 1 | 3 | 0.0109 | 0.011161 | 0.0116 | 0.01099 | 0.011 | 0.010771 | 0.0112 | 0.0112 | 0.0109 | 0.0113 | 0.0109 | 0.0121 | 0.0112 | 0.010876 | 0.0125 | 0.011 | 0.0113 | 0.0116 | 0.012 | 0.0111 | 0.0107 | 0.0111 | 0.0116 | 0.0112 | 0.0111 | 16
| 1 | 4 | 0.0111 | 0.01096 | 0.0116 | 0.01089 | 0.0108 | 0.011051 | 0.0114 | 0.0109 | 0.0114 | 0.0112 | 0.0115 | 0.012 | 0.0112 | 0.012203 | 0.0108 | 0.0111 | 0.0105 | 0.0117 | 0.0123 | 0.0115 | 0.0107 | 0.0109 | 0.0114 | 0.0111 | 0.0112 | 0.011145
| 1 | 5 | 0.0109 | 0.01154 | 0.0116 | 0.0112 | 0.011 | 0.011001 | 0.0115 | 0.011 | 0.0111 | 0.0111 | 0.0111 | 0.0127 | 0.0115 | 0.012073 | 0.0082 | 0.011 | 0.0114 | 0.0116 | 0.0119 | 0.0113 | 0.0106 | 0.0111 | 0.0115 | 0.0111 | 0.0112 | 35
| 1 | 6 | 0.0112 | 0.011049 | 0.0116 | 0.01117 | 0.0106 | 0.010818 | 0.0111 | 0.0112 | 0.0112 | 0.0111 | 0.0114 | 0.0115 | 0.0115 | 0.012102 | 0.0121 | 0.0111 | 0.0109 | 0.0118 | 0.0116 | 0.0114 | 0.0108 | 0.0111 | 0.0118 | 0.0109 | 0.011
| Mean |  | 0.011033 | 0.011198 | 0.011633 | 0.010998 | 0.01085 | 0.010896 | 0.011317 | 0.011033 | 0.011167 | 0.011183 | 0.011233 | 0.012033 | 0.011317 | 0.011619 | 0.0113 | 0.011083 | 0.011033 | 0.01175 | 0.011983 | 0.0113 | 0.010683 | 0.011117 | 0.011617 | 0.011 | 0.011083
| Median |  | 0.011 | 0.011161 | 0.0116 | 0.01096 | 0.0109 | 0.010901 | 0.01135 | 0.011 | 0.0112 | 0.0112 | 0.01125 | 0.01205 | 0.0113 | 0.011674 | 0.012 | 0.0111 | 0.0111 | 0.01175 | 0.012 | 0.0113 | 0.0107 | 0.0111 | 0.0116 | 0.01105 | 0.0111
| Std Dev. |  | 0.00015055 | 0.00020581 | 5.164e-05 | 0.00015626 | 0.00021679 | 0.00013116 | 0.0001472 | 0.00013663 | 0.0001633 | 7.5277e-05 | 0.00023381 | 0.00041793 | 0.00017224 | 0.0005727 | 0.001631 | 7.5277e-05 | 0.0004274 | 0.00015166 | 0.00023166 | 0.00014142 | 7.5277e-05 | 0.00020412 | 0.00016021 | 0.00017889 | 0.0001169
| Rel.Std.Dev. |  | 1.365% | 1.838% | 0.4439% | 1.421% | 1.998% | 1.204% | 1.301% | 1.238% | 1.462% | 0.6731% | 2.081% | 3.473% | 1.522% | 4.929% | 14.43% | 0.6792% | 3.874% | 1.291% | 1.933% | 1.252% | 0.7046% | 1.836% | 1.379% | 1.626% | 1.055%
| PDM3 |  | -1.004% | 0.4696% | 4.379% | -1.318% | -2.649% | -2.237% | 1.538% | -1.004% | 0.192% | 0.3415% | 0.7901% | 7.968% | 1.538% | 4.249% | 1.388% | -0.5557% | -1.004% | 5.426% | 7.519% | 1.388% | -4.145% | -0.2566% | 4.23% | -1.303% | -0.5557%
| Z-Score (Absolute) |  | 0.67 | 0.06 | 2.02 | 0.83 | 1.5 | 1.29 | 0.6 | 0.67 | 0.07 | 0 | 0.22 | 3.82 | 0.6 | 1.96 | 0.52 | 0.45 | 0.67 | 2.55 | 3.6 | 0.52 | 2.25 | 0.3 | 1.95 | 0.82 | 0.45
| Analytical results for Mo in OREAS 922b (Certified Value 5.01 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5 | 5.5 | 5.02 | 5.2 | 4.9006 | 5.53 | 4.73 | 4.66 | 4.87 | 5.27 | 6 | 5.15 | 4.6681 | 5.22 | 5.3 | 5.07 | 4.8 | 5 | 4.6 | 4.28 | 4.86 | 5.3 | 4.92 | 5 |  | 1
| 1 | 2 | 5 | 5.5 | 4.99 | 5 | 4.8801 | 5.37 | 4.82 | 4.83 | 5.07 | 5.07 | 6 | 5.23 | 4.7344 | 4.92 | 5.3 | 4.88 | 4.8 | 5.1 | 4.8 | 4.55 | 5.47 | 5.5 | 5.09 | 5.1 |  | 31
| 1 | 3 | 5 | 5.5 | 4.89 | 5.3 | 4.8829 | 5.08 | 4.88 | 4.58 | 5.08 | 5.34 | 6 | 5.32 | 4.7154 | 5.21 | 5 | 5 | 4.8 | 5 | 4.6 | 4.51 | 4.73 | 5.4 | 5.19 | 5 |  | 16
| 1 | 4 | 5 | 5.5 | 4.84 | 4.9 | 5.0227 | 5.26 | 4.56 | 4.91 | 5 | 5.3 | 6 | 5.3 | 4.8178 | 4.28 | 5.2 | 4.83 | 4.9 | 5.1 | 5.3 | 4.41 | 4.73 | 5.1 | 5.15 | 5.1 |  | 5.011
| 1 | 5 | 5.1 | 5.5 | 4.99 | 5.3 | 4.9032 | 5.29 | 4.68 | 4.63 | 4.88 | 5.46 | 6 | 5.15 | 4.8027 | 3.52 | 5.3 | 5.16 | 4.7 | 5 | 5.2 | 4.57 | 4.89 | 5.1 | 5.19 | 5 |  | 36
| 1 | 6 | 5 | 5.5 | 4.92 | 5.1 | 4.8532 | 5.22 | 4.9 | 4.7 | 4.83 | 4.92 | 6 | 5.54 | 4.6916 | 5.09 | 5.2 | 5 | 4.7 | 4.9 | 5.2 | 4.29 | 5.16 | 5.2 | 5.24 | 5
| Mean |  | 5.0167 | 5.5 | 4.9417 | 5.1333 | 4.9071 | 5.2917 | 4.7617 | 4.7183 | 4.955 | 5.2267 | 6 | 5.2817 | 4.7383 | 4.7067 | 5.2167 | 4.99 | 4.7833 | 5.0167 | 4.95 | 4.435 | 4.9733 | 5.2667 | 5.13 | 5.0333
| Median |  | 5 | 5.5 | 4.955 | 5.15 | 4.8917 | 5.275 | 4.775 | 4.68 | 4.94 | 5.285 | 6 | 5.265 | 4.7249 | 5.005 | 5.25 | 5 | 4.8 | 5 | 5 | 4.46 | 4.875 | 5.25 | 5.17 | 5
| Std Dev. |  | 0.040825 | 0 | 0.06969 | 0.1633 | 0.059392 | 0.15092 | 0.13029 | 0.1264 | 0.10895 | 0.19654 | 0 | 0.14552 | 0.060189 | 0.67816 | 0.1169 | 0.121 | 0.075277 | 0.075277 | 0.32094 | 0.12865 | 0.28974 | 0.1633 | 0.11437 | 0.05164
| Rel.Std.Dev. |  | 0.8138% | 0% | 1.41% | 3.181% | 1.21% | 2.852% | 2.736% | 2.679% | 2.199% | 3.76% | 0% | 2.755% | 1.27% | 14.41% | 2.241% | 2.425% | 1.574% | 1.501% | 6.484% | 2.901% | 5.826% | 3.101% | 2.229% | 1.026%
| PDM3 |  | 0.113% | 9.758% | -1.384% | 2.441% | -2.073% | 5.601% | -4.976% | -5.841% | -1.118% | 4.304% | 19.74% | 5.401% | -5.441% | -6.073% | 4.104% | -0.4192% | -4.543% | 0.113% | -1.217% | -11.49% | -0.7518% | 5.102% | 2.375% | 0.4456%
| Z-Score (Absolute) |  | 0.09 | 1.66 | 0.16 | 0.47 | 0.27 | 0.98 | 0.75 | 0.89 | 0.11 | 0.77 | NA | 0.95 | 0.82 | 0.92 | 0.74 | 0 | 0.67 | 0.09 | 0.13 | 1.81 | 0.05 | 0.9 | 0.46 | 0.14
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 922b (Certified Value 0.096 wt.%) | [26 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.089 | 0.24591 | 0.1445 | 0.14 | 0.0927 | 0.095083 | 0.1 | 0.09 | 0.1 | 0.21 | 0.16 | 0.09 | 0.091 | 0.1136 | 0.1 | 0.092 | 0.178 | 0.0656 | 0.099 | 0.1 | 0.14 | 0.108 | 0.0915 | 0.05 | 0.11 | 1
| 1 | 2 | 0.081 | 0.24597 | 0.146 | 0.13 | 0.0865 | 0.095618 | 0.1 | 0.1 | 0.1 | 0.21 | 0.16 | 0.1 | 0.095 | 0.11044 | 0.1 | 0.0923 | 0.159 | 0.0663 | 0.1 | 0.09 | 0.14 | 0.109 | 0.0902 | 0.05 | 0.1 | 21
| 1 | 3 | 0.088 | 0.24658 | 0.1464 | 0.1 | 0.0884 | 0.095412 | 0.11 | 0.09 | 0.1 | 0.21 | 0.16 | 0.09 | 0.093 | 0.12061 | 0.1 | 0.0911 | 0.175 | 0.0639 | 0.097 | 0.09 | 0.14 | 0.105 | 0.0908 | 0.05 | 0.11 | 16
| 1 | 4 | 0.09 | 0.24475 | 0.1451 | 0.1 | 0.0884 | 0.093586 | 0.1 | 0.1 | 0.1 | 0.22 | 0.16 | 0.09 | 0.092 | 0.10967 | 0.08 | 0.0906 | 0.174 | 0.0655 | 0.098 | 0.1 | 0.14 | 0.108 | 0.089 | 0.05 | 0.1 | 0.096132
| 1 | 5 | 0.086 | 0.24454 | 0.145 | 0.11 | 0.0892 | 0.094744 | 0.12 | 0.1 | 0.1 | 0.21 | 0.16 | 0.1 | 0.094 | 0.1126 | 0.07 | 0.0911 | 0.178 | 0.0643 | 0.099 | 0.09 | 0.14 | 0.104 | 0.09 | 0.05 | 0.11 | 37
| 1 | 6 | 0.088 | 0.24465 | 0.145 | 0.1 | 0.0857 | 0.096267 | 0.11 | 0.1 | 0.1 | 0.21 | 0.16 | 0.09 | 0.093 | 0.12269 | 0.13 | 0.0912 | 0.166 | 0.0649 | 0.095 | 0.1 | 0.14 | 0.105 | 0.0918 | 0.05 | 0.11
| Mean |  | 0.087 | 0.2454 | 0.14533 | 0.11333 | 0.088483 | 0.095118 | 0.10667 | 0.096667 | 0.1 | 0.21167 | 0.16 | 0.093333 | 0.093 | 0.11493 | 0.096667 | 0.091383 | 0.17167 | 0.065083 | 0.098 | 0.095 | 0.14 | 0.1065 | 0.09055 | 0.05 | 0.10667
| Median |  | 0.088 | 0.24533 | 0.14505 | 0.105 | 0.0884 | 0.095247 | 0.105 | 0.1 | 0.1 | 0.21 | 0.16 | 0.09 | 0.093 | 0.1131 | 0.1 | 0.09115 | 0.1745 | 0.0652 | 0.0985 | 0.095 | 0.14 | 0.1065 | 0.0905 | 0.05 | 0.11
| Std Dev. |  | 0.0032249 | 0.00086046 | 0.00071461 | 0.017512 | 0.002449 | 0.00091041 | 0.008165 | 0.005164 | 1.5202e-17 | 0.0040825 | 0 | 0.005164 | 0.0014142 | 0.0054324 | 0.020656 | 0.00063692 | 0.007607 | 0.00089088 | 0.0017889 | 0.0054772 | 0 | 0.0020736 | 0.0010349 | 7.6012e-18 | 0.005164
| Rel.Std.Dev. |  | 3.707% | 0.3506% | 0.4917% | 15.45% | 2.768% | 0.9571% | 7.655% | 5.342% | 1.52e-14% | 1.929% | 0% | 5.533% | 1.521% | 4.727% | 21.37% | 0.697% | 4.431% | 1.369% | 1.825% | 5.766% | 0% | 1.947% | 1.143% | 1.52e-14% | 4.841%
| PDM3 |  | -9.5% | 155.3% | 51.18% | 17.89% | -7.957% | -1.055% | 10.96% | 0.5557% | 4.023% | 120.2% | 66.44% | -2.912% | -3.258% | 19.56% | 0.5557% | -4.94% | 78.57% | -32.3% | 1.943% | -1.178% | 45.63% | 10.78% | -5.807% | -47.99% | 10.96%
| Z-Score (Absolute) |  | 0.75 | 10.7 | 3.47 | 1.16 | 0.64 | 0.16 | 0.67 | 0.05 | NA | 8.26 | 4.53 | 0.29 | 0.31 | 1.27 | 0.05 | 0.43 | 5.37 | 2.33 | 0.05 | 0.17 | 3.08 | 0.66 | 0.49 | 3.42 | 0.67
Results from laboratory 2.09 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb in OREAS 922b (Certified Value 5.75 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 4.4 | 6.1701 | 5.1 | 11.17 | 6.42 | 5.9865 | 6.5 | 6.7 | 5.4 | 6.1 | 6.7 | 10.9 | 4.6 | 5.2752 | 6.7 | 8.31 | 5.29 | 4.4 | < 5 | 4.1 | 6.1 | 4.76 | 7.5 | 6.6 |  | 1
| 1 | 2 | 4.7 | 6.2369 | 5.1 | 9.43 | 4.87 | 6.4361 | 6.8 | 6.8 | 5.7 | 5.5 | 6.1 | 9.2 | 4.9 | 5.1105 | 6.5 | 6.95 | 4.96 | 4.7 | < 5 | 4 | 5.9 | 5.39 | 7.2 | 6.2 |  | 15
| 1 | 3 | 4.4 | 6.349 | 5.2 | 10.82 | 4.35 | 5.6893 | 6.4 | 6.9 | 5.3 | 6.1 | 6.5 | 8.2 | 4.9 | 5.3184 | 6.9 | 8.76 | 4.97 | 5 | < 5 | 4.3 | 5.9 | 5.04 | 7.3 | 5.6 |  | 16
| 1 | 4 | 4.4 | 6.3635 | 5.1 | 10.43 | 6.63 | 6.0072 | 6.4 | 6.2 | 5.5 | 5.9 | 6 | 9.7 | 4.9 | 5.0254 | 5.2 | 8.2 | 5.26 | 5 | < 5 | 4.3 | 5.8 | 4.43 | 7.4 | 5.5 |  | 5.7452
| 1 | 5 | 4.4 | 6.1201 | 5.1 | 11.25 | 4.72 | 5.8347 | 6.7 | 7.2 | 5.3 | 6.2 | 6.7 | 10 | 4.7 | 5.0135 | 4.3 | 7.38 | 6.23 | 4.9 | < 5 | 4 | 6.2 | 4.31 | 6.7 | 6.2 |  | 38
| 1 | 6 | 4.2 | 6.2324 | 5.1 | 10.99 | 4.49 | 5.8057 | 6.5 | 7.2 | 5.7 | 5.9 | 6.3 | 10.2 | 4.8 | 5.2883 | 6.4 | 8.5 | 5.52 | 4.4 | < 5 | 4.4 | 6 | 4.59 | 7.1 | 5.7
| Mean |  | 4.4167 | 6.2453 | 5.1167 | 10.6817 | 5.2467 | 5.9599 | 6.55 | 6.8333 | 5.4833 | 5.95 | 6.3833 | 9.7 | 4.8 | 5.1719 | 6 | 8.0167 | 5.3717 | 4.7333 |  | 4.1833 | 5.9833 | 4.7533 | 7.2 | 5.9667
| Median |  | 4.4 | 6.2347 | 5.1 | 10.905 | 4.795 | 5.9106 | 6.5 | 6.85 | 5.45 | 6 | 6.4 | 9.85 | 4.85 | 5.1929 | 6.45 | 8.255 | 5.275 | 4.8 |  | 4.2 | 5.95 | 4.675 | 7.25 | 5.95
| Std Dev. |  | 0.16021 | 0.096221 | 0.040825 | 0.67907 | 1.0086 | 0.26167 | 0.16432 | 0.37238 | 0.18348 | 0.251 | 0.29944 | 0.9252 | 0.12649 | 0.13857 | 1.0237 | 0.69985 | 0.47089 | 0.28048 |  | 0.17224 | 0.1472 | 0.40382 | 0.28284 | 0.43205
| Rel.Std.Dev. |  | 3.627% | 1.541% | 0.7979% | 6.357% | 19.22% | 4.391% | 2.509% | 5.449% | 3.346% | 4.218% | 4.691% | 9.538% | 2.635% | 2.679% | 17.06% | 8.73% | 8.766% | 5.926% |  | 4.117% | 2.46% | 8.495% | 3.928% | 7.241%
| PDM3 |  | -23.12% | 8.706% | -10.94% | 85.92% | -8.677% | 3.738% | 14.01% | 18.94% | -4.558% | 3.565% | 11.11% | 68.84% | -16.45% | -9.979% | 4.435% | 39.54% | -6.501% | -17.61% |  | -27.19% | 4.145% | -17.26% | 25.32% | 3.855%
| Z-Score (Absolute) |  | 1.28 | 0.24 | 0.7 | 3.93 | 0.59 | 0 | 0.49 | 0.73 | 0.39 | 0 | 0.36 | 3.12 | 0.96 | 0.65 | 0.04 | 1.71 | 0.49 | 1.02 | 2.87 | 1.47 | 0.02 | 1 | 1.04 | 0.01
| Analytical results for Nd in OREAS 922b (Certified Value 35 ppm) | [26 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.18; 2.08: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS |  | 1
2
| 1 | 1 | 2.01: 34.319; 2.02: 34.8; 2.03: 33.9512; 2.04: 35.8; 2.05: 35.17; 2.06: 34.3; 2.07: 38; 2.08: 34.8 |  | 1
| 1 | 2 | 2.01: 34.414; 2.02: 35.6; 2.03: 35.67; 2.04: 36.6; 2.05: 34.23; 2.06: 32.3; 2.07: 38; 2.08: 34.3 |  | 4
| 1 | 3 | 2.01: 34.515; 2.02: 35.6; 2.03: 34.4823; 2.04: 35.7; 2.05: 34.49; 2.06: 35.1; 2.07: 38; 2.08: 35 |  | 16
| 1 | 4 | 2.01: 34.675; 2.02: 35.4; 2.03: 33.5948; 2.04: 35; 2.05: 35.39; 2.06: 33.3; 2.07: 38; 2.08: 35.4 |  | 34.9836
| 1 | 5 | 2.01: 34.379; 2.02: 35.8; 2.03: 35.003; 2.04: 36.5; 2.05: 34.59; 2.06: 36; 2.07: 37; 2.08: 33.8 |  | 39
| 1 | 6 | 2.01: 34.535; 2.02: 35.5; 2.03: 34.3941; 2.04: 36; 2.05: 35.47; 2.06: 35.5; 2.07: 37; 2.08: 35.4
| Mean |  | 2.01: 34.4728; 2.02: 35.45; 2.03: 34.5159; 2.04: 35.9333; 2.05: 34.89; 2.06: 34.4167; 2.07: 37.6667; 2.08: 34.7833
| Median |  | 2.01: 34.4645; 2.02: 35.55; 2.03: 34.4382; 2.04: 35.9; 2.05: 34.88; 2.06: 34.7; 2.07: 38; 2.08: 34.9
| Std Dev. |  | 2.01: 0.12841; 2.02: 0.34496; 2.03: 0.74193; 2.04: 0.58538; 2.05: 0.51969; 2.06: 1.4063; 2.07: 0.5164; 2.08: 0.63377
| Rel.Std.Dev. |  | 2.01: 0.3725%; 2.02: 0.9731%; 2.03: 2.15%; 2.04: 1.629%; 2.05: 1.49%; 2.06: 4.086%; 2.07: 1.371%; 2.08: 1.822%
| PDM3 |  | 2.01: -1.46%; 2.02: 1.333%; 2.03: -1.337%; 2.04: 2.715%; 2.05: -0.2676%; 2.06: -1.621%; 2.07: 7.669%; 2.08: -0.5725%
| Z-Score (Absolute) |  | 2.01: 0.63; 2.02: 1.06; 2.03: 0.55; 2.04: 1.89; 2.05: 0.09; 2.06: 0.72; 2.07: 4.87; 2.08: 0.09
| Analytical results for Ni in OREAS 922b (Certified Value 12.2 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 12 | 11.22 | 13.4 | 14.01 | 13.3 | 11.6836 | 12.9 | 11.2 | 11 | 12 | 12.7 | 14 | 12 | 11.4479 | 19.5 | 12 | 11.5 | 11.7 | 12.2 | 12.5 | 12.3 | 11.9 | 12.4 | 12.3 | 12.3 | 1
| 1 | 2 | 12 | 10.33 | 13.7 | 11.76 | 11.4 | 11.7457 | 12.5 | 10.8 | 11.2 | 12 | 12.6 | 15 | 12 | 11.6683 | 17.8 | 13.7 | 11.7 | 12 | 12.5 | 12.9 | 14.4 | 11.6 | 12.3 | 12.3 | 12.7 | 22
| 1 | 3 | 12 | 10.72 | 13.5 | 13.6 | 14.8 | 11.0913 | 12.1 | 11 | 11.2 | 12.3 | 12.7 | 13 | 12 | 12.3361 | 19.5 | 11.9 | 12.5 | 11.8 | 12.2 | 13.3 | 11.9 | 11.6 | 12.8 | 12.8 | 12.5 | 16
| 1 | 4 | 12 | 10.47 | 13.7 | 11.7 | 11.3 | 12.2825 | 12.1 | 10 | 11.8 | 11.8 | 12.8 | 13 | 12 | 12.7012 | 13.5 | 12.9 | 11.5 | 12.3 | 12.8 | 14.1 | 12.2 | 11.9 | 11.9 | 12.3 | 12.7 | 12.2469
| 1 | 5 | 12 | 12.33 | 13.4 | 11.72 | 12.1 | 12.3969 | 12.7 | 10.3 | 11.6 | 11.8 | 13.3 | 13 | 14 | 12.5602 | 16.5 | 12.1 | 12.6 | 12.4 | 12.3 | 12.5 | 13.7 | 14.1 | 13 | 13.1 | 12.7 | 40
| 1 | 6 | 12 | 11.7 | 13.7 | 11.52 | 11.8 | 11.4188 | 11.9 | 11.4 | 11.2 | 11 | 12 | 14 | 11 | 12.2867 | 17.8 | 12.8 | 12.4 | 11.4 | 12.2 | 12.8 | 11.7 | 12.2 | 12.4 | 12.5 | 12.2
| Mean |  | 12 | 11.1283 | 13.5667 | 12.385 | 12.45 | 11.7698 | 12.3667 | 10.7833 | 11.3333 | 11.8167 | 12.6833 | 13.6667 | 12.1667 | 12.1668 | 17.4333 | 12.5667 | 12.0333 | 11.9333 | 12.3667 | 13.0167 | 12.7 | 12.2167 | 12.4667 | 12.55 | 12.5167
| Median |  | 12 | 10.97 | 13.6 | 11.74 | 11.95 | 11.7146 | 12.3 | 10.9 | 11.2 | 11.9 | 12.7 | 13.5 | 12 | 12.3114 | 17.8 | 12.45 | 12.05 | 11.9 | 12.25 | 12.85 | 12.25 | 11.9 | 12.4 | 12.4 | 12.6
| Std Dev. |  | 0 | 0.77675 | 0.15055 | 1.1106 | 1.3576 | 0.49955 | 0.39328 | 0.53821 | 0.30111 | 0.44008 | 0.41673 | 0.8165 | 0.98319 | 0.49972 | 2.2429 | 0.69761 | 0.52026 | 0.37771 | 0.24221 | 0.608 | 1.09 | 0.94956 | 0.38816 | 0.33317 | 0.22286
| Rel.Std.Dev. |  | 0% | 6.98% | 1.11% | 8.967% | 10.9% | 4.244% | 3.18% | 4.991% | 2.657% | 3.724% | 3.286% | 5.974% | 8.081% | 4.107% | 12.87% | 5.551% | 4.323% | 3.165% | 1.959% | 4.671% | 8.582% | 7.773% | 3.114% | 2.655% | 1.781%
| PDM3 |  | -2.016% | -9.134% | 10.78% | 1.128% | 1.658% | -3.896% | 0.9778% | -11.95% | -7.46% | -3.513% | 3.564% | 11.59% | -0.6552% | -0.6545% | 42.35% | 2.611% | -1.744% | -2.56% | 0.9778% | 6.285% | 3.7% | -0.247% | 1.794% | 2.475% | 2.203%
| Z-Score (Absolute) |  | NA | 2.66 | 2.54 | 0.02 | 0.16 | 1.29 | 0.02 | 3.4 | 2.23 | 1.19 | 0.66 | NA | NA | 0.45 | 10.8 | 0.41 | 0.73 | 0.94 | 0.02 | 1.37 | 0.69 | 0.34 | 0.19 | 0.37 | 0.3
Results from laboratories 2.01, 2.12 and 2.13 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 922b (Certified Value 0.033 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.034 | 0.0278 | 0.03 | 0.0336 | 0.034402 | 0.035 | 0.032 | 0.033 | 0.034 | 0.031 | 0.037 | 0.034 | 0.030883 | 0.03 | 0.0336 | 0.035 | 0.0356 | 0.035 | 0.0351 | 0.03 | 0.0327 | 0.0329 | 0.034 | 0.032 |  | 1
| 1 | 2 | 0.034 | 0.0278 | 0.03 | 0.0287 | 0.035387 | 0.035 | 0.032 | 0.034 | 0.035 | 0.033 | 0.037 | 0.036 | 0.031262 | 0.032 | 0.0332 | 0.032 | 0.0353 | 0.035 | 0.0351 | 0.03 | 0.034 | 0.0332 | 0.0332 | 0.033 |  | 0
| 1 | 3 | 0.033 | 0.028 | 0.04 | 0.0317 | 0.034214 | 0.034 | 0.032 | 0.033 | 0.034 | 0.032 | 0.038 | 0.034 | 0.034348 | 0.031 | 0.0328 | 0.034 | 0.0354 | 0.035 | 0.0353 | 0.03 | 0.0333 | 0.0346 | 0.0347 | 0.033 |  | 16
| 1 | 4 | 0.035 | 0.028 | 0.03 | 0.0315 | 0.0351 | 0.035 | 0.032 | 0.033 | 0.035 | 0.032 | 0.037 | 0.035 | 0.031664 | 0.025 | 0.0337 | 0.032 | 0.0364 | 0.035 | 0.0353 | 0.03 | 0.0336 | 0.0301 | 0.0345 | 0.033 |  | 0.033391
| 1 | 5 | 0.033 | 0.028 | 0.03 | 0.032 | 0.034819 | 0.035 | 0.032 | 0.034 | 0.034 | 0.031 | 0.037 | 0.036 | 0.03272 | 0.019 | 0.0328 | 0.035 | 0.035 | 0.035 | 0.0354 | 0.03 | 0.0333 | 0.0322 | 0.0341 | 0.033 |  | 41
| 1 | 6 | 0.034 | 0.0278 | 0.03 | 0.0317 | 0.035784 | 0.034 | 0.032 | 0.034 | 0.034 | 0.032 | 0.036 | 0.035 | 0.034429 | 0.031 | 0.0333 | 0.033 | 0.0361 | 0.034 | 0.0353 | 0.03 | 0.0332 | 0.0328 | 0.0333 | 0.033
| Mean |  | 0.033833 | 0.0279 | 0.031667 | 0.031533 | 0.034951 | 0.034667 | 0.032 | 0.0335 | 0.034333 | 0.031833 | 0.037 | 0.035 | 0.032551 | 0.028 | 0.033233 | 0.0335 | 0.035633 | 0.034833 | 0.03525 | 0.03 | 0.03335 | 0.032633 | 0.033967 | 0.032833
| Median |  | 0.034 | 0.0279 | 0.03 | 0.0317 | 0.034959 | 0.035 | 0.032 | 0.0335 | 0.034 | 0.032 | 0.037 | 0.035 | 0.032192 | 0.0305 | 0.03325 | 0.0335 | 0.0355 | 0.035 | 0.0353 | 0.03 | 0.0333 | 0.03285 | 0.03405 | 0.033
| Std Dev. |  | 0.00075277 | 0.00010954 | 0.0040825 | 0.0015858 | 0.0005947 | 0.0005164 | 0 | 0.00054772 | 0.0005164 | 0.00075277 | 0.00063246 | 0.00089443 | 0.0015503 | 0.0050596 | 0.00038297 | 0.0013784 | 0.00052409 | 0.00040825 | 0.00012247 | 0 | 0.00043243 | 0.001476 | 0.0006121 | 0.00040825
| Rel.Std.Dev. |  | 2.225% | 0.3926% | 12.89% | 5.029% | 1.702% | 1.49% | 0% | 1.635% | 1.504% | 2.365% | 1.709% | 2.556% | 4.763% | 18.07% | 1.152% | 4.115% | 1.471% | 1.172% | 0.3474% | 0% | 1.297% | 4.523% | 1.802% | 1.243%
| PDM3 |  | 1.324% | -16.45% | -5.165% | -5.565% | 4.671% | 3.819% | -4.167% | 0.3252% | 2.821% | -4.666% | 10.81% | 4.817% | -2.516% | -16.15% | -0.4734% | 0.3252% | 6.714% | 4.318% | 5.566% | -10.16% | -0.124% | -2.27% | 1.723% | -1.671%
| Z-Score (Absolute) |  | 0.19 | 2.63 | 0.84 | 0.9 | 0.73 | 0.59 | 0.68 | 0.04 | 0.43 | 0.76 | 1.7 | 0.75 | 0.42 | 2.58 | 0.09 | 0.04 | 1.05 | 0.67 | 0.87 | 1.63 | 0.04 | 0.38 | 0.26 | 0.28
| Analytical results for Pb in OREAS 922b (Certified Value 43.5 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 44 | 37.06 | 45.15 | 55.9 | 41 | 48.2345 | 49 | 40.7 | 40.9 | 41.1 | 46.2 | 42 | 45 | 35.9512 | 66.4 | 44 | 39.81 | 43 | 45.8 | 44 | 45 | 40.9 | 53.4 | 42.4 | 43.1 | 1
| 1 | 2 | 45 | 35.92 | 45.34 | 55.9 | 41.3 | 47.322 | 43.2 | 41.1 | 41.1 | 42.5 | 44.3 | 42 | 43.8 | 37.6699 | 66.2 | 43.2 | 39 | 43 | 51.9 | 44.1 | 45 | 41.1 | 51.4 | 42.5 | 41.8 | 17
| 1 | 3 | 44 | 38.25 | 44.73 | 57.8 | 41.4 | 46.6748 | 48.2 | 42.7 | 39.9 | 42.5 | 47.1 | 44 | 42.8 | 40.6991 | 75.3 | 42.5 | 41.6 | 55 | 55.5 | 42.4 | 45 | 39.8 | 50.5 | 43.4 | 41.9 | 16
| 1 | 4 | 44 | 36.41 | 45.46 | 56.2 | 40.7 | 48.4209 | 43.1 | 40.3 | 42.5 | 42 | 45.2 | 44 | 44.2 | 35.5432 | 54 | 43.4 | 39.67 | 60 | 56.8 | 44.6 | 46 | 39.5 | 52.3 | 42.4 | 42.2 | 43.5078
| 1 | 5 | 43 | 35.31 | 45.7 | 56.7 | 41.1 | 48.0789 | 45.8 | 41.6 | 40.6 | 41.5 | 47.5 | 47 | 45.1 | 38.477 | 49.1 | 43.5 | 42.35 | 68 | 46.6 | 42 | 45 | 41.7 | 47 | 42.6 | 41.6 | 42
| 1 | 6 | 45 | 38.62 | 44.94 | 56.4 | 41.3 | 47.3296 | 45.3 | 42.1 | 40.6 | 41.3 | 42.5 | 42 | 43.7 | 39.8031 | 66.6 | 44.3 | 41.27 | 42 | 46.1 | 43.2 | 46 | 40.3 | 48.5 | 44.2 | 42
| Mean |  | 44.1667 | 36.9283 | 45.22 | 56.4833 | 41.1333 | 47.6768 | 45.7667 | 41.4167 | 40.9333 | 41.8167 | 45.4667 | 43.5 | 44.1 | 38.0239 | 62.9333 | 43.4833 | 40.6167 | 51.8333 | 50.45 | 43.3833 | 45.3333 | 40.55 | 50.5167 | 42.9167 | 42.1
| Median |  | 44 | 36.735 | 45.245 | 56.3 | 41.2 | 47.7043 | 45.55 | 41.35 | 40.75 | 41.75 | 45.7 | 43 | 44 | 38.0735 | 66.3 | 43.45 | 40.54 | 49 | 49.25 | 43.6 | 45 | 40.6 | 50.95 | 42.55 | 41.95
| Std Dev. |  | 0.75277 | 1.3062 | 0.35378 | 0.71391 | 0.2582 | 0.67481 | 2.4614 | 0.89536 | 0.86872 | 0.608 | 1.8747 | 1.9748 | 0.86718 | 2.0544 | 9.5941 | 0.63061 | 1.3083 | 10.8704 | 4.9658 | 1.0284 | 0.5164 | 0.83367 | 2.3962 | 0.73326 | 0.52915
| Rel.Std.Dev. |  | 1.704% | 3.537% | 0.7824% | 1.264% | 0.6277% | 1.415% | 5.378% | 2.162% | 2.122% | 1.454% | 4.123% | 4.54% | 1.966% | 5.403% | 15.24% | 1.45% | 3.221% | 20.97% | 9.843% | 2.371% | 1.139% | 2.056% | 4.743% | 1.709% | 1.257%
| PDM3 |  | 1.514% | -15.12% | 3.935% | 29.82% | -5.458% | 9.582% | 5.192% | -4.806% | -5.917% | -3.887% | 4.502% | -0.01789% | 1.361% | -12.6% | 44.65% | -0.0562% | -6.645% | 19.14% | 15.96% | -0.286% | 4.196% | -6.798% | 16.11% | -1.359% | -3.236%
| Z-Score (Absolute) |  | 0.2 | 1.95 | 0.51 | 3.86 | 0.7 | 1.24 | 0.67 | 0.62 | 0.76 | 0.5 | 0.59 | 0 | 0.18 | 1.63 | 5.78 | 0 | 0.86 | 2.48 | 2.07 | 0.03 | 0.55 | 0.88 | 2.09 | 0.17 | 0.42
| Analytical results for Pr in OREAS 922b (Certified Value 9.46 ppm) | [26 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 8.5805; 2.02: 9.5; 2.03: 9.3581; 2.04: 10.4; 2.05: 9.48; 2.06: 9.3; 2.07: 9.48 | 1
| 1 | 2 | 2.01: 8.8448; 2.02: 9.5; 2.03: 9.7181; 2.04: 10.2; 2.05: 9.5; 2.06: 8.7; 2.07: 9.37 | 6
| 1 | 3 | 2.01: 8.5276; 2.02: 9.5; 2.03: 9.4546; 2.04: 10; 2.05: 9.32; 2.06: 9.6; 2.07: 9.42 | 16
| 1 | 4 | 2.01: 8.5817; 2.02: 9.5; 2.03: 9.233; 2.04: 9.9; 2.05: 9.56; 2.06: 9; 2.07: 9.68 | 9.4587
| 1 | 5 | 2.01: 8.6082; 2.02: 9.5; 2.03: 9.5012; 2.04: 10.4; 2.05: 9.52; 2.06: 9.6; 2.07: 9.01 | 43
| 1 | 6 | 2.01: 8.4125; 2.02: 9.5; 2.03: 9.2946; 2.04: 10.3; 2.05: 9.75; 2.06: 9.5; 2.07: 9.71
| Mean |  | 2.01: 8.5925; 2.02: 9.5; 2.03: 9.4266; 2.04: 10.2; 2.05: 9.5217; 2.06: 9.2833; 2.07: 9.445
| Median |  | 2.01: 8.5811; 2.02: 9.5; 2.03: 9.4063; 2.04: 10.25; 2.05: 9.51; 2.06: 9.4; 2.07: 9.45
| Std Dev. |  | 2.01: 0.14197; 2.02: 0; 2.03: 0.17373; 2.04: 0.20976; 2.05: 0.13891; 2.06: 0.3656; 2.07: 0.25383
| Rel.Std.Dev. |  | 2.01: 1.652%; 2.02: 0%; 2.03: 1.843%; 2.04: 2.056%; 2.05: 1.459%; 2.06: 3.938%; 2.07: 2.687%
| PDM3 |  | 2.01: -9.157%; 2.02: 0.4371%; 2.03: -0.3389%; 2.04: 7.838%; 2.05: 0.6662%; 2.06: -1.854%; 2.07: -0.1444%
| Z-Score (Absolute) |  | 2.01: 7.5; 2.02: 0.48; 2.03: 0.16; 2.04: 6.64; 2.05: 0.67; 2.06: 1.42; 2.07: 0
| Analytical results for Rb in OREAS 922b (Certified Value 152 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 0
0
| 1 | 1 | 152 | 163.6 | 142.13 | 153.88 | 143.8102 | 152 | 146 | 152.5 | 158.5 | 154 | 140 | 154 | 130.6701 | 153 | 157.27 | 146.1 | 158.9 | 148 | 153 | 158.13 | 144 | 169.7 |  |  |  | 1
| 1 | 2 | 154 | 165 | 144.48 | 147.31 | 147.5416 | 157 | 145 | 156 | 165.5 | 151 | 142 | 158 | 128.2153 | 153 | 158.72 | 122.4 | 159.7 | 155 | 149 | 159.52 | 147.1 | 172.7 |  |  |  | 37
| 1 | 3 | 150 | 165 | 144.62 | 151.81 | 141.8898 | 155.5 | 154.5 | 150 | 160 | 161 | 130 | 158 | 155.3077 | 157 | 157.37 | 150 | 158.9 | 157 | 149 | 158.52 | 149.8 | 172.2 |  |  |  | 16
| 1 | 4 | 153 | 165 | 141.02 | 152.02 | 145.1394 | 153 | 145 | 157 | 161 | 155.5 | 132 | 158 | 141.0226 | 125 | 162.36 | 142.6 | 162.1 | 155 | 147 | 153.56 | 147.2 | 169.7 |  |  |  | 152.2434
| 1 | 5 | 152 | 164.8 | 142.44 | 152.67 | 145.5272 | 162 | 146.5 | 152 | 159 | 159 | 141 | 154 | 126.3487 | 102 | 162.48 | 136.1 | 157.1 | 151 | 150 | 154.54 | 149.7 | 169.7 |  |  |  | 44
| 1 | 6 | 155 | 165 | 142.95 | 149.68 | 145.7667 | 154 | 152 | 153 | 156.5 | 145 | 131 | 158 | 157.4059 | 148 | 161.01 | 139.5 | 155.5 | 165 | 150 | 157.04 | 153.1 | 166.1
| Mean |  | 152.6667 | 164.7333 | 142.94 | 151.2283 | 144.9458 | 155.5833 | 148.1667 | 153.4167 | 160.0833 | 154.25 | 136 | 156.6667 | 139.8284 | 139.6667 | 159.8683 | 139.45 | 158.7 | 155.1667 | 149.6667 | 156.885 | 148.4833 | 170.0167
| Median |  | 152.5 | 165 | 142.695 | 151.915 | 145.3333 | 154.75 | 146.25 | 152.75 | 159.5 | 154.75 | 136 | 158 | 135.8464 | 150.5 | 159.865 | 141.05 | 158.9 | 155 | 149.5 | 157.585 | 148.45 | 169.7
| Std Dev. |  | 1.7512 | 0.56095 | 1.399 | 2.359 | 1.9192 | 3.6113 | 4.0579 | 2.6157 | 3.0564 | 5.7598 | 5.5498 | 2.0656 | 13.7873 | 21.7225 | 2.3935 | 9.6682 | 2.2557 | 5.8109 | 1.9664 | 2.3556 | 3.1045 | 2.3498
| Rel.Std.Dev. |  | 1.147% | 0.3405% | 0.9788% | 1.56% | 1.324% | 2.321% | 2.739% | 1.705% | 1.909% | 3.734% | 4.081% | 1.318% | 9.86% | 15.55% | 1.497% | 6.933% | 1.421% | 3.745% | 1.314% | 1.502% | 2.091% | 1.382%
| PDM3 |  | 0.278% | 8.204% | -6.111% | -0.6668% | -4.793% | 2.194% | -2.678% | 0.7706% | 5.15% | 1.318% | -10.67% | 2.905% | -8.155% | -8.261% | 5.008% | -8.403% | 4.241% | 1.92% | -1.693% | 3.049% | -2.47% | 11.67%
| Z-Score (Absolute) |  | 0.05 | 1.5 | 1.29 | 0.23 | 1.04 | 0.33 | 0.62 | 0.05 | 0.9 | 0.15 | 2.18 | 0.46 | 1.69 | 1.71 | 0.87 | 1.74 | 0.72 | 0.27 | 0.43 | 0.49 | 0.58 | 2.17
| Analytical results for Re in OREAS 922b (Certified Value < 0.002 ppm) | [26 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [9 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.002 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [9 empty] | 1
| 1 | 2 | < 0.05 | < 0.002 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [9 empty] | 23
| 1 | 3 | < 0.05 | < 0.002 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [9 empty] | 16
| 1 | 4 | < 0.05 | < 0.002 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [9 empty] | < 0.002
| 1 | 5 | < 0.05 | < 0.002 | < 0.01 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005 | [9 empty] | 45
| 1 | 6 | < 0.05 | < 0.002 | < 0.01 | < 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | < 0.002 | < 0.005
| Mean |  | 2.06: 0.002; 2.09: 0.002
| Median |  | 2.06: 0.002; 2.09: 0.002
| Std Dev. |  | 2.06: 0
Rel.Std.Dev. | bar: 2.06
PDM3
| Z-Score (Absolute) |  | 96.76 | 0.34 | 15.85 | 0.34 | 1.01 | 0.34 | 0.34 | 0.34 | 2.36 | 96.76 | 0.34 | 0.34 | 1.69 | 5.73 | 0.34 | 5.73
| Analytical results for S in OREAS 922b (Certified Value 0.274 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.267 | 0.29291 | 0.265 | 0.14 | 0.24 | 0.28694 | 0.29 | 0.28 | 0.28 | 0.27 | 0.27 | 0.27 | 0.277 | 0.25211 | 0.3 | 0.2696 | 0.26 | 0.2732 | 0.2 | 0.28 | 0.25 | 0.26 | 0.29 | 0.28 | 0.28 | 1
| 1 | 2 | 0.268 | 0.28865 | 0.265 | 0.11 | 0.25 | 0.28046 | 0.29 | 0.27 | 0.28 | 0.27 | 0.29 | 0.27 | 0.285 | 0.25931 | 0.29 | 0.2648 | 0.23 | 0.2747 | 0.3 | 0.28 | 0.25 | 0.28 | 0.27 | 0.27 | 0.28 | 0
| 1 | 3 | 0.266 | 0.28633 | 0.265 | 0.14 | 0.25 | 0.28787 | 0.28 | 0.28 | 0.27 | 0.27 | 0.28 | 0.27 | 0.273 | 0.27475 | 0.3 | 0.2621 | 0.25 | 0.2659 | 0.3 | 0.28 | 0.26 | 0.28 | 0.28 | 0.29 | 0.27 | 16
| 1 | 4 | 0.269 | 0.2808 | 0.264 | 0.19 | 0.24 | 0.30637 | 0.29 | 0.27 | 0.28 | 0.28 | 0.28 | 0.27 | 0.276 | 0.24005 | 0.3 | 0.2658 | 0.23 | 0.273 | 0.2 | 0.29 | 0.25 | 0.28 | 0.26 | 0.28 | 0.27 | 0.27391
| 1 | 5 | 0.267 | 0.2852 | 0.265 | 0.15 | 0.27 | 0.29329 | 0.29 | 0.27 | 0.28 | 0.27 | 0.28 | 0.27 | 0.285 | 0.23113 | 0.3 | 0.2638 | 0.25 | 0.2695 | 0.3 | 0.28 | 0.25 | 0.27 | 0.26 | 0.28 | 0.28 | 46
| 1 | 6 | 0.272 | 0.28309 | 0.265 | 0.14 | 0.24 | 0.28381 | 0.28 | 0.28 | 0.28 | 0.27 | 0.28 | 0.27 | 0.283 | 0.26415 | 0.31 | 0.2648 | 0.23 | 0.2783 | 0.2 | 0.28 | 0.25 | 0.28 | 0.27 | 0.28 | 0.27
| Mean |  | 0.26817 | 0.28616 | 0.26483 | 0.145 | 0.24833 | 0.28979 | 0.28667 | 0.275 | 0.27833 | 0.27167 | 0.28 | 0.27 | 0.27983 | 0.25358 | 0.3 | 0.26515 | 0.24167 | 0.27243 | 0.25 | 0.28167 | 0.25167 | 0.275 | 0.27167 | 0.28 | 0.275
| Median |  | 0.2675 | 0.28576 | 0.265 | 0.14 | 0.245 | 0.2874 | 0.29 | 0.275 | 0.28 | 0.27 | 0.28 | 0.27 | 0.28 | 0.25571 | 0.3 | 0.2648 | 0.24 | 0.2731 | 0.25 | 0.28 | 0.25 | 0.28 | 0.27 | 0.28 | 0.275
| Std Dev. |  | 0.002137 | 0.004263 | 0.00040825 | 0.025884 | 0.01169 | 0.009185 | 0.005164 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0063246 | 0 | 0.0051543 | 0.016012 | 0.0063246 | 0.0025138 | 0.013292 | 0.0042838 | 0.054772 | 0.0040825 | 0.0040825 | 0.0083666 | 0.01169 | 0.0063246 | 0.0054772
| Rel.Std.Dev. |  | 0.7969% | 1.49% | 0.1542% | 17.85% | 4.708% | 3.17% | 1.801% | 1.992% | 1.467% | 1.503% | 2.259% | 0% | 1.842% | 6.314% | 2.108% | 0.9481% | 5.5% | 1.572% | 21.91% | 1.449% | 1.622% | 3.042% | 4.303% | 2.259% | 1.992%
| PDM3 |  | -2.099% | 4.472% | -3.315% | -47.06% | -9.339% | 5.796% | 4.655% | 0.3961% | 1.613% | -0.8208% | 2.222% | -1.429% | 2.161% | -7.423% | 9.523% | -3.2% | -11.77% | -0.5409% | -8.731% | 2.83% | -8.122% | 0.3961% | -0.8208% | 2.222% | 0.3961%
| Z-Score (Absolute) |  | 0.53 | 1.18 | 0.84 | 12.2 | 2.41 | 1.52 | 1.23 | 0.12 | 0.44 | 0.19 | 0.6 | 0.35 | 0.58 | 1.91 | 2.49 | 0.81 | 3.04 | 0.12 | NA | 0.75 | 2.09 | 0.12 | 0.19 | 0.6 | 0.12
Results from laboratory 2.19 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 922b (Certified Value 3.75 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.8 | 4.154 | 3 | 3.04 | 3.75 | 3.7691 | 4.38 | 3.48 | 3.62 | 3.67 | 4.03 | 4.2 | 4 | 2.9494 | 6 | 3.9 | 3.86 | 3.7 | 3.6 | < 5 | 3.75 | 3.61 | 3.65 | 3.4 | 3.6 | 1
| 1 | 2 | 3.8 | 4.1755 | 3 | 3.05 | 3.68 | 3.8322 | 4.07 | 3.54 | 3.62 | 3.85 | 3.76 | 3.9 | 3.97 | 3.1666 | 7 | 3.87 | 3.45 | 3.67 | 3.7 | < 5 | 3.64 | 3.74 | 3.58 | 3.4 | 3.9 | 39
| 1 | 3 | 3.9 | 4.1417 | 3 | 2.94 | 3.8 | 3.8848 | 3.61 | 3.74 | 3.47 | 3.79 | 4.17 | 3.6 | 3.89 | 3.7298 | 9 | 3.98 | 3.94 | 3.65 | 3.7 | < 5 | 5.34 | 3.5 | 3.85 | 3.5 | 3.9 | 16
| 1 | 4 | 3.8 | 4.0996 | 3 | 3.64 | 3.7 | 3.8308 | 4.22 | 3.5 | 3.71 | 3.82 | 3.9 | 3.4 | 3.91 | 3.4639 | 6 | 3.91 | 3.61 | 3.68 | 3.8 | < 5 | 3.56 | 3.47 | 3.7 | 3.3 | 3.7 | 3.7549
| 1 | 5 | 3.7 | 4.1144 | 3 | 3.15 | 3.77 | 3.8084 | 3.89 | 3.62 | 3.48 | 3.73 | 4.14 | 3.8 | 3.94 | 3.4497 | 5 | 3.97 | 4.2 | 3.68 | 3.6 | < 5 | 4 | 3.59 | 3.6 | 3.4 | 3.7 | 47
| 1 | 6 | 3.7 | 4.1287 | 3 | 3.04 | 3.6 | 3.8608 | 3.71 | 3.72 | 3.55 | 3.68 | 3.77 | 3.7 | 3.95 | 3.8396 | 6 | 3.88 | 3.92 | 3.59 | 3.5 | < 5 | 4.03 | 3.64 | 3.47 | 3.6 | 3.7
| Mean |  | 3.7833 | 4.1357 | 3 | 3.1433 | 3.7167 | 3.831 | 3.98 | 3.6 | 3.575 | 3.7567 | 3.9617 | 3.7667 | 3.9433 | 3.4332 | 6.5 | 3.9183 | 3.83 | 3.6617 | 3.65 |  | 4.0533 | 3.5917 | 3.6417 | 3.4333 | 3.75
| Median |  | 3.8 | 4.1352 | 3 | 3.045 | 3.725 | 3.8315 | 3.98 | 3.58 | 3.585 | 3.76 | 3.965 | 3.75 | 3.945 | 3.4568 | 6 | 3.905 | 3.89 | 3.675 | 3.65 |  | 3.875 | 3.6 | 3.625 | 3.4 | 3.7
| Std Dev. |  | 0.075277 | 0.027426 | 0 | 0.25224 | 0.072296 | 0.04032 | 0.29786 | 0.11171 | 0.092682 | 0.074744 | 0.17949 | 0.27325 | 0.039833 | 0.33449 | 1.3784 | 0.046224 | 0.26488 | 0.038687 | 0.10488 |  | 0.65805 | 0.09786 | 0.12797 | 0.10328 | 0.12247
| Rel.Std.Dev. |  | 1.99% | 0.6632% | 0% | 8.025% | 1.945% | 1.052% | 7.484% | 3.103% | 2.593% | 1.99% | 4.531% | 7.254% | 1.01% | 9.743% | 21.21% | 1.18% | 6.916% | 1.057% | 2.873% |  | 16.23% | 2.725% | 3.514% | 3.008% | 3.266%
| PDM3 |  | 0.7576% | 10.14% | -20.1% | -16.29% | -1.018% | 2.027% | 5.995% | -4.125% | -4.791% | 0.04745% | 5.507% | 0.3138% | 5.019% | -8.568% | 73.11% | 4.353% | 2% | -2.483% | -2.793% |  | 7.948% | -4.347% | -3.015% | -8.564% | -0.1301%
| Z-Score (Absolute) |  | 0.12 | 1.58 | NA | 2.52 | 0.15 | 0.32 | 0.94 | 0.63 | 0.74 | 0.01 | 0.86 | 0.06 | 0.78 | 1.32 | 11.34 | 0.68 | 0.32 | 0.38 | 0.43 | 5.17 | 1.24 | 0.67 | 0.46 | 1.32 | 0.01
Results from laboratory 2.03 were removed due to their 0.01 ppm reading resolution.
| Analytical results for Sc in OREAS 922b (Certified Value 11.5 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 12 | 14.4042 | 12 | 10.28 | 12.2 | 11.0993 | 11.7 | 11.2 | 12 | 12.2 | 12.1 | 13 | 12 | 9.8542 | 10.9 | 11.8 | 11.4 | 12 | 11.3 | 10.8 | 11.1 | 12 | 12.2 | 12.3 |  | 1
| 1 | 2 | 12 | 14.3656 | 12 | 10.43 | 11.1 | 11.3908 | 11.7 | 11 | 12.6 | 13 | 11.7 | 13 | 11.9 | 10.148 | 11 | 12.4 | 10.2 | 12 | 11 | 11.8 | 11.4 | 11.6 | 12.3 | 12.4 |  | 19
| 1 | 3 | 12 | 14.322 | 12 | 10.67 | 11.6 | 11.2803 | 10.6 | 11.8 | 12 | 12.6 | 12.7 | 13 | 11.1 | 11.4255 | 11.3 | 11.8 | 10.9 | 12 | 11.6 | 11.5 | 11.4 | 11.7 | 11.9 | 12.4 |  | 16
| 1 | 4 | 12 | 14.3514 | 12 | 10.18 | 11.5 | 11.1037 | 11.4 | 11.1 | 12.8 | 12.6 | 11.9 | 13 | 11.6 | 11.3741 | 9 | 11.8 | 10.3 | 12 | 12.2 | 11.7 | 11.1 | 11.5 | 12.5 | 12.3 |  | 11.5431
| 1 | 5 | 12 | 14.3152 | 12 | 10.6 | 11.3 | 11.1259 | 11.4 | 11 | 12 | 12 | 12.5 | 13 | 12 | 10.7305 | 7.4 | 11.9 | 11.1 | 12 | 11.5 | 10.6 | 11.7 | 11.3 | 12.5 | 12.4 |  | 48
| 1 | 6 | 12 | 14.3936 | 12 | 10.76 | 11.1 | 11.3552 | 10.8 | 11.9 | 12.2 | 11.9 | 11.4 | 12 | 11.3 | 11.7086 | 10.5 | 12.2 | 10.6 | 11 | 11.5 | 10.2 | 11.1 | 11.9 | 12.4 | 12.1
| Mean |  | 12 | 14.3587 | 12 | 10.4867 | 11.4667 | 11.2259 | 11.2667 | 11.3333 | 12.2667 | 12.3833 | 12.05 | 12.8333 | 11.65 | 10.8735 | 10.0167 | 11.9833 | 10.75 | 11.8333 | 11.5167 | 11.1 | 11.3 | 11.6667 | 12.3 | 12.3167
| Median |  | 12 | 14.3585 | 12 | 10.515 | 11.4 | 11.2031 | 11.4 | 11.15 | 12.1 | 12.4 | 12 | 13 | 11.75 | 11.0523 | 10.7 | 11.85 | 10.75 | 12 | 11.5 | 11.15 | 11.25 | 11.65 | 12.35 | 12.35
| Std Dev. |  | 0 | 0.03641 | 0 | 0.22853 | 0.41312 | 0.13254 | 0.46332 | 0.40825 | 0.35024 | 0.42151 | 0.48888 | 0.40825 | 0.38341 | 0.75346 | 1.5171 | 0.25626 | 0.46797 | 0.40825 | 0.39707 | 0.65727 | 0.24495 | 0.2582 | 0.22804 | 0.1169
| Rel.Std.Dev. |  | 0% | 0.2536% | 0% | 2.179% | 3.603% | 1.181% | 4.112% | 3.602% | 2.855% | 3.404% | 4.057% | 3.181% | 3.291% | 6.929% | 15.15% | 2.138% | 4.353% | 3.45% | 3.448% | 5.921% | 2.168% | 2.213% | 1.854% | 0.9492%
| PDM3 |  | 3.959% | 24.39% | 3.959% | -9.152% | -0.6617% | -2.748% | -2.394% | -1.817% | 6.269% | 7.28% | 4.392% | 11.18% | 0.9265% | -5.801% | -13.22% | 3.814% | -6.87% | 2.515% | -0.2286% | -3.838% | -2.106% | 1.071% | 6.558% | 6.702%
| Z-Score (Absolute) |  | NA | 3.96 | 0.67 | 1.44 | 0.07 | 0.41 | 0.35 | 0.26 | 1.05 | 1.21 | 0.74 | NA | 0.19 | 0.9 | 2.09 | 0.65 | 1.07 | NA | 0 | 0.58 | 0.3 | 0.21 | 1.09 | 1.12
Results from laboratories 2.01, 2.12 and 2.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 922b (Certified Value 4.9 ppm) | [26 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 6 | 3.9811 | 7 | < 5 | 4.5 | 4.6891 | 7 | 5 | 5 | 5 | 6 | 4.4 | 5 | 4.492 | 5 | 4.8 | 4 | 5 | 5 | 5 | 2 | 4.8 | 5.1 | 5 |  | 1
| 1 | 2 | 6 | 4.0403 | 7 | < 5 | 4.5 | 5.157 | 6 | 5 | 5 | 5 | 5 | 4.5 | 5 | 4.4281 | 4 | 4.5 | 5.3 | 5.12 | 5 | 5 | 3 | 4.4 | 4.7 | 5 |  | 41
| 1 | 3 | 6 | 3.9617 | 7 | < 5 | 4.6 | 4.9028 | 5 | 5 | 5 | 5 | 5 | 4.6 | 4 | 4.588 | 5 | 4.4 | 4.2 | 5.11 | 5 | 5 | 3 | 4.6 | 4.6 | 5 |  | 16
| 1 | 4 | 6 | 3.9413 | 7 | < 5 | 4.3 | 5.3015 | 6 | 6 | 5 | 5 | 6 | 4.8 | 5 | 4.2208 | 4 | 5.9 | 5.7 | 5.03 | 5 | 5 | 4 | 4.4 | 4.8 | 5 |  | 4.9004
| 1 | 5 | 5 | 3.9015 | 7 | < 5 | 4.9 | 4.9477 | 5 | 6 | 5 | 5 | 6 | 4.8 | 4 | 4.2717 | 3 | 5.1 | 4.2 | 4.99 | 5 | 5 | 2 | 4.4 | 5 | 5 |  | 49
| 1 | 6 | 5 | 4.0106 | 7 | < 5 | 4.9 | 5.1008 | 5 | 6 | 5 | 5 | 5 | 4.3 | 5 | 4.3772 | 5 | 5 | 4.4 | 4.99 | 5 | 5 | 3 | 4.7 | 4.8 | 5
| Mean |  | 5.6667 | 3.9728 | 7 |  | 4.6167 | 5.0165 | 5.6667 | 5.5 | 5 | 5 | 5.5 | 4.5667 | 4.6667 | 4.3963 | 4.3333 | 4.95 | 4.6333 | 5.04 | 5 | 5 | 2.8333 | 4.55 | 4.8333 | 5
| Median |  | 6 | 3.9714 | 7 |  | 4.55 | 5.0242 | 5.5 | 5.5 | 5 | 5 | 5.5 | 4.55 | 5 | 4.4027 | 4.5 | 4.9 | 4.3 | 5.015 | 5 | 5 | 3 | 4.5 | 4.8 | 5
| Std Dev. |  | 0.5164 | 0.049504 | 0 |  | 0.24014 | 0.21582 | 0.8165 | 0.54772 | 0 | 0 | 0.54772 | 0.20656 | 0.5164 | 0.13683 | 0.8165 | 0.53944 | 0.69474 | 0.06 | 0 | 0 | 0.75277 | 0.17607 | 0.18619 | 0
| Rel.Std.Dev. |  | 9.113% | 1.246% | 0% |  | 5.202% | 4.302% | 14.41% | 9.959% | 0% | 0% | 9.959% | 4.523% | 11.07% | 3.112% | 18.84% | 10.9% | 14.99% | 1.19% | 0% | 0% | 26.57% | 3.87% | 3.852% | 0%
| PDM3 |  | 15.64% | -18.93% | 42.84% |  | -5.79% | 2.368% | 15.64% | 12.24% | 2.032% | 2.032% | 12.24% | -6.811% | -4.77% | -10.29% | -11.57% | 1.012% | -5.45% | 2.848% | 2.032% | 2.032% | -42.18% | -7.151% | -1.369% | 2.032%
| Z-Score (Absolute) |  | 1.22 | 1.76 | 3.56 | 4.35 | 0.63 | 0.07 | 1.22 | 0.92 | 0.04 | 0.04 | 0.92 | 0.72 | 0.54 | 1.02 | 1.13 | 0.04 | 0.6 | 0.11 | 0.04 | 0.04 | 3.77 | 0.75 | 0.25 | 0.04
| Analytical results for Sm in OREAS 922b (Certified Value 6.48 ppm) | [26 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 6.1569; 2.02: 6.4; 2.03: 5.695; 2.04: 5.7; 2.05: 6.47; 2.06: 6.2; 2.07: 6.9 |  | 1
| 1 | 2 | 2.01: 6.6781; 2.02: 6.54; 2.03: 5.7735; 2.04: 5.6; 2.05: 6.47; 2.06: 6; 2.07: 6.7 |  | 4
| 1 | 3 | 2.01: 6.4014; 2.02: 6.34; 2.03: 5.6413; 2.04: 4.6; 2.05: 6.36; 2.06: 6.5; 2.07: 6.6 |  | 16
| 1 | 4 | 2.01: 6.5012; 2.02: 6.42; 2.03: 5.7583; 2.04: 6.4; 2.05: 6.35; 2.06: 6.1; 2.07: 6.85 |  | 6.4812
| 1 | 5 | 2.01: 6.5687; 2.02: 6.64; 2.03: 5.7838; 2.04: 4.7; 2.05: 6.38; 2.06: 6.4; 2.07: 6.53 |  | 50
| 1 | 6 | 2.01: 6.2582; 2.02: 6.54; 2.03: 5.8571; 2.04: 5; 2.05: 6.68; 2.06: 6.6; 2.07: 6.9
| Mean |  | 2.01: 6.4274; 2.02: 6.48; 2.03: 5.7515; 2.04: 5.3333; 2.05: 6.4517; 2.06: 6.3; 2.07: 6.7467
| Median |  | 2.01: 6.4513; 2.02: 6.48; 2.03: 5.7659; 2.04: 5.3; 2.05: 6.425; 2.06: 6.3; 2.07: 6.775
| Std Dev. |  | 2.01: 0.19531; 2.02: 0.11171; 2.03: 0.074915; 2.04: 0.69186; 2.05: 0.12384; 2.06: 0.23664; 2.07: 0.16021
| Rel.Std.Dev. |  | 2.01: 3.039%; 2.02: 1.724%; 2.03: 1.303%; 2.04: 12.97%; 2.05: 1.92%; 2.06: 3.756%; 2.07: 2.375%
| PDM3 |  | 2.01: -0.8291%; 2.02: -0.01774%; 2.03: -11.26%; 2.04: -17.71%; 2.05: -0.4549%; 2.06: -2.795%; 2.07: 4.097%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.28; 2.03: 3.58; 2.04: 5.79; 2.05: 0.13; 2.06: 0.67; 2.07: 1.69
| Analytical results for Sn in OREAS 922b (Certified Value 5.76 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.5 | 5.7993 | 7 | 3.96 | 5.7 | 5.695 | 5.9 | 5.8 | 5.9 | 5.6 | 5.8 | 6 | 5.9 | 5.4786 | 4.9 | 5.8 | 5.9 | 5.4 | < 10 | 5.2 | 5.64 | 6.2 | 5.5 | 5.8 |  | 1
| 1 | 2 | 5.4 | 5.6503 | 6.9 | 3.8 | 5.5 | 5.7735 | 6 | 5.7 | 6 | 5.8 | 5.6 | 6 | 5.7 | 5.3722 | 5 | 6.1 | 5.6 | 5.5 | < 10 | 5.8 | 5.49 | 6.1 | 5.5 | 6.1 |  | 26
| 1 | 3 | 5.5 | 5.8075 | 6.7 | 3.83 | 5.7 | 5.6413 | 6.2 | 6 | 5.8 | 5.7 | 5.9 | 6 | 5.7 | 5.5803 | 5 | 6.1 | 6 | 5.4 | < 10 | 5.8 | 5.3 | 6.1 | 5.7 | 5.9 |  | 16
| 1 | 4 | 5.4 | 5.8641 | 7 | 3.86 | 5.5 | 5.7583 | 5.9 | 5.7 | 6.2 | 5.7 | 5.8 | 6 | 5.8 | 5.4446 | 4.1 | 6.1 | 5.6 | 5.6 | < 10 | 5.9 | 5.23 | 6.1 | 5.5 | 5.9 |  | 5.7562
| 1 | 5 | 5.5 | 5.9425 | 7.1 | 4.06 | 5.6 | 5.7838 | 6.6 | 5.8 | 5.9 | 5.5 | 6 | 6 | 5.9 | 5.3898 | 3.3 | 6.1 | 6.1 | 5.4 | < 10 | 5.7 | 5.47 | 6 | 5.7 | 6 |  | 51
| 1 | 6 | 5.5 | 5.8878 | 6.8 | 3.95 | 5.5 | 5.8571 | 6.1 | 6.1 | 5.9 | 5.6 | 5.4 | 6 | 5.8 | 5.6309 | 4.9 | 6.1 | 6.1 | 5.3 | < 10 | 5.9 | 5.38 | 6.1 | 5.8 | 5.8
| Mean |  | 5.4667 | 5.8252 | 6.9167 | 3.91 | 5.5833 | 5.7515 | 6.1167 | 5.85 | 5.95 | 5.65 | 5.75 | 6 | 5.8 | 5.4828 | 4.5333 | 6.05 | 5.8833 | 5.4333 |  | 5.7167 | 5.4183 | 6.1 | 5.6167 | 5.9167
| Median |  | 5.5 | 5.8358 | 6.95 | 3.905 | 5.55 | 5.7659 | 6.05 | 5.8 | 5.9 | 5.65 | 5.8 | 6 | 5.8 | 5.4616 | 4.9 | 6.1 | 5.95 | 5.4 |  | 5.8 | 5.425 | 6.1 | 5.6 | 5.9
| Std Dev. |  | 0.05164 | 0.10076 | 0.1472 | 0.09757 | 0.098319 | 0.074915 | 0.26394 | 0.16432 | 0.13784 | 0.10488 | 0.21679 | 0 | 0.089443 | 0.10374 | 0.69474 | 0.12247 | 0.23166 | 0.10328 |  | 0.26394 | 0.14689 | 0.063246 | 0.13292 | 0.1169
| Rel.Std.Dev. |  | 0.9446% | 1.73% | 2.128% | 2.495% | 1.761% | 1.303% | 4.315% | 2.809% | 2.317% | 1.856% | 3.77% | 0% | 1.542% | 1.892% | 15.33% | 2.024% | 3.938% | 1.901% |  | 4.617% | 2.711% | 1.037% | 2.366% | 1.976%
| PDM3 |  | -5.03% | 1.2% | 20.16% | -32.07% | -3.003% | -0.08149% | 6.262% | 1.63% | 3.367% | -1.845% | -0.1078% | 4.235% | 0.7609% | -4.75% | -21.24% | 5.104% | 2.209% | -5.609% |  | -0.6868% | -5.87% | 5.973% | -2.424% | 2.788%
| Z-Score (Absolute) |  | 0.96 | 0.25 | 3.93 | 6.2 | 0.56 | 0.01 | 1.24 | 0.34 | 0.67 | 0.34 | 0 | NA | 0.17 | 0.9 | 4.1 | 1.01 | 0.45 | 1.07 | 2.53 | 0.11 | 1.12 | 1.18 | 0.45 | 0.56
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 922b (Certified Value 43.1 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 43.5 | 41.48 | 44 | 41.89 | 43.32 | 42.9471 | 45.3 | 42.6 | 42.9 | 42.5 | 45.1 | 44.1 | 41.4 | 39.7847 | 33.3 | 43.46 | 45 | 43 | 41.1 | 39.2 | 43.4 | 44.61 | 39.1 | 45 |  | 1
| 1 | 2 | 43.6 | 42.15 | 44 | 42.11 | 41.74 | 42.9798 | 45.5 | 43 | 44.5 | 44.2 | 45.1 | 46 | 42.9 | 38.7117 | 33.7 | 43.87 | 40 | 44 | 41 | 39 | 42.9 | 44.47 | 40.3 | 45 |  | 27
| 1 | 3 | 43.3 | 42.64 | 44 | 42.78 | 42.1 | 42.9978 | 43.8 | 44.4 | 42.5 | 42.5 | 46.6 | 45.3 | 41.6 | 42.713 | 34.2 | 43.36 | 42 | 44 | 39.7 | 38.7 | 42.2 | 44.75 | 41.3 | 45 |  | 16
| 1 | 4 | 43.6 | 41.38 | 43 | 41.29 | 42.93 | 42.8211 | 44.3 | 42.6 | 45.1 | 42.6 | 45.5 | 45.1 | 42.1 | 41.3966 | 27.5 | 44.27 | 41 | 45 | 41.9 | 38.7 | 41.7 | 43.16 | 40.3 | 46 |  | 43.0725
| 1 | 5 | 43.3 | 43.08 | 43 | 43.1 | 41.82 | 43.5685 | 45.3 | 42 | 43 | 42.5 | 46.6 | 46.6 | 43 | 40.4555 | 22.8 | 44.37 | 45 | 44 | 40.4 | 38.8 | 42.5 | 43.5 | 41.2 | 45 |  | 52
| 1 | 6 | 43.4 | 40.92 | 43 | 42.71 | 41.04 | 43.9717 | 43.5 | 43.9 | 43.2 | 42.3 | 42.2 | 44.7 | 43 | 41.4393 | 33 | 44.9 | 42 | 43 | 40.1 | 38.8 | 42.5 | 44.2 | 41.9 | 46
| Mean |  | 43.45 | 41.9417 | 43.5 | 42.3133 | 42.1583 | 43.2143 | 44.6167 | 43.0833 | 43.5333 | 42.7667 | 45.1833 | 45.3 | 42.3333 | 40.7501 | 30.75 | 44.0383 | 42.5 | 43.8333 | 40.7 | 38.8667 | 42.5333 | 44.115 | 40.6833 | 45.3333
| Median |  | 43.45 | 41.815 | 43.5 | 42.41 | 41.96 | 42.9888 | 44.8 | 42.8 | 43.1 | 42.5 | 45.3 | 45.2 | 42.5 | 40.9261 | 33.15 | 44.07 | 42 | 44 | 40.7 | 38.8 | 42.5 | 44.335 | 40.75 | 45
| Std Dev. |  | 0.13784 | 0.82453 | 0.54772 | 0.67257 | 0.83528 | 0.4532 | 0.86352 | 0.89981 | 1.025 | 0.70899 | 1.6142 | 0.89666 | 0.73121 | 1.4085 | 4.6029 | 0.58806 | 2.0736 | 0.75277 | 0.79246 | 0.19664 | 0.58195 | 0.6437 | 0.99281 | 0.5164
| Rel.Std.Dev. |  | 0.3172% | 1.966% | 1.259% | 1.589% | 1.981% | 1.049% | 1.935% | 2.089% | 2.355% | 1.658% | 3.573% | 1.979% | 1.727% | 3.456% | 14.97% | 1.335% | 4.879% | 1.717% | 1.947% | 0.5059% | 1.368% | 1.459% | 2.44% | 1.139%
| PDM3 |  | 0.8765% | -2.625% | 0.9926% | -1.762% | -2.122% | 0.3294% | 3.585% | 0.02521% | 1.07% | -0.71% | 4.901% | 5.172% | -1.716% | -5.392% | -28.61% | 2.242% | -1.329% | 1.766% | -5.508% | -9.764% | -1.252% | 2.42% | -5.547% | 5.249%
| Z-Score (Absolute) |  | 0.37 | 0.7 | 0.41 | 0.44 | 0.55 | 0.21 | 1.21 | 0.11 | 0.43 | 0.11 | 1.61 | 1.69 | 0.42 | 1.55 | 8.68 | 0.79 | 0.3 | 0.65 | 1.59 | 2.89 | 0.28 | 0.85 | 1.6 | 1.72
| Analytical results for Ta in OREAS 922b (Certified Value 0.46 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.31 | 1.0478 | 0.4 | 0.61 | 0.47 | 0.60123 | 0.47 | 0.53 | 0.36 | 0.4 | 0.51 | 0.8 | 0.31 | 0.36902 | 0.56 | 0.66 | 0.6 | 0.3 | 0.35 | 0.32 | 0.31 | 0.59 | 0.6 |  |  | 1
| 1 | 2 | 0.33 | 1.0535 | 0.4 | < 0.5 | 0.38 | 0.58658 | 0.49 | 0.52 | 0.38 | 0.38 | 0.46 | 0.4 | 0.34 | 0.35212 | 0.55 | 0.53 | 0.5 | 0.3 | 0.38 | 0.3 | 0.35 | 0.57 | 0.6 |  |  | 10
| 1 | 3 | 0.3 | 0.99573 | 0.4 | 0.63 | 0.31 | 0.58265 | 0.49 | 0.52 | 0.35 | 0.42 | 0.51 | 0.3 | 0.33 | 0.48526 | 0.54 | 0.71 | 0.6 | 0.3 | 0.39 | 0.32 | 0.34 | 0.57 | 0.6 |  |  | 16
| 1 | 4 | 0.32 | 1.0685 | 0.4 | 0.62 | 0.5 | 0.5232 | 0.48 | 0.48 | 0.35 | 0.41 | 0.44 | 0.5 | 0.34 | 0.42445 | 0.44 | 0.66 | 0.7 | 0.4 | 0.37 | 0.34 | 0.34 | 0.57 | 0.5 |  |  | 0.45637
| 1 | 5 | 0.31 | 1.0384 | 0.5 | 0.72 | 0.57 | 0.60044 | 0.49 | 0.54 | 0.34 | 0.43 | 0.52 | 0.6 | 0.31 | 0.37828 | 0.34 | 0.56 | 0.9 | 0.4 | 0.39 | 0.36 | 0.33 | 0.52 | 0.6 |  |  | 53
| 1 | 6 | 0.29 | 0.98894 | 0.4 | 0.74 | 0.31 | 0.54448 | 0.51 | 0.55 | 0.37 | 0.44 | 0.49 | 0.7 | 0.34 | 0.44166 | 0.53 | 0.67 | 0.9 | 0.3 | 0.41 | 0.48 | 0.34 | 0.57 | 0.5
| Mean |  | 0.31 | 1.0321 | 0.41667 | 0.664 | 0.42333 | 0.5731 | 0.48833 | 0.52333 | 0.35833 | 0.41333 | 0.48833 | 0.55 | 0.32833 | 0.40847 | 0.49333 | 0.63167 | 0.7 | 0.33333 | 0.38167 | 0.35333 | 0.335 | 0.565 | 0.56667
| Median |  | 0.31 | 1.0431 | 0.4 | 0.63 | 0.425 | 0.58461 | 0.49 | 0.525 | 0.355 | 0.415 | 0.5 | 0.55 | 0.335 | 0.40137 | 0.535 | 0.66 | 0.65 | 0.3 | 0.385 | 0.33 | 0.34 | 0.57 | 0.6
| Std Dev. |  | 0.014142 | 0.032422 | 0.040825 | 0.061074 | 0.1069 | 0.032 | 0.013292 | 0.024221 | 0.01472 | 0.021602 | 0.031885 | 0.18708 | 0.01472 | 0.050786 | 0.086641 | 0.070261 | 0.16733 | 0.05164 | 0.020412 | 0.06532 | 0.013784 | 0.023452 | 0.05164
| Rel.Std.Dev. |  | 4.562% | 3.141% | 9.798% | 9.198% | 25.25% | 5.584% | 2.722% | 4.628% | 4.108% | 5.226% | 6.529% | 34.02% | 4.483% | 12.43% | 17.56% | 11.12% | 23.9% | 15.49% | 5.348% | 18.49% | 4.115% | 4.151% | 9.113%
| PDM3 |  | -32.07% | 126.2% | -8.699% | 45.5% | -7.238% | 25.58% | 7.005% | 14.67% | -21.48% | -9.43% | 7.005% | 20.52% | -28.05% | -10.5% | 8.1% | 38.41% | 53.39% | -26.96% | -16.37% | -22.58% | -26.59% | 23.8% | 24.17%
| Z-Score (Absolute) |  | 0.98 | 3.89 | NA | 0.94 | 0.22 | 0.79 | 0.22 | 0.46 | 0.66 | 0.29 | 0.22 | NA | 0.86 | 0.32 | 0.25 | 1.19 | NA | NA | 0.5 | 0.69 | 0.81 | 0.74 | NA
Results from laboratories 2.03, 2.12, 2.17, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 922b (Certified Value 0.63 ppm) | [26 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.16; 2.08: 2.17; 2.09: 2.20; 2.10: 2.21; 2.11: 2.22 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*OES/MS; 2.11: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 0.566; 2.02: 0.65; 2.03: 0.83; 2.04: 0.62881; 2.05: 0.7; 2.06: 0.6; 2.07: 0.65; 2.08: 0.6; 2.09: 0.57; 2.10: 0.55; 2.11: 0.67 |  |  | 1
| 1 | 2 | 2.01: 0.571; 2.02: 0.65; 2.03: 0.81; 2.04: 0.6785; 2.05: 0.7; 2.06: 0.61; 2.07: 0.93; 2.08: 0.5; 2.09: 0.54; 2.10: 0.56; 2.11: 0.71 |  |  | 11
| 1 | 3 | 2.01: 0.5836; 2.02: 0.65; 2.03: 0.83; 2.04: 0.63232; 2.05: 0.7; 2.06: 0.54; 2.07: 0.66; 2.08: 0.6; 2.09: 0.54; 2.10: 0.58; 2.11: 0.65 |  |  | 16
| 1 | 4 | 2.01: 0.5569; 2.02: 0.65; 2.03: 0.8; 2.04: 0.61045; 2.05: 0.8; 2.06: 0.56; 2.07: 0.63; 2.08: 0.5; 2.09: 0.58; 2.10: 0.55; 2.11: 0.65 |  |  | 0.6339
| 1 | 5 | 2.01: 0.5387; 2.02: 0.65; 2.03: 0.85; 2.04: 0.64664; 2.05: 0.8; 2.06: 0.6; 2.07: 0.63; 2.08: 0.6; 2.09: 0.54; 2.10: 0.58; 2.11: 0.67 |  |  | 54
| 1 | 6 | 2.01: 0.6148; 2.02: 0.65; 2.03: 0.84; 2.04: 0.63874; 2.05: 0.8; 2.06: 0.58; 2.07: 0.65; 2.08: 0.6; 2.09: 0.52; 2.10: 0.58; 2.11: 0.71
| Mean |  | 2.01: 0.57183; 2.02: 0.65; 2.03: 0.82667; 2.04: 0.63924; 2.05: 0.75; 2.06: 0.58167; 2.07: 0.69167; 2.08: 0.56667; 2.09: 0.54833; 2.10: 0.56667; 2.11: 0.67667
| Median |  | 2.01: 0.5685; 2.02: 0.65; 2.03: 0.83; 2.04: 0.63553; 2.05: 0.75; 2.06: 0.59; 2.07: 0.65; 2.08: 0.6; 2.09: 0.54; 2.10: 0.57; 2.11: 0.67
| Std Dev. |  | 2.01: 0.025846; 2.02: 0; 2.03: 0.018619; 2.04: 0.022723; 2.05: 0.054772; 2.06: 0.027142; 2.07: 0.11737; 2.08: 0.05164; 2.09: 0.022286; 2.10: 0.015055; 2.11: 0.027325
| Rel.Std.Dev. |  | 2.01: 4.52%; 2.02: 0%; 2.03: 2.252%; 2.04: 3.555%; 2.05: 7.303%; 2.06: 4.666%; 2.07: 16.97%; 2.08: 9.113%; 2.09: 4.064%; 2.10: 2.657%; 2.11: 4.038%
| PDM3 |  | 2.01: -9.791%; 2.02: 2.54%; 2.03: 30.41%; 2.04: 0.8434%; 2.05: 18.32%; 2.06: -8.24%; 2.07: 9.113%; 2.08: -10.61%; 2.09: -13.5%; 2.10: -10.61%; 2.11: 6.747%
| Z-Score (Absolute) |  | 2.01: 0.79; 2.02: 0.13; 2.03: 2.2; 2.04: 0; 2.05: NA; 2.06: 0.67; 2.07: 0.61; 2.08: NA; 2.09: 1.06; 2.10: 0.85; 2.11: 0.44
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 922b (Certified Value 0.084 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | < 0.1 | 0.1 | 2.26 | < 0.2 | 0.07 | 0.07 | 0.08 | 0.08 | 0.1 | < 0.1 | 0.06 | < 0.1 | 0.12 | < 0.05 | 0.08 | < 0.5 | 4.15 | 0.06 | 0.09 | < 0.1 | [5 empty] | 1
| 1 | 2 | < 0.1 | 0.1 | 2.16 | < 0.2 | 0.08 | 0.11 | 0.1 | 0.07 | 0.05 | < 0.1 | 0.08 | < 0.1 | 0.04 | 0.09 | 0.07 | < 0.5 | 4.33 | 0.06 | 0.08 | 0.1 | [5 empty] | 12
| 1 | 3 | < 0.1 | 0.1 | 2.2 | < 0.2 | 0.07 | 0.07 | 0.09 | 0.08 | 0.11 | < 0.1 | 0.08 | < 0.1 | 0.29 | 0.11 | 0.07 | < 0.5 | 4.55 | 0.06 | 0.1 | 0.1 | [5 empty] | 16
| 1 | 4 | < 0.1 | 0.1 | 4.24 | < 0.2 | 0.06 | 0.07 | 0.1 | 0.09 | 0.1 | < 0.1 | 0.07 | < 0.1 | 0.93 | 0.12 | 0.07 | < 0.5 | 4.45 | 0.07 | 0.09 | 0.1 | [5 empty] | 0.083654
| 1 | 5 | < 0.1 | 0.1 | 2.3 | < 0.2 | 0.06 | 0.07 | 0.08 | 0.07 | 0.09 | < 0.1 | 0.07 | < 0.1 | 0.03 | 0.08 | 0.08 | < 0.5 | 4.26 | 0.07 | 0.08 | < 0.1 | [5 empty] | 55
| 1 | 6 | < 0.1 | 0.1 | 2.33 | < 0.2 | 0.06 | 0.07 | 0.07 | 0.07 | 0.1 | < 0.1 | 0.1 | < 0.1 | 0.12 | 0.1 | 0.07 | < 0.5 | 4.32 | 0.09 | 0.11 | < 0.1
| Mean |  |  | 0.1 | 2.5817 |  | 0.066667 | 0.076667 | 0.086667 | 0.076667 | 0.091667 |  | 0.076667 |  | 0.255 | 0.1 | 0.073333 |  | 4.3433 | 0.068333 | 0.091667 | 0.1
| Median |  |  | 0.1 | 2.28 |  | 0.065 | 0.07 | 0.085 | 0.075 | 0.1 |  | 0.075 |  | 0.12 | 0.1 | 0.07 |  | 4.325 | 0.065 | 0.09 | 0.1
| Std Dev. |  |  | 1.5202e-17 | 0.81482 |  | 0.008165 | 0.01633 | 0.012111 | 0.008165 | 0.02137 |  | 0.013663 |  | 0.34355 | 0.015811 | 0.005164 |  | 0.14081 | 0.01169 | 0.01169 | 1.6997e-17
| Rel.Std.Dev. |  |  | 1.52e-14% | 31.56% |  | 12.25% | 21.3% | 13.97% | 10.65% | 23.31% |  | 17.82% |  | 134.7% | 15.81% | 7.042% |  | 3.242% | 17.11% | 12.75% | 1.7e-14%
| PDM3 |  |  | 19.54% | 2986% |  | -20.31% | -8.352% | 3.602% | -8.352% | 9.579% |  | -8.352% |  | 204.8% | 19.54% | -12.34% |  | 5092% | -18.31% | 9.579% | 19.54%
| Z-Score (Absolute) |  | 1.51 | 0.87 | 119 | 0.87 | 0.71 | 0.24 | 0.24 | 0.24 | 0.48 | 1.51 | 0.24 | 1.51 | 8.25 | 0.28 | 0.4 | 8.01 | 202.85 | 0.63 | 0.48 | 0.32
| Analytical results for Th in OREAS 922b (Certified Value 14.3 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 14.4 | 14.432 | 13.3 | 15.01 | 14.03 | 14.551 | 13.85 | 13.45 | 12.7 | 13.85 | 14.1 | 16.1 | 15.05 | 14.8959 | < 20 | 14.24 | 12.9 | 15.6 | 14.51 | 12.9 | 14.64 | 14.41 | 13.2 | 15.7 |  | 1
| 1 | 2 | 14.4 | 14.323 | 13.5 | 14.91 | 13.98 | 14.7729 | 13.75 | 13.95 | 12.5 | 14.5 | 13.5 | 16.6 | 14.26 | 14.8829 | < 20 | 14.55 | 12 | 15.3 | 14.43 | 12.7 | 14.4 | 14.53 | 13.5 | 15.7 |  | 13
| 1 | 3 | 14.3 | 14.537 | 13.9 | 15.25 | 14.22 | 14.3832 | 14.7 | 14.55 | 13 | 14.1 | 14.8 | 15.8 | 13.54 | 15.645 | < 20 | 14.23 | 13.4 | 15.7 | 14.99 | 13 | 14.07 | 14.51 | 13.8 | 15.5 |  | 16
| 1 | 4 | 14.4 | 14.359 | 13.6 | 14.83 | 13.89 | 13.9155 | 13.9 | 13.55 | 14.25 | 14 | 13.85 | 16.4 | 14 | 15.2935 | < 20 | 14.55 | 12.2 | 15.9 | 15.39 | 13.1 | 13.85 | 14.21 | 13.5 | 15.2 |  | 14.2597
| 1 | 5 | 14.4 | 14.334 | 13.9 | 15.15 | 14.1 | 14.4208 | 15.2 | 13.75 | 13.25 | 14.1 | 14.65 | 16.8 | 14.86 | 15.4697 | < 20 | 14.38 | 13.4 | 15.5 | 15.71 | 13.1 | 14.15 | 14.56 | 13.8 | 15.5 |  | 56
| 1 | 6 | 14.5 | 14.619 | 13.8 | 15.09 | 14.06 | 14.7749 | 14.5 | 14.25 | 13.85 | 14.15 | 13.15 | 16.5 | 13.56 | 15.5535 | < 20 | 14.78 | 13.2 | 15.2 | 14.61 | 13.4 | 14.39 | 14.34 | 14.1 | 14.8
| Mean |  | 14.4 | 14.434 | 13.6667 | 15.04 | 14.0467 | 14.4697 | 14.3167 | 13.9167 | 13.2583 | 14.1167 | 14.0083 | 16.3667 | 14.2117 | 15.2901 |  | 14.455 | 12.85 | 15.5333 | 14.94 | 13.0333 | 14.25 | 14.4267 | 13.65 | 15.4
| Median |  | 14.4 | 14.3955 | 13.7 | 15.05 | 14.045 | 14.4859 | 14.2 | 13.85 | 13.125 | 14.1 | 13.975 | 16.45 | 14.13 | 15.3816 |  | 14.465 | 13.05 | 15.55 | 14.8 | 13.05 | 14.27 | 14.46 | 13.65 | 15.5
| Std Dev. |  | 0.063246 | 0.12065 | 0.24221 | 0.15531 | 0.11165 | 0.31895 | 0.5785 | 0.42269 | 0.67633 | 0.21602 | 0.64297 | 0.36148 | 0.6397 | 0.33131 |  | 0.21267 | 0.61237 | 0.2582 | 0.51996 | 0.23381 | 0.28164 | 0.13426 | 0.31464 | 0.34641
| Rel.Std.Dev. |  | 0.4392% | 0.8359% | 1.772% | 1.033% | 0.7949% | 2.204% | 4.041% | 3.037% | 5.101% | 1.53% | 4.59% | 2.209% | 4.501% | 2.167% |  | 1.471% | 4.766% | 1.662% | 3.48% | 1.794% | 1.976% | 0.9307% | 2.305% | 2.249%
| PDM3 |  | 0.9838% | 1.222% | -4.159% | 5.472% | -1.494% | 1.473% | 0.3994% | -2.406% | -7.022% | -1.003% | -1.763% | 14.78% | -0.337% | 7.226% |  | 1.369% | -9.886% | 8.932% | 4.771% | -8.6% | -0.06815% | 1.171% | -4.276% | 7.997%
| Z-Score (Absolute) |  | 0.16 | 0.21 | 0.85 | 1.04 | 0.32 | 0.26 | 0.05 | 0.5 | 1.41 | 0.23 | 0.38 | 2.86 | 0.1 | 1.38 | 5.87 | 0.24 | 1.97 | 1.71 | 0.9 | 1.71 | 0.05 | 0.2 | 0.87 | 1.53
| Analytical results for Ti in OREAS 922b (Certified Value 0.224 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.212 | 0.21023 | 0.21 | 0.28 | 0.228 | 0.22076 | 0.251 | 0.255 | 0.222 | 0.239 | 0.24 | 0.321 | 0.197 | 0.20258 | 0.244 | 0.2706 | 0.229 | 0.223 | 0.2 | 0.2 | 0.23 | 0.2058 | 0.25 | 0.237 |  | 1
| 1 | 2 | 0.21 | 0.20659 | 0.212 | 0.25 | 0.203 | 0.21478 | 0.253 | 0.259 | 0.226 | 0.24 | 0.241 | 0.326 | 0.204 | 0.20141 | 0.242 | 0.2474 | 0.21 | 0.227 | 0.19 | 0.19 | 0.23 | 0.204 | 0.25 | 0.235 |  | 32
| 1 | 3 | 0.213 | 0.20775 | 0.21 | 0.28 | 0.2075 | 0.21846 | 0.241 | 0.254 | 0.216 | 0.238 | 0.237 | 0.324 | 0.206 | 0.216 | 0.251 | 0.2773 | 0.212 | 0.23 | 0.19 | 0.2 | 0.23 | 0.2159 | 0.26 | 0.225 |  | 16
| 1 | 4 | 0.226 | 0.20773 | 0.21 | 0.27 | 0.2399 | 0.21719 | 0.248 | 0.248 | 0.217 | 0.243 | 0.237 | 0.325 | 0.2 | 0.21341 | 0.195 | 0.2634 | 0.216 | 0.232 | 0.2 | 0.2 | 0.23 | 0.2038 | 0.24 | 0.223 |  | 0.22422
| 1 | 5 | 0.212 | 0.21042 | 0.21 | 0.28 | 0.2024 | 0.2185 | 0.25 | 0.258 | 0.22 | 0.231 | 0.239 | 0.323 | 0.209 | 0.20798 | 0.16 | 0.2555 | 0.238 | 0.231 | 0.19 | 0.2 | 0.22 | 0.202 | 0.26 | 0.236 |  | 57
| 1 | 6 | 0.222 | 0.20813 | 0.21 | 0.28 | 0.1964 | 0.22372 | 0.245 | 0.261 | 0.225 | 0.235 | 0.248 | 0.317 | 0.209 | 0.21675 | 0.238 | 0.2712 | 0.219 | 0.215 | 0.19 | 0.19 | 0.23 | 0.2085 | 0.25 | 0.231
| Mean |  | 0.21583 | 0.20847 | 0.21033 | 0.27333 | 0.21287 | 0.2189 | 0.248 | 0.25583 | 0.221 | 0.23767 | 0.24033 | 0.32267 | 0.20417 | 0.20969 | 0.22167 | 0.26423 | 0.22067 | 0.22633 | 0.19333 | 0.19667 | 0.22833 | 0.20667 | 0.25167 | 0.23117
| Median |  | 0.2125 | 0.20794 | 0.21 | 0.28 | 0.20525 | 0.21848 | 0.249 | 0.2565 | 0.221 | 0.2385 | 0.2395 | 0.3235 | 0.205 | 0.2107 | 0.24 | 0.267 | 0.2175 | 0.2285 | 0.19 | 0.2 | 0.23 | 0.2049 | 0.25 | 0.233
| Std Dev. |  | 0.0065243 | 0.0015255 | 0.0008165 | 0.012111 | 0.017127 | 0.0030636 | 0.0043818 | 0.0046224 | 0.0040988 | 0.0041793 | 0.0040825 | 0.003266 | 0.0048751 | 0.0067174 | 0.036203 | 0.011137 | 0.010801 | 0.0064395 | 0.005164 | 0.005164 | 0.0040825 | 0.0050294 | 0.0075277 | 0.005947
| Rel.Std.Dev. |  | 3.023% | 0.7317% | 0.3882% | 4.431% | 8.046% | 1.4% | 1.767% | 1.807% | 1.855% | 1.758% | 1.699% | 1.012% | 2.388% | 3.204% | 16.33% | 4.215% | 4.895% | 2.845% | 2.671% | 2.626% | 1.788% | 2.434% | 2.991% | 2.573%
| PDM3 |  | -3.739% | -7.021% | -6.192% | 21.91% | -5.062% | -2.371% | 10.61% | 14.1% | -1.435% | 5.998% | 7.188% | 43.91% | -8.943% | -6.48% | -1.138% | 17.85% | -1.584% | 0.9437% | -13.77% | -12.29% | 1.836% | -7.828% | 12.24% | 3.099%
| Z-Score (Absolute) |  | 0.27 | 0.63 | 0.54 | 2.55 | 0.41 | 0.12 | 1.31 | 1.69 | 0.02 | 0.8 | 0.93 | 4.96 | 0.84 | 0.57 | 0.02 | 2.1 | 0.03 | 0.24 | 1.37 | 1.21 | 0.34 | 0.72 | 1.49 | 0.48
| Analytical results for Tl in OREAS 922b (Certified Value 0.94 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.9 | 0.8654 | 1 | < 0.5 | 0.99 | 0.98236 | 0.94 | 0.95 | 0.9 | 0.91 | 0.91 | 0.96 | 0.98 | 0.95376 | 0.99 | 1.04 | 0.89 | 0.888 | 1 | 0.84 | 1.14 | 0.94 | 0.94 | 0.7 |  | 1
| 1 | 2 | 0.94 | 0.8832 | 1 | < 0.5 | 0.97 | 0.99019 | 0.93 | 0.95 | 0.89 | 0.93 | 0.87 | 0.99 | 0.95 | 0.94013 | 0.96 | 1 | 0.86 | 0.88 | 1 | 0.9 | 1.17 | 0.98 | 0.94 | 0.9 |  | 14
| 1 | 3 | 0.93 | 0.8648 | 1 | < 0.5 | 0.99 | 0.96633 | 0.98 | 1 | 0.91 | 0.92 | 0.97 | 0.97 | 0.93 | 0.98899 | 1.01 | 1.01 | 0.97 | 0.891 | 1 | 0.92 | 1.06 | 0.99 | 0.95 | 0.9 |  | 16
| 1 | 4 | 0.95 | 0.8581 | 1 | < 0.5 | 0.94 | 0.98548 | 0.92 | 0.94 | 0.91 | 0.93 | 0.85 | 0.99 | 0.95 | 0.94784 | 0.87 | 1 | 0.89 | 0.896 | 1 | 0.9 | 1.07 | 0.94 | 0.94 | 0.9 |  | 0.94234
| 1 | 5 | 0.93 | 0.8369 | 1 | < 0.5 | 0.96 | 0.99765 | 1.03 | 0.94 | 0.88 | 0.92 | 0.94 | 1.01 | 0.98 | 0.95169 | 0.65 | 1.04 | 0.92 | 0.882 | 1 | 0.92 | 1.11 | 0.97 | 0.95 | 0.8 |  | 58
| 1 | 6 | 0.9 | 0.9007 | 1 | < 0.5 | 0.97 | 0.97325 | 0.98 | 0.98 | 0.91 | 0.92 | 0.95 | 0.97 | 0.94 | 0.99303 | 1 | 1.06 | 0.92 | 0.862 | 1 | 0.9 | 1.14 | 0.97 | 0.98 | 0.9
| Mean |  | 0.925 | 0.86818 | 1 |  | 0.97 | 0.98254 | 0.96333 | 0.96 | 0.9 | 0.92167 | 0.915 | 0.98167 | 0.955 | 0.96257 | 0.91333 | 1.025 | 0.90833 | 0.88317 | 1 | 0.89667 | 1.115 | 0.965 | 0.95 | 0.85
| Median |  | 0.93 | 0.8651 | 1 |  | 0.97 | 0.98392 | 0.96 | 0.95 | 0.905 | 0.92 | 0.925 | 0.98 | 0.95 | 0.95272 | 0.975 | 1.025 | 0.905 | 0.885 | 1 | 0.9 | 1.125 | 0.97 | 0.945 | 0.9
| Std Dev. |  | 0.020736 | 0.021828 | 0 |  | 0.018974 | 0.011352 | 0.041312 | 0.024495 | 0.012649 | 0.0075277 | 0.047223 | 0.018348 | 0.020736 | 0.02255 | 0.13866 | 0.0251 | 0.037639 | 0.011907 | 0 | 0.029439 | 0.043243 | 0.020736 | 0.015492 | 0.083666
| Rel.Std.Dev. |  | 2.242% | 2.514% | 0% |  | 1.956% | 1.155% | 4.288% | 2.552% | 1.405% | 0.8168% | 5.161% | 1.869% | 2.171% | 2.343% | 15.18% | 2.449% | 4.144% | 1.348% | 0% | 3.283% | 3.878% | 2.149% | 1.631% | 9.843%
| PDM3 |  | -1.84% | -7.869% | 6.119% |  | 2.935% | 4.266% | 2.228% | 1.874% | -4.493% | -2.194% | -2.901% | 4.173% | 1.343% | 2.147% | -3.078% | 8.772% | -3.609% | -6.279% | 6.119% | -4.847% | 18.32% | 2.405% | 0.8128% | -9.799%
| Z-Score (Absolute) |  | 0.52 | 1.7 | NA | 14.5 | 0.41 | 0.67 | 0.28 | 0.21 | 1.04 | 0.59 | 0.73 | 0.66 | 0.1 | 0.26 | 0.76 | 1.55 | 0.86 | 1.38 | NA | 1.11 | 3.42 | 0.31 | 0 | NA
Results from laboratories 2.03, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 922b (Certified Value 0.23 ppm) | [26 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS |  | 2
2
| 1 | 1 | 2.01: 0.1963; 2.02: 0.252; 2.03: 0.24603; 2.04: 0.3; 2.05: 0.27; 2.06: 0.2; 2.07: 0.21 |  | 1
| 1 | 2 | 2.01: 0.1759; 2.02: 0.252; 2.03: 0.25226; 2.04: 0.3; 2.05: 0.25; 2.06: 0.2; 2.07: 0.19 |  | 5
| 1 | 3 | 2.01: 0.184; 2.02: 0.252; 2.03: 0.25136; 2.04: 0.3; 2.05: 0.27; 2.06: 0.2; 2.07: 0.2 |  | 16
| 1 | 4 | 2.01: 0.1728; 2.02: 0.252; 2.03: 0.23732; 2.04: 0.3; 2.05: 0.27; 2.06: 0.2; 2.07: 0.19 |  | 0.2275
| 1 | 5 | 2.01: 0.1667; 2.02: 0.252; 2.03: 0.24484; 2.04: 0.3; 2.05: 0.26; 2.06: 0.2; 2.07: 0.19 |  | 59
| 1 | 6 | 2.01: 0.1794; 2.02: 0.252; 2.03: 0.23622; 2.04: 0.3; 2.05: 0.27; 2.06: 0.2; 2.07: 0.2
| Mean |  | 2.01: 0.17918; 2.02: 0.252; 2.03: 0.24467; 2.04: 0.3; 2.05: 0.265; 2.06: 0.2; 2.07: 0.19667
| Median |  | 2.01: 0.17765; 2.02: 0.252; 2.03: 0.24543; 2.04: 0.3; 2.05: 0.27; 2.06: 0.2; 2.07: 0.195
| Std Dev. |  | 2.01: 0.010233; 2.02: 0; 2.03: 0.0067778; 2.04: 0; 2.05: 0.0083666; 2.06: 3.0405e-17; 2.07: 0.008165
| Rel.Std.Dev. |  | 2.01: 5.711%; 2.02: 0%; 2.03: 2.77%; 2.04: 0%; 2.05: 3.157%; 2.06: 1.52e-14%; 2.07: 4.152%
| PDM3 |  | 2.01: -21.24%; 2.02: 10.77%; 2.03: 7.546%; 2.04: 31.87%; 2.05: 16.48%; 2.06: -12.09%; 2.07: -13.55%
| Z-Score (Absolute) |  | 2.01: 2.17; 2.02: 0.24; 2.03: 0; 2.04: NA; 2.05: 0.67; 2.06: NA; 2.07: 1.59
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 922b (Certified Value 4.29 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 4.35 | 3.7081 | 4.2 | 4.62 | 4.55 | 4.3383 | 4.1 | 3.8 | 3.8 | 4.5 | 4.3 | 4.5 | 4.46 | 4.0587 | 4.27 | 4.4 | 4.2 | 4.16 | 4.5 | 4.3 | 3.82 | 4.42 | 4.57 | 3.9 | 4.5 | 1
| 1 | 2 | 4.37 | 3.8574 | 4.2 | 4.58 | 4.39 | 4.4504 | 4.1 | 3.8 | 3.9 | 4.6 | 4.2 | 4.6 | 4.23 | 4.0604 | 4.23 | 4.45 | 3.9 | 4.16 | 4.5 | 4.2 | 3.8 | 4.22 | 4.82 | 3.8 | 4.4 | 16
| 1 | 3 | 4.33 | 3.6626 | 4.3 | 4.68 | 4.46 | 4.3006 | 4.4 | 4.1 | 3.9 | 4.6 | 4.5 | 4.5 | 4.1 | 4.208 | 4.33 | 4.47 | 4.3 | 4.16 | 4.4 | 4.2 | 3.9 | 4.2 | 4.68 | 4.2 | 4.4 | 16
| 1 | 4 | 4.41 | 3.6562 | 4.3 | 4.57 | 4.31 | 4.1292 | 4 | 3.8 | 4.2 | 4.5 | 4.2 | 4.5 | 4.25 | 4.2339 | 3.5 | 4.49 | 4.1 | 4.21 | 4.6 | 4.4 | 3.85 | 4.17 | 4.64 | 4.1 | 4.4 | 4.2872
| 1 | 5 | 4.34 | 3.6816 | 4.3 | 4.66 | 4.36 | 4.3652 | 4.5 | 3.8 | 4 | 4.5 | 4.6 | 4.6 | 4.45 | 4.2323 | 2.86 | 4.46 | 4.5 | 4.11 | 4.5 | 4.3 | 3.92 | 4.25 | 4.71 | 4 | 4.4 | 60
| 1 | 6 | 4.37 | 3.6226 | 4.3 | 4.66 | 4.29 | 4.3524 | 4.4 | 4 | 4.1 | 4.5 | 4.1 | 4.6 | 4.18 | 4.2609 | 4.19 | 4.61 | 4.3 | 4.08 | 4.3 | 4.2 | 3.94 | 4.22 | 4.57 | 4 | 4.3
| Mean |  | 4.3617 | 3.6981 | 4.2667 | 4.6283 | 4.3933 | 4.3227 | 4.25 | 3.8833 | 3.9833 | 4.5333 | 4.3167 | 4.55 | 4.2783 | 4.1757 | 3.8967 | 4.48 | 4.2167 | 4.1467 | 4.4667 | 4.2667 | 3.8717 | 4.2467 | 4.665 | 4 | 4.4
| Median |  | 4.36 | 3.6721 | 4.3 | 4.64 | 4.375 | 4.3453 | 4.25 | 3.8 | 3.95 | 4.5 | 4.25 | 4.55 | 4.24 | 4.2201 | 4.21 | 4.465 | 4.25 | 4.16 | 4.5 | 4.25 | 3.875 | 4.22 | 4.66 | 4 | 4.4
| Std Dev. |  | 0.028577 | 0.083024 | 0.05164 | 0.045789 | 0.097707 | 0.10693 | 0.20736 | 0.13292 | 0.1472 | 0.05164 | 0.19408 | 0.054772 | 0.14634 | 0.091518 | 0.59268 | 0.070427 | 0.20412 | 0.045461 | 0.10328 | 0.08165 | 0.056716 | 0.088919 | 0.094816 | 0.14142 | 0.063246
| Rel.Std.Dev. |  | 0.6552% | 2.245% | 1.21% | 0.9893% | 2.224% | 2.474% | 4.879% | 3.423% | 3.695% | 1.139% | 4.496% | 1.204% | 3.421% | 2.192% | 15.21% | 1.572% | 4.841% | 1.096% | 2.312% | 1.914% | 1.465% | 2.094% | 2.032% | 3.536% | 1.437%
| PDM3 |  | 1.737% | -13.74% | -0.4792% | 7.957% | 2.475% | 0.8275% | -0.868% | -9.421% | -7.088% | 5.741% | 0.687% | 6.13% | -0.2071% | -2.601% | -9.11% | 4.497% | -1.645% | -3.278% | 4.186% | -0.4792% | -9.693% | -0.9457% | 8.812% | -6.699% | 2.631%
| Z-Score (Absolute) |  | 0.48 | 2.88 | 0 | 1.83 | 0.64 | 0.28 | 0.08 | 1.94 | 1.43 | 1.35 | 0.25 | 1.43 | 0.06 | 0.46 | 1.87 | 1.08 | 0.25 | 0.61 | 1.01 | 0 | 2 | 0.1 | 2.01 | 1.35 | 0.67
| Analytical results for V in OREAS 922b (Certified Value 184 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 190 | 162.18 | 184 | 181.4 | 184 | 173.3664 | 185 | 189 | 186 | 184 | 171 | 197 | 190 | 177.0115 | 181 | 194 | 184 | 194 | 180 | 189 | 170 | 192 | 186 | 174 | 193 | 1
| 1 | 2 | 190 | 161.06 | 184 | 179.2 | 177 | 175.3418 | 185 | 192 | 187 | 186 | 176 | 208 | 197 | 172.1673 | 182 | 187 | 170 | 194 | 181 | 190 | 170 | 191 | 185 | 169 | 194 | 36
| 1 | 3 | 193 | 161.73 | 184 | 185.6 | 186 | 174.6634 | 184 | 192 | 182 | 188 | 168 | 196 | 192 | 188.3226 | 184 | 192 | 181 | 189 | 180 | 180 | 170 | 190 | 181 | 178 | 193 | 16
| 1 | 4 | 192 | 159.56 | 183 | 185 | 171 | 173.2478 | 185 | 187 | 187 | 192 | 172 | 198 | 196 | 175.653 | 147 | 190 | 172 | 192 | 187 | 190 | 172 | 192 | 176 | 173 | 193 | 183.6626
| 1 | 5 | 193 | 158.05 | 183 | 183.7 | 187 | 175.8612 | 193 | 193 | 185 | 187 | 170 | 211 | 199 | 179.6761 | 121 | 188 | 185 | 189 | 181 | 178 | 170 | 190 | 182 | 173 | 192 | 61
| 1 | 6 | 190 | 163.96 | 183 | 184.9 | 184 | 174.6627 | 187 | 193 | 185 | 182 | 174 | 198 | 200 | 191.0711 | 177 | 194 | 177 | 189 | 179 | 185 | 171 | 192 | 187 | 170 | 191
| Mean |  | 191.3333 | 161.09 | 183.5 | 183.3 | 181.5 | 174.5239 | 186.5 | 191 | 185.3333 | 186.5 | 171.8333 | 201.3333 | 195.6667 | 180.6502 | 165.3333 | 190.8333 | 178.1667 | 191.1667 | 181.3333 | 185.3333 | 170.5 | 191.1667 | 182.8333 | 172.8333 | 192.6667
| Median |  | 191 | 161.395 | 183.5 | 184.3 | 184 | 174.663 | 185 | 192 | 185.5 | 186.5 | 171.5 | 198 | 196.5 | 178.3438 | 179 | 191 | 179 | 190.5 | 180.5 | 187 | 170 | 191.5 | 183.5 | 173 | 193
| Std Dev. |  | 1.5055 | 2.0697 | 0.54772 | 2.5028 | 6.2209 | 1.0454 | 3.3317 | 2.4495 | 1.8619 | 3.4496 | 2.8577 | 6.4395 | 3.9328 | 7.4643 | 25.7268 | 2.9944 | 6.2423 | 2.4833 | 2.8752 | 5.2789 | 0.83666 | 0.98319 | 4.0702 | 3.1885 | 1.0328
| Rel.Std.Dev. |  | 0.7869% | 1.285% | 0.2985% | 1.365% | 3.428% | 0.599% | 1.786% | 1.282% | 1.005% | 1.85% | 1.663% | 3.198% | 2.01% | 4.132% | 15.56% | 1.569% | 3.504% | 1.299% | 1.586% | 2.848% | 0.4907% | 0.5143% | 2.226% | 1.845% | 0.5361%
| PDM3 |  | 4.177% | -12.29% | -0.08851% | -0.1974% | -1.177% | -4.976% | 1.545% | 3.995% | 0.9097% | 1.545% | -6.441% | 9.621% | 6.536% | -1.64% | -9.98% | 3.904% | -2.992% | 4.086% | -1.268% | 0.9097% | -7.167% | 4.086% | -0.4515% | -5.896% | 4.903%
| Z-Score (Absolute) |  | 0.7 | 2.01 | 0 | 0.02 | 0.18 | 0.81 | 0.27 | 0.67 | 0.16 | 0.27 | 1.05 | 1.6 | 1.09 | 0.26 | 1.63 | 0.66 | 0.48 | 0.69 | 0.19 | 0.16 | 1.17 | 0.69 | 0.06 | 0.96 | 0.82
| Analytical results for W in OREAS 922b (Certified Value 2.8 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.7 | 2.1056 | 3.2 | < 5 | 3.1 | 2.7615 | 3.1 | 2.8 | 2.5 | 2.5 | 2.8 | 3 | 2.8 | 2.1003 | 2.9 | 3.1 | 2.7 | 2.7 | 4.3 | 3 | 2.48 | 3.1 | 3.2 | 2.5 |  | 1
| 1 | 2 | 2.7 | 2.1194 | 3.2 | < 5 | 2.8 | 2.7615 | 3.1 | 2.8 | 2.6 | 2.6 | 2.6 | 2.8 | 2.9 | 2.1978 | 2.7 | 3.1 | 2.6 | 2.6 | 4.9 | 3 | 2.46 | 2.9 | 3.2 | 2.6 |  | 17
| 1 | 3 | 2.6 | 2.1514 | 3.2 | < 5 | 2.9 | 2.6737 | 3.1 | 3 | 2.6 | 2.6 | 2.9 | 2.9 | 2.9 | 2.2573 | 3.01 | 3.2 | 2.7 | 2.7 | 4.7 | 3 | 2.4 | 3 | 3.2 | 2.6 |  | 16
| 1 | 4 | 2.7 | 2.1131 | 3.1 | < 5 | 3.1 | 2.6884 | 3 | 2.8 | 2.7 | 2.5 | 2.7 | 3 | 2.9 | 2.4721 | 2.88 | 3.2 | 2.7 | 2.9 | 4.5 | 4 | 2.44 | 2.9 | 3.1 | 2.4 |  | 2.8047
| 1 | 5 | 2.7 | 2.0915 | 3 | < 5 | 2.9 | 2.7516 | 3.4 | 3 | 2.5 | 2.4 | 2.9 | 2.7 | 2.9 | 2.3394 | 1.83 | 3.1 | 3.1 | 2.7 | 4.8 | 3 | 2.51 | 2.9 | 3 | 2.5 |  | 62
| 1 | 6 | 2.8 | 2.0967 | 3.1 | < 5 | 2.9 | 2.7101 | 3.3 | 2.9 | 2.6 | 2.8 | 2.7 | 3.2 | 2.9 | 2.5586 | 2.97 | 3.2 | 2.8 | 2.5 | 4.5 | 3 | 2.45 | 2.9 | 3.1 | 2.5
| Mean |  | 2.7 | 2.1129 | 3.1333 |  | 2.95 | 2.7245 | 3.1667 | 2.8833 | 2.5833 | 2.5667 | 2.7667 | 2.9333 | 2.8833 | 2.3209 | 2.715 | 3.15 | 2.7667 | 2.6833 | 4.6167 | 3.1667 | 2.4567 | 2.95 | 3.1333 | 2.5167
| Median |  | 2.7 | 2.1093 | 3.15 |  | 2.9 | 2.7308 | 3.1 | 2.85 | 2.6 | 2.55 | 2.75 | 2.95 | 2.9 | 2.2983 | 2.89 | 3.15 | 2.7 | 2.7 | 4.6 | 3 | 2.455 | 2.9 | 3.15 | 2.5
| Std Dev. |  | 0.063246 | 0.021449 | 0.08165 |  | 0.12247 | 0.038882 | 0.15055 | 0.098319 | 0.075277 | 0.13663 | 0.12111 | 0.17512 | 0.040825 | 0.1718 | 0.44653 | 0.054772 | 0.17512 | 0.13292 | 0.22286 | 0.40825 | 0.037238 | 0.083666 | 0.08165 | 0.075277
| Rel.Std.Dev. |  | 2.342% | 1.015% | 2.606% |  | 4.152% | 1.427% | 4.754% | 3.41% | 2.914% | 5.323% | 4.377% | 5.97% | 1.416% | 7.402% | 16.45% | 1.739% | 6.33% | 4.953% | 4.827% | 12.89% | 1.516% | 2.836% | 2.606% | 2.991%
| PDM3 |  | -3.733% | -24.66% | 11.72% |  | 5.181% | -2.86% | 12.91% | 2.804% | -7.893% | -8.487% | -1.356% | 4.586% | 2.804% | -17.25% | -3.198% | 12.31% | -1.356% | -4.327% | 64.6% | 12.91% | -12.41% | 5.181% | 11.72% | -10.27%
| Z-Score (Absolute) |  | 0.25 | 2.4 | 1.35 | 0.98 | 0.67 | 0.16 | 1.47 | 0.43 | 0.67 | 0.74 | 0 | 0.61 | 0.43 | 1.64 | 0.19 | 1.41 | 0 | 0.31 | 6.8 | NA | 1.14 | 0.67 | 1.35 | 0.92
Results from laboratory 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 922b (Certified Value 16.4 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 16.5 | 17.2522 | 15 | 19.4 | 16.4 | 15.9129 | 16.5 | 15.8 | 16.3 | 17.2 | 16.5 | 20.2 | 14.5 | 16.1183 | 16.6 | 16.27 | 14.3 | 17 | 14.1 | 13.6 | 17.2 | 16.82 | 16 | 16.6 |  | 1
| 1 | 2 | 16.3 | 17.5896 | 14.8 | 18.1 | 15.57 | 15.9995 | 16.6 | 15.8 | 16.7 | 19.6 | 18.2 | 21.4 | 14.4 | 16.329 | 15.1 | 15.84 | 13 | 16.7 | 14.2 | 14.5 | 17 | 16.37 | 15.3 | 17.9 |  | 18
| 1 | 3 | 16.6 | 17.4212 | 14.8 | 20.8 | 14.93 | 16.1263 | 16.4 | 16.6 | 16 | 16.4 | 18.8 | 20.4 | 14.6 | 16.0214 | 17.5 | 16.75 | 14.4 | 16.2 | 13.7 | 15.2 | 16.6 | 20.28 | 15.8 | 16.1 |  | 16
| 1 | 4 | 16.5 | 17.5448 | 14.7 | 18.8 | 15.65 | 15.813 | 16.1 | 15.9 | 17.3 | 16.6 | 16.7 | 20.4 | 14.5 | 16.0936 | 12.5 | 16.84 | 13.6 | 16.6 | 14.3 | 14.5 | 16 | 16.71 | 15.3 | 15.6 |  | 16.4057
| 1 | 5 | 16.6 | 17.4338 | 15 | 20 | 15.69 | 16.0614 | 18.1 | 15.8 | 16.3 | 17 | 18.1 | 21.4 | 14.2 | 16.3507 | 11.1 | 16.92 | 14.3 | 16.7 | 14.1 | 14.3 | 16.8 | 17.23 | 15.6 | 16.5 |  | 63
| 1 | 6 | 16 | 17.3234 | 14.9 | 22.6 | 14.72 | 16.3976 | 17 | 16.2 | 16.5 | 16.4 | 17 | 20.6 | 14.4 | 16.2382 | 17.8 | 16.74 | 14.1 | 16.2 | 14.7 | 13.7 | 16.8 | 16.42 | 16.3 | 15.4
| Mean |  | 16.4167 | 17.4275 | 14.8667 | 19.95 | 15.4933 | 16.0518 | 16.7833 | 16.0167 | 16.5167 | 17.2 | 17.55 | 20.7333 | 14.4333 | 16.1919 | 15.1 | 16.56 | 13.95 | 16.5667 | 14.1833 | 14.3 | 16.7333 | 17.305 | 15.7167 | 16.35
| Median |  | 16.5 | 17.4275 | 14.85 | 19.7 | 15.61 | 16.0305 | 16.55 | 15.85 | 16.4 | 16.8 | 17.55 | 20.5 | 14.45 | 16.1782 | 15.85 | 16.745 | 14.2 | 16.65 | 14.15 | 14.4 | 16.8 | 16.765 | 15.7 | 16.3
| Std Dev. |  | 0.23166 | 0.12781 | 0.12111 | 1.5997 | 0.60109 | 0.20212 | 0.70828 | 0.32506 | 0.44907 | 1.2198 | 0.93968 | 0.53166 | 0.13663 | 0.13436 | 2.759 | 0.41929 | 0.54681 | 0.31411 | 0.32506 | 0.58992 | 0.41312 | 1.4902 | 0.39707 | 0.8961
| Rel.Std.Dev. |  | 1.411% | 0.7334% | 0.8146% | 8.018% | 3.88% | 1.259% | 4.22% | 2.03% | 2.719% | 7.092% | 5.354% | 2.564% | 0.9466% | 0.8298% | 18.27% | 2.532% | 3.92% | 1.896% | 2.292% | 4.125% | 2.469% | 8.611% | 2.526% | 5.481%
| PDM3 |  | 0.06711% | 6.229% | -9.381% | 21.6% | -5.561% | -2.157% | 2.302% | -2.371% | 0.6767% | 4.842% | 6.975% | 26.38% | -12.02% | -1.303% | -7.959% | 0.9408% | -14.97% | 0.9814% | -13.55% | -12.83% | 1.997% | 5.482% | -4.2% | -0.3393%
| Z-Score (Absolute) |  | 0.03 | 0.83 | 1.2 | 2.82 | 0.7 | 0.26 | 0.32 | 0.29 | 0.11 | 0.65 | 0.92 | 3.44 | 1.54 | 0.15 | 1.01 | 0.14 | 1.92 | 0.14 | 1.74 | 1.65 | 0.28 | 0.73 | 0.53 | 0.03
| Analytical results for Yb in OREAS 922b (Certified Value 1.67 ppm) | [26 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.16; 2.08: 2.17; 2.09: 2.20; 2.10: 2.21; 2.11: 2.22 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*OES/MS; 2.11: 4A*OES/MS |  | 2
3
| 1 | 1 | 2.01: 1.6254; 2.02: 1.7; 2.03: 2.45; 2.04: 1.8779; 2.05: 2.2; 2.06: 1.6; 2.07: 1.83; 2.08: 1.6; 2.09: 1.6; 2.10: 1.6; 2.11: 1.82 |  | 1
| 1 | 2 | 2.01: 1.6518; 2.02: 1.7; 2.03: 2.29; 2.04: 1.8167; 2.05: 2.4; 2.06: 1.6; 2.07: 1.69; 2.08: 1.5; 2.09: 1.61; 2.10: 1.7; 2.11: 1.75 |  | 19
| 1 | 3 | 2.01: 1.6045; 2.02: 1.7; 2.03: 2.48; 2.04: 1.7906; 2.05: 2.2; 2.06: 1.5; 2.07: 1.8; 2.08: 1.6; 2.09: 1.56; 2.10: 1.8; 2.11: 1.72 |  | 16
| 1 | 4 | 2.01: 1.6572; 2.02: 1.7; 2.03: 2.31; 2.04: 1.652; 2.05: 2.3; 2.06: 1.6; 2.07: 1.74; 2.08: 1.4; 2.09: 1.66; 2.10: 1.8; 2.11: 1.7 |  | 1.6735
| 1 | 5 | 2.01: 1.6427; 2.02: 1.7; 2.03: 2.44; 2.04: 1.7742; 2.05: 2.3; 2.06: 1.5; 2.07: 1.73; 2.08: 1.6; 2.09: 1.56; 2.10: 1.7; 2.11: 1.84 |  | 64
| 1 | 6 | 2.01: 1.6551; 2.02: 1.7; 2.03: 2.49; 2.04: 1.7502; 2.05: 2.3; 2.06: 1.6; 2.07: 1.81; 2.08: 1.6; 2.09: 1.5; 2.10: 1.7; 2.11: 1.75
| Mean |  | 2.01: 1.6395; 2.02: 1.7; 2.03: 2.41; 2.04: 1.7769; 2.05: 2.2833; 2.06: 1.5667; 2.07: 1.7667; 2.08: 1.55; 2.09: 1.5817; 2.10: 1.7167; 2.11: 1.7633
| Median |  | 2.01: 1.6473; 2.02: 1.7; 2.03: 2.445; 2.04: 1.7824; 2.05: 2.3; 2.06: 1.6; 2.07: 1.77; 2.08: 1.6; 2.09: 1.58; 2.10: 1.7; 2.11: 1.75
| Std Dev. |  | 2.01: 0.020698; 2.02: 0; 2.03: 0.087407; 2.04: 0.075246; 2.05: 0.075277; 2.06: 0.05164; 2.07: 0.05465; 2.08: 0.083666; 2.09: 0.054559; 2.10: 0.075277; 2.11: 0.055377
| Rel.Std.Dev. |  | 2.01: 1.262%; 2.02: 0%; 2.03: 3.627%; 2.04: 4.235%; 2.05: 3.297%; 2.06: 3.296%; 2.07: 3.093%; 2.08: 5.398%; 2.09: 3.449%; 2.10: 4.385%; 2.11: 3.141%
| PDM3 |  | 2.01: -2.034%; 2.02: 1.584%; 2.03: 44.01%; 2.04: 6.181%; 2.05: 36.44%; 2.06: -6.383%; 2.07: 5.568%; 2.08: -7.379%; 2.09: -5.487%; 2.10: 2.58%; 2.11: 5.369%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.15; 2.03: 6.05; 2.04: 0.53; 2.05: 4.95; 2.06: 1.31; 2.07: 0.44; 2.08: 1.46; 2.09: 1.18; 2.10: 0; 2.11: 0.41
| Analytical results for Zn in OREAS 922b (Certified Value 75 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 74 | 69.41 | 80 | 87.37 | 74 | 79.8774 | 82 | 72 | 74 | 77 | 69 | 80 | 73 | 68.0505 | 74 | 74 | 78.4 | 77 | 77 | 72 | 79 | 75.3 | 79 | 74 | 74 | 1
| 1 | 2 | 75 | 69.91 | 80 | 81.25 | 71 | 77.5169 | 75 | 72 | 74 | 75 | 71 | 81 | 74 | 69.7464 | 71 | 76 | 73.3 | 77 | 94 | 71 | 79 | 71.2 | 74 | 71 | 74 | 40
| 1 | 3 | 73 | 69.75 | 80 | 83.9 | 73 | 77.2299 | 105 | 73 | 72 | 75 | 69 | 82 | 73 | 72.4249 | 76 | 75 | 77.8 | 92 | 113 | 69 | 78 | 71.8 | 74 | 75 | 74 | 16
| 1 | 4 | 75 | 69.85 | 79 | 83.8 | 75 | 83.1297 | 76 | 72 | 75 | 76 | 71 | 81 | 73 | 72.0735 | 80 | 76 | 71 | 100 | 116 | 70 | 79 | 69.2 | 77 | 74 | 71 | 74.7683
| 1 | 5 | 73 | 70.78 | 80 | 82.89 | 73 | 80.68 | 93 | 73 | 74 | 74 | 70 | 80 | 75 | 70.7197 | 67 | 75 | 76.3 | 107 | 77 | 72 | 79 | 73.8 | 90 | 73 | 74 | 65
| 1 | 6 | 74 | 71.61 | 79 | 82.23 | 71 | 77.5602 | 96 | 73 | 74 | 75 | 70 | 80 | 74 | 67.712 | 81 | 78 | 73.2 | 76 | 79 | 70 | 78 | 72.6 | 74 | 72 | 77
| Mean |  | 74 | 70.2183 | 79.6667 | 83.5733 | 72.8333 | 79.3324 | 87.8333 | 72.5 | 73.8333 | 75.3333 | 70 | 80.6667 | 73.6667 | 70.1212 | 74.8333 | 75.6667 | 75 | 88.1667 | 92.6667 | 70.6667 | 78.6667 | 72.3167 | 78 | 73.1667 | 74
| Median |  | 74 | 69.88 | 80 | 83.345 | 73 | 78.7188 | 87.5 | 72.5 | 74 | 75 | 70 | 80.5 | 73.5 | 70.233 | 75 | 75.5 | 74.8 | 84.5 | 86.5 | 70.5 | 79 | 72.2 | 75.5 | 73.5 | 74
| Std Dev. |  | 0.89443 | 0.81923 | 0.5164 | 2.1091 | 1.6021 | 2.3405 | 12.0568 | 0.54772 | 0.98319 | 1.0328 | 0.89443 | 0.8165 | 0.8165 | 1.9854 | 5.3448 | 1.3663 | 2.9401 | 13.4672 | 18.096 | 1.2111 | 0.5164 | 2.117 | 6.229 | 1.472 | 1.8974
| Rel.Std.Dev. |  | 1.209% | 1.167% | 0.6482% | 2.524% | 2.2% | 2.95% | 13.73% | 0.7555% | 1.332% | 1.371% | 1.278% | 1.012% | 1.108% | 2.831% | 7.142% | 1.806% | 3.92% | 15.27% | 19.53% | 1.714% | 0.6564% | 2.927% | 7.986% | 2.012% | 2.564%
| PDM3 |  | -1.028% | -6.085% | 6.551% | 11.78% | -2.588% | 6.104% | 17.47% | -3.034% | -1.25% | 0.7557% | -6.377% | 7.889% | -1.473% | -6.215% | 0.08698% | 1.202% | 0.3099% | 17.92% | 23.94% | -5.486% | 5.214% | -3.279% | 4.322% | -2.142% | -1.028%
| Z-Score (Absolute) |  | 0.18 | 0.98 | 1.03 | 1.86 | 0.43 | 0.96 | 2.77 | 0.5 | 0.21 | 0.11 | 1.03 | 1.24 | 0.25 | 1 | 0 | 0.18 | 0.04 | 2.84 | 3.8 | 0.89 | 0.82 | 0.54 | 0.67 | 0.35 | 0.18
| Analytical results for Zr in OREAS 922b (Certified Value 85 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 83.7 | 79.85 | 89.2 | 92.5 | 86 | 84.9567 | 89.7 | 85.3 | 82.9 | 82.4 | 88.5 | 93 | 79 | 80.8271 | 91.1 | 85.2 | 84.3 | 82.8 | 81 | 80.2 | 86.7 | 86.8 | 69.7 | 83.2 |  | 1
| 1 | 2 | 84.5 | 78.16 | 89.6 | 85.69 | 81.6 | 83.5918 | 89.1 | 84.8 | 85.5 | 86.4 | 85.9 | 93 | 79.4 | 78.3638 | 84.2 | 85.7 | 83.3 | 84.8 | 85 | 81.2 | 82.4 | 87.6 | 71.8 | 81.8 |  | 41
| 1 | 3 | 85.8 | 78.15 | 89.6 | 93.2 | 82.9 | 84.9526 | 86.5 | 90.8 | 82.9 | 88.4 | 90.5 | 90 | 76.6 | 86.1624 | 85.3 | 85.7 | 88.2 | 84.1 | 81 | 81.1 | 83.7 | 86.9 | 73.9 | 82.5 |  | 16
| 1 | 4 | 83.1 | 77.85 | 89.5 | 91.42 | 85.5 | 85.5456 | 87.7 | 86 | 86.1 | 85.8 | 87 | 89 | 78.1 | 85.6453 | 68.5 | 86.3 | 84.2 | 86.6 | 86 | 81.7 | 82.9 | 84.3 | 70.8 | 81.2 |  | 85.0321
| 1 | 5 | 84 | 79.33 | 89.6 | 94.04 | 82.5 | 86.1293 | 91.9 | 89.2 | 82.6 | 82.7 | 84 | 95 | 82.3 | 81.4691 | 56.5 | 85.8 | 91.3 | 85.9 | 97 | 81 | 81.9 | 83.4 | 72.6 | 82.2 |  | 66
| 1 | 6 | 85 | 76.69 | 89.5 | 97.14 | 81.4 | 86.0245 | 87.2 | 89 | 82.7 | 82.4 | 83.5 | 93 | 80.1 | 85.422 | 81.4 | 87.3 | 88.8 | 81 | 84 | 80.2 | 80.1 | 85.7 | 74.7 | 81.8
| Mean |  | 84.35 | 78.3383 | 89.5 | 92.3317 | 83.3167 | 85.2001 | 88.6833 | 87.5167 | 83.7833 | 84.6833 | 86.5667 | 92.1667 | 79.25 | 82.9816 | 77.8333 | 86 | 86.6833 | 84.2 | 85.6667 | 80.9 | 82.95 | 85.7833 | 72.25 | 82.1167
| Median |  | 84.25 | 78.155 | 89.55 | 92.85 | 82.7 | 85.2511 | 88.4 | 87.5 | 82.9 | 84.25 | 86.45 | 93 | 79.2 | 83.4456 | 82.8 | 85.75 | 86.25 | 84.45 | 84.5 | 81.05 | 82.65 | 86.25 | 72.2 | 82
| Std Dev. |  | 0.96488 | 1.1217 | 0.15492 | 3.7885 | 1.9712 | 0.93483 | 1.9743 | 2.4661 | 1.5779 | 2.5443 | 2.6771 | 2.2286 | 1.9212 | 3.207 | 12.8626 | 0.72664 | 3.2059 | 2.0601 | 5.9217 | 0.5933 | 2.198 | 1.6412 | 1.8791 | 0.68823
| Rel.Std.Dev. |  | 1.144% | 1.432% | 0.1731% | 4.103% | 2.366% | 1.097% | 2.226% | 2.818% | 1.883% | 3.005% | 3.092% | 2.418% | 2.424% | 3.865% | 16.53% | 0.8449% | 3.698% | 2.447% | 6.913% | 0.7334% | 2.65% | 1.913% | 2.601% | 0.8381%
| PDM3 |  | -0.8022% | -7.872% | 5.254% | 8.584% | -2.017% | 0.1975% | 4.294% | 2.922% | -1.469% | -0.4102% | 1.805% | 8.39% | -6.8% | -2.411% | -8.466% | 1.138% | 1.942% | -0.9786% | 0.7463% | -4.859% | -2.449% | 0.8835% | -15.03% | -3.429%
| Z-Score (Absolute) |  | 0.05 | 1.98 | 1.59 | 2.5 | 0.38 | 0.22 | 1.33 | 0.96 | 0.23 | 0.05 | 0.66 | 2.45 | 1.68 | 0.49 | 2.14 | 0.47 | 0.69 | 0.1 | 0.37 | 1.16 | 0.5 | 0.41 | 3.92 | 0.77
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C714: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C825: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C898: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1083: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1119: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1137: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 922b (Certified Value 0.973 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 0.94 | 0.98 | 0.95 | 0.94 | 0.95 | 0.96348 | 0.96 | 0.99 | 1.07 | 0.95 | 0.99 | 0.966 | 1.02 | 1.0199 | 1.09 | 0.923 | 0.9 | 1 | 0.98 | 0.94 | 0.9 | 1.1 | 1 | 0.9 | 1
| 1 | 2 | 0.93 | 0.92 | 0.95 | 0.94 | 0.91 | 0.93034 | 0.99 | 0.94 | 1.03 | 0.95 | 0.92 | 0.954 | 1.05 | 1.0616 | 1.3 | 0.918 | 0.9 | 1 | 1.06 | 0.9 | 1 | 1.1 | 1 | 0.9 | 8
| 1 | 3 | 0.93 | 1.04 | 0.95 | 0.96 | 0.92 | 0.95899 | 0.98 | 1 | 1.04 | 0.94 | 0.95 | 0.96 | 1.02 | 1.0936 | 1.05 | 0.874 | 0.9 | 1 | 1.03 | 0.9 | 1 | 1.1 | 1 | 0.9 | 16
| 1 | 4 | 0.93 | 0.98 | 0.95 | 1 | 0.89 | 0.95764 | 1 | 0.95 | 1.04 | 0.97 | 0.93 | 0.908 | 1.04 | 1.0733 | 1.14 | 0.896 | 0.9 | 0.9 | 0.97 | 0.87 | 1 | 1.1 | 1.01 | 0.9 | 0.97259
| 1 | 5 | 0.94 | 1.01 | 0.95 | 0.97 | 0.88 | 0.92998 | 0.99 | 0.93 | 1.03 | 0.96 | 0.95 | 1 | 1.01 | 1.0268 | 1.01 | 0.906 | 0.9 | 1 | 1.06 | 0.91 | 1.1 | 1.1 | 1.02 | 0.9 | 68
| 1 | 6 | 0.95 | 0.99 | 0.95 | 0.96 | 0.9 | 0.93458 | 1.02 | 0.97 | 1.01 | 0.97 | 0.95 | 1 | 1.07 | 1.0636 | 1.12 | 0.928 | 0.9 | 1 | 0.96 | 0.99 | 1 | 1.1 | 1.03 | 0.9
| Mean |  | 0.93667 | 0.98667 | 0.95 | 0.96167 | 0.90833 | 0.94583 | 0.99 | 0.96333 | 1.0367 | 0.95667 | 0.94833 | 0.96467 | 1.035 | 1.0565 | 1.1183 | 0.9075 | 0.9 | 0.98333 | 1.01 | 0.91833 | 1 | 1.1 | 1.01 | 0.9
| Median |  | 0.935 | 0.985 | 0.95 | 0.96 | 0.905 | 0.94611 | 0.99 | 0.96 | 1.035 | 0.955 | 0.95 | 0.963 | 1.03 | 1.0626 | 1.105 | 0.912 | 0.9 | 1 | 1.005 | 0.905 | 1 | 1.1 | 1.005 | 0.9
| Std Dev. |  | 0.008165 | 0.039833 | 1.2162e-16 | 0.022286 | 0.024833 | 0.015759 | 0.02 | 0.028048 | 0.019664 | 0.012111 | 0.024014 | 0.034192 | 0.022583 | 0.028126 | 0.10068 | 0.020137 | 0 | 0.040825 | 0.045607 | 0.041673 | 0.063246 | 2.4324e-16 | 0.012649 | 0
| Rel.Std.Dev. |  | 0.8717% | 4.037% | 1.28e-14% | 2.317% | 2.734% | 1.666% | 2.02% | 2.912% | 1.897% | 1.266% | 2.532% | 3.544% | 2.182% | 2.662% | 9.003% | 2.219% | 0% | 4.152% | 4.516% | 4.538% | 6.325% | 2.211e-14% | 1.252% | 0%
| PDM3 |  | -3.693% | 1.448% | -2.322% | -1.123% | -6.606% | -2.75% | 1.791% | -0.9513% | 6.589% | -1.637% | -2.494% | -0.8142% | 6.417% | 8.624% | 14.99% | -6.692% | -7.463% | 1.105% | 3.847% | -5.578% | 2.819% | 13.1% | 3.847% | -7.463%
| Z-Score (Absolute) |  | 0.45 | 0.37 | 0.23 | 0.04 | 0.92 | 0.3 | 0.43 | 0.01 | 1.2 | 0.12 | 0.26 | 0.01 | 1.17 | 1.53 | 2.55 | 0.93 | 1.06 | 0.32 | 0.76 | 0.75 | 0.59 | 2.25 | 0.76 | 1.06
| Analytical results for Al in OREAS 922b (Certified Value 0.793 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.06 | 0.86717 | 0.69 | 0.73 | 0.90853 | 0.82 | 0.81 | 0.87 | 0.74 | 0.81 | 1.09 | 0.807 | 0.94564 | 0.77 | 0.77 | 0.78 | 0.79 | 0.79 | 0.89 | 1.255 | 0.73 | 0.78 |  |  | 1
| 1 | 2 | 1.03 | 0.85758 | 0.68 | 0.748 | 0.88488 | 0.76 | 0.8 | 0.88 | 0.74 | 0.79 | 0.94 | 0.798 | 0.99279 | 0.79 | 0.78 | 0.75 | 0.78 | 0.78 | 0.92 | 1.1738 | 0.8 | 0.78 |  |  | 0
| 1 | 3 | 1.02 | 0.85271 | 0.67 | 0.745 | 0.89066 | 0.74 | 0.79 | 0.86 | 0.74 | 0.78 | 0.96 | 0.808 | 0.96061 | 0.76 | 0.74 | 0.77 | 0.79 | 0.79 | 0.93 | 1.1973 | 0.74 | 0.78 |  |  | 16
| 1 | 4 | 1.04 | 0.86721 | 0.69 | 0.748 | 0.90112 | 0.75 | 0.78 | 0.88 | 0.73 | 0.78 | 0.91 | 0.808 | 0.99594 | 0.7 | 0.77 | 0.77 | 0.8 | 0.78 | 0.93 | 1.2604 | 0.76 | 0.77 |  |  | 0.79305
| 1 | 5 | 1.01 | 0.85863 | 0.67 | 0.745 | 0.89658 | 0.75 | 0.81 | 0.87 | 0.75 | 0.79 | 1.09 | 0.812 | 0.97011 | 0.76 | 0.7 | 0.77 | 0.8 | 0.79 | 0.93 | 1.222 | 0.77 | 0.77 |  |  | 69
| 1 | 6 | 1.03 | 0.86884 | 0.7 | 0.753 | 0.86704 | 0.77 | 0.8 | 0.88 | 0.74 | 0.79 | 0.99 | 0.836 | 0.95472 | 0.81 | 0.77 | 0.75 | 0.77 | 0.79 | 0.91 | 1.206 | 0.78 | 0.76
| Mean |  | 1.0317 | 0.86202 | 0.68333 | 0.74483 | 0.89147 | 0.765 | 0.79833 | 0.87333 | 0.74 | 0.79 | 0.99667 | 0.8115 | 0.96997 | 0.765 | 0.755 | 0.765 | 0.78833 | 0.78667 | 0.91833 | 1.2191 | 0.76333 | 0.77333
| Median |  | 1.03 | 0.8629 | 0.685 | 0.7465 | 0.89362 | 0.755 | 0.8 | 0.875 | 0.74 | 0.79 | 0.975 | 0.808 | 0.96536 | 0.765 | 0.77 | 0.77 | 0.79 | 0.79 | 0.925 | 1.214 | 0.765 | 0.775
| Std Dev. |  | 0.017224 | 0.006601 | 0.012111 | 0.0078337 | 0.014497 | 0.02881 | 0.01169 | 0.008165 | 0.0063246 | 0.010954 | 0.076855 | 0.012865 | 0.020528 | 0.037283 | 0.030166 | 0.012247 | 0.01169 | 0.005164 | 0.016021 | 0.033769 | 0.02582 | 0.008165
| Rel.Std.Dev. |  | 1.67% | 0.7658% | 1.772% | 1.052% | 1.626% | 3.766% | 1.464% | 0.9349% | 0.8547% | 1.387% | 7.711% | 1.585% | 2.116% | 4.874% | 3.996% | 1.601% | 1.483% | 0.6564% | 1.745% | 2.77% | 3.383% | 1.056%
| PDM3 |  | 30.09% | 8.698% | -13.83% | -6.079% | 12.41% | -3.536% | 0.6667% | 10.12% | -6.689% | -0.3841% | 25.68% | 2.327% | 22.31% | -3.536% | -4.797% | -3.536% | -0.5942% | -0.8044% | 15.8% | 53.72% | -3.747% | -2.486%
| Z-Score (Absolute) |  | 4.17 | 1.25 | 1.82 | 0.76 | 1.76 | 0.42 | 0.16 | 1.45 | 0.84 | 0.01 | 3.56 | 0.38 | 3.11 | 0.42 | 0.59 | 0.42 | 0.01 | 0.04 | 2.22 | 7.39 | 0.44 | 0.27
| Analytical results for As in OREAS 922b (Certified Value 52 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 1
| 1 | 1 | 50 | 50.15 | 39.4 | 49.81 | 54.0366 | 49.3 | 53.3 | 59.1 | 49.8 | 53.9 | 52.9 | 48 | 50.4194 | 59.6 | 50.6 | 51.6 | 52.2 | 50 | 55 | 48.2 | 53 | 49.7 | 53.3 |  | 1
| 1 | 2 | 49 | 50.97 | 37.1 | 49.87 | 52.6228 | 48.9 | 52.5 | 55.7 | 52.5 | 52.8 | 53.4 | 48 | 50.827 | 0.7 | 51.5 | 51.5 | 52 | 54 | 55 | 48.7 | 54 | 51.9 | 52.3 |  | 9
| 1 | 3 | 50 | 51.17 | 39.4 | 50.66 | 52.6347 | 52.3 | 55 | 58 | 50.2 | 52.7 | 53.4 | 48 | 49.943 | 60.9 | 51 | 52.2 | 53.4 | 53 | 55 | 50.2 | 53 | 49.3 | 51.8 |  | 16
| 1 | 4 | 50 | 50.11 | 42.6 | 49.58 | 53.5747 | 51.7 | 53.4 | 58.6 | 50.8 | 52.9 | 50.1 | 48 | 51.5312 | 60.5 | 50.5 | 51.9 | 52.7 | 53 | 56 | 50.4 | 54 | 50.2 | 50.8 |  | 51.6941
| 1 | 5 | 49 | 52.23 | 39.3 | 49.94 | 53.423 | 50.7 | 55 | 57.9 | 49.5 | 54.2 | 55.9 | 48 | 51.6312 | 54.3 | 50.5 | 54.2 | 52.8 | 54 | 56 | 50.9 | 54 | 50.6 | 50.4 |  | 70
| 1 | 6 | 50 | 50.65 | 41.9 | 52.03 | 51.8712 | 52.5 | 55.2 | 57.5 | 51.9 | 54.7 | 55.1 | 51 | 50.2236 | 59.6 | 52.4 | 52.5 | 52.4 | 51 | 56 | 49.2 | 57 | 50.7 | 51.1
| Mean |  | 49.6667 | 50.88 | 39.95 | 50.315 | 53.0272 | 50.9 | 54.0667 | 57.8 | 50.7833 | 53.5333 | 53.4667 | 48.5 | 50.7626 | 49.2667 | 51.0833 | 52.3167 | 52.5833 | 52.5 | 55.5 | 49.6 | 54.1667 | 50.4 | 51.6167
| Median |  | 50 | 50.81 | 39.4 | 49.905 | 53.0289 | 51.2 | 54.2 | 57.95 | 50.5 | 53.4 | 53.4 | 48 | 50.6232 | 59.6 | 50.8 | 52.05 | 52.55 | 53 | 55.5 | 49.7 | 54 | 50.4 | 51.45
| Std Dev. |  | 0.5164 | 0.78656 | 1.9987 | 0.91588 | 0.7911 | 1.5336 | 1.1413 | 1.173 | 1.1957 | 0.84538 | 2.0126 | 1.2247 | 0.69702 | 23.9129 | 0.75211 | 0.99482 | 0.49967 | 1.6432 | 0.54772 | 1.0602 | 1.472 | 0.90774 | 1.0722
| Rel.Std.Dev. |  | 1.04% | 1.546% | 5.003% | 1.82% | 1.492% | 3.013% | 2.111% | 2.029% | 2.354% | 1.579% | 3.764% | 2.525% | 1.373% | 48.54% | 1.472% | 1.902% | 0.9502% | 3.13% | 0.9869% | 2.137% | 2.717% | 1.801% | 2.077%
| PDM3 |  | -3.922% | -1.575% | -22.72% | -2.668% | 2.579% | -1.536% | 4.59% | 11.81% | -1.762% | 3.558% | 3.429% | -6.179% | -1.802% | -4.696% | -1.181% | 1.204% | 1.72% | 1.559% | 7.362% | -4.051% | 4.783% | -2.503% | -0.1498%
| Z-Score (Absolute) |  | 0.64 | 0.09 | 5.06 | 0.35 | 0.88 | 0.08 | 1.36 | 3.05 | 0.14 | 1.11 | 1.08 | 1.17 | 0.15 | 0.83 | 0 | 0.56 | 0.68 | 0.64 | 2.01 | 0.67 | 1.4 | 0.31 | 0.24
| Analytical results for Au in OREAS 922b (Certified Value < 0.02 ppm) | [25 empty] | Certified
Au | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  | 1
| No | No | 2.01: 2.05; 2.02: 2.07; 2.04: 2.08; 2.05: 2.09; 2.06: 2.10; 2.07: 2.11; 2.08: 2.12; 2.09: 2.14; 2.10: 2.17; 2.11: 2.19; 2.12: 2.24 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*OES/MS; 2.09: AR*OES/MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*OES/MS |  | 3
|  |  | 2.01: 0.5g; 2.02: 0.5g; 2.04: 0.5g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 01g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g |  | 3
| 1 | 1 | 2.01: 0.0043; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0068; 2.11: 0.009; 2.12: 0.0046 |  | 1
| 1 | 2 | 2.01: 0.0042; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0088; 2.11: 0.0089; 2.12: 0.0044 |  | 10
| 1 | 3 | 2.01: 0.0042; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0083; 2.11: 0.0097; 2.12: 0.0044 |  | 16
| 1 | 4 | 2.01: 0.0038; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0069; 2.11: 0.0109; 2.12: 0.0046 |  | < 0.02
| 1 | 5 | 2.01: 0.004; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0135; 2.11: 0.0085; 2.12: 0.0043 |  | 71
| 1 | 6 | 2.01: 0.0042; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0088; 2.11: 0.0078; 2.12: 0.005
| Mean |  | 2.01: 0.0041167; 2.10: 0.00885; 2.11: 0.0091333; 2.12: 0.00455
| Median |  | 2.01: 0.0042; 2.10: 0.00855; 2.11: 0.00895; 2.12: 0.0045
| Std Dev. |  | 2.01: 0.00018348; 2.10: 0.0024468; 2.11: 0.0010671; 2.12: 0.000251
| Rel.Std.Dev. |  | 2.01: 4.457%; 2.10: 27.65%; 2.11: 11.68%; 2.12: 5.516%
PDM3
| Z-Score (Absolute) |  | 2.01: 3.9; 2.02: 0.67; 2.04: 0.67; 2.05: 0.67; 2.06: 0.67; 2.07: 0.67; 2.08: 6.91; 2.09: 6.72; 2.10: 0.22; 2.11: 0; 2.12: 3.57
| Analytical results for B in OREAS 922b (Certified Value < 10 ppm) | [25 empty] | Certified
B | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  | 1
| No | No | 2.01: 2.06; 2.02: 2.07; 2.04: 2.08; 2.05: 2.09; 2.06: 2.10; 2.07: 2.11; 2.08: 2.12; 2.09: 2.14; 2.10: 2.17; 2.11: 2.19; 2.12: 2.20; 2.13: 2.24 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*OES/MS; 2.09: AR*OES/MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*OES/MS |  | 1
|  |  | 2.01: 0.2g; 2.02: 0.5g; 2.04: 0.5g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 01g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.5g |  | 1
| 1 | 1 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 7; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 1
| 1 | 2 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 6; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 1
| 1 | 3 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 7; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 16
| 1 | 4 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 6; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | < 10
| 1 | 5 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 7; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 72
| 1 | 6 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 7; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10
| Mean |  | 2.08: 6.6667
| Median |  | 2.08: 7
| Std Dev. |  | 2.08: 0.5164
| Rel.Std.Dev. |  | 2.08: 7.746%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.13
| Analytical results for Ba in OREAS 922b (Certified Value 188 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 0
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  |  |  | 0
| 1 | 1 | 255 | 183.9 | 166.06 | 200.9917 | 190 | 190 | 210 | 170 | 190 | 226 | 182 | 217.7743 | 120 | 200 | 181 | 176 | 177 | 305 | 161 | 187 |  |  |  |  | 1
| 1 | 2 | 249 | 182.3 | 170.16 | 207.294 | 170 | 180 | 210 | 170 | 180 | 218 | 177 | 235.3775 | 60 | 204.7 | 170 | 192 | 176 | 316 | 176 | 186 |  |  |  |  | 27
| 1 | 3 | 250 | 175 | 169.53 | 212.1005 | 170 | 180 | 210 | 170 | 180 | 230 | 180 | 244.0793 | 130 | 189.2 | 172 | 195 | 176 | 307 | 165 | 187 |  |  |  |  | 16
| 1 | 4 | 250 | 184.9 | 172.58 | 202.5099 | 170 | 180 | 210 | 170 | 180 | 220 | 180 | 231.8948 | 130 | 202.7 | 179 | 184 | 177 | 311 | 170 | 180 |  |  |  |  | 187.6625
| 1 | 5 | 248 | 181.7 | 168.16 | 195.9902 | 170 | 180 | 210 | 170 | 180 | 225 | 181 | 212.0024 | 120 | 181.4 | 179 | 191 | 177 | 312 | 171 | 179 |  |  |  |  | 73
| 1 | 6 | 248 | 187.9 | 172.81 | 199.0733 | 180 | 180 | 210 | 170 | 180 | 230 | 188 | 221.8791 | 130 | 198.9 | 174 | 179 | 176 | 309 | 174 | 183
| Mean |  | 250 | 182.6167 | 169.8833 | 202.9933 | 175 | 181.6667 | 210 | 170 | 181.6667 | 224.8333 | 181.3333 | 227.1679 | 115 | 196.15 | 175.8333 | 186.1667 | 176.5 | 310 | 169.5 | 183.6667
| Median |  | 249.5 | 183.1 | 169.845 | 201.7508 | 170 | 180 | 210 | 170 | 180 | 225.5 | 180.5 | 226.887 | 125 | 199.45 | 176.5 | 187.5 | 176.5 | 310 | 170.5 | 184.5
| Std Dev. |  | 2.6077 | 4.3296 | 2.5933 | 5.8324 | 8.3666 | 4.0825 | 0 | 0 | 4.0825 | 4.9967 | 3.6697 | 12.0161 | 27.3861 | 8.9917 | 4.446 | 7.679 | 0.54772 | 3.8987 | 5.6125 | 3.559
| Rel.Std.Dev. |  | 1.043% | 2.371% | 1.527% | 2.873% | 4.781% | 2.247% | 0% | 0% | 2.247% | 2.222% | 2.024% | 5.29% | 23.81% | 4.584% | 2.529% | 4.125% | 0.3103% | 1.258% | 3.311% | 1.938%
| PDM3 |  | 33.22% | -2.689% | -9.474% | 8.169% | -6.747% | -3.195% | 11.9% | -9.412% | -3.195% | 19.81% | -3.373% | 21.05% | -38.72% | 4.523% | -6.303% | -0.7971% | -5.948% | 65.19% | -9.678% | -2.129%
| Z-Score (Absolute) |  | 3.75 | 0.03 | 0.68 | 1.15 | 0.39 | 0.03 | 1.54 | 0.67 | 0.03 | 2.36 | 0.04 | 2.49 | 3.71 | 0.77 | 0.35 | 0.22 | 0.31 | 7.07 | 0.7 | 0.08
| Analytical results for Be in OREAS 922b (Certified Value 0.47 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.7 | 0.41664 | < 2 | 0.45 | 0.51246 | 0.47 | 0.47 | 0.51 | 0.43 | 0.47 | 0.5 | 0.43 | < 0.5 | 0.51 | 0.47 | 0.6 | 0.46 | < 1 | [6 empty] | 1
| 1 | 2 | 0.6 | 0.40236 | < 2 | 0.45 | 0.53088 | 0.45 | 0.47 | 0.52 | 0.43 | 0.45 | 0.5 | 0.44 | < 0.5 | 1.1 | 0.44 | 0.6 | 0.5 | < 1 | [6 empty] | 28
| 1 | 3 | 0.6 | 0.38736 | < 2 | 0.47 | 0.50463 | 0.44 | 0.46 | 0.51 | 0.42 | 0.45 | 0.6 | 0.42 | < 0.5 | 0.45 | 0.48 | 0.6 | 0.47 | < 1 | [6 empty] | 16
| 1 | 4 | 0.6 | 0.41664 | < 2 | 0.48 | 0.52979 | 0.44 | 0.45 | 0.52 | 0.42 | 0.45 | 0.5 | 0.36 | 0.50437 | 0.62 | 0.44 | 0.6 | 0.48 | < 1 | [6 empty] | 0.46974
| 1 | 5 | 0.6 | 0.43476 | < 2 | 0.51 | < 0.5 | 0.44 | 0.47 | 0.51 | 0.43 | 0.46 | 0.5 | 0.41 | < 0.5 | 0.51 | 0.47 | 0.6 | 0.48 | 1 | [6 empty] | 74
| 1 | 6 | 0.6 | 0.39444 | < 2 | 0.54 | < 0.5 | 0.45 | 0.46 | 0.51 | 0.43 | 0.45 | 0.5 | 0.46 | < 0.5 | 0.76 | 0.46 | 0.6 | 0.5 | < 1
| Mean |  | 0.61667 | 0.4087 |  | 0.48333 | 0.51944 | 0.44833 | 0.46333 | 0.51333 | 0.42667 | 0.455 | 0.51667 | 0.42 | 0.50437 | 0.65833 | 0.46 | 0.6 | 0.48167 | 1
| Median |  | 0.6 | 0.4095 |  | 0.475 | 0.52113 | 0.445 | 0.465 | 0.51 | 0.43 | 0.45 | 0.5 | 0.425 | 0.50437 | 0.565 | 0.465 | 0.6 | 0.48 | 1
| Std Dev. |  | 0.040825 | 0.017344 |  | 0.03559 | 0.01299 | 0.01169 | 0.008165 | 0.005164 | 0.005164 | 0.0083666 | 0.040825 | 0.034059 |  | 0.2426 | 0.016733 | 0 | 0.016021
| Rel.Std.Dev. |  | 6.62% | 4.244% |  | 7.364% | 2.501% | 2.608% | 1.762% | 1.006% | 1.21% | 1.839% | 7.902% | 8.109% |  | 36.85% | 3.638% | 0% | 3.326%
| PDM3 |  | 31.28% | -13% |  | 2.893% | 10.58% | -4.558% | -1.365% | 9.279% | -9.17% | -3.139% | 9.989% | -10.59% | 7.371% | 40.15% | -2.074% | 27.73% | 2.538% | 112.9%
| Z-Score (Absolute) |  | NA | 1.04 | 10.92 | 0.47 | 0.61 | 0.24 | 0.07 | 1.08 | 0.67 | 0.1 | NA | 0.81 | 3.39 | 4.01 | 0 | NA | 0.44 | 2.49
Results from laboratories 2.01, 2.12 and 2.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 922b (Certified Value 12 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 2
| 1 | 1 | 11.6 | 11.469 | 12.3 | 10 | 12.12 | 13.2624 | 12.9 | 12.8 | 12.8 | 11.25 | 13.75 | 12.6 | 11.67 | 12.6246 | 11.9 | 10.5 | 11.98 | 13.3 | 10.91 | 11.6 | 12.3 | 13 | 11.95 | 10.1 | 1
| 1 | 2 | 12 | 11.487 | 12.3 | 10.6 | 12.46 | 13.6771 | 11.85 | 12.8 | 12.95 | 11.15 | 13.25 | 12.8 | 11.54 | 12.6966 | 9.81 | 10.68 | 10.98 | 13.2 | 11.57 | 12.4 | 12.7 | 13 | 12.04 | 10.3 | 29
| 1 | 3 | 12 | 11.459 | 12.4 | 9.8 | 12.34 | 13.9666 | 11.75 | 12.55 | 12.75 | 11.25 | 13.65 | 12.7 | 12.23 | 12.9288 | 11.7 | 10.9 | 11.98 | 13.2 | 11.53 | 12.5 | 12.8 | 13 | 12.32 | 10.1 | 16
| 1 | 4 | 11.7 | 11.598 | 12.4 | 11.1 | 12.28 | 12.898 | 11.55 | 12.6 | 12.65 | 11.2 | 13.5 | 12.1 | 11.88 | 12.8834 | 12.4 | 10.43 | 11.98 | 13.3 | 11.47 | 12.5 | 12.8 | 13 | 12.04 | 10.3 | 12.0429
| 1 | 5 | 11.9 | 11.448 | 12.3 | 10.7 | 12.44 | 12.5456 | 11.85 | 12.85 | 12.95 | 11.4 | 13.25 | 12.6 | 11.86 | 12.4103 | 11 | 11.28 | 11.98 | 13.5 | 11.36 | 12 | 12.7 | 12 | 12.08 | 10.3 | 75
| 1 | 6 | 11.7 | 11.4 | 12.5 | 8.3 | 12.43 | 12.5502 | 11.9 | 13.1 | 12.55 | 11 | 13.6 | 13.1 | 12 | 12.4926 | 12.4 | 10.71 | 11.98 | 13.3 | 10.94 | 12.2 | 12.6 | 13 | 12.33 | 10.4
| Mean |  | 11.8167 | 11.4768 | 12.3667 | 10.0833 | 12.345 | 13.15 | 11.9667 | 12.7833 | 12.775 | 11.2083 | 13.5 | 12.65 | 11.8633 | 12.6727 | 11.535 | 10.75 | 11.8133 | 13.3 | 11.2967 | 12.2 | 12.65 | 12.8333 | 12.1267 | 10.25
| Median |  | 11.8 | 11.464 | 12.35 | 10.3 | 12.385 | 13.0802 | 11.85 | 12.8 | 12.775 | 11.225 | 13.55 | 12.65 | 11.87 | 12.6606 | 11.8 | 10.695 | 11.98 | 13.3 | 11.415 | 12.3 | 12.7 | 13 | 12.06 | 10.3
| Std Dev. |  | 0.17224 | 0.06618 | 0.08165 | 0.99482 | 0.12988 | 0.59092 | 0.47399 | 0.19664 | 0.16047 | 0.13197 | 0.20976 | 0.32711 | 0.24304 | 0.20695 | 0.99184 | 0.30816 | 0.40825 | 0.10954 | 0.29663 | 0.35214 | 0.18708 | 0.40825 | 0.15946 | 0.12247
| Rel.Std.Dev. |  | 1.458% | 0.5766% | 0.6602% | 9.866% | 1.052% | 4.494% | 3.961% | 1.538% | 1.256% | 1.177% | 1.554% | 2.586% | 2.049% | 1.633% | 8.599% | 2.867% | 3.456% | 0.8236% | 2.626% | 2.886% | 1.479% | 3.181% | 1.315% | 1.195%
| PDM3 |  | -1.878% | -4.7% | 2.689% | -16.27% | 2.509% | 9.193% | -0.6328% | 6.148% | 6.079% | -6.93% | 12.1% | 5.041% | -1.491% | 5.23% | -4.217% | -10.74% | -1.906% | 10.44% | -6.196% | 1.305% | 5.041% | 6.564% | 0.6957% | -14.89%
| Z-Score (Absolute) |  | 0.35 | 0.74 | 0.27 | 2.33 | 0.25 | 1.17 | 0.18 | 0.75 | 0.74 | 1.05 | 1.57 | 0.6 | 0.3 | 0.62 | 0.67 | 1.57 | 0.36 | 1.34 | 0.95 | 0.08 | 0.6 | NA | 0 | 2.14
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 922b (Certified Value 0.036 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.03 | 0.035094 | 0.04 | 0.035597 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.033 | 0.041378 | 0.04 | 0.04 | 0.0515 | 0.03 | 0.04 | 0.03 | 0.036 | 0.03 | 0.03 | 0.03 |  |  | 1
| 1 | 2 | 0.03 | 0.034723 | 0.04 | 0.034521 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.033 | 0.044381 | 0.04 | 0.04 | 0.0512 | 0.03 | 0.04 | 0.04 | 0.036 | 0.04 | 0.03 | 0.03 |  |  | 0
| 1 | 3 | 0.04 | 0.035236 | 0.03 | 0.03429 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.034 | 0.046186 | 0.04 | 0.04 | 0.0514 | 0.04 | 0.04 | 0.04 | 0.036 | 0.04 | 0.03 | 0.03 |  |  | 16
| 1 | 4 | 0.03 | 0.034719 | 0.04 | 0.03532 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.034 | 0.044653 | 0.04 | 0.04 | 0.0508 | 0.05 | 0.04 | 0.04 | 0.036 | 0.04 | 0.03 | 0.03 |  |  | 0.036287
| 1 | 5 | 0.03 | 0.035276 | 0.04 | 0.035137 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.034 | 0.042529 | 0.04 | 0.04 | 0.0516 | 0.03 | 0.04 | 0.04 | 0.038 | 0.04 | 0.03 | 0.03 |  |  | 76
| 1 | 6 | 0.03 | 0.035978 | 0.04 | 0.033898 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.035 | 0.041554 | 0.04 | 0.04 | 0.0519 | 0.03 | 0.04 | 0.04 | 0.036 | 0.04 | 0.03 | 0.03
| Mean |  | 0.031667 | 0.035171 | 0.038333 | 0.034794 | 0.031667 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.033833 | 0.043447 | 0.04 | 0.04 | 0.0514 | 0.035 | 0.04 | 0.038333 | 0.036333 | 0.038333 | 0.03 | 0.03
| Median |  | 0.03 | 0.035165 | 0.04 | 0.034829 | 0.03 | 0.03 | 0.04 | 0.03 | 0.03 | 0.04 | 0.034 | 0.043455 | 0.04 | 0.04 | 0.05145 | 0.03 | 0.04 | 0.04 | 0.036 | 0.04 | 0.03 | 0.03
| Std Dev. |  | 0.0040825 | 0.00046421 | 0.0040825 | 0.00065895 | 0.0040825 | 0 | 0 | 0 | 0 | 0 | 0.00075277 | 0.0019254 | 0 | 0 | 0.00037417 | 0.0083666 | 0 | 0.0040825 | 0.0008165 | 0.0040825 | 0 | 0
| Rel.Std.Dev. |  | 12.89% | 1.32% | 10.65% | 1.894% | 12.89% | 0% | 0% | 0% | 0% | 0% | 2.225% | 4.432% | 0% | 0% | 0.7279% | 23.9% | 0% | 10.65% | 2.247% | 10.65% | 0% | 0%
| PDM3 |  | -12.73% | -3.075% | 5.64% | -4.115% | -12.73% | -17.33% | 10.23% | -17.33% | -17.33% | 10.23% | -6.762% | 19.73% | 10.23% | 10.23% | 41.65% | -3.546% | 10.23% | 5.64% | 0.128% | 5.64% | -17.33% | -17.33%
| Z-Score (Absolute) |  | 0.65 | 0.09 | 0.41 | 0.15 | 0.65 | 0.91 | 0.67 | 0.91 | 0.91 | 0.67 | 0.3 | 1.22 | 0.67 | 0.67 | 2.48 | 0.12 | 0.67 | 0.41 | 0.09 | 0.41 | 0.91 | 0.91
| Analytical results for Cd in OREAS 922b (Certified Value 0.072 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.08 | < 0.1 | 0.1 | < 1 | 0.091 | < 0.2 | 0.09 | 0.08 | 0.07 | 0.06 | 0.06 | 0.02 | 0.08 | < 0.05 | 0.74 | 0.06 | < 0.1 | 0.08 | 0.09 | 0.07 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.07 | < 0.1 | 0.1 | < 1 | 0.075 | < 0.2 | 0.07 | 0.06 | 0.07 | 0.06 | 0.07 | 0.07 | 0.09 | < 0.05 | 1.6 | 0.05 | < 0.1 | 0.06 | 0.08 | 0.08 | < 0.1 |  |  |  | 11
| 1 | 3 | 0.07 | < 0.1 | 0.1 | < 1 | 0.065 | < 0.2 | 0.07 | 0.08 | 0.07 | 0.06 | 0.07 | 0.04 | 0.08 | 0.062516 | 0.65 | 0.05 | < 0.1 | 0.08 | 0.08 | 0.08 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.07 | < 0.1 | 0.1 | < 1 | 0.072 | < 0.2 | 0.08 | 0.06 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | < 0.05 | 0.59 | 0.06 | < 0.1 | 0.07 | 0.11 | 0.08 | < 0.1 |  |  |  | 0.072215
| 1 | 5 | 0.07 | < 0.1 | 0.1 | < 1 | 0.077 | < 0.2 | 0.07 | 0.08 | 0.07 | 0.06 | 0.06 | 0.12 | 0.08 | 0.051024 | 0.46 | 0.04 | < 0.1 | 0.08 | 0.09 | 0.08 | < 0.1 |  |  |  | 77
| 1 | 6 | 0.07 | < 0.1 | 0.1 | < 1 | 0.076 | < 0.2 | 0.07 | 0.06 | 0.07 | 0.06 | 0.07 | 0.1 | 0.09 | 0.059842 | 0.61 | 0.05 | < 0.1 | 0.07 | 0.17 | 0.07 | < 0.1
| Mean |  | 0.071667 |  | 0.1 |  | 0.076 |  | 0.075 | 0.07 | 0.068333 | 0.06 | 0.066667 | 0.07 | 0.083333 | 0.057794 | 0.775 | 0.051667 |  | 0.073333 | 0.10333 | 0.076667
| Median |  | 0.07 |  | 0.1 |  | 0.0755 |  | 0.07 | 0.07 | 0.07 | 0.06 | 0.07 | 0.07 | 0.08 | 0.059842 | 0.63 | 0.05 |  | 0.075 | 0.09 | 0.08
| Std Dev. |  | 0.0040825 |  | 1.5202e-17 |  | 0.0085323 |  | 0.0083666 | 0.010954 | 0.0040825 | 0 | 0.005164 | 0.036878 | 0.005164 | 0.0060135 | 0.41428 | 0.0075277 |  | 0.008165 | 0.034448 | 0.005164
| Rel.Std.Dev. |  | 5.696% |  | 1.52e-14% |  | 11.23% |  | 11.16% | 15.65% | 5.974% | 0% | 7.746% | 52.68% | 6.197% | 10.41% | 53.46% | 14.57% |  | 11.13% | 33.34% | 6.736%
| PDM3 |  | -0.7592% |  | 38.48% |  | 5.241% |  | 3.857% | -3.067% | -5.375% | -16.91% | -7.683% | -3.067% | 15.4% | -19.97% | 973.2% | -28.45% |  | 1.549% | 43.09% | 6.165%
| Z-Score (Absolute) |  | 0.05 | 1.2 | NA | 24.8 | 0.3 | 1.69 | 0.24 | 0.05 | 0.14 | 0.63 | 0.24 | 0.05 | 0.72 | 1.7 | 40.7 | 1.11 | 1.2 | 0.14 | 1.88 | 0.34 | 1.2
Results from laboratory 2.03 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 922b (Certified Value 60 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 0
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 62.6 | 64.852 | 57.9 | 55.878 | 61.6614 | 57.2 | 59.7 | 61.7 | 57.9 | 60.7 | 61.1 | 61.9 | 64.0217 | 60.3 | 55 | 59.7 | 55.3 | 68 | 59 | 55 |  |  |  |  | 1
| 1 | 2 | 63.7 | 64.597 | 59.6 | 55.804 | 61.8969 | 55.9 | 55.6 | 61.9 | 56.6 | 59.6 | 61.5 | 58.15 | 64.3393 | 86.6 | 58 | 62.1 | 57.7 | 70 | 58.38 | 58 |  |  |  |  | 12
| 1 | 3 | 62.8 | 64.889 | 55.5 | 57.882 | 63.3604 | 54.9 | 57.9 | 61 | 57.4 | 59.1 | 61.9 | 60.78 | 66.0876 | 62.7 | 57 | 61.9 | 58.3 | 68 | 55.03 | 54 |  |  |  |  | 16
| 1 | 4 | 63.5 | 64.777 | 57.5 | 58.443 | 60.6582 | 54.8 | 58 | 60.6 | 58.3 | 58.5 | 60.7 | 58.82 | 65.7519 | 63.2 | 56 | 59.3 | 59 | 69 | 57.56 | 56 |  |  |  |  | 60.1333
| 1 | 5 | 62.9 | 64.863 | 58.2 | 57.358 | 59.2685 | 56.7 | 60.4 | 61.2 | 56.5 | 61.1 | 63.9 | 59.6 | 62.148 | 57.8 | 57 | 59.2 | 59.7 | 68 | 58.58 | 56 |  |  |  |  | 78
| 1 | 6 | 63.5 | 64.599 | 58.1 | 56.306 | 59.6596 | 56.2 | 57.7 | 60.3 | 59.8 | 61.8 | 64.7 | 63.09 | 62.1637 | 61.5 | 54 | 60.5 | 57.8 | 75 | 58.22 | 59
| Mean |  | 63.1667 | 64.7628 | 57.8 | 56.9452 | 61.0842 | 55.95 | 58.2167 | 61.1167 | 57.75 | 60.1333 | 62.3 | 60.39 | 64.0853 | 65.35 | 56.1667 | 60.45 | 57.9667 | 69.6667 | 57.795 | 56.3333
| Median |  | 63.2 | 64.8145 | 58 | 56.832 | 61.1598 | 56.05 | 57.95 | 61.1 | 57.65 | 60.15 | 61.7 | 60.19 | 64.1805 | 62.1 | 56.5 | 60.1 | 58.05 | 68.5 | 58.3 | 56
| Std Dev. |  | 0.45461 | 0.133 | 1.3327 | 1.1083 | 1.5286 | 0.96073 | 1.6869 | 0.61779 | 1.2276 | 1.2691 | 1.6199 | 1.8883 | 1.6912 | 10.5881 | 1.472 | 1.2865 | 1.5095 | 2.7325 | 1.4347 | 1.8619
| Rel.Std.Dev. |  | 0.7197% | 0.2054% | 2.306% | 1.946% | 2.502% | 1.717% | 2.898% | 1.011% | 2.126% | 2.111% | 2.6% | 3.127% | 2.639% | 16.2% | 2.621% | 2.128% | 2.604% | 3.922% | 2.482% | 3.305%
| PDM3 |  | 5.044% | 7.699% | -3.88% | -5.302% | 1.581% | -6.957% | -3.187% | 1.635% | -3.963% | 9.485e-05% | 3.603% | 0.4269% | 6.572% | 8.675% | -6.596% | 0.5267% | -3.603% | 15.85% | -3.888% | -6.319%
| Z-Score (Absolute) |  | 0.79 | 1.22 | 0.67 | 0.9 | 0.22 | 1.17 | 0.55 | 0.23 | 0.68 | 0.03 | 0.55 | 0.03 | 1.04 | 1.38 | 1.11 | 0.05 | 0.62 | 2.55 | 0.67 | 1.06
| Analytical results for Co in OREAS 922b (Certified Value 3.41 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 3.5 | 2.9436 | 4 | 3 | 3.4 | 3.6521 | 3.2 | 3.4 | 3.6 | 3.4 | 3.4 | 3.5 | 3.3 | 3.5903 | 4.1 | 3.4 | 3.6 | 3.3 | 3.7 | 3.03 | 4 | 3.3 | 3.4 |  | 1
| 1 | 2 | 3.5 | 2.9091 | 4 | 3.1 | 3.47 | 3.6203 | 3.2 | 3.4 | 3.6 | 3.4 | 3.3 | 3.6 | 3.3 | 3.7949 | 19.9 | 3.4 | 3.5 | 3.6 | 3.9 | 3.13 | 4 | 3.4 | 3.2 |  | 32
| 1 | 3 | 3.5 | 2.8988 | 4 | 2.9 | 3.45 | 3.5265 | 3.2 | 3.4 | 3.6 | 3.4 | 3.4 | 3.7 | 3.4 | 4.6022 | 4.2 | 3.5 | 3.6 | 3.6 | 3.8 | 3.23 | 4 | 3.4 | 3.2 |  | 16
| 1 | 4 | 3.5 | 3.0139 | 4.1 | 3 | 3.47 | 3.7422 | 3.2 | 3.4 | 3.6 | 3.4 | 3.3 | 3.4 | 3.3 | 4.5312 | 4.3 | 3.4 | 3.6 | 3.5 | 3.8 | 3.23 | 4 | 3.5 | 3 |  | 3.4127
| 1 | 5 | 3.4 | 3.0158 | 4.1 | 3 | 3.45 | 3.484 | 3.2 | 3.5 | 3.6 | 3.4 | 3.4 | 3.5 | 3.3 | 3.8417 | 3.8 | 3.8 | 3.6 | 3.7 | 3.7 | 3.1 | 4 | 3.5 | 3.1 |  | 79
| 1 | 6 | 3.4 | 2.9544 | 4 | 3.2 | 3.46 | 3.4582 | 3.3 | 3.4 | 3.5 | 3.5 | 3.4 | 3.6 | 3.4 | 4.0496 | 4.3 | 3.5 | 3.5 | 3.4 | 3.6 | 3.13 | 4 | 3.5 | 3.1
| Mean |  | 3.4667 | 2.9559 | 4.0333 | 3.0333 | 3.45 | 3.5806 | 3.2167 | 3.4167 | 3.5833 | 3.4167 | 3.3667 | 3.55 | 3.3333 | 4.0683 | 6.7667 | 3.5 | 3.5667 | 3.5167 | 3.75 | 3.1417 | 4 | 3.4333 | 3.1667
| Median |  | 3.5 | 2.949 | 4 | 3 | 3.455 | 3.5734 | 3.2 | 3.4 | 3.6 | 3.4 | 3.4 | 3.55 | 3.3 | 3.9457 | 4.25 | 3.45 | 3.6 | 3.55 | 3.75 | 3.13 | 4 | 3.45 | 3.15
| Std Dev. |  | 0.05164 | 0.050131 | 0.05164 | 0.10328 | 0.026077 | 0.10959 | 0.040825 | 0.040825 | 0.040825 | 0.040825 | 0.05164 | 0.10488 | 0.05164 | 0.41335 | 6.4367 | 0.15492 | 0.05164 | 0.1472 | 0.10488 | 0.077567 | 0 | 0.08165 | 0.13663
| Rel.Std.Dev. |  | 1.49% | 1.696% | 1.28% | 3.405% | 0.7558% | 3.061% | 1.269% | 1.195% | 1.139% | 1.195% | 1.534% | 2.954% | 1.549% | 10.16% | 95.12% | 4.426% | 1.448% | 4.186% | 2.797% | 2.469% | 0% | 2.378% | 4.315%
| PDM3 |  | 1.581% | -13.38% | 18.19% | -11.12% | 1.092% | 4.918% | -5.745% | 0.1158% | 4.999% | 0.1158% | -1.349% | 4.023% | -2.326% | 19.21% | 98.28% | 2.558% | 4.511% | 3.046% | 9.883% | -7.942% | 17.21% | 0.6041% | -7.21%
| Z-Score (Absolute) |  | 0.05 | 2.74 | 3.14 | 2.32 | 0.05 | 0.67 | 1.32 | 0.23 | 0.68 | 0.23 | 0.5 | 0.5 | 0.68 | 3.33 | 18.05 | 0.23 | 0.59 | 0.32 | 1.59 | 1.73 | NA | 0.14 | 1.59
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 922b (Certified Value 19.7 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 24 | 28.875 | 18.5 | 19.6 | 19.5752 | 20 | 20 | 23 | 24 | 19 | 23 | 22 | 24.0254 | 17 | 20.4 | 21 | 20 | 17 | 19.4 | 24 | 19 | 22 |  |  | 1
| 1 | 2 | 24 | 25.407 | 18.5 | 18 | 21.1835 | 21 | 20 | 22 | 17 | 19 | 30 | 19 | 23.1764 | 32 | 20.4 | 21 | 19 | 18 | 18.95 | 25 | 20 | 19 |  |  | 33
| 1 | 3 | 24 | 25.641 | 18.6 | 19.2 | 20.0001 | 18 | 20 | 21 | 18 | 18 | 24 | 19 | 23.2227 | 19 | 20.9 | 20 | 20 | 18 | 19.65 | 25 | 22 | 20 |  |  | 16
| 1 | 4 | 24 | 25.859 | 18.5 | 18.9 | 21.4184 | 19 | 20 | 21 | 18 | 18 | 21 | 19 | 23.569 | 21 | 20.7 | 22 | 19 | 18 | 19.7 | 25 | 20 | 18 |  |  | 19.7065
| 1 | 5 | 24 | 27.038 | 18.7 | 19.9 | 18.9191 | 19 | 19 | 23 | 19 | 20 | 24 | 20 | 21.2427 | 17 | 26 | 20 | 20 | 18 | 20.35 | 25 | 19 | 18 |  |  | 80
| 1 | 6 | 24 | 26.708 | 18.6 | 20.2 | 18.3201 | 19 | 19 | 21 | 18 | 21 | 25 | 20 | 21.5594 | 18 | 21.8 | 20 | 19 | 17 | 20.4 | 28 | 20 | 20
| Mean |  | 24 | 26.588 | 18.5667 | 19.3 | 19.9027 | 19.3333 | 19.6667 | 21.8333 | 19 | 19.1667 | 24.5 | 19.8333 | 22.7993 | 20.6667 | 21.7 | 20.6667 | 19.5 | 17.6667 | 19.7417 | 25.3333 | 20 | 19.5
| Median |  | 24 | 26.2835 | 18.55 | 19.4 | 19.7877 | 19 | 20 | 21.5 | 18 | 19 | 24 | 19.5 | 23.1995 | 18.5 | 20.8 | 20.5 | 19.5 | 18 | 19.675 | 25 | 20 | 19.5
| Std Dev. |  | 0 | 1.2861 | 0.08165 | 0.78994 | 1.2269 | 1.0328 | 0.5164 | 0.98319 | 2.5298 | 1.169 | 3.0166 | 1.169 | 1.1294 | 5.7504 | 2.1689 | 0.8165 | 0.54772 | 0.5164 | 0.55805 | 1.3663 | 1.0954 | 1.5166
| Rel.Std.Dev. |  | 0% | 4.837% | 0.4398% | 4.093% | 6.165% | 5.342% | 2.626% | 4.503% | 13.31% | 6.099% | 12.31% | 5.894% | 4.954% | 27.82% | 9.995% | 3.951% | 2.809% | 2.923% | 2.827% | 5.393% | 5.477% | 7.777%
| PDM3 |  | 21.79% | 34.92% | -5.784% | -2.063% | 0.9958% | -1.894% | -0.2021% | 10.79% | -3.585% | -2.739% | 24.32% | 0.6436% | 15.69% | 4.872% | 10.12% | 4.872% | -1.048% | -10.35% | 0.1785% | 28.55% | 1.489% | -1.048%
| Z-Score (Absolute) |  | 3.49 | 5.67 | 1.1 | 0.48 | 0.03 | 0.45 | 0.17 | 1.66 | 0.73 | 0.59 | 3.91 | 0.03 | 2.47 | 0.67 | 1.55 | 0.67 | 0.31 | 1.86 | 0.11 | 4.61 | 0.11 | 0.31
| Analytical results for Cs in OREAS 922b (Certified Value 1.11 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 1.77 | 1.8397 | 0.98 | 1.289 | 0.86 | 1.02 | 0.96 | 0.94 | 1 | 1.26 | 1.04 | 1.3693 | 1.1 | 1.1 | 1.28 | 1.32 | 0.93 | 0.9 | [6 empty] | 1
| 1 | 2 | 1.75 | 1.8306 | 0.97 | 1.4233 | 0.92 | 1.04 | 0.98 | 0.91 | 1.02 | 1.25 | 0.98 | 1.3732 | 1.7 | 1.14 | 1.39 | 1.34 | 0.97 | 0.8 | [6 empty] | 14
| 1 | 3 | 1.77 | 1.8538 | 0.98 | 1.468 | 0.89 | 1.04 | 0.95 | 0.89 | 0.97 | 1.28 | 0.99 | 1.4679 | 1.11 | 1.18 | 1.14 | 1.36 | 0.94 | 0.9 | [6 empty] | 16
| 1 | 4 | 1.72 | 1.8546 | 0.98 | 1.298 | 0.9 | 1.05 | 0.97 | 0.95 | 0.98 | 1.16 | 0.99 | 1.4709 | 1.15 | 1.19 | 1.28 | 1.33 | 0.99 | 0.8 | [6 empty] | 1.1112
| 1 | 5 | 1.69 | 1.8341 | 0.98 | 1.2473 | 0.95 | 1.05 | 0.94 | 0.91 | 1.06 | 1.26 | 0.97 | 1.3608 | 0.91 | 1.14 | 1.34 | 1.36 | 0.99 | 0.8 | [6 empty] | 81
| 1 | 6 | 1.71 | 1.8173 | 0.99 | 1.345 | 0.92 | 1.06 | 0.94 | 0.95 | 1.04 | 1.26 | 1.06 | 1.3878 | 1.11 | 1.19 | 1.31 | 1.3 | 1 | 0.8
| Mean |  | 1.735 | 1.8384 | 0.98 | 1.3451 | 0.90667 | 1.0433 | 0.95667 | 0.925 | 1.0117 | 1.245 | 1.005 | 1.405 | 1.18 | 1.1567 | 1.29 | 1.335 | 0.97 | 0.83333
| Median |  | 1.735 | 1.8369 | 0.98 | 1.3215 | 0.91 | 1.045 | 0.955 | 0.925 | 1.01 | 1.26 | 0.99 | 1.3805 | 1.11 | 1.16 | 1.295 | 1.335 | 0.98 | 0.8
| Std Dev. |  | 0.033317 | 0.014297 | 0.0063246 | 0.085023 | 0.030768 | 0.013663 | 0.01633 | 0.0251 | 0.034881 | 0.042778 | 0.036194 | 0.050659 | 0.26848 | 0.036148 | 0.08438 | 0.023452 | 0.028983 | 0.05164
| Rel.Std.Dev. |  | 1.92% | 0.7777% | 0.6454% | 6.321% | 3.394% | 1.31% | 1.707% | 2.713% | 3.448% | 3.436% | 3.601% | 3.606% | 22.75% | 3.125% | 6.541% | 1.757% | 2.988% | 6.197%
| PDM3 |  | 56.14% | 65.44% | -11.81% | 21.05% | -18.41% | -6.108% | -13.91% | -16.76% | -8.958% | 12.04% | -9.558% | 26.44% | 6.191% | 4.091% | 16.09% | 20.14% | -12.71% | -25.01%
| Z-Score (Absolute) |  | 2.19 | 2.58 | 0.67 | 0.71 | 0.95 | 0.43 | 0.76 | 0.88 | 0.55 | 0.33 | 0.57 | 0.94 | 0.09 | 0 | 0.5 | 0.67 | 0.71 | NA
Results from laboratory 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 922b (Certified Value 0.221 wt.%) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 0.218 | 0.21412 | 0.22407 | 0.2229 | 0.22423 | 0.2239 | 0.226 | 0.221 | 0.233 | 0.214 | 0.226 | 0.227 | 0.2174 | 0.21711 | 0.213 | 0.21517 | 0.2182 | 0.21458 | 0.2237 | 0.2208 | 0.213 | 0.2199 | 0.21838 | 0.22277 | 1
| 1 | 2 | 0.216 | 0.21948 | 0.22453 | 0.2283 | 0.22829 | 0.21798 | 0.218 | 0.219 | 0.23 | 0.214 | 0.215 | 0.221 | 0.2212 | 0.22149 | 0.226 | 0.22153 | 0.2158 | 0.21592 | 0.2256 | 0.2224 | 0.2166 | 0.226 | 0.22065 | 0.22317 | 35
| 1 | 3 | 0.215 | 0.21589 | 0.22572 | 0.2229 | 0.23206 | 0.21837 | 0.213 | 0.218 | 0.228 | 0.21 | 0.213 | 0.22 | 0.2199 | 0.21988 | 0.217 | 0.21873 | 0.2181 | 0.21609 | 0.2249 | 0.2282 | 0.2225 | 0.226 | 0.21786 | 0.22041 | 16
| 1 | 4 | 0.218 | 0.21495 | 0.22445 | 0.2275 | 0.22731 | 0.22424 | 0.21 | 0.214 | 0.23 | 0.213 | 0.217 | 0.212 | 0.2229 | 0.22348 | 0.22 | 0.21549 | 0.2164 | 0.21342 | 0.2278 | 0.228 | 0.2227 | 0.2249 | 0.2223 | 0.22316 | 0.22066
| 1 | 5 | 0.216 | 0.21482 | 0.22983 | 0.2176 | 0.22882 | 0.21618 | 0.212 | 0.221 | 0.229 | 0.217 | 0.22 | 0.23 | 0.2212 | 0.2151 | 0.223 | 0.21773 | 0.2125 | 0.21422 | 0.2245 | 0.2227 | 0.2207 | 0.2233 | 0.22512 | 0.22314 | 82
| 1 | 6 | 0.221 | 0.21422 | 0.22742 | 0.2261 | 0.23287 | 0.21379 | 0.218 | 0.221 | 0.23 | 0.213 | 0.218 | 0.225 | 0.2263 | 0.21562 | 0.215 | 0.22607 | 0.2203 | 0.21438 | 0.2255 | 0.2202 | 0.2185 | 0.2537 | 0.22248 | 0.22005
| Mean |  | 0.21733 | 0.21558 | 0.226 | 0.22422 | 0.22893 | 0.21908 | 0.21617 | 0.219 | 0.23 | 0.2135 | 0.21817 | 0.2225 | 0.22148 | 0.21878 | 0.219 | 0.21912 | 0.21688 | 0.21477 | 0.22533 | 0.22372 | 0.219 | 0.22897 | 0.22113 | 0.22212
| Median |  | 0.217 | 0.21488 | 0.22512 | 0.2245 | 0.22856 | 0.21818 | 0.2155 | 0.22 | 0.23 | 0.2135 | 0.2175 | 0.223 | 0.2212 | 0.21849 | 0.2185 | 0.21823 | 0.21725 | 0.21448 | 0.2252 | 0.22255 | 0.2196 | 0.22545 | 0.22148 | 0.22295
| Std Dev. |  | 0.0021602 | 0.0020144 | 0.0022405 | 0.003956 | 0.0031783 | 0.0041942 | 0.0058109 | 0.0027568 | 0.0016733 | 0.0022583 | 0.004535 | 0.0063482 | 0.0029849 | 0.0033731 | 0.0049396 | 0.0041231 | 0.0026649 | 0.0010371 | 0.0013952 | 0.0035239 | 0.003764 | 0.01233 | 0.0027422 | 0.0014735
| Rel.Std.Dev. |  | 0.994% | 0.9344% | 0.9913% | 1.764% | 1.388% | 1.914% | 2.688% | 1.259% | 0.7275% | 1.058% | 2.079% | 2.853% | 1.348% | 1.542% | 2.256% | 1.882% | 1.229% | 0.4829% | 0.6192% | 1.575% | 1.719% | 5.385% | 1.24% | 0.6634%
| PDM3 |  | -1.507% | -2.302% | 2.421% | 1.612% | 3.748% | -0.7169% | -2.036% | -0.752% | 4.233% | -3.245% | -1.13% | 0.8342% | 0.3734% | -0.852% | -0.752% | -0.6977% | -1.711% | -2.67% | 2.118% | 1.386% | -0.752% | 3.765% | 0.2141% | 0.6604%
| Z-Score (Absolute) |  | 0.4 | 0.8 | 1.56 | 1.16 | 2.23 | 0 | 0.66 | 0.02 | 2.47 | 1.27 | 0.21 | 0.77 | 0.54 | 0.07 | 0.02 | 0 | 0.5 | 0.98 | 1.41 | 1.05 | 0.02 | 2.24 | 0.46 | 0.68
| Analytical results for Dy in OREAS 922b (Indicative Value 2 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 1.9193; 2.02: 1.8661; 2.03: 2.2 | 1
| 1 | 2 | 2.01: 1.8294; 2.02: 2.0188; 2.03: 2.2 | 3
| 1 | 3 | 2.01: 1.8306; 2.02: 2.109; 2.03: 2 | 16
| 1 | 4 | 2.01: 1.9163; 2.02: 1.899; 2.03: 2.1 | 1.996
| 1 | 5 | 2.01: 1.912; 2.02: 1.7587; 2.03: 2.2 | 9
| 1 | 6 | 2.01: 1.8586; 2.02: 1.9098; 2.03: 2.4
| Mean |  | 2.01: 1.8777; 2.02: 1.9269; 2.03: 2.1833
| Median |  | 2.01: 1.8853; 2.02: 1.9044; 2.03: 2.2
| Std Dev. |  | 2.01: 0.043158; 2.02: 0.12218; 2.03: 0.13292
| Rel.Std.Dev. |  | 2.01: 2.298%; 2.02: 6.341%; 2.03: 6.088%
| PDM3 |  | 2.01: -5.926%; 2.02: -3.461%; 2.03: 9.387%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0; 2.03: 3.51
| Analytical results for Er in OREAS 922b (Indicative Value 0.82 ppm) | [25 empty] | Indicative
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 0.8216; 2.02: 0.72196; 2.03: 0.9 | 1
| 1 | 2 | 2.01: 0.7916; 2.02: 0.77528; 2.03: 0.9 | 4
| 1 | 3 | 2.01: 0.799; 2.02: 0.77224; 2.03: 0.9 | 16
| 1 | 4 | 2.01: 0.8372; 2.02: 0.76701; 2.03: 0.8 | 0.8186
| 1 | 5 | 2.01: 0.8122; 2.02: 0.71341; 2.03: 0.9 | 10
| 1 | 6 | 2.01: 0.8076; 2.02: 0.71577; 2.03: 1
| Mean |  | 2.01: 0.81153; 2.02: 0.74428; 2.03: 0.9
| Median |  | 2.01: 0.8099; 2.02: 0.74449; 2.03: 0.9
| Std Dev. |  | 2.01: 0.016312; 2.02: 0.030076; 2.03: 0.063246
| Rel.Std.Dev. |  | 2.01: 2.01%; 2.02: 4.041%; 2.03: 7.027%
| PDM3 |  | 2.01: -0.8637%; 2.02: -9.08%; 2.03: 9.943%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.03: 0.89
| Analytical results for Eu in OREAS 922b (Indicative Value 0.88 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 0.8503; 2.02: 0.80476; 2.03: 0.9 | 1
| 1 | 2 | 2.01: 0.8366; 2.02: 0.89474; 2.03: 1 | 5
| 1 | 3 | 2.01: 0.8323; 2.02: 0.81384; 2.03: 0.9 | 16
| 1 | 4 | 2.01: 0.879; 2.02: 0.83567; 2.03: 1 | 0.88265
| 1 | 5 | 2.01: 0.8835; 2.02: 0.83567; 2.03: 1 | 11
| 1 | 6 | 2.01: 0.8212; 2.02: 0.80016; 2.03: 1
| Mean |  | 2.01: 0.85048; 2.02: 0.83081; 2.03: 0.96667
| Median |  | 2.01: 0.84345; 2.02: 0.82475; 2.03: 1
| Std Dev. |  | 2.01: 0.025629; 2.02: 0.034757; 2.03: 0.05164
| Rel.Std.Dev. |  | 2.01: 3.013%; 2.02: 4.183%; 2.03: 5.342%
| PDM3 |  | 2.01: -3.645%; 2.02: -5.874%; 2.03: 9.518%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.03: 3.98
| Analytical results for Fe in OREAS 922b (Certified Value 2.86 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2.97 | 2.9247 | 2.72 | 3.018 | 2.9218 | 2.85 | 2.9 | 3.02 | 2.72 | 2.86 | 3.01 | 2.88 | 2.9732 | 3.2 | 2.86 | 2.7667 | 2.74 | 2.79 | 2.85 | 2.71 | 2.79 | 2.76 | 2.87 |  | 1
| 1 | 2 | 2.94 | 2.8185 | 2.79 | 3.021 | 2.8356 | 2.7 | 2.81 | 3.03 | 2.73 | 2.73 | 2.99 | 2.91 | 3.0762 | 3.34 | 2.97 | 2.7523 | 2.77 | 2.91 | 2.87 | 2.8 | 2.87 | 2.75 | 2.88 |  | 0
| 1 | 3 | 2.94 | 2.8622 | 2.72 | 3.008 | 2.8252 | 2.69 | 2.78 | 3 | 2.69 | 2.76 | 3.01 | 2.9 | 3.0173 | 3.17 | 2.94 | 2.7423 | 2.83 | 2.91 | 2.87 | 2.83 | 2.87 | 2.88 | 2.86 |  | 16
| 1 | 4 | 2.97 | 2.937 | 2.78 | 2.979 | 2.8791 | 2.71 | 2.72 | 3.01 | 2.7 | 2.76 | 2.86 | 2.93 | 3.0597 | 2.95 | 2.89 | 2.7259 | 2.8 | 2.87 | 2.86 | 2.87 | 2.84 | 2.86 | 2.9 |  | 2.8617
| 1 | 5 | 2.96 | 2.9564 | 2.65 | 2.976 | 2.883 | 2.66 | 2.82 | 2.99 | 2.73 | 2.77 | 3.02 | 2.92 | 2.922 | 3.16 | 3 | 2.6933 | 2.8 | 2.91 | 2.86 | 2.83 | 2.81 | 2.89 | 2.88 |  | 83
| 1 | 6 | 2.93 | 2.8226 | 2.77 | 3.056 | 2.7912 | 2.76 | 2.79 | 3.03 | 2.71 | 2.78 | 3.06 | 3 | 2.9168 | 3.37 | 2.99 | 2.7999 | 2.76 | 2.86 | 2.84 | 2.82 | 3.18 | 2.89 | 2.86
| Mean |  | 2.9517 | 2.8869 | 2.7383 | 3.0097 | 2.856 | 2.7283 | 2.8033 | 3.0133 | 2.7133 | 2.7767 | 2.9917 | 2.9233 | 2.9942 | 3.1983 | 2.9417 | 2.7467 | 2.7833 | 2.875 | 2.8583 | 2.81 | 2.8933 | 2.8383 | 2.875
| Median |  | 2.95 | 2.8935 | 2.745 | 3.013 | 2.8573 | 2.705 | 2.8 | 3.015 | 2.715 | 2.765 | 3.01 | 2.915 | 2.9953 | 3.185 | 2.955 | 2.7473 | 2.785 | 2.89 | 2.86 | 2.825 | 2.855 | 2.87 | 2.875
| Std Dev. |  | 0.017224 | 0.0603 | 0.052694 | 0.029737 | 0.047243 | 0.067946 | 0.058878 | 0.01633 | 0.01633 | 0.044121 | 0.068532 | 0.041312 | 0.068084 | 0.15065 | 0.056362 | 0.036236 | 0.03266 | 0.047223 | 0.01169 | 0.054037 | 0.14404 | 0.065549 | 0.015166
| Rel.Std.Dev. |  | 0.5835% | 2.089% | 1.924% | 0.988% | 1.654% | 2.49% | 2.1% | 0.5419% | 0.6018% | 1.589% | 2.291% | 1.413% | 2.274% | 4.71% | 1.916% | 1.319% | 1.173% | 1.643% | 0.409% | 1.923% | 4.978% | 2.309% | 0.5275%
| PDM3 |  | 3.145% | 0.882% | -4.31% | 5.172% | -0.1986% | -4.659% | -2.038% | 5.3% | -5.183% | -2.97% | 4.543% | 2.155% | 4.632% | 11.77% | 2.796% | -4.016% | -2.737% | 0.4662% | -0.1162% | -1.805% | 1.107% | -0.8151% | 0.4662%
| Z-Score (Absolute) |  | 0.67 | 0.1 | 1.2 | 1.18 | 0.17 | 1.29 | 0.63 | 1.22 | 1.42 | 0.86 | 1.03 | 0.43 | 1.05 | 2.84 | 0.59 | 1.13 | 0.81 | 0 | 0.15 | 0.57 | 0.16 | 0.32 | 0
| Analytical results for Ga in OREAS 922b (Certified Value 2.42 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 3 | 3.2559 | 2.3 | < 5 | 2 | 2.9667 | 2.24 | 2.58 | 2.41 | 2.24 | 2.33 | 1.82 | 2.2 | 2.3883 | < 10 | 2.6 | 2 | 2.52 | 1.9 | 3.04 | 2.51 | 5 |  |  | 1
| 1 | 2 | 3 | 3.2094 | 2.3 | < 5 | 2.1 | 2.9721 | 2.15 | 2.52 | 2.41 | 2.25 | 2.33 | 1.5 | 2.1 | 2.5636 | < 10 | 2.6 | 2 | 2.47 | 2 | 3.42 | 2.68 | 4 |  |  | 18
| 1 | 3 | 3 | 3.2593 | 2.2 | < 5 | 2 | 3.167 | 2.01 | 2.65 | 2.37 | 2.23 | 2.3 | 1.63 | 2.2 | 2.5384 | < 10 | 2.5 | 2 | 2.42 | 2 | 3.38 | 2.57 | 5 |  |  | 16
| 1 | 4 | 3 | 3.3001 | 2.3 | < 5 | 2 | 3.1357 | 2.09 | 2.53 | 2.37 | 2.19 | 2.3 | 1.36 | 2.1 | 2.7135 | < 10 | 2.7 | 2 | 2.47 | 2 | 3.46 | 2.6 | 4 |  |  | 2.4167
| 1 | 5 | 2.9 | 3.3474 | 2.2 | < 5 | 2 | 2.8834 | 2.11 | 2.58 | 2.38 | 2.21 | 2.37 | 1.49 | 2.2 | 2.2782 | < 10 | 2.7 | 2 | 2.49 | 1.9 | 3.37 | 2.65 | 4 |  |  | 84
| 1 | 6 | 3 | 3.2748 | 2.4 | < 5 | 2.1 | 2.9336 | 2.08 | 2.56 | 2.34 | 2.29 | 2.34 | 1.58 | 2.2 | 2.3702 | < 10 | 2.6 | 2 | 2.31 | 2 | 3.4 | 2.68 | 4
| Mean |  | 2.9833 | 3.2745 | 2.2833 |  | 2.0333 | 3.0097 | 2.1133 | 2.57 | 2.38 | 2.235 | 2.3283 | 1.5633 | 2.1667 | 2.4754 |  | 2.6167 | 2 | 2.4467 | 1.9667 | 3.345 | 2.615 | 4.3333
| Median |  | 3 | 3.267 | 2.3 |  | 2 | 2.9694 | 2.1 | 2.57 | 2.375 | 2.235 | 2.33 | 1.54 | 2.2 | 2.4633 |  | 2.6 | 2 | 2.47 | 2 | 3.39 | 2.625 | 4
| Std Dev. |  | 0.040825 | 0.04645 | 0.075277 |  | 0.05164 | 0.11456 | 0.077115 | 0.046476 | 0.026833 | 0.034496 | 0.026394 | 0.15578 | 0.05164 | 0.15874 |  | 0.075277 | 0 | 0.074476 | 0.05164 | 0.15281 | 0.06775 | 0.5164
| Rel.Std.Dev. |  | 1.368% | 1.419% | 3.297% |  | 2.54% | 3.806% | 3.649% | 1.808% | 1.127% | 1.543% | 1.134% | 9.964% | 2.383% | 6.413% |  | 2.877% | 0% | 3.044% | 2.626% | 4.568% | 2.591% | 11.92%
| PDM3 |  | 23.45% | 35.49% | -5.519% |  | -15.86% | 24.54% | -12.55% | 6.343% | -1.519% | -7.519% | -3.657% | -35.31% | -10.35% | 2.427% |  | 8.274% | -17.24% | 1.239% | -18.62% | 38.41% | 8.205% | 79.31%
| Z-Score (Absolute) |  | 1.35 | 2.11 | 0.46 | 0.1 | 1.11 | 1.42 | 0.9 | 0.28 | 0.21 | 0.59 | 0.34 | 2.33 | 0.76 | 0.04 | 6.58 | 0.4 | NA | 0.04 | 1.28 | 2.29 | 0.4 | NA
Results from laboratories 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 922b (Indicative Value 4.07 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  |  | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g |  |  | 2
| 1 | 1 | 2.01: 4.4961; 2.02: 3.5415; 2.03: 4.1 |  |  | 1
| 1 | 2 | 2.01: 4.343; 2.02: 3.6662; 2.03: 4.2 |  |  | 6
| 1 | 3 | 2.01: 4.3346; 2.02: 3.6827; 2.03: 4.1 |  |  | 16
| 1 | 4 | 2.01: 4.5588; 2.02: 3.598; 2.03: 4.1 |  |  | 4.0669
| 1 | 5 | 2.01: 4.549; 2.02: 3.4766; 2.03: 4.2 |  |  | 12
| 1 | 6 | 2.01: 4.3582; 2.02: 3.5987; 2.03: 4.3
| Mean |  | 2.01: 4.4399; 2.02: 3.5939; 2.03: 4.1667
| Median |  | 2.01: 4.4272; 2.02: 3.5984; 2.03: 4.15
| Std Dev. |  | 2.01: 0.10616; 2.02: 0.077026; 2.03: 0.08165
| Rel.Std.Dev. |  | 2.01: 2.391%; 2.02: 2.143%; 2.03: 1.96%
| PDM3 |  | 2.01: 9.174%; 2.02: -11.63%; 2.03: 2.454%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 1.41; 2.03: 0
| Analytical results for Ge in OREAS 922b (Certified Value 0.1 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
|  |  | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [9 empty] | 2
| 1 | 1 | 1.1804 | < 5 | < 0.2 | 0.12 | 0.11 | 0.13 | 0.09 | 0.08 | < 0.1 | < 0.1 | 0.12 | 0.15 | < 0.1 | 0.05 | 1.6 | [9 empty] | 1
| 1 | 2 | 1.121 | < 5 | < 0.2 | 0.1 | 0.09 | 0.12 | 0.1 | 0.07 | < 0.1 | < 0.1 | 0.19 | 0.15 | < 0.1 | 0.05 | 1.5 | [9 empty] | 41
| 1 | 3 | 1.1522 | < 5 | < 0.2 | 0.1 | 0.1 | 0.12 | 0.09 | 0.07 | < 0.1 | < 0.1 | 0.13 | 0.15 | < 0.1 | 0.05 | 1.5 | [9 empty] | 16
| 1 | 4 | 1.1814 | < 5 | < 0.2 | 0.1 | 0.1 | 0.12 | 0.08 | 0.08 | < 0.1 | < 0.1 | 0.12 | 0.17 | < 0.1 | 0.06 | 1.4 | [9 empty] | 0.10272
| 1 | 5 | 1.1938 | < 5 | < 0.2 | 0.1 | 0.12 | 0.13 | 0.09 | 0.08 | < 0.1 | < 0.1 | 0.12 | 0.16 | < 0.1 | 0.05 | 1.4 | [9 empty] | 85
| 1 | 6 | 1.1353 | < 5 | < 0.2 | 0.1 | 0.12 | 0.12 | 0.09 | 0.07 | < 0.1 | < 0.1 | 0.1 | 0.16 | < 0.1 | 0.05 | 1.5
| Mean |  | 2.01: 1.1607; 2.04: 0.10333; 2.05: 0.10667; 2.06: 0.12333; 2.07: 0.09; 2.08: 0.075; 2.11: 0.13; 2.12: 0.15667; 2.14: 0.051667; 2.15: 1.4833
| Median |  | 2.01: 1.1663; 2.04: 0.1; 2.05: 0.105; 2.06: 0.12; 2.07: 0.09; 2.08: 0.075; 2.11: 0.12; 2.12: 0.155; 2.14: 0.05; 2.15: 1.5
| Std Dev. |  | 2.01: 0.029002; 2.04: 0.008165; 2.05: 0.012111; 2.06: 0.005164; 2.07: 0.0063246; 2.08: 0.0054772; 2.11: 0.030984; 2.12: 0.008165; 2.14: 0.0040825; 2.15: 0.075277
| Rel.Std.Dev. |  | 2.01: 2.499%; 2.04: 7.902%; 2.05: 11.35%; 2.06: 4.187%; 2.07: 7.027%; 2.08: 7.303%; 2.11: 23.83%; 2.12: 5.212%; 2.14: 7.902%; 2.15: 5.075%
| PDM3 |  | 2.01: 1030%; 2.04: 0.5949%; 2.05: 3.84%; 2.06: 20.06%; 2.07: -12.39%; 2.08: -26.99%; 2.11: 26.55%; 2.12: 52.51%; 2.14: -49.7%; 2.15: 1344%
| Z-Score (Absolute) |  | 18.23 | 41.29 | 0.03 | 0.03 | 0.09 | 0.37 | 0.2 | 0.46 | 0.89 | 0.89 | 0.49 | 0.95 | 0.89 | 0.86 | NA
Results from laboratory 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 922b (Certified Value 0.32 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 2
|  |  | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [7 empty] | 2
| 1 | 1 | 0.34 | 0.36 | 0.2976 | 0.26 | 0.28 | 0.28 | 0.29 | 0.32 | 0.3 | 0.34 | 0.36737 | 0.55 | 0.21 | 0.37 | 0.27 | 0.32 | 0.3 | [7 empty] | 1
| 1 | 2 | 0.34 | 0.37 | 0.3348 | 0.27 | 0.33 | 0.28 | 0.29 | 0.32 | 0.3 | 0.35 | 0.37808 | 0.48 | 0.25 | 0.39 | 0.32 | 0.32 | 0.3 | [7 empty] | 25
| 1 | 3 | 0.32 | 0.35 | 0.3217 | 0.26 | 0.29 | 0.27 | 0.29 | 0.32 | 0.3 | 0.35 | 0.37136 | 0.57 | 0.25 | 0.36 | 0.35 | 0.35 | 0.3 | [7 empty] | 16
| 1 | 4 | 0.33 | 0.38 | 0.33528 | 0.27 | 0.33 | 0.27 | 0.3 | 0.32 | 0.3 | 0.37 | 0.37637 | 0.53 | 0.23 | 0.38 | 0.37 | 0.32 | 0.3 | [7 empty] | 0.31807
| 1 | 5 | 0.34 | 0.36 | 0.26537 | 0.26 | 0.32 | 0.27 | 0.29 | 0.33 | 0.3 | 0.35 | 0.36803 | 0.52 | 0.25 | 0.37 | 0.34 | 0.31 | 0.3 | [7 empty] | 86
| 1 | 6 | 0.33 | 0.37 | 0.28308 | 0.25 | 0.3 | 0.27 | 0.29 | 0.32 | 0.3 | 0.37 | 0.36851 | 0.57 | 0.22 | 0.37 | 0.33 | 0.34 | 0.3
| Mean |  | 0.33333 | 0.365 | 0.30631 | 0.26167 | 0.30833 | 0.27333 | 0.29167 | 0.32167 | 0.3 | 0.355 | 0.37162 | 0.53667 | 0.235 | 0.37333 | 0.33 | 0.32667 | 0.3
| Median |  | 0.335 | 0.365 | 0.30965 | 0.26 | 0.31 | 0.27 | 0.29 | 0.32 | 0.3 | 0.35 | 0.36994 | 0.54 | 0.24 | 0.37 | 0.335 | 0.32 | 0.3
| Std Dev. |  | 0.008165 | 0.010488 | 0.028913 | 0.0075277 | 0.02137 | 0.005164 | 0.0040825 | 0.0040825 | 0 | 0.012247 | 0.0045826 | 0.034448 | 0.017607 | 0.010328 | 0.034059 | 0.015055 | 0
| Rel.Std.Dev. |  | 2.449% | 2.873% | 9.439% | 2.877% | 6.931% | 1.889% | 1.4% | 1.269% | 0% | 3.45% | 1.233% | 6.419% | 7.492% | 2.766% | 10.32% | 4.609% | 0%
| PDM3 |  | 4.8% | 14.76% | -3.697% | -17.73% | -3.06% | -14.06% | -8.3% | 1.132% | -5.68% | 11.61% | 16.84% | 68.73% | -26.12% | 17.38% | 3.752% | 2.704% | -5.68%
| Z-Score (Absolute) |  | 0.13 | 0.74 | 0.39 | 1.25 | 0.35 | 1.03 | 0.67 | 0.1 | NA | 0.55 | 0.87 | 4.05 | 1.77 | 0.9 | 0.06 | 0 | NA
Results from laboratories 2.12 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 922b (Certified Value 0.023 ppm) | [25 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | NA | < 5 | 0.03 | < 1 | 0.03 | 0.04 | 0.02 | 0.02 | 0.03 | 0.03 | 0.03 | < 0.05 | < 0.01 | 0.018 | 0.18 | 0.02 | 0.13 | < 0.1 | 0.022 | 0.315 |  |  |  |  | 1
| 1 | 2 | NA | < 5 | 0.02 | < 1 | 0.03 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.05 | < 0.01 | 0.01 | 0.25 | 0.02 | 0.13 | < 0.1 | 0.024 | 0.415 |  |  |  |  | 23
| 1 | 3 | NA | < 5 | 0.03 | < 1 | 0.02 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.05 | < 0.01 | 0.01 | 0.29 | 0.02 | 0.13 | < 0.1 | 0.023 | 0.373 |  |  |  |  | 16
| 1 | 4 | NA | < 5 | 0.03 | < 1 | 0.04 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.05 | < 0.01 | 0.022 | 0.39 | 0.02 | 0.13 | < 0.1 | 0.026 | 0.375 |  |  |  |  | 0.022682
| 1 | 5 | NA | < 5 | 0.02 | < 1 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | < 0.05 | < 0.01 | 0.024 | 0.71 | 0.02 | 0.15 | < 0.1 | 0.025 | 0.379 |  |  |  |  | 87
| 1 | 6 | NA | < 5 | 0.03 | < 1 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.03 | 0.03 | < 0.05 | < 0.01 | 0.01 | 0.15 | 0.02 | 0.14 | < 0.1 | 0.023 | 0.397
| Mean |  |  |  | 0.026667 |  | 0.026667 | 0.026667 | 0.02 | 0.02 | 0.023333 | 0.023333 | 0.023333 |  |  | 0.015667 | 0.32833 | 0.02 | 0.135 |  | 0.023833 | 0.37567
| Median |  |  |  | 0.03 |  | 0.025 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |  |  | 0.014 | 0.27 | 0.02 | 0.13 |  | 0.0235 | 0.377
| Std Dev. |  |  |  | 0.005164 |  | 0.008165 | 0.010328 | 0 | 0 | 0.005164 | 0.005164 | 0.005164 |  |  | 0.0065013 | 0.20537 | 0 | 0.0083666 |  | 0.001472 | 0.033768
| Rel.Std.Dev. |  |  |  | 19.36% |  | 30.62% | 38.73% | 0% | 0% | 22.13% | 22.13% | 22.13% |  |  | 41.5% | 62.55% | 0% | 6.197% |  | 6.176% | 8.989%
| PDM3 |  |  |  | 17.57% |  | 17.57% | 17.57% | -11.82% | -11.82% | 2.872% | 2.872% | 2.872% |  |  | -30.93% | 1348% | -11.82% | 495.2% |  | 5.077% | 1556%
| Z-Score (Absolute) |  | NA | 333.78 | 0.22 | 64.06 | 0.22 | 0.22 | 0.67 | 0.67 | 0.22 | 0.22 | 0.22 | 0 | 2.7 | 1.26 | 40.91 | 0.67 | 14.83 | 3.37 | 0.16 | 47.29
| Analytical results for Ho in OREAS 922b (Indicative Value 0.3 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | 2.03: Lab; 2.04: Lab; 2.05: Lab |  |  |  |  | 1
| No | No | 2.03: 2.02; 2.04: 2.06; 2.05: 2.12 |  |  |  |  | ppm
|  |  | 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS |  |  |  |  | 2
|  |  | 2.03: 0.5g; 2.04: 0.2g; 2.05: 0.5g |  |  |  |  | 2
| 1 | 1 | 2.03: 0.36699; 2.04: 0.29206; 2.05: 0.3 |  |  |  |  | 1
| 1 | 2 | 2.03: 0.36153; 2.04: 0.32502; 2.05: 0.3 |  |  |  |  | 3
| 1 | 3 | 2.03: 0.35646; 2.04: 0.32192; 2.05: 0.3 |  |  |  |  | 16
| 1 | 4 | 2.03: 0.37869; 2.04: 0.29651; 2.05: 0.3 |  |  |  |  | 0.29881
| 1 | 5 | 2.03: 0.37609; 2.04: 0.2669; 2.05: 0.3 |  |  |  |  | 9
| 1 | 6 | 2.03: 0.35906; 2.04: 0.28335; 2.05: 0.3
| Mean |  | 2.03: 0.36647; 2.04: 0.29763; 2.05: 0.3
| Median |  | 2.03: 0.36426; 2.04: 0.29428; 2.05: 0.3
| Std Dev. |  | 2.03: 0.0091832; 2.04: 0.022456; 2.05: 0
| Rel.Std.Dev. |  | 2.03: 2.506%; 2.04: 7.545%; 2.05: 0%
| PDM3 |  | 2.03: 22.64%; 2.04: -0.3973%; 2.05: 0.3973%
| Z-Score (Absolute) |  | 2.03: 18.88; 2.04: 0.67; 2.05: 0
| Analytical results for In in OREAS 922b (Certified Value 0.2 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 0.21 | 0.19152 | < 5 | 0.2 | 0.19473 | 0.189 | 0.208 | 0.216 | 0.193 | 0.207 | 0.19 | 0.2 | 0.20233 | 0.018 | 0.2 | 0.2 | 0.17 | 0.19 | 0.18 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.19152 | < 5 | 0.21 | 0.20126 | 0.194 | 0.201 | 0.204 | 0.195 | 0.193 | 0.2 | 0.205 | 0.20054 | 0.024 | 0.22 | 0.21 | 0.16 | 0.202 | 0.19 | [5 empty] | 11
| 1 | 3 | 0.21 | 0.18648 | < 5 | 0.2 | 0.19789 | 0.185 | 0.205 | 0.215 | 0.193 | 0.196 | 0.2 | 0.197 | 0.2071 | 0.014 | 0.22 | 0.21 | 0.17 | 0.199 | 0.19 | [5 empty] | 16
| 1 | 4 | 0.21 | 0.21518 | < 5 | 0.2 | 0.19368 | 0.189 | 0.204 | 0.214 | 0.2 | 0.193 | 0.19 | 0.208 | 0.20573 | 0.014 | 0.2 | 0.21 | 0.14 | 0.202 | 0.19 | [5 empty] | 0.2002
| 1 | 5 | 0.2 | 0.20902 | < 5 | 0.19 | 0.18473 | 0.2 | 0.212 | 0.214 | 0.19 | 0.198 | 0.21 | 0.189 | 0.20179 | 0.017 | 0.23 | 0.21 | 0.15 | 0.202 | 0.18 | [5 empty] | 88
| 1 | 6 | 0.2 | 0.19138 | < 5 | 0.2 | 0.19617 | 0.193 | 0.205 | 0.208 | 0.202 | 0.199 | 0.21 | 0.207 | 0.20417 | 0.015 | 0.2 | 0.2 | 0.16 | 0.201 | 0.18
| Mean |  | 0.205 | 0.19752 |  | 0.2 | 0.19474 | 0.19167 | 0.20583 | 0.21183 | 0.1955 | 0.19767 | 0.2 | 0.201 | 0.20361 | 0.017 | 0.21167 | 0.20667 | 0.15833 | 0.19933 | 0.185
| Median |  | 0.205 | 0.19152 |  | 0.2 | 0.19545 | 0.191 | 0.205 | 0.214 | 0.194 | 0.197 | 0.2 | 0.2025 | 0.20325 | 0.016 | 0.21 | 0.21 | 0.16 | 0.2015 | 0.185
| Std Dev. |  | 0.0054772 | 0.011625 |  | 0.0063246 | 0.0055794 | 0.0052026 | 0.0037639 | 0.0047504 | 0.0045935 | 0.0052026 | 0.0089443 | 0.0072388 | 0.0025056 | 0.0037947 | 0.013292 | 0.005164 | 0.01169 | 0.0047188 | 0.0054772
| Rel.Std.Dev. |  | 2.672% | 5.886% |  | 3.162% | 2.865% | 2.714% | 1.829% | 2.243% | 2.35% | 2.632% | 4.472% | 3.601% | 1.231% | 22.32% | 6.279% | 2.499% | 7.383% | 2.367% | 2.961%
| PDM3 |  | 2.395% | -1.342% |  | -0.1021% | -2.728% | -4.264% | 2.812% | 5.809% | -2.35% | -1.268% | -0.1021% | 0.3974% | 1.701% | -91.51% | 5.725% | 3.228% | -20.91% | -0.435% | -7.594%
| Z-Score (Absolute) |  | 0.64 | 0.32 | 294.97 | 0 | 0.67 | 1.07 | 0.75 | 1.52 | 0.58 | 0.3 | 0 | 0.13 | 0.46 | 23.47 | 1.5 | 0.85 | 5.34 | 0.09 | 1.92
| Analytical results for K in OREAS 922b (Certified Value 0.281 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.377 | 0.2813 | 0.29 | 0.275 | 0.30182 | 0.28 | 0.28 | 0.31 | 0.26 | 0.27 | 0.3 | 0.258 | 0.30335 | 0.28 | 0.26 | 0.27 | 0.28 | 0.28 | 0.315 | 0.4741 | 0.24 | 0.3 |  |  | 1
| 1 | 2 | 0.361 | 0.275 | 0.28 | 0.276 | 0.2942 | 0.26 | 0.27 | 0.31 | 0.25 | 0.27 | 0.3 | 0.254 | 0.32706 | 0.38 | 0.27 | 0.27 | 0.28 | 0.27 | 0.325 | 0.4378 | 0.26 | 0.29 |  |  | 0
| 1 | 3 | 0.363 | 0.2785 | 0.28 | 0.273 | 0.29646 | 0.26 | 0.27 | 0.3 | 0.25 | 0.26 | 0.38 | 0.258 | 0.33908 | 0.29 | 0.25 | 0.28 | 0.29 | 0.27 | 0.33 | 0.4445 | 0.25 | 0.29 |  |  | 16
| 1 | 4 | 0.363 | 0.2813 | 0.28 | 0.276 | 0.29935 | 0.25 | 0.27 | 0.31 | 0.25 | 0.26 | 0.3 | 0.258 | 0.3426 | 0.3 | 0.27 | 0.28 | 0.28 | 0.27 | 0.325 | 0.4635 | 0.25 | 0.29 |  |  | 0.28104
| 1 | 5 | 0.358 | 0.281 | 0.28 | 0.278 | 0.29759 | 0.26 | 0.28 | 0.3 | 0.25 | 0.27 | 0.31 | 0.257 | 0.30041 | 0.27 | 0.24 | 0.28 | 0.28 | 0.28 | 0.33 | 0.4542 | 0.26 | 0.29 |  |  | 89
| 1 | 6 | 0.355 | 0.2769 | 0.29 | 0.279 | 0.28934 | 0.27 | 0.27 | 0.31 | 0.25 | 0.27 | 0.35 | 0.267 | 0.29512 | 0.29 | 0.27 | 0.27 | 0.28 | 0.28 | 0.32 | 0.4482 | 0.26 | 0.29
| Mean |  | 0.36283 | 0.279 | 0.28333 | 0.27617 | 0.29646 | 0.26333 | 0.27333 | 0.30667 | 0.25167 | 0.26667 | 0.32333 | 0.25867 | 0.31794 | 0.30167 | 0.26 | 0.275 | 0.28167 | 0.275 | 0.32417 | 0.45372 | 0.25333 | 0.29167
| Median |  | 0.362 | 0.27975 | 0.28 | 0.276 | 0.29702 | 0.26 | 0.27 | 0.31 | 0.25 | 0.27 | 0.305 | 0.258 | 0.31521 | 0.29 | 0.265 | 0.275 | 0.28 | 0.275 | 0.325 | 0.4512 | 0.255 | 0.29
| Std Dev. |  | 0.0076004 | 0.0026548 | 0.005164 | 0.002137 | 0.0043389 | 0.010328 | 0.005164 | 0.005164 | 0.0040825 | 0.005164 | 0.033862 | 0.0043665 | 0.020878 | 0.039707 | 0.012649 | 0.0054772 | 0.0040825 | 0.0054772 | 0.0058452 | 0.013268 | 0.008165 | 0.0040825
| Rel.Std.Dev. |  | 2.095% | 0.9515% | 1.823% | 0.7738% | 1.464% | 3.922% | 1.889% | 1.684% | 1.622% | 1.936% | 10.47% | 1.688% | 6.567% | 13.16% | 4.865% | 1.992% | 1.449% | 1.992% | 1.803% | 2.924% | 3.223% | 1.4%
| PDM3 |  | 29.1% | -0.7276% | 0.8143% | -1.736% | 5.484% | -6.302% | -2.744% | 9.117% | -10.45% | -5.116% | 15.05% | -7.962% | 13.13% | 7.338% | -7.488% | -2.151% | 0.2213% | -2.151% | 15.34% | 61.44% | -9.86% | 3.779%
| Z-Score (Absolute) |  | 2.98 | 0.05 | 0.11 | 0.15 | 0.58 | 0.61 | 0.25 | 0.95 | 1.04 | 0.49 | 1.55 | 0.78 | 1.36 | 0.77 | 0.73 | 0.19 | 0.05 | 0.19 | 1.58 | 6.26 | 0.98 | 0.41
| Analytical results for La in OREAS 922b (Certified Value 28.5 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 31.5 | 30.5468 | 30.24 | 27.8 | 27.138 | 27.7932 | 28.4 | 29.4 | 30.1 | 27.8 | 29.5 | 34.3 | 28 | 31.4051 | 28.4 | 28.2 | 28 | 26.6 | 26.9 | 27.9 | 35 | 28.9 | 28 |  | 1
| 1 | 2 | 31.2 | 30.4136 | 30.23 | 28 | 26.791 | 28.5124 | 26.1 | 27.4 | 30.2 | 28 | 29.1 | 34.6 | 27.1 | 31.865 | 48.3 | 27.5 | 27 | 28.8 | 27.8 | 29 | 36 | 27.6 | 28 |  | 12
| 1 | 3 | 31.4 | 30.4524 | 30.34 | 26.7 | 27.509 | 28.8687 | 26 | 28.6 | 29.5 | 27.7 | 28.8 | 34.5 | 29.1 | 32.7833 | 29.5 | 27.9 | 27 | 28.2 | 27.8 | 29.3 | 35 | 27.7 | 28 |  | 16
| 1 | 4 | 31.3 | 30.4768 | 30.25 | 28.1 | 27.808 | 27.5375 | 26.2 | 28.2 | 30 | 28.5 | 29 | 34.2 | 27.4 | 32.9507 | 29.8 | 28.6 | 28 | 27.9 | 26.5 | 28.8 | 35 | 27.7 | 28 |  | 28.5128
| 1 | 5 | 31.2 | 30.3992 | 30.52 | 28 | 26.979 | 26.9494 | 26.8 | 29 | 29.4 | 27.2 | 30 | 35.4 | 27.8 | 29.8375 | 27.3 | 28.7 | 28 | 28.3 | 27.1 | 29.5 | 35 | 28.2 | 27 |  | 90
| 1 | 6 | 31.2 | 30.4296 | 30.27 | 27.7 | 26.961 | 27.181 | 27.1 | 29.7 | 29.6 | 29.1 | 30 | 35.4 | 28.8 | 30.4878 | 29 | 27.9 | 28 | 27.1 | 27.4 | 29 | 38 | 28.2 | 27
| Mean |  | 31.3 | 30.4531 | 30.3083 | 27.7167 | 27.1977 | 27.807 | 26.7667 | 28.7167 | 29.8 | 28.05 | 29.4 | 34.7333 | 28.0333 | 31.5549 | 32.05 | 28.1333 | 27.6667 | 27.8167 | 27.25 | 28.9167 | 35.6667 | 28.05 | 27.6667
| Median |  | 31.25 | 30.441 | 30.26 | 27.9 | 27.0585 | 27.6653 | 26.5 | 28.8 | 29.8 | 27.9 | 29.3 | 34.55 | 27.9 | 31.6351 | 29.25 | 28.05 | 28 | 28.05 | 27.25 | 29 | 35 | 27.95 | 28
| Std Dev. |  | 0.12649 | 0.05359 | 0.11089 | 0.51929 | 0.38528 | 0.75165 | 0.90921 | 0.84004 | 0.34059 | 0.66558 | 0.51769 | 0.53541 | 0.78145 | 1.2381 | 8.0102 | 0.45898 | 0.5164 | 0.81833 | 0.51672 | 0.55648 | 1.2111 | 0.49295 | 0.5164
| Rel.Std.Dev. |  | 0.4041% | 0.176% | 0.3659% | 1.874% | 1.417% | 2.703% | 3.397% | 2.925% | 1.143% | 2.373% | 1.761% | 1.541% | 2.788% | 3.924% | 24.99% | 1.631% | 1.866% | 2.942% | 1.896% | 1.924% | 3.395% | 1.757% | 1.866%
| PDM3 |  | 9.775% | 6.805% | 6.297% | -2.792% | -4.612% | -2.475% | -6.124% | 0.715% | 4.514% | -1.623% | 3.112% | 21.82% | -1.682% | 10.67% | 12.41% | -1.331% | -2.968% | -2.441% | -4.429% | 1.416% | 25.09% | -1.623% | -2.968%
| Z-Score (Absolute) |  | 2.42 | 1.77 | 1.66 | 0.32 | 0.71 | 0.25 | 1.04 | 0.45 | 1.27 | 0.06 | 0.97 | 5.04 | 0.08 | 2.61 | 2.99 | 0 | 0.36 | 0.24 | 0.67 | 0.6 | 5.75 | 0.06 | 0.36
| Analytical results for Li in OREAS 922b (Certified Value 3.81 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 4.5 | < 5 | 3.43 | < 10 | 3.6 | 3.6 | 4 | 3.4 | 3.6 | 3.9 | 3.2 | 3.419 | 4.5 | 4 | 4 | 3.5 | 4.5 | 4.1 | [6 empty] | 1
| 1 | 2 | 4.5 | 5 | 3.36 | < 10 | 3.5 | 3.7 | 4 | 3.4 | 3.4 | 4 | 3.3 | 3.7029 | 23.6 | 3.8 | 4 | 3.7 | 4.6 | 3.1 | [6 empty] | 0
| 1 | 3 | 4.6 | 5.1 | 3.44 | < 10 | 3.3 | 3.6 | 4 | 3.4 | 3.4 | 4 | 3.2 | 4.0497 | 4.1 | 3.9 | 4 | 4.2 | 4.5 | 3.8 | [6 empty] | 16
| 1 | 4 | 4.6 | < 5 | 3.46 | < 10 | 3.4 | 3.6 | 4 | 3.3 | 3.4 | 3.7 | 3.4 | 3.957 | 4.7 | 3.8 | 4 | 3.5 | 4.7 | 3.8 | [6 empty] | 3.8097
| 1 | 5 | 4.5 | 5.2 | 3.52 | < 10 | 3.4 | 3.7 | 4 | 3.5 | 3.5 | 4.1 | 3.3 | 3.4804 | 4 | 4 | 4 | 3.7 | 4.8 | 3.4 | [6 empty] | 91
| 1 | 6 | 4.5 | 5.8 | 3.51 | < 10 | 3.5 | 3.7 | 4 | 3.3 | 3.4 | 4.1 | 3.3 | 3.6868 | 4.7 | 3.7 | 4 | 3.9 | 5 | 4.3
| Mean |  | 4.5333 | 5.275 | 3.4533 |  | 3.45 | 3.65 | 4 | 3.3833 | 3.45 | 3.9667 | 3.2833 | 3.716 | 7.6 | 3.8667 | 4 | 3.75 | 4.6833 | 3.75
| Median |  | 4.5 | 5.15 | 3.45 |  | 3.45 | 3.65 | 4 | 3.4 | 3.4 | 4 | 3.3 | 3.6949 | 4.6 | 3.85 | 4 | 3.7 | 4.65 | 3.8
| Std Dev. |  | 0.05164 | 0.3594 | 0.058538 |  | 0.10488 | 0.054772 | 0 | 0.075277 | 0.083666 | 0.15055 | 0.075277 | 0.25066 | 7.844 | 0.12111 | 0 | 0.26646 | 0.19408 | 0.44159
| Rel.Std.Dev. |  | 1.139% | 6.813% | 1.695% |  | 3.04% | 1.501% | 0% | 2.225% | 2.425% | 3.795% | 2.293% | 6.745% | 103.2% | 3.132% | 0% | 7.106% | 4.144% | 11.78%
| PDM3 |  | 18.99% | 38.46% | -9.354% |  | -9.442% | -4.192% | 4.995% | -11.19% | -9.442% | 4.12% | -13.82% | -2.46% | 99.49% | 1.495% | 4.995% | -1.567% | 22.93% | -1.567%
| Z-Score (Absolute) |  | 1.76 | 1.35 | 0.67 | 2.81 | 0.67 | 0.22 | NA | 0.82 | 0.67 | 0.49 | 1.05 | 0.08 | 8.65 | 0.26 | NA | 0 | 2.1 | 0
Results from laboratory 2.09 and 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 922b (Certified Value 0.098 ppm) | [25 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | [6 empty] | 1
| No | No | 2.01: 2.02; 2.04: 2.04; 2.05: 2.06; 2.06: 2.12; 2.07: 2.13; 2.08: 2.21 | [6 empty] | ppm
|  |  | 2.01: AR*MS; 2.04: AR*OES; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*OES/MS | [6 empty] | 3
|  |  | 2.01: 0.5g; 2.04: 0.25g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.25g; 2.08: 0.25g | [6 empty] | 3
| 1 | 1 | 2.01: 0.0902; 2.04: < 5; 2.05: 0.10301; 2.06: 0.1; 2.07: 0.11; 2.08: 0.09 | [6 empty] | 1
| 1 | 2 | 2.01: 0.082; 2.04: < 5; 2.05: 0.11193; 2.06: 0.1; 2.07: 0.1; 2.08: 0.09 | [6 empty] | 13
| 1 | 3 | 2.01: 0.0855; 2.04: < 5; 2.05: 0.11214; 2.06: 0.1; 2.07: 0.1; 2.08: 0.1 | [6 empty] | 16
| 1 | 4 | 2.01: 0.0942; 2.04: < 5; 2.05: 0.10861; 2.06: 0.1; 2.07: 0.1; 2.08: 0.09 | [6 empty] | 0.097565
| 1 | 5 | 2.01: 0.0892; 2.04: < 5; 2.05: 0.090798; 2.06: 0.1; 2.07: 0.1; 2.08: 0.09 | [6 empty] | 92
| 1 | 6 | 2.01: 0.0852; 2.04: < 5; 2.05: 0.10417; 2.06: 0.1; 2.07: 0.11; 2.08: 0.09
| Mean |  | 2.01: 0.087717; 2.05: 0.10511; 2.06: 0.1; 2.07: 0.10333; 2.08: 0.091667
| Median |  | 2.01: 0.08735; 2.05: 0.10639; 2.06: 0.1; 2.07: 0.1; 2.08: 0.09
| Std Dev. |  | 2.01: 0.0043435; 2.05: 0.0079785; 2.06: 1.5202e-17; 2.07: 0.005164; 2.08: 0.0040825
| Rel.Std.Dev. |  | 2.01: 4.952%; 2.05: 7.591%; 2.06: 1.52e-14%; 2.07: 4.997%; 2.08: 4.454%
| PDM3 |  | 2.01: -10.09%; 2.05: 7.733%; 2.06: 2.495%; 2.07: 5.912%; 2.08: -6.046%
| Z-Score (Absolute) |  | 2.01: 1.4; 2.04: 240.59; 2.05: 0.35; 2.06: 0.17; 2.07: 0.17; 2.08: 1
| Analytical results for Mg in OREAS 922b (Certified Value 0.13 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.14 | 0.1252 | 0.13 | 0.121 | 0.13894 | 0.13 | 0.12 | 0.13 | 0.11 | 0.12 | 0.15 | 0.128 | 0.13978 | 0.14 | 0.12 | 0.13 | 0.13 | 0.13 | 0.127 | 0.14 | 0.12 | 0.14 |  |  | 1
| 1 | 2 | 0.14 | 0.1211 | 0.13 | 0.123 | 0.13536 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.14 | 0.129 | 0.14531 | 0.14 | 0.13 | 0.12 | 0.13 | 0.13 | 0.132 | 0.15 | 0.13 | 0.13 |  |  | 0
| 1 | 3 | 0.14 | 0.1246 | 0.13 | 0.121 | 0.13484 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.15 | 0.129 | 0.14895 | 0.13 | 0.12 | 0.12 | 0.13 | 0.13 | 0.134 | 0.14 | 0.12 | 0.13 |  |  | 16
| 1 | 4 | 0.14 | 0.1241 | 0.13 | 0.123 | 0.13799 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.14 | 0.13 | 0.15158 | 0.13 | 0.13 | 0.13 | 0.13 | 0.13 | 0.134 | 0.14 | 0.13 | 0.13 |  |  | 0.12973
| 1 | 5 | 0.14 | 0.1257 | 0.13 | 0.122 | 0.13764 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.16 | 0.13 | 0.14005 | 0.13 | 0.12 | 0.13 | 0.13 | 0.13 | 0.133 | 0.14 | 0.13 | 0.14 |  |  | 93
| 1 | 6 | 0.14 | 0.1224 | 0.13 | 0.124 | 0.13319 | 0.12 | 0.12 | 0.13 | 0.11 | 0.12 | 0.15 | 0.133 | 0.13617 | 0.14 | 0.13 | 0.13 | 0.13 | 0.13 | 0.132 | 0.16 | 0.13 | 0.13
| Mean |  | 0.14 | 0.12385 | 0.13 | 0.12233 | 0.13633 | 0.115 | 0.12 | 0.13 | 0.11 | 0.12 | 0.14833 | 0.12983 | 0.14364 | 0.135 | 0.125 | 0.12667 | 0.13 | 0.13 | 0.132 | 0.145 | 0.12667 | 0.13333
| Median |  | 0.14 | 0.12435 | 0.13 | 0.1225 | 0.1365 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.15 | 0.1295 | 0.14268 | 0.135 | 0.125 | 0.13 | 0.13 | 0.13 | 0.1325 | 0.14 | 0.13 | 0.13
| Std Dev. |  | 0 | 0.0017627 | 0 | 0.0012111 | 0.0022031 | 0.0083666 | 0 | 0 | 0 | 0 | 0.0075277 | 0.0017224 | 0.005962 | 0.0054772 | 0.0054772 | 0.005164 | 0 | 0 | 0.0026077 | 0.0083666 | 0.005164 | 0.005164
| Rel.Std.Dev. |  | 0% | 1.423% | 0% | 0.99% | 1.616% | 7.275% | 0% | 0% | 0% | 0% | 5.075% | 1.327% | 4.151% | 4.057% | 4.382% | 4.077% | 0% | 0% | 1.976% | 5.77% | 4.077% | 3.873%
| PDM3 |  | 7.919% | -4.53% | 0.2109% | -5.699% | 5.087% | -11.35% | -7.498% | 0.2109% | -15.21% | -7.498% | 14.34% | 0.08239% | 10.72% | 4.065% | -3.643% | -2.359% | 0.2109% | 0.2109% | 1.753% | 11.77% | -2.359% | 2.78%
| Z-Score (Absolute) |  | 1.21 | 0.74 | 0 | 0.93 | 0.77 | 1.81 | 1.21 | 0 | 2.42 | 1.21 | 2.22 | 0.02 | 1.65 | 0.6 | 0.6 | 0.4 | 0 | 0 | 0.24 | 1.81 | 0.4 | 0.4
| Analytical results for Mn in OREAS 922b (Certified Value 0.009 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0086 | 0.0085 | 0.00907 | 0.00863 | 0.0089802 | 0.0088 | 0.0091 | 0.0093 | 0.0084 | 0.0087 | 0.0092 | 0.0087 | 0.0089562 | 0.0099 | 0.0089 | 0.0086 | 0.0083 | 0.0087 | 0.0083 | 0.0092 | 0.0085 | 0.0083 |  |  | 1
| 1 | 2 | 0.0085 | 0.0085 | 0.00926 | 0.00871 | 0.0090282 | 0.0084 | 0.009 | 0.0092 | 0.0083 | 0.0083 | 0.0091 | 0.0087 | 0.0090078 | 0.0783 | 0.0092 | 0.0085 | 0.0087 | 0.0088 | 0.0083 | 0.0094 | 0.0085 | 0.008 |  |  | 30
| 1 | 3 | 0.0085 | 0.0084 | 0.00899 | 0.00871 | 0.0090488 | 0.0082 | 0.0089 | 0.0091 | 0.0083 | 0.0083 | 0.009 | 0.0088 | 0.010086 | 0.0103 | 0.0091 | 0.0086 | 0.0087 | 0.0087 | 0.0087 | 0.0092 | 0.0084 | 0.008 |  |  | 16
| 1 | 4 | 0.0086 | 0.0085 | 0.00937 | 0.0086 | 0.0092539 | 0.0083 | 0.0088 | 0.0092 | 0.0083 | 0.0084 | 0.0084 | 0.0088 | 0.010046 | 0.0116 | 0.0089 | 0.0087 | 0.0085 | 0.0088 | 0.0085 | 0.0093 | 0.0085 | 0.0078 |  |  | 0.0087381
| 1 | 5 | 0.0085 | 0.0085 | 0.00902 | 0.00856 | 0.0088347 | 0.0082 | 0.009 | 0.0091 | 0.0084 | 0.0085 | 0.0091 | 0.0088 | 0.0093377 | 0.0093 | 0.0095 | 0.0086 | 0.0087 | 0.0088 | 0.0085 | 0.0092 | 0.0086 | 0.0077 |  |  | 94
| 1 | 6 | 0.0085 | 0.0085 | 0.0093 | 0.00877 | 0.0084713 | 0.0084 | 0.0091 | 0.0092 | 0.0083 | 0.0085 | 0.0092 | 0.009 | 0.0094024 | 0.0107 | 0.0092 | 0.0085 | 0.0081 | 0.0088 | 0.0085 | 0.0098 | 0.0087 | 0.0079
| Mean |  | 0.0085333 | 0.0084833 | 0.0091683 | 0.0086633 | 0.0089362 | 0.0083833 | 0.0089833 | 0.0091833 | 0.0083333 | 0.00845 | 0.009 | 0.0088 | 0.0094727 | 0.021683 | 0.0091333 | 0.0085833 | 0.0085 | 0.0087667 | 0.0084667 | 0.00935 | 0.0085333 | 0.00795
| Median |  | 0.0085 | 0.0085 | 0.009165 | 0.00867 | 0.0090042 | 0.00835 | 0.009 | 0.0092 | 0.0083 | 0.00845 | 0.0091 | 0.0088 | 0.00937 | 0.0105 | 0.00915 | 0.0086 | 0.0086 | 0.0088 | 0.0085 | 0.00925 | 0.0085 | 0.00795
| Std Dev. |  | 5.164e-05 | 4.0825e-05 | 0.00016117 | 7.9415e-05 | 0.00026474 | 0.00022286 | 0.0001169 | 7.5277e-05 | 5.164e-05 | 0.00015166 | 0.00030332 | 0.00010954 | 0.00049223 | 0.027747 | 0.00022509 | 7.5277e-05 | 0.00025298 | 5.164e-05 | 0.00015055 | 0.00023452 | 0.00010328 | 0.00020736
| Rel.Std.Dev. |  | 0.6052% | 0.4812% | 1.758% | 0.9167% | 2.963% | 2.658% | 1.301% | 0.8197% | 0.6197% | 1.795% | 3.37% | 1.245% | 5.196% | 128% | 2.465% | 0.877% | 2.976% | 0.589% | 1.778% | 2.508% | 1.21% | 2.608%
| PDM3 |  | -2.343% | -2.916% | 4.924% | -0.8557% | 2.267% | -4.06% | 2.806% | 5.095% | -4.632% | -3.297% | 2.997% | 0.7084% | 8.407% | 148.1% | 4.523% | -1.771% | -2.725% | 0.3269% | -3.106% | 7.003% | -2.343% | -9.019%
| Z-Score (Absolute) |  | 0.46 | 0.59 | 1.15 | 0.13 | 0.56 | 0.84 | 0.68 | 1.18 | 0.97 | 0.67 | 0.72 | 0.21 | 1.92 | 32.79 | 1.06 | 0.33 | 0.54 | 0.13 | 0.63 | 1.61 | 0.46 | 1.93
| Analytical results for Mo in OREAS 922b (Certified Value 4.73 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 4.8 | 4.65 | 5 | 4.6 | 4.7159 | 4.39 | 4.88 | 4.75 | 4.63 | 4.64 | 4.98 | 4.87 | 4.1333 | 5.24 | 3.82 | 3.5 | 4.6 | 4.27 | 4.96 | 4.41 | 5 | 4.55 | 4.5 |  | 1
| 1 | 2 | 4.7 | 4.7 | 5.5 | 4.65 | 4.7697 | 4.62 | 4.92 | 4.74 | 4.45 | 4.53 | 5.14 | 4.75 | 4.4427 | 1.05 | 4.02 | 3.8 | 4.7 | 4.79 | 5.24 | 4.5 | 5 | 4.89 | 4.4 |  | 31
| 1 | 3 | 4.8 | 4.66 | 4.9 | 4.67 | 4.7623 | 4.59 | 4.93 | 4.78 | 4.37 | 4.54 | 4.89 | 4.79 | 4.077 | 5.3 | 3.98 | 3.8 | 4.8 | 4.08 | 5.24 | 4.41 | 5 | 4.65 | 4.4 |  | 16
| 1 | 4 | 4.8 | 4.65 | 5.5 | 4.59 | 4.7904 | 4.64 | 5.01 | 4.66 | 4.5 | 4.51 | 4.74 | 4.7 | 4.5197 | 5.59 | 3.98 | 3.8 | 4.7 | 4.16 | 5.11 | 4.58 | 5 | 4.79 | 4.3 |  | 4.7338
| 1 | 5 | 4.7 | 4.67 | 5.1 | 4.78 | 4.4528 | 4.66 | 4.93 | 4.74 | 4.42 | 4.71 | 5.33 | 4.82 | 4.4447 | 4.78 | 3.95 | 4.2 | 4.7 | 4.85 | 5.03 | 4.82 | 5 | 4.77 | 4.3 |  | 95
| 1 | 6 | 4.7 | 4.67 | 5.5 | 4.69 | 4.5983 | 4.63 | 4.9 | 4.58 | 4.53 | 4.74 | 5.4 | 4.96 | 4.2465 | 5.57 | 3.95 | 3.8 | 4.6 | 4.28 | 4.92 | 4.5 | 5 | 4.78 | 4.3
| Mean |  | 4.75 | 4.6667 | 5.25 | 4.6633 | 4.6816 | 4.5883 | 4.9283 | 4.7083 | 4.4833 | 4.6117 | 5.08 | 4.815 | 4.3106 | 4.5883 | 3.95 | 3.8167 | 4.6833 | 4.405 | 5.0833 | 4.5367 | 5 | 4.7383 | 4.3667
| Median |  | 4.75 | 4.665 | 5.3 | 4.66 | 4.7391 | 4.625 | 4.925 | 4.74 | 4.475 | 4.59 | 5.06 | 4.805 | 4.3446 | 5.27 | 3.965 | 3.8 | 4.7 | 4.275 | 5.07 | 4.5 | 5 | 4.775 | 4.35
| Std Dev. |  | 0.054772 | 0.018619 | 0.28107 | 0.069186 | 0.13158 | 0.099883 | 0.04446 | 0.074409 | 0.091579 | 0.099079 | 0.25698 | 0.091815 | 0.18401 | 1.7581 | 0.068702 | 0.22286 | 0.075277 | 0.33038 | 0.1375 | 0.15293 | 0 | 0.11974 | 0.08165
| Rel.Std.Dev. |  | 1.153% | 0.399% | 5.354% | 1.484% | 2.811% | 2.177% | 0.9021% | 1.58% | 2.043% | 2.148% | 5.059% | 1.907% | 4.269% | 38.32% | 1.739% | 5.839% | 1.607% | 7.5% | 2.705% | 3.371% | 0% | 2.527% | 1.87%
| PDM3 |  | 0.343% | -1.417% | 10.91% | -1.488% | -1.103% | -3.072% | 4.11% | -0.5372% | -5.29% | -2.579% | 7.314% | 1.716% | -8.939% | -3.072% | -16.56% | -19.37% | -1.065% | -6.945% | 7.385% | -4.164% | 5.624% | 0.09652% | -7.755%
| Z-Score (Absolute) |  | 0.41 | 0.01 | 2.83 | 0.01 | 0.08 | 0.37 | 1.28 | 0.21 | 0.88 | 0.26 | 2.01 | 0.73 | 1.72 | 0.37 | 3.46 | 4.11 | 0.09 | 1.26 | 2.03 | 0.62 | NA | 0.36 | 1.45
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 922b (Certified Value 0.01 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.01 | 0.01 | 0.008 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.016 | 0.016 | 0.010802 | < 0.01 | 0.007 | 0.008 | 0.01 | 0.02 | 0.013 | < 0.01 | < 0.01 |  |  |  |  | 1
| 1 | 2 | 0.01 | 0.01 | 0.007 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.014 | 0.015 | 0.011247 | 0.01 | 0.007 | 0.007 | 0.01 | 0.02 | 0.013 | < 0.01 | < 0.01 |  |  |  |  | 21
| 1 | 3 | 0.01 | 0.01 | 0.007 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.015 | 0.016 | 0.012112 | < 0.01 | 0.007 | 0.007 | 0.01 | 0.02 | 0.013 | < 0.01 | 0.01 |  |  |  |  | 16
| 1 | 4 | 0.01 | 0.01 | 0.007 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.015 | 0.016 | 0.01228 | < 0.01 | 0.007 | 0.008 | 0.01 | 0.02 | 0.013 | < 0.01 | 0.01 |  |  |  |  | 0.009753
| 1 | 5 | 0.01 | 0.01 | 0.007 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.016 | 0.016 | 0.011533 | < 0.01 | 0.007 | 0.008 | 0.01 | 0.02 | 0.013 | < 0.01 | 0.01 |  |  |  |  | 96
| 1 | 6 | 0.01 | 0.01 | 0.008 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.015 | 0.018 | 0.011281 | < 0.01 | 0.007 | 0.008 | 0.01 | 0.02 | 0.013 | 0.01 | 0.01
| Mean |  | 0.01 | 0.01 | 0.0073333 |  |  | 0.01 | 0.01 | 0.03 | 0.01 | 0.015167 | 0.016167 | 0.011542 | 0.01 | 0.007 | 0.0076667 | 0.01 | 0.02 | 0.013 | 0.01 | 0.01
| Median |  | 0.01 | 0.01 | 0.007 |  |  | 0.01 | 0.01 | 0.03 | 0.01 | 0.015 | 0.016 | 0.011407 | 0.01 | 0.007 | 0.008 | 0.01 | 0.02 | 0.013 | 0.01 | 0.01
| Std Dev. |  | 0 | 0 | 0.0005164 |  |  | 0 | 0 | 0 | 0 | 0.00075277 | 0.00098319 | 0.00056074 |  | 9.5015e-19 | 0.0005164 | 0 | 0 | 0 |  | 0
| Rel.Std.Dev. |  | 0% | 0% | 7.042% |  |  | 0% | 0% | 0% | 0% | 4.963% | 6.082% | 4.858% |  | 1.357e-14% | 6.736% | 0% | 0% | 0% |  | 0%
| PDM3 |  | 2.532% | 2.532% | -24.81% |  |  | 2.532% | 2.532% | 207.6% | 2.532% | 55.51% | 65.76% | 18.35% | 2.532% | -28.23% | -21.39% | 2.532% | 105.1% | 33.29% | 2.532% | 2.532%
| Z-Score (Absolute) |  | 0 | 0 | 0.63 | 1.19 | 1.19 | 0 | 0 | 4.76 | 0 | 1.23 | 1.47 | 0.37 | 0.99 | 0.71 | 0.56 | 0 | 2.38 | 0.71 | 0.99 | 0.4
| Analytical results for Nb in OREAS 922b (Certified Value 0.069 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [9 empty] | 3
| 1 | 1 | < 0.2 | < 5 | 0.06 | 0.08 | 0.06 | 0.07 | 0.08 | 0.1 | 0.12 | < 0.2 | < 0.05 | 0.27 | 0.14 | 0.07 | 0.1 | [9 empty] | 1
| 1 | 2 | < 0.2 | < 5 | 0.07 | 0.08 | 0.07 | 0.06 | 0.09 | 0.1 | 0.12 | < 0.2 | 0.11 | 0.3 | 0.16 | 0.06 | 0.1 | [9 empty] | 15
| 1 | 3 | < 0.2 | < 5 | 0.06 | 0.08 | 0.06 | 0.06 | 0.07 | 0.1 | 0.13 | < 0.2 | 0.05 | 0.29 | 0.1 | 0.06 | 0.1 | [9 empty] | 16
| 1 | 4 | < 0.2 | < 5 | 0.07 | 0.09 | 0.07 | 0.07 | 0.07 | 0.1 | 0.12 | < 0.2 | 0.06 | 0.29 | 0.13 | 0.07 | 0.1 | [9 empty] | 0.06881
| 1 | 5 | < 0.2 | < 5 | 0.08 | 0.08 | 0.07 | 0.06 | 0.08 | 0.1 | 0.11 | < 0.2 | < 0.05 | 0.3 | 0.13 | 0.07 | 0.1 | [9 empty] | 97
| 1 | 6 | < 0.2 | < 5 | 0.06 | 0.08 | 0.06 | 0.07 | 0.08 | 0.1 | 0.13 | < 0.2 | 0.06 | 0.3 | 0.15 | 0.08 | 0.1
| Mean |  |  |  | 0.066667 | 0.081667 | 0.065 | 0.065 | 0.078333 | 0.1 | 0.12167 |  | 0.07 | 0.29167 | 0.135 | 0.068333 | 0.1
| Median |  |  |  | 0.065 | 0.08 | 0.065 | 0.065 | 0.08 | 0.1 | 0.12 |  | 0.06 | 0.295 | 0.135 | 0.07 | 0.1
| Std Dev. |  |  |  | 0.008165 | 0.0040825 | 0.0054772 | 0.0054772 | 0.0075277 | 1.5202e-17 | 0.0075277 |  | 0.02708 | 0.01169 | 0.020736 | 0.0075277 | 1.5202e-17
| Rel.Std.Dev. |  |  |  | 12.25% | 4.999% | 8.427% | 8.427% | 9.61% | 1.52e-14% | 6.187% |  | 38.69% | 4.008% | 15.36% | 11.02% | 1.52e-14%
| PDM3 |  |  |  | -3.114% | 18.69% | -5.536% | -5.536% | 13.84% | 45.33% | 76.82% |  | 1.73% | 323.9% | 96.19% | -0.692% | 45.33%
| Z-Score (Absolute) |  | 0.67 | 88.95 | 0.55 | 0 | 0.61 | 0.61 | 0.12 | NA | 1.47 | 0.67 | 0.98 | 7.72 | 1.96 | 0.49 | NA
Results from laboratories 2.12 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 922b (Indicative Value 25.7 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.07: Lab | [9 empty] | 1
| No | No | 2.01: 2.02; 2.04: 2.06; 2.07: 2.12 | [9 empty] | ppm
|  |  | 2.01: AR*MS; 2.04: AR*MS; 2.07: AR*OES/MS | [9 empty] | 1
|  |  | 2.01: 0.5g; 2.04: 0.2g; 2.07: 0.5g | [9 empty] | 1
| 1 | 1 | 2.01: 23.168; 2.04: 24.8697; 2.07: 26.9 | [9 empty] | 1
| 1 | 2 | 2.01: 23.341; 2.04: 26.5871; 2.07: 28 | [9 empty] | 4
| 1 | 3 | 2.01: 23.353; 2.04: 26.1638; 2.07: 28.2 | [9 empty] | 16
| 1 | 4 | 2.01: 23.363; 2.04: 24.8783; 2.07: 27.8 | [9 empty] | 25.6786
| 1 | 5 | 2.01: 23.291; 2.04: 24.8533; 2.07: 29.2 | [9 empty] | 10
| 1 | 6 | 2.01: 23.414; 2.04: 25.4319; 2.07: 29.4
| Mean |  | 2.01: 23.3217; 2.04: 25.464; 2.07: 28.25
| Median |  | 2.01: 23.347; 2.04: 25.1551; 2.07: 28.1
| Std Dev. |  | 2.01: 0.085022; 2.04: 0.75119; 2.07: 0.92898
| Rel.Std.Dev. |  | 2.01: 0.3646%; 2.04: 2.95%; 2.07: 3.288%
| PDM3 |  | 2.01: -9.178%; 2.04: -0.8355%; 2.07: 10.01%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.04: 0; 2.07: 0.88
| Analytical results for Ni in OREAS 922b (Certified Value 9.31 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 10 | 7.26 | 9.4 | 8.62 | 9.7 | 9.2012 | 9.3 | 9 | 9.9 | 11.4 | 9.3 | 9.6 | 10 | 11.3486 | 12.3 | 9.3 | 10 | 9 | 9.4 | 9.2 | 10 | 9.1 | 9.6 |  | 1
| 1 | 2 | 10 | 7.33 | 9.3 | 8.54 | 9.2 | 9.8707 | 9.4 | 9.3 | 9.3 | 8.4 | 9.5 | 12.2 | 8.1 | 11.3133 | 36.9 | 9.3 | 9.9 | 9.3 | 9.9 | 9.3 | 10 | 9 | 9.1 |  | 22
| 1 | 3 | 10 | 7.33 | 9.2 | 8.09 | 9.3 | 9.1762 | 8.2 | 9.5 | 9.3 | 8.8 | 9.2 | 10.3 | 8.6 | 11.7395 | 12.7 | 9.6 | 9.2 | 9.7 | 14.3 | 9.2 | 10 | 9 | 9.6 |  | 16
| 1 | 4 | 10 | 7.36 | 9.4 | 8.88 | 9.5 | 9.64 | 8.6 | 9.6 | 9.3 | 8.6 | 9.2 | 9.2 | 8.9 | 10.8761 | 14.8 | 9.5 | 10.1 | 9.2 | 9.9 | 9.2 | 10 | 9.3 | 8.8 |  | 9.3093
| 1 | 5 | 10 | 7.39 | 9.6 | 8.46 | 10 | 8.7422 | 8.7 | 9.4 | 9.7 | 9.1 | 9.9 | 10 | 9.1 | 10.5001 | 11.2 | 9.9 | 8.9 | 10.2 | 9.5 | 9.3 | 10 | 9.4 | 8.7 |  | 98
| 1 | 6 | 10 | 7.44 | 9.3 | 9.68 | 10.2 | 8.8428 | 8.4 | 9.3 | 9.1 | 8.6 | 10.4 | 10.8 | 9.1 | 10.7469 | 12.5 | 9.7 | 9.3 | 9 | 9.9 | 9.3 | 11 | 9.4 | 9.4
| Mean |  | 10 | 7.3517 | 9.3667 | 8.7117 | 9.65 | 9.2455 | 8.7667 | 9.35 | 9.4333 | 9.15 | 9.5833 | 10.35 | 8.9667 | 11.0874 | 16.7333 | 9.55 | 9.5667 | 9.4 | 10.4833 | 9.25 | 10.1667 | 9.2 | 9.2
| Median |  | 10 | 7.345 | 9.35 | 8.58 | 9.6 | 9.1887 | 8.65 | 9.35 | 9.3 | 8.7 | 9.4 | 10.15 | 9 | 11.0947 | 12.6 | 9.55 | 9.6 | 9.25 | 9.9 | 9.25 | 10 | 9.2 | 9.25
| Std Dev. |  | 0 | 0.061128 | 0.13663 | 0.5392 | 0.3937 | 0.44018 | 0.48442 | 0.20736 | 0.30111 | 1.1274 | 0.47924 | 1.0616 | 0.6314 | 0.45825 | 9.9488 | 0.23452 | 0.49666 | 0.46904 | 1.883 | 0.054772 | 0.40825 | 0.18974 | 0.39497
| Rel.Std.Dev. |  | 0% | 0.8315% | 1.459% | 6.189% | 4.08% | 4.761% | 5.526% | 2.218% | 3.192% | 12.32% | 5.001% | 10.26% | 7.042% | 4.133% | 59.45% | 2.456% | 5.192% | 4.99% | 17.96% | 0.5921% | 4.016% | 2.062% | 4.293%
| PDM3 |  | 7.419% | -21.03% | 0.6162% | -6.42% | 3.66% | -0.6854% | -5.829% | 0.4372% | 1.332% | -1.711% | 2.944% | 11.18% | -3.681% | 19.1% | 79.75% | 2.586% | 2.765% | 0.9743% | 12.61% | -0.637% | 9.21% | -1.174% | -1.174%
| Z-Score (Absolute) |  | NA | 6.27 | 0 | 2.04 | 0.88 | 0.38 | 1.87 | 0.05 | 0.21 | 0.67 | 0.67 | 3.06 | 1.24 | 5.36 | 22.93 | 0.57 | 0.62 | 0.1 | 3.48 | 0.36 | NA | 0.52 | 0.52
Results from laboratories 2.01 and 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 922b (Certified Value 0.029 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.029 | 0.0259 | 0.03 | 0.0269 | 0.031028 | 0.029 | 0.028 | 0.03 | 0.026 | 0.029 | 0.032 | 0.029 | 0.030608 | 0.026 | 0.028 | 0.027 | 0.0287 | 0.03 | 0.0287 | 0.0293 | 0.0277 | 0.034 |  |  | 1
| 1 | 2 | 0.0289 | 0.0251 | 0.03 | 0.0275 | 0.030086 | 0.028 | 0.028 | 0.031 | 0.026 | 0.027 | 0.032 | 0.029 | 0.031491 | 0.068 | 0.029 | 0.026 | 0.029 | 0.03 | 0.0301 | 0.0306 | 0.0281 | 0.031 |  |  | 0
| 1 | 3 | 0.0291 | 0.0254 | 0.03 | 0.0273 | 0.030344 | 0.027 | 0.028 | 0.03 | 0.026 | 0.027 | 0.032 | 0.03 | 0.032175 | 0.027 | 0.028 | 0.026 | 0.0285 | 0.03 | 0.0305 | 0.03 | 0.0286 | 0.031 |  |  | 16
| 1 | 4 | 0.029 | 0.0255 | 0.03 | 0.0278 | 0.03185 | 0.027 | 0.027 | 0.03 | 0.026 | 0.027 | 0.031 | 0.029 | 0.03199 | 0.027 | 0.029 | 0.027 | 0.029 | 0.03 | 0.0301 | 0.029 | 0.0284 | 0.03 |  |  | 0.02879
| 1 | 5 | 0.0291 | 0.0254 | 0.03 | 0.0266 | 0.02896 | 0.027 | 0.029 | 0.03 | 0.027 | 0.028 | 0.032 | 0.029 | 0.029739 | 0.025 | 0.028 | 0.027 | 0.0287 | 0.03 | 0.0304 | 0.0291 | 0.0289 | 0.031 |  |  | 99
| 1 | 6 | 0.0292 | 0.0251 | 0.03 | 0.028 | 0.029009 | 0.028 | 0.028 | 0.03 | 0.026 | 0.028 | 0.032 | 0.029 | 0.029861 | 0.028 | 0.029 | 0.027 | 0.0288 | 0.03 | 0.0299 | 0.0321 | 0.0284 | 0.03
| Mean |  | 0.02905 | 0.0254 | 0.03 | 0.02735 | 0.030213 | 0.027667 | 0.028 | 0.030167 | 0.026167 | 0.027667 | 0.031833 | 0.029167 | 0.030977 | 0.0335 | 0.0285 | 0.026667 | 0.028783 | 0.03 | 0.02995 | 0.030017 | 0.02835 | 0.031167
| Median |  | 0.02905 | 0.0254 | 0.03 | 0.0274 | 0.030215 | 0.0275 | 0.028 | 0.03 | 0.026 | 0.0275 | 0.032 | 0.029 | 0.03105 | 0.027 | 0.0285 | 0.027 | 0.02875 | 0.03 | 0.0301 | 0.02965 | 0.0284 | 0.031
| Std Dev. |  | 0.00010488 | 0.00029665 | 0 | 0.00053198 | 0.0011311 | 0.0008165 | 0.00063246 | 0.00040825 | 0.00040825 | 0.0008165 | 0.00040825 | 0.00040825 | 0.001062 | 0.016932 | 0.00054772 | 0.0005164 | 0.00019408 | 0 | 0.00065038 | 0.001189 | 0.00041352 | 0.001472
| Rel.Std.Dev. |  | 0.361% | 1.168% | 0% | 1.945% | 3.744% | 2.951% | 2.259% | 1.353% | 1.56% | 2.951% | 1.282% | 1.4% | 3.428% | 50.54% | 1.922% | 1.936% | 0.6743% | 0% | 2.172% | 3.961% | 1.459% | 4.723%
| PDM3 |  | 0.9042% | -11.77% | 4.204% | -5.001% | 4.943% | -3.901% | -2.743% | 4.783% | -9.111% | -3.901% | 10.57% | 1.309% | 7.598% | 16.36% | -1.006% | -7.374% | -0.0221% | 4.204% | 4.03% | 4.262% | -1.527% | 8.256%
| Z-Score (Absolute) |  | 0.04 | 2.26 | 0.54 | 1.07 | 0.67 | 0.88 | 0.68 | 0.65 | 1.79 | 0.88 | 1.66 | 0.04 | 1.14 | 2.68 | 0.37 | 1.49 | 0.2 | 0.54 | 0.51 | 0.55 | 0.46 | 1.25
| Analytical results for Pb in OREAS 922b (Certified Value 29.6 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 1
| 1 | 1 | 34.2 | 22.7 | 26.6 | 60 | 31.1 | 33.0841 | 32.2 | 29.1 | 31.4 | 27.9 | 29.3 | 36.8 | 28.1 | 23.4433 | 50 | 27.3 | 34 | 28.2 | 27.5 | 33.4 | 32.8 | 37 | 26.8 | 28 | 1
| 1 | 2 | 33.5 | 20.61 | 26.8 | 56.6 | 25.5 | 32.1269 | 29.3 | 28.6 | 31.8 | 28.3 | 28.9 | 36.3 | 29.2 | 23.405 | 47 | 26.47 | 33 | 27.4 | 26 | 33.3 | 33.7 | 36 | 27.7 | 28.3 | 17
| 1 | 3 | 33.7 | 23.41 | 26.6 | 61.2 | 26 | 31.4987 | 28.3 | 28.3 | 32 | 27 | 29 | 37.3 | 28.3 | 24.2585 | 49 | 26.3 | 34 | 27.6 | 26.5 | 32.5 | 33.4 | 37 | 26.4 | 27.9 | 16
| 1 | 4 | 33.7 | 23.57 | 26.5 | 59 | 25.8 | 31.7865 | 30.3 | 28.6 | 31.5 | 27.6 | 28.9 | 36.9 | 28.3 | 27.3625 | 46 | 26.78 | 35 | 27.8 | 26.4 | 33.3 | 33 | 37 | 27.4 | 28 | 29.6025
| 1 | 5 | 33.5 | 25.26 | 26.4 | 59.1 | 27.2 | 31.6498 | 28.5 | 29.1 | 31.8 | 27.7 | 28.9 | 37.7 | 28.2 | 23.2045 | 45 | 26.81 | 38 | 27.9 | 25.7 | 33.3 | 32.9 | 37 | 27.6 | 28.6 | 100
| 1 | 6 | 33.5 | 23.4 | 26.7 | 58.1 | 25.9 | 31.4349 | 28.6 | 29.4 | 30.7 | 27.4 | 29.1 | 38.3 | 29.1 | 27.3965 | 48 | 26.11 | 33 | 27.3 | 25.3 | 33.6 | 32 | 36 | 28.2 | 29
| Mean |  | 33.6833 | 23.1583 | 26.6 | 59 | 26.9167 | 31.9302 | 29.5333 | 28.85 | 31.5333 | 27.65 | 29.0167 | 37.2167 | 28.5333 | 24.8451 | 47.5 | 26.6283 | 34.5 | 27.7 | 26.2333 | 33.2333 | 32.9667 | 36.6667 | 27.35 | 28.3
| Median |  | 33.6 | 23.405 | 26.6 | 59.05 | 25.95 | 31.7181 | 28.95 | 28.85 | 31.65 | 27.65 | 28.95 | 37.1 | 28.3 | 23.8509 | 47.5 | 26.625 | 34 | 27.7 | 26.2 | 33.3 | 32.95 | 37 | 27.5 | 28.15
| Std Dev. |  | 0.27142 | 1.5108 | 0.14142 | 1.5761 | 2.1311 | 0.61659 | 1.4976 | 0.41352 | 0.46332 | 0.44159 | 0.16021 | 0.7111 | 0.48442 | 1.9961 | 1.8708 | 0.42631 | 1.8708 | 0.33466 | 0.76333 | 0.37771 | 0.58195 | 0.5164 | 0.65038 | 0.42895
| Rel.Std.Dev. |  | 0.8058% | 6.524% | 0.5317% | 2.671% | 7.917% | 1.931% | 5.071% | 1.433% | 1.469% | 1.597% | 0.5521% | 1.911% | 1.698% | 8.034% | 3.939% | 1.601% | 5.423% | 1.208% | 2.91% | 1.137% | 1.765% | 1.408% | 2.378% | 1.516%
| PDM3 |  | 13.79% | -21.77% | -10.14% | 99.31% | -9.073% | 7.863% | -0.2335% | -2.542% | 6.523% | -6.596% | -1.979% | 25.72% | -3.612% | -16.07% | 60.46% | -10.05% | 16.54% | -6.427% | -11.38% | 12.27% | 11.36% | 23.86% | -7.609% | -4.4%
| Z-Score (Absolute) |  | 1.21 | 1.47 | 0.59 | 7.65 | 0.51 | 0.76 | 0.15 | 0.02 | 0.66 | 0.33 | 0.02 | 2.11 | 0.1 | 1.04 | 4.72 | 0.59 | 1.42 | 0.31 | 0.69 | 1.09 | 1.03 | 1.97 | 0.4 | 0.16
| Analytical results for Pd in OREAS 922b (Indicative Value 3.33 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab; 2.02: Lab | 1
| No | No | 2.01: 2.05; 2.02: 2.14 | ppb
|  |  | 2.01: AR*MS; 2.02: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 01g | 2
| 1 | 1 | 2.01: 3; 2.02: < 10 | 1
| 1 | 2 | 2.01: 3; 2.02: < 10 | 5
| 1 | 3 | 2.01: 3; 2.02: < 10 | 16
| 1 | 4 | 2.01: 4; 2.02: < 10 | 3.3333
| 1 | 5 | 2.01: 3; 2.02: < 10 | 11
| 1 | 6 | 2.01: 4; 2.02: < 10
| Mean |  | 2.01: 3.3333
| Median |  | 2.01: 3
| Std Dev. |  | 2.01: 0.5164
| Rel.Std.Dev. |  | 2.01: 15.49%
| PDM3 |  | 2.01: 1.11e-13%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.67
| Analytical results for Pr in OREAS 922b (Indicative Value 6.91 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 6.2009; 2.02: 6.3207; 2.03: 7.8 | 1
| 1 | 2 | 2.01: 6.3264; 2.02: 6.9306; 2.03: 7.8 | 6
| 1 | 3 | 2.01: 6.3364; 2.02: 7.0264; 2.03: 7.9 | 16
| 1 | 4 | 2.01: 6.3212; 2.02: 6.4152; 2.03: 7.7 | 6.9059
| 1 | 5 | 2.01: 6.3514; 2.02: 6.06; 2.03: 8 | 12
| 1 | 6 | 2.01: 6.2438; 2.02: 6.3737; 2.03: 8.2
| Mean |  | 2.01: 6.2967; 2.02: 6.5211; 2.03: 7.9
| Median |  | 2.01: 6.3238; 2.02: 6.3944; 2.03: 7.85
| Std Dev. |  | 2.01: 0.060044; 2.02: 0.3765; 2.03: 0.17889
| Rel.Std.Dev. |  | 2.01: 0.9536%; 2.02: 5.774%; 2.03: 2.264%
| PDM3 |  | 2.01: -8.822%; 2.02: -5.572%; 2.03: 14.39%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0; 2.03: 4.14
| Analytical results for Pt in OREAS 922b (Indicative Value 1.8 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.02: Lab | 1
| No | No | 2.01: 2.05; 2.02: 2.14 | ppb
|  |  | 2.01: AR*MS; 2.02: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 01g | 2
| 1 | 1 | 2.01: 2; 2.02: < 5 | 1
| 1 | 2 | 2.02: < 5 | 3
| 1 | 3 | 2.01: 3; 2.02: < 5 | 16
| 1 | 4 | 2.01: 2; 2.02: < 5 | 1.8
| 1 | 5 | 2.01: 1; 2.02: < 5 | 9
| 1 | 6 | 2.01: 1; 2.02: < 5
| Mean |  | 2.01: 1.8
| Median |  | 2.01: 2
| Std Dev. |  | 2.01: 0.83666
| Rel.Std.Dev. |  | 2.01: 46.48%
PDM3 | bar: 2.01
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.67
| Analytical results for Rb in OREAS 922b (Certified Value 16.7 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 25.8 | 23.8 | 15.37 | 17.0878 | 16.6 | 16.2 | 18.4 | 15.8 | 16.5 | 20.3 | 15.47 | 18.3161 | 14.9 | 14.7 | 16.7 | 16.9 | 14.9 | 18.4 | [6 empty] | 1
| 1 | 2 | 25 | 25.3 | 15.38 | 17.1077 | 15.2 | 15.7 | 18.8 | 16 | 16.5 | 19.6 | 14.91 | 19.4656 | 18.8 | 16.2 | 17.6 | 17.2 | 16 | 17.6 | [6 empty] | 37
| 1 | 3 | 25.5 | 22.6 | 15.45 | 17.6631 | 14.7 | 16.1 | 18.2 | 15.6 | 16.2 | 21.3 | 15.56 | 20.1445 | 15 | 15.8 | 17.2 | 17.8 | 15.8 | 17.9 | [6 empty] | 16
| 1 | 4 | 24.9 | 22.1 | 15.21 | 17.4338 | 15 | 15.9 | 18.4 | 16.2 | 16.6 | 19.7 | 15.07 | 20.3335 | 15.4 | 15.3 | 17 | 17.4 | 15.6 | 17.2 | [6 empty] | 16.7091
| 1 | 5 | 24.5 | 21.8 | 15.26 | 16.8148 | 15.4 | 16.3 | 18.4 | 15.6 | 17.1 | 21.7 | 14.84 | 17.6355 | 13.9 | 15.6 | 17.4 | 17.5 | 15.6 | 16.9 | [6 empty] | 101
| 1 | 6 | 24.8 | 23.6 | 15.53 | 16.8864 | 15 | 16.6 | 17.7 | 16 | 17.2 | 21.2 | 15.75 | 17.846 | 14.9 | 14.9 | 17.5 | 17.3 | 16 | 17.1
| Mean |  | 25.0833 | 23.2 | 15.3667 | 17.1656 | 15.3167 | 16.1333 | 18.3167 | 15.8667 | 16.6833 | 20.6333 | 15.2667 | 18.9569 | 15.4833 | 15.4167 | 17.2333 | 17.35 | 15.65 | 17.5167
| Median |  | 24.95 | 23.1 | 15.375 | 17.0977 | 15.1 | 16.15 | 18.4 | 15.9 | 16.55 | 20.75 | 15.27 | 18.8909 | 14.95 | 15.45 | 17.3 | 17.35 | 15.7 | 17.4
| Std Dev. |  | 0.47924 | 1.3008 | 0.1181 | 0.32553 | 0.67057 | 0.31411 | 0.36009 | 0.24221 | 0.38687 | 0.88919 | 0.37655 | 1.1794 | 1.6987 | 0.56362 | 0.33862 | 0.30166 | 0.40866 | 0.56362
| Rel.Std.Dev. |  | 1.911% | 5.607% | 0.7685% | 1.896% | 4.378% | 1.947% | 1.966% | 1.527% | 2.319% | 4.31% | 2.466% | 6.222% | 10.97% | 3.656% | 1.965% | 1.739% | 2.611% | 3.218%
| PDM3 |  | 50.12% | 38.85% | -8.034% | 2.732% | -8.333% | -3.446% | 9.621% | -5.042% | -0.1543% | 23.49% | -8.633% | 13.45% | -7.336% | -7.735% | 3.137% | 3.836% | -6.339% | 4.833%
| Z-Score (Absolute) |  | 3.88 | 2.99 | 0.74 | 0.11 | 0.77 | 0.38 | 0.66 | 0.5 | 0.11 | 1.77 | 0.79 | 0.97 | 0.69 | 0.72 | 0.15 | 0.2 | 0.61 | 0.28
| Analytical results for Re in OREAS 922b (Certified Value < 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [10 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | < 0.001 | 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.001 | 0.001 | < 0.001 | 0.001 | 0.001 | < 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 23
| 1 | 3 | < 0.05 | < 0.01 | 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.001 | 0.001 | < 0.001 | 0.001 | 0.001 | < 0.001 | < 0.05 | < 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | < 0.001
| 1 | 5 | < 0.05 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | 0.001 | < 0.001 | < 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 102
| 1 | 6 | < 0.05 | < 0.01 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.001 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005
| Mean |  | 2.05: 0.001; 2.06: 0.001; 2.08: 0.001; 2.09: 0.001; 2.12: 0.001
| Median |  | 2.05: 0.001; 2.06: 0.001; 2.08: 0.001; 2.09: 0.001; 2.12: 0.001
| Std Dev. |  | 2.06: 0; 2.08: 0; 2.09: 0; 2.12: 0
Rel.Std.Dev. | bar: 2.06; 2.08 to 2.09; 2.12
PDM3
| Z-Score (Absolute) |  | 48.89 | 8.43 | 0.51 | 0 | 0.67 | 0 | 0.34 | 0.67 | 48.89 | 0.17 | 0.67 | 18.54 | 0.67 | 3.37
| Analytical results for S in OREAS 922b (Certified Value 0.271 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 3
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 0.26 | 0.27495 | 0.2875 | 0.18 | 0.26 | 0.28221 | 0.27 | 0.28 | 0.29 | 0.26 | 0.28 | 0.288 | 0.28 | 0.25514 | 0.28 | 0.25 | 0.2787 | 0.26 | 0.27 | 0.27 | 0.25 | 0.2789 | 0.26 | 0.26 | 1
| 1 | 2 | 0.26 | 0.2744 | 0.27233 | 0.15 | 0.27 | 0.27415 | 0.26 | 0.28 | 0.29 | 0.26 | 0.26 | 0.282 | 0.29 | 0.26208 | 0.3 | 0.26 | 0.2749 | 0.26 | 0.27 | 0.28 | 0.25 | 0.2613 | 0.27 | 0.27 | 0
| 1 | 3 | 0.26 | 0.27376 | 0.28937 | 0.16 | 0.26 | 0.27575 | 0.26 | 0.28 | 0.29 | 0.26 | 0.26 | 0.284 | 0.29 | 0.26572 | 0.28 | 0.25 | 0.2742 | 0.27 | 0.26 | 0.28 | 0.26 | 0.2667 | 0.27 | 0.26 | 16
| 1 | 4 | 0.26 | 0.27793 | 0.27819 | 0.18 | 0.26 | 0.28 | 0.26 | 0.27 | 0.29 | 0.26 | 0.27 | 0.277 | 0.29 | 0.26305 | 0.26 | 0.25 | 0.2739 | 0.26 | 0.27 | 0.28 | 0.26 | 0.2647 | 0.28 | 0.26 | 0.27082
| 1 | 5 | 0.27 | 0.27466 | 0.27385 | 0.16 | 0.26 | 0.27936 | 0.26 | 0.28 | 0.29 | 0.26 | 0.27 | 0.281 | 0.28 | 0.26015 | 0.28 | 0.25 | 0.2727 | 0.26 | 0.27 | 0.28 | 0.26 | 0.2762 | 0.28 | 0.26 | 103
| 1 | 6 | 0.26 | 0.27492 | 0.28264 | 0.16 | 0.27 | 0.27286 | 0.27 | 0.28 | 0.29 | 0.26 | 0.27 | 0.279 | 0.29 | 0.25721 | 0.3 | 0.26 | 0.2774 | 0.26 | 0.27 | 0.28 | 0.25 | 0.2805 | 0.28 | 0.26
| Mean |  | 0.26167 | 0.2751 | 0.28065 | 0.165 | 0.26333 | 0.27739 | 0.26333 | 0.27833 | 0.29 | 0.26 | 0.26833 | 0.28183 | 0.28667 | 0.26056 | 0.28333 | 0.25333 | 0.2753 | 0.26167 | 0.26833 | 0.27833 | 0.255 | 0.27138 | 0.27333 | 0.26167
| Median |  | 0.26 | 0.27479 | 0.28042 | 0.16 | 0.26 | 0.27755 | 0.26 | 0.28 | 0.29 | 0.26 | 0.27 | 0.2815 | 0.29 | 0.26111 | 0.28 | 0.25 | 0.27455 | 0.26 | 0.27 | 0.28 | 0.255 | 0.27145 | 0.275 | 0.26
| Std Dev. |  | 0.0040825 | 0.0014521 | 0.0070467 | 0.012247 | 0.005164 | 0.0036774 | 0.005164 | 0.0040825 | 0 | 0 | 0.0075277 | 0.0038687 | 0.005164 | 0.0038938 | 0.015055 | 0.005164 | 0.002283 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0081374 | 0.008165 | 0.0040825
| Rel.Std.Dev. |  | 1.56% | 0.5278% | 2.511% | 7.423% | 1.961% | 1.326% | 1.961% | 1.467% | 0% | 0% | 2.805% | 1.373% | 1.801% | 1.494% | 5.314% | 2.038% | 0.8293% | 1.56% | 1.521% | 1.467% | 2.148% | 2.998% | 2.987% | 1.56%
| PDM3 |  | -3.38% | 1.581% | 3.629% | -39.07% | -2.765% | 2.425% | -2.765% | 2.774% | 7.082% | -3.996% | -0.9185% | 4.066% | 5.851% | -3.79% | 4.62% | -6.457% | 1.654% | -3.38% | -0.9185% | 2.774% | -5.842% | 0.2077% | 0.9277% | -3.38%
| Z-Score (Absolute) |  | 0.66 | 0.42 | 0.87 | 8.48 | 0.53 | 0.61 | 0.53 | 0.69 | 1.63 | 0.8 | 0.12 | 0.97 | 1.36 | 0.75 | 1.09 | 1.34 | 0.44 | 0.66 | 0.12 | 0.69 | 1.2 | 0.12 | 0.28 | 0.66
| Analytical results for Sb in OREAS 922b (Certified Value 2.27 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 2
| 1 | 1 | 3.18 | 2.5371 | 1.5 | < 5 | 1.86 | 2.2801 | 1.66 | 2.33 | 1.49 | 1.79 | 2.15 | 2.61 | 2.61 | 2.386 | 3.39 | 2.44 | 2.53 | 1.7 | < 5 | 3.19 | 2.16 | 3 | 1.95 | 2.11 | 1
| 1 | 2 | 3.17 | 2.5456 | 1.56 | < 5 | 1.92 | 2.4686 | 1.92 | 2.3 | 1.52 | 1.82 | 2.1 | 2.62 | 2.52 | 2.5193 | 2.01 | 2.43 | 4.51 | 1.7 | < 5 | 3.37 | 2.15 | 3 | 1.84 | 2.14 | 39
| 1 | 3 | 3.16 | 2.5145 | 1.54 | < 5 | 1.86 | 2.4671 | 1.88 | 2.48 | 1.45 | 1.77 | 2.07 | 2.62 | 2.7 | 2.4316 | 5 | 2.43 | 2.64 | 1.7 | < 5 | 3.02 | 2.23 | 3 | 1.91 | 2.13 | 16
| 1 | 4 | 3.12 | 2.6134 | 1.52 | < 5 | 2.04 | 2.2243 | 1.93 | 2.48 | 1.52 | 1.93 | 2.1 | 2.6 | 2.65 | 2.4838 | 5.51 | 2.37 | 2.78 | 1.8 | < 5 | 3.39 | 2.16 | 3 | 2.05 | 2.12 | 2.2749
| 1 | 5 | 3.14 | 2.5699 | 1.51 | < 5 | 1.89 | 2.2065 | 1.94 | 2.4 | 1.49 | 1.74 | 2.18 | 2.78 | 2.55 | 2.4022 | 2.83 | 2.58 | 3.55 | 1.8 | < 5 | 3.48 | 2.31 | 3 | 1.99 | 2.11 | 104
| 1 | 6 | 3.12 | 2.5267 | 1.57 | < 5 | 1.9 | 2.327 | 1.9 | 2.41 | 1.48 | 1.84 | 2.25 | 2.74 | 2.76 | 2.458 | 5.85 | 2.45 | 2.65 | 1.7 | < 5 | 3.31 | 2.28 | 3 | 1.95 | 2.15
| Mean |  | 3.1483 | 2.5512 | 1.5333 |  | 1.9117 | 2.3289 | 1.8717 | 2.4 | 1.4917 | 1.815 | 2.1417 | 2.6617 | 2.6317 | 2.4468 | 4.0983 | 2.45 | 3.11 | 1.7333 |  | 3.2933 | 2.215 | 3 | 1.9483 | 2.1267
| Median |  | 3.15 | 2.5413 | 1.53 |  | 1.895 | 2.3035 | 1.91 | 2.405 | 1.49 | 1.805 | 2.125 | 2.62 | 2.63 | 2.4448 | 4.195 | 2.435 | 2.715 | 1.7 |  | 3.34 | 2.195 | 3 | 1.95 | 2.125
| Std Dev. |  | 0.025626 | 0.035765 | 0.028048 |  | 0.067057 | 0.11568 | 0.10591 | 0.074565 | 0.026394 | 0.066558 | 0.066156 | 0.077567 | 0.090646 | 0.050318 | 1.5713 | 0.06957 | 0.77864 | 0.05164 |  | 0.16476 | 0.06892 | 0 | 0.07111 | 0.01633
| Rel.Std.Dev. |  | 0.8139% | 1.402% | 1.829% |  | 3.508% | 4.967% | 5.659% | 3.107% | 1.769% | 3.667% | 3.089% | 2.914% | 3.444% | 2.056% | 38.34% | 2.84% | 25.04% | 2.979% |  | 5.003% | 3.112% | 0% | 3.65% | 0.7679%
| PDM3 |  | 38.4% | 12.15% | -32.6% |  | -15.97% | 2.377% | -17.72% | 5.501% | -34.43% | -20.22% | -5.855% | 17% | 15.68% | 7.559% | 80.16% | 7.699% | 36.71% | -23.8% |  | 44.77% | -2.632% | 31.88% | -14.35% | -6.515%
| Z-Score (Absolute) |  | 1.85 | 0.37 | 2.14 | 0.25 | 1.2 | 0.18 | 1.3 | 0 | 2.24 | 1.44 | 0.64 | 0.65 | 0.57 | 0.12 | 4.19 | 0.12 | 1.75 | 1.64 | 0.25 | 2.2 | 0.46 | NA | 1.11 | 0.67
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 922b (Certified Value 2.03 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2 | 2.2701 | 1.8 | 1.7 | 1.89 | 2.0566 | 2 | 2.1 | 2.1 | 2 | 2 | 2 | 1.8 | 1.9961 | 2.1 | 2 | 1.8 | 1.8 | 2.1 | 1.8 | 3 | 2.1 | 2.2 |  | 1
| 1 | 2 | 2 | 2.2296 | 1.8 | 1.7 | 1.93 | 2.2669 | 2.1 | 2.1 | 2.1 | 2 | 2 | 2.2 | 1.7 | 1.9856 | 3.3 | 1.9 | 1.8 | 1.8 | 2.3 | 1.8 | 3 | 2.2 | 2.1 |  | 19
| 1 | 3 | 2 | 2.2865 | 1.8 | 1.6 | 1.91 | 2.1874 | 2 | 2.2 | 2.1 | 2 | 2 | 2.5 | 1.8 | 2.2995 | 2.5 | 1.9 | 1.9 | 1.9 | 2.2 | 2 | 3 | 2.1 | 2.2 |  | 16
| 1 | 4 | 2 | 2.3318 | 1.8 | 1.7 | 1.88 | 2.2263 | 2.1 | 2.1 | 2.1 | 2 | 2 | 2.1 | 1.7 | 2.0651 | 2.4 | 2.1 | 1.8 | 2 | 2.1 | 1.9 | 2 | 2.1 | 2.1 |  | 2.0295
| 1 | 5 | 2 | 2.3458 | 1.8 | 1.7 | 1.93 | 1.9687 | 2.1 | 2.1 | 2.1 | 2 | 2.1 | 2.5 | 1.7 | 2.2705 | 2 | 2 | 1.8 | 1.9 | 2.1 | 2 | 3 | 2.1 | 2.1 |  | 105
| 1 | 6 | 2 | 2.2271 | 1.8 | 1.7 | 1.93 | 2.0176 | 2.1 | 2.2 | 2.1 | 2 | 2.2 | 2.4 | 1.8 | 2.0147 | 2.6 | 2 | 1.8 | 1.8 | 2.1 | 1.9 | 3 | 2.2 | 2.1
| Mean |  | 2 | 2.2818 | 1.8 | 1.6833 | 1.9117 | 2.1206 | 2.0667 | 2.1333 | 2.1 | 2 | 2.05 | 2.2833 | 1.75 | 2.1052 | 2.4833 | 1.9833 | 1.8167 | 1.8667 | 2.15 | 1.9 | 2.8333 | 2.1333 | 2.1333
| Median |  | 2 | 2.2783 | 1.8 | 1.7 | 1.92 | 2.122 | 2.1 | 2.1 | 2.1 | 2 | 2 | 2.3 | 1.75 | 2.0399 | 2.45 | 2 | 1.8 | 1.85 | 2.1 | 1.9 | 3 | 2.1 | 2.1
| Std Dev. |  | 0 | 0.049942 | 0 | 0.040825 | 0.022286 | 0.12232 | 0.05164 | 0.05164 | 0 | 0 | 0.083666 | 0.2137 | 0.054772 | 0.14218 | 0.46224 | 0.075277 | 0.040825 | 0.08165 | 0.083666 | 0.089443 | 0.40825 | 0.05164 | 0.05164
| Rel.Std.Dev. |  | 0% | 2.189% | 0% | 2.425% | 1.166% | 5.768% | 2.499% | 2.421% | 0% | 0% | 4.081% | 9.359% | 3.13% | 6.754% | 18.61% | 3.795% | 2.247% | 4.374% | 3.891% | 4.708% | 14.41% | 2.421% | 2.421%
| PDM3 |  | -1.452% | 12.44% | -11.31% | -17.06% | -5.804% | 4.489% | 1.833% | 5.118% | 3.475% | -1.452% | 1.012% | 12.51% | -13.77% | 3.734% | 22.36% | -2.273% | -10.49% | -8.022% | 5.939% | -6.379% | 39.61% | 5.118% | 5.118%
| Z-Score (Absolute) |  | NA | 0.99 | 1.41 | 1.99 | 0.85 | 0.18 | 0.08 | 0.25 | 0.08 | NA | 0.17 | 0.99 | 1.65 | 0.11 | 1.99 | 0.5 | 1.32 | 1.08 | 0.33 | 0.91 | NA | 0.25 | 0.25
Results from laboratories 2.01, 2.10 and 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 922b (Certified Value 4.45 ppm) | [25 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 4 | 3.231 | 4.5 | < 5 | 4.5898 | 4.4 | 4.9 | 4.6 | 4.7 | 4.9 | 4.3 | 4 | 4.259 | 2.3 | 4.1 | 2.86 | 4.4 | < 5 | 5 | 4.1 | 6.6 |  |  |  | 1
| 1 | 2 | 5 | 3.1444 | 4.3 | < 5 | 4.6197 | 4.2 | 5 | 4.5 | 4.8 | 4.4 | 4.5 | 3 | 4.438 | 1.7 | 3.7 | 2.95 | 4.5 | < 5 | 5 | 4.3 | 6.2 |  |  |  | 41
| 1 | 3 | 5 | 3.1452 | 4.3 | < 5 | 4.53 | 4.4 | 4.7 | 4.4 | 4.8 | 4.4 | 4.6 | 4 | 3.8901 | 2 | 3.6 | 2.66 | 4.5 | < 5 | 5 | 4.2 | 6.2 |  |  |  | 16
| 1 | 4 | 5 | 3.1461 | 4.6 | < 5 | 4.0855 | 4.5 | 4.1 | 4.7 | 4.8 | 4.7 | 4.3 | 4 | 3.8213 | 2.3 | 4 | 4.05 | 4.5 | < 5 | 5 | 4.3 | 5.8 |  |  |  | 4.455
| 1 | 5 | 5 | 3.1159 | 4.5 | < 5 | 4.4583 | 4.6 | 5.3 | 4.7 | 4.7 | 4.7 | 4.9 | 3 | 4.1618 | 2.2 | 3.5 | 3.66 | 4.6 | < 5 | 5 | 4.3 | 5.6 |  |  |  | 106
| 1 | 6 | 5 | 3.1869 | 4.6 | < 5 | 4.2669 | 4.5 | 5.4 | 4.7 | 4.9 | 4.6 | 4.8 | 4 | 4.4388 | 2 | 3.5 | 3.35 | 4.4 | < 5 | 5 | 4.5 | 6.2
| Mean |  | 4.8333 | 3.1616 | 4.4667 |  | 4.425 | 4.4333 | 4.9 | 4.6 | 4.7833 | 4.6167 | 4.5667 | 3.6667 | 4.1681 | 2.0833 | 3.7333 | 3.255 | 4.4833 |  | 5 | 4.2833 | 6.1
| Median |  | 5 | 3.1456 | 4.5 |  | 4.4942 | 4.45 | 4.95 | 4.65 | 4.8 | 4.65 | 4.55 | 4 | 4.2104 | 2.1 | 3.65 | 3.15 | 4.5 |  | 5 | 4.3 | 6.2
| Std Dev. |  | 0.40825 | 0.040866 | 0.13663 |  | 0.20853 | 0.13663 | 0.46904 | 0.12649 | 0.075277 | 0.19408 | 0.25033 | 0.5164 | 0.26533 | 0.23166 | 0.2582 | 0.53069 | 0.075277 |  | 0 | 0.13292 | 0.35214
| Rel.Std.Dev. |  | 8.447% | 1.293% | 3.059% |  | 4.713% | 3.082% | 9.572% | 2.75% | 1.574% | 4.204% | 5.482% | 14.08% | 6.366% | 11.12% | 6.916% | 16.3% | 1.679% |  | 0% | 3.103% | 5.773%
| PDM3 |  | 8.493% | -29.03% | 0.2622% |  | -0.6724% | -0.4861% | 9.989% | 3.255% | 7.37% | 3.629% | 2.507% | -17.7% | -6.439% | -53.24% | -16.2% | -26.94% | 0.6363% |  | 12.23% | -3.853% | 36.93%
| Z-Score (Absolute) |  | NA | 3.23 | 0.08 | 4.92 | 0.02 | 0 | 1.19 | 0.42 | 0.89 | 0.47 | 0.34 | NA | 0.67 | 5.98 | 1.78 | 3 | 0.13 | NA | NA | 0.38 | 4.24
Results from laboratories 2.01, 2.13, 2.20 and 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 922b (Indicative Value 4.78 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  |  |  | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g |  |  |  | 2
| 1 | 1 | 2.01: 4.546; 2.02: 4.7439; 2.03: 5.2 |  |  |  | 1
| 1 | 2 | 2.01: 4.347; 2.02: 5.0361; 2.03: 5.1 |  |  |  | 4
| 1 | 3 | 2.01: 4.3582; 2.02: 5.0856; 2.03: 4.9 |  |  |  | 16
| 1 | 4 | 2.01: 4.5675; 2.02: 4.7551; 2.03: 4.6 |  |  |  | 4.778
| 1 | 5 | 2.01: 4.5395; 2.02: 4.593; 2.03: 4.6 |  |  |  | 10
| 1 | 6 | 2.01: 4.433; 2.02: 4.7989; 2.03: 5.8
| Mean |  | 2.01: 4.4652; 2.02: 4.8354; 2.03: 5.0333
| Median |  | 2.01: 4.4863; 2.02: 4.777; 2.03: 5
| Std Dev. |  | 2.01: 0.098966; 2.02: 0.18858; 2.03: 0.45019
| Rel.Std.Dev. |  | 2.01: 2.216%; 2.02: 3.9%; 2.03: 8.944%
| PDM3 |  | 2.01: -6.546%; 2.02: 1.202%; 2.03: 5.344%
| Z-Score (Absolute) |  | 2.01: 1.26; 2.02: 0; 2.03: 0.67
| Analytical results for Sn in OREAS 922b (Certified Value 2.42 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 2.5 | 2.008 | < 5 | 2.43 | 2.3625 | 2.3 | 2.5 | 2.6 | 2.3 | 2.4 | 2.55 | 2.5 | 2.4787 | < 0.2 | 1.3 | 2.6 | 2 | 2.5 | 2.3 | [5 empty] | 1
| 1 | 2 | 2.5 | 2.0243 | < 5 | 2.54 | 2.4503 | 2.3 | 2.4 | 2.6 | 2.3 | 2.3 | 2.57 | 2.5 | 2.4303 | 0.4 | 1.5 | 2.6 | 2.3 | 2.6 | 2.3 | [5 empty] | 26
| 1 | 3 | 2.5 | 1.9875 | < 5 | 2.53 | 2.6398 | 2.3 | 2.6 | 2.6 | 2.3 | 2.3 | 2.54 | 2.5 | 2.5328 | < 0.2 | 1.4 | 2.7 | 2.4 | 2.4 | 2.3 | [5 empty] | 16
| 1 | 4 | 2.5 | 2.0291 | < 5 | 2.48 | 2.3698 | 2.4 | 2.4 | 2.5 | 2.3 | 2.3 | 2.55 | 2.5 | 2.5589 | < 0.2 | 1.6 | 2.7 | 2.3 | 2.5 | 2.3 | [5 empty] | 2.4202
| 1 | 5 | 2.5 | 2.0228 | < 5 | 2.48 | 2.2936 | 2.4 | 2.4 | 2.6 | 2.2 | 2.4 | 2.65 | 2.5 | 2.4658 | < 0.2 | 1.4 | 2.7 | 2.3 | 2.5 | 2.2 | [5 empty] | 107
| 1 | 6 | 2.5 | 2.0147 | < 5 | 2.48 | 2.3581 | 2.4 | 2.4 | 2.5 | 2.3 | 2.3 | 2.66 | 2.6 | 2.5059 | < 0.2 | 1.4 | 2.6 | 2.3 | 2.6 | 2.3
| Mean |  | 2.5 | 2.0144 |  | 2.49 | 2.4123 | 2.35 | 2.45 | 2.5667 | 2.2833 | 2.3333 | 2.5867 | 2.5167 | 2.4954 | 0.4 | 1.4333 | 2.65 | 2.2667 | 2.5167 | 2.2833
| Median |  | 2.5 | 2.0187 |  | 2.48 | 2.3661 | 2.35 | 2.4 | 2.6 | 2.3 | 2.3 | 2.56 | 2.5 | 2.4923 | 0.4 | 1.4 | 2.65 | 2.3 | 2.5 | 2.3
| Std Dev. |  | 0 | 0.015161 |  | 0.04 | 0.12206 | 0.054772 | 0.083666 | 0.05164 | 0.040825 | 0.05164 | 0.053914 | 0.040825 | 0.046767 |  | 0.10328 | 0.054772 | 0.13663 | 0.075277 | 0.040825
| Rel.Std.Dev. |  | 0% | 0.7526% |  | 1.606% | 5.06% | 2.331% | 3.415% | 2.012% | 1.788% | 2.213% | 2.084% | 1.622% | 1.874% |  | 7.206% | 2.067% | 6.028% | 2.991% | 1.788%
| PDM3 |  | 3.297% | -16.77% |  | 2.884% | -0.3254% | -2.901% | 1.231% | 6.052% | -5.655% | -3.59% | 6.878% | 3.986% | 3.107% | -83.47% | -40.78% | 9.495% | -6.344% | 3.986% | -5.655%
| Z-Score (Absolute) |  | 0.29 | 2.52 | 0.29 | 0.23 | 0.22 | 0.58 | 0 | 0.67 | 0.96 | 0.67 | 0.79 | 0.39 | 0.26 | 13.29 | 5.88 | 1.16 | 1.06 | 0.39 | 0.96
| Analytical results for Sr in OREAS 922b (Certified Value 9.04 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 12.8 | 7.3478 | 7 | 12.1 | 6.86 | 9.9547 | 9.6 | 8.8 | 11 | 8.5 | 9.6 | 14.2 | 8.7 | 12.4411 | 4.5 | 9.3 | 8 | 8.8 | 8 | 11.2 | 13 | 7.8 | 10 |  | 1
| 1 | 2 | 13.4 | 7.023 | 6 | 11.9 | 6.84 | 9.9738 | 8.5 | 8.6 | 11.2 | 8.4 | 9.6 | 14.2 | 8.2 | 12.9012 | 15.8 | 9.1 | 7 | 9.2 | 7.9 | 10.6 | 14 | 7.9 | 9 |  | 27
| 1 | 3 | 12.8 | 7.0948 | 6 | 10.9 | 7.13 | 10.3688 | 8.3 | 8.8 | 10.8 | 8.5 | 9.3 | 14.2 | 8.5 | 13.6987 | 4.7 | 8.2 | 8 | 9.9 | 7.8 | 10.9 | 14 | 8.1 | 10 |  | 16
| 1 | 4 | 13.4 | 7.5416 | 6 | 11.4 | 7.1 | 9.9738 | 8.5 | 8.6 | 11 | 8.5 | 9.4 | 14.2 | 8.4 | 13.9773 | 4.7 | 9.2 | 8 | 8.9 | 7.7 | 11 | 13 | 8.1 | 9 |  | 9.0363
| 1 | 5 | 13.1 | 7.4087 | 7 | 11.7 | 6.93 | 9.6175 | 8.5 | 8.9 | 11 | 8.5 | 9.9 | 15.4 | 8.6 | 12.004 | 4.2 | 7.6 | 8 | 9.9 | 7.9 | 10.6 | 14 | 8.1 | 10 |  | 108
| 1 | 6 | 13.5 | 7.0997 | 6 | 12.2 | 6.89 | 9.9972 | 8.8 | 8.8 | 11 | 8.7 | 9.8 | 15.1 | 8.9 | 11.9004 | 4.8 | 8.6 | 7 | 8.9 | 7.9 | 10.9 | 14 | 8.3 | 9
| Mean |  | 13.1667 | 7.2526 | 6.3333 | 11.7 | 6.9583 | 9.981 | 8.7 | 8.75 | 11 | 8.5167 | 9.6 | 14.55 | 8.55 | 12.8204 | 6.45 | 8.6667 | 7.6667 | 9.2667 | 7.8667 | 10.8667 | 13.6667 | 8.05 | 9.5
| Median |  | 13.25 | 7.2238 | 6 | 11.8 | 6.91 | 9.9738 | 8.5 | 8.8 | 11 | 8.5 | 9.6 | 14.2 | 8.55 | 12.6711 | 4.7 | 8.85 | 8 | 9.05 | 7.9 | 10.9 | 14 | 8.1 | 9.5
| Std Dev. |  | 0.31411 | 0.20876 | 0.5164 | 0.4858 | 0.12545 | 0.23813 | 0.46904 | 0.12247 | 0.12649 | 0.098319 | 0.22804 | 0.55045 | 0.2429 | 0.86869 | 4.5855 | 0.66833 | 0.5164 | 0.50859 | 0.10328 | 0.23381 | 0.5164 | 0.17607 | 0.54772
| Rel.Std.Dev. |  | 2.386% | 2.878% | 8.154% | 4.152% | 1.803% | 2.386% | 5.391% | 1.4% | 1.15% | 1.154% | 2.375% | 3.783% | 2.841% | 6.776% | 71.09% | 7.712% | 6.736% | 5.488% | 1.313% | 2.152% | 3.779% | 2.187% | 5.766%
| PDM3 |  | 45.71% | -19.74% | -29.91% | 29.48% | -23% | 10.45% | -3.722% | -3.169% | 21.73% | -5.751% | 6.238% | 61.02% | -5.382% | 41.88% | -28.62% | -4.091% | -15.16% | 2.549% | -12.94% | 20.26% | 51.24% | -10.92% | 5.131%
| Z-Score (Absolute) |  | 2.42 | 0.82 | NA | 1.62 | 0.98 | 0.67 | 0.03 | 0 | 1.23 | 0.13 | 0.47 | 3.18 | 0.11 | 2.23 | 1.26 | 0.05 | NA | 0.28 | 0.48 | 1.16 | NA | 0.38 | NA
Results from laboratories 2.03, 2.19, 2.23 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 922b (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 10
| 1 | 3 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.02 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | < 0.01
| 1 | 5 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 109
| 1 | 6 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | 0.01 | < 0.01 | < 0.05 | < 0.3 | 0.03 | < 0.1
| Mean |  | 2.15: 0.013333; 2.24: 0.03
| Median |  | 2.15: 0.01; 2.24: 0.03
| Std Dev. |  | 2.15: 0.0057735
| Rel.Std.Dev. |  | 2.15: 43.3%
PDM3
| Z-Score (Absolute) |  | 2.56 | 403.1 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 2.56 | 0.67 | 2.56 | 0 | 0.67 | 2.56 | 22.79 | 0 | 6.61
| Analytical results for Tb in OREAS 922b (Certified Value 0.42 ppm) | [25 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | [7 empty] | 1
| No | No | 2.01: 2.02; 2.04: 2.04; 2.06: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.21; 2.10: 2.22 | [7 empty] | ppm
|  |  | 2.01: AR*MS; 2.04: AR*OES; 2.06: AR*MS; 2.07: AR*OES/MS; 2.08: AR*MS; 2.09: AR*OES/MS; 2.10: AR*OES/MS | [7 empty] | 2
|  |  | 2.01: 0.5g; 2.04: 0.25g; 2.06: 0.2g; 2.07: 0.5g; 2.08: 0.25g; 2.09: 0.25g; 2.10: 0.2g | [7 empty] | 3
| 1 | 1 | 2.01: 0.4058; 2.04: < 5; 2.06: 0.458; 2.07: 0.5; 2.08: 0.41; 2.09: 0.36; 2.10: 0.47 | [7 empty] | 1
| 1 | 2 | 2.01: 0.3835; 2.04: < 5; 2.06: 0.48927; 2.07: 0.5; 2.08: 0.39; 2.09: 0.36; 2.10: 0.44 | [7 empty] | 11
| 1 | 3 | 2.01: 0.3937; 2.04: < 5; 2.06: 0.49412; 2.07: 0.4; 2.08: 0.4; 2.09: 0.38; 2.10: 0.46 | [7 empty] | 16
| 1 | 4 | 2.01: 0.42; 2.04: < 5; 2.06: 0.45445; 2.07: 0.4; 2.08: 0.38; 2.09: 0.37; 2.10: 0.44 | [7 empty] | 0.41632
| 1 | 5 | 2.01: 0.4174; 2.04: < 5; 2.06: 0.48565; 2.07: 0.5; 2.08: 0.39; 2.09: 0.37; 2.10: 0.44 | [7 empty] | 110
| 1 | 6 | 2.01: 0.3898; 2.04: < 5; 2.06: 0.44383; 2.07: 0.5; 2.08: 0.4; 2.09: 0.38; 2.10: 0.44
| Mean |  | 2.01: 0.4017; 2.06: 0.47089; 2.07: 0.46667; 2.08: 0.395; 2.09: 0.37; 2.10: 0.44833
| Median |  | 2.01: 0.39975; 2.06: 0.47183; 2.07: 0.5; 2.08: 0.395; 2.09: 0.37; 2.10: 0.44
| Std Dev. |  | 2.01: 0.015067; 2.06: 0.02128; 2.07: 0.05164; 2.08: 0.010488; 2.09: 0.0089443; 2.10: 0.013292
| Rel.Std.Dev. |  | 2.01: 3.751%; 2.06: 4.519%; 2.07: 11.07%; 2.08: 2.655%; 2.09: 2.417%; 2.10: 2.965%
| PDM3 |  | 2.01: -3.511%; 2.06: 13.11%; 2.07: 12.09%; 2.08: -5.12%; 2.09: -11.13%; 2.10: 7.69%
| Z-Score (Absolute) |  | 2.01: 0.41; 2.04: 36.88; 2.06: 0.82; 2.07: NA; 2.08: 0.53; 2.09: 0.98; 2.10: 0.41
Results from laboratory 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 922b (Certified Value 0.059 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.05 | 0.04 | < 5 | 0.05 | < 0.2 | 0.04 | 0.03 | 0.04 | 0.06 | 0.05 | 0.06 | 0.05 | 0.0641 | 0.03 | 0.06 | 1.52 | < 0.2 | < 5 | 0.06 | 0.06 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.06 | 0.04 | < 5 | 0.06 | < 0.2 | 0.06 | 0.05 | 0.06 | 0.06 | 0.06 | 0.04 | 0.07 | 0.069768 | 0.07 | 0.04 | 1.13 | < 0.2 | < 5 | 0.07 | 0.06 | < 0.1 |  |  |  | 12
| 1 | 3 | 0.07 | 0.03 | < 5 | 0.06 | < 0.2 | 0.06 | 0.05 | 0.07 | 0.06 | 0.07 | 0.03 | 0.05 | 0.068916 | 0.02 | 0.05 | 1.74 | < 0.2 | < 5 | 0.05 | 0.06 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.06 | 0.04 | < 5 | 0.05 | < 0.2 | 0.06 | 0.08 | 0.07 | 0.06 | 0.06 | 0.05 | 0.06 | 0.081954 | 0.03 | 0.06 | 1.56 | < 0.2 | < 5 | 0.06 | 0.07 | < 0.1 |  |  |  | 0.058656
| 1 | 5 | 0.05 | 0.04 | < 5 | 0.05 | < 0.2 | 0.06 | 0.04 | 0.06 | 0.07 | 0.05 | 0.05 | 0.05 | 0.082044 | 0.02 | 0.04 | 2.29 | < 0.2 | < 5 | 0.07 | 0.06 | < 0.1 |  |  |  | 111
| 1 | 6 | 0.06 | 0.04 | < 5 | 0.05 | < 0.2 | 0.07 | 0.1 | 0.08 | 0.06 | 0.06 | 0.07 | 0.07 | 0.068351 | 0.04 | 0.07 | 1.79 | < 0.2 | < 5 | 0.07 | 0.06 | < 0.1
| Mean |  | 0.058333 | 0.038333 |  | 0.053333 |  | 0.058333 | 0.058333 | 0.063333 | 0.061667 | 0.058333 | 0.05 | 0.058333 | 0.072522 | 0.035 | 0.053333 | 1.6717 |  |  | 0.063333 | 0.061667
| Median |  | 0.06 | 0.04 |  | 0.05 |  | 0.06 | 0.05 | 0.065 | 0.06 | 0.06 | 0.05 | 0.055 | 0.069342 | 0.03 | 0.055 | 1.65 |  |  | 0.065 | 0.06
| Std Dev. |  | 0.0075277 | 0.0040825 |  | 0.005164 |  | 0.0098319 | 0.026394 | 0.013663 | 0.0040825 | 0.0075277 | 0.014142 | 0.0098319 | 0.0075967 | 0.018708 | 0.012111 | 0.38207 |  |  | 0.008165 | 0.0040825
| Rel.Std.Dev. |  | 12.9% | 10.65% |  | 9.682% |  | 16.85% | 45.25% | 21.57% | 6.62% | 12.9% | 28.28% | 16.85% | 10.48% | 53.45% | 22.71% | 22.86% |  |  | 12.89% | 6.62%
| PDM3 |  | -0.5493% | -34.65% |  | -9.074% |  | -0.5493% | -0.5493% | 7.975% | 5.134% | -0.5493% | -14.76% | -0.5493% | 23.64% | -40.33% | -9.074% | 2750% |  |  | 7.975% | 5.134%
| Z-Score (Absolute) |  | 0 | 2.7 | 329.29 | 0.67 | 5.62 | 0 | 0 | 0.67 | 0.45 | 0 | 1.12 | 0 | 1.91 | 3.15 | 0.67 | 217.58 | 5.62 | 329.29 | 0.67 | 0.45 | 1.12
| Analytical results for Th in OREAS 922b (Certified Value 9.91 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 10.3 | 10.478 | 10.5 | 11 | 9.505 | 9.9829 | 9 | 9.9 | 9.7 | 9.5 | 10 | 10.7 | 9.33 | 11.096 | < 20 | 8.7 | 9.4 | 9.2 | 9.8 | 9.6 | 9.2 |  |  |  | 1
| 1 | 2 | 10.5 | 10.58 | 10.45 | 11 | 9.77 | 10.8487 | 9.5 | 9.9 | 9.8 | 9.5 | 9.7 | 10.8 | 9.64 | 11.2076 | < 20 | 8.8 | 9.4 | 8.7 | 10.2 | 10.2 | 9.5 |  |  |  | 13
| 1 | 3 | 10.2 | 10.336 | 10.45 | 9.9 | 9.778 | 10.3396 | 9.3 | 9.7 | 9.7 | 9.4 | 9.8 | 10.8 | 9.46 | 11.8042 | < 20 | 9 | 9.4 | 9 | 10.3 | 10 | 9.2 |  |  |  | 16
| 1 | 4 | 10.3 | 10.409 | 10.5 | 11.1 | 9.748 | 10.2121 | 9.5 | 9.7 | 9.6 | 9.4 | 9.6 | 10.6 | 9.28 | 11.5419 | < 20 | 8.9 | 9.5 | 9 | 10.3 | 10 | 9.4 |  |  |  | 9.9094
| 1 | 5 | 10.2 | 10.552 | 10.4 | 11 | 9.427 | 9.2036 | 9.6 | 9.8 | 9.9 | 9.2 | 9.7 | 11.1 | 9.4 | 10.6182 | < 20 | 9.1 | 9.5 | 8.8 | 9.7 | 10.1 | 9.4 |  |  |  | 112
| 1 | 6 | 10.3 | 10.396 | 10.3 | 11 | 9.59 | 9.5322 | 9.5 | 10.1 | 9.4 | 9.4 | 9.9 | 11.3 | 9.63 | 10.8349 | < 20 | 8.8 | 9.3 | 8.5 | 10 | 10.3 | 9.4
| Mean |  | 10.3 | 10.4585 | 10.4333 | 10.8333 | 9.6363 | 10.0199 | 9.4 | 9.85 | 9.6833 | 9.4 | 9.7833 | 10.8833 | 9.4567 | 11.1838 |  | 8.8833 | 9.4167 | 8.8667 | 10.05 | 10.0333 | 9.35
| Median |  | 10.3 | 10.4435 | 10.45 | 11 | 9.669 | 10.0975 | 9.5 | 9.85 | 9.7 | 9.4 | 9.75 | 10.8 | 9.43 | 11.1518 |  | 8.85 | 9.4 | 8.9 | 10.1 | 10.05 | 9.4
| Std Dev. |  | 0.10954 | 0.095129 | 0.075277 | 0.45898 | 0.15075 | 0.58851 | 0.21909 | 0.15166 | 0.17224 | 0.10954 | 0.1472 | 0.26394 | 0.15108 | 0.43884 |  | 0.1472 | 0.075277 | 0.25033 | 0.25884 | 0.24221 | 0.12247
| Rel.Std.Dev. |  | 1.064% | 0.9096% | 0.7215% | 4.237% | 1.564% | 5.873% | 2.331% | 1.54% | 1.779% | 1.165% | 1.505% | 2.425% | 1.598% | 3.924% |  | 1.657% | 0.7994% | 2.823% | 2.576% | 2.414% | 1.31%
| PDM3 |  | 3.941% | 5.541% | 5.287% | 9.324% | -2.756% | 1.114% | -5.141% | -0.5997% | -2.282% | -5.141% | -1.272% | 9.828% | -4.569% | 12.86% |  | -10.35% | -4.973% | -10.52% | 1.419% | 1.25% | -5.645%
| Z-Score (Absolute) |  | 0.67 | 0.91 | 0.87 | 1.47 | 0.32 | 0.25 | 0.67 | 0 | 0.25 | 0.67 | 0.1 | 1.55 | 0.59 | 2 | 0.22 | 1.45 | 0.65 | 1.47 | 0.3 | 0.27 | 0.75
| Analytical results for Ti in OREAS 922b (Certified Value 0.011 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0175 | 0.01118 | 0.02 | 0.0092 | 0.013865 | 0.013 | 0.012 | 0.015 | 0.012 | 0.012 | 0.02 | 0.011 | 0.019755 | 0.006 | 0.012 | 0.011 | 0.01 | 0.01 | 0.018 | 0.0185 | 0.01 | 0.013 |  |  | 1
| 1 | 2 | 0.0165 | 0.00871 | 0.02 | 0.0092 | 0.013786 | 0.012 | 0.012 | 0.015 | 0.012 | 0.012 | 0.02 | 0.01 | 0.020821 | 0.039 | 0.012 | 0.01 | 0.01 | 0.01 | 0.018 | 0.019 | 0.011 | 0.013 |  |  | 32
| 1 | 3 | 0.0167 | 0.0091 | 0.02 | 0.0096 | 0.013684 | 0.011 | 0.012 | 0.015 | 0.011 | 0.012 | 0.02 | 0.011 | 0.020523 | 0.007 | 0.011 | 0.01 | 0.01 | 0.01 | 0.019 | 0.0184 | 0.01 | 0.013 |  |  | 16
| 1 | 4 | 0.0168 | 0.01001 | 0.02 | 0.0093 | 0.014207 | 0.011 | 0.012 | 0.015 | 0.012 | 0.012 | 0.02 | 0.01 | 0.020969 | 0.006 | 0.012 | 0.011 | 0.01 | < 0.01 | 0.018 | 0.0188 | 0.011 | 0.012 |  |  | 0.011405
| 1 | 5 | 0.0164 | 0.01196 | 0.02 | 0.0093 | 0.013041 | 0.012 | 0.012 | 0.015 | 0.012 | 0.012 | 0.02 | 0.01 | 0.018959 | 0.007 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.018 | 0.0189 | 0.011 | 0.012 |  |  | 113
| 1 | 6 | 0.0159 | 0.01014 | 0.02 | 0.0092 | 0.013028 | 0.012 | 0.012 | 0.015 | 0.012 | 0.012 | 0.02 | 0.011 | 0.018992 | 0.006 | 0.011 | 0.009 | 0.01 | 0.01 | 0.018 | 0.0178 | 0.011 | 0.012
| Mean |  | 0.016633 | 0.010183 | 0.02 | 0.0093 | 0.013602 | 0.011833 | 0.012 | 0.015 | 0.011833 | 0.012 | 0.02 | 0.0105 | 0.020003 | 0.011833 | 0.011333 | 0.010167 | 0.01 | 0.01 | 0.018167 | 0.018567 | 0.010667 | 0.0125
| Median |  | 0.0166 | 0.010075 | 0.02 | 0.00925 | 0.013735 | 0.012 | 0.012 | 0.015 | 0.012 | 0.012 | 0.02 | 0.0105 | 0.020139 | 0.0065 | 0.0115 | 0.01 | 0.01 | 0.01 | 0.018 | 0.01865 | 0.011 | 0.0125
| Std Dev. |  | 0.00052789 | 0.0012269 | 0 | 0.00015492 | 0.00047309 | 0.00075277 | 1.9003e-18 | 0 | 0.00040825 | 1.9003e-18 | 0 | 0.00054772 | 0.00089963 | 0.013318 | 0.0008165 | 0.00075277 | 0 | 0 | 0.00040825 | 0.00044121 | 0.0005164 | 0.00054772
| Rel.Std.Dev. |  | 3.174% | 12.05% | 0% | 1.666% | 3.478% | 6.361% | 1.584e-14% | 0% | 3.45% | 1.584e-14% | 0% | 5.216% | 4.497% | 112.5% | 7.204% | 7.404% | 0% | 0% | 2.247% | 2.376% | 4.841% | 4.382%
| PDM3 |  | 45.84% | -10.71% | 75.36% | -18.46% | 19.26% | 3.752% | 5.213% | 31.52% | 3.752% | 5.213% | 75.36% | -7.938% | 75.38% | 3.752% | -0.6319% | -10.86% | -12.32% | -12.32% | 59.28% | 62.79% | -6.477% | 9.597%
| Z-Score (Absolute) |  | 1.83 | 0.67 | 3.13 | 1.01 | 0.65 | 0.03 | 0.03 | 1.19 | 0.03 | 0.03 | 3.13 | 0.55 | 3.13 | 0.03 | 0.23 | 0.68 | 0.74 | 1.39 | 2.42 | 2.57 | 0.48 | 0.23
| Analytical results for Tl in OREAS 922b (Certified Value 0.14 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  | 2
| 1 | 1 | 0.19 | 0.1679 | 0.2 | < 5 | 0.12 | 0.16152 | 0.12 | 0.12 | 0.14 | 0.12 | 0.13 | 0.16 | 0.11 | 0.15117 | 0.14 | 0.13 | 0.118 | 0.1 | 0.04 | 0.12 | 0.21 | 0.12 | 0.12 |  | 1
| 1 | 2 | 0.2 | 0.1698 | 0.2 | < 5 | 0.12 | 0.18002 | 0.13 | 0.12 | 0.14 | 0.12 | 0.13 | 0.14 | 0.12 | 0.16085 | 0.13 | 0.13 | 1.101 | 0.1 | 0.05 | 0.13 | 0.23 | 0.12 | 0.13 |  | 14
| 1 | 3 | 0.19 | 0.17 | 0.2 | < 5 | 0.12 | 0.18293 | 0.13 | 0.12 | 0.14 | 0.11 | 0.14 | 0.15 | 0.13 | 0.16941 | 0.14 | 0.13 | 0.116 | 0.1 | 0.05 | 0.12 | 0.22 | 0.12 | 0.13 |  | 16
| 1 | 4 | 0.19 | 0.1757 | 0.2 | < 5 | 0.12 | 0.17293 | 0.13 | 0.14 | 0.14 | 0.11 | 0.13 | 0.14 | 0.12 | 0.17288 | 0.18 | 0.15 | 0.116 | 0.1 | 0.05 | 0.13 | 0.22 | 0.12 | 0.13 |  | 0.13549
| 1 | 5 | 0.19 | 0.1753 | 0.2 | < 5 | 0.12 | 0.15759 | 0.13 | 0.13 | 0.14 | 0.12 | 0.14 | 0.15 | 0.11 | 0.14729 | 0.13 | 0.14 | 0.125 | 0.1 | 0.05 | 0.12 | 0.19 | 0.12 | 0.13 |  | 114
| 1 | 6 | 0.19 | 0.1685 | 0.2 | < 5 | 0.11 | 0.15689 | 0.12 | 0.15 | 0.13 | 0.12 | 0.14 | 0.15 | 0.12 | 0.14889 | 0.16 | 0.14 | 0.117 | 0.1 | 0.05 | 0.13 | 0.23 | 0.12 | 0.12
| Mean |  | 0.19167 | 0.1712 | 0.2 |  | 0.11833 | 0.16865 | 0.12667 | 0.13 | 0.13833 | 0.11667 | 0.135 | 0.14833 | 0.11833 | 0.15841 | 0.14667 | 0.13667 | 0.28217 | 0.1 | 0.048333 | 0.125 | 0.21667 | 0.12 | 0.12667
| Median |  | 0.19 | 0.1699 | 0.2 |  | 0.12 | 0.16722 | 0.13 | 0.125 | 0.14 | 0.12 | 0.135 | 0.15 | 0.12 | 0.15601 | 0.14 | 0.135 | 0.1175 | 0.1 | 0.05 | 0.125 | 0.22 | 0.12 | 0.13
| Std Dev. |  | 0.0040825 | 0.0034246 | 3.0405e-17 |  | 0.0040825 | 0.011516 | 0.005164 | 0.012649 | 0.0040825 | 0.005164 | 0.0054772 | 0.0075277 | 0.0075277 | 0.01098 | 0.019664 | 0.008165 | 0.40116 | 1.5202e-17 | 0.0040825 | 0.0054772 | 0.015055 | 0 | 0.005164
| Rel.Std.Dev. |  | 2.13% | 2% | 1.52e-14% |  | 3.45% | 6.829% | 4.077% | 9.73% | 2.951% | 4.426% | 4.057% | 5.075% | 6.361% | 6.931% | 13.41% | 5.974% | 142.2% | 1.52e-14% | 8.447% | 4.382% | 6.949% | 0% | 4.077%
| PDM3 |  | 41.46% | 26.36% | 47.61% |  | -12.66% | 24.47% | -6.512% | -4.052% | 2.099% | -13.89% | -0.3615% | 9.479% | -12.66% | 16.92% | 8.249% | 0.8686% | 108.3% | -26.19% | -64.33% | -7.742% | 59.91% | -11.43% | -6.512%
| Z-Score (Absolute) |  | 2.02 | 1.27 | NA | 86.92 | 0.67 | 1.18 | 0.37 | 0.25 | 0.06 | 0.74 | 0.06 | 0.43 | 0.67 | 0.8 | 0.37 | 0 | 5.35 | NA | 3.25 | 0.43 | 2.94 | 0.61 | 0.37
Results from laboratories 2.03 and 2.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 922b (Indicative Value 0.098 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  | 3
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g |  | 3
| 1 | 1 | 2.01: 0.0931; 2.02: 0.099809; 2.03: 0.1 |  | 1
| 1 | 2 | 2.01: 0.0906; 2.02: 0.11212; 2.03: 0.1 |  | 5
| 1 | 3 | 2.01: 0.0928; 2.02: 0.10448; 2.03: 0.1 |  | 16
| 1 | 4 | 2.01: 0.0944; 2.02: 0.10086; 2.03: 0.1 |  | 0.09772
| 1 | 5 | 2.01: 0.0942; 2.02: 0.098367; 2.03: 0.1 |  | 11
| 1 | 6 | 2.01: 0.0908; 2.02: 0.099033; 2.03: 0.1
| Mean |  | 2.01: 0.09265; 2.02: 0.10244; 2.03: 0.1
| Median |  | 2.01: 0.09295; 2.02: 0.10033; 2.03: 0.1
| Std Dev. |  | 2.01: 0.0016319; 2.02: 0.0052043; 2.03: 1.5202e-17
| Rel.Std.Dev. |  | 2.01: 1.761%; 2.02: 5.08%; 2.03: 1.52e-14%
| PDM3 |  | 2.01: -5.188%; 2.02: 4.835%; 2.03: 2.333%
| Z-Score (Absolute) |  | 2.01: 2.03; 2.02: 0.67; 2.03: 0
| Analytical results for U in OREAS 922b (Certified Value 2.27 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2.52 | 1.9521 | 2.2 | < 5 | 2.266 | 2.4324 | 2.13 | 2.38 | 2.33 | 2.09 | 2.27 | 2.3 | 2.2 | 2.7163 | 2.45 | 2.18 | 2.43 | 2.2 | 2.28 | 2.31 | 2.19 | 2.1 | 2.1 |  | 1
| 1 | 2 | 2.47 | 1.8824 | 2 | < 5 | 2.365 | 2.5677 | 2.26 | 2.37 | 2.36 | 2.16 | 2.18 | 2.3 | 2.2 | 2.7255 | 1.65 | 2.21 | 2.54 | 2.2 | 2.13 | 2.44 | 2.19 | 2.24 | 2.1 |  | 16
| 1 | 3 | 2.5 | 1.8954 | 2.1 | < 5 | 2.29 | 2.585 | 2.18 | 2.37 | 2.36 | 2.08 | 2.24 | 2.4 | 2.33 | 2.5693 | 2.48 | 2.22 | 2.56 | 2.2 | 2.22 | 2.51 | 2.29 | 2.14 | 2.1 |  | 16
| 1 | 4 | 2.49 | 1.9774 | 2.2 | < 5 | 2.339 | 2.4427 | 2.23 | 2.3 | 2.33 | 2.1 | 2.21 | 2.3 | 2.23 | 2.6454 | 2.56 | 2.17 | 2.56 | 2.3 | 2.16 | 2.46 | 2.26 | 2.24 | 2.1 |  | 2.2736
| 1 | 5 | 2.48 | 1.9616 | 2.2 | < 5 | 2.257 | 2.1148 | 2.24 | 2.37 | 2.31 | 2.13 | 2.21 | 2.4 | 2.28 | 2.5529 | 2.3 | 2.31 | 2.54 | 2.3 | 2.16 | 2.34 | 2.26 | 2.21 | 2.1 |  | 115
| 1 | 6 | 2.5 | 1.9383 | 2.2 | < 5 | 2.289 | 2.3277 | 2.2 | 2.41 | 2.31 | 2.08 | 2.26 | 2.4 | 2.29 | 2.6412 | 2.54 | 2.15 | 2.6 | 2.2 | 2.07 | 2.42 | 2.22 | 2.23 | 2.1
| Mean |  | 2.4933 | 1.9345 | 2.15 |  | 2.301 | 2.4117 | 2.2067 | 2.3667 | 2.3333 | 2.1067 | 2.2283 | 2.35 | 2.255 | 2.6418 | 2.33 | 2.2067 | 2.5383 | 2.2333 | 2.17 | 2.4133 | 2.235 | 2.1933 | 2.1
| Median |  | 2.495 | 1.9452 | 2.2 |  | 2.2895 | 2.4376 | 2.215 | 2.37 | 2.33 | 2.095 | 2.225 | 2.35 | 2.255 | 2.6433 | 2.465 | 2.195 | 2.55 | 2.2 | 2.16 | 2.43 | 2.24 | 2.22 | 2.1
| Std Dev. |  | 0.017512 | 0.037795 | 0.083666 |  | 0.042346 | 0.17376 | 0.047188 | 0.036148 | 0.022509 | 0.032042 | 0.034303 | 0.054772 | 0.053198 | 0.071725 | 0.3456 | 0.056804 | 0.057417 | 0.05164 | 0.072664 | 0.075277 | 0.041352 | 0.059217 | 0
| Rel.Std.Dev. |  | 0.7023% | 1.954% | 3.891% |  | 1.84% | 7.205% | 2.138% | 1.527% | 0.9647% | 1.521% | 1.539% | 2.331% | 2.359% | 2.715% | 14.83% | 2.574% | 2.262% | 2.312% | 3.349% | 3.119% | 1.85% | 2.7% | 0%
| PDM3 |  | 9.666% | -14.91% | -5.435% |  | 1.207% | 6.078% | -2.943% | 4.095% | 2.629% | -7.341% | -1.99% | 3.362% | -0.8167% | 16.19% | 2.482% | -2.943% | 11.65% | -1.77% | -4.555% | 6.147% | -1.696% | -3.529% | -7.634%
| Z-Score (Absolute) |  | 1.69 | 2.27 | 0.75 | 1.74 | 0.33 | 1.11 | 0.34 | 0.79 | 0.56 | 1.05 | 0.19 | 0.67 | 0 | 2.75 | 0.53 | 0.34 | 2.01 | 0.15 | 0.6 | 1.12 | 0.14 | 0.44 | 1.1
| Analytical results for V in OREAS 922b (Certified Value 22.3 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 1
| 1 | 1 | 29 | 23.75 | 19.8 | 18.8 | 23.0007 | 22 | 22 | 24 | 20 | 22 | 26 | 21 | 23.7387 | 24 | 23 | 21 | 22 | 20 | 21.5 | 33 | 20 | 25 |  |  | 1
| 1 | 2 | 28 | 24 | 20.6 | 19.5 | 22.8292 | 20 | 22 | 24 | 20 | 21 | 24 | 20 | 28.265 | 34 | 23 | 20 | 22 | 20 | 22.1 | 35 | 22 | 24 |  |  | 36
| 1 | 3 | 28 | 23.75 | 19.3 | 19.3 | 23.1506 | 20 | 22 | 24 | 20 | 21 | 26 | 21 | 28.3599 | 24 | 22 | 21 | 22 | 21 | 22.5 | 34 | 20 | 24 |  |  | 16
| 1 | 4 | 28 | 23.75 | 20 | 18.8 | 23.3854 | 20 | 21 | 24 | 20 | 21 | 23 | 20 | 26.4556 | 25 | 23 | 21 | 22 | 21 | 21.8 | 34 | 21 | 23 |  |  | 22.2827
| 1 | 5 | 27 | 24.25 | 19.5 | 19.2 | 21.6925 | 20 | 22 | 24 | 20 | 21 | 26 | 20 | 22.7116 | 22 | 21 | 20 | 22 | 20 | 22.4 | 35 | 21 | 23 |  |  | 116
| 1 | 6 | 27 | 24 | 20.1 | 19.7 | 21.8374 | 21 | 22 | 24 | 20 | 22 | 25 | 21 | 23.5913 | 24 | 23 | 20 | 22 | 20 | 22.4 | 39 | 21 | 23
| Mean |  | 27.8333 | 23.9167 | 19.8833 | 19.2167 | 22.6493 | 20.5 | 21.8333 | 24 | 20 | 21.3333 | 25 | 20.5 | 25.5203 | 25.5 | 22.5 | 20.5 | 22 | 20.3333 | 22.1167 | 35 | 20.8333 | 23.6667
| Median |  | 28 | 23.875 | 19.9 | 19.25 | 22.9149 | 20 | 22 | 24 | 20 | 21 | 25.5 | 20.5 | 25.0971 | 24 | 23 | 20.5 | 22 | 20 | 22.25 | 34.5 | 21 | 23.5
| Std Dev. |  | 0.75277 | 0.20412 | 0.46224 | 0.3656 | 0.71043 | 0.83666 | 0.40825 | 0 | 0 | 0.5164 | 1.2649 | 0.54772 | 2.5002 | 4.2778 | 0.83666 | 0.54772 | 0 | 0.5164 | 0.39707 | 2.0976 | 0.75277 | 0.8165
| Rel.Std.Dev. |  | 2.705% | 0.8535% | 2.325% | 1.903% | 3.137% | 4.081% | 1.87% | 0% | 0% | 2.421% | 5.06% | 2.672% | 9.797% | 16.78% | 3.718% | 2.672% | 0% | 2.54% | 1.795% | 5.993% | 3.613% | 3.45%
| PDM3 |  | 24.91% | 7.333% | -10.77% | -13.76% | 1.645% | -8% | -2.017% | 7.707% | -10.24% | -4.26% | 12.19% | -8% | 14.53% | 14.44% | 0.9753% | -8% | -1.269% | -8.748% | -0.745% | 57.07% | -6.504% | 6.211%
| Z-Score (Absolute) |  | 2.34 | 0.75 | 0.88 | 1.15 | 0.24 | 0.63 | 0.09 | 0.79 | 0.83 | 0.29 | 1.19 | 0.63 | 1.4 | 1.39 | 0.18 | 0.63 | 0.02 | 0.7 | 0.02 | 5.24 | 0.5 | 0.65
| Analytical results for W in OREAS 922b (Certified Value 0.75 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 0.7 | 0.55931 | 1.12 | < 5 | 0.73 | < 1 | 0.74 | 0.86 | 0.57 | 0.74 | 0.77 | 0.8 | 0.78 | 0.83917 | 0.73 | 0.66 | 0.5 | 0.92 | 1.1 | 0.75 | 0.72 | 0.4 |  |  | 1
| 1 | 2 | 0.7 | 0.51129 | 1.1 | < 5 | 0.74 | < 1 | 0.74 | 0.92 | 0.58 | 0.69 | 0.75 | 0.8 | 0.79 | 0.82425 | 0.9 | 0.71 | 0.5 | 1.12 | 1.1 | 0.9 | 0.72 | 0.4 |  |  | 17
| 1 | 3 | 0.7 | 0.53677 | 1.09 | < 5 | 0.76 | < 1 | 0.72 | 0.91 | 0.57 | 0.69 | 0.74 | 0.8 | 0.77 | 0.83645 | 0.77 | 0.77 | 0.5 | 0.99 | 0.9 | 0.8 | 0.71 | 0.4 |  |  | 16
| 1 | 4 | 0.7 | 0.6515 | 1.02 | < 5 | 0.76 | < 1 | 0.71 | 0.89 | 0.57 | 0.7 | 0.76 | 0.7 | 0.78 | 0.8921 | 1.05 | 0.72 | 0.6 | 1.06 | 1 | 0.75 | 0.73 | 0.4 |  |  | 0.74785
| 1 | 5 | 0.7 | 0.5871 | 1.05 | < 5 | 0.77 | < 1 | 0.76 | 0.93 | 0.58 | 0.82 | 0.8 | 0.9 | 0.77 | 0.91361 | 0.61 | 0.74 | 0.5 | 1.05 | 1 | 0.8 | 0.74 | 0.4 |  |  | 117
| 1 | 6 | 0.7 | 0.60803 | 1.07 | < 5 | 0.78 | < 1 | 0.7 | 0.93 | 0.57 | 0.68 | 0.79 | 0.9 | 0.78 | 0.89248 | 0.94 | 0.71 | 0.5 | 1.04 | 1.1 | 0.75 | 0.73 | 0.5
| Mean |  | 0.7 | 0.57567 | 1.075 |  | 0.75667 |  | 0.72833 | 0.90667 | 0.57333 | 0.72 | 0.76833 | 0.81667 | 0.77833 | 0.86634 | 0.83333 | 0.71833 | 0.51667 | 1.03 | 1.0333 | 0.79167 | 0.725 | 0.41667
| Median |  | 0.7 | 0.5732 | 1.08 |  | 0.76 |  | 0.73 | 0.915 | 0.57 | 0.695 | 0.765 | 0.8 | 0.78 | 0.86563 | 0.835 | 0.715 | 0.5 | 1.045 | 1.05 | 0.775 | 0.725 | 0.4
| Std Dev. |  | 1.2162e-16 | 0.050702 | 0.036194 |  | 0.018619 |  | 0.022286 | 0.027325 | 0.005164 | 0.053292 | 0.023166 | 0.075277 | 0.0075277 | 0.037374 | 0.15958 | 0.03656 | 0.040825 | 0.068118 | 0.08165 | 0.058452 | 0.010488 | 0.040825
| Rel.Std.Dev. |  | 1.737e-14% | 8.808% | 3.367% |  | 2.461% |  | 3.06% | 3.014% | 0.9007% | 7.402% | 3.015% | 9.218% | 0.9672% | 4.314% | 19.15% | 5.09% | 7.902% | 6.613% | 7.902% | 7.383% | 1.447% | 9.798%
| PDM3 |  | -6.398% | -23.02% | 43.75% |  | 1.179% |  | -2.609% | 21.24% | -23.34% | -3.724% | 2.739% | 9.202% | 4.077% | 15.84% | 11.43% | -3.946% | -30.91% | 37.73% | 38.17% | 5.859% | -3.055% | -44.28%
| Z-Score (Absolute) |  | NA | 2 | 3.18 | 17.96 | 0.12 | 2.78 | 0.41 | 1.44 | 2.02 | 0.5 | 0 | NA | 0.1 | 1.02 | 0.67 | 0.52 | NA | 2.71 | NA | 0.24 | 0.45 | NA
Results from laboratories 2.01, 2.12, 2.19, 2.21 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 922b (Certified Value 7.29 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 7.93 | 7.9117 | 7.5 | 6.36 | 7.5563 | 7.23 | 7.55 | 7.99 | 6.91 | 7.42 | 7.67 | 7.23 | 8.2045 | 7.35 | 6.39 | 7.16 | 6.7 | 7.32 | 6.5 | [5 empty] | 1
| 1 | 2 | 8.01 | 7.7221 | 7.9 | 6.43 | 7.4124 | 6.8 | 7.29 | 8.24 | 6.92 | 7.24 | 7.79 | 7.1 | 8.2599 | 10.8 | 6.86 | 7.5 | 6.6 | 7.22 | 6.5 | [5 empty] | 18
| 1 | 3 | 7.95 | 7.7428 | 7 | 6.34 | 7.497 | 6.79 | 7.64 | 7.96 | 6.73 | 7.27 | 7.61 | 7.17 | 8.297 | 7.47 | 6.71 | 7.52 | 6.9 | 7.16 | 6.5 | [5 empty] | 16
| 1 | 4 | 8.06 | 7.729 | 7.3 | 6.39 | 7.5156 | 6.84 | 7.47 | 7.94 | 6.94 | 7.37 | 7.41 | 7.09 | 8.4211 | 7.54 | 6.68 | 7.21 | 6.8 | 7.04 | 6.3 | [5 empty] | 7.2863
| 1 | 5 | 7.97 | 7.6579 | 7.4 | 6.39 | 7.3976 | 7.06 | 7.68 | 7.94 | 6.86 | 7.39 | 8.12 | 6.9 | 7.8978 | 6.89 | 6.88 | 7.44 | 6.8 | 7.26 | 6.2 | [5 empty] | 118
| 1 | 6 | 7.89 | 7.7825 | 7.4 | 6.5 | 7.2751 | 6.94 | 7.64 | 7.89 | 7 | 7.58 | 7.83 | 7.22 | 7.9648 | 7.46 | 6.48 | 7.35 | 6.6 | 7.38 | 6.3
| Mean |  | 7.9683 | 7.7577 | 7.4167 | 6.4017 | 7.4423 | 6.9433 | 7.545 | 7.9933 | 6.8933 | 7.3783 | 7.7383 | 7.1183 | 8.1742 | 7.9183 | 6.6667 | 7.3633 | 6.7333 | 7.23 | 6.3833
| Median |  | 7.96 | 7.7359 | 7.4 | 6.39 | 7.4547 | 6.89 | 7.595 | 7.95 | 6.915 | 7.38 | 7.73 | 7.135 | 8.2322 | 7.465 | 6.695 | 7.395 | 6.75 | 7.24 | 6.4
| Std Dev. |  | 0.060139 | 0.085539 | 0.29269 | 0.057067 | 0.10217 | 0.17351 | 0.14625 | 0.12517 | 0.092014 | 0.12123 | 0.23903 | 0.12189 | 0.20226 | 1.431 | 0.19816 | 0.15108 | 0.12111 | 0.1205 | 0.13292
| Rel.Std.Dev. |  | 0.7547% | 1.103% | 3.946% | 0.8914% | 1.373% | 2.499% | 1.938% | 1.566% | 1.335% | 1.643% | 3.089% | 1.712% | 2.474% | 18.07% | 2.972% | 2.052% | 1.799% | 1.667% | 2.082%
| PDM3 |  | 9.36% | 6.469% | 1.789% | -12.14% | 2.141% | -4.707% | 3.55% | 9.703% | -5.394% | 1.263% | 6.204% | -2.306% | 12.19% | 8.674% | -8.504% | 1.057% | -7.589% | -0.773% | -12.39%
| Z-Score (Absolute) |  | 0.91 | 0.59 | 0.06 | 1.51 | 0.1 | 0.67 | 0.26 | 0.95 | 0.75 | 0 | 0.56 | 0.4 | 1.23 | 0.84 | 1.1 | 0.02 | 1 | 0.23 | 1.54
| Analytical results for Yb in OREAS 922b (Certified Value 0.71 ppm) | [25 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | [5 empty] | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.21; 2.08: 2.22 | [5 empty] | ppm
|  |  | 2.01: AR*MS; 2.02: AR*OES; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*OES/MS; 2.08: AR*OES/MS | [5 empty] | 2
|  |  | 2.01: 0.5g; 2.02: 0.25g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 0.25g; 2.08: 0.2g | [5 empty] | 3
| 1 | 1 | 2.01: 0.6996; 2.02: < 5; 2.04: 0.74993; 2.05: 0.8; 2.06: 0.8; 2.07: 0.6; 2.08: 0.6 | [5 empty] | 1
| 1 | 2 | 2.01: 0.6835; 2.02: < 5; 2.04: 0.79804; 2.05: 0.7; 2.06: 0.7; 2.07: 0.6; 2.08: 0.7 | [5 empty] | 19
| 1 | 3 | 2.01: 0.6971; 2.02: < 5; 2.04: 0.82376; 2.05: 0.7; 2.06: 0.8; 2.07: 0.7; 2.08: 0.7 | [5 empty] | 16
| 1 | 4 | 2.01: 0.709; 2.02: < 5; 2.04: 0.7451; 2.05: 0.7; 2.06: 0.8; 2.07: 0.6; 2.08: 0.7 | [5 empty] | 0.70815
| 1 | 5 | 2.01: 0.7114; 2.02: < 5; 2.04: 0.75931; 2.05: 0.7; 2.06: 0.7; 2.07: 0.6; 2.08: 0.7 | [5 empty] | 119
| 1 | 6 | 2.01: 0.7059; 2.02: < 5; 2.04: 0.71062; 2.05: 0.8; 2.06: 0.8; 2.07: 0.6; 2.08: 0.6
| Mean |  | 2.01: 0.70108; 2.04: 0.76446; 2.05: 0.73333; 2.06: 0.76667; 2.07: 0.61667; 2.08: 0.66667
| Median |  | 2.01: 0.70275; 2.04: 0.75462; 2.05: 0.7; 2.06: 0.8; 2.07: 0.6; 2.08: 0.7
| Std Dev. |  | 2.01: 0.010194; 2.04: 0.040386; 2.05: 0.05164; 2.06: 0.05164; 2.07: 0.040825; 2.08: 0.05164
| Rel.Std.Dev. |  | 2.01: 1.454%; 2.04: 5.283%; 2.05: 7.042%; 2.06: 6.736%; 2.07: 6.62%; 2.08: 7.746%
| PDM3 |  | 2.01: -0.9974%; 2.04: 7.952%; 2.05: 3.557%; 2.06: 8.264%; 2.07: -12.92%; 2.08: -5.857%
| Z-Score (Absolute) |  | 2.01: 0.65; 2.02: 35.74; 2.04: 0.63; 2.05: 0; 2.06: 0.67; 2.07: 2.36; 2.08: 1.35
| Analytical results for Zn in OREAS 922b (Certified Value 66 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 1
| 1 | 1 | 66 | 61.4768 | 59.6 | 81.8 | 74.6 | 69.4696 | 71 | 63 | 68 | 63 | 66 | 63.6 | 66 | 49.1925 | 69 | 71.4 | 70 | 65 | 66 | 74 | 63.9 | 66 | 65 | 68 | 1
| 1 | 2 | 66 | 60.0992 | 59.4 | 84.9 | 67.2 | 66.3884 | 67 | 63 | 68 | 64 | 63 | 62.7 | 67 | 51.8384 | 74 | 73.3 | 71 | 68 | 66 | 73 | 65.8 | 66 | 66 | 66 | 40
| 1 | 3 | 66 | 62.1936 | 59.5 | 86.8 | 68.2 | 67.5977 | 65 | 62 | 67 | 63 | 63 | 60.6 | 67 | 55.1403 | 72 | 72.4 | 70 | 68 | 65 | 75 | 66.8 | 65 | 65 | 66 | 16
| 1 | 4 | 67 | 60.8048 | 59.6 | 88.7 | 67.3 | 66.4444 | 67 | 61 | 67 | 63 | 63 | 61.7 | 67 | 56.0219 | 81 | 71.8 | 71 | 66 | 64 | 74 | 66.1 | 66 | 65 | 65 | 65.8353
| 1 | 5 | 66 | 65.1056 | 59.5 | 89.7 | 67.9 | 66.9239 | 64 | 63 | 67 | 64 | 64 | 65.1 | 67 | 46.8686 | 64 | 71.5 | 72 | 69 | 66 | 75 | 65.1 | 66 | 66 | 64 | 120
| 1 | 6 | 66 | 63.3024 | 59.5 | 81.2 | 69.1 | 64.6038 | 66 | 62 | 68 | 62 | 64 | 63.8 | 69 | 47.8653 | 80 | 73.9 | 68 | 65 | 62 | 75 | 64.1 | 65 | 67 | 64
| Mean |  | 66.1667 | 62.1637 | 59.5167 | 85.5167 | 69.05 | 66.9046 | 66.6667 | 62.3333 | 67.5 | 63.1667 | 63.8333 | 62.9167 | 67.1667 | 51.1545 | 73.3333 | 72.3833 | 70.3333 | 66.8333 | 64.8333 | 74.3333 | 65.3 | 65.6667 | 65.6667 | 65.5
| Median |  | 66 | 61.8352 | 59.5 | 85.85 | 68.05 | 66.6842 | 66.5 | 62.5 | 67.5 | 63 | 63.5 | 63.15 | 67 | 50.5154 | 73 | 72.1 | 70.5 | 67 | 65.5 | 74.5 | 65.45 | 66 | 65.5 | 65.5
| Std Dev. |  | 0.40825 | 1.8184 | 0.075277 | 3.5244 | 2.8048 | 1.602 | 2.4221 | 0.8165 | 0.54772 | 0.75277 | 1.169 | 1.6068 | 0.98319 | 3.8232 | 6.5013 | 1.0226 | 1.3663 | 1.7224 | 1.6021 | 0.8165 | 1.1472 | 0.5164 | 0.8165 | 1.5166
| Rel.Std.Dev. |  | 0.617% | 2.925% | 0.1265% | 4.121% | 4.062% | 2.394% | 3.633% | 1.31% | 0.8114% | 1.192% | 1.831% | 2.554% | 1.464% | 7.474% | 8.865% | 1.413% | 1.943% | 2.577% | 2.471% | 1.098% | 1.757% | 0.7864% | 1.243% | 2.315%
| PDM3 |  | 0.5033% | -5.577% | -9.598% | 29.89% | 4.883% | 1.624% | 1.263% | -5.319% | 2.529% | -4.054% | -3.041% | -4.433% | 2.022% | -22.3% | 11.39% | 9.946% | 6.832% | 1.516% | -1.522% | 12.91% | -0.8131% | -0.2562% | -0.2562% | -0.5093%
| Z-Score (Absolute) |  | 0.07 | 1.05 | 1.79 | 5.47 | 0.87 | 0.28 | 0.21 | 1 | 0.44 | 0.77 | 0.58 | 0.84 | 0.35 | 4.12 | 2.07 | 1.8 | 1.23 | 0.26 | 0.3 | 2.35 | 0.17 | 0.07 | 0.07 | 0.12
| Analytical results for Zr in OREAS 922b (Certified Value 11.2 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 1
| 1 | 1 | 12.5 | 14.2 | 12.54 | 12.0729 | 8.8 | 10.7 | 11 | 10.8 | 11.4 | 10.4 | 12.2 | 11.6738 | 19 | 10.4 | 12.6 | 7.5 | 10.3 | 9.6 | [6 empty] | 1
| 1 | 2 | 12.6 | 14.6 | 12.68 | 11.8523 | 9.6 | 10.9 | 11.1 | 10.9 | 11 | 10.4 | 12 | 11.8149 | 19.4 | 10.2 | 11.8 | 9.6 | 10.7 | 9.7 | [6 empty] | 41
| 1 | 3 | 12.6 | 13.6 | 12.66 | 11.9935 | 9.5 | 10.8 | 10.7 | 10.9 | 11 | 10.6 | 12.3 | 12.5725 | 19.7 | 10.4 | 11.8 | 10 | 10.4 | 9.7 | [6 empty] | 16
| 1 | 4 | 12.6 | 13.2 | 12.56 | 11.9881 | 9.5 | 11 | 10.8 | 11.4 | 10.9 | 9.7 | 12.2 | 12.4161 | 18.5 | 10.3 | 11.3 | 9.9 | 10.9 | 9.4 | [6 empty] | 11.2207
| 1 | 5 | 12.7 | 13.4 | 12.77 | 11.8455 | 9.9 | 11.1 | 10.8 | 10.9 | 11.3 | 10.9 | 11.7 | 11.3415 | 18.5 | 9.9 | 12.5 | 8.6 | 10.8 | 9.3 | [6 empty] | 121
| 1 | 6 | 12.5 | 14.3 | 12.67 | 11.6886 | 9.5 | 11.2 | 10.8 | 11.4 | 11.2 | 10.7 | 12.3 | 11.8748 | 19 | 10.1 | 12 | 9.2 | 10.7 | 9.6
| Mean |  | 12.5833 | 13.8833 | 12.6467 | 11.9068 | 9.4667 | 10.95 | 10.8667 | 11.05 | 11.1333 | 10.45 | 12.1167 | 11.9489 | 19.0167 | 10.2167 | 12 | 9.1333 | 10.6333 | 9.55
| Median |  | 12.6 | 13.9 | 12.665 | 11.9202 | 9.5 | 10.95 | 10.8 | 10.9 | 11.1 | 10.5 | 12.2 | 11.8448 | 19 | 10.25 | 11.9 | 9.4 | 10.7 | 9.6
| Std Dev. |  | 0.075277 | 0.56006 | 0.084774 | 0.13859 | 0.36148 | 0.18708 | 0.15055 | 0.27386 | 0.19664 | 0.41352 | 0.23166 | 0.46371 | 0.47924 | 0.19408 | 0.4858 | 0.95009 | 0.23381 | 0.16432
| Rel.Std.Dev. |  | 0.5982% | 4.034% | 0.6703% | 1.164% | 3.818% | 1.709% | 1.385% | 2.478% | 1.766% | 3.957% | 1.912% | 3.881% | 2.52% | 1.9% | 4.048% | 10.4% | 2.199% | 1.721%
| PDM3 |  | 12.14% | 23.73% | 12.71% | 6.114% | -15.63% | -2.413% | -3.155% | -1.522% | -0.7789% | -6.869% | 7.985% | 6.49% | 69.48% | -8.948% | 6.945% | -18.6% | -5.235% | -14.89%
| Z-Score (Absolute) |  | 1.13 | 2.11 | 1.18 | 0.62 | 1.23 | 0.11 | 0.17 | 0.03 | 0.03 | 0.49 | 0.78 | 0.65 | 5.99 | 0.66 | 0.69 | 1.48 | 0.35 | 1.17
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1042: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1188: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Ag in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ag | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 8
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 14
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 922b (Certified Value 12.4 wt.%) | [17 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 12.43 | 12.301 | 12.14 | 11.94 | 12.4102 | 12.45 | 12.25 | 12.5 | 12.4 | 12.2 | 12.28 | 12.09 | 12.4 | 11.62 | 12.36 |  | 1
| 1 | 2 | 12.4 | 12.281 | 12.37 | 12.4 | 12.3059 | 12.45 | 12.15 | 12.35 | 12.45 | 12.37 | 12.29 | 11.8 | 12.41 | 11.639 | 12.41 |  | 0
| 1 | 3 | 12.44 | 12.386 | 12.21 | 12.74 | 12.3685 | 12.4 | 12.4 | 12.55 | 12.5 | 12.97 | 12.28 | 11.99 | 12.41 | 11.677 | 12.46 |  | 16
| 1 | 4 | 12.37 | 12.241 | 12.27 | 12.65 | 12.38 | 12.3 | 12.4 | 12.5 | 12.25 | 12.58 | 12.29 | 11.94 | 12.41 | 11.564 | 12.5 |  | 12.3955
| 1 | 5 | 12.44 | 12.252 | 12.11 | 12.79 | 12.5004 | 12.3 | 12.35 | 12.5 | 12.45 | 12.88 | 12.27 | 12.01 | 12.45 | 11.696 | 12.48 |  | 123
| 1 | 6 | 12.42 | 12.302 | 12.43 | 12.69 | 12.3746 | 12.45 | 12.55 | 12.45 | 12.4 | 13.22 | 12.27 | 11.99 | 12.44 | 11.583 | 12.46
| Mean |  | 12.4167 | 12.2938 | 12.255 | 12.535 | 12.3899 | 12.3917 | 12.35 | 12.475 | 12.4083 | 12.7033 | 12.28 | 11.97 | 12.42 | 11.6298 | 12.445
| Median |  | 12.425 | 12.291 | 12.24 | 12.67 | 12.3773 | 12.425 | 12.375 | 12.5 | 12.425 | 12.73 | 12.28 | 11.99 | 12.41 | 11.6295 | 12.46
| Std Dev. |  | 0.027325 | 0.051612 | 0.12677 | 0.32142 | 0.063975 | 0.073598 | 0.13784 | 0.06892 | 0.08612 | 0.38702 | 0.0089443 | 0.09654 | 0.02 | 0.051596 | 0.051284
| Rel.Std.Dev. |  | 0.2201% | 0.4198% | 1.034% | 2.564% | 0.5163% | 0.5939% | 1.116% | 0.5525% | 0.6941% | 3.047% | 0.07284% | 0.8065% | 0.161% | 0.4437% | 0.4121%
| PDM3 |  | 0.1711% | -0.8198% | -1.133% | 1.126% | -0.04451% | -0.03055% | -0.3667% | 0.6417% | 0.1039% | 2.484% | -0.9314% | -3.432% | 0.198% | -6.177% | 0.3997%
| Z-Score (Absolute) |  | 0.2 | 0.79 | 1.11 | 1.16 | 0.01 | 0 | 0.34 | 0.67 | 0.13 | 2.52 | 0.9 | 3.41 | 0.23 | 6.16 | 0.43
| Analytical results for As in OREAS 922b (Indicative Value 83 ppm) | [17 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.19 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | 200 | 50 | < 10 | 55 | [8 empty] | 1
| 1 | 2 | < 100 | 200 | < 100 | < 100 | 100 | 50 | < 10 | 57 | [8 empty] | 9
| 1 | 3 | < 100 | 100 | 100 | 100 | 100 | 50 | < 10 | 53 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | 100 | 100 | 50 | < 10 | 50 | [8 empty] | 82.5
| 1 | 5 | < 100 | 100 | < 100 | < 100 | 100 | 50 | < 10 | 55 | [8 empty] | 15
| 1 | 6 | < 100 | 200 | 100 | 100 | 200 | 50 | < 10 | 50
| Mean |  |  | 150 | 100 | 100 | 133.3333 | 50 |  | 53.3333
| Median |  |  | 150 | 100 | 100 | 100 | 50 |  | 54
| Std Dev. |  |  | 57.735 | 0 | 0 | 51.6398 | 0 |  | 2.8752
| Rel.Std.Dev. |  |  | 38.49% | 0% | 0% | 38.73% | 0% |  | 5.391%
| PDM3 |  |  | 81.82% | 21.21% | 21.21% | 61.62% | -39.39% |  | -35.35%
| Z-Score (Absolute) |  | 0.54 | 3.06 | 0.36 | 0.81 | 3.96 | 0.54 | 2.97 | 0.36
| Analytical results for Au in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Au | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 10
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 16
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 922b (Certified Value 0.221 wt.%) | [17 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 | 2.25 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.2222 | 0.23 | 0.22162 | 0.21 | 0.21 | 0.22 | 0.23 | 0.21 | 0.24 | 0.2233 | 0.201 | 0.22 | 0.23826 |  |  |  | 1
| 1 | 2 | 0.2222 | 0.23 | 0.21685 | 0.21 | 0.21 | 0.23 | 0.23 | 0.22 | 0.22 | 0.2155 | 0.1898 | 0.22 | 0.24005 |  |  |  | 0
| 1 | 3 | 0.2333 | 0.22 | 0.21595 | 0.2 | 0.23 | 0.22 | 0.23 | 0.21 | 0.22 | 0.2222 | 0.201 | 0.21 | 0.23938 |  |  |  | 16
| 1 | 4 | 0.2389 | 0.21 | 0.21879 | 0.2 | 0.22 | 0.23 | 0.22 | 0.21 | 0.24 | 0.2233 | 0.1898 | 0.22 | 0.23781 |  |  |  | 0.22129
| 1 | 5 | 0.2367 | 0.22 | 0.21781 | 0.21 | 0.22 | 0.23 | 0.22 | 0.22 | 0.25 | 0.2311 | 0.201 | 0.23 | 0.2396 |  |  |  | 124
| 1 | 6 | 0.2423 | 0.21 | 0.21873 | 0.2 | 0.23 | 0.22 | 0.23 | 0.21 | 0.26 | 0.2222 | 0.201 | 0.21 | 0.23938
| Mean |  | 0.2326 | 0.22 | 0.21829 | 0.205 | 0.22 | 0.225 | 0.22667 | 0.21333 | 0.23833 | 0.22293 | 0.19727 | 0.21833 | 0.23908
| Median |  | 0.235 | 0.22 | 0.21827 | 0.205 | 0.22 | 0.225 | 0.23 | 0.21 | 0.24 | 0.22275 | 0.201 | 0.22 | 0.23938
| Std Dev. |  | 0.008572 | 0.0089443 | 0.0019679 | 0.0054772 | 0.0089443 | 0.0054772 | 0.005164 | 0.005164 | 0.016021 | 0.0049657 | 0.0057837 | 0.0075277 | 0.00085751
| Rel.Std.Dev. |  | 3.685% | 4.066% | 0.9015% | 2.672% | 4.066% | 2.434% | 2.278% | 2.421% | 6.722% | 2.227% | 2.932% | 3.448% | 0.3587%
| PDM3 |  | 5.109% | -0.5846% | -1.356% | -7.363% | -0.5846% | 1.675% | 2.428% | -3.597% | 7.7% | 0.7409% | -10.86% | -1.338% | 8.037%
| Z-Score (Absolute) |  | 1.27 | 0 | 0.17 | 1.52 | 0 | 0.51 | 0.67 | 0.67 | 1.85 | 0.3 | 2.3 | 0.17 | 1.93
| Analytical results for Bi in OREAS 922b (Certified Value < 100 ppm) | [17 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | 100 | < 10 | 32 | [9 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 34 | [9 empty] | 29
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 36 | [9 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 33 | [9 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 35 | [9 empty] | 125
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 33
| Mean | [5 empty] | 100 |  | 33.8333
| Median | [5 empty] | 100 |  | 33.5
| Std Dev. | [7 empty] | 1.472
| Rel.Std.Dev. | [7 empty] | 4.351%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 922b (Certified Value 0.054 wt.%) | [17 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.05 | 0.025 | 0.06 | 0.24 | 0.028948 | 0.05 | 0.05 | 0.06 | 0.05 | 0.09 | 0.06 | < 0.01 | 0.06 | 0.042 | 0.05 |  | 1
| 1 | 2 | 0.06 | 0.023 | 0.06 | 0.3 | 0.026053 | 0.06 | 0.05 | 0.06 | 0.06 | 0.09 | 0.06 | < 0.01 | 0.06 | 0.042 | 0.05 |  | 0
| 1 | 3 | 0.05 | 0.024 | 0.05 | 0.31 | 0.027025 | 0.05 | 0.04 | 0.06 | 0.06 | 0.09 | 0.06 | < 0.01 | 0.06 | 0.042 | 0.05 |  | 16
| 1 | 4 | 0.05 | 0.028 | 0.06 | 0.33 | 0.027987 | 0.05 | 0.07 | 0.06 | 0.05 | 0.09 | 0.06 | < 0.01 | 0.06 | 0.042 | 0.05 |  | 0.054497
| 1 | 5 | 0.05 | 0.025 | 0.06 | 0.26 | 0.030879 | 0.05 | 0.04 | 0.06 | 0.05 | 0.1 | 0.06 | < 0.01 | 0.06 | 0.056 | 0.05 |  | 126
| 1 | 6 | 0.06 | 0.025 | 0.06 | 0.16 | 0.031839 | 0.05 | 0.05 | 0.06 | 0.06 | 0.1 | 0.06 | < 0.01 | 0.06 | 0.056 | 0.05
| Mean |  | 0.053333 | 0.025 | 0.058333 | 0.26667 | 0.028788 | 0.051667 | 0.05 | 0.06 | 0.055 | 0.093333 | 0.06 |  | 0.06 | 0.046667 | 0.05
| Median |  | 0.05 | 0.025 | 0.06 | 0.28 | 0.028467 | 0.05 | 0.05 | 0.06 | 0.055 | 0.09 | 0.06 |  | 0.06 | 0.042 | 0.05
| Std Dev. |  | 0.005164 | 0.0016733 | 0.0040825 | 0.06186 | 0.0022332 | 0.0040825 | 0.010954 | 0 | 0.0054772 | 0.005164 | 0 |  | 0 | 0.0072296 | 7.6012e-18
| Rel.Std.Dev. |  | 9.682% | 6.693% | 6.999% | 23.2% | 7.757% | 7.902% | 21.91% | 0% | 9.959% | 5.533% | 0% |  | 0% | 15.49% | 1.52e-14%
| PDM3 |  | -2.136% | -54.13% | 7.039% | 389.3% | -47.17% | -5.194% | -8.252% | 10.1% | 0.9224% | 71.26% | 10.1% |  | 10.1% | -14.37% | -8.252%
| Z-Score (Absolute) |  | 0 | 2.87 | 0.51 | 21.58 | 2.48 | 0.17 | 0.34 | 0.67 | 0.17 | 4.05 | 0.67 | 4.89 | 0.67 | 0.68 | 0.34
| Analytical results for Cd in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.20 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 11
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 17
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 922b (Indicative Value 79 ppm) | [17 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 80 | 80 | < 80 | 80 | 35 | [11 empty] | 1
| 1 | 2 | < 80 | < 80 | 80 | 160 | 30 | [11 empty] | 12
| 1 | 3 | < 80 | 160 | < 80 | 80 | 30 | [11 empty] | 16
| 1 | 4 | < 80 | 80 | 80 | 160 | 30 | [11 empty] | 78.5309
| 1 | 5 | < 80 | 80 | < 80 | 80 | 27 | [11 empty] | 18
| 1 | 6 | < 80 | 160 | 80 | 160 | 29
| Mean |  |  | 112 | 80 | 120 | 30.1667
| Median |  |  | 80 | 80 | 120 | 30
| Std Dev. |  |  | 43.8178 | 0 | 43.8178 | 2.6394
| Rel.Std.Dev. |  |  | 39.12% | 0% | 36.51% | 8.75%
| PDM3 |  |  | 42.62% | 1.871% | 52.81% | -61.59%
| Z-Score (Absolute) |  | 0.44 | 0.89 | 0 | 1.33 | 0.67
| Analytical results for Cl in OREAS 922b (Indicative Value 19.3 ppm) | [17 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.20 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 30 | 30 | [14 empty] | 1
| 1 | 2 | < 30 | 14 | [14 empty] | 13
| 1 | 3 | < 30 | 19 | [14 empty] | 16
| 1 | 4 | < 30 | 21 | [14 empty] | 19.3333
| 1 | 5 | < 30 | 18 | [14 empty] | 19
| 1 | 6 | < 30 | 14
| Mean |  |  | 19.3333
| Median |  |  | 18.5
| Std Dev. |  |  | 5.9217
| Rel.Std.Dev. |  |  | 30.63%
| PDM3 |  |  | 1.554e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 922b (Certified Value < 100 ppm) | [17 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.19 | 2.20 | 2.25 | [7 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | < 100 | 37 | < 10 | 100 | 19 | [7 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | < 100 | < 37 | < 10 | 100 | 16 | [7 empty] | 32
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | < 100 | 37 | < 10 | 100 | 18 | [7 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | < 100 | < 37 | < 10 | 100 | 21 | [7 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | < 100 | 51 | < 10 | 100 | 17 | [7 empty] | 127
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | 44 | < 10 | 100 | 14
| Mean | [6 empty] | 42.25 |  | 100 | 17.5
| Median | [6 empty] | 40.5 |  | 100 | 17.5
| Std Dev. | [6 empty] | 6.702 |  | 0 | 2.429
| Rel.Std.Dev. | [6 empty] | 15.86% |  | 0% | 13.88%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Cr2O3 in OREAS 922b (Certified Value 141 ppm) | [17 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 | 2.25 |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 0
0
| 1 | 1 | < 100 | < 100 | 144.7382 | < 150 | 150 | 150 | 150 | 200 | 100 | < 50 | 117 | 146 | 400 | 137.4 |  |  | 1
| 1 | 2 | < 100 | < 100 | 135.0889 | < 150 | 150 | 150 | 150 | 200 | 100 | < 50 | 102 | 146 | < 100 | 128.6 |  |  | 0
| 1 | 3 | < 100 | < 100 | 125.4715 | < 150 | 150 | 150 | 150 | 200 | 100 | < 50 | 102 | 146 | < 100 | 131.5 |  |  | 16
| 1 | 4 | < 100 | < 100 | 125.4603 | < 150 | 150 | 150 | 150 | 200 | 100 | < 50 | 102 | 146 | < 100 | 133 |  |  | 141.3358
| 1 | 5 | < 100 | < 100 | 135.0949 | < 150 | 150 | 150 | 150 | 100 | 100 | < 50 | 117 | 146 | < 100 | 130.1 |  |  | 128
| 1 | 6 | < 100 | < 100 | 125.4252 | < 150 | 150 | 150 | 150 | 200 | 100 | < 50 | 117 | 146 | < 100 | 128.6
| Mean |  |  |  | 131.8798 |  | 150 | 150 | 150 | 183.3333 | 100 |  | 109.5 | 146 | 400 | 131.5333
| Median |  |  |  | 130.2802 |  | 150 | 150 | 150 | 200 | 100 |  | 109.5 | 146 | 400 | 130.8
| Std Dev. |  |  |  | 7.8729 |  | 0 | 0 | 0 | 40.8248 | 0 |  | 8.2158 | 0 |  | 3.3417
| Rel.Std.Dev. |  |  |  | 5.97% |  | 0% | 0% | 0% | 22.27% | 0% |  | 7.503% | 0% |  | 2.541%
| PDM3 |  |  |  | -6.69% |  | 6.13% | 6.13% | 6.13% | 29.71% | -29.25% |  | -22.52% | 3.3% | 183% | -6.936%
| Z-Score (Absolute) |  | 1.86 | 1.86 | 0.3 | 1.25 | 0.67 | 0.67 | 0.67 | NA | NA | 2.52 | 0.29 | 0.67 | 0.32 | 0.29
Results from laboratories 2.12 and 2.16 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 14
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 20
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Cu in OREAS 922b (Certified Value 0.22 wt.%) | [17 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.19 | 2.20 | 2.25 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 0.26 | 0.22508 | 0.224 | 0.222 | 0.22 | 0.222 | 0.2413 | 0.21 | 0.222 | 0.212 | 0.23 | 0.1901 |  |  |  |  | 1
| 1 | 2 | 0.25 | 0.21814 | 0.228 | 0.22 | 0.216 | 0.209 | 0.2436 | 0.21 | 0.221 | 0.217 | 0.22 | 0.1901 |  |  |  |  | 35
| 1 | 3 | 0.25 | 0.22514 | 0.225 | 0.227 | 0.221 | 0.205 | 0.2333 | 0.22 | 0.22 | 0.218 | 0.21 | 0.1905 |  |  |  |  | 16
| 1 | 4 | 0.25 | 0.2282 | 0.225 | 0.225 | 0.22 | 0.2 | 0.2628 | 0.21 | 0.222 | 0.216 | 0.23 | 0.1885 |  |  |  |  | 0.21955
| 1 | 5 | 0.24 | 0.22972 | 0.223 | 0.224 | 0.219 | 0.206 | 0.2572 | 0.21 | 0.225 | 0.219 | 0.21 | 0.1891 |  |  |  |  | 129
| 1 | 6 | 0.24 | 0.22197 | 0.225 | 0.231 | 0.232 | 0.22 | 0.2748 | 0.22 | 0.22 | 0.219 | 0.22 | 0.1897
| Mean |  | 0.24833 | 0.22471 | 0.225 | 0.22483 | 0.22133 | 0.21033 | 0.25217 | 0.21333 | 0.22167 | 0.21683 | 0.22 | 0.18967
| Median |  | 0.25 | 0.22511 | 0.225 | 0.2245 | 0.22 | 0.2075 | 0.2504 | 0.21 | 0.2215 | 0.2175 | 0.22 | 0.1899
| Std Dev. |  | 0.0075277 | 0.0041997 | 0.0016733 | 0.0038687 | 0.0055015 | 0.0087788 | 0.01548 | 0.005164 | 0.0018619 | 0.0026394 | 0.0089443 | 0.00074207
| Rel.Std.Dev. |  | 3.031% | 1.869% | 0.7437% | 1.721% | 2.486% | 4.174% | 6.139% | 2.421% | 0.84% | 1.217% | 4.066% | 0.3912%
| PDM3 |  | 13.11% | 2.352% | 2.484% | 2.409% | 0.8144% | -4.196% | 14.86% | -2.83% | 0.9662% | -1.235% | 0.207% | -13.61%
| Z-Score (Absolute) |  | 4.43 | 0.53 | 0.58 | 0.55 | 0.03 | 1.84 | 5.06 | 1.35 | 0.03 | 0.77 | 0.25 | 5.26
| Analytical results for Dy in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Dy | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 3
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 21
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Er | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 4
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 22
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Eu | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 5
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 23
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 922b (Certified Value 4.49 wt.%) | [17 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 4.47 | 4.644 | 4.4191 | 5.1 | 4.4425 | 4.575 | 4.418 | 4.546 | 4.504 | 4.49 | 4.418 | 4.49 | 4.76 | 4.489 | 4.561 | 4.48 | 1
| 1 | 2 | 4.42 | 4.583 | 4.4591 | 5.18 | 4.3894 | 4.561 | 4.375 | 4.518 | 4.532 | 4.53 | 4.418 | 4.49 | 4.66 | 4.475 | 4.546 | 4.46 | 0
| 1 | 3 | 4.45 | 4.613 | 4.4291 | 5 | 4.404 | 4.518 | 4.518 | 4.575 | 4.532 | 4.54 | 4.432 | 4.5 | 4.75 | 4.518 | 4.546 | 4.46 | 16
| 1 | 4 | 4.44 | 4.86 | 4.4891 | 5.09 | 4.4162 | 4.504 | 4.489 | 4.546 | 4.489 | 4.52 | 4.389 | 4.49 | 4.73 | 4.461 | 4.518 | 4.49 | 4.4937
| 1 | 5 | 4.42 | 4.619 | 4.4091 | 5.04 | 4.4726 | 4.518 | 4.432 | 4.561 | 4.504 | 4.52 | 4.432 | 4.5 | 4.75 | 4.518 | 4.561 | 4.46 | 130
| 1 | 6 | 4.41 | 4.616 | 4.465 | 5.31 | 4.4034 | 4.546 | 4.504 | 4.561 | 4.518 | 4.6 | 4.375 | 4.48 | 4.78 | 4.489 | 4.575 | 4.49
| Mean |  | 4.435 | 4.6558 | 4.445 | 5.12 | 4.4214 | 4.537 | 4.456 | 4.5512 | 4.5132 | 4.5333 | 4.4107 | 4.4917 | 4.7383 | 4.4917 | 4.5512 | 4.4733
| Median |  | 4.43 | 4.6175 | 4.4441 | 5.095 | 4.4101 | 4.532 | 4.4605 | 4.5535 | 4.511 | 4.525 | 4.418 | 4.49 | 4.75 | 4.489 | 4.5535 | 4.47
| Std Dev. |  | 0.022583 | 0.10189 | 0.030834 | 0.11118 | 0.030796 | 0.027971 | 0.056246 | 0.01957 | 0.017233 | 0.036697 | 0.023492 | 0.0075277 | 0.041673 | 0.022888 | 0.01957 | 0.015055
| Rel.Std.Dev. |  | 0.5092% | 2.188% | 0.6937% | 2.171% | 0.6965% | 0.6165% | 1.262% | 0.43% | 0.3818% | 0.8095% | 0.5326% | 0.1676% | 0.8795% | 0.5096% | 0.43% | 0.3366%
| PDM3 |  | -1.307% | 3.608% | -1.083% | 13.94% | -1.61% | 0.9633% | -0.8393% | 1.279% | 0.4329% | 0.8817% | -1.848% | -0.04556% | 5.444% | -0.04556% | 1.279% | -0.4535%
| Z-Score (Absolute) |  | 0.93 | 2.12 | 0.79 | 8.53 | 1.12 | 0.48 | 0.64 | 0.67 | 0.15 | 0.43 | 1.27 | 0.15 | 3.26 | 0.15 | 0.67 | 0.4
| Analytical results for Ga in OREAS 922b (Indicative Value 13 ppm) | [17 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | 13 | [14 empty] | 1
| 1 | 2 | < 100 | 13 | [14 empty] | 18
| 1 | 3 | < 100 | 13 | [14 empty] | 16
| 1 | 4 | < 100 | 13 | [14 empty] | 13
| 1 | 5 | < 100 | 13 | [14 empty] | 24
| 1 | 6 | < 100 | 13
| Mean |  |  | 13
| Median |  |  | 13
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Gd in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Gd | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 6
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 25
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 41
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 14
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Hf in OREAS 922b (Certified Value < 80 ppm) | [17 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 170 | < 80 | < 80 | < 80 | < 80 | < 10 | [10 empty] | 1
| 1 | 2 | < 170 | < 80 | < 80 | < 80 | < 80 | < 10 | [10 empty] | 25
| 1 | 3 | < 170 | < 80 | < 80 | < 80 | < 80 | < 10 | [10 empty] | 16
| 1 | 4 | < 170 | < 80 | 80 | < 80 | < 80 | < 10 | [10 empty] | < 80
| 1 | 5 | < 170 | < 80 | < 80 | < 80 | < 80 | < 10 | [10 empty] | 131
| 1 | 6 | < 170 | < 80 | < 80 | < 80 | < 80 | < 10
| Mean |  | 2.08: 80
| Median |  | 2.08: 80
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 8.09 | 0 | 1.35 | 0 | 0 | 7.14
| Analytical results for Hg in OREAS 922b (Indicative Value < 100 ppm) | [17 empty] | Indicative
Hg | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.04 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100 | [10 empty] | 1
| 1 | 2 | 2.06: < 100 | [10 empty] | 23
| 1 | 3 | 2.06: < 100 | [10 empty] | 16
| 1 | 4 | 2.06: < 100 | [10 empty] | < 100
| 1 | 5 | 2.06: < 100 | [10 empty] | 15
| 1 | 6 | 2.06: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Ho in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ho | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 3
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 16
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for In in OREAS 922b (Indicative Value < 100 ppm) | [17 empty] | Indicative
In | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.04 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100 | [10 empty] | 1
| 1 | 2 | 2.06: < 100 | [10 empty] | 11
| 1 | 3 | 2.06: < 100 | [10 empty] | 16
| 1 | 4 | 2.06: < 100 | [10 empty] | < 100
| 1 | 5 | 2.06: < 100 | [10 empty] | 17
| 1 | 6 | 2.06: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for K2O in OREAS 922b (Certified Value 3.19 wt.%) | [17 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 3.15 | 3.235 | 3.15 | 4.38 | 3.2273 | 3.16 | 3.17 | 3.17 | 3.22 | 3.21 | 3.18 | 3.311 | 3.17 | 3.1681 | 3.18 |  | 1
| 1 | 2 | 3.13 | 3.205 | 3.2 | 4.56 | 3.2046 | 3.26 | 3.14 | 3.1 | 3.24 | 3.13 | 3.17 | 3.255 | 3.17 | 3.1681 | 3.14 |  | 0
| 1 | 3 | 3.13 | 3.231 | 3.19 | 4.34 | 3.2251 | 3.23 | 3.23 | 3.16 | 3.23 | 3.15 | 3.18 | 3.301 | 3.19 | 3.1561 | 3.26 |  | 16
| 1 | 4 | 3.16 | 3.292 | 3.23 | 4.19 | 3.212 | 3.18 | 3.22 | 3.19 | 3.2 | 3.12 | 3.18 | 3.277 | 3.17 | 3.1079 | 3.27 |  | 3.1918
| 1 | 5 | 3.15 | 3.222 | 3.22 | 4.19 | 3.2569 | 3.24 | 3.18 | 3.17 | 3.24 | 3.22 | 3.16 | 3.329 | 3.19 | 3.1681 | 3.24 |  | 132
| 1 | 6 | 3.15 | 3.215 | 3.17 | 4.2 | 3.22 | 3.22 | 3.25 | 3.17 | 3.23 | 3.11 | 3.17 | 3.31 | 3.19 | 3.144 | 3.27
| Mean |  | 3.145 | 3.2333 | 3.1933 | 4.31 | 3.2243 | 3.215 | 3.1983 | 3.16 | 3.2267 | 3.1567 | 3.1733 | 3.2972 | 3.18 | 3.1521 | 3.2267
| Median |  | 3.15 | 3.2265 | 3.195 | 4.27 | 3.2226 | 3.225 | 3.2 | 3.17 | 3.23 | 3.14 | 3.175 | 3.3055 | 3.18 | 3.1621 | 3.25
| Std Dev. |  | 0.012247 | 0.030716 | 0.030111 | 0.14778 | 0.018056 | 0.037815 | 0.041673 | 0.030984 | 0.015055 | 0.047188 | 0.008165 | 0.026716 | 0.010954 | 0.023678 | 0.054283
| Rel.Std.Dev. |  | 0.3894% | 0.95% | 0.9429% | 3.429% | 0.56% | 1.176% | 1.303% | 0.9805% | 0.4666% | 1.495% | 0.2573% | 0.8103% | 0.3445% | 0.7512% | 1.682%
| PDM3 |  | -1.467% | 1.3% | 0.04709% | 35.03% | 1.018% | 0.7259% | 0.2037% | -0.9972% | 1.091% | -1.102% | -0.5795% | 3.3% | -0.3706% | -1.246% | 1.091%
| Z-Score (Absolute) |  | 1.27 | 0.83 | 0.12 | 26.46 | 0.62 | 0.4 | 0 | 0.91 | 0.67 | 0.99 | 0.59 | 2.35 | 0.44 | 1.1 | 0.67
| Analytical results for La in OREAS 922b (Indicative Value 55 ppm) | [17 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 90 | < 90 | < 90 | < 90 | 35 | [11 empty] | 1
| 1 | 2 | < 90 | < 90 | 90 | < 90 | 32 | [11 empty] | 12
| 1 | 3 | < 90 | 90 | 90 | < 90 | 34 | [11 empty] | 16
| 1 | 4 | < 90 | 90 | 90 | < 90 | 32 | [11 empty] | 54.5608
| 1 | 5 | < 90 | < 90 | 90 | 90 | 36 | [11 empty] | 18
| 1 | 6 | < 90 | 90 | < 90 | < 90 | 32
| Mean |  |  | 90 | 90 | 90 | 33.5
| Median |  |  | 90 | 90 | 90 | 33
| Std Dev. |  | 2.07: 0; 2.08: 0; 2.10: 1.7607
| Rel.Std.Dev. |  | 2.07: 0%; 2.08: 0%; 2.10: 5.256%
| PDM3 |  |  | 64.95% | 64.95% | 64.95% | -38.6%
| Z-Score (Absolute) |  | 0.34 | 0.67 | 1.01 | 0 | 0.77
| Analytical results for Lu in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Lu | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 13
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 19
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for MgO in OREAS 922b (Certified Value 0.704 wt.%) | [17 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.76 | 0.671 | 0.63 | 1.32 | 0.68733 | 0.7 | 0.72 | 0.72 | 0.72 | 0.79 | 0.68 | 0.67 | 0.72 | 0.763 | 0.72 |  | 1
| 1 | 2 | 0.75 | 0.662 | 0.63 | 1.37 | 0.67813 | 0.71 | 0.73 | 0.71 | 0.72 | 0.66 | 0.67 | 0.65 | 0.71 | 0.746 | 0.72 |  | 0
| 1 | 3 | 0.72 | 0.676 | 0.65 | 1.39 | 0.68546 | 0.7 | 0.74 | 0.72 | 0.73 | 0.65 | 0.67 | 0.7 | 0.72 | 0.763 | 0.75 |  | 16
| 1 | 4 | 0.7 | 0.693 | 0.63 | 1.4 | 0.67824 | 0.71 | 0.75 | 0.74 | 0.7 | 0.6 | 0.68 | 0.68 | 0.71 | 0.763 | 0.73 |  | 0.70402
| 1 | 5 | 0.73 | 0.647 | 0.63 | 1.22 | 0.69043 | 0.7 | 0.74 | 0.72 | 0.74 | 0.72 | 0.69 | 0.68 | 0.72 | 0.779 | 0.73 |  | 133
| 1 | 6 | 0.76 | 0.652 | 0.62 | 1.22 | 0.68418 | 0.7 | 0.75 | 0.73 | 0.72 | 0.68 | 0.68 | 0.68 | 0.72 | 0.779 | 0.73
| Mean |  | 0.73667 | 0.66683 | 0.63167 | 1.32 | 0.68396 | 0.70333 | 0.73833 | 0.72333 | 0.72167 | 0.68333 | 0.67833 | 0.67667 | 0.71667 | 0.7655 | 0.73
| Median |  | 0.74 | 0.6665 | 0.63 | 1.345 | 0.68482 | 0.7 | 0.74 | 0.72 | 0.72 | 0.67 | 0.68 | 0.68 | 0.72 | 0.763 | 0.73
| Std Dev. |  | 0.024221 | 0.016869 | 0.0098319 | 0.082219 | 0.004946 | 0.005164 | 0.01169 | 0.010328 | 0.013292 | 0.06532 | 0.0075277 | 0.01633 | 0.005164 | 0.012357 | 0.010954
| Rel.Std.Dev. |  | 3.288% | 2.53% | 1.557% | 6.229% | 0.7231% | 0.7342% | 1.583% | 1.428% | 1.842% | 9.559% | 1.11% | 2.413% | 0.7206% | 1.614% | 1.501%
| PDM3 |  | 4.637% | -5.283% | -10.28% | 87.49% | -2.85% | -0.09806% | 4.873% | 2.743% | 2.506% | -2.939% | -3.649% | -3.886% | 1.796% | 8.732% | 3.69%
| Z-Score (Absolute) |  | 0.41 | 1.03 | 1.75 | 12.44 | 0.67 | 0.27 | 0.45 | 0.14 | 0.1 | 0.69 | 0.79 | 0.82 | 0 | 1.01 | 0.27
| Analytical results for MnO in OREAS 922b (Indicative Value 0.013 wt.%) | [17 empty] | Indicative
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.01 | 0.007 | 0.01 | 0.0096492 | 0.013 | 0.013 | 0.013 | 0.013 | 0.021 | 0.02 | 0.006 | 0.01 | 0.0258 | 0.02 |  |  | 1
| 1 | 2 | 0.02 | 0.007 | 0.01 | 0.010614 | 0.013 | 0.013 | 0.013 | 0.013 | 0.021 | 0.02 | < 0.005 | 0.01 | 0.0258 | 0.02 |  |  | 14
| 1 | 3 | 0.01 | 0.007 | 0.01 | 0.010617 | 0.013 | 0.013 | 0.013 | < 0.013 | 0.019 | 0.02 | < 0.005 | 0.01 | 0.0258 | 0.02 |  |  | 16
| 1 | 4 | 0.01 | 0.007 | 0.01 | 0.010616 | 0.013 | 0.013 | 0.013 | 0.013 | 0.02 | 0.02 | < 0.005 | 0.01 | 0.0258 | 0.02 |  |  | 0.012879
| 1 | 5 | 0.01 | 0.007 | 0.01 | 0.01158 | 0.013 | 0.013 | 0.013 | < 0.013 | 0.02 | 0.02 | < 0.005 | 0.01 | 0.0258 | 0.02 |  |  | 20
| 1 | 6 | 0.01 | 0.007 | 0.01 | 0.011578 | 0.013 | 0.013 | 0.013 | 0.013 | 0.024 | 0.02 | 0.005 | 0.01 | 0.0258 | 0.02
| Mean |  | 0.011667 | 0.007 | 0.01 | 0.010776 | 0.013 | 0.013 | 0.013 | 0.013 | 0.020833 | 0.02 | 0.0055 | 0.01 | 0.0258 | 0.02
| Median |  | 0.01 | 0.007 | 0.01 | 0.010616 | 0.013 | 0.013 | 0.013 | 0.013 | 0.0205 | 0.02 | 0.0055 | 0.01 | 0.0258 | 0.02
| Std Dev. |  | 0.0040825 | 9.5015e-19 | 0 | 0.00072599 | 0 | 0 | 0 | 0 | 0.0017224 | 0 | 0.00070711 | 0 | 0 | 0
| Rel.Std.Dev. |  | 34.99% | 1.357e-14% | 0% | 6.737% | 0% | 0% | 0% | 0% | 8.268% | 0% | 12.86% | 0% | 0% | 0%
| PDM3 |  | -9.412% | -45.65% | -22.35% | -16.33% | 0.9409% | 0.9409% | 0.9409% | 0.9409% | 61.76% | 55.29% | -57.29% | -22.35% | 100.3% | 55.29%
| Z-Score (Absolute) |  | 0.18 | 1.56 | 0.67 | 0.45 | 0.18 | 0.18 | 0.18 | 0.45 | 2.52 | 2.27 | 2.59 | 0.67 | 3.99 | 2.27
| Analytical results for Mo in OREAS 922b (Certified Value < 50 ppm) | [17 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 50 | < 50 | 26 | 15 | [9 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | < 50 | 23 | 15 | [9 empty] | 31
| 1 | 3 | < 100 | < 50 | < 50 | < 50 | < 50 | 22 | 14 | [9 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 50 | < 50 | 24 | 15 | [9 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 50 | < 50 | 27 | 14 | [9 empty] | 134
| 1 | 6 | < 100 | < 50 | < 50 | < 50 | < 50 | 26 | 14
| Mean |  | 2.25: 24.6667 | 14.5
| Median |  | 2.25: 25 | 14.5
| Std Dev. |  | 2.25: 1.9664 | 0.54772
| Rel.Std.Dev. |  | 2.25: 7.972% | 3.777%
PDM3
| Z-Score (Absolute) | [7 empty] | NA | bar: 2.06 to 2.25
| Analytical results for Na2O in OREAS 922b (Certified Value 0.124 wt.%) | [17 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.16 | 2.18 | 2.19 | 2.24 | [7 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 3
3
| 1 | 1 | 0.17 | 0.079 | 0.1 | 0.15 | 0.086843 | 0.14 | 0.14 | 0.14 | 0.12 | [7 empty] | 1
| 1 | 2 | 0.17 | 0.079 | 0.116 | 0.15 | 0.082018 | 0.14 | 0.12 | 0.14 | 0.11 | [7 empty] | 0
| 1 | 3 | 0.15 | 0.08 | 0.097 | 0.15 | 0.080109 | 0.14 | 0.17 | 0.14 | 0.1 | [7 empty] | 16
| 1 | 4 | 0.15 | 0.078 | 0.106 | 0.15 | 0.076241 | 0.13 | 0.26 | 0.14 | 0.11 | [7 empty] | 0.12378
| 1 | 5 | 0.15 | 0.079 | 0.123 | 0.15 | 0.088777 | 0.15 | 0.16 | 0.13 | 0.11 | [7 empty] | 135
| 1 | 6 | 0.18 | 0.08 | 0.105 | 0.15 | 0.083938 | 0.14 | 0.13 | 0.14 | 0.11
| Mean |  | 0.16167 | 0.079167 | 0.10783 | 0.15 | 0.082988 | 0.14 | 0.16333 | 0.13833 | 0.11
| Median |  | 0.16 | 0.079 | 0.1055 | 0.15 | 0.082978 | 0.14 | 0.15 | 0.14 | 0.11
| Std Dev. |  | 0.013292 | 0.00075277 | 0.0098675 | 0 | 0.0045612 | 0.0063246 | 0.050859 | 0.0040825 | 0.0063246
| Rel.Std.Dev. |  | 8.222% | 0.9509% | 9.151% | 0% | 5.496% | 4.518% | 31.14% | 2.951% | 5.75%
| PDM3 |  | 30.61% | -36.04% | -12.88% | 21.19% | -32.95% | 13.11% | 31.96% | 11.76% | -11.13%
| Z-Score (Absolute) |  | 0.63 | 1.6 | 0.82 | 0.31 | 1.49 | 0.04 | 0.67 | 0 | 0.76
| Analytical results for Nb in OREAS 922b (Indicative Value 48.8 ppm) | [17 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 1
| 1 | 2 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 15
| 1 | 3 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 48.8333
| 1 | 5 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 21
| 1 | 6 | < 50 | < 50 | 60 | < 50 | 46
| Mean |  | 2.08: 51.6667; 2.10: 46
| Median |  | 2.08: 50; 2.10: 46
| Std Dev. |  | 2.08: 4.0825; 2.10: 0
| Rel.Std.Dev. |  | 2.08: 7.902%; 2.10: 0%
| PDM3 |  | 2.08: 5.802%; 2.10: -5.802%
Z-Score (Absolute) | bar: 2.06 to 2.10
| Analytical results for Nd in OREAS 922b (Indicative Value 61 ppm) | [17 empty] | Indicative
Nd | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: 63 | [10 empty] | 1
| 1 | 2 | 2.06: 61 | [10 empty] | 4
| 1 | 3 | 2.06: 64 | [10 empty] | 16
| 1 | 4 | 2.06: 63 | [10 empty] | 61
| 1 | 5 | 2.06: 59 | [10 empty] | 22
| 1 | 6 | 2.06: 56
| Mean |  | 2.06: 61
| Median |  | 2.06: 62
| Std Dev. |  | 2.06: 3.0332
| Rel.Std.Dev. |  | 2.06: 4.972%
PDM3 | bar: 2.06
Z-Score (Absolute) | bar: 2.06
| Analytical results for Ni in OREAS 922b (Certified Value < 50 ppm) | [17 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 196 | < 50 | < 50 | 140 | < 50 | 31 | < 100 | < 10 | < 100 | < 10 | [5 empty] | 1
| 1 | 2 | < 100 | < 196 | < 50 | < 50 | < 50 | < 50 | < 24 | < 100 | < 10 | < 100 | < 10 | [5 empty] | 22
| 1 | 3 | < 100 | < 196 | < 50 | < 50 | 80 | < 50 | 31 | < 100 | < 10 | < 100 | < 10 | [5 empty] | 16
| 1 | 4 | < 100 | < 196 | < 50 | < 50 | 50 | < 50 | < 24 | < 100 | < 10 | < 100 | < 10 | [5 empty] | < 50
| 1 | 5 | < 100 | < 196 | < 50 | < 50 | 120 | < 50 | < 24 | < 100 | < 10 | < 100 | < 10 | [5 empty] | 136
| 1 | 6 | 400 | < 196 | < 50 | < 50 | 50 | < 50 | 31 | < 100 | < 10 | < 100 | < 10
| Mean |  | 2.06: 400; 2.10: 88 | 31
| Median |  | 2.06: 400; 2.10: 80 | 31
| Std Dev. |  | 2.10: 40.8656 | 0
| Rel.Std.Dev. |  | 2.10: 46.44% | 0%
PDM3
| Z-Score (Absolute) |  | 2.81 | 2.47 | 0 | 0 | 1.77 | 0 | 0.11 | 0.84 | 0.67 | 0.84 | 0.67
| Analytical results for P2O5 in OREAS 922b (Certified Value 0.076 wt.%) | [17 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.077 | 0.081 | 0.07 | 0.17 | 0.069552 | 0.06 | 0.06 | 0.08 | 0.08 | 0.08 | 0.08 | 0.028 | 0.0733 | 0.0229 | 0.07 |  | 1
| 1 | 2 | 0.077 | 0.082 | 0.08 | 0.18 | 0.070574 | 0.07 | 0.06 | 0.08 | 0.08 | 0.08 | 0.078 | 0.032 | 0.0733 | 0.0229 | 0.07 |  | 0
| 1 | 3 | 0.078 | 0.081 | 0.08 | 0.18 | 0.0665 | 0.06 | 0.06 | 0.07 | 0.08 | 0.09 | 0.078 | 0.035 | 0.0756 | 0.0229 | 0.08 |  | 16
| 1 | 4 | 0.075 | 0.079 | 0.08 | 0.2 | 0.06854 | 0.06 | 0.06 | 0.08 | 0.08 | 0.08 | 0.08 | 0.032 | 0.0756 | 0.0229 | 0.08 |  | 0.075809
| 1 | 5 | 0.077 | 0.08 | 0.07 | 0.19 | 0.067509 | 0.06 | 0.06 | 0.08 | 0.08 | 0.08 | 0.08 | 0.028 | 0.0756 | 0.0229 | 0.08 |  | 137
| 1 | 6 | 0.078 | 0.079 | 0.07 | 0.17 | 0.071589 | 0.06 | 0.06 | 0.08 | 0.08 | 0.08 | 0.08 | 0.03 | 0.0756 | 0.0229 | 0.08
| Mean |  | 0.077 | 0.080333 | 0.075 | 0.18167 | 0.069044 | 0.061667 | 0.06 | 0.078333 | 0.08 | 0.081667 | 0.079333 | 0.030833 | 0.074833 | 0.0229 | 0.076667
| Median |  | 0.077 | 0.0805 | 0.075 | 0.18 | 0.069046 | 0.06 | 0.06 | 0.08 | 0.08 | 0.08 | 0.08 | 0.031 | 0.0756 | 0.0229 | 0.08
| Std Dev. |  | 0.0010954 | 0.0012111 | 0.0054772 | 0.01169 | 0.0019057 | 0.0040825 | 0 | 0.0040825 | 0 | 0.0040825 | 0.0010328 | 0.0027142 | 0.0011877 | 0 | 0.005164
| Rel.Std.Dev. |  | 1.423% | 1.508% | 7.303% | 6.435% | 2.76% | 6.62% | 0% | 5.212% | 0% | 4.999% | 1.302% | 8.803% | 1.587% | 0% | 6.736%
| PDM3 |  | 1.571% | 5.968% | -1.068% | 139.6% | -8.924% | -18.66% | -20.85% | 3.329% | 5.528% | 7.726% | 4.648% | -59.33% | -1.287% | -69.79% | 1.131%
| Z-Score (Absolute) |  | 0.06 | 0.67 | 0.31 | 19.31 | 1.4 | 2.76 | 3.07 | 0.31 | 0.61 | 0.92 | 0.49 | 8.43 | 0.33 | 9.88 | 0
| Analytical results for Pb in OREAS 922b (Indicative Value 43.2 ppm) | [17 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 1
1
| 1 | 1 | < 100 | < 50 | < 50 | 60 | 150 | < 30 | < 100 | 60 | < 10 | 51 | [6 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | 140 | < 30 | < 100 | < 50 | < 10 | 51 | [6 empty] | 17
| 1 | 3 | < 100 | < 50 | 80 | 70 | 90 | < 30 | < 100 | 70 | < 10 | 52 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | 60 | 120 | < 30 | < 100 | 60 | < 10 | 50 | [6 empty] | 43.2
| 1 | 5 | < 100 | < 50 | < 50 | < 50 | 140 | 40 | < 100 | 70 | < 10 | 51 | [6 empty] | 23
| 1 | 6 | < 100 | < 50 | 70 | 60 | 180 | < 30 | < 100 | 70 | < 10 | 51
| Mean |  |  |  | 75 | 62.5 | 136.6667 | 40 |  | 66 |  | 51
| Median |  |  |  | 75 | 60 | 140 | 40 |  | 70 |  | 51
| Std Dev. |  | 2.08: 7.0711; 2.09: 5; 2.10: 30.1109 |  | 5.4772 |  | 0.63246
| Rel.Std.Dev. |  | 2.08: 9.428%; 2.09: 8%; 2.10: 22.03% |  | 8.299% |  | 1.24%
| PDM3 |  |  |  | 73.61% | 44.68% | 216.4% | -7.407% |  | 52.78% |  | 18.06%
| Z-Score (Absolute) |  | 0 | 1.93 | 0.64 | 0 | 6.68 | 2.38 | 0 | 0.71 | 3.47 | 0.08
| Analytical results for Pr in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Pr | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 6
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 24
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Rb in OREAS 922b (Certified Value 125 ppm) | [17 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 100 | 100 | 110 | 130 | 110 | 170 | [10 empty] | 1
| 1 | 2 | < 100 | 90 | 110 | 150 | 110 | 169 | [10 empty] | 37
| 1 | 3 | < 100 | 100 | 130 | 130 | 80 | 169 | [10 empty] | 16
| 1 | 4 | < 100 | 90 | 120 | 140 | 100 | 167 | [10 empty] | 124.6667
| 1 | 5 | < 100 | 100 | 110 | 150 | 100 | 170 | [10 empty] | 138
| 1 | 6 | < 100 | 100 | 160 | 140 | 110 | 169
| Mean |  |  | 96.6667 | 123.3333 | 140 | 101.6667 | 169
| Median |  |  | 100 | 115 | 140 | 105 | 169
| Std Dev. |  |  | 5.164 | 19.6638 | 8.9443 | 11.6905 | 1.0954
| Rel.Std.Dev. |  |  | 5.342% | 15.94% | 6.389% | 11.5% | 0.6482%
| PDM3 |  |  | -22.46% | -1.07% | 12.3% | -18.45% | 35.56%
| Z-Score (Absolute) |  | 1.95 | 0.49 | 0.34 | 0.86 | 0.34 | 1.76
| Analytical results for Sb in OREAS 922b (Certified Value < 50 ppm) | [17 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 39
| 1 | 3 | < 100 | < 50 | < 50 | 50 | < 50 | < 10 | 23 | [9 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 139
| 1 | 6 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23
| Mean |  | 2.09: 50 | 23
| Median |  | 2.09: 50 | 23
| Std Dev. | [7 empty] | 0
| Rel.Std.Dev. | [7 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 8.43 | 0 | 0 | 1.4 | 0 | 6.74 | 0.67
| Analytical results for Sc in OREAS 922b (Indicative Value 15 ppm) | [17 empty] | Indicative
Sc | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: 15 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: 16 | [10 empty] | 19
| 1 | 3 | 2.06: < 100; 2.07: 14 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: 14 | [10 empty] | 15
| 1 | 5 | 2.06: < 100; 2.07: 16 | [10 empty] | 25
| 1 | 6 | 2.06: < 100; 2.07: 15
| Mean |  | 2.07: 15
| Median |  | 2.07: 15
| Std Dev. |  | 2.07: 0.89443
| Rel.Std.Dev. |  | 2.07: 5.963%
PDM3 | bar: 2.07
| Z-Score (Absolute) |  | 2.06: 0.67; 2.07: 0.67
| Analytical results for Se in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Se | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: < 10 | [10 empty] | 41
| 1 | 3 | 2.06: < 100; 2.07: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 100; 2.07: < 10 | [10 empty] | 14
| 1 | 6 | 2.06: < 100; 2.07: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.07
| Analytical results for SiO2 in OREAS 922b (Certified Value 74.44 wt.%) | [17 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 2
3
| 1 | 1 | 73.9 | 74.069 | 73.91 | 74.96 | 74.5851 | 75.3 | 73.9 | 75.2 | 74.8 | 74.78 | 74.5 | 72.73 | 74.51 | 73.096 | 74.13 |  | 1
| 1 | 2 | 74.36 | 74.129 | 75.41 | 73.92 | 73.9808 | 75 | 73.1 | 74.3 | 75.3 | 73.28 | 74.47 | 71.11 | 74.43 | 72.818 | 74.13 |  | 0
| 1 | 3 | 73.87 | 73.828 | 74.46 | 73.71 | 74.2653 | 75.1 | 74.8 | 75.5 | 75.2 | 74.47 | 74.52 | 72.48 | 74.73 | 72.776 | 73.98 |  | 16
| 1 | 4 | 73.93 | 73.677 | 74.85 | 73.98 | 74.3001 | 74.4 | 74.9 | 75.5 | 74.5 | 73.6 | 74.37 | 71.63 | 74.35 | 71.856 | 73.95 |  | 74.4396
| 1 | 5 | 74.19 | 74.112 | 73.84 | 74.61 | 75.1622 | 74.4 | 74.1 | 75.2 | 75 | 75.37 | 74.3 | 72.89 | 74.81 | 72.776 | 73.94 |  | 140
| 1 | 6 | 74.38 | 74.044 | 72.02 | 74.42 | 74.3531 | 74.9 | 75.2 | 74.7 | 75 | 73.52 | 74.47 | 72.58 | 74.76 | 72.626 | 73.92
| Mean |  | 74.105 | 73.9765 | 74.0817 | 74.2667 | 74.4411 | 74.85 | 74.3333 | 75.0667 | 74.9667 | 74.17 | 74.4383 | 72.2367 | 74.5983 | 72.658 | 74.0083
| Median |  | 74.06 | 74.0565 | 74.185 | 74.2 | 74.3266 | 74.95 | 74.45 | 75.2 | 75 | 74.035 | 74.47 | 72.53 | 74.62 | 72.776 | 73.965
| Std Dev. |  | 0.23484 | 0.18247 | 1.1693 | 0.47631 | 0.40278 | 0.37283 | 0.78145 | 0.4761 | 0.28752 | 0.82965 | 0.085186 | 0.70494 | 0.19292 | 0.42179 | 0.096212
| Rel.Std.Dev. |  | 0.3169% | 0.2467% | 1.578% | 0.6413% | 0.5411% | 0.4981% | 1.051% | 0.6342% | 0.3835% | 1.119% | 0.1144% | 0.9759% | 0.2586% | 0.5805% | 0.13%
| PDM3 |  | -0.4495% | -0.6221% | -0.4809% | -0.2323% | 0.002022% | 0.5513% | -0.1428% | 0.8424% | 0.708% | -0.3622% | -0.001717% | -2.959% | 0.2132% | -2.393% | -0.5794%
| Z-Score (Absolute) |  | 0.42 | 0.76 | 0.48 | 0 | 0.46 | 1.52 | 0.17 | 2.09 | 1.83 | 0.25 | 0.45 | 5.3 | 0.87 | 4.2 | 0.67
| Analytical results for Sm in OREAS 922b (Indicative Value 28.5 ppm) | [17 empty] | Indicative
Sm | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: 27 | [10 empty] | 1
| 1 | 2 | 2.06: 28 | [10 empty] | 4
| 1 | 3 | 2.06: 26 | [10 empty] | 16
| 1 | 4 | 2.06: 23 | [10 empty] | 28.5
| 1 | 5 | 2.06: 35 | [10 empty] | 15
| 1 | 6 | 2.06: 32
| Mean |  | 2.06: 28.5
| Median |  | 2.06: 27.5
| Std Dev. |  | 2.06: 4.3243
| Rel.Std.Dev. |  | 2.06: 15.17%
PDM3 | bar: 2.06
Z-Score (Absolute) | bar: 2.06
| Analytical results for Sn in OREAS 922b (Certified Value < 50 ppm) | [17 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | [8 empty] | 1
| 1 | 2 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | [8 empty] | 26
| 1 | 3 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | [8 empty] | 16
| 1 | 4 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | [8 empty] | < 50
| 1 | 5 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10 | [8 empty] | 141
| 1 | 6 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 1.69 | 3.63 | 0 | 0 | 0 | 0 | 1.69 | 1.35
| Analytical results for SO3 in OREAS 922b (Certified Value 0.667 wt.%) | [17 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.16 | 2.18 | 2.19 | 2.20 | 2.25 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 0.7192 | 0.6363 | 0.87 | 0.674 | 0.649 | 0.674 | 0.674 | 0.68 | 0.469 | 0.5719 | 0.6742 | 0.59483 |  |  |  |  | 1
| 1 | 2 | 0.7067 | 0.635 | 0.84 | 0.699 | 0.649 | 0.649 | 0.674 | 0.68 | 0.441 | 0.6068 | 0.6742 | 0.58384 |  |  |  |  | 0
| 1 | 3 | 0.7092 | 0.6336 | 0.86 | 0.674 | 0.674 | 0.674 | 0.674 | 0.68 | 0.474 | 0.6118 | 0.6742 | 0.59558 |  |  |  |  | 16
| 1 | 4 | 0.7042 | 0.6432 | 0.89 | 0.674 | 0.674 | 0.699 | 0.674 | 0.68 | 0.474 | 0.6193 | 0.6493 | 0.59458 |  |  |  |  | 0.66734
| 1 | 5 | 0.7117 | 0.6356 | 0.85 | 0.724 | 0.674 | 0.649 | 0.674 | 0.68 | 0.406 | 0.6168 | 0.6742 | 0.57435 |  |  |  |  | 142
| 1 | 6 | 0.6992 | 0.6362 | 0.85 | 0.674 | 0.699 | 0.674 | 0.674 | 0.68 | 0.46 | 0.6318 | 0.6742 | 0.57785
| Mean |  | 0.70837 | 0.63665 | 0.86 | 0.6865 | 0.66983 | 0.66983 | 0.674 | 0.68 | 0.454 | 0.60973 | 0.67005 | 0.58684
| Median |  | 0.70795 | 0.6359 | 0.855 | 0.674 | 0.674 | 0.674 | 0.674 | 0.68 | 0.4645 | 0.6143 | 0.6742 | 0.58921
| Std Dev. |  | 0.0068313 | 0.0033572 | 0.017889 | 0.020917 | 0.018819 | 0.018819 | 0 | 0 | 0.026586 | 0.02036 | 0.010165 | 0.0094441
| Rel.Std.Dev. |  | 0.9644% | 0.5273% | 2.08% | 3.047% | 2.81% | 2.81% | 0% | 0% | 5.856% | 3.339% | 1.517% | 1.609%
| PDM3 |  | 6.148% | -4.599% | 28.87% | 2.871% | 0.3739% | 0.3739% | 0.9983% | 1.897% | -31.97% | -8.632% | 0.4064% | -12.06%
| Z-Score (Absolute) |  | 1.03 | 0.9 | 5.11 | 0.45 | 0 | 0 | 0.11 | 0.27 | 5.82 | 1.63 | 0 | 2.24
| Analytical results for SrO in OREAS 922b (Certified Value 63 ppm) | [17 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.18 | 2.19 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | 90 | 106 | < 100 | 54.4 | [6 empty] | 1
| 1 | 2 | 200 | < 50 | < 120 | < 120 | < 120 | < 120 | 110 | 71 | < 100 | 54.4 | [6 empty] | 0
| 1 | 3 | < 100 | < 50 | < 120 | 120 | < 120 | < 120 | 100 | 95 | < 100 | 54.4 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | 90 | 83 | < 100 | 53.2 | [6 empty] | 62.5794
| 1 | 5 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | 110 | 95 | < 100 | 54.4 | [6 empty] | 143
| 1 | 6 | < 100 | < 50 | < 120 | 120 | < 120 | < 120 | 110 | 83 | < 100 | 54.4
| Mean |  | 2.06: 200; 2.09: 120 | 101.6667 | 88.8333 |  | 54.2
| Median |  | 2.06: 200; 2.09: 120 | 105 | 89 |  | 54.4
| Std Dev. |  | 2.09: 0 | 9.8319 | 12.3031 |  | 0.4899
| Rel.Std.Dev. | [7 empty] | 9.671% | 13.85% |  | 0.9039% | bar: 2.09
| PDM3 |  | 2.06: 219.6%; 2.09: 91.76% | 62.46% | 41.95% |  | -13.39%
| Z-Score (Absolute) |  | 0.86 | 1.84 | 0 | 1.06 | 0 | 0 | 2.29 | 1.59 | 0.49 | 0.27
| Analytical results for Ta in OREAS 922b (Indicative Value 14.7 ppm) | [17 empty] | Indicative
Ta | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: 14 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: 15 | [10 empty] | 10
| 1 | 3 | 2.06: < 100; 2.07: 16 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: 14 | [10 empty] | 14.6667
| 1 | 5 | 2.06: < 100; 2.07: 15 | [10 empty] | 16
| 1 | 6 | 2.06: < 100; 2.07: 14
| Mean |  | 2.07: 14.6667
| Median |  | 2.07: 14.5
| Std Dev. |  | 2.07: 0.8165
| Rel.Std.Dev. |  | 2.07: 5.567%
| PDM3 |  | 2.07: -2.331e-13%
| Z-Score (Absolute) |  | 2.06: 0.67; 2.07: 0.67
| Analytical results for Tb in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Tb | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 11
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 17
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Te in OREAS 922b (Indicative Value < 100 ppm) | [17 empty] | Indicative
Te | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.04 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100 | [10 empty] | 1
| 1 | 2 | 2.06: < 100 | [10 empty] | 12
| 1 | 3 | 2.06: < 100 | [10 empty] | 16
| 1 | 4 | 2.06: < 100 | [10 empty] | < 100
| 1 | 5 | 2.06: < 100 | [10 empty] | 18
| 1 | 6 | 2.06: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Th in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Th | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 13
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 19
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for TiO2 in OREAS 922b (Certified Value 0.528 wt.%) | [17 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.52 | 0.556 | 0.52 | 0.78 | 0.55287 | 0.51 | 0.52 | 0.51 | 0.52 | 0.53 | 0.52 | 0.57 | 0.52 | 0.5171 | 0.52 |  | 1
| 1 | 2 | 0.57 | 0.552 | 0.53 | 0.95 | 0.53667 | 0.51 | 0.51 | 0.52 | 0.53 | 0.52 | 0.52 | 0.55 | 0.52 | 0.5004 | 0.52 |  | 0
| 1 | 3 | 0.56 | 0.551 | 0.53 | 0.97 | 0.55504 | 0.51 | 0.53 | 0.52 | 0.52 | 0.53 | 0.51 | 0.56 | 0.52 | 0.5171 | 0.52 |  | 16
| 1 | 4 | 0.53 | 0.55 | 0.52 | 0.95 | 0.53879 | 0.5 | 0.52 | 0.51 | 0.52 | 0.53 | 0.52 | 0.54 | 0.52 | 0.5171 | 0.51 |  | 0.52814
| 1 | 5 | 0.55 | 0.557 | 0.52 | 0.81 | 0.54885 | 0.5 | 0.52 | 0.51 | 0.53 | 0.53 | 0.51 | 0.56 | 0.52 | 0.5171 | 0.51 |  | 144
| 1 | 6 | 0.56 | 0.553 | 0.54 | 0.8 | 0.5599 | 0.51 | 0.53 | 0.5 | 0.53 | 0.51 | 0.52 | 0.57 | 0.52 | 0.5171 | 0.51
| Mean |  | 0.54833 | 0.55317 | 0.52667 | 0.87667 | 0.54869 | 0.50667 | 0.52167 | 0.51167 | 0.525 | 0.525 | 0.51667 | 0.55833 | 0.52 | 0.51432 | 0.515
| Median |  | 0.555 | 0.5525 | 0.525 | 0.88 | 0.55086 | 0.51 | 0.52 | 0.51 | 0.525 | 0.53 | 0.52 | 0.56 | 0.52 | 0.5171 | 0.515
| Std Dev. |  | 0.019408 | 0.0027869 | 0.008165 | 0.088468 | 0.0092299 | 0.005164 | 0.0075277 | 0.0075277 | 0.0054772 | 0.0083666 | 0.005164 | 0.01169 | 0 | 0.0068177 | 0.0054772
| Rel.Std.Dev. |  | 3.539% | 0.5038% | 1.55% | 10.09% | 1.682% | 1.019% | 1.443% | 1.471% | 1.043% | 1.594% | 0.9995% | 2.094% | 0% | 1.326% | 1.064%
| PDM3 |  | 3.823% | 4.739% | -0.279% | 65.99% | 3.89% | -4.066% | -1.226% | -3.119% | -0.5946% | -0.5946% | -2.172% | 5.717% | -1.541% | -2.617% | -2.488%
| Z-Score (Absolute) |  | 1.47 | 1.78 | 0.11 | 22.23 | 1.5 | 1.16 | 0.21 | 0.84 | 0 | 0 | 0.53 | 2.11 | 0.32 | 0.67 | 0.63
| Analytical results for Tl in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Tl | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: < 10 | [10 empty] | 14
| 1 | 3 | 2.06: < 100; 2.07: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 100; 2.07: < 10 | [10 empty] | 20
| 1 | 6 | 2.06: < 100; 2.07: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.07
| Analytical results for Tm in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Tm | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 5
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 21
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for U in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
U | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 16
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 22
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for V in OREAS 922b (Certified Value 161 ppm) | [17 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | 199.9892 | 200 | 200 | 100 | 100 | 162 | 123 | 190 | 223 | [6 empty] | 1
| 1 | 2 | < 100 | 199.9892 | 100 | 200 | 100 | 100 | 168 | 118 | 180 | 223 | [6 empty] | 36
| 1 | 3 | < 100 | 178.4139 | 200 | 200 | 100 | 100 | 157 | 140 | 190 | 222 | [6 empty] | 16
| 1 | 4 | < 100 | 200.022 | 100 | 200 | 100 | 100 | 157 | 134 | 190 | 224 | [6 empty] | 161.1741
| 1 | 5 | < 100 | 199.998 | 100 | 200 | 100 | 100 | 162 | 129 | 200 | 223 | [6 empty] | 145
| 1 | 6 | < 100 | 178.3481 | 100 | 200 | 200 | 100 | 174 | 196 | 190 | 221
| Mean |  |  | 192.7934 | 133.3333 | 200 | 116.6667 | 100 | 163.3333 | 140 | 190 | 222.6667
| Median |  |  | 199.9892 | 100 | 200 | 100 | 100 | 162 | 131.5 | 190 | 223
| Std Dev. |  |  | 11.1638 | 51.6398 | 0 | 40.8248 | 0 | 6.6232 | 28.5167 | 6.3246 | 1.0328
| Rel.Std.Dev. |  |  | 5.791% | 38.73% | 0% | 34.99% | 0% | 4.055% | 20.37% | 3.329% | 0.4638%
| PDM3 |  |  | 19.62% | -17.27% | 24.09% | -27.61% | -37.96% | 1.34% | -13.14% | 17.88% | 38.15%
| Z-Score (Absolute) |  | 1.73 | 0.7 | 0.31 | 0.82 | 0.6 | 0.88 | 0.2 | 0.2 | 0.65 | 1.21
| Analytical results for W in OREAS 922b (Indicative Value 10.1 ppm) | [17 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 1
1
| 1 | 1 | < 100 | < 10 | < 10 | < 10 | 20 | 13 | [10 empty] | 1
| 1 | 2 | < 100 | < 10 | < 10 | < 10 | < 10 | 16 | [10 empty] | 17
| 1 | 3 | < 100 | < 10 | < 10 | 10 | 10 | 14 | [10 empty] | 16
| 1 | 4 | < 100 | < 10 | 20 | < 10 | < 10 | 19 | [10 empty] | 10.125
| 1 | 5 | < 100 | < 10 | < 10 | < 10 | 20 | 14 | [10 empty] | 23
| 1 | 6 | < 100 | < 10 | < 10 | < 10 | 10 | 17
| Mean |  |  |  | 20 | 10 | 15 | 15.5
| Median |  |  |  | 20 | 10 | 15 | 15
| Std Dev. |  | 2.10: 5.7735; 2.25: 2.2583
| Rel.Std.Dev. |  | 2.10: 38.49%; 2.25: 14.57%
| PDM3 |  |  |  | 97.53% | -1.235% | 48.15% | 53.09%
| Z-Score (Absolute) |  | 6.54 | 0.74 | 0.34 | 0.61 | 0.34 | 0.96
| Analytical results for Y in OREAS 922b (Indicative Value 36.3 ppm) | [17 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 39 | < 39 | < 39 | < 39 | 30 | [11 empty] | 1
| 1 | 2 | < 39 | 39 | 47 | < 39 | 31 | [11 empty] | 18
| 1 | 3 | < 39 | 55 | 55 | < 39 | 30 | [11 empty] | 16
| 1 | 4 | < 39 | < 39 | 47 | < 39 | 30 | [11 empty] | 36.3043
| 1 | 5 | < 39 | 39 | < 39 | < 39 | 30 | [11 empty] | 24
| 1 | 6 | < 39 | 71 | 39 | < 39 | 30
| Mean |  | 2.07: 51; 2.08: 47; 2.10: 30.1667
| Median |  | 2.07: 47; 2.08: 47; 2.10: 30
| Std Dev. |  | 2.07: 15.3188; 2.08: 6.532; 2.10: 0.40825
| Rel.Std.Dev. |  | 2.07: 30.04%; 2.08: 13.9%; 2.10: 1.353%
| PDM3 |  | 2.07: 40.48%; 2.08: 29.46%; 2.10: -16.91%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 0.51 | 0.67 | 0
| Analytical results for Yb in OREAS 922b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Yb | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 19
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 25
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Zn in OREAS 922b (Certified Value 75 ppm) | [17 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 160 | 80 | 50 | 60 | 60 | 100 | < 100 | 80 | < 50 | 97 | [5 empty] | 1
| 1 | 2 | < 100 | < 160 | 100 | 50 | 60 | 60 | 100 | < 100 | 70 | < 50 | 97 | [5 empty] | 40
| 1 | 3 | < 100 | < 160 | 70 | 70 | 60 | 60 | 100 | < 100 | 80 | < 50 | 96 | [5 empty] | 16
| 1 | 4 | < 100 | < 160 | 60 | 70 | 60 | 50 | 90 | < 100 | 80 | < 50 | 96 | [5 empty] | 74.5238
| 1 | 5 | < 100 | < 160 | 70 | 60 | 60 | 60 | 100 | < 100 | 80 | < 50 | 98 | [5 empty] | 146
| 1 | 6 | < 100 | < 160 | 70 | 70 | 60 | 50 | 100 | < 100 | 80 | < 50 | 96
| Mean |  |  |  | 75 | 61.6667 | 60 | 56.6667 | 98.3333 |  | 78.3333 |  | 96.6667
| Median |  |  |  | 70 | 65 | 60 | 60 | 100 |  | 80 |  | 96.5
| Std Dev. |  |  |  | 13.784 | 9.8319 | 0 | 5.164 | 4.0825 |  | 4.0825 |  | 0.8165
| Rel.Std.Dev. |  |  |  | 18.38% | 15.94% | 0% | 9.113% | 4.152% |  | 5.212% |  | 0.8447%
| PDM3 |  |  |  | 0.639% | -17.25% | -19.49% | -23.96% | 31.95% |  | 5.112% |  | 29.71%
| Z-Score (Absolute) |  | 0.59 | 0.94 | 0.67 | 0 | 0.08 | 0.25 | 1.85 | 0.59 | 0.84 | 1.85 | 1.77
| Analytical results for Zr in OREAS 922b (Certified Value 112 ppm) | [17 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.19 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 370 | 200 | 100 | 100 | 100 | 120 | 138 | [8 empty] | 1
| 1 | 2 | < 100 | < 370 | 100 | 100 | 100 | 100 | 110 | 138 | [8 empty] | 41
| 1 | 3 | < 100 | < 370 | 100 | 100 | 100 | 100 | 150 | 136 | [8 empty] | 16
| 1 | 4 | < 100 | < 370 | 100 | 100 | 100 | 100 | 150 | 137 | [8 empty] | 111.7222
| 1 | 5 | < 100 | < 370 | 100 | 100 | 100 | 100 | 130 | 137 | [8 empty] | 147
| 1 | 6 | 300 | < 370 | 100 | 100 | 100 | 100 | 140 | 136
| Mean |  | 300 |  | 116.6667 | 100 | 100 | 100 | 133.3333 | 137
| Median |  | 300 |  | 100 | 100 | 100 | 100 | 135 | 137
| Std Dev. |  |  |  | 40.8248 | 0 | 0 | 0 | 16.3299 | 0.89443
| Rel.Std.Dev. |  |  |  | 34.99% | 0% | 0% | 0% | 12.25% | 0.6529%
| PDM3 |  | 168.5% |  | 4.426% | -10.49% | -10.49% | -10.49% | 19.34% | 22.63%
| Z-Score (Absolute) |  | 0.9 | 4.14 | 0.45 | 0.45 | 0.45 | 0.45 | 1.35 | 1.55
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C396: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C414: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C432: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C450: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C468: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C486: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C504: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C522: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C540: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C558: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C576: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C594: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C612: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C630: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C648: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C666: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C684: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C702: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C720: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C738: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C756: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C774: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C792: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C810: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C828: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C846: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C864: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C882: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C900: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C936: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C954: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1026: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 922b (Certified Value 3.59 wt.%) | [18 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*Furnace | BF*XRF | 2
3
| 1 | 1 | 3.435 | 3.7 | 3.56 | 3.3 | 3.6763 | 3.34 | 3.56 | 3.69 | 3.33 | 3.87 | 3.7464 | 3.42 | 4.28 | 3.4 | 3.38 | 3.97 | 3.75 | 1
| 1 | 2 | 3.5506 | 3.76 | 3.48 | 3.66 | 3.6763 | 3.33 | 3.62 | 3.69 | 3.61 | 3.85 | 3.7543 | 3.44 | 4.54 | 3.4 | 3.45 | 3.99 | 3.84 | 0
| 1 | 3 | 3.3962 | 3.9 | 3.48 | 3.43 | 3.6507 | 3.3 | 3.65 | 3.61 | 3.4 | 3.79 | 3.757 | 3.48 | 4.28 | 3.4 | 3.39 | 3.9 | 3.76 | 16
| 1 | 4 | 3.372 | 3.82 | 3.48 | 3.08 | 3.6597 | 3.39 | 3.76 | 3.65 | 3.42 | 3.94 | 3.7812 | 3.46 | 4.59 | 3.5 | 3.43 | 3.91 | 3.73 | 3.5917
| 1 | 5 | 3.4259 | 3.76 | 3.51 | 3.25 | 3.6718 | 3.28 | 3.6 | 3.68 | 3.48 | 3.88 | 3.7599 | 3.39 | 4.73 | 3.4 | 3.46 | 3.97 | 3.8 | 149
| 1 | 6 | 3.4591 | 3.79 | 3.54 | 3.52 | 3.688 | 3.28 | 3.54 | 3.82 | 3.45 | 3.89 | 3.748 | 3.38 | 4.36 | 3.4 | 3.45 | 3.98 | 3.8
| Mean |  | 3.4398 | 3.7883 | 3.5083 | 3.3733 | 3.6705 | 3.32 | 3.6217 | 3.69 | 3.4483 | 3.87 | 3.7578 | 3.4283 | 4.4633 | 3.4167 | 3.4267 | 3.9533 | 3.78
| Median |  | 3.4305 | 3.775 | 3.495 | 3.365 | 3.6741 | 3.315 | 3.61 | 3.685 | 3.435 | 3.875 | 3.7557 | 3.43 | 4.45 | 3.4 | 3.44 | 3.97 | 3.78
| Std Dev. |  | 0.062208 | 0.067651 | 0.034881 | 0.20666 | 0.013294 | 0.042426 | 0.078592 | 0.070711 | 0.09411 | 0.049396 | 0.012594 | 0.0392 | 0.1849 | 0.040825 | 0.033862 | 0.038297 | 0.040497
| Rel.Std.Dev. |  | 1.808% | 1.786% | 0.9942% | 6.126% | 0.3622% | 1.278% | 2.17% | 1.916% | 2.729% | 1.276% | 0.3351% | 1.143% | 4.143% | 1.195% | 0.9882% | 0.9687% | 1.071%
| PDM3 |  | -4.229% | 5.475% | -2.321% | -6.08% | 2.193% | -7.564% | 0.8345% | 2.737% | -3.991% | 7.749% | 4.625% | -4.548% | 24.27% | -4.873% | -4.595% | 10.07% | 5.243%
| Z-Score (Absolute) |  | 0.67 | 0.62 | 0.42 | 0.92 | 0.18 | 1.12 | 0 | 0.25 | 0.64 | 0.92 | 0.5 | 0.72 | 3.12 | 0.76 | 0.72 | 1.23 | 0.59
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 922b (Indicative Value 0.98 ppm) | [22 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.04 | 2.13 | 2.24 | [18 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | [18 empty] | 3
3
| 1 | 1 | 1.2 | 1 | < 10 | [18 empty] | 1
| 1 | 2 | 1.19 | < 1 | < 10 | [18 empty] | 8
| 1 | 3 | 1.21 | 1 | < 10 | [18 empty] | 16
| 1 | 4 | 1.28 | < 1 | < 10 | [18 empty] | 0.98
| 1 | 5 | 1.2 | < 1 | < 10 | [18 empty] | 27
| 1 | 6 | 1.18 | 1 | < 10
| Mean |  | 1.21 | 1
| Median |  | 1.2 | 1
| Std Dev. |  | 0.035777 | 0
| Rel.Std.Dev. |  | 2.957% | 0%
| PDM3 |  | 23.47% | 2.041%
| Z-Score (Absolute) |  | 0 | 0.67 | 5.56
| Analytical results for Al in OREAS 922b (Certified Value 6.46 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 6.36 | 6.1736 | 5.73 | 6.46 | 6.4669 | 6.377 | 6.563 | 6.483 | 6.483 | 6.51 | 6.47 | 6.38 | 6.4489 | 6.58 | 6.37 | 6.51 | 6.57 | 6.703 | 6.54 | 6.43 |  | 1
| 1 | 2 | 6.32 | 6.1738 | 5.66 | 6.73 | 6.4748 | 6.43 | 6.51 | 6.536 | 6.298 | 6.377 | 6.39 | 6.3 | 6.2207 | 6.58 | 6.69 | 6.467 | 6.58 | 6.754 | 6.58 | 6.52 |  | 0
| 1 | 3 | 6.38 | 6.1513 | 5.81 | 6.67 | 6.4536 | 6.43 | 6.642 | 6.563 | 6.404 | 6.483 | 6.39 | 6.39 | 6.0542 | 6.59 | 6.57 | 6.499 | 6.58 | 6.737 | 6.62 | 6.24 |  | 16
| 1 | 4 | 6.37 | 6.172 | 5.66 | 6.52 | 6.5235 | 6.404 | 6.483 | 6.483 | 6.272 | 6.43 | 6.44 | 6.39 | 6.5379 | 6.61 | 6.61 | 6.573 | 6.57 | 6.708 | 6.57 | 6.4 |  | 6.4632
| 1 | 5 | 6.3 | 6.1905 | 5.72 | 6.37 | 6.4738 | 6.457 | 6.536 | 6.51 | 6.272 | 6.51 | 6.45 | 6.43 | 6.2369 | 6.63 | 6.69 | 6.499 | 6.58 | 6.726 | 6.53 | 6.18 |  | 151
| 1 | 6 | 6.34 | 6.1615 | 5.81 | 6.55 | 6.5502 | 6.457 | 6.43 | 6.43 | 6.351 | 6.51 | 6.43 | 6.38 | 6.2525 | 6.56 | 6.43 | 6.404 | 6.57 | 6.72 | 6.51 | 6.48
| Mean |  | 6.345 | 6.1704 | 5.7317 | 6.55 | 6.4905 | 6.4258 | 6.5273 | 6.5008 | 6.3467 | 6.47 | 6.4283 | 6.3783 | 6.2918 | 6.5917 | 6.56 | 6.492 | 6.575 | 6.7247 | 6.5583 | 6.375
| Median |  | 6.35 | 6.1728 | 5.725 | 6.535 | 6.4743 | 6.43 | 6.523 | 6.4965 | 6.3245 | 6.4965 | 6.435 | 6.385 | 6.2447 | 6.585 | 6.59 | 6.499 | 6.575 | 6.723 | 6.555 | 6.415
| Std Dev. |  | 0.030822 | 0.013216 | 0.067355 | 0.13282 | 0.037682 | 0.031109 | 0.072437 | 0.046559 | 0.084156 | 0.055169 | 0.032506 | 0.042622 | 0.17398 | 0.024833 | 0.13372 | 0.055411 | 0.0054772 | 0.018886 | 0.039707 | 0.13561
| Rel.Std.Dev. |  | 0.4858% | 0.2142% | 1.175% | 2.028% | 0.5806% | 0.4841% | 1.11% | 0.7162% | 1.326% | 0.8527% | 0.5057% | 0.6682% | 2.765% | 0.3767% | 2.038% | 0.8535% | 0.0833% | 0.2808% | 0.6054% | 2.127%
| PDM3 |  | -1.829% | -4.53% | -11.32% | 1.342% | 0.4214% | -0.5788% | 0.9917% | 0.5817% | -1.804% | 0.1046% | -0.5401% | -1.314% | -2.652% | 1.987% | 1.497% | 0.445% | 1.729% | 4.045% | 1.471% | -1.365%
| Z-Score (Absolute) |  | 1.04 | 2.39 | 5.78 | 0.54 | 0.08 | 0.42 | 0.36 | 0.16 | 1.03 | 0.08 | 0.4 | 0.79 | 1.46 | 0.86 | 0.62 | 0.09 | 0.73 | 1.89 | 0.6 | 0.81
| Analytical results for As in OREAS 922b (Certified Value 55 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES | PF*OES | PF*MS | PF*OES/MS | [5 empty] | 0
1
| 1 | 1 | 50 | 23.3 | 200 | < 100 | < 100 | 200 | 100 | 56 | 48 | 56.2939 | 59 | < 100 | 60 | 54 | 54 | 54 | [5 empty] | 1
| 1 | 2 | 50 | 29 | 100 | < 100 | 100 | 100 | < 100 | 56 | 50 | 49.1906 | 63 | < 100 | 60 | 56 | 56 | 56 | [5 empty] | 9
| 1 | 3 | 50 | 21.7 | 100 | 100 | 100 | 100 | < 100 | 54 | 47 | 51.319 | 57 | < 100 | 60 | 55 | 57 | 54 | [5 empty] | 16
| 1 | 4 | 50 | 29 | 100 | < 100 | 100 | 100 | < 100 | 54 | 50 | 59.9644 | 63 | < 100 | 61 | 56 | 54 | 51 | [5 empty] | 54.8555
| 1 | 5 | 50 | 26.3 | 100 | < 100 | 100 | < 100 | < 100 | 54 | 53 | 54.1463 | 59 | < 100 | 60 | 56 | 54 | 52 | [5 empty] | 152
| 1 | 6 | 50 | 22.6 | 100 | < 100 | 100 | 100 | < 100 | 55 | 50 | 57.2841 | 59 | < 100 | 61 | 56 | 56 | 54
| Mean |  | 50 | 25.3167 | 116.6667 | 100 | 100 | 120 | 100 | 54.8333 | 49.6667 | 54.6997 | 60 |  | 60.3333 | 55.5 | 55.1667 | 53.5
| Median |  | 50 | 24.8 | 100 | 100 | 100 | 100 | 100 | 54.5 | 50 | 55.2201 | 59 |  | 60 | 56 | 55 | 54
| Std Dev. |  | 0 | 3.2443 | 40.8248 |  | 0 | 44.7214 |  | 0.98319 | 2.0656 | 3.9755 | 2.4495 |  | 0.5164 | 0.83666 | 1.3292 | 1.7607
| Rel.Std.Dev. |  | 0% | 12.81% | 34.99% |  | 0% | 37.27% |  | 1.793% | 4.159% | 7.268% | 4.082% |  | 0.8559% | 1.507% | 2.409% | 3.291%
| PDM3 |  | -8.851% | -53.85% | 112.7% | 82.3% | 82.3% | 118.8% | 82.3% | -0.04045% | -9.459% | -0.284% | 9.378% |  | 9.986% | 1.175% | 0.5672% | -2.471%
| Z-Score (Absolute) |  | 0.67 | 4.22 | NA | NA | NA | NA | NA | 0.02 | 0.72 | 0 | 0.76 | 0.67 | 0.81 | 0.11 | 0.07 | 0.17
Results from laboratories 2.07, 2.08, 2.09, 2.10 and 2.11 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 922b (Indicative Value 117 ppm) | [22 empty] | Indicative
B | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  | 1
| No | No | 2.01: 2.12; 2.02: 2.14; 2.04: 2.16; 2.05: 2.18; 2.06: 2.23 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*MS; 2.06: PF*OES |  | 0
0
| 1 | 1 | 2.01: 110; 2.02: < 50; 2.04: 133; 2.05: 108; 2.06: 122 |  | 1
| 1 | 2 | 2.01: 110; 2.02: < 50; 2.04: 126; 2.05: 102; 2.06: 122 |  | 1
| 1 | 3 | 2.01: 110; 2.02: < 50; 2.04: 129; 2.05: 102; 2.06: 130 |  | 16
| 1 | 4 | 2.01: 110; 2.02: < 50; 2.04: 133; 2.05: 111; 2.06: 121 |  | 117.0333
| 1 | 5 | 2.01: 110; 2.02: < 50; 2.04: 128; 2.05: 101; 2.06: 123 |  | 28
| 1 | 6 | 2.01: 110; 2.02: < 50; 2.04: 134; 2.05: 101; 2.06: 131
| Mean |  | 2.01: 110; 2.04: 130.5; 2.05: 104.1667; 2.06: 124.8333
| Median |  | 2.01: 110; 2.04: 131; 2.05: 102; 2.06: 122.5
| Std Dev. |  | 2.01: 0; 2.04: 3.2711; 2.05: 4.2622; 2.06: 4.446
| Rel.Std.Dev. |  | 2.01: 0%; 2.04: 2.507%; 2.05: 4.092%; 2.06: 3.562%
| PDM3 |  | 2.01: -6.01%; 2.04: 11.51%; 2.05: -10.99%; 2.06: 6.665%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 3.86; 2.04: 0.93; 2.05: 0.27; 2.06: 0.67
| Analytical results for Ba in OREAS 922b (Certified Value 0.195 wt.%) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [7 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 3
3
| 1 | 1 | 0.209 | 0.1864 | 0.1952 | 0.19943 | 0.204 | 0.1946 | 0.18179 | 0.2015 | 0.2 | 0.1926 | 0.1944 | 0.1937 | 0.1989 | 0.1839 | [7 empty] | 1
| 1 | 2 | 0.201 | 0.18 | 0.19519 | 0.20037 | 0.2 | 0.1944 | 0.17798 | 0.2034 | 0.19 | 0.1931 | 0.1954 | 0.1985 | 0.2027 | 0.1897 | [7 empty] | 27
| 1 | 3 | 0.204 | 0.1844 | 0.19629 | 0.1992 | 0.194 | 0.1949 | 0.17951 | 0.2075 | 0.2 | 0.1955 | 0.1943 | 0.1906 | 0.2002 | 0.1859 | [7 empty] | 16
| 1 | 4 | 0.206 | 0.1808 | 0.19006 | 0.20204 | 0.194 | 0.1952 | 0.18207 | 0.2005 | 0.2 | 0.1902 | 0.1954 | 0.1974 | 0.2015 | 0.1825 | [7 empty] | 0.19505
| 1 | 5 | 0.208 | 0.184 | 0.19282 | 0.19954 | 0.2 | 0.1953 | 0.18465 | 0.1995 | 0.2 | 0.1957 | 0.1951 | 0.1999 | 0.1989 | 0.1804 | [7 empty] | 153
| 1 | 6 | 0.208 | 0.186 | 0.19666 | 0.20134 | 0.208 | 0.1933 | 0.18382 | 0.2019 | 0.2 | 0.1957 | 0.1954 | 0.1917 | 0.1982 | 0.1895
| Mean |  | 0.206 | 0.1836 | 0.19437 | 0.20032 | 0.2 | 0.19462 | 0.18164 | 0.20238 | 0.19833 | 0.1938 | 0.195 | 0.1953 | 0.20007 | 0.18532
| Median |  | 0.207 | 0.1842 | 0.1952 | 0.19995 | 0.2 | 0.19475 | 0.18193 | 0.2017 | 0.2 | 0.1943 | 0.19525 | 0.19555 | 0.19955 | 0.1849
| Std Dev. |  | 0.0030332 | 0.0026533 | 0.0025008 | 0.0011534 | 0.0055136 | 0.00073052 | 0.0025283 | 0.0028315 | 0.0040825 | 0.0022361 | 0.00051769 | 0.003832 | 0.0017443 | 0.0037728
| Rel.Std.Dev. |  | 1.472% | 1.445% | 1.287% | 0.5758% | 2.757% | 0.3754% | 1.392% | 1.399% | 2.058% | 1.154% | 0.2655% | 1.962% | 0.8719% | 2.036%
| PDM3 |  | 5.612% | -5.872% | -0.3502% | 2.7% | 2.536% | -0.2237% | -6.878% | 3.758% | 1.682% | -0.6424% | -0.02722% | 0.1266% | 2.57% | -4.992%
| Z-Score (Absolute) |  | 1.5 | 1.59 | 0.11 | 0.71 | 0.67 | 0.07 | 1.87 | 1 | 0.44 | 0.19 | 0.02 | 0.02 | 0.68 | 1.36
| Analytical results for Be in OREAS 922b (Certified Value 2.51 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*OES | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 2 | 3.3 | 2.4 | < 3 | < 5 | 2.6447 | 3 | 2.4 | 2 | < 5 | < 5 | 2 | < 2 | [8 empty] | 1
| 1 | 2 | 3 | 3.3 | 2.3 | < 3 | < 5 | 2.2193 | 3 | 2.9 | 2 | < 5 | < 5 | 3 | < 2 | [8 empty] | 28
| 1 | 3 | 2 | 3.4 | 2.3 | < 3 | < 5 | 2.5535 | 3 | 2.2 | 2 | < 5 | < 5 | 2 | < 2 | [8 empty] | 16
| 1 | 4 | 2 | 3.3 | 2.9 | < 3 | < 5 | 2.3175 | 3 | 1.2 | 2 | < 5 | < 5 | 3 | < 2 | [8 empty] | 2.5087
| 1 | 5 | 2 | 3.4 | 1.7 | 3 | < 5 | 2.359 | 3 | 2.5 | 2 | < 5 | < 5 | 3 | < 2 | [8 empty] | 154
| 1 | 6 | 3 | 3.3 | 2.4 | < 3 | < 5 | 2.4453 | 3 | 2.4 | 2 | < 5 | < 5 | 2 | < 2
| Mean |  | 2.01: 2.3333; 2.02: 3.3333; 2.04: 2.3333; 2.05: 3; 2.07: 2.4232; 2.08: 3; 2.09: 2.2667; 2.10: 2; 2.13: 2.5
| Median |  | 2.01: 2; 2.02: 3.3; 2.04: 2.35; 2.05: 3; 2.07: 2.4022; 2.08: 3; 2.09: 2.4; 2.10: 2; 2.13: 2.5
| Std Dev. |  | 2.01: 0.5164; 2.02: 0.05164; 2.04: 0.38297; 2.07: 0.15707; 2.08: 0; 2.09: 0.57155; 2.10: 0; 2.13: 0.54772
| Rel.Std.Dev. |  | 2.01: 22.13%; 2.02: 1.549%; 2.04: 16.41%; 2.07: 6.482%; 2.08: 0%; 2.09: 25.22%; 2.10: 0%; 2.13: 21.91%
| PDM3 |  | 2.01: -6.992%; 2.02: 32.87%; 2.04: -6.992%; 2.05: 19.58%; 2.07: -3.409%; 2.08: 19.58%; 2.09: -9.649%; 2.10: -20.28%; 2.13: -0.3482%
| Z-Score (Absolute) |  | 0.67 | 6.83 | 0.67 | 5.05 | 0.58 | 0 | 4.33 | 1.17 | 3.18 | 0.58 | 0.58 | 0.58 | 10.68
| Analytical results for Bi in OREAS 922b (Certified Value 12.8 ppm) | [22 empty] | Certified
Bi | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.23; 2.09: 2.24 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES/MS |  | 1
2
| 1 | 1 | 2.01: 12.3; 2.02: 9.1; 2.04: 12; 2.05: 13.1; 2.06: 12.7555; 2.07: 13; 2.08: 13.1; 2.09: 12.7 |  | 1
| 1 | 2 | 2.01: 12.5; 2.02: 9.8; 2.04: 12; 2.05: 12.9; 2.06: 11.7164; 2.07: 13; 2.08: 13; 2.09: 13.3 |  | 29
| 1 | 3 | 2.01: 12.6; 2.02: 9.9; 2.04: 12; 2.05: 12.6; 2.06: 11.5785; 2.07: 13; 2.08: 13.1; 2.09: 12.7 |  | 16
| 1 | 4 | 2.01: 12.6; 2.02: 9.9; 2.04: 12; 2.05: 13.1; 2.06: 12.972; 2.07: 13.7; 2.08: 12.6; 2.09: 12.8 |  | 12.8155
| 1 | 5 | 2.01: 12.8; 2.02: 7.3; 2.04: 13; 2.05: 12; 2.06: 12.0497; 2.07: 13.7; 2.08: 13.2; 2.09: 12.6 |  | 155
| 1 | 6 | 2.01: 12.8; 2.02: 7.8; 2.04: 13; 2.05: 13.2; 2.06: 12.0845; 2.07: 13.1; 2.08: 13.3; 2.09: 13.8
| Mean |  | 2.01: 12.6; 2.02: 8.9667; 2.04: 12.3333; 2.05: 12.8167; 2.06: 12.1928; 2.07: 13.25; 2.08: 13.05; 2.09: 12.9833
| Median |  | 2.01: 12.6; 2.02: 9.45; 2.04: 12; 2.05: 13; 2.06: 12.0671; 2.07: 13.05; 2.08: 13.1; 2.09: 12.75
| Std Dev. |  | 2.01: 0.18974; 2.02: 1.1483; 2.04: 0.5164; 2.05: 0.4535; 2.06: 0.55863; 2.07: 0.35071; 2.08: 0.2429; 2.09: 0.47081
| Rel.Std.Dev. |  | 2.01: 1.506%; 2.02: 12.81%; 2.04: 4.187%; 2.05: 3.538%; 2.06: 4.582%; 2.07: 2.647%; 2.08: 1.861%; 2.09: 3.626%
| PDM3 |  | 2.01: -1.681%; 2.02: -30.03%; 2.04: -3.762%; 2.05: 0.009422%; 2.06: -4.859%; 2.07: 3.391%; 2.08: 1.83%; 2.09: 1.31%
| Z-Score (Absolute) |  | 2.01: 0.63; 2.02: 11.13; 2.04: NA; 2.05: 0; 2.06: 1.8; 2.07: 1.25; 2.08: 0.67; 2.09: 0.48
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 922b (Certified Value 0.053 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS |  |  |  | 3
3
| 1 | 1 | < 0.1 | 0.067779 | 0.04 | 0.04 | < 0.05 | < 0.05 | < 0.05 | 0.057 | 0.064 | 0.08 | < 0.1 | < 0.1 |  | 0.13 | 0.043 | 0.05 | 0.098 | < 0.1 |  |  |  | 1
| 1 | 2 | < 0.1 | 0.066896 | 0.04 | 0.04 | < 0.05 | < 0.05 | 0.05 | < 0.05 | 0.05 | 0.05 | < 0.1 | < 0.1 | 0.2 | 0.21 | 0.043 | 0.05 | 0.103 | < 0.1 |  |  |  | 0
| 1 | 3 | < 0.1 | 0.066552 | 0.05 | 0.04 | < 0.05 | < 0.05 | 0.05 | < 0.05 | 0.064 | 0.06 | < 0.1 | < 0.1 |  | 0.2 | 0.043 | 0.05 | 0.098 | < 0.1 |  |  |  | 16
| 1 | 4 | < 0.1 | 0.067067 | 0.05 | 0.04 | < 0.05 | < 0.05 | 0.1 | < 0.05 | 0.05 | 0.06 | < 0.1 | < 0.1 |  | 0.21 | 0.043 | 0.05 | 0.106 | < 0.1 |  |  |  | 0.053466
| 1 | 5 | < 0.1 | 0.067182 | 0.05 | 0.04 | < 0.05 | < 0.05 | 0.079 | < 0.05 | 0.05 | 0.08 | < 0.1 | < 0.1 |  | 0.19 | 0.043 | 0.05 | 0.113 | < 0.1 |  |  |  | 156
| 1 | 6 | < 0.1 | 0.067112 | 0.05 | 0.04 | < 0.05 | < 0.05 | 0.057 | 0.05 | 0.057 | 0.1 | < 0.1 | < 0.1 |  | 0.17 | 0.043 | 0.05 | 0.102 | < 0.1
| Mean |  |  | 0.067098 | 0.046667 | 0.04 |  |  | 0.0672 | 0.0535 | 0.055833 | 0.071667 |  |  | 0.2 | 0.185 | 0.043 | 0.05 | 0.10333
| Median |  |  | 0.067089 | 0.05 | 0.04 |  |  | 0.057 | 0.0535 | 0.0535 | 0.07 |  |  | 0.2 | 0.195 | 0.043 | 0.05 | 0.1025
| Std Dev. |  |  | 0.00040271 | 0.005164 | 0 |  |  | 0.021856 | 0.0049497 | 0.0068823 | 0.018348 |  |  |  | 0.030822 | 7.6012e-18 | 7.6012e-18 | 0.0056451
| Rel.Std.Dev. |  |  | 0.6002% | 11.07% | 0% |  |  | 32.52% | 9.252% | 12.33% | 25.6% |  |  |  | 16.66% | 1.768e-14% | 1.52e-14% | 5.463%
| PDM3 |  |  | 25.5% | -12.72% | -25.19% |  |  | 25.69% | 0.06338% | 4.428% | 34.04% |  |  | 274.1% | 246% | -19.58% | -6.483% | 93.27%
| Z-Score (Absolute) |  | 0 | 1.15 | 0.22 | 0.67 | 1.68 | 1.68 | 0.68 | 1.04 | 0.4 | 1.46 | 0 | 0 | NA | 9.1 | 0.48 | 0 | 3.6 | 0
Results from laboratory 2.16 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 922b (Certified Value < 10 ppm) | [22 empty] | Certified
Cd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.12; 2.04: 2.13; 2.05: 2.14; 2.07: 2.17; 2.08: 2.21; 2.09: 2.24 |  |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: PF*OES/MS; 2.04: PF*MS; 2.05: PF*MS; 2.07: PF*OES; 2.08: PF*OES; 2.09: PF*OES/MS |  |  |  | 1
1
| 1 | 1 | 2.01: < 1; 2.02: < 2; 2.04: < 0.2; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: 2 |  |  |  | 1
| 1 | 2 | 2.01: < 1; 2.02: < 2; 2.04: < 0.2; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: < 1 |  |  |  | 11
| 1 | 3 | 2.01: < 1; 2.02: < 2; 2.04: < 0.2; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: 3 |  |  |  | 16
| 1 | 4 | 2.01: < 1; 2.02: < 2; 2.04: < 0.2; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: 2 |  |  |  | < 10
| 1 | 5 | 2.01: < 1; 2.02: < 2; 2.04: < 0.2; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: 1 |  |  |  | 157
| 1 | 6 | 2.01: < 1; 2.02: < 2; 2.04: < 0.2; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: < 1
| Mean |  | 2.09: 2
| Median |  | 2.09: 2
| Std Dev. |  | 2.09: 0.8165
| Rel.Std.Dev. |  | 2.09: 40.82%
PDM3
| Z-Score (Absolute) |  | 2.01: 0.48; 2.02: 0.24; 2.04: 0.67; 2.05: 1.69; 2.07: 11.32; 2.08: 1.69; 2.09: 0
| Analytical results for Ce in OREAS 922b (Certified Value 81 ppm) | [22 empty] | Certified
Ce | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  |  | 0
1
| 1 | 1 | 2.01: 81; 2.02: 77.1; 2.04: 81.5496; 2.05: 86.5; 2.07: 81.9; 2.08: 85; 2.09: 76; 2.10: 78.2; 2.11: 80.1 |  |  |  | 1
| 1 | 2 | 2.01: 82.5; 2.02: 79.6; 2.04: 79.6893; 2.05: 80.2; 2.07: 80.5; 2.08: 82; 2.09: 76.9; 2.10: 79.8; 2.11: 83.9 |  |  |  | 12
| 1 | 3 | 2.01: 82.7; 2.02: 77.9; 2.04: 79.5967; 2.05: 81.1; 2.07: 79.6; 2.08: 84; 2.09: 78.5; 2.10: 79.4; 2.11: 81.4 |  |  |  | 16
| 1 | 4 | 2.01: 79.3; 2.02: 76.1; 2.04: 77.6862; 2.05: 82.3; 2.07: 85.7; 2.08: 79; 2.09: 76.5; 2.10: 78.1; 2.11: 80.5 |  |  |  | 80.5077
| 1 | 5 | 2.01: 83.1; 2.02: 77.1; 2.04: 80.8058; 2.05: 83.7; 2.07: 79.5; 2.08: 81; 2.09: 78.3; 2.10: 79.6; 2.11: 80.8 |  |  |  | 158
| 1 | 6 | 2.01: 81.5; 2.02: 78.2; 2.04: 80.3274; 2.05: 90; 2.07: 83.8; 2.08: 85; 2.09: 78.4; 2.10: 79.8; 2.11: 81.9
| Mean |  | 2.01: 81.6833; 2.02: 77.6667; 2.04: 79.9425; 2.05: 83.9667; 2.07: 81.8333; 2.08: 82.6667; 2.09: 77.4333; 2.10: 79.15; 2.11: 81.4333
| Median |  | 2.01: 82; 2.02: 77.5; 2.04: 80.0083; 2.05: 83; 2.07: 81.2; 2.08: 83; 2.09: 77.6; 2.10: 79.5; 2.11: 81.1
| Std Dev. |  | 2.01: 1.4063; 2.02: 1.1978; 2.04: 1.3229; 2.05: 3.6898; 2.07: 2.4913; 2.08: 2.4221; 2.09: 1.0985; 2.10: 0.7893; 2.11: 1.3677
| Rel.Std.Dev. |  | 2.01: 1.722%; 2.02: 1.542%; 2.04: 1.655%; 2.05: 4.394%; 2.07: 3.044%; 2.08: 2.93%; 2.09: 1.419%; 2.10: 0.9972%; 2.11: 1.68%
| PDM3 |  | 2.01: 1.46%; 2.02: -3.529%; 2.04: -0.702%; 2.05: 4.296%; 2.07: 1.647%; 2.08: 2.682%; 2.09: -3.819%; 2.10: -1.686%; 2.11: 1.15%
| Z-Score (Absolute) |  | 2.01: 0.11; 2.02: 1.7; 2.04: 0.67; 2.05: 1.15; 2.07: 0.18; 2.08: 0.56; 2.09: 1.81; 2.10: 1.03; 2.11: 0
| Analytical results for Co in OREAS 922b (Certified Value 3.7 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  | ppm
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS |  |  | 2
3
| 1 | 1 | 4 | 8.2 | 4 | 4.0568 | < 20 | < 20 | < 20 | < 20 | < 20 | 3.2 | 3.8 | 3.3115 | 3 | < 10 | 3.6 | 8 | < 10 | 4 | 4 |  |  | 1
| 1 | 2 | 4 | 7.5 | 3.7 | 4.2485 | < 20 | < 20 | < 20 | < 20 | < 20 | 3.9 | 3.8 | 3.3075 | 4 | < 10 | 3.5 | 8 | < 10 | 4 | 5 |  |  | 32
| 1 | 3 | 4 | 8.3 | 3.7 | 4.1761 | < 20 | < 20 | < 20 | < 20 | < 20 | 3.7 | 3.5 | 3.3071 | 4 | < 10 | 3.7 | 8 | < 10 | 4 | 4 |  |  | 16
| 1 | 4 | 4 | 11 | 3.8 | 4.1554 | < 20 | < 20 | < 20 | < 20 | < 20 | 3.6 | 3.6 | 3.3711 | 4 | < 10 | 3.6 | 8 | < 10 | 4 | 4 |  |  | 3.7023
| 1 | 5 | 4 | 8.8 | 3.7 | 4.0191 | < 20 | < 20 | < 20 | < 20 | < 20 | 3.8 | 3.5 | 3.3601 | 4 | < 10 | 3.7 | 8 | < 10 | 4 | 4 |  |  | 159
| 1 | 6 | 4 | 8.5 | 3.6 | 4.0037 | < 20 | < 20 | < 20 | < 20 | < 20 | 4.6 | 3.8 | 3.2673 | 4 | < 10 | 3.6 | 8 | < 10 | 4 | 4
| Mean |  | 4 | 8.7167 | 3.75 | 4.1099 | [5 empty] | 3.8 | 3.6667 | 3.3208 | 3.8333 |  | 3.6167 | 8 |  | 4 | 4.1667
| Median |  | 4 | 8.4 | 3.7 | 4.1061 | [5 empty] | 3.75 | 3.7 | 3.3095 | 4 |  | 3.6 | 8 |  | 4 | 4
| Std Dev. |  | 0 | 1.199 | 0.13784 | 0.097966 | [5 empty] | 0.46043 | 0.15055 | 0.038439 | 0.40825 |  | 0.075277 | 0 |  | 0 | 0.40825
| Rel.Std.Dev. |  | 0% | 13.76% | 3.676% | 2.384% | [5 empty] | 12.12% | 4.106% | 1.158% | 10.65% |  | 2.081% | 0% |  | 0% | 9.798%
| PDM3 |  | 8.04% | 135.4% | 1.287% | 11.01% | [5 empty] | 2.638% | -0.9635% | -10.31% | 3.538% |  | -2.314% | 116.1% |  | 8.04% | 12.54%
| Z-Score (Absolute) |  | NA | 1.64 | 0.55 | 0.39 | 2.2 | 2.2 | 2.2 | 2.2 | 2.2 | 0.53 | 0.59 | 0.74 | NA | 0 | 0.61 | NA | 0 | NA | NA
Results from laboratories 2.01, 2.16, 2.20, 2.23 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr2O3 in OREAS 922b (Certified Value 130 ppm) | [22 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES |  |  |  |  | 0
0
| 1 | 1 | 102 | 130.1 | 129.7463 | 100 | 200 | 200 | 100 | 300 | 132 | 149.1 | 105.5771 | 128.6 | 150 | 120 | 121.3 | 149.1 | 131.5 |  |  |  |  | 1
| 1 | 2 | 117 | 134.5 | 131.2622 | 100 | 200 | 200 | 100 | 200 | 146 | 137.4 | 112.6551 | 137.4 | 150 | 120 | 121.3 | 147.6 | 127.2 |  |  |  |  | 0
| 1 | 3 | 102 | 138.8 | 133.6693 | 200 | 200 | 200 | 100 | 200 | 146 | 124.2 | 104.5007 | 128.6 | 150 | 130 | 121.3 | 149.1 | 119.8 |  |  |  |  | 16
| 1 | 4 | 117 | 134.5 | 132.845 | 100 | 200 | 200 | 100 | 200 | 146 | 122.8 | 113.7682 | 131.5 | 150 | 120 | 122.8 | 149.1 | 121.3 |  |  |  |  | 130.3621
| 1 | 5 | 117 | 127.2 | 129.4318 | 100 | 200 | 200 | 100 | 100 | 132 | 127.2 | 120.9286 | 130.1 | 150 | 120 | 121.3 | 149.1 | 140.3 |  |  |  |  | 160
| 1 | 6 | 117 | 130.1 | 130.9532 | 100 | 200 | 200 | 100 | 200 | 132 | 137.4 | 117.0895 | 127.2 | 150 | 120 | 122.8 | 150.5 | 125.7
| Mean |  | 112 | 132.5333 | 131.318 | 116.6667 | 200 | 200 | 100 | 200 | 139 | 133.0167 | 112.4199 | 130.5667 | 150 | 121.6667 | 121.8 | 149.0833 | 127.6333
| Median |  | 117 | 132.3 | 131.1077 | 100 | 200 | 200 | 100 | 200 | 139 | 132.3 | 113.2116 | 129.35 | 150 | 120 | 121.3 | 149.1 | 126.45
| Std Dev. |  | 7.746 | 4.1784 | 1.6748 | 40.8248 | 0 | 0 | 0 | 63.2456 | 7.6681 | 10.1296 | 6.4127 | 3.6566 | 0 | 4.0825 | 0.7746 | 0.91742 | 7.4939
| Rel.Std.Dev. |  | 6.916% | 3.153% | 1.275% | 34.99% | 0% | 0% | 0% | 31.62% | 5.517% | 7.615% | 5.704% | 2.801% | 0% | 3.355% | 0.636% | 0.6154% | 5.871%
| PDM3 |  | -14.09% | 1.666% | 0.7332% | -10.51% | 53.42% | 53.42% | -23.29% | 53.42% | 6.626% | 2.036% | -13.76% | 0.1569% | 15.06% | -6.67% | -6.568% | 14.36% | -2.093%
| Z-Score (Absolute) |  | 1.73 | 0.11 | 0 | NA | NA | NA | NA | NA | 0.67 | 0.15 | 1.69 | 0.07 | 1.33 | NA | 0.85 | 1.59 | 0.33
Results from laboratory 2.19 were removed due to their 10 ppm reading resolution. Results from laboratories 2.07, 2.08, 2.09, 2.10 and 2.11 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 922b (Certified Value 4.78 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 4.48 | 4.9 | 5.251 | 5.8 | 4.9 | 4.3367 | 4.9 | 8 | 4.7 | 4.3 | 4.5 | [10 empty] | 1
| 1 | 2 | 4.61 | 4.7 | 5.2706 | 5.2 | 5 | 3.9252 | 4.9 | 7 | 4.7 | 4.1 | 4.9 | [10 empty] | 14
| 1 | 3 | 4.38 | 4.8 | 5.179 | 5.4 | 4.9 | 4.041 | 4.9 | 6 | 4.8 | 4.4 | 4.8 | [10 empty] | 16
| 1 | 4 | 4.57 | 5.1 | 5.3405 | 5.6 | 4.8 | 4.541 | 4.9 | 6 | 4.8 | 4.2 | 4.7 | [10 empty] | 4.7758
| 1 | 5 | 4.55 | 4.8 | 5.3269 | 5.4 | 4.5 | 4.1556 | 4.6 | 5 | 4.8 | 4.6 | 4.7 | [10 empty] | 161
| 1 | 6 | 4.54 | 4.8 | 5.4156 | 5.2 | 4.9 | 4.3972 | 5.6 | 6 | 4.8 | 3.9 | 4.7
| Mean |  | 4.5217 | 4.85 | 5.2973 | 5.4333 | 4.8333 | 4.2328 | 4.9667 | 6.3333 | 4.7667 | 4.25 | 4.7167
| Median |  | 4.545 | 4.8 | 5.2987 | 5.4 | 4.9 | 4.2462 | 4.9 | 6 | 4.8 | 4.25 | 4.7
| Std Dev. |  | 0.081343 | 0.13784 | 0.081985 | 0.23381 | 0.17512 | 0.23246 | 0.33267 | 1.0328 | 0.05164 | 0.2429 | 0.13292
| Rel.Std.Dev. |  | 1.799% | 2.842% | 1.548% | 4.303% | 3.623% | 5.492% | 6.698% | 16.31% | 1.083% | 5.715% | 2.818%
| PDM3 |  | -5.322% | 1.553% | 10.92% | 13.77% | 1.204% | -11.37% | 3.996% | 32.61% | -0.1921% | -11.01% | -1.239%
| Z-Score (Absolute) |  | 0.84 | 0.15 | 1.51 | 1.92 | 0.1 | 1.72 | 0.51 | NA | 0.1 | 1.67 | 0.25
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 922b (Certified Value 0.22 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.221 | 0.20849 | 0.217 | 0.221 | 0.211 | 0.221 | 0.217 | 0.209 | 0.227 | 0.2203 | 0.22006 | 0.2223 | 0.223 | 0.225 | 0.2258 | 0.225 | 0.227 |  |  |  |  | 1
| 1 | 2 | 0.221 | 0.20076 | 0.2134 | 0.222 | 0.211 | 0.223 | 0.212 | 0.21 | 0.226 | 0.2195 | 0.22065 | 0.219 | 0.234 | 0.2252 | 0.2296 | 0.2279 | 0.2358 |  |  |  |  | 35
| 1 | 3 | 0.221 | 0.20381 | 0.2214 | 0.213 | 0.214 | 0.224 | 0.213 | 0.211 | 0.221 | 0.2196 | 0.22355 | 0.2215 | 0.229 | 0.2258 | 0.2217 | 0.2285 | 0.2237 |  |  |  |  | 16
| 1 | 4 | 0.222 | 0.20734 | 0.2143 | 0.21 | 0.212 | 0.224 | 0.204 | 0.214 | 0.219 | 0.2194 | 0.22065 | 0.2217 | 0.23 | 0.2264 | 0.2274 | 0.227 | 0.2185 |  |  |  |  | 0.2197
| 1 | 5 | 0.22 | 0.20629 | 0.2117 | 0.224 | 0.213 | 0.224 | 0.212 | 0.214 | 0.224 | 0.2212 | 0.21604 | 0.2221 | 0.234 | 0.2252 | 0.2281 | 0.2257 | 0.2219 |  |  |  |  | 162
| 1 | 6 | 0.221 | 0.20663 | 0.214 | 0.222 | 0.204 | 0.218 | 0.208 | 0.212 | 0.227 | 0.2185 | 0.21706 | 0.2193 | 0.225 | 0.2255 | 0.225 | 0.2239 | 0.233
| Mean |  | 0.221 | 0.20556 | 0.2153 | 0.21867 | 0.21083 | 0.22233 | 0.211 | 0.21167 | 0.224 | 0.21975 | 0.21967 | 0.22098 | 0.22917 | 0.22552 | 0.22627 | 0.22633 | 0.22665
| Median |  | 0.221 | 0.20646 | 0.21415 | 0.2215 | 0.2115 | 0.2235 | 0.212 | 0.2115 | 0.225 | 0.21955 | 0.22035 | 0.2216 | 0.2295 | 0.22535 | 0.2266 | 0.22635 | 0.22535
| Std Dev. |  | 0.00063246 | 0.0028116 | 0.0034456 | 0.0057155 | 0.0035449 | 0.0024221 | 0.0044721 | 0.0020656 | 0.0033466 | 0.00091378 | 0.0027232 | 0.0014511 | 0.004535 | 0.00051543 | 0.0027725 | 0.0017716 | 0.0066612
| Rel.Std.Dev. |  | 0.2862% | 1.368% | 1.6% | 2.614% | 1.681% | 1.089% | 2.119% | 0.9759% | 1.494% | 0.4158% | 1.24% | 0.6567% | 1.979% | 0.2286% | 1.225% | 0.7828% | 2.939%
| PDM3 |  | 0.5934% | -6.436% | -2.001% | -0.4687% | -4.034% | 1.2% | -3.958% | -3.655% | 1.959% | 0.02444% | -0.01314% | 0.5858% | 4.311% | 2.649% | 2.991% | 3.021% | 3.165%
| Z-Score (Absolute) |  | 0 | 1.97 | 0.73 | 0.3 | 1.3 | 0.17 | 1.27 | 1.19 | 0.39 | 0.16 | 0.17 | 0 | 1.04 | 0.58 | 0.67 | 0.68 | 0.72
| Analytical results for Dy in OREAS 922b (Certified Value 4.79 ppm) | [22 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  |  | 1
| No | No | 2.01: 2.04; 2.05: 2.05; 2.06: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.18; 2.10: 2.19; 2.11: 2.22; 2.12: 2.24 |  |  |  |  | ppm
|  |  | 2.01: PF*MS; 2.05: BF*MS; 2.06: PF*OES/MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: BF*OES/MS; 2.11: BF*MS; 2.12: PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 4.6; 2.05: 4.8; 2.06: 5.0767; 2.07: 5.1; 2.08: 4.84; 2.09: 5; 2.10: 4.56; 2.11: 4.74; 2.12: 4.8 |  |  |  |  | 1
| 1 | 2 | 2.01: 4.4; 2.05: 5.1; 2.06: 5.1853; 2.07: 4.6; 2.08: 4.76; 2.09: 5; 2.10: 4.58; 2.11: 4.8; 2.12: 4.74 |  |  |  |  | 3
| 1 | 3 | 2.01: 4.7; 2.05: 5.2; 2.06: 4.8978; 2.07: 4.7; 2.08: 4.54; 2.09: 5; 2.10: 4.69; 2.11: 4.57; 2.12: 4.72 |  |  |  |  | 16
| 1 | 4 | 2.01: 4.3; 2.05: 4.6; 2.06: 5.0083; 2.07: 5.2; 2.08: 4.73; 2.09: 4; 2.10: 4.55; 2.11: 4.86; 2.12: 4.75 |  |  |  |  | 4.788
| 1 | 5 | 2.01: 4.7; 2.05: 5; 2.06: 5.1093; 2.07: 4.7; 2.08: 4.74; 2.09: 5; 2.10: 4.57; 2.11: 4.91; 2.12: 4.74 |  |  |  |  | 163
| 1 | 6 | 2.01: 4.7; 2.05: 5; 2.06: 4.915; 2.07: 5; 2.08: 4.88; 2.09: 5; 2.10: 4.66; 2.11: 4.68; 2.12: 4.82
| Mean |  | 2.01: 4.5667; 2.05: 4.95; 2.06: 5.0321; 2.07: 4.8833; 2.08: 4.7483; 2.09: 4.8333; 2.10: 4.6017; 2.11: 4.76; 2.12: 4.7617
| Median |  | 2.01: 4.65; 2.05: 5; 2.06: 5.0425; 2.07: 4.85; 2.08: 4.75; 2.09: 5; 2.10: 4.575; 2.11: 4.77; 2.12: 4.745
| Std Dev. |  | 2.01: 0.17512; 2.05: 0.21679; 2.06: 0.11291; 2.07: 0.24833; 2.08: 0.11805; 2.09: 0.40825; 2.10: 0.058452; 2.11: 0.1241; 2.12: 0.0392
| Rel.Std.Dev. |  | 2.01: 3.835%; 2.05: 4.38%; 2.06: 2.244%; 2.07: 5.085%; 2.08: 2.486%; 2.09: 8.447%; 2.10: 1.27%; 2.11: 2.607%; 2.12: 0.8232%
| PDM3 |  | 2.01: -4.622%; 2.05: 3.384%; 2.06: 5.098%; 2.07: 1.992%; 2.08: -0.8278%; 2.09: 0.9475%; 2.10: -3.891%; 2.11: -0.5841%; 2.12: -0.5493%
| Z-Score (Absolute) |  | 2.01: 0.93; 2.05: 0.91; 2.06: 1.3; 2.07: 0.59; 2.08: 0.06; 2.09: NA; 2.10: 0.76; 2.11: 0; 2.12: 0
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 922b (Certified Value 2.75 ppm) | [22 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  |  |  | 1
| No | No | 2.01: 2.04; 2.05: 2.05; 2.06: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.18; 2.10: 2.19; 2.11: 2.22; 2.12: 2.24 |  |  |  |  | ppm
|  |  | 2.01: PF*MS; 2.05: BF*MS; 2.06: PF*OES/MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: BF*OES/MS; 2.11: BF*MS; 2.12: PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 2.6; 2.05: 2.8; 2.06: 2.8126; 2.07: 2.6; 2.08: 2.89; 2.09: 3; 2.10: 2.82; 2.11: 2.81; 2.12: 2.87 |  |  |  |  | 1
| 1 | 2 | 2.01: 2.4; 2.05: 2.9; 2.06: 2.7143; 2.07: 2.5; 2.08: 2.82; 2.09: 3; 2.10: 2.79; 2.11: 2.76; 2.12: 3.02 |  |  |  |  | 4
| 1 | 3 | 2.01: 2.6; 2.05: 2.7; 2.06: 2.828; 2.07: 2.7; 2.08: 2.58; 2.09: 3; 2.10: 2.85; 2.11: 2.82; 2.12: 2.83 |  |  |  |  | 16
| 1 | 4 | 2.01: 2.4; 2.05: 2.5; 2.06: 2.8178; 2.07: 2.8; 2.08: 2.69; 2.09: 3; 2.10: 2.76; 2.11: 2.71; 2.12: 3 |  |  |  |  | 2.7486
| 1 | 5 | 2.01: 2.6; 2.05: 2.8; 2.06: 2.6879; 2.07: 2.5; 2.08: 2.61; 2.09: 3; 2.10: 2.92; 2.11: 2.88; 2.12: 2.92 |  |  |  |  | 164
| 1 | 6 | 2.01: 2.7; 2.05: 2.8; 2.06: 2.7198; 2.07: 2.7; 2.08: 2.72; 2.09: 3; 2.10: 2.9; 2.11: 2.63; 2.12: 2.85
| Mean |  | 2.01: 2.55; 2.05: 2.75; 2.06: 2.7634; 2.07: 2.6333; 2.08: 2.7183; 2.09: 3; 2.10: 2.84; 2.11: 2.7683; 2.12: 2.915
| Median |  | 2.01: 2.6; 2.05: 2.8; 2.06: 2.7662; 2.07: 2.65; 2.08: 2.705; 2.09: 3; 2.10: 2.835; 2.11: 2.785; 2.12: 2.895
| Std Dev. |  | 2.01: 0.12247; 2.05: 0.13784; 2.06: 0.06254; 2.07: 0.12111; 2.08: 0.11957; 2.09: 0; 2.10: 0.06229; 2.11: 0.088863; 2.12: 0.079687
| Rel.Std.Dev. |  | 2.01: 4.803%; 2.05: 5.012%; 2.06: 2.263%; 2.07: 4.599%; 2.08: 4.399%; 2.09: 0%; 2.10: 2.193%; 2.11: 3.21%; 2.12: 2.734%
| PDM3 |  | 2.01: -7.224%; 2.05: 0.05269%; 2.06: 0.5408%; 2.07: -4.192%; 2.08: -1.099%; 2.09: 9.148%; 2.10: 3.327%; 2.11: 0.7197%; 2.12: 6.056%
| Z-Score (Absolute) |  | 2.01: 2.29; 2.05: 0.07; 2.06: 0.07; 2.07: 1.37; 2.08: 0.43; 2.09: NA; 2.10: 0.92; 2.11: 0.13; 2.12: 1.75
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 922b (Certified Value 1.3 ppm) | [22 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  |  |  | 1
| No | No | 2.01: 2.04; 2.05: 2.05; 2.06: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.18; 2.10: 2.19; 2.11: 2.24 |  |  |  |  | ppm
|  |  | 2.01: PF*MS; 2.05: BF*MS; 2.06: PF*OES/MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: BF*OES/MS; 2.11: PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.01: 1.3; 2.05: 1.2; 2.06: 1.8003; 2.07: 1.4; 2.08: 1.27; 2.09: < 2; 2.10: 1.19; 2.11: 1.29 |  |  |  |  | 1
| 1 | 2 | 2.01: 1.4; 2.05: 1.3; 2.06: 1.7365; 2.07: 1.3; 2.08: 1.17; 2.09: < 2; 2.10: 1.21; 2.11: 1.35 |  |  |  |  | 5
| 1 | 3 | 2.01: 1.4; 2.05: 1.2; 2.06: 1.7911; 2.07: 1.3; 2.08: 1.21; 2.09: < 2; 2.10: 1.25; 2.11: 1.31 |  |  |  |  | 16
| 1 | 4 | 2.01: 1.4; 2.05: 1.4; 2.06: 1.8014; 2.07: 1.3; 2.08: 1.26; 2.09: < 2; 2.10: 1.24; 2.11: 1.34 |  |  |  |  | 1.2989
| 1 | 5 | 2.01: 1.3; 2.05: 1.4; 2.06: 1.7906; 2.07: 1.7; 2.08: 1.22; 2.09: < 2; 2.10: 1.3; 2.11: 1.3 |  |  |  |  | 165
| 1 | 6 | 2.01: 1.4; 2.05: 1.3; 2.06: 1.7459; 2.07: 1.3; 2.08: 1.3; 2.09: < 2; 2.10: 1.28; 2.11: 1.35
| Mean |  | 2.01: 1.3667; 2.05: 1.3; 2.06: 1.7776; 2.07: 1.3833; 2.08: 1.2383; 2.10: 1.245; 2.11: 1.3233
| Median |  | 2.01: 1.4; 2.05: 1.3; 2.06: 1.7909; 2.07: 1.3; 2.08: 1.24; 2.10: 1.245; 2.11: 1.325
| Std Dev. |  | 2.01: 0.05164; 2.05: 0.089443; 2.06: 0.02873; 2.07: 0.16021; 2.08: 0.047081; 2.10: 0.041352; 2.11: 0.026583
| Rel.Std.Dev. |  | 2.01: 3.779%; 2.05: 6.88%; 2.06: 1.616%; 2.07: 11.58%; 2.08: 3.802%; 2.10: 3.321%; 2.11: 2.009%
| PDM3 |  | 2.01: 5.218%; 2.05: 0.08554%; 2.06: 36.86%; 2.07: 6.501%; 2.08: -4.662%; 2.10: -4.149%; 2.11: 1.882%
| Z-Score (Absolute) |  | 2.01: 0.54; 2.05: 0.11; 2.06: 4.54; 2.07: 0.7; 2.08: 0.71; 2.09: 3.04; 2.10: 0.65; 2.11: 0.11
| Analytical results for Fe in OREAS 922b (Certified Value 3.14 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 3.15 | 3.1288 | 3.09 | 3.11 | 3.1826 | 3.148 | 3.148 | 3.106 | 3.078 | 3.029 | 3.22 | 3.1 | 2.9967 | 3.17 | 3.18 | 3.3093 | 3.064 | 3.17 | 3.497 | 3.12 | 3.15 | 1
| 1 | 2 | 3.2 | 3.138 | 3.04 | 3.21 | 3.1936 | 3.099 | 3.127 | 3.168 | 3.113 | 3.029 | 3.19 | 3.08 | 2.8385 | 3.13 | 3.23 | 3.3669 | 3.099 | 3.16 | 3.497 | 3.15 | 3.32 | 0
| 1 | 3 | 3.14 | 3.1313 | 3.04 | 3.19 | 3.1922 | 3.029 | 3.162 | 3.182 | 3.071 | 3.057 | 3.19 | 3.12 | 2.8021 | 3.16 | 3.12 | 3.3137 | 3.106 | 3.17 | 3.513 | 3.17 | 3.22 | 16
| 1 | 4 | 3.17 | 3.1672 | 3.06 | 3.13 | 3.2204 | 3.05 | 3.113 | 3.134 | 2.973 | 3.092 | 3.2 | 3.11 | 3.0472 | 3.17 | 3.24 | 3.1875 | 3.148 | 3.16 | 3.502 | 3.14 | 3.18 | 3.1447
| 1 | 5 | 3.12 | 3.1508 | 3.05 | 3.06 | 3.1833 | 3.168 | 3.168 | 3.141 | 3.022 | 3.113 | 3.2 | 3.12 | 2.9465 | 3.18 | 3.31 | 3.1882 | 3.127 | 3.18 | 3.524 | 3.13 | 3.11 | 166
| 1 | 6 | 3.18 | 3.1577 | 3.02 | 3.16 | 3.1869 | 3.085 | 3.029 | 3.113 | 3.043 | 3.05 | 3.22 | 3.09 | 2.9211 | 3.14 | 3.23 | 3.3501 | 3.064 | 3.18 | 3.501 | 3.11 | 3.25
| Mean |  | 3.16 | 3.1456 | 3.05 | 3.1433 | 3.1932 | 3.0965 | 3.1245 | 3.1407 | 3.05 | 3.0617 | 3.2033 | 3.1033 | 2.9253 | 3.1583 | 3.2183 | 3.286 | 3.1013 | 3.17 | 3.5057 | 3.1367 | 3.205
| Median |  | 3.16 | 3.1444 | 3.045 | 3.145 | 3.1896 | 3.092 | 3.1375 | 3.1375 | 3.057 | 3.0535 | 3.2 | 3.105 | 2.9338 | 3.165 | 3.23 | 3.3115 | 3.1025 | 3.17 | 3.5015 | 3.135 | 3.2
| Std Dev. |  | 0.028983 | 0.015384 | 0.023664 | 0.055015 | 0.014095 | 0.054077 | 0.051204 | 0.029891 | 0.0489 | 0.034197 | 0.013663 | 0.01633 | 0.092879 | 0.019408 | 0.063692 | 0.079033 | 0.033619 | 0.0089443 | 0.010727 | 0.021602 | 0.075033
| Rel.Std.Dev. |  | 0.9172% | 0.4891% | 0.7759% | 1.75% | 0.4414% | 1.746% | 1.639% | 0.9517% | 1.603% | 1.117% | 0.4265% | 0.5262% | 3.175% | 0.6145% | 1.979% | 2.405% | 1.084% | 0.2822% | 0.306% | 0.6887% | 2.341%
| PDM3 |  | 0.4876% | 0.0305% | -3.01% | -0.04242% | 1.543% | -1.532% | -0.6413% | -0.1272% | -3.01% | -2.639% | 1.866% | -1.314% | -6.975% | 0.4346% | 2.343% | 4.493% | -1.378% | 0.8056% | 11.48% | -0.2544% | 1.919%
| Z-Score (Absolute) |  | 0.24 | 0.03 | 1.34 | 0 | 0.71 | 0.67 | 0.27 | 0.04 | 1.34 | 1.17 | 0.86 | 0.57 | 3.12 | 0.21 | 1.07 | 2.04 | 0.61 | 0.38 | 5.19 | 0.1 | 0.88
| Analytical results for Ga in OREAS 922b (Certified Value 19.1 ppm) | [22 empty] | Certified
Ga | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.16; 2.10: 2.19; 2.11: 2.23; 2.12: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: BF*OES/MS; 2.11: PF*MS; 2.12: PF*OES/MS | 1
2
| 1 | 1 | 2.01: 19; 2.02: 15.9; 2.04: 18.7; 2.05: 19.5972; 2.06: 22.7; 2.07: 20; 2.08: 17.6611; 2.09: 19; 2.10: 17.5; 2.11: 19; 2.12: 21 | 1
| 1 | 2 | 2.01: 20; 2.02: 18.6; 2.04: 18.4; 2.05: 19.9138; 2.06: 19; 2.07: 20; 2.08: 17.1556; 2.09: 19; 2.10: 17.4; 2.11: 18; 2.12: 22 | 18
| 1 | 3 | 2.01: 20; 2.02: 16.2; 2.04: 19.3; 2.05: 19.3777; 2.06: 21.7; 2.07: 20; 2.08: 17.3812; 2.09: 19; 2.10: 18.1; 2.11: 19; 2.12: 21 | 16
| 1 | 4 | 2.01: 20; 2.02: 16.6; 2.04: 18.9; 2.05: 19.698; 2.06: 19.8; 2.07: 20; 2.08: 18.5525; 2.09: 19; 2.10: 17.7; 2.11: 20; 2.12: 21 | 19.1347
| 1 | 5 | 2.01: 20; 2.02: 16.9; 2.04: 19.2; 2.05: 19.1672; 2.06: 17.3; 2.07: 19; 2.08: 17.7688; 2.09: 19; 2.10: 17.7; 2.11: 19; 2.12: 21 | 167
| 1 | 6 | 2.01: 19; 2.02: 19.9; 2.04: 19.3; 2.05: 20.014; 2.06: 20.5; 2.07: 21; 2.08: 18.404; 2.09: 19; 2.10: 17.9; 2.11: 19; 2.12: 21
| Mean |  | 2.01: 19.6667; 2.02: 17.35; 2.04: 18.9667; 2.05: 19.628; 2.06: 20.1667; 2.07: 20; 2.08: 17.8205; 2.09: 19; 2.10: 17.7167; 2.11: 19; 2.12: 21.1667
| Median |  | 2.01: 20; 2.02: 16.75; 2.04: 19.05; 2.05: 19.6476; 2.06: 20.15; 2.07: 20; 2.08: 17.7149; 2.09: 19; 2.10: 17.7; 2.11: 19; 2.12: 21
| Std Dev. |  | 2.01: 0.5164; 2.02: 1.5656; 2.04: 0.36697; 2.05: 0.31984; 2.06: 1.9284; 2.07: 0.63246; 2.08: 0.55489; 2.09: 0; 2.10: 0.25626; 2.11: 0.63246; 2.12: 0.40825
| Rel.Std.Dev. |  | 2.01: 2.626%; 2.02: 9.023%; 2.04: 1.935%; 2.05: 1.63%; 2.06: 9.562%; 2.07: 3.162%; 2.08: 3.114%; 2.09: 0%; 2.10: 1.446%; 2.11: 3.329%; 2.12: 1.929%
| PDM3 |  | 2.01: 2.78%; 2.02: -9.327%; 2.04: -0.8782%; 2.05: 2.578%; 2.06: 5.393%; 2.07: 4.522%; 2.08: -6.868%; 2.09: -0.704%; 2.10: -7.411%; 2.11: -0.704%; 2.12: 10.62%
| Z-Score (Absolute) |  | 2.01: 0.45; 2.02: 1.11; 2.04: 0.02; 2.05: 0.42; 2.06: 0.79; 2.07: 0.67; 2.08: 0.8; 2.09: 0; 2.10: 0.87; 2.11: 0; 2.12: 1.46
| Analytical results for Gd in OREAS 922b (Certified Value 5.51 ppm) | [22 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.18; 2.08: 2.19; 2.09: 2.24 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: BF*OES/MS; 2.09: PF*OES/MS | 2
3
| 1 | 1 | 2.01: 5.7; 2.02: 5.47; 2.04: 5.3073; 2.05: 6.2; 2.06: 5.68; 2.07: 5; 2.08: 5.68; 2.09: 8 | 1
| 1 | 2 | 2.01: 5.6; 2.02: 5.41; 2.04: 5.6421; 2.05: 6; 2.06: 5.91; 2.07: 5; 2.08: 5.44; 2.09: 9 | 6
| 1 | 3 | 2.01: 5.8; 2.02: 5.82; 2.04: 5.3668; 2.05: 7.1; 2.06: 5.64; 2.07: 6; 2.08: 5.57; 2.09: 8 | 16
| 1 | 4 | 2.01: 5.5; 2.02: 5.46; 2.04: 5.2891; 2.05: 6.6; 2.06: 5.85; 2.07: 5; 2.08: 5.49; 2.09: 8 | 5.511
| 1 | 5 | 2.01: 5.7; 2.02: 5.41; 2.04: 5.4118; 2.05: 7.2; 2.06: 5.41; 2.07: 5; 2.08: 5.49; 2.09: 8 | 168
| 1 | 6 | 2.01: 5.8; 2.02: 5.47; 2.04: 5.3012; 2.05: 6.5; 2.06: 5.94; 2.07: 6; 2.08: 5.52; 2.09: 8
| Mean |  | 2.01: 5.6833; 2.02: 5.5067; 2.04: 5.3864; 2.05: 6.6; 2.06: 5.7383; 2.07: 5.3333; 2.08: 5.5317; 2.09: 8.1667
| Median |  | 2.01: 5.7; 2.02: 5.465; 2.04: 5.337; 2.05: 6.55; 2.06: 5.765; 2.07: 5; 2.08: 5.505; 2.09: 8
| Std Dev. |  | 2.01: 0.1169; 2.02: 0.15603; 2.04: 0.13371; 2.05: 0.47749; 2.06: 0.20173; 2.07: 0.5164; 2.08: 0.084242; 2.09: 0.40825
| Rel.Std.Dev. |  | 2.01: 2.057%; 2.02: 2.834%; 2.04: 2.482%; 2.05: 7.235%; 2.06: 3.516%; 2.07: 9.682%; 2.08: 1.523%; 2.09: 4.999%
| PDM3 |  | 2.01: 3.127%; 2.02: -0.07831%; 2.04: -2.261%; 2.05: 19.76%; 2.06: 4.125%; 2.07: -3.224%; 2.08: 0.3753%; 2.09: 48.19%
| Z-Score (Absolute) |  | 2.01: 0.29; 2.02: 0.39; 2.04: 0.85; 2.05: 3.8; 2.06: 0.5; 2.07: 1.05; 2.08: 0.29; 2.09: 9.81
| Analytical results for Ge in OREAS 922b (Certified Value 2.05 ppm) | [22 empty] | Certified
Ge | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.06; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.23; 2.09: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES/MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES/MS | 2
2
| 1 | 1 | 2.01: 2; 2.02: 2.2199; 2.04: 2.4; 2.05: 2; 2.06: 1.837; 2.07: 2; 2.08: 2; 2.09: 2 | 1
| 1 | 2 | 2.01: 2; 2.02: 2.2252; 2.04: 2.3; 2.05: 2; 2.06: 1.523; 2.07: 2; 2.08: 2; 2.09: 3 | 41
| 1 | 3 | 2.01: 2; 2.02: 2.2158; 2.04: 2.5; 2.05: 2; 2.06: 1.4405; 2.07: 2; 2.08: 2; 2.09: 3 | 16
| 1 | 4 | 2.01: 2; 2.02: 2.2447; 2.04: 2.5; 2.05: 2; 2.06: 1.8516; 2.07: 2; 2.08: 2; 2.09: 3 | 2.0527
| 1 | 5 | 2.01: 2; 2.02: 2.2277; 2.04: 2.4; 2.05: 2; 2.06: 1.627; 2.07: 2; 2.08: 2; 2.09: 3 | 169
| 1 | 6 | 2.01: 2; 2.02: 2.1704; 2.04: 2.7; 2.05: 2; 2.06: 1.8287; 2.07: 3; 2.08: 2; 2.09: 2
| Mean |  | 2.01: 2; 2.02: 2.2173; 2.04: 2.4667; 2.05: 2; 2.06: 1.6846; 2.07: 2.1667; 2.08: 2; 2.09: 2.6667
| Median |  | 2.01: 2; 2.02: 2.2226; 2.04: 2.45; 2.05: 2; 2.06: 1.7279; 2.07: 2; 2.08: 2; 2.09: 3
| Std Dev. |  | 2.01: 0; 2.02: 0.025034; 2.04: 0.13663; 2.05: 0; 2.06: 0.17939; 2.07: 0.40825; 2.08: 0; 2.09: 0.5164
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 1.129%; 2.04: 5.539%; 2.05: 0%; 2.06: 10.65%; 2.07: 18.84%; 2.08: 0%; 2.09: 19.36%
| PDM3 |  | 2.01: -2.565%; 2.02: 8.02%; 2.04: 20.17%; 2.05: -2.565%; 2.06: -17.93%; 2.07: 5.554%; 2.08: -2.565%; 2.09: 29.91%
| Z-Score (Absolute) |  | 2.01: 0.52; 2.02: 0.83; 2.04: 2.38; 2.05: 0.52; 2.06: 2.47; 2.07: 0.52; 2.08: 0.52; 2.09: 3.62
| Analytical results for Hf in OREAS 922b (Certified Value 3.63 ppm) | [22 empty] | Certified
Hf | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.19; 2.08: 2.22 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: BF*OES/MS; 2.08: BF*MS | 2
2
| 1 | 1 | 2.01: 2; 2.02: 3.9; 2.04: 3.7648; 2.05: < 10; 2.06: 4; 2.07: 3.7; 2.08: 2.7 | 1
| 1 | 2 | 2.01: 1.7; 2.02: 4.1; 2.04: 3.6987; 2.05: < 10; 2.06: 4; 2.07: 3.7; 2.08: 2.7 | 25
| 1 | 3 | 2.01: 1.7; 2.02: 3.8; 2.04: 3.6054; 2.05: < 10; 2.06: 4; 2.07: 3.8; 2.08: 2.8 | 16
| 1 | 4 | 2.01: 1.7; 2.02: 4; 2.04: 3.7362; 2.05: < 10; 2.06: 4; 2.07: 3.6; 2.08: 2.7 | 3.6265
| 1 | 5 | 2.01: 1.8; 2.02: 4; 2.04: 3.7086; 2.05: < 10; 2.06: 4; 2.07: 3.7; 2.08: 2.8 | 170
| 1 | 6 | 2.01: 2; 2.02: 4.1; 2.04: 3.6826; 2.05: < 10; 2.06: 4; 2.07: 3.7; 2.08: 2.8
| Mean |  | 2.01: 1.8167; 2.02: 3.9833; 2.04: 3.6994; 2.06: 4; 2.07: 3.7; 2.08: 2.75
| Median |  | 2.01: 1.75; 2.02: 4; 2.04: 3.7036; 2.06: 4; 2.07: 3.7; 2.08: 2.75
| Std Dev. |  | 2.01: 0.1472; 2.02: 0.1169; 2.04: 0.054497; 2.06: 0; 2.07: 0.063246; 2.08: 0.054772
| Rel.Std.Dev. |  | 2.01: 8.103%; 2.02: 2.935%; 2.04: 1.473%; 2.06: 0%; 2.07: 1.709%; 2.08: 1.992%
| PDM3 |  | 2.01: -49.91%; 2.02: 9.838%; 2.04: 2.009%; 2.06: 10.3%; 2.07: 2.025%; 2.08: -24.17%
| Z-Score (Absolute) |  | 2.01: 4.23; 2.02: 0.64; 2.04: 0; 2.05: 2.92; 2.06: 0.67; 2.07: 0; 2.08: 2.14
| Analytical results for Hg in OREAS 922b (Indicative Value < 5 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: PF*OES | 2
2
| 1 | 1 | 2.01: < 5 | 1
| 1 | 2 | 2.01: < 5 | 23
| 1 | 3 | 2.01: < 5 | 16
| 1 | 4 | 2.01: < 5 | < 5
| 1 | 5 | 2.01: < 5 | 29
| 1 | 6 | 2.01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Ho in OREAS 922b (Certified Value 0.97 ppm) | [22 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.18; 2.08: 2.19; 2.09: 2.22; 2.10: 2.24 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: BF*OES/MS; 2.09: BF*MS; 2.10: PF*OES/MS | 2
3
| 1 | 1 | 2.01: 0.8; 2.02: 1; 2.04: 1.0068; 2.05: 0.9; 2.06: 0.94; 2.07: < 2; 2.08: 0.93; 2.09: 0.98; 2.10: 0.98 | 1
| 1 | 2 | 2.01: 0.7; 2.02: 1; 2.04: 1.0175; 2.05: 1; 2.06: 0.97; 2.07: < 2; 2.08: 0.93; 2.09: 0.93; 2.10: 1.01 | 3
| 1 | 3 | 2.01: 0.8; 2.02: 1; 2.04: 0.97011; 2.05: 1; 2.06: 0.96; 2.07: < 2; 2.08: 0.97; 2.09: 0.97; 2.10: 0.99 | 16
| 1 | 4 | 2.01: 0.7; 2.02: 0.9; 2.04: 1.0022; 2.05: 1; 2.06: 0.94; 2.07: < 2; 2.08: 0.91; 2.09: 0.9; 2.10: 1 | 0.96641
| 1 | 5 | 2.01: 0.8; 2.02: 1; 2.04: 0.97941; 2.05: 0.9; 2.06: 0.94; 2.07: < 2; 2.08: 0.93; 2.09: 0.98; 2.10: 0.96 | 171
| 1 | 6 | 2.01: 0.8; 2.02: 1; 2.04: 1.0166; 2.05: 1.1; 2.06: 0.98; 2.07: < 2; 2.08: 0.94; 2.09: 0.93; 2.10: 0.97
| Mean |  | 2.01: 0.76667; 2.02: 0.98333; 2.04: 0.99877; 2.05: 0.98333; 2.06: 0.955; 2.08: 0.935; 2.09: 0.94833; 2.10: 0.985
| Median |  | 2.01: 0.8; 2.02: 1; 2.04: 1.0045; 2.05: 1; 2.06: 0.95; 2.08: 0.93; 2.09: 0.95; 2.10: 0.985
| Std Dev. |  | 2.01: 0.05164; 2.02: 0.040825; 2.04: 0.019705; 2.05: 0.075277; 2.06: 0.017607; 2.08: 0.019748; 2.09: 0.033116; 2.10: 0.018708
| Rel.Std.Dev. |  | 2.01: 6.736%; 2.02: 4.152%; 2.04: 1.973%; 2.05: 7.655%; 2.06: 1.844%; 2.08: 2.112%; 2.09: 3.492%; 2.10: 1.899%
| PDM3 |  | 2.01: -20.67%; 2.02: 1.751%; 2.04: 3.349%; 2.05: 1.751%; 2.06: -1.18%; 2.08: -3.25%; 2.09: -1.87%; 2.10: 1.924%
| Z-Score (Absolute) |  | 2.01: NA; 2.02: 0; 2.04: 0.62; 2.05: NA; 2.06: 1.15; 2.07: 0.67; 2.08: 1.96; 2.09: 1.42; 2.10: 0.07
Results from laboratories 2.04 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 922b (Certified Value 0.27 ppm) | [22 empty] | Certified
In | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.12; 2.04: 2.13; 2.05: 2.14; 2.06: 2.16; 2.07: 2.23; 2.08: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES/MS; 2.04: PF*MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*OES/MS | 2
2
| 1 | 1 | 2.01: 0.3; 2.02: 0.2; 2.04: 0.3; 2.05: 0.24032; 2.06: 0.3; 2.07: 0.3; 2.08: 0.3 | 1
| 1 | 2 | 2.01: 0.3; 2.02: 0.4; 2.04: 0.3; 2.05: 0.23128; 2.06: 0.3; 2.07: 0.3; 2.08: 0.3 | 11
| 1 | 3 | 2.01: 0.3; 2.02: < 0.2; 2.04: 0.3; 2.05: 0.23692; 2.06: 0.2; 2.07: 0.3; 2.08: 0.3 | 16
| 1 | 4 | 2.01: 0.3; 2.02: 0.3; 2.04: 0.3; 2.05: 0.23875; 2.06: 0.3; 2.07: 0.2; 2.08: < 0.3 | 0.26713
| 1 | 5 | 2.01: 0.3; 2.02: 0.3; 2.04: 0.2; 2.05: 0.24215; 2.06: 0.2; 2.07: 0.2; 2.08: < 0.3 | 172
| 1 | 6 | 2.01: 0.3; 2.02: 0.4; 2.04: 0.3; 2.05: 0.2301; 2.06: 0.2; 2.07: 0.3; 2.08: 0.3
| Mean |  | 2.01: 0.3; 2.02: 0.32; 2.04: 0.28333; 2.05: 0.23659; 2.06: 0.25; 2.07: 0.26667; 2.08: 0.3
| Median |  | 2.01: 0.3; 2.02: 0.3; 2.04: 0.3; 2.05: 0.23783; 2.06: 0.25; 2.07: 0.3; 2.08: 0.3
| Std Dev. |  | 2.01: 0; 2.02: 0.083666; 2.04: 0.040825; 2.05: 0.0048957; 2.06: 0.054772; 2.07: 0.05164; 2.08: 0
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 26.15%; 2.04: 14.41%; 2.05: 2.069%; 2.06: 21.91%; 2.07: 19.36%; 2.08: 0%
| PDM3 |  | 2.01: 12.3%; 2.02: 19.79%; 2.04: 6.065%; 2.05: -11.43%; 2.06: -6.413%; 2.07: -0.174%; 2.08: 12.3%
| Z-Score (Absolute) |  | 2.01: 1.35; 2.02: 0.67; 2.04: 0.67; 2.05: 1.22; 2.06: 0.67; 2.07: 0; 2.08: 0.67
| Analytical results for K in OREAS 922b (Certified Value 2.62 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 2.53 | 2.4857 | 2.32 | 2.54 | 2.6861 | 2.623 | 2.582 | 2.698 | 2.69 | 2.681 | 2.7 | 2.6 | 2.1801 | 2.72 | 2.4 | 2.7822 | 2.615 | 2.65 | 2.749 | 2.64 | 2.61 | 1
| 1 | 2 | 2.5 | 2.4649 | 2.29 | 2.63 | 2.7177 | 2.723 | 2.549 | 2.69 | 2.557 | 2.565 | 2.7 | 2.6 | 2.3298 | 2.86 | 2.6 | 2.8012 | 2.598 | 2.66 | 2.787 | 2.64 | 2.55 | 0
| 1 | 3 | 2.54 | 2.4683 | 2.34 | 2.61 | 2.7033 | 2.706 | 2.573 | 2.706 | 2.573 | 2.623 | 2.7 | 2.6 | 2.4249 | 2.72 | 2.6 | 2.7773 | 2.607 | 2.67 | 2.757 | 2.66 | 2.63 | 16
| 1 | 4 | 2.55 | 2.4613 | 2.31 | 2.56 | 2.6956 | 2.715 | 2.54 | 2.706 | 2.573 | 2.681 | 2.7 | 2.6 | 2.1639 | 2.74 | 2.6 | 2.7363 | 2.607 | 2.67 | 2.771 | 2.63 | 2.56 | 2.6236
| 1 | 5 | 2.55 | 2.4699 | 2.28 | 2.5 | 2.6952 | 2.723 | 2.482 | 2.739 | 2.54 | 2.64 | 2.7 | 2.6 | 2.497 | 2.73 | 2.6 | 2.6388 | 2.59 | 2.68 | 2.752 | 2.63 | 2.54 | 173
| 1 | 6 | 2.52 | 2.481 | 2.33 | 2.57 | 2.7342 | 2.715 | 2.474 | 2.673 | 2.598 | 2.648 | 2.8 | 2.6 | 2.4661 | 2.71 | 2.4 | 2.779 | 2.59 | 2.68 | 2.745 | 2.62 | 2.65
| Mean |  | 2.5317 | 2.4718 | 2.3117 | 2.5683 | 2.7054 | 2.7008 | 2.5333 | 2.702 | 2.5885 | 2.6397 | 2.7167 | 2.6 | 2.3437 | 2.7467 | 2.5333 | 2.7525 | 2.6012 | 2.6683 | 2.7602 | 2.6367 | 2.59
| Median |  | 2.535 | 2.4691 | 2.315 | 2.565 | 2.6995 | 2.715 | 2.5445 | 2.702 | 2.573 | 2.644 | 2.7 | 2.6 | 2.3774 | 2.725 | 2.6 | 2.7782 | 2.6025 | 2.67 | 2.7545 | 2.635 | 2.585
| Std Dev. |  | 0.019408 | 0.0095005 | 0.023166 | 0.047081 | 0.017633 | 0.038649 | 0.045579 | 0.021918 | 0.053324 | 0.043219 | 0.040825 | 0 | 0.14448 | 0.056451 | 0.10328 | 0.059601 | 0.010187 | 0.01169 | 0.01593 | 0.013663 | 0.046043
| Rel.Std.Dev. |  | 0.7666% | 0.3843% | 1.002% | 1.833% | 0.6518% | 1.431% | 1.799% | 0.8112% | 2.06% | 1.637% | 1.503% | 0% | 6.165% | 2.055% | 4.077% | 2.165% | 0.3916% | 0.4381% | 0.5771% | 0.5182% | 1.778%
| PDM3 |  | -3.505% | -5.785% | -11.89% | -2.107% | 3.116% | 2.943% | -3.441% | 2.988% | -1.338% | 0.6118% | 3.547% | -0.9001% | -10.67% | 4.69% | -3.441% | 4.911% | -0.8556% | 1.704% | 5.205% | 0.4975% | -1.281%
| Z-Score (Absolute) |  | 0.67 | 1.25 | 2.81 | 0.32 | 1.01 | 0.97 | 0.66 | 0.98 | 0.12 | 0.38 | 1.12 | 0.01 | 2.5 | 1.41 | 0.66 | 1.47 | 0 | 0.65 | 1.54 | 0.34 | 0.11
| Analytical results for La in OREAS 922b (Certified Value 43 ppm) | [22 empty] | Certified
La | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.20; 2.11: 2.21; 2.12: 2.22; 2.13: 2.24 | ppm
|  |  | 2.01: PF*OES; 2.02: PF*MS; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*OES; 2.11: PF*OES; 2.12: BF*MS; 2.13: PF*OES/MS | 1
2
| 1 | 1 | 2.01: 38.27; 2.02: 43.1; 2.04: 42.8; 2.05: 48.2818; 2.06: 43.6; 2.07: 42.2; 2.08: 44; 2.09: 42.8; 2.10: 38; 2.11: 34; 2.12: 42.5; 2.13: 43.13 | 1
| 1 | 2 | 2.01: 37.92; 2.02: 43.7; 2.04: 43.7; 2.05: 48.3909; 2.06: 44; 2.07: 41.6; 2.08: 44; 2.09: 43; 2.10: 40; 2.11: 34; 2.12: 42.9; 2.13: 45.4 | 12
| 1 | 3 | 2.01: 38.44; 2.02: 44; 2.04: 43.6; 2.05: 46.5893; 2.06: 42.9; 2.07: 41; 2.08: 45; 2.09: 43.7; 2.10: 41; 2.11: 34; 2.12: 42.2; 2.13: 44.48 | 16
| 1 | 4 | 2.01: 37.91; 2.02: 42.3; 2.04: 42.5; 2.05: 45.9599; 2.06: 42.7; 2.07: 42.6; 2.08: 44; 2.09: 42.9; 2.10: 41; 2.11: 33; 2.12: 42.1; 2.13: 43.46 | 43.025
| 1 | 5 | 2.01: 37.88; 2.02: 44.5; 2.04: 43.7; 2.05: 48.1933; 2.06: 44.2; 2.07: 39.7; 2.08: 45; 2.09: 43.2; 2.10: 40; 2.11: 35; 2.12: 44.5; 2.13: 43.84 | 174
| 1 | 6 | 2.01: 38.28; 2.02: 43.7; 2.04: 43.4; 2.05: 48.2751; 2.06: 47; 2.07: 43.1; 2.08: 44; 2.09: 44.2; 2.10: 41; 2.11: 33; 2.12: 42.8; 2.13: 44.56
| Mean |  | 2.01: 38.1167; 2.02: 43.55; 2.04: 43.2833; 2.05: 47.6151; 2.06: 44.0667; 2.07: 41.7; 2.08: 44.3333; 2.09: 43.3; 2.10: 40.1667; 2.11: 33.8333; 2.12: 42.8333; 2.13: 44.145
| Median |  | 2.01: 38.095; 2.02: 43.7; 2.04: 43.5; 2.05: 48.2342; 2.06: 43.8; 2.07: 41.9; 2.08: 44; 2.09: 43.1; 2.10: 40.5; 2.11: 34; 2.12: 42.65; 2.13: 44.16
| Std Dev. |  | 2.01: 0.24172; 2.02: 0.76354; 2.04: 0.51153; 2.05: 1.0591; 2.06: 1.5539; 2.07: 1.2264; 2.08: 0.5164; 2.09: 0.54406; 2.10: 1.169; 2.11: 0.75277; 2.12: 0.8756; 2.13: 0.8309
| Rel.Std.Dev. |  | 2.01: 0.6341%; 2.02: 1.753%; 2.04: 1.182%; 2.05: 2.224%; 2.06: 3.526%; 2.07: 2.941%; 2.08: 1.165%; 2.09: 1.256%; 2.10: 2.91%; 2.11: 2.225%; 2.12: 2.044%; 2.13: 1.882%
| PDM3 |  | 2.01: -11.41%; 2.02: 1.22%; 2.04: 0.6004%; 2.05: 10.67%; 2.06: 2.421%; 2.07: -3.08%; 2.08: 3.041%; 2.09: 0.6392%; 2.10: -6.643%; 2.11: -21.36%; 2.12: -0.4455%; 2.13: 2.603%
| Z-Score (Absolute) |  | 2.01: 3.68; 2.02: 0.18; 2.04: 0.01; 2.05: 3.08; 2.06: 0.55; 2.07: 1.13; 2.08: 0.74; 2.09: 0.01; 2.10: 2.22; 2.11: 6.73; 2.12: 0.33; 2.13: 0.61
| Analytical results for Li in OREAS 922b (Certified Value 25.1 ppm) | [22 empty] | Certified
Li | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.18; 2.09: 2.20; 2.10: 2.21; 2.11: 2.23; 2.12: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*OES; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES; 2.10: PF*OES; 2.11: PF*MS; 2.12: PF*OES/MS | 1
2
| 1 | 1 | 2.01: 24; 2.02: 42.951; 2.04: 25; 2.05: 23; 2.06: 62.3144; 2.07: 27; 2.08: 26; 2.09: 26; 2.10: 23; 2.11: 28; 2.12: 22 | 1
| 1 | 2 | 2.01: 26; 2.02: 41.7185; 2.04: 22; 2.05: 24; 2.06: 28.9681; 2.07: 26; 2.08: 23; 2.09: 26; 2.10: 25; 2.11: 26; 2.12: 24 | 0
| 1 | 3 | 2.01: 25; 2.02: 41.642; 2.04: 24; 2.05: 22; 2.06: 22.361; 2.07: 25; 2.08: 23; 2.09: 26; 2.10: 23; 2.11: 27; 2.12: 24 | 16
| 1 | 4 | 2.01: 25; 2.02: 43.2315; 2.04: 26; 2.05: 25; 2.06: 27.2966; 2.07: 29; 2.08: 21; 2.09: 26; 2.10: 24; 2.11: 30; 2.12: 24 | 25.0743
| 1 | 5 | 2.01: 26; 2.02: 42.1435; 2.04: 25; 2.05: 24; 2.06: 27.3415; 2.07: 26; 2.08: 23; 2.09: 26; 2.10: 24; 2.11: 43; 2.12: 24 | 175
| 1 | 6 | 2.01: 25; 2.02: 41.54; 2.04: 24; 2.05: 25; 2.06: 30.2484; 2.07: 34; 2.08: 21; 2.09: 26; 2.10: 23; 2.11: 26; 2.12: 24
| Mean |  | 2.01: 25.1667; 2.02: 42.2044; 2.04: 24.3333; 2.05: 23.8333; 2.06: 33.0883; 2.07: 27.8333; 2.08: 22.8333; 2.09: 26; 2.10: 23.6667; 2.11: 30; 2.12: 23.6667
| Median |  | 2.01: 25; 2.02: 41.931; 2.04: 24.5; 2.05: 24; 2.06: 28.1548; 2.07: 26.5; 2.08: 23; 2.09: 26; 2.10: 23.5; 2.11: 27.5; 2.12: 24
| Std Dev. |  | 2.01: 0.75277; 2.02: 0.72248; 2.04: 1.3663; 2.05: 1.169; 2.06: 14.5661; 2.07: 3.3116; 2.08: 1.8348; 2.09: 0; 2.10: 0.8165; 2.11: 6.5422; 2.12: 0.8165
| Rel.Std.Dev. |  | 2.01: 2.991%; 2.02: 1.712%; 2.04: 5.615%; 2.05: 4.905%; 2.06: 44.02%; 2.07: 11.9%; 2.08: 8.036%; 2.09: 0%; 2.10: 3.45%; 2.11: 21.81%; 2.12: 3.45%
| PDM3 |  | 2.01: 0.3683%; 2.02: 68.32%; 2.04: -2.955%; 2.05: -4.949%; 2.06: 31.96%; 2.07: 11%; 2.08: -8.937%; 2.09: 3.692%; 2.10: -5.614%; 2.11: 19.64%; 2.12: -5.614%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 7.66; 2.04: 0.37; 2.05: 0.6; 2.06: 3.56; 2.07: 1.2; 2.08: 1.05; 2.09: 0.37; 2.10: 0.67; 2.11: 2.17; 2.12: 0.67
| Analytical results for Lu in OREAS 922b (Certified Value 0.42 ppm) | [22 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.13; 2.06: 2.18; 2.07: 2.19; 2.08: 2.22; 2.09: 2.24 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: BF*OES/MS; 2.08: BF*MS; 2.09: PF*OES/MS | 2
2
| 1 | 1 | 2.01: < 0.5; 2.02: 0.4; 2.04: 0.41554; 2.05: 0.39; 2.06: < 2; 2.07: 0.37; 2.08: 0.37; 2.09: 0.48 | 1
| 1 | 2 | 2.01: < 0.5; 2.02: 0.4; 2.04: 0.40395; 2.05: 0.4; 2.06: < 2; 2.07: 0.39; 2.08: 0.41; 2.09: 0.51 | 13
| 1 | 3 | 2.01: < 0.5; 2.02: 0.4; 2.04: 0.43007; 2.05: 0.39; 2.06: < 2; 2.07: 0.4; 2.08: 0.38; 2.09: 0.49 | 16
| 1 | 4 | 2.01: < 0.5; 2.02: 0.4; 2.04: 0.41772; 2.05: 0.39; 2.06: < 2; 2.07: 0.39; 2.08: 0.44; 2.09: 0.49 | 0.41575
| 1 | 5 | 2.01: < 0.5; 2.02: 0.4; 2.04: 0.41229; 2.05: 0.38; 2.06: < 2; 2.07: 0.39; 2.08: 0.39; 2.09: 0.49 | 176
| 1 | 6 | 2.01: < 0.5; 2.02: 0.4; 2.04: 0.42291; 2.05: 0.4; 2.06: < 2; 2.07: 0.38; 2.08: 0.37; 2.09: 0.48
| Mean |  | 2.02: 0.4; 2.04: 0.41708; 2.05: 0.39167; 2.07: 0.38667; 2.08: 0.39333; 2.09: 0.49
| Median |  | 2.02: 0.4; 2.04: 0.41663; 2.05: 0.39; 2.07: 0.39; 2.08: 0.385; 2.09: 0.49
| Std Dev. |  | 2.02: 6.0809e-17; 2.04: 0.0089533; 2.05: 0.0075277; 2.07: 0.010328; 2.08: 0.027325; 2.09: 0.010954
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.04: 2.147%; 2.05: 1.922%; 2.07: 2.671%; 2.08: 6.947%; 2.09: 2.236%
| PDM3 |  | 2.02: -3.788%; 2.04: 0.3205%; 2.05: -5.793%; 2.07: -6.995%; 2.08: -5.392%; 2.09: 17.86%
| Z-Score (Absolute) |  | 2.01: 4.07; 2.02: NA; 2.04: 0.67; 2.05: 0.05; 2.06: 17.23; 2.07: 0.19; 2.08: 0; 2.09: 2.74
Results from laboratory 2.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 922b (Certified Value 0.418 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.41 | 0.41532 | 0.41 | 0.41 | 0.46079 | 0.392 | 0.422 | 0.41 | 0.41 | 0.41 | 0.42 | 0.42 | 0.41082 | 0.43 | 0.43 | 0.422 | 0.4 | 0.458 | 0.43 | 0.43 |  | 1
| 1 | 2 | 0.41 | 0.41475 | 0.4 | 0.43 | 0.46142 | 0.398 | 0.41 | 0.422 | 0.428 | 0.404 | 0.41 | 0.42 | 0.37601 | 0.42 | 0.45 | 0.428 | 0.41 | 0.457 | 0.44 | 0.44 |  | 0
| 1 | 3 | 0.41 | 0.41249 | 0.42 | 0.43 | 0.46324 | 0.404 | 0.422 | 0.428 | 0.44 | 0.41 | 0.41 | 0.42 | 0.38482 | 0.42 | 0.42 | 0.422 | 0.4 | 0.451 | 0.44 | 0.43 |  | 16
| 1 | 4 | 0.41 | 0.41617 | 0.41 | 0.42 | 0.46605 | 0.398 | 0.398 | 0.422 | 0.428 | 0.41 | 0.42 | 0.42 | 0.41754 | 0.42 | 0.44 | 0.434 | 0.41 | 0.461 | 0.43 | 0.42 |  | 0.41755
| 1 | 5 | 0.41 | 0.4179 | 0.41 | 0.41 | 0.45753 | 0.404 | 0.41 | 0.422 | 0.422 | 0.41 | 0.42 | 0.42 | 0.3955 | 0.42 | 0.44 | 0.428 | 0.41 | 0.463 | 0.43 | 0.42 |  | 177
| 1 | 6 | 0.41 | 0.41539 | 0.41 | 0.42 | 0.46068 | 0.404 | 0.41 | 0.41 | 0.44 | 0.416 | 0.42 | 0.41 | 0.39907 | 0.42 | 0.44 | 0.422 | 0.41 | 0.457 | 0.43 | 0.44
| Mean |  | 0.41 | 0.41534 | 0.41 | 0.42 | 0.46162 | 0.4 | 0.412 | 0.419 | 0.428 | 0.41 | 0.41667 | 0.41833 | 0.39729 | 0.42167 | 0.43667 | 0.426 | 0.40667 | 0.45783 | 0.43333 | 0.43
| Median |  | 0.41 | 0.41535 | 0.41 | 0.42 | 0.46111 | 0.401 | 0.41 | 0.422 | 0.428 | 0.41 | 0.42 | 0.42 | 0.39729 | 0.42 | 0.44 | 0.425 | 0.41 | 0.4575 | 0.43 | 0.43
| Std Dev. |  | 6.0809e-17 | 0.0017717 | 0.0063246 | 0.0089443 | 0.0028485 | 0.004899 | 0.0090333 | 0.0073485 | 0.011384 | 0.0037947 | 0.005164 | 0.0040825 | 0.015542 | 0.0040825 | 0.010328 | 0.004899 | 0.005164 | 0.0041191 | 0.005164 | 0.0089443
| Rel.Std.Dev. |  | 1.483e-14% | 0.4266% | 1.543% | 2.13% | 0.6171% | 1.225% | 2.193% | 1.754% | 2.66% | 0.9255% | 1.239% | 0.9759% | 3.912% | 0.9682% | 2.365% | 1.15% | 1.27% | 0.8997% | 1.192% | 2.08%
| PDM3 |  | -1.808% | -0.5299% | -1.808% | 0.5868% | 10.55% | -4.203% | -1.329% | 0.3473% | 2.503% | -1.808% | -0.2115% | 0.1876% | -4.851% | 0.9859% | 4.578% | 2.024% | -2.606% | 9.648% | 3.78% | 2.982%
| Z-Score (Absolute) |  | 0.67 | 0.26 | 0.67 | 0.1 | 3.31 | 1.44 | 0.51 | 0.03 | 0.73 | 0.67 | 0.16 | 0.03 | 1.66 | 0.23 | 1.39 | 0.57 | 0.93 | 3.02 | 1.13 | 0.87
| Analytical results for Mn in OREAS 922b (Certified Value 0.01 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 3
3
| 1 | 1 | 0.011 | 0.010951 | 0.01052 | 0.011 | 0.008 | 0.008 | 0.008 | 0.015 | 0.015 | 0.0113 | 0.0113 | 0.011247 | 0.0109 | 0.01 | 0.008 | 0.0141 | 0.0108 | 0.0104 | 0.0102 |  |  | 1
| 1 | 2 | 0.011 | 0.01069 | 0.01052 | 0.012 | 0.015 | 0.008 | 0.008 | 0.008 | 0.015 | 0.0107 | 0.0113 | 0.011507 | 0.0108 | 0.01 | 0.008 | 0.014 | 0.0107 | 0.011 | 0.0106 |  |  | 30
| 1 | 3 | 0.011 | 0.010858 | 0.01057 | 0.012 | 0.008 | 0.008 | 0.008 | 0.008 | 0.015 | 0.0118 | 0.0111 | 0.011261 | 0.011 | 0.01 | 0.008 | 0.014 | 0.0108 | 0.0112 | 0.0107 |  |  | 16
| 1 | 4 | 0.011 | 0.01133 | 0.01033 | 0.011 | 0.008 | 0.008 | 0.008 | 0.008 | 0.015 | 0.012 | 0.0113 | 0.011224 | 0.0109 | 0.01 | 0.008 | 0.0141 | 0.0109 | 0.011 | 0.0116 |  |  | 0.010488
| 1 | 5 | 0.011 | 0.010959 | 0.01072 | 0.011 | 0.008 | 0.008 | 0.008 | 0.008 | 0.015 | 0.0119 | 0.0113 | 0.011273 | 0.0109 | 0.01 | 0.008 | 0.014 | 0.0108 | 0.0112 | 0.01 |  |  | 178
| 1 | 6 | 0.011 | 0.011145 | 0.0106 | 0.011 | 0.015 | 0.008 | 0.008 | 0.008 | 0.015 | 0.0125 | 0.0112 | 0.011476 | 0.0109 | 0.01 | 0.008 | 0.014 | 0.0108 | 0.0107 | 0.0104
| Mean |  | 0.011 | 0.010989 | 0.010543 | 0.011333 | 0.010333 | 0.008 | 0.008 | 0.0091667 | 0.015 | 0.0117 | 0.01125 | 0.011331 | 0.0109 | 0.01 | 0.008 | 0.014033 | 0.0108 | 0.010917 | 0.010583
| Median |  | 0.011 | 0.010955 | 0.010545 | 0.011 | 0.008 | 0.008 | 0.008 | 0.008 | 0.015 | 0.01185 | 0.0113 | 0.011267 | 0.0109 | 0.01 | 0.008 | 0.014 | 0.0108 | 0.011 | 0.0105
| Std Dev. |  | 0 | 0.00022319 | 0.00012785 | 0.0005164 | 0.0036148 | 0 | 0 | 0.0028577 | 0 | 0.0006229 | 8.3666e-05 | 0.00012557 | 6.3246e-05 | 0 | 0 | 5.164e-05 | 6.3246e-05 | 0.00031252 | 0.00056006
| Rel.Std.Dev. |  | 0% | 2.031% | 1.213% | 4.556% | 34.98% | 0% | 0% | 31.18% | 0% | 5.324% | 0.7437% | 1.108% | 0.5802% | 0% | 0% | 0.368% | 0.5856% | 2.863% | 5.292%
| PDM3 |  | 4.879% | 4.772% | 0.5249% | 8.057% | -1.477% | -23.72% | -23.72% | -12.6% | 43.02% | 11.55% | 7.263% | 8.038% | 3.926% | -4.655% | -23.72% | 33.8% | 2.972% | 4.084% | 0.9063%
| Z-Score (Absolute) |  | 0.16 | 0.14 | 0.56 | 0.67 | 0.89 | 4.91 | 4.91 | 2.9 | 7.14 | 1.24 | 0.54 | 0.67 | 0 | 1.4 | 4.91 | 4.88 | 0.16 | 0.03 | 0.49
| Analytical results for Mo in OREAS 922b (Certified Value 5.19 ppm) | [22 empty] | Certified
Mo | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.16; 2.10: 2.21; 2.11: 2.23; 2.12: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*OES; 2.11: PF*MS; 2.12: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 5; 2.02: 4.1; 2.04: 6; 2.05: 7; 2.07: 5; 2.08: 4.5152; 2.09: 6; 2.10: < 10; 2.11: 6; 2.12: < 10 |  |  | 1
| 1 | 2 | 2.01: 5; 2.02: 4.1; 2.04: 5; 2.05: 7; 2.07: 6; 2.08: 4.321; 2.09: 6; 2.10: < 10; 2.11: 5; 2.12: < 10 |  |  | 31
| 1 | 3 | 2.01: 5; 2.02: 4.2; 2.04: 6; 2.05: 7; 2.07: 5; 2.08: 4.6862; 2.09: 5; 2.10: < 10; 2.11: 6; 2.12: < 10 |  |  | 16
| 1 | 4 | 2.01: 5; 2.02: 4.6; 2.04: 6; 2.05: 7; 2.07: 5; 2.08: 4.1775; 2.09: 5; 2.10: < 10; 2.11: 5; 2.12: < 10 |  |  | 5.1865
| 1 | 5 | 2.01: 5; 2.02: 4.2; 2.04: 5; 2.05: 8; 2.07: 5; 2.08: 4.3967; 2.09: 5; 2.10: < 10; 2.11: 5; 2.12: < 10 |  |  | 179
| 1 | 6 | 2.01: 5; 2.02: 4.3; 2.04: 5; 2.05: 8; 2.07: 5; 2.08: 4.6187; 2.09: 6; 2.10: < 10; 2.11: 6; 2.12: < 10
| Mean |  | 2.01: 5; 2.02: 4.25; 2.04: 5.5; 2.05: 7.3333; 2.07: 5.1667; 2.08: 4.4526; 2.09: 5.5; 2.11: 5.5
| Median |  | 2.01: 5; 2.02: 4.2; 2.04: 5.5; 2.05: 7; 2.07: 5; 2.08: 4.456; 2.09: 5.5; 2.11: 5.5
| Std Dev. |  | 2.01: 0; 2.02: 0.18708; 2.04: 0.54772; 2.05: 0.5164; 2.07: 0.40825; 2.08: 0.19091; 2.09: 0.54772; 2.11: 0.54772
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 4.402%; 2.04: 9.959%; 2.05: 7.042%; 2.07: 7.902%; 2.08: 4.288%; 2.09: 9.959%; 2.11: 9.959%
| PDM3 |  | 2.01: -3.597%; 2.02: -18.06%; 2.04: 6.044%; 2.05: 41.39%; 2.07: -0.3831%; 2.08: -14.15%; 2.09: 6.044%; 2.11: 6.044%
| Z-Score (Absolute) |  | 2.01: 0.13; 2.02: 1.35; 2.04: 0.67; 2.05: 3.64; 2.07: 0.13; 2.08: 1.02; 2.09: 0.67; 2.10: 0.13; 2.11: 0.67; 2.12: 0.13
| Analytical results for Na in OREAS 922b (Indicative Value 0.084 wt.%) | [22 empty] | Indicative
Na | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.19; 2.05: 2.20 |  |  | wt.%
|  |  | 2.01: PF*OES; 2.02: BF*OES; 2.04: BF*OES/MS; 2.05: BF*OES |  |  | 3
3
| 1 | 1 | 2.01: 0.11; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05 |  |  | 1
| 1 | 2 | 2.01: 0.1; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05 |  |  | 21
| 1 | 3 | 2.01: 0.11; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05 |  |  | 16
| 1 | 4 | 2.01: 0.1; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05 |  |  | 0.083505
| 1 | 5 | 2.01: 0.11; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05 |  |  | 27
| 1 | 6 | 2.01: 0.1; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05
| Mean |  | 2.01: 0.105; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05
| Median |  | 2.01: 0.105; 2.02: 0.09; 2.04: 0.089; 2.05: 0.05
| Std Dev. |  | 2.01: 0.0054772; 2.02: 1.5202e-17; 2.04: 1.5202e-17; 2.05: 7.6012e-18
| Rel.Std.Dev. |  | 2.01: 5.216%; 2.02: 1.689e-14%; 2.04: 1.708e-14%; 2.05: 1.52e-14%
| PDM3 |  | 2.01: 25.74%; 2.02: 7.778%; 2.04: 6.58%; 2.05: -40.12%
| Z-Score (Absolute) |  | 2.01: 1.31; 2.02: 0.04; 2.04: 0.04; 2.05: 3.33
| Analytical results for Nb in OREAS 922b (Certified Value 11 ppm) | [22 empty] | Certified
Nb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.19; 2.10: 2.20; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: PF*OES; 2.11: PF*OES/MS |  |  | 1
2
| 1 | 1 | 2.01: 11.2; 2.02: 11.9; 2.04: 9.8945; 2.05: 11.7; 2.07: 11; 2.08: 10.1578; 2.09: 10.7; 2.10: 34; 2.11: 13 |  |  | 1
| 1 | 2 | 2.01: 9.6; 2.02: 11.7; 2.04: 9.4897; 2.05: 12.9; 2.07: 11; 2.08: 10.2297; 2.09: 11.1; 2.10: 33; 2.11: 13 |  |  | 15
| 1 | 3 | 2.01: 11; 2.02: 12; 2.04: 9.6128; 2.05: 12.5; 2.07: 11; 2.08: 10.0245; 2.09: 11.4; 2.10: 33; 2.11: 14 |  |  | 16
| 1 | 4 | 2.01: 10.5; 2.02: 11.2; 2.04: 9.9811; 2.05: 12.3; 2.07: 12; 2.08: 10.2309; 2.09: 11; 2.10: 34; 2.11: 13 |  |  | 10.9675
| 1 | 5 | 2.01: 11.2; 2.02: 11.6; 2.04: 9.2692; 2.05: 12.1; 2.07: 11; 2.08: 9.9882; 2.09: 11; 2.10: 33; 2.11: 13 |  |  | 180
| 1 | 6 | 2.01: 11.2; 2.02: 11.6; 2.04: 9.5907; 2.05: 12.4; 2.07: 12; 2.08: 10.2402; 2.09: 10.9; 2.10: 33; 2.11: 12
| Mean |  | 2.01: 10.7833; 2.02: 11.6667; 2.04: 9.6397; 2.05: 12.3167; 2.07: 11.3333; 2.08: 10.1452; 2.09: 11.0167; 2.10: 33.3333; 2.11: 13
| Median |  | 2.01: 11.1; 2.02: 11.65; 2.04: 9.6017; 2.05: 12.35; 2.07: 11; 2.08: 10.1938; 2.09: 11; 2.10: 33; 2.11: 13
| Std Dev. |  | 2.01: 0.64005; 2.02: 0.28048; 2.04: 0.26242; 2.05: 0.40208; 2.07: 0.5164; 2.08: 0.11214; 2.09: 0.23166; 2.10: 0.5164; 2.11: 0.63246
| Rel.Std.Dev. |  | 2.01: 5.936%; 2.02: 2.404%; 2.04: 2.722%; 2.05: 3.265%; 2.07: 4.556%; 2.08: 1.105%; 2.09: 2.103%; 2.10: 1.549%; 2.11: 4.865%
| PDM3 |  | 2.01: -1.679%; 2.02: 6.375%; 2.04: -12.11%; 2.05: 12.3%; 2.07: 3.336%; 2.08: -7.497%; 2.09: 0.4485%; 2.10: 203.9%; 2.11: 18.53%
| Z-Score (Absolute) |  | 2.01: 0.1; 2.02: 0.68; 2.04: 1.12; 2.05: 1.26; 2.07: NA; 2.08: 0.67; 2.09: 0.1; 2.10: NA; 2.11: NA
Results from laboratories 2.13, 2.20 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 922b (Certified Value 36 ppm) | [22 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 1
2
| 1 | 1 | 2.01: 36.83; 2.02: 35.4; 2.04: 38.5788; 2.05: 43.4; 2.07: 35.3; 2.08: 36; 2.09: 34.1; 2.10: 34.5; 2.11: 36.1 |  |  | 1
| 1 | 2 | 2.01: 37.11; 2.02: 36.6; 2.04: 37.0608; 2.05: 40.6; 2.07: 34.9; 2.08: 36; 2.09: 34; 2.10: 34.8; 2.11: 37.6 |  |  | 4
| 1 | 3 | 2.01: 36.96; 2.02: 36.2; 2.04: 37.9227; 2.05: 40.8; 2.07: 35.2; 2.08: 37; 2.09: 34.7; 2.10: 34.4; 2.11: 36.4 |  |  | 16
| 1 | 4 | 2.01: 35.74; 2.02: 35.9; 2.04: 37.1527; 2.05: 40.3; 2.07: 36; 2.08: 35; 2.09: 34.5; 2.10: 34.3; 2.11: 36.4 |  |  | 36.0147
| 1 | 5 | 2.01: 37.75; 2.02: 35.5; 2.04: 38.9379; 2.05: 43.5; 2.07: 34.9; 2.08: 36; 2.09: 34.9; 2.10: 35.8; 2.11: 35.9 |  |  | 181
| 1 | 6 | 2.01: 36.92; 2.02: 36.1; 2.04: 38.9451; 2.05: 47.1; 2.07: 35.9; 2.08: 37; 2.09: 34.9; 2.10: 34.3; 2.11: 36.3
| Mean |  | 2.01: 36.885; 2.02: 35.95; 2.04: 38.0997; 2.05: 42.6167; 2.07: 35.3667; 2.08: 36.1667; 2.09: 34.5167; 2.10: 34.6833; 2.11: 36.45
| Median |  | 2.01: 36.94; 2.02: 36; 2.04: 38.2507; 2.05: 42.1; 2.07: 35.25; 2.08: 36; 2.09: 34.6; 2.10: 34.45; 2.11: 36.35
| Std Dev. |  | 2.01: 0.65111; 2.02: 0.45056; 2.04: 0.8549; 2.05: 2.6164; 2.07: 0.48028; 2.08: 0.75277; 2.09: 0.392; 2.10: 0.57764; 2.11: 0.59582
| Rel.Std.Dev. |  | 2.01: 1.765%; 2.02: 1.253%; 2.04: 2.244%; 2.05: 6.139%; 2.07: 1.358%; 2.08: 2.081%; 2.09: 1.136%; 2.10: 1.665%; 2.11: 1.635%
| PDM3 |  | 2.01: 2.416%; 2.02: -0.1798%; 2.04: 5.789%; 2.05: 18.33%; 2.07: -1.799%; 2.08: 0.4218%; 2.09: -4.16%; 2.10: -3.697%; 2.11: 1.209%
| Z-Score (Absolute) |  | 2.01: 0.61; 2.02: 0.18; 2.04: 1.63; 2.05: 5.44; 2.07: 0.67; 2.08: 0; 2.09: 1.39; 2.10: 1.25; 2.11: 0.24
| Analytical results for Ni in OREAS 922b (Indicative Value 20.9 ppm) | [22 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 1
1
| 1 | 1 | < 20 | 14.5 | < 20 | 20 | 40 | 30 | 20 | 30 | 22 | < 20 |  | 80 | < 20 | 16 | 12 | 55 | [5 empty] | 1
| 1 | 2 | < 20 | 11.93 | < 20 | 30 | 70 | < 20 | 40 | 30 | 22 | < 20 |  | < 50 | < 20 | 16 | 12 | < 50 | [5 empty] | 22
| 1 | 3 | < 20 | 13.52 | < 20 | 30 | 90 | 20 | 20 | 30 | 17 | < 20 | 24 | < 50 | < 20 | 16 | 12 | 77 | [5 empty] | 16
| 1 | 4 | < 20 | 11.2 | < 20 | 30 | 80 | < 20 | < 20 | 40 | 19 | < 20 |  | < 50 | < 20 | 16 | 11 | < 50 | [5 empty] | 20.9137
| 1 | 5 | < 20 | 11.8 | < 20 | 30 | 30 | < 20 | 20 | 20 | 20 | < 20 |  | < 50 | < 20 | 16 | 12 | 51 | [5 empty] | 28
| 1 | 6 | < 20 | 11.39 | < 20 | 20 | 40 | < 20 | 20 | 20 | 32 | < 20 |  | < 50 | < 20 | 16 | 12 | < 50
| Mean |  |  | 12.39 |  | 26.6667 | 58.3333 | 25 | 24 | 28.3333 | 22 |  | 24 | 80 |  | 16 | 11.8333 | 61
| Median |  |  | 11.865 |  | 30 | 55 | 25 | 20 | 30 | 21 |  | 24 | 80 |  | 16 | 12 | 55
| Std Dev. |  | 2.02: 1.3195; 2.05: 5.164; 2.07: 24.8328; 2.08: 7.0711; 2.09: 8.9443; 2.10: 7.5277; 2.11: 5.2536; 2.17: 0; 2.19: 0.40825; 2.20: 14
| Rel.Std.Dev. |  | 2.02: 10.65%; 2.05: 19.36%; 2.07: 42.57%; 2.08: 28.28%; 2.09: 37.27%; 2.10: 26.57%; 2.11: 23.88%; 2.17: 0%; 2.19: 3.45%; 2.20: 22.95%
| PDM3 |  |  | -40.76% |  | 27.51% | 178.9% | 19.54% | 14.76% | 35.48% | 5.194% |  | 14.76% | 282.5% |  | -23.5% | -43.42% | 191.7%
| Z-Score (Absolute) |  | 0.71 | 0.52 | 0.71 | 0.63 | 3.2 | 0.31 | 0.23 | 0.77 | 0.26 | 0.71 | 0.42 | 1.24 | 0.71 | 0.23 | 0.57 | 1.96
| Analytical results for P in OREAS 922b (Certified Value 0.032 wt.%) | [22 empty] | Certified
P | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.07: 2.13; 2.08: 2.19; 2.09: 2.20; 2.10: 2.21 |  |  | wt.%
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*OES; 2.05: PF*OES/MS; 2.07: PF*OES; 2.08: BF*OES/MS; 2.09: PF*OES; 2.10: PF*OES |  |  | 3
3
| 1 | 1 | 2.01: 0.04; 2.02: 0.03; 2.04: 0.03; 2.05: 0.034584; 2.07: 0.03; 2.08: 0.031; 2.09: 0.0265; 2.10: 0.033 |  |  | 1
| 1 | 2 | 2.01: 0.03; 2.02: 0.03; 2.04: 0.03; 2.05: 0.034326; 2.07: 0.03; 2.08: 0.031; 2.09: 0.0265; 2.10: 0.036 |  |  | 0
| 1 | 3 | 2.01: 0.03; 2.02: 0.04; 2.04: 0.03; 2.05: 0.03666; 2.07: 0.03; 2.08: 0.031; 2.09: 0.0269; 2.10: 0.033 |  |  | 16
| 1 | 4 | 2.01: 0.03; 2.02: 0.03; 2.04: 0.03; 2.05: 0.034155; 2.07: 0.03; 2.08: 0.035; 2.09: 0.0268; 2.10: 0.027 |  |  | 0.031513
| 1 | 5 | 2.01: 0.04; 2.02: 0.03; 2.04: 0.03; 2.05: 0.036298; 2.07: 0.03; 2.08: 0.035; 2.09: 0.0269; 2.10: 0.028 |  |  | 182
| 1 | 6 | 2.01: 0.03; 2.02: 0.04; 2.04: 0.03; 2.05: 0.035485; 2.07: 0.03; 2.08: 0.031; 2.09: 0.0265; 2.10: 0.032
| Mean |  | 2.01: 0.033333; 2.02: 0.033333; 2.04: 0.03; 2.05: 0.035251; 2.07: 0.03; 2.08: 0.032333; 2.09: 0.026683; 2.10: 0.0315
| Median |  | 2.01: 0.03; 2.02: 0.03; 2.04: 0.03; 2.05: 0.035035; 2.07: 0.03; 2.08: 0.031; 2.09: 0.02665; 2.10: 0.0325
| Std Dev. |  | 2.01: 0.005164; 2.02: 0.005164; 2.04: 0; 2.05: 0.001062; 2.07: 0; 2.08: 0.0020656; 2.09: 0.00020412; 2.10: 0.0033912
| Rel.Std.Dev. |  | 2.01: 15.49%; 2.02: 15.49%; 2.04: 0%; 2.05: 3.013%; 2.07: 0%; 2.08: 6.388%; 2.09: 0.765%; 2.10: 10.77%
| PDM3 |  | 2.01: 5.777%; 2.02: 5.777%; 2.04: -4.801%; 2.05: 11.86%; 2.07: -4.801%; 2.08: 2.604%; 2.09: -15.33%; 2.10: -0.04061%
| Z-Score (Absolute) |  | 2.01: 0.64; 2.02: 0.64; 2.04: 0.71; 2.05: 1.42; 2.07: 0.71; 2.08: 0.1; 2.09: 2.05; 2.10: 0.1
| Analytical results for Pb in OREAS 922b (Certified Value 46.4 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.21 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 40 | 55.9 | < 100 | 100 | < 100 | < 100 | 100 | 45.7 | 44 | 37.5122 | 45 | < 100 | 51 | 61 | < 100 | [6 empty] | 1
| 1 | 2 | 40 | 55.4 | < 100 | < 100 | < 100 | < 100 | < 100 | 44.6 | 44 | 33.9467 | 43 | < 100 | 51 | 42 | < 100 | [6 empty] | 17
| 1 | 3 | 40 | 55.1 | < 100 | 100 | < 100 | < 100 | < 100 | 43.8 | 45 | 41.0676 | 44 | < 100 | 49 | 44 | < 100 | [6 empty] | 16
| 1 | 4 | 40 | 55.1 | < 100 | 100 | < 100 | < 100 | < 100 | 45.9 | 44 | 36.0159 | 39 | < 100 | 50 | 43 | < 100 | [6 empty] | 46.3563
| 1 | 5 | 40 | 56.9 | < 100 | < 100 | < 100 | < 100 | < 100 | 45.7 | 45 | 41.5701 | 41 | < 100 | 50 | 58 | < 100 | [6 empty] | 183
| 1 | 6 | 40 | 52.8 | < 100 | 100 | < 100 | < 100 | < 100 | 49.2 | 44 | 36.8101 | 42 | < 100 | 50 | 49 | < 100
| Mean |  | 2.01: 40; 2.02: 55.2; 2.05: 100; 2.09: 100; 2.10: 45.8167; 2.11: 44.3333; 2.12: 37.8204; 2.13: 42.3333; 2.16: 50.1667; 2.17: 49.5
| Median |  | 2.01: 40; 2.02: 55.25; 2.05: 100; 2.09: 100; 2.10: 45.7; 2.11: 44; 2.12: 37.1612; 2.13: 42.5; 2.16: 50; 2.17: 46.5
| Std Dev. |  | 2.01: 0; 2.02: 1.3565; 2.05: 0; 2.10: 1.8454; 2.11: 0.5164; 2.12: 2.9659; 2.13: 2.1602; 2.16: 0.75277; 2.17: 8.167
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 2.457%; 2.05: 0%; 2.10: 4.028%; 2.11: 1.165%; 2.12: 7.842%; 2.13: 5.103%; 2.16: 1.501%; 2.17: 16.5%
| PDM3 |  | 2.01: -13.71%; 2.02: 19.08%; 2.05: 115.7%; 2.09: 115.7%; 2.10: -1.164%; 2.11: -4.364%; 2.12: -18.41%; 2.13: -8.678%; 2.16: 8.22%; 2.17: 6.782%
| Z-Score (Absolute) |  | NA | 1.5 | 0 | 9.6 | 0 | 0 | 2.4 | 1.2 | 1.63 | 3.51 | 2.21 | 0 | 0.05 | 0.14 | 0
Results from laboratory 2.01 were removed due to their 10 ppm reading resolution.
| Analytical results for Pr in OREAS 922b (Certified Value 9.89 ppm) | [22 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: 9.9; 2.02: 10; 2.04: 10.0014; 2.05: 11.1; 2.07: 10.49; 2.08: 10; 2.09: 9.39; 2.10: 9.72; 2.11: 9.74 |  |  | 1
| 1 | 2 | 2.01: 10; 2.02: 10.2; 2.04: 9.9607; 2.05: 10.5; 2.07: 10.29; 2.08: 10; 2.09: 9.5; 2.10: 9.69; 2.11: 10.24 |  |  | 6
| 1 | 3 | 2.01: 10; 2.02: 10.2; 2.04: 9.8445; 2.05: 10.5; 2.07: 10.11; 2.08: 10; 2.09: 9.62; 2.10: 9.5; 2.11: 9.84 |  |  | 16
| 1 | 4 | 2.01: 9.8; 2.02: 9.9; 2.04: 9.7622; 2.05: 10.6; 2.07: 10.7; 2.08: 10; 2.09: 9.39; 2.10: 9.4; 2.11: 9.83 |  |  | 9.8868
| 1 | 5 | 2.01: 10.1; 2.02: 9.9; 2.04: 9.7767; 2.05: 11; 2.07: 10.01; 2.08: 10; 2.09: 9.5; 2.10: 9.94; 2.11: 9.81 |  |  | 184
| 1 | 6 | 2.01: 9.9; 2.02: 9.9; 2.04: 9.8925; 2.05: 11.2; 2.07: 10.51; 2.08: 10; 2.09: 9.71; 2.10: 9.66; 2.11: 10.01
| Mean |  | 2.01: 9.95; 2.02: 10.0167; 2.04: 9.873; 2.05: 10.8167; 2.07: 10.3517; 2.08: 10; 2.09: 9.5183; 2.10: 9.6517; 2.11: 9.9117
| Median |  | 2.01: 9.95; 2.02: 9.95; 2.04: 9.8685; 2.05: 10.8; 2.07: 10.39; 2.08: 10; 2.09: 9.5; 2.10: 9.675; 2.11: 9.835
| Std Dev. |  | 2.01: 0.10488; 2.02: 0.1472; 2.04: 0.096845; 2.05: 0.31885; 2.07: 0.26248; 2.08: 0; 2.09: 0.12703; 2.10: 0.18745; 2.11: 0.18389
| Rel.Std.Dev. |  | 2.01: 1.054%; 2.02: 1.47%; 2.04: 0.9809%; 2.05: 2.948%; 2.07: 2.536%; 2.08: 0%; 2.09: 1.335%; 2.10: 1.942%; 2.11: 1.855%
| PDM3 |  | 2.01: 0.6396%; 2.02: 1.314%; 2.04: -0.1392%; 2.05: 9.406%; 2.07: 4.702%; 2.08: 1.145%; 2.09: -3.726%; 2.10: -2.378%; 2.11: 0.2519%
| Z-Score (Absolute) |  | 2.01: 0.07; 2.02: 0.32; 2.04: 0.21; 2.05: 3.27; 2.07: 1.55; 2.08: NA; 2.09: 1.52; 2.10: 1.03; 2.11: 0.07
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 922b (Certified Value 154 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.22 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 157 | 163.8 | 156.402 | 154 | 157 | 147.4996 | 156.5 | 105 | 151.9 | 151 | 154.4 | 155 | [9 empty] | 1
| 1 | 2 | 155 | 161.4 | 155.8071 | 156 | 155 | 147.092 | 160.7 | 165 | 154.1 | 152 | 153.8 | 161 | [9 empty] | 37
| 1 | 3 | 152 | 162.3 | 156.108 | 154 | 157 | 151.5932 | 157.6 | 105 | 155.8 | 150 | 152.7 | 156 | [9 empty] | 16
| 1 | 4 | 153 | 155.4 | 156.8201 | 152 | 155 | 156.1523 | 159.9 | 181 | 152.2 | 154 | 152.4 | 155 | [9 empty] | 154.4527
| 1 | 5 | 151 | 161.1 | 153.5191 | 154 | 150 | 154.7986 | 157.9 | 140 | 156.1 | 155 | 152.8 | 154 | [9 empty] | 185
| 1 | 6 | 151 | 162.1 | 157.8827 | 153 | 157 | 155.8887 | 161 | 116 | 153.1 | 152 | 150.7 | 156
| Mean |  | 153.1667 | 161.0167 | 156.0898 | 153.8333 | 155.1667 | 152.1707 | 158.9333 | 135.3333 | 153.8667 | 152.3333 | 152.8 | 156.1667
| Median |  | 152.5 | 161.75 | 156.255 | 154 | 156 | 153.1959 | 158.9 | 128 | 153.6 | 152 | 152.75 | 155.5
| Std Dev. |  | 2.4014 | 2.9075 | 1.4515 | 1.3292 | 2.7142 | 4.1117 | 1.849 | 32.2532 | 1.7896 | 1.8619 | 1.2759 | 2.4833
| Rel.Std.Dev. |  | 1.568% | 1.806% | 0.9299% | 0.864% | 1.749% | 2.702% | 1.163% | 23.83% | 1.163% | 1.222% | 0.835% | 1.59%
| PDM3 |  | -0.8327% | 4.25% | 1.06% | -0.401% | 0.4622% | -1.477% | 2.901% | -12.38% | -0.3794% | -1.372% | -1.07% | 1.11%
| Z-Score (Absolute) |  | 0.29 | 3.02 | 0.95 | 0.01 | 0.56 | 0.71 | 2.15 | 7.81 | 0.01 | 0.64 | 0.44 | 0.98
| Analytical results for Re in OREAS 922b (Indicative Value < 0.1 ppm) | [22 empty] | Indicative
Re | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.14; 2.04: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 1
| 1 | 2 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 23
| 1 | 3 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 16
| 1 | 4 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | < 0.1
| 1 | 5 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 29
| 1 | 6 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.04
| Analytical results for S in OREAS 922b (Certified Value 0.267 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.16 | 2.20 | 2.23 | 2.24 | [8 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 3
3
| 1 | 1 | 0.27 | 0.14 | 0.27 | 0.25 | 0.27 | 0.29 | 0.28 | 0.24 | 0.29812 | 0.27 | 0.26 | 0.26 | 0.25 | [8 empty] | 1
| 1 | 2 | 0.26 | 0.13 | 0.28 | 0.24 | 0.29 | 0.28 | 0.28 | 0.25 | 0.27374 | 0.26 | 0.26 | 0.27 | 0.25 | [8 empty] | 0
| 1 | 3 | 0.27 | 0.14 | 0.28 | 0.25 | 0.28 | 0.29 | 0.27 | 0.23 | 0.29894 | 0.26 | 0.26 | 0.27 | 0.25 | [8 empty] | 16
| 1 | 4 | 0.26 | 0.19 | 0.28 | 0.25 | 0.29 | 0.28 | 0.27 | 0.24 | 0.27152 | 0.27 | 0.26 | 0.26 | 0.22 | [8 empty] | 0.26664
| 1 | 5 | 0.26 | 0.15 | 0.28 | 0.24 | 0.28 | 0.27 | 0.28 | 0.24 | 0.29234 | 0.26 | 0.26 | 0.27 | 0.24 | [8 empty] | 186
| 1 | 6 | 0.26 | 0.15 | 0.28 | 0.25 | 0.29 | 0.28 | 0.28 | 0.26 | 0.29336 | 0.26 | 0.26 | 0.26 | 0.26
| Mean |  | 0.26333 | 0.15 | 0.27833 | 0.24667 | 0.28333 | 0.28167 | 0.27667 | 0.24333 | 0.288 | 0.26333 | 0.26 | 0.265 | 0.245
| Median |  | 0.26 | 0.145 | 0.28 | 0.25 | 0.285 | 0.28 | 0.28 | 0.24 | 0.29285 | 0.26 | 0.26 | 0.265 | 0.25
| Std Dev. |  | 0.005164 | 0.020976 | 0.0040825 | 0.005164 | 0.008165 | 0.0075277 | 0.005164 | 0.010328 | 0.012203 | 0.005164 | 0 | 0.0054772 | 0.013784
| Rel.Std.Dev. |  | 1.961% | 13.98% | 1.467% | 2.094% | 2.882% | 2.673% | 1.866% | 4.244% | 4.237% | 1.961% | 0% | 2.067% | 5.626%
| PDM3 |  | -1.24% | -43.74% | 4.386% | -7.49% | 6.261% | 5.636% | 3.761% | -8.741% | 8.013% | -1.24% | -2.49% | -0.6148% | -8.116%
| Z-Score (Absolute) |  | 0 | 4.59 | 0.61 | 0.67 | 0.81 | 0.74 | 0.54 | 0.81 | 1 | 0 | 0.13 | 0.07 | 0.74
| Analytical results for Sb in OREAS 922b (Certified Value 4.03 ppm) | [22 empty] | Certified
Sb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.12; 2.05: 2.13; 2.07: 2.14; 2.08: 2.16; 2.09: 2.21; 2.10: 2.23; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES; 2.10: PF*MS; 2.11: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 5; 2.02: 4.3; 2.04: 4; 2.05: 3.7; 2.07: 3.8519; 2.08: 3.8; 2.09: < 50; 2.10: 3.3; 2.11: 5 |  |  | 1
| 1 | 2 | 2.01: < 5; 2.02: 4.2; 2.04: 5; 2.05: 3.7; 2.07: 3.357; 2.08: 6.8; 2.09: < 50; 2.10: 3.5; 2.11: 5 |  |  | 39
| 1 | 3 | 2.01: < 5; 2.02: 4.1; 2.04: 4; 2.05: 3.6; 2.07: 3.3805; 2.08: 3.9; 2.09: < 50; 2.10: 3.6; 2.11: 4 |  |  | 16
| 1 | 4 | 2.01: < 5; 2.02: 4.1; 2.04: 4; 2.05: 3.7; 2.07: 3.6175; 2.08: 47.8; 2.09: < 50; 2.10: 3.4; 2.11: 5 |  |  | 4.0329
| 1 | 5 | 2.01: < 5; 2.02: 4.1; 2.04: 4; 2.05: 3.6; 2.07: 3.7257; 2.08: 27.7; 2.09: < 50; 2.10: 3.8; 2.11: 5 |  |  | 187
| 1 | 6 | 2.01: < 5; 2.02: 4.6; 2.04: 5; 2.05: 3.9; 2.07: 3.8912; 2.08: 3.8; 2.09: < 50; 2.10: 3.6; 2.11: 5
| Mean |  | 2.02: 4.2333; 2.04: 4.3333; 2.05: 3.7; 2.07: 3.6373; 2.08: 15.6333; 2.10: 3.5333; 2.11: 4.8333
| Median |  | 2.02: 4.15; 2.04: 4; 2.05: 3.7; 2.07: 3.6716; 2.08: 5.35; 2.10: 3.55; 2.11: 5
| Std Dev. |  | 2.02: 0.19664; 2.04: 0.5164; 2.05: 0.10954; 2.07: 0.22943; 2.08: 18.3088; 2.10: 0.17512; 2.11: 0.40825
| Rel.Std.Dev. |  | 2.02: 4.645%; 2.04: 11.92%; 2.05: 2.961%; 2.07: 6.308%; 2.08: 117.1%; 2.10: 4.956%; 2.11: 8.447%
| PDM3 |  | 2.02: 4.97%; 2.04: 7.45%; 2.05: -8.254%; 2.07: -9.809%; 2.08: 287.6%; 2.10: -12.39%; 2.11: 19.85%
| Z-Score (Absolute) |  | 2.01: 1.95; 2.02: 0; 2.04: 0.11; 2.05: 0.6; 2.07: 0.67; 2.08: 12.81; 2.09: 23.34; 2.10: 0.79; 2.11: 0.67
| Analytical results for Sc in OREAS 922b (Certified Value 10.8 ppm) | [22 empty] | Certified
Sc | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.13; 2.07: 2.14; 2.08: 2.16; 2.09: 2.18; 2.10: 2.19; 2.11: 2.20; 2.12: 2.21; 2.13: 2.23 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*OES; 2.05: PF*OES; 2.07: PF*OES; 2.08: PF*OES; 2.09: PF*MS; 2.10: BF*OES/MS; 2.11: PF*OES; 2.12: PF*OES; 2.13: PF*OES |  |  | 1
2
| 1 | 1 | 2.01: < 20; 2.02: 10.1; 2.04: 11; 2.05: 11; 2.07: < 20; 2.08: 10; 2.09: 11; 2.10: 11; 2.11: 12; 2.12: 7; 2.13: 11 |  |  | 1
| 1 | 2 | 2.01: < 20; 2.02: 10; 2.04: 11; 2.05: 11; 2.07: < 20; 2.09: 11; 2.10: 12; 2.11: 12; 2.12: 6; 2.13: 11 |  |  | 19
| 1 | 3 | 2.01: < 20; 2.02: 10.4; 2.04: 11; 2.05: 11; 2.07: < 20; 2.09: 10; 2.10: 11; 2.11: 12; 2.12: 6; 2.13: 11 |  |  | 16
| 1 | 4 | 2.01: < 20; 2.02: 9.7; 2.04: 11; 2.05: 11; 2.07: < 20; 2.09: 10; 2.10: 11; 2.11: 12; 2.12: 7; 2.13: 11 |  |  | 10.8063
| 1 | 5 | 2.01: < 20; 2.02: 10.2; 2.04: 10; 2.05: 11; 2.07: < 20; 2.09: 10; 2.10: 11; 2.11: 12; 2.12: 6; 2.13: 11 |  |  | 188
| 1 | 6 | 2.01: < 20; 2.02: 10.3; 2.04: 11; 2.05: 11; 2.07: < 20; 2.09: 11; 2.10: 11; 2.11: 12; 2.12: 6; 2.13: 10
| Mean |  | 2.02: 10.1167; 2.04: 10.8333; 2.05: 11; 2.08: 10; 2.09: 10.5; 2.10: 11.1667; 2.11: 12; 2.12: 6.3333; 2.13: 10.8333
| Median |  | 2.02: 10.15; 2.04: 11; 2.05: 11; 2.08: 10; 2.09: 10.5; 2.10: 11; 2.11: 12; 2.12: 6; 2.13: 11
| Std Dev. |  | 2.02: 0.24833; 2.04: 0.40825; 2.05: 0; 2.09: 0.54772; 2.10: 0.40825; 2.11: 0; 2.12: 0.5164; 2.13: 0.40825
| Rel.Std.Dev. |  | 2.02: 2.455%; 2.04: 3.768%; 2.05: 0%; 2.09: 5.216%; 2.10: 3.656%; 2.11: 0%; 2.12: 8.154%; 2.13: 3.768%
| PDM3 |  | 2.02: -6.381%; 2.04: 0.2506%; 2.05: 1.793%; 2.08: -7.461%; 2.09: -2.834%; 2.10: 3.335%; 2.11: 11.05%; 2.12: -41.39%; 2.13: 0.2506%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.52; 2.04: 0.45; 2.05: 0.67; 2.07: 0.67; 2.08: 0.67; 2.09: 0; 2.10: 0.9; 2.11: 2.02; 2.12: 5.62; 2.13: 0.45
| Analytical results for Se in OREAS 922b (Indicative Value < 10 ppm) | [22 empty] | Indicative
Se | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab |  |  | 1
| No | No | 2.01: 2.12; 2.02: 2.14; 2.04: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*OES/MS |  |  | 1
1
| 1 | 1 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 1
| 1 | 2 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 41
| 1 | 3 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 16
| 1 | 4 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | < 10
| 1 | 5 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 27
| 1 | 6 | 2.01: < 8; 2.02: < 20; 2.04: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.04
| Analytical results for Si in OREAS 922b (Certified Value 35.21 wt.%) | [22 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 32.3641 | 34.02 | 34.5 | 35.71 | 37.49 | 35.71 | 35.01 | > 30 | > 30 | 29.1101 | 35.4 | 31.9 | 34.667 | 35.22 | 41.119 | 36.2 | 34.3 |  |  |  |  | 1
| 1 | 2 | 32.7577 | 35.41 | 35.99 | 35.39 | 37.68 | 35.11 | 34.69 | > 30 | > 30 | 30.386 | 35.5 | 34.8 | 34.742 | 35.19 | 41.67 | 36.5 | 35.1 |  |  |  |  | 0
| 1 | 3 | 32.972 | 35.25 | 36.09 | 35.99 | 37.49 | 35.11 | 34.92 | > 30 | > 30 | 29.5578 | 35.6 | 32 | 34.555 | 35.21 | 41.976 | 37 | 34.6 |  |  |  |  | 16
| 1 | 4 | 32.4428 | 34.41 | 35.99 | 35.11 | 37.49 | 34.41 | 35.01 | > 30 | > 30 | 29.7087 | 35.6 | 32.4 | 34.499 | 35.18 | 42.649 | 36.6 | 34.4 |  |  |  |  | 35.2128
| 1 | 5 | 32.5453 | 33.66 | 36.09 | 35.39 | 37.68 | 34.5 | 35.39 | > 30 | > 30 | 30.5041 | 35.7 | 32.7 | 34.634 | 35.22 | 42.929 | 36.5 | 33.3 |  |  |  |  | 189
| 1 | 6 | 32.5638 | 34.56 | 36.18 | 34.78 | 37.21 | 35.2 | 35.39 | > 30 | > 30 | 31.1436 | 35.4 | 31.7 | 34.779 | 35.21 | 41.685 | 36.4 | 34.9
| Mean |  | 32.6076 | 34.5517 | 35.8067 | 35.395 | 37.5067 | 35.0067 | 35.0683 |  |  | 30.0684 | 35.5333 | 32.5833 | 34.646 | 35.205 | 42.0047 | 36.5333 | 34.4333
| Median |  | 32.5545 | 34.485 | 36.04 | 35.39 | 37.49 | 35.11 | 35.01 |  |  | 30.0473 | 35.55 | 32.2 | 34.6505 | 35.21 | 41.8305 | 36.5 | 34.5
| Std Dev. |  | 0.22252 | 0.68151 | 0.64413 | 0.42726 | 0.17259 | 0.48294 | 0.27528 |  |  | 0.7423 | 0.12111 | 1.1444 | 0.10713 | 0.016432 | 0.67364 | 0.26583 | 0.6314
| Rel.Std.Dev. |  | 0.6824% | 1.972% | 1.799% | 1.207% | 0.4602% | 1.38% | 0.785% |  |  | 2.469% | 0.3408% | 3.512% | 0.3092% | 0.04667% | 1.604% | 0.7276% | 1.834%
| PDM3 |  | -7.398% | -1.878% | 1.686% | 0.5173% | 6.514% | -0.5855% | -0.4104% |  |  | -14.61% | 0.9102% | -7.467% | -1.61% | -0.02224% | 19.29% | 3.75% | -2.214%
| Z-Score (Absolute) |  | 2.62 | 0.55 | 0.79 | 0.35 | 2.59 | 0.07 | 0 | NA | NA | 5.31 | 0.5 | 2.64 | 0.45 | 0.15 | 7.37 | 1.56 | 0.67
| Analytical results for Sm in OREAS 922b (Certified Value 6.84 ppm) | [22 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: 6.9; 2.02: 7.4; 2.04: 6.601; 2.05: 7.2; 2.07: 7.3; 2.08: 7; 2.09: 6.43; 2.10: 6.41; 2.11: 6.71 |  |  | 1
| 1 | 2 | 2.01: 7.1; 2.02: 7.6; 2.04: 6.6159; 2.05: 6.9; 2.07: 7.3; 2.08: 7; 2.09: 6.49; 2.10: 6.41; 2.11: 7.07 |  |  | 4
| 1 | 3 | 2.01: 7.1; 2.02: 7.2; 2.04: 6.6307; 2.05: 6.6; 2.07: 6.8; 2.08: 7; 2.09: 6.49; 2.10: 6.26; 2.11: 6.93 |  |  | 16
| 1 | 4 | 2.01: 6.8; 2.02: 6.9; 2.04: 6.69; 2.05: 7; 2.07: 7.5; 2.08: 7; 2.09: 6.34; 2.10: 6.45; 2.11: 6.79 |  |  | 6.838
| 1 | 5 | 2.01: 7; 2.02: 6.8; 2.04: 6.9429; 2.05: 7; 2.07: 6.7; 2.08: 7; 2.09: 6.47; 2.10: 6.86; 2.11: 6.71 |  |  | 190
| 1 | 6 | 2.01: 7; 2.02: 7.5; 2.04: 6.5276; 2.05: 7.1; 2.07: 7.2; 2.08: 7; 2.09: 6.5; 2.10: 6.44; 2.11: 7.02
| Mean |  | 2.01: 6.9833; 2.02: 7.2333; 2.04: 6.668; 2.05: 6.9667; 2.07: 7.1333; 2.08: 7; 2.09: 6.4533; 2.10: 6.4717; 2.11: 6.8717
| Median |  | 2.01: 7; 2.02: 7.3; 2.04: 6.6233; 2.05: 7; 2.07: 7.25; 2.08: 7; 2.09: 6.48; 2.10: 6.425; 2.11: 6.86
| Std Dev. |  | 2.01: 0.1169; 2.02: 0.3266; 2.04: 0.14447; 2.05: 0.20656; 2.07: 0.31411; 2.08: 0; 2.09: 0.060882; 2.10: 0.20233; 2.11: 0.15728
| Rel.Std.Dev. |  | 2.01: 1.674%; 2.02: 4.515%; 2.04: 2.167%; 2.05: 2.965%; 2.07: 4.403%; 2.08: 0%; 2.09: 0.9434%; 2.10: 3.126%; 2.11: 2.289%
| PDM3 |  | 2.01: 2.126%; 2.02: 5.782%; 2.04: -2.486%; 2.05: 1.882%; 2.07: 4.32%; 2.08: 2.37%; 2.09: -5.625%; 2.10: -5.357%; 2.11: 0.493%
| Z-Score (Absolute) |  | 2.01: 0.19; 2.02: 0.91; 2.04: 0.73; 2.05: 0.14; 2.07: 0.62; 2.08: NA; 2.09: 1.35; 2.10: 1.3; 2.11: 0.14
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 922b (Certified Value 6.27 ppm) | [22 empty] | Certified
Sn | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.17; 2.10: 2.18; 2.11: 2.19; 2.12: 2.21; 2.13: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*OES; 2.13: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: < 5; 2.02: 7; 2.04: 6.434; 2.05: 6.7; 2.07: 7; 2.08: < 100; 2.09: < 50; 2.10: 10; 2.11: 5; 2.12: < 50; 2.13: 10 |  |  | 1
| 1 | 2 | 2.01: < 5; 2.02: 6; 2.04: 6.5446; 2.05: 6.4; 2.07: 7; 2.08: < 100; 2.09: < 50; 2.10: 11; 2.11: 5; 2.12: < 50; 2.13: 12 |  |  | 26
| 1 | 3 | 2.01: < 5; 2.02: 7; 2.04: 6.5273; 2.05: 5.5; 2.07: 6; 2.08: < 100; 2.09: < 50; 2.10: 10; 2.11: 6; 2.12: < 50; 2.13: 12 |  |  | 16
| 1 | 4 | 2.01: < 5; 2.02: 7; 2.04: 6.3181; 2.05: 6.1; 2.07: 7; 2.08: < 100; 2.09: < 50; 2.10: 22; 2.11: 5; 2.12: < 50; 2.13: 7 |  |  | 6.2725
| 1 | 5 | 2.01: < 5; 2.02: 7; 2.04: 6.3464; 2.05: 6.4; 2.07: 6; 2.08: < 100; 2.09: 100; 2.10: 12; 2.11: 6; 2.12: < 50; 2.13: 10 |  |  | 191
| 1 | 6 | 2.01: < 5; 2.02: 6; 2.04: 6.3049; 2.05: 5.6; 2.07: 7; 2.08: < 100; 2.09: < 50; 2.10: 11; 2.11: 6; 2.12: < 50; 2.13: 9
| Mean |  | 2.02: 6.6667; 2.04: 6.4126; 2.05: 6.1167; 2.07: 6.6667; 2.09: 100; 2.10: 12.6667; 2.11: 5.5; 2.13: 10
| Median |  | 2.02: 7; 2.04: 6.3902; 2.05: 6.25; 2.07: 7; 2.09: 100; 2.10: 11; 2.11: 5.5; 2.13: 10
| Std Dev. |  | 2.02: 0.5164; 2.04: 0.10579; 2.05: 0.47924; 2.07: 0.5164; 2.10: 4.6332; 2.11: 0.54772; 2.13: 1.8974
| Rel.Std.Dev. |  | 2.02: 7.746%; 2.04: 1.65%; 2.05: 7.835%; 2.07: 7.746%; 2.10: 36.58%; 2.11: 9.959%; 2.13: 18.97%
| PDM3 |  | 2.02: 6.284%; 2.04: 2.233%; 2.05: -2.485%; 2.07: 6.284%; 2.09: 1494%; 2.10: 101.9%; 2.11: -12.32%; 2.13: 59.43%
| Z-Score (Absolute) |  | 2.01: 0.84; 2.02: 0; 2.04: 0.05; 2.05: 0.11; 2.07: 0; 2.08: 8.77; 2.09: 6.24; 2.10: 1.21; 2.11: 0.24; 2.12: 3.71; 2.13: 0.67
| Analytical results for Sr in OREAS 922b (Certified Value 47.6 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 40 | 45.89 | 41.7 | 47.3 | 48.2443 | 60 | 44 | 44.1197 | 53 | < 50 | 44.3 | 26 | 44 | 46 | 54 | [6 empty] | 1
| 1 | 2 | 40 | 46.73 | 40.8 | 46.8 | 48.5344 | 55 | 44 | 39.2537 | 60 | 50 | 44.7 | 26 | 44 | 46 | 59 | [6 empty] | 27
| 1 | 3 | 40 | 48.61 | 42.4 | 45.6 | 49.0929 | 58 | 43 | 43.3157 | 52 | 50 | 45 | 26 | 43 | 45 | 57 | [6 empty] | 16
| 1 | 4 | 40 | 50.54 | 40.7 | 45.6 | 48.4501 | 54 | 44 | 39.891 | 51 | 50 | 43.2 | 26 | 44 | 45 | 57 | [6 empty] | 47.6273
| 1 | 5 | 40 | 48.75 | 41.7 | 45.2 | 48.0797 | 56 | 45 | 43.8556 | 53 | 50 | 44.5 | 25 | 44 | 44 | 54 | [6 empty] | 192
| 1 | 6 | 40 | 46.52 | 42.5 | 46 | 48.51 | 57 | 45 | 42.6845 | 54 | 50 | 44.8 | 26 | 44 | 45 | 54
| Mean |  | 40 | 47.84 | 41.6333 | 46.0833 | 48.4852 | 56.6667 | 44.1667 | 42.1867 | 53.8333 | 50 | 44.4167 | 25.8333 | 43.8333 | 45.1667 | 55.8333
| Median |  | 40 | 47.67 | 41.7 | 45.8 | 48.48 | 56.5 | 44 | 43.0001 | 53 | 50 | 44.6 | 26 | 44 | 45 | 55.5
| Std Dev. |  | 0 | 1.76 | 0.76333 | 0.80602 | 0.34527 | 2.1602 | 0.75277 | 2.0936 | 3.1885 | 0 | 0.64317 | 0.40825 | 0.40825 | 0.75277 | 2.137
| Rel.Std.Dev. |  | 0% | 3.679% | 1.833% | 1.749% | 0.7121% | 3.812% | 1.704% | 4.963% | 5.923% | 0% | 1.448% | 1.58% | 0.9314% | 1.667% | 3.827%
| PDM3 |  | -16.01% | 0.4465% | -12.59% | -3.242% | 1.801% | 18.98% | -7.266% | -11.42% | 13.03% | 4.982% | -6.741% | -45.76% | -7.966% | -5.166% | 17.23%
| Z-Score (Absolute) |  | NA | 0.59 | 0.98 | 0.15 | 0.76 | 2.83 | 0.34 | 0.84 | 2.11 | 0.08 | 0.27 | 4.98 | 0.42 | 0.08 | 2.62
Results from laboratory 2.01 were removed due to their 10 ppm reading resolution.
| Analytical results for Ta in OREAS 922b (Certified Value 0.97 ppm) | [22 empty] | Certified
Ta | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.16; 2.10: 2.18; 2.11: 2.19; 2.12: 2.23; 2.13: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*MS; 2.13: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 0.5; 2.02: 0.9; 2.04: 1.0547; 2.05: 3; 2.07: 1; 2.08: 0.88092; 2.09: 1.2; 2.10: 0.7; 2.11: 0.8; 2.12: 1.3; 2.13: < 3 |  |  | 1
| 1 | 2 | 2.01: < 0.5; 2.02: 0.9; 2.04: 0.99182; 2.05: 2.9; 2.07: 1; 2.08: 0.82136; 2.09: 1.2; 2.10: 0.8; 2.11: 0.8; 2.12: 0.9; 2.13: < 3 |  |  | 10
| 1 | 3 | 2.01: < 0.5; 2.02: 1; 2.04: 1.0319; 2.05: 2.2; 2.07: 0.9; 2.08: 0.86295; 2.09: 1; 2.10: 0.7; 2.11: 0.9; 2.12: 1; 2.13: < 3 |  |  | 16
| 1 | 4 | 2.01: < 0.5; 2.02: 1; 2.04: 1.0603; 2.05: 2.2; 2.07: 1; 2.08: 0.80084; 2.09: 1; 2.10: 1.3; 2.11: 0.9; 2.12: 0.9; 2.13: < 3 |  |  | 0.96677
| 1 | 5 | 2.01: < 0.5; 2.02: 1; 2.04: 0.97904; 2.05: 2.4; 2.07: 0.9; 2.08: 0.8923; 2.09: 1.1; 2.10: 0.7; 2.11: 0.8; 2.12: 2.3; 2.13: < 3 |  |  | 193
| 1 | 6 | 2.01: < 0.5; 2.02: 1; 2.04: 1.0074; 2.05: 3.3; 2.07: 1; 2.08: 0.82136; 2.09: 1.7; 2.10: 1.2; 2.11: 0.8; 2.12: 1.4; 2.13: < 3
| Mean |  | 2.02: 0.96667; 2.04: 1.0209; 2.05: 2.6667; 2.07: 0.96667; 2.08: 0.84662; 2.09: 1.2; 2.10: 0.9; 2.11: 0.83333; 2.12: 1.3
| Median |  | 2.02: 1; 2.04: 1.0196; 2.05: 2.65; 2.07: 1; 2.08: 0.84215; 2.09: 1.15; 2.10: 0.75; 2.11: 0.8; 2.12: 1.15
| Std Dev. |  | 2.02: 0.05164; 2.04: 0.033452; 2.05: 0.46332; 2.07: 0.05164; 2.08: 0.037155; 2.09: 0.26077; 2.10: 0.27568; 2.11: 0.05164; 2.12: 0.53292
| Rel.Std.Dev. |  | 2.02: 5.342%; 2.04: 3.277%; 2.05: 17.37%; 2.07: 5.342%; 2.08: 4.389%; 2.09: 21.73%; 2.10: 30.63%; 2.11: 6.197%; 2.12: 40.99%
| PDM3 |  | 2.02: -0.01042%; 2.04: 5.594%; 2.05: 175.8%; 2.07: -0.01042%; 2.08: -12.43%; 2.09: 24.12%; 2.10: -6.906%; 2.11: -13.8%; 2.12: 34.47%
| Z-Score (Absolute) |  | 2.01: 3.62; 2.02: 0; 2.04: 0.27; 2.05: 8.6; 2.07: 0; 2.08: 0.61; 2.09: 1.18; 2.10: 0.34; 2.11: 0.67; 2.12: 1.69; 2.13: 2.7
| Analytical results for Tb in OREAS 922b (Certified Value 0.9 ppm) | [22 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 0.9; 2.02: 0.8; 2.04: 0.92273; 2.05: 0.7; 2.07: 0.78; 2.08: < 2; 2.09: 0.8; 2.10: 0.9; 2.11: 1.14 |  |  | 1
| 1 | 2 | 2.01: 0.8; 2.02: 0.8; 2.04: 0.87771; 2.05: 0.9; 2.07: 0.8; 2.08: < 2; 2.09: 0.8; 2.10: 0.87; 2.11: 1.18 |  |  | 11
| 1 | 3 | 2.01: 0.9; 2.02: 0.8; 2.04: 0.85177; 2.05: 0.7; 2.07: 0.78; 2.08: < 2; 2.09: 0.81; 2.10: 0.9; 2.11: 1.1 |  |  | 16
| 1 | 4 | 2.01: 0.8; 2.02: 0.8; 2.04: 0.87586; 2.05: 0.8; 2.07: 0.82; 2.08: < 2; 2.09: 0.78; 2.10: 0.87; 2.11: 1.16 |  |  | 0.89995
| 1 | 5 | 2.01: 0.9; 2.02: 0.8; 2.04: 0.87425; 2.05: 0.9; 2.07: 0.76; 2.08: < 2; 2.09: 0.79; 2.10: 0.91; 2.11: 1.15 |  |  | 194
| 1 | 6 | 2.01: 0.9; 2.02: 0.8; 2.04: 0.89628; 2.05: 0.9; 2.07: 0.82; 2.08: < 2; 2.09: 0.8; 2.10: 0.88; 2.11: 1.1
| Mean |  | 2.01: 0.86667; 2.02: 0.8; 2.04: 0.8831; 2.05: 0.81667; 2.07: 0.79333; 2.09: 0.79667; 2.10: 0.88833; 2.11: 1.1383
| Median |  | 2.01: 0.9; 2.02: 0.8; 2.04: 0.87679; 2.05: 0.85; 2.07: 0.79; 2.09: 0.8; 2.10: 0.89; 2.11: 1.145
| Std Dev. |  | 2.01: 0.05164; 2.02: 1.2162e-16; 2.04: 0.024022; 2.05: 0.098319; 2.07: 0.024221; 2.09: 0.010328; 2.10: 0.017224; 2.11: 0.032506
| Rel.Std.Dev. |  | 2.01: 5.958%; 2.02: 1.52e-14%; 2.04: 2.72%; 2.05: 12.04%; 2.07: 3.053%; 2.09: 1.296%; 2.10: 1.939%; 2.11: 2.856%
| PDM3 |  | 2.01: -3.699%; 2.02: -11.11%; 2.04: -1.873%; 2.05: -9.255%; 2.07: -11.85%; 2.09: -11.48%; 2.10: -1.291%; 2.11: 26.49%
| Z-Score (Absolute) |  | 2.01: NA; 2.02: NA; 2.04: 0.02; 2.05: NA; 2.07: 0.69; 2.08: 0.85; 2.09: 0.66; 2.10: 0.02; 2.11: 1.88
Results from laboratories 2.04, 2.05 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 922b (Indicative Value < 1 ppm) | [22 empty] | Indicative
Te | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab |  |  | 1
| No | No | 2.01: 2.12; 2.02: 2.13; 2.04: 2.14; 2.05: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*MS; 2.05: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 8; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 1
| 1 | 2 | 2.01: 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 12
| 1 | 3 | 2.01: 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 16
| 1 | 4 | 2.01: 7; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | < 1
| 1 | 5 | 2.01: < 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 28
| 1 | 6 | 2.01: 8; 2.02: < 1; 2.04: < 1; 2.05: < 5
| Mean |  | 2.01: 7
| Median |  | 2.01: 7
Std Dev. | bar: 2.01
| Rel.Std.Dev. |  | 2.01: 14.29%
PDM3
| Z-Score (Absolute) |  | 2.01: 3.26; 2.02: 0.67; 2.04: 0.67; 2.05: 0.67
| Analytical results for Th in OREAS 922b (Certified Value 13.9 ppm) | [22 empty] | Certified
Th | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.14; 2.10: 2.16; 2.11: 2.19; 2.12: 2.23; 2.13: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*MS; 2.13: PF*OES/MS |  |  | 1
2
| 1 | 1 | 2.01: 14.1; 2.02: 12.9; 2.04: 12.6; 2.05: 14.9212; 2.07: 16.3; 2.08: 14.4; 2.09: 13.2778; 2.10: 13.8; 2.11: 13.5; 2.12: 14.6; 2.13: 14 |  |  | 1
| 1 | 2 | 2.01: 13.9; 2.02: 13.7; 2.04: 14.7; 2.05: 15.2569; 2.07: 14.4; 2.08: 14.4; 2.09: 12.4285; 2.10: 14.5; 2.11: 13.4; 2.12: 14.4; 2.13: 15 |  |  | 13
| 1 | 3 | 2.01: 13.8; 2.02: 13.5; 2.04: 14.5; 2.05: 15.0536; 2.07: 15.1; 2.08: 13.8; 2.09: 12.9425; 2.10: 13.7; 2.11: 13.7; 2.12: 14.3; 2.13: 14 |  |  | 16
| 1 | 4 | 2.01: 13.7; 2.02: 12.4; 2.04: 14.1; 2.05: 15.2704; 2.07: 15.3; 2.08: 14.6; 2.09: 12.6485; 2.10: 14.1; 2.11: 13.6; 2.12: 14.3; 2.13: 15 |  |  | 13.9161
| 1 | 5 | 2.01: 14.1; 2.02: 13.2; 2.04: 13.6; 2.05: 15.1569; 2.07: 15.8; 2.08: 13.9; 2.09: 13.0038; 2.10: 14.4; 2.11: 13.7; 2.12: 14.2; 2.13: 15 |  |  | 195
| 1 | 6 | 2.01: 14.3; 2.02: 12.5; 2.04: 14; 2.05: 15.6; 2.07: 15.6; 2.08: 14.7; 2.09: 12.8112; 2.10: 14; 2.11: 13.6; 2.12: 13.9; 2.13: 15
| Mean |  | 2.01: 13.9833; 2.02: 13.0333; 2.04: 13.9167; 2.05: 15.2098; 2.07: 15.4167; 2.08: 14.3; 2.09: 12.852; 2.10: 14.0833; 2.11: 13.5833; 2.12: 14.2833; 2.13: 14.6667
| Median |  | 2.01: 14; 2.02: 13.05; 2.04: 14.05; 2.05: 15.2069; 2.07: 15.45; 2.08: 14.4; 2.09: 12.8769; 2.10: 14.05; 2.11: 13.6; 2.12: 14.3; 2.13: 15
| Std Dev. |  | 2.01: 0.22286; 2.02: 0.52789; 2.04: 0.75211; 2.05: 0.2318; 2.07: 0.64936; 2.08: 0.36878; 2.09: 0.29494; 2.10: 0.31885; 2.11: 0.1169; 2.12: 0.23166; 2.13: 0.5164
| Rel.Std.Dev. |  | 2.01: 1.594%; 2.02: 4.05%; 2.04: 5.404%; 2.05: 1.524%; 2.07: 4.212%; 2.08: 2.579%; 2.09: 2.295%; 2.10: 2.264%; 2.11: 0.8606%; 2.12: 1.622%; 2.13: 3.521%
| PDM3 |  | 2.01: 0.4829%; 2.02: -6.344%; 2.04: 0.003836%; 2.05: 9.296%; 2.07: 10.78%; 2.08: 2.758%; 2.09: -7.646%; 2.10: 1.201%; 2.11: -2.391%; 2.12: 2.639%; 2.13: 5.393%
| Z-Score (Absolute) |  | 2.01: 0.09; 2.02: 1.88; 2.04: 0.22; 2.05: 2.21; 2.07: 2.6; 2.08: 0.5; 2.09: 2.22; 2.10: 0.09; 2.11: 0.85; 2.12: 0.47; 2.13: NA
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 922b (Certified Value 0.31 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.31 | 0.29752 | 0.29 | 0.308 | 0.31975 | 0.3 | 0.312 | 0.312 | 0.312 | 0.396 | 0.32 | 0.31 | 0.31727 | 0.32 | 0.31 | 0.312 | 0.31 | 0.288 | 0.31 | 0.312 |  | 1
| 1 | 2 | 0.31 | 0.29899 | 0.26 | 0.321 | 0.31883 | 0.306 | 0.312 | 0.3 | 0.294 | 0.3 | 0.31 | 0.31 | 0.28726 | 0.32 | 0.32 | 0.312 | 0.32 | 0.292 | 0.32 | 0.312 |  | 32
| 1 | 3 | 0.31 | 0.29164 | 0.28 | 0.318 | 0.32038 | 0.312 | 0.312 | 0.3 | 0.3 | 0.3 | 0.31 | 0.31 | 0.29729 | 0.32 | 0.31 | 0.312 | 0.31 | 0.278 | 0.32 | 0.319 |  | 16
| 1 | 4 | 0.31 | 0.29415 | 0.27 | 0.311 | 0.32367 | 0.306 | 0.3 | 0.318 | 0.288 | 0.3 | 0.31 | 0.31 | 0.29265 | 0.32 | 0.31 | 0.312 | 0.32 | 0.289 | 0.31 | 0.324 |  | 0.30958
| 1 | 5 | 0.31 | 0.29414 | 0.28 | 0.305 | 0.31804 | 0.306 | 0.312 | 0.318 | 0.294 | 0.306 | 0.31 | 0.31 | 0.31006 | 0.32 | 0.32 | 0.312 | 0.32 | 0.302 | 0.31 | 0.306 |  | 196
| 1 | 6 | 0.31 | 0.29775 | 0.29 | 0.312 | 0.31941 | 0.312 | 0.3 | 0.342 | 0.294 | 0.306 | 0.31 | 0.3 | 0.3085 | 0.32 | 0.31 | 0.306 | 0.32 | 0.285 | 0.31 | 0.319
| Mean |  | 0.31 | 0.2957 | 0.27833 | 0.3125 | 0.32002 | 0.307 | 0.308 | 0.315 | 0.297 | 0.318 | 0.31167 | 0.30833 | 0.30217 | 0.32 | 0.31333 | 0.311 | 0.31667 | 0.289 | 0.31333 | 0.31533
| Median |  | 0.31 | 0.29584 | 0.28 | 0.3115 | 0.31958 | 0.306 | 0.312 | 0.315 | 0.294 | 0.303 | 0.31 | 0.31 | 0.3029 | 0.32 | 0.31 | 0.312 | 0.32 | 0.2885 | 0.31 | 0.3155
| Std Dev. |  | 0 | 0.0028157 | 0.01169 | 0.0060249 | 0.0019607 | 0.0045166 | 0.0061968 | 0.015531 | 0.0082704 | 0.038325 | 0.0040825 | 0.0040825 | 0.011551 | 0 | 0.005164 | 0.0024495 | 0.005164 | 0.0079498 | 0.005164 | 0.0065013
| Rel.Std.Dev. |  | 0% | 0.9522% | 4.2% | 1.928% | 0.6127% | 1.471% | 2.012% | 4.93% | 2.785% | 12.05% | 1.31% | 1.324% | 3.823% | 0% | 1.648% | 0.7876% | 1.631% | 2.751% | 1.648% | 2.062%
| PDM3 |  | 0.135% | -4.485% | -10.09% | 0.9425% | 3.37% | -0.8341% | -0.511% | 1.75% | -4.064% | 2.719% | 0.6734% | -0.4034% | -2.393% | 3.365% | 1.212% | 0.458% | 2.288% | -6.648% | 1.212% | 1.858%
| Z-Score (Absolute) |  | 0.19 | 2.43 | 5.16 | 0.2 | 1.38 | 0.7 | 0.54 | 0.55 | 2.27 | 1.02 | 0.07 | 0.45 | 1.42 | 1.38 | 0.33 | 0.07 | 0.86 | 3.48 | 0.33 | 0.65
| Analytical results for Tl in OREAS 922b (Certified Value 0.99 ppm) | [22 empty] | Certified
Tl | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.23; 2.09: 2.24 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES/MS |  | 2
3
| 1 | 1 | 2.01: 0.9; 2.02: < 1; 2.04: 1.1; 2.05: 1; 2.06: 0.99697; 2.07: 1.1; 2.08: 1; 2.09: 1.02 |  | 1
| 1 | 2 | 2.01: 0.9; 2.02: < 1; 2.04: 1.1; 2.05: 1; 2.06: 0.83812; 2.07: 1; 2.08: 1; 2.09: 1.02 |  | 14
| 1 | 3 | 2.01: 0.9; 2.02: < 1; 2.04: 0.9; 2.05: 1; 2.06: 0.88126; 2.07: 1.1; 2.08: 0.9; 2.09: 1 |  | 16
| 1 | 4 | 2.01: 0.9; 2.02: < 1; 2.04: 1.1; 2.05: 1; 2.06: 1.0076; 2.07: 1; 2.08: 0.9; 2.09: 1 |  | 0.9917
| 1 | 5 | 2.01: 0.9; 2.02: < 1; 2.04: 1; 2.05: 1; 2.06: 0.95366; 2.07: 1.1; 2.08: 1; 2.09: 1.02 |  | 197
| 1 | 6 | 2.01: 0.9; 2.02: < 1; 2.04: 1.1; 2.05: 1; 2.06: 0.9017; 2.07: 1.2; 2.08: 1; 2.09: 1.08
| Mean |  | 2.01: 0.9; 2.04: 1.05; 2.05: 1; 2.06: 0.92988; 2.07: 1.0833; 2.08: 0.96667; 2.09: 1.0233
| Median |  | 2.01: 0.9; 2.04: 1.1; 2.05: 1; 2.06: 0.92768; 2.07: 1.1; 2.08: 1; 2.09: 1.02
| Std Dev. |  | 2.01: 0; 2.04: 0.083666; 2.05: 0; 2.06: 0.06735; 2.07: 0.075277; 2.08: 0.05164; 2.09: 0.029439
| Rel.Std.Dev. |  | 2.01: 0%; 2.04: 7.968%; 2.05: 0%; 2.06: 7.243%; 2.07: 6.949%; 2.08: 5.342%; 2.09: 2.877%
| PDM3 |  | 2.01: -9.246%; 2.04: 5.879%; 2.05: 0.8373%; 2.06: -6.234%; 2.07: 9.24%; 2.08: -2.524%; 2.09: 3.19%
| Z-Score (Absolute) |  | 2.01: 0.94; 2.02: 5.43; 2.04: 0.75; 2.05: 0.19; 2.06: 0.6; 2.07: 1.12; 2.08: 0.19; 2.09: 0.45
| Analytical results for Tm in OREAS 922b (Certified Value 0.4 ppm) | [22 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.18; 2.08: 2.19; 2.09: 2.22; 2.10: 2.24 |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: BF*OES/MS; 2.09: BF*MS; 2.10: PF*OES/MS |  | 2
3
| 1 | 1 | 2.01: < 1; 2.02: 0.4; 2.04: 0.4727; 2.05: 0.4; 2.06: 0.41; 2.07: < 2; 2.08: 0.39; 2.09: 0.41; 2.10: 0.4 |  | 1
| 1 | 2 | 2.01: < 1; 2.02: 0.4; 2.04: 0.48534; 2.05: 0.4; 2.06: 0.39; 2.07: < 2; 2.08: 0.38; 2.09: 0.44; 2.10: 0.43 |  | 5
| 1 | 3 | 2.01: < 1; 2.02: 0.4; 2.04: 0.46516; 2.05: 0.4; 2.06: 0.39; 2.07: < 2; 2.08: 0.4; 2.09: 0.41; 2.10: 0.41 |  | 16
| 1 | 4 | 2.01: < 1; 2.02: 0.4; 2.04: 0.48722; 2.05: 0.4; 2.06: 0.37; 2.07: < 2; 2.08: 0.4; 2.09: 0.39; 2.10: 0.41 |  | 0.40111
| 1 | 5 | 2.01: < 1; 2.02: 0.4; 2.04: 0.46662; 2.05: 0.4; 2.06: 0.39; 2.07: < 2; 2.08: 0.39; 2.09: 0.43; 2.10: 0.4 |  | 198
| 1 | 6 | 2.01: < 1; 2.02: 0.4; 2.04: 0.47044; 2.05: 0.5; 2.06: 0.38; 2.07: < 2; 2.08: 0.42; 2.09: 0.39; 2.10: 0.41
| Mean |  | 2.02: 0.4; 2.04: 0.47458; 2.05: 0.41667; 2.06: 0.38833; 2.08: 0.39667; 2.09: 0.41167; 2.10: 0.41
| Median |  | 2.02: 0.4; 2.04: 0.47157; 2.05: 0.4; 2.06: 0.39; 2.08: 0.395; 2.09: 0.41; 2.10: 0.41
| Std Dev. |  | 2.02: 6.0809e-17; 2.04: 0.0094701; 2.05: 0.040825; 2.06: 0.013292; 2.08: 0.013663; 2.09: 0.020412; 2.10: 0.010954
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.04: 1.995%; 2.05: 9.798%; 2.06: 3.423%; 2.08: 3.444%; 2.09: 4.958%; 2.10: 2.672%
| PDM3 |  | 2.02: -0.277%; 2.04: 18.32%; 2.05: 3.878%; 2.06: -3.186%; 2.08: -1.108%; 2.09: 2.632%; 2.10: 2.216%
| Z-Score (Absolute) |  | 2.01: 3.97; 2.02: 0.52; 2.04: 2.83; 2.05: 0.22; 2.06: 1.05; 2.07: 26.45; 2.08: 0.67; 2.09: 0; 2.10: 0.07
| Analytical results for U in OREAS 922b (Certified Value 4.54 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.19 | 2.22 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 4.3 | 4.4 | 4.7 | 5.4976 | 4.8 | 4.58 | 4.4451 | 4.2 | 4.5 | 4.57 | 4.7 | 6 | [9 empty] | 1
| 1 | 2 | 4.4 | 4.4 | 4.7 | 5.678 | 4.8 | 4.57 | 3.9811 | 5 | 4.5 | 4.57 | 4.7 | 6 | [9 empty] | 16
| 1 | 3 | 4.1 | 4.5 | 4.6 | 5.5046 | 4.3 | 5 | 3.9442 | 4.4 | 4.5 | 4.59 | 4.8 | 6 | [9 empty] | 16
| 1 | 4 | 4.1 | 4.5 | 4.5 | 5.5175 | 4.8 | 4.59 | 4.5982 | 4.6 | 4.4 | 4.59 | 4.6 | 6 | [9 empty] | 4.5355
| 1 | 5 | 4.1 | 4.5 | 4.6 | 5.4761 | 4.9 | 4.94 | 4.3164 | 4.5 | 4.6 | 4.61 | 4.5 | 5 | [9 empty] | 199
| 1 | 6 | 4.4 | 4.5 | 4.6 | 5.4046 | 4.8 | 4.6 | 4.3568 | 4.6 | 4.5 | 4.7 | 4.5 | 5
| Mean |  | 4.2333 | 4.4667 | 4.6167 | 5.5131 | 4.7333 | 4.7133 | 4.2736 | 4.55 | 4.5 | 4.605 | 4.6333 | 5.6667
| Median |  | 4.2 | 4.5 | 4.6 | 5.5011 | 4.8 | 4.595 | 4.3366 | 4.55 | 4.5 | 4.59 | 4.65 | 6
| Std Dev. |  | 0.15055 | 0.05164 | 0.075277 | 0.090192 | 0.21602 | 0.19997 | 0.25984 | 0.26646 | 0.063246 | 0.048888 | 0.12111 | 0.5164
| Rel.Std.Dev. |  | 3.556% | 1.156% | 1.631% | 1.636% | 4.564% | 4.243% | 6.08% | 5.856% | 1.405% | 1.062% | 2.614% | 9.113%
| PDM3 |  | -6.661% | -1.517% | 1.79% | 21.55% | 4.363% | 3.922% | -5.773% | 0.3206% | -0.7819% | 1.533% | 2.158% | 24.94%
| Z-Score (Absolute) |  | 2.31 | 0.86 | 0.07 | 5.65 | 0.8 | 0.67 | 2.06 | 0.34 | 0.65 | 0 | 0.18 | NA
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 922b (Certified Value 193 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 190 | 183.055 | 205.9 | 190 | 216 | 190 | 181.1279 | 197 | 190 | 206 | 180 | 196 | 193 | 192 | [7 empty] | 1
| 1 | 2 | 180 | 178.9655 | 199.4 | 196 | 211 | 191 | 179.0383 | 193 | 200 | 207 | 179 | 195 | 198 | 200 | [7 empty] | 36
| 1 | 3 | 180 | 179.88 | 204.9 | 197 | 205 | 190 | 177.9499 | 197 | 190 | 209 | 179 | 195 | 199 | 198 | [7 empty] | 16
| 1 | 4 | 190 | 181.4495 | 203 | 193 | 196 | 190 | 187.1901 | 193 | 190 | 209 | 179 | 196 | 192 | 192 | [7 empty] | 192.6098
| 1 | 5 | 180 | 181.611 | 203.3 | 185 | 204 | 190 | 184.7851 | 193 | 190 | 205 | 180 | 196 | 196 | 192 | [7 empty] | 200
| 1 | 6 | 180 | 183.055 | 204.9 | 191 | 201 | 188 | 175.7148 | 200 | 190 | 206 | 180 | 196 | 195 | 197
| Mean |  | 183.3333 | 181.336 | 203.5667 | 192 | 205.5 | 189.8333 | 180.9677 | 195.5 | 191.6667 | 207 | 179.5 | 195.6667 | 195.5 | 195.1667
| Median |  | 180 | 181.5302 | 204.1 | 192 | 204.5 | 190 | 180.0831 | 195 | 190 | 206.5 | 179.5 | 196 | 195.5 | 194.5
| Std Dev. |  | 5.164 | 1.6576 | 2.3132 | 4.3818 | 7.1204 | 0.98319 | 4.3299 | 2.9496 | 4.0825 | 1.6733 | 0.54772 | 0.5164 | 2.7386 | 3.6009
| Rel.Std.Dev. |  | 2.817% | 0.9141% | 1.136% | 2.282% | 3.465% | 0.5179% | 2.393% | 1.509% | 2.13% | 0.8084% | 0.3051% | 0.2639% | 1.401% | 1.845%
| PDM3 |  | -4.816% | -5.853% | 5.689% | -0.3166% | 6.692% | -1.441% | -6.044% | 1.501% | -0.4897% | 7.471% | -6.806% | 1.587% | 1.501% | 1.327%
| Z-Score (Absolute) |  | 1.01 | 1.2 | 0.98 | 0.16 | 1.17 | 0.37 | 1.24 | 0.19 | 0.19 | 1.32 | 1.38 | 0.2 | 0.19 | 0.16
| Analytical results for W in OREAS 922b (Certified Value 4.07 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.21 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 5 | 3 | 5.4248 | 4.6 | 3 | 2.5307 | 5 | < 100 | 3.6 | < 50 | 3 | < 5 | [9 empty] | 1
| 1 | 2 | < 5 | 4 | 5.4371 | 7.3 | 3 | 2.3662 | 5 | < 100 | 3.4 | < 50 | 3 | 5 | [9 empty] | 17
| 1 | 3 | < 5 | 4 | 5.5442 | 3.5 | 3 | 2.5389 | 3 | < 100 | 3.9 | < 50 | 4 | < 5 | [9 empty] | 16
| 1 | 4 | < 5 | 3 | 5.1776 | 4.2 | 3 | 2.5003 | 7 | < 100 | 3.7 | < 50 | 4 | 6 | [9 empty] | 4.0707
| 1 | 5 | < 5 | 4 | 5.4655 | 3.5 | 3 | 2.4481 | 6 | < 100 | 3.7 | < 50 | 4 | < 5 | [9 empty] | 201
| 1 | 6 | < 5 | 3 | 5.4173 | 4.6 | 3 | 2.4047 | 6 | < 100 | 3.5 | < 50 | 4 | < 5
| Mean |  | 2.02: 3.5; 2.04: 5.4111; 2.05: 4.6167; 2.06: 3; 2.07: 2.4648; 2.08: 5.3333; 2.10: 3.6333; 2.12: 3.6667; 2.13: 5.5
| Median |  | 2.02: 3.5; 2.04: 5.431; 2.05: 4.4; 2.06: 3; 2.07: 2.4742; 2.08: 5.5; 2.10: 3.65; 2.12: 4; 2.13: 5.5
| Std Dev. |  | 2.02: 0.54772; 2.04: 0.12336; 2.05: 1.4049; 2.06: 0; 2.07: 0.070295; 2.08: 1.3663; 2.10: 0.17512; 2.12: 0.5164; 2.13: 0.70711
| Rel.Std.Dev. |  | 2.02: 15.65%; 2.04: 2.28%; 2.05: 30.43%; 2.06: 0%; 2.07: 2.852%; 2.08: 25.62%; 2.10: 4.82%; 2.12: 14.08%; 2.13: 12.86%
| PDM3 |  | 2.02: -14.02%; 2.04: 32.93%; 2.05: 13.41%; 2.06: -26.3%; 2.07: -39.45%; 2.08: 31.02%; 2.10: -10.74%; 2.12: -9.924%; 2.13: 35.11%
| Z-Score (Absolute) |  | 0.73 | 0.1 | 1.12 | 0.62 | 0.41 | 0.76 | 1.07 | 29.53 | 0.01 | 13.6 | 0.01 | 0.1
| Analytical results for Y in OREAS 922b (Certified Value 26.6 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*OES | BF*MS | PF*MS | PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 25 | 24.3 | 27.9 | 28.8942 | 28.6 | 28.5 | 25.603 | 26.4 | 27 | 25.5 | 27 | 22 | 26.5 | 27 | 27 | [6 empty] | 1
| 1 | 2 | 25.4 | 22.5 | 27.9 | 28.8659 | 29.5 | 28.6 | 24.5163 | 27 | 27 | 25.7 | 26 | 21 | 25.5 | 26.3 | 29 | [6 empty] | 18
| 1 | 3 | 24.7 | 23.9 | 27.8 | 26.9153 | 27.3 | 27 | 24.3788 | 26.5 | 28 | 26.4 | 27 | 21 | 26.1 | 26.4 | 27 | [6 empty] | 16
| 1 | 4 | 24.6 | 22.4 | 27 | 26.6493 | 28.4 | 29.1 | 25.1718 | 27 | 26 | 25.7 | 26 | 21 | 25.3 | 26.1 | 28 | [6 empty] | 26.5616
| 1 | 5 | 24.6 | 25 | 27 | 28.4261 | 28.6 | 26.2 | 24.36 | 26.6 | 27 | 26.4 | 26 | 22 | 26.8 | 26.7 | 28 | [6 empty] | 202
| 1 | 6 | 25.4 | 24.2 | 27.5 | 27.9738 | 28.3 | 29.9 | 25.3221 | 27 | 27 | 25.9 | 26 | 22 | 25.6 | 26.6 | 27
| Mean |  | 24.95 | 23.7167 | 27.5167 | 27.9541 | 28.45 | 28.2167 | 24.892 | 26.75 | 27 | 25.9333 | 26.3333 | 21.5 | 25.9667 | 26.5167 | 27.6667
| Median |  | 24.85 | 24.05 | 27.65 | 28.1999 | 28.5 | 28.55 | 24.844 | 26.8 | 27 | 25.8 | 26 | 21.5 | 25.85 | 26.5 | 27.5
| Std Dev. |  | 0.37815 | 1.0458 | 0.42622 | 0.97157 | 0.7064 | 1.3703 | 0.53969 | 0.28107 | 0.63246 | 0.38297 | 0.5164 | 0.54772 | 0.59889 | 0.31885 | 0.8165
| Rel.Std.Dev. |  | 1.516% | 4.409% | 1.549% | 3.476% | 2.483% | 4.856% | 2.168% | 1.051% | 2.342% | 1.477% | 1.961% | 2.548% | 2.306% | 1.202% | 2.951%
| PDM3 |  | -6.067% | -10.71% | 3.596% | 5.242% | 7.109% | 6.231% | -6.286% | 0.7092% | 1.65% | -2.365% | -0.8595% | -19.06% | -2.24% | -0.1693% | 4.16%
| Z-Score (Absolute) |  | 0.92 | 1.64 | 0.59 | 0.84 | 1.13 | 1 | 0.95 | 0.14 | 0.28 | 0.34 | 0.11 | 2.94 | 0.32 | 0 | 0.67
| Analytical results for Yb in OREAS 922b (Certified Value 2.72 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 2.5 | 2.54 | 2.8842 | 3.3 | 2.6 | 3 | 2.6 | 2.65 | 2.62 | [12 empty] | 1
| 1 | 2 | 2.4 | 2.79 | 2.8687 | 2.2 | 2.7 | 3 | 2.62 | 2.71 | 2.83 | [12 empty] | 19
| 1 | 3 | 2.5 | 2.87 | 2.8667 | 2.6 | 2.5 | 3 | 2.78 | 2.61 | 2.73 | [12 empty] | 16
| 1 | 4 | 2.3 | 2.69 | 2.8735 | 2.7 | 2.7 | 2 | 2.54 | 2.57 | 2.76 | [12 empty] | 2.7204
| 1 | 5 | 2.5 | 2.58 | 2.8487 | 3.1 | 2.6 | 2 | 2.73 | 2.71 | 2.77 | [12 empty] | 203
| 1 | 6 | 2.5 | 2.73 | 2.8167 | 3.4 | 2.9 | 2 | 2.63 | 2.63 | 2.73
| Mean |  | 2.45 | 2.7 | 2.8597 | 2.8833 | 2.6667 | 2.5 | 2.65 | 2.6467 | 2.74
| Median |  | 2.5 | 2.71 | 2.8677 | 2.9 | 2.65 | 2.5 | 2.625 | 2.64 | 2.745
| Std Dev. |  | 0.083666 | 0.1249 | 0.024046 | 0.46224 | 0.13663 | 0.54772 | 0.088544 | 0.055737 | 0.069282
| Rel.Std.Dev. |  | 3.415% | 4.626% | 0.8409% | 16.03% | 5.123% | 21.91% | 3.341% | 2.106% | 2.529%
| PDM3 |  | -9.941% | -0.7513% | 5.12% | 5.988% | -1.977% | -8.103% | -2.589% | -2.712% | 0.7191%
| Z-Score (Absolute) |  | 3.37 | 0.24 | 2.55 | 2.89 | 0.24 | NA | 0.48 | 0.53 | 0.82
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 922b (Certified Value 76 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 70 | 100 | 60 | 100 | 110 | 70 | 80 | 70 | 74 | 80.1114 | 79.45 | 70 | 66 | 95 | 78 | 73 | [5 empty] | 1
| 1 | 2 | 70 | 99.2 | 70 | 100 | 110 | 70 | 90 | 90 | 74 | 83.4141 | 78.33 | 80 | 66 | 90 | 85 | 75 | [5 empty] | 40
| 1 | 3 | 80 | 107.9 | 70 | 100 | 110 | 70 | 70 | 60 | 74 | 80.93 | 76.94 | 70 | 66 | 88 | 82 | 82 | [5 empty] | 16
| 1 | 4 | 80 | 100.2 | 80 | 100 | 110 | 90 | 70 | 70 | 74 | 78.155 | 78.3 | 80 | 66 | 90 | 81 | 70 | [5 empty] | 75.5121
| 1 | 5 | 70 | 99.8 | 70 | 100 | 100 | 70 | 70 | 60 | 75 | 77.933 | 76.42 | 80 | 67 | 93 | 80 | 73 | [5 empty] | 204
| 1 | 6 | 70 | 101.6 | 70 | 100 | 100 | 60 | 70 | 60 | 74 | 82.1575 | 74.8 | 80 | 67 | 92 | 84 | 79
| Mean |  | 73.3333 | 101.45 | 70 | 100 | 106.6667 | 71.6667 | 75 | 68.3333 | 74.1667 | 80.4502 | 77.3733 | 76.6667 | 66.3333 | 91.3333 | 81.6667 | 75.3333
| Median |  | 70 | 100.1 | 70 | 100 | 110 | 70 | 70 | 65 | 74 | 80.5207 | 77.62 | 80 | 66 | 91 | 81.5 | 74
| Std Dev. |  | 5.164 | 3.2581 | 6.3246 | 0 | 5.164 | 9.8319 | 8.3666 | 11.6905 | 0.40825 | 2.1748 | 1.6616 | 5.164 | 0.5164 | 2.5033 | 2.582 | 4.4121
| Rel.Std.Dev. |  | 7.042% | 3.212% | 9.035% | 0% | 4.841% | 13.72% | 11.16% | 17.11% | 0.5504% | 2.703% | 2.148% | 6.736% | 0.7785% | 2.741% | 3.162% | 5.857%
| PDM3 |  | -2.885% | 34.35% | -7.3% | 32.43% | 41.26% | -5.092% | -0.6781% | -9.507% | -1.782% | 6.54% | 2.465% | 1.529% | -12.16% | 20.95% | 8.15% | -0.2367%
| Z-Score (Absolute) |  | 0.36 | 3.39 | 0.8 | 3.2 | 4.09 | 0.58 | 0.13 | 1.02 | 0.24 | 0.59 | 0.18 | 0.09 | 1.29 | 2.04 | 0.76 | 0.09
| Analytical results for Zr in OREAS 922b (Certified Value 130 ppm) | [22 empty] | Certified
Zr | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab | [6 empty] | 1
| No | No | 2.01: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.13; 2.12: 2.18; 2.13: 2.19; 2.14: 2.20 | [6 empty] | ppm
|  |  | 2.01: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.12: PF*MS; 2.13: BF*OES/MS; 2.14: BF*OES | [6 empty] | 0
0
| 1 | 1 | 2.01: 94.3; 2.04: 126; 2.05: 130.7969; 2.06: 147; 2.12: 133; 2.13: 133.1; 2.14: 122 | [6 empty] | 1
| 1 | 2 | 2.01: 89; 2.04: 134; 2.05: 128.8319; 2.06: 137; 2.12: 127; 2.13: 133.7; 2.14: 121 | [6 empty] | 41
| 1 | 3 | 2.01: 93.2; 2.04: 135; 2.05: 129.3212; 2.06: 129; 2.12: 129; 2.13: 136.1; 2.14: 121 | [6 empty] | 16
| 1 | 4 | 2.01: 91.7; 2.04: 135; 2.05: 131.7397; 2.06: 146; 2.12: 121; 2.13: 128.6; 2.14: 122 | [6 empty] | 130.4863
| 1 | 5 | 2.01: 96.4; 2.04: 126; 2.05: 133.5801; 2.06: 129; 2.12: 118; 2.13: 133.9; 2.14: 122 | [6 empty] | 205
| 1 | 6 | 2.01: 96.2; 2.04: 133; 2.05: 136.0365; 2.06: 146; 2.12: 131; 2.13: 129.8; 2.14: 122
| Mean |  | 2.01: 93.4667; 2.04: 131.5; 2.05: 131.7177; 2.06: 139; 2.12: 126.5; 2.13: 132.5333; 2.14: 121.6667
| Median |  | 2.01: 93.75; 2.04: 133.5; 2.05: 131.2683; 2.06: 141.5; 2.12: 128; 2.13: 133.4; 2.14: 122
| Std Dev. |  | 2.01: 2.8254; 2.04: 4.3243; 2.05: 2.7218; 2.06: 8.5557; 2.12: 5.8566; 2.13: 2.8005; 2.14: 0.5164
| Rel.Std.Dev. |  | 2.01: 3.023%; 2.04: 3.288%; 2.05: 2.066%; 2.06: 6.155%; 2.12: 4.63%; 2.13: 2.113%; 2.14: 0.4244%
| PDM3 |  | 2.01: -28.37%; 2.04: 0.7769%; 2.05: 0.9437%; 2.06: 6.525%; 2.12: -3.055%; 2.13: 1.569%; 2.14: -6.759%
| Z-Score (Absolute) |  | 2.01: 5.13; 2.04: 0; 2.05: 0.03; 2.06: 1.01; 2.12: 0.67; 2.13: 0.14; 2.14: 1.33
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C294: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C331: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1042: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 922b (Certified Value 0.766 wt.%) | [24 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.77 | 0.78 | 0.73 | 0.78 | 0.76853 | 0.75 | 0.86 | 0.78 | 0.8 | 0.76 | 0.75 | 0.766 | 0.84188 | 0.74 | 0.75 | 0.84 | 0.86 | 0.75 | 0.75 | 0.793 | 0.78 | 0.79 | 0.78 | 1
| 1 | 2 | 0.79 | 0.77 | 0.75 | 0.77 | 0.76033 | 0.75 | 0.86 | 0.77 | 0.79 | 0.76 | 0.76 | 0.767 | 0.84361 | 0.84 | 0.78 | 0.84 | 0.82 | 0.74 | 0.76 | 0.83 | 0.75 | 0.76 | 0.79 | 0
| 1 | 3 | 0.78 | 0.76 | 0.72 | 0.78 | 0.76834 | 0.77 | 0.8 | 0.77 | 0.79 | 0.77 | 0.78 | 0.76 | 0.84077 | 0.87 | 0.77 | 0.84 | 0.84 | 0.73 | 0.74 | 0.78 | 0.75 | 0.79 | 0.78 | 16
| 1 | 4 | 0.79 | 0.77 | 0.72 | 0.77 | 0.76139 | 0.74 | 0.8 | 0.77 | 0.79 | 0.76 | 0.75 | 0.764 | 0.82341 | 0.79 | 0.77 | 0.84 | 0.84 | 0.74 | 0.74 | 0.784 | 0.75 | 0.75 | 0.77 | 0.76584
| 1 | 5 | 0.78 | 0.77 | 0.75 | 0.77 | 0.76711 | 0.75 | 0.83 | 0.77 | 0.78 | 0.77 | 0.77 | 0.756 | 0.82458 | 0.79 | 0.78 | 0.84 | 0.84 | 0.75 | 0.74 | 0.777 | 0.76 | 0.79 | 0.78 | 207
| 1 | 6 | 0.77 | 0.76 | 0.72 | 0.78 | 0.7562 | 0.75 | 0.84 | 0.77 | 0.79 | 0.77 | 0.8 | 0.751 | 0.842 | 0.74 | 0.77 | 0.83 | 0.83 | 0.74 | 0.75 | 0.786 | 0.75 | 0.79 | 0.79
| Mean |  | 0.78 | 0.76833 | 0.73167 | 0.775 | 0.76365 | 0.75167 | 0.83167 | 0.77167 | 0.79 | 0.765 | 0.76833 | 0.76067 | 0.83604 | 0.795 | 0.77 | 0.83833 | 0.83833 | 0.74167 | 0.74667 | 0.79167 | 0.75667 | 0.77833 | 0.78167
| Median |  | 0.78 | 0.77 | 0.725 | 0.775 | 0.76425 | 0.75 | 0.835 | 0.77 | 0.79 | 0.765 | 0.765 | 0.762 | 0.84132 | 0.79 | 0.77 | 0.84 | 0.84 | 0.74 | 0.745 | 0.785 | 0.75 | 0.79 | 0.78
| Std Dev. |  | 0.0089443 | 0.0075277 | 0.01472 | 0.0054772 | 0.0050878 | 0.0098319 | 0.027142 | 0.0040825 | 0.0063246 | 0.0054772 | 0.019408 | 0.0062503 | 0.0093823 | 0.05244 | 0.010954 | 0.0040825 | 0.013292 | 0.0075277 | 0.008165 | 0.019562 | 0.012111 | 0.018348 | 0.0075277
| Rel.Std.Dev. |  | 1.147% | 0.9797% | 2.012% | 0.7067% | 0.6662% | 1.308% | 3.264% | 0.529% | 0.8006% | 0.716% | 2.526% | 0.8217% | 1.122% | 6.596% | 1.423% | 0.487% | 1.585% | 1.015% | 1.094% | 2.471% | 1.601% | 2.357% | 0.963%
| PDM3 |  | 1.849% | 0.3254% | -4.462% | 1.196% | -0.2862% | -1.851% | 8.595% | 0.7607% | 3.155% | -0.1098% | 0.3254% | -0.6757% | 9.166% | 3.807% | 0.543% | 9.466% | 9.466% | -3.157% | -2.504% | 3.372% | -1.198% | 1.631% | 2.066%
| Z-Score (Absolute) |  | 0.37 | 0.15 | 1.8 | 0.15 | 0.36 | 0.9 | 2.7 | 0 | 0.82 | 0.3 | 0.15 | 0.49 | 2.89 | 1.05 | 0.07 | 3 | 3 | 1.35 | 1.12 | 0.9 | 0.67 | 0.3 | 0.45
| Analytical results for S in OREAS 922b (Certified Value 0.264 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.28 | 0.28 | 0.31 | 0.29 | 0.26701 | 0.26 | 0.28 | 0.24 | 0.26 | 0.27 | 0.25 | 0.242 | 0.29 | 0.27 | 0.25 | 0.26 | 0.24 | 0.26 | 0.28 | 0.278 | 0.26 | 0.27 | 0.27 | 1
| 1 | 2 | 0.28 | 0.26 | 0.31 | 0.29 | 0.26658 | 0.26 | 0.27 | 0.25 | 0.26 | 0.27 | 0.26 | 0.245 | 0.28236 | 0.27 | 0.26 | 0.26 | 0.23 | 0.25 | 0.29 | 0.28 | 0.26 | 0.26 | 0.27 | 0
| 1 | 3 | 0.28 | 0.28 | 0.31 | 0.28 | 0.2684 | 0.26 | 0.27 | 0.24 | 0.26 | 0.26 | 0.26 | 0.245 | 0.27119 | 0.27 | 0.26 | 0.27 | 0.24 | 0.26 | 0.29 | 0.274 | 0.26 | 0.27 | 0.28 | 16
| 1 | 4 | 0.28 | 0.28 | 0.3 | 0.25 | 0.27145 | 0.25 | 0.26 | 0.24 | 0.26 | 0.26 | 0.25 | 0.246 | 0.2794 | 0.28 | 0.26 | 0.25 | 0.25 | 0.26 | 0.28 | 0.262 | 0.26 | 0.26 | 0.27 | 0.26444
| 1 | 5 | 0.28 | 0.29 | 0.31 | 0.27 | 0.25622 | 0.26 | 0.27 | 0.25 | 0.26 | 0.27 | 0.26 | 0.237 | 0.27537 | 0.28 | 0.26 | 0.24 | 0.23 | 0.26 | 0.28 | 0.274 | 0.26 | 0.27 | 0.27 | 208
| 1 | 6 | 0.28 | 0.3 | 0.31 | 0.26 | 0.26969 | 0.26 | 0.27 | 0.24 | 0.26 | 0.26 | 0.26 | 0.243 | 0.28338 | 0.27 | 0.26 | 0.26 | 0.23 | 0.26 | 0.28 | 0.257 | 0.26 | 0.28 | 0.27
| Mean |  | 0.28 | 0.28167 | 0.30833 | 0.27333 | 0.26656 | 0.25833 | 0.27 | 0.24333 | 0.26 | 0.265 | 0.25667 | 0.243 | 0.28028 | 0.27333 | 0.25833 | 0.25667 | 0.23667 | 0.25833 | 0.28333 | 0.27083 | 0.26 | 0.26833 | 0.27167
| Median |  | 0.28 | 0.28 | 0.31 | 0.275 | 0.26771 | 0.26 | 0.27 | 0.24 | 0.26 | 0.265 | 0.26 | 0.244 | 0.28088 | 0.27 | 0.26 | 0.26 | 0.235 | 0.26 | 0.28 | 0.274 | 0.26 | 0.27 | 0.27
| Std Dev. |  | 0 | 0.013292 | 0.0040825 | 0.01633 | 0.0053693 | 0.0040825 | 0.0063246 | 0.005164 | 0 | 0.0054772 | 0.005164 | 0.0032863 | 0.0065721 | 0.005164 | 0.0040825 | 0.010328 | 0.008165 | 0.0040825 | 0.005164 | 0.0092177 | 0 | 0.0075277 | 0.0040825
| Rel.Std.Dev. |  | 0% | 4.719% | 1.324% | 5.974% | 2.014% | 1.58% | 2.342% | 2.122% | 0% | 2.067% | 2.012% | 1.352% | 2.345% | 1.889% | 1.58% | 4.024% | 3.45% | 1.58% | 1.823% | 3.403% | 0% | 2.805% | 1.503%
| PDM3 |  | 5.883% | 6.513% | 16.6% | 3.362% | 0.8006% | -2.31% | 2.101% | -7.983% | -1.68% | 0.2107% | -2.941% | -8.109% | 5.99% | 3.362% | -2.31% | -2.941% | -10.5% | -2.31% | 7.143% | 2.417% | -1.68% | 1.471% | 2.732%
| Z-Score (Absolute) |  | 1.1 | 1.24 | 3.42 | 0.56 | 0 | 0.67 | 0.28 | 1.9 | 0.54 | 0.13 | 0.81 | 1.93 | 1.12 | 0.56 | 0.67 | 0.81 | 2.45 | 0.67 | 1.37 | 0.35 | 0.54 | 0.15 | 0.42
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 922b (Indicative Value 1.45 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.5 | 8
| Mean |  | 1.45 | 16
| Median |  | 1.45 | 1.45
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 922b (Indicative Value 64 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 64.4 | 1
| 1 | 2 | 64.4 | 9
| Mean |  | 64.4 | 16
| Median |  | 64.4 | 64.4
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 922b (Indicative Value 0.203 wt.%) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.201 | 1
| 1 | 2 | 0.205 | 27
| Mean |  | 0.203 | 16
| Median |  | 0.203 | 0.203
| Std Dev. |  | 0.0028284 | 33
| Rel.Std.Dev. |  | 1.393%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 922b (Indicative Value 2.4 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.4 | 28
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 922b (Indicative Value 13.9 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.7 | 1
| 1 | 2 | 14 | 29
| Mean |  | 13.85 | 16
| Median |  | 13.85 | 13.85
| Std Dev. |  | 0.21213 | 35
| Rel.Std.Dev. |  | 1.532%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 922b (Indicative Value < 0.1 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.1 | 1
| 1 | 2 | < 0.1 | 11
| Mean |  |  | 16
| Median |  |  | < 0.1
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 922b (Indicative Value 75 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75.2 | 1
| 1 | 2 | 75.7 | 12
| Mean |  | 75.45 | 16
| Median |  | 75.45 | 75.45
| Std Dev. |  | 0.35355 | 37
| Rel.Std.Dev. |  | 0.4686%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 922b (Indicative Value 4.2 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4.2 | 32
| Mean |  | 4.2 | 16
| Median |  | 4.2 | 4.2
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 922b (Indicative Value 85 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84 | 1
| 1 | 2 | 85 | 33
| Mean |  | 84.5 | 16
| Median |  | 84.5 | 84.5
| Std Dev. |  | 0.70711 | 39
| Rel.Std.Dev. |  | 0.8368%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 922b (Indicative Value 4.95 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.01 | 1
| 1 | 2 | 4.89 | 14
| Mean |  | 4.95 | 16
| Median |  | 4.95 | 4.95
| Std Dev. |  | 0.084853 | 40
| Rel.Std.Dev. |  | 1.714%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 922b (Indicative Value 0.236 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.236 | 1
| 1 | 2 | 0.236 | 35
| Mean |  | 0.236 | 16
| Median |  | 0.236 | 0.236
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 922b (Indicative Value 4.73 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.72 | 1
| 1 | 2 | 4.74 | 3
| Mean |  | 4.73 | 16
| Median |  | 4.73 | 4.73
| Std Dev. |  | 0.014142 | 42
| Rel.Std.Dev. |  | 0.299%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 922b (Indicative Value 2.88 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.81 | 1
| 1 | 2 | 2.94 | 4
| Mean |  | 2.875 | 16
| Median |  | 2.875 | 2.875
| Std Dev. |  | 0.091924 | 43
| Rel.Std.Dev. |  | 3.197%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 922b (Indicative Value 1.26 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.27 | 5
| Mean |  | 1.26 | 16
| Median |  | 1.26 | 1.26
| Std Dev. |  | 0.014142 | 44
| Rel.Std.Dev. |  | 1.122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 922b (Indicative Value 18.2 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.2 | 1
| 1 | 2 | 18.1 | 18
| Mean |  | 18.15 | 16
| Median |  | 18.15 | 18.15
| Std Dev. |  | 0.070711 | 45
| Rel.Std.Dev. |  | 0.3896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 922b (Indicative Value 5.63 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.65 | 1
| 1 | 2 | 5.6 | 6
| Mean |  | 5.625 | 16
| Median |  | 5.625 | 5.625
| Std Dev. |  | 0.035355 | 46
| Rel.Std.Dev. |  | 0.6285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 922b (Indicative Value 1.88 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.95 | 1
| 1 | 2 | 1.8 | 41
| Mean |  | 1.875 | 16
| Median |  | 1.875 | 1.875
| Std Dev. |  | 0.10607 | 47
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 922b (Indicative Value 3.78 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.71 | 1
| 1 | 2 | 3.84 | 25
| Mean |  | 3.775 | 16
| Median |  | 3.775 | 3.775
| Std Dev. |  | 0.091924 | 31
| Rel.Std.Dev. |  | 2.435%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 922b (Indicative Value 0.99 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.98 | 1
| 1 | 2 | 0.99 | 3
| Mean |  | 0.985 | 16
| Median |  | 0.985 | 0.985
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 0.7179%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 922b (Indicative Value 0.25 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.25 | 11
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 922b (Indicative Value 42.7 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 43 | 1
| 1 | 2 | 42.4 | 12
| Mean |  | 42.7 | 16
| Median |  | 42.7 | 42.7
| Std Dev. |  | 0.42426 | 34
| Rel.Std.Dev. |  | 0.9936%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 922b (Indicative Value 0.4 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 13
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 922b (Indicative Value 0.012 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0116 | 1
| 1 | 2 | 0.0116 | 30
| Mean |  | 0.0116 | 16
| Median |  | 0.0116 | 0.0116
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 922b (Indicative Value 4.8 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 4.8 | 31
| Mean |  | 4.8 | 16
| Median |  | 4.8 | 4.8
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 922b (Indicative Value 11.6 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.7 | 1
| 1 | 2 | 11.5 | 15
| Mean |  | 11.6 | 16
| Median |  | 11.6 | 11.6
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 1.219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 922b (Indicative Value 36.4 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.5 | 1
| 1 | 2 | 36.3 | 4
| Mean |  | 36.4 | 16
| Median |  | 36.4 | 36.4
| Std Dev. |  | 0.14142 | 39
| Rel.Std.Dev. |  | 0.3885%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 922b (Indicative Value 16 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14 | 1
| 1 | 2 | 18 | 22
| Mean |  | 16 | 16
| Median |  | 16 | 16
| Std Dev. |  | 2.8284 | 40
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 922b (Indicative Value 46.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 47 | 1
| 1 | 2 | 46 | 17
| Mean |  | 46.5 | 16
| Median |  | 46.5 | 46.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 922b (Indicative Value 9.89 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.84 | 1
| 1 | 2 | 9.94 | 6
| Mean |  | 9.89 | 16
| Median |  | 9.89 | 9.89
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 0.715%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 922b (Indicative Value 158 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 156 | 1
| 1 | 2 | 159 | 37
| Mean |  | 157.5 | 16
| Median |  | 157.5 | 157.5
| Std Dev. |  | 2.1213 | 43
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 922b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 23
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 922b (Indicative Value 4.85 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.5 | 1
| 1 | 2 | 5.2 | 39
| Mean |  | 4.85 | 16
| Median |  | 4.85 | 4.85
| Std Dev. |  | 0.49497 | 45
| Rel.Std.Dev. |  | 10.21%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 922b (Indicative Value 12.1 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.4 | 1
| 1 | 2 | 11.8 | 19
| Mean |  | 12.1 | 16
| Median |  | 12.1 | 12.1
| Std Dev. |  | 0.42426 | 46
| Rel.Std.Dev. |  | 3.506%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 922b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 922b (Indicative Value 6.98 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.93 | 1
| 1 | 2 | 7.03 | 4
| Mean |  | 6.98 | 16
| Median |  | 6.98 | 6.98
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 1.013%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 922b (Indicative Value 6.6 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 6.6 | 26
| Mean |  | 6.6 | 16
| Median |  | 6.6 | 6.6
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 922b (Indicative Value 43.5 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 43.6 | 1
| 1 | 2 | 43.4 | 27
| Mean |  | 43.5 | 16
| Median |  | 43.5 | 43.5
| Std Dev. |  | 0.14142 | 33
| Rel.Std.Dev. |  | 0.3251%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 922b (Indicative Value 0.95 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.96 | 1
| 1 | 2 | 0.94 | 10
| Mean |  | 0.95 | 16
| Median |  | 0.95 | 0.95
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 1.489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 922b (Indicative Value 0.81 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 0.81 | 11
| Mean |  | 0.805 | 16
| Median |  | 0.805 | 0.805
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.8784%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 922b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 12
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 922b (Indicative Value 14.1 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.1 | 1
| 1 | 2 | 14 | 13
| Mean |  | 14.05 | 16
| Median |  | 14.05 | 14.05
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.5033%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 922b (Indicative Value 0.323 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.322 | 1
| 1 | 2 | 0.324 | 32
| Mean |  | 0.323 | 16
| Median |  | 0.323 | 0.323
| Std Dev. |  | 0.0014142 | 38
| Rel.Std.Dev. |  | 0.4378%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 922b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 14
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 922b (Indicative Value 0.4 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.39 | 1
| 1 | 2 | 0.4 | 5
| Mean |  | 0.395 | 16
| Median |  | 0.395 | 0.395
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 922b (Indicative Value 4.53 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.51 | 1
| 1 | 2 | 4.54 | 16
| Mean |  | 4.525 | 16
| Median |  | 4.525 | 4.525
| Std Dev. |  | 0.021213 | 41
| Rel.Std.Dev. |  | 0.4688%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 922b (Indicative Value 197 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 197 | 1
| 1 | 2 | 196 | 36
| Mean |  | 196.5 | 16
| Median |  | 196.5 | 196.5
| Std Dev. |  | 0.70711 | 42
| Rel.Std.Dev. |  | 0.3599%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 922b (Indicative Value 3.5 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3.5 | 17
| Mean |  | 3.5 | 16
| Median |  | 3.5 | 3.5
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 922b (Indicative Value 26.9 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.5 | 1
| 1 | 2 | 27.2 | 18
| Mean |  | 26.85 | 16
| Median |  | 26.85 | 26.85
| Std Dev. |  | 0.49497 | 44
| Rel.Std.Dev. |  | 1.843%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 922b (Indicative Value 2.8 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.82 | 1
| 1 | 2 | 2.78 | 19
| Mean |  | 2.8 | 16
| Median |  | 2.8 | 2.8
| Std Dev. |  | 0.028284 | 45
| Rel.Std.Dev. |  | 1.01%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 922b (Indicative Value 70 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 70 | 40
| Mean |  | 70 | 16
| Median |  | 70 | 70
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 922b (Indicative Value 131 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 129 | 1
| 1 | 2 | 133 | 41
| Mean |  | 131 | 16
| Median |  | 131 | 131
| Std Dev. |  | 2.8284 | 47
| Rel.Std.Dev. |  | 2.159%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.