## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 163b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 4.2584 | 4.1389 | 4.378 | 4.1592 | 4.3577
Al, Aluminium (wt.%) | 0.37241 | 0.35896 | 0.38586 | 0.36237 | 0.38245
As, Arsenic (ppm) | 167.8738 | 161.2318 | 174.5158 | 163.6923 | 172.0553
Ba, Barium (ppm) | 38.5096 | 36.0295 | 40.9897 | 36.9723 | 40.047
Be, Beryllium (ppm) | 0.28554 | 0.24712 | 0.32395 | 0.26155 | 0.30952
Bi, Bismuth (ppm) | 17.2788 | 16.5001 | 18.0576 | 16.8051 | 17.7526
Ca, Calcium (wt.%) | 2.6413 | 2.5523 | 2.7302 | 2.5838 | 2.6987
Cd, Cadmium (ppm) | 0.45372 | 0.41981 | 0.48763 | 0.41579 | 0.49164
Ce, Cerium (ppm) | 9.9871 | 9.4588 | 10.5154 | 9.572 | 10.4022
Co, Cobalt (ppm) | 236.518 | 228.2705 | 244.7654 | 231.5013 | 241.5346
Cr, Chromium (ppm) | 12.8579 | 11.2671 | 14.4488 | 11.8905 | 13.8254
Cs, Caesium (ppm) | 0.96904 | 0.89913 | 1.039 | 0.92486 | 1.0132
Cu, Copper (wt.%) | 1.6834 | 1.6491 | 1.7176 | 1.6659 | 1.7009
Dy, Dysprosium (ppm) | 0.85778 | 0.74263 | 0.97292 | 0.81298 | 0.90257
Er, Erbium (ppm) | 0.3766 | 0.32814 | 0.42506 | 0.32817 | 0.42504
Eu, Europium (ppm) | 0.15976 | 0.13312 | 0.1864 | IND | IND
Fe, Iron (wt.%) | 4.0998 | 3.9796 | 4.22 | 4.0392 | 4.1604
Ga, Gallium (ppm) | 1.4653 | 1.3538 | 1.5767 | 1.3519 | 1.5786
Gd, Gadolinium (ppm) | 1.2306 | 1.1345 | 1.3267 | 1.1922 | 1.269
Ge, Germanium (ppm) | 0.067381 | 0.037787 | 0.096975 | IND | IND
Hf, Hafnium (ppm) | 0.36787 | 0.32489 | 0.41085 | 0.34507 | 0.39067
Ho, Holmium (ppm) | 0.13305 | 0.10202 | 0.16408 | IND | IND
In, Indium (ppm) | 1.0624 | 1.0139 | 1.111 | 1.0319 | 1.093
K, Potassium (wt.%) | 0.0897 | 0.088585 | 0.090815 | 0.087054 | 0.092347
La, Lanthanum (ppm) | 4.7353 | 4.4912 | 4.9794 | 4.5397 | 4.931
Li, Lithium (ppm) | 100.3113 | 95.8164 | 104.8062 | 97.3749 | 103.2477
Lu, Lutetium (ppm) | < 0.1 | IND | IND | IND | IND
Mg, Magnesium (wt.%) | 3.5119 | 3.3805 | 3.6434 | 3.458 | 3.5658
Mn, Manganese (wt.%) | 0.047956 | 0.046516 | 0.049395 | 0.047095 | 0.048816
Mo, Molybdenum (ppm) | 2.0963 | 1.9547 | 2.2379 | 1.9749 | 2.2176
Na, Sodium (wt.%) | 0.016527 | 0.014986 | 0.018067 | 0.016006 | 0.017047
Nb, Niobium (ppm) | 1.4499 | 1.2183 | 1.6815 | 1.2149 | 1.6849
Nd, Neodymium (ppm) | 5.181 | 4.7433 | 5.6187 | 4.8398 | 5.5221
Ni, Nickel (ppm) | 24.349 | 23.0063 | 25.6918 | 23.2854 | 25.4126
P, Phosphorus (wt.%) | 0.017672 | 0.016924 | 0.018419 | 0.017056 | 0.018288
Pb, Lead (ppm) | 55.685 | 53.5463 | 57.8237 | 53.668 | 57.702
Pr, Praseodymium (ppm) | 1.3111 | 1.2091 | 1.4131 | 1.2093 | 1.4129
Rb, Rubidium (ppm) | 4.6992 | 4.4375 | 4.9609 | 4.4314 | 4.9671
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 3.3486 | 3.2381 | 3.4591 | 3.2835 | 3.4136
Sb, Antimony (ppm) | 8.464 | 8.0624 | 8.8656 | 8.102 | 8.826
Sc, Scandium (ppm) | 1.0166 | 0.93405 | 1.0992 | IND | IND
Se, Selenium (ppm) | 1.6879 | 1.4399 | 1.9359 | IND | IND
Sm, Samarium (ppm) | 1.2851 | 1.1659 | 1.4043 | 1.2249 | 1.3454
Sn, Tin (ppm) | 2.41 | 2.2298 | 2.5901 | 2.2684 | 2.5515
Sr, Strontium (ppm) | 9.3935 | 9.0009 | 9.7862 | 9.0233 | 9.7638
Ta, Tantalum (ppm) | 2.0539 | 1.5635 | 2.5443 | 1.8293 | 2.2785
Tb, Terbium (ppm) | 0.16284 | 0.13721 | 0.18847 | IND | IND
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 0.9187 | 0.86191 | 0.97549 | 0.86712 | 0.97028
Ti, Titanium (wt.%) | 0.020513 | 0.019434 | 0.021591 | 0.019739 | 0.021287
Tl, Thallium (ppm) | 1.5948 | 1.5269 | 1.6628 | 1.5428 | 1.6469
Tm, Thulium (ppm) | < 0.1 | IND | IND | IND | IND
U, Uranium (ppm) | 0.69952 | 0.6448 | 0.75425 | 0.67299 | 0.72606
V, Vanadium (ppm) | 10.6751 | 9.912 | 11.4382 | 9.8229 | 11.5273
W, Tungsten (ppm) | 2.0495 | 1.8731 | 2.2258 | 1.9183 | 2.1807
Y, Yttrium (ppm) | 4.671 | 4.3493 | 4.9927 | 4.4381 | 4.9039
Yb, Ytterbium (ppm) | 0.29084 | 0.26508 | 0.3166 | 0.24079 | 0.3409
Zn, Zinc (ppm) | 26.5764 | 24.9955 | 28.1572 | 25.1406 | 28.0122
Zr, Zirconium (ppm) | 12.0591 | 10.986 | 13.1323 | 11.2956 | 12.8227
Aqua Regia Digestion
Ag, Silver (ppm) | 4.2003 | 4.0177 | 4.3829 | 4.1017 | 4.2989
Al, Aluminium (wt.%) | 0.12605 | 0.11846 | 0.13364 | 0.12072 | 0.13138
As, Arsenic (ppm) | 164.6445 | 160.8394 | 168.4497 | 160.7033 | 168.5858
Au, Gold (ppm) | < 0.02 | IND | IND | IND | IND
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 26.4313 | 24.1938 | 28.6687 | 25.1551 | 27.7074
Be, Beryllium (ppm) | 0.10972 | 0.099964 | 0.11948 | IND | IND
Bi, Bismuth (ppm) | 17.0565 | 16.387 | 17.726 | 16.631 | 17.482
Ca, Calcium (wt.%) | 2.5317 | 2.4441 | 2.6194 | 2.4886 | 2.5749
Cd, Cadmium (ppm) | 0.44055 | 0.41726 | 0.46385 | 0.41237 | 0.46874
Ce, Cerium (ppm) | 4.4804 | 4.1074 | 4.8533 | 4.267 | 4.6937
Co, Cobalt (ppm) | 226.7943 | 220.5644 | 233.0242 | 222.1181 | 231.4705
Cr, Chromium (ppm) | 10.588 | 9.5712 | 11.6047 | 9.8652 | 11.3107
Cs, Caesium (ppm) | 0.7837 | 0.74888 | 0.81853 | 0.75937 | 0.80804
Cu, Copper (wt.%) | 1.6779 | 1.6566 | 1.6992 | 1.6606 | 1.6952
Fe, Iron (wt.%) | 3.9546 | 3.854 | 4.0551 | 3.8794 | 4.0297
Ga, Gallium (ppm) | 0.62648 | 0.58382 | 0.66913 | 0.59328 | 0.65967
Ge, Germanium (ppm) | 0.040952 | 0.030085 | 0.051819 | IND | IND
Hf, Hafnium (ppm) | 0.04899 | 0.038641 | 0.059339 | IND | IND
Hg, Mercury (ppm) | 0.079382 | 0.067075 | 0.091689 | 0.073282 | 0.085482
In, Indium (ppm) | 1.0208 | 0.9853 | 1.0562 | 0.99147 | 1.0501
K, Potassium (wt.%) | 0.014803 | 0.012671 | 0.016936 | 0.014373 | 0.015234
La, Lanthanum (ppm) | 1.826 | 1.6708 | 1.9811 | 1.7049 | 1.9471
Li, Lithium (ppm) | 4.6573 | 4.3577 | 4.957 | 4.44 | 4.8747
Mg, Magnesium (wt.%) | 1.4078 | 1.3611 | 1.4546 | 1.3773 | 1.4383
Mn, Manganese (wt.%) | 0.046404 | 0.045111 | 0.047698 | 0.045566 | 0.047243
Mo, Molybdenum (ppm) | 2.0909 | 1.9587 | 2.2232 | 1.9922 | 2.1897
Na, Sodium (wt.%) | 0.0097305 | 0.008506 | 0.010955 | IND | IND
Nb, Niobium (ppm) | < 0.05 | IND | IND | IND | IND
Ni, Nickel (ppm) | 23.2926 | 22.2121 | 24.3731 | 22.5076 | 24.0776
P, Phosphorus (wt.%) | 0.016458 | 0.015532 | 0.017384 | 0.015978 | 0.016939
Pb, Lead (ppm) | 54.0453 | 51.9397 | 56.1509 | 52.2166 | 55.874
Rb, Rubidium (ppm) | 1.2458 | 1.1573 | 1.3344 | 1.1365 | 1.3552
Re, Rhenium (ppm) | 0.001 | 0.001 | 0.001 | IND | IND
S, Sulphur (wt.%) | 3.3028 | 3.2087 | 3.3969 | 3.2481 | 3.3575
Sb, Antimony (ppm) | 7.2982 | 6.7664 | 7.8301 | 7.0489 | 7.5475
Sc, Scandium (ppm) | 0.79583 | 0.70663 | 0.88502 | IND | IND
Se, Selenium (ppm) | 1.6814 | 1.4361 | 1.9267 | IND | IND
Sn, Tin (ppm) | 1.6332 | 1.547 | 1.7195 | 1.5518 | 1.7146
Sr, Strontium (ppm) | 7.0009 | 6.5996 | 7.4021 | 6.7438 | 7.2579
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.12829 | 0.10822 | 0.14837 | IND | IND
Te, Tellurium (ppm) | < 0.05 | IND | IND | IND | IND
Th, Thorium (ppm) | 0.47351 | 0.40733 | 0.5397 | 0.4456 | 0.50142
Ti, Titanium (wt.%) | < 0.005 | IND | IND | IND | IND
Tl, Thallium (ppm) | 1.4849 | 1.4017 | 1.568 | 1.4426 | 1.5271
U, Uranium (ppm) | 0.43846 | 0.40711 | 0.4698 | 0.41696 | 0.45996
V, Vanadium (ppm) | 5.0128 | 4.4566 | 5.569 | 4.5982 | 5.4275
W, Tungsten (ppm) | 1.5715 | 1.4292 | 1.7138 | 1.4403 | 1.7027
Y, Yttrium (ppm) | 2.9072 | 2.7919 | 3.0226 | 2.8078 | 3.0067
Yb, Ytterbium (ppm) | 0.16426 | 0.13055 | 0.19797 | IND | IND
Zn, Zinc (ppm) | 24.5001 | 23.6122 | 25.388 | 23.3941 | 25.6062
Zr, Zirconium (ppm) | 1.5564 | 1.3735 | 1.7393 | 1.4927 | 1.62
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 0.69869 | 0.66932 | 0.72806 | 0.67457 | 0.72282
BaO, Barium oxide (ppm) | 63.997 | 31.358 | 96.636 | IND | IND
Bi, Bismuth (ppm) | < 100 | IND | IND | IND | IND
CaO, Calcium oxide (wt.%) | 3.6928 | 3.6359 | 3.7497 | 3.6369 | 3.7487
Ce, Cerium (ppm) | < 80 | IND | IND | IND | IND
Co, Cobalt (ppm) | 236.3398 | 200.9898 | 271.6897 | IND | IND
Cr2O3, Chromium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Cu, Copper (wt.%) | 1.6751 | 1.6385 | 1.7116 | 1.6497 | 1.7004
Fe2O3, Iron(III) oxide (wt.%) | 5.8129 | 5.7444 | 5.8815 | 5.7684 | 5.8575
Hf, Hafnium (ppm) | < 80 | IND | IND | IND | IND
K2O, Potassium oxide (wt.%) | 0.10549 | 0.096145 | 0.11483 | IND | IND
MgO, Magnesium oxide (wt.%) | 5.7472 | 5.6331 | 5.8613 | 5.6731 | 5.8212
MnO, Manganese oxide (wt.%) | 0.061585 | 0.051906 | 0.071264 | IND | IND
Mo, Molybdenum (ppm) | < 50 | IND | IND | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.038248 | 0.03451 | 0.041985 | IND | IND
Rb, Rubidium (ppm) | < 50 | IND | IND | IND | IND
Sb, Antimony (ppm) | < 50 | IND | IND | IND | IND
SiO2, Silicon dioxide (wt.%) | 73.6264 | 72.9929 | 74.2598 | 73.2739 | 73.9789
Sn, Tin (ppm) | < 50 | IND | IND | IND | IND
SO3, Sulphur trioxide (wt.%) | 8.3861 | 8.1969 | 8.5754 | 8.3 | 8.4723
SrO, Strontium oxide (ppm) | < 120 | IND | IND | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.056392 | 0.043512 | 0.069273 | IND | IND
V, Vanadium (ppm) | < 100 | IND | IND | IND | IND
Zn, Zinc (ppm) | < 50 | IND | IND | IND | IND
Zr, Zirconium (ppm) | < 100 | IND | IND | IND | IND
Thermogravimetry
LOI1000, Loss On Ignition @1000 °C (wt.%) | 4.9939 | 4.3936 | 5.5942 | 4.8122 | 5.1756
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 0.37944 | 0.36667 | 0.39222 | 0.36979 | 0.3891
As, Arsenic (ppm) | 164.6967 | 154.3302 | 175.0633 | 161.2542 | 168.1393
Ba, Barium (ppm) | 37.7402 | 35.2489 | 40.2315 | 35.8105 | 39.6699
Be, Beryllium (ppm) | < 1 | IND | IND | IND | IND
Bi, Bismuth (ppm) | 17.5957 | 16.8392 | 18.3522 | 17.001 | 18.1904
Ca, Calcium (wt.%) | 2.6103 | 2.5087 | 2.712 | 2.5659 | 2.6548
Cd, Cadmium (ppm) | < 10 | IND | IND | IND | IND
Ce, Cerium (ppm) | 10.4998 | 9.8958 | 11.1038 | 10.1266 | 10.8731
Co, Cobalt (ppm) | 239.2711 | 229.9503 | 248.5919 | 233.3319 | 245.2102
Cr2O3, Chromium(III) oxide (ppm) | < 100 | IND | IND | IND | IND
Cs, Caesium (ppm) | 0.98463 | 0.87952 | 1.0898 | IND | IND
Cu, Copper (wt.%) | 1.6602 | 1.6086 | 1.7117 | 1.6382 | 1.6821
Dy, Dysprosium (ppm) | 1.3233 | 1.1838 | 1.4628 | 1.2136 | 1.4329
Er, Erbium (ppm) | 0.58143 | 0.51801 | 0.64485 | 0.54954 | 0.61332
Eu, Europium (ppm) | 0.1831 | 0.1597 | 0.20649 | IND | IND
Fe, Iron (wt.%) | 4.0565 | 3.9766 | 4.1363 | 3.9933 | 4.1196
Ga, Gallium (ppm) | 1.5385 | 0.52242 | 2.5547 | IND | IND
Gd, Gadolinium (ppm) | 1.4397 | 1.2497 | 1.6297 | 1.3559 | 1.5235
Ge, Germanium (ppm) | 1.0349 | 0.89286 | 1.177 | IND | IND
Ho, Holmium (ppm) | 0.23194 | 0.20899 | 0.2549 | 0.21458 | 0.24931
In, Indium (ppm) | 1.1293 | 0.96949 | 1.2891 | IND | IND
K, Potassium (wt.%) | 0.10762 | 0.086862 | 0.12838 | 0.10178 | 0.11346
La, Lanthanum (ppm) | 5.2607 | 4.8464 | 5.675 | 5.0855 | 5.436
Li, Lithium (ppm) | 105.6452 | 99.8244 | 111.466 | 101.8569 | 109.4335
Mg, Magnesium (wt.%) | 3.5042 | 3.4236 | 3.5849 | 3.4568 | 3.5516
Mn, Manganese (wt.%) | 0.047588 | 0.046501 | 0.048676 | 0.047164 | 0.048013
Mo, Molybdenum (ppm) | 2.2642 | 1.4371 | 3.0912 | IND | IND
Nb, Niobium (ppm) | 1.9153 | 1.5994 | 2.2312 | IND | IND
Nd, Neodymium (ppm) | 5.4963 | 5.1801 | 5.8125 | 5.1517 | 5.8409
Ni, Nickel (ppm) | 29.3903 | 21.6379 | 37.1427 | 26.5458 | 32.2348
P, Phosphorus (wt.%) | 0.019507 | 0.017725 | 0.021289 | IND | IND
Pb, Lead (ppm) | 55.8834 | 48.3488 | 63.418 | 53.3092 | 58.4576
Pr, Praseodymium (ppm) | 1.3689 | 1.263 | 1.4748 | 1.2806 | 1.4571
Rb, Rubidium (ppm) | 4.8892 | 4.3319 | 5.4465 | 4.4697 | 5.3087
S, Sulphur (wt.%) | 3.3566 | 3.265 | 3.4483 | 3.2941 | 3.4192
Sb, Antimony (ppm) | 8.2988 | 6.6432 | 9.9544 | 7.324 | 9.2736
Sc, Scandium (ppm) | < 5 | IND | IND | IND | IND
Si, Silicon (wt.%) | 34.7556 | 33.8076 | 35.7037 | 34.2756 | 35.2357
Sm, Samarium (ppm) | 1.4821 | 1.3513 | 1.6129 | 1.3636 | 1.6005
Sn, Tin (ppm) | 3.0359 | 2.4349 | 3.637 | IND | IND
Sr, Strontium (ppm) | 9.6895 | 8.2694 | 11.1096 | 9.1082 | 10.2707
Ta, Tantalum (ppm) | 2.2975 | 1.6363 | 2.9586 | 1.9745 | 2.6204
Tb, Terbium (ppm) | 0.24288 | 0.21821 | 0.26754 | 0.22461 | 0.26115
Th, Thorium (ppm) | 0.90608 | 0.76809 | 1.0441 | IND | IND
Ti, Titanium (wt.%) | 0.029636 | 0.02872 | 0.030552 | 0.028222 | 0.03105
Tl, Thallium (ppm) | 1.6747 | 1.5569 | 1.7925 | IND | IND
Tm, Thulium (ppm) | 0.074333 | 0.059052 | 0.089615 | IND | IND
U, Uranium (ppm) | 0.72855 | 0.64138 | 0.81572 | 0.66601 | 0.79109
V, Vanadium (ppm) | 11.7208 | 9.1698 | 14.2718 | IND | IND
W, Tungsten (ppm) | 2.5489 | 1.517 | 3.5808 | IND | IND
Y, Yttrium (ppm) | 7.0013 | 6.5635 | 7.4392 | 6.6398 | 7.3629
Yb, Ytterbium (ppm) | 0.42482 | 0.32948 | 0.52016 | 0.38504 | 0.4646
Zn, Zinc (ppm) | 31.3796 | 26.1213 | 36.638 | 29.4431 | 33.3161
Zr, Zirconium (ppm) | 21.6982 | 19.2612 | 24.1351 | 19.5788 | 23.8176
Infrared Combustion
C, Carbon (wt.%) | 1.6441 | 1.6147 | 1.6734 | 1.6266 | 1.6615
S, Sulphur (wt.%) | 3.3731 | 3.3103 | 3.4359 | 3.3238 | 3.4223
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 163b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 19 | Hg | ppm | 0.078333
Aqua Regia Digestion
Dy | ppm | 0.59357 | Ho | ppm | 0.071951 | Pr | ppm | 0.63767
Er | ppm | 0.2082 | Lu | ppm | 0.02 | Pt | ppb | < 5
Eu | ppm | 0.10123 | Nd | ppm | 2.6657 | Sm | ppm | 0.78311
Gd | ppm | 0.8643 | Pd | ppb | < 10 | Tm | ppm | < 0.05
Borate Fusion XRF
Ag | ppm | < 10 | Hg | ppm | < 100 | Se | ppm | < 10
As | ppm | 186 | Ho | ppm | < 10 | Sm | ppm | 25.5
Au | ppm | < 10 | In | ppm | < 100 | Ta | ppm | 63.5
Cd | ppm | < 10 | La | ppm | < 90 | Tb | ppm | < 10
Cl | ppm | 316.6667 | Lu | ppm | < 10 | Te | ppm | < 100
Cs | ppm | < 10 | Na2O | wt.% | 0.073292 | Th | ppm | < 10
Dy | ppm | 20.6667 | Nb | ppm | 44.6667 | Tl | ppm | < 10
Er | ppm | < 10 | Nd | ppm | 41 | Tm | ppm | < 10
Eu | ppm | < 10 | Ni | ppm | < 100 | U | ppm | < 10
Ga | ppm | < 10 | Pb | ppm | 62.7778 | W | ppm | 23.3333
Gd | ppm | < 10 | Pr | ppm | < 10 | Y | ppm | < 39
Ge | ppm | < 10 | Sc | ppm | < 10 | Yb | ppm | < 10
Borate / Peroxide Fusion ICP
Ag | ppm | 4.3767 | Hg | ppm | < 5 | Re | ppm | < 0.1
B | ppm | < 30 | Lu | ppm | 0.058222 | Se | ppm | < 10
Hf | ppm | 0.60327 | Na | wt.% | 0.022168 | Te | ppm | < 1
Laser Ablation ICP-MS
Ag | ppm | 3.65 | Hf | ppm | 0.645 | Sm | ppm | 1.44
As | ppm | 166.5 | Ho | ppm | 0.225 | Sn | ppm | 3
Ba | ppm | 36 | In | ppm | 0.85 | Sr | ppm | 8.75
Be | ppm | 0.4 | La | ppm | 4.76 | Ta | ppm | 2.33
Bi | ppm | 16 | Lu | ppm | 0.055 | Tb | ppm | 0.225
Cd | ppm | 0.35 | Mn | wt.% | 0.04515 | Te | ppm | < 0.2
Ce | ppm | 9.58 | Mo | ppm | 1.5 | Th | ppm | 0.915
Co | ppm | 227 | Nb | ppm | 1.755 | Ti | wt.% | 0.02935
Cr | ppm | 10 | Nd | ppm | 5.19 | Tl | ppm | 0.4
Cs | ppm | 0.895 | Ni | ppm | 26 | Tm | ppm | 0.075
Cu | wt.% | 1.52 | Pb | ppm | 47.5 | U | ppm | 0.685
Dy | ppm | 1.235 | Pr | ppm | 1.275 | V | ppm | 10.65
Er | ppm | 0.555 | Rb | ppm | 4.325 | W | ppm | 2
Eu | ppm | 0.16 | Re | ppm | < 0.01 | Y | ppm | 6.76
Ga | ppm | 1.35 | Sb | ppm | 9.7 | Yb | ppm | 0.44
Gd | ppm | 1.56 | Sc | ppm | 1.6 | Zn | ppm | 20
Ge | ppm | 0.775 | Se | ppm | < 5 | Zr | ppm | 21.5
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 163b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 4.2584 | 0.18264 | 3.8932 | 4.6237 | 3.7105 | 4.8063 | 4.289% | 8.578% | 12.87% | 4.0455 | 4.4713
Al, wt.% | 0.37241 | 0.014448 | 0.34351 | 0.40131 | 0.32907 | 0.41576 | 3.88% | 7.759% | 11.64% | 0.35379 | 0.39103
As, ppm | 167.8738 | 7.265 | 153.3439 | 182.4037 | 146.0789 | 189.6687 | 4.328% | 8.655% | 12.98% | 159.4801 | 176.2675
Ba, ppm | 38.5096 | 2.4317 | 33.6462 | 43.3731 | 31.2145 | 45.8048 | 6.315% | 12.63% | 18.94% | 36.5842 | 40.4351
Be, ppm | 0.28554 | 0.027095 | 0.23135 | 0.33973 | 0.20425 | 0.36682 | 9.489% | 18.98% | 28.47% | 0.27126 | 0.29981
Bi, ppm | 17.2788 | 0.86356 | 15.5517 | 19.0059 | 14.6881 | 19.8695 | 4.998% | 9.996% | 14.99% | 16.4149 | 18.1428
Ca, wt.% | 2.6413 | 0.062998 | 2.5153 | 2.7673 | 2.4523 | 2.8303 | 2.385% | 4.77% | 7.155% | 2.5092 | 2.7733
Cd, ppm | 0.45372 | 0.032764 | 0.38819 | 0.51925 | 0.35543 | 0.55201 | 7.221% | 14.44% | 21.66% | 0.43103 | 0.4764
Ce, ppm | 9.9871 | 0.42694 | 9.1332 | 10.841 | 8.7063 | 11.2679 | 4.275% | 8.55% | 12.82% | 9.4877 | 10.4865
Co, ppm | 236.518 | 7.7522 | 221.0135 | 252.0224 | 213.2613 | 259.7746 | 3.278% | 6.555% | 9.833% | 224.6921 | 248.3439
Cr, ppm | 12.8579 | 1.8222 | 9.2136 | 16.5023 | 7.3914 | 18.3244 | 14.17% | 28.34% | 42.51% | 12.215 | 13.5008
Cs, ppm | 0.96904 | 0.048513 | 0.87202 | 1.0661 | 0.82351 | 1.1146 | 5.006% | 10.01% | 15.02% | 0.92059 | 1.0175
Cu, wt.% | 1.6834 | 0.034827 | 1.6137 | 1.753 | 1.5789 | 1.7878 | 2.069% | 4.138% | 6.207% | 1.5992 | 1.7675
Dy, ppm | 0.85778 | 0.075889 | 0.706 | 1.0096 | 0.63011 | 1.0854 | 8.847% | 17.69% | 26.54% | 0.81489 | 0.90067
Er, ppm | 0.3766 | 0.030777 | 0.31505 | 0.43816 | 0.28427 | 0.46893 | 8.172% | 16.34% | 24.52% | 0.35777 | 0.39543
Eu, ppm | 0.15976 | 0.011894 | 0.13598 | 0.18355 | 0.12408 | 0.19545 | 7.445% | 14.89% | 22.33% | 0.15178 | 0.16775
Fe, wt.% | 4.0998 | 0.14439 | 3.811 | 4.3886 | 3.6666 | 4.533 | 3.522% | 7.044% | 10.57% | 3.8948 | 4.3048
Ga, ppm | 1.4653 | 0.15607 | 1.1531 | 1.7774 | 0.99708 | 1.9335 | 10.65% | 21.3% | 31.95% | 1.392 | 1.5385
Gd, ppm | 1.2306 | 0.039112 | 1.1524 | 1.3088 | 1.1133 | 1.348 | 3.178% | 6.356% | 9.535% | 1.1691 | 1.2922
Ge, ppm | 0.067381 | 0.023094 | 0.021193 | 0.11357 | 0 | 0.13666 | 34.27% | 68.55% | 102.8% | 0.064012 | 0.07075
Hf, ppm | 0.36787 | 0.039867 | 0.28814 | 0.44761 | 0.24827 | 0.48747 | 10.84% | 21.67% | 32.51% | 0.34948 | 0.38626
Ho, ppm | 0.13305 | 0.019243 | 0.094563 | 0.17154 | 0.07532 | 0.19078 | 14.46% | 28.93% | 43.39% | 0.1264 | 0.1397
In, ppm | 1.0624 | 0.052461 | 0.9575 | 1.1673 | 0.90504 | 1.2198 | 4.938% | 9.876% | 14.81% | 1.0093 | 1.1155
K, wt.% | 0.0897 | 0.0034223 | 0.082856 | 0.096545 | 0.079433 | 0.099967 | 3.815% | 7.631% | 11.45% | 0.085215 | 0.094185
La, ppm | 4.7353 | 0.27562 | 4.1841 | 5.2866 | 3.9085 | 5.5622 | 5.821% | 11.64% | 17.46% | 4.4985 | 4.9721
Li, ppm | 100.3113 | 5.1059 | 90.0995 | 110.5232 | 84.9935 | 115.6291 | 5.09% | 10.18% | 15.27% | 95.2958 | 105.3269
Lu, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Mg, wt.% | 3.5119 | 0.11426 | 3.2834 | 3.7404 | 3.1691 | 3.8547 | 3.254% | 6.507% | 9.761% | 3.3363 | 3.6875
Mn, wt.% | 0.047956 | 0.0019292 | 0.044097 | 0.051814 | 0.042168 | 0.053743 | 4.023% | 8.046% | 12.07% | 0.045558 | 0.050353
Mo, ppm | 2.0963 | 0.12412 | 1.848 | 2.3445 | 1.7239 | 2.4686 | 5.921% | 11.84% | 17.76% | 1.9914 | 2.2011
Na, wt.% | 0.016527 | 0.0035779 | 0.0093708 | 0.023682 | 0.005793 | 0.02726 | 21.65% | 43.3% | 64.95% | 0.0157 | 0.017353
Nb, ppm | 1.4499 | 0.22998 | 0.98996 | 1.9099 | 0.75999 | 2.1398 | 15.86% | 31.72% | 47.58% | 1.3774 | 1.5224
Nd, ppm | 5.181 | 0.37769 | 4.4256 | 5.9363 | 4.0479 | 6.314 | 7.29% | 14.58% | 21.87% | 4.9219 | 5.44
Ni, ppm | 24.349 | 1.3488 | 21.6515 | 27.0466 | 20.3027 | 28.3954 | 5.539% | 11.08% | 16.62% | 23.1316 | 25.5665
P, wt.% | 0.017672 | 0.00077284 | 0.016126 | 0.019217 | 0.015353 | 0.01999 | 4.373% | 8.747% | 13.12% | 0.016788 | 0.018555
Pb, ppm | 55.685 | 2.5914 | 50.5022 | 60.8679 | 47.9108 | 63.4593 | 4.654% | 9.307% | 13.96% | 52.9008 | 58.4693
Pr, ppm | 1.3111 | 0.064142 | 1.1828 | 1.4394 | 1.1187 | 1.5035 | 4.892% | 9.785% | 14.68% | 1.2455 | 1.3767
Rb, ppm | 4.6992 | 0.36331 | 3.9726 | 5.4259 | 3.6093 | 5.7892 | 7.731% | 15.46% | 23.19% | 4.4643 | 4.9342
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 3.3486 | 0.1306 | 3.0874 | 3.6098 | 2.9568 | 3.7404 | 3.9% | 7.801% | 11.7% | 3.1812 | 3.516
Sb, ppm | 8.464 | 0.74997 | 6.964 | 9.9639 | 6.2141 | 10.7139 | 8.861% | 17.72% | 26.58% | 8.0408 | 8.8872
Sc, ppm | 1.0166 | 0.0557 | 0.90522 | 1.128 | 0.84952 | 1.1837 | 5.479% | 10.96% | 16.44% | 0.96579 | 1.0675
Se, ppm | 1.6879 | 0.32778 | 1.0323 | 2.3434 | 0.70454 | 2.6712 | 19.42% | 38.84% | 58.26% | 1.6035 | 1.7723
Sm, ppm | 1.2851 | 0.11025 | 1.0646 | 1.5056 | 0.95439 | 1.6159 | 8.579% | 17.16% | 25.74% | 1.2209 | 1.3494
Sn, ppm | 2.41 | 0.17052 | 2.0689 | 2.751 | 1.8984 | 2.9215 | 7.076% | 14.15% | 21.23% | 2.2895 | 2.5305
Sr, ppm | 9.3935 | 0.60563 | 8.1823 | 10.6048 | 7.5767 | 11.2104 | 6.447% | 12.89% | 19.34% | 8.9239 | 9.8632
Ta, ppm | 2.0539 | 0.62899 | 0.79595 | 3.3119 | 0.16697 | 3.9409 | 30.62% | 61.25% | 91.87% | 1.9512 | 2.1566
Tb, ppm | 0.16284 | 0.026771 | 0.1093 | 0.21638 | 0.082532 | 0.24315 | 16.44% | 32.88% | 49.32% | 0.1547 | 0.17099
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 0.9187 | 0.04425 | 0.8302 | 1.0072 | 0.78595 | 1.0514 | 4.817% | 9.633% | 14.45% | 0.87276 | 0.96463
Ti, wt.% | 0.020513 | 0.0017718 | 0.016969 | 0.024056 | 0.015197 | 0.025828 | 8.637% | 17.27% | 25.91% | 0.019487 | 0.021538
Tl, ppm | 1.5948 | 0.095578 | 1.4037 | 1.786 | 1.3081 | 1.8816 | 5.993% | 11.99% | 17.98% | 1.5151 | 1.6746
Tm, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
U, ppm | 0.69952 | 0.050071 | 0.59938 | 0.79966 | 0.54931 | 0.84974 | 7.158% | 14.32% | 21.47% | 0.66455 | 0.7345
V, ppm | 10.6751 | 0.61879 | 9.4375 | 11.9126 | 8.8187 | 12.5314 | 5.797% | 11.59% | 17.39% | 10.1413 | 11.2088
W, ppm | 2.0495 | 0.17677 | 1.6959 | 2.403 | 1.5192 | 2.5798 | 8.625% | 17.25% | 25.88% | 1.947 | 2.152
Y, ppm | 4.671 | 0.52472 | 3.6216 | 5.7205 | 3.0969 | 6.2452 | 11.23% | 22.47% | 33.7% | 4.4375 | 4.9046
Yb, ppm | 0.29084 | 0.026797 | 0.23725 | 0.34444 | 0.21045 | 0.37124 | 9.214% | 18.43% | 27.64% | 0.2763 | 0.30539
Zn, ppm | 26.5764 | 2.5152 | 21.546 | 31.6067 | 19.0309 | 34.1219 | 9.464% | 18.93% | 28.39% | 25.2475 | 27.9052
Zr, ppm | 12.0591 | 1.1412 | 9.7768 | 14.3415 | 8.6356 | 15.4827 | 9.463% | 18.93% | 28.39% | 11.4562 | 12.6621
Aqua Regia Digestion
Ag, ppm | 4.2003 | 0.19418 | 3.8119 | 4.5887 | 3.6177 | 4.7828 | 4.623% | 9.246% | 13.87% | 3.9903 | 4.4103
Al, wt.% | 0.12605 | 0.0094327 | 0.10719 | 0.14492 | 0.097753 | 0.15435 | 7.483% | 14.97% | 22.45% | 0.11975 | 0.13235
As, ppm | 164.6445 | 4.4792 | 155.6862 | 173.6028 | 151.207 | 178.082 | 2.721% | 5.441% | 8.162% | 156.4123 | 172.8768
Au, ppm | < 0.02 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 26.4313 | 2.6613 | 21.1087 | 31.7538 | 18.4474 | 34.4151 | 10.07% | 20.14% | 30.21% | 25.1097 | 27.7528
Be, ppm | 0.10972 | 0.0083178 | 0.093085 | 0.12636 | 0.084768 | 0.13467 | 7.581% | 15.16% | 22.74% | 0.10424 | 0.11521
Bi, ppm | 17.0565 | 0.88586 | 15.2848 | 18.8282 | 14.3989 | 19.7141 | 5.194% | 10.39% | 15.58% | 16.2037 | 17.9093
Ca, wt.% | 2.5317 | 0.11264 | 2.3065 | 2.757 | 2.1938 | 2.8697 | 4.449% | 8.898% | 13.35% | 2.4052 | 2.6583
Cd, ppm | 0.44055 | 0.029157 | 0.38224 | 0.49887 | 0.35308 | 0.52803 | 6.618% | 13.24% | 19.85% | 0.41853 | 0.46258
Ce, ppm | 4.4804 | 0.58284 | 3.3147 | 5.6461 | 2.7319 | 6.2289 | 13.01% | 26.02% | 39.03% | 4.2564 | 4.7044
Co, ppm | 226.7943 | 12.2683 | 202.2576 | 251.331 | 189.9893 | 263.5993 | 5.409% | 10.82% | 16.23% | 215.4546 | 238.134
Cr, ppm | 10.588 | 0.9634 | 8.6612 | 12.5147 | 7.6978 | 13.4781 | 9.099% | 18.2% | 27.3% | 10.0586 | 11.1173
Cs, ppm | 0.7837 | 0.045835 | 0.69203 | 0.87537 | 0.6462 | 0.92121 | 5.848% | 11.7% | 17.55% | 0.74452 | 0.82289
Cu, wt.% | 1.6779 | 0.019692 | 1.6385 | 1.7173 | 1.6188 | 1.737 | 1.174% | 2.347% | 3.521% | 1.594 | 1.7618
Fe, wt.% | 3.9546 | 0.13865 | 3.6773 | 4.2319 | 3.5386 | 4.3705 | 3.506% | 7.012% | 10.52% | 3.7568 | 4.1523
Ga, ppm | 0.62648 | 0.047737 | 0.531 | 0.72195 | 0.48327 | 0.76969 | 7.62% | 15.24% | 22.86% | 0.59515 | 0.6578
Ge, ppm | 0.040952 | 0.014607 | 0.011739 | 0.070166 | 0 | 0.084773 | 35.67% | 71.34% | 107% | 0.038905 | 0.043
Hf, ppm | 0.04899 | 0.0064125 | 0.036165 | 0.061815 | 0.029752 | 0.068227 | 13.09% | 26.18% | 39.27% | 0.04654 | 0.051439
Hg, ppm | 0.079382 | 0.011359 | 0.056664 | 0.1021 | 0.045305 | 0.11346 | 14.31% | 28.62% | 42.93% | 0.075413 | 0.083351
In, ppm | 1.0208 | 0.042801 | 0.93516 | 1.1064 | 0.89236 | 1.1492 | 4.193% | 8.386% | 12.58% | 0.96973 | 1.0718
K, wt.% | 0.014803 | 0.0046594 | 0.0054845 | 0.024122 | 0.00082512 | 0.028781 | 31.48% | 62.95% | 94.43% | 0.014063 | 0.015543
La, ppm | 1.826 | 0.19544 | 1.4351 | 2.2168 | 1.2397 | 2.4123 | 10.7% | 21.41% | 32.11% | 1.7347 | 1.9173
Li, ppm | 4.6573 | 0.39254 | 3.8723 | 5.4424 | 3.4797 | 5.835 | 8.428% | 16.86% | 25.29% | 4.4245 | 4.8902
Mg, wt.% | 1.4078 | 0.052398 | 1.303 | 1.5126 | 1.2506 | 1.565 | 3.722% | 7.444% | 11.17% | 1.3374 | 1.4782
Mn, wt.% | 0.046404 | 0.0016324 | 0.04314 | 0.049669 | 0.041507 | 0.051302 | 3.518% | 7.036% | 10.55% | 0.044084 | 0.048725
Mo, ppm | 2.0909 | 0.14512 | 1.8007 | 2.3812 | 1.6556 | 2.5263 | 6.94% | 13.88% | 20.82% | 1.9864 | 2.1955
Na, wt.% | 0.0097305 | 0.0020267 | 0.0056772 | 0.013784 | 0.0036505 | 0.015811 | 20.83% | 41.66% | 62.48% | 0.009244 | 0.010217
Nb, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ni, ppm | 23.2926 | 1.3666 | 20.5594 | 26.0258 | 19.1928 | 27.3924 | 5.867% | 11.73% | 17.6% | 22.1279 | 24.4572
P, wt.% | 0.016458 | 0.00093832 | 0.014582 | 0.018335 | 0.013643 | 0.019273 | 5.701% | 11.4% | 17.1% | 0.015635 | 0.017281
Pb, ppm | 54.0453 | 1.9237 | 50.1979 | 57.8927 | 48.2742 | 59.8164 | 3.559% | 7.119% | 10.68% | 51.343 | 56.7475
Rb, ppm | 1.2458 | 0.092157 | 1.0615 | 1.4302 | 0.96937 | 1.5223 | 7.397% | 14.79% | 22.19% | 1.1835 | 1.3081
Re, ppm | 0.001 | 0 | 0.001 | 0.001 | 0.001 | 0.001 | 0% | 0% | 0% | 0.00095 | 0.00105
S, wt.% | 3.3028 | 0.11584 | 3.0711 | 3.5345 | 2.9553 | 3.6503 | 3.507% | 7.015% | 10.52% | 3.1377 | 3.468
Sb, ppm | 7.2982 | 0.97018 | 5.3579 | 9.2386 | 4.3877 | 10.2088 | 13.29% | 26.59% | 39.88% | 6.9333 | 7.6631
Sc, ppm | 0.79583 | 0.12106 | 0.55372 | 1.0379 | 0.43266 | 1.159 | 15.21% | 30.42% | 45.63% | 0.75604 | 0.83562
Se, ppm | 1.6814 | 0.17378 | 1.3338 | 2.0289 | 1.16 | 2.2027 | 10.34% | 20.67% | 31.01% | 1.5973 | 1.7654
Sn, ppm | 1.6332 | 0.097841 | 1.4375 | 1.8289 | 1.3397 | 1.9268 | 5.991% | 11.98% | 17.97% | 1.5516 | 1.7149
Sr, ppm | 7.0009 | 0.4684 | 6.0641 | 7.9377 | 5.5957 | 8.4061 | 6.691% | 13.38% | 20.07% | 6.6508 | 7.3509
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.12829 | 0.013006 | 0.10228 | 0.1543 | 0.089276 | 0.16731 | 10.14% | 20.27% | 30.41% | 0.12188 | 0.13471
Te, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 0.47351 | 0.057303 | 0.35891 | 0.58812 | 0.3016 | 0.64542 | 12.1% | 24.2% | 36.3% | 0.44984 | 0.49719
Ti, wt.% | < 0.005 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tl, ppm | 1.4849 | 0.11187 | 1.2611 | 1.7086 | 1.1493 | 1.8204 | 7.534% | 15.07% | 22.6% | 1.4106 | 1.5591
U, ppm | 0.43846 | 0.033632 | 0.37119 | 0.50572 | 0.33756 | 0.53935 | 7.67% | 15.34% | 23.01% | 0.41653 | 0.46038
V, ppm | 5.0128 | 0.5503 | 3.9122 | 6.1134 | 3.3619 | 6.6637 | 10.98% | 21.96% | 32.93% | 4.7622 | 5.2635
W, ppm | 1.5715 | 0.1802 | 1.2111 | 1.9319 | 1.0309 | 2.1121 | 11.47% | 22.93% | 34.4% | 1.4929 | 1.6501
Y, ppm | 2.9072 | 0.15907 | 2.5891 | 3.2254 | 2.43 | 3.3844 | 5.471% | 10.94% | 16.41% | 2.7619 | 3.0526
Yb, ppm | 0.16426 | 0.039697 | 0.084865 | 0.24365 | 0.045169 | 0.28335 | 24.17% | 48.33% | 72.5% | 0.15605 | 0.17247
Zn, ppm | 24.5001 | 1.266 | 21.9682 | 27.0321 | 20.7022 | 28.2981 | 5.167% | 10.33% | 15.5% | 23.2751 | 25.7251
Zr, ppm | 1.5564 | 0.19113 | 1.1741 | 1.9387 | 0.983 | 2.1298 | 12.28% | 24.56% | 36.84% | 1.4786 | 1.6342
Borate Fusion XRF
Al2O3, wt.% | 0.69869 | 0.038083 | 0.62253 | 0.77486 | 0.58444 | 0.81294 | 5.451% | 10.9% | 16.35% | 0.66376 | 0.73363
BaO, ppm | 63.997 | 22.0408 | 19.9153 | 108.0787 | 0 | 130.1195 | 34.44% | 68.88% | 103.3% | 60.7972 | 67.1969
Bi, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
CaO, wt.% | 3.6928 | 0.068185 | 3.5564 | 3.8292 | 3.4883 | 3.8974 | 1.846% | 3.693% | 5.539% | 3.5082 | 3.8775
Ce, ppm | < 80 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Co, ppm | 236.3398 | 44.2665 | 147.8068 | 324.8728 | 103.5403 | 369.1393 | 18.73% | 37.46% | 56.19% | 224.5228 | 248.1568
Cr2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cu, wt.% | 1.6751 | 0.031611 | 1.6118 | 1.7383 | 1.5802 | 1.7699 | 1.887% | 3.774% | 5.662% | 1.5913 | 1.7588
Fe2O3, wt.% | 5.8129 | 0.049422 | 5.7141 | 5.9118 | 5.6647 | 5.9612 | 0.8502% | 1.7% | 2.551% | 5.5223 | 6.1036
Hf, ppm | < 80 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
K2O, wt.% | 0.10549 | 0.0065424 | 0.092403 | 0.11857 | 0.08586 | 0.12511 | 6.202% | 12.4% | 18.61% | 0.10021 | 0.11076
MgO, wt.% | 5.7472 | 0.14451 | 5.4581 | 6.0362 | 5.3136 | 6.1807 | 2.514% | 5.029% | 7.543% | 5.4598 | 6.0345
MnO, wt.% | 0.061585 | 0.0067175 | 0.04815 | 0.07502 | 0.041433 | 0.081738 | 10.91% | 21.82% | 32.72% | 0.058506 | 0.064664
Mo, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
P2O5, wt.% | 0.038248 | 0.0033854 | 0.031477 | 0.045019 | 0.028092 | 0.048404 | 8.851% | 17.7% | 26.55% | 0.036336 | 0.04016
Rb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sb, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SiO2, wt.% | 73.6264 | 0.47863 | 72.6691 | 74.5836 | 72.1905 | 75.0622 | 0.6501% | 1.3% | 1.95% | 69.945 | 77.3077
Sn, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
SO3, wt.% | 8.3861 | 0.17574 | 8.0347 | 8.7376 | 7.8589 | 8.9134 | 2.096% | 4.191% | 6.287% | 7.9668 | 8.8055
SrO, ppm | < 120 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
TiO2, wt.% | 0.056392 | 0.011016 | 0.03436 | 0.078425 | 0.023344 | 0.089441 | 19.53% | 39.07% | 58.6% | 0.053573 | 0.059212
V, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Zn, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Zr, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Thermogravimetry
LOI1000, wt.% | 4.9939 | 1.039 | 2.9159 | 7.072 | 1.8768 | 8.111 | 20.81% | 41.61% | 62.42% | 4.7442 | 5.2436
Borate / Peroxide Fusion ICP
Al, wt.% | 0.37944 | 0.012672 | 0.3541 | 0.40479 | 0.34143 | 0.41746 | 3.34% | 6.679% | 10.02% | 0.36047 | 0.39841
As, ppm | 164.6967 | 10.6993 | 143.298 | 186.0954 | 132.5987 | 196.7948 | 6.496% | 12.99% | 19.49% | 156.4619 | 172.9316
Ba, ppm | 37.7402 | 1.9212 | 33.8979 | 41.5825 | 31.9768 | 43.5037 | 5.09% | 10.18% | 15.27% | 35.8532 | 39.6272
Be, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Bi, ppm | 17.5957 | 0.66296 | 16.2698 | 18.9216 | 15.6068 | 19.5846 | 3.768% | 7.535% | 11.3% | 16.7159 | 18.4755
Ca, wt.% | 2.6103 | 0.097546 | 2.4152 | 2.8054 | 2.3177 | 2.903 | 3.737% | 7.474% | 11.21% | 2.4798 | 2.7409
Cd, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 10.4998 | 0.32474 | 9.8504 | 11.1493 | 9.5256 | 11.4741 | 3.093% | 6.186% | 9.278% | 9.9749 | 11.0248
Co, ppm | 239.2711 | 12.9675 | 213.3361 | 265.206 | 200.3686 | 278.1735 | 5.42% | 10.84% | 16.26% | 227.3075 | 251.2346
Cr2O3, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cs, ppm | 0.98463 | 0.079139 | 0.82636 | 1.1429 | 0.74722 | 1.2221 | 8.037% | 16.07% | 24.11% | 0.9354 | 1.0339
Cu, wt.% | 1.6602 | 0.071925 | 1.5163 | 1.804 | 1.4444 | 1.8759 | 4.332% | 8.665% | 13% | 1.5771 | 1.7432
Dy, ppm | 1.3233 | 0.095563 | 1.1321 | 1.5144 | 1.0366 | 1.61 | 7.222% | 14.44% | 21.67% | 1.2571 | 1.3894
Er, ppm | 0.58143 | 0.040722 | 0.49999 | 0.66287 | 0.45926 | 0.7036 | 7.004% | 14.01% | 21.01% | 0.55236 | 0.6105
Eu, ppm | 0.1831 | 0.028072 | 0.12695 | 0.23924 | 0.098882 | 0.26731 | 15.33% | 30.66% | 46% | 0.17394 | 0.19225
Fe, wt.% | 4.0565 | 0.078318 | 3.8998 | 4.2131 | 3.8215 | 4.2914 | 1.931% | 3.861% | 5.792% | 3.8536 | 4.2593
Ga, ppm | 1.5385 | 0.4119 | 0.71476 | 2.3623 | 0.30286 | 2.7742 | 26.77% | 53.54% | 80.32% | 1.4616 | 1.6155
Gd, ppm | 1.4397 | 0.13461 | 1.1705 | 1.7089 | 1.0359 | 1.8435 | 9.35% | 18.7% | 28.05% | 1.3677 | 1.5117
Ge, ppm | 1.0349 | 0.17847 | 0.67799 | 1.3919 | 0.49952 | 1.5703 | 17.24% | 34.49% | 51.73% | 0.98318 | 1.0867
Ho, ppm | 0.23194 | 0.01218 | 0.20758 | 0.2563 | 0.1954 | 0.26848 | 5.251% | 10.5% | 15.75% | 0.22035 | 0.24354
In, ppm | 1.1293 | 0.11886 | 0.89157 | 1.367 | 0.77271 | 1.4859 | 10.53% | 21.05% | 31.58% | 1.0728 | 1.1858
K, wt.% | 0.10762 | 0.027755 | 0.05211 | 0.16313 | 0.024355 | 0.19088 | 25.79% | 51.58% | 77.37% | 0.10224 | 0.113
La, ppm | 5.2607 | 0.29287 | 4.675 | 5.8465 | 4.3821 | 6.1393 | 5.567% | 11.13% | 16.7% | 4.9977 | 5.5237
Li, ppm | 105.6452 | 7.4791 | 90.6871 | 120.6033 | 83.208 | 128.0824 | 7.079% | 14.16% | 21.24% | 100.3629 | 110.9275
Mg, wt.% | 3.5042 | 0.1293 | 3.2456 | 3.7628 | 3.1163 | 3.8921 | 3.69% | 7.38% | 11.07% | 3.329 | 3.6794
Mn, wt.% | 0.047588 | 0.0019426 | 0.043703 | 0.051474 | 0.041761 | 0.053416 | 4.082% | 8.164% | 12.25% | 0.045209 | 0.049968
Mo, ppm | 2.2642 | 0.44341 | 1.3773 | 3.151 | 0.93391 | 3.5944 | 19.58% | 39.17% | 58.75% | 2.1509 | 2.3774
Nb, ppm | 1.9153 | 0.45052 | 1.0143 | 2.8164 | 0.56375 | 3.2669 | 23.52% | 47.04% | 70.57% | 1.8196 | 2.0111
Nd, ppm | 5.4963 | 0.29399 | 4.9083 | 6.0843 | 4.6143 | 6.3783 | 5.349% | 10.7% | 16.05% | 5.2215 | 5.7711
Ni, ppm | 29.3903 | 5.6803 | 18.0296 | 40.751 | 12.3493 | 46.4313 | 19.33% | 38.65% | 57.98% | 27.9208 | 30.8598
P, wt.% | 0.019507 | 0.0010058 | 0.017496 | 0.021519 | 0.01649 | 0.022525 | 5.156% | 10.31% | 15.47% | 0.018532 | 0.020483
Pb, ppm | 55.8834 | 6.8083 | 42.2667 | 69.5001 | 35.4584 | 76.3085 | 12.18% | 24.37% | 36.55% | 53.0892 | 58.6776
Pr, ppm | 1.3689 | 0.054078 | 1.2607 | 1.477 | 1.2067 | 1.5311 | 3.951% | 7.901% | 11.85% | 1.3004 | 1.4373
Rb, ppm | 4.8892 | 0.33048 | 4.2282 | 5.5502 | 3.8978 | 5.8806 | 6.759% | 13.52% | 20.28% | 4.6447 | 5.1337
S, wt.% | 3.3566 | 0.10275 | 3.1511 | 3.5621 | 3.0484 | 3.6649 | 3.061% | 6.122% | 9.183% | 3.1888 | 3.5244
Sb, ppm | 8.2988 | 1.1186 | 6.0616 | 10.5359 | 4.9431 | 11.6545 | 13.48% | 26.96% | 40.44% | 7.8839 | 8.7137
Sc, ppm | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Si, wt.% | 34.7556 | 1.2863 | 32.1831 | 37.3282 | 30.8969 | 38.6144 | 3.701% | 7.402% | 11.1% | 33.0179 | 36.4934
Sm, ppm | 1.4821 | 0.10683 | 1.2684 | 1.6957 | 1.1616 | 1.8026 | 7.208% | 14.42% | 21.62% | 1.408 | 1.5562
Sn, ppm | 3.0359 | 0.42432 | 2.1873 | 3.8846 | 1.763 | 4.3089 | 13.98% | 27.95% | 41.93% | 2.8841 | 3.1877
Sr, ppm | 9.6895 | 0.92003 | 7.8494 | 11.5295 | 6.9294 | 12.4495 | 9.495% | 18.99% | 28.49% | 9.205 | 10.1739
Ta, ppm | 2.2975 | 0.39964 | 1.4982 | 3.0968 | 1.0985 | 3.4964 | 17.39% | 34.79% | 52.18% | 2.1826 | 2.4123
Tb, ppm | 0.24288 | 0.032149 | 0.17858 | 0.30718 | 0.14643 | 0.33933 | 13.24% | 26.47% | 39.71% | 0.23073 | 0.25502
Th, ppm | 0.90608 | 0.085462 | 0.73515 | 1.077 | 0.64969 | 1.1625 | 9.432% | 18.86% | 28.3% | 0.86078 | 0.95138
Ti, wt.% | 0.029636 | 0.00089671 | 0.027842 | 0.031429 | 0.026946 | 0.032326 | 3.026% | 6.051% | 9.077% | 0.028154 | 0.031118
Tl, ppm | 1.6747 | 0.097404 | 1.4799 | 1.8695 | 1.3825 | 1.9669 | 5.816% | 11.63% | 17.45% | 1.591 | 1.7584
Tm, ppm | 0.074333 | 0.013817 | 0.046699 | 0.10197 | 0.032881 | 0.11579 | 18.59% | 37.18% | 55.77% | 0.070617 | 0.07805
U, ppm | 0.72855 | 0.067157 | 0.59424 | 0.86286 | 0.52708 | 0.93002 | 9.218% | 18.44% | 27.65% | 0.69212 | 0.76498
V, ppm | 11.7208 | 2.2698 | 7.1812 | 16.2605 | 4.9113 | 18.5303 | 19.37% | 38.73% | 58.1% | 11.1348 | 12.3069
W, ppm | 2.5489 | 0.61283 | 1.3232 | 3.7746 | 0.71041 | 4.3874 | 24.04% | 48.09% | 72.13% | 2.4215 | 2.6764
Y, ppm | 7.0013 | 0.2715 | 6.4583 | 7.5443 | 6.1868 | 7.8158 | 3.878% | 7.756% | 11.63% | 6.6513 | 7.3514
Yb, ppm | 0.42482 | 0.048515 | 0.32779 | 0.52185 | 0.27928 | 0.57037 | 11.42% | 22.84% | 34.26% | 0.40358 | 0.44606
Zn, ppm | 31.3796 | 7.0379 | 17.3038 | 45.4555 | 10.2658 | 52.4934 | 22.43% | 44.86% | 67.29% | 29.8106 | 32.9486
Zr, ppm | 21.6982 | 1.6964 | 18.3053 | 25.091 | 16.6089 | 26.7875 | 7.818% | 15.64% | 23.46% | 20.6133 | 22.7831
Infrared Combustion
C, wt.% | 1.6441 | 0.048104 | 1.5478 | 1.7403 | 1.4997 | 1.7884 | 2.926% | 5.852% | 8.778% | 1.5619 | 1.7263
S, wt.% | 3.3731 | 0.066919 | 3.2392 | 3.5069 | 3.1723 | 3.5738 | 1.984% | 3.968% | 5.952% | 3.2044 | 3.5417
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 163b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
PSF*XRF | pyro-sulfate fusion with X-ray fluorescence spectroscopy
Text Values:
N.A. | Not Applicable (Lab 2.19)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 163b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
Inspectorate (BV), Lima, Peru
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek, Townsville, QLD, Australia
Intertek Genalysis, Adelaide, SA, Australia
MSALABS, Vancouver, BC, Canada
Paragon Geochemical Laboratories, Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## 4-Acid
| Analytical results for Ag in OREAS 163b (Certified Value 4.26 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 4.2 | 4.89 | 4 | 4.29 | 4.21 | 4.2667 | 4.44 | 4.17 | 4.03 | 3.92 | 4.09 | 4.3 | 4.37 | 4.3037 | 4.1 | 4.49 | 4.457 | 4.3 | 4.3 | 4.4 | 4.01 | 4.3 | 4.52 | 4.33 | 4 | 1
| 1 | 2 | 4.3 | 4.84 | 4.1 | 4.24 | 4.32 | 4.2448 | 4.5 | 4.16 | 4.26 | 4.06 | 4.21 | 4.2 | 4.31 | 4.3521 | 4.1 | 4.46 | 4.255 | 4.2 | 4.2 | 4.6 | 4.05 | 4.4 | 4.62 | 4.27 | 4 | 8
| 1 | 3 | 4.2 | 4.72 | 4.1 | 4.02 | 4.28 | 4.3008 | 4.59 | 4.28 | 4.45 | 3.95 | 4.02 | 4.2 | 4.34 | 4.3776 | 3.8 | 4.45 | 4.094 | 4.2 | 4.1 | 4.4 | 4.03 | 4.3 | 4.5 | 4.48 | 4.1 | 16
| 1 | 4 | 4.3 | 4.78 | 4.2 | 4.23 | 4.26 | 4.247 | 4.6 | 4.15 | 4.32 | 3.86 | 3.99 | 4.2 | 4.42 | 4.31 | 4.2 | 4.48 | 4.4 | 4.3 | 4.3 | 4.6 | 4.03 | 4.3 | 4.59 | 4.48 | 4.1 | 4.2584
| 1 | 5 | 4.3 | 4.59 | 4 | 4.26 | 4.17 | 4.3186 | 4.46 | 4.28 | 4.03 | 3.96 | 4.25 | 4.3 | 4.47 | 4.4193 | 3.2 | 4.49 | 4.736 | 4.2 | 4.1 | 4.5 | 4.03 | 4.2 | 4.54 | 4.33 | 4 | 7
| 1 | 6 | 4.2 | 5.08 | 4.1 | 4.26 | 4.3 | 4.2407 | 4.44 | 4.08 | 4.27 | 3.93 | 4.05 | 4.2 | 4.39 | 4.4358 | 3.9 | 4.48 | 4.118 | 4.2 | 4.3 | 4.5 | 4.03 | 4.4 | 4.42 | 4.63 | 4.1
| Mean |  | 4.25 | 4.8167 | 4.0833 | 4.2167 | 4.2567 | 4.2698 | 4.505 | 4.1867 | 4.2267 | 3.9467 | 4.1017 | 4.2333 | 4.3833 | 4.3664 | 3.8833 | 4.475 | 4.3433 | 4.2333 | 4.2167 | 4.5 | 4.03 | 4.3167 | 4.5317 | 4.42 | 4.05
| Median |  | 4.25 | 4.81 | 4.1 | 4.25 | 4.27 | 4.2568 | 4.48 | 4.165 | 4.265 | 3.94 | 4.07 | 4.2 | 4.38 | 4.3649 | 4 | 4.48 | 4.3275 | 4.2 | 4.25 | 4.5 | 4.03 | 4.3 | 4.53 | 4.405 | 4.05
| Std Dev. |  | 0.054772 | 0.16573 | 0.075277 | 0.098522 | 0.056804 | 0.032694 | 0.073144 | 0.078909 | 0.16669 | 0.065625 | 0.10553 | 0.05164 | 0.057155 | 0.054881 | 0.3656 | 0.016432 | 0.24127 | 0.05164 | 0.098319 | 0.089443 | 0.012649 | 0.075277 | 0.070545 | 0.13416 | 0.054772
| Rel.Std.Dev. |  | 1.289% | 3.441% | 1.844% | 2.337% | 1.334% | 0.7657% | 1.624% | 1.885% | 3.944% | 1.663% | 2.573% | 1.22% | 1.304% | 1.257% | 9.415% | 0.3672% | 5.555% | 1.22% | 2.332% | 1.988% | 0.3139% | 1.744% | 1.557% | 3.035% | 1.352%
| PDM3 |  | -0.1978% | 13.11% | -4.112% | -0.9806% | -0.04129% | 0.2666% | 5.79% | -1.685% | -0.7458% | -7.321% | -3.681% | -0.5892% | 2.933% | 2.536% | -8.808% | 5.086% | 1.994% | -0.5892% | -0.9806% | 5.673% | -5.364% | 1.368% | 6.416% | 3.794% | -4.894%
| Z-Score (Absolute) |  | 0 | 2.87 | 0.84 | 0.17 | 0.03 | 0.1 | 1.29 | 0.32 | 0.12 | 1.53 | 0.75 | 0.08 | 0.67 | 0.59 | 1.85 | 1.14 | 0.47 | 0.08 | 0.17 | 1.26 | 1.11 | 0.34 | 1.42 | 0.86 | 1.01
| Analytical results for Al in OREAS 163b (Certified Value 0.372 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.399 | 0.35001 | 0.375 | 0.38 | 0.3341 | 0.38537 | 0.38 | 0.36 | 0.36 | 0.38 | 0.37 | 0.38 | 0.37 | 0.39111 | 0.36 | 0.3908 | 0.38 | 0.3496 | 0.37 | 0.37 | 0.38 | 0.37 | 0.3695 | 0.38 | 0.48 | 1
| 1 | 2 | 0.396 | 0.34226 | 0.372 | 0.35 | 0.3608 | 0.37619 | 0.38 | 0.36 | 0.4 | 0.38 | 0.37 | 0.38 | 0.37 | 0.39489 | 0.37 | 0.3869 | 0.46 | 0.3479 | 0.37 | 0.37 | 0.37 | 0.36 | 0.3762 | 0.38 | 0.47 | 0
| 1 | 3 | 0.397 | 0.35659 | 0.378 | 0.36 | 0.3507 | 0.37608 | 0.4 | 0.38 | 0.39 | 0.38 | 0.36 | 0.38 | 0.39 | 0.38523 | 0.35 | 0.375 | 0.37 | 0.3464 | 0.37 | 0.37 | 0.38 | 0.37 | 0.3585 | 0.39 | 0.47 | 16
| 1 | 4 | 0.388 | 0.35492 | 0.374 | 0.37 | 0.3315 | 0.38598 | 0.37 | 0.37 | 0.39 | 0.37 | 0.36 | 0.37 | 0.38 | 0.39085 | 0.37 | 0.3837 | 0.39 | 0.3533 | 0.37 | 0.37 | 0.38 | 0.37 | 0.3687 | 0.38 | 0.46 | 0.37241
| 1 | 5 | 0.392 | 0.35479 | 0.375 | 0.34 | 0.3511 | 0.39066 | 0.39 | 0.37 | 0.36 | 0.37 | 0.37 | 0.39 | 0.39 | 0.38405 | 0.27 | 0.3969 | 0.37 | 0.3532 | 0.37 | 0.37 | 0.38 | 0.37 | 0.3689 | 0.38 | 0.48 | 8
| 1 | 6 | 0.398 | 0.36005 | 0.375 | 0.34 | 0.345 | 0.37947 | 0.38 | 0.37 | 0.38 | 0.37 | 0.38 | 0.38 | 0.39 | 0.39298 | 0.34 | 0.3742 | 0.36 | 0.3491 | 0.37 | 0.37 | 0.38 | 0.36 | 0.3597 | 0.39 | 0.44
| Mean |  | 0.395 | 0.3531 | 0.37483 | 0.35667 | 0.34553 | 0.38229 | 0.38333 | 0.36833 | 0.38 | 0.375 | 0.36833 | 0.38 | 0.38167 | 0.38985 | 0.34333 | 0.38458 | 0.38833 | 0.34992 | 0.37 | 0.37 | 0.37833 | 0.36667 | 0.36692 | 0.38333 | 0.46667
| Median |  | 0.3965 | 0.35485 | 0.375 | 0.355 | 0.34785 | 0.38242 | 0.38 | 0.37 | 0.385 | 0.375 | 0.37 | 0.38 | 0.385 | 0.39098 | 0.355 | 0.3853 | 0.375 | 0.34935 | 0.37 | 0.37 | 0.38 | 0.37 | 0.3688 | 0.38 | 0.47
| Std Dev. |  | 0.0041952 | 0.0062242 | 0.0019408 | 0.01633 | 0.011124 | 0.0059475 | 0.010328 | 0.0075277 | 0.016733 | 0.0054772 | 0.0075277 | 0.0063246 | 0.0098319 | 0.0043083 | 0.037771 | 0.0089011 | 0.03656 | 0.0028081 | 6.0809e-17 | 6.0809e-17 | 0.0040825 | 0.005164 | 0.0066767 | 0.005164 | 0.015055
| Rel.Std.Dev. |  | 1.062% | 1.763% | 0.5178% | 4.578% | 3.219% | 1.556% | 2.694% | 2.044% | 4.403% | 1.461% | 2.044% | 1.664% | 2.576% | 1.105% | 11% | 2.314% | 9.415% | 0.8025% | 1.643e-14% | 1.643e-14% | 1.079% | 1.408% | 1.82% | 1.347% | 3.226%
| PDM3 |  | 6.066% | -5.184% | 0.6505% | -4.228% | -7.217% | 2.653% | 2.933% | -1.095% | 2.038% | 0.6952% | -1.095% | 2.038% | 2.485% | 4.683% | -7.808% | 3.269% | 4.275% | -6.04% | -0.6474% | -0.6474% | 1.59% | -1.542% | -1.475% | 2.933% | 25.31%
| Z-Score (Absolute) |  | 1.62 | 1.77 | 0.01 | 1.48 | 2.38 | 0.59 | 0.67 | 0.54 | 0.4 | 0 | 0.54 | 0.4 | 0.54 | 1.2 | 2.56 | 0.78 | 1.08 | 2.03 | 0.4 | 0.4 | 0.27 | 0.67 | 0.65 | 0.67 | 7.42
| Analytical results for As in OREAS 163b (Certified Value 168 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 168 | 165 | 161.8 | 162.9937 | 167.5 | 162.5 | 162 | 170.5 | 167.5 | 170 | 157 | 165.8083 | 161 | 171.4 | 171.5 | 151.2 | 186 | 158 | 167 | 168 | 177.9 | 171.5 | 20.1 |  |  | 1
| 1 | 2 | 167 | 154.4 | 168.4 | 161.5078 | 167.5 | 164 | 184 | 172 | 171.5 | 165 | 159 | 169.7099 | 163 | 174.8 | 175.9 | 147.9 | 182 | 150 | 169 | 168 | 180.9 | 173.1 | 21.4 |  |  | 9
| 1 | 3 | 165 | 151.5 | 163.3 | 164.2795 | 169 | 170 | 171 | 170.5 | 167 | 167 | 158 | 170.4411 | 162 | 175.9 | 168.8 | 150 | 179 | 165 | 168 | 168 | 175.1 | 175.2 | 22.5 |  |  | 16
| 1 | 4 | 167 | 161.7 | 161.2 | 163.0302 | 166.5 | 170.5 | 180 | 167 | 167.5 | 156 | 159 | 166.707 | 179 | 172.4 | 176.4 | 151.1 | 184 | 165 | 169 | 163 | 175.2 | 169.1 | 19.3 |  |  | 167.8738
| 1 | 5 | 169 | 146.6 | 163.7 | 160.7232 | 168.5 | 173.5 | 165.5 | 169.5 | 175 | 160 | 159 | 167.9925 | 129 | 169.4 | 168.2 | 153.1 | 181 | 158 | 168 | 163 | 179 | 173.5 | 22.1 |  |  | 9
| 1 | 6 | 166 | 149.7 | 165.7 | 160.9569 | 162 | 167 | 177 | 168 | 173.5 | 162 | 162 | 169.0504 | 172 | 175.1 | 165.2 | 150.4 | 187 | 163 | 169 | 164 | 174.9 | 169.8 | 24.6
| Mean |  | 167 | 154.8167 | 164.0167 | 162.2486 | 166.8333 | 167.9167 | 173.25 | 169.5833 | 170.3333 | 163.3333 | 159 | 168.2849 | 161 | 173.1667 | 171 | 150.6167 | 183.1667 | 159.8333 | 168.3333 | 165.6667 | 177.1667 | 172.0333 | 21.6667
| Median |  | 167 | 152.95 | 163.5 | 162.2508 | 167.5 | 168.5 | 174 | 170 | 169.5 | 163.5 | 159 | 168.5215 | 162.5 | 173.6 | 170.15 | 150.75 | 183 | 160.5 | 168.5 | 166 | 176.55 | 172.3 | 21.75
| Std Dev. |  | 1.4142 | 7.1547 | 2.6649 | 1.4025 | 2.5232 | 4.1883 | 8.5659 | 1.828 | 3.4737 | 5.0465 | 1.6733 | 1.7874 | 17.1697 | 2.5208 | 4.4663 | 1.7058 | 3.0605 | 5.7764 | 0.8165 | 2.582 | 2.4945 | 2.3321 | 1.8769
| Rel.Std.Dev. |  | 0.8468% | 4.621% | 1.625% | 0.8644% | 1.512% | 2.494% | 4.944% | 1.078% | 2.039% | 3.09% | 1.052% | 1.062% | 10.66% | 1.456% | 2.612% | 1.133% | 1.671% | 3.614% | 0.485% | 1.559% | 1.408% | 1.356% | 8.663%
| PDM3 |  | -0.5205% | -7.778% | -2.298% | -3.351% | -0.6198% | 0.02553% | 3.203% | 1.018% | 1.465% | -2.705% | -5.286% | 0.2449% | -4.095% | 3.153% | 1.862% | -10.28% | 9.11% | -4.79% | 0.2737% | -1.315% | 5.536% | 2.478% | -87.09%
| Z-Score (Absolute) |  | 0 | 1.73 | 0.42 | 0.67 | 0.02 | 0.13 | 0.89 | 0.37 | 0.47 | 0.52 | 1.14 | 0.18 | 0.85 | 0.88 | 0.57 | 2.33 | 2.29 | 1.02 | 0.19 | 0.19 | 1.44 | 0.71 | 20.63
| Analytical results for B in OREAS 163b (Indicative Value 19 ppm) | [26 empty] | Indicative
B | Round | Replicate | 2.01: Lab |  |  | 1
| No | No | 2.01: 2.20 |  |  | ppm
|  |  | 2.01: 4A*OES |  |  | 1
1
| 1 | 1 | 2.01: 19 |  |  | 1
| 1 | 2 | 2.01: 19 |  |  | 1
| 1 | 3 | 2.01: 19 |  |  | 16
| 1 | 4 | 2.01: 19 |  |  | 19
| 1 | 5 | 2.01: 19 |  |  | 7
| 1 | 6 | 2.01: 19
| Mean |  | 2.01: 19
| Median |  | 2.01: 19
| Std Dev. |  | 2.01: 0
Rel.Std.Dev. | bar: 2.01
PDM3 | bar: 2.01
Z-Score (Absolute) | bar: 2.01
| Analytical results for Ba in OREAS 163b (Certified Value 38.5 ppm) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 36 | 65 | 41.12 | 36.3 | 36.5495 | 40 | 40 | 40 | 40 | 40 | 40 | 34 | 38.3914 | 50 | 37.5 | 40 | 40 | 40 | 95 | 36 | 37 |  |  |  |  | 1
| 1 | 2 | 36 | 65 | 48.15 | 37.9 | 37.6515 | 40 | 40 | 40 | 40 | 40 | 41 | 42 | 39.173 | 40 | 39.7 | 41 | 38 | 42 | 94 | 36 | 36 |  |  |  |  | 27
| 1 | 3 | 36 | 65 | 47.73 | 37.5 | 38.1912 | 40 | 40 | 40 | 40 | 30 | 39 | 40 | 39.172 | 30 | 38.7 | 40 | 38 | 41 | 93 | 36 | 38 |  |  |  |  | 16
| 1 | 4 | 36 | 65 | 40.93 | 36.9 | 36.952 | 40 | 40 | 40 | 40 | 30 | 35 | 41 | 38.8099 | 40 | 37 | 41 | 39 | 43 | 93 | 36 | 37 |  |  |  |  | 38.5096
| 1 | 5 | 36 | 65 | 47 | 36.2 | 37.6149 | 40 | 40 | 40 | 40 | 40 | 38 | 41 | 38.1276 | 40 | 37.1 | 39 | 39 | 39 | 93 | 36 | 37 |  |  |  |  | 10
| 1 | 6 | 36 | 65 | 46.61 | 36.4 | 36.6235 | 40 | 40 | 40 | 40 | 40 | 41 | 35 | 38.4061 | 30 | 36.4 | 40 | 37 | 40 | 93 | 37 | 35
| Mean |  | 36 | 65 | 45.2567 | 36.8667 | 37.2638 | 40 | 40 | 40 | 40 | 36.6667 | 39 | 38.8333 | 38.68 | 38.3333 | 37.7333 | 40.1667 | 38.5 | 40.8333 | 93.5 | 36.1667 | 36.6667
| Median |  | 36 | 65 | 46.805 | 36.65 | 37.2835 | 40 | 40 | 40 | 40 | 40 | 39.5 | 40.5 | 38.608 | 40 | 37.3 | 40 | 38.5 | 40.5 | 93 | 36 | 37
| Std Dev. |  | 0 | 0 | 3.3224 | 0.70048 | 0.6559 | 0 | 0 | 0 | 0 | 5.164 | 2.2804 | 3.4303 | 0.43943 | 7.5277 | 1.2307 | 0.75277 | 1.0488 | 1.472 | 0.83666 | 0.40825 | 1.0328
| Rel.Std.Dev. |  | 0% | 0% | 7.341% | 1.9% | 1.76% | 0% | 0% | 0% | 0% | 14.08% | 5.847% | 8.833% | 1.136% | 19.64% | 3.262% | 1.874% | 2.724% | 3.605% | 0.8948% | 1.129% | 2.817%
| PDM3 |  | -6.517% | 68.79% | 17.52% | -4.266% | -3.235% | 3.87% | 3.87% | 3.87% | 3.87% | -4.786% | 1.273% | 0.8405% | 0.4423% | -0.4578% | -2.016% | 4.303% | -0.02506% | 6.034% | 142.8% | -6.084% | -4.786%
| Z-Score (Absolute) |  | 1 | 9.79 | 2.45 | 0.67 | 0.53 | NA | NA | NA | NA | NA | 0.12 | 0.06 | 0 | NA | 0.35 | 0.55 | 0.07 | 0.8 | 20.39 | 0.93 | 0.75
| Analytical results for Be in OREAS 163b (Certified Value 0.29 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | < 0.5 | 0.1704 | < 1 | < 0.2 | 0.26 | 0.27259 | 0.3 | 0.29 | 0.31 | 0.3 | 0.28 | 0.4 | 0.25 | < 0.5 | 0.21 | 0.34 | < 1 | < 1 | < 5 | < 1 | < 0.5 | 0.25 | 0.29 | < 1 |  | 1
| 1 | 2 | < 0.5 | 0.16596 | < 1 | < 0.2 | 0.3 | 0.2988 | 0.3 | 0.3 | 0.32 | 0.27 | 0.26 | 0.4 | 0.26 | < 0.5 | 0.33 | 0.32 | < 1 | < 1 | < 5 | < 1 | < 0.5 | 0.26 | 0.3 | < 1 |  | 28
| 1 | 3 | < 0.5 | 0.17628 | < 1 | < 0.2 | 0.3 | 0.2641 | 0.3 | 0.37 | 0.3 | 0.34 | 0.27 | 0.3 | 0.24 | < 0.5 | 0.46 | 0.29 | < 1 | < 1 | < 5 | < 1 | < 0.5 | 0.27 | 0.28 | < 1 |  | 16
| 1 | 4 | < 0.5 | 0.18672 | < 1 | < 0.2 | 0.34 | 0.26239 | 0.29 | 0.3 | 0.29 | 0.27 | 0.26 | 0.3 | 0.27 | < 0.5 | 0.48 | 0.25 | < 1 | < 1 | < 5 | < 1 | < 0.5 | 0.26 | 0.3 | < 1 |  | 0.28554
| 1 | 5 | < 0.5 | 0.20972 | < 1 | < 0.2 | 0.31 | 0.26982 | 0.3 | 0.28 | 0.3 | 0.31 | 0.27 | 0.4 | 0.26 | < 0.5 | 0.06 | 0.28 | < 1 | < 1 | < 5 | < 1 | < 0.5 | 0.28 | 0.29 | < 1 |  | 11
| 1 | 6 | < 0.5 | 0.1956 | < 1 | < 0.2 | 0.32 | 0.28282 | 0.33 | 0.33 | 0.32 | 0.3 | 0.3 | 0.4 | 0.29 | < 0.5 | 0.21 | 0.31 | < 1 | < 1 | < 5 | < 1 | < 0.5 | 0.28 | 0.3 | < 1
| Mean |  | 2.02: 0.18411; 2.05: 0.305; 2.06: 0.27509; 2.07: 0.30333; 2.08: 0.31167; 2.09: 0.30667; 2.10: 0.29833; 2.11: 0.27333; 2.12: 0.36667; 2.13: 0.26167; 2.15: 0.29167; 2.16: 0.29833; 2.23: 0.26667; 2.24: 0.29333
| Median |  | 2.02: 0.1815; 2.05: 0.305; 2.06: 0.2712; 2.07: 0.3; 2.08: 0.3; 2.09: 0.305; 2.10: 0.3; 2.11: 0.27; 2.12: 0.4; 2.13: 0.26; 2.15: 0.27; 2.16: 0.3; 2.23: 0.265; 2.24: 0.295
| Std Dev. |  | 2.02: 0.016573; 2.05: 0.026646; 2.06: 0.013696; 2.07: 0.013663; 2.08: 0.033116; 2.09: 0.012111; 2.10: 0.026394; 2.11: 0.015055; 2.12: 0.05164; 2.13: 0.017224; 2.15: 0.16266; 2.16: 0.031885; 2.23: 0.012111; 2.24: 0.008165
| Rel.Std.Dev. |  | 2.02: 9.002%; 2.05: 8.736%; 2.06: 4.979%; 2.07: 4.504%; 2.08: 10.63%; 2.09: 3.949%; 2.10: 8.847%; 2.11: 5.508%; 2.12: 14.08%; 2.13: 6.582%; 2.15: 55.77%; 2.16: 10.69%; 2.23: 4.541%; 2.24: 2.784%
| PDM3 |  | 2.02: -35.52%; 2.05: 6.817%; 2.06: -3.659%; 2.07: 6.233%; 2.08: 9.152%; 2.09: 7.401%; 2.10: 4.482%; 2.11: -4.273%; 2.12: 28.41%; 2.13: -8.359%; 2.15: 2.147%; 2.16: 4.482%; 2.23: -6.608%; 2.24: 2.731%
| Z-Score (Absolute) |  | 0.98 | 2.38 | 4.35 | 4.17 | 0.2 | 0.44 | 0.16 | 0.34 | 0.23 | 0.05 | 0.48 | NA | 0.73 | 0.98 | 0.09 | 0.05 | 4.35 | 4.35 | NA | 4.35 | 0.98 | 0.62 | 0.05 | 4.35
Results from laboratory 2.12 and 2.20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 163b (Certified Value 17.3 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 17.1 | 17.305 | 17.5 | 13.27 | 16.8 | 18.9437 | 16.2 | 16.15 | 17.65 | 16.6 | 14.9 | 17.9 | 18.26 | 18.3534 | 15.8 | 17.43 | 16.95 | 10.74 | 18.1 | 18.76 | 17.7 | 16.59 | 16.8 | 17.17 | 17.2 | 1
| 1 | 2 | 17.5 | 17.509 | 17.4 | 11.33 | 16.96 | 19.1612 | 16.65 | 15.75 | 18.6 | 17 | 15.5 | 17.7 | 17.62 | 18.0152 | 13.5 | 17.57 | 16.15 | 10.63 | 17.8 | 20.17 | 17.8 | 16.95 | 16.97 | 16.83 | 17 | 29
| 1 | 3 | 16.9 | 17.398 | 17.6 | 13.27 | 16.85 | 19.6475 | 17.2 | 16.5 | 19 | 16.75 | 15.1 | 18 | 18.71 | 17.5568 | 13.4 | 17.58 | 16.32 | 10.63 | 17.5 | 19.66 | 18 | 17.12 | 16.57 | 17.69 | 17.2 | 16
| 1 | 4 | 16.7 | 17.48 | 17.5 | 14.58 | 16.82 | 18.6015 | 16.8 | 16.7 | 18.3 | 16.3 | 14.85 | 16.8 | 18.22 | 17.1235 | 15.7 | 17.42 | 17.3 | 10.85 | 18.3 | 17.55 | 17.5 | 16.51 | 16.47 | 17.57 | 17.1 | 17.2788
| 1 | 5 | 16.6 | 17.31 | 17.5 | 13.22 | 16.72 | 17.8364 | 16.35 | 16.85 | 17.9 | 15.9 | 16.05 | 17.3 | 17.94 | 17.6032 | 14.4 | 17.43 | 16.76 | 10.79 | 18.3 | 17.54 | 18.2 | 16.37 | 16.9 | 17.07 | 17.3 | 12
| 1 | 6 | 17.1 | 17.638 | 17.6 | 11.71 | 17.36 | 18.1731 | 16.95 | 16.65 | 18.05 | 16.15 | 15.2 | 17.6 | 17.97 | 18.5298 | 12.7 | 17.69 | 16 | 12.61 | 18.2 | 17.33 | 17.6 | 17.02 | 16.79 | 17.11 | 17.8
| Mean |  | 16.9833 | 17.44 | 17.5167 | 12.8967 | 16.9183 | 18.7272 | 16.6917 | 16.4333 | 18.25 | 16.45 | 15.2667 | 17.55 | 18.12 | 17.8636 | 14.25 | 17.52 | 16.58 | 11.0417 | 18.0333 | 18.5017 | 17.8 | 16.76 | 16.75 | 17.24 | 17.2667
| Median |  | 17 | 17.439 | 17.5 | 13.245 | 16.835 | 18.7726 | 16.725 | 16.575 | 18.175 | 16.45 | 15.15 | 17.65 | 18.095 | 17.8092 | 13.95 | 17.5 | 16.54 | 10.765 | 18.15 | 18.155 | 17.75 | 16.77 | 16.795 | 17.14 | 17.2
| Std Dev. |  | 0.32506 | 0.12842 | 0.075277 | 1.1899 | 0.22999 | 0.66311 | 0.37339 | 0.41069 | 0.49193 | 0.40743 | 0.44907 | 0.44159 | 0.36949 | 0.53223 | 1.2818 | 0.11063 | 0.50529 | 0.77326 | 0.32042 | 1.2161 | 0.26077 | 0.3088 | 0.19277 | 0.32564 | 0.28048
| Rel.Std.Dev. |  | 1.914% | 0.7363% | 0.4297% | 9.227% | 1.359% | 3.541% | 2.237% | 2.499% | 2.696% | 2.477% | 2.942% | 2.516% | 2.039% | 2.979% | 8.995% | 0.6315% | 3.048% | 7.003% | 1.777% | 6.573% | 1.465% | 1.843% | 1.151% | 1.889% | 1.624%
| PDM3 |  | -1.71% | 0.9329% | 1.377% | -25.36% | -2.086% | 8.383% | -3.398% | -4.893% | 5.621% | -4.797% | -11.65% | 1.569% | 4.868% | 3.385% | -17.53% | 1.396% | -4.044% | -36.1% | 4.367% | 7.077% | 3.016% | -3.003% | -3.06% | -0.2246% | -0.07029%
| Z-Score (Absolute) |  | 0.28 | 0.22 | 0.3 | 4.7 | 0.35 | 1.61 | 0.59 | 0.87 | 1.09 | 0.85 | 2.13 | 0.34 | 0.95 | 0.67 | 3.23 | 0.3 | 0.71 | 6.7 | 0.86 | 1.36 | 0.61 | 0.52 | 0.53 | 0 | 0.03
| Analytical results for Ca in OREAS 163b (Certified Value 2.64 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 2.68 | 2.5759 | 2.62 | 2.69 | 2.4291 | 2.5948 | 2.75 | 2.64 | 2.56 | 2.66 | 2.62 | 2.77 | 2.605 | 2.3649 | 2.55 | 2.7136 | 2.68 | 2.6523 | 2.64 | 2.69 | 2.62 | 2.595 | 2.6903 | 2.68 | 2.84 | 1
| 1 | 2 | 2.68 | 2.5776 | 2.64 | 2.48 | 2.6383 | 2.5374 | 2.77 | 2.64 | 2.76 | 2.66 | 2.69 | 2.68 | 2.597 | 2.6697 | 2.56 | 2.6514 | 2.72 | 2.6392 | 2.61 | 2.64 | 2.64 | 2.549 | 2.7388 | 2.67 | 2.84 | 0
| 1 | 3 | 2.69 | 2.6143 | 2.58 | 2.42 | 2.5371 | 2.5536 | 2.74 | 2.73 | 2.74 | 2.71 | 2.61 | 2.69 | 2.736 | 2.5582 | 2.42 | 2.6119 | 2.68 | 2.6403 | 2.58 | 2.65 | 2.63 | 2.55 | 2.6175 | 2.75 | 2.87 | 16
| 1 | 4 | 2.67 | 2.5029 | 2.58 | 2.64 | 2.4265 | 2.5767 | 2.63 | 2.68 | 2.7 | 2.63 | 2.63 | 2.51 | 2.653 | 2.4373 | 2.68 | 2.6811 | 2.69 | 2.6826 | 2.59 | 2.6 | 2.62 | 2.578 | 2.6152 | 2.65 | 2.86 | 2.6413
| 1 | 5 | 2.68 | 2.5314 | 2.66 | 2.39 | 2.5641 | 2.5625 | 2.79 | 2.66 | 2.56 | 2.65 | 2.7 | 2.62 | 2.702 | 2.6346 | 2.02 | 2.677 | 2.61 | 2.6415 | 2.6 | 2.6 | 2.62 | 2.583 | 2.6649 | 2.66 | 2.82 | 13
| 1 | 6 | 2.7 | 2.5989 | 2.68 | 2.43 | 2.5081 | 2.5766 | 2.69 | 2.65 | 2.65 | 2.6 | 2.76 | 2.69 | 2.657 | 2.7682 | 2.5 | 2.6132 | 2.6 | 2.6326 | 2.59 | 2.65 | 2.62 | 2.54 | 2.6294 | 2.75 | 2.88
| Mean |  | 2.6833 | 2.5668 | 2.6267 | 2.5083 | 2.5172 | 2.5669 | 2.7283 | 2.6667 | 2.6617 | 2.6517 | 2.6683 | 2.66 | 2.6583 | 2.5721 | 2.455 | 2.658 | 2.6633 | 2.6481 | 2.6017 | 2.6383 | 2.625 | 2.5658 | 2.6593 | 2.6933 | 2.8517
| Median |  | 2.68 | 2.5768 | 2.63 | 2.455 | 2.5226 | 2.5696 | 2.745 | 2.655 | 2.675 | 2.655 | 2.66 | 2.685 | 2.655 | 2.5964 | 2.525 | 2.6642 | 2.68 | 2.6409 | 2.595 | 2.645 | 2.62 | 2.564 | 2.6471 | 2.675 | 2.85
| Std Dev. |  | 0.010328 | 0.042022 | 0.041312 | 0.12576 | 0.081648 | 0.02016 | 0.058793 | 0.034448 | 0.087274 | 0.03656 | 0.058452 | 0.087636 | 0.053991 | 0.15047 | 0.22941 | 0.040402 | 0.04761 | 0.018065 | 0.02137 | 0.034303 | 0.0083666 | 0.022338 | 0.048782 | 0.045019 | 0.022286
| Rel.Std.Dev. |  | 0.3849% | 1.637% | 1.573% | 5.014% | 3.244% | 0.7854% | 2.155% | 1.292% | 3.279% | 1.379% | 2.191% | 3.295% | 2.031% | 5.85% | 9.345% | 1.52% | 1.788% | 0.6822% | 0.8214% | 1.3% | 0.3187% | 0.8706% | 1.834% | 1.671% | 0.7815%
| PDM3 |  | 1.593% | -2.818% | -0.5527% | -5.033% | -4.697% | -2.814% | 3.296% | 0.9617% | 0.7724% | 0.3938% | 1.025% | 0.7093% | 0.6462% | -2.617% | -7.052% | 0.6349% | 0.8355% | 0.2582% | -1.499% | -0.111% | -0.6158% | -2.856% | 0.6847% | 1.971% | 7.966%
| Z-Score (Absolute) |  | 0.8 | 2.15 | 0.63 | 3.62 | 3.4 | 2.14 | 1.94 | 0.38 | 0.25 | 0 | 0.42 | 0.21 | 0.17 | 2.01 | 4.97 | 0.16 | 0.3 | 0.09 | 1.26 | 0.34 | 0.67 | 2.17 | 0.19 | 1.05 | 5.06
| Analytical results for Cd in OREAS 163b (Certified Value 0.45 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.43 | 0.5 | < 1 | 0.46 | 0.4133 | 0.49 | 0.43 | 0.41 | 0.45 | 0.45 | < 0.3 | 0.49 | 0.45532 | 3 | 0.46 | 0.45 | 0.6 | 0.5 | 0.47 | 0.39 | 0.48 | 0.44 | 0.4 |  |  | 1
| 1 | 2 | 0.44 | 0.5 | < 1 | 0.43 | 0.42999 | 0.52 | 0.46 | 0.45 | 0.45 | 0.46 | < 0.3 | 0.47 | 0.45881 | 2.9 | 0.47 | 0.37 | 0.44 | 0.4 | 0.48 | 0.42 | 0.46 | 0.43 | 0.4 |  |  | 11
| 1 | 3 | 0.43 | 0.5 | < 1 | 0.44 | 0.42615 | 0.52 | 0.5 | 0.46 | 0.43 | 0.47 | 0.5 | 0.47 | 0.48598 | 2.8 | 0.51 | 0.44 | 0.49 | 0.4 | 0.47 | 0.4 | 0.5 | 0.45 | 0.4 |  |  | 16
| 1 | 4 | 0.41 | 0.5 | < 1 | 0.44 | 0.3966 | 0.51 | 0.5 | 0.46 | 0.46 | 0.43 | 0.4 | 0.48 | 0.44754 | 3.1 | 0.47 | 0.42 | 0.47 | 0.4 | 0.48 | 0.38 | 0.49 | 0.44 | 0.4 |  |  | 0.45372
| 1 | 5 | 0.4 | 0.5 | < 1 | 0.45 | 0.4174 | 0.48 | 0.47 | 0.42 | 0.42 | 0.47 | < 0.3 | 0.5 | 0.46032 | 2.3 | 0.48 | 0.41 | 0.54 | 0.5 | 0.46 | 0.42 | 0.47 | 0.44 | 0.4 |  |  | 14
| 1 | 6 | 0.42 | 0.5 | < 1 | 0.45 | 0.41074 | 0.48 | 0.48 | 0.45 | 0.46 | 0.43 | 0.5 | 0.47 | 0.48716 | 2.8 | 0.51 | 0.44 | 0.46 | 0.5 | 0.45 | 0.39 | 0.48 | 0.47 | 0.4
| Mean |  | 0.42167 | 0.5 |  | 0.445 | 0.4157 | 0.5 | 0.47333 | 0.44167 | 0.445 | 0.45167 | 0.46667 | 0.48 | 0.46585 | 2.8167 | 0.48333 | 0.42167 | 0.5 | 0.45 | 0.46833 | 0.4 | 0.48 | 0.445 | 0.4
| Median |  | 0.425 | 0.5 |  | 0.445 | 0.41535 | 0.5 | 0.475 | 0.45 | 0.45 | 0.455 | 0.5 | 0.475 | 0.45956 | 2.85 | 0.475 | 0.43 | 0.48 | 0.45 | 0.47 | 0.395 | 0.48 | 0.44 | 0.4
| Std Dev. |  | 0.01472 | 0 |  | 0.010488 | 0.011928 | 0.018974 | 0.026583 | 0.02137 | 0.016432 | 0.018348 | 0.057735 | 0.012649 | 0.016646 | 0.27869 | 0.021602 | 0.029269 | 0.059666 | 0.054772 | 0.01169 | 0.016733 | 0.014142 | 0.013784 | 6.0809e-17
| Rel.Std.Dev. |  | 3.491% | 0% |  | 2.357% | 2.87% | 3.795% | 5.616% | 4.838% | 3.693% | 4.062% | 12.37% | 2.635% | 3.573% | 9.894% | 4.469% | 6.941% | 11.93% | 12.17% | 2.496% | 4.183% | 2.946% | 3.098% | 1.52e-14%
| PDM3 |  | -7.064% | 10.2% |  | -1.922% | -8.38% | 10.2% | 4.323% | -2.656% | -1.922% | -0.4523% | 2.854% | 5.792% | 2.675% | 520.8% | 6.527% | -7.064% | 10.2% | -0.8196% | 3.221% | -11.84% | 5.792% | -1.922% | -11.84%
| Z-Score (Absolute) |  | 1.43 | NA | 1.1 | 0.67 | 1.62 | 1.1 | 0.24 | 0.78 | 0.67 | 0.46 | NA | 0.46 | 0 | 76.01 | 0.57 | 1.43 | 1.1 | NA | 0.08 | 2.13 | 0.46 | 0.67 | NA
Results from laboratories 2.03, 2.12, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 163b (Certified Value 9.99 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 9.3384 | 11 | 7.5 | 9.74 | 9.5701 | 10.1 | 9.48 | 10.1 | 10.4 | 9.72 | 10 | 9.38 | 9.7 | 9.71 | 10.6 | 11 | 9.87 | 8.92 | 9.9 | 10.34 | 10.3 | 7 |  |  |  | 1
| 1 | 2 | 9.3221 | 11 | 8.8 | 10.13 | 10.0113 | 10.35 | 9.68 | 11.25 | 10.65 | 10.2 | 9.7 | 9.77 | 8.44 | 9.89 | 9.81 | 11 | 9.95 | 8.92 | 9.9 | 10.58 | 10.18 | 6 |  |  |  | 31
| 1 | 3 | 9.6007 | 11 | 7.5 | 10.1 | 10.6971 | 10.7 | 9.49 | 11.05 | 10.55 | 9.67 | 9.9 | 9.57 | 8.11 | 10.14 | 9.52 | 10 | 9.67 | 8.83 | 10 | 10.06 | 10.61 | 4 |  |  |  | 16
| 1 | 4 | 9.5015 | 11 | 8.8 | 9.61 | 10.0029 | 10.85 | 9.51 | 11 | 10.3 | 9.66 | 9.1 | 9.68 | 9.61 | 9.74 | 10.62 | 11 | 10.34 | 8.51 | 9.6 | 9.88 | 10.07 | 5 |  |  |  | 9.9871
| 1 | 5 | 9.2002 | 11 | 8.4 | 9.9 | 10.2396 | 10.1 | 10.15 | 10.25 | 10.3 | 10.2 | 9.6 | 10.13 | 8.79 | 9.79 | 10.4 | 10 | 9.34 | 8.79 | 9.6 | 10.15 | 10.17 | 6 |  |  |  | 15
| 1 | 6 | 9.4446 | 11 | 9.3 | 9.99 | 10.377 | 9.92 | 9.61 | 10.8 | 10.25 | 9.88 | 9.7 | 9.96 | 7.8 | 10.18 | 9.52 | 10 | 9.85 | 9.35 | 9.9 | 10.16 | 10.04 | 5
| Mean |  | 9.4013 | 11 | 8.3833 | 9.9117 | 10.1497 | 10.3367 | 9.6533 | 10.7417 | 10.4083 | 9.8883 | 9.6667 | 9.7483 | 8.7417 | 9.9083 | 10.0783 | 10.5 | 9.8367 | 8.8867 | 9.8167 | 10.195 | 10.2283 | 5.5
| Median |  | 9.3915 | 11 | 8.6 | 9.945 | 10.1255 | 10.225 | 9.56 | 10.9 | 10.35 | 9.8 | 9.7 | 9.725 | 8.615 | 9.84 | 10.105 | 10.5 | 9.86 | 8.875 | 9.9 | 10.155 | 10.175 | 5.5
| Std Dev. |  | 0.14307 | 0 | 0.7414 | 0.20488 | 0.38368 | 0.36914 | 0.25555 | 0.46413 | 0.15943 | 0.25396 | 0.31411 | 0.26947 | 0.78121 | 0.20469 | 0.5224 | 0.54772 | 0.32922 | 0.27252 | 0.17224 | 0.24081 | 0.20837 | 1.0488
| Rel.Std.Dev. |  | 1.522% | 0% | 8.844% | 2.067% | 3.78% | 3.571% | 2.647% | 4.321% | 1.532% | 2.568% | 3.249% | 2.764% | 8.937% | 2.066% | 5.183% | 5.216% | 3.347% | 3.067% | 1.755% | 2.362% | 2.037% | 19.07%
| PDM3 |  | -5.866% | 10.14% | -16.06% | -0.7553% | 1.628% | 3.5% | -3.342% | 7.555% | 4.218% | -0.9889% | -3.208% | -2.391% | -12.47% | -0.7887% | 0.9135% | 5.136% | -1.506% | -11.02% | -1.707% | 2.082% | 2.415% | -44.93%
| Z-Score (Absolute) |  | 1.26 | NA | 3.88 | 0.06 | 0.67 | 1.16 | 0.61 | 2.2 | 1.34 | 0 | 0.57 | 0.36 | 2.96 | 0.05 | 0.49 | NA | 0.13 | 2.58 | 0.18 | 0.79 | 0.88 | NA
Results from laboratories 2.03, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 163b (Certified Value 237 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 236 | 244.2911 | 244 | 196.5 | 236.2 | 237.0233 | 244 | 234 | 232 | 239 | 227 | 241 | 209 | 242.0768 | 227 | 233.3 | 252.2 | 236.6 | 243.3 | 241 | 219 | 228 | 246 | 232.6 | 228.7 | 1
| 1 | 2 | 235 | 241.909 | 246 | 183.2 | 241.4 | 235.1746 | 245 | 235 | 248 | 240 | 229 | 240 | 209 | 244.3265 | 225 | 236.7 | 250.8 | 238.1 | 238.9 | 238 | 233 | 230 | 247.3 | 231.2 | 234.5 | 32
| 1 | 3 | 235 | 245.3719 | 240 | 185.4 | 245.8 | 244.1761 | 241 | 239 | 245 | 243 | 223 | 238 | 230 | 240.9225 | 215 | 232.3 | 250.2 | 236.1 | 238.3 | 242 | 227 | 228 | 244.4 | 239.4 | 232.3 | 16
| 1 | 4 | 231 | 245.3902 | 244 | 204 | 231.5 | 241.1959 | 236 | 236 | 244 | 233 | 224 | 238 | 233 | 242.869 | 234 | 232.8 | 255.9 | 235.3 | 240 | 241 | 220 | 221 | 239.7 | 229.9 | 230.8 | 236.518
| 1 | 5 | 234 | 243.2929 | 240 | 178.9 | 240.7 | 236.756 | 250 | 237 | 231 | 236 | 234 | 244 | 223 | 241.9239 | 171 | 229.2 | 243.6 | 238.3 | 243.2 | 242 | 219 | 222 | 242.3 | 229.8 | 234.8 | 16
| 1 | 6 | 236 | 243.4257 | 244 | 185.7 | 240.5 | 244.5332 | 240 | 236 | 238 | 234 | 238 | 243 | 215 | 243.7288 | 212 | 237.5 | 241.9 | 234.9 | 238.2 | 237 | 231 | 228 | 240.1 | 241.2 | 227.4
| Mean |  | 234.5 | 243.9468 | 243 | 188.95 | 239.35 | 239.8099 | 242.6667 | 236.1667 | 239.6667 | 237.5 | 229.1667 | 240.6667 | 219.8333 | 242.6412 | 214 | 233.6333 | 249.1 | 236.55 | 240.3167 | 240.1667 | 224.8333 | 226.1667 | 243.3 | 234.0167 | 231.4167
| Median |  | 235 | 243.8584 | 244 | 185.55 | 240.6 | 239.1096 | 242.5 | 236 | 241 | 237.5 | 228 | 240.5 | 219 | 242.4729 | 220 | 233.05 | 250.5 | 236.35 | 239.45 | 241 | 223.5 | 228 | 243.35 | 231.9 | 231.55
| Std Dev. |  | 1.8708 | 1.3479 | 2.4495 | 9.3887 | 4.9091 | 4.0465 | 4.8028 | 1.7224 | 7.1181 | 3.8341 | 5.8452 | 2.5033 | 10.4387 | 1.2539 | 22.5566 | 3.0526 | 5.3299 | 1.411 | 2.3609 | 2.137 | 6.3377 | 3.7103 | 3.1209 | 5.005 | 3.0209
| Rel.Std.Dev. |  | 0.7978% | 0.5526% | 1.008% | 4.969% | 2.051% | 1.687% | 1.979% | 0.7293% | 2.97% | 1.614% | 2.551% | 1.04% | 4.748% | 0.5168% | 10.54% | 1.307% | 2.14% | 0.5965% | 0.9824% | 0.8898% | 2.819% | 1.641% | 1.283% | 2.139% | 1.305%
| PDM3 |  | -0.8532% | 3.141% | 2.741% | -20.11% | 1.197% | 1.392% | 2.6% | -0.1485% | 1.331% | 0.4152% | -3.108% | 1.754% | -7.054% | 2.589% | -9.521% | -1.22% | 5.32% | 0.01354% | 1.606% | 1.543% | -4.94% | -4.377% | 2.867% | -1.058% | -2.157%
| Z-Score (Absolute) |  | 0.39 | 0.85 | 0.72 | 6.37 | 0.24 | 0.3 | 0.68 | 0.17 | 0.28 | 0 | 1.09 | 0.42 | 2.32 | 0.67 | 3.08 | 0.51 | 1.52 | 0.12 | 0.37 | 0.35 | 1.66 | 1.49 | 0.76 | 0.46 | 0.8
| Analytical results for Cr in OREAS 163b (Certified Value 12.9 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
1
| 1 | 1 | 15 | 15.182 | 12.4 | < 0.5 | 13 | 12.7567 | 14 | 11 | 12 | 12 | 12 | 19 | 15 | 11.9862 | 10 | 15 | 17 | 9 | 12 | 12 | 9.1 | 13 | 14 | 17 |  | 1
| 1 | 2 | 13 | 15.072 | 12.5 | < 0.5 | 13 | 13.4357 | 13 | 12 | 9 | 12 | 12 | 21 | 13 | 13.7779 | 10 | 14 | 16 | 10 | 14 | 11 | 10.5 | 13 | 12 | 13 |  | 33
| 1 | 3 | 14 | 15.864 | 12 | < 0.5 | 14 | 12.0386 | 20 | 12 | 21 | 12 | 12 | 18 | 13 | 13.9108 | 30 | 13 | 15 | 10 | 13 | 12 | 9.2 | 16 | 12 | 14 |  | 16
| 1 | 4 | 11 | 15.468 | 12.5 | < 0.5 | 13 | 12.0459 | 12 | 17 | 9 | 12 | 13 | 18 | 16 | 12.6885 | 11 | 15 | 16 | 10 | 12 | 12 | 9.9 | 13 | 12 | 17 |  | 12.8579
| 1 | 5 | 12 | 16.062 | 12.5 | < 0.5 | 13 | 11.9249 | 13 | 11 | 11 | 14 | 13 | 18 | 12 | 13.0066 | 6 | 16 | 16 | 11 | 14 | 12 | 11.1 | 14 | 13 | 14 |  | 17
| 1 | 6 | 12 | 15.446 | 12.4 | < 0.5 | 13 | 12.9449 | 13 | 13 | 10 | 13 | 14 | 17 | 15 | 14.4885 | 8 | 15 | 15 | 10 | 13 | 11 | 9.1 | 23 | 13 | 13
| Mean |  | 12.8333 | 15.5157 | 12.3833 |  | 13.1667 | 12.5245 | 14.1667 | 12.6667 | 12 | 12.5 | 12.6667 | 18.5 | 14 | 13.3098 | 12.5 | 14.6667 | 15.8333 | 10 | 13 | 11.6667 | 9.8167 | 15.3333 | 12.6667 | 14.6667
| Median |  | 12.5 | 15.457 | 12.45 |  | 13 | 12.4013 | 13 | 12 | 10.5 | 12 | 12.5 | 18 | 14 | 13.3923 | 10 | 15 | 16 | 10 | 13 | 12 | 9.55 | 13.5 | 12.5 | 14
| Std Dev. |  | 1.472 | 0.38344 | 0.19408 |  | 0.40825 | 0.61411 | 2.9269 | 2.2509 | 4.5607 | 0.83666 | 0.8165 | 1.3784 | 1.5492 | 0.91648 | 8.7579 | 1.0328 | 0.75277 | 0.63246 | 0.89443 | 0.5164 | 0.84004 | 3.9328 | 0.8165 | 1.8619
| Rel.Std.Dev. |  | 11.47% | 2.471% | 1.567% |  | 3.101% | 4.903% | 20.66% | 17.77% | 38.01% | 6.693% | 6.446% | 7.451% | 11.07% | 6.886% | 70.06% | 7.042% | 4.754% | 6.325% | 6.88% | 4.426% | 8.557% | 25.65% | 6.446% | 12.69%
| PDM3 |  | -0.1913% | 20.67% | -3.691% |  | 2.401% | -2.593% | 10.18% | -1.488% | -6.672% | -2.784% | -1.488% | 43.88% | 8.882% | 3.514% | -2.784% | 14.07% | 23.14% | -22.23% | 1.105% | -9.265% | -23.65% | 19.25% | -1.488% | 14.07%
| Z-Score (Absolute) |  | 0.06 | 2.03 | 0.27 | 9.2 | 0.31 | 0.17 | 1.04 | 0.06 | 0.55 | 0.18 | 0.06 | 4.23 | 0.92 | 0.41 | 0.18 | 1.41 | 2.27 | 2.02 | 0.18 | 0.8 | 2.16 | 1.9 | 0.06 | 1.41
| Analytical results for Cs in OREAS 163b (Certified Value 0.97 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1 | 1.0046 | 0.99 | 0.99 | 0.94206 | 0.92 | 0.9 | 0.91 | 0.96 | 0.88 | 1.02 | 0.95 | 0.964 | 0.91 | 0.99 | 1 | < 1 | 0.92 | 0.98 | 0.91 | 0.9 |  |  |  |  | 1
| 1 | 2 | 1 | 1.0479 | 0.99 | 0.94 | 1.0175 | 0.93 | 1 | 0.96 | 0.95 | 0.91 | 1.07 | 0.94 | 0.9744 | 0.75 | 1.04 | 0.9 | < 1 | 1 | 1.04 | 0.9 | 1 |  |  |  |  | 13
| 1 | 3 | 1 | 1.0146 | 1.02 | 0.97 | 1.0222 | 1 | 1 | 1 | 0.96 | 0.88 | 1.07 | 0.96 | 1.0391 | 0.79 | 1.03 | 1 | < 1 | 0.94 | 1.02 | 0.92 | 0.9 |  |  |  |  | 16
| 1 | 4 | 1 | 1.0129 | 0.982 | 0.98 | 0.97627 | 0.97 | 1 | 1.02 | 0.91 | 0.87 | 0.91 | 1.02 | 1.0251 | 0.94 | 1.01 | 1 | < 1 | 0.92 | 0.97 | 0.94 | 1 |  |  |  |  | 0.96904
| 1 | 5 | 1 | 1.0337 | 0.99 | 0.95 | 0.99456 | 0.93 | 0.91 | 0.93 | 0.96 | 0.91 | 0.97 | 0.97 | 1.0366 | 0.82 | 0.94 | 1 | < 1 | 0.83 | 0.99 | 0.93 | 1 |  |  |  |  | 18
| 1 | 6 | 1 | 1.0493 | 1.026 | 0.96 | 0.98218 | 0.91 | 0.91 | 0.95 | 0.94 | 0.89 | 1.01 | 0.96 | 1.0234 | 0.71 | 0.99 | 1 | < 1 | 0.92 | 0.98 | 0.97 | 1
| Mean |  | 1 | 1.0272 | 0.99967 | 0.965 | 0.98913 | 0.94333 | 0.95333 | 0.96167 | 0.94667 | 0.89 | 1.0083 | 0.96667 | 1.0104 | 0.82 | 1 | 0.98333 |  | 0.92167 | 0.99667 | 0.92833 | 0.96667
| Median |  | 1 | 1.0241 | 0.99 | 0.965 | 0.98837 | 0.93 | 0.955 | 0.955 | 0.955 | 0.885 | 1.015 | 0.96 | 1.0242 | 0.805 | 1 | 1 |  | 0.92 | 0.985 | 0.925 | 1
| Std Dev. |  | 0 | 0.019141 | 0.018435 | 0.018708 | 0.029524 | 0.034448 | 0.051251 | 0.041673 | 0.019664 | 0.016733 | 0.061455 | 0.028048 | 0.03269 | 0.089889 | 0.035777 | 0.040825 |  | 0.054559 | 0.027325 | 0.024833 | 0.05164
| Rel.Std.Dev. |  | 0% | 1.863% | 1.844% | 1.939% | 2.985% | 3.652% | 5.376% | 4.333% | 2.077% | 1.88% | 6.095% | 2.901% | 3.235% | 10.96% | 3.578% | 4.152% |  | 5.92% | 2.742% | 2.675% | 5.342%
| PDM3 |  | 3.194% | 5.997% | 3.16% | -0.4174% | 2.073% | -2.653% | -1.621% | -0.7613% | -2.309% | -8.157% | 4.054% | -0.2454% | 4.27% | -15.38% | 3.194% | 1.475% |  | -4.889% | 2.85% | -4.201% | -0.2454%
| Z-Score (Absolute) |  | NA | 1.21 | 0.69 | 0.03 | 0.49 | 0.38 | 0.19 | 0.03 | 0.32 | 1.39 | 0.85 | 0.06 | 0.89 | 2.71 | 0.69 | NA | 8.76 | 0.79 | 0.63 | 0.66 | NA
Results from laboratories 2.01, 2.17 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 163b (Certified Value 1.68 wt.%) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*AAS | 4A*MS | 4A*OES | 4A*OES | 4A*AAS | 4A*AAS | 4A*MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 1.73 | 1.6707 | 1.668 | 1.67 | 1.5954 | 1.7191 | 1.67 | 1.66 | 1.705 | 1.685 | 1.69 | 1.71 | 1.71 | 1.6324 | 1.8 | 1.6554 | 1.692 | 1.6847 | 1.706 | 1.71 | 1.72 | 1.679 | 1.6848 | 1.659 | 1.65 | 1
| 1 | 2 | 1.74 | 1.6369 | 1.688 | 1.64 | 1.649 | 1.7154 | 1.68 | 1.665 | 1.695 | 1.675 | 1.715 | 1.67 | 1.7 | 1.6901 | 1.72 | 1.6673 | 1.682 | 1.6727 | 1.704 | 1.71 | 1.74 | 1.687 | 1.7053 | 1.653 | 1.65 | 35
| 1 | 3 | 1.74 | 1.647 | 1.664 | 1.63 | 1.6268 | 1.7119 | 1.68 | 1.65 | 1.64 | 1.7 | 1.685 | 1.72 | 1.7 | 1.7775 | 1.78 | 1.6931 | 1.701 | 1.6688 | 1.726 | 1.7 | 1.74 | 1.679 | 1.6787 | 1.652 | 1.65 | 16
| 1 | 4 | 1.74 | 1.6692 | 1.638 | 1.63 | 1.58 | 1.7105 | 1.675 | 1.675 | 1.655 | 1.66 | 1.675 | 1.7 | 1.7 | 1.7341 | 1.72 | 1.6307 | 1.68 | 1.6812 | 1.663 | 1.71 | 1.71 | 1.676 | 1.6733 | 1.703 | 1.66 | 1.6834
| 1 | 5 | 1.74 | 1.6607 | 1.687 | 1.62 | 1.6056 | 1.7147 | 1.665 | 1.655 | 1.72 | 1.695 | 1.675 | 1.67 | 1.7 | 1.6499 | 1.67 | 1.6186 | 1.668 | 1.68 | 1.683 | 1.71 | 1.77 | 1.678 | 1.6822 | 1.685 | 1.66 | 19
| 1 | 6 | 1.73 | 1.6471 | 1.675 | 1.62 | 1.6272 | 1.7205 | 1.66 | 1.66 | 1.685 | 1.68 | 1.705 | 1.68 | 1.71 | 1.7717 | 1.8 | 1.6677 | 1.675 | 1.6781 | 1.705 | 1.71 | 1.73 | 1.678 | 1.6484 | 1.695 | 1.64
| Mean |  | 1.7367 | 1.6553 | 1.67 | 1.635 | 1.614 | 1.7154 | 1.6717 | 1.6608 | 1.6833 | 1.6825 | 1.6908 | 1.6917 | 1.7033 | 1.7093 | 1.7483 | 1.6555 | 1.683 | 1.6776 | 1.6978 | 1.7083 | 1.735 | 1.6795 | 1.6788 | 1.6745 | 1.6517
| Median |  | 1.74 | 1.6539 | 1.6715 | 1.63 | 1.6162 | 1.715 | 1.6725 | 1.66 | 1.69 | 1.6825 | 1.6875 | 1.69 | 1.7 | 1.7121 | 1.75 | 1.6613 | 1.681 | 1.679 | 1.7045 | 1.71 | 1.735 | 1.6785 | 1.6805 | 1.672 | 1.65
| Std Dev. |  | 0.005164 | 0.013664 | 0.01845 | 0.018708 | 0.025048 | 0.0039052 | 0.008165 | 0.008612 | 0.030441 | 0.014405 | 0.016253 | 0.02137 | 0.005164 | 0.061604 | 0.053072 | 0.027124 | 0.011866 | 0.0058389 | 0.02183 | 0.0040825 | 0.020736 | 0.0038341 | 0.018449 | 0.02259 | 0.0075277
| Rel.Std.Dev. |  | 0.2973% | 0.8255% | 1.105% | 1.144% | 1.552% | 0.2277% | 0.4884% | 0.5185% | 1.808% | 0.8562% | 0.9613% | 1.263% | 0.3032% | 3.604% | 3.036% | 1.638% | 0.705% | 0.3481% | 1.286% | 0.239% | 1.195% | 0.2283% | 1.099% | 1.349% | 0.4558%
| PDM3 |  | 3.167% | -1.669% | -0.7934% | -2.873% | -4.121% | 1.901% | -0.6944% | -1.338% | -0.001313% | -0.05082% | 0.4442% | 0.4937% | 1.187% | 1.541% | 3.86% | -1.656% | -0.02111% | -0.3434% | 0.8601% | 1.484% | 3.068% | -0.229% | -0.2718% | -0.5261% | -1.882%
| Z-Score (Absolute) |  | 1.75 | 0.88 | 0.4 | 1.54 | 2.22 | 1.06 | 0.35 | 0.7 | 0.03 | 0 | 0.27 | 0.3 | 0.67 | 0.87 | 2.13 | 0.87 | 0.02 | 0.16 | 0.5 | 0.84 | 1.7 | 0.1 | 0.12 | 0.26 | 1
| Analytical results for Dy in OREAS 163b (Certified Value 0.86 ppm) | [26 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 0.7802; 2.02: 0.934; 2.03: 0.87455; 2.04: 1.2; 2.05: 1.05; 2.06: 1; 2.07: 0.77 | 1
| 1 | 2 | 2.01: 0.8075; 2.02: 0.938; 2.03: 0.88218; 2.04: 1.2; 2.05: 0.85; 2.06: 1; 2.07: 0.76 | 3
| 1 | 3 | 2.01: 0.7576; 2.02: 0.935; 2.03: 0.95719; 2.04: 1.4; 2.05: 1.35; 2.06: 1; 2.07: 0.81 | 16
| 1 | 4 | 2.01: 0.756; 2.02: 0.934; 2.03: 0.90066; 2.04: 1.2; 2.05: 0.83; 2.06: 1.1; 2.07: 0.84 | 0.85778
| 1 | 5 | 2.01: 0.8207; 2.02: 0.934; 2.03: 0.87489; 2.04: 1.3; 2.05: 0.81; 2.06: 1; 2.07: 0.79 | 20
| 1 | 6 | 2.01: 0.7591; 2.02: 0.936; 2.03: 0.88574; 2.04: 1.2; 2.05: 0.89; 2.06: 0.9; 2.07: 0.78
| Mean |  | 2.01: 0.78018; 2.02: 0.93517; 2.03: 0.89587; 2.04: 1.25; 2.05: 0.96333; 2.06: 1; 2.07: 0.79167
| Median |  | 2.01: 0.76965; 2.02: 0.9345; 2.03: 0.88396; 2.04: 1.2; 2.05: 0.87; 2.06: 1; 2.07: 0.785
| Std Dev. |  | 2.01: 0.028026; 2.02: 0.0016021; 2.03: 0.03152; 2.04: 0.083666; 2.05: 0.2081; 2.06: 0.063246; 2.07: 0.029269
| Rel.Std.Dev. |  | 2.01: 3.592%; 2.02: 0.1713%; 2.03: 3.518%; 2.04: 6.693%; 2.05: 21.6%; 2.06: 6.325%; 2.07: 3.697%
| PDM3 |  | 2.01: -9.046%; 2.02: 9.022%; 2.03: 4.441%; 2.04: 45.73%; 2.05: 12.31%; 2.06: 16.58%; 2.07: -7.707%
| Z-Score (Absolute) |  | 2.01: 1.16; 2.02: 0.39; 2.03: 0; 2.04: NA; 2.05: 0.67; 2.06: NA; 2.07: 1.04
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 163b (Certified Value 0.38 ppm) | [26 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 0.3024; 2.02: 0.4; 2.03: 0.35302; 2.04: 0.5; 2.05: 0.37; 2.06: 0.4; 2.07: 0.29 | 1
| 1 | 2 | 2.01: 0.3217; 2.02: 0.4; 2.03: 0.4038; 2.04: 0.5; 2.05: 0.36; 2.06: 0.4; 2.07: 0.35 | 4
| 1 | 3 | 2.01: 0.3478; 2.02: 0.4; 2.03: 0.42481; 2.04: 0.7; 2.05: 0.37; 2.06: 0.4; 2.07: 0.36 | 16
| 1 | 4 | 2.01: 0.3109; 2.02: 0.4; 2.03: 0.37645; 2.04: 0.6; 2.05: 0.38; 2.06: 0.4; 2.07: 0.31 | 0.3766
| 1 | 5 | 2.01: 0.3539; 2.02: 0.4; 2.03: 0.37277; 2.04: 0.6; 2.05: 0.38; 2.06: 0.4; 2.07: 0.32 | 21
| 1 | 6 | 2.01: 0.3251; 2.02: 0.4; 2.03: 0.37544; 2.04: 0.6; 2.05: 0.37; 2.06: 0.4; 2.07: 0.31
| Mean |  | 2.01: 0.32697; 2.02: 0.4; 2.03: 0.38438; 2.04: 0.58333; 2.05: 0.37167; 2.06: 0.4; 2.07: 0.32333
| Median |  | 2.01: 0.3234; 2.02: 0.4; 2.03: 0.37595; 2.04: 0.6; 2.05: 0.37; 2.06: 0.4; 2.07: 0.315
| Std Dev. |  | 2.01: 0.02026; 2.02: 6.0809e-17; 2.03: 0.025585; 2.04: 0.075277; 2.05: 0.0075277; 2.06: 6.0809e-17; 2.07: 0.026583
| Rel.Std.Dev. |  | 2.01: 6.196%; 2.02: 1.52e-14%; 2.03: 6.656%; 2.04: 12.9%; 2.05: 2.025%; 2.06: 1.52e-14%; 2.07: 8.222%
| PDM3 |  | 2.01: -13.18%; 2.02: 6.213%; 2.03: 2.066%; 2.04: 54.89%; 2.05: -1.311%; 2.06: 6.213%; 2.07: -14.14%
| Z-Score (Absolute) |  | 2.01: 2.48; 2.02: 0.67; 2.03: 0; 2.04: 8.59; 2.05: 0.55; 2.06: 0.67; 2.07: 2.64
| Analytical results for Eu in OREAS 163b (Certified Value 0.16 ppm) | [26 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 0.1096; 2.02: 0.15; 2.03: 0.16251; 2.04: 0.17; 2.05: 0.16; 2.06: 0.2; 2.07: 0.16 | 1
| 1 | 2 | 2.01: 0.1293; 2.02: 0.15; 2.03: 0.17274; 2.04: 0.18; 2.05: 0.15; 2.06: 0.2; 2.07: 0.14 | 5
| 1 | 3 | 2.01: 0.108; 2.02: 0.15; 2.03: 0.16699; 2.04: 0.16; 2.05: 0.15; 2.06: 0.2; 2.07: 0.17 | 16
| 1 | 4 | 2.01: 0.1058; 2.02: 0.15; 2.03: 0.16957; 2.04: 0.17; 2.05: 0.16; 2.06: 0.2; 2.07: 0.15 | 0.15976
| 1 | 5 | 2.01: 0.1352; 2.02: 0.15; 2.03: 0.16604; 2.04: 0.17; 2.05: 0.17; 2.06: 0.2; 2.07: 0.13 | 22
| 1 | 6 | 2.01: 0.1124; 2.02: 0.15; 2.03: 0.16507; 2.04: 0.2; 2.05: 0.18; 2.06: 0.2; 2.07: 0.15
| Mean |  | 2.01: 0.11672; 2.02: 0.15; 2.03: 0.16715; 2.04: 0.175; 2.05: 0.16167; 2.06: 0.2; 2.07: 0.15
| Median |  | 2.01: 0.111; 2.02: 0.15; 2.03: 0.16651; 2.04: 0.17; 2.05: 0.16; 2.06: 0.2; 2.07: 0.15
| Std Dev. |  | 2.01: 0.012365; 2.02: 0; 2.03: 0.0035825; 2.04: 0.013784; 2.05: 0.01169; 2.06: 3.0405e-17; 2.07: 0.014142
| Rel.Std.Dev. |  | 2.01: 10.59%; 2.02: 0%; 2.03: 2.143%; 2.04: 7.877%; 2.05: 7.231%; 2.06: 1.52e-14%; 2.07: 9.428%
| PDM3 |  | 2.01: -26.94%; 2.02: -6.112%; 2.03: 4.625%; 2.04: 9.537%; 2.05: 1.191%; 2.06: 25.18%; 2.07: -6.112%
| Z-Score (Absolute) |  | 2.01: 3.08; 2.02: 0.46; 2.03: 0.89; 2.04: 1.51; 2.05: 0.46; 2.06: NA; 2.07: 0.46
Results from laboratory 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 163b (Certified Value 4.1 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 4.19 | 4.2341 | 4.24 | 4.39 | 3.89 | 4.1258 | 4.1 | 3.91 | 3.81 | 3.96 | 3.87 | 4.8 | 4 | 4.0517 | 4.28 | 4.25 | 4.12 | 4.0103 | 3.97 | 4.25 | 4.16 | 4.02 | 3.98 | 4.14 | 4.27 | 1
| 1 | 2 | 4.2 | 4.2145 | 4.26 | 4.03 | 4.06 | 4.0895 | 4.15 | 3.93 | 4.09 | 3.97 | 3.94 | 4.51 | 3.98 | 4.1267 | 4.24 | 4.15 | 4.1 | 3.9728 | 3.96 | 4.21 | 4.16 | 3.95 | 4.03 | 4.12 | 4.26 | 0
| 1 | 3 | 4.19 | 4.2515 | 4.26 | 3.97 | 4.01 | 4.0718 | 4.13 | 4.02 | 4.08 | 4 | 3.85 | 4.52 | 4.21 | 4.0139 | 4.03 | 4.1 | 4.15 | 3.9805 | 3.89 | 4.18 | 4.16 | 3.97 | 3.91 | 4.27 | 4.25 | 16
| 1 | 4 | 4.21 | 4.247 | 4.28 | 4.33 | 3.84 | 4.085 | 3.98 | 3.95 | 4.05 | 3.87 | 3.84 | 4.17 | 4.06 | 4.1948 | 4.44 | 4.2 | 4.18 | 3.9956 | 3.95 | 4.2 | 4.17 | 4.02 | 3.91 | 4.12 | 4.25 | 4.0998
| 1 | 5 | 4.21 | 4.2249 | 4.28 | 3.92 | 4.04 | 4.0899 | 4.17 | 4 | 3.86 | 3.91 | 4.01 | 4.34 | 4.16 | 4.0033 | 3.28 | 4.19 | 4.04 | 3.9762 | 3.94 | 4.19 | 4.17 | 4 | 4 | 4.14 | 4.29 | 23
| 1 | 6 | 4.19 | 4.2641 | 4.28 | 3.98 | 3.99 | 3.9846 | 4.03 | 3.95 | 3.98 | 3.88 | 4.06 | 4.51 | 4.07 | 4.1626 | 4.1 | 4.1 | 4.04 | 3.9829 | 3.95 | 4.2 | 4.17 | 3.93 | 3.93 | 4.27 | 4.25
| Mean |  | 4.1983 | 4.2393 | 4.2667 | 4.1033 | 3.9717 | 4.0744 | 4.0933 | 3.96 | 3.9783 | 3.9317 | 3.9283 | 4.475 | 4.08 | 4.0922 | 4.0617 | 4.165 | 4.105 | 3.9864 | 3.9433 | 4.205 | 4.165 | 3.9817 | 3.96 | 4.1767 | 4.2617
| Median |  | 4.195 | 4.2405 | 4.27 | 4.005 | 4 | 4.0872 | 4.115 | 3.95 | 4.015 | 3.935 | 3.905 | 4.51 | 4.065 | 4.0892 | 4.17 | 4.17 | 4.11 | 3.9817 | 3.95 | 4.2 | 4.165 | 3.985 | 3.955 | 4.14 | 4.255
| Std Dev. |  | 0.0098319 | 0.01829 | 0.01633 | 0.20275 | 0.087502 | 0.047524 | 0.073937 | 0.041952 | 0.11856 | 0.052694 | 0.091086 | 0.2104 | 0.089666 | 0.080426 | 0.40892 | 0.059582 | 0.057184 | 0.014085 | 0.028048 | 0.02429 | 0.0054772 | 0.037639 | 0.050596 | 0.072847 | 0.016021
| Rel.Std.Dev. |  | 0.2342% | 0.4314% | 0.3827% | 4.941% | 2.203% | 1.166% | 1.806% | 1.059% | 2.98% | 1.34% | 2.319% | 4.702% | 2.198% | 1.965% | 10.07% | 1.431% | 1.393% | 0.3533% | 0.7113% | 0.5776% | 0.1315% | 0.9453% | 1.278% | 1.744% | 0.3759%
| PDM3 |  | 2.403% | 3.404% | 4.07% | 0.08587% | -3.126% | -0.6192% | -0.158% | -3.41% | -2.963% | -4.101% | -4.183% | 9.151% | -0.4833% | -0.1867% | -0.9304% | 1.59% | 0.1265% | -2.767% | -3.817% | 2.566% | 1.59% | -2.882% | -3.41% | 1.875% | 3.948%
| Z-Score (Absolute) |  | 0.65 | 0.9 | 1.06 | 0.07 | 0.74 | 0.11 | 0.01 | 0.81 | 0.69 | 0.98 | 1 | 2.34 | 0.07 | 0 | 0.19 | 0.44 | 0.08 | 0.65 | 0.91 | 0.69 | 0.44 | 0.67 | 0.81 | 0.52 | 1.03
| Analytical results for Ga in OREAS 163b (Certified Value 1.47 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 1.5 | 1.4699 | 1.7 | < 5 | 1.33 | 1.5552 | 1.48 | 1.39 | 1.33 | 1.4 | 1.32 | 2.1 | 1.33 | 1.5778 | 12.3 | 1.51 | 1.58 | < 5 | 1.5 | 1.5 | 1.66 | 1.47 | 1.59 | 2 |  | 1
| 1 | 2 | 1.5 | 1.4649 | 1.7 | < 5 | 1.21 | 1.4467 | 1.54 | 1.54 | 1.46 | 1.44 | 1.33 | 1.8 | 1.4 | 1.6277 | 8.33 | 1.24 | 1.41 | < 5 | 1.6 | 1.6 | 1.64 | 1.5 | 1.59 | 1 |  | 17
| 1 | 3 | 1.5 | 1.4635 | 1.7 | < 5 | 1.23 | 1.4233 | 1.61 | 1.55 | 1.41 | 1.39 | 1.28 | 1.9 | 1.39 | 1.6374 | 7.61 | 1.33 | 1.46 | < 5 | 1.6 | 1.6 | 1.57 | 1.48 | 1.56 | 1 |  | 16
| 1 | 4 | 1.4 | 1.4805 | 1.7 | < 5 | 1.24 | 1.4997 | 1.51 | 1.43 | 1.4 | 1.37 | 1.34 | 1.8 | 1.37 | 1.6195 | 6.23 | 1.34 | 1.68 | < 5 | 1.5 | 1.5 | 1.45 | 1.32 | 1.55 | 1 |  | 1.4653
| 1 | 5 | 1.5 | 1.4976 | 1.7 | < 5 | 1.23 | 1.5202 | 1.47 | 1.56 | 1.33 | 1.38 | 1.42 | 1.7 | 1.46 | 1.587 | 8.16 | 1.38 | 1.46 | < 5 | 1.5 | 1.5 | 1.55 | 1.55 | 1.47 | 1 |  | 24
| 1 | 6 | 1.5 | 1.4917 | 1.7 | < 5 | 1.29 | 1.4509 | 1.46 | 1.52 | 1.39 | 1.37 | 1.28 | 1.8 | 1.39 | 1.6036 | 7.03 | 1.34 | 1.45 | < 5 | 1.5 | 1.5 | 1.63 | 1.42 | 1.55 | 2
| Mean |  | 1.4833 | 1.478 | 1.7 |  | 1.255 | 1.4827 | 1.5117 | 1.4983 | 1.3867 | 1.3917 | 1.3283 | 1.85 | 1.39 | 1.6089 | 8.2767 | 1.3567 | 1.5067 |  | 1.5333 | 1.5333 | 1.5833 | 1.4567 | 1.5517 | 1.3333
| Median |  | 1.5 | 1.4752 | 1.7 |  | 1.235 | 1.4753 | 1.495 | 1.53 | 1.395 | 1.385 | 1.325 | 1.8 | 1.39 | 1.6116 | 7.885 | 1.34 | 1.46 |  | 1.5 | 1.5 | 1.6 | 1.475 | 1.555 | 1
| Std Dev. |  | 0.040825 | 0.014346 | 0 |  | 0.045497 | 0.050544 | 0.056362 | 0.070828 | 0.050067 | 0.026394 | 0.051543 | 0.13784 | 0.042426 | 0.023449 | 2.1162 | 0.088242 | 0.10231 |  | 0.05164 | 0.05164 | 0.077889 | 0.079162 | 0.044008 | 0.5164
| Rel.Std.Dev. |  | 2.752% | 0.9706% | 0% |  | 3.625% | 3.409% | 3.728% | 4.727% | 3.611% | 1.897% | 3.88% | 7.451% | 3.052% | 1.458% | 25.57% | 6.504% | 6.79% |  | 3.368% | 3.368% | 4.919% | 5.434% | 2.836% | 38.73%
| PDM3 |  | 1.232% | 0.8695% | 16.02% |  | -14.35% | 1.186% | 3.166% | 2.256% | -5.365% | -5.024% | -9.346% | 26.26% | -5.138% | 9.799% | 464.9% | -7.412% | 2.825% |  | 4.644% | 4.644% | 8.057% | -0.5878% | 5.896% | -9.005%
| Z-Score (Absolute) |  | 0.12 | 0.15 | 1.23 | 6.19 | 1.54 | 0.12 | 0.06 | 0.03 | 0.72 | 0.69 | 1.08 | 2.16 | 0.7 | 0.66 | 42.06 | 0.91 | 0.03 | 6.19 | 0.19 | 0.19 | 0.5 | 0.28 | 0.31 | 1.05
| Analytical results for Gd in OREAS 163b (Certified Value 1.23 ppm) | [26 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 1.181; 2.02: 1.23; 2.03: 1.1521; 2.04: 1.4; 2.05: 1.25; 2.06: 1.4; 2.07: 1.21 |  | 1
| 1 | 2 | 2.01: 1.1603; 2.02: 1.22; 2.03: 1.2654; 2.04: 1.7; 2.05: 1.24; 2.06: 1.3; 2.07: 1.16 |  | 6
| 1 | 3 | 2.01: 1.174; 2.02: 1.23; 2.03: 1.2456; 2.04: 1.4; 2.05: 1.31; 2.06: 1.4; 2.07: 1.24 |  | 16
| 1 | 4 | 2.01: 1.227; 2.02: 1.22; 2.03: 1.2201; 2.04: 1.5; 2.05: 1.29; 2.06: 1.4; 2.07: 1.3 |  | 1.2306
| 1 | 5 | 2.01: 1.2351; 2.02: 1.23; 2.03: 1.2437; 2.04: 1.5; 2.05: 1.23; 2.06: 1.4; 2.07: 1.22 |  | 25
| 1 | 6 | 2.01: 1.2072; 2.02: 1.22; 2.03: 1.2286; 2.04: 1.4; 2.05: 1.31; 2.06: 1.3; 2.07: 1.18
| Mean |  | 2.01: 1.1974; 2.02: 1.225; 2.03: 1.2259; 2.04: 1.4833; 2.05: 1.2717; 2.06: 1.3667; 2.07: 1.2183
| Median |  | 2.01: 1.1941; 2.02: 1.225; 2.03: 1.2361; 2.04: 1.45; 2.05: 1.27; 2.06: 1.4; 2.07: 1.215
| Std Dev. |  | 2.01: 0.030287; 2.02: 0.0054772; 2.03: 0.039396; 2.04: 0.1169; 2.05: 0.036009; 2.06: 0.05164; 2.07: 0.04916
| Rel.Std.Dev. |  | 2.01: 2.529%; 2.02: 0.4471%; 2.03: 3.214%; 2.04: 7.881%; 2.05: 2.832%; 2.06: 3.779%; 2.07: 4.035%
| PDM3 |  | 2.01: -2.697%; 2.02: -0.4569%; 2.03: -0.383%; 2.04: 20.54%; 2.05: 3.335%; 2.06: 11.05%; 2.07: -0.9986%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.02; 2.03: 0; 2.04: 6.1; 2.05: 1.08; 2.06: 3.33; 2.07: 0.18
| Analytical results for Ge in OREAS 163b (Certified Value 0.067 ppm) | [26 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.23 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | 0.9 | 0.8608 | 0.9 | < 0.2 | 0.05 | 0.05 | 0.06 | 0.06 | < 0.05 | < 0.1 | < 0.1 | 0.18 | 1 | < 0.05 | 0.8 | < 0.05 | 2.1 | [8 empty] | 1
| 1 | 2 | 0.9 | 0.8962 | 0.9 | < 0.2 | < 0.05 | 0.06 | 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.1 | 0.1 | 1 | < 0.05 | 0.8 | < 0.05 | 2.1 | [8 empty] | 19
| 1 | 3 | 0.9 | 0.8869 | 0.9 | < 0.2 | 0.05 | 0.07 | 0.06 | 0.07 | < 0.05 | < 0.1 | < 0.1 | 0.1 | 1.1 | < 0.05 | 0.7 | < 0.05 | 2.1 | [8 empty] | 16
| 1 | 4 | 0.9 | 0.8719 | 0.9 | < 0.2 | < 0.05 | 0.09 | 0.05 | 0.07 | < 0.05 | < 0.1 | < 0.1 | 0.14 | 1 | < 0.05 | 0.8 | < 0.05 | 2.1 | [8 empty] | 0.067381
| 1 | 5 | 0.9 | 0.85065 | 0.9 | < 0.2 | < 0.05 | 0.07 | 0.05 | 0.08 | < 0.05 | < 0.1 | < 0.1 | 0.09 | 1.1 | 0.07 | 0.8 | < 0.05 | 2.2 | [8 empty] | 26
| 1 | 6 | 0.9 | 0.8684 | 0.9 | < 0.2 | < 0.05 | 0.06 | 0.05 | 0.08 | < 0.05 | < 0.1 | < 0.1 | 0.12 | 1.2 | < 0.05 | 0.8 | 0.05 | 2.1
| Mean |  | 0.9 | 0.87248 | 0.9 |  | 0.05 | 0.066667 | 0.053333 | 0.071667 |  |  |  | 0.12167 | 1.0667 | 0.07 | 0.78333 | 0.05 | 2.1167
| Median |  | 0.9 | 0.87015 | 0.9 |  | 0.05 | 0.065 | 0.05 | 0.07 |  |  |  | 0.11 | 1.05 | 0.07 | 0.8 | 0.05 | 2.1
| Std Dev. |  | 0 | 0.016736 | 0 |  | 0 | 0.013663 | 0.005164 | 0.0075277 |  |  |  | 0.033714 | 0.08165 |  | 0.040825 |  | 0.040825
| Rel.Std.Dev. |  | 0% | 1.918% | 0% |  | 0% | 20.49% | 9.682% | 10.5% |  |  |  | 27.71% | 7.655% |  | 5.212% |  | 1.929%
| PDM3 |  | 1236% | 1195% | 1236% |  | -25.8% | -1.06% | -20.85% | 6.36% |  |  |  | 80.57% | 1483% | 3.887% | 1063% | -25.8% | 3041%
| Z-Score (Absolute) |  | 13.14 | 12.71 | 13.14 | 0.45 | 0.61 | 0.08 | 0.29 | 0 | 0.74 | 0.34 | 0.34 | 0.79 | 15.79 | 0.62 | 11.29 | 0.67 | 32.45
| Analytical results for Hf in OREAS 163b (Certified Value 0.37 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.4 | 0.4 | < 0.5 | 0.36 | 0.37313 | 0.4 | 0.4 | 0.3 | 0.4 | 0.4 | 0.4 | 0.34 | 0.38406 | 0.4 | 0.39 | 0.36 | 0.3 | 0.25 | 0.35 | 0.34 | 0.4 | 0.3 |  |  |  | 1
| 1 | 2 | 0.4 | 0.4 | < 0.5 | 0.33 | 0.38898 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.34 | 0.41597 | 0.3 | 0.37 | 0.36 | 0.3 | 0.24 | 0.31 | 0.37 | 0.4 | 0.3 |  |  |  | 23
| 1 | 3 | 0.4 | 0.4 | < 0.5 | 0.37 | 0.35822 | 0.4 | 0.3 | 0.4 | 0.4 | 0.3 | 0.4 | 0.34 | 0.41614 | 0.3 | 0.41 | 0.38 | 0.3 | 0.28 | 0.37 | 0.37 | 0.4 | 0.3 |  |  |  | 16
| 1 | 4 | 0.4 | 0.4 | < 0.5 | 0.38 | 0.36261 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.34 | 0.35273 | 0.3 | 0.38 | 0.41 | 0.3 | 0.25 | 0.32 | 0.39 | 0.4 | 0.3 |  |  |  | 0.36787
| 1 | 5 | 0.4 | 0.4 | < 0.5 | 0.37 | 0.35299 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.5 | 0.35 | 0.37079 | 0.3 | 0.35 | 0.4 | 0.3 | 0.25 | 0.32 | 0.37 | 0.4 | 0.3 |  |  |  | 27
| 1 | 6 | 0.4 | 0.4 | < 0.5 | 0.33 | 0.36595 | 0.4 | 0.4 | 0.4 | 0.3 | 0.3 | 0.4 | 0.34 | 0.44299 | 0.3 | 0.39 | 0.36 | 0.3 | 0.25 | 0.34 | 0.36 | 0.4 | 0.3
| Mean |  | 0.4 | 0.4 |  | 0.35667 | 0.36698 | 0.4 | 0.38333 | 0.38333 | 0.38333 | 0.36667 | 0.41667 | 0.34167 | 0.39711 | 0.31667 | 0.38167 | 0.37833 | 0.3 | 0.25333 | 0.335 | 0.36667 | 0.4 | 0.3
| Median |  | 0.4 | 0.4 |  | 0.365 | 0.36428 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.34 | 0.40001 | 0.3 | 0.385 | 0.37 | 0.3 | 0.25 | 0.33 | 0.37 | 0.4 | 0.3
| Std Dev. |  | 6.0809e-17 | 6.0809e-17 |  | 0.021602 | 0.012763 | 6.0809e-17 | 0.040825 | 0.040825 | 0.040825 | 0.05164 | 0.040825 | 0.0040825 | 0.033632 | 0.040825 | 0.020412 | 0.022286 | 0 | 0.013663 | 0.022583 | 0.01633 | 6.0809e-17 | 0
| Rel.Std.Dev. |  | 1.52e-14% | 1.52e-14% |  | 6.057% | 3.478% | 1.52e-14% | 10.65% | 10.65% | 10.65% | 14.08% | 9.798% | 1.195% | 8.469% | 12.89% | 5.348% | 5.891% | 0% | 5.393% | 6.741% | 4.454% | 1.52e-14% | 0%
| PDM3 |  | 8.734% | 8.734% |  | -3.046% | -0.2423% | 8.734% | 4.203% | 4.203% | 4.203% | -0.3275% | 13.26% | -7.123% | 7.949% | -13.92% | 3.75% | 2.844% | -18.45% | -31.14% | -8.936% | -0.3275% | 8.734% | -18.45%
| Z-Score (Absolute) |  | 0.67 | 0.67 | 3.02 | 0.39 | 0.14 | 0.67 | 0.26 | 0.26 | 0.26 | 0.15 | 1.09 | 0.76 | 0.6 | 1.38 | 0.22 | 0.14 | 1.79 | 2.94 | 0.93 | 0.15 | 0.67 | 1.79
| Analytical results for Hg in OREAS 163b (Indicative Value 0.078 ppm) | [26 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.03: Lab |  |  |  | 1
| No | No | 2.01: 2.06; 2.03: 2.18 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS |  |  |  | 3
3
| 1 | 1 | 2.01: < 1; 2.03: 0.09 |  |  |  | 1
| 1 | 2 | 2.01: < 1; 2.03: 0.07 |  |  |  | 1
| 1 | 3 | 2.01: < 1; 2.03: 0.07 |  |  |  | 16
| 1 | 4 | 2.01: < 1; 2.03: 0.08 |  |  |  | 0.078333
| 1 | 5 | 2.01: < 1; 2.03: 0.09 |  |  |  | 7
| 1 | 6 | 2.01: < 1; 2.03: 0.07
| Mean |  | 2.03: 0.078333
| Median |  | 2.03: 0.075
| Std Dev. |  | 2.03: 0.0098319
| Rel.Std.Dev. |  | 2.03: 12.55%
| PDM3 |  | 2.03: 4.441e-14%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 0.67
| Analytical results for Ho in OREAS 163b (Certified Value 0.13 ppm) | [26 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab |  |  |  | 1
| No | No | 2.01: 2.02; 2.03: 2.03; 2.04: 2.06; 2.05: 2.12; 2.06: 2.16; 2.07: 2.17; 2.08: 2.20 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.01: 0.10738; 2.03: 0.145; 2.04: 0.15397; 2.05: 0.2; 2.06: 0.14; 2.07: 0.2; 2.08: 0.1 |  |  |  | 1
| 1 | 2 | 2.01: 0.11908; 2.03: 0.146; 2.04: 0.15082; 2.05: 0.2; 2.06: 0.14; 2.07: 0.2; 2.08: 0.1 |  |  |  | 3
| 1 | 3 | 2.01: 0.11843; 2.03: 1.45; 2.04: 0.16648; 2.05: 0.2; 2.06: 0.15; 2.07: 0.1; 2.08: 0.11 |  |  |  | 16
| 1 | 4 | 2.01: 0.11934; 2.03: 0.146; 2.04: 0.15067; 2.05: 0.2; 2.06: 0.14; 2.07: 0.2; 2.08: 0.1 |  |  |  | 0.13305
| 1 | 5 | 2.01: 0.13894; 2.03: 0.145; 2.04: 0.14922; 2.05: 0.2; 2.06: 0.13; 2.07: 0.2; 2.08: 0.1 |  |  |  | 28
| 1 | 6 | 2.01: 0.12558; 2.03: 0.145; 2.04: 0.14917; 2.05: 0.2; 2.06: 0.14; 2.07: 0.1; 2.08: 0.12
| Mean |  | 2.01: 0.12146; 2.03: 0.36283; 2.04: 0.15339; 2.05: 0.2; 2.06: 0.14; 2.07: 0.16667; 2.08: 0.105
| Median |  | 2.01: 0.11921; 2.03: 0.1455; 2.04: 0.15075; 2.05: 0.2; 2.06: 0.14; 2.07: 0.2; 2.08: 0.1
| Std Dev. |  | 2.01: 0.010393; 2.03: 0.5326; 2.04: 0.0066474; 2.05: 3.0405e-17; 2.06: 0.0063246; 2.07: 0.05164; 2.08: 0.0083666
| Rel.Std.Dev. |  | 2.01: 8.557%; 2.03: 146.8%; 2.04: 4.334%; 2.05: 1.52e-14%; 2.06: 4.518%; 2.07: 30.98%; 2.08: 7.968%
| PDM3 |  | 2.01: -8.712%; 2.03: 172.7%; 2.04: 15.29%; 2.05: 50.32%; 2.06: 5.224%; 2.07: 25.27%; 2.08: -21.08%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.03: 8.1; 2.04: 0.49; 2.05: NA; 2.06: 0; 2.07: NA; 2.08: 1.27
Results from laboratories 2.12 and 2.17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 163b (Certified Value 1.06 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.19 | 1.1531 | 1.13 | < 1 | 1.02 | 0.94212 | 1.09 | 1.035 | 0.985 | 1.02 | 1.02 | 1.2 | 1.02 | 1.0595 | 0.944 | 1.08 | 1.04 | 1.05 | 1.06 | 1.03 | 1.08 | 1.07 | 1.024 | 1.05 |  | 1
| 1 | 2 | 1.15 | 1.1368 | 1.11 | < 1 | 1.03 | 0.98558 | 1.13 | 1.045 | 1.01 | 1.025 | 1.02 | 1.1 | 1.032 | 1.1159 | 0.826 | 1.08 | 0.92 | 1.04 | 1.14 | 1.08 | 1.04 | 1.09 | 1.007 | 1.07 |  | 10
| 1 | 3 | 1.17 | 1.1321 | 1.11 | < 1 | 1.03 | 1.0246 | 1.125 | 1.105 | 1.055 | 1 | 1.035 | 1.1 | 1.02 | 1.0919 | 0.809 | 1.07 | 1.01 | 1.02 | 1.06 | 1.04 | 1.06 | 1.11 | 1.05 | 1.09 |  | 16
| 1 | 4 | 1.14 | 1.156 | 1.12 | < 1 | 0.99 | 0.98088 | 1.135 | 1.08 | 0.997 | 0.969 | 1.035 | 1.1 | 1.042 | 1.1404 | 0.944 | 1.1 | 1.04 | 1.04 | 1.11 | 1.02 | 1.04 | 1.06 | 1.06 | 1.06 |  | 1.0624
| 1 | 5 | 1.15 | 1.1351 | 1.12 | < 1 | 1.05 | 0.95625 | 1.09 | 1.06 | 0.994 | 0.997 | 1.095 | 1.1 | 1.051 | 1.0973 | 0.857 | 1.09 | 1.04 | 1.04 | 1.04 | 1.08 | 1.01 | 1.06 | 1.022 | 1.08 |  | 29
| 1 | 6 | 1.13 | 1.1192 | 1.11 | < 1 | 1.07 | 0.97464 | 1.075 | 1.035 | 1.005 | 0.998 | 1.04 | 1.1 | 1.039 | 1.1602 | 0.767 | 1.1 | 0.98 | 1.05 | 1.08 | 1.05 | 1.05 | 1.05 | 1.088 | 1.07
| Mean |  | 1.155 | 1.1387 | 1.1167 |  | 1.0317 | 0.97734 | 1.1075 | 1.06 | 1.0077 | 1.0015 | 1.0408 | 1.1167 | 1.034 | 1.1109 | 0.85783 | 1.0867 | 1.005 | 1.04 | 1.0817 | 1.05 | 1.0467 | 1.0733 | 1.0418 | 1.07
| Median |  | 1.15 | 1.1359 | 1.115 |  | 1.03 | 0.97776 | 1.1075 | 1.0525 | 1.001 | 0.999 | 1.035 | 1.1 | 1.0355 | 1.1066 | 0.8415 | 1.085 | 1.025 | 1.04 | 1.07 | 1.045 | 1.045 | 1.065 | 1.037 | 1.07
| Std Dev. |  | 0.021679 | 0.013763 | 0.008165 |  | 0.027142 | 0.028297 | 0.025446 | 0.027928 | 0.024768 | 0.019927 | 0.027824 | 0.040825 | 0.012442 | 0.036117 | 0.072799 | 0.012111 | 0.048062 | 0.010954 | 0.037103 | 0.025298 | 0.023381 | 0.022509 | 0.029842 | 0.014142
| Rel.Std.Dev. |  | 1.877% | 1.209% | 0.7312% |  | 2.631% | 2.895% | 2.298% | 2.635% | 2.458% | 1.99% | 2.673% | 3.656% | 1.203% | 3.251% | 8.486% | 1.114% | 4.782% | 1.053% | 3.43% | 2.409% | 2.234% | 2.097% | 2.864% | 1.322%
| PDM3 |  | 8.714% | 7.179% | 5.106% |  | -2.895% | -8.008% | 4.243% | -0.2279% | -5.154% | -5.734% | -2.032% | 5.106% | -2.675% | 4.56% | -19.26% | 2.282% | -5.405% | -2.11% | 1.811% | -1.169% | -1.483% | 1.027% | -1.938% | 0.7134%
| Z-Score (Absolute) |  | 1.82 | 1.54 | 1.17 | 9.36 | 0.28 | 1.21 | 1.01 | 0.2 | 0.69 | 0.8 | 0.13 | 1.17 | 0.24 | 1.07 | 3.25 | 0.65 | 0.74 | 0.14 | 0.57 | 0.03 | 0.03 | 0.43 | 0.11 | 0.37
| Analytical results for K in OREAS 163b (Certified Value 0.09 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.078 | 0.09376 | 0.0906 | 0.08 | 0.08 | 0.087559 | 0.09 | 0.09 | 0.08 | 0.09 | 0.09 | 0.11 | 0.09 | 0.095258 | 0.11 | 0.0833 | 0.09 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0914 | 0.09 | 0.09 |  | 1
| 1 | 2 | 0.074 | 0.09349 | 0.0914 | 0.08 | 0.0916 | 0.088746 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.11 | 0.09 | 0.093086 | 0.09 | 0.0839 | 0.09 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0938 | 0.09 | 0.09 |  | 0
| 1 | 3 | 0.077 | 0.09395 | 0.092 | 0.08 | 0.0852 | 0.088757 | 0.1 | 0.09 | 0.09 | 0.09 | 0.09 | 0.11 | 0.09 | 0.10276 | 0.08 | 0.0805 | 0.09 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0879 | 0.09 | 0.09 |  | 16
| 1 | 4 | 0.075 | 0.09296 | 0.0916 | 0.08 | 0.0795 | 0.088049 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.093192 | 0.1 | 0.0812 | 0.09 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0857 | 0.09 | 0.09 |  | 0.0897
| 1 | 5 | 0.076 | 0.092924 | 0.0915 | 0.08 | 0.0864 | 0.092672 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.08947 | 0.09 | 0.082 | 0.08 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0911 | 0.09 | 0.09 |  | 30
| 1 | 6 | 0.078 | 0.09288 | 0.0912 | 0.08 | 0.0828 | 0.084278 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.10527 | 0.08 | 0.0804 | 0.09 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0877 | 0.09 | 0.1
| Mean |  | 0.076333 | 0.093327 | 0.091383 | 0.08 | 0.08425 | 0.088343 | 0.091667 | 0.09 | 0.088333 | 0.09 | 0.09 | 0.105 | 0.09 | 0.096506 | 0.091667 | 0.081883 | 0.088333 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0896 | 0.09 | 0.091667
| Median |  | 0.0765 | 0.093225 | 0.09145 | 0.08 | 0.084 | 0.088397 | 0.09 | 0.09 | 0.09 | 0.09 | 0.09 | 0.105 | 0.09 | 0.094225 | 0.09 | 0.0816 | 0.09 | 0.09 | 0.09 | 0.08 | 0.08 | 0.0895 | 0.09 | 0.09
| Std Dev. |  | 0.001633 | 0.00046885 | 0.00046655 | 0 | 0.0045227 | 0.0026943 | 0.0040825 | 1.5202e-17 | 0.0040825 | 1.5202e-17 | 1.5202e-17 | 0.0054772 | 1.5202e-17 | 0.0061598 | 0.01169 | 0.0014607 | 0.0040825 | 1.5202e-17 | 1.5202e-17 | 0 | 0 | 0.0029947 | 1.5202e-17 | 0.0040825
| Rel.Std.Dev. |  | 2.139% | 0.5024% | 0.5105% | 0% | 5.368% | 3.05% | 4.454% | 1.689e-14% | 4.622% | 1.689e-14% | 1.689e-14% | 5.216% | 1.689e-14% | 6.383% | 12.75% | 1.784% | 4.622% | 1.689e-14% | 1.689e-14% | 0% | 0% | 3.342% | 1.689e-14% | 4.454%
| PDM3 |  | -14.9% | 4.044% | 1.876% | -10.81% | -6.076% | -1.513% | 2.192% | 0.3342% | -1.524% | 0.3342% | 0.3342% | 17.06% | 0.3342% | 7.588% | 2.192% | -8.714% | -1.524% | 0.3342% | 0.3342% | -10.81% | -10.81% | -0.1117% | 0.3342% | 2.192%
| Z-Score (Absolute) |  | 5.53 | 1.35 | 0.56 | 4.05 | 2.33 | 0.67 | 0.67 | 0 | 0.67 | 0 | 0 | 6.07 | 0 | 2.63 | 0.67 | 3.28 | 0.67 | 0 | 0 | 4.05 | 4.05 | 0.16 | 0 | 0.67
| Analytical results for La in OREAS 163b (Certified Value 4.74 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 4.69 | 4.7 | 4.3 | 4.7 | 5.7369 | 4.7 | 4.5 | 5 | 4.8 | 4.5 | 5.8 | 4.49 | 5 | 4.57 | 5.2 | 5 | 5 | 3.9 | 4.8 | 4.79 | 4.6 | 7 |  |  |  | 1
| 1 | 2 | 4.74 | 4.7 | 4.06 | 5.03 | 5.8495 | 4.8 | 4.5 | 5.3 | 5.4 | 4.6 | 5.6 | 4.54 | 4.4 | 4.77 | 5 | 4.9 | 4.9 | 4.1 | 4.8 | 4.98 | 4.4 | 7 |  |  |  | 11
| 1 | 3 | 4.91 | 4.6 | 4.01 | 4.75 | 6.0887 | 5 | 4.4 | 5.4 | 5.1 | 4.6 | 5.6 | 4.48 | 4.2 | 4.78 | 4.9 | 4.9 | 4.9 | 3.9 | 4.8 | 4.66 | 4.7 | 7 |  |  |  | 16
| 1 | 4 | 4.75 | 4.6 | 4.48 | 4.59 | 5.9905 | 5 | 4.5 | 5.3 | 4.7 | 4.5 | 5.1 | 4.57 | 4.9 | 4.64 | 5.2 | 4.9 | 5.1 | 3.9 | 4.7 | 4.59 | 4.6 | 7 |  |  |  | 4.7353
| 1 | 5 | 4.37 | 4.7 | 4.21 | 4.74 | 6.3842 | 4.7 | 4.7 | 5.1 | 5 | 4.7 | 5.3 | 4.69 | 4.5 | 4.78 | 5.3 | 4.9 | 4.9 | 4 | 4.6 | 4.78 | 4.4 | 7 |  |  |  | 31
| 1 | 6 | 4.71 | 4.7 | 4.13 | 4.69 | 5.9036 | 4.6 | 4.6 | 5.1 | 4.7 | 4.5 | 5.6 | 4.66 | 4 | 4.97 | 4.9 | 4.8 | 5.1 | 4.2 | 4.8 | 4.66 | 4.8 | 6
| Mean |  | 4.695 | 4.6667 | 4.1983 | 4.75 | 5.9922 | 4.8 | 4.5333 | 5.2 | 4.95 | 4.5667 | 5.5 | 4.5717 | 4.5 | 4.7517 | 5.0833 | 4.9 | 4.9833 | 4 | 4.75 | 4.7433 | 4.5833 | 6.8333
| Median |  | 4.725 | 4.7 | 4.17 | 4.72 | 5.947 | 4.75 | 4.5 | 5.2 | 4.9 | 4.55 | 5.6 | 4.555 | 4.45 | 4.775 | 5.1 | 4.9 | 4.95 | 3.95 | 4.8 | 4.72 | 4.6 | 7
| Std Dev. |  | 0.17729 | 0.05164 | 0.17268 | 0.14846 | 0.22641 | 0.16733 | 0.10328 | 0.15492 | 0.27386 | 0.08165 | 0.25298 | 0.087044 | 0.38987 | 0.1379 | 0.17224 | 0.063246 | 0.098319 | 0.12649 | 0.083666 | 0.13924 | 0.16021 | 0.40825
| Rel.Std.Dev. |  | 3.776% | 1.107% | 4.113% | 3.125% | 3.778% | 3.486% | 2.278% | 2.979% | 5.533% | 1.788% | 4.6% | 1.904% | 8.664% | 2.902% | 3.388% | 1.291% | 1.973% | 3.162% | 1.761% | 2.935% | 3.495% | 5.974%
| PDM3 |  | -0.8514% | -1.45% | -11.34% | 0.3101% | 26.54% | 1.366% | -4.265% | 9.813% | 4.534% | -3.561% | 16.15% | -3.456% | -4.969% | 0.3453% | 7.349% | 3.478% | 5.238% | -15.53% | 0.3101% | 0.1693% | -3.21% | 44.31%
| Z-Score (Absolute) |  | 0.2 | 0.31 | 2.03 | 0 | 4.57 | 0.18 | 0.8 | 1.66 | 0.74 | 0.67 | 2.76 | 0.66 | 0.92 | 0.01 | 1.23 | 0.55 | 0.86 | 2.76 | 0 | 0.02 | 0.61 | NA
| Analytical results for Li in OREAS 163b (Certified Value 100 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 101 | 101.4405 | 95 | 110.8 | 97.6 | 106.185 | 108 | 100.5 | 100.5 | 107 | 101 | 105 | 102 | 102.9292 | 89.1 | 97.8 | 88.6 | 97.3 | 97 | 78 | 94.3 | 92.7 | 94.5 | 97.4 |  | 1
| 1 | 2 | 99.4 | 100.4905 | 95 | 103.4 | 98.3 | 100.9059 | 111 | 101 | 109.5 | 109.5 | 103 | 104 | 100 | 105.7184 | 85.2 | 98.9 | 91.2 | 96.3 | 97 | 91 | 96 | 95 | 91.4 | 98.8 |  | 0
| 1 | 3 | 101 | 100.1865 | 95 | 99.69 | 95.7 | 102.6942 | 109 | 106 | 109 | 108.5 | 100 | 103 | 105 | 106.6777 | 79.9 | 98.8 | 109.1 | 98.6 | 96 | 83 | 94 | 93.5 | 97.9 | 97.8 |  | 16
| 1 | 4 | 101 | 100.791 | 95 | 106.2 | 93.5 | 105.5381 | 105.5 | 103.5 | 109 | 105 | 98.5 | 103 | 103 | 104.373 | 91.9 | 99 | 92.4 | 96.4 | 97 | 78 | 95.7 | 92.6 | 94.3 | 96.6 |  | 100.3113
| 1 | 5 | 102 | 101.521 | 95 | 99.28 | 97 | 106.1395 | 111 | 101.5 | 103 | 105 | 104 | 108 | 105 | 105.4459 | 87.5 | 97.4 | 88.8 | 100.8 | 96 | 81 | 95.6 | 93.4 | 91.5 | 96.4 |  | 32
| 1 | 6 | 102 | 102.063 | 95 | 100.1 | 95.7 | 103.2966 | 104 | 103 | 105 | 102 | 107 | 104 | 103 | 104.029 | 73.2 | 100.2 | 108.4 | 98.1 | 98 | 86 | 95.3 | 91.6 | 91.7 | 98.6
| Mean |  | 101.0667 | 101.0821 | 95 | 103.245 | 96.3 | 104.1265 | 108.0833 | 102.5833 | 106 | 106.1667 | 102.25 | 104.5 | 103 | 104.8622 | 84.4667 | 98.6833 | 96.4167 | 97.9167 | 96.8333 | 82.8333 | 95.15 | 93.1333 | 93.55 | 97.6
| Median |  | 101 | 101.1158 | 95 | 101.75 | 96.35 | 104.4174 | 108.5 | 102.25 | 107 | 106 | 102 | 104 | 103 | 104.9094 | 86.35 | 98.85 | 91.8 | 97.7 | 97 | 82 | 95.45 | 93.05 | 93 | 97.6
| Std Dev. |  | 0.95219 | 0.70994 | 0 | 4.5626 | 1.717 | 2.1631 | 2.8708 | 2.0351 | 3.755 | 2.7325 | 3.0619 | 1.8708 | 1.8974 | 1.3446 | 6.8448 | 0.98877 | 9.6636 | 1.6798 | 0.75277 | 5.0365 | 0.81179 | 1.1413 | 2.5548 | 0.99599
| Rel.Std.Dev. |  | 0.9421% | 0.7023% | 0% | 4.419% | 1.783% | 2.077% | 2.656% | 1.984% | 3.542% | 2.574% | 2.994% | 1.79% | 1.842% | 1.282% | 8.103% | 1.002% | 10.02% | 1.716% | 0.7774% | 6.08% | 0.8532% | 1.225% | 2.731% | 1.02%
| PDM3 |  | 0.753% | 0.7684% | -5.295% | 2.925% | -3.999% | 3.803% | 7.748% | 2.265% | 5.671% | 5.837% | 1.933% | 4.176% | 2.68% | 4.537% | -15.8% | -1.623% | -3.883% | -2.387% | -3.467% | -17.42% | -5.145% | -7.156% | -6.74% | -2.703%
| Z-Score (Absolute) |  | 0.21 | 0.21 | 0.84 | 0.58 | 0.62 | 0.73 | 1.41 | 0.47 | 1.06 | 1.08 | 0.41 | 0.8 | 0.54 | 0.86 | 2.66 | 0.21 | 0.6 | 0.34 | 0.52 | 2.94 | 0.81 | 1.16 | 1.09 | 0.39
| Analytical results for Lu in OREAS 163b (Certified Value < 0.1 ppm) | [26 empty] | Certified
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.16; 2.08: 2.17; 2.09: 2.20; 2.10: 2.21 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*OES/MS |  | 2
2
| 1 | 1 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 1; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.04; 2.07: 0.04; 2.08: < 0.1; 2.09: 0.02; 2.10: 0.04 |  | 1
| 1 | 2 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 1; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.04; 2.07: 0.04; 2.08: < 0.1; 2.09: 0.02; 2.10: 0.04 |  | 4
| 1 | 3 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 1; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.04; 2.07: 0.05; 2.08: < 0.1; 2.09: 0.02; 2.10: 0.04 |  | 16
| 1 | 4 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 1; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.03; 2.08: < 0.1; 2.09: 0.02; 2.10: 0.04 |  | < 0.1
| 1 | 5 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 1; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.04; 2.08: < 0.1; 2.09: 0.02; 2.10: 0.04 |  | 33
| 1 | 6 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 1; 2.04: < 0.05; 2.05: < 0.1; 2.06: 0.05; 2.07: 0.04; 2.08: < 0.1; 2.09: 0.02; 2.10: 0.04
| Mean |  | 2.06: 0.045; 2.07: 0.04; 2.09: 0.02; 2.10: 0.04
| Median |  | 2.06: 0.045; 2.07: 0.04; 2.09: 0.02; 2.10: 0.04
| Std Dev. |  | 2.06: 0.0054772; 2.07: 0.0063246; 2.09: 0; 2.10: 0
| Rel.Std.Dev. |  | 2.06: 12.17%; 2.07: 15.81%; 2.09: 0%; 2.10: 0%
PDM3
| Z-Score (Absolute) |  | 2.01: 0.34; 2.02: 0.34; 2.03: 61.02; 2.04: 3.03; 2.05: 0.34; 2.06: 0.34; 2.07: 1.01; 2.08: 0.34; 2.09: 3.71; 2.10: 1.01
| Analytical results for Mg in OREAS 163b (Certified Value 3.51 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.51 | 3.5972 | 3.54 | 3.64 | 3.2313 | 3.7151 | 3.6 | 3.5 | 3.37 | 3.48 | 3.48 | 4.2 | 3.504 | 3.3563 | 3.63 | 3.4632 | 3.79 | 3.4788 | 3.53 | 3.49 | 3.46 | 3.42 | 3.438 | 3.6 | 3.4 | 1
| 1 | 2 | 3.51 | 3.5802 | 3.53 | 3.52 | 3.5063 | 3.6445 | 3.64 | 3.49 | 3.72 | 3.51 | 3.54 | 4.06 | 3.484 | 3.4273 | 3.61 | 3.4785 | 3.8 | 3.4557 | 3.52 | 3.49 | 3.45 | 3.37 | 3.5421 | 3.57 | 3.39 | 0
| 1 | 3 | 3.55 | 3.575 | 3.53 | 3.31 | 3.3881 | 3.6866 | 3.61 | 3.62 | 3.62 | 3.52 | 3.47 | 4.15 | 3.683 | 3.3497 | 3.39 | 3.4315 | 3.66 | 3.4505 | 3.47 | 3.48 | 3.45 | 3.38 | 3.3955 | 3.69 | 3.4 | 16
| 1 | 4 | 3.57 | 3.562 | 3.52 | 3.57 | 3.2414 | 3.7284 | 3.49 | 3.55 | 3.65 | 3.4 | 3.46 | 3.86 | 3.549 | 3.5065 | 3.8 | 3.5167 | 3.84 | 3.4913 | 3.53 | 3.5 | 3.44 | 3.43 | 3.3653 | 3.57 | 3.38 | 3.5119
| 1 | 5 | 3.52 | 3.5637 | 3.54 | 3.25 | 3.3911 | 3.6779 | 3.66 | 3.43 | 3.41 | 3.45 | 3.61 | 4.02 | 3.642 | 3.5323 | 2.75 | 3.5182 | 3.71 | 3.4591 | 3.51 | 3.41 | 3.44 | 3.42 | 3.4521 | 3.59 | 3.38 | 34
| 1 | 6 | 3.54 | 3.5587 | 3.5 | 3.32 | 3.3527 | 3.7124 | 3.49 | 3.56 | 3.57 | 3.41 | 3.68 | 4.07 | 3.57 | 3.3805 | 3.46 | 3.4334 | 3.57 | 3.4418 | 3.52 | 3.46 | 3.45 | 3.37 | 3.3498 | 3.7 | 3.4
| Mean |  | 3.5333 | 3.5728 | 3.5267 | 3.435 | 3.3518 | 3.6942 | 3.5817 | 3.525 | 3.5567 | 3.4617 | 3.54 | 4.06 | 3.572 | 3.4254 | 3.44 | 3.4736 | 3.7283 | 3.4629 | 3.5133 | 3.4717 | 3.4483 | 3.3983 | 3.4238 | 3.62 | 3.3917
| Median |  | 3.53 | 3.5693 | 3.53 | 3.42 | 3.3704 | 3.6995 | 3.605 | 3.525 | 3.595 | 3.465 | 3.51 | 4.065 | 3.5595 | 3.4039 | 3.535 | 3.4708 | 3.75 | 3.4574 | 3.52 | 3.485 | 3.45 | 3.4 | 3.4167 | 3.595 | 3.395
| Std Dev. |  | 0.024221 | 0.014501 | 0.015055 | 0.16159 | 0.10338 | 0.030762 | 0.07414 | 0.065955 | 0.13852 | 0.050365 | 0.088769 | 0.11781 | 0.077617 | 0.078137 | 0.36704 | 0.038386 | 0.10147 | 0.018572 | 0.022509 | 0.033116 | 0.0075277 | 0.027869 | 0.07028 | 0.05933 | 0.0098319
| Rel.Std.Dev. |  | 0.6855% | 0.4059% | 0.4269% | 4.704% | 3.084% | 0.8327% | 2.07% | 1.871% | 3.895% | 1.455% | 2.508% | 2.902% | 2.173% | 2.281% | 10.67% | 1.105% | 2.722% | 0.5363% | 0.6407% | 0.9539% | 0.2183% | 0.8201% | 2.053% | 1.639% | 0.2899%
| PDM3 |  | 0.6097% | 1.733% | 0.4198% | -2.19% | -4.559% | 5.189% | 1.986% | 0.3724% | 1.274% | -1.431% | 0.7995% | 15.61% | 1.711% | -2.462% | -2.048% | -1.092% | 6.162% | -1.397% | 0.04019% | -1.146% | -1.811% | -3.234% | -2.509% | 3.077% | -3.424%
| Z-Score (Absolute) |  | 0.21 | 0.62 | 0.14 | 0.81 | 1.68 | 1.88 | 0.71 | 0.12 | 0.45 | 0.54 | 0.28 | 5.67 | 0.61 | 0.91 | 0.76 | 0.41 | 2.23 | 0.52 | 0 | 0.43 | 0.67 | 1.19 | 0.93 | 1.11 | 1.26
| Analytical results for Mn in OREAS 163b (Certified Value 0.048 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.0491 | 0.049387 | 0.049 | 0.04972 | 0.0451 | 0.046181 | 0.0476 | 0.0453 | 0.0452 | 0.0456 | 0.0463 | 0.0552 | 0.0487 | 0.048765 | 0.0457 | 0.0488 | 0.0501 | 0.0486 | 0.0466 | 0.0497 | 0.0476 | 0.0451 | 0.0488 | 0.0482 | 0.0505 | 1
| 1 | 2 | 0.0491 | 0.049464 | 0.049 | 0.04607 | 0.0472 | 0.046401 | 0.048 | 0.0453 | 0.0477 | 0.0457 | 0.0474 | 0.0527 | 0.0485 | 0.049525 | 0.0406 | 0.049 | 0.0507 | 0.0486 | 0.0465 | 0.05 | 0.0477 | 0.046 | 0.0493 | 0.0481 | 0.0499 | 30
| 1 | 3 | 0.0496 | 0.05015 | 0.0492 | 0.04518 | 0.0463 | 0.046705 | 0.0478 | 0.0475 | 0.0476 | 0.046 | 0.0462 | 0.0534 | 0.051 | 0.049721 | 0.0397 | 0.0484 | 0.0484 | 0.0486 | 0.0457 | 0.0494 | 0.0477 | 0.0449 | 0.048 | 0.0496 | 0.0498 | 16
| 1 | 4 | 0.0488 | 0.04994 | 0.0488 | 0.04915 | 0.0443 | 0.04561 | 0.0457 | 0.0458 | 0.0453 | 0.0445 | 0.0464 | 0.0505 | 0.0494 | 0.049184 | 0.0472 | 0.0496 | 0.0509 | 0.0485 | 0.0463 | 0.0503 | 0.0476 | 0.0435 | 0.0475 | 0.048 | 0.0498 | 0.047956
| 1 | 5 | 0.0493 | 0.049581 | 0.0492 | 0.04491 | 0.0469 | 0.046393 | 0.0484 | 0.0451 | 0.0452 | 0.045 | 0.0474 | 0.0513 | 0.0504 | 0.049218 | 0.0426 | 0.0494 | 0.0491 | 0.0482 | 0.0462 | 0.0487 | 0.0477 | 0.0446 | 0.0489 | 0.0483 | 0.0504 | 35
| 1 | 6 | 0.0493 | 0.049914 | 0.049 | 0.04487 | 0.0464 | 0.045568 | 0.0461 | 0.0457 | 0.0461 | 0.0446 | 0.0488 | 0.0525 | 0.0494 | 0.048871 | 0.0378 | 0.0483 | 0.0473 | 0.0481 | 0.0465 | 0.0492 | 0.0477 | 0.0449 | 0.0479 | 0.0497 | 0.049
| Mean |  | 0.0492 | 0.049739 | 0.049033 | 0.04665 | 0.046033 | 0.046143 | 0.047267 | 0.045783 | 0.046183 | 0.045233 | 0.047083 | 0.0526 | 0.049567 | 0.049214 | 0.042267 | 0.048917 | 0.049417 | 0.048433 | 0.0463 | 0.04955 | 0.047667 | 0.044833 | 0.0484 | 0.04865 | 0.0499
| Median |  | 0.0492 | 0.049748 | 0.049 | 0.045625 | 0.04635 | 0.046287 | 0.0477 | 0.0455 | 0.0457 | 0.0453 | 0.0469 | 0.0526 | 0.0494 | 0.049201 | 0.0416 | 0.0489 | 0.0496 | 0.04855 | 0.0464 | 0.04955 | 0.0477 | 0.0449 | 0.0484 | 0.04825 | 0.04985
| Std Dev. |  | 0.00026833 | 0.00030492 | 0.00015055 | 0.0022076 | 0.001113 | 0.00046073 | 0.0010985 | 0.00088185 | 0.0011856 | 0.00062183 | 0.0010008 | 0.0016444 | 0.00096885 | 0.00036701 | 0.0036209 | 0.00052313 | 0.0014091 | 0.00022509 | 0.00032863 | 0.00057533 | 5.164e-05 | 0.00080911 | 0.00069857 | 0.00078166 | 0.00053666
| Rel.Std.Dev. |  | 0.5454% | 0.613% | 0.307% | 4.732% | 2.418% | 0.9985% | 2.324% | 1.926% | 2.567% | 1.375% | 2.126% | 3.126% | 1.955% | 0.7457% | 8.567% | 1.069% | 2.852% | 0.4647% | 0.7098% | 1.161% | 0.1083% | 1.805% | 1.443% | 1.607% | 1.075%
| PDM3 |  | 2.595% | 3.72% | 2.247% | -2.722% | -4.008% | -3.78% | -1.436% | -4.53% | -3.696% | -5.677% | -1.819% | 9.685% | 3.36% | 2.624% | -11.86% | 2.004% | 3.047% | 0.9963% | -3.452% | 3.325% | -0.6024% | -6.511% | 0.9268% | 1.448% | 4.055%
| Z-Score (Absolute) |  | 0.41 | 0.69 | 0.32 | 0.9 | 1.21 | 1.16 | 0.58 | 1.34 | 1.14 | 1.62 | 0.67 | 2.15 | 0.6 | 0.42 | 3.14 | 0.26 | 0.52 | 0.02 | 1.08 | 0.59 | 0.38 | 1.83 | 0 | 0.13 | 0.77
| Analytical results for Mo in OREAS 163b (Certified Value 2.1 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.1 | 2.1 | 2.72 | 2.1 | 2.2236 | 2.22 | 2 | 2.02 | 2.05 | 2 | 2 | 2.13 | 1.7653 | 2.04 | 2.2 | 2 | 2 | 2 | 2.1 | 1.73 | 1.87 | 2.3 | 2.11 | 2.3 |  | 1
| 1 | 2 | 2.1 | 2.1 | 2.51 | 2.1 | 2.2579 | 2.06 | 2.03 | 2.11 | 2.1 | 2.33 | 2 | 2.09 | 1.8125 | 1.84 | 2.1 | 1.96 | 1.9 | 2.1 | 2.1 | 2.09 | 1.97 | 2.3 | 2.05 | 2.1 |  | 31
| 1 | 3 | 2.1 | 2.1 | 2.5 | 2.2 | 2.2252 | 2.41 | 2.03 | 2.5 | 2.05 | 2.08 | 2 | 2.18 | 1.8104 | 1.71 | 2.1 | 2.1 | 2 | 2 | 2 | 1.73 | 1.89 | 2.2 | 2.08 | 2.2 |  | 16
| 1 | 4 | 2.1 | 2.1 | 2.87 | 2.2 | 2.2849 | 2.23 | 2.12 | 2.1 | 2 | 2.1 | 3 | 2.27 | 1.7699 | 2.02 | 2.2 | 2.19 | 1.9 | 2.1 | 2 | 1.72 | 1.88 | 2.3 | 2.1 | 2.2 |  | 2.0963
| 1 | 5 | 2.1 | 2.1 | 2.71 | 2 | 2.3461 | 2.15 | 2.16 | 2.03 | 2.09 | 2.22 | 2 | 2.12 | 1.8724 | 1.91 | 2 | 2.2 | 2 | 2.2 | 1.9 | 1.72 | 1.82 | 2.3 | 2.04 | 2.2 |  | 36
| 1 | 6 | 2.1 | 2.1 | 2.98 | 2.2 | 2.3453 | 2.04 | 2.14 | 2.01 | 2.03 | 2.11 | 2 | 2.32 | 1.8618 | 1.62 | 2.1 | 1.99 | 2 | 2.1 | 2 | 1.83 | 1.91 | 2.4 | 2.1 | 2.1
| Mean |  | 2.1 | 2.1 | 2.715 | 2.1333 | 2.2805 | 2.185 | 2.08 | 2.1283 | 2.0533 | 2.14 | 2.1667 | 2.185 | 1.8154 | 1.8567 | 2.1167 | 2.0733 | 1.9667 | 2.0833 | 2.0167 | 1.8033 | 1.89 | 2.3 | 2.08 | 2.1833
| Median |  | 2.1 | 2.1 | 2.715 | 2.15 | 2.2714 | 2.185 | 2.075 | 2.065 | 2.05 | 2.105 | 2 | 2.155 | 1.8115 | 1.875 | 2.1 | 2.05 | 2 | 2.1 | 2 | 1.73 | 1.885 | 2.3 | 2.09 | 2.2
| Std Dev. |  | 0 | 0 | 0.19108 | 0.08165 | 0.055378 | 0.13546 | 0.067823 | 0.18691 | 0.037238 | 0.11679 | 0.40825 | 0.091378 | 0.044763 | 0.16789 | 0.075277 | 0.10539 | 0.05164 | 0.075277 | 0.075277 | 0.14665 | 0.049396 | 0.063246 | 0.028983 | 0.075277
| Rel.Std.Dev. |  | 0% | 0% | 7.038% | 3.827% | 2.428% | 6.2% | 3.261% | 8.782% | 1.814% | 5.457% | 18.84% | 4.182% | 2.466% | 9.042% | 3.556% | 5.083% | 2.626% | 3.613% | 3.733% | 8.132% | 2.614% | 2.75% | 1.393% | 3.448%
| PDM3 |  | 0.1785% | 0.1785% | 29.52% | 1.769% | 8.789% | 4.233% | -0.7756% | 1.53% | -2.048% | 2.087% | 3.359% | 4.233% | -13.4% | -11.43% | 0.9735% | -1.094% | -6.182% | -0.6166% | -3.797% | -13.97% | -9.839% | 9.719% | -0.7756% | 4.154%
| Z-Score (Absolute) |  | 0 | 0 | 4.98 | 0.27 | 1.46 | 0.69 | 0.16 | 0.23 | 0.38 | 0.32 | NA | 0.69 | 2.3 | 1.97 | 0.13 | 0.22 | 1.08 | 0.13 | 0.67 | 2.4 | 1.7 | 1.62 | 0.16 | 0.67
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 163b (Certified Value 0.017 wt.%) | [26 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 3
3
| 1 | 1 | 0.016 | 0.016747 | 0.07 | < 0.01 | 0.0128 | 0.018541 | 0.02 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.015 | 0.055822 | 0.02 | 0.0146 | 0.017 | 0.0102 | 0.015 | 0.02 | 0.01 | 0.015 | 0.013 | 0.02 | 0.02 | 1
| 1 | 2 | 0.015 | 0.016816 | 0.0696 | < 0.01 | 0.0138 | 0.019035 | 0.01 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016 | 0.059537 | 0.02 | 0.0141 | 0.018 | 0.0104 | 0.015 | 0.02 | 0.01 | 0.015 | 0.0138 | 0.02 | 0.02 | 23
| 1 | 3 | 0.016 | 0.017143 | 0.07 | < 0.01 | 0.0134 | 0.019382 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016 | 0.058324 | 0.02 | 0.0138 | 0.017 | 0.0102 | 0.014 | 0.02 | 0.01 | 0.015 | 0.0128 | 0.02 | 0.02 | 16
| 1 | 4 | 0.016 | 0.016519 | 0.07 | < 0.01 | 0.0132 | 0.018817 | 0.01 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.018 | 0.061417 | 0.02 | 0.0142 | 0.018 | 0.0104 | 0.015 | 0.02 | 0.01 | 0.015 | 0.0123 | 0.02 | 0.02 | 0.016527
| 1 | 5 | 0.016 | 0.016719 | 0.07 | < 0.01 | 0.0133 | 0.018444 | 0.01 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016 | 0.056078 | 0.01 | 0.0143 | 0.017 | 0.0105 | 0.015 | 0.02 | 0.01 | 0.015 | 0.0131 | 0.02 | 0.02 | 37
| 1 | 6 | 0.015 | 0.016555 | 0.07 | < 0.01 | 0.0133 | 0.018286 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016 | 0.060212 | 0.02 | 0.0138 | 0.017 | 0.0104 | 0.015 | 0.02 | 0.01 | 0.015 | 0.0128 | 0.02 | 0.02
| Mean |  | 0.015667 | 0.01675 | 0.069933 |  | 0.0133 | 0.018751 | 0.015 | 0.015 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016167 | 0.058565 | 0.018333 | 0.014133 | 0.017333 | 0.01035 | 0.014833 | 0.02 | 0.01 | 0.015 | 0.012967 | 0.02 | 0.02
| Median |  | 0.016 | 0.016733 | 0.07 |  | 0.0133 | 0.018679 | 0.015 | 0.015 | 0.02 | 0.02 | 0.02 | 0.02 | 0.016 | 0.058931 | 0.02 | 0.01415 | 0.017 | 0.0104 | 0.015 | 0.02 | 0.01 | 0.015 | 0.0129 | 0.02 | 0.02
| Std Dev. |  | 0.0005164 | 0.00022413 | 0.0001633 |  | 0.00032249 | 0.00040943 | 0.0054772 | 0.0054772 | 0 | 0 | 0 | 0 | 0.00098319 | 0.0022611 | 0.0040825 | 0.00030768 | 0.0005164 | 0.00012247 | 0.00040825 | 0 | 0 | 0 | 0.00049261 | 0 | 0
| Rel.Std.Dev. |  | 3.296% | 1.338% | 0.2335% |  | 2.425% | 2.184% | 36.51% | 36.51% | 0% | 0% | 0% | 0% | 6.082% | 3.861% | 22.27% | 2.177% | 2.979% | 1.183% | 2.752% | 0% | 0% | 0% | 3.799% | 0% | 0%
| PDM3 |  | -5.203% | 1.351% | 323.2% |  | -19.52% | 13.46% | -9.237% | -9.237% | 21.02% | 21.02% | 21.02% | 21.02% | -2.178% | 254.4% | 10.93% | -14.48% | 4.882% | -37.37% | -10.25% | 21.02% | -39.49% | -9.237% | -21.54% | 21.02% | 21.02%
| Z-Score (Absolute) |  | 0.22 | 0 | 11.03 | 2.44 | 0.72 | 0.42 | 0.36 | 0.36 | 0.67 | 0.67 | 0.67 | 0.67 | 0.12 | 8.68 | 0.33 | 0.54 | 0.12 | 1.33 | 0.4 | 0.67 | 1.4 | 0.36 | 0.78 | 0.67 | 0.67
| Analytical results for Nb in OREAS 163b (Certified Value 1.45 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 1.3 | 1.5158 | 1.4 | < 5 | 1.47 | 1.709 | 1.6 | 1.3 | 1.3 | 1.4 | 1.6 | 1.8 | 1.4 | 1.4915 | 1.7 | 1.48 | 1.38 | 1.1 | 1.4 | 1.2 | 1.5 | 1.84 | 1.3 | 1.3 |  | 1
| 1 | 2 | 1.3 | 1.5932 | 1.3 | < 5 | 1.28 | 1.5731 | 1.5 | 2 | 1.9 | 2.1 | 1.5 | 1.8 | 1.1 | 2.0047 | 1.3 | 1.84 | 1.38 | 1.2 | 1.5 | 1.7 | 1.5 | 1.48 | 1.4 | 1.2 |  | 14
| 1 | 3 | 1.3 | 1.5222 | 1.3 | < 5 | 1.41 | 1.5656 | 1.5 | 1.9 | 1.9 | 1.8 | 1.4 | 1.8 | 1 | 2.0515 | 1.3 | 1.9 | 1.27 | 0.9 | 1.4 | 1.1 | 1.4 | 1.35 | 1.3 | 1.4 |  | 16
| 1 | 4 | 1.3 | 1.619 | 1.4 | < 5 | 1.45 | 1.4758 | 1.3 | 1.4 | 1.6 | 1.5 | 1.4 | 1.8 | 1.1 | 1.8178 | 1.5 | 1.55 | 1.55 | 1.2 | 1.5 | 1.2 | 1.3 | 1.31 | 1.3 | 1.5 |  | 1.4499
| 1 | 5 | 1.3 | 1.6356 | 1.4 | < 5 | 1.3 | 1.4735 | 1.4 | 1.5 | 1.4 | 1.6 | 1.2 | 1.8 | 1.3 | 1.5804 | 1.3 | 1.64 | 1.4 | 1.1 | 1.6 | 1.1 | 1.3 | 1.32 | 1.4 | 1.4 |  | 38
| 1 | 6 | 1.3 | 1.4999 | 1.4 | < 5 | 1.95 | 1.7019 | 1.4 | 1.5 | 1.6 | 1.2 | 1.4 | 2.3 | 1.2 | 1.7261 | 1.1 | 1.73 | 1.32 | 1 | 1.6 | 1.1 | 1.3 | 1.56 | 1.4 | 1.2
| Mean |  | 1.3 | 1.5643 | 1.3667 |  | 1.4767 | 1.5832 | 1.45 | 1.6 | 1.6167 | 1.6 | 1.4167 | 1.8833 | 1.1833 | 1.7787 | 1.3667 | 1.69 | 1.3833 | 1.0833 | 1.5 | 1.2333 | 1.3833 | 1.4767 | 1.35 | 1.3333
| Median |  | 1.3 | 1.5577 | 1.4 |  | 1.43 | 1.5694 | 1.45 | 1.5 | 1.6 | 1.55 | 1.4 | 1.8 | 1.15 | 1.7719 | 1.3 | 1.685 | 1.38 | 1.1 | 1.5 | 1.15 | 1.35 | 1.415 | 1.35 | 1.35
| Std Dev. |  | 0 | 0.058623 | 0.05164 |  | 0.24459 | 0.10384 | 0.10488 | 0.28284 | 0.24833 | 0.31623 | 0.13292 | 0.20412 | 0.1472 | 0.22433 | 0.20656 | 0.16395 | 0.094798 | 0.1169 | 0.089443 | 0.23381 | 0.098319 | 0.20363 | 0.054772 | 0.12111
| Rel.Std.Dev. |  | 0% | 3.748% | 3.779% |  | 16.56% | 6.559% | 7.233% | 17.68% | 15.36% | 19.76% | 9.382% | 10.84% | 12.44% | 12.61% | 15.11% | 9.701% | 6.853% | 10.79% | 5.963% | 18.96% | 7.107% | 13.79% | 4.057% | 9.083%
| PDM3 |  | -10.34% | 7.888% | -5.742% |  | 1.845% | 9.19% | 0.005669% | 10.35% | 11.5% | 10.35% | -2.293% | 29.89% | -18.39% | 22.67% | -5.742% | 16.56% | -4.592% | -25.28% | 3.454% | -14.94% | -4.592% | 1.845% | -6.891% | -8.041%
| Z-Score (Absolute) |  | 0.88 | 0.54 | 0.52 | 5.6 | 0.07 | 0.65 | 0.07 | 0.74 | 0.83 | 0.74 | 0.25 | 2.27 | 1.51 | 1.7 | 0.52 | 1.22 | 0.43 | 2.05 | 0.2 | 1.24 | 0.43 | 0.07 | 0.61 | 0.7
| Analytical results for Nd in OREAS 163b (Certified Value 5.18 ppm) | [26 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.18; 2.08: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 4.3142; 2.02: 5.5; 2.03: 5.0537; 2.04: 5.7; 2.05: 4.94; 2.06: 5.6; 2.07: 5; 2.08: 5.1 |  | 1
| 1 | 2 | 2.01: 4.6892; 2.02: 5.5; 2.03: 5.3181; 2.04: 5.3; 2.05: 5.02; 2.06: 5.3; 2.07: 5; 2.08: 5 |  | 5
| 1 | 3 | 2.01: 4.6362; 2.02: 5.5; 2.03: 5.2409; 2.04: 5.8; 2.05: 5.25; 2.06: 5.3; 2.07: 5; 2.08: 4.9 |  | 16
| 1 | 4 | 2.01: 4.6792; 2.02: 5.5; 2.03: 5.2712; 2.04: 5; 2.05: 5.01; 2.06: 5.8; 2.07: 6; 2.08: 4.9 |  | 5.181
| 1 | 5 | 2.01: 4.3014; 2.02: 5.5; 2.03: 5.2266; 2.04: 5.3; 2.05: 5.36; 2.06: 5.6; 2.07: 5; 2.08: 4.8 |  | 39
| 1 | 6 | 2.01: 4.5232; 2.02: 5.5; 2.03: 5.1865; 2.04: 5.8; 2.05: 5.28; 2.06: 5.2; 2.07: 6; 2.08: 4.9
| Mean |  | 2.01: 4.5239; 2.02: 5.5; 2.03: 5.2162; 2.04: 5.4833; 2.05: 5.1433; 2.06: 5.4667; 2.07: 5.3333; 2.08: 4.9333
| Median |  | 2.01: 4.5797; 2.02: 5.5; 2.03: 5.2337; 2.04: 5.5; 2.05: 5.135; 2.06: 5.45; 2.07: 5; 2.08: 4.9
| Std Dev. |  | 2.01: 0.1775; 2.02: 0; 2.03: 0.091018; 2.04: 0.33116; 2.05: 0.17397; 2.06: 0.23381; 2.07: 0.5164; 2.08: 0.10328
| Rel.Std.Dev. |  | 2.01: 3.924%; 2.02: 0%; 2.03: 1.745%; 2.04: 6.039%; 2.05: 3.382%; 2.06: 4.277%; 2.07: 9.682%; 2.08: 2.094%
| PDM3 |  | 2.01: -12.68%; 2.02: 6.158%; 2.03: 0.6795%; 2.04: 5.836%; 2.05: -0.7263%; 2.06: 5.515%; 2.07: 2.941%; 2.08: -4.78%
| Z-Score (Absolute) |  | 2.01: 1.75; 2.02: 0.72; 2.03: 0; 2.04: 0.67; 2.05: 0.18; 2.06: 0.63; 2.07: NA; 2.08: 0.71
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 163b (Certified Value 24.3 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 26 | 23.34 | 26.5 | 22.7 | 23.5 | 24.8465 | 22.6 | 21.4 | 26 | 22.2 | 23.5 | 25 | 24 | 23.668 | 26.2 | 24.3 | 25.9 | 23.9 | 25 | 25.8 | 26.2 | 23.8 | 25.2 | 24.3 | 25 | 1
| 1 | 2 | 25 | 23.75 | 26.4 | 22.26 | 24.2 | 24.7169 | 23.6 | 21.6 | 23.4 | 22.8 | 24.6 | 23 | 24 | 26.3677 | 22.5 | 24.8 | 25.6 | 23.1 | 24.6 | 25.5 | 31.1 | 25 | 25.5 | 23.1 | 26.9 | 16
| 1 | 3 | 25 | 24.27 | 26.4 | 22.32 | 24.3 | 24.6315 | 27.6 | 22.4 | 32.9 | 23 | 23.5 | 24 | 25 | 23.762 | 23 | 24.8 | 24.1 | 23.6 | 24.7 | 26.7 | 24.2 | 23.9 | 26.5 | 23.6 | 25.6 | 16
| 1 | 4 | 25 | 23.81 | 26.4 | 23.49 | 23.1 | 24.143 | 23.7 | 24.5 | 23.4 | 21.7 | 23.1 | 24 | 26 | 27.2471 | 25.4 | 23.7 | 25.5 | 23.7 | 25.1 | 26.1 | 23.7 | 22.6 | 25.3 | 24 | 27.2 | 24.349
| 1 | 5 | 25 | 22.16 | 26.4 | 21.68 | 23.8 | 23.4706 | 22.7 | 23.3 | 24.5 | 23.8 | 25.4 | 24 | 25 | 27.0338 | 23 | 24.1 | 25.1 | 24.3 | 26.3 | 27.1 | 23.5 | 24 | 24.9 | 23.8 | 24.6 | 40
| 1 | 6 | 25 | 23.96 | 25.4 | 25.98 | 24.2 | 25.2277 | 22.8 | 23.3 | 23.4 | 22.8 | 24.5 | 23 | 25 | 27.8054 | 21.3 | 25.1 | 23.2 | 23.5 | 25 | 25.4 | 24.2 | 23.1 | 29.2 | 24.3 | 27.8
| Mean |  | 25.1667 | 23.5483 | 26.25 | 23.0717 | 23.85 | 24.506 | 23.8333 | 22.75 | 25.6 | 22.7167 | 24.1 | 23.8333 | 24.8333 | 25.9807 | 23.5667 | 24.4667 | 24.9 | 23.6833 | 25.1167 | 26.1 | 25.4833 | 23.7333 | 26.1 | 23.85 | 26.1833
| Median |  | 25 | 23.78 | 26.4 | 22.51 | 24 | 24.6742 | 23.2 | 22.85 | 23.95 | 22.8 | 24 | 24 | 25 | 26.7007 | 23 | 24.55 | 25.3 | 23.65 | 25 | 25.95 | 24.2 | 23.85 | 25.4 | 23.9 | 26.25
| Std Dev. |  | 0.40825 | 0.74435 | 0.41833 | 1.5448 | 0.47645 | 0.6166 | 1.9044 | 1.1777 | 3.7197 | 0.71671 | 0.87407 | 0.75277 | 0.75277 | 1.8146 | 1.8554 | 0.52409 | 1.0412 | 0.40208 | 0.61128 | 0.67823 | 2.9144 | 0.82381 | 1.6137 | 0.45935 | 1.2968
| Rel.Std.Dev. |  | 1.622% | 3.161% | 1.594% | 6.696% | 1.998% | 2.516% | 7.99% | 5.177% | 14.53% | 3.155% | 3.627% | 3.158% | 3.031% | 6.984% | 7.873% | 2.142% | 4.181% | 1.698% | 2.434% | 2.599% | 11.44% | 3.471% | 6.183% | 1.926% | 4.953%
| PDM3 |  | 3.358% | -3.288% | 7.807% | -5.246% | -2.05% | 0.6447% | -2.118% | -6.567% | 5.138% | -6.704% | -1.023% | -2.118% | 1.989% | 6.701% | -3.213% | 0.4831% | 2.263% | -2.734% | 3.153% | 7.191% | 4.658% | -2.529% | 7.191% | -2.05% | 7.533%
| Z-Score (Absolute) |  | 0.6 | 0.79 | 1.54 | 1.2 | 0.53 | 0.03 | 0.55 | 1.48 | 0.98 | 1.51 | 0.32 | 0.55 | 0.32 | 1.3 | 0.77 | 0 | 0.37 | 0.67 | 0.56 | 1.41 | 0.88 | 0.63 | 1.41 | 0.53 | 1.48
| Analytical results for P in OREAS 163b (Certified Value 0.018 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.018 | 0.0125 | 0.02 | 0.0154 | 0.018364 | 0.017 | 0.017 | 0.016 | 0.018 | 0.016 | 0.021 | 0.017 | 0.017072 | 0.016 | 0.0184 | 0.019 | 0.0177 | 0.018 | 0.0218 | 0.02 | 0.017 | 0.0173 | 0.0183 | 0.018 |  | 1
| 1 | 2 | 0.018 | 0.0125 | 0.02 | 0.0177 | 0.018494 | 0.018 | 0.017 | 0.019 | 0.018 | 0.017 | 0.021 | 0.018 | 0.018622 | 0.016 | 0.0181 | 0.019 | 0.018 | 0.018 | 0.0218 | 0.02 | 0.017 | 0.018 | 0.0184 | 0.018 |  | 0
| 1 | 3 | 0.018 | 0.0125 | 0.02 | 0.0167 | 0.01878 | 0.018 | 0.017 | 0.017 | 0.018 | 0.016 | 0.021 | 0.018 | 0.0204 | 0.013 | 0.0179 | 0.018 | 0.0175 | 0.017 | 0.0206 | 0.02 | 0.0171 | 0.0177 | 0.0184 | 0.018 |  | 16
| 1 | 4 | 0.018 | 0.0125 | 0.02 | 0.0153 | 0.018233 | 0.017 | 0.017 | 0.019 | 0.017 | 0.017 | 0.021 | 0.017 | 0.018676 | 0.017 | 0.0179 | 0.019 | 0.0181 | 0.018 | 0.0212 | 0.02 | 0.0171 | 0.0178 | 0.0177 | 0.018 |  | 0.017672
| 1 | 5 | 0.018 | 0.0125 | 0.02 | 0.0169 | 0.01768 | 0.018 | 0.018 | 0.017 | 0.018 | 0.017 | 0.021 | 0.018 | 0.018311 | 0.013 | 0.0182 | 0.018 | 0.0175 | 0.018 | 0.0207 | 0.02 | 0.0172 | 0.0162 | 0.018 | 0.018 |  | 41
| 1 | 6 | 0.018 | 0.0125 | 0.02 | 0.0155 | 0.018175 | 0.017 | 0.017 | 0.019 | 0.017 | 0.017 | 0.021 | 0.018 | 0.020279 | 0.011 | 0.0175 | 0.018 | 0.018 | 0.018 | 0.0208 | 0.02 | 0.017 | 0.0173 | 0.0184 | 0.018
| Mean |  | 0.018 | 0.0125 | 0.02 | 0.01625 | 0.018288 | 0.0175 | 0.017167 | 0.017833 | 0.017667 | 0.016667 | 0.021 | 0.017667 | 0.018893 | 0.014333 | 0.018 | 0.0185 | 0.0178 | 0.017833 | 0.02115 | 0.02 | 0.017067 | 0.017383 | 0.0182 | 0.018
| Median |  | 0.018 | 0.0125 | 0.02 | 0.0161 | 0.018299 | 0.0175 | 0.017 | 0.018 | 0.018 | 0.017 | 0.021 | 0.018 | 0.018649 | 0.0145 | 0.018 | 0.0185 | 0.01785 | 0.018 | 0.021 | 0.02 | 0.01705 | 0.0175 | 0.01835 | 0.018
| Std Dev. |  | 0 | 1.9003e-18 | 0 | 0.00099146 | 0.00036753 | 0.00054772 | 0.00040825 | 0.0013292 | 0.0005164 | 0.0005164 | 0 | 0.0005164 | 0.0012622 | 0.0023381 | 0.00030984 | 0.00054772 | 0.00026833 | 0.00040825 | 0.00054314 | 0 | 8.165e-05 | 0.00064317 | 0.00028983 | 0
| Rel.Std.Dev. |  | 0% | 1.52e-14% | 0% | 6.101% | 2.01% | 3.13% | 2.378% | 7.453% | 2.923% | 3.098% | 0% | 2.923% | 6.681% | 16.31% | 1.721% | 2.961% | 1.507% | 2.289% | 2.568% | 0% | 0.4784% | 3.7% | 1.592% | 0%
| PDM3 |  | 1.858% | -29.27% | 13.18% | -8.045% | 3.487% | -0.9719% | -2.858% | 0.9144% | -0.02873% | -5.687% | 18.83% | -0.02873% | 6.912% | -18.89% | 1.858% | 4.687% | 0.7258% | 0.9144% | 19.68% | 13.18% | -3.424% | -1.632% | 2.989% | 1.858%
| Z-Score (Absolute) |  | 0.2 | 6.41 | 2.61 | 1.9 | 0.55 | 0.4 | 0.8 | 0 | 0.2 | 1.4 | 3.81 | 0.2 | 1.27 | 4.21 | 0.2 | 0.8 | 0.04 | 0 | 3.99 | 2.61 | 0.92 | 0.54 | 0.44 | 0.2
| Analytical results for Pb in OREAS 163b (Certified Value 56 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 56 | 50.88 | 50.2 | 66.03 | 52.7 | 54.2487 | 58 | 54 | 54.6 | 54.7 | 55 | 58 | 54.6 | 53.2172 | 97.1 | 56.2 | 54.9 | 74 | 58.8 | 58 | 57 | 55 | 59.6 | 54 | 57.8 | 1
| 1 | 2 | 58 | 49.06 | 50.6 | 67.26 | 53 | 54.6484 | 58.3 | 52 | 58.7 | 57.1 | 58.3 | 57 | 54.9 | 54.7469 | 66.7 | 57.2 | 53.31 | 56 | 59.7 | 58 | 57 | 55 | 59.1 | 53.1 | 57.6 | 17
| 1 | 3 | 56 | 50.87 | 50.6 | 67.36 | 52.9 | 58.1508 | 59.4 | 55 | 55.5 | 54.4 | 55.7 | 57 | 57 | 54.7525 | 68.6 | 57.3 | 54.33 | 58 | 57.5 | 57 | 56 | 53.8 | 58.6 | 55.8 | 57.9 | 16
| 1 | 4 | 55 | 51.13 | 49.8 | 71.46 | 52.3 | 53.9485 | 58.1 | 61.5 | 59.1 | 55.7 | 54.6 | 57 | 56.9 | 52.6879 | 79.4 | 56.3 | 57.62 | 58 | 58.6 | 57 | 56 | 52.9 | 58.2 | 55.5 | 58.9 | 55.685
| 1 | 5 | 57 | 48.5 | 50 | 65.47 | 52.4 | 53.7537 | 57.2 | 54.8 | 54.8 | 53.9 | 58.6 | 59 | 57.8 | 53.5626 | 85.4 | 57.2 | 55.68 | 59 | 57.6 | 56 | 57 | 53.3 | 59.7 | 54.1 | 58.7 | 42
| 1 | 6 | 57 | 56.93 | 49.6 | 68.81 | 53.7 | 53.1728 | 58.4 | 56.3 | 57.4 | 55.8 | 55.7 | 57 | 56.1 | 55.4833 | 66.1 | 57.3 | 54.47 | 63 | 58.3 | 56 | 57 | 55.3 | 58.6 | 53.9 | 57.7
| Mean |  | 56.5 | 51.2283 | 50.1333 | 67.7317 | 52.8333 | 54.6538 | 58.2333 | 55.6 | 56.6833 | 55.2667 | 56.3167 | 57.5 | 56.2167 | 54.0751 | 77.2167 | 56.9167 | 55.0517 | 61.3333 | 58.4167 | 57 | 56.6667 | 54.2167 | 58.9667 | 54.4 | 58.1
| Median |  | 56.5 | 50.875 | 50.1 | 67.31 | 52.8 | 54.0986 | 58.2 | 54.9 | 56.45 | 55.2 | 55.7 | 57 | 56.5 | 54.1548 | 74 | 57.2 | 54.685 | 58.5 | 58.45 | 57 | 57 | 54.4 | 58.85 | 54.05 | 57.85
| Std Dev. |  | 1.0488 | 2.9971 | 0.41312 | 2.1649 | 0.50465 | 1.7829 | 0.71181 | 3.2193 | 1.9854 | 1.1639 | 1.7081 | 0.83666 | 1.2608 | 1.0785 | 12.454 | 0.51929 | 1.4769 | 6.6232 | 0.81833 | 0.89443 | 0.5164 | 1.0147 | 0.60222 | 1.0354 | 0.55498
| Rel.Std.Dev. |  | 1.856% | 5.851% | 0.824% | 3.196% | 0.9552% | 3.262% | 1.222% | 5.79% | 3.503% | 2.106% | 3.033% | 1.455% | 2.243% | 1.994% | 16.13% | 0.9124% | 2.683% | 10.8% | 1.401% | 1.569% | 0.9113% | 1.872% | 1.021% | 1.903% | 0.9552%
| PDM3 |  | 1.464% | -8.003% | -9.97% | 21.63% | -5.121% | -1.852% | 4.576% | -0.1527% | 1.793% | -0.7513% | 1.134% | 3.259% | 0.9547% | -2.891% | 38.67% | 2.212% | -1.137% | 10.14% | 4.906% | 2.361% | 1.763% | -2.637% | 5.893% | -2.308% | 4.337%
| Z-Score (Absolute) |  | 0 | 2.05 | 2.48 | 4.37 | 1.43 | 0.72 | 0.67 | 0.35 | 0.07 | 0.48 | 0.07 | 0.39 | 0.11 | 0.94 | 8.06 | 0.16 | 0.56 | 1.88 | 0.75 | 0.19 | 0.06 | 0.89 | 0.96 | 0.82 | 0.62
| Analytical results for Pr in OREAS 163b (Certified Value 1.31 ppm) | [26 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS | 2
3
| 1 | 1 | 2.01: 0.9913; 2.02: 1.31; 2.03: 1.271; 2.04: 1.4; 2.05: 1.21; 2.06: 1.3; 2.07: 1.17 | 1
| 1 | 2 | 2.01: 1.0393; 2.02: 1.31; 2.03: 1.3166; 2.04: 1.4; 2.05: 1.22; 2.06: 1.2; 2.07: 1.13 | 18
| 1 | 3 | 2.01: 1.0098; 2.02: 1.31; 2.03: 1.3939; 2.04: 1.4; 2.05: 1.25; 2.06: 1.3; 2.07: 1.16 | 16
| 1 | 4 | 2.01: 1.042; 2.02: 1.31; 2.03: 1.3028; 2.04: 1.4; 2.05: 1.26; 2.06: 1.3; 2.07: 1.18 | 1.3111
| 1 | 5 | 2.01: 1.0078; 2.02: 1.31; 2.03: 1.3346; 2.04: 1.3; 2.05: 1.26; 2.06: 1.3; 2.07: 1.1 | 43
| 1 | 6 | 2.01: 1.0203; 2.02: 1.31; 2.03: 1.274; 2.04: 1.5; 2.05: 1.28; 2.06: 1.3; 2.07: 1.13
| Mean |  | 2.01: 1.0184; 2.02: 1.31; 2.03: 1.3155; 2.04: 1.4; 2.05: 1.2467; 2.06: 1.2833; 2.07: 1.145
| Median |  | 2.01: 1.0151; 2.02: 1.31; 2.03: 1.3097; 2.04: 1.4; 2.05: 1.255; 2.06: 1.3; 2.07: 1.145
| Std Dev. |  | 2.01: 0.019586; 2.02: 2.4324e-16; 2.03: 0.045539; 2.04: 0.063246; 2.05: 0.026583; 2.06: 0.040825; 2.07: 0.030166
| Rel.Std.Dev. |  | 2.01: 1.923%; 2.02: 1.857e-14%; 2.03: 3.462%; 2.04: 4.518%; 2.05: 2.132%; 2.06: 3.181%; 2.07: 2.635%
| PDM3 |  | 2.01: -22.32%; 2.02: -0.0836%; 2.03: 0.3344%; 2.04: 6.781%; 2.05: -4.914%; 2.06: -2.118%; 2.07: -12.67%
| Z-Score (Absolute) |  | 2.01: 4.87; 2.02: 0.49; 2.03: 0.59; 2.04: 2.15; 2.05: 0.67; 2.06: 0; 2.07: 2.54
| Analytical results for Rb in OREAS 163b (Certified Value 4.7 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 4.8 | 4.9 | 6 | 4.59 | 4.5692 | 4.3 | 4.3 | 4.3 | 4.6 | 4.4 | 5.2 | 4.7 | 5.4117 | 5.4 | 4.8 | 5.5 | 4.9 | 4.6 | 4.4 | 5.03 | 4.3 | 5.6 |  |  |  | 1
| 1 | 2 | 4.6 | 4.9 | < 5 | 4.8 | 4.6871 | 4.5 | 4.2 | 4.4 | 4.8 | 4.3 | 5.2 | 4.8 | 5.536 | 4.4 | 5.01 | 4.7 | 4.8 | 4.7 | 4.8 | 5.03 | 4.2 | 5.1 |  |  |  | 37
| 1 | 3 | 4.7 | 4.9 | 6.9 | 4.57 | 4.8405 | 5.1 | 4.5 | 5.1 | 4.7 | 4.3 | 5.4 | 4.7 | 5.4227 | 3.7 | 4.81 | 4.5 | 4.6 | 4.6 | 4.4 | 4.79 | 4.5 | 5.1 |  |  |  | 16
| 1 | 4 | 4.7 | 4.9 | 5.2 | 4.44 | 4.5037 | 4.6 | 4.6 | 4.4 | 4.7 | 4.3 | 4.8 | 5.1 | 5.5625 | 4.6 | 4.91 | 5.3 | 4.9 | 4.4 | 4.6 | 4.62 | 4.4 | 4.5 |  |  |  | 4.6992
| 1 | 5 | 4.8 | 4.9 | 6.4 | 4.73 | 4.6011 | 4.4 | 4.6 | 4.4 | 4.8 | 4.4 | 4.9 | 4.8 | 5.604 | 4.1 | 4.49 | 5.3 | 4.7 | 4.6 | 4.3 | 4.81 | 4.3 | 4.7 |  |  |  | 44
| 1 | 6 | 4.9 | 4.9 | 5.5 | 4.62 | 4.4726 | 4.3 | 4.3 | 4.6 | 4.6 | 4.2 | 5 | 4.7 | 5.592 | 3.7 | 4.97 | 4.5 | 4.9 | 4.8 | 4.4 | 4.78 | 4.5 | 5.1
| Mean |  | 4.75 | 4.9 | 6 | 4.625 | 4.6124 | 4.5333 | 4.4167 | 4.5333 | 4.7 | 4.3167 | 5.0833 | 4.8 | 5.5215 | 4.3167 | 4.8317 | 4.9667 | 4.8 | 4.6167 | 4.4833 | 4.8433 | 4.3667 | 5.0167
| Median |  | 4.75 | 4.9 | 6 | 4.605 | 4.5851 | 4.45 | 4.4 | 4.4 | 4.7 | 4.3 | 5.1 | 4.75 | 5.5493 | 4.25 | 4.86 | 5 | 4.85 | 4.6 | 4.4 | 4.8 | 4.35 | 5.1
| Std Dev. |  | 0.10488 | 9.7295e-16 | 0.68191 | 0.12661 | 0.13488 | 0.30111 | 0.17224 | 0.29439 | 0.089443 | 0.075277 | 0.22286 | 0.15492 | 0.084247 | 0.64317 | 0.18723 | 0.45019 | 0.12649 | 0.13292 | 0.18348 | 0.15971 | 0.12111 | 0.38166
| Rel.Std.Dev. |  | 2.208% | 1.986e-14% | 11.37% | 2.738% | 2.924% | 6.642% | 3.9% | 6.494% | 1.903% | 1.744% | 4.384% | 3.227% | 1.526% | 14.9% | 3.875% | 9.064% | 2.635% | 2.879% | 4.093% | 3.297% | 2.773% | 7.608%
| PDM3 |  | 1.08% | 4.272% | 27.68% | -1.58% | -1.849% | -3.53% | -6.013% | -3.53% | 0.01637% | -8.141% | 8.174% | 2.144% | 17.5% | -8.141% | 2.818% | 5.691% | 2.144% | -1.757% | -4.594% | 3.067% | -7.077% | 6.755%
| Z-Score (Absolute) |  | 0.09 | 0.62 | 2.43 | 0.35 | 0.4 | 0.67 | 1.08 | 0.67 | 0.09 | 1.44 | 1.26 | 0.26 | 2.8 | 1.44 | 0.38 | 0.85 | 0.26 | 0.38 | 0.85 | 0.42 | 1.26 | 1.03
| Analytical results for Re in OREAS 163b (Certified Value < 0.002 ppm) | [26 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.23 | 2.24 | 2.25 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.005 | < 0.01 | 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.001 | < 0.05 | < 0.002 | < 0.002 | < 0.003 | < 0.005 | 0.002 | < 0.002 | < 0.005 | [8 empty] | 1
| 1 | 2 | < 0.05 | < 0.005 | < 0.01 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | 0.004 | < 0.003 | < 0.005 |  | < 0.002 | < 0.005 | [8 empty] | 38
| 1 | 3 | < 0.05 | < 0.005 | < 0.01 | 0.002 | < 0.002 | < 0.002 | 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | 0.003 | < 0.003 | < 0.005 |  | < 0.002 | < 0.005 | [8 empty] | 16
| 1 | 4 | < 0.05 | < 0.005 | < 0.01 | 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | 0.002 | < 0.003 | < 0.005 |  | < 0.002 | < 0.005 | [8 empty] | < 0.002
| 1 | 5 | < 0.05 | < 0.005 | < 0.01 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | 0.003 | < 0.003 | < 0.005 |  | < 0.002 | < 0.005 | [8 empty] | 45
| 1 | 6 | < 0.05 | < 0.005 | < 0.01 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.001 | < 0.05 | < 0.002 | 0.003 | 0.003 | < 0.005 |  | < 0.002 | < 0.005
| Mean |  | 2.05: 0.002; 2.06: 0.002; 2.08: 0.002; 2.09: 0.002; 2.13: 0.003; 2.14: 0.003; 2.16: 0.002
| Median |  | 2.05: 0.002; 2.06: 0.002; 2.08: 0.002; 2.09: 0.002; 2.13: 0.003; 2.14: 0.003; 2.16: 0.002
| Std Dev. |  | 2.05: 0; 2.13: 0.00070711
| Rel.Std.Dev. |  | 2.05: 0%; 2.13: 23.57%
PDM3
| Z-Score (Absolute) |  | 23.43 | 0.67 | 3.2 | 0 | 0.67 | 0.84 | 0.67 | 0.67 | 1.35 | 23.43 | 0.84 | 0.84 | 0.08 | 0.67 | 0.17 | 0.84 | 0.67
| Analytical results for S in OREAS 163b (Certified Value 3.35 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 3.51 | 3.673 | 3.3 | 2.9 | 3.22 | 3.5523 | 3.39 | 3.29 | 3.22 | 3.28 | 3.31 | 3.43 | 3.271 | 3.286 | 3.49 | 3.5181 | 3.46 | 3.285 | 3.2 | 3.15 | 3.19 | 3.41 | 3.35 | 3.62 | 3.13 | 1
| 1 | 2 | 3.49 | 3.652 | 3.3 | 2.77 | 3.47 | 3.4704 | 3.42 | 3.26 | 3.46 | 3.3 | 3.41 | 3.4 | 3.283 | 3.6642 | 3.46 | 3.4294 | 3.53 | 3.2921 | 3.2 | 3.12 | 3.19 | 3.39 | 3.47 | 3.62 | 3.07 | 0
| 1 | 3 | 3.52 | 3.6536 | 3.3 | 2.69 | 3.34 | 3.4898 | 3.39 | 3.41 | 3.35 | 3.31 | 3.35 | 3.42 | 3.403 | 3.5325 | 3.27 | 3.3854 | 3.4 | 3.2778 | 3.2 | 3.14 | 3.19 | 3.39 | 3.3 | 3.72 | 3.11 | 16
| 1 | 4 | 3.5 | 3.6924 | 3.3 | 2.93 | 3.21 | 3.5492 | 3.29 | 3.31 | 3.41 | 3.21 | 3.39 | 3.41 | 3.303 | 3.3052 | 3.66 | 3.4685 | 3.54 | 3.3188 | 3.2 | 3.14 | 3.19 | 3.42 | 3.35 | 3.59 | 3.17 | 3.3486
| 1 | 5 | 3.49 | 3.697 | 3.3 | 2.63 | 3.46 | 3.5864 | 3.44 | 3.37 | 3.27 | 3.23 | 3.48 | 3.57 | 3.419 | 3.6137 | 2.62 | 3.4638 | 3.39 | 3.2929 | 3.2 | 3.14 | 3.18 | 3.4 | 3.56 | 3.6 | 3.09 | 46
| 1 | 6 | 3.5 | 3.665 | 3.3 | 2.57 | 3.33 | 3.5003 | 3.29 | 3.35 | 3.34 | 3.22 | 3.54 | 3.48 | 3.342 | 3.5401 | 3.33 | 3.3845 | 3.34 | 3.2349 | 3.2 | 3.09 | 3.19 | 3.35 | 3.3 | 3.71 | 3.17
| Mean |  | 3.5017 | 3.6722 | 3.3 | 2.7483 | 3.3383 | 3.5247 | 3.37 | 3.3317 | 3.3417 | 3.2583 | 3.4133 | 3.4517 | 3.3368 | 3.4903 | 3.305 | 3.4416 | 3.4433 | 3.2836 | 3.2 | 3.13 | 3.1883 | 3.3933 | 3.3883 | 3.6433 | 3.1233
| Median |  | 3.5 | 3.669 | 3.3 | 2.73 | 3.335 | 3.5248 | 3.39 | 3.33 | 3.345 | 3.255 | 3.4 | 3.425 | 3.3225 | 3.5363 | 3.395 | 3.4466 | 3.43 | 3.2885 | 3.2 | 3.14 | 3.19 | 3.395 | 3.35 | 3.62 | 3.12
| Std Dev. |  | 0.01169 | 0.019129 | 0 | 0.14538 | 0.11197 | 0.04458 | 0.064807 | 0.055287 | 0.087958 | 0.043551 | 0.084538 | 0.064317 | 0.062503 | 0.15857 | 0.36204 | 0.052223 | 0.080664 | 0.027586 | 4.8648e-16 | 0.021909 | 0.0040825 | 0.024221 | 0.10458 | 0.056804 | 0.041312
| Rel.Std.Dev. |  | 0.3339% | 0.5209% | 0% | 5.29% | 3.354% | 1.265% | 1.923% | 1.659% | 2.632% | 1.337% | 2.477% | 1.863% | 1.873% | 4.543% | 10.95% | 1.517% | 2.343% | 0.8401% | 1.52e-14% | 0.7% | 0.128% | 0.7138% | 3.086% | 1.559% | 1.323%
| PDM3 |  | 4.572% | 9.663% | -1.451% | -17.93% | -0.306% | 5.261% | 0.6397% | -0.5051% | -0.2064% | -2.695% | 1.934% | 3.079% | -0.3508% | 4.232% | -1.301% | 2.778% | 2.83% | -1.941% | -4.437% | -6.528% | -4.785% | 1.337% | 1.187% | 8.802% | -6.727%
| Z-Score (Absolute) |  | 1.08 | 2.23 | 0.28 | 4 | 0.02 | 1.24 | 0.19 | 0.07 | 0 | 0.56 | 0.48 | 0.74 | 0.03 | 1 | 0.25 | 0.67 | 0.69 | 0.39 | 0.96 | 1.43 | 1.03 | 0.35 | 0.31 | 2.04 | 1.47
| Analytical results for Sb in OREAS 163b (Certified Value 8.46 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 8.5 | 8.461 | 9.5 | < 5 | 8.76 | 9.3268 | 9.11 | 8.18 | 8.19 | 8.66 | 8.64 | 6.4 | 7.76 | 6.9998 | 12 | 9.34 | 9.67 | 8.64 | 8.2 | 3.81 | 6.04 | 7.82 | 8.8 | 7.9 | 5.2 | 1
| 1 | 2 | 8.3 | 8.5024 | 9.5 | < 5 | 8.7 | 9.7177 | 9.45 | 8.4 | 8.58 | 8.95 | 8.74 | 6.2 | 7.77 | 7.4515 | 8 | 9.41 | 8.46 | 8.53 | 8 | 3.95 | 7.54 | 8.23 | 8.71 | 7.9 | 5.2 | 39
| 1 | 3 | 8.3 | 8.3856 | 9.5 | < 5 | 8.92 | 10.1222 | 9.76 | 8.49 | 8.65 | 8.71 | 8.59 | 9.7 | 8.05 | 7.3828 | 8.74 | 9.41 | 8.4 | 8.49 | 7.9 | 4.05 | 6.67 | 8.35 | 8.99 | 8.2 | 5.4 | 16
| 1 | 4 | 8.2 | 8.295 | 9.5 | < 5 | 8.55 | 9.4515 | 9.65 | 8.52 | 8.24 | 8.54 | 8.83 | 8.9 | 6.95 | 6.9535 | 7.89 | 9.4 | 9.1 | 8.69 | 8.2 | 4.2 | 7.96 | 7.93 | 8.5 | 8.2 | 5.3 | 8.464
| 1 | 5 | 8.4 | 8.428 | 9.5 | < 5 | 8.9 | 9.4376 | 9.42 | 8.59 | 8.18 | 8.55 | 9.1 | 8.5 | 7.32 | 7.1723 | 9.69 | 9.2 | 9.68 | 8.74 | 8 | 3.85 | 7.08 | 7.92 | 8.41 | 7.9 | 5.3 | 47
| 1 | 6 | 8.2 | 8.5972 | 9.5 | < 5 | 8.77 | 9.2851 | 9.22 | 8.59 | 8.29 | 8.56 | 8.57 | 6.6 | 7.65 | 7.5007 | 7.7 | 9.52 | 8.34 | 8.65 | 8.1 | 3.88 | 7.49 | 8.17 | 8.2 | 8.5 | 5.6
| Mean |  | 8.3167 | 8.4449 | 9.5 |  | 8.7667 | 9.5568 | 9.435 | 8.4617 | 8.355 | 8.6617 | 8.745 | 7.7167 | 7.5833 | 7.2434 | 9.0033 | 9.38 | 8.9417 | 8.6233 | 8.0667 | 3.9567 | 7.13 | 8.07 | 8.6017 | 8.1 | 5.3333
| Median |  | 8.3 | 8.4445 | 9.5 |  | 8.765 | 9.4446 | 9.435 | 8.505 | 8.265 | 8.61 | 8.69 | 7.55 | 7.705 | 7.2775 | 8.37 | 9.405 | 8.78 | 8.645 | 8.05 | 3.915 | 7.285 | 8.05 | 8.605 | 8.05 | 5.3
| Std Dev. |  | 0.1169 | 0.10299 | 0 |  | 0.13619 | 0.31546 | 0.24664 | 0.1551 | 0.20637 | 0.1569 | 0.19947 | 1.4986 | 0.38955 | 0.23551 | 1.6411 | 0.10564 | 0.63057 | 0.095429 | 0.12111 | 0.1461 | 0.6912 | 0.20909 | 0.28646 | 0.24495 | 0.15055
| Rel.Std.Dev. |  | 1.406% | 1.22% | 0% |  | 1.553% | 3.301% | 2.614% | 1.833% | 2.47% | 1.811% | 2.281% | 19.42% | 5.137% | 3.251% | 18.23% | 1.126% | 7.052% | 1.107% | 1.501% | 3.693% | 9.694% | 2.591% | 3.33% | 3.024% | 2.823%
| PDM3 |  | -1.741% | -0.2259% | 12.24% |  | 3.576% | 12.91% | 11.47% | -0.02741% | -1.288% | 2.336% | 3.32% | -8.829% | -10.4% | -14.42% | 6.372% | 10.82% | 5.644% | 1.883% | -4.694% | -53.25% | -15.76% | -4.655% | 1.627% | -4.3% | -36.99%
| Z-Score (Absolute) |  | 0.17 | 0 | 1.43 | 8.07 | 0.44 | 1.51 | 1.34 | 0.02 | 0.12 | 0.29 | 0.41 | 0.99 | 1.17 | 1.63 | 0.76 | 1.27 | 0.67 | 0.24 | 0.51 | 6.09 | 1.78 | 0.51 | 0.21 | 0.47 | 4.22
| Analytical results for Sc in OREAS 163b (Certified Value 1.02 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1 | 0.95472 | 1 | < 0.5 | 1 | 1.0444 | 1 | 1 | 1.1 | 1.1 | 1 | 1 | 1 | 1.1138 | 1.1 | 1 | 1 | 1 | 0.7 | 1.2 | 0.9 | 1.1 | 0.9 | 1.4 |  | 1
| 1 | 2 | 1 | 0.98566 | 1 | < 0.5 | 1 | 1.0405 | 1 | 1 | 1 | 1.2 | 1.1 | 1 | 1 | 1.1382 | 0.7 | 0.9 | 1 | 1 | 0.7 | 1.3 | 1 | 1.1 | 0.9 | 1.3 |  | 19
| 1 | 3 | 1 | 0.97954 | 1 | < 0.5 | 1 | 1.104 | 1 | 1 | 1.1 | 1.1 | 1 | 1 | 0.9 | 1.0197 | 0.7 | 1.1 | 1.2 | < 1 | 0.7 | 1.3 | 0.9 | 1.1 | 1 | 1.3 |  | 16
| 1 | 4 | 1 | 0.98972 | 1 | < 0.5 | 1 | 1.0497 | 1 | 1.1 | 1 | 1.1 | 1 | 1 | 1 | 1.0141 | 0.9 | 0.9 | 1 | 1 | 0.7 | 1.2 | 0.9 | 1.1 | 1 | 1.3 |  | 1.0166
| 1 | 5 | 1 | 1.0049 | 1 | < 0.5 | 1 | 1.0077 | 1 | 1.1 | 1 | 1.1 | 1 | 1 | 1 | 1.0957 | 0.7 | 1 | 1.1 | 1 | 0.7 | 1.2 | 0.9 | 1.1 | 1 | 1.3 |  | 48
| 1 | 6 | 1 | 0.97614 | 1 | < 0.5 | 1.1 | 1.0063 | 1 | 1 | 1 | 1.1 | 1 | 1 | 1 | 1.1298 | 0.8 | 1 | 1.1 | 1 | 0.7 | 1.3 | 1 | 1 | 1.1 | 1.3
| Mean |  | 1 | 0.98179 | 1 |  | 1.0167 | 1.0421 | 1 | 1.0333 | 1.0333 | 1.1167 | 1.0167 | 1 | 0.98333 | 1.0852 | 0.81667 | 0.98333 | 1.0667 | 1 | 0.7 | 1.25 | 0.93333 | 1.0833 | 0.98333 | 1.3167
| Median |  | 1 | 0.9826 | 1 |  | 1 | 1.0424 | 1 | 1 | 1 | 1.1 | 1 | 1 | 1 | 1.1047 | 0.75 | 1 | 1.05 | 1 | 0.7 | 1.25 | 0.9 | 1.1 | 1 | 1.3
| Std Dev. |  | 0 | 0.016635 | 0 |  | 0.040825 | 0.035667 | 0 | 0.05164 | 0.05164 | 0.040825 | 0.040825 | 0 | 0.040825 | 0.054877 | 0.16021 | 0.075277 | 0.08165 | 0 | 1.2162e-16 | 0.054772 | 0.05164 | 0.040825 | 0.075277 | 0.040825
| Rel.Std.Dev. |  | 0% | 1.694% | 0% |  | 4.016% | 3.423% | 0% | 4.997% | 4.997% | 3.656% | 4.016% | 0% | 4.152% | 5.057% | 19.62% | 7.655% | 7.655% | 0% | 1.737e-14% | 4.382% | 5.533% | 3.768% | 7.655% | 3.101%
| PDM3 |  | -1.635% | -3.426% | -1.635% |  | 0.004547% | 2.508% | -1.635% | 1.644% | 1.644% | 9.841% | 0.004547% | -1.635% | -3.274% | 6.748% | -19.67% | -3.274% | 4.923% | -1.635% | -31.14% | 22.96% | -8.193% | 6.562% | -3.274% | 29.51%
| Z-Score (Absolute) |  | 0 | 0.33 | 0 | 13.41 | 0.3 | 0.75 | 0 | 0.6 | 0.6 | 2.09 | 0.3 | 0 | 0.3 | 1.52 | 3.28 | 0.3 | 1.19 | 1.49 | 5.36 | 4.47 | 1.19 | 1.49 | 0.3 | 5.66
| Analytical results for Se in OREAS 163b (Certified Value 1.69 ppm) | [26 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 2 | 1.3343 | 2 | < 5 | 1.4 | 2.1325 | 3 | 3 | 2 | 2 | 2 | 1.9 | < 1 | < 2 | 3 | 1.9 | 1.3 | 1.51 | 2 | < 5 | 2 | 1.5 | 1.2 | 1 |  | 1
| 1 | 2 | 2 | 1.2954 | 2 | < 5 | 1.5 | 1.9515 | 3 | 3 | 2 | 2 | 2 | 2 | < 1 | < 2 | < 1 | 2 | 1.4 | 1.44 | 2 | < 5 | 5 | 1.5 | 1.4 | 2 |  | 41
| 1 | 3 | 2 | 1.3144 | 2 | < 5 | 1.6 | 2.044 | 4 | 3 | 2 | 2 | 2 | 1.9 | < 1 | < 2 | 1 | 2.1 | 1.6 | 1.43 | 2 | < 5 | 4 | 1.6 | 1.7 | 2 |  | 16
| 1 | 4 | 1 | 1.3527 | 2 | < 5 | 1.6 | 2.0295 | 2 | 3 | 2 | 2 | 2 | 1.5 | < 1 | < 2 | 3 | 1.4 | 1.4 | 1.42 | 2 | < 5 | 4 | 1.5 | 1.6 | 2 |  | 1.6879
| 1 | 5 | 2 | 1.3034 | 2 | < 5 | 1.4 | 2.0302 | 3 | 4 | 2 | 2 | 1 | 1.7 | < 1 | < 2 | < 1 | 1.6 | 1.3 | 1.46 | 2 | < 5 | 5 | 1.4 | 1.6 | 2 |  | 49
| 1 | 6 | 1 | 1.3542 | 2 | < 5 | 1.8 | 2.009 | 3 | 3 | 2 | 2 | 2 | 1.7 | < 1 | < 2 | 1 | 1.9 | 1.4 | 2.07 | 2 | < 5 | 5 | 1.5 | 1.5 | 2
| Mean |  | 1.6667 | 1.3257 | 2 |  | 1.55 | 2.0328 | 3 | 3.1667 | 2 | 2 | 1.8333 | 1.7833 |  |  | 2 | 1.8167 | 1.4 | 1.555 | 2 |  | 4.1667 | 1.5 | 1.5 | 1.8333
| Median |  | 2 | 1.3243 | 2 |  | 1.55 | 2.0299 | 3 | 3 | 2 | 2 | 2 | 1.8 |  |  | 2 | 1.9 | 1.4 | 1.45 | 2 |  | 4.5 | 1.5 | 1.55 | 2
| Std Dev. |  | 0.5164 | 0.025133 | 0 |  | 0.15166 | 0.058734 | 0.63246 | 0.40825 | 0 | 0 | 0.40825 | 0.18348 |  |  | 1.1547 | 0.26394 | 0.10954 | 0.2543 | 0 |  | 1.169 | 0.063246 | 0.17889 | 0.40825
| Rel.Std.Dev. |  | 30.98% | 1.896% | 0% |  | 9.784% | 2.889% | 21.08% | 12.89% | 0% | 0% | 22.27% | 10.29% |  |  | 57.74% | 14.53% | 7.825% | 16.35% | 0% |  | 28.06% | 4.216% | 11.93% | 22.27%
| PDM3 |  | -1.256% | -21.46% | 18.49% |  | -8.169% | 20.44% | 77.74% | 87.61% | 18.49% | 18.49% | 8.618% | 5.656% |  |  | 18.49% | 7.63% | -17.06% | -7.872% | 18.49% |  | 146.9% | -11.13% | -11.13% | 8.618%
| Z-Score (Absolute) |  | 0.36 | 1.12 | 0.39 | 1.52 | 0.62 | 0.47 | 2.64 | 3.02 | 0.39 | 0.39 | 0.02 | 0.09 | 2.98 | 1.85 | 0.73 | 0.02 | 0.96 | 0.61 | 0.39 | 1.52 | 5.26 | 0.73 | 0.73 | 0.02
| Analytical results for Sm in OREAS 163b (Certified Value 1.29 ppm) | [26 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 1.1556; 2.02: 1.34; 2.03: 2.6681; 2.04: 1.1; 2.05: 1.34; 2.06: 1.3; 2.07: 1.28 |  | 1
| 1 | 2 | 2.01: 1.2071; 2.02: 1.344; 2.03: 2.6701; 2.04: 1.5; 2.05: 1.42; 2.06: 1.3; 2.07: 1.2 |  | 5
| 1 | 3 | 2.01: 1.1379; 2.02: 1.34; 2.03: 2.7396; 2.04: 1.1; 2.05: 1.37; 2.06: 1.4; 2.07: 1.27 |  | 16
| 1 | 4 | 2.01: 1.2023; 2.02: 1.34; 2.03: 2.7104; 2.04: 1.4; 2.05: 1.27; 2.06: 1.3; 2.07: 1.34 |  | 1.2851
| 1 | 5 | 2.01: 1.1826; 2.02: 1.348; 2.03: 2.7675; 2.04: 1.4; 2.05: 1.3; 2.06: 1.3; 2.07: 1.16 |  | 50
| 1 | 6 | 2.01: 1.1672; 2.02: 1.38; 2.03: 2.7363; 2.04: 1; 2.05: 1.42; 2.06: 1.4; 2.07: 1.25
| Mean |  | 2.01: 1.1755; 2.02: 1.3487; 2.03: 2.7153; 2.04: 1.25; 2.05: 1.3533; 2.06: 1.3333; 2.07: 1.25
| Median |  | 2.01: 1.1749; 2.02: 1.342; 2.03: 2.7233; 2.04: 1.25; 2.05: 1.355; 2.06: 1.3; 2.07: 1.26
| Std Dev. |  | 2.01: 0.027003; 2.02: 0.01568; 2.03: 0.040117; 2.04: 0.20736; 2.05: 0.06186; 2.06: 0.05164; 2.07: 0.063246
| Rel.Std.Dev. |  | 2.01: 2.297%; 2.02: 1.163%; 2.03: 1.477%; 2.04: 16.59%; 2.05: 4.571%; 2.06: 3.873%; 2.07: 5.06%
| PDM3 |  | 2.01: -8.535%; 2.02: 4.944%; 2.03: 111.3%; 2.04: -2.734%; 2.05: 5.307%; 2.06: 3.751%; 2.07: -2.734%
| Z-Score (Absolute) |  | 2.01: 1.28; 2.02: 0.12; 2.03: 11.18; 2.04: 0.67; 2.05: 0.16; 2.06: 0; 2.07: 0.67
| Analytical results for Sn in OREAS 163b (Certified Value 2.41 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.3 | 1.8201 | 4.2 | < 5 | 2.1 | 2.6681 | 2.5 | 2.3 | 2.7 | 2.5 | 2.3 | 3 | 2.4 | 2.1521 | 2.3 | 2.5 | 3.2 | 2.3 | < 10 | 2.4 | 2.28 | 2.5 | 2.2 | 2.5 |  | 1
| 1 | 2 | 2.4 | 1.9009 | 4.2 | < 5 | 2.3 | 2.6701 | 2.6 | 2.4 | 2.7 | 2.3 | 2.3 | 3 | 2.4 | 2.3971 | 2 | 2.6 | 3.6 | 2.2 | < 10 | 2.6 | 2.5 | 2.6 | 2.2 | 2.4 |  | 26
| 1 | 3 | 2.3 | 1.7267 | 4.3 | < 5 | 2.2 | 2.7396 | 2.6 | 2.5 | 3.1 | 2.3 | 2.1 | 3 | 2.4 | 2.6414 | 1.8 | 2.5 | 2.4 | 2.2 | < 10 | 2.4 | 2.3 | 2.5 | 2.4 | 2.5 |  | 16
| 1 | 4 | 2.4 | 1.7751 | 4.3 | < 5 | 2.3 | 2.7104 | 2.6 | 2.4 | 2.7 | 2.5 | 2.2 | 3 | 2.4 | 2.3898 | 2.2 | 2.7 | 3.6 | 2.3 | < 10 | 2.4 | 2.28 | 2.6 | 2.4 | 2.5 |  | 2.41
| 1 | 5 | 2.4 | 1.6308 | 4.3 | < 5 | 2.4 | 2.7675 | 2.5 | 2.3 | 2.7 | 2.3 | 2.4 | 3 | 2.6 | 2.3641 | 2 | 2.4 | 3 | 2.2 | < 10 | 2.4 | 2.22 | 2.5 | 2.3 | 2.5 |  | 51
| 1 | 6 | 2.4 | 1.9243 | 4.2 | < 5 | 2.4 | 2.7363 | 2.4 | 2.3 | 2.7 | 2.3 | 2.2 | 3 | 2.4 | 2.7047 | 1.8 | 2.4 | 2.3 | 2.3 | < 10 | 2.3 | 2.24 | 2.6 | 2.4 | 2.5
| Mean |  | 2.3667 | 1.7963 | 4.25 |  | 2.2833 | 2.7153 | 2.5333 | 2.3667 | 2.7667 | 2.3667 | 2.25 | 3 | 2.4333 | 2.4415 | 2.0167 | 2.5167 | 3.0167 | 2.25 |  | 2.4167 | 2.3033 | 2.55 | 2.3167 | 2.4833
| Median |  | 2.4 | 1.7976 | 4.25 |  | 2.3 | 2.7233 | 2.55 | 2.35 | 2.7 | 2.3 | 2.25 | 3 | 2.4 | 2.3934 | 2 | 2.5 | 3.1 | 2.25 |  | 2.4 | 2.28 | 2.55 | 2.35 | 2.5
| Std Dev. |  | 0.05164 | 0.11007 | 0.054772 |  | 0.1169 | 0.040117 | 0.08165 | 0.08165 | 0.1633 | 0.10328 | 0.10488 | 0 | 0.08165 | 0.2018 | 0.20412 | 0.1169 | 0.56716 | 0.054772 |  | 0.098319 | 0.10073 | 0.054772 | 0.098319 | 0.040825
| Rel.Std.Dev. |  | 2.182% | 6.127% | 1.289% |  | 5.12% | 1.477% | 3.223% | 3.45% | 5.902% | 4.364% | 4.661% | 0% | 3.355% | 8.265% | 10.12% | 4.645% | 18.8% | 2.434% |  | 4.068% | 4.373% | 2.148% | 4.244% | 1.644%
| PDM3 |  | -1.796% | -25.46% | 76.35% |  | -5.254% | 12.67% | 5.12% | -1.796% | 14.8% | -1.796% | -6.637% | 24.48% | 0.9701% | 1.31% | -16.32% | 4.428% | 25.18% | -6.637% |  | 0.2785% | -4.424% | 5.811% | -3.871% | 3.045%
| Z-Score (Absolute) |  | 0.39 | 3.68 | 10.5 | 0.39 | 0.87 | 1.63 | 0.58 | 0.39 | 1.93 | 0.39 | 1.06 | NA | 0 | 0.05 | 2.41 | 0.48 | 3.37 | 1.06 | 14.83 | 0.1 | 0.75 | 0.67 | 0.67 | 0.29
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 163b (Certified Value 9.39 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 9.2 | 8.36 | 11 | 10.28 | 8.62 | 10.2311 | 8.7 | 8.8 | 10 | 9.2 | 8.7 | 11.2 | 9.5 | 10.5109 | 6 | 9.07 | 10 | 9 | 18.4 | 11.1 | 8.9 | 9.77 | 9.5 | 11 |  | 1
| 1 | 2 | 9.5 | 8.48 | 11 | 10.22 | 8.94 | 10.2625 | 9 | 9.1 | 10.2 | 9.4 | 8.9 | 9.7 | 9.5 | 9.5184 | 6 | 9.32 | 10 | 9 | 18.2 | 10.3 | 9.4 | 9.39 | 9.3 | 10 |  | 27
| 1 | 3 | 9.1 | 8.58 | 11 | 9.5 | 8.85 | 9.795 | 10.2 | 9 | 10.2 | 9.4 | 8.6 | 10.1 | 10 | 9.3421 | 6 | 9.29 | 10 | 9 | 18.4 | 10.9 | 9.1 | 9.57 | 9.4 | 11 |  | 16
| 1 | 4 | 9.1 | 8.16 | 11 | 9.86 | 8.56 | 10.1298 | 9.1 | 8.9 | 10.9 | 9 | 8.7 | 9.3 | 9.9 | 10.487 | 7 | 9.15 | 10 | 9 | 18.1 | 10.1 | 8.9 | 9.42 | 9.4 | 10 |  | 9.3935
| 1 | 5 | 9.3 | 7.49 | 11 | 9.24 | 8.78 | 9.9492 | 8.9 | 9.5 | 9.8 | 9.4 | 9 | 9.7 | 9.9 | 9.2119 | 5 | 8.95 | 10 | 9 | 18 | 10.1 | 8.8 | 9.47 | 8.8 | 11 |  | 52
| 1 | 6 | 9.4 | 8.35 | 11 | 9.53 | 8.88 | 10.2912 | 8.7 | 9 | 10.4 | 8.9 | 8.6 | 9.2 | 9.8 | 10.1778 | 6 | 9.14 | 10 | 9 | 18.8 | 10 | 9.1 | 9.48 | 9.2 | 10
| Mean |  | 9.2667 | 8.2367 | 11 | 9.7717 | 8.7717 | 10.1098 | 9.1 | 9.05 | 10.25 | 9.2167 | 8.75 | 9.8667 | 9.7667 | 9.8747 | 6 | 9.1533 | 10 | 9 | 18.3167 | 10.4167 | 9.0333 | 9.5167 | 9.2667 | 10.5
| Median |  | 9.25 | 8.355 | 11 | 9.695 | 8.815 | 10.1804 | 8.95 | 9 | 10.2 | 9.3 | 8.7 | 9.7 | 9.85 | 9.8481 | 6 | 9.145 | 10 | 9 | 18.3 | 10.2 | 9 | 9.475 | 9.35 | 10.5
| Std Dev. |  | 0.1633 | 0.39205 | 0 | 0.42002 | 0.15105 | 0.19808 | 0.56214 | 0.2429 | 0.37815 | 0.22286 | 0.16432 | 0.72847 | 0.21602 | 0.58677 | 0.63246 | 0.13779 | 0 | 0 | 0.28577 | 0.46655 | 0.21602 | 0.13852 | 0.25033 | 0.54772
| Rel.Std.Dev. |  | 1.762% | 4.76% | 0% | 4.298% | 1.722% | 1.959% | 6.177% | 2.684% | 3.689% | 2.418% | 1.878% | 7.383% | 2.212% | 5.942% | 10.54% | 1.505% | 0% | 0% | 1.56% | 4.479% | 2.391% | 1.456% | 2.701% | 5.216%
| PDM3 |  | -1.351% | -12.32% | 17.1% | 4.025% | -6.62% | 7.625% | -3.125% | -3.657% | 9.117% | -1.883% | -6.851% | 5.037% | 3.972% | 5.122% | -36.13% | -2.557% | 6.456% | -4.19% | 94.99% | 10.89% | -3.835% | 1.311% | -1.351% | 11.78%
| Z-Score (Absolute) |  | 0 | 1.38 | NA | 0.68 | 0.66 | 1.13 | 0.22 | 0.29 | 1.32 | 0.07 | 0.69 | 0.81 | 0.67 | 0.82 | 4.38 | 0.15 | NA | NA | 12.14 | 1.54 | 0.31 | 0.34 | 0 | NA
Results from laboratories 2.03, 2.17, 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 163b (Certified Value 2.05 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 2.07 | 3.1566 | 1.5 | < 5 | 2.07 | 2.4871 | 2.34 | 1.74 | 2.04 | 2.15 | 2.48 | 1.1 | 0.83 | 1.6444 | 0.17 | 2.7 | 1.3 | 1.5 | 1.99 | 1.12 | 3.16 | 2.6 | 2.8 |  |  | 1
| 1 | 2 | 2.01 | 3.1807 | 1.5 | < 5 | 1.97 | 2.6882 | 2.51 | 2.3 | 2.47 | 2.64 | 2.37 | 1 | 0.73 | 2.2665 | 0.19 | 2.69 | 1.1 | 1.4 | 2.09 | 1.95 | 2.42 | 2.7 | 2 |  |  | 10
| 1 | 3 | 2.11 | 3.1776 | 1.6 | < 5 | 1.94 | 2.4077 | 3.02 | 2.59 | 3.32 | 3.21 | 2.04 | 1.3 | 0.81 | 2.1422 | 0.14 | 3.1 | 1.3 | 1.1 | 1.93 | 1.03 | 2.4 | 2.52 | 2.6 |  |  | 16
| 1 | 4 | 2.17 | 3.179 | 1.5 | < 5 | 2.14 | 2.0706 | 1.83 | 1.71 | 1.99 | 3.12 | 1.9 | 1.4 | 0.87 | 1.7573 | 0.17 | 2.41 | 1.2 | 1.7 | 1.98 | 1.65 | 2.83 | 2.52 | 2.8 |  |  | 2.0539
| 1 | 5 | 2.07 | 3.1803 | 1.5 | < 5 | 2.07 | 2.0626 | 1.98 | 1.82 | 2.19 | 2.6 | 1.73 | 1.4 | 0.82 | 1.9639 | 0.25 | 2.7 | 1.2 | 1.5 | 2.07 | 1.18 | 1.99 | 2.6 | 2.3 |  |  | 53
| 1 | 6 | 1.92 | 3.1966 | 1.6 | < 5 | 1.79 | 2.364 | 2.3 | 2.24 | 2.07 | 1.85 | 2.7 | 1.6 | 0.79 | 2.2775 | 0.16 | 3.08 | 1.5 | 1.2 | 1.99 | 1.53 | 2.58 | 2.82 | 2.5
| Mean |  | 2.0583 | 3.1785 | 1.5333 |  | 1.9967 | 2.3467 | 2.33 | 2.0667 | 2.3467 | 2.595 | 2.2033 | 1.3 | 0.80833 | 2.0086 | 0.18 | 2.78 | 1.2667 | 1.4 | 2.0083 | 1.41 | 2.5633 | 2.6267 | 2.5
| Median |  | 2.07 | 3.1797 | 1.5 |  | 2.02 | 2.3858 | 2.32 | 2.03 | 2.13 | 2.62 | 2.205 | 1.35 | 0.815 | 2.0531 | 0.17 | 2.7 | 1.25 | 1.45 | 1.99 | 1.355 | 2.5 | 2.6 | 2.55
| Std Dev. |  | 0.085888 | 0.012794 | 0.05164 |  | 0.12485 | 0.24385 | 0.42 | 0.36142 | 0.50694 | 0.53084 | 0.37249 | 0.21909 | 0.046655 | 0.26627 | 0.037947 | 0.26465 | 0.13663 | 0.21909 | 0.060139 | 0.35917 | 0.40083 | 0.1157 | 0.30984
| Rel.Std.Dev. |  | 4.173% | 0.4025% | 3.368% |  | 6.253% | 10.39% | 18.03% | 17.49% | 21.6% | 20.46% | 16.91% | 16.85% | 5.772% | 13.26% | 21.08% | 9.52% | 10.79% | 15.65% | 2.994% | 25.47% | 15.64% | 4.405% | 12.39%
| PDM3 |  | 0.2146% | 54.75% | -25.35% |  | -2.788% | 14.25% | 13.44% | 0.6203% | 14.25% | 26.34% | 7.274% | -36.71% | -60.64% | -2.204% | -91.24% | 35.35% | -38.33% | -31.84% | -2.22% | -31.35% | 24.8% | 27.89% | 21.72%
| Z-Score (Absolute) |  | 0.01 | 1.51 | 0.72 | 0.59 | 0.1 | 0.38 | 0.36 | 0 | 0.38 | 0.72 | 0.19 | 1.04 | 1.71 | 0.08 | 2.56 | 0.97 | 1.09 | 0.91 | 0.08 | 0.89 | 0.67 | 0.76 | 0.59
| Analytical results for Tb in OREAS 163b (Certified Value 0.16 ppm) | [26 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.16; 2.08: 2.17; 2.09: 2.20; 2.10: 2.21; 2.11: 2.22 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*OES/MS; 2.11: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 0.1149; 2.02: 0.18; 2.03: < 0.5; 2.04: 0.17509; 2.05: 0.2; 2.06: 0.17; 2.07: 0.17; 2.08: 0.2; 2.09: 0.14; 2.10: 0.15; 2.11: 0.2 |  |  | 1
| 1 | 2 | 2.01: 0.1175; 2.02: 0.18; 2.03: < 0.5; 2.04: 0.18282; 2.05: 0.2; 2.06: 0.17; 2.07: 0.17; 2.08: 0.2; 2.09: 0.14; 2.10: 0.14; 2.11: 0.22 |  |  | 11
| 1 | 3 | 2.01: 0.1149; 2.02: 0.18; 2.03: < 0.5; 2.04: 0.18138; 2.05: 0.3; 2.06: 0.16; 2.07: 0.17; 2.08: 0.2; 2.09: 0.13; 2.10: 0.15; 2.11: 0.2 |  |  | 16
| 1 | 4 | 2.01: 0.1229; 2.02: 0.18; 2.03: < 0.5; 2.04: 0.1763; 2.05: 0.2; 2.06: 0.17; 2.07: 0.17; 2.08: 0.2; 2.09: 0.14; 2.10: 0.14; 2.11: 0.2 |  |  | 0.16284
| 1 | 5 | 2.01: 0.1315; 2.02: 0.18; 2.03: < 0.5; 2.04: 0.17689; 2.05: 0.2; 2.06: 0.17; 2.07: 0.17; 2.08: 0.2; 2.09: 0.13; 2.10: 0.14; 2.11: 0.2 |  |  | 54
| 1 | 6 | 2.01: 0.1257; 2.02: 0.18; 2.03: < 0.5; 2.04: 0.17659; 2.05: 0.2; 2.06: 0.17; 2.07: 0.17; 2.08: 0.2; 2.09: 0.13; 2.10: 0.14; 2.11: 0.22
| Mean |  | 2.01: 0.12123; 2.02: 0.18; 2.04: 0.17818; 2.05: 0.21667; 2.06: 0.16833; 2.07: 0.17; 2.08: 0.2; 2.09: 0.135; 2.10: 0.14333; 2.11: 0.20667
| Median |  | 2.01: 0.1202; 2.02: 0.18; 2.04: 0.17674; 2.05: 0.2; 2.06: 0.17; 2.07: 0.17; 2.08: 0.2; 2.09: 0.135; 2.10: 0.14; 2.11: 0.2
| Std Dev. |  | 2.01: 0.0066677; 2.02: 3.0405e-17; 2.04: 0.0031282; 2.05: 0.040825; 2.06: 0.0040825; 2.07: 0; 2.08: 3.0405e-17; 2.09: 0.0054772; 2.10: 0.005164; 2.11: 0.010328
| Rel.Std.Dev. |  | 2.01: 5.5%; 2.02: 1.689e-14%; 2.04: 1.756%; 2.05: 18.84%; 2.06: 2.425%; 2.07: 0%; 2.08: 1.52e-14%; 2.09: 4.057%; 2.10: 3.603%; 2.11: 4.997%
| PDM3 |  | 2.01: -25.55%; 2.02: 10.54%; 2.04: 9.417%; 2.05: 33.05%; 2.06: 3.371%; 2.07: 4.395%; 2.08: 22.82%; 2.09: -17.1%; 2.10: -11.98%; 2.11: 26.91%
| Z-Score (Absolute) |  | 2.01: 1.23; 2.02: 0.25; 2.03: 2.02; 2.04: 0.21; 2.05: NA; 2.06: 0.04; 2.07: 0; 2.08: NA; 2.09: 0.89; 2.10: 0.67; 2.11: 0.93
| Analytical results for Te in OREAS 163b (Certified Value < 0.05 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | < 0.1 | < 0.05 | < 0.5 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.1 | < 0.01 | < 0.05 | < 0.03 | < 0.5 | 5 | < 0.05 | < 0.05 | < 0.1 | [5 empty] | 1
| 1 | 2 | < 0.1 | < 0.05 | < 0.5 | < 0.2 | 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.1 | < 0.01 | 0.06 | < 0.03 | < 0.5 | 5 | < 0.05 | < 0.05 | < 0.1 | [5 empty] | 12
| 1 | 3 | < 0.1 | < 0.05 | < 0.5 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.1 | < 0.01 | 0.06 | < 0.03 | < 0.5 | 5 | < 0.05 | < 0.05 | 0.1 | [5 empty] | 16
| 1 | 4 | < 0.1 | < 0.05 | < 0.5 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.1 | 0.07 | 0.07 | < 0.03 | < 0.5 | 5 | < 0.05 | < 0.05 | < 0.1 | [5 empty] | < 0.05
| 1 | 5 | < 0.1 | < 0.05 | < 0.5 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.1 | < 0.01 | < 0.05 | < 0.03 | < 0.5 | 5 | < 0.05 | < 0.05 | < 0.1 | [5 empty] | 55
| 1 | 6 | < 0.1 | < 0.05 | < 0.5 | < 0.2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.1 | < 0.01 | < 0.05 | < 0.03 | < 0.5 | 5 | < 0.05 | < 0.05 | < 0.1
| Mean |  | 2.07: 0.05; 2.15: 0.07; 2.17: 0.063333; 2.20: 5; 2.25: 0.1
| Median |  | 2.07: 0.05; 2.15: 0.07; 2.17: 0.06; 2.20: 5; 2.25: 0.1
| Std Dev. |  | 2.17: 0.0057735; 2.20: 0
| Rel.Std.Dev. |  | 2.17: 9.116%; 2.20: 0%
PDM3
| Z-Score (Absolute) |  | 1.32 | 0.12 | 12.88 | 4.21 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 1.32 | 0.12 | 1.32 | 0.65 | 0.99 | 0.7 | 12.88 | 287.42 | 0.12 | 0.12 | 1.81
| Analytical results for Th in OREAS 163b (Certified Value 0.92 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 0.92 | 0.779 | 0.9 | < 0.5 | 0.88 | 0.91729 | 0.95 | 0.96 | 0.99 | 0.86 | 0.87 | 1.1 | 0.89 | 0.96886 | 2.2 | 0.96 | 1 | 0.9 | 0.99 | 1 | 0.72 | 0.94 | 0.9 | 1.1 |  | 1
| 1 | 2 | 0.94 | 0.7505 | 0.9 | < 0.5 | 0.92 | 0.91427 | 0.95 | 0.86 | 0.9 | 0.88 | 0.89 | 1.1 | 0.9 | 1.1473 | 1.7 | 0.91 | 0.9 | 0.9 | 1.05 | 0.9 | 0.51 | 0.94 | 0.8 | 1.1 |  | 13
| 1 | 3 | 0.96 | 0.7831 | 0.9 | < 0.5 | 0.9 | 0.91916 | 0.96 | 0.9 | 0.91 | 0.9 | 0.86 | 1.1 | 0.92 | 1.0499 | 1.5 | 0.95 | 0.9 | 0.9 | 0.96 | 1 | 0.74 | 0.86 | 0.9 | 1.2 |  | 16
| 1 | 4 | 0.92 | 0.7051 | 0.9 | < 0.5 | 0.87 | 0.92815 | 1 | 0.98 | 0.91 | 0.84 | 0.85 | 1 | 0.9 | 0.95552 | 1.8 | 0.97 | 0.9 | 0.9 | 1.02 | 0.9 | 0.55 | 0.9 | 0.9 | 1.1 |  | 0.9187
| 1 | 5 | 0.97 | 0.786 | 0.9 | < 0.5 | 0.87 | 0.94429 | 0.93 | 0.96 | 0.95 | 0.83 | 0.93 | 1.1 | 0.93 | 1.0591 | 1.8 | 0.92 | 0.9 | 0.9 | 0.98 | 1 | 0.62 | 0.91 | 0.9 | 1.1 |  | 56
| 1 | 6 | 0.95 | 0.7349 | 0.9 | < 0.5 | 0.91 | 0.90613 | 0.9 | 0.91 | 0.86 | 0.85 | 0.87 | 1.1 | 0.92 | 1.1716 | 1.4 | 0.96 | 0.9 | 0.9 | 0.98 | 1.1 | 0.7 | 0.9 | 1 | 1.1
| Mean |  | 0.94333 | 0.75643 | 0.9 |  | 0.89167 | 0.92155 | 0.94833 | 0.92833 | 0.92 | 0.86 | 0.87833 | 1.0833 | 0.91 | 1.0587 | 1.7333 | 0.945 | 0.91667 | 0.9 | 0.99667 | 0.98333 | 0.64 | 0.90833 | 0.9 | 1.1167
| Median |  | 0.945 | 0.76475 | 0.9 |  | 0.89 | 0.91822 | 0.95 | 0.935 | 0.91 | 0.855 | 0.87 | 1.1 | 0.91 | 1.0545 | 1.75 | 0.955 | 0.9 | 0.9 | 0.985 | 1 | 0.66 | 0.905 | 0.9 | 1.1
| Std Dev. |  | 0.020656 | 0.032337 | 0 |  | 0.02137 | 0.01323 | 0.033116 | 0.045789 | 0.044721 | 0.026077 | 0.028577 | 0.040825 | 0.015492 | 0.088766 | 0.28048 | 0.02429 | 0.040825 | 0 | 0.03266 | 0.075277 | 0.095289 | 0.029944 | 0.063246 | 0.040825
| Rel.Std.Dev. |  | 2.19% | 4.275% | 0% |  | 2.397% | 1.436% | 3.492% | 4.932% | 4.861% | 3.032% | 3.254% | 3.768% | 1.702% | 8.384% | 16.18% | 2.57% | 4.454% | 0% | 3.277% | 7.655% | 14.89% | 3.297% | 7.027% | 3.656%
| PDM3 |  | 2.681% | -17.66% | -2.035% |  | -2.942% | 0.3102% | 3.226% | 1.049% | 0.1416% | -6.389% | -4.394% | 17.92% | -0.9469% | 15.24% | 88.67% | 2.863% | -0.2212% | -2.035% | 8.487% | 7.035% | -30.34% | -1.128% | -2.035% | 21.55%
| Z-Score (Absolute) |  | 0.59 | 3.85 | 0.44 | 15.91 | 0.63 | 0.08 | 0.71 | 0.24 | 0.04 | 1.39 | 0.95 | 3.93 | 0.2 | 3.34 | 19.4 | 0.63 | 0.04 | 0.44 | 1.86 | 1.55 | 6.62 | 0.24 | 0.44 | 4.72
| Analytical results for Ti in OREAS 163b (Certified Value 0.021 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.0197 | 0.019925 | 0.02 | 0.03 | 0.0189 | 0.021681 | 0.021 | 0.021 | 0.021 | 0.022 | 0.02 | 0.0299 | 0.017 | 0.020229 | 0.02 | 0.0236 | 0.025 | 0.017 | 0.02 | 0.02 | 0.02 | 0.0201 | 0.02 | 0.021 |  | 1
| 1 | 2 | 0.0201 | 0.020065 | 0.02 | 0.03 | 0.0179 | 0.021845 | 0.022 | 0.022 | 0.024 | 0.023 | 0.021 | 0.03 | 0.017 | 0.021274 | 0.02 | 0.0231 | 0.024 | 0.017 | 0.02 | 0.02 | 0.02 | 0.0194 | 0.02 | 0.021 |  | 32
| 1 | 3 | 0.0199 | 0.020238 | 0.0202 | 0.03 | 0.0207 | 0.021374 | 0.022 | 0.023 | 0.021 | 0.023 | 0.021 | 0.0298 | 0.018 | 0.020681 | 0.02 | 0.0227 | 0.021 | 0.016 | 0.02 | 0.02 | 0.02 | 0.0198 | 0.02 | 0.021 |  | 16
| 1 | 4 | 0.0196 | 0.019321 | 0.02 | 0.03 | 0.0173 | 0.021439 | 0.021 | 0.023 | 0.023 | 0.022 | 0.021 | 0.0298 | 0.017 | 0.02141 | 0.02 | 0.0236 | 0.025 | 0.017 | 0.02 | 0.02 | 0.02 | 0.0181 | 0.02 | 0.021 |  | 0.020513
| 1 | 5 | 0.0198 | 0.019472 | 0.0202 | 0.02 | 0.0187 | 0.021848 | 0.022 | 0.023 | 0.021 | 0.022 | 0.021 | 0.0306 | 0.018 | 0.020065 | 0.02 | 0.0231 | 0.024 | 0.017 | 0.02 | 0.02 | 0.02 | 0.0197 | 0.03 | 0.022 |  | 57
| 1 | 6 | 0.0198 | 0.019622 | 0.0202 | 0.03 | 0.0192 | 0.0212 | 0.022 | 0.022 | 0.023 | 0.022 | 0.022 | 0.0293 | 0.017 | 0.021088 | 0.02 | 0.0228 | 0.02 | 0.017 | 0.02 | 0.02 | 0.02 | 0.0177 | 0.03 | 0.022
| Mean |  | 0.019817 | 0.019774 | 0.0201 | 0.028333 | 0.018783 | 0.021565 | 0.021667 | 0.022333 | 0.022167 | 0.022333 | 0.021 | 0.0299 | 0.017333 | 0.020791 | 0.02 | 0.02315 | 0.023167 | 0.016833 | 0.02 | 0.02 | 0.02 | 0.019133 | 0.023333 | 0.021333
| Median |  | 0.0198 | 0.019773 | 0.0201 | 0.03 | 0.0188 | 0.02156 | 0.022 | 0.0225 | 0.022 | 0.022 | 0.021 | 0.02985 | 0.017 | 0.020884 | 0.02 | 0.0231 | 0.024 | 0.017 | 0.02 | 0.02 | 0.02 | 0.01955 | 0.02 | 0.021
| Std Dev. |  | 0.00017224 | 0.00035867 | 0.00010954 | 0.0040825 | 0.0011703 | 0.00026732 | 0.0005164 | 0.0008165 | 0.0013292 | 0.0005164 | 0.00063246 | 0.00041952 | 0.0005164 | 0.00055857 | 0 | 0.00038341 | 0.002137 | 0.00040825 | 0 | 0 | 0 | 0.00098928 | 0.005164 | 0.0005164
| Rel.Std.Dev. |  | 0.8692% | 1.814% | 0.545% | 14.41% | 6.231% | 1.24% | 2.383% | 3.656% | 5.996% | 2.312% | 3.012% | 1.403% | 2.979% | 2.687% | 0% | 1.656% | 9.224% | 2.425% | 0% | 0% | 0% | 5.17% | 22.13% | 2.421%
| PDM3 |  | -3.393% | -3.602% | -2.012% | 38.13% | -8.431% | 5.128% | 5.626% | 8.876% | 8.063% | 8.876% | 2.376% | 45.76% | -15.5% | 1.357% | -2.499% | 12.86% | 12.94% | -17.94% | -2.499% | -2.499% | -2.499% | -6.724% | 13.75% | 4.001%
| Z-Score (Absolute) |  | 0.61 | 0.63 | 0.45 | 4.19 | 1.19 | 0.38 | 0.43 | 0.81 | 0.72 | 0.81 | 0.06 | 5.08 | 2.01 | 0.06 | 0.5 | 1.27 | 1.28 | 2.29 | 0.5 | 0.5 | 0.5 | 0.99 | 1.37 | 0.25
| Analytical results for Tl in OREAS 163b (Certified Value 1.59 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.57 | 1.4964 | 1.5 | < 5 | 1.65 | 1.5641 | 1.55 | 1.66 | 1.54 | 1.54 | 1.41 | 1.67 | 1.56 | 1.6916 | 1.53 | 1.72 | 1.73 | 1.479 | 1.6 | 1.66 | 1.66 | 1.63 | 1.61 | 1.3 |  | 1
| 1 | 2 | 1.57 | 1.5648 | 1.5 | < 5 | 1.66 | 1.7025 | 1.59 | 1.62 | 1.57 | 1.59 | 1.46 | 1.69 | 1.63 | 1.7678 | 1.35 | 1.72 | 1.63 | 1.461 | 1.6 | 1.67 | 1.65 | 1.57 | 1.61 | 1.4 |  | 14
| 1 | 3 | 1.57 | 1.5109 | 1.5 | < 5 | 1.69 | 1.6941 | 1.63 | 1.65 | 1.6 | 1.53 | 1.41 | 1.68 | 1.64 | 1.7625 | 1.37 | 1.73 | 1.58 | 1.445 | 1.6 | 1.73 | 1.75 | 1.56 | 1.6 | 1.3 |  | 16
| 1 | 4 | 1.55 | 1.5364 | 1.5 | < 5 | 1.61 | 1.6161 | 1.59 | 1.74 | 1.54 | 1.49 | 1.45 | 1.56 | 1.6 | 1.7085 | 1.61 | 1.74 | 1.78 | 1.485 | 1.6 | 1.67 | 1.61 | 1.56 | 1.61 | 1.2 |  | 1.5948
| 1 | 5 | 1.59 | 1.5431 | 1.5 | < 5 | 1.63 | 1.5623 | 1.52 | 1.68 | 1.56 | 1.5 | 1.55 | 1.63 | 1.61 | 1.7854 | 1.5 | 1.75 | 1.75 | 1.475 | 1.6 | 1.72 | 1.62 | 1.61 | 1.58 | 1.2 |  | 58
| 1 | 6 | 1.57 | 1.5325 | 1.5 | < 5 | 1.65 | 1.5596 | 1.51 | 1.6 | 1.47 | 1.52 | 1.43 | 1.7 | 1.62 | 1.7892 | 1.31 | 1.76 | 1.53 | 1.469 | 1.6 | 1.65 | 1.71 | 1.6 | 1.59 | 1.5
| Mean |  | 1.57 | 1.5307 | 1.5 |  | 1.6483 | 1.6164 | 1.565 | 1.6583 | 1.5467 | 1.5283 | 1.4517 | 1.655 | 1.61 | 1.7508 | 1.445 | 1.7367 | 1.6667 | 1.469 | 1.6 | 1.6833 | 1.6667 | 1.5883 | 1.6 | 1.3167
| Median |  | 1.57 | 1.5345 | 1.5 |  | 1.65 | 1.5901 | 1.57 | 1.655 | 1.55 | 1.525 | 1.44 | 1.675 | 1.615 | 1.7651 | 1.435 | 1.735 | 1.68 | 1.472 | 1.6 | 1.67 | 1.655 | 1.585 | 1.605 | 1.3
| Std Dev. |  | 0.012649 | 0.02417 | 0 |  | 0.027142 | 0.066824 | 0.046368 | 0.04916 | 0.043665 | 0.035449 | 0.052313 | 0.05244 | 0.028284 | 0.040983 | 0.11862 | 0.01633 | 0.10132 | 0.014367 | 2.4324e-16 | 0.033267 | 0.053914 | 0.029269 | 0.012649 | 0.1169
| Rel.Std.Dev. |  | 0.8057% | 1.579% | 0% |  | 1.647% | 4.134% | 2.963% | 2.964% | 2.823% | 2.319% | 3.604% | 3.169% | 1.757% | 2.341% | 8.209% | 0.9403% | 6.079% | 0.978% | 1.52e-14% | 1.976% | 3.235% | 1.843% | 0.7906% | 8.879%
| PDM3 |  | -1.557% | -4.022% | -5.946% |  | 3.355% | 1.355% | -1.871% | 3.982% | -3.02% | -4.17% | -8.977% | 3.773% | 0.9511% | 9.782% | -9.395% | 8.893% | 4.504% | -7.89% | 0.3241% | 5.549% | 4.504% | -0.4075% | 0.3241% | -17.44%
| Z-Score (Absolute) |  | 0.3 | 0.7 | 1.01 | 9.1 | 0.49 | 0.17 | 0.35 | 0.59 | 0.54 | 0.72 | 1.5 | 0.56 | 0.1 | 1.53 | 1.57 | 1.38 | 0.67 | 1.33 | 0 | 0.84 | 0.67 | 0.12 | 0 | 2.87
| Analytical results for Tm in OREAS 163b (Certified Value < 0.1 ppm) | [26 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.06; 2.04: 2.12; 2.05: 2.16; 2.06: 2.17; 2.07: 2.20 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*MS |  | 2
2
| 1 | 1 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 0.05; 2.04: < 0.1; 2.05: 0.04; 2.06: < 0.1; 2.07: 0.02 |  | 1
| 1 | 2 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 0.05; 2.04: < 0.1; 2.05: 0.05; 2.06: < 0.1; 2.07: 0.02 |  | 6
| 1 | 3 | 2.01: < 0.1; 2.02: < 0.1; 2.03: 0.050092; 2.04: < 0.1; 2.05: 0.05; 2.06: < 0.1; 2.07: 0.02 |  | 16
| 1 | 4 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 0.05; 2.04: < 0.1; 2.05: 0.04; 2.06: < 0.1; 2.07: 0.02 |  | < 0.1
| 1 | 5 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 0.05; 2.04: < 0.1; 2.05: 0.05; 2.06: < 0.1; 2.07: 0.02 |  | 59
| 1 | 6 | 2.01: < 0.1; 2.02: < 0.1; 2.03: < 0.05; 2.04: < 0.1; 2.05: 0.05; 2.06: < 0.1; 2.07: 0.02
| Mean |  | 2.03: 0.050092; 2.05: 0.046667; 2.07: 0.02
| Median |  | 2.03: 0.050092; 2.05: 0.05; 2.07: 0.02
| Std Dev. |  | 2.05: 0.005164; 2.07: 0
| Rel.Std.Dev. |  | 2.05: 11.07%; 2.07: 0%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.07
| Analytical results for U in OREAS 163b (Certified Value 0.7 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 2
3
| 1 | 1 | 0.7 | 0.5894 | 0.7 | < 0.5 | 0.73 | 0.74405 | 0.7 | 0.6 | 0.7 | 0.7 | 0.6 | 0.8 | 0.65 | 0.75075 | 0.7 | 0.71 | 0.7 | 0.65 | 0.7 | 0.7 | 0.67 | 0.71 | 0.73 | 0.6 | 0.8 | 1
| 1 | 2 | 0.71 | 0.5723 | 0.7 | < 0.5 | 0.72 | 0.7618 | 0.7 | 0.6 | 0.7 | 0.8 | 0.7 | 0.7 | 0.66 | 0.76968 | 0.61 | 0.7 | 0.7 | 0.64 | 0.7 | 0.7 | 0.67 | 0.68 | 0.75 | 0.6 | 0.8 | 16
| 1 | 3 | 0.71 | 0.5602 | 0.8 | < 0.5 | 0.71 | 0.75919 | 0.7 | 0.6 | 0.7 | 0.7 | 0.7 | 0.8 | 0.7 | 0.74903 | 0.62 | 0.7 | 0.7 | 0.64 | 0.7 | 0.8 | 0.67 | 0.71 | 0.72 | 0.6 | 0.8 | 16
| 1 | 4 | 0.7 | 0.6029 | 0.7 | < 0.5 | 0.71 | 0.74718 | 0.7 | 0.7 | 0.7 | 0.7 | 0.6 | 0.7 | 0.7 | 0.74081 | 0.73 | 0.72 | 0.7 | 0.64 | 0.7 | 0.7 | 0.66 | 0.69 | 0.73 | 0.6 | 0.8 | 0.69952
| 1 | 5 | 0.71 | 0.6307 | 0.7 | < 0.5 | 0.67 | 0.71736 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.7 | 0.74 | 0.75809 | 0.64 | 0.71 | 0.7 | 0.66 | 0.7 | 0.7 | 0.67 | 0.69 | 0.73 | 0.6 | 0.8 | 60
| 1 | 6 | 0.7 | 0.5975 | 0.7 | < 0.5 | 0.7 | 0.74656 | 0.6 | 0.6 | 0.7 | 0.7 | 0.7 | 0.8 | 0.72 | 0.77369 | 0.56 | 0.74 | 0.7 | 0.69 | 0.7 | 0.7 | 0.68 | 0.69 | 0.71 | 0.7 | 0.7
| Mean |  | 0.705 | 0.59217 | 0.71667 |  | 0.70667 | 0.74602 | 0.68333 | 0.63333 | 0.7 | 0.71667 | 0.66667 | 0.75 | 0.695 | 0.75701 | 0.64333 | 0.71333 | 0.7 | 0.65333 | 0.7 | 0.71667 | 0.67 | 0.695 | 0.72833 | 0.61667 | 0.78333
| Median |  | 0.705 | 0.59345 | 0.7 |  | 0.71 | 0.74687 | 0.7 | 0.6 | 0.7 | 0.7 | 0.7 | 0.75 | 0.7 | 0.75442 | 0.63 | 0.71 | 0.7 | 0.645 | 0.7 | 0.7 | 0.67 | 0.69 | 0.73 | 0.6 | 0.8
| Std Dev. |  | 0.0054772 | 0.024709 | 0.040825 |  | 0.020656 | 0.015805 | 0.040825 | 0.05164 | 1.2162e-16 | 0.040825 | 0.05164 | 0.054772 | 0.034496 | 0.012689 | 0.062183 | 0.015055 | 1.2162e-16 | 0.019664 | 1.2162e-16 | 0.040825 | 0.0063246 | 0.012247 | 0.013292 | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 0.7769% | 4.173% | 5.696% |  | 2.923% | 2.119% | 5.974% | 8.154% | 1.737e-14% | 5.696% | 7.746% | 7.303% | 4.964% | 1.676% | 9.666% | 2.111% | 1.737e-14% | 3.01% | 1.737e-14% | 5.696% | 0.944% | 1.762% | 1.825% | 6.62% | 5.212%
| PDM3 |  | 0.7829% | -15.35% | 2.451% |  | 1.021% | 6.647% | -2.314% | -9.462% | 0.06817% | 2.451% | -4.697% | 7.216% | -0.6466% | 8.218% | -8.033% | 1.974% | 0.06817% | -6.603% | 0.06817% | 2.451% | -4.22% | -0.6466% | 4.119% | -11.84% | 11.98%
| Z-Score (Absolute) |  | 0.12 | 2.57 | 0.4 | 10.71 | 0.16 | 1.1 | 0.4 | 1.59 | 0 | 0.4 | 0.79 | 1.19 | 0.12 | 1.36 | 1.35 | 0.32 | 0 | 1.11 | 0 | 0.4 | 0.71 | 0.12 | 0.67 | 1.98 | 1.98
| Analytical results for V in OREAS 163b (Certified Value 10.7 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 12 | 10.7 | 11 | 8.46 | 10 | 10.4398 | 11 | 10 | 10 | 10 | 10 | 15 | 10 | 12.1636 | 12 | 11 | 11 | 11 | 10 | 11 | 11 | 10 | 10 | 11 | 11 | 1
| 1 | 2 | 12 | 10.41 | 11 | 9.84 | 12 | 10.3761 | 11 | 10 | 11 | 11 | 10 | 15 | 10 | 12.9177 | 10 | 11 | 11 | 10 | 9 | 11 | 11 | 10.8 | 11 | 11 | 11 | 36
| 1 | 3 | 12 | 11.61 | 11 | 9.55 | 11 | 10.7172 | 11 | 11 | 10 | 10 | 10 | 14 | 11 | 11.7234 | 10 | 11 | 11 | 10 | 10 | 11 | 11 | 10 | 10 | 11 | 11 | 16
| 1 | 4 | 12 | 11.24 | 11 | 9.35 | 11 | 10.3651 | 10 | 10 | 11 | 10 | 10 | 14 | 11 | 12.9557 | 11 | 11 | 11 | 11 | 10 | 11 | 11 | 10.1 | 10 | 11 | 11 | 10.6751
| 1 | 5 | 12 | 10.18 | 11 | 8.86 | 11 | 10.392 | 11 | 12 | 10 | 10 | 10 | 14 | 11 | 11.88 | 10 | 11 | 11 | 11 | 10 | 11 | 11 | 9.9 | 10 | 11 | 11 | 61
| 1 | 6 | 12 | 11.15 | 11 | 9.46 | 11 | 10.1302 | 11 | 10 | 11 | 10 | 10 | 14 | 11 | 12.91 | 9 | 10 | 10 | 11 | 10 | 11 | 11 | 10.6 | 10 | 11 | 11
| Mean |  | 12 | 10.8817 | 11 | 9.2533 | 11 | 10.4034 | 10.8333 | 10.5 | 10.5 | 10.1667 | 10 | 14.3333 | 10.6667 | 12.4251 | 10.3333 | 10.8333 | 10.8333 | 10.6667 | 9.8333 | 11 | 11 | 10.2333 | 10.1667 | 11 | 11
| Median |  | 12 | 10.925 | 11 | 9.405 | 11 | 10.3841 | 11 | 10 | 10.5 | 10 | 10 | 14 | 11 | 12.5368 | 10 | 11 | 11 | 11 | 10 | 11 | 11 | 10.05 | 10 | 11 | 11
| Std Dev. |  | 0 | 0.54382 | 0 | 0.50365 | 0.63246 | 0.18803 | 0.40825 | 0.83666 | 0.54772 | 0.40825 | 0 | 0.5164 | 0.5164 | 0.5687 | 1.0328 | 0.40825 | 0.40825 | 0.5164 | 0.40825 | 0 | 0 | 0.37238 | 0.40825 | 0 | 0
| Rel.Std.Dev. |  | 0% | 4.998% | 0% | 5.443% | 5.75% | 1.807% | 3.768% | 7.968% | 5.216% | 4.016% | 0% | 3.603% | 4.841% | 4.577% | 9.995% | 3.768% | 3.768% | 4.841% | 4.152% | 0% | 0% | 3.639% | 4.016% | 0% | 0%
| PDM3 |  | 12.41% | 1.935% | 3.044% | -13.32% | 3.044% | -2.545% | 1.482% | -1.64% | -1.64% | -4.763% | -6.324% | 34.27% | -0.07881% | 16.39% | -3.201% | 1.482% | 1.482% | -0.07881% | -7.885% | 3.044% | 3.044% | -4.138% | -4.763% | 3.044% | 3.044%
| Z-Score (Absolute) |  | 2.36 | 0.1 | 0.34 | 3.2 | 0.34 | 0.87 | 0 | 0.67 | 0.67 | 1.35 | 1.69 | 7.08 | 0.34 | 3.22 | 1.01 | 0 | 0 | 0.34 | 2.02 | 0.34 | 0.34 | 1.21 | 1.35 | 0.34 | 0.34
| Analytical results for W in OREAS 163b (Certified Value 2.05 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2 | 1.5595 | 2.3 | < 5 | 2 | 1.9629 | 2.2 | 2 | 2 | 1.8 | 1.8 | 2.2 | 2 | 2.4564 | 2.24 | 2.1 | 2.3 | 2 | 2.8 | 3 | 1.75 | 2.2 | 2.2 | 2.2 |  | 1
| 1 | 2 | 2 | 1.5444 | 2.4 | < 5 | 2.1 | 2.0571 | 2.2 | 2 | 1.9 | 1.8 | 1.8 | 2.1 | 2 | 2.053 | 1.59 | 2 | 2 | 2 | 2.9 | 3 | 1.88 | 2.2 | 2.2 | 2 |  | 17
| 1 | 3 | 2 | 1.4568 | 2.3 | < 5 | 2.1 | 2.1729 | 2.3 | 2.3 | 2.1 | 1.8 | 1.8 | 2.3 | 2.1 | 1.9411 | 1.42 | 2.2 | 1.9 | 2 | 2.7 | 2 | 1.78 | 2.1 | 2.2 | 2 |  | 16
| 1 | 4 | 2 | 1.6139 | 2.2 | < 5 | 2.1 | 2.1586 | 2.3 | 2.1 | 1.8 | 1.7 | 1.8 | 1.9 | 2.1 | 1.94 | 1.77 | 2.1 | 2.2 | 2 | 2.8 | 2 | 1.61 | 2.1 | 2.2 | 2.1 |  | 2.0495
| 1 | 5 | 2.1 | 1.5826 | 2.3 | < 5 | 2.1 | 1.9191 | 2.2 | 2.1 | 1.9 | 1.8 | 1.9 | 2.1 | 2.1 | 2.0926 | 1.45 | 2.1 | 2.1 | 2 | 2.8 | 2 | 1.68 | 2.2 | 2.1 | 2 |  | 62
| 1 | 6 | 2 | 1.4921 | 2.4 | < 5 | 2.1 | 1.9501 | 2.2 | 1.9 | 1.9 | 1.8 | 1.8 | 2 | 2 | 2.4903 | 1.29 | 2.2 | 2 | 2 | 2.8 | 2 | 1.75 | 2.2 | 2.1 | 1.8
| Mean |  | 2.0167 | 1.5415 | 2.3167 |  | 2.0833 | 2.0368 | 2.2333 | 2.0667 | 1.9333 | 1.7833 | 1.8167 | 2.1 | 2.05 | 2.1622 | 1.6267 | 2.1167 | 2.0833 | 2 | 2.8 | 2.3333 | 1.7417 | 2.1667 | 2.1667 | 2.0167
| Median |  | 2 | 1.5519 | 2.3 |  | 2.1 | 2.01 | 2.2 | 2.05 | 1.9 | 1.8 | 1.8 | 2.1 | 2.05 | 2.0728 | 1.52 | 2.1 | 2.05 | 2 | 2.8 | 2 | 1.75 | 2.2 | 2.2 | 2
| Std Dev. |  | 0.040825 | 0.058094 | 0.075277 |  | 0.040825 | 0.11008 | 0.05164 | 0.13663 | 0.10328 | 0.040825 | 0.040825 | 0.14142 | 0.054772 | 0.24867 | 0.34215 | 0.075277 | 0.1472 | 0 | 0.063246 | 0.5164 | 0.091524 | 0.05164 | 0.05164 | 0.13292
| Rel.Std.Dev. |  | 2.024% | 3.769% | 3.249% |  | 1.96% | 5.405% | 2.312% | 6.611% | 5.342% | 2.289% | 2.247% | 6.734% | 2.672% | 11.5% | 21.03% | 3.556% | 7.065% | 0% | 2.259% | 22.13% | 5.255% | 2.383% | 2.383% | 6.591%
| PDM3 |  | -1.601% | -24.78% | 13.04% |  | 1.652% | -0.6201% | 8.971% | 0.8385% | -5.667% | -12.99% | -11.36% | 2.465% | 0.0253% | 5.501% | -20.63% | 3.278% | 1.652% | -2.414% | 36.62% | 13.85% | -15.02% | 5.718% | 5.718% | -1.601%
| Z-Score (Absolute) |  | 0.43 | 3.92 | 1.78 | 3.13 | 0.06 | 0.28 | 1.16 | 0.06 | 1.04 | 2.15 | 1.9 | 0.18 | 0.18 | 0.64 | 3.3 | 0.31 | 0.06 | NA | 5.33 | NA | 2.45 | 0.67 | 0.67 | 0.43
Results from laboratory 2.19 and 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 163b (Certified Value 4.67 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 5 | 4.2157 | 4.5 | < 5 | 3.87 | 4.5993 | 4.7 | 4.9 | 4.9 | 5.1 | 5.1 | 7 | 4 | 5.8063 | 4.67 | 4.03 | 5 | 5 | 4.4 | 4.3 | 5.5 | 4.73 | 4.4 | 5.3 |  | 1
| 1 | 2 | 4.9 | 4.2736 | 4.5 | < 5 | 4.27 | 4.6139 | 4.7 | 5.2 | 5.7 | 5.3 | 5.2 | 6.7 | 4 | 7.0122 | 4.29 | 4.13 | 4.4 | 5.1 | 4.3 | 4.4 | 5.7 | 5.17 | 4.2 | 5.5 |  | 18
| 1 | 3 | 5.1 | 4.237 | 4.5 | < 5 | 4.01 | 4.7072 | 4.6 | 5.3 | 5.9 | 4.9 | 5 | 7.2 | 3.9 | 6.6765 | 3.61 | 3.98 | 4.5 | 4.8 | 4.2 | 4.5 | 5.6 | 4.63 | 4.4 | 5.4 |  | 16
| 1 | 4 | 4.9 | 4.1605 | 4.5 | < 5 | 3.88 | 4.615 | 4.5 | 5.1 | 5.8 | 4.8 | 5 | 6.2 | 4.1 | 5.7418 | 4.71 | 4 | 4.8 | 5.3 | 4.3 | 4.2 | 5.3 | 5.25 | 4.6 | 5.2 |  | 4.671
| 1 | 5 | 5 | 4.2205 | 4.4 | < 5 | 3.9 | 4.625 | 4.4 | 5.2 | 5 | 5.2 | 5.2 | 6.8 | 4.1 | 6.9878 | 3.53 | 4.02 | 4.8 | 4.7 | 4.1 | 4.3 | 5.3 | 5.16 | 4.4 | 5.3 |  | 63
| 1 | 6 | 5.1 | 4.251 | 4.4 | < 5 | 4.1 | 4.5801 | 4.4 | 4.8 | 5.8 | 5.2 | 4.9 | 6.7 | 4.1 | 7.2792 | 3.4 | 4.12 | 4.3 | 5.2 | 4 | 4.2 | 5.5 | 4.99 | 4.4 | 5.1
| Mean |  | 5 | 4.2264 | 4.4667 |  | 4.005 | 4.6234 | 4.55 | 5.0833 | 5.5167 | 5.0833 | 5.0667 | 6.7667 | 4.0333 | 6.584 | 4.035 | 4.0467 | 4.6333 | 5.0167 | 4.2167 | 4.3167 | 5.4833 | 4.9883 | 4.4 | 5.3
| Median |  | 5 | 4.2287 | 4.5 |  | 3.955 | 4.6145 | 4.55 | 5.15 | 5.75 | 5.15 | 5.05 | 6.75 | 4.05 | 6.8322 | 3.95 | 4.025 | 4.65 | 5.05 | 4.25 | 4.3 | 5.5 | 5.075 | 4.4 | 5.3
| Std Dev. |  | 0.089443 | 0.03857 | 0.05164 |  | 0.15758 | 0.043911 | 0.13784 | 0.19408 | 0.4446 | 0.19408 | 0.12111 | 0.33862 | 0.08165 | 0.6561 | 0.59376 | 0.06314 | 0.27325 | 0.23166 | 0.1472 | 0.1169 | 0.16021 | 0.25538 | 0.12649 | 0.14142
| Rel.Std.Dev. |  | 1.789% | 0.9126% | 1.156% |  | 3.934% | 0.9497% | 3.029% | 3.818% | 8.059% | 3.818% | 2.39% | 5.004% | 2.024% | 9.965% | 14.72% | 1.56% | 5.898% | 4.618% | 3.491% | 2.708% | 2.922% | 5.119% | 2.875% | 2.668%
| PDM3 |  | 7.043% | -9.519% | -4.375% |  | -14.26% | -1.019% | -2.591% | 8.827% | 18.1% | 8.827% | 8.47% | 44.86% | -13.65% | 40.95% | -13.62% | -13.37% | -0.8069% | 7.4% | -9.727% | -7.586% | 17.39% | 6.793% | -5.802% | 13.47%
| Z-Score (Absolute) |  | 0.56 | 0.61 | 0.24 | 3.21 | 0.94 | 0.01 | 0.12 | 0.69 | 1.34 | 0.69 | 0.66 | 3.23 | 0.9 | 2.95 | 0.9 | 0.88 | 0.01 | 0.59 | 0.62 | 0.47 | 1.29 | 0.54 | 0.34 | 1.01
| Analytical results for Yb in OREAS 163b (Certified Value 0.29 ppm) | [26 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.03; 2.03: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.16; 2.08: 2.17; 2.09: 2.20; 2.10: 2.21; 2.11: 2.22 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*OES/MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*MS; 2.07: 4A*MS; 2.08: 4A*MS; 2.09: 4A*MS; 2.10: 4A*OES/MS; 2.11: 4A*OES/MS |  | 2
3
| 1 | 1 | 2.01: 0.2375; 2.02: 0.29; 2.03: < 0.5; 2.04: 0.29702; 2.05: 0.4; 2.06: 0.3; 2.07: 0.33; 2.08: 0.3; 2.09: 0.24; 2.10: 0.3; 2.11: 0.33 |  | 1
| 1 | 2 | 2.01: 0.2608; 2.02: 0.29; 2.03: < 0.5; 2.04: 0.33867; 2.05: 0.4; 2.06: 0.3; 2.07: 0.29; 2.08: 0.3; 2.09: 0.28; 2.10: 0.3; 2.11: 0.34 |  | 19
| 1 | 3 | 2.01: 0.2453; 2.02: 0.29; 2.03: < 0.5; 2.04: 0.33856; 2.05: 0.5; 2.06: 0.3; 2.07: 0.26; 2.08: 0.3; 2.09: 0.28; 2.10: 0.3; 2.11: 0.32 |  | 16
| 1 | 4 | 2.01: 0.2306; 2.02: 0.29; 2.03: < 0.5; 2.04: 0.30362; 2.05: 0.4; 2.06: 0.3; 2.07: 0.27; 2.08: 0.3; 2.09: 0.26; 2.10: 0.3; 2.11: 0.33 |  | 0.29084
| 1 | 5 | 2.01: 0.2641; 2.02: 0.29; 2.03: < 0.5; 2.04: 0.31154; 2.05: 0.4; 2.06: 0.3; 2.07: 0.25; 2.08: 0.3; 2.09: 0.26; 2.10: 0.3; 2.11: 0.32 |  | 64
| 1 | 6 | 2.01: 0.243; 2.02: 0.29; 2.03: < 0.5; 2.04: 0.30688; 2.05: 0.4; 2.06: 0.3; 2.07: 0.27; 2.08: 0.3; 2.09: 0.28; 2.10: 0.3; 2.11: 0.34
| Mean |  | 2.01: 0.24688; 2.02: 0.29; 2.04: 0.31605; 2.05: 0.41667; 2.06: 0.3; 2.07: 0.27833; 2.08: 0.3; 2.09: 0.26667; 2.10: 0.3; 2.11: 0.33
| Median |  | 2.01: 0.24415; 2.02: 0.29; 2.04: 0.30921; 2.05: 0.4; 2.06: 0.3; 2.07: 0.27; 2.08: 0.3; 2.09: 0.27; 2.10: 0.3; 2.11: 0.33
| Std Dev. |  | 2.01: 0.013121; 2.02: 0; 2.04: 0.018109; 2.05: 0.040825; 2.06: 0; 2.07: 0.028577; 2.08: 0; 2.09: 0.01633; 2.10: 0; 2.11: 0.0089443
| Rel.Std.Dev. |  | 2.01: 5.315%; 2.02: 0%; 2.04: 5.73%; 2.05: 9.798%; 2.06: 0%; 2.07: 10.27%; 2.08: 0%; 2.09: 6.124%; 2.10: 0%; 2.11: 2.71%
| PDM3 |  | 2.01: -15.12%; 2.02: -0.2904%; 2.04: 8.667%; 2.05: 43.26%; 2.06: 3.148%; 2.07: -4.302%; 2.08: 3.148%; 2.09: -8.313%; 2.10: 3.148%; 2.11: 13.46%
| Z-Score (Absolute) |  | 2.01: 1.65; 2.02: 0.31; 2.03: 1.56; 2.04: 0.5; 2.05: 3.63; 2.06: 0; 2.07: 0.67; 2.08: 0; 2.09: 1.04; 2.10: 0; 2.11: 0.93
| Analytical results for Zn in OREAS 163b (Certified Value 26.6 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 27 | 26.89 | 31 | < 0.5 | 25 | 28.0921 | 26 | 27 | 26 | 29 | 25 | 27 | 25 | 26.1261 | 26 | 27 | 26.8 | 52 | 29 | 27 | 27 | 22.3 | 25 | 25 | 23 | 1
| 1 | 2 | 28 | 20.83 | 30 | < 0.5 | 25 | 27.2471 | 26 | 27 | 26 | 30 | 27 | 26 | 25 | 25.5532 | 34 | 27 | 26.9 | 31 | 30 | 27 | 28 | 23 | 24 | 25 | 23 | 40
| 1 | 3 | 27 | 22.04 | 31 | < 0.5 | 25 | 27.5731 | 26 | 28 | 26 | 28 | 25 | 28 | 26 | 25.3195 | 25 | 27 | 24.7 | 38 | 30 | 27 | 29 | 22.6 | 25 | 26 | 24 | 16
| 1 | 4 | 27 | 21.69 | 32 | < 0.5 | 25 | 27.1756 | 25 | 32 | 26 | 34 | 26 | 27 | 26 | 25.9875 | 33 | 25 | 27.2 | 35 | 27 | 27 | 27 | 21.8 | 24 | 25 | 23 | 26.5764
| 1 | 5 | 28 | 23.22 | 29 | < 0.5 | 25 | 28.6932 | 27 | 28 | 26 | 29 | 26 | 28 | 26 | 26.5017 | 33 | 27 | 25.8 | 45 | 27 | 27 | 29 | 20.5 | 25 | 25 | 22 | 65
| 1 | 6 | 27 | 25.05 | 31 | < 0.5 | 26 | 26.536 | 27 | 30 | 25 | 32 | 26 | 27 | 25 | 26.9136 | 32 | 26 | 24.5 | 31 | 28 | 26 | 28 | 23 | 24 | 26 | 25
| Mean |  | 27.3333 | 23.2867 | 30.6667 |  | 25.1667 | 27.5529 | 26.1667 | 28.6667 | 25.8333 | 30.3333 | 25.8333 | 27.1667 | 25.5 | 26.0669 | 30.5 | 26.5 | 25.9833 | 38.6667 | 28.5 | 26.8333 | 28 | 22.2 | 24.5 | 25.3333 | 23.3333
| Median |  | 27 | 22.63 | 31 |  | 25 | 27.4101 | 26 | 28 | 26 | 29.5 | 26 | 27 | 25.5 | 26.0568 | 32.5 | 27 | 26.3 | 36.5 | 28.5 | 27 | 28 | 22.45 | 24.5 | 25 | 23
| Std Dev. |  | 0.5164 | 2.2907 | 1.0328 |  | 0.40825 | 0.75649 | 0.75277 | 1.9664 | 0.40825 | 2.2509 | 0.75277 | 0.75277 | 0.54772 | 0.58964 | 3.937 | 0.83666 | 1.172 | 8.3586 | 1.3784 | 0.40825 | 0.89443 | 0.94868 | 0.54772 | 0.5164 | 1.0328
| Rel.Std.Dev. |  | 1.889% | 9.837% | 3.368% |  | 1.622% | 2.746% | 2.877% | 6.859% | 1.58% | 7.421% | 2.914% | 2.771% | 2.148% | 2.262% | 12.91% | 3.157% | 4.511% | 21.62% | 4.837% | 1.521% | 3.194% | 4.273% | 2.236% | 2.038% | 4.426%
| PDM3 |  | 2.848% | -12.38% | 15.39% |  | -5.304% | 3.674% | -1.542% | 7.865% | -2.796% | 14.14% | -2.796% | 2.221% | -4.05% | -1.917% | 14.76% | -0.2874% | -2.231% | 45.49% | 7.238% | 0.9669% | 5.357% | -16.47% | -7.813% | -4.677% | -12.2%
| Z-Score (Absolute) |  | 0.57 | 1.4 | 2.19 | 12.61 | 0.49 | 0.67 | 0 | 1.22 | 0.16 | 2.03 | 0.16 | 0.49 | 0.32 | 0.05 | 2.11 | 0.16 | 0.09 | 6.08 | 1.14 | 0.32 | 0.89 | 1.93 | 0.81 | 0.41 | 1.38
| Analytical results for Zr in OREAS 163b (Certified Value 12.1 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 12.1 | 11.92 | 10 | 21.73 | 11.6 | 13.0937 | 11.6 | 11.6 | 12.6 | 12.6 | 14.7 | 15 | 11.4 | 11.9948 | 13.4 | 13.2 | 13.1 | 11.5 | 8.2 | 9.6 | 11.6 | 13.1 | 12.2 | 13 |  | 1
| 1 | 2 | 12.3 | 12.24 | 10.5 | 20.29 | 11.7 | 12.7323 | 11.7 | 11.8 | 13.8 | 13.3 | 12 | 14 | 12.9 | 12.9981 | 10.8 | 13.4 | 12.2 | 11.4 | 8.7 | 9.6 | 11.1 | 13.2 | 11.9 | 12.9 |  | 41
| 1 | 3 | 13 | 12.39 | 10.1 | 30.7 | 12.7 | 12.9904 | 12 | 12.2 | 12.9 | 12 | 11.6 | 15 | 11.8 | 12.4902 | 10 | 13.1 | 11.6 | 10.8 | 9 | 9.6 | 12.2 | 14.4 | 12.3 | 13.2 |  | 16
| 1 | 4 | 12.5 | 11.51 | 10.3 | 21.23 | 11.8 | 13.1456 | 12.7 | 13 | 13.8 | 11.9 | 11.9 | 13 | 11.8 | 14.0177 | 12.2 | 12.3 | 12.4 | 11.3 | 7.9 | 9.6 | 10.8 | 14.4 | 12.4 | 13.7 |  | 12.0591
| 1 | 5 | 12.5 | 10.55 | 10.4 | 16.34 | 11.8 | 13.359 | 11.4 | 13 | 11.9 | 12.6 | 12.4 | 16 | 12.7 | 12.6303 | 10 | 12.8 | 12.6 | 10.9 | 8 | 9.6 | 11.1 | 12.5 | 11.4 | 13.2 |  | 66
| 1 | 6 | 13.3 | 12.05 | 10 | 20.54 | 12.3 | 12.893 | 11.4 | 11.8 | 14.4 | 19.6 | 11.3 | 15 | 11.7 | 14.4258 | 9.5 | 13.4 | 11 | 12 | 8 | 9.7 | 11.1 | 12.1 | 11.7 | 13.2
| Mean |  | 12.6167 | 11.7767 | 10.2167 | 21.805 | 11.9833 | 13.0357 | 11.8 | 12.2333 | 13.2333 | 13.6667 | 12.3167 | 14.6667 | 12.05 | 13.0928 | 10.9833 | 13.0333 | 12.15 | 11.3167 | 8.3 | 9.6167 | 11.3167 | 13.2833 | 11.9833 | 13.2
| Median |  | 12.5 | 11.985 | 10.2 | 20.885 | 11.8 | 13.042 | 11.65 | 12 | 13.35 | 12.6 | 11.95 | 15 | 11.8 | 12.8142 | 10.4 | 13.15 | 12.3 | 11.35 | 8.1 | 9.6 | 11.1 | 13.15 | 12.05 | 13.2
| Std Dev. |  | 0.44907 | 0.6726 | 0.2137 | 4.7585 | 0.42622 | 0.21654 | 0.49396 | 0.62503 | 0.92664 | 2.95 | 1.2254 | 1.0328 | 0.60249 | 0.94063 | 1.5158 | 0.42269 | 0.74766 | 0.43551 | 0.44721 | 0.040825 | 0.50365 | 0.95376 | 0.38687 | 0.27568
| Rel.Std.Dev. |  | 3.559% | 5.711% | 2.092% | 21.82% | 3.557% | 1.661% | 4.186% | 5.109% | 7.002% | 21.59% | 9.949% | 7.042% | 5% | 7.184% | 13.8% | 3.243% | 6.154% | 3.848% | 5.388% | 0.4245% | 4.451% | 7.18% | 3.228% | 2.088%
| PDM3 |  | 4.623% | -2.342% | -15.28% | 80.82% | -0.6286% | 8.098% | -2.149% | 1.445% | 9.737% | 13.33% | 2.136% | 21.62% | -0.07575% | 8.572% | -8.921% | 8.079% | 0.7535% | -6.157% | -31.17% | -20.25% | -6.157% | 10.15% | -0.6286% | 9.461%
| Z-Score (Absolute) |  | 0.36 | 0.28 | 1.48 | 7.37 | 0.13 | 0.68 | 0.27 | 0.06 | 0.83 | 1.16 | 0.13 | NA | 0.08 | 0.72 | 0.89 | 0.67 | 0 | 0.64 | 2.94 | 1.93 | 0.64 | 0.87 | 0.13 | 0.8
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C273: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 163b (Certified Value 4.2 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 4.25 | 4.8116 | 4.447 | 4.44 | 4.25 | 4.1122 | 4.49 | 4.22 | 4.5 | 4.08 | 4.15 | 4.21 | 4.56 | 4.3286 | 4.13 | 4.094 | 4.1 | 4.2 | 4.09 | 3.99 | 4.3 | 4.7 | 4.35 | 4.1 | 1
| 1 | 2 | 4.23 | 4.9128 | 4.6 | 4.26 | 4.17 | 4.1079 | 4.39 | 4.07 | 4.3 | 4.07 | 4 | 4.11 | 4.54 | 3.9648 | 4.01 | 4.125 | 4 | 4.1 | 4.19 | 3.99 | 4.4 | 4.7 | 4.49 | 4 | 8
| 1 | 3 | 4.27 | 4.9588 | 4.641 | 4.62 | 4.26 | 4.0749 | 4.32 | 4.13 | 4.36 | 4.04 | 4.06 | 4.11 | 4.46 | 4.4104 | 4.19 | 3.883 | 4.1 | 3.9 | 4.16 | 3.98 | 4.4 | 4.7 | 4.21 | 4.1 | 16
| 1 | 4 | 4.22 | 4.6736 | 4.695 | 4.43 | 4.19 | 4.092 | 4.39 | 3.95 | 4.47 | 4.09 | 3.99 | 4.17 | 4.46 | 3.8706 | 4.01 | 3.833 | 4.1 | 4.1 | 4.08 | 3.98 | 4.4 | 4.8 | 4.43 | 4.1 | 4.2003
| 1 | 5 | 4.3 | 4.8852 | 4.536 | 4.1 | 4.2 | 4.1956 | 4.32 | 4.25 | 4.49 | 4.08 | 4.06 | 4.27 | 4.47 | 4.1942 | 4.11 | 3.999 | 4.1 | 4 | 4.16 | 3.95 | 4.4 | 4.7 | 4.22 | 4.2 | 68
| 1 | 6 | 4.26 | 4.968 | 4.6 | 4.19 | 4.52 | 4.0767 | 4.28 | 3.98 | 4.34 | 4.13 | 3.89 | 4.01 | 4.45 | 4.3575 | 3.83 | 3.987 | 4 | 3.9 | 4.32 | 3.94 | 4.4 | 4.8 | 4.44 | 4
| Mean |  | 4.255 | 4.8683 | 4.5865 | 4.34 | 4.265 | 4.1099 | 4.365 | 4.1 | 4.41 | 4.0817 | 4.025 | 4.1467 | 4.49 | 4.1877 | 4.0467 | 3.9868 | 4.0667 | 4.0333 | 4.1667 | 3.9717 | 4.3833 | 4.7333 | 4.3567 | 4.0833
| Median |  | 4.255 | 4.899 | 4.6 | 4.345 | 4.225 | 4.0999 | 4.355 | 4.1 | 4.415 | 4.08 | 4.03 | 4.14 | 4.465 | 4.2614 | 4.06 | 3.993 | 4.1 | 4.05 | 4.16 | 3.98 | 4.4 | 4.7 | 4.39 | 4.1
| Std Dev. |  | 0.02881 | 0.11092 | 0.08616 | 0.19131 | 0.12973 | 0.044716 | 0.075033 | 0.12296 | 0.086718 | 0.029269 | 0.08735 | 0.090701 | 0.047329 | 0.22293 | 0.12738 | 0.11417 | 0.05164 | 0.12111 | 0.086641 | 0.02137 | 0.040825 | 0.05164 | 0.1186 | 0.075277
| Rel.Std.Dev. |  | 0.6771% | 2.278% | 1.879% | 4.408% | 3.042% | 1.088% | 1.719% | 2.999% | 1.966% | 0.7171% | 2.17% | 2.187% | 1.054% | 5.324% | 3.148% | 2.864% | 1.27% | 3.003% | 2.079% | 0.5381% | 0.9314% | 1.091% | 2.722% | 1.844%
| PDM3 |  | 1.302% | 15.9% | 9.195% | 3.326% | 1.54% | -2.153% | 3.921% | -2.388% | 4.993% | -2.824% | -4.173% | -1.277% | 6.897% | -0.3001% | -3.658% | -5.082% | -3.181% | -3.975% | -0.8007% | -5.443% | 4.358% | 12.69% | 3.723% | -2.785%
| Z-Score (Absolute) |  | 0.35 | 3.15 | 1.86 | 0.74 | 0.4 | 0.31 | 0.86 | 0.35 | 1.06 | 0.44 | 0.69 | 0.14 | 1.43 | 0.05 | 0.59 | 0.87 | 0.5 | 0.66 | 0.05 | 0.94 | 0.94 | 2.53 | 0.82 | 0.43
| Analytical results for Al in OREAS 163b (Certified Value 0.126 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.156 | 0.13743 | 0.09 | 0.126 | 0.12918 | 0.12 | 0.12 | 0.13 | 0.11 | 0.12 | 0.13 | 0.132 | 0.1402 | 0.12 | 0.12 | 0.12 | 0.12 | 0.13 | 0.13 | 0.1423 | 0.12 | 0.12 |  |  | 1
| 1 | 2 | 0.154 | 0.12905 | 0.09 | 0.121 | 0.12955 | 0.12 | 0.12 | 0.13 | 0.11 | 0.12 | 0.12 | 0.134 | 0.13535 | 0.07 | 0.12 | 0.12 | 0.13 | 0.12 | 0.14 | 0.1462 | 0.12 | 0.12 |  |  | 0
| 1 | 3 | 0.153 | 0.1326 | 0.1 | 0.122 | 0.12989 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.11 | 0.136 | 0.14947 | 0.08 | 0.13 | 0.12 | 0.13 | 0.13 | 0.14 | 0.1421 | 0.12 | 0.12 |  |  | 16
| 1 | 4 | 0.155 | 0.135 | 0.09 | 0.126 | 0.12905 | 0.12 | 0.12 | 0.13 | 0.11 | 0.12 | 0.12 | 0.129 | 0.14001 | 0.08 | 0.12 | 0.12 | 0.12 | 0.13 | 0.14 | 0.1448 | 0.12 | 0.12 |  |  | 0.12605
| 1 | 5 | 0.159 | 0.1341 | 0.09 | 0.119 | 0.13265 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12 | 0.11 | 0.131 | 0.14308 | 0.08 | 0.13 | 0.12 | 0.12 | 0.13 | 0.14 | 0.1432 | 0.12 | 0.12 |  |  | 69
| 1 | 6 | 0.151 | 0.1305 | 0.09 | 0.126 | 0.12679 | 0.11 | 0.12 | 0.13 | 0.11 | 0.12 | 0.11 | 0.131 | 0.14948 | 0.08 | 0.12 | 0.12 | 0.13 | 0.13 | 0.14 | 0.1448 | 0.12 | 0.12
| Mean |  | 0.15467 | 0.13311 | 0.091667 | 0.12333 | 0.12952 | 0.11667 | 0.12 | 0.13 | 0.11167 | 0.12 | 0.11667 | 0.13217 | 0.14293 | 0.085 | 0.12333 | 0.12 | 0.125 | 0.12833 | 0.13833 | 0.1439 | 0.12 | 0.12
| Median |  | 0.1545 | 0.13335 | 0.09 | 0.124 | 0.12937 | 0.12 | 0.12 | 0.13 | 0.11 | 0.12 | 0.115 | 0.1315 | 0.14164 | 0.08 | 0.12 | 0.12 | 0.125 | 0.13 | 0.14 | 0.144 | 0.12 | 0.12
| Std Dev. |  | 0.0027325 | 0.0030574 | 0.0040825 | 0.0030768 | 0.0018832 | 0.005164 | 0 | 0 | 0.0040825 | 0 | 0.008165 | 0.0024833 | 0.0056415 | 0.017607 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0016248 | 0 | 0
| Rel.Std.Dev. |  | 1.767% | 2.297% | 4.454% | 2.495% | 1.454% | 4.426% | 0% | 0% | 3.656% | 0% | 6.999% | 1.879% | 3.947% | 20.71% | 4.187% | 0% | 4.382% | 3.181% | 2.951% | 1.129% | 0% | 0%
| PDM3 |  | 22.7% | 5.603% | -27.28% | -2.156% | 2.753% | -7.445% | -4.8% | 3.133% | -11.41% | -4.8% | -7.445% | 4.852% | 13.39% | -32.57% | -2.156% | -4.8% | -0.8337% | 1.811% | 9.744% | 14.16% | -4.8% | -4.8%
| Z-Score (Absolute) |  | 3.17 | 0.99 | 3.2 | 0 | 0.63 | 0.67 | 0.34 | 0.67 | 1.18 | 0.34 | 0.67 | 0.89 | 1.98 | 3.88 | 0 | 0.34 | 0.17 | 0.51 | 1.52 | 2.08 | 0.34 | 0.34
| Analytical results for As in OREAS 163b (Certified Value 165 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 0
| 1 | 1 | 163 | 189.5 | 160.7 | 165.22 | 162.8504 | 169.5 | 175 | 180 | 163 | 171 | 163 | 161 | 157.8067 | 180 | 166.4 | 163.7 | 180.9 | 152 | 165 | 158 | 179 | 169.5 | 163.5 |  | 1
| 1 | 2 | 162 | 190.87 | 159.8 | 160.21 | 166.0367 | 168 | 173.5 | 177.5 | 162 | 166.5 | 161 | 156 | 146.9759 | 191 | 168.9 | 160.6 | 176.9 | 145 | 167 | 164 | 182 | 169.1 | 165.1 |  | 9
| 1 | 3 | 161 | 185.16 | 165.3 | 163.21 | 164.7574 | 166 | 173.5 | 178 | 160 | 167.5 | 164 | 161 | 159.5991 | 193 | 170.5 | 163.1 | 166.7 | 155 | 167 | 160 | 178 | 163.2 | 162.1 |  | 16
| 1 | 4 | 166 | 186.16 | 166 | 162.94 | 163.5593 | 167.5 | 168 | 180.5 | 158 | 168.5 | 162 | 160 | 147.2196 | 189 | 154.9 | 165.4 | 172.6 | 153 | 166 | 155 | 182 | 168.6 | 165.6 |  | 164.6445
| 1 | 5 | 163 | 185.93 | 173.1 | 161.31 | 165.3481 | 168 | 173.5 | 181 | 161.5 | 168 | 168 | 164 | 156.0204 | 193 | 170.4 | 166.8 | 170.8 | 151 | 165 | 157 | 181 | 168 | 162.2 |  | 70
| 1 | 6 | 161 | 187.86 | 171.5 | 166.06 | 161.4852 | 163.5 | 170.5 | 178 | 162 | 165.5 | 158 | 149 | 157.7882 | 183 | 165.5 | 160.5 | 166.3 | 154 | 165 | 157 | 180 | 165.7 | 163.1
| Mean |  | 162.6667 | 187.58 | 166.0667 | 163.1583 | 164.0062 | 167.0833 | 172.3333 | 179.1667 | 161.0833 | 167.8333 | 162.6667 | 158.5 | 154.235 | 188.1667 | 166.1 | 163.35 | 172.3667 | 151.6667 | 165.8333 | 158.5 | 180.3333 | 167.35 | 163.6
| Median |  | 162.5 | 187.01 | 165.65 | 163.075 | 164.1584 | 167.75 | 173.5 | 179 | 161.75 | 167.75 | 162.5 | 160.5 | 156.9043 | 190 | 167.65 | 163.4 | 171.7 | 152.5 | 165.5 | 157.5 | 180.5 | 168.3 | 163.3
| Std Dev. |  | 1.8619 | 2.2443 | 5.4342 | 2.2284 | 1.6958 | 2.0837 | 2.582 | 1.5055 | 1.8005 | 1.8886 | 3.3267 | 5.3198 | 5.6437 | 5.4559 | 5.858 | 2.5288 | 5.741 | 3.559 | 0.98319 | 3.1464 | 1.633 | 2.4337 | 1.4642
| Rel.Std.Dev. |  | 1.145% | 1.196% | 3.272% | 1.366% | 1.034% | 1.247% | 1.498% | 0.8403% | 1.118% | 1.125% | 2.045% | 3.356% | 3.659% | 2.899% | 3.527% | 1.548% | 3.331% | 2.347% | 0.5929% | 1.985% | 0.9055% | 1.454% | 0.895%
| PDM3 |  | -1.201% | 13.93% | 0.8638% | -0.9027% | -0.3877% | 1.481% | 4.67% | 8.82% | -2.163% | 1.937% | -1.201% | -3.732% | -6.322% | 14.29% | 0.884% | -0.7863% | 4.69% | -7.882% | 0.722% | -3.732% | 9.529% | 1.643% | -0.6344%
| Z-Score (Absolute) |  | 0.67 | 4.63 | 0.05 | 0.57 | 0.39 | 0.27 | 1.38 | 2.84 | 1.01 | 0.43 | 0.67 | 1.56 | 2.47 | 4.76 | 0.06 | 0.53 | 1.39 | 3.02 | 0 | 1.56 | 3.09 | 0.32 | 0.48
| Analytical results for Au in OREAS 163b (Certified Value < 0.02 ppm) | [25 empty] | Certified
Au | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  | 1
| No | No | 2.01: 2.05; 2.02: 2.07; 2.04: 2.08; 2.05: 2.09; 2.06: 2.10; 2.07: 2.11; 2.08: 2.12; 2.09: 2.14; 2.10: 2.17; 2.11: 2.19; 2.12: 2.24 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*OES/MS; 2.09: AR*OES/MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*OES/MS |  | 3
|  |  | 2.01: 0.5g; 2.02: 0.5g; 2.04: 0.5g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 01g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g |  | 3
| 1 | 1 | 2.01: 0.003; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0092; 2.11: 0.0041; 2.12: 0.0007 |  | 1
| 1 | 2 | 2.01: 0.0027; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: 0.0014; 2.09: < 0.001; 2.10: 0.0075; 2.11: 0.0039; 2.12: 0.0011 |  | 10
| 1 | 3 | 2.01: 0.005; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0086; 2.11: 0.0036; 2.12: 0.0008 |  | 16
| 1 | 4 | 2.01: 0.0029; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: 0.0012; 2.09: < 0.001; 2.10: 0.0087; 2.11: 0.0043; 2.12: 0.0012 |  | < 0.02
| 1 | 5 | 2.01: 0.0025; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: < 0.0005; 2.09: < 0.001; 2.10: 0.0087; 2.11: 0.0046; 2.12: 0.0007 |  | 71
| 1 | 6 | 2.01: 0.0029; 2.02: < 0.02; 2.04: < 0.02; 2.05: < 0.02; 2.06: < 0.02; 2.07: < 0.02; 2.08: 0.0012; 2.09: < 0.001; 2.10: 0.0086; 2.11: 0.0037; 2.12: 0.0005
| Mean |  | 2.01: 0.0031667; 2.08: 0.0012667; 2.10: 0.00855; 2.11: 0.0040333; 2.12: 0.00083333
| Median |  | 2.01: 0.0029; 2.08: 0.0012; 2.10: 0.00865; 2.11: 0.004; 2.12: 0.00075
| Std Dev. |  | 2.01: 0.00091579; 2.08: 0.00011547; 2.10: 0.00056125; 2.11: 0.00037771; 2.12: 0.00026583
| Rel.Std.Dev. |  | 2.01: 28.92%; 2.08: 9.116%; 2.10: 6.564%; 2.11: 9.365%; 2.12: 31.9%
PDM3
| Z-Score (Absolute) |  | 2.01: 2.5; 2.02: 0.67; 2.04: 0.67; 2.05: 0.67; 2.06: 0.67; 2.07: 0.67; 2.08: 3.62; 2.09: 3.74; 2.10: 0; 2.11: 2.1; 2.12: 3.59
| Analytical results for B in OREAS 163b (Certified Value < 10 ppm) | [25 empty] | Certified
B | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  | 1
| No | No | 2.01: 2.06; 2.02: 2.07; 2.04: 2.08; 2.05: 2.09; 2.06: 2.10; 2.07: 2.11; 2.08: 2.12; 2.09: 2.14; 2.10: 2.17; 2.11: 2.19; 2.12: 2.20; 2.13: 2.24 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*OES/MS; 2.09: AR*OES/MS; 2.10: AR*MS; 2.11: AR*MS; 2.12: AR*MS; 2.13: AR*OES/MS |  | 1
|  |  | 2.01: 0.2g; 2.02: 0.5g; 2.04: 0.5g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 01g; 2.10: 0.5g; 2.11: 0.5g; 2.12: 0.5g; 2.13: 0.5g |  | 1
| 1 | 1 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 4; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 1
| 1 | 2 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 5; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 1
| 1 | 3 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 5; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 16
| 1 | 4 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 5; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | < 10
| 1 | 5 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 4; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10 |  | 72
| 1 | 6 | 2.01: < 2; 2.02: < 10; 2.04: < 10; 2.05: < 10; 2.06: < 10; 2.07: < 10; 2.08: 5; 2.09: < 10; 2.10: < 20; 2.11: < 20; 2.12: < 10; 2.13: < 10
| Mean |  | 2.08: 4.6667
| Median |  | 2.08: 5
| Std Dev. |  | 2.08: 0.5164
| Rel.Std.Dev. |  | 2.08: 11.07%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.13
| Analytical results for Ba in OREAS 163b (Certified Value 26.4 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 27 | 25.9 | 22.25 | 27.1096 | 30 | 30 | 30 | 20 | 30 | 16.2 | 26 | 27.1844 | < 10 | 29 | 28 | 22 | 40 | 29 | 24 | 27 |  |  |  |  | 1
| 1 | 2 | 28 | 29.3 | 22.5 | 27.7161 | 30 | 30 | 30 | 30 | 30 | 15 | 27 | 30.4789 | 10 | 28.8 | 27 | 21 | 38 | 30 | 24 | 28 |  |  |  |  | 27
| 1 | 3 | 27 | 28.1 | 22.23 | 26.8532 | 30 | 30 | 30 | 30 | 30 | 15.5 | 26 | 30.3531 | 10 | 28.8 | 26 | 23 | 40 | 30 | 24 | 26 |  |  |  |  | 16
| 1 | 4 | 27 | 28.5 | 21.26 | 26.5884 | 20 | 30 | 30 | 30 | 30 | 18.1 | 26 | 27.6335 | 10 | 27.2 | 27 | 22 | 40 | 30 | 22 | 28 |  |  |  |  | 26.4313
| 1 | 5 | 29 | 29.9 | 21.77 | 28.8807 | 30 | 30 | 30 | 30 | 30 | 12.8 | 26 | 26.5997 | 10 | 29.7 | 26 | 21 | 38 | 30 | 24 | 26 |  |  |  |  | 73
| 1 | 6 | 28 | 27.7 | 21.8 | 26.8789 | 20 | 30 | 30 | 30 | 30 | 14.7 | 26 | 27.5351 | < 10 | 28.6 | 25 | 23 | 39 | 29 | 23 | 28
| Mean |  | 27.6667 | 28.2333 | 21.9683 | 27.3378 | 26.6667 | 30 | 30 | 28.3333 | 30 | 15.3833 | 26.1667 | 28.2975 | 10 | 28.6833 | 26.5 | 22 | 39.1667 | 29.6667 | 23.5 | 27.1667
| Median |  | 27.5 | 28.3 | 22.015 | 26.9942 | 30 | 30 | 30 | 30 | 30 | 15.25 | 26 | 27.5843 | 10 | 28.8 | 26.5 | 22 | 39.5 | 30 | 24 | 27.5
| Std Dev. |  | 0.8165 | 1.3952 | 0.44719 | 0.84647 | 5.164 | 0 | 0 | 4.0825 | 0 | 1.752 | 0.40825 | 1.6809 | 0 | 0.82077 | 1.0488 | 0.89443 | 0.98319 | 0.5164 | 0.83666 | 0.98319
| Rel.Std.Dev. |  | 2.951% | 4.942% | 2.036% | 3.096% | 19.36% | 0% | 0% | 14.41% | 0% | 11.39% | 1.56% | 5.94% | 0% | 2.861% | 3.958% | 4.066% | 2.51% | 1.741% | 3.56% | 3.619%
| PDM3 |  | 4.674% | 6.818% | -16.89% | 3.43% | 0.8907% | 13.5% | 13.5% | 7.196% | 13.5% | -41.8% | -1.001% | 7.061% | -62.17% | 8.521% | 0.2601% | -16.77% | 48.18% | 12.24% | -11.09% | 2.782%
| Z-Score (Absolute) |  | 0.22 | 0.53 | 2.83 | 0.05 | NA | NA | NA | NA | NA | 6.36 | 0.58 | 0.56 | NA | 0.77 | 0.4 | 2.81 | 6.38 | 1.29 | 2.01 | 0.05
Results from laboratories 2.07, 2.08, 2.09, 2.10, 2.11 and 2.15 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 163b (Certified Value 0.11 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.5 | < 0.1 | < 2 | 0.12 | 0.11808 | 0.11 | 0.11 | 0.12 | 0.11 | 0.1 | 0.2 | 0.12 | < 0.5 | 0.2 | 0.12 | 0.1 | 0.1 | < 1 | [6 empty] | 1
| 1 | 2 | < 0.5 | < 0.1 | < 2 | 0.1 | 0.12269 | 0.11 | 0.11 | 0.11 | 0.1 | 0.11 | 0.2 | 0.09 | < 0.5 | < 0.05 | 0.12 | 0.1 | 0.1 | < 1 | [6 empty] | 28
| 1 | 3 | < 0.5 | < 0.1 | < 2 | 0.1 | 0.12086 | 0.11 | 0.11 | 0.12 | 0.1 | 0.1 | 0.3 | 0.12 | < 0.5 | 0.13 | 0.11 | 0.1 | 0.11 | < 1 | [6 empty] | 16
| 1 | 4 | < 0.5 | < 0.1 | < 2 | 0.11 | 0.10529 | 0.11 | 0.11 | 0.12 | 0.1 | 0.1 | 0.2 | 0.1 | < 0.5 | 0.07 | 0.11 | 0.1 | 0.1 | < 1 | [6 empty] | 0.10972
| 1 | 5 | < 0.5 | < 0.1 | < 2 | 0.11 | 0.11306 | 0.11 | 0.11 | 0.12 | 0.1 | 0.1 | 0.2 | 0.13 | < 0.5 | 0.06 | 0.11 | 0.1 | 0.1 | < 1 | [6 empty] | 74
| 1 | 6 | < 0.5 | < 0.1 | < 2 | 0.1 | 0.11329 | 0.11 | 0.11 | 0.12 | 0.1 | 0.11 | 0.2 | 0.11 | < 0.5 | 0.06 | 0.12 | 0.1 | 0.12 | < 1
| Mean |  |  |  |  | 0.10667 | 0.11554 | 0.11 | 0.11 | 0.11833 | 0.10167 | 0.10333 | 0.21667 | 0.11167 |  | 0.104 | 0.115 | 0.1 | 0.105
| Median |  |  |  |  | 0.105 | 0.11568 | 0.11 | 0.11 | 0.12 | 0.1 | 0.1 | 0.2 | 0.115 |  | 0.07 | 0.115 | 0.1 | 0.1
| Std Dev. |  |  |  |  | 0.008165 | 0.0063591 | 0 | 0 | 0.0040825 | 0.0040825 | 0.005164 | 0.040825 | 0.01472 |  | 0.061074 | 0.0054772 | 1.5202e-17 | 0.0083666
| Rel.Std.Dev. |  |  |  |  | 7.655% | 5.504% | 0% | 0% | 3.45% | 4.016% | 4.997% | 18.84% | 13.18% |  | 58.72% | 4.763% | 1.52e-14% | 7.968%
| PDM3 |  |  |  |  | -2.784% | 5.308% | 0.2541% | 0.2541% | 7.849% | -7.341% | -5.822% | 97.47% | 1.773% |  | -5.214% | 4.811% | -8.86% | -4.303%
| Z-Score (Absolute) |  | 12.51 | 5.47 | 79.94 | 0.37 | 0.42 | 0.07 | 0.07 | 0.67 | 0.82 | 0.67 | NA | 0.07 | 12.51 | 1.8 | 0.37 | NA | 0.52 | 34.99
Results from laboratories 2.12 and 2.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 163b (Certified Value 17.1 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 2
| 1 | 1 | 16.3 | 15.392 | 17.3 | 11.8 | 17.28 | 17.1088 | 16.65 | 17.95 | 17.4 | 15.9 | 19.1 | 17.1 | 17.07 | 17.1583 | 19 | 15.42 | 11.99 | 18.7 | 17.91 | 18.5 | 18 | 18 | 17.43 | 16.7 | 1
| 1 | 2 | 16.3 | 15.667 | 17.3 | 10.7 | 16.92 | 18.9229 | 17.05 | 17.8 | 17.55 | 16.95 | 18.2 | 17.1 | 16.74 | 16.9273 | 17.2 | 15.61 | 12.99 | 18.2 | 16.29 | 17.5 | 18.1 | 18 | 17.44 | 16.5 | 29
| 1 | 3 | 16.3 | 15.275 | 17.5 | 11.1 | 17.13 | 17.5414 | 16.75 | 18 | 17.45 | 16.45 | 18.7 | 16.9 | 17.16 | 18.1635 | 17 | 15.18 | 14.99 | 17.3 | 16.92 | 17.9 | 17.6 | 18 | 16.88 | 17.4 | 16
| 1 | 4 | 16.4 | 15.353 | 17.3 | 14 | 16.08 | 16.2702 | 16.7 | 17.3 | 17.65 | 17 | 18.8 | 17.2 | 15.87 | 17.0999 | 17.1 | 15.39 | 12.99 | 17.9 | 16.52 | 18 | 17.7 | 18 | 17.86 | 17.2 | 17.0565
| 1 | 5 | 16.4 | 15.283 | 17.4 | 11.7 | 16.6 | 18.1379 | 16.95 | 17.95 | 17.55 | 16.05 | 18.9 | 17.6 | 16.55 | 17.2668 | 16.9 | 15.46 | 12.99 | 17.9 | 16.35 | 18.3 | 17.5 | 18 | 17.19 | 16.4 | 75
| 1 | 6 | 16.1 | 15.361 | 17.5 | 11.7 | 16.77 | 15.9039 | 16.6 | 17.35 | 17.45 | 16.85 | 18.7 | 16.6 | 16.78 | 17.9749 | 16.3 | 15.35 | 12.99 | 17.3 | 16.23 | 17.8 | 18 | 18 | 17.28 | 16
| Mean |  | 16.3 | 15.3885 | 17.3833 | 11.8333 | 16.7967 | 17.3142 | 16.7833 | 17.725 | 17.5083 | 16.5333 | 18.7333 | 17.0833 | 16.695 | 17.4318 | 17.25 | 15.4017 | 13.1567 | 17.8833 | 16.7033 | 18 | 17.8167 | 18 | 17.3467 | 16.7
| Median |  | 16.3 | 15.357 | 17.35 | 11.7 | 16.845 | 17.3251 | 16.725 | 17.875 | 17.5 | 16.65 | 18.75 | 17.1 | 16.76 | 17.2126 | 17.05 | 15.405 | 12.99 | 17.9 | 16.435 | 17.95 | 17.85 | 18 | 17.355 | 16.6
| Std Dev. |  | 0.10954 | 0.14395 | 0.098319 | 1.1448 | 0.4273 | 1.1348 | 0.17795 | 0.31741 | 0.091742 | 0.4761 | 0.30111 | 0.33116 | 0.46211 | 0.50933 | 0.91378 | 0.14077 | 0.98319 | 0.53821 | 0.6413 | 0.35777 | 0.24833 | 0 | 0.32432 | 0.52154
| Rel.Std.Dev. |  | 0.6721% | 0.9354% | 0.5656% | 9.675% | 2.544% | 6.554% | 1.06% | 1.791% | 0.524% | 2.88% | 1.607% | 1.938% | 2.768% | 2.922% | 5.297% | 0.914% | 7.473% | 3.01% | 3.839% | 1.988% | 1.394% | 0% | 1.87% | 3.123%
| PDM3 |  | -4.435% | -9.779% | 1.916% | -30.62% | -1.523% | 1.511% | -1.602% | 3.919% | 2.649% | -3.067% | 9.831% | 0.1573% | -2.119% | 2.2% | 1.134% | -9.702% | -22.86% | 4.848% | -2.071% | 5.532% | 4.457% | 5.532% | 1.701% | -2.09%
| Z-Score (Absolute) |  | 1.24 | 2.69 | 0.48 | 8.33 | 0.45 | 0.37 | 0.48 | 1.02 | 0.67 | 0.87 | 2.62 | 0 | 0.62 | 0.55 | 0.26 | 2.67 | 6.23 | 1.27 | 0.6 | 1.45 | 1.16 | NA | 0.42 | 0.61
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 163b (Certified Value 2.53 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 2.47 | 2.5188 | 2.35 | 2.4404 | 2.59 | 2.61 | 2.65 | 2.51 | 2.57 | 2.27 | 2.723 | 2.5379 | 2.98 | 2.38 | 2.7487 | 2.56 | 2.51 | 2.63 | 2.489 | 2.56 | 2.62 | 2.18 |  |  | 1
| 1 | 2 | 2.46 | 2.4443 | 2.39 | 2.4504 | 2.58 | 2.6 | 2.65 | 2.52 | 2.54 | 2.25 | 2.756 | 2.5229 | 3.03 | 2.29 | 2.7067 | 2.49 | 2.54 | 2.64 | 2.533 | 2.56 | 2.57 | 2.18 |  |  | 0
| 1 | 3 | 2.45 | 2.5485 | 2.47 | 2.434 | 2.51 | 2.59 | 2.66 | 2.52 | 2.52 | 2.46 | 2.74 | 2.5436 | 3.27 | 2.28 | 2.7553 | 2.4 | 2.55 | 2.63 | 2.49 | 2.54 | 2.5 | 2.19 |  |  | 16
| 1 | 4 | 2.46 | 2.6735 | 2.38 | 2.426 | 2.55 | 2.51 | 2.69 | 2.48 | 2.58 | 2.61 | 2.606 | 2.5483 | 3.26 | 2.26 | 2.7779 | 2.45 | 2.5 | 2.63 | 2.494 | 2.56 | 2.58 | 2.19 |  |  | 2.5317
| 1 | 5 | 2.47 | 2.5249 | 2.46 | 2.4444 | 2.54 | 2.6 | 2.7 | 2.51 | 2.51 | 2.48 | 2.65 | 2.5268 | 3.15 | 2.29 | 2.8071 | 2.39 | 2.52 | 2.62 | 2.489 | 2.55 | 2.6 | 2.18 |  |  | 76
| 1 | 6 | 2.42 | 2.5641 | 2.45 | 2.4123 | 2.49 | 2.56 | 2.64 | 2.54 | 2.57 | 2.31 | 2.71 | 2.5434 | 3.24 | 2.39 | 2.6981 | 2.41 | 2.52 | 2.63 | 2.489 | 2.57 | 2.54 | 2.13
| Mean |  | 2.455 | 2.5457 | 2.4167 | 2.4346 | 2.5433 | 2.5783 | 2.665 | 2.5133 | 2.5483 | 2.3967 | 2.6975 | 2.5371 | 3.155 | 2.315 | 2.749 | 2.45 | 2.5233 | 2.63 | 2.4973 | 2.5567 | 2.5683 | 2.175
| Median |  | 2.46 | 2.5367 | 2.42 | 2.4372 | 2.545 | 2.595 | 2.655 | 2.515 | 2.555 | 2.385 | 2.7165 | 2.5406 | 3.195 | 2.29 | 2.752 | 2.43 | 2.52 | 2.63 | 2.4895 | 2.56 | 2.575 | 2.18
| Std Dev. |  | 0.018708 | 0.074985 | 0.049666 | 0.013794 | 0.038816 | 0.037639 | 0.02429 | 0.019664 | 0.029269 | 0.1425 | 0.05773 | 0.010157 | 0.1247 | 0.055408 | 0.041547 | 0.065422 | 0.018619 | 0.0063246 | 0.01758 | 0.010328 | 0.043089 | 0.022583
| Rel.Std.Dev. |  | 0.762% | 2.946% | 2.055% | 0.5666% | 1.526% | 1.46% | 0.9114% | 0.7824% | 1.149% | 5.946% | 2.14% | 0.4003% | 3.952% | 2.393% | 1.511% | 2.67% | 0.7379% | 0.2405% | 0.704% | 0.404% | 1.678% | 1.038%
| PDM3 |  | -3.031% | 0.5505% | -4.545% | -3.837% | 0.4577% | 1.84% | 5.263% | -0.7272% | 0.6552% | -5.335% | 6.547% | 0.2134% | 24.62% | -8.561% | 8.58% | -3.229% | -0.3322% | 3.881% | -1.359% | 0.9844% | 1.445% | -14.09%
| Z-Score (Absolute) |  | 0.66 | 0.04 | 0.95 | 0.81 | 0.02 | 0.29 | 0.96 | 0.21 | 0.06 | 1.11 | 1.21 | 0.02 | 4.74 | 1.74 | 1.61 | 0.7 | 0.13 | 0.69 | 0.33 | 0.13 | 0.22 | 2.81
| Analytical results for Cd in OREAS 163b (Certified Value 0.44 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.43 | 0.3574 | 0.46 | < 1 | 0.452 | 0.42347 | 0.45 | 0.45 | 0.45 | 0.41 | 0.45 | 0.52 | 0.49 | 0.38472 | 2.62 | 0.43 | 0.5 | 0.45 | 0.43 | 0.5 | 0.46 | 0.4 |  |  | 1
| 1 | 2 | 0.46 | 0.3438 | 0.45 | < 1 | 0.443 | 0.39505 | 0.46 | 0.49 | 0.43 | 0.43 | 0.42 | 0.46 | 0.48 | 0.3654 | 2.71 | 0.41 | 0.4 | 0.46 | 0.43 | 0.5 | 0.45 | 0.4 |  |  | 11
| 1 | 3 | 0.44 | 0.339 | 0.44 | < 1 | 0.446 | 0.3976 | 0.43 | 0.47 | 0.44 | 0.41 | 0.43 | 0.54 | 0.46 | 0.37167 | 2.82 | 0.46 | 0.4 | 0.45 | 0.41 |  | 0.47 | 0.4 |  |  | 16
| 1 | 4 | 0.43 | 0.3522 | 0.48 | < 1 | 0.447 | 0.45085 | 0.45 | 0.47 | 0.45 | 0.42 | 0.44 | 0.44 | 0.47 | 0.37587 | 2.49 | 0.39 | 0.4 | 0.43 | 0.42 | 0.5 | 0.47 | 0.4 |  |  | 0.44055
| 1 | 5 | 0.46 | 0.3539 | 0.46 | < 1 | 0.461 | 0.39537 | 0.44 | 0.47 | 0.44 | 0.42 | 0.43 | 0.49 | 0.5 | 0.3866 | 2.67 | 0.42 | 0.4 | 0.45 | 0.41 | 0.5 | 0.46 | 0.4 |  |  | 77
| 1 | 6 | 0.44 | 0.3705 | 0.45 | < 1 | 0.466 | 0.4163 | 0.42 | 0.42 | 0.43 | 0.42 | 0.45 | 0.42 | 0.48 | 0.36936 | 2.22 | 0.44 | 0.4 | 0.45 | 0.42 |  | 0.46 | 0.4
| Mean |  | 0.44333 | 0.3528 | 0.45667 |  | 0.4525 | 0.41311 | 0.44167 | 0.46167 | 0.44 | 0.41833 | 0.43667 | 0.47833 | 0.48 | 0.3756 | 2.5883 | 0.425 | 0.41667 | 0.44833 | 0.42 | 0.5 | 0.46167 | 0.4
| Median |  | 0.44 | 0.35305 | 0.455 |  | 0.4495 | 0.40695 | 0.445 | 0.47 | 0.44 | 0.42 | 0.435 | 0.475 | 0.48 | 0.37377 | 2.645 | 0.425 | 0.4 | 0.45 | 0.42 | 0.5 | 0.46 | 0.4
| Std Dev. |  | 0.013663 | 0.011022 | 0.013663 |  | 0.0091378 | 0.022017 | 0.01472 | 0.024014 | 0.0089443 | 0.0075277 | 0.012111 | 0.046655 | 0.014142 | 0.0085159 | 0.21047 | 0.02429 | 0.040825 | 0.0098319 | 0.0089443 | 0 | 0.0075277 | 6.0809e-17
| Rel.Std.Dev. |  | 3.082% | 3.124% | 2.992% |  | 2.019% | 5.33% | 3.333% | 5.202% | 2.033% | 1.799% | 2.773% | 9.754% | 2.946% | 2.267% | 8.131% | 5.715% | 9.798% | 2.193% | 2.13% | 0% | 1.631% | 1.52e-14%
| PDM3 |  | 0.6307% | -19.92% | 3.657% |  | 2.711% | -6.23% | 0.2524% | 4.792% | -0.1259% | -5.044% | -0.8826% | 8.575% | 8.954% | -14.74% | 487.5% | -3.531% | -5.422% | 1.766% | -4.666% | 13.49% | 4.792% | -9.205%
| Z-Score (Absolute) |  | 0 | 3.33 | 0.49 | 2.08 | 0.34 | 1.11 | 0.06 | 0.67 | 0.12 | 0.92 | 0.25 | 1.29 | 1.35 | 2.49 | 78.89 | 0.67 | NA | 0.18 | 0.86 | NA | 0.67 | NA
Results from laboratories 2.19, 2.23 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 163b (Certified Value 4.48 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 2
| 1 | 1 | 6.13 | 4.6399 | 5.3 | 3.72 | 5.1238 | 4.34 | 4.02 | 4.78 | 4.5 | 4.36 | 4.26 | 4.38 | 4.609 | 4.72 | 3.61 | 4.93 | 5.4 | 4.02 | 5 | [5 empty] | 1
| 1 | 2 | 6.09 | 4.8912 | 5.4 | 3.593 | 5.0925 | 4.19 | 4.01 | 4.59 | 4.46 | 4.28 | 4.19 | 4.27 | 4.699 | 4.47 | 3.68 | 4.61 | 5.6 | 4.15 | 5 | [5 empty] | 31
| 1 | 3 | 6.04 | 4.791 | 6.2 | 3.668 | 4.9928 | 4.18 | 3.89 | 4.8 | 4.43 | 4.42 | 3.92 | 4.6 | 4.5569 | 4.39 | 3.75 | 4.69 | 5.7 | 4.02 | 4 | [5 empty] | 16
| 1 | 4 | 6.19 | 4.7561 | 6.1 | 3.474 | 4.78 | 4.12 | 3.63 | 4.83 | 4.55 | 4.37 | 3.99 | 4.29 | 4.6575 | 4.49 | 3.6 | 4.85 | 5.4 | 4.15 | 4 | [5 empty] | 4.4804
| 1 | 5 | 6.48 | 4.8953 | 5.4 | 3.474 | 5.4388 | 4.24 | 4.03 | 4.68 | 4.42 | 4.19 | 3.76 | 4.46 | 4.7768 | 4.47 | 3.47 | 4.77 | 5.4 | 4.07 | 5 | [5 empty] | 78
| 1 | 6 | 6.26 | 4.9648 | 5.3 | 3.546 | 4.7827 | 4.09 | 3.88 | 4.64 | 4.49 | 4.13 | 3.92 | 4.2 | 4.7654 | 4.31 | 3.65 | 4.78 | 5.6 | 4.08 | 4
| Mean |  | 6.1983 | 4.8231 | 5.6167 | 3.5792 | 5.0351 | 4.1933 | 3.91 | 4.72 | 4.475 | 4.2917 | 4.0067 | 4.3667 | 4.6774 | 4.475 | 3.6267 | 4.7717 | 5.5167 | 4.0817 | 4.5
| Median |  | 6.16 | 4.8411 | 5.4 | 3.5695 | 5.0426 | 4.185 | 3.95 | 4.73 | 4.475 | 4.32 | 3.955 | 4.335 | 4.6783 | 4.47 | 3.63 | 4.775 | 5.5 | 4.075 | 4.5
| Std Dev. |  | 0.15791 | 0.11748 | 0.41673 | 0.10113 | 0.24685 | 0.089368 | 0.15245 | 0.096954 | 0.048477 | 0.11303 | 0.18651 | 0.14583 | 0.086825 | 0.13766 | 0.09395 | 0.11321 | 0.13292 | 0.058452 | 0.54772
| Rel.Std.Dev. |  | 2.548% | 2.436% | 7.42% | 2.826% | 4.903% | 2.131% | 3.899% | 2.054% | 1.083% | 2.634% | 4.655% | 3.34% | 1.856% | 3.076% | 2.591% | 2.373% | 2.409% | 1.432% | 12.17%
| PDM3 |  | 38.34% | 7.648% | 25.36% | -20.11% | 12.38% | -6.407% | -12.73% | 5.348% | -0.12% | -4.212% | -10.57% | -2.538% | 4.398% | -0.12% | -19.05% | 6.501% | 23.13% | -8.899% | 0.438%
| Z-Score (Absolute) |  | 3.13 | 0.63 | 2.08 | 1.63 | 1.02 | 0.51 | 1.03 | 0.45 | 0 | 0.33 | 0.85 | 0.2 | 0.37 | 0 | 1.54 | 0.54 | 1.89 | 0.72 | NA
Results from laboratory 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 163b (Certified Value 227 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 0
| 1 | 1 | 237 | 198.984 | 238.7 | 170.8 | 237.22 | 219.6033 | 221 | 222 | 245 | 219 | 228 | 212 | 223 | 223.1768 | 210 | 241.3 | 243.9 | 221 | 240 | 215 | 249 | 239.9 | 218.9 |  | 1
| 1 | 2 | 232 | 198.961 | 239 | 173.8 | 227.02 | 232.4089 | 221 | 222 | 240 | 219 | 220 | 211 | 223 | 218.2889 | 226 | 236.9 | 232.7 | 226 | 228 | 213 | 252 | 235.9 | 214.2 |  | 32
| 1 | 3 | 234 | 198.878 | 242 | 183.5 | 226.23 | 228.9405 | 218 | 224 | 241 | 219 | 221 | 214 | 221 | 222.9957 | 243 | 237.5 | 224.5 | 225 | 239 | 215 | 246 | 231.6 | 210.5 |  | 16
| 1 | 4 | 236 | 199.049 | 240 | 178.8 | 230.88 | 224.6464 | 220 | 216 | 243 | 216 | 225 | 216 | 222 | 219.0302 | 244 | 221.8 | 237.9 | 222 | 234 | 210 | 251 | 237.1 | 220.2 |  | 226.7943
| 1 | 5 | 239 | 199.507 | 236.8 | 185.6 | 229.55 | 230.785 | 222 | 223 | 245 | 220 | 225 | 218 | 218 | 221.2532 | 226 | 238.5 | 233.8 | 221 | 243 | 209 | 249 | 237.6 | 217.1 |  | 79
| 1 | 6 | 234 | 198.646 | 243 | 185.6 | 233.2 | 218.0998 | 216 | 218 | 242 | 219 | 221 | 208 | 219 | 225.2147 | 231 | 239.7 | 224.9 | 224 | 239 | 211 | 249 | 234.4 | 217
| Mean |  | 235.3333 | 199.0042 | 239.9167 | 179.6833 | 230.6833 | 225.7473 | 219.6667 | 220.8333 | 242.6667 | 218.6667 | 223.3333 | 213.1667 | 221 | 221.6599 | 230 | 235.95 | 232.95 | 223.1667 | 237.1667 | 212.1667 | 249.3333 | 236.0833 | 216.3167
| Median |  | 235 | 198.9725 | 239.5 | 181.15 | 230.215 | 226.7935 | 220.5 | 222 | 242.5 | 219 | 223 | 213 | 221.5 | 222.1245 | 228.5 | 238 | 233.25 | 223 | 239 | 212 | 249 | 236.5 | 217.05
| Std Dev. |  | 2.5033 | 0.28331 | 2.2754 | 6.307 | 4.0899 | 5.9565 | 2.2509 | 3.1252 | 2.0656 | 1.3663 | 3.1411 | 3.6009 | 2.0976 | 2.652 | 12.6333 | 7.1094 | 7.5019 | 2.137 | 5.3448 | 2.5626 | 2.0656 | 2.859 | 3.4971
| Rel.Std.Dev. |  | 1.064% | 0.1424% | 0.9484% | 3.51% | 1.773% | 2.639% | 1.025% | 1.415% | 0.8512% | 0.6248% | 1.406% | 1.689% | 0.9491% | 1.196% | 5.493% | 3.013% | 3.22% | 0.9576% | 2.254% | 1.208% | 0.8284% | 1.211% | 1.617%
| PDM3 |  | 3.765% | -12.25% | 5.786% | -20.77% | 1.715% | -0.4616% | -3.143% | -2.628% | 6.999% | -3.584% | -1.526% | -6.009% | -2.555% | -2.264% | 1.413% | 4.037% | 2.714% | -1.6% | 4.573% | -6.45% | 9.938% | 4.096% | -4.62%
| Z-Score (Absolute) |  | 0.84 | 1.71 | 1.16 | 3.06 | 0.52 | 0.17 | 0.26 | 0.18 | 1.36 | 0.33 | 0 | 0.71 | 0.16 | 0.12 | 0.47 | 0.88 | 0.67 | 0.01 | 0.97 | 0.78 | 1.82 | 0.89 | 0.49
| Analytical results for Cr in OREAS 163b (Certified Value 10.6 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 11 | 13.838 | 10.5 | 10.4 | 10.9627 | 10 | 11 | 11 | 10 | 10 | 11 | 10 | 12.7992 | 12 | 11.9 | 12 | 9 | 10 | 10.75 | 19 | 11 | 9 |  |  | 1
| 1 | 2 | 11 | 14.124 | 10.2 | 9.9 | 10.3125 | 10 | 13 | 11 | 9 | 11 | 10 | 11 | 10.151 | 13 | 12.2 | 11 | 10 | 10 | 10.3 | 12 | 10 | 9 |  |  | 33
| 1 | 3 | 11 | 13.365 | 10.5 | 10.9 | 10.6072 | 30 | 11 | 11 | 10 | 10 | 11 | 12 | 11.6522 | 12 | 11.2 | 11 | 11 | 10 | 10.4 | 16 | 11 | 10 |  |  | 16
| 1 | 4 | 11 | 13.2 | 10.2 | 10.3 | 10.2492 | 9 | 11 | 12 | 13 | 10 | 10 | 9 | 10.7587 | 12 | 11.9 | 12 | 10 | 10 | 10.45 | 12 | 11 | 9 |  |  | 10.588
| 1 | 5 | 11 | 14.829 | 10.3 | 15.3 | 9.8112 | 11 | 10 | 12 | 10 | 10 | 10 | 12 | 11.5469 | 12 | 10.5 | 11 | 9 | 10 | 10 | 11 | 11 | 9 |  |  | 80
| 1 | 6 | 10 | 14.311 | 10.3 | 10.4 | 9.7455 | 9 | 17 | 11 | 9 | 9 | 9 | 10 | 12.7377 | 14 | 12 | 10 | 9 | 11 | 11.2 | 18 | 10 | 10
| Mean |  | 10.8333 | 13.9445 | 10.3333 | 11.2 | 10.2814 | 13.1667 | 12.1667 | 11.3333 | 10.1667 | 10 | 10.1667 | 10.6667 | 11.6076 | 12.5 | 11.6167 | 11.1667 | 9.6667 | 10.1667 | 10.5167 | 14.6667 | 10.6667 | 9.3333
| Median |  | 11 | 13.981 | 10.3 | 10.4 | 10.2809 | 10 | 11 | 11 | 10 | 10 | 10 | 10.5 | 11.5996 | 12 | 11.9 | 11 | 9.5 | 10 | 10.425 | 14 | 11 | 9
| Std Dev. |  | 0.40825 | 0.60833 | 0.13663 | 2.0337 | 0.46488 | 8.2805 | 2.5626 | 0.5164 | 1.472 | 0.63246 | 0.75277 | 1.2111 | 1.0531 | 0.83666 | 0.64317 | 0.75277 | 0.8165 | 0.40825 | 0.41312 | 3.4448 | 0.5164 | 0.5164
| Rel.Std.Dev. |  | 3.768% | 4.363% | 1.322% | 18.16% | 4.522% | 62.89% | 21.06% | 4.556% | 14.48% | 6.325% | 7.404% | 11.35% | 9.072% | 6.693% | 5.537% | 6.741% | 8.447% | 4.016% | 3.928% | 23.49% | 4.841% | 5.533%
| PDM3 |  | 2.318% | 31.7% | -2.405% | 5.781% | -2.895% | 24.36% | 14.91% | 7.04% | -3.979% | -5.553% | -3.979% | 0.7435% | 9.63% | 18.06% | 9.716% | 5.466% | -8.701% | -3.979% | -0.6733% | 38.52% | 0.7435% | -11.85%
| Z-Score (Absolute) |  | 0.1 | 3.69 | 0.48 | 0.52 | 0.54 | 2.79 | 1.64 | 0.67 | 0.67 | 0.87 | 0.67 | 0.1 | 0.99 | 2.02 | 1 | 0.48 | 1.25 | 0.67 | 0.27 | 4.53 | 0.1 | 1.64
| Analytical results for Cs in OREAS 163b (Certified Value 0.78 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.79 | 0.74354 | 0.85 | 0.77081 | 0.79 | 0.83 | 0.82 | 0.75 | 0.81 | 0.76 | 0.85 | 0.69609 | 0.9 | 0.75 | 0.87 | 0.74 | 0.77 | 0.7 | [6 empty] | 1
| 1 | 2 | 0.79 | 0.76286 | 0.86 | 0.77721 | 0.78 | 0.83 | 0.77 | 0.77 | 0.76 | 0.73 | 0.82 | 0.72991 | 0.87 | 0.75 | 0.88 | 0.75 | 0.78 | 0.8 | [6 empty] | 13
| 1 | 3 | 0.81 | 0.73248 | 0.85 | 0.76717 | 0.78 | 0.86 | 0.79 | 0.74 | 0.75 | 0.7 | 0.84 | 0.78235 | 0.88 | 0.73 | 0.88 | 0.78 | 0.77 | 0.7 | [6 empty] | 16
| 1 | 4 | 0.81 | 0.77798 | 0.84 | 0.73589 | 0.78 | 0.78 | 0.8 | 0.72 | 0.76 | 0.73 | 0.79 | 0.69716 | 0.85 | 0.78 | 0.85 | 0.77 | 0.81 | 0.8 | [6 empty] | 0.7837
| 1 | 5 | 0.83 | 0.77728 | 0.83 | 0.78262 | 0.8 | 0.83 | 0.79 | 0.76 | 0.77 | 0.72 | 0.87 | 0.75869 | 0.92 | 0.73 | 0.91 | 0.76 | 0.78 | 0.8 | [6 empty] | 81
| 1 | 6 | 0.79 | 0.75712 | 0.86 | 0.74751 | 0.73 | 0.81 | 0.77 | 0.73 | 0.76 | 0.71 | 0.83 | 0.80686 | 0.85 | 0.76 | 0.85 | 0.75 | 0.78 | 0.8
| Mean |  | 0.80333 | 0.75854 | 0.84833 | 0.76353 | 0.77667 | 0.82333 | 0.79 | 0.745 | 0.76833 | 0.725 | 0.83333 | 0.74518 | 0.87833 | 0.75 | 0.87333 | 0.75833 | 0.78167 | 0.76667
| Median |  | 0.8 | 0.75999 | 0.85 | 0.76899 | 0.78 | 0.83 | 0.79 | 0.745 | 0.76 | 0.725 | 0.835 | 0.7443 | 0.875 | 0.75 | 0.875 | 0.755 | 0.78 | 0.8
| Std Dev. |  | 0.01633 | 0.018187 | 0.01169 | 0.018102 | 0.024221 | 0.026583 | 0.018974 | 0.018708 | 0.02137 | 0.020736 | 0.027325 | 0.045427 | 0.027869 | 0.018974 | 0.022509 | 0.01472 | 0.01472 | 0.05164
| Rel.Std.Dev. |  | 2.033% | 2.398% | 1.378% | 2.371% | 3.119% | 3.229% | 2.402% | 2.511% | 2.781% | 2.86% | 3.279% | 6.096% | 3.173% | 2.53% | 2.577% | 1.941% | 1.883% | 6.736%
| PDM3 |  | 2.505% | -3.21% | 8.247% | -2.574% | -0.8979% | 5.057% | 0.8034% | -4.939% | -1.961% | -7.491% | 6.333% | -4.916% | 12.07% | -4.301% | 11.44% | -3.237% | -0.2599% | -2.174%
| Z-Score (Absolute) |  | 0.67 | 0.46 | 1.81 | 0.33 | 0 | 1.18 | 0.34 | 0.8 | 0.21 | 1.31 | 1.43 | 0.8 | 2.57 | 0.67 | 2.44 | 0.46 | 0.13 | NA
Results from laboratory 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 163b (Certified Value 1.68 wt.%) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*OES | AR*OES | AR*OES | AR*AAS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*MS | AR*OES | AR*OES/MS | 2
|  |  | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 01g |  | 0.25g | 0.25g | 01g | 0.4g | 0.25g | 3
| 1 | 1 | 1.66 | 1.6688 | 1.7234 | 1.56 | 1.6937 | 1.7 | 1.665 | 1.685 | 1.675 | 1.675 | 1.71 | 1.71 | 1.66 | 1.6962 | 1.61 | 1.671 | 1.6737 | 1.659 | 1.69 | > 1 | 1.666 | > 1 | 1.66 | > 1 | 1
| 1 | 2 | 1.65 | 1.6844 | 1.7234 | 1.58 | 1.6433 | 1.6937 | 1.675 | 1.67 | 1.64 | 1.655 | 1.7 | 1.71 | 1.66 | 1.7136 | 1.66 | 1.676 | 1.6423 | 1.664 | 1.69 | > 1 | 1.661 | > 1 | 1.659 | > 1 | 35
| 1 | 3 | 1.66 | 1.6913 | 1.749 | 1.62 | 1.691 | 1.6859 | 1.68 | 1.685 | 1.685 | 1.655 | 1.715 | 1.7 | 1.66 | 1.7137 | 1.67 | 1.679 | 1.6777 | 1.646 | 1.7 | > 1 | 1.666 | > 1 | 1.666 | > 1 | 16
| 1 | 4 | 1.69 | 1.6708 | 1.7321 | 1.57 | 1.6577 | 1.6787 | 1.67 | 1.655 | 1.68 | 1.67 | 1.7 | 1.72 | 1.67 | 1.7083 | 1.66 | 1.669 | 1.6944 | 1.673 | 1.69 | > 1 | 1.663 | > 1 | 1.654 | > 1 | 1.6779
| 1 | 5 | 1.67 | 1.6972 | 1.7205 | 1.62 | 1.6881 | 1.6968 | 1.68 | 1.68 | 1.69 | 1.665 | 1.705 | 1.69 | 1.68 | 1.6949 | 1.64 | 1.684 | 1.7067 | 1.65 | 1.7 | > 1 | 1.659 | > 1 | 1.677 | > 1 | 82
| 1 | 6 | 1.65 | 1.6947 | 1.7578 | 1.62 | 1.691 | 1.6955 | 1.685 | 1.66 | 1.68 | 1.655 | 1.715 | 1.7 | 1.67 | 1.7095 | 1.69 | 1.679 | 1.642 | 1.632 | 1.69 | > 1 | 1.672 | > 1 | 1.687 | > 1
| Mean |  | 1.6633 | 1.6845 | 1.7343 | 1.595 | 1.6775 | 1.6918 | 1.6758 | 1.6725 | 1.675 | 1.6625 | 1.7075 | 1.705 | 1.6667 | 1.706 | 1.655 | 1.6763 | 1.6728 | 1.654 | 1.6933 |  | 1.6645 |  | 1.6672
| Median |  | 1.66 | 1.6879 | 1.7278 | 1.6 | 1.6896 | 1.6946 | 1.6775 | 1.675 | 1.68 | 1.66 | 1.7075 | 1.705 | 1.665 | 1.7089 | 1.66 | 1.6775 | 1.6757 | 1.6545 | 1.69 |  | 1.6645 |  | 1.663
| Std Dev. |  | 0.015055 | 0.012213 | 0.015483 | 0.028107 | 0.021464 | 0.007959 | 0.0073598 | 0.012942 | 0.017889 | 0.0088034 | 0.006892 | 0.010488 | 0.008165 | 0.0084194 | 0.027386 | 0.0055737 | 0.026531 | 0.014491 | 0.005164 |  | 0.0045935 |  | 0.012513
| Rel.Std.Dev. |  | 0.9051% | 0.725% | 0.8927% | 1.762% | 1.28% | 0.4705% | 0.4392% | 0.7738% | 1.068% | 0.5295% | 0.4036% | 0.6151% | 0.4899% | 0.4935% | 1.655% | 0.3325% | 1.586% | 0.8761% | 0.305% |  | 0.276% |  | 0.7505%
| PDM3 |  | -0.8687% | 0.3948% | 3.364% | -4.941% | -0.02584% | 0.8255% | -0.1237% | -0.3223% | -0.1734% | -0.9183% | 1.764% | 1.615% | -0.67% | 1.676% | -1.365% | -0.09389% | -0.3045% | -1.425% | 0.9193% |  | -0.7991% |  | -0.6402%
| Z-Score (Absolute) |  | 0.67 | 0.55 | 3.43 | 4.62 | 0.14 | 0.97 | 0.05 | 0.14 | 0 | 0.72 | 1.88 | 1.73 | 0.48 | 1.79 | 1.16 | 0.08 | 0.13 | 1.21 | 1.06 | NA | 0.61 | NA | 0.45 | NA
| Analytical results for Dy in OREAS 163b (Indicative Value 0.59 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 0.5104; 2.02: 0.60875; 2.03: 0.6 | 1
| 1 | 2 | 2.01: 0.5546; 2.02: 0.64105; 2.03: 0.6 | 3
| 1 | 3 | 2.01: 0.5489; 2.02: 0.6012; 2.03: 0.7 | 16
| 1 | 4 | 2.01: 0.5219; 2.02: 0.55794; 2.03: 0.7 | 0.59357
| 1 | 5 | 2.01: 0.5235; 2.02: 0.60357; 2.03: 0.7 | 9
| 1 | 6 | 2.01: 0.5525; 2.02: 0.5599; 2.03: 0.6
| Mean |  | 2.01: 0.5353; 2.02: 0.5954; 2.03: 0.65
| Median |  | 2.01: 0.5362; 2.02: 0.60239; 2.03: 0.65
| Std Dev. |  | 2.01: 0.018932; 2.02: 0.031707; 2.03: 0.054772
| Rel.Std.Dev. |  | 2.01: 3.537%; 2.02: 5.325%; 2.03: 8.427%
| PDM3 |  | 2.01: -9.817%; 2.02: 0.3092%; 2.03: 9.507%
| Z-Score (Absolute) |  | 2.01: 0.74; 2.02: 0; 2.03: 0.67
| Analytical results for Er in OREAS 163b (Indicative Value 0.21 ppm) | [25 empty] | Indicative
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 0.1903; 2.02: 0.22394; 2.03: 0.2 | 1
| 1 | 2 | 2.01: 0.2074; 2.02: 0.22948; 2.03: 0.2 | 4
| 1 | 3 | 2.01: 0.1939; 2.02: 0.22507; 2.03: 0.2 | 16
| 1 | 4 | 2.01: 0.187; 2.02: 0.21731; 2.03: 0.2 | 0.2082
| 1 | 5 | 2.01: 0.1986; 2.02: 0.24466; 2.03: 0.2 | 10
| 1 | 6 | 2.01: 0.2066; 2.02: 0.2233; 2.03: 0.2
| Mean |  | 2.01: 0.1973; 2.02: 0.22729; 2.03: 0.2
| Median |  | 2.01: 0.19625; 2.02: 0.22451; 2.03: 0.2
| Std Dev. |  | 2.01: 0.008448; 2.02: 0.0093629; 2.03: 3.0405e-17
| Rel.Std.Dev. |  | 2.01: 4.282%; 2.02: 4.119%; 2.03: 1.52e-14%
| PDM3 |  | 2.01: -5.235%; 2.02: 9.172%; 2.03: -3.938%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 6.82; 2.03: 0
| Analytical results for Eu in OREAS 163b (Indicative Value 0.1 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: < 0.1; 2.02: 0.11109; 2.03: 0.1 | 1
| 1 | 2 | 2.01: 0.1011; 2.02: 0.11754; 2.03: 0.1 | 5
| 1 | 3 | 2.01: 0.1069; 2.02: 0.11973; 2.03: 0.1 | 16
| 1 | 4 | 2.01: < 0.1; 2.02: 0.12151; 2.03: 0.1 | 0.10123
| 1 | 5 | 2.01: 0.1028; 2.02: 0.12534; 2.03: 0.1 | 11
| 1 | 6 | 2.01: 0.1012; 2.02: 0.11491; 2.03: 0.1
| Mean |  | 2.01: 0.103; 2.02: 0.11835; 2.03: 0.1
| Median |  | 2.01: 0.102; 2.02: 0.11863; 2.03: 0.1
| Std Dev. |  | 2.01: 0.0027142; 2.02: 0.0050168; 2.03: 1.5202e-17
| Rel.Std.Dev. |  | 2.01: 2.635%; 2.02: 4.239%; 2.03: 1.52e-14%
| PDM3 |  | 2.01: 1.75%; 2.02: 16.92%; 2.03: -1.214%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.84; 2.03: 0
| Analytical results for Fe in OREAS 163b (Certified Value 3.95 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 3.97 | 4.1589 | 3.68 | 4.142 | 3.8707 | 3.83 | 4.02 | 4.18 | 3.76 | 3.89 | 3.8 | 4.1 | 3.9666 | 4.24 | 3.96 | 3.9373 | 3.91 | 3.98 | 4.05 | 3.9 | 4.06 | 4.09 | 3.96 |  | 1
| 1 | 2 | 3.96 | 4.1624 | 3.74 | 4.07 | 3.8953 | 3.81 | 4.01 | 4.12 | 3.77 | 3.81 | 3.81 | 4.16 | 3.8934 | 3.94 | 3.86 | 3.8838 | 3.85 | 4.01 | 4.01 | 3.94 | 4.13 | 4.03 | 3.97 |  | 0
| 1 | 3 | 3.88 | 4.2258 | 3.84 | 4.022 | 3.8599 | 3.76 | 4.03 | 4.12 | 3.77 | 3.8 | 3.8 | 4.13 | 4.0415 | 4.31 | 3.91 | 3.9552 | 3.64 | 3.97 | 4.04 | 3.92 | 4.01 | 3.93 | 3.99 |  | 16
| 1 | 4 | 3.94 | 4.1748 | 3.72 | 3.948 | 3.8637 | 3.81 | 3.93 | 4.17 | 3.71 | 3.87 | 3.9 | 3.93 | 3.9627 | 4.31 | 3.71 | 3.9719 | 3.81 | 3.83 | 4.08 | 3.88 | 4.08 | 4.04 | 3.99 |  | 3.9546
| 1 | 5 | 3.97 | 4.1553 | 3.84 | 4.024 | 3.9079 | 3.82 | 4.01 | 4.21 | 3.79 | 3.84 | 3.88 | 4 | 4.0226 | 4 | 3.91 | 4.0406 | 3.71 | 3.86 | 4.06 | 3.87 | 4.11 | 4.06 | 3.98 |  | 83
| 1 | 6 | 3.89 | 4.1891 | 3.83 | 4.134 | 3.8496 | 3.74 | 3.93 | 4.13 | 3.76 | 3.97 | 3.68 | 4.08 | 3.9264 | 4.08 | 3.97 | 3.9024 | 3.61 | 3.93 | 4.02 | 3.94 | 4.11 | 3.97 | 3.93
| Mean |  | 3.935 | 4.1777 | 3.775 | 4.0567 | 3.8745 | 3.795 | 3.9883 | 4.155 | 3.76 | 3.8633 | 3.8117 | 4.0667 | 3.9689 | 4.1467 | 3.8867 | 3.9485 | 3.755 | 3.93 | 4.0433 | 3.9083 | 4.0833 | 4.02 | 3.97
| Median |  | 3.95 | 4.1686 | 3.785 | 4.047 | 3.8672 | 3.81 | 4.01 | 4.15 | 3.765 | 3.855 | 3.805 | 4.09 | 3.9646 | 4.16 | 3.91 | 3.9463 | 3.76 | 3.95 | 4.045 | 3.91 | 4.095 | 4.035 | 3.975
| Std Dev. |  | 0.040373 | 0.026629 | 0.070356 | 0.074185 | 0.022397 | 0.036194 | 0.045789 | 0.037283 | 0.026833 | 0.062503 | 0.077567 | 0.086178 | 0.056048 | 0.1617 | 0.095219 | 0.055698 | 0.12029 | 0.071274 | 0.02582 | 0.029944 | 0.043665 | 0.05933 | 0.022804
| Rel.Std.Dev. |  | 1.026% | 0.6374% | 1.864% | 1.829% | 0.5781% | 0.9537% | 1.148% | 0.8973% | 0.7136% | 1.618% | 2.035% | 2.119% | 1.412% | 3.9% | 2.45% | 1.411% | 3.203% | 1.814% | 0.6386% | 0.7662% | 1.069% | 1.476% | 0.5744%
| PDM3 |  | -0.4947% | 5.643% | -4.541% | 2.582% | -2.024% | -4.035% | 0.8539% | 5.068% | -4.92% | -2.307% | -3.613% | 2.835% | 0.3616% | 4.858% | -1.717% | -0.1525% | -5.046% | -0.6211% | 2.245% | -1.169% | 3.256% | 1.655% | 0.3903%
| Z-Score (Absolute) |  | 0.1 | 1.63 | 1.23 | 0.77 | 0.53 | 1.09 | 0.28 | 1.47 | 1.34 | 0.61 | 0.97 | 0.84 | 0.14 | 1.41 | 0.44 | 0 | 1.38 | 0.13 | 0.67 | 0.29 | 0.96 | 0.51 | 0.15
| Analytical results for Ga in OREAS 163b (Certified Value 0.63 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.8 | 0.65638 | 0.8 | < 5 | 0.6 | 0.63818 | 0.61 | 0.7 | 0.67 | 0.6 | 0.64 | 0.49 | 0.6 | 0.6348 | < 10 | 0.7 | < 1 | 0.6 | 0.7 | 0.69 | 0.67 | < 1 |  |  | 1
| 1 | 2 | 0.8 | 0.66163 | 0.9 | < 5 | 0.6 | 0.6774 | 0.61 | 0.67 | 0.64 | 0.62 | 0.58 | 0.55 | 0.6 | 0.54818 | < 10 | 0.7 | < 1 | 0.65 | 0.6 | 0.63 | 0.67 | < 1 |  |  | 17
| 1 | 3 | 0.8 | 0.67163 | 0.9 | < 5 | 0.6 | 0.6816 | 0.55 | 0.67 | 0.65 | 0.6 | 0.6 | 0.54 | 0.6 | 0.6532 | < 10 | 0.7 | < 1 | 0.64 | 0.7 | 0.84 | 0.67 | < 1 |  |  | 16
| 1 | 4 | 0.8 | 0.66275 | 0.8 | < 5 | 0.6 | 0.60045 | 0.61 | 0.63 | 0.67 | 0.59 | 0.62 | 0.53 | 0.6 | 0.62092 | < 10 | 0.7 | < 1 | 0.58 | 0.7 | 0.63 | 0.68 | < 1 |  |  | 0.62648
| 1 | 5 | 1.6 | 0.66375 | 0.8 | < 5 | 0.6 | 0.66008 | 0.61 | 0.66 | 0.67 | 0.61 | 0.63 | 0.53 | 0.6 | 0.63038 | < 10 | 0.7 | < 1 | 0.59 | 0.7 | 0.6 | 0.66 | < 1 |  |  | 84
| 1 | 6 | 0.8 | 0.65087 | 0.9 | < 5 | 0.6 | 0.62886 | 0.57 | 0.69 | 0.66 | 0.6 | 0.59 | 0.52 | 0.6 | 0.52922 | < 10 | 0.7 | < 1 | 0.61 | 0.7 | 0.72 | 0.66 | < 1
| Mean |  | 0.93333 | 0.66117 | 0.85 |  | 0.6 | 0.64776 | 0.59333 | 0.67 | 0.66 | 0.60333 | 0.61 | 0.52667 | 0.6 | 0.60278 |  | 0.7 |  | 0.61167 | 0.68333 | 0.685 | 0.66833
| Median |  | 0.8 | 0.66219 | 0.85 |  | 0.6 | 0.64913 | 0.61 | 0.67 | 0.665 | 0.6 | 0.61 | 0.53 | 0.6 | 0.62565 |  | 0.7 |  | 0.605 | 0.7 | 0.66 | 0.67
| Std Dev. |  | 0.3266 | 0.0070377 | 0.054772 |  | 0 | 0.031186 | 0.026583 | 0.024495 | 0.012649 | 0.010328 | 0.023664 | 0.020656 | 0 | 0.051089 |  | 1.2162e-16 |  | 0.027869 | 0.040825 | 0.087807 | 0.0075277
| Rel.Std.Dev. |  | 34.99% | 1.064% | 6.444% |  | 0% | 4.814% | 4.48% | 3.656% | 1.917% | 1.712% | 3.879% | 3.922% | 0% | 8.476% |  | 1.737e-14% |  | 4.556% | 5.974% | 12.82% | 1.126%
| PDM3 |  | 48.98% | 5.537% | 35.68% |  | -4.226% | 3.398% | -5.29% | 6.947% | 5.351% | -3.694% | -2.63% | -15.93% | -4.226% | -3.782% |  | 11.74% |  | -2.364% | 9.076% | 9.342% | 6.681%
| Z-Score (Absolute) |  | NA | 0.57 | NA | 33.63 | NA | 0.32 | 0.65 | 0.72 | 0.54 | 0.47 | 0.35 | 1.85 | NA | 0.48 | 78.58 | NA | 2.33 | 0.32 | NA | 0.99 | 0.69 | 2.33
Results from laboratories 2.01, 2.05, 2.13, 2.17 and 2.21 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.03 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 163b (Indicative Value 0.86 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  |  | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g |  |  | 2
| 1 | 1 | 2.01: 0.8218; 2.02: 0.94768; 2.03: 0.9 |  |  | 1
| 1 | 2 | 2.01: 0.8857; 2.02: 1.0376; 2.03: 0.8 |  |  | 6
| 1 | 3 | 2.01: 0.8603; 2.02: 0.93909; 2.03: 0.9 |  |  | 16
| 1 | 4 | 2.01: 0.8542; 2.02: 0.86734; 2.03: 0.9 |  |  | 0.8643
| 1 | 5 | 2.01: 0.8809; 2.02: 1.03; 2.03: 0.9 |  |  | 12
| 1 | 6 | 2.01: 0.8687; 2.02: 0.86344; 2.03: 0.8
| Mean |  | 2.01: 0.86193; 2.02: 0.94752; 2.03: 0.86667
| Median |  | 2.01: 0.8645; 2.02: 0.94339; 2.03: 0.9
| Std Dev. |  | 2.01: 0.022994; 2.02: 0.075477; 2.03: 0.05164
| Rel.Std.Dev. |  | 2.01: 2.668%; 2.02: 7.966%; 2.03: 5.958%
| PDM3 |  | 2.01: -0.2738%; 2.02: 9.629%; 2.03: 0.2738%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 11.52; 2.03: 0
| Analytical results for Ge in OREAS 163b (Certified Value 0.041 ppm) | [25 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [9 empty] | 3
| 1 | 1 | 1.0841 | < 5 | < 0.2 | 0.05 | 0.06 | 0.05 | 0.06 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.5 | [9 empty] | 1
| 1 | 2 | 1.0776 | < 5 | < 0.2 | 0.05 | < 0.05 | 0.05 | 0.06 | 0.05 | < 0.1 | < 0.1 | 0.06 | < 0.05 | < 0.1 | < 0.05 | 0.6 | [9 empty] | 19
| 1 | 3 | 1.0828 | < 5 | < 0.2 | 0.05 | < 0.05 | 0.05 | 0.05 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.5 | [9 empty] | 16
| 1 | 4 | 1.1973 | < 5 | < 0.2 | < 0.05 | < 0.05 | 0.05 | 0.05 | < 0.05 | < 0.1 | < 0.1 | 0.05 | 0.05 | < 0.1 | < 0.05 | 0.5 | [9 empty] | 0.040952
| 1 | 5 | 1.1638 | < 5 | < 0.2 | 0.05 | 0.05 | 0.05 | 0.05 | < 0.05 | < 0.1 | < 0.1 | 0.07 | < 0.05 | < 0.1 | < 0.05 | 0.5 | [9 empty] | 85
| 1 | 6 | 1.1087 | < 5 | < 0.2 | 0.05 | 0.06 | 0.05 | 0.05 | < 0.05 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | 0.5
| Mean |  | 2.01: 1.1191; 2.04: 0.05; 2.05: 0.056667; 2.06: 0.05; 2.07: 0.053333; 2.08: 0.05; 2.11: 0.06; 2.12: 0.05; 2.15: 0.51667
| Median |  | 2.01: 1.0964; 2.04: 0.05; 2.05: 0.06; 2.06: 0.05; 2.07: 0.05; 2.08: 0.05; 2.11: 0.06; 2.12: 0.05; 2.15: 0.5
| Std Dev. |  | 2.01: 0.049973; 2.04: 0; 2.05: 0.0057735; 2.06: 7.6012e-18; 2.07: 0.005164; 2.11: 0.01; 2.15: 0.040825
| Rel.Std.Dev. |  | 2.01: 4.466%; 2.04: 0%; 2.05: 10.19%; 2.06: 1.52e-14%; 2.07: 9.682%; 2.11: 16.67%; 2.15: 7.902%
| PDM3 |  | 2.01: 2633%; 2.04: 22.09%; 2.05: 38.37%; 2.06: 22.09%; 2.07: 30.23%; 2.08: 22.09%; 2.11: 46.51%; 2.12: 22.09%; 2.15: 1162%
| Z-Score (Absolute) |  | 78.64 | 180.22 | 3.68 | 0.31 | 0.67 | 0 | 0.25 | 1.53 | 0 | 0 | 0.55 | 1.53 | 0 | 1.84 | 34.33
| Analytical results for Hf in OREAS 163b (Certified Value 0.049 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.05 | 0.06 | < 0.1 | 0.03 | 0.03 | 0.05 | 0.05 | 0.05 | < 0.1 | 0.05 | < 0.05 | 0.08 | 0.11 | 0.05 | < 0.08 | 0.06 | < 0.1 | [7 empty] | 1
| 1 | 2 | 0.05 | 0.05 | < 0.1 | 0.04 | 0.03 | 0.05 | 0.05 | 0.05 | < 0.1 | 0.05 | < 0.05 | 0.07 | 0.11 | 0.05 | < 0.08 | 0.05 | < 0.1 | [7 empty] | 23
| 1 | 3 | 0.05 | 0.05 | < 0.1 | 0.03 | 0.04 | 0.05 | 0.04 | 0.04 | < 0.1 | 0.05 | < 0.05 | 0.06 | 0.11 | < 0.05 | < 0.08 | 0.05 | < 0.1 | [7 empty] | 16
| 1 | 4 | 0.05 | 0.05 | < 0.1 | 0.03 | 0.04 | 0.05 | 0.04 | 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.08 | 0.11 | 0.05 | < 0.08 | 0.06 | < 0.1 | [7 empty] | 0.04899
| 1 | 5 | 0.05 | 0.05 | < 0.1 | 0.03 | 0.05 | 0.05 | 0.05 | 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.07 | 0.11 | 0.05 | < 0.08 | 0.06 | < 0.1 | [7 empty] | 86
| 1 | 6 | 0.05 | 0.04 | < 0.1 | 0.03 | 0.04 | 0.05 | 0.04 | 0.04 | < 0.1 | 0.05 | 0.053232 | 0.05 | 0.1 | 0.05 | < 0.08 | 0.06 | < 0.1
| Mean |  | 0.05 | 0.05 |  | 0.031667 | 0.038333 | 0.05 | 0.045 | 0.046667 |  | 0.05 | 0.053232 | 0.068333 | 0.10833 | 0.05 |  | 0.056667
| Median |  | 0.05 | 0.05 |  | 0.03 | 0.04 | 0.05 | 0.045 | 0.05 |  | 0.05 | 0.053232 | 0.07 | 0.11 | 0.05 |  | 0.06
| Std Dev. |  | 7.6012e-18 | 0.0063246 |  | 0.0040825 | 0.0075277 | 7.6012e-18 | 0.0054772 | 0.005164 |  | 0 |  | 0.01169 | 0.0040825 | 0 |  | 0.005164
| Rel.Std.Dev. |  | 1.52e-14% | 12.65% |  | 12.89% | 19.64% | 1.52e-14% | 12.17% | 11.07% |  | 0% |  | 17.11% | 3.768% | 0% |  | 9.113%
| PDM3 |  | 2.062% | 2.062% |  | -35.36% | -21.75% | 2.062% | -8.144% | -4.742% |  | 2.062% | 8.66% | 39.48% | 121.1% | 2.062% |  | 15.67%
| Z-Score (Absolute) |  | 0 | 0 | 0 | 2.47 | 1.57 | 0 | 0.67 | 0.45 | 0 | 1.12 | 2.74 | 2.47 | 7.87 | 0.56 | 1.35 | 0.9 | 0
| Analytical results for Hg in OREAS 163b (Certified Value 0.079 ppm) | [25 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.122 | < 5 | 0.09 | < 1 | 0.09 | 0.06 | 0.08 | 0.07 | 0.08 | 0.08 | 0.09 | 0.077354 | 0.04 | 0.053 | 0.08 | 0.08 | 0.08 | < 0.1 | 0.083 | 0.327 |  |  |  |  | 1
| 1 | 2 | 0.117 | < 5 | 0.09 | < 1 | 0.09 | 0.1 | 0.07 | 0.08 | 0.06 | 0.09 | 0.1 | 0.064275 | 0.04 | 0.026 | 0.08 | 0.08 | 0.07 | < 0.1 | 0.084 | 0.34 |  |  |  |  | 1
| 1 | 3 | 0.12 | < 5 | 0.1 | < 1 | 0.08 | 0.11 | 0.08 | 0.08 | 0.08 | 0.09 | 0.09 | 0.070135 | 0.03 | 0.064 | 0.08 | 0.08 | 0.08 | < 0.1 | 0.08 | 0.294 |  |  |  |  | 16
| 1 | 4 | 0.12 | < 5 | 0.08 | < 1 | 0.09 | 0.09 | 0.07 | 0.08 | 0.07 | 0.08 | 0.08 | 0.069155 | 0.03 | 0.06 | 0.08 | 0.09 | 0.08 | < 0.1 | 0.088 | 0.31 |  |  |  |  | 0.079382
| 1 | 5 | 0.12 | < 5 | 0.08 | < 1 | 0.09 | 0.09 | 0.06 | 0.07 | 0.08 | 0.07 | 0.1 | 0.069676 | 0.03 | 0.051 | 0.09 | 0.08 | 0.08 | < 0.1 | 0.086 | 0.336 |  |  |  |  | 87
| 1 | 6 | 0.12 | < 5 | 0.09 | < 1 | 0.08 | 0.1 | 0.08 | 0.07 | 0.07 | 0.07 | 0.09 | 0.076699 | 0.03 | 0.056 | 0.02 | 0.07 | 0.07 | < 0.1 | 0.087 | 0.322
| Mean |  | 0.11983 |  | 0.088333 |  | 0.086667 | 0.091667 | 0.073333 | 0.075 | 0.073333 | 0.08 | 0.091667 | 0.071216 | 0.033333 | 0.051667 | 0.071667 | 0.08 | 0.076667 |  | 0.084667 | 0.3215
| Median |  | 0.12 |  | 0.09 |  | 0.09 | 0.095 | 0.075 | 0.075 | 0.075 | 0.08 | 0.09 | 0.069905 | 0.03 | 0.0545 | 0.08 | 0.08 | 0.08 |  | 0.085 | 0.3245
| Std Dev. |  | 0.0016021 |  | 0.0075277 |  | 0.005164 | 0.017224 | 0.008165 | 0.0054772 | 0.008165 | 0.0089443 | 0.0075277 | 0.0049739 | 0.005164 | 0.013426 | 0.025626 | 0.0063246 | 0.005164 |  | 0.0029439 | 0.017155
| Rel.Std.Dev. |  | 1.337% |  | 8.522% |  | 5.958% | 18.79% | 11.13% | 7.303% | 11.13% | 11.18% | 8.212% | 6.984% | 15.49% | 25.99% | 35.76% | 7.906% | 6.736% |  | 3.477% | 5.336%
| PDM3 |  | 50.96% |  | 11.28% |  | 9.177% | 15.48% | -7.62% | -5.52% | -7.62% | 0.7784% | 15.48% | -10.29% | -58.01% | -34.91% | -9.719% | 0.7784% | -3.421% |  | 6.657% | 305%
| Z-Score (Absolute) |  | 3.14 | 190.66 | 0.66 | 33.09 | 0.53 | 0.92 | 0.53 | 0.39 | 0.53 | 0 | 0.92 | 0.69 | 3.68 | 2.23 | 0.66 | 0 | 0.26 | 2.36 | 0.37 | 19.03
| Analytical results for Ho in OREAS 163b (Indicative Value 0.072 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | 2.03: Lab; 2.04: Lab; 2.05: Lab |  |  |  |  | 1
| No | No | 2.03: 2.02; 2.04: 2.06; 2.05: 2.12 |  |  |  |  | ppm
|  |  | 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*OES/MS |  |  |  |  | 3
|  |  | 2.03: 0.5g; 2.04: 0.2g; 2.05: 0.5g |  |  |  |  | 3
| 1 | 1 | 2.03: 0.10049; 2.04: 0.098065; 2.05: < 0.1 |  |  |  |  | 1
| 1 | 2 | 2.03: 0.1027; 2.04: 0.096823; 2.05: < 0.1 |  |  |  |  | 3
| 1 | 3 | 2.03: < 0.1; 2.04: 0.095794; 2.05: 0.1 |  |  |  |  | 16
| 1 | 4 | 2.03: < 0.1; 2.04: < 0.1; 2.05: < 0.1 |  |  |  |  | 0.071951
| 1 | 5 | 2.03: 0.10049; 2.04: 0.10075; 2.05: < 0.1 |  |  |  |  | 9
| 1 | 6 | 2.03: < 0.1; 2.04: < 0.1; 2.05: < 0.1
| Mean |  | 2.03: 0.10123; 2.04: 0.097858; 2.05: 0.1
| Median |  | 2.03: 0.10049; 2.04: 0.097444; 2.05: 0.1
| Std Dev. |  | 2.03: 0.0012759; 2.04: 0.0021401
| Rel.Std.Dev. |  | 2.03: 1.26%; 2.04: 2.187%
| PDM3 |  | 2.03: 40.69%; 2.04: 36.01%; 2.05: 38.98%
| Z-Score (Absolute) |  | 2.03: 0; 2.04: 0.67; 2.05: 1.85
| Analytical results for In in OREAS 163b (Certified Value 1.02 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 1.05 | 0.94934 | < 5 | 1.06 | 0.99335 | 1.04 | 1.105 | 1.145 | 1.03 | 1.06 | 1.03 | 1.094 | 1.0259 | 0.064 | 0.98 | 1.03 | 0.96 | 1.034 | 1.02 | [5 empty] | 1
| 1 | 2 | 1.04 | 0.98672 | < 5 | 1.02 | 0.97425 | 1.035 | 1.115 | 1.085 | 1.03 | 1.025 | 1 | 1.08 | 0.92778 | 0.063 | 1.04 | 0.99 | 0.96 | 1.059 | 1.02 | [5 empty] | 10
| 1 | 3 | 1.04 | 0.93282 | < 5 | 1.04 | 0.95604 | 0.992 | 1.105 | 1.11 | 1.02 | 1.045 | 1.01 | 1.054 | 1.0199 | 0.065 | 0.99 | 0.99 | 1.05 | 1.035 | 1.02 | [5 empty] | 16
| 1 | 4 | 1.05 | 0.9289 | < 5 | 1.01 | 0.96773 | 1.035 | 1.05 | 1.115 | 1.035 | 1.04 | 1.07 | 1.067 | 0.92424 | 0.062 | 1.02 | 1 | 1 | 1.105 | 1.02 | [5 empty] | 1.0208
| 1 | 5 | 1.04 | 0.95718 | < 5 | 1.05 | 0.99599 | 1.04 | 1.165 | 1.13 | 1.035 | 1.045 | 1.03 | 1.096 | 0.97109 | 0.072 | 1.03 | 1.02 | 0.9 | 1.042 | 1 | [5 empty] | 88
| 1 | 6 | 1.07 | 0.96558 | < 5 | 1.05 | 0.94781 | 0.975 | 1.065 | 1.1 | 1.04 | 0.992 | 1.02 | 1.072 | 1.0113 | 0.059 | 1.01 | 0.98 | 0.97 | 1.045 | 1.01
| Mean |  | 1.0483 | 0.95342 |  | 1.0383 | 0.97253 | 1.0195 | 1.1008 | 1.1142 | 1.0317 | 1.0345 | 1.0267 | 1.0772 | 0.98004 | 0.064167 | 1.0117 | 1.0017 | 0.97333 | 1.0533 | 1.015
| Median |  | 1.045 | 0.95326 |  | 1.045 | 0.97099 | 1.035 | 1.105 | 1.1125 | 1.0325 | 1.0425 | 1.025 | 1.076 | 0.99119 | 0.0635 | 1.015 | 0.995 | 0.965 | 1.0435 | 1.02
| Std Dev. |  | 0.01169 | 0.021509 |  | 0.019408 | 0.019455 | 0.028487 | 0.040548 | 0.021311 | 0.0068313 | 0.023654 | 0.024221 | 0.016204 | 0.046035 | 0.0043551 | 0.023166 | 0.019408 | 0.049666 | 0.026868 | 0.0083666
| Rel.Std.Dev. |  | 1.115% | 2.256% |  | 1.869% | 2.001% | 2.794% | 3.683% | 1.913% | 0.6622% | 2.286% | 2.359% | 1.504% | 4.697% | 6.787% | 2.29% | 1.938% | 5.103% | 2.551% | 0.8243%
| PDM3 |  | 2.701% | -6.597% |  | 1.721% | -4.725% | -0.1238% | 7.844% | 9.15% | 1.068% | 1.346% | 0.5783% | 5.526% | -3.99% | -93.71% | -0.8912% | -1.871% | -4.647% | 3.191% | -0.5647%
| Z-Score (Absolute) |  | 0.55 | 1.85 | 37.26 | 0.3 | 1.37 | 0.18 | 1.88 | 2.21 | 0.13 | 0.2 | 0 | 1.28 | 1.18 | 24.34 | 0.38 | 0.63 | 1.35 | 0.67 | 0.3
| Analytical results for K in OREAS 163b (Certified Value 0.015 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.014 | 0.02115 | 0.01 | 0.014 | 0.013655 | 0.02 | 0.02 | 0.02 | 0.01 | 0.01 | 0.02 | 0.012 | 0.021435 | 0.03 | 0.02 | 0.02 | 0.01 | < 0.01 | 0.014 | 0.0223 | 0.01 | 0.02 |  |  | 1
| 1 | 2 | 0.015 | 0.0212 | 0.01 | 0.014 | 0.014379 | 0.01 | 0.02 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.020557 | 0.01 | 0.02 | 0.02 | 0.01 | < 0.01 | 0.013 | 0.0233 | 0.01 | 0.02 |  |  | 0
| 1 | 3 | 0.014 | 0.0212 | 0.02 | 0.014 | 0.01356 | 0.01 | 0.02 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.023482 | 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | 0.013 | 0.0225 | 0.01 | 0.02 |  |  | 16
| 1 | 4 | 0.015 | 0.02114 | 0.01 | 0.013 | 0.013283 | 0.02 | 0.01 | 0.02 | 0.01 | 0.01 | 0.02 | 0.012 | 0.020217 | 0.01 | 0.01 | 0.02 | 0.01 | < 0.01 | 0.011 | 0.0219 | 0.01 | 0.02 |  |  | 0.014803
| 1 | 5 | 0.014 | 0.02119 | 0.01 | 0.013 | 0.014499 | 0.02 | 0.02 | 0.02 | 0.01 | 0.01 | 0.02 | 0.013 | 0.021971 | 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | 0.012 | 0.0218 | 0.01 | 0.02 |  |  | 89
| 1 | 6 | 0.014 | 0.02124 | 0.01 | 0.014 | 0.013162 | 0.01 | 0.01 | 0.02 | 0.01 | 0.02 | 0.02 | 0.012 | 0.023069 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.0224 | 0.01 | 0.02
| Mean |  | 0.014333 | 0.021187 | 0.011667 | 0.013667 | 0.013756 | 0.015 | 0.016667 | 0.02 | 0.01 | 0.011667 | 0.02 | 0.0125 | 0.021789 | 0.013333 | 0.016667 | 0.015 | 0.01 |  | 0.012167 | 0.022367 | 0.01 | 0.02
| Median |  | 0.014 | 0.021195 | 0.01 | 0.014 | 0.013607 | 0.015 | 0.02 | 0.02 | 0.01 | 0.01 | 0.02 | 0.0125 | 0.021703 | 0.01 | 0.02 | 0.015 | 0.01 |  | 0.0125 | 0.02235 | 0.01 | 0.02
| Std Dev. |  | 0.0005164 | 3.6697e-05 | 0.0040825 | 0.0005164 | 0.00055967 | 0.0054772 | 0.005164 | 0 | 0 | 0.0040825 | 0 | 0.00054772 | 0.0013156 | 0.008165 | 0.005164 | 0.0054772 | 0 |  | 0.001472 | 0.00053541 | 0 | 0
| Rel.Std.Dev. |  | 3.603% | 0.1732% | 34.99% | 3.779% | 4.068% | 36.51% | 30.98% | 0% | 0% | 34.99% | 0% | 4.382% | 6.038% | 61.24% | 30.98% | 36.51% | 0% |  | 12.1% | 2.394% | 0% | 0%
| PDM3 |  | -3.174% | 43.12% | -21.19% | -7.678% | -7.072% | 1.329% | 12.59% | 35.11% | -32.45% | -21.19% | 35.11% | -15.56% | 47.19% | -9.93% | 12.59% | 1.329% | -32.45% |  | -17.81% | 51.09% | -32.45% | 35.11%
| Z-Score (Absolute) |  | 0.07 | 1.84 | 0.61 | 0.1 | 0.07 | 0.25 | 0.67 | 1.53 | 1.04 | 0.61 | 1.53 | 0.4 | 1.99 | 0.18 | 0.67 | 0.25 | 1.04 | 2.33 | 0.48 | 2.14 | 1.04 | 1.53
| Analytical results for La in OREAS 163b (Certified Value 1.83 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 2.73 | 2.5875 | 2.8 | 2.2 | 1.591 | 2.7501 | 1.8 | 1.8 | 2 | 1.8 | 1.9 | 2 | 1.8 | 2.3786 | 1.9 | 1.7 | 2 | 1.7 | 1.8 | 2.5 | 1.8 | 2 |  |  | 1
| 1 | 2 | 2.78 | 2.5602 | 2.8 | 2.1 | 1.509 | 2.5568 | 1.7 | 1.7 | 1.9 | 1.8 | 1.9 | 2 | 1.8 | 2.4045 | 1.8 | 1.6 | 2 | 1.7 | 1.7 | 2.7 | 1.8 | 2 |  |  | 11
| 1 | 3 | 2.75 | 2.5372 | 2.8 | 2.3 | 1.585 | 2.7662 | 1.7 | 1.7 | 2 | 1.8 | 1.9 | 1.9 | 1.8 | 2.2128 | 1.9 | 1.7 | 1 | 1.7 | 1.7 | 2.7 | 1.8 | 2 |  |  | 16
| 1 | 4 | 2.68 | 2.5331 | 2.8 | 2.2 | 1.483 | 2.6347 | 1.7 | 1.6 | 2 | 1.9 | 1.9 | 1.9 | 1.9 | 2.1824 | 1.8 | 1.6 | 2 | 1.6 | 1.8 | 2.5 | 1.8 | 2 |  |  | 1.826
| 1 | 5 | 2.63 | 2.5895 | 2.8 | 2.5 | 1.471 | 2.8068 | 1.8 | 1.8 | 1.9 | 1.8 | 1.8 | 1.7 | 1.8 | 2.2673 | 1.7 | 1.6 | 2 | 1.6 | 1.7 | 2.5 | 1.8 | 2 |  |  | 90
| 1 | 6 | 2.59 | 2.6504 | 2.8 | 2.1 | 1.534 | 2.6285 | 1.7 | 1.7 | 1.9 | 1.8 | 1.9 | 1.8 | 1.8 | 2.1391 | 1.7 | 1.6 | 1 | 1.6 | 1.7 | 2.6 | 1.8 | 2
| Mean |  | 2.6933 | 2.5763 | 2.8 | 2.2333 | 1.5288 | 2.6905 | 1.7333 | 1.7167 | 1.95 | 1.8167 | 1.8833 | 1.8833 | 1.8167 | 2.2641 | 1.8 | 1.6333 | 1.6667 | 1.65 | 1.7333 | 2.5833 | 1.8 | 2
| Median |  | 2.705 | 2.5739 | 2.8 | 2.2 | 1.5215 | 2.6924 | 1.7 | 1.7 | 1.95 | 1.8 | 1.9 | 1.9 | 1.8 | 2.24 | 1.8 | 1.6 | 2 | 1.65 | 1.7 | 2.55 | 1.8 | 2
| Std Dev. |  | 0.073394 | 0.043487 | 4.8648e-16 | 0.15055 | 0.050764 | 0.097607 | 0.05164 | 0.075277 | 0.054772 | 0.040825 | 0.040825 | 0.1169 | 0.040825 | 0.10749 | 0.089443 | 0.05164 | 0.5164 | 0.054772 | 0.05164 | 0.098319 | 0 | 0
| Rel.Std.Dev. |  | 2.725% | 1.688% | 1.737e-14% | 6.741% | 3.32% | 3.628% | 2.979% | 4.385% | 2.809% | 2.247% | 2.168% | 6.207% | 2.247% | 4.747% | 4.969% | 3.162% | 30.98% | 3.32% | 2.979% | 3.806% | 0% | 0%
| PDM3 |  | 47.5% | 41.09% | 53.34% | 22.31% | -16.27% | 47.35% | -5.074% | -5.986% | 6.792% | -0.5097% | 3.141% | 3.141% | -0.5097% | 24% | -1.423% | -10.55% | -8.725% | -9.637% | -5.074% | 41.48% | -1.423% | 9.531%
| Z-Score (Absolute) |  | 3.41 | 2.94 | 3.84 | 1.55 | 1.3 | 3.4 | 0.47 | 0.54 | 0.4 | 0.13 | 0.13 | 0.13 | 0.13 | 1.68 | 0.2 | 0.88 | NA | 0.81 | 0.47 | 2.97 | 0.2 | NA
Results from laboratories 2.19 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 163b (Certified Value 4.66 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 5.4 | < 5 | 3.57 | < 10 | 4.6 | 4.9 | 5.5 | 4.8 | 4.8 | 4.9 | 4.3 | 4.2973 | 4.2 | 4.6 | 4 | 4.9 | 4.3 | 4.3 | [6 empty] | 1
| 1 | 2 | 5.2 | < 5 | 3.49 | < 10 | 4.8 | 5 | 5.3 | 4.7 | 4.5 | 4.7 | 4.3 | 4.162 | 4.4 | 4.9 | 4 | 4.9 | 4.2 | 4.8 | [6 empty] | 0
| 1 | 3 | 5.3 | < 5 | 3.63 | < 10 | 4.5 | 5 | 5.4 | 4.8 | 4.4 | 4.4 | 4.6 | 4.7565 | 3.7 | 4.4 | 4 | 4.9 | 4.1 | 4.4 | [6 empty] | 16
| 1 | 4 | 5.2 | < 5 | 3.52 | < 10 | 4.6 | 4.8 | 5.5 | 4.6 | 4.7 | 4.7 | 4.4 | 3.9643 | 3.5 | 4.3 | 4 | 4.8 | 4.5 | 4.5 | [6 empty] | 4.6573
| 1 | 5 | 5.3 | < 5 | 3.61 | < 10 | 4.7 | 5 | 5.5 | 4.8 | 4.6 | 4.5 | 4.5 | 4.3976 | 4.4 | 4.3 | 4 | 4.9 | 4.4 | 4.7 | [6 empty] | 91
| 1 | 6 | 5.1 | < 5 | 3.54 | < 10 | 4.5 | 4.7 | 5.4 | 4.6 | 4.7 | 4.6 | 4.3 | 4.5588 | 4.5 | 4.4 | 4 | 4.6 | 4.3 | 4.8
| Mean |  | 5.25 |  | 3.56 |  | 4.6167 | 4.9 | 5.4333 | 4.7167 | 4.6167 | 4.6333 | 4.4 | 4.3561 | 4.1167 | 4.4833 | 4 | 4.8333 | 4.3 | 4.5833
| Median |  | 5.25 |  | 3.555 |  | 4.6 | 4.95 | 5.45 | 4.75 | 4.65 | 4.65 | 4.35 | 4.3474 | 4.3 | 4.4 | 4 | 4.9 | 4.3 | 4.6
| Std Dev. |  | 0.10488 |  | 0.053666 |  | 0.1169 | 0.12649 | 0.08165 | 0.098319 | 0.1472 | 0.17512 | 0.12649 | 0.28204 | 0.41673 | 0.23166 | 0 | 0.12111 | 0.14142 | 0.2137
| Rel.Std.Dev. |  | 1.998% |  | 1.507% |  | 2.532% | 2.581% | 1.503% | 2.085% | 3.188% | 3.78% | 2.875% | 6.475% | 10.12% | 5.167% | 0% | 2.506% | 3.289% | 4.662%
| PDM3 |  | 12.73% |  | -23.56% |  | -0.8733% | 5.21% | 16.66% | 1.274% | -0.8733% | -0.5154% | -5.525% | -6.469% | -11.61% | -3.736% | -14.11% | 3.779% | -7.673% | -1.589%
| Z-Score (Absolute) |  | 1.64 | 5.48 | 2.73 | 0.99 | 0 | 0.73 | 2.11 | 0.26 | 0 | 0.04 | 0.56 | 0.67 | 1.29 | 0.35 | NA | 0.56 | 0.82 | 0.09
Results from laboratory 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 163b (Indicative Value 0.02 ppm) | [25 empty] | Indicative
Lu | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | [6 empty] | 1
| No | No | 2.01: 2.02; 2.04: 2.04; 2.05: 2.06; 2.06: 2.12; 2.07: 2.13; 2.08: 2.21 | [6 empty] | ppm
|  |  | 2.01: AR*MS; 2.04: AR*OES; 2.05: AR*MS; 2.06: AR*OES/MS; 2.07: AR*MS; 2.08: AR*OES/MS | [6 empty] | 3
|  |  | 2.01: 0.5g; 2.04: 0.25g; 2.05: 0.2g; 2.06: 0.5g; 2.07: 0.25g; 2.08: 0.25g | [6 empty] | 3
| 1 | 1 | 2.01: < 0.05; 2.04: < 5; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.02; 2.08: 0.02 | [6 empty] | 1
| 1 | 2 | 2.01: < 0.05; 2.04: < 5; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.02; 2.08: 0.02 | [6 empty] | 4
| 1 | 3 | 2.01: < 0.05; 2.04: < 5; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.02; 2.08: 0.02 | [6 empty] | 16
| 1 | 4 | 2.01: < 0.05; 2.04: < 5; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.02; 2.08: 0.02 | [6 empty] | 0.02
| 1 | 5 | 2.01: < 0.05; 2.04: < 5; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.02; 2.08: 0.02 | [6 empty] | 10
| 1 | 6 | 2.01: < 0.05; 2.04: < 5; 2.05: < 0.05; 2.06: < 0.1; 2.07: 0.02; 2.08: 0.02
| Mean |  | 2.07: 0.02; 2.08: 0.02
| Median |  | 2.07: 0.02; 2.08: 0.02
| Std Dev. |  | 2.07: 0; 2.08: 0
Rel.Std.Dev. | bar: 2.07 to 2.08
PDM3 | bar: 2.07 to 2.08
| Z-Score (Absolute) |  | 2.01: 0; 2.04: 333.78; 2.05: 0; 2.06: 3.37; 2.07: 0.67; 2.08: 0.67
| Analytical results for Mg in OREAS 163b (Certified Value 1.41 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.42 | 1.3745 | 1.33 | 1.529 | 1.395 | 1.37 | 1.43 | 1.52 | 1.34 | 1.39 | 1.2 | 1.513 | 1.4439 | 1.24 | 1.49 | 1.44 | 1.41 | 1.52 | 1.383 | 1.43 | 1.41 | 1.43 |  |  | 1
| 1 | 2 | 1.41 | 1.338 | 1.35 | 1.361 | 1.4169 | 1.37 | 1.42 | 1.5 | 1.35 | 1.35 | 1.26 | 1.531 | 1.3901 | 1.33 | 1.47 | 1.42 | 1.46 | 1.53 | 1.399 | 1.46 | 1.39 | 1.43 |  |  | 0
| 1 | 3 | 1.41 | 1.356 | 1.39 | 1.417 | 1.3717 | 1.33 | 1.44 | 1.5 | 1.35 | 1.35 | 1.27 | 1.528 | 1.4014 | 1.42 | 1.47 | 1.34 | 1.39 | 1.52 | 1.396 | 1.42 | 1.36 | 1.45 |  |  | 16
| 1 | 4 | 1.41 | 1.375 | 1.34 | 1.436 | 1.3852 | 1.36 | 1.39 | 1.51 | 1.32 | 1.37 | 1.23 | 1.45 | 1.398 | 1.43 | 1.37 | 1.38 | 1.38 | 1.52 | 1.382 | 1.44 | 1.4 | 1.44 |  |  | 1.4078
| 1 | 5 | 1.45 | 1.373 | 1.39 | 1.375 | 1.4204 | 1.36 | 1.43 | 1.53 | 1.35 | 1.37 | 1.29 | 1.473 | 1.43 | 1.32 | 1.47 | 1.35 | 1.36 | 1.51 | 1.379 | 1.45 | 1.4 | 1.44 |  |  | 92
| 1 | 6 | 1.41 | 1.374 | 1.38 | 1.46 | 1.3531 | 1.33 | 1.4 | 1.5 | 1.35 | 1.37 | 1.24 | 1.507 | 1.4403 | 1.35 | 1.48 | 1.34 | 1.42 | 1.52 | 1.389 | 1.46 | 1.37 | 1.41
| Mean |  | 1.4183 | 1.3651 | 1.3633 | 1.4297 | 1.3904 | 1.3533 | 1.4183 | 1.51 | 1.3433 | 1.3667 | 1.2483 | 1.5003 | 1.4173 | 1.3483 | 1.4583 | 1.3783 | 1.4033 | 1.52 | 1.388 | 1.4433 | 1.3883 | 1.4333
| Median |  | 1.41 | 1.3735 | 1.365 | 1.4265 | 1.3901 | 1.36 | 1.425 | 1.505 | 1.35 | 1.37 | 1.25 | 1.51 | 1.4157 | 1.34 | 1.47 | 1.365 | 1.4 | 1.52 | 1.386 | 1.445 | 1.395 | 1.435
| Std Dev. |  | 0.016021 | 0.015117 | 0.026583 | 0.061141 | 0.026034 | 0.018619 | 0.019408 | 0.012649 | 0.012111 | 0.015055 | 0.031885 | 0.032222 | 0.023486 | 0.070261 | 0.044008 | 0.043089 | 0.035024 | 0.0063246 | 0.0080994 | 0.01633 | 0.019408 | 0.013663
| Rel.Std.Dev. |  | 1.13% | 1.107% | 1.95% | 4.277% | 1.872% | 1.376% | 1.368% | 0.8377% | 0.9015% | 1.102% | 2.554% | 2.148% | 1.657% | 5.211% | 3.018% | 3.126% | 2.496% | 0.4161% | 0.5835% | 1.131% | 1.398% | 0.9532%
| PDM3 |  | 0.7455% | -3.037% | -3.161% | 1.551% | -1.24% | -3.872% | 0.7455% | 7.257% | -4.582% | -2.924% | -11.33% | 6.57% | 0.6713% | -4.227% | 3.587% | -2.096% | -0.32% | 7.967% | -1.409% | 2.521% | -1.385% | 1.811%
| Z-Score (Absolute) |  | 0.44 | 0.65 | 0.68 | 0.67 | 0.13 | 0.88 | 0.44 | 2.3 | 1.09 | 0.61 | 3.02 | 2.1 | 0.42 | 0.99 | 1.25 | 0.38 | 0.13 | 2.5 | 0.18 | 0.94 | 0.17 | 0.74
| Analytical results for Mn in OREAS 163b (Certified Value 0.046 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0456 | 0.047 | 0.04562 | 0.0488 | 0.046564 | 0.0453 | 0.0484 | 0.0482 | 0.0449 | 0.0455 | 0.0421 | 0.0471 | 0.046165 | 0.0468 | 0.0477 | 0.05 | 0.0472 | 0.047 | 0.0456 | 0.051 | 0.0476 | 0.0443 |  |  | 1
| 1 | 2 | 0.0455 | 0.0475 | 0.04629 | 0.04714 | 0.047943 | 0.0454 | 0.0488 | 0.0481 | 0.0453 | 0.0446 | 0.043 | 0.0476 | 0.045106 | 0.0494 | 0.0471 | 0.0492 | 0.0468 | 0.0469 | 0.0461 | 0.0517 | 0.0467 | 0.0443 |  |  | 30
| 1 | 3 | 0.045 | 0.0472 | 0.04772 | 0.04771 | 0.046369 | 0.0451 | 0.0487 | 0.0478 | 0.0454 | 0.0441 | 0.0441 | 0.0474 | 0.046201 | 0.0529 | 0.0474 | 0.0462 | 0.0466 | 0.0468 | 0.0457 | 0.0506 | 0.0458 | 0.0432 |  |  | 16
| 1 | 4 | 0.045 | 0.0472 | 0.04597 | 0.04645 | 0.046404 | 0.0448 | 0.0478 | 0.0486 | 0.0447 | 0.0454 | 0.0441 | 0.0451 | 0.045165 | 0.0532 | 0.0447 | 0.048 | 0.0461 | 0.0468 | 0.0454 | 0.0514 | 0.0472 | 0.0443 |  |  | 0.046404
| 1 | 5 | 0.047 | 0.0473 | 0.04753 | 0.04739 | 0.047007 | 0.0448 | 0.048 | 0.0489 | 0.0448 | 0.0446 | 0.043 | 0.0456 | 0.04556 | 0.0495 | 0.0478 | 0.047 | 0.0459 | 0.0468 | 0.0462 | 0.051 | 0.0472 | 0.0431 |  |  | 93
| 1 | 6 | 0.0452 | 0.0475 | 0.04731 | 0.04831 | 0.045381 | 0.0444 | 0.0478 | 0.0479 | 0.0458 | 0.046 | 0.041 | 0.0468 | 0.046342 | 0.0502 | 0.048 | 0.0463 | 0.0473 | 0.0468 | 0.0451 | 0.0516 | 0.0462 | 0.0433
| Mean |  | 0.04555 | 0.047283 | 0.04674 | 0.047633 | 0.046611 | 0.044967 | 0.04825 | 0.04825 | 0.04515 | 0.045033 | 0.042883 | 0.0466 | 0.045756 | 0.050333 | 0.047117 | 0.047783 | 0.04665 | 0.04685 | 0.045683 | 0.051217 | 0.046783 | 0.04375
| Median |  | 0.04535 | 0.04725 | 0.0468 | 0.04755 | 0.046484 | 0.04495 | 0.0482 | 0.04815 | 0.0451 | 0.045 | 0.043 | 0.04695 | 0.045862 | 0.04985 | 0.04755 | 0.0475 | 0.0467 | 0.0468 | 0.04565 | 0.0512 | 0.04695 | 0.0438
| Std Dev. |  | 0.00075299 | 0.00019408 | 0.00088985 | 0.00083997 | 0.00084207 | 0.00037238 | 0.00044609 | 0.00042308 | 0.00042308 | 0.00071181 | 0.0011957 | 0.0010178 | 0.00055109 | 0.0024031 | 0.0012254 | 0.0015677 | 0.00056833 | 8.3666e-05 | 0.00041673 | 0.00042151 | 0.0006824 | 0.00060581
| Rel.Std.Dev. |  | 1.653% | 0.4105% | 1.904% | 1.763% | 1.807% | 0.8281% | 0.9245% | 0.8769% | 0.9371% | 1.581% | 2.788% | 2.184% | 1.204% | 4.774% | 2.601% | 3.281% | 1.218% | 0.1786% | 0.9122% | 0.823% | 1.459% | 1.385%
| PDM3 |  | -1.841% | 1.894% | 0.7232% | 2.648% | 0.4459% | -3.098% | 3.977% | 3.977% | -2.703% | -2.955% | -7.588% | 0.4215% | -1.396% | 8.467% | 1.535% | 2.972% | 0.5292% | 0.9602% | -1.554% | 10.37% | 0.8166% | -5.72%
| Z-Score (Absolute) |  | 0.74 | 0.38 | 0.03 | 0.6 | 0.05 | 1.11 | 1 | 1 | 0.99 | 1.07 | 2.45 | 0.06 | 0.6 | 2.34 | 0.27 | 0.7 | 0.03 | 0.1 | 0.65 | 2.9 | 0.06 | 1.89
| Analytical results for Mo in OREAS 163b (Certified Value 2.09 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 2.1 | 2.35 | 3.8 | 2.06 | 1.979 | 2.33 | 2.2 | 2.13 | 1.97 | 2.14 | 2.51 | 2.22 | 1.6256 | 2.3 | 2.04 | 1.5 | 2.1 | 1.97 | 2.21 | 1.98 | 2 | 2.08 | 2 |  | 1
| 1 | 2 | 2.1 | 2.28 | 3.4 | 2.05 | 1.9971 | 2.3 | 2.21 | 2.02 | 2.05 | 2.12 | 2.44 | 2.09 | 1.4607 | 2.15 | 2 | 1.6 | 2 | 2 | 2.08 | 1.89 | 2 | 2.1 | 2 |  | 31
| 1 | 3 | 2.1 | 2.3 | 3.7 | 2.12 | 2.1172 | 2.63 | 2.2 | 2.1 | 1.95 | 2.17 | 2.36 | 2.16 | 1.6065 | 2.04 | 2.03 | 1.7 | 2 | 1.75 | 2.1 | 2.05 | 2 | 2.1 | 2 |  | 16
| 1 | 4 | 2 | 2.3 | 3.5 | 2.01 | 2.0191 | 2.21 | 2.11 | 2.15 | 2.14 | 2.09 | 2.43 | 2.07 | 1.3914 | 1.99 | 2.04 | 1.6 | 2 | 1.77 | 2.09 | 1.8 | 2 | 2.12 | 2 |  | 2.0909
| 1 | 5 | 2.1 | 2.29 | 3.4 | 2.18 | 2.2464 | 2.28 | 2.08 | 2.14 | 1.99 | 2.13 | 2.41 | 2.21 | 1.6776 | 2.05 | 1.85 | 1.6 | 2 | 1.79 | 2.19 | 1.91 | 2 | 2.09 | 1.9 |  | 94
| 1 | 6 | 2.1 | 2.33 | 3.7 | 2.15 | 2.0507 | 2.17 | 2.31 | 2.08 | 1.97 | 1.89 | 2.36 | 2.03 | 1.7945 | 2.01 | 1.93 | < 0.2 | 1.9 | 1.87 | 2.07 | 1.88 | 2 | 2.02 | 2
| Mean |  | 2.0833 | 2.3083 | 3.5833 | 2.095 | 2.0682 | 2.32 | 2.185 | 2.1033 | 2.0117 | 2.09 | 2.4183 | 2.13 | 1.5927 | 2.09 | 1.9817 | 1.6 | 2 | 1.8583 | 2.1233 | 1.9183 | 2 | 2.085 | 1.9833
| Median |  | 2.1 | 2.3 | 3.6 | 2.09 | 2.0349 | 2.29 | 2.2 | 2.115 | 1.98 | 2.125 | 2.42 | 2.125 | 1.6161 | 2.045 | 2.015 | 1.6 | 2 | 1.83 | 2.095 | 1.9 | 2 | 2.095 | 2
| Std Dev. |  | 0.040825 | 0.026394 | 0.17224 | 0.065345 | 0.099885 | 0.16297 | 0.08167 | 0.048442 | 0.071671 | 0.10139 | 0.056362 | 0.07823 | 0.14639 | 0.11679 | 0.07679 | 0.070711 | 0.063246 | 0.10666 | 0.060553 | 0.086583 | 0 | 0.034496 | 0.040825
| Rel.Std.Dev. |  | 1.96% | 1.143% | 4.807% | 3.119% | 4.83% | 7.025% | 3.738% | 2.303% | 3.563% | 4.851% | 2.331% | 3.673% | 9.191% | 5.588% | 3.875% | 4.419% | 3.162% | 5.74% | 2.852% | 4.513% | 0% | 1.655% | 2.058%
| PDM3 |  | -0.364% | 10.4% | 71.37% | 0.194% | -1.086% | 10.95% | 4.498% | 0.5925% | -3.791% | -0.04515% | 15.66% | 1.868% | -23.83% | -0.04515% | -5.226% | -23.48% | -4.349% | -11.12% | 1.549% | -8.255% | -4.349% | -0.2843% | -5.147%
| Z-Score (Absolute) |  | 0.03 | 1.61 | 10.9 | 0.05 | 0.14 | 1.69 | 0.71 | 0.12 | 0.55 | 0.02 | 2.41 | 0.31 | 3.61 | 0.02 | 0.77 | 5.38 | 0.64 | 1.67 | 0.26 | 1.23 | NA | 0.02 | 0.76
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 163b (Certified Value 0.01 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.012 | 0.014 | 0.01073 | < 0.01 | 0.009 | 0.006 | 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 |  |  |  |  | 1
| 1 | 2 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.015 | 0.010454 | < 0.01 | 0.008 | 0.006 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 |  |  |  |  | 23
| 1 | 3 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.012 | 0.014 | 0.011448 | < 0.01 | 0.008 | 0.005 | 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 |  |  |  |  | 16
| 1 | 4 | < 0.01 | 0.02 | 0.006 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.009 | 0.014 | < 0.01 | < 0.01 | 0.009 | 0.006 | 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 |  |  |  |  | 0.0097305
| 1 | 5 | < 0.01 | 0.01 | 0.005 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.014 | 0.011891 | < 0.01 | 0.009 | 0.005 | 0.01 | 0.01 | < 0.01 | 0.01 | < 0.01 |  |  |  |  | 95
| 1 | 6 | < 0.01 | 0.01 | 0.006 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.012 | 0.013 | 0.010268 | < 0.01 | 0.009 | 0.006 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01
| Mean |  |  | 0.011667 | 0.0058333 |  |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.010833 | 0.014 | 0.010958 |  | 0.0086667 | 0.0056667 | 0.01 | 0.01 |  | 0.01 | 0.01
| Median |  |  | 0.01 | 0.006 |  |  | 0.01 | 0.01 | 0.01 | 0.01 | 0.011 | 0.014 | 0.01073 |  | 0.009 | 0.006 | 0.01 | 0.01 |  | 0.01 | 0.01
| Std Dev. |  |  | 0.0040825 | 0.00040825 |  |  | 0 | 0 | 0 | 0 | 0.0013292 | 0.00063246 | 0.00068776 |  | 0.0005164 | 0.0005164 | 0 | 0 |  | 0 | 0
| Rel.Std.Dev. |  |  | 34.99% | 6.999% |  |  | 0% | 0% | 0% | 0% | 12.27% | 4.518% | 6.276% |  | 5.958% | 9.113% | 0% | 0% |  | 0% | 0%
| PDM3 |  |  | 19.9% | -40.05% |  |  | 2.769% | 2.769% | 2.769% | 2.769% | 11.33% | 43.88% | 12.62% |  | -10.93% | -41.76% | 2.769% | 2.769% |  | 2.769% | 2.769%
| Z-Score (Absolute) |  | 2.24 | 0.76 | 1.87 | 2.24 | 2.24 | 0.01 | 0.01 | 0.01 | 0.01 | 0.38 | 1.81 | 0.01 | 2.24 | 0.59 | 1.94 | 0.01 | 0.01 | 2.24 | 0.01 | 1.49
| Analytical results for Nb in OREAS 163b (Certified Value < 0.05 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [9 empty] | 3
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [9 empty] | 3
| 1 | 1 | < 0.2 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | < 0.05 | 0.12 | < 0.05 | < 0.05 | < 0.1 | [9 empty] | 1
| 1 | 2 | < 0.2 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | < 0.05 | 0.13 | < 0.05 | < 0.05 | < 0.1 | [9 empty] | 14
| 1 | 3 | < 0.2 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | < 0.05 | 0.12 | < 0.05 | < 0.05 | < 0.1 | [9 empty] | 16
| 1 | 4 | < 0.2 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | < 0.05 | 0.11 | < 0.05 | < 0.05 | < 0.1 | [9 empty] | < 0.05
| 1 | 5 | < 0.2 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | < 0.05 | 0.12 | < 0.05 | < 0.05 | < 0.1 | [9 empty] | 96
| 1 | 6 | < 0.2 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.05 | < 0.2 | < 0.05 | 0.12 | < 0.05 | < 0.05 | < 0.1
| Mean |  | 2.20: 0.12
| Median |  | 2.20: 0.12
| Std Dev. |  | 2.20: 0.0063246
| Rel.Std.Dev. |  | 2.20: 5.27%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.25
| Analytical results for Nd in OREAS 163b (Indicative Value 2.67 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.07: Lab | [9 empty] | 1
| No | No | 2.01: 2.02; 2.04: 2.06; 2.07: 2.12 | [9 empty] | ppm
|  |  | 2.01: AR*MS; 2.04: AR*MS; 2.07: AR*OES/MS | [9 empty] | 2
|  |  | 2.01: 0.5g; 2.04: 0.2g; 2.07: 0.5g | [9 empty] | 2
| 1 | 1 | 2.01: 2.6807; 2.04: 2.8728; 2.07: 2.5 | [9 empty] | 1
| 1 | 2 | 2.01: 2.7788; 2.04: 2.9684; 2.07: 2.51 | [9 empty] | 5
| 1 | 3 | 2.01: 2.7964; 2.04: 2.8796; 2.07: 2.28 | [9 empty] | 16
| 1 | 4 | 2.01: 2.6814; 2.04: 2.6936; 2.07: 2.32 | [9 empty] | 2.6657
| 1 | 5 | 2.01: 2.8213; 2.04: 3.1407; 2.07: 2.31 | [9 empty] | 11
| 1 | 6 | 2.01: 2.8415; 2.04: 2.6982; 2.07: 2.21
| Mean |  | 2.01: 2.7667; 2.04: 2.8755; 2.07: 2.355
| Median |  | 2.01: 2.7876; 2.04: 2.8762; 2.07: 2.315
| Std Dev. |  | 2.01: 0.06968; 2.04: 0.16941; 2.07: 0.12243
| Rel.Std.Dev. |  | 2.01: 2.519%; 2.04: 5.892%; 2.07: 5.199%
| PDM3 |  | 2.01: 3.787%; 2.04: 7.87%; 2.07: -11.66%
| Z-Score (Absolute) |  | 2.01: 0; 2.04: 0.67; 2.07: 2.55
| Analytical results for Ni in OREAS 163b (Certified Value 23.3 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 2
| 1 | 1 | 24 | 22.63 | 23.6 | 21.32 | 23.9 | 25.2684 | 22 | 24.4 | 23.7 | 21.1 | 23.5 | 21.7 | 23 | 23.4647 | 23.8 | 23.9 | 24.2 | 24.4 | 25.7 | 22.9 | 28 | 23.4 | 21.5 |  | 1
| 1 | 2 | 25 | 22.99 | 23.7 | 20.06 | 23 | 24.3553 | 22 | 24.5 | 23.2 | 21.1 | 23.2 | 21.5 | 22.4 | 22.6471 | 25.1 | 23.9 | 23.7 | 24.6 | 23.8 | 23 | 25 | 23 | 22 |  | 16
| 1 | 3 | 24 | 22.79 | 23.5 | 21.04 | 22.6 | 25.8238 | 30.2 | 23.5 | 23.4 | 20.4 | 23.2 | 22.1 | 23.7 | 23.5642 | 25.3 | 23.2 | 22.9 | 24.8 | 24 | 24.7 | 26 | 23.2 | 21.6 |  | 16
| 1 | 4 | 24 | 23.49 | 23.5 | 22.2 | 23 | 25.1293 | 21.2 | 22.8 | 24.5 | 22 | 23.1 | 21.6 | 21.6 | 23.1626 | 25.8 | 23.4 | 23.9 | 23.6 | 24.6 | 22.6 | 25 | 23.5 | 21.9 |  | 23.2926
| 1 | 5 | 24 | 23.63 | 24.9 | 21.51 | 25.5 | 23.5658 | 21.9 | 23.5 | 24.7 | 20.6 | 23.4 | 21.7 | 23.2 | 23.0202 | 25.3 | 22.6 | 23.2 | 24.1 | 24.6 | 23.2 | 25 | 23.5 | 21.3 |  | 97
| 1 | 6 | 23 | 22.84 | 24.3 | 21.34 | 23.5 | 23.5561 | 20.4 | 26.5 | 23.5 | 20.8 | 23.5 | 20.4 | 21.7 | 23.9615 | 25.6 | 23.5 | 22.5 | 24.9 | 25.3 | 23.9 | 27 | 22.4 | 21.7
| Mean |  | 24 | 23.0617 | 23.9167 | 21.245 | 23.5833 | 24.6165 | 22.95 | 24.2 | 23.8333 | 21 | 23.3167 | 21.5 | 22.6 | 23.3034 | 25.15 | 23.4167 | 23.4 | 24.4 | 24.6667 | 23.3833 | 26 | 23.1667 | 21.6667
| Median |  | 24 | 22.915 | 23.65 | 21.33 | 23.25 | 24.7423 | 21.95 | 23.95 | 23.6 | 20.95 | 23.3 | 21.65 | 22.7 | 23.3136 | 25.3 | 23.45 | 23.45 | 24.5 | 24.6 | 23.1 | 25.5 | 23.3 | 21.65
| Std Dev. |  | 0.63246 | 0.40519 | 0.56716 | 0.69902 | 1.042 | 0.94255 | 3.6065 | 1.2931 | 0.6186 | 0.56214 | 0.17224 | 0.57619 | 0.84617 | 0.45999 | 0.7064 | 0.48751 | 0.64498 | 0.4858 | 0.73121 | 0.77825 | 1.2649 | 0.42269 | 0.2582
| Rel.Std.Dev. |  | 2.635% | 1.757% | 2.371% | 3.29% | 4.418% | 3.829% | 15.71% | 5.343% | 2.596% | 2.677% | 0.7387% | 2.68% | 3.744% | 1.974% | 2.809% | 2.082% | 2.756% | 1.991% | 2.964% | 3.328% | 4.865% | 1.825% | 1.192%
| PDM3 |  | 3.037% | -0.9913% | 2.679% | -8.791% | 1.248% | 5.684% | -1.471% | 3.896% | 2.322% | -9.842% | 0.1035% | -7.696% | -2.973% | 0.04644% | 7.974% | 0.5328% | 0.4612% | 4.754% | 5.899% | 0.3897% | 11.62% | -0.5405% | -6.98%
| Z-Score (Absolute) |  | NA | 0.34 | 0.54 | 2.2 | 0.2 | 1.25 | 0.45 | 0.83 | 0.45 | 2.45 | 0.08 | 1.94 | 0.81 | 0.09 | 1.8 | 0.03 | 0.01 | 1.03 | 1.31 | 0.01 | 2.67 | 0.23 | 1.77
Results from laboratory 2.01 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 163b (Certified Value 0.016 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0164 | 0.012 | 0.02 | 0.0154 | 0.016107 | 0.017 | 0.017 | 0.017 | 0.016 | 0.016 | 0.019 | 0.018 | 0.01723 | 0.016 | 0.017 | 0.016 | 0.0153 | 0.02 | 0.0163 | 0.0177 | 0.0166 | 0.017 |  |  | 1
| 1 | 2 | 0.0166 | 0.0119 | 0.02 | 0.0148 | 0.017219 | 0.017 | 0.017 | 0.018 | 0.016 | 0.016 | 0.019 | 0.018 | 0.016518 | 0.017 | 0.016 | 0.016 | 0.0159 | 0.02 | 0.0173 | 0.0183 | 0.0164 | 0.017 |  |  | 0
| 1 | 3 | 0.0165 | 0.0119 | 0.02 | 0.0148 | 0.015813 | 0.016 | 0.017 | 0.018 | 0.016 | 0.016 | 0.019 | 0.017 | 0.017317 | 0.016 | 0.016 | 0.015 | 0.0146 | 0.02 | 0.0168 | 0.0176 | 0.0158 | 0.016 |  |  | 16
| 1 | 4 | 0.0166 | 0.0121 | 0.02 | 0.0145 | 0.015547 | 0.017 | 0.016 | 0.018 | 0.016 | 0.016 | 0.019 | 0.017 | 0.016934 | 0.016 | 0.016 | 0.016 | 0.0144 | 0.02 | 0.0166 | 0.0174 | 0.0161 | 0.017 |  |  | 0.016458
| 1 | 5 | 0.0164 | 0.0122 | 0.02 | 0.0151 | 0.016349 | 0.017 | 0.017 | 0.018 | 0.016 | 0.016 | 0.02 | 0.017 | 0.016893 | 0.019 | 0.016 | 0.015 | 0.0146 | 0.02 | 0.0167 | 0.0178 | 0.0165 | 0.017 |  |  | 98
| 1 | 6 | 0.0164 | 0.012 | 0.02 | 0.0146 | 0.015311 | 0.016 | 0.016 | 0.018 | 0.016 | 0.016 | 0.019 | 0.017 | 0.016646 | 0.013 | 0.018 | 0.015 | 0.0149 | 0.02 | 0.0164 | 0.0183 | 0.0164 | 0.017
| Mean |  | 0.016483 | 0.012017 | 0.02 | 0.014867 | 0.016058 | 0.016667 | 0.016667 | 0.017833 | 0.016 | 0.016 | 0.019167 | 0.017333 | 0.016923 | 0.016167 | 0.0165 | 0.0155 | 0.01495 | 0.02 | 0.016683 | 0.01785 | 0.0163 | 0.016833
| Median |  | 0.01645 | 0.012 | 0.02 | 0.0148 | 0.01596 | 0.017 | 0.017 | 0.018 | 0.016 | 0.016 | 0.019 | 0.017 | 0.016913 | 0.016 | 0.016 | 0.0155 | 0.01475 | 0.02 | 0.01665 | 0.01775 | 0.0164 | 0.017
| Std Dev. |  | 9.8319e-05 | 0.0001169 | 0 | 0.00033267 | 0.00068053 | 0.0005164 | 0.0005164 | 0.00040825 | 0 | 0 | 0.00040825 | 0.0005164 | 0.00031323 | 0.0019408 | 0.00083666 | 0.00054772 | 0.00056125 | 0 | 0.00035449 | 0.00037283 | 0.00029665 | 0.00040825
| Rel.Std.Dev. |  | 0.5965% | 0.9729% | 0% | 2.238% | 4.238% | 3.098% | 3.098% | 2.289% | 0% | 0% | 2.13% | 2.979% | 1.851% | 12% | 5.071% | 3.534% | 3.754% | 0% | 2.125% | 2.089% | 1.82% | 2.425%
| PDM3 |  | 0.1528% | -26.99% | 21.52% | -9.67% | -2.434% | 1.267% | 1.267% | 8.355% | -2.784% | -2.784% | 16.46% | 5.317% | 2.824% | -1.771% | 0.2541% | -5.822% | -9.164% | 21.52% | 1.368% | 8.457% | -0.9611% | 2.279%
| Z-Score (Absolute) |  | 0.12 | 5.28 | 3.95 | 1.98 | 0.61 | 0.1 | 0.1 | 1.44 | 0.67 | 0.67 | 2.99 | 0.87 | 0.39 | 0.48 | 0.1 | 1.25 | 1.89 | 3.95 | 0.12 | 1.46 | 0.33 | 0.29
| Analytical results for Pb in OREAS 163b (Certified Value 54 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 1
| 1 | 1 | 55.2 | 47.75 | 56.805 | 68.8 | 54.2 | 51.0904 | 53.7 | 53.9 | 54.4 | 52.3 | 52.8 | 56.6 | 54 | 52.6917 | 77 | 54.41 | 53 | 56.8 | 54 | 55.5 | 55 | 57 | 55.3 | 57.1 | 1
| 1 | 2 | 55.4 | 49.29 | 55.818 | 72.4 | 52.8 | 54.5013 | 53.3 | 52.8 | 52.8 | 52.6 | 50.6 | 59.3 | 55 | 51.9246 | 73 | 52.31 | 52 | 56 | 54 | 55.9 | 55.1 | 56 | 55.6 | 55.8 | 17
| 1 | 3 | 54.7 | 50.92 | 56.63 | 70.3 | 52.6 | 53.6884 | 52.5 | 54.2 | 53.3 | 53.1 | 52.2 | 60.7 | 55.4 | 53.5964 | 70 | 53.64 | 52 | 53 | 53 | 55.4 | 52.5 | 56 | 53.9 | 56.5 | 16
| 1 | 4 | 55 | 45.02 | 56.95 | 72.9 | 49.9 | 53.732 | 52.5 | 52 | 54.3 | 52.9 | 52.8 | 55.2 | 51.5 | 48.2784 | 78 | 52.41 | 55 | 54.9 | 52 | 55.8 | 52.6 | 58 | 57.3 | 56.8 | 54.0453
| 1 | 5 | 54.8 | 47.05 | 56.257 | 70.3 | 51.4 | 56.7015 | 52.7 | 54.1 | 54.2 | 52.5 | 52.8 | 58.1 | 54.5 | 54.7647 | 57 | 52.52 | 54 | 53.4 | 53 | 55.2 | 52.9 | 57 | 55 | 54.9 | 99
| 1 | 6 | 54 | 51.78 | 59.25 | 75 | 51.4 | 51.1776 | 51.4 | 52 | 52.5 | 51.9 | 55.7 | 54 | 54.7 | 48.6133 | 58 | 52.03 | 55 | 52.4 | 53 | 55.4 | 54 | 57 | 54.4 | 54.2
| Mean |  | 54.85 | 48.635 | 56.9517 | 71.6167 | 52.05 | 53.4819 | 52.6833 | 53.1667 | 53.5833 | 52.55 | 52.8167 | 57.3167 | 54.1833 | 51.6449 | 68.8333 | 52.8867 | 53.5 | 54.4167 | 53.1667 | 55.5333 | 53.6833 | 56.8333 | 55.25 | 55.8833
| Median |  | 54.9 | 48.52 | 56.7175 | 71.35 | 52 | 53.7102 | 52.6 | 53.35 | 53.75 | 52.55 | 52.8 | 57.35 | 54.6 | 52.3081 | 71.5 | 52.465 | 53.5 | 54.15 | 53 | 55.45 | 53.45 | 57 | 55.15 | 56.15
| Std Dev. |  | 0.48888 | 2.5261 | 1.1979 | 2.2409 | 1.4802 | 2.1232 | 0.79099 | 1.0328 | 0.82805 | 0.42778 | 1.6497 | 2.5294 | 1.3963 | 2.6548 | 9.2394 | 0.9291 | 1.3784 | 1.7623 | 0.75277 | 0.26583 | 1.1856 | 0.75277 | 1.1777 | 1.1409
| Rel.Std.Dev. |  | 0.8913% | 5.194% | 2.103% | 3.129% | 2.844% | 3.97% | 1.501% | 1.943% | 1.545% | 0.8141% | 3.124% | 4.413% | 2.577% | 5.14% | 13.42% | 1.757% | 2.576% | 3.239% | 1.416% | 0.4787% | 2.209% | 1.325% | 2.132% | 2.042%
| PDM3 |  | 1.489% | -10.01% | 5.378% | 32.51% | -3.692% | -1.042% | -2.52% | -1.626% | -0.8547% | -2.767% | -2.273% | 6.053% | 0.2555% | -4.441% | 27.36% | -2.144% | -1.009% | 0.6872% | -1.626% | 2.753% | -0.6697% | 5.159% | 2.229% | 3.401%
| Z-Score (Absolute) |  | 0.71 | 2.93 | 1.95 | 10.54 | 0.93 | 0.09 | 0.56 | 0.27 | 0.03 | 0.64 | 0.48 | 2.16 | 0.32 | 1.17 | 8.91 | 0.44 | 0.08 | 0.46 | 0.27 | 1.11 | 0.03 | 1.88 | 0.95 | 1.32
| Analytical results for Pd in OREAS 163b (Indicative Value < 10 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.14 | ppb
|  |  | 2.01: AR*OES/MS | 1
|  |  | 2.01: 01g | 1
| 1 | 1 | 2.01: < 10 | 1
| 1 | 2 | 2.01: < 10 | 6
| 1 | 3 | 2.01: < 10 | 16
| 1 | 4 | 2.01: < 10 | < 10
| 1 | 5 | 2.01: < 10 | 12
| 1 | 6 | 2.01: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Pr in OREAS 163b (Indicative Value 0.64 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g | 2
| 1 | 1 | 2.01: 0.6134; 2.02: 0.66204; 2.03: 0.6 | 1
| 1 | 2 | 2.01: 0.6339; 2.02: 0.65544; 2.03: 0.6 | 18
| 1 | 3 | 2.01: 0.6424; 2.02: 0.6366; 2.03: 0.6 | 16
| 1 | 4 | 2.01: 0.6291; 2.02: 0.60342; 2.03: 0.5 | 0.63767
| 1 | 5 | 2.01: 0.647; 2.02: 0.67552; 2.03: 0.5 | 9
| 1 | 6 | 2.01: 0.6564; 2.02: 0.59681; 2.03: 0.5
| Mean |  | 2.01: 0.63703; 2.02: 0.6383; 2.03: 0.55
| Median |  | 2.01: 0.63815; 2.02: 0.64602; 2.03: 0.55
| Std Dev. |  | 2.01: 0.015063; 2.02: 0.032199; 2.03: 0.054772
| Rel.Std.Dev. |  | 2.01: 2.365%; 2.02: 5.044%; 2.03: 9.959%
| PDM3 |  | 2.01: -0.09963%; 2.02: 0.09963%; 2.03: -13.75%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.03: 46.19
| Analytical results for Pt in OREAS 163b (Indicative Value < 5 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.14 | ppb
|  |  | 2.01: AR*OES/MS | 2
|  |  | 2.01: 01g | 2
| 1 | 1 | 2.01: < 5 | 1
| 1 | 2 | 2.01: < 5 | 4
| 1 | 3 | 2.01: < 5 | 16
| 1 | 4 | 2.01: < 5 | < 5
| 1 | 5 | 2.01: < 5 | 10
| 1 | 6 | 2.01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Rb in OREAS 163b (Certified Value 1.25 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 1.62 | 9.4 | 1.29 | 1.238 | 1.2 | 1.3 | 1.5 | 1.2 | 1.3 | 1.2 | 1.19 | 1.2484 | 1.2 | 1.3 | 1.4 | 1 | 1.3 | 1.3 | [6 empty] | 1
| 1 | 2 | 1.61 | 9.8 | 1.23 | 1.1311 | 1.2 | 1.3 | 1.4 | 1.2 | 1.2 | 1.2 | 1.15 | 1.206 | 1.3 | 1.2 | 1.4 | 1.3 | 1.2 | 1.4 | [6 empty] | 37
| 1 | 3 | 1.65 | 7.8 | 1.29 | 1.1501 | 1.2 | 1.2 | 1.4 | 1.2 | 1.2 | 1.2 | 1.2 | 1.3404 | 1.4 | 1.1 | 1.3 | 1.1 | 1.2 | 1.3 | [6 empty] | 16
| 1 | 4 | 1.64 | 7.6 | 1.23 | 1.1846 | 1.2 | 1.2 | 1.4 | 1.2 | 1.2 | 1.2 | 1.15 | 1.1587 | 1.3 | 1.2 | 1.4 | 1.2 | 1.2 | 1.3 | [6 empty] | 1.2458
| 1 | 5 | 1.74 | 9.5 | 1.26 | 1.1984 | 1.2 | 1.3 | 1.4 | 1.2 | 1.3 | 1.1 | 1.2 | 1.2356 | 1.2 | 1.2 | 1.4 | 1.3 | 1.2 | 1.4 | [6 empty] | 100
| 1 | 6 | 1.66 | 8.8 | 1.3 | 1.136 | 1.2 | 1.3 | 1.4 | 1.2 | 1.3 | 1.2 | 1.17 | 1.3135 | 1 | 1.1 | 1.4 | 1.3 | 1.2 | 1.3
| Mean |  | 1.6533 | 8.8167 | 1.2667 | 1.173 | 1.2 | 1.2667 | 1.4167 | 1.2 | 1.25 | 1.1833 | 1.1767 | 1.2504 | 1.2333 | 1.1833 | 1.3833 | 1.2 | 1.2167 | 1.3333
| Median |  | 1.645 | 9.1 | 1.275 | 1.1674 | 1.2 | 1.3 | 1.4 | 1.2 | 1.25 | 1.2 | 1.18 | 1.242 | 1.25 | 1.2 | 1.4 | 1.25 | 1.2 | 1.3
| Std Dev. |  | 0.046332 | 0.9261 | 0.031411 | 0.041599 | 0 | 0.05164 | 0.040825 | 0 | 0.054772 | 0.040825 | 0.023381 | 0.06736 | 0.13663 | 0.075277 | 0.040825 | 0.12649 | 0.040825 | 0.05164
| Rel.Std.Dev. |  | 2.802% | 10.5% | 2.48% | 3.546% | 0% | 4.077% | 2.882% | 0% | 4.382% | 3.45% | 1.987% | 5.387% | 11.08% | 6.361% | 2.951% | 10.54% | 3.355% | 3.873%
| PDM3 |  | 32.71% | 607.7% | 1.672% | -5.844% | -3.679% | 1.672% | 13.71% | -3.679% | 0.3339% | -5.017% | -5.552% | 0.3674% | -1.004% | -5.017% | 11.04% | -3.679% | -2.342% | 7.023%
| Z-Score (Absolute) |  | 5.55 | 102.16 | 0.34 | 0.93 | 0.56 | 0.34 | 2.36 | 0.56 | 0.11 | 0.79 | 0.88 | 0.12 | 0.11 | 0.79 | 1.91 | 0.56 | 0.34 | 1.24
| Analytical results for Re in OREAS 163b (Certified Value 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.20 | 2.21 | 2.24 | 2.25 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | [10 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | 0.001 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | < 0.001 | < 0.001 | < 0.02 | 0.001 | < 0.005 | [10 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 38
| 1 | 3 | < 0.05 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 0.001
| 1 | 5 | < 0.05 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005 | [10 empty] | 101
| 1 | 6 | < 0.05 | < 0.01 | 0.001 | 0.001 | 0.001 | 0.001 | 0.001 | 0.002 | < 0.05 | 0.001 | < 0.001 | < 0.02 | < 0.001 | < 0.005
| Mean |  | 2.05: 0.001; 2.06: 0.001; 2.07: 0.001; 2.08: 0.001; 2.09: 0.001; 2.10: 0.002; 2.12: 0.001; 2.15: 0.001
| Median |  | 2.05: 0.001; 2.06: 0.001; 2.07: 0.001; 2.08: 0.001; 2.09: 0.001; 2.10: 0.002; 2.12: 0.001; 2.15: 0.001
| Std Dev. |  | 2.05: 0; 2.06: 0; 2.07: 0; 2.08: 0; 2.09: 0; 2.10: 0; 2.12: 0
Rel.Std.Dev. | bar: 2.05 to 2.10; 2.12
| PDM3 |  | 2.05: 0%; 2.06: 0%; 2.07: 0%; 2.08: 0%; 2.09: 0%; 2.10: 100%; 2.12: 0%; 2.15: 0%
| Z-Score (Absolute) |  | 35.31 | 5.88 | 0.12 | 0.12 | 0 | 0 | 0 | 1.47 | 35.31 | 0.12 | 0.74 | 13.24 | 0.61 | 2.21
| Analytical results for S in OREAS 163b (Certified Value 3.3 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 3
| 1 | 1 | 3.32 | 3.9411 | 3.258 | 2.24 | 3.36 | 3.2999 | 3.26 | 3.43 | 3.43 | 3.2 | 3.27 | 3.446 | 3.16 | 3.3248 | 3.49 | 3.39 | 3.3596 | 3.33 | 3.15 | 3.49 | 3.15 | 3.3224 | 3.45 | 3.32 | 1
| 1 | 2 | 3.36 | 3.9592 | 3.1788 | 2.25 | 3.19 | 3.3109 | 3.25 | 3.42 | 3.4 | 3.19 | 3.22 | 3.344 | 3.18 | 3.0817 | 3.32 | 3.48 | 3.3015 | 3.3 | 3.09 | 3.44 | 3.19 | 3.3587 | 3.39 | 3.35 | 0
| 1 | 3 | 3.39 | 3.9649 | 3.24 | 2.43 | 3.24 | 3.305 | 3.2 | 3.44 | 3.39 | 3.19 | 3.21 | 3.342 | 3.08 | 3.3702 | 3.47 | 3.55 | 3.3049 | 3.13 | 3.12 | 3.48 | 3.17 | 3.2824 | 3.33 | 3.35 | 16
| 1 | 4 | 3.36 | 3.9526 | 3.245 | 2.35 | 3.27 | 3.2905 | 3.25 | 3.37 | 3.43 | 3.16 | 3.27 | 3.364 | 3.03 | 3.0919 | 3.37 | 3.15 | 3.3985 | 3.2 | 3.15 | 3.5 | 3.13 | 3.3606 | 3.36 | 3.35 | 3.3028
| 1 | 5 | 3.34 | 3.9317 | 3.3206 | 2.36 | 3.4 | 3.3225 | 3.25 | 3.44 | 3.45 | 3.2 | 3.23 | 3.396 | 3.07 | 3.2875 | 3.46 | 3.5 | 3.4551 | 3.15 | 3.1 | 3.5 | 3.14 | 3.3108 | 3.44 | 3.34 | 102
| 1 | 6 | 3.36 | 3.9445 | 3.256 | 2.49 | 3.25 | 3.2863 | 3.18 | 3.39 | 3.39 | 3.19 | 3.32 | 3.315 | 3.1 | 3.3444 | 3.23 | 3.44 | 3.3302 | 3.13 | 3.16 | 3.48 | 3.15 | 3.3739 | 3.34 | 3.29
| Mean |  | 3.355 | 3.949 | 3.2498 | 2.3533 | 3.285 | 3.3025 | 3.2317 | 3.415 | 3.415 | 3.1883 | 3.2533 | 3.3678 | 3.1033 | 3.2501 | 3.39 | 3.4183 | 3.3583 | 3.2067 | 3.1283 | 3.4817 | 3.155 | 3.3348 | 3.385 | 3.3333
| Median |  | 3.36 | 3.9486 | 3.2505 | 2.355 | 3.26 | 3.3025 | 3.25 | 3.425 | 3.415 | 3.19 | 3.25 | 3.354 | 3.09 | 3.3061 | 3.415 | 3.46 | 3.3449 | 3.175 | 3.135 | 3.485 | 3.15 | 3.3405 | 3.375 | 3.345
| Std Dev. |  | 0.023452 | 0.012262 | 0.045344 | 0.098116 | 0.079183 | 0.013348 | 0.033116 | 0.02881 | 0.0251 | 0.01472 | 0.041312 | 0.046786 | 0.056804 | 0.12938 | 0.10218 | 0.14219 | 0.059758 | 0.088242 | 0.029269 | 0.022286 | 0.021679 | 0.03533 | 0.050892 | 0.024221
| Rel.Std.Dev. |  | 0.699% | 0.3105% | 1.395% | 4.169% | 2.41% | 0.4042% | 1.025% | 0.8436% | 0.735% | 0.4617% | 1.27% | 1.389% | 1.83% | 3.981% | 3.014% | 4.159% | 1.779% | 2.752% | 0.9356% | 0.6401% | 0.6871% | 1.059% | 1.503% | 0.7266%
| PDM3 |  | 1.58% | 19.56% | -1.607% | -28.75% | -0.5394% | -0.008793% | -2.154% | 3.397% | 3.397% | -3.466% | -1.498% | 1.969% | -6.04% | -1.597% | 2.64% | 3.498% | 1.68% | -2.911% | -5.283% | 5.415% | -4.475% | 0.9684% | 2.488% | 0.924%
| Z-Score (Absolute) |  | 0.32 | 5.38 | 0.58 | 8.22 | 0.28 | 0.13 | 0.73 | 0.83 | 0.83 | 1.1 | 0.55 | 0.43 | 1.83 | 0.58 | 0.61 | 0.86 | 0.34 | 0.95 | 1.61 | 1.39 | 1.39 | 0.14 | 0.57 | 0.13
| Analytical results for Sb in OREAS 163b (Certified Value 7.3 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 2
| 1 | 1 | 7.2 | 5.4277 | 9.31 | < 5 | 7.84 | 8.0566 | 7.45 | 8.3 | 6.32 | 6.7 | 7.39 | 8.16 | 7.09 | 5.509 | 7.9 | 7.64 | 7.46 | 7.1 | 3.12 | 7.44 | 6.84 | 6 | 8.07 | 9.67 | 1
| 1 | 2 | 7.33 | 5.5193 | 9.3 | < 5 | 7.36 | 7.8501 | 7.34 | 8.06 | 6.24 | 6.66 | 7.19 | 7.9 | 7.14 | 4.9871 | 8 | 7.56 | 7.51 | 7 | 3.06 | 7.13 | 7.01 | 6 | 8.1 | 9.62 | 39
| 1 | 3 | 7.09 | 5.4506 | 9.35 | < 5 | 7.51 | 7.7333 | 7.3 | 8.03 | 6.38 | 6.66 | 7.3 | 7.9 | 7.02 | 5.4367 | 9.3 | 7.46 | 7.76 | 6.5 | 3 | 7.03 | 6.88 | 6 | 8.02 | 9.54 | 16
| 1 | 4 | 7.05 | 5.3906 | 9.5 | < 5 | 7.62 | 7.6241 | 7.25 | 7.79 | 6.5 | 6.75 | 7.25 | 8.21 | 7.04 | 4.8128 | 10.5 | 7.48 | 7.71 | 6.7 | 3.05 | 7.22 | 6.78 | 6 | 8.21 | 9.64 | 7.2982
| 1 | 5 | 7.25 | 5.5471 | 9.39 | < 5 | 7.37 | 7.9332 | 7.32 | 8.13 | 6.08 | 6.69 | 7.26 | 7.79 | 7.26 | 5.5817 | 10.1 | 7.36 | 7.83 | 6.5 | 3 | 7.29 | 6.99 | 6 | 8.24 | 9.59 | 103
| 1 | 6 | 7.11 | 5.4033 | 9.51 | < 5 | 7.77 | 7.6121 | 7.24 | 7.94 | 6.12 | 6.63 | 6.98 | 7.67 | 6.78 | 5.8366 | 7.31 | 7.34 | < 0.05 | 6.6 | 2.99 | 7.04 | 6.97 | 6 | 8.15 | 9.56
| Mean |  | 7.1717 | 5.4564 | 9.3933 |  | 7.5783 | 7.8016 | 7.3167 | 8.0417 | 6.2733 | 6.6817 | 7.2283 | 7.9383 | 7.055 | 5.3607 | 8.8517 | 7.4733 | 7.654 | 6.7333 | 3.0367 | 7.1917 | 6.9117 | 6 | 8.1317 | 9.6033
| Median |  | 7.155 | 5.4391 | 9.37 |  | 7.565 | 7.7917 | 7.31 | 8.045 | 6.28 | 6.675 | 7.255 | 7.9 | 7.065 | 5.4729 | 8.65 | 7.47 | 7.71 | 6.65 | 3.025 | 7.175 | 6.925 | 6 | 8.125 | 9.605
| Std Dev. |  | 0.10704 | 0.063544 | 0.092232 |  | 0.20134 | 0.17708 | 0.07607 | 0.17244 | 0.15933 | 0.041673 | 0.13848 | 0.20971 | 0.15972 | 0.38544 | 1.3026 | 0.11501 | 0.16103 | 0.2582 | 0.050067 | 0.15817 | 0.092394 | 0 | 0.084242 | 0.049261
| Rel.Std.Dev. |  | 1.492% | 1.165% | 0.9819% |  | 2.657% | 2.27% | 1.04% | 2.144% | 2.54% | 0.6237% | 1.916% | 2.642% | 2.264% | 7.19% | 14.72% | 1.539% | 2.104% | 3.835% | 1.649% | 2.199% | 1.337% | 0% | 1.036% | 0.513%
| PDM3 |  | -1.734% | -25.24% | 28.71% |  | 3.838% | 6.897% | 0.2528% | 10.19% | -14.04% | -8.448% | -0.9575% | 8.771% | -3.333% | -26.55% | 21.29% | 2.399% | 4.875% | -7.74% | -58.39% | -1.46% | -5.296% | -17.79% | 11.42% | 31.58%
| Z-Score (Absolute) |  | 0.02 | 1.57 | 1.99 | 4.24 | 0.35 | 0.55 | 0.11 | 0.77 | 0.83 | 0.46 | 0.03 | 0.67 | 0.12 | 1.65 | 1.5 | 0.25 | 0.73 | 0.41 | 3.75 | 0 | 0.25 | NA | 0.85 | 2.18
Results from laboratory 2.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 163b (Certified Value 0.8 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | < 1 | 0.68096 | 0.6 | < 1 | 0.79 | 0.83809 | 1 | 0.9 | 0.9 | 0.9 | 0.8 | 0.7 | 0.7 | < 1 | 1 | 0.8 | 0.7 | 0.7 | 1.1 | 0.7 | 0.9 | 0.7 |  |  | 1
| 1 | 2 | < 1 | 0.69972 | 0.6 | < 1 | 0.74 | 0.84248 | 1 | 0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 0.7 | < 1 | 0.9 | 0.8 | 0.7 | 0.6 | 1 | 0.6 | 0.9 | 0.8 |  |  | 19
| 1 | 3 | < 1 | 0.6881 | 0.6 | < 1 | 0.77 | 0.84118 | 1 | 0.9 | 0.9 | 0.9 | 0.9 | 0.7 | 0.7 | < 1 | 0.8 | 0.8 | 0.7 | 0.7 | 0.9 | 0.7 | 0.9 | 0.7 |  |  | 16
| 1 | 4 | < 1 | 0.7105 | 0.6 | < 1 | 0.74 | 0.84809 | 1 | 0.8 | 0.9 | 0.9 | 0.8 | 1 | 0.7 | < 1 | 0.6 | 0.8 | 0.7 | 0.6 | 1 | 0.8 | 0.9 | 0.7 |  |  | 0.79583
| 1 | 5 | < 1 | 0.72842 | 0.5 | < 1 | 0.76 | 0.85639 | 1 | 0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 0.7 | < 1 | 0.9 | 0.7 | 0.7 | 0.6 | 1 | 0.7 | 0.9 | 0.7 |  |  | 104
| 1 | 6 | < 1 | 0.7119 | 0.6 | < 1 | 0.77 | 0.83352 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 0.7 | < 1 | 0.7 | 0.8 | 0.7 | 0.6 | 1 | 0.7 | 0.9 | 0.8
| Mean |  |  | 0.70327 | 0.58333 |  | 0.76167 | 0.84329 | 0.98333 | 0.88333 | 0.9 | 0.9 | 0.81667 | 0.8 | 0.7 |  | 0.81667 | 0.78333 | 0.7 | 0.63333 | 1 | 0.7 | 0.9 | 0.73333
| Median |  |  | 0.70511 | 0.6 |  | 0.765 | 0.84183 | 1 | 0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 0.7 |  | 0.85 | 0.8 | 0.7 | 0.6 | 1 | 0.7 | 0.9 | 0.7
| Std Dev. |  |  | 0.017319 | 0.040825 |  | 0.019408 | 0.0080286 | 0.040825 | 0.040825 | 0 | 0 | 0.040825 | 0.10954 | 1.2162e-16 |  | 0.1472 | 0.040825 | 1.2162e-16 | 0.05164 | 0.063246 | 0.063246 | 0 | 0.05164
| Rel.Std.Dev. |  |  | 2.463% | 6.999% |  | 2.548% | 0.9521% | 4.152% | 4.622% | 0% | 0% | 4.999% | 13.69% | 1.737e-14% |  | 18.02% | 5.212% | 1.737e-14% | 8.154% | 6.325% | 9.035% | 0% | 7.042%
| PDM3 |  |  | -11.63% | -26.7% |  | -4.292% | 5.964% | 23.56% | 11% | 13.09% | 13.09% | 2.619% | 0.5243% | -12.04% |  | 2.619% | -1.57% | -12.04% | -20.42% | 25.66% | -12.04% | 13.09% | -7.853%
| Z-Score (Absolute) |  | 1.45 | 0.32 | 0.99 | 1.45 | 0 | 0.45 | 1.23 | 0.67 | 0.77 | 0.77 | 0.3 | 0.21 | 0.34 | 1.45 | NA | 0.12 | 0.34 | 0.71 | 1.32 | 0.34 | 0.77 | 0.16
Results from laboratory 2.15 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 163b (Certified Value 1.68 ppm) | [25 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 1 | 0.9065 | 1.6 | < 5 | 1.773 | 1.8 | 1.6 | 1.8 | 1.7 | 1.9 | 1.6 | 1 | 1.3414 | 2 | 1.5 | < 0.05 | 1.7 | < 5 | 2 | 1.5 | 1.5 |  |  |  | 1
| 1 | 2 | 2 | 1.0104 | 1.6 | < 5 | 1.8587 | 1.8 | 2 | 1.7 | 1.8 | 1.7 | 1.7 | < 1 | 1.2027 | 0.7 | 1.8 | < 0.05 | 1.5 | < 5 | 2 | 1.6 | 2 |  |  |  | 41
| 1 | 3 | 1 | 0.9693 | 1.6 | < 5 | 1.7188 | 1.8 | 2 | 1.7 | 1.6 | 1.7 | 1.7 | < 1 | 1.7242 | 1.1 | 1.8 | < 0.05 | 1.5 | < 5 | 2 | 1.5 | 1.8 |  |  |  | 16
| 1 | 4 | 1 | 0.9383 | 1.6 | < 5 | 1.8839 | 1.9 | 1.6 | 1.6 | 1.9 | 1.9 | 1.7 | 1 | 1.2664 | 0.5 | 1.4 | < 0.05 | 1.6 | < 5 | 2 | 1.6 | 2 |  |  |  | 1.6814
| 1 | 5 | 1 | 1.0157 | 1.6 | < 5 | 1.9209 | 1.6 | 1.4 | 1.9 | 1.7 | 1.9 | 1.7 | 1 | 1.5188 | 1 | 2 | 0.34 | 1.5 | < 5 | 2 | 1.5 | 1.5 |  |  |  | 105
| 1 | 6 | 1 | 0.9579 | 1.6 | < 5 | 1.7928 | 1.7 | 1.7 | 1.6 | 1.8 | 1.7 | 1.7 | 2 | 1.6452 | 0.9 | 1.8 | < 0.05 | 1.4 | < 5 | 2 | 1.5 | 1.8
| Mean |  | 1.1667 | 0.96635 | 1.6 |  | 1.8247 | 1.7667 | 1.7167 | 1.7167 | 1.75 | 1.8 | 1.6833 | 1.25 | 1.4498 | 1.0333 | 1.7167 | 0.34 | 1.5333 |  | 2 | 1.5333 | 1.7667
| Median |  | 1 | 0.9636 | 1.6 |  | 1.8258 | 1.8 | 1.65 | 1.7 | 1.75 | 1.8 | 1.7 | 1 | 1.4301 | 0.95 | 1.8 | 0.34 | 1.5 |  | 2 | 1.5 | 1.8
| Std Dev. |  | 0.40825 | 0.042011 | 2.4324e-16 |  | 0.075931 | 0.10328 | 0.24014 | 0.1169 | 0.10488 | 0.10954 | 0.040825 | 0.5 | 0.21199 | 0.52026 | 0.22286 |  | 0.10328 |  | 0 | 0.05164 | 0.22509
| Rel.Std.Dev. |  | 34.99% | 4.347% | 1.52e-14% |  | 4.161% | 5.846% | 13.99% | 6.81% | 5.993% | 6.086% | 2.425% | 40% | 14.62% | 50.35% | 12.98% |  | 6.736% |  | 0% | 3.368% | 12.74%
| PDM3 |  | -30.61% | -42.53% | -4.84% |  | 8.524% | 5.073% | 2.099% | 2.099% | 4.082% | 7.055% | 0.1167% | -25.66% | -13.77% | -38.54% | 2.099% | -79.78% | -8.805% |  | 18.95% | -8.805% | 5.073%
| Z-Score (Absolute) |  | NA | 4.5 | 0.7 | 4.7 | 0.65 | 0.3 | 0 | 0 | 0.2 | 0.5 | 0.2 | NA | 1.6 | 4.1 | 0 | 9.84 | 1.1 | 4.7 | NA | 1.1 | 0.3
Results from laboratories 2.01, 2.13 and 2.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 163b (Indicative Value 0.78 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  |  |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 |  |  |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  |  |  | 2
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g |  |  |  | 2
| 1 | 1 | 2.01: 0.757; 2.02: 0.88064; 2.03: 0.8 |  |  |  | 1
| 1 | 2 | 2.01: 0.8016; 2.02: 0.88869; 2.03: 0.6 |  |  |  | 5
| 1 | 3 | 2.01: 0.7879; 2.02: 0.86047; 2.03: 0.6 |  |  |  | 16
| 1 | 4 | 2.01: 0.7922; 2.02: 0.84408; 2.03: 0.6 |  |  |  | 0.78311
| 1 | 5 | 2.01: 0.8256; 2.02: 0.92342; 2.03: 0.7 |  |  |  | 11
| 1 | 6 | 2.01: 0.8098; 2.02: 0.82461; 2.03: 0.8
| Mean |  | 2.01: 0.79568; 2.02: 0.87032; 2.03: 0.68333
| Median |  | 2.01: 0.7969; 2.02: 0.87056; 2.03: 0.65
| Std Dev. |  | 2.01: 0.023231; 2.02: 0.035037; 2.03: 0.098319
| Rel.Std.Dev. |  | 2.01: 2.92%; 2.02: 4.026%; 2.03: 14.39%
| PDM3 |  | 2.01: 1.605%; 2.02: 11.14%; 2.03: -12.74%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.03: 1.02
| Analytical results for Sn in OREAS 163b (Certified Value 1.63 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 2.2 | 1.1801 | < 5 | 1.77 | 1.7098 | 1.7 | 1.7 | 1.8 | 1.6 | 1.6 | 1.6 | 1.8 | 1.6583 | < 0.2 | 2.6 | 1.7 | 1.5 | 1.7 | 1.5 | [5 empty] | 1
| 1 | 2 | 2.1 | 1.2053 | < 5 | 1.7 | 1.7103 | 1.7 | 1.6 | 1.7 | 1.6 | 1.5 | 1.59 | 1.7 | 1.497 | < 0.2 | 2.4 | 1.7 | 1.6 | 1.7 | 1.5 | [5 empty] | 26
| 1 | 3 | 2.1 | 1.1804 | < 5 | 1.67 | 1.6616 | 1.6 | 1.6 | 1.7 | 1.5 | 1.5 | 1.6 | 1.7 | 1.3563 | < 0.2 | 2.4 | 1.7 | 1.6 | 1.7 | 1.5 | [5 empty] | 16
| 1 | 4 | 2.1 | 1.1761 | < 5 | 1.71 | 1.7165 | 1.7 | 1.6 | 1.8 | 1.7 | 1.5 | 1.62 | 1.7 | 1.4764 | < 0.2 | 2.5 | 1.7 | 1.6 | 1.7 | 1.5 | [5 empty] | 1.6332
| 1 | 5 | 2.1 | 1.2254 | < 5 | 1.74 | 1.7618 | 1.7 | 1.6 | 1.8 | 1.5 | 1.5 | 1.63 | 1.7 | 1.6529 | 0.4 | 2.3 | 1.7 | 1.5 | 1.7 | 1.5 | [5 empty] | 106
| 1 | 6 | 2 | 1.2019 | < 5 | 1.7 | 1.7136 | 1.6 | 1.6 | 1.7 | 1.6 | 1.4 | 1.58 | 1.7 | 1.767 | < 0.2 | 2.4 | 1.7 | 1.5 | 1.6 | 1.5
| Mean |  | 2.1 | 1.1949 |  | 1.715 | 1.7122 | 1.6667 | 1.6167 | 1.75 | 1.5833 | 1.5 | 1.6033 | 1.7167 | 1.568 | 0.4 | 2.4333 | 1.7 | 1.55 | 1.6833 | 1.5
| Median |  | 2.1 | 1.1911 |  | 1.705 | 1.7119 | 1.7 | 1.6 | 1.75 | 1.6 | 1.5 | 1.6 | 1.7 | 1.575 | 0.4 | 2.4 | 1.7 | 1.55 | 1.7 | 1.5
| Std Dev. |  | 0.063246 | 0.01934 |  | 0.035071 | 0.031777 | 0.05164 | 0.040825 | 0.054772 | 0.075277 | 0.063246 | 0.018619 | 0.040825 | 0.15048 |  | 0.10328 | 0 | 0.054772 | 0.040825 | 0
| Rel.Std.Dev. |  | 3.012% | 1.619% |  | 2.045% | 1.856% | 3.098% | 2.525% | 3.13% | 4.754% | 4.216% | 1.161% | 2.378% | 9.597% |  | 4.244% | 0% | 3.534% | 2.425% | 0%
| PDM3 |  | 28.58% | -26.84% |  | 5.007% | 4.838% | 2.047% | -1.014% | 7.15% | -3.055% | -8.158% | -1.831% | 5.109% | -3.995% | -75.51% | 48.99% | 4.088% | -5.096% | 3.068% | -8.158%
| Z-Score (Absolute) |  | 3.51 | 3.82 | 6.74 | 0.39 | 0.37 | 0 | 0.4 | 0.67 | 0.67 | 1.35 | 0.51 | 0.4 | 0.8 | 12.27 | 6.2 | 0.27 | 0.94 | 0.13 | 1.35
| Analytical results for Sr in OREAS 163b (Certified Value 7 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  | 3
| 1 | 1 | 6.9 | 3.4456 | 8.3 | 8.6 | 7.35 | 6.7781 | 7 | 7 | 7.7 | 6.8 | 7.3 | 6 | 7.5 | 6.2691 | 7.2 | 7.4 | 7 | 7.4 | 8 | 6.4 | 7 | 6.9 | 7 |  | 1
| 1 | 2 | 6.9 | 3.4943 | 8.5 | 8.8 | 6.85 | 7.0004 | 7 | 6.8 | 7.3 | 6.8 | 6.6 | 6.5 | 7.6 | 5.8636 | 7.3 | 7.2 | 7 | 7.3 | 7.7 | 6.5 | 7 | 6.9 | 7 |  | 27
| 1 | 3 | 7 | 3.4707 | 8.7 | 9.2 | 6.82 | 6.7758 | 6.7 | 6.8 | 7.4 | 6.9 | 6.8 | 6.5 | 7.6 | 5.876 | 7 | 7.3 | 7 | 7.5 | 7.9 | 6.8 | 6 | 6.7 | 7 |  | 16
| 1 | 4 | 6.8 | 3.399 | 8.8 | 8.6 | 7.02 | 6.9456 | 6.9 | 6.5 | 7.6 | 6.7 | 7 | 6.7 | 7.2 | 5.9668 | 7.1 | 7 | 7 | 7.7 | 8 | 6.5 | 7 | 7 | 7 |  | 7.0009
| 1 | 5 | 6.8 | 3.512 | 8.5 | 8.8 | 6.97 | 7.2616 | 6.9 | 6.9 | 7.8 | 6.8 | 7.2 | 6.9 | 7.4 | 6.2339 | 7 | 7.4 | 7 | 7.4 | 8.1 | 6.3 | 6 | 6.8 | 7 |  | 107
| 1 | 6 | 6.8 | 3.449 | 9.1 | 8.6 | 7.08 | 7.0982 | 6.6 | 6.7 | 7.5 | 6.9 | 7.3 | 6.4 | 7.5 | 6.3291 | 6.9 | 7.6 | 7 | 7.3 | 7.9 | 6.2 | 7 | 6.8 | 7
| Mean |  | 6.8667 | 3.4617 | 8.65 | 8.7667 | 7.015 | 6.9766 | 6.85 | 6.7833 | 7.55 | 6.8167 | 7.0333 | 6.5 | 7.4667 | 6.0897 | 7.0833 | 7.3167 | 7 | 7.4333 | 7.9333 | 6.45 | 6.6667 | 6.85 | 7
| Median |  | 6.85 | 3.4598 | 8.6 | 8.7 | 6.995 | 6.973 | 6.9 | 6.8 | 7.55 | 6.8 | 7.1 | 6.5 | 7.5 | 6.1003 | 7.05 | 7.35 | 7 | 7.4 | 7.95 | 6.45 | 7 | 6.85 | 7
| Std Dev. |  | 0.08165 | 0.040064 | 0.28107 | 0.23381 | 0.1917 | 0.18829 | 0.16432 | 0.17224 | 0.18708 | 0.075277 | 0.28752 | 0.30332 | 0.15055 | 0.21081 | 0.1472 | 0.20412 | 0 | 0.15055 | 0.13663 | 0.20736 | 0.5164 | 0.10488 | 0
| Rel.Std.Dev. |  | 1.189% | 1.157% | 3.249% | 2.667% | 2.733% | 2.699% | 2.399% | 2.539% | 2.478% | 1.104% | 4.088% | 4.666% | 2.016% | 3.462% | 2.078% | 2.79% | 0% | 2.025% | 1.722% | 3.215% | 7.746% | 1.531% | 0%
| PDM3 |  | -1.917% | -50.55% | 23.56% | 25.22% | 0.2019% | -0.3463% | -2.155% | -3.107% | 7.844% | -2.631% | 0.4638% | -7.154% | 6.653% | -13.01% | 1.178% | 4.511% | -0.01235% | 6.177% | 13.32% | -7.869% | -4.774% | -2.155% | -0.01235%
| Z-Score (Absolute) |  | 0.23 | 6.28 | 2.94 | 3.15 | 0.03 | 0.03 | 0.26 | 0.38 | 0.99 | 0.32 | 0.07 | 0.88 | 0.84 | 1.61 | 0.16 | 0.57 | NA | 0.78 | 1.67 | 0.97 | NA | 0.26 | NA
Results from laboratories 2.19, 2.23 and 2.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 163b (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | 0.02 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 10
| 1 | 3 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | < 0.01
| 1 | 5 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1 | [7 empty] | 108
| 1 | 6 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | 0.01 | < 0.05 | < 0.3 | < 0.01 | < 0.1
| Mean |  | 2.20: 0.0125
| Median |  | 2.20: 0.01
| Std Dev. |  | 2.20: 0.005
| Rel.Std.Dev. |  | 2.20: 40%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.25
| Analytical results for Tb in OREAS 163b (Certified Value 0.13 ppm) | [25 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | [7 empty] | 1
| No | No | 2.01: 2.02; 2.04: 2.04; 2.06: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.21; 2.10: 2.22 | [7 empty] | ppm
|  |  | 2.01: AR*MS; 2.04: AR*OES; 2.06: AR*MS; 2.07: AR*OES/MS; 2.08: AR*MS; 2.09: AR*OES/MS; 2.10: AR*OES/MS | [7 empty] | 2
|  |  | 2.01: 0.5g; 2.04: 0.25g; 2.06: 0.2g; 2.07: 0.5g; 2.08: 0.25g; 2.09: 0.25g; 2.10: 0.2g | [7 empty] | 2
| 1 | 1 | 2.01: 0.1019; 2.04: < 5; 2.06: 0.13978; 2.07: 0.1; 2.08: 0.13; 2.09: 0.14; 2.10: 0.13 | [7 empty] | 1
| 1 | 2 | 2.01: 0.1055; 2.04: < 5; 2.06: 0.14677; 2.07: 0.1; 2.08: 0.13; 2.09: 0.13; 2.10: 0.14 | [7 empty] | 11
| 1 | 3 | 2.01: 0.1038; 2.04: < 5; 2.06: 0.13966; 2.07: 0.1; 2.08: 0.13; 2.09: 0.13; 2.10: 0.14 | [7 empty] | 16
| 1 | 4 | 2.01: 0.1033; 2.04: < 5; 2.06: 0.13108; 2.07: 0.1; 2.08: 0.13; 2.09: 0.13; 2.10: 0.14 | [7 empty] | 0.12829
| 1 | 5 | 2.01: 0.1056; 2.04: < 5; 2.06: 0.14231; 2.07: 0.1; 2.08: 0.13; 2.09: 0.14; 2.10: 0.13 | [7 empty] | 109
| 1 | 6 | 2.01: 0.108; 2.04: < 5; 2.06: 0.1311; 2.07: 0.1; 2.08: 0.13; 2.09: 0.13; 2.10: 0.13
| Mean |  | 2.01: 0.10468; 2.06: 0.13845; 2.07: 0.1; 2.08: 0.13; 2.09: 0.13333; 2.10: 0.135
| Median |  | 2.01: 0.10465; 2.06: 0.13972; 2.07: 0.1; 2.08: 0.13; 2.09: 0.13; 2.10: 0.135
| Std Dev. |  | 2.01: 0.0021424; 2.06: 0.0062572; 2.07: 1.5202e-17; 2.08: 0; 2.09: 0.005164; 2.10: 0.0054772
| Rel.Std.Dev. |  | 2.01: 2.047%; 2.06: 4.519%; 2.07: 1.52e-14%; 2.08: 0%; 2.09: 3.873%; 2.10: 4.057%
| PDM3 |  | 2.01: -18.4%; 2.06: 7.917%; 2.07: -22.05%; 2.08: 1.33%; 2.09: 3.929%; 2.10: 5.228%
| Z-Score (Absolute) |  | 2.01: 4.71; 2.04: 377.6; 2.06: 0.68; 2.07: NA; 2.08: 0.67; 2.09: 0.13; 2.10: 0.13
Results from laboratory 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 163b (Certified Value < 0.05 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 3
|  |  | 01g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  |  | 3
| 1 | 1 | < 0.05 | < 0.02 | < 5 | < 0.2 | 0.01 | < 0.01 | 0.01 | 0.01 | < 0.01 | < 0.02 | < 0.05 | < 0.05 | < 0.01 | 0.03 | 2.22 | < 0.2 | < 5 | < 0.05 | < 0.01 | < 0.1 |  |  |  |  | 1
| 1 | 2 | < 0.05 | < 0.02 | < 5 | < 0.2 | 0.01 | 0.01 | 0.02 | 0.01 | < 0.01 | < 0.02 | < 0.05 | < 0.05 | < 0.01 | < 0.02 | 3.07 | < 0.2 | < 5 | < 0.05 | < 0.01 | < 0.1 |  |  |  |  | 12
| 1 | 3 | < 0.05 | < 0.02 | < 5 | < 0.2 | 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | < 0.02 | < 0.05 | < 0.05 | < 0.01 | < 0.02 | 3.3 | < 0.2 | < 5 | < 0.05 | < 0.01 | < 0.1 |  |  |  |  | 16
| 1 | 4 | < 0.05 | < 0.02 | < 5 | < 0.2 | 0.02 | 0.03 | 0.01 | 0.01 | 0.01 | < 0.02 | < 0.05 | < 0.05 | < 0.01 | 0.02 | 2.19 | < 0.2 | < 5 | < 0.05 | 0.01 | < 0.1 |  |  |  |  | < 0.05
| 1 | 5 | < 0.05 | < 0.02 | < 5 | < 0.2 | 0.01 | < 0.01 | 0.02 | 0.01 | 0.01 | < 0.02 | < 0.05 | < 0.05 | < 0.01 | < 0.02 | 2.23 | < 0.2 | < 5 | < 0.05 | < 0.01 | < 0.1 |  |  |  |  | 110
| 1 | 6 | < 0.05 | < 0.02 | < 5 | < 0.2 | 0.01 | 0.02 | < 0.01 | 0.01 | 0.01 | < 0.02 | < 0.05 | < 0.05 | < 0.01 | < 0.02 | 1.85 | < 0.2 | < 5 | < 0.05 | < 0.01 | < 0.1
| Mean |  | 2.07: 0.011667; 2.08: 0.02; 2.09: 0.014; 2.10: 0.01; 2.11: 0.01; 2.17: 0.025; 2.18: 2.4767; 2.24: 0.01
| Median |  | 2.07: 0.01; 2.08: 0.02; 2.09: 0.01; 2.10: 0.01; 2.11: 0.01; 2.17: 0.025; 2.18: 2.225; 2.24: 0.01
| Std Dev. |  | 2.07: 0.0040825; 2.08: 0.008165; 2.09: 0.0054772; 2.10: 0; 2.11: 0; 2.17: 0.0070711; 2.18: 0.57123
| Rel.Std.Dev. |  | 2.07: 34.99%; 2.08: 40.82%; 2.09: 39.12%; 2.10: 0%; 2.11: 0%; 2.17: 28.28%; 2.18: 23.06%
PDM3
| Z-Score (Absolute) |  | 0.34 | 0.67 | 167.23 | 5.39 | 0.56 | 0.34 | 0.51 | 0.67 | 0.84 | 0.67 | 0.34 | 0.34 | 1.01 | 0.34 | 165.66 | 5.39 | 167.23 | 0.34 | 0.96 | 2.02
| Analytical results for Th in OREAS 163b (Certified Value 0.47 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 0.59 | 0.6884 | 0.4 | < 5 | 0.425 | 0.50444 | 0.5 | 0.5 | 0.5 | 0.4 | 0.5 | 0.5 | 0.44 | 0.5413 | 0.6 | 0.5 | 0.4 | 0.5 | 0.5 | 0.5 | 0.5 |  |  |  | 1
| 1 | 2 | 0.6 | 0.7141 | 0.5 | < 5 | 0.423 | 0.55277 | 0.5 | 0.5 | 0.4 | 0.5 | 0.4 | 0.5 | 0.46 | 0.53092 | 0.4 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 | 0.5 |  |  |  | 13
| 1 | 3 | 0.58 | 0.7212 | 0.5 | < 5 | 0.417 | 0.50779 | 0.5 | 0.5 | 0.5 | 0.4 | 0.4 | 0.5 | 0.46 | 0.56191 | 0.4 | 0.4 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 |  |  |  | 16
| 1 | 4 | 0.58 | 0.7894 | 0.5 | < 5 | 0.426 | 0.48073 | 0.5 | 0.4 | 0.5 | 0.5 | 0.5 | 0.5 | 0.44 | 0.51559 | 0.4 | 0.5 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 |  |  |  | 0.47351
| 1 | 5 | 0.6 | 0.7806 | 0.4 | < 5 | 0.404 | 0.53192 | 0.5 | 0.5 | 0.4 | 0.4 | 0.4 | 0.5 | 0.46 | 0.5353 | 0.4 | 0.4 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5 |  |  |  | 111
| 1 | 6 | 0.59 | 0.795 | 0.5 | < 5 | 0.417 | 0.4723 | 0.5 | 0.5 | 0.5 | 0.4 | 0.4 | 0.5 | 0.46 | 0.57348 | 0.4 | 0.4 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5
| Mean |  | 0.59 | 0.74812 | 0.46667 |  | 0.41867 | 0.50832 | 0.5 | 0.48333 | 0.46667 | 0.43333 | 0.43333 | 0.5 | 0.45333 | 0.54308 | 0.43333 | 0.43333 | 0.4 | 0.43333 | 0.5 | 0.5 | 0.5
| Median |  | 0.59 | 0.7509 | 0.5 |  | 0.42 | 0.50611 | 0.5 | 0.5 | 0.5 | 0.4 | 0.4 | 0.5 | 0.46 | 0.5383 | 0.4 | 0.4 | 0.4 | 0.4 | 0.5 | 0.5 | 0.5
| Std Dev. |  | 0.0089443 | 0.045619 | 0.05164 |  | 0.008165 | 0.030354 | 0 | 0.040825 | 0.05164 | 0.05164 | 0.05164 | 0 | 0.010328 | 0.021194 | 0.08165 | 0.05164 | 6.0809e-17 | 0.05164 | 0 | 0 | 0
| Rel.Std.Dev. |  | 1.516% | 6.098% | 11.07% |  | 1.95% | 5.971% | 0% | 8.447% | 11.07% | 11.92% | 11.92% | 0% | 2.278% | 3.902% | 18.84% | 11.92% | 1.52e-14% | 11.92% | 0% | 0% | 0%
| PDM3 |  | 24.6% | 57.99% | -1.446% |  | -11.58% | 7.351% | 5.594% | 2.074% | -1.446% | -8.485% | -8.485% | 5.594% | -4.262% | 14.69% | -8.485% | -8.485% | -15.52% | -8.485% | 5.594% | 5.594% | 5.594%
| Z-Score (Absolute) |  | 1.44 | 3.57 | 0.22 | 27.2 | 0.87 | 0.34 | 0.22 | 0 | 0.22 | 0.67 | 0.67 | 0.22 | 0.4 | 0.81 | 0.67 | 0.67 | 1.12 | 0.67 | 0.22 | 0.22 | 0.22
| Analytical results for Ti in OREAS 163b (Certified Value < 0.005 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 0.0024 | 0.0013 | < 0.01 | 0.0015 | 0.001498 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.01 | 0.002 | 0.0021712 | < 0.005 | 0.002 | 0.001 | < 0.005 | < 0.01 | < 0.01 | 0.0025 | < 0.005 | < 0.005 |  |  | 1
| 1 | 2 | 0.0024 | 0.0013 | < 0.01 | 0.0014 | 0.0015454 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.01 | 0.002 | 0.0022044 | < 0.005 | 0.002 | 0.001 | < 0.005 | < 0.01 | < 0.01 | 0.0028 | < 0.005 | < 0.005 |  |  | 32
| 1 | 3 | 0.0025 | 0.0013 | < 0.01 | 0.0014 | 0.0015716 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.01 | 0.002 | 0.0022673 | < 0.005 | 0.002 | 0.001 | < 0.005 | < 0.01 | < 0.01 | 0.0024 | < 0.005 | < 0.005 |  |  | 16
| 1 | 4 | 0.0024 | 0.0013 | < 0.01 | 0.0014 | 0.0014897 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.01 | 0.002 | 0.0021762 | < 0.005 | 0.002 | 0.001 | < 0.005 | < 0.01 | < 0.01 | 0.0026 | < 0.005 | < 0.005 |  |  | < 0.005
| 1 | 5 | 0.0025 | 0.0013 | < 0.01 | 0.0014 | 0.0015631 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.01 | 0.002 | 0.002235 | < 0.005 | 0.002 | 0.001 | < 0.005 | < 0.01 | < 0.01 | 0.0025 | < 0.005 | < 0.005 |  |  | 112
| 1 | 6 | 0.0025 | 0.0013 | < 0.01 | 0.0014 | 0.0014519 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.01 | 0.002 | 0.0022875 | < 0.005 | 0.002 | 0.001 | < 0.005 | < 0.01 | < 0.01 | 0.0025 | < 0.005 | < 0.005
| Mean |  | 2.01: 0.00245; 2.03: 0.0013; 2.05: 0.0014167; 2.06: 0.00152; 2.13: 0.002; 2.14: 0.0022236; 2.17: 0.002; 2.19: 0.001; 2.23: 0.00255
| Median |  | 2.01: 0.00245; 2.03: 0.0013; 2.05: 0.0014; 2.06: 0.0015217; 2.13: 0.002; 2.14: 0.0022197; 2.17: 0.002; 2.19: 0.001; 2.23: 0.0025
| Std Dev. |  | 2.01: 5.4772e-05; 2.03: 0; 2.05: 4.0825e-05; 2.06: 4.7327e-05; 2.13: 0; 2.14: 4.7929e-05; 2.17: 0; 2.19: 0; 2.23: 0.00013784
| Rel.Std.Dev. |  | 2.01: 2.236%; 2.03: 0%; 2.05: 2.882%; 2.06: 3.114%; 2.13: 0%; 2.14: 2.156%; 2.17: 0%; 2.19: 0%; 2.23: 5.406%
PDM3
| Z-Score (Absolute) |  | 0.21 | 4.96 | 10.33 | 4.48 | 4.05 | 0 | 0 | 0 | 0 | 0 | 10.33 | 2.07 | 1.14 | 0 | 2.07 | 6.2 | 0 | 10.33 | 10.33 | 0.21 | 0 | 0
| Analytical results for Tl in OREAS 163b (Certified Value 1.48 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 1.55 | 1.3857 | 1.63 | < 5 | 1.5 | 1.6015 | 1.53 | 1.58 | 1.55 | 1.36 | 1.54 | 1.33 | 1.56 | 1.4671 | 1.67 | 1.63 | 1.314 | 1.5 | 1.49 | 1.53 | 1.46 | 1.41 | 1.32 |  | 1
| 1 | 2 | 1.55 | 1.4041 | 1.7 | < 5 | 1.45 | 1.6841 | 1.51 | 1.6 | 1.54 | 1.39 | 1.47 | 1.26 | 1.6 | 1.4111 | 1.67 | 1.58 | 1.317 | 1.5 | 1.52 | 1.45 | 1.5 | 1.43 | 1.29 |  | 14
| 1 | 3 | 1.54 | 1.3867 | 1.72 | < 5 | 1.48 | 1.6501 | 1.47 | 1.58 | 1.52 | 1.37 | 1.54 | 1.28 | 1.52 | 1.5156 | 1.65 | 1.65 | 1.361 | 1.4 | 1.5 | 1.5 | 1.47 | 1.38 | 1.36 |  | 16
| 1 | 4 | 1.55 | 1.3387 | 1.71 | < 5 | 1.4 | 1.5476 | 1.49 | 1.5 | 1.54 | 1.38 | 1.52 | 1.32 | 1.48 | 1.3427 | 1.66 | 1.58 | 1.35 | 1.5 | 1.46 | 1.49 | 1.4 | 1.47 | 1.32 |  | 1.4849
| 1 | 5 | 1.57 | 1.4035 | 1.68 | < 5 | 1.45 | 1.6082 | 1.51 | 1.56 | 1.56 | 1.38 | 1.54 | 1.25 | 1.57 | 1.4901 | 1.67 | 1.54 | 1.37 | 1.4 | 1.51 | 1.53 | 1.43 | 1.43 | 1.28 |  | 113
| 1 | 6 | 1.56 | 1.3741 | 1.8 | < 5 | 1.46 | 1.4854 | 1.46 | 1.54 | 1.54 | 1.36 | 1.61 | 1.26 | 1.55 | 1.3297 | 1.63 | 1.51 | < 0.008 | 1.4 | 1.49 | 1.48 | 1.46 | 1.4 | 1.3
| Mean |  | 1.5533 | 1.3821 | 1.7067 |  | 1.4567 | 1.5962 | 1.495 | 1.56 | 1.5417 | 1.3733 | 1.5367 | 1.2833 | 1.5467 | 1.4261 | 1.6583 | 1.5817 | 1.3424 | 1.45 | 1.495 | 1.4967 | 1.4533 | 1.42 | 1.3117
| Median |  | 1.55 | 1.3862 | 1.705 |  | 1.455 | 1.6049 | 1.5 | 1.57 | 1.54 | 1.375 | 1.54 | 1.27 | 1.555 | 1.4391 | 1.665 | 1.58 | 1.35 | 1.45 | 1.495 | 1.495 | 1.46 | 1.42 | 1.31
| Std Dev. |  | 0.010328 | 0.024181 | 0.055737 |  | 0.033862 | 0.071352 | 0.026646 | 0.035777 | 0.013292 | 0.012111 | 0.045019 | 0.033862 | 0.041793 | 0.07781 | 0.016021 | 0.052694 | 0.025579 | 0.054772 | 0.020736 | 0.030768 | 0.034448 | 0.030984 | 0.028577
| Rel.Std.Dev. |  | 0.6649% | 1.75% | 3.266% |  | 2.325% | 4.47% | 1.782% | 2.293% | 0.8622% | 0.8818% | 2.93% | 2.639% | 2.702% | 5.456% | 0.9661% | 3.332% | 1.905% | 3.777% | 1.387% | 2.056% | 2.37% | 2.182% | 2.179%
| PDM3 |  | 4.612% | -6.918% | 14.94% |  | -1.898% | 7.496% | 0.6834% | 5.061% | 3.826% | -7.51% | 3.49% | -13.57% | 4.163% | -3.96% | 11.68% | 6.52% | -9.594% | -2.347% | 0.6834% | 0.7957% | -2.123% | -4.368% | -11.66%
| Z-Score (Absolute) |  | 0.57 | 1.1 | 2.07 | 9.83 | 0.37 | 0.99 | 0 | 0.64 | 0.46 | 1.19 | 0.41 | 2.07 | 0.51 | 0.67 | 1.6 | 0.85 | 3.67 | 0.44 | 0 | 0.02 | 0.41 | 0.73 | 1.79
| Analytical results for Tm in OREAS 163b (Indicative Value < 0.05 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab |  | 1
| No | No | 2.01: 2.02; 2.02: 2.06; 2.03: 2.12 |  | ppm
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS |  | 3
|  |  | 2.01: 0.5g; 2.02: 0.2g; 2.03: 0.5g |  | 3
| 1 | 1 | 2.01: < 0.05; 2.02: < 0.05; 2.03: < 0.1 |  | 1
| 1 | 2 | 2.01: < 0.05; 2.02: < 0.05; 2.03: < 0.1 |  | 6
| 1 | 3 | 2.01: < 0.05; 2.02: < 0.05; 2.03: < 0.1 |  | 16
| 1 | 4 | 2.01: < 0.05; 2.02: < 0.05; 2.03: < 0.1 |  | < 0.05
| 1 | 5 | 2.01: < 0.05; 2.02: < 0.05; 2.03: < 0.1 |  | 12
| 1 | 6 | 2.01: < 0.05; 2.02: < 0.05; 2.03: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.03
| Analytical results for U in OREAS 163b (Certified Value 0.44 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  | 3
| 1 | 1 | 0.45 | 0.3343 | 0.43 | < 5 | 0.46 | 0.49719 | 0.41 | 0.46 | 0.48 | 0.39 | 0.42 | 0.5 | 0.41 | 0.46607 | 0.5 | 0.43 | 0.44 | 0.4 | 0.43 | 0.49 | 0.41 | 0.41 | 0.4 |  | 1
| 1 | 2 | 0.46 | 0.3426 | 0.41 | < 5 | 0.444 | 0.51605 | 0.43 | 0.43 | 0.41 | 0.39 | 0.4 | 0.4 | 0.4 | 0.49727 | 0.5 | 0.44 | 0.45 | 0.4 | 0.44 | 0.45 | 0.4 | 0.4 | 0.4 |  | 16
| 1 | 3 | 0.44 | 0.3337 | 0.44 | < 5 | 0.433 | 0.48769 | 0.42 | 0.45 | 0.43 | 0.38 | 0.41 | 0.4 | 0.41 | 0.50181 | 0.49 | 0.41 | 0.46 | 0.4 | 0.43 | 0.5 | 0.42 | 0.39 | 0.4 |  | 16
| 1 | 4 | 0.45 | 0.3263 | 0.45 | < 5 | 0.426 | 0.46538 | 0.43 | 0.45 | 0.43 | 0.42 | 0.41 | 0.4 | 0.4 | 0.45034 | 0.49 | 0.43 | 0.46 | 0.5 | 0.39 | 0.53 | 0.4 | 0.41 | 0.4 |  | 0.43846
| 1 | 5 | 0.46 | 0.3223 | 0.43 | < 5 | 0.411 | 0.52622 | 0.44 | 0.43 | 0.43 | 0.4 | 0.42 | 0.4 | 0.41 | 0.45652 | 0.5 | 0.41 | 0.46 | 0.4 | 0.42 | 0.47 | 0.41 | 0.41 | 0.4 |  | 114
| 1 | 6 | 0.43 | 0.3307 | 0.48 | < 5 | 0.433 | 0.46064 | 0.4 | 0.43 | 0.43 | 0.38 | 0.41 | 0.4 | 0.42 | 0.45113 | 0.48 | 0.42 | < 0.05 | 0.4 | 0.43 | 0.49 | 0.4 | 0.43 | 0.4
| Mean |  | 0.44833 | 0.33165 | 0.44 |  | 0.4345 | 0.49219 | 0.42167 | 0.44167 | 0.435 | 0.39333 | 0.41167 | 0.41667 | 0.40833 | 0.47052 | 0.49333 | 0.42333 | 0.454 | 0.41667 | 0.42333 | 0.48833 | 0.40667 | 0.40833 | 0.4
| Median |  | 0.45 | 0.3322 | 0.435 |  | 0.433 | 0.49244 | 0.425 | 0.44 | 0.43 | 0.39 | 0.41 | 0.4 | 0.41 | 0.4613 | 0.495 | 0.425 | 0.46 | 0.4 | 0.43 | 0.49 | 0.405 | 0.41 | 0.4
| Std Dev. |  | 0.01169 | 0.0070438 | 0.023664 |  | 0.01655 | 0.026409 | 0.01472 | 0.013292 | 0.023452 | 0.015055 | 0.0075277 | 0.040825 | 0.0075277 | 0.02321 | 0.008165 | 0.012111 | 0.0089443 | 0.040825 | 0.017512 | 0.027142 | 0.008165 | 0.013292 | 6.0809e-17
| Rel.Std.Dev. |  | 2.608% | 2.124% | 5.378% |  | 3.809% | 5.365% | 3.491% | 3.009% | 5.391% | 3.828% | 1.829% | 9.798% | 1.844% | 4.933% | 1.655% | 2.861% | 1.97% | 9.798% | 4.137% | 5.558% | 2.008% | 3.255% | 1.52e-14%
| PDM3 |  | 2.253% | -24.36% | 0.352% |  | -0.9024% | 12.26% | -3.829% | 0.7321% | -0.7883% | -10.29% | -6.11% | -4.97% | -6.87% | 7.314% | 12.52% | -3.449% | 3.545% | -4.97% | -3.449% | 11.38% | -7.25% | -6.87% | -8.771%
| Z-Score (Absolute) |  | 0.64 | 3.19 | 0.36 | 67.85 | 0.18 | 2.07 | 0.24 | 0.42 | 0.2 | 1.17 | 0.57 | NA | 0.67 | 1.36 | 2.11 | 0.18 | 1.52 | NA | 0.18 | 1.95 | 0.73 | 0.67 | NA
Results from laboratories 2.12, 2.19 and 2.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 163b (Certified Value 5.01 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 6 | 5 | 6 | 4.8 | 4.8193 | 5 | 5 | 5 | 4 | 5 | 5 | 5 | 4.7274 | 6 | 6 | 5 | 5 | 5 | 4.9 | 7 | 5 | 5 |  |  | 1
| 1 | 2 | 6 | 5 | 6.9 | 4.6 | 5.0326 | 5 | 5 | 5 | 4 | 5 | 5 | 5 | 4.3417 | 5 | 6 | 4 | 5 | 5 | 4.9 | 7 | 6 | 5 |  |  | 36
| 1 | 3 | 6 | 5 | 7.3 | 4.7 | 4.9816 | 5 | 5 | 5 | 4 | 5 | 6 | 5 | 5.0474 | 6 | 6 | 4 | 5 | 5 | 5.1 | 6 | 5 | 5 |  |  | 16
| 1 | 4 | 6 | 5 | 6.5 | 4.8 | 4.852 | 4 | 5 | 5 | 4 | 5 | 6 | 4 | 4.2258 | 5 | 6 | 5 | 5 | 5 | 5 | 7 | 6 | 5 |  |  | 5.0128
| 1 | 5 | 6 | 5 | 7.2 | 4.7 | 5.0187 | 5 | 5 | 5 | 4 | 5 | 5 | 4 | 5.0231 | 5 | 6 | 4 | 5 | 6 | 4.8 | 6 | 5 | 5 |  |  | 115
| 1 | 6 | 6 | 5 | 6.5 | 4.6 | 4.6492 | 5 | 5 | 5 | 4 | 4 | 5 | 5 | 4.7212 | 5 | 6 | 4 | 5 | 5 | 5.2 | 7 | 6 | 5
| Mean |  | 6 | 5 | 6.7333 | 4.7 | 4.8922 | 4.8333 | 5 | 5 | 4 | 4.8333 | 5.3333 | 4.6667 | 4.6811 | 5.3333 | 6 | 4.3333 | 5 | 5.1667 | 4.9833 | 6.6667 | 5.5 | 5
| Median |  | 6 | 5 | 6.7 | 4.7 | 4.9168 | 5 | 5 | 5 | 4 | 5 | 5 | 5 | 4.7243 | 5 | 6 | 4 | 5 | 5 | 4.95 | 7 | 5.5 | 5
| Std Dev. |  | 0 | 0 | 0.49261 | 0.089443 | 0.14811 | 0.40825 | 0 | 0 | 0 | 0.40825 | 0.5164 | 0.5164 | 0.33982 | 0.5164 | 0 | 0.5164 | 0 | 0.40825 | 0.1472 | 0.5164 | 0.54772 | 0
| Rel.Std.Dev. |  | 0% | 0% | 7.316% | 1.903% | 3.027% | 8.447% | 0% | 0% | 0% | 8.447% | 9.682% | 11.07% | 7.259% | 9.682% | 0% | 11.92% | 0% | 7.902% | 2.954% | 7.746% | 9.959% | 0%
| PDM3 |  | 19.69% | -0.256% | 34.32% | -6.241% | -2.406% | -3.581% | -0.256% | -0.256% | -20.2% | -3.581% | 6.394% | -6.906% | -6.618% | 6.394% | 19.69% | -13.56% | -0.256% | 3.069% | -0.5885% | 32.99% | 9.718% | -0.256%
| Z-Score (Absolute) |  | 2.18 | 0 | 3.78 | 0.65 | 0.23 | 0.36 | 0 | 0 | 2.18 | 0.36 | 0.73 | 0.73 | 0.69 | 0.73 | 2.18 | 1.45 | 0 | 0.36 | 0.04 | 3.63 | 1.09 | 0
| Analytical results for W in OREAS 163b (Certified Value 1.57 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.19 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g |  |  | 2
| 1 | 1 | 1.6 | 0.58003 | 1.87 | < 5 | 1.67 | 1.4213 | 1.56 | 1.74 | 1.31 | 1.41 | 1.53 | 1.9 | 1.58 | 1.5787 | 1.55 | 1.7 | 1.4 | 2.65 | 1.6 | 1.55 | 1.56 | 1.2 |  |  | 1
| 1 | 2 | 1.5 | 0.66018 | 1.84 | < 5 | 1.62 | 1.5376 | 1.56 | 1.8 | 1.39 | 1.46 | 1.53 | 1.8 | 1.59 | 1.4759 | 1.2 | 1.75 | 1.4 | 2.43 | 1.6 | 1.7 | 1.67 | 1.5 |  |  | 17
| 1 | 3 | 1.5 | 0.58276 | 2.09 | < 5 | 1.58 | 1.4553 | 1.6 | 1.76 | 1.33 | 1.44 | 1.66 | 1.8 | 1.53 | 1.7093 | 1.02 | 1.59 | 1.3 | 2.3 | 1.6 | 1.5 | 1.54 | 1.2 |  |  | 16
| 1 | 4 | 1.5 | 0.64194 | 1.94 | < 5 | 1.69 | 1.3994 | 1.49 | 1.68 | 1.36 | 1.56 | 1.54 | 1.8 | 1.47 | 1.4186 | 1.03 | 1.68 | 1.3 | 2.35 | 1.6 | 1.65 | 1.56 | 1.4 |  |  | 1.5715
| 1 | 5 | 1.6 | 0.65972 | 1.92 | < 5 | 1.87 | 1.5628 | 1.5 | 1.74 | 1.3 | 1.44 | 1.56 | 1.8 | 1.55 | 1.5796 | 0.98 | 1.6 | 1.3 | 2.51 | 1.6 | 1.5 | 1.61 | 1.2 |  |  | 116
| 1 | 6 | 1.5 | 0.65885 | 2.03 | < 5 | 1.91 | 1.4499 | 1.44 | 1.92 | 1.94 | 1.47 | 1.54 | 1.7 | 1.56 | 1.7058 | 0.97 | 1.68 | 1.3 | 2.37 | 1.7 | 1.65 | 1.53 | 1.2
| Mean |  | 1.5333 | 0.63058 | 1.9483 |  | 1.7233 | 1.4711 | 1.525 | 1.7733 | 1.4383 | 1.4633 | 1.56 | 1.8 | 1.5467 | 1.578 | 1.125 | 1.6667 | 1.3333 | 2.435 | 1.6167 | 1.5917 | 1.5783 | 1.2833
| Median |  | 1.5 | 0.65039 | 1.93 |  | 1.68 | 1.4526 | 1.53 | 1.75 | 1.345 | 1.45 | 1.54 | 1.8 | 1.555 | 1.5792 | 1.025 | 1.68 | 1.3 | 2.4 | 1.6 | 1.6 | 1.56 | 1.2
| Std Dev. |  | 0.05164 | 0.038718 | 0.095376 |  | 0.1353 | 0.065043 | 0.058566 | 0.08165 | 0.24799 | 0.05164 | 0.0502 | 0.063246 | 0.043205 | 0.11782 | 0.22421 | 0.06121 | 0.05164 | 0.12771 | 0.040825 | 0.08612 | 0.052694 | 0.13292
| Rel.Std.Dev. |  | 3.368% | 6.14% | 4.895% |  | 7.851% | 4.421% | 3.84% | 4.604% | 17.24% | 3.529% | 3.218% | 3.514% | 2.793% | 7.467% | 19.93% | 3.673% | 3.873% | 5.245% | 2.525% | 5.411% | 3.339% | 10.36%
| PDM3 |  | -2.429% | -59.87% | 23.98% |  | 9.662% | -6.391% | -2.959% | 12.84% | -8.474% | -6.883% | -0.7319% | 14.54% | -1.58% | 0.4125% | -28.41% | 6.056% | -15.16% | 54.95% | 2.874% | 1.283% | 0.4347% | -18.34%
| Z-Score (Absolute) |  | 0.2 | 5.36 | 2.16 | 5.31 | 0.88 | 0.56 | 0.25 | 1.17 | 0.75 | 0.6 | 0.05 | 1.32 | 0.13 | 0.05 | 2.53 | 0.56 | 1.34 | 4.94 | 0.27 | 0.13 | 0.05 | 1.63
| Analytical results for Y in OREAS 163b (Certified Value 2.91 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [5 empty] | 3
| 1 | 1 | 2.81 | 2.679 | < 5 | 3 | 2.8971 | 2.9 | 3.24 | 3.18 | 2.95 | 2.99 | 2.58 | 3.02 | 2.8222 | 3.88 | 2.92 | 3.09 | 2.7 | 2.92 | 2.5 | [5 empty] | 1
| 1 | 2 | 2.83 | 2.7639 | < 5 | 2.88 | 2.942 | 2.91 | 3.18 | 3.07 | 3.01 | 2.9 | 2.4 | 2.86 | 2.6483 | 4.03 | 2.84 | 3.02 | 2.6 | 3.01 | 2.5 | [5 empty] | 18
| 1 | 3 | 2.8 | 2.6978 | < 5 | 3.01 | 2.9023 | 2.79 | 3.18 | 3.16 | 2.65 | 2.95 | 2.45 | 2.97 | 2.8985 | 3.96 | 2.96 | 2.98 | 2.7 | 2.91 | 2.5 | [5 empty] | 16
| 1 | 4 | 2.86 | 2.7562 | < 5 | 2.94 | 2.8814 | 2.87 | 3.04 | 3.21 | 2.73 | 2.95 | 2.52 | 2.84 | 2.6883 | 3.91 | 2.86 | 3.06 | 2.5 | 3.11 | 2.5 | [5 empty] | 2.9072
| 1 | 5 | 2.8 | 2.7766 | < 5 | 2.91 | 2.9641 | 2.91 | 3.23 | 3.18 | 2.95 | 2.97 | 2.54 | 2.98 | 2.8049 | 3.94 | 2.86 | 3.15 | 2.5 | 2.95 | 2.5 | [5 empty] | 117
| 1 | 6 | 2.8 | 2.7111 | < 5 | 2.93 | 2.8405 | 2.75 | 3.13 | 3.1 | 2.98 | 3.09 | 2.39 | 2.92 | 2.8476 | 3.78 | 2.88 | 3.05 | 2.6 | 2.95 | 2.5
| Mean |  | 2.8167 | 2.7308 |  | 2.945 | 2.9045 | 2.855 | 3.1667 | 3.15 | 2.8783 | 2.975 | 2.48 | 2.9317 | 2.785 | 3.9167 | 2.8867 | 3.0583 | 2.6 | 2.975 | 2.5
| Median |  | 2.805 | 2.7336 |  | 2.935 | 2.8997 | 2.885 | 3.18 | 3.17 | 2.95 | 2.96 | 2.485 | 2.945 | 2.8135 | 3.925 | 2.87 | 3.055 | 2.6 | 2.95 | 2.5
| Std Dev. |  | 0.024221 | 0.039985 |  | 0.050892 | 0.043938 | 0.068629 | 0.073666 | 0.053666 | 0.14972 | 0.063797 | 0.07823 | 0.07111 | 0.096562 | 0.084063 | 0.045019 | 0.058452 | 0.089443 | 0.074766 | 0
| Rel.Std.Dev. |  | 0.8599% | 1.464% |  | 1.728% | 1.513% | 2.404% | 2.326% | 1.704% | 5.202% | 2.144% | 3.154% | 2.426% | 3.467% | 2.146% | 1.56% | 1.911% | 3.44% | 2.513% | 0%
| PDM3 |  | -3.115% | -6.07% |  | 1.299% | -0.09282% | -1.797% | 8.923% | 8.35% | -0.9943% | 2.331% | -14.7% | 0.8402% | -4.206% | 34.72% | -0.7077% | 5.197% | -10.57% | 2.331% | -14.01%
| Z-Score (Absolute) |  | 0.46 | 1.03 | 2.56 | 0.39 | 0.12 | 0.21 | 1.86 | 1.75 | 0.06 | 0.59 | 2.7 | 0.3 | 0.67 | 6.83 | 0 | 1.14 | 1.9 | 0.59 | 2.56
| Analytical results for Yb in OREAS 163b (Certified Value 0.16 ppm) | [25 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | [5 empty] | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.21; 2.08: 2.22 | [5 empty] | ppm
|  |  | 2.01: AR*MS; 2.02: AR*OES; 2.04: AR*MS; 2.05: AR*OES/MS; 2.06: AR*MS; 2.07: AR*OES/MS; 2.08: AR*OES/MS | [5 empty] | 2
|  |  | 2.01: 0.5g; 2.02: 0.25g; 2.04: 0.2g; 2.05: 0.5g; 2.06: 0.25g; 2.07: 0.25g; 2.08: 0.2g | [5 empty] | 2
| 1 | 1 | 2.01: 0.142; 2.02: < 5; 2.04: 0.1673; 2.05: 0.1; 2.06: 0.2; 2.07: 0.2; 2.08: 0.1 | [5 empty] | 1
| 1 | 2 | 2.01: 0.1408; 2.02: < 5; 2.04: 0.17772; 2.05: 0.1; 2.06: 0.2; 2.07: 0.2; 2.08: 0.1 | [5 empty] | 19
| 1 | 3 | 2.01: 0.1333; 2.02: < 5; 2.04: 0.16895; 2.05: 0.2; 2.06: 0.2; 2.07: 0.2; 2.08: 0.2 | [5 empty] | 16
| 1 | 4 | 2.01: 0.136; 2.02: < 5; 2.04: 0.15308; 2.05: 0.1; 2.06: 0.2; 2.07: 0.2; 2.08: 0.2 | [5 empty] | 0.16426
| 1 | 5 | 2.01: 0.1389; 2.02: < 5; 2.04: 0.16314; 2.05: 0.2; 2.06: 0.2; 2.07: 0.2; 2.08: 0.1 | [5 empty] | 118
| 1 | 6 | 2.01: 0.1395; 2.02: < 5; 2.04: 0.15262; 2.05: 0.1; 2.06: 0.2; 2.07: 0.2; 2.08: 0.2
| Mean |  | 2.01: 0.13842; 2.04: 0.1638; 2.05: 0.13333; 2.06: 0.2; 2.07: 0.2; 2.08: 0.15
| Median |  | 2.01: 0.1392; 2.04: 0.16522; 2.05: 0.1; 2.06: 0.2; 2.07: 0.2; 2.08: 0.15
| Std Dev. |  | 2.01: 0.0032233; 2.04: 0.0097238; 2.05: 0.05164; 2.06: 3.0405e-17; 2.07: 3.0405e-17; 2.08: 0.054772
| Rel.Std.Dev. |  | 2.01: 2.329%; 2.04: 5.936%; 2.05: 38.73%; 2.06: 1.52e-14%; 2.07: 1.52e-14%; 2.08: 36.51%
| PDM3 |  | 2.01: -15.73%; 2.04: -0.2785%; 2.05: -18.83%; 2.06: 21.76%; 2.07: 21.76%; 2.08: -8.681%
| Z-Score (Absolute) |  | 2.01: 0.56; 2.02: 51.7; 2.04: 0; 2.05: 0.67; 2.06: 0.8; 2.07: 0.8; 2.08: 0.31
| Analytical results for Zn in OREAS 163b (Certified Value 24.5 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | 1
|  |  | 01g | 0.5g | 0.5g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.25g | 2
| 1 | 1 | 25 | 25.53 | 21.5 | 53.6 | 25.7 | 23.4936 | 24 | 25 | 25 | 24 | 24 | 25 | 25 | 9.1791 | 28 | 23.3 | 25 | 25 | 25 | 24 | 29.8 | 26 | 23 | 24 | 1
| 1 | 2 | 25 | 25.37 | 21.7 | 57.1 | 25.4 | 24.2734 | 24 | 26 | 24 | 24 | 23 | 28.4 | 25 | 7.9903 | 24 | 23 | 25 | 25 | 26 | 23 | 26.4 | 26 | 23 | 24 | 40
| 1 | 3 | 25 | 30.43 | 21.6 | 55.8 | 25.1 | 23.674 | 24 | 25 | 25 | 24 | 23 | 37.2 | 25 | 9.2998 | 23 | 23 | 26 | 24 | 26 | 24 | 25.4 | 26 | 23 | 24 | 16
| 1 | 4 | 25 | 25.65 | 21.6 | 60 | 24.4 | 23.3842 | 24 | 25 | 24 | 24 | 23 | 25.5 | 24 | 7.6519 | 25 | 23 | 26 | 24 | 25 | 24 | 24.9 | 26 | 23 | 25 | 24.5001
| 1 | 5 | 25 | 27.66 | 20.8 | 57 | 25 | 23.9134 | 24 | 26 | 25 | 24 | 23 | 26.6 | 24 | 7.7712 | 26 | 22.7 | 25 | 24 | 25 | 23 | 25.2 | 26 | 24 | 24 | 119
| 1 | 6 | 25 | 30.28 | 22.4 | 52.2 | 25.4 | 23.3237 | 25 | 25 | 24 | 24 | 24 | 23.7 | 25 | 8.5417 | 24 | 22.6 | 26 | 24 | 26 | 24 | 25.1 | 26 | 23 | 24
| Mean |  | 25 | 27.4867 | 21.6 | 55.95 | 25.1667 | 23.6771 | 24.1667 | 25.3333 | 24.5 | 24 | 23.3333 | 27.7333 | 24.6667 | 8.4057 | 25 | 22.9333 | 25.5 | 24.3333 | 25.5 | 23.6667 | 26.1333 | 26 | 23.1667 | 24.1667
| Median |  | 25 | 26.655 | 21.6 | 56.4 | 25.25 | 23.5838 | 24 | 25 | 24.5 | 24 | 23 | 26.05 | 25 | 8.266 | 24.5 | 23 | 25.5 | 24 | 25.5 | 24 | 25.3 | 26 | 23 | 24
| Std Dev. |  | 0 | 2.3739 | 0.5099 | 2.7725 | 0.45019 | 0.36215 | 0.40825 | 0.5164 | 0.54772 | 0 | 0.5164 | 4.9005 | 0.5164 | 0.71545 | 1.7889 | 0.25033 | 0.54772 | 0.5164 | 0.54772 | 0.5164 | 1.8715 | 0 | 0.40825 | 0.40825
| Rel.Std.Dev. |  | 0% | 8.637% | 2.361% | 4.955% | 1.789% | 1.53% | 1.689% | 2.038% | 2.236% | 0% | 2.213% | 17.67% | 2.094% | 8.512% | 7.155% | 1.092% | 2.148% | 2.122% | 2.148% | 2.182% | 7.162% | 0% | 1.762% | 1.689%
| PDM3 |  | 2.04% | 12.19% | -11.84% | 128.4% | 2.721% | -3.359% | -1.361% | 3.401% | -0.0005362% | -2.041% | -4.762% | 13.2% | 0.6797% | -65.69% | 2.04% | -6.395% | 4.081% | -0.6808% | 4.081% | -3.402% | 6.666% | 6.122% | -5.443% | -1.361%
| Z-Score (Absolute) |  | 0.31 | 2.14 | 2.19 | 23.07 | 0.43 | 0.67 | 0.31 | 0.55 | 0.06 | 0.43 | 0.92 | 2.32 | 0.06 | 11.9 | 0.31 | 1.21 | 0.67 | 0.18 | 0.67 | 0.67 | 1.14 | 1.04 | 1.04 | 0.31
| Analytical results for Zr in OREAS 163b (Certified Value 1.56 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.20 | 2.21 | 2.22 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.25g | [6 empty] | 2
| 1 | 1 | 1.7 | < 1 | 1.62 | 1.4088 | 1.3 | 1.4 | 1.6 | 1.4 | 1.4 | 1.7 | 1.5 | 1.3849 | 3.4 | < 5 | 1.7 | 1.6 | 1.6 | 2 | [6 empty] | 1
| 1 | 2 | 1.7 | < 1 | 1.61 | 1.4241 | 1.3 | 1.4 | 1.5 | 1.4 | 1.3 | 1.5 | 1.5 | 1.3763 | 2.9 | < 5 | 1.7 | 2.8 | 1.6 | 1.9 | [6 empty] | 41
| 1 | 3 | 1.8 | < 1 | 1.6 | 1.4324 | 1.2 | 1.4 | 1.6 | 1.3 | 1.4 | 1.4 | 1.5 | 1.5492 | 2.4 | < 5 | 1.9 | 1.7 | 1.6 | 1.8 | [6 empty] | 16
| 1 | 4 | 1.8 | < 1 | 1.6 | 1.4025 | 1.3 | 1.3 | 1.6 | 1.4 | 1.4 | 1.5 | 1.5 | 1.2855 | 2.5 | < 5 | 1.7 | 2 | 1.6 | 1.8 | [6 empty] | 1.5564
| 1 | 5 | 1.8 | < 1 | 1.58 | 1.4479 | 1.3 | 1.4 | 1.6 | 1.4 | 1.4 | 1.5 | 1.6 | 1.6121 | 2.5 | < 5 | 1.9 | 1.9 | 1.6 | 2 | [6 empty] | 120
| 1 | 6 | 1.8 | < 1 | 1.61 | 1.4038 | 1.2 | 1.4 | 1.6 | 1.4 | 1.3 | 1.5 | 1.5 | 1.628 | 2.4 | < 5 | 1.8 | 1.8 | 1.6 | 1.8
| Mean |  | 1.7667 |  | 1.6033 | 1.4199 | 1.2667 | 1.3833 | 1.5833 | 1.3833 | 1.3667 | 1.5167 | 1.5167 | 1.4727 | 2.6833 |  | 1.7833 | 1.9667 | 1.6 | 1.8833
| Median |  | 1.8 |  | 1.605 | 1.4164 | 1.3 | 1.4 | 1.6 | 1.4 | 1.4 | 1.5 | 1.5 | 1.467 | 2.5 |  | 1.75 | 1.85 | 1.6 | 1.85
| Std Dev. |  | 0.05164 |  | 0.013663 | 0.018147 | 0.05164 | 0.040825 | 0.040825 | 0.040825 | 0.05164 | 0.098319 | 0.040825 | 0.14243 | 0.39707 |  | 0.098319 | 0.43205 | 2.4324e-16 | 0.098319
| Rel.Std.Dev. |  | 2.923% |  | 0.8521% | 1.278% | 4.077% | 2.951% | 2.578% | 2.951% | 3.779% | 6.483% | 2.692% | 9.671% | 14.8% |  | 5.513% | 21.97% | 1.52e-14% | 5.22%
| PDM3 |  | 13.51% |  | 3.016% | -8.768% | -18.62% | -11.12% | 1.731% | -11.12% | -12.19% | -2.553% | -2.553% | -5.38% | 72.41% |  | 14.58% | 26.36% | 2.802% | 21.01%
| Z-Score (Absolute) |  | 0.83 | 4.05 | 0.21 | 0.5 | 1.09 | 0.64 | 0.13 | 0.64 | 0.71 | 0.13 | 0.13 | 0.3 | 4.37 | 3.66 | 0.9 | 1.61 | 0.19 | 1.28
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C784: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C820: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C838: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C857: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C987: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1059: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1077: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1132: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Ag in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ag | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 8
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 14
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Al2O3 in OREAS 163b (Certified Value 0.699 wt.%) | [17 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.67 | 0.672 | 0.71 | 1.05 | 0.73955 | 0.7 | 0.72 | 0.68 | 0.73 | 1.08 | 0.48 | 0.71 | 0.227 | 0.67 |  |  | 1
| 1 | 2 | 0.64 | 0.699 | 0.71 | 0.76 | 0.74266 | 0.7 | 0.7 | 0.71 | 0.71 | 1.05 | 0.48 | 0.71 | 0.208 | 0.68 |  |  | 0
| 1 | 3 | 0.63 | 0.665 | 0.71 | 0.68 | 0.73693 | 0.71 | 0.68 | 0.71 | 0.73 | 0.95 | 0.49 | 0.71 | 0.208 | 0.66 |  |  | 16
| 1 | 4 | 0.65 | 0.645 | 0.73 | 0.58 | 0.74107 | 0.69 | 0.69 | 0.71 | 0.7 | 1.12 | 0.45 | 0.72 | 0.208 | 0.65 |  |  | 0.69869
| 1 | 5 | 0.65 | 0.682 | 0.71 | 0.85 | 0.73816 | 0.69 | 0.71 | 0.7 | 0.72 | 0.94 | 0.43 | 0.72 | 0.208 | 0.68 |  |  | 122
| 1 | 6 | 0.64 | 0.677 | 0.72 | 0.9 | 0.73037 | 0.69 | 0.69 | 0.7 | 0.73 | 1.09 | 0.45 | 0.72 | 0.208 | 0.64
| Mean |  | 0.64667 | 0.67333 | 0.715 | 0.80333 | 0.73812 | 0.69667 | 0.69833 | 0.70167 | 0.72 | 1.0383 | 0.46333 | 0.715 | 0.21117 | 0.66333
| Median |  | 0.645 | 0.6745 | 0.71 | 0.805 | 0.73885 | 0.695 | 0.695 | 0.705 | 0.725 | 1.065 | 0.465 | 0.715 | 0.208 | 0.665
| Std Dev. |  | 0.013663 | 0.018007 | 0.0083666 | 0.16693 | 0.0043098 | 0.008165 | 0.01472 | 0.01169 | 0.012649 | 0.075741 | 0.023381 | 0.0054772 | 0.0077567 | 0.01633
| Rel.Std.Dev. |  | 2.113% | 2.674% | 1.17% | 20.78% | 0.5839% | 1.172% | 2.108% | 1.666% | 1.757% | 7.294% | 5.046% | 0.766% | 3.673% | 2.462%
| PDM3 |  | -7.446% | -3.63% | 2.334% | 14.98% | 5.643% | -0.29% | -0.05148% | 0.4256% | 3.05% | 48.61% | -33.69% | 2.334% | -69.78% | -5.061%
| Z-Score (Absolute) |  | 1.14 | 0.57 | 0.32 | 2.2 | 0.81 | 0.07 | 0.04 | 0.04 | 0.43 | 7.2 | 5.04 | 0.32 | 10.41 | 0.78
| Analytical results for As in OREAS 163b (Indicative Value 186 ppm) | [17 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.19 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | 300 | 200 | 100 | 300 | N.A. | < 10 | 179 | [8 empty] | 1
| 1 | 2 | < 100 | 300 | 200 | 100 | 300 | N.A. | < 10 | 172 | [8 empty] | 9
| 1 | 3 | < 100 | 300 | 100 | 100 | 300 | N.A. | < 10 | 178 | [8 empty] | 16
| 1 | 4 | < 100 | 300 | 100 | 200 | 300 | N.A. | < 10 | 176 | [8 empty] | 186
| 1 | 5 | < 100 | 200 | 200 | 200 | 300 | N.A. | < 10 | 182 | [8 empty] | 15
| 1 | 6 | < 100 | 200 | 100 | 200 | 300 | N.A. | < 10 | 177
| Mean |  |  | 266.6667 | 150 | 150 | 300 |  |  | 177.3333
| Median |  |  | 300 | 150 | 150 | 300 |  |  | 177.5
| Std Dev. |  |  | 51.6398 | 54.7723 | 54.7723 | 0 |  |  | 3.3267
| Rel.Std.Dev. |  |  | 19.36% | 36.51% | 36.51% | 0% |  |  | 1.876%
| PDM3 |  |  | 43.37% | -19.35% | -19.35% | 61.29% |  |  | -4.659%
| Z-Score (Absolute) |  | 0.67 | 0.79 | 0 | 0 | 1.01 | NA | 0.98 | 0.18
| Analytical results for Au in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Au | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 2
2
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 10
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 16
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 163b (Certified Value 64 ppm) | [17 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 | 2.25 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 110 | < 100 | < 50 | < 100 | < 100 | < 100 | 100 |  | 100 | < 110 | 893 | < 100 | 59.2 |  |  |  | 1
| 1 | 2 | < 110 | < 100 | 47.8289 | < 100 | < 100 | 100 | 100 |  | < 100 | < 110 | 893 | < 100 | 63.6 |  |  |  | 0
| 1 | 3 | < 110 | < 100 | < 50 | < 100 | < 100 | < 100 | 100 |  | < 100 | < 110 | 893 | 100 | 67 |  |  |  | 16
| 1 | 4 | < 110 | < 100 | 66.9926 | < 100 | < 100 | 100 | < 100 | 110 | < 100 | < 110 | 782 | < 100 | 56.9 |  |  |  | 63.997
| 1 | 5 | 112 | < 100 | 47.8533 | < 100 | < 100 | < 100 | 100 |  | < 100 | < 110 | 782 | < 100 | 63.6 |  |  |  | 123
| 1 | 6 | < 110 | < 100 | 57.196 | < 100 | < 100 | 100 | 100 |  | < 100 | 110 | 893 | < 100 | 65.9
| Mean |  | 112 |  | 54.9677 |  |  | 100 | 100 | 110 | 100 | 110 | 856 | 100 | 62.7
| Median |  | 112 |  | 52.5247 |  |  | 100 | 100 | 110 | 100 | 110 | 893 | 100 | 63.6
| Std Dev. |  |  |  | 9.1495 |  |  | 0 | 0 |  |  |  | 57.3202 |  | 3.9049
| Rel.Std.Dev. |  |  |  | 16.65% |  |  | 0% | 0% |  |  |  | 6.696% |  | 6.228%
| PDM3 |  | 75.01% |  | -14.11% |  |  | 56.26% | 56.26% | 71.88% | 56.26% | 71.88% | 1238% | 56.26% | -2.027%
| Z-Score (Absolute) |  | 0.13 | 0.67 | 0.94 | 0.67 | 0.67 | 0.65 | 1.54 | 2.6 | 0.23 | 0.13 | 42.09 | 0.23 | 0
| Analytical results for Bi in OREAS 163b (Certified Value < 100 ppm) | [17 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 100 | < 100 | < 100 | 100 | < 10 | 36 | [9 empty] | 1
| 1 | 2 | < 100 | < 100 | < 100 | < 100 | 100 | < 10 | 36 | [9 empty] | 29
| 1 | 3 | < 100 | < 100 | < 100 | < 100 | 100 | < 10 | 37 | [9 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | < 100 | 100 | < 10 | 35 | [9 empty] | < 100
| 1 | 5 | < 100 | < 100 | < 100 | < 100 | 100 | < 10 | 37 | [9 empty] | 124
| 1 | 6 | < 100 | < 100 | < 100 | < 100 | < 100 | < 10 | 37
| Mean | [5 empty] | 100 |  | 36.3333
| Median | [5 empty] | 100 |  | 36.5
| Std Dev. | [5 empty] | 0 |  | 0.8165
| Rel.Std.Dev. | [5 empty] | 0% |  | 2.247%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for CaO in OREAS 163b (Certified Value 3.69 wt.%) | [17 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 3.61 | 3.707 | 3.84 | 4.1 | 3.6544 | 3.66 | 3.6 | 3.64 | 3.69 | 3.53 | 3.69 | 3.68 | 3.7638 | 3.71 |  |  | 1
| 1 | 2 | 3.63 | 3.687 | 3.81 | 4.29 | 3.7026 | 3.68 | 3.58 | 3.66 | 3.69 | 3.49 | 3.73 | 3.67 | 3.7499 | 3.76 |  |  | 0
| 1 | 3 | 3.62 | 3.717 | 3.78 | 4.59 | 3.6559 | 3.65 | 3.55 | 3.64 | 3.71 | 3.52 | 3.77 | 3.65 | 3.8198 | 3.72 |  |  | 16
| 1 | 4 | 3.61 | 3.675 | 3.83 | 3.88 | 3.686 | 3.66 | 3.57 | 3.69 | 3.68 | 3.44 | 3.72 | 3.72 | 3.7219 | 3.72 |  |  | 3.6928
| 1 | 5 | 3.61 | 3.75 | 3.85 | 4.44 | 3.661 | 3.66 | 3.61 | 3.67 | 3.72 | 3.59 | 3.72 | 3.66 | 3.7778 | 3.72 |  |  | 125
| 1 | 6 | 3.62 | 3.698 | 3.85 | 4.66 | 3.6195 | 3.65 | 3.55 | 3.71 | 3.67 | 3.51 | 3.74 | 3.67 | 3.8058 | 3.73
| Mean |  | 3.6167 | 3.7057 | 3.8267 | 4.3267 | 3.6633 | 3.66 | 3.5767 | 3.6683 | 3.6933 | 3.5133 | 3.7283 | 3.675 | 3.7732 | 3.7267
| Median |  | 3.615 | 3.7025 | 3.835 | 4.365 | 3.6585 | 3.66 | 3.575 | 3.665 | 3.69 | 3.515 | 3.725 | 3.67 | 3.7708 | 3.72
| Std Dev. |  | 0.008165 | 0.026242 | 0.027325 | 0.29851 | 0.028704 | 0.010954 | 0.025033 | 0.027869 | 0.018619 | 0.049261 | 0.026394 | 0.02429 | 0.036107 | 0.017512
| Rel.Std.Dev. |  | 0.2258% | 0.7082% | 0.7141% | 6.899% | 0.7836% | 0.2993% | 0.6999% | 0.7597% | 0.5041% | 1.402% | 0.7079% | 0.661% | 0.9569% | 0.4699%
| PDM3 |  | -2.062% | 0.3481% | 3.625% | 17.16% | -0.8006% | -0.8886% | -3.145% | -0.6629% | 0.01408% | -4.86% | 0.9619% | -0.4824% | 2.176% | 0.9167%
| Z-Score (Absolute) |  | 1.05 | 0.33 | 2.22 | 10 | 0.33 | 0.38 | 1.67 | 0.25 | 0.14 | 2.66 | 0.69 | 0.14 | 1.39 | 0.66
| Analytical results for Cd in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.20 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 11
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 17
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ce in OREAS 163b (Certified Value < 80 ppm) | [17 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 80 | < 80 | < 80 | < 80 | < 10 | [11 empty] | 1
| 1 | 2 | < 80 | < 80 | < 80 | 80 | < 10 | [11 empty] | 31
| 1 | 3 | < 80 | < 80 | < 80 | < 80 | < 10 | [11 empty] | 16
| 1 | 4 | < 80 | < 80 | < 80 | 80 | < 10 | [11 empty] | < 80
| 1 | 5 | < 80 | < 80 | < 80 | 80 | < 10 | [11 empty] | 126
| 1 | 6 | < 80 | < 80 | < 80 | < 80 | < 10
| Mean |  |  |  |  | 80
| Median |  |  |  |  | 80
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Cl in OREAS 163b (Indicative Value 317 ppm) | [17 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.20 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 0
0
| 1 | 1 | 530 | 78 | [14 empty] | 1
| 1 | 2 | 550 | 92 | [14 empty] | 12
| 1 | 3 | 540 | 85 | [14 empty] | 16
| 1 | 4 | 550 | 80 | [14 empty] | 316.6667
| 1 | 5 | 570 | 80 | [14 empty] | 18
| 1 | 6 | 540 | 105
| Mean |  | 546.6667 | 86.6667
| Median |  | 545 | 82.5
| Std Dev. |  | 13.6626 | 10.3086
| Rel.Std.Dev. |  | 2.499% | 11.89%
| PDM3 |  | 72.63% | -72.63%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 163b (Certified Value 236 ppm) | [17 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.19 | 2.20 | 2.25 | [7 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | < 100 | 200 | 300 | 200 | 200 | 264 | N.A. | < 50 | 290 | [7 empty] | 1
| 1 | 2 | < 100 | 200 | 200 | 200 | 200 | 257 | N.A. | < 50 | 287 | [7 empty] | 32
| 1 | 3 | < 100 | 200 | 200 | 200 | 300 | 279 | N.A. | < 50 | 290 | [7 empty] | 16
| 1 | 4 | 200 | 200 | 200 | 200 | 300 | 257 | N.A. | < 50 | 290 | [7 empty] | 236.3398
| 1 | 5 | < 100 | 200 | 300 | 200 | 300 | 264 | N.A. | < 50 | 288 | [7 empty] | 127
| 1 | 6 | < 100 | 200 | 200 | 200 | 300 | 272 | N.A. | < 50 | 288
| Mean |  | 200 | 200 | 233.3333 | 200 | 266.6667 | 265.5 |  |  | 288.8333
| Median |  | 200 | 200 | 200 | 200 | 300 | 264 |  |  | 289
| Std Dev. |  |  | 0 | 51.6398 | 0 | 51.6398 | 8.6429 |  |  | 1.3292
| Rel.Std.Dev. |  |  | 0% | 22.13% | 0% | 19.36% | 3.255% |  |  | 0.4602%
| PDM3 |  | -15.38% | -15.38% | -1.272% | -15.38% | 12.83% | 12.34% |  |  | 22.21%
| Z-Score (Absolute) |  | 1.93 | 0.23 | 0.23 | 0.23 | 0.68 | 0.67 | NA | 2.61 | 0.98
| Analytical results for Cr2O3 in OREAS 163b (Certified Value < 100 ppm) | [17 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 | 2.25 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 50 | < 150 | < 150 | 150 | < 150 | < 100 | < 50 | < 100 | < 15 | < 100 | 19 |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 50 | < 150 | < 150 | 150 | < 150 | < 100 | < 50 | < 100 | < 15 | < 100 | 21.9 |  |  |  | 0
| 1 | 3 | < 100 | < 100 | < 50 | < 150 | < 150 | 150 | < 150 | 100 | < 50 | < 100 | < 15 | < 100 | 20.5 |  |  |  | 16
| 1 | 4 | < 100 | < 100 | 270.6772 | < 150 | < 150 | 150 | < 150 | 100 | < 50 | < 100 | < 15 | < 100 | 21.9 |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | 67.6713 | < 150 | < 150 | 150 | < 150 | < 100 | < 50 | < 100 | < 15 | < 100 | 21.9 |  |  |  | 128
| 1 | 6 | < 100 | < 100 | 115.5476 | < 150 | < 150 | 150 | < 150 | 100 | < 50 | < 100 | < 15 | < 100 | 21.9
| Mean |  |  |  | 151.2987 |  |  | 150 |  | 100 |  |  |  |  | 21.1833
| Median |  |  |  | 115.5476 |  |  | 150 |  | 100 |  |  |  |  | 21.9
| Std Dev. |  |  |  | 106.12 |  |  | 0 |  | 0 |  |  |  |  | 1.2073
| Rel.Std.Dev. |  |  |  | 70.14% |  |  | 0% |  | 0% |  |  |  |  | 5.699%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 1.11 | 0.67 | 0.67 | 2.81 | 0.67 | 0.73 | 0.73 | 0 | 1.25 | 0 | 0.84
| Analytical results for Cs in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Cs | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 13
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 19
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Cu in OREAS 163b (Certified Value 1.68 wt.%) | [17 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.19 | 2.20 | 2.25 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  |  |  | 2
3
| 1 | 1 | 1.64 | 1.7795 | 1.675 | 1.705 | 1.675 | 1.665 | 1.6936 | 1.65 | 1.65 | 1.74 | 1.68 | 1.3666 |  |  |  |  | 1
| 1 | 2 | 1.65 | 1.7838 | 1.66 | 1.675 | 1.655 | 1.67 | 1.7575 | 1.64 | 1.67 | 1.7 | 1.65 | 1.362 |  |  |  |  | 35
| 1 | 3 | 1.63 | 1.7825 | 1.675 | 1.665 | 1.65 | 1.685 | 1.7255 | 1.65 | 1.66 | 1.72 | 1.68 | 1.3607 |  |  |  |  | 16
| 1 | 4 | 1.59 | 1.7893 | 1.675 | 1.685 | 1.71 | 1.68 | 1.7335 | 1.65 | 1.67 | 1.68 | 1.67 | 1.3705 |  |  |  |  | 1.6751
| 1 | 5 | 1.61 | 1.7762 | 1.66 | 1.7 | 1.685 | 1.695 | 1.7095 | 1.63 | 1.67 | 1.72 | 1.69 | 1.3647 |  |  |  |  | 129
| 1 | 6 | 1.66 | 1.7607 | 1.65 | 1.66 | 1.71 | 1.645 | 1.6936 | 1.66 | 1.67 | 1.74 | 1.66 | 1.3527
| Mean |  | 1.63 | 1.7787 | 1.6658 | 1.6817 | 1.6808 | 1.6733 | 1.7189 | 1.6467 | 1.665 | 1.7167 | 1.6717 | 1.3629
| Median |  | 1.635 | 1.781 | 1.6675 | 1.68 | 1.68 | 1.675 | 1.7175 | 1.65 | 1.67 | 1.72 | 1.675 | 1.3634
| Std Dev. |  | 0.026077 | 0.0098294 | 0.010685 | 0.018348 | 0.025965 | 0.017512 | 0.024961 | 0.010328 | 0.0083666 | 0.023381 | 0.01472 | 0.0060725
| Rel.Std.Dev. |  | 1.6% | 0.5526% | 0.6414% | 1.091% | 1.545% | 1.047% | 1.452% | 0.6272% | 0.5025% | 1.362% | 0.8805% | 0.4456%
| PDM3 |  | -2.69% | 6.187% | -0.5504% | 0.3949% | 0.3451% | -0.1026% | 2.616% | -1.695% | -0.6001% | 2.484% | -0.2021% | -18.64%
| Z-Score (Absolute) |  | 1.64 | 4.09 | 0.26 | 0.35 | 0.32 | 0.03 | 1.79 | 1 | 0.29 | 1.7 | 0.03 | 11.93
| Analytical results for Dy in OREAS 163b (Indicative Value 20.7 ppm) | [17 empty] | Indicative
Dy | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | 20 | [15 empty] | 1
| 1 | 2 | 21 | [15 empty] | 3
| 1 | 3 | 20 | [15 empty] | 16
| 1 | 4 | 21 | [15 empty] | 20.6667
| 1 | 5 | 21 | [15 empty] | 20
| 1 | 6 | 21
| Mean |  | 20.6667
| Median |  | 21
| Std Dev. |  | 0.5164
| Rel.Std.Dev. |  | 2.499%
| PDM3 |  | -1.554e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Er | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 4
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 21
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Eu | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 5
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 22
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 163b (Certified Value 5.81 wt.%) | [17 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.18 | 2.19 | 2.20 | 2.24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 5.94 | 5.925 | 5.8488 | 5.42 | 5.7786 | 5.805 | 5.847 | 5.776 | 5.776 | 5.66 | 5.747 | 5.805 | 5.83 | 5.83 | 5.776 | 5.82 | 1
| 1 | 2 | 5.87 | 5.901 | 5.8588 | 5.4 | 5.8158 | 5.805 | 5.776 | 5.79 | 5.79 | 5.7 | 5.733 | 5.819 | 5.86 | 5.82 | 5.719 | 5.85 | 0
| 1 | 3 | 5.82 | 5.991 | 5.7888 | 5.58 | 5.7706 | 5.819 | 5.733 | 5.776 | 5.847 | 5.69 | 5.79 | 5.819 | 5.86 | 5.83 | 5.762 | 5.81 | 16
| 1 | 4 | 5.8 | 5.888 | 5.8988 | 5.66 | 5.8312 | 5.79 | 5.776 | 5.89 | 5.819 | 5.68 | 5.762 | 5.805 | 5.83 | 5.85 | 5.69 | 5.83 | 5.8129
| 1 | 5 | 5.87 | 5.906 | 5.8088 | 5.67 | 5.7656 | 5.79 | 5.833 | 5.847 | 5.819 | 5.63 | 5.719 | 5.805 | 5.82 | 5.82 | 5.79 | 5.78 | 130
| 1 | 6 | 5.8 | 5.898 | 5.8388 | 5.55 | 5.7263 | 5.762 | 5.733 | 5.876 | 5.733 | 5.67 | 5.819 | 5.805 | 5.84 | 5.83 | 5.762 | 5.79
| Mean |  | 5.85 | 5.9182 | 5.8404 | 5.5467 | 5.7814 | 5.7952 | 5.783 | 5.8258 | 5.7973 | 5.6717 | 5.7617 | 5.8097 | 5.84 | 5.83 | 5.7498 | 5.8133
| Median |  | 5.845 | 5.9035 | 5.8438 | 5.565 | 5.7746 | 5.7975 | 5.776 | 5.8185 | 5.8045 | 5.675 | 5.7545 | 5.805 | 5.835 | 5.83 | 5.762 | 5.815
| Std Dev. |  | 0.054406 | 0.037712 | 0.038679 | 0.11553 | 0.037589 | 0.01957 | 0.048361 | 0.051639 | 0.040133 | 0.024833 | 0.037276 | 0.0072296 | 0.016733 | 0.010954 | 0.037749 | 0.02582
| Rel.Std.Dev. |  | 0.93% | 0.6372% | 0.6623% | 2.083% | 0.6502% | 0.3377% | 0.8363% | 0.8864% | 0.6923% | 0.4378% | 0.647% | 0.1244% | 0.2865% | 0.1879% | 0.6565% | 0.4441%
| PDM3 |  | 0.6375% | 1.81% | 0.4731% | -4.581% | -0.5434% | -0.3058% | -0.5151% | 0.2218% | -0.2685% | -2.43% | -0.8821% | -0.05635% | 0.4655% | 0.2934% | -1.086% | 0.006729%
| Z-Score (Absolute) |  | 0.99 | 2.45 | 0.79 | 5.48 | 0.47 | 0.18 | 0.43 | 0.48 | 0.13 | 2.81 | 0.89 | 0.13 | 0.78 | 0.57 | 1.14 | 0.21
| Analytical results for Ga in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ga | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 17
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 23
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Gd in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Gd | Round | Replicate | Lab | [15 empty] | 1
| No | No | 2.25 | [15 empty] | ppm
|  |  | BF*XRF | [15 empty] | 1
1
| 1 | 1 | < 10 | [15 empty] | 1
| 1 | 2 | < 10 | [15 empty] | 6
| 1 | 3 | < 10 | [15 empty] | 16
| 1 | 4 | < 10 | [15 empty] | < 10
| 1 | 5 | < 10 | [15 empty] | 24
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.04 | 2.25 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [14 empty] | 1
| 1 | 2 | < 100 | < 10 | [14 empty] | 19
| 1 | 3 | < 100 | < 10 | [14 empty] | 16
| 1 | 4 | < 100 | < 10 | [14 empty] | < 10
| 1 | 5 | < 100 | < 10 | [14 empty] | 25
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Hf in OREAS 163b (Certified Value < 80 ppm) | [17 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 170 | < 80 | < 80 | < 80 | < 80 | < 10 | [10 empty] | 1
| 1 | 2 | < 170 | < 80 | < 80 | < 80 | < 80 | 11 | [10 empty] | 23
| 1 | 3 | < 170 | < 80 | < 80 | < 80 | < 80 | 11 | [10 empty] | 16
| 1 | 4 | < 170 | < 80 | < 80 | < 80 | < 80 | 11 | [10 empty] | < 80
| 1 | 5 | < 170 | < 80 | < 80 | < 80 | < 80 | 11 | [10 empty] | 131
| 1 | 6 | < 170 | < 80 | < 80 | < 80 | < 80 | 11
| Mean |  | 2.25: 11
| Median |  | 2.25: 11
| Std Dev. |  | 2.25: 0
Rel.Std.Dev. | bar: 2.25
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.25
| Analytical results for Hg in OREAS 163b (Indicative Value < 100 ppm) | [17 empty] | Indicative
Hg | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.04 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100 | [10 empty] | 1
| 1 | 2 | 2.06: < 100 | [10 empty] | 1
| 1 | 3 | 2.06: < 100 | [10 empty] | 16
| 1 | 4 | 2.06: < 100 | [10 empty] | < 100
| 1 | 5 | 2.06: < 100 | [10 empty] | 14
| 1 | 6 | 2.06: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Ho in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Ho | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 3
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 15
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for In in OREAS 163b (Indicative Value < 100 ppm) | [17 empty] | Indicative
In | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.04 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100 | [10 empty] | 1
| 1 | 2 | 2.06: < 100 | [10 empty] | 10
| 1 | 3 | 2.06: < 100 | [10 empty] | 16
| 1 | 4 | 2.06: < 100 | [10 empty] | < 100
| 1 | 5 | 2.06: < 100 | [10 empty] | 16
| 1 | 6 | 2.06: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for K2O in OREAS 163b (Certified Value 0.105 wt.%) | [17 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.11 | 0.108 | 0.11 | 0.18 | 0.096561 | 0.11 | 0.1 | 0.1 | 0.11 | 0.1 | 0.096 | 0.08 | 0.1084 | 0.11 |  |  | 1
| 1 | 2 | 0.1 | 0.109 | 0.11 | 0.18 | 0.0956 | 0.11 | 0.1 | 0.11 | 0.11 | 0.1 | 0.095 | 0.08 | 0.1084 | 0.11 |  |  | 0
| 1 | 3 | 0.11 | 0.112 | 0.11 | 0.17 | 0.099296 | 0.1 | 0.1 | 0.1 | 0.11 | 0.11 | 0.1 | 0.08 | 0.1084 | 0.12 |  |  | 16
| 1 | 4 | 0.1 | 0.108 | 0.11 | 0.18 | 0.095646 | 0.11 | 0.1 | 0.11 | 0.11 | 0.1 | 0.096 | 0.08 | 0.1084 | 0.11 |  |  | 0.10549
| 1 | 5 | 0.1 | 0.107 | 0.1 | 0.18 | 0.095649 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.095 | 0.09 | 0.1084 | 0.12 |  |  | 132
| 1 | 6 | 0.1 | 0.109 | 0.1 | 0.17 | 0.097233 | 0.1 | 0.11 | 0.11 | 0.12 | 0.1 | 0.097 | 0.08 | 0.1084 | 0.12
| Mean |  | 0.10333 | 0.10883 | 0.10667 | 0.17667 | 0.096664 | 0.10667 | 0.10333 | 0.10667 | 0.11167 | 0.10333 | 0.0965 | 0.081667 | 0.1084 | 0.115
| Median |  | 0.1 | 0.1085 | 0.11 | 0.18 | 0.096105 | 0.11 | 0.1 | 0.11 | 0.11 | 0.1 | 0.096 | 0.08 | 0.1084 | 0.115
| Std Dev. |  | 0.005164 | 0.0017224 | 0.005164 | 0.005164 | 0.0014467 | 0.005164 | 0.005164 | 0.005164 | 0.0040825 | 0.005164 | 0.0018708 | 0.0040825 | 1.5202e-17 | 0.0054772
| Rel.Std.Dev. |  | 4.997% | 1.583% | 4.841% | 2.923% | 1.497% | 4.841% | 4.997% | 4.841% | 3.656% | 4.997% | 1.939% | 4.999% | 1.402e-14% | 4.763%
| PDM3 |  | -2.042% | 3.172% | 1.118% | 67.48% | -8.364% | 1.118% | -2.042% | 1.118% | 5.858% | -2.042% | -8.52% | -22.58% | 2.761% | 9.018%
| Z-Score (Absolute) |  | 0.67 | 0.44 | 0 | 14.16 | 2.02 | 0 | 0.67 | 0 | 1.01 | 0.67 | 2.06 | 5.06 | 0.35 | 1.69
| Analytical results for La in OREAS 163b (Indicative Value < 90 ppm) | [17 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 0
0
| 1 | 1 | < 90 | < 90 | < 90 | < 90 | 24 | [11 empty] | 1
| 1 | 2 | 90 | < 90 | < 90 | < 90 | 24 | [11 empty] | 11
| 1 | 3 | < 90 | < 90 | < 90 | < 90 | 24 | [11 empty] | 16
| 1 | 4 | < 90 | < 90 | < 90 | < 90 | 24 | [11 empty] | < 90
| 1 | 5 | < 90 | < 90 | < 90 | < 90 | 24 | [11 empty] | 17
| 1 | 6 | < 90 | < 90 | < 90 | 90 | 24
| Mean |  | 2.06: 90; 2.09: 90; 2.10: 24
| Median |  | 2.06: 90; 2.09: 90; 2.10: 24
| Std Dev. |  | 2.10: 0
Rel.Std.Dev. | bar: 2.10
PDM3
| Z-Score (Absolute) |  | 0.67 | 0 | 0 | 0.67 | 1.77
| Analytical results for Lu in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Lu | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 4
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 18
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for MgO in OREAS 163b (Certified Value 5.75 wt.%) | [17 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 5.88 | 5.59 | 5.79 | 4.67 | 6.1991 | 5.74 | 5.88 | 5.75 | 5.86 | 5.63 | 5.65 | 5.83 | 5.422 | 5.86 |  |  | 1
| 1 | 2 | 5.91 | 5.586 | 5.81 | 4.99 | 6.2446 | 5.75 | 5.84 | 5.78 | 5.92 | 5.56 | 5.71 | 5.81 | 5.439 | 5.88 |  |  | 0
| 1 | 3 | 5.9 | 5.605 | 5.74 | 5.06 | 6.1945 | 5.72 | 5.8 | 5.73 | 5.93 | 5.68 | 5.74 | 5.83 | 5.439 | 5.89 |  |  | 16
| 1 | 4 | 5.88 | 5.562 | 5.75 | 4.91 | 6.2415 | 5.74 | 5.82 | 5.88 | 5.93 | 5.42 | 5.65 | 5.82 | 5.389 | 5.86 |  |  | 5.7472
| 1 | 5 | 5.85 | 5.659 | 5.73 | 5.3 | 6.1761 | 5.7 | 5.95 | 5.83 | 5.94 | 5.77 | 5.64 | 5.87 | 5.455 | 5.89 |  |  | 133
| 1 | 6 | 5.84 | 5.597 | 5.79 | 5.04 | 6.1506 | 5.7 | 5.81 | 5.85 | 5.84 | 5.6 | 5.68 | 5.89 | 5.472 | 5.86
| Mean |  | 5.8767 | 5.5998 | 5.7683 | 4.995 | 6.2011 | 5.725 | 5.85 | 5.8033 | 5.9033 | 5.61 | 5.6783 | 5.8417 | 5.436 | 5.8733
| Median |  | 5.88 | 5.5935 | 5.77 | 5.015 | 6.1968 | 5.73 | 5.83 | 5.805 | 5.925 | 5.615 | 5.665 | 5.83 | 5.439 | 5.87
| Std Dev. |  | 0.027325 | 0.032419 | 0.032506 | 0.20599 | 0.036754 | 0.021679 | 0.056569 | 0.059217 | 0.042269 | 0.11798 | 0.039707 | 0.031252 | 0.028566 | 0.015055
| Rel.Std.Dev. |  | 0.465% | 0.5789% | 0.5635% | 4.124% | 0.5927% | 0.3787% | 0.967% | 1.02% | 0.716% | 2.103% | 0.6993% | 0.535% | 0.5255% | 0.2563%
| PDM3 |  | 2.253% | -2.564% | 0.3683% | -13.09% | 7.898% | -0.3856% | 1.789% | 0.9773% | 2.717% | -2.387% | -1.198% | 1.644% | -5.414% | 2.195%
| Z-Score (Absolute) |  | 0.62 | 1.26 | 0.12 | 5.38 | 2.82 | 0.41 | 0.44 | 0.12 | 0.8 | 1.2 | 0.73 | 0.38 | 2.38 | 0.59
| Analytical results for MnO in OREAS 163b (Certified Value 0.062 wt.%) | [17 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.07 | 0.007 | 0.06 | 0.058926 | 0.065 | 0.065 | 0.052 | 0.052 | 0.069 | 0.056 | 0.07 | 0.142 | 0.07 |  |  |  | 1
| 1 | 2 | 0.06 | 0.008 | 0.06 | 0.059907 | 0.065 | 0.065 | 0.052 | 0.052 | 0.069 | 0.05 | 0.07 | 0.142 | 0.07 |  |  |  | 0
| 1 | 3 | 0.06 | 0.016 | 0.05 | 0.058795 | 0.065 | 0.065 | 0.052 | 0.052 | 0.07 | 0.062 | 0.07 | 0.142 | 0.07 |  |  |  | 16
| 1 | 4 | 0.06 | 0.013 | 0.06 | 0.059936 | 0.065 | 0.065 | 0.065 | 0.052 | 0.068 | 0.057 | 0.07 | 0.142 | 0.07 |  |  |  | 0.061585
| 1 | 5 | 0.06 | 0.019 | 0.07 | 0.058971 | 0.065 | 0.065 | 0.065 | 0.052 | 0.069 | 0.054 | 0.07 | 0.142 | 0.07 |  |  |  | 134
| 1 | 6 | 0.06 | 0.007 | 0.06 | 0.058737 | 0.052 | 0.065 | 0.065 | 0.052 | 0.068 | 0.047 | 0.06 | 0.142 | 0.06
| Mean |  | 0.061667 | 0.011667 | 0.06 | 0.059212 | 0.062833 | 0.065 | 0.0585 | 0.052 | 0.068833 | 0.054333 | 0.068333 | 0.142 | 0.068333
| Median |  | 0.06 | 0.0105 | 0.06 | 0.058948 | 0.065 | 0.065 | 0.0585 | 0.052 | 0.069 | 0.055 | 0.07 | 0.142 | 0.07
| Std Dev. |  | 0.0040825 | 0.0051251 | 0.0063246 | 0.0005561 | 0.0053072 | 0 | 0.0071204 | 0 | 0.00075277 | 0.0053166 | 0.0040825 | 0 | 0.0040825
| Rel.Std.Dev. |  | 6.62% | 43.93% | 10.54% | 0.9392% | 8.447% | 0% | 12.17% | 0% | 1.094% | 9.785% | 5.974% | 0% | 5.974%
| PDM3 |  | 0.1323% | -81.06% | -2.574% | -3.854% | 2.027% | 5.545% | -5.01% | -15.56% | 11.77% | -11.78% | 10.96% | 130.6% | 10.96%
| Z-Score (Absolute) |  | 0 | 5.06 | 0.17 | 0.25 | 0.08 | 0.29 | 0.36 | 1.01 | 0.72 | 0.74 | 0.67 | 8.13 | 0.67
| Analytical results for Mo in OREAS 163b (Certified Value < 50 ppm) | [17 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 50 | < 50 | 16 | 17 | [9 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | < 50 | 17 | 17 | [9 empty] | 31
| 1 | 3 | < 100 | < 50 | < 50 | < 50 | < 50 | 17 | 17 | [9 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 50 | < 50 | 19 | 17 | [9 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 50 | < 50 | 16 | 18 | [9 empty] | 135
| 1 | 6 | < 100 | < 50 | < 50 | < 50 | < 50 | 16 | 17
| Mean |  | 2.25: 16.8333 | 17.1667
| Median |  | 2.25: 16.5 | 17
| Std Dev. |  | 2.25: 1.169 | 0.40825
| Rel.Std.Dev. |  | 2.25: 6.945% | 2.378%
PDM3
| Z-Score (Absolute) | [7 empty] | NA | bar: 2.06 to 2.25
| Analytical results for Na2O in OREAS 163b (Indicative Value 0.073 wt.%) | [17 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.18 | 2.19 | 2.24 | [8 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [8 empty] | 3
3
| 1 | 1 | 0.08 | 0.078 | < 0.003 | 0.06 | < 0.01 | 0.06 | N.A. | < 0.01 | [8 empty] | 1
| 1 | 2 | 0.1 | 0.078 | < 0.003 | 0.06 | < 0.01 | 0.07 | N.A. | < 0.01 | [8 empty] | 13
| 1 | 3 | 0.11 | 0.08 | < 0.003 | 0.06 | < 0.01 | < 0.01 | N.A. | < 0.01 | [8 empty] | 16
| 1 | 4 | 0.07 | 0.079 | < 0.003 | 0.06 | < 0.01 | 0.07 | N.A. | < 0.01 | [8 empty] | 0.073292
| 1 | 5 | 0.08 | 0.079 | 0.004 | 0.06 | < 0.01 | 0.08 | N.A. | < 0.01 | [8 empty] | 19
| 1 | 6 | 0.09 | 0.079 | 0.008 | 0.06 | < 0.01 | 0.05 | N.A. | < 0.01
| Mean |  | 0.088333 | 0.078833 | 0.006 | 0.06 |  | 0.066
| Median |  | 0.085 | 0.079 | 0.006 | 0.06 |  | 0.07
| Std Dev. |  | 0.01472 | 0.00075277 | 0.0028284 | 0 |  | 0.011402
| Rel.Std.Dev. |  | 16.66% | 0.9549% | 47.14% | 0% |  | 17.28%
| PDM3 |  | 20.52% | 7.561% | -91.81% | -18.14% |  | -9.949%
| Z-Score (Absolute) |  | 0.67 | 0.48 | 1.1 | 0.09 | 1.05 | 0 | NA | 1.05
| Analytical results for Nb in OREAS 163b (Indicative Value 44.7 ppm) | [17 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 1
| 1 | 2 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 14
| 1 | 3 | < 50 | < 50 | 50 | < 50 | 46 | [11 empty] | 16
| 1 | 4 | < 50 | < 50 | 60 | < 50 | 46 | [11 empty] | 44.6667
| 1 | 5 | < 50 | < 50 | < 50 | < 50 | 46 | [11 empty] | 20
| 1 | 6 | < 50 | < 50 | < 50 | < 50 | 46
| Mean |  | 2.08: 52.5; 2.10: 46
| Median |  | 2.08: 50; 2.10: 46
| Std Dev. |  | 2.08: 5; 2.10: 0
| Rel.Std.Dev. |  | 2.08: 9.524%; 2.10: 0%
| PDM3 |  | 2.08: 17.54%; 2.10: 2.985%
Z-Score (Absolute) | bar: 2.06 to 2.10
| Analytical results for Nd in OREAS 163b (Indicative Value 41 ppm) | [17 empty] | Indicative
Nd | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: 38 | [10 empty] | 1
| 1 | 2 | 2.06: 38 | [10 empty] | 5
| 1 | 3 | 2.06: 38 | [10 empty] | 16
| 1 | 4 | 2.06: 44 | [10 empty] | 41
| 1 | 5 | 2.06: 47 | [10 empty] | 21
| 1 | 6 | 2.06: 41
| Mean |  | 2.06: 41
| Median |  | 2.06: 39.5
| Std Dev. |  | 2.06: 3.7947
| Rel.Std.Dev. |  | 2.06: 9.255%
PDM3 | bar: 2.06
Z-Score (Absolute) | bar: 2.06
| Analytical results for Ni in OREAS 163b (Indicative Value < 100 ppm) | [17 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 196 | < 50 | < 50 | 80 | < 50 | 24 | < 100 | N.A. | < 100 | 18 | [5 empty] | 1
| 1 | 2 | < 100 | < 196 | < 50 | < 50 | 160 | < 50 | 24 | < 100 | N.A. | < 100 | 18 | [5 empty] | 16
| 1 | 3 | < 100 | < 196 | < 50 | < 50 | 190 | < 50 | 55 | < 100 | N.A. | < 100 | 18 | [5 empty] | 16
| 1 | 4 | 200 | < 196 | < 50 | < 50 | 170 | < 50 | 55 | < 100 | N.A. | < 100 | 17 | [5 empty] | < 100
| 1 | 5 | < 100 | < 196 | < 50 | < 50 | 120 | < 50 | 55 | < 100 | N.A. | < 100 | 15 | [5 empty] | 22
| 1 | 6 | < 100 | < 196 | < 50 | < 50 | 160 | < 50 | < 24 | < 100 | N.A. | < 100 | 19
| Mean |  | 2.06: 200; 2.10: 146.6667 | 42.6 |  |  |  | 17.5
| Median |  | 2.06: 200; 2.10: 160 | 55 |  |  |  | 18
| Std Dev. |  | 2.10: 39.833 | 16.9794 |  |  |  | 1.3784
| Rel.Std.Dev. |  | 2.10: 27.16% | 39.86% |  |  |  | 7.877%
PDM3
| Z-Score (Absolute) |  | 1.13 | 1.97 | 0.67 | 0.67 | 3.72 | 0.67 | 0.23 | 0.23 | NA | 0.23 | 0.95
| Analytical results for P2O5 in OREAS 163b (Certified Value 0.038 wt.%) | [17 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.033 | 0.041 | 0.04 | 0.1 | 0.031743 | 0.03 | 0.02 | 0.04 | 0.04 | 0.04 | < 0.005 | 0.03 | < 0.0023 | 0.04 |  |  | 1
| 1 | 2 | 0.034 | 0.04 | 0.04 | 0.1 | 0.034823 | 0.03 | 0.03 | 0.04 | 0.04 | 0.04 | < 0.005 | 0.04 | < 0.0023 | 0.04 |  |  | 0
| 1 | 3 | 0.033 | 0.042 | 0.04 | 0.12 | 0.032694 | 0.03 | 0.02 | 0.04 | 0.04 | 0.04 | < 0.005 | 0.04 | < 0.0023 | 0.04 |  |  | 16
| 1 | 4 | 0.033 | 0.039 | 0.05 | 0.09 | 0.030741 | 0.03 | 0.04 | 0.04 | 0.04 | 0.04 | < 0.005 | 0.04 | < 0.0023 | 0.04 |  |  | 0.038248
| 1 | 5 | 0.033 | 0.04 | 0.04 | 0.1 | 0.031767 | 0.03 | 0.03 | 0.04 | 0.04 | 0.04 | < 0.005 | 0.04 | < 0.0023 | 0.04 |  |  | 136
| 1 | 6 | 0.034 | 0.041 | 0.04 | 0.17 | 0.03062 | 0.03 | 0.02 | 0.04 | 0.04 | 0.04 | < 0.005 | 0.04 | < 0.0023 | 0.04
| Mean |  | 0.033333 | 0.0405 | 0.041667 | 0.11333 | 0.032065 | 0.03 | 0.026667 | 0.04 | 0.04 | 0.04 |  | 0.038333 |  | 0.04
| Median |  | 0.033 | 0.0405 | 0.04 | 0.1 | 0.031755 | 0.03 | 0.025 | 0.04 | 0.04 | 0.04 |  | 0.04 |  | 0.04
| Std Dev. |  | 0.0005164 | 0.0010488 | 0.0040825 | 0.029439 | 0.0015516 | 0 | 0.008165 | 0 | 0 | 0 |  | 0.0040825 |  | 0
| Rel.Std.Dev. |  | 1.549% | 2.59% | 9.798% | 25.98% | 4.839% | 0% | 30.62% | 0% | 0% | 0% |  | 10.65% |  | 0%
| PDM3 |  | -12.85% | 5.888% | 8.938% | 196.3% | -16.17% | -21.56% | -30.28% | 4.581% | 4.581% | 4.581% |  | 0.2233% |  | 4.581%
| Z-Score (Absolute) |  | 0.94 | 0.22 | 0.4 | 12 | 1.15 | 1.48 | 2.02 | 0.13 | 0.13 | 0.13 | 5.93 | 0.13 | 6.15 | 0.13
| Analytical results for Pb in OREAS 163b (Indicative Value 63 ppm) | [17 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | 50 | 200 | < 30 | < 100 | < 100 | < 10 | 64 | [6 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | 50 | 200 | < 30 | < 100 | < 100 | < 10 | 63 | [6 empty] | 17
| 1 | 3 | < 100 | < 50 | < 50 | 60 | 170 | < 30 | 100 | < 100 | < 10 | 63 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | 80 | 210 | < 30 | < 100 | < 100 | < 10 | 65 | [6 empty] | 62.7778
| 1 | 5 | < 100 | < 50 | < 50 | 60 | 180 | < 30 | < 100 | < 100 | < 10 | 62 | [6 empty] | 23
| 1 | 6 | < 100 | < 50 | < 50 | 100 | 190 | < 30 | < 100 | < 100 | < 10 | 63
| Mean |  | 2.09: 66.6667; 2.10: 191.6667 | 100 |  |  | 63.3333
| Median |  | 2.09: 60; 2.10: 195 | 100 |  |  | 63
| Std Dev. |  | 2.09: 19.6638; 2.10: 14.7196 |  |  |  | 1.0328
| Rel.Std.Dev. |  | 2.09: 29.5%; 2.10: 7.68% |  |  |  | 1.631%
| PDM3 |  | 2.09: 6.195%; 2.10: 205.3% | 59.29% |  |  | 0.885%
| Z-Score (Absolute) |  | 0 | 0.81 | 0.81 | 0.54 | 4.59 | 1.13 | 0.27 | 0 | 1.46 | 0.43
| Analytical results for Pr in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Pr | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 18
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 24
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Rb in OREAS 163b (Certified Value < 50 ppm) | [17 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | [10 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | [10 empty] | 37
| 1 | 3 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | [10 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | [10 empty] | < 50
| 1 | 5 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | [10 empty] | 137
| 1 | 6 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.25
| Analytical results for Sb in OREAS 163b (Certified Value < 50 ppm) | [17 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [9 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 1
| 1 | 2 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 39
| 1 | 3 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 16
| 1 | 4 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | < 50
| 1 | 5 | < 100 | 50 | < 50 | < 50 | < 50 | < 10 | 23 | [9 empty] | 138
| 1 | 6 | < 100 | < 50 | < 50 | < 50 | < 50 | < 10 | 23
| Mean |  | 2.07: 50 | 23
| Median |  | 2.07: 50 | 23
| Std Dev. | [7 empty] | 0
| Rel.Std.Dev. | [7 empty] | 0%
PDM3
| Z-Score (Absolute) |  | 8.43 | 1.4 | 0 | 0 | 0 | 6.74 | 0.67
| Analytical results for Sc in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Sc | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: < 10 | [10 empty] | 19
| 1 | 3 | 2.06: < 100; 2.07: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 100; 2.07: < 10 | [10 empty] | 25
| 1 | 6 | 2.06: < 100; 2.07: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.07
| Analytical results for Se in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Se | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: < 10 | [10 empty] | 41
| 1 | 3 | 2.06: < 100; 2.07: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 100; 2.07: < 10 | [10 empty] | 14
| 1 | 6 | 2.06: < 100; 2.07: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.07
| Analytical results for SiO2 in OREAS 163b (Certified Value 73.63 wt.%) | [17 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 73.11 | 72.99 | 73.97 | 61.23 | 73.7217 | 73.6 | 73.5 | 73.2 | 73.8 | 70.66 | 67.7 | 73.92 | 73.717 | 73.32 |  |  | 1
| 1 | 2 | 73.48 | 72.98 | 74.32 | 61.31 | 74.3421 | 74 | 73 | 73.8 | 74.1 | 70.55 | 68.24 | 73.89 | 73.096 | 73.59 |  |  | 0
| 1 | 3 | 73.68 | 72.869 | 73.3 | 60.95 | 73.6356 | 73.8 | 72.5 | 73.3 | 74.2 | 70.67 | 68.78 | 73.79 | 74.273 | 73.42 |  |  | 16
| 1 | 4 | 73.62 | 72.965 | 74.6 | 61.4 | 74.2517 | 73.6 | 73.1 | 74.7 | 74.1 | 69.33 | 68.25 | 73.91 | 73.375 | 73.64 |  |  | 73.6264
| 1 | 5 | 73.33 | 73.036 | 73.4 | 60.31 | 73.6701 | 73.6 | 74 | 74.1 | 74.2 | 71.86 | 67.95 | 73.87 | 73.289 | 73.63 |  |  | 139
| 1 | 6 | 73.33 | 72.928 | 74.07 | 60.41 | 73.1381 | 73.2 | 72.7 | 74.5 | 73.7 | 70.25 | 68.03 | 73.77 | 74.273 | 73.53
| Mean |  | 73.425 | 72.9613 | 73.9433 | 60.935 | 73.7932 | 73.6333 | 73.1333 | 73.9333 | 74.0167 | 70.5533 | 68.1583 | 73.8583 | 73.6705 | 73.5217
| Median |  | 73.405 | 72.9725 | 74.02 | 61.09 | 73.6959 | 73.6 | 73.05 | 73.95 | 74.1 | 70.605 | 68.135 | 73.88 | 73.546 | 73.56
| Std Dev. |  | 0.21135 | 0.057263 | 0.50969 | 0.47124 | 0.44402 | 0.26583 | 0.5465 | 0.61536 | 0.2137 | 0.81498 | 0.36647 | 0.063377 | 0.50815 | 0.12766
| Rel.Std.Dev. |  | 0.2878% | 0.07848% | 0.6893% | 0.7734% | 0.6017% | 0.361% | 0.7473% | 0.8323% | 0.2887% | 1.155% | 0.5377% | 0.08581% | 0.6898% | 0.1736%
| PDM3 |  | -0.2735% | -0.9033% | 0.4305% | -17.24% | 0.2266% | 0.009464% | -0.6696% | 0.4169% | 0.5301% | -4.174% | -7.427% | 0.3151% | 0.05994% | -0.1422%
| Z-Score (Absolute) |  | 0.28 | 1.15 | 0.68 | 23.63 | 0.4 | 0.1 | 0.83 | 0.66 | 0.82 | 5.65 | 10.13 | 0.52 | 0.17 | 0.1
| Analytical results for Sm in OREAS 163b (Indicative Value 25.5 ppm) | [17 empty] | Indicative
Sm | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: 29 | [10 empty] | 1
| 1 | 2 | 2.06: 22 | [10 empty] | 5
| 1 | 3 | 2.06: 22 | [10 empty] | 16
| 1 | 4 | 2.06: 28 | [10 empty] | 25.5
| 1 | 5 | 2.06: 28 | [10 empty] | 15
| 1 | 6 | 2.06: 24
| Mean |  | 2.06: 25.5
| Median |  | 2.06: 26
| Std Dev. |  | 2.06: 3.2094
| Rel.Std.Dev. |  | 2.06: 12.59%
PDM3 | bar: 2.06
Z-Score (Absolute) | bar: 2.06
| Analytical results for Sn in OREAS 163b (Certified Value < 50 ppm) | [17 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.20 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 1500 | < 10 | [8 empty] | 1
| 1 | 2 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 1600 | < 10 | [8 empty] | 26
| 1 | 3 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 1500 | < 10 | [8 empty] | 16
| 1 | 4 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 1500 | < 10 | [8 empty] | < 50
| 1 | 5 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 1600 | < 10 | [8 empty] | 140
| 1 | 6 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 1600 | < 10
| Mean | [7 empty] | 1550
| Median | [7 empty] | 1550
| Std Dev. | [7 empty] | 54.7723
| Rel.Std.Dev. | [7 empty] | 3.534%
PDM3
| Z-Score (Absolute) |  | 1.69 | 3.63 | 0 | 0 | 0 | 0 | 102.83 | 1.35
| Analytical results for SO3 in OREAS 163b (Certified Value 8.39 wt.%) | [17 empty] | Certified
SO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.18 | 2.19 | 2.20 | 2.25 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 2
3
| 1 | 1 | 8.3406 | 8.3917 | 21.33 | 8.565 | 8.54 | 8.341 | 8.615 | 7.967 | N.A. | 10.2385 | 6.7007 | [5 empty] | 1
| 1 | 2 | 8.3156 | 8.4302 | 21.14 | 8.191 | 8.465 | 8.416 | 8.565 | 8.062 | N.A. | 10.4133 | 6.7349 | [5 empty] | 0
| 1 | 3 | 8.2657 | 8.4424 | 21.22 | 8.54 | 8.465 | 8.341 | 8.49 | 8.085 | N.A. | 10.3384 | 6.7576 | [5 empty] | 16
| 1 | 4 | 8.2807 | 8.4162 | 21.25 | 8.54 | 8.49 | 8.515 | 8.565 | 8.023 | N.A. | 10.2884 | 6.7104 | [5 empty] | 8.3861
| 1 | 5 | 8.3131 | 8.3717 | 21.15 | 8.615 | 8.615 | 8.366 | 8.465 | 7.979 | N.A. | 10.2635 | 6.7082 | [5 empty] | 141
| 1 | 6 | 8.3656 | 8.3989 | 21.15 | 8.341 | 8.416 | 8.441 | 8.49 | 8.044 | N.A. | 10.3883 | 6.7576
| Mean |  | 8.3135 | 8.4085 | 21.2067 | 8.4653 | 8.4985 | 8.4033 | 8.5317 | 8.0267 |  | 10.3217 | 6.7283
| Median |  | 8.3143 | 8.4076 | 21.185 | 8.54 | 8.4775 | 8.391 | 8.5275 | 8.0335 |  | 10.3134 | 6.7227
| Std Dev. |  | 0.036916 | 0.026108 | 0.075011 | 0.16383 | 0.069882 | 0.068137 | 0.058452 | 0.04647 |  | 0.070033 | 0.025497
| Rel.Std.Dev. |  | 0.444% | 0.3105% | 0.3537% | 1.935% | 0.8223% | 0.8108% | 0.6851% | 0.5789% |  | 0.6785% | 0.379%
| PDM3 |  | -0.8657% | 0.2667% | 152.9% | 0.9442% | 1.34% | 0.2049% | 1.735% | -4.287% |  | 23.08% | -19.77%
| Z-Score (Absolute) |  | 0.76 | 0.18 | 78.81 | 0.18 | 0.38 | 0.21 | 0.59 | 2.53 | NA | 11.63 | 10.55
| Analytical results for SrO in OREAS 163b (Certified Value < 120 ppm) | [17 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.18 | 2.19 | 2.24 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | < 50 | N.A. | < 100 | < 12 | [6 empty] | 1
| 1 | 2 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | 60 | N.A. | < 100 | < 12 | [6 empty] | 0
| 1 | 3 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | < 50 | N.A. | < 100 | 11.8 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | < 50 | N.A. | < 100 | 11.8 | [6 empty] | < 120
| 1 | 5 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | < 50 | N.A. | < 100 | 11.8 | [6 empty] | 142
| 1 | 6 | < 100 | < 50 | < 120 | < 120 | < 120 | < 120 | 50 | N.A. | < 100 | < 12
| Mean | [7 empty] | 55 |  |  | 11.8
| Median | [7 empty] | 55 |  |  | 11.8
| Std Dev. | [7 empty] | 7.0711 |  |  | 2.1756e-15
| Rel.Std.Dev. | [7 empty] | 12.86% |  |  | 1.844e-14%
PDM3
| Z-Score (Absolute) |  | 0 | 1.85 | 0.67 | 0.67 | 0.67 | 0.67 | 1.11 | NA | 0 | 3.04
| Analytical results for Ta in OREAS 163b (Indicative Value 64 ppm) | [17 empty] | Indicative
Ta | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100; 2.07: 65 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: 68 | [10 empty] | 10
| 1 | 3 | 2.06: < 100; 2.07: 61 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: 60 | [10 empty] | 63.5
| 1 | 5 | 2.06: < 100; 2.07: 65 | [10 empty] | 16
| 1 | 6 | 2.06: < 100; 2.07: 62
| Mean |  | 2.07: 63.5
| Median |  | 2.07: 63.5
| Std Dev. |  | 2.07: 3.0166
| Rel.Std.Dev. |  | 2.07: 4.751%
PDM3 | bar: 2.07
| Z-Score (Absolute) |  | 2.06: 0.67; 2.07: 0.67
| Analytical results for Tb in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Tb | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 11
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 17
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Te in OREAS 163b (Indicative Value < 100 ppm) | [17 empty] | Indicative
Te | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.04 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 0
0
| 1 | 1 | 2.06: < 100 | [10 empty] | 1
| 1 | 2 | 2.06: < 100 | [10 empty] | 12
| 1 | 3 | 2.06: < 100 | [10 empty] | 16
| 1 | 4 | 2.06: < 100 | [10 empty] | < 100
| 1 | 5 | 2.06: < 100 | [10 empty] | 18
| 1 | 6 | 2.06: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Th in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Th | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 13
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 19
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for TiO2 in OREAS 163b (Certified Value 0.056 wt.%) | [17 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.20 | 2.24 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF |  |  | 3
3
| 1 | 1 | 0.06 | 0.064 | 0.04 | < 0.01 | 0.025343 | 0.06 | 0.05 | 0.04 | 0.07 | 0.06 | 0.04 | 0.06 | 0.0334 | 0.08 |  |  | 1
| 1 | 2 | 0.05 | 0.064 | 0.04 | < 0.01 | 0.030419 | 0.06 | 0.05 | 0.06 | 0.07 | 0.05 | 0.04 | 0.06 | 0.05 | 0.07 |  |  | 0
| 1 | 3 | 0.07 | 0.065 | 0.04 | < 0.01 | 0.026298 | 0.06 | 0.05 | 0.06 | 0.05 | 0.06 | 0.04 | 0.07 | 0.05 | 0.07 |  |  | 16
| 1 | 4 | 0.08 | 0.064 | 0.05 | < 0.01 | 0.023333 | 0.06 | 0.05 | 0.05 | 0.06 | 0.05 | 0.04 | 0.07 | 0.05 | 0.08 |  |  | 0.056392
| 1 | 5 | 0.06 | 0.062 | 0.04 | < 0.01 | 0.023333 | 0.06 | 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | 0.07 | 0.05 | 0.07 |  |  | 143
| 1 | 6 | 0.07 | 0.061 | 0.04 | < 0.01 | 0.025262 | 0.06 | 0.04 | 0.05 | 0.07 | 0.05 | 0.04 | 0.06 | 0.05 | 0.07
| Mean |  | 0.065 | 0.063333 | 0.041667 |  | 0.025665 | 0.06 | 0.048333 | 0.051667 | 0.061667 | 0.055 | 0.041667 | 0.065 | 0.047233 | 0.073333
| Median |  | 0.065 | 0.064 | 0.04 |  | 0.025302 | 0.06 | 0.05 | 0.05 | 0.065 | 0.055 | 0.04 | 0.065 | 0.05 | 0.07
| Std Dev. |  | 0.010488 | 0.0015055 | 0.0040825 |  | 0.0026133 | 0 | 0.0040825 | 0.0075277 | 0.0098319 | 0.0054772 | 0.0040825 | 0.0054772 | 0.0067769 | 0.005164
| Rel.Std.Dev. |  | 16.14% | 2.377% | 9.798% |  | 10.18% | 0% | 8.447% | 14.57% | 15.94% | 9.959% | 9.798% | 8.427% | 14.35% | 7.042%
| PDM3 |  | 15.26% | 12.31% | -26.11% |  | -54.49% | 6.397% | -14.29% | -8.38% | 9.353% | -2.469% | -26.11% | 15.26% | -16.24% | 30.04%
| Z-Score (Absolute) |  | 0.73 | 0.62 | 0.73 | 3.01 | 1.72 | 0.41 | 0.31 | 0.1 | 0.52 | 0.1 | 0.73 | 0.73 | 0.38 | 1.24
| Analytical results for Tl in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Tl | Round | Replicate | 2.06: Lab; 2.07: Lab | [10 empty] | 1
| No | No | 2.06: 2.04; 2.07: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF; 2.07: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 100; 2.07: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 100; 2.07: < 10 | [10 empty] | 14
| 1 | 3 | 2.06: < 100; 2.07: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 100; 2.07: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 100; 2.07: < 10 | [10 empty] | 20
| 1 | 6 | 2.06: < 100; 2.07: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.07
| Analytical results for Tm in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Tm | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 6
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 21
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for U in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
U | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 16
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 22
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for V in OREAS 163b (Certified Value < 100 ppm) | [17 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.18 | 2.19 | 2.25 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | < 110 | < 100 | < 100 | < 100 | < 100 | < 17 | < 28 | N.A. | 16 | [6 empty] | 1
| 1 | 2 | < 100 | < 110 | < 100 | < 100 | < 100 | < 100 | < 17 | < 28 | N.A. | 14 | [6 empty] | 36
| 1 | 3 | < 100 | < 110 | < 100 | < 100 | < 100 | < 100 | < 17 | < 28 | N.A. | 16 | [6 empty] | 16
| 1 | 4 | < 100 | < 110 | < 100 | < 100 | < 100 | < 100 | < 17 | < 28 | N.A. | 15 | [6 empty] | < 100
| 1 | 5 | < 100 | < 110 | < 100 | < 100 | < 100 | < 100 | < 17 | < 28 | N.A. | 15 | [6 empty] | 144
| 1 | 6 | < 100 | < 110 | < 100 | < 100 | < 100 | < 100 | < 17 | < 28 | N.A. | 15
| Mean | [10 empty] | 15.1667
| Median | [10 empty] | 15
| Std Dev. | [10 empty] | 0.75277
| Rel.Std.Dev. | [10 empty] | 4.963%
PDM3
| Z-Score (Absolute) | [7 empty] | NA | NA | NA | NA | bar: 2.06 to 2.25
| Analytical results for W in OREAS 163b (Indicative Value 23.3 ppm) | [17 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [10 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [10 empty] | 1
1
| 1 | 1 | < 100 | < 10 | < 10 | < 10 | 10 | 23 | [10 empty] | 1
| 1 | 2 | < 100 | < 10 | < 10 | < 10 | 20 | 23 | [10 empty] | 17
| 1 | 3 | < 100 | < 10 | < 10 | < 10 | 30 | 29 | [10 empty] | 16
| 1 | 4 | < 100 | < 10 | < 10 | < 10 | 20 | 26 | [10 empty] | 23.3333
| 1 | 5 | < 100 | < 10 | < 10 | < 10 | 30 | 27 | [10 empty] | 23
| 1 | 6 | < 100 | < 10 | < 10 | < 10 | 10 | 32
| Mean |  | 2.10: 20; 2.25: 26.6667
| Median |  | 2.10: 20; 2.25: 26.5
| Std Dev. |  | 2.10: 8.9443; 2.25: 3.5024
| Rel.Std.Dev. |  | 2.10: 44.72%; 2.25: 13.13%
| PDM3 |  | 2.10: -14.29%; 2.25: 14.29%
| Z-Score (Absolute) |  | 3.37 | 0.67 | 0.67 | 0.67 | 0.67 | 1.27
| Analytical results for Y in OREAS 163b (Indicative Value < 39 ppm) | [17 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 2.07 | 2.08 | 2.09 | 2.10 | 2.25 | [11 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | [11 empty] | 1
1
| 1 | 1 | < 39 | < 39 | < 39 | < 39 | < 10 | [11 empty] | 1
| 1 | 2 | < 39 | < 39 | < 39 | < 39 | < 10 | [11 empty] | 18
| 1 | 3 | < 39 | < 39 | < 39 | < 39 | < 10 | [11 empty] | 16
| 1 | 4 | < 39 | < 39 | 39 | < 39 | < 10 | [11 empty] | < 39
| 1 | 5 | < 39 | < 39 | < 39 | < 39 | < 10 | [11 empty] | 24
| 1 | 6 | < 39 | < 39 | < 39 | < 39 | < 10
| Mean |  | 2.08: 39
| Median |  | 2.08: 39
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06 to 2.10
| Analytical results for Yb in OREAS 163b (Indicative Value < 10 ppm) | [17 empty] | Indicative
Yb | Round | Replicate | 2.06: Lab | [10 empty] | 1
| No | No | 2.06: 2.25 | [10 empty] | ppm
|  |  | 2.06: BF*XRF | [10 empty] | 1
1
| 1 | 1 | 2.06: < 10 | [10 empty] | 1
| 1 | 2 | 2.06: < 10 | [10 empty] | 19
| 1 | 3 | 2.06: < 10 | [10 empty] | 16
| 1 | 4 | 2.06: < 10 | [10 empty] | < 10
| 1 | 5 | 2.06: < 10 | [10 empty] | 25
| 1 | 6 | 2.06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.06
| Analytical results for Zn in OREAS 163b (Certified Value < 50 ppm) | [17 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.19 | 2.20 | 2.25 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | PSF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | < 100 | < 50 | 67 | [5 empty] | 1
| 1 | 2 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | < 100 | < 50 | 67 | [5 empty] | 40
| 1 | 3 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 10 | < 100 | < 100 | < 50 | 68 | [5 empty] | 16
| 1 | 4 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | < 10 | < 100 | < 100 | < 50 | 67 | [5 empty] | < 50
| 1 | 5 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 20 | < 100 | < 100 | < 50 | 67 | [5 empty] | 145
| 1 | 6 | < 100 | < 160 | < 50 | < 50 | < 50 | < 50 | 50 | < 100 | < 100 | < 50 | 67
| Mean | [7 empty] | 24 |  |  |  | 67.1667
| Median | [7 empty] | 20 |  |  |  | 67
| Std Dev. | [7 empty] | 15.1658 |  |  |  | 0.40825
| Rel.Std.Dev. | [7 empty] | 63.19% |  |  |  | 0.6078%
PDM3
| Z-Score (Absolute) |  | 4.05 | 8.95 | 0 | 0 | 0 | 0 | 0.67 | 4.05 | 4.05 | 0 | 6.82
| Analytical results for Zr in OREAS 163b (Certified Value < 100 ppm) | [17 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.19 | 2.25 | [8 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 370 | < 100 | < 100 | < 100 | < 100 | N.A. | 22 | [8 empty] | 1
| 1 | 2 | < 100 | < 370 | < 100 | < 100 | < 100 | < 100 | N.A. | 22 | [8 empty] | 41
| 1 | 3 | < 100 | < 370 | < 100 | < 100 | < 100 | < 100 | N.A. | 22 | [8 empty] | 16
| 1 | 4 | < 100 | < 370 | < 100 | < 100 | < 100 | < 100 | N.A. | 22 | [8 empty] | < 100
| 1 | 5 | < 100 | < 370 | < 100 | < 100 | < 100 | < 100 | N.A. | 22 | [8 empty] | 146
| 1 | 6 | < 100 | < 370 | < 100 | < 100 | < 100 | < 100 | N.A. | 22
| Mean | [8 empty] | 22
| Median | [8 empty] | 22
| Std Dev. | [8 empty] | 0
| Rel.Std.Dev. | [8 empty] | 0%
PDM3
| Z-Score (Absolute) | [7 empty] | NA | NA | bar: 2.06 to 2.25
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 163b (Certified Value 4.99 wt.%) | [18 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.16 | 2.18 | 2.19 | 2.20 | 2.22 | 2.24 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | BF*OES/MS | LOI*TGA | LOI*Furnace | BF*XRF | 2
2
| 1 | 1 | 5.1361 | 4.06 | 3.71 | 4.22 | 3.5632 | 5.63 | 5.41 | 6.42 | 5.64 | 7.31 | 4.9927 | 3.93 | 4.92 | 6.8 | 7.16 | 3.62 | 5.72 | 1
| 1 | 2 | 4.996 | 4.09 | 3.75 | 4.43 | 3.5381 | 5.62 | 5.14 | 5.44 | 5.38 | 7.14 | 4.9262 | 3.93 | 4.56 | 6.9 | 7.09 | 3.61 | 5.73 | 0
| 1 | 3 | 4.9159 | 4.1 | 3.78 | 4.52 | 3.788 | 5.8 | 5.33 | 5.58 | 5.51 | 6.67 | 4.805 | 3.9 | 4.85 | 6.9 | 7.16 | 3.69 | 5.75 | 16
| 1 | 4 | 5.0351 | 4 | 3.62 | 4.96 | 3.4894 | 5.75 | 5.34 | 4.79 | 5.45 | 6.96 | 5.2055 | 3.83 | 4.61 | 6.5 | 7.13 | 3.58 | 5.8 | 4.9939
| 1 | 5 | 5.1326 | 4.07 | 3.83 | 4.68 | 3.4864 | 5.57 | 5.48 | 6.26 | 5.43 | 6.41 | 5.0632 | 3.86 | 4.94 | 6.7 | 7.03 | 3.69 | 5.7 | 148
| 1 | 6 | 4.9431 | 4.08 | 3.83 | 4.34 | 3.8885 | 5.73 | 5.31 | 5.36 | 5.38 | 7.29 | 4.7712 | 3.93 | 5.13 | 6.9 | 7.03 | 3.51 | 5.65
| Mean |  | 5.0265 | 4.0667 | 3.7533 | 4.525 | 3.6256 | 5.6833 | 5.335 | 5.6417 | 5.465 | 6.9633 | 4.9606 | 3.8967 | 4.835 | 6.7833 | 7.1 | 3.6167 | 5.725
| Median |  | 5.0155 | 4.075 | 3.765 | 4.475 | 3.5506 | 5.68 | 5.335 | 5.51 | 5.44 | 7.05 | 4.9595 | 3.915 | 4.885 | 6.85 | 7.11 | 3.615 | 5.725
| Std Dev. |  | 0.093229 | 0.03559 | 0.080166 | 0.2644 | 0.17026 | 0.089368 | 0.11432 | 0.60618 | 0.098539 | 0.36065 | 0.16297 | 0.04274 | 0.21529 | 0.16021 | 0.06 | 0.068605 | 0.0501
| Rel.Std.Dev. |  | 1.855% | 0.8752% | 2.136% | 5.843% | 4.696% | 1.572% | 2.143% | 10.74% | 1.803% | 5.179% | 3.285% | 1.097% | 4.453% | 2.362% | 0.8451% | 1.897% | 0.8751%
| PDM3 |  | 0.6516% | -18.57% | -24.84% | -9.39% | -27.4% | 13.81% | 6.83% | 12.97% | 9.433% | 39.44% | -0.6664% | -21.97% | -3.182% | 35.83% | 42.17% | -27.58% | 14.64%
| Z-Score (Absolute) |  | 0 | 0.93 | 1.23 | 0.48 | 1.35 | 0.63 | 0.3 | 0.59 | 0.42 | 1.87 | 0.06 | 1.09 | 0.18 | 1.7 | 2 | 1.36 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 163b (Indicative Value 4.38 ppm) | [22 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 2.04 | 2.13 | 2.24 | [18 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES/MS | [18 empty] | 2
2
| 1 | 1 | 4.54 | 4 | < 10 | [18 empty] | 1
| 1 | 2 | 4.58 | 4 | < 10 | [18 empty] | 8
| 1 | 3 | 4.46 | 5 | < 10 | [18 empty] | 16
| 1 | 4 | 4.59 | 4 | < 10 | [18 empty] | 4.3767
| 1 | 5 | 4.51 | 4 | < 10 | [18 empty] | 27
| 1 | 6 | 4.84 | 4 | < 10
| Mean |  | 4.5867 | 4.1667
| Median |  | 4.56 | 4
| Std Dev. |  | 0.13292 | 0.40825
| Rel.Std.Dev. |  | 2.898% | 9.798%
| PDM3 |  | 4.798% | -4.798%
| Z-Score (Absolute) |  | 0 | 0.69 | 0.67
| Analytical results for Al in OREAS 163b (Certified Value 0.379 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.39 | 0.36543 | 0.34 | 0.37 | 0.38509 | 0.397 | 0.381 | 0.381 | 0.37 | 0.392 | 0.37 | 0.37 | 0.41029 | 0.39 | 0.38 | 0.381 | 0.38 | 0.364 | 0.39 | 0.37 |  | 1
| 1 | 2 | 0.4 | 0.36739 | 0.32 | 0.38 | 0.38808 | 0.381 | 0.376 | 0.36 | 0.37 | 0.36 | 0.37 | 0.38 | 0.40679 | 0.39 | 0.4 | 0.381 | 0.38 | 0.368 | 0.41 | 0.36 |  | 0
| 1 | 3 | 0.39 | 0.36881 | 0.36 | 0.38 | 0.36264 | 0.386 | 0.392 | 0.37 | 0.349 | 0.381 | 0.37 | 0.39 | 0.40489 | 0.38 | 0.38 | 0.386 | 0.38 | 0.367 | 0.38 | 0.36 |  | 16
| 1 | 4 | 0.39 | 0.36353 | 0.35 | 0.38 | 0.39039 | 0.381 | 0.402 | 0.381 | 0.381 | 0.392 | 0.37 | 0.38 | 0.36323 | 0.38 | 0.38 | 0.381 | 0.39 | 0.365 | 0.39 | 0.37 |  | 0.37944
| 1 | 5 | 0.39 | 0.36761 | 0.34 | 0.38 | 0.38466 | 0.386 | 0.402 | 0.37 | 0.36 | 0.381 | 0.38 | 0.38 | 0.40279 | 0.38 | 0.37 | 0.386 | 0.39 | 0.363 | 0.35 | 0.36 |  | 150
| 1 | 6 | 0.4 | 0.36507 | 0.35 | 0.38 | 0.39153 | 0.392 | 0.392 | 0.381 | 0.37 | 0.36 | 0.37 | 0.38 | 0.35604 | 0.38 | 0.37 | 0.386 | 0.38 | 0.369 | 0.38 | 0.37
| Mean |  | 0.39333 | 0.36631 | 0.34333 | 0.37833 | 0.38373 | 0.38717 | 0.39083 | 0.37383 | 0.36667 | 0.37767 | 0.37167 | 0.38 | 0.39067 | 0.38333 | 0.38 | 0.3835 | 0.38333 | 0.366 | 0.38333 | 0.365
| Median |  | 0.39 | 0.36641 | 0.345 | 0.38 | 0.38659 | 0.386 | 0.392 | 0.3755 | 0.37 | 0.381 | 0.37 | 0.38 | 0.40384 | 0.38 | 0.38 | 0.3835 | 0.38 | 0.366 | 0.385 | 0.365
| Std Dev. |  | 0.005164 | 0.0019563 | 0.013663 | 0.0040825 | 0.010693 | 0.0063061 | 0.010666 | 0.0086583 | 0.010912 | 0.014542 | 0.0040825 | 0.0063246 | 0.024273 | 0.005164 | 0.010954 | 0.0027386 | 0.005164 | 0.0023664 | 0.019664 | 0.0054772
| Rel.Std.Dev. |  | 1.313% | 0.5341% | 3.979% | 1.079% | 2.787% | 1.629% | 2.729% | 2.316% | 2.976% | 3.85% | 1.098% | 1.664% | 6.213% | 1.347% | 2.883% | 0.7141% | 1.347% | 0.6466% | 5.13% | 1.501%
| PDM3 |  | 3.661% | -3.462% | -9.516% | -0.2923% | 1.13% | 2.036% | 3.002% | -1.478% | -3.367% | -0.468% | -2.049% | 0.147% | 2.959% | 1.025% | 0.147% | 1.069% | 1.025% | -3.543% | 1.025% | -3.806%
| Z-Score (Absolute) |  | 1.36 | 1.4 | 3.74 | 0.17 | 0.38 | 0.74 | 1.1 | 0.61 | 1.33 | 0.25 | 0.85 | 0 | 1.09 | 0.34 | 0 | 0.38 | 0.34 | 1.43 | 0.34 | 1.53
| Analytical results for As in OREAS 163b (Certified Value 165 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 160 | 145.5 | 200 | 200 | 200 | 200 | 200 | 168 | 152 | 173.5192 | 178 | 200 | 168 | 149 | 177 | 175 | [5 empty] | 1
| 1 | 2 | 160 | 150.4 | 200 | 100 | 200 | 200 | 200 | 162 | 146 | 176.0658 | 177 | 200 | 169 | 151 | 181 | 174 | [5 empty] | 9
| 1 | 3 | 170 | 159.6 | 200 | 200 | 200 | 200 | 200 | 170 | 153 | 174.5908 | 179 | 200 | 168 | 150 | 176 | 169 | [5 empty] | 16
| 1 | 4 | 170 | 153.7 | 300 | 200 | 200 | 200 | 200 | 161 | 144 | 165.9109 | 175 | 200 | 168 | 150 | 176 | 176 | [5 empty] | 164.6967
| 1 | 5 | 170 | 150.5 | 200 | 100 | 200 | 200 | 200 | 164 | 145 | 171.9865 | 175 | 200 | 168 | 151 | 156 | 164 | [5 empty] | 151
| 1 | 6 | 160 | 154.7 | 300 | 200 | 200 | 200 | 200 | 170 | 153 | 165.131 | 174 | 200 | 169 | 149 | 176 | 173
| Mean |  | 165 | 152.4 | 233.3333 | 166.6667 | 200 | 200 | 200 | 165.8333 | 148.8333 | 171.2007 | 176.3333 | 200 | 168.3333 | 150 | 173.6667 | 171.8333
| Median |  | 165 | 152.1 | 200 | 200 | 200 | 200 | 200 | 166 | 149 | 172.7529 | 176 | 200 | 168 | 150 | 176 | 173.5
| Std Dev. |  | 5.4772 | 4.7758 | 51.6398 | 51.6398 | 0 | 0 | 0 | 4.0208 | 4.2622 | 4.6039 | 1.9664 | 0 | 0.5164 | 0.89443 | 8.8694 | 4.535
| Rel.Std.Dev. |  | 3.32% | 3.134% | 22.13% | 30.98% | 0% | 0% | 0% | 2.425% | 2.864% | 2.689% | 1.115% | 0% | 0.3068% | 0.5963% | 5.107% | 2.639%
| PDM3 |  | 0.1841% | -7.466% | 41.67% | 1.196% | 21.44% | 21.44% | 21.44% | 0.6901% | -9.632% | 3.949% | 7.065% | 21.44% | 2.208% | -8.924% | 5.446% | 4.333%
| Z-Score (Absolute) |  | 0.25 | 1.75 | NA | NA | NA | NA | NA | 0.15 | 2.17 | 0.49 | 1.1 | NA | 0.15 | 2.03 | 0.78 | 0.57
Results from laboratories 2.07, 2.08, 2.09, 2.10, 2.11 and 2.17 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 163b (Indicative Value < 30 ppm) | [22 empty] | Indicative
B | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab |  | 1
| No | No | 2.01: 2.12; 2.02: 2.14; 2.04: 2.18 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*MS |  | 1
1
| 1 | 1 | 2.01: 10; 2.02: < 50; 2.04: < 30 |  | 1
| 1 | 2 | 2.01: 10; 2.02: < 50; 2.04: < 30 |  | 1
| 1 | 3 | 2.01: < 10; 2.02: < 50; 2.04: < 30 |  | 16
| 1 | 4 | 2.01: < 10; 2.02: < 50; 2.04: < 30 |  | < 30
| 1 | 5 | 2.01: 10; 2.02: < 50; 2.04: < 30 |  | 28
| 1 | 6 | 2.01: 10; 2.02: < 50; 2.04: < 30
| Mean |  | 2.01: 10
| Median |  | 2.01: 10
| Std Dev. |  | 2.01: 0
Rel.Std.Dev. | bar: 2.01
PDM3
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 1.01; 2.04: 0
| Analytical results for Ba in OREAS 163b (Certified Value 37.7 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 1
2
| 1 | 1 | 37 | 36.8 | 36.2 | 43.0911 | 40 | 36 | 33.2478 | 41 | < 100 | 36 | 37 | 75 | 38 | 28 | [7 empty] | 1
| 1 | 2 | 38 | 48.2 | 37.6 | 44.5438 | 41 | 38 | 43.2222 | 35 | < 100 | 37 | 37 | 76 | 38 | 30 | [7 empty] | 27
| 1 | 3 | 38 | 38.6 | 34.6 | 42.8171 | 37 | 38 | 36.7005 | 37 | < 100 | 37 | 38 | 76 | 37 | 30 | [7 empty] | 16
| 1 | 4 | 38 | 38.8 | 36.4 | 43.4759 | 37 | 37 | 34.8128 | 39 | < 100 | 38 | 37 | 75 | 37 | 31 | [7 empty] | 37.7402
| 1 | 5 | 39 | 37.8 | 37 | 46.285 | 37 | 38 | 45.5911 | 42 | < 100 | 38 | 37 | 77 | 32 | 29 | [7 empty] | 152
| 1 | 6 | 38 | 36.5 | 38.5 | 46.0183 | 41 | 37 | 44.2769 | 38 | < 100 | 37 | 38 | 77 | 36 | 31
| Mean |  | 38 | 39.45 | 36.7167 | 44.3719 | 38.8333 | 37.3333 | 39.6419 | 38.6667 |  | 37.1667 | 37.3333 | 76 | 36.3333 | 29.8333
| Median |  | 38 | 38.2 | 36.7 | 44.0098 | 38.5 | 37.5 | 39.9614 | 38.5 |  | 37 | 37 | 76 | 37 | 30
| Std Dev. |  | 0.63246 | 4.3853 | 1.3333 | 1.5007 | 2.0412 | 0.8165 | 5.3395 | 2.582 |  | 0.75277 | 0.5164 | 0.89443 | 2.2509 | 1.169
| Rel.Std.Dev. |  | 1.664% | 11.12% | 3.631% | 3.382% | 5.256% | 2.187% | 13.47% | 6.678% |  | 2.025% | 1.383% | 1.177% | 6.195% | 3.919%
| PDM3 |  | 0.6884% | 4.53% | -2.712% | 17.57% | 2.896% | -1.078% | 5.039% | 2.455% |  | -1.52% | -1.078% | 101.4% | -3.728% | -20.95%
| Z-Score (Absolute) |  | 0.18 | 0.61 | 0.88 | 3.29 | 0.27 | 0.54 | 0.71 | 0.18 | 6.36 | 0.64 | 0.54 | 20.52 | 1.09 | 4.63
| Analytical results for Be in OREAS 163b (Certified Value < 1 ppm) | [22 empty] | Certified
Be | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.12; 2.06: 2.13; 2.07: 2.14; 2.08: 2.18; 2.09: 2.19; 2.10: 2.20; 2.11: 2.21; 2.12: 2.24 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*OES; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: PF*OES; 2.11: PF*OES; 2.12: PF*OES/MS |  | 2
2
| 1 | 1 | 2.01: < 1; 2.02: < 2; 2.05: < 3; 2.06: < 5; 2.07: < 1; 2.08: 0.3; 2.09: < 1; 2.10: < 5; 2.11: < 5; 2.12: < 2 |  | 1
| 1 | 2 | 2.01: < 1; 2.02: < 2; 2.04: 0.5; 2.05: < 3; 2.06: < 5; 2.07: < 1; 2.08: 0.8; 2.09: < 1; 2.10: < 5; 2.11: < 5; 2.12: < 2 |  | 28
| 1 | 3 | 2.01: < 1; 2.02: < 2; 2.05: < 3; 2.06: < 5; 2.07: < 1; 2.08: 0.5; 2.09: < 1; 2.10: < 5; 2.11: < 5; 2.12: < 2 |  | 16
| 1 | 4 | 2.01: < 1; 2.02: < 2; 2.05: < 3; 2.06: < 5; 2.07: < 1; 2.08: 0.8; 2.09: < 1; 2.10: < 5; 2.11: < 5; 2.12: < 2 |  | < 1
| 1 | 5 | 2.01: < 1; 2.02: < 2; 2.05: < 3; 2.06: < 5; 2.07: < 1; 2.08: 1.2; 2.09: < 1; 2.10: < 5; 2.11: < 5; 2.12: < 2 |  | 153
| 1 | 6 | 2.01: < 1; 2.02: < 2; 2.05: < 3; 2.06: < 5; 2.07: < 1; 2.08: 0.5; 2.09: < 1; 2.10: < 5; 2.11: < 5; 2.12: < 2
| Mean |  | 2.04: 0.5; 2.08: 0.68333
| Median |  | 2.04: 0.5; 2.08: 0.65
| Std Dev. |  | 2.08: 0.31885
| Rel.Std.Dev. |  | 2.08: 46.66%
PDM3
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0; 2.04: 0.67; 2.05: 0.67; 2.06: 2.02; 2.07: 0.67; 2.08: 0.43; 2.09: 0.67; 2.10: 2.02; 2.11: 2.02; 2.12: 0
| Analytical results for Bi in OREAS 163b (Certified Value 17.6 ppm) | [22 empty] | Certified
Bi | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.23; 2.09: 2.24 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES/MS |  | 1
2
| 1 | 1 | 2.01: 17.1; 2.02: 13.8; 2.04: 17; 2.05: 17.5; 2.06: 17.18; 2.07: 18.4; 2.08: 17.6; 2.09: 18 |  | 1
| 1 | 2 | 2.01: 17; 2.02: 14; 2.04: 16; 2.05: 17.4; 2.06: 18.1389; 2.07: 18; 2.08: 18.4; 2.09: 18.1 |  | 29
| 1 | 3 | 2.01: 16.9; 2.02: 14.1; 2.04: 16; 2.05: 18.2; 2.06: 17.4031; 2.07: 18.5; 2.08: 17.7; 2.09: 17.7 |  | 16
| 1 | 4 | 2.01: 16.9; 2.02: 15.2; 2.04: 16; 2.05: 16.9; 2.06: 16.1465; 2.07: 18.5; 2.08: 17.5; 2.09: 18.3 |  | 17.5957
| 1 | 5 | 2.01: 16.8; 2.02: 13.8; 2.04: 17; 2.05: 17.1; 2.06: 18.0185; 2.07: 18.2; 2.08: 15.9; 2.09: 17.9 |  | 154
| 1 | 6 | 2.01: 16.8; 2.02: 14.5; 2.04: 16; 2.05: 17.6; 2.06: 15.938; 2.07: 18.1; 2.08: 17.4; 2.09: 18.4
| Mean |  | 2.01: 16.9167; 2.02: 14.2333; 2.04: 16.3333; 2.05: 17.45; 2.06: 17.1375; 2.07: 18.2833; 2.08: 17.4167; 2.09: 18.0667
| Median |  | 2.01: 16.9; 2.02: 14.05; 2.04: 16; 2.05: 17.45; 2.06: 17.2915; 2.07: 18.3; 2.08: 17.55; 2.09: 18.05
| Std Dev. |  | 2.01: 0.1169; 2.02: 0.53914; 2.04: 0.5164; 2.05: 0.45056; 2.06: 0.92434; 2.07: 0.2137; 2.08: 0.82321; 2.09: 0.2582
| Rel.Std.Dev. |  | 2.01: 0.6911%; 2.02: 3.788%; 2.04: 3.162%; 2.05: 2.582%; 2.06: 5.394%; 2.07: 1.169%; 2.08: 4.727%; 2.09: 1.429%
| PDM3 |  | 2.01: -3.859%; 2.02: -19.11%; 2.04: -7.174%; 2.05: -0.828%; 2.06: -2.604%; 2.07: 3.908%; 2.08: -1.017%; 2.09: 2.677%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 4.29; 2.04: NA; 2.05: 0.04; 2.06: 0.38; 2.07: 1.17; 2.08: 0; 2.09: 0.88
Results from laboratory 2.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 163b (Certified Value 2.61 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.5 | 2.5576 | 2.42 | 2.57 | 2.623 | 2.559 | 2.637 | 2.637 | 2.652 | 2.78 | 2.6 | 2.6196 | 2.7 | 2.5 | 2.659 | 2.64 | 3.076 | 2.8 | 2.5 |  |  | 1
| 1 | 2 | 2.5 | 2.5359 | 2.44 | 2.62 | 2.652 | 2.466 | 2.609 | 2.587 | 2.609 | 2.77 | 2.6 | 2.6512 | 2.7 | 2.49 | 2.666 | 2.65 | 3.069 | 2.7 | 2.5 |  |  | 0
| 1 | 3 | 2.5 | 2.5364 | 2.47 | 2.63 | 2.63 | 2.58 | 2.687 | 2.601 | 2.716 | 2.68 | 2.6 | 2.5981 | 2.7 | 2.41 | 2.694 | 2.66 | 3.054 | 2.7 | 2.4 |  |  | 16
| 1 | 4 | 2.6 | 2.5371 | 2.48 | 2.6 | 2.673 | 2.623 | 2.752 | 2.637 | 2.694 | 2.63 | 2.6 | 2.3197 | 2.7 | 2.52 | 2.652 | 2.65 | 3.051 | 2.7 | 2.6 |  |  | 2.6103
| 1 | 5 | 2.6 | 2.5394 | 2.46 | 2.62 | 2.652 | 2.687 | 2.773 | 2.652 | 2.709 | 2.8 | 2.6 | 2.5638 | 2.7 | 2.48 | 2.687 | 2.65 | 3.052 | 2.4 | 2.5 |  |  | 155
| 1 | 6 | 2.5 | 2.5605 | 2.46 | 2.62 | 2.63 | 2.673 | 2.73 | 2.659 | 2.673 | 2.82 | 2.6 | 2.3643 | 2.7 | 2.46 | 2.666 | 2.65 | 3.069 | 2.7 | 2.6
| Mean |  | 2.5333 | 2.5445 | 2.455 | 2.61 | 2.6433 | 2.598 | 2.698 | 2.6288 | 2.6755 | 2.7467 | 2.6 | 2.5195 | 2.7 | 2.4767 | 2.6707 | 2.65 | 3.0618 | 2.6667 | 2.5167
| Median |  | 2.5 | 2.5382 | 2.46 | 2.62 | 2.641 | 2.6015 | 2.7085 | 2.637 | 2.6835 | 2.775 | 2.6 | 2.581 | 2.7 | 2.485 | 2.666 | 2.65 | 3.0615 | 2.7 | 2.5
| Std Dev. |  | 0.05164 | 0.011387 | 0.021679 | 0.021909 | 0.01897 | 0.08178 | 0.065312 | 0.028653 | 0.040213 | 0.074744 | 0 | 0.1411 | 4.8648e-16 | 0.038297 | 0.016367 | 0.0063246 | 0.010759 | 0.13663 | 0.075277
| Rel.Std.Dev. |  | 2.038% | 0.4475% | 0.8831% | 0.8394% | 0.7177% | 3.148% | 2.421% | 1.09% | 1.503% | 2.721% | 0% | 5.6% | 1.802e-14% | 1.546% | 0.6128% | 0.2387% | 0.3514% | 5.123% | 2.991%
| PDM3 |  | -2.95% | -2.523% | -5.951% | -0.01304% | 1.264% | -0.4727% | 3.358% | 0.7085% | 2.496% | 5.223% | -0.3961% | -3.482% | 3.435% | -5.121% | 2.311% | 1.519% | 17.3% | 2.158% | -3.589%
| Z-Score (Absolute) |  | 0.93 | 0.82 | 1.7 | 0.18 | 0.14 | 0.3 | 0.67 | 0 | 0.45 | 1.15 | 0.28 | 1.07 | 0.69 | 1.49 | 0.41 | 0.21 | 4.23 | 0.37 | 1.1
| Analytical results for Cd in OREAS 163b (Certified Value < 10 ppm) | [22 empty] | Certified
Cd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.12; 2.04: 2.13; 2.05: 2.14; 2.07: 2.17; 2.08: 2.21; 2.09: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: PF*OES/MS; 2.04: PF*MS; 2.05: PF*MS; 2.07: PF*OES; 2.08: PF*OES; 2.09: PF*OES/MS |  |  | 1
1
| 1 | 1 | 2.01: < 1; 2.02: < 2; 2.04: 0.4; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: < 1 |  |  | 1
| 1 | 2 | 2.01: < 1; 2.02: < 2; 2.04: 0.5; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: < 1 |  |  | 11
| 1 | 3 | 2.01: < 1; 2.02: < 2; 2.04: 0.5; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: 1 |  |  | 16
| 1 | 4 | 2.01: < 1; 2.02: < 2; 2.04: 0.5; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: < 1 |  |  | < 10
| 1 | 5 | 2.01: < 1; 2.02: < 2; 2.04: 0.5; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: < 1 |  |  | 156
| 1 | 6 | 2.01: < 1; 2.02: < 2; 2.04: 0.5; 2.05: < 10; 2.07: < 50; 2.08: < 10; 2.09: 1
| Mean |  | 2.04: 0.48333; 2.09: 1
| Median |  | 2.04: 0.5; 2.09: 1
| Std Dev. |  | 2.04: 0.040825; 2.09: 0
| Rel.Std.Dev. |  | 2.04: 8.447%; 2.09: 0%
PDM3
| Z-Score (Absolute) |  | 2.01: 0.65; 2.02: 0; 2.04: 0.67; 2.05: 5.22; 2.07: 31.32; 2.08: 5.22; 2.09: 0.44
| Analytical results for Ce in OREAS 163b (Certified Value 10.5 ppm) | [22 empty] | Certified
Ce | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 1
2
| 1 | 1 | 2.01: 11.5; 2.02: 10.2; 2.04: 10.7498; 2.05: 10.3; 2.07: 10.7; 2.08: 12; 2.09: 10.2; 2.10: 11.8; 2.11: 11 |  |  | 1
| 1 | 2 | 2.01: 11.3; 2.02: 10; 2.04: 10.3191; 2.05: 10.4; 2.07: 10.4; 2.08: 12; 2.09: 10.8; 2.10: 11.5; 2.11: 11.4 |  |  | 31
| 1 | 3 | 2.01: 11.6; 2.02: 10.1; 2.04: 10.42; 2.05: 10.4; 2.07: 10.7; 2.08: 11; 2.09: 10.5; 2.10: 10.9; 2.11: 10.9 |  |  | 16
| 1 | 4 | 2.01: 11.1; 2.02: 10.2; 2.04: 10.7808; 2.05: 9.3; 2.07: 10; 2.08: 12; 2.09: 10.4; 2.10: 11.5; 2.11: 11 |  |  | 10.4998
| 1 | 5 | 2.01: 11.5; 2.02: 10.4; 2.04: 10.0877; 2.05: 10.5; 2.07: 10.5; 2.08: 13; 2.09: 10.7; 2.10: 11.8; 2.11: 10.6 |  |  | 157
| 1 | 6 | 2.01: 11.2; 2.02: 10.8; 2.04: 10.037; 2.05: 10.5; 2.07: 10.7; 2.08: 12; 2.09: 10.6; 2.10: 11.5; 2.11: 10.9
| Mean |  | 2.01: 11.3667; 2.02: 10.2833; 2.04: 10.3991; 2.05: 10.2333; 2.07: 10.5; 2.08: 12; 2.09: 10.5333; 2.10: 11.5; 2.11: 10.9667
| Median |  | 2.01: 11.4; 2.02: 10.2; 2.04: 10.3695; 2.05: 10.4; 2.07: 10.6; 2.08: 12; 2.09: 10.55; 2.10: 11.5; 2.11: 10.95
| Std Dev. |  | 2.01: 0.19664; 2.02: 0.28577; 2.04: 0.31735; 2.05: 0.46332; 2.07: 0.27568; 2.08: 0.63246; 2.09: 0.21602; 2.10: 0.32863; 2.11: 0.2582
| Rel.Std.Dev. |  | 2.01: 1.73%; 2.02: 2.779%; 2.04: 3.052%; 2.05: 4.528%; 2.07: 2.626%; 2.08: 5.27%; 2.09: 2.051%; 2.10: 2.858%; 2.11: 2.354%
| PDM3 |  | 2.01: 8.256%; 2.02: -2.062%; 2.04: -0.9598%; 2.05: -2.538%; 2.07: 0.001484%; 2.08: 14.29%; 2.09: 0.3189%; 2.10: 9.525%; 2.11: 4.446%
| Z-Score (Absolute) |  | 2.01: 2.22; 2.02: 0.61; 2.04: 0.31; 2.05: 0.74; 2.07: 0.04; 2.08: NA; 2.09: 0.04; 2.10: 2.57; 2.11: 1.17
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 163b (Certified Value 239 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  | ppm
|  |  | PF*OES/MS | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 230 | 182.2 | 238.1 | 239.8693 | 250 | 210 | 220 | 240 | 240 | 257 | 229 | 235.2993 | 246 | 250 | 238.8 | 242 | 221 | 247 | 233 |  |  | 1
| 1 | 2 | 232 | 189.9 | 240.7 | 242.4751 | 250 | 230 | 230 | 270 | 220 | 277 | 219 | 254.0344 | 253 | 250 | 240 | 241 | 222 | 245 | 239 |  |  | 32
| 1 | 3 | 227 | 195.8 | 234.7 | 242.3866 | 250 | 220 | 220 | 240 | 230 | 240 | 235 | 261.2124 | 254 | 250 | 243.2 | 242 | 220 | 248 | 234 |  |  | 16
| 1 | 4 | 232 | 202.4 | 235.7 | 245.1929 | 250 | 220 | 220 | 250 | 240 | 258 | 221 | 253.9383 | 253 | 250 | 244.4 | 242 | 222 | 251 | 235 |  |  | 239.2711
| 1 | 5 | 230 | 189.9 | 233.4 | 239.1463 | 260 | 220 | 240 | 220 | 240 | 240 | 220 | 267.3876 | 252 | 250 | 248.7 | 241 | 220 | 219 | 229 |  |  | 158
| 1 | 6 | 228 | 192.2 | 241.6 | 243.7858 | 250 | 230 | 220 | 270 | 240 | 238 | 225 | 237.8472 | 249 | 250 | 242.4 | 241 | 221 | 244 | 235
| Mean |  | 229.8333 | 192.0667 | 237.3667 | 242.1427 | 251.6667 | 221.6667 | 225 | 248.3333 | 235 | 251.6667 | 224.8333 | 251.6199 | 251.1667 | 250 | 242.9167 | 241.5 | 221 | 242.3333 | 234.1667
| Median |  | 230 | 191.05 | 236.9 | 242.4309 | 250 | 220 | 220 | 245 | 240 | 248.5 | 223 | 253.9863 | 252.5 | 250 | 242.8 | 241.5 | 221 | 246 | 234.5
| Std Dev. |  | 2.0412 | 6.745 | 3.3224 | 2.2944 | 4.0825 | 7.5277 | 8.3666 | 19.4079 | 8.3666 | 15.2927 | 6.2102 | 12.7113 | 3.0605 | 0 | 3.5057 | 0.54772 | 0.89443 | 11.6905 | 3.2506
| Rel.Std.Dev. |  | 0.8881% | 3.512% | 1.4% | 0.9475% | 1.622% | 3.396% | 3.718% | 7.815% | 3.56% | 6.077% | 2.762% | 5.052% | 1.219% | 0% | 1.443% | 0.2268% | 0.4047% | 4.824% | 1.388%
| PDM3 |  | -3.944% | -19.73% | -0.7959% | 1.2% | 5.181% | -7.358% | -5.964% | 3.787% | -1.785% | 5.181% | -6.034% | 5.161% | 4.972% | 4.484% | 1.524% | 0.9316% | -7.636% | 1.28% | -2.133%
| Z-Score (Absolute) |  | 0.81 | 3.45 | 0.29 | 0.04 | 0.71 | 1.38 | 1.15 | 0.48 | 0.45 | 0.71 | 1.16 | 0.71 | 0.67 | 0.59 | 0.1 | 0 | 1.43 | 0.06 | 0.51
| Analytical results for Cr2O3 in OREAS 163b (Certified Value < 100 ppm) | [22 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.01 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.19 | 2.20 | 2.21 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | [7 empty] | 0
0
| 1 | 1 | < 73 | 24.0064 | < 100 | 200 | 300 | < 100 | 100 | 73 | 23.4 | < 73 | < 150 | < 20 | 40.9 | 33.6 | [7 empty] | 1
| 1 | 2 | < 73 | 23.2407 | < 100 | 100 | 100 | < 100 | 100 | 58 | 27.8 | < 73 | < 150 | < 20 | 40.9 | 36.5 | [7 empty] | 0
| 1 | 3 | < 73 | 24.4958 | < 100 | 100 | 100 | < 100 | 100 | 44 | 24.8 | < 73 | < 150 | 30 | 40.9 | 35.1 | [7 empty] | 16
| 1 | 4 | < 73 | 24.5528 | < 100 | 100 | 100 | < 100 | 100 | 58 | 23.4 | < 73 | < 150 | 20 | 40.9 | 35.1 | [7 empty] | < 100
| 1 | 5 | < 73 | 23.4713 | < 100 | 100 | 100 | < 100 | 200 | 117 | 21.9 | < 73 | < 150 | 20 | 42.4 | 33.6 | [7 empty] | 159
| 1 | 6 | < 73 | 23.7088 | < 100 | 100 | 100 | < 100 | 500 | 58 | 21.9 | < 73 | < 150 | 30 | 42.4 | 35.1
| Mean |  |  | 23.9126 |  | 116.6667 | 133.3333 |  | 183.3333 | 68 | 23.8667 |  |  | 25 | 41.4 | 34.8333
| Median |  |  | 23.8576 |  | 100 | 100 |  | 100 | 58 | 23.4 |  |  | 25 | 40.9 | 35.1
| Std Dev. |  |  | 0.53788 |  | 40.8248 | 81.6497 |  | 160.2082 | 25.6982 | 2.2142 |  |  | 5.7735 | 0.7746 | 1.0985
| Rel.Std.Dev. |  |  | 2.249% |  | 34.99% | 61.24% |  | 87.39% | 37.79% | 9.277% |  |  | 23.09% | 1.871% | 3.154%
PDM3
| Z-Score (Absolute) |  | 0.28 | 0.67 | 0.13 | 2.19 | 2.71 | 0.13 | 4.25 | 0.7 | 0.67 | 0.28 | 0.85 | 0.79 | 0.13 | 0.34
| Analytical results for Cs in OREAS 163b (Certified Value 0.98 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.23 | 2.24 | [10 empty] | ppm
|  |  | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*OES/MS | [10 empty] | 2
3
| 1 | 1 | 1.01 | 1 | 0.96931 | 3.6 | 0.9 | 0.89157 | 1 | 4 | 1.1 | 0.4 | 1 | [10 empty] | 1
| 1 | 2 | 0.96 | 1 | 1.0074 | 2.5 | 0.9 | 0.87601 | 1 | 3 | 1 | 0.7 | 1 | [10 empty] | 13
| 1 | 3 | 1.06 | 0.9 | 1.0079 | 2.3 | 1.1 | 0.85028 | 1 | 5 | 1 | 0.2 | 1 | [10 empty] | 16
| 1 | 4 | 1.01 | 1 | 1.005 | 2.3 | 1.2 | 0.78853 | 0.9 | 3 | 1 | 0.3 | 1 | [10 empty] | 0.98463
| 1 | 5 | 1.05 | 1 | 0.98555 | 2.8 | 1 | 0.78327 | 1 | 8 | 1.1 |  | 1 | [10 empty] | 160
| 1 | 6 | 1.03 | 1 | 0.95814 | 1.6 | 1 | 0.81941 | 1 | < 2 | 1.1 | 0.3 | 1
| Mean |  | 1.02 | 0.98333 | 0.98889 | 2.5167 | 1.0167 | 0.83485 | 0.98333 | 4.6 | 1.05 | 0.38 | 1
| Median |  | 1.02 | 1 | 0.99526 | 2.4 | 1 | 0.83485 | 1 | 4 | 1.05 | 0.3 | 1
| Std Dev. |  | 0.035777 | 0.040825 | 0.02147 | 0.66156 | 0.1169 | 0.045154 | 0.040825 | 2.0736 | 0.054772 | 0.19235 | 0
| Rel.Std.Dev. |  | 3.508% | 4.152% | 2.171% | 26.29% | 11.5% | 5.409% | 4.152% | 45.08% | 5.216% | 50.62% | 0%
| PDM3 |  | 3.592% | -0.132% | 0.4321% | 155.6% | 3.253% | -15.21% | -0.132% | 367.2% | 6.639% | -61.41% | 1.561%
| Z-Score (Absolute) |  | 0.67 | 0.56 | 0.37 | 51.14 | 0.56 | 5.57 | 0.56 | 101.15 | 1.69 | 20.9 | 0
| Analytical results for Cu in OREAS 163b (Certified Value 1.66 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.67 | 1.5135 | 1.514 | 1.69 | 1.565 | 1.655 | 1.635 | 1.605 | > 1 | 1.67 | 1.7666 | 1.7157 | 1.701 | 1.76 | 1.5782 | 1.71 | 1.7454 |  |  |  |  | 1
| 1 | 2 | 1.7 | 1.5215 | 1.534 | 1.665 | 1.575 | 1.64 | 1.6 | 1.61 | > 1 | 1.69 | 1.7594 | 1.7081 | 1.764 | 1.77 | 1.5821 | 1.7275 | 1.7301 |  |  |  |  | 35
| 1 | 3 | 1.72 | 1.54 | 1.565 | 1.665 | 1.595 | 1.67 | 1.615 | 1.64 | > 1 | 1.66 | 1.7217 | 1.7146 | 1.712 | 1.76 | 1.5705 | 1.6828 | 1.721 |  |  |  |  | 16
| 1 | 4 | 1.73 | 1.5337 | 1.555 | 1.72 | 1.665 | 1.66 | 1.64 | 1.63 | > 1 | 1.68 | 1.6925 | 1.708 | 1.707 | 1.76 | 1.5913 | 1.6749 | 1.7575 |  |  |  |  | 1.6602
| 1 | 5 | 1.7 | 1.5211 | 1.527 | 1.705 | 1.655 | 1.685 | 1.61 | 1.62 | > 1 | 1.66 | 1.7237 | 1.7069 | 1.701 | 1.76 | 1.5738 | 1.5094 | 1.7064 |  |  |  |  | 161
| 1 | 6 | 1.72 | 1.5211 | 1.548 | 1.675 | 1.665 | 1.665 | 1.6 | 1.605 | > 1 | 1.67 | 1.7171 | 1.7068 | 1.689 | 1.77 | 1.5544 | 1.6791 | 1.7732
| Mean |  | 1.7067 | 1.5252 | 1.5405 | 1.6867 | 1.62 | 1.6625 | 1.6167 | 1.6183 |  | 1.6717 | 1.7302 | 1.71 | 1.7123 | 1.7633 | 1.5751 | 1.664 | 1.7389
| Median |  | 1.71 | 1.5213 | 1.541 | 1.6825 | 1.625 | 1.6625 | 1.6125 | 1.615 |  | 1.67 | 1.7227 | 1.7081 | 1.704 | 1.76 | 1.576 | 1.6809 | 1.7377
| Std Dev. |  | 0.021602 | 0.0097309 | 0.018939 | 0.022509 | 0.046797 | 0.015083 | 0.017224 | 0.014376 |  | 0.01169 | 0.027864 | 0.0040276 | 0.026455 | 0.005164 | 0.012429 | 0.078421 | 0.024572
| Rel.Std.Dev. |  | 1.266% | 0.638% | 1.229% | 1.335% | 2.889% | 0.9073% | 1.065% | 0.8883% |  | 0.6993% | 1.61% | 0.2355% | 1.545% | 0.2929% | 0.7891% | 4.713% | 1.413%
| PDM3 |  | 2.801% | -8.132% | -7.208% | 1.597% | -2.419% | 0.1411% | -2.62% | -2.519% |  | 0.6932% | 4.218% | 3.003% | 3.143% | 6.215% | -5.126% | 0.2284% | 4.745%
| Z-Score (Absolute) |  | 0.54 | 1.98 | 1.76 | 0.26 | 0.66 | 0.07 | 0.71 | 0.69 | NA | 0.05 | 0.86 | 0.59 | 0.62 | 1.32 | 1.29 | 0.05 | 0.99
| Analytical results for Dy in OREAS 163b (Certified Value 1.32 ppm) | [22 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | 1 | 1.4 | 1.4886 | 1.3 | 1.23 | < 2 | 1.23 | 1.4 | 1.31 | [12 empty] | 1
| 1 | 2 | 1 | 1.4 | 1.4934 | 1.3 | 1.28 | < 2 | 1.29 | 1.36 | 1.28 | [12 empty] | 3
| 1 | 3 | 1 | 1.2 | 1.4824 | 1.2 | 1.29 | < 2 | 1.26 | 1.29 | 1.26 | [12 empty] | 16
| 1 | 4 | 1 | 1.3 | 1.4589 | 1.5 | 1.11 | < 2 | 1.34 | 1.35 | 1.32 | [12 empty] | 1.3233
| 1 | 5 | 1 | 1.4 | 1.4034 | 1.8 | 1.2 | < 2 | 1.24 | 1.43 | 1.2 | [12 empty] | 162
| 1 | 6 | 1 | 1.4 | 1.4706 | 1.3 | 1.22 | < 2 | 1.31 | 1.27 | 1.29
| Mean |  | 2.01: 1; 2.06: 1.35; 2.07: 1.4662; 2.08: 1.4; 2.09: 1.2217; 2.11: 1.2783; 2.12: 1.35; 2.13: 1.2767
| Median |  | 2.01: 1; 2.06: 1.4; 2.07: 1.4765; 2.08: 1.3; 2.09: 1.225; 2.11: 1.275; 2.12: 1.355; 2.13: 1.285
| Std Dev. |  | 2.01: 0; 2.06: 0.083666; 2.07: 0.033217; 2.08: 0.21909; 2.09: 0.064936; 2.11: 0.042622; 2.12: 0.061644; 2.13: 0.043205
| Rel.Std.Dev. |  | 2.01: 0%; 2.06: 6.197%; 2.07: 2.266%; 2.08: 15.65%; 2.09: 5.315%; 2.11: 3.334%; 2.12: 4.566%; 2.13: 3.384%
| PDM3 |  | 2.01: -24.43%; 2.06: 2.02%; 2.07: 10.8%; 2.08: 5.799%; 2.09: -7.678%; 2.11: -3.396%; 2.12: 2.02%; 2.13: -3.522%
| Z-Score (Absolute) |  | NA | 0.39 | 1.66 | 0.94 | 1.01 | 3.44 | 0.39 | 0.39 | 0.41
Results from laboratory 2.04 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Er in OREAS 163b (Certified Value 0.58 ppm) | [22 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
3
| 1 | 1 | < 0.5 | 0.6 | 0.5391 | 0.5 | 0.56 | < 2 | 0.58 | 0.53 | 0.63 | [12 empty] | 1
| 1 | 2 | < 0.5 | 0.6 | 0.52978 | 0.8 | 0.55 | < 2 | 0.61 | 0.6 | 0.66 | [12 empty] | 4
| 1 | 3 | < 0.5 | 0.6 | 0.53531 | 0.6 | 0.56 | < 2 | 0.6 | 0.58 | 0.59 | [12 empty] | 16
| 1 | 4 | < 0.5 | 0.6 | 0.5338 | 0.6 | 0.53 | < 2 | 0.6 | 0.63 | 0.66 | [12 empty] | 0.58143
| 1 | 5 | < 0.5 | 0.5 | 0.54468 | 0.6 | 0.54 | < 2 | 0.6 | 0.62 | 0.6 | [12 empty] | 163
| 1 | 6 | < 0.5 | 0.5 | 0.53022 | 0.6 | 0.57 | < 2 | 0.58 | 0.6 | 0.65
| Mean |  | 2.06: 0.56667; 2.07: 0.53548; 2.08: 0.61667; 2.09: 0.55167; 2.11: 0.595; 2.12: 0.59333; 2.13: 0.63167
| Median |  | 2.06: 0.6; 2.07: 0.53455; 2.08: 0.6; 2.09: 0.555; 2.11: 0.6; 2.12: 0.6; 2.13: 0.64
| Std Dev. |  | 2.06: 0.05164; 2.07: 0.0056721; 2.08: 0.098319; 2.09: 0.01472; 2.11: 0.012247; 2.12: 0.03559; 2.13: 0.030605
| Rel.Std.Dev. |  | 2.06: 9.113%; 2.07: 1.059%; 2.08: 15.94%; 2.09: 2.668%; 2.11: 2.058%; 2.12: 5.998%; 2.13: 4.845%
| PDM3 |  | 2.06: -2.539%; 2.07: -7.903%; 2.08: 6.06%; 2.09: -5.119%; 2.11: 2.334%; 2.12: 2.047%; 2.13: 8.64%
| Z-Score (Absolute) |  | 5.56 | NA | 0.94 | NA | 0.67 | 6.58 | 0.03 | 0 | 0.62
Results from laboratories 2.05 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 163b (Certified Value 0.18 ppm) | [22 empty] | Certified
Eu | Round | Replicate | 2.01: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | [7 empty] | 1
| No | No | 2.01: 2.04; 2.06: 2.05; 2.07: 2.06; 2.08: 2.12; 2.09: 2.13; 2.10: 2.18; 2.11: 2.19; 2.12: 2.24 | [7 empty] | ppm
|  |  | 2.01: PF*MS; 2.06: BF*MS; 2.07: PF*OES/MS; 2.08: PF*OES/MS; 2.09: PF*MS; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*OES/MS | [7 empty] | 2
2
| 1 | 1 | 2.01: < 0.5; 2.06: 0.2; 2.07: 0.2029; 2.08: 0.1; 2.09: 0.19; 2.10: < 2; 2.11: 0.18; 2.12: 0.18 | [7 empty] | 1
| 1 | 2 | 2.01: < 0.5; 2.06: 0.2; 2.07: 0.20327; 2.08: 0.1; 2.09: 0.17; 2.10: < 2; 2.11: 0.19; 2.12: 0.19 | [7 empty] | 5
| 1 | 3 | 2.01: < 0.5; 2.06: 0.2; 2.07: 0.20763; 2.08: 0.2; 2.09: 0.18; 2.10: < 2; 2.11: 0.17; 2.12: 0.17 | [7 empty] | 16
| 1 | 4 | 2.01: < 0.5; 2.06: 0.2; 2.07: 0.19722; 2.08: 0.1; 2.09: 0.17; 2.10: < 2; 2.11: 0.18; 2.12: 0.2 | [7 empty] | 0.1831
| 1 | 5 | 2.01: < 0.5; 2.06: 0.2; 2.07: 0.20159; 2.08: 0.2; 2.09: 0.17; 2.10: < 2; 2.11: 0.2; 2.12: 0.18 | [7 empty] | 164
| 1 | 6 | 2.01: < 0.5; 2.06: 0.2; 2.07: 0.20888; 2.08: 0.2; 2.09: 0.19; 2.10: < 2; 2.11: 0.18; 2.12: 0.18
| Mean |  | 2.06: 0.2; 2.07: 0.20358; 2.08: 0.15; 2.09: 0.17833; 2.11: 0.18333; 2.12: 0.18333
| Median |  | 2.06: 0.2; 2.07: 0.20308; 2.08: 0.15; 2.09: 0.175; 2.11: 0.18; 2.12: 0.18
| Std Dev. |  | 2.06: 3.0405e-17; 2.07: 0.0042297; 2.08: 0.054772; 2.09: 0.0098319; 2.11: 0.010328; 2.12: 0.010328
| Rel.Std.Dev. |  | 2.06: 1.52e-14%; 2.07: 2.078%; 2.08: 36.51%; 2.09: 5.513%; 2.11: 5.633%; 2.12: 5.633%
| PDM3 |  | 2.06: 9.232%; 2.07: 11.19%; 2.08: -18.08%; 2.09: -2.602%; 2.11: 0.129%; 2.12: 0.129%
| Z-Score (Absolute) |  | 2.01: 3.12; 2.06: 0.45; 2.07: 0.64; 2.08: 2.23; 2.09: 0.71; 2.10: 43.17; 2.11: 0.45; 2.12: 0.45
| Analytical results for Fe in OREAS 163b (Certified Value 4.06 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 4.09 | 4.3502 | 3.96 | 4.01 | 4.0278 | 3.966 | 3.868 | 4.022 | 4.008 | 3.987 | 4.14 | 3.95 | 4.0383 | 4.15 | 4.14 | 4.3898 | 4.036 | 4.04 | 4.327 | 4.15 | 4.01 | 1
| 1 | 2 | 4.11 | 4.369 | 3.99 | 4.1 | 4.0985 | 3.938 | 3.868 | 3.917 | 3.938 | 3.973 | 4.11 | 3.99 | 4.1615 | 4.16 | 4.14 | 4.3643 | 4.099 | 4.04 | 4.391 | 4.13 | 4.03 | 0
| 1 | 3 | 4.15 | 4.3252 | 4.05 | 4.09 | 4.0755 | 3.952 | 3.987 | 3.917 | 3.917 | 4.092 | 4.14 | 4.06 | 4.0088 | 4.14 | 3.98 | 4.1674 | 4.113 | 4.06 | 4.337 | 4.05 | 3.96 | 16
| 1 | 4 | 4.2 | 4.3216 | 4.08 | 4.08 | 4.0627 | 3.952 | 4.092 | 3.959 | 4.029 | 4.057 | 4.13 | 3.98 | 3.6574 | 4.13 | 4.13 | 3.9228 | 4.099 | 4.06 | 4.343 | 4.12 | 4.1 | 4.0565
| 1 | 5 | 4.12 | 4.339 | 4.04 | 4.07 | 4.058 | 3.987 | 4.099 | 3.994 | 3.917 | 4.029 | 4.17 | 4.03 | 4.0112 | 4.15 | 4.09 | 4.3182 | 4.113 | 4.05 | 4.325 | 3.64 | 4.01 | 165
| 1 | 6 | 4.16 | 4.3468 | 4.05 | 4.07 | 4.0661 | 3.966 | 4.092 | 3.98 | 4.008 | 4.099 | 4.17 | 3.98 | 3.667 | 4.13 | 4.05 | 4.1505 | 4.099 | 4.05 | 4.324 | 4.03 | 4.07
| Mean |  | 4.1383 | 4.342 | 4.0283 | 4.07 | 4.0648 | 3.9602 | 4.001 | 3.9648 | 3.9695 | 4.0395 | 4.1433 | 3.9983 | 3.924 | 4.1433 | 4.0883 | 4.2188 | 4.0932 | 4.05 | 4.3412 | 4.02 | 4.03
| Median |  | 4.135 | 4.3429 | 4.045 | 4.075 | 4.0644 | 3.959 | 4.0395 | 3.9695 | 3.973 | 4.043 | 4.14 | 3.985 | 4.01 | 4.145 | 4.11 | 4.2428 | 4.099 | 4.05 | 4.332 | 4.085 | 4.02
| Std Dev. |  | 0.039707 | 0.017444 | 0.04446 | 0.031623 | 0.023108 | 0.01681 | 0.11112 | 0.042329 | 0.051009 | 0.05271 | 0.023381 | 0.039707 | 0.21044 | 0.012111 | 0.063692 | 0.17619 | 0.028833 | 0.0089443 | 0.025538 | 0.19204 | 0.049396
| Rel.Std.Dev. |  | 0.9595% | 0.4018% | 1.104% | 0.777% | 0.5685% | 0.4245% | 2.777% | 1.068% | 1.285% | 1.305% | 0.5643% | 0.9931% | 5.363% | 0.2923% | 1.558% | 4.176% | 0.7044% | 0.2208% | 0.5883% | 4.777% | 1.226%
| PDM3 |  | 2.018% | 7.038% | -0.6933% | 0.3339% | 0.2053% | -2.374% | -1.367% | -2.259% | -2.144% | -0.418% | 2.142% | -1.433% | -3.264% | 2.142% | 0.7859% | 4.003% | 0.905% | -0.1591% | 7.019% | -0.8987% | -0.6522%
| Z-Score (Absolute) |  | 1.15 | 3.81 | 0.28 | 0.26 | 0.19 | 1.17 | 0.64 | 1.11 | 1.05 | 0.14 | 1.22 | 0.67 | 1.64 | 1.22 | 0.5 | 2.2 | 0.56 | 0 | 3.8 | 0.39 | 0.26
| Analytical results for Ga in OREAS 163b (Certified Value 1.54 ppm) | [22 empty] | Certified
Ga | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.16; 2.10: 2.19; 2.11: 2.23; 2.12: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: BF*OES/MS; 2.11: PF*MS; 2.12: PF*OES/MS | 2
2
| 1 | 1 | 2.01: 2; 2.02: < 5; 2.04: 1.3; 2.05: 3.2045; 2.06: 1.4; 2.07: 2; 2.08: 1.6379; 2.09: 2; 2.10: 1; 2.11: 2; 2.12: < 5 | 1
| 1 | 2 | 2.01: 1; 2.02: < 5; 2.04: 1.3; 2.05: 3.2408; 2.06: 1.4; 2.07: 2; 2.08: 1.7525; 2.09: 2; 2.10: 1.1; 2.11: 1; 2.12: < 5 | 17
| 1 | 3 | 2.01: 1; 2.02: < 5; 2.04: 1.4; 2.05: 3.1504; 2.06: 1.4; 2.07: 2; 2.08: 1.4704; 2.09: 1; 2.10: 1.1; 2.11: 2; 2.12: < 5 | 16
| 1 | 4 | 2.01: 1; 2.02: < 5; 2.04: 1.5; 2.05: 3.2547; 2.06: 1.1; 2.07: 2; 2.08: 1.5887; 2.09: 2; 2.10: 1.1; 2.11: 2; 2.12: < 5 | 1.5385
| 1 | 5 | 2.01: 1; 2.02: < 5; 2.04: 1.4; 2.05: 3.1124; 2.06: 1.9; 2.07: 1; 2.08: 1.8237; 2.09: 1; 2.10: 1.1; 2.11: 1; 2.12: < 5 | 166
| 1 | 6 | 2.01: 2; 2.02: < 5; 2.04: 1.6; 2.05: 3.0892; 2.06: 2.1; 2.07: 2; 2.08: 1.5458; 2.09: 1; 2.10: 1.1; 2.11: 1; 2.12: < 5
| Mean |  | 2.01: 1.3333; 2.04: 1.4167; 2.05: 3.1753; 2.06: 1.55; 2.07: 1.8333; 2.08: 1.6365; 2.09: 1.5; 2.10: 1.0833; 2.11: 1.5
| Median |  | 2.01: 1; 2.04: 1.4; 2.05: 3.1774; 2.06: 1.4; 2.07: 2; 2.08: 1.6133; 2.09: 1.5; 2.10: 1.1; 2.11: 1.5
| Std Dev. |  | 2.01: 0.5164; 2.04: 0.1169; 2.05: 0.06849; 2.06: 0.37283; 2.07: 0.40825; 2.08: 0.1316; 2.09: 0.54772; 2.10: 0.040825; 2.11: 0.54772
| Rel.Std.Dev. |  | 2.01: 38.73%; 2.04: 8.252%; 2.05: 2.157%; 2.06: 24.05%; 2.07: 22.27%; 2.08: 8.042%; 2.09: 36.51%; 2.10: 3.768%; 2.11: 36.51%
| PDM3 |  | 2.01: -13.34%; 2.04: -7.922%; 2.05: 106.4%; 2.06: 0.7444%; 2.07: 19.16%; 2.08: 6.366%; 2.09: -2.505%; 2.10: -29.59%; 2.11: -2.505%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 2.96; 2.04: 0.41; 2.05: 5.06; 2.06: 0; 2.07: 0.88; 2.08: 0.27; 2.09: 0.16; 2.10: 1.45; 2.11: 0.16; 2.12: 2.96
| Analytical results for Gd in OREAS 163b (Certified Value 1.44 ppm) | [22 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.18; 2.08: 2.19; 2.09: 2.24 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: BF*OES/MS; 2.09: PF*OES/MS | 2
3
| 1 | 1 | 2.01: 1.4; 2.02: 1.55; 2.04: 1.3132; 2.05: 1.9; 2.06: 1.59; 2.07: < 2; 2.08: 1.58; 2.09: < 3 | 1
| 1 | 2 | 2.01: 1.4; 2.02: 1.33; 2.04: 1.2025; 2.05: 2.1; 2.06: 1.63; 2.07: < 2; 2.08: 1.59; 2.09: < 3 | 6
| 1 | 3 | 2.01: 1.4; 2.02: 1.57; 2.04: 1.2255; 2.05: 1.6; 2.06: 1.56; 2.07: < 2; 2.08: 1.54; 2.09: < 3 | 16
| 1 | 4 | 2.01: 1.4; 2.02: 1.34; 2.04: 1.2374; 2.05: 1.5; 2.06: 1.49; 2.07: < 2; 2.08: 1.6; 2.09: < 3 | 1.4397
| 1 | 5 | 2.01: 1.4; 2.02: 1.35; 2.04: 1.2257; 2.05: 1.9; 2.06: 1.51; 2.07: < 2; 2.08: 1.54; 2.09: < 3 | 167
| 1 | 6 | 2.01: 1.4; 2.02: 1.51; 2.04: 1.2431; 2.05: 2.2; 2.06: 1.61; 2.07: < 2; 2.08: 1.54; 2.09: < 3
| Mean |  | 2.01: 1.4; 2.02: 1.4417; 2.04: 1.2412; 2.05: 1.8667; 2.06: 1.565; 2.08: 1.565
| Median |  | 2.01: 1.4; 2.02: 1.43; 2.04: 1.2316; 2.05: 1.9; 2.06: 1.575; 2.08: 1.56
| Std Dev. |  | 2.01: 2.4324e-16; 2.02: 0.11321; 2.04: 0.037913; 2.05: 0.27325; 2.06: 0.055767; 2.08: 0.028107
| Rel.Std.Dev. |  | 2.01: 1.737e-14%; 2.02: 7.853%; 2.04: 3.054%; 2.05: 14.64%; 2.06: 3.563%; 2.08: 1.796%
| PDM3 |  | 2.01: -2.758%; 2.02: 0.1363%; 2.04: -13.78%; 2.05: 29.66%; 2.06: 8.703%; 2.08: 8.703%
| Z-Score (Absolute) |  | 2.01: 0.51; 2.02: 0.21; 2.04: 1.64; 2.05: 2.83; 2.06: 0.67; 2.07: 3.37; 2.08: 0.67; 2.09: 0.21
| Analytical results for Ge in OREAS 163b (Certified Value 1.03 ppm) | [22 empty] | Certified
Ge | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.06; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.23; 2.09: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES/MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES/MS | 2
2
| 1 | 1 | 2.01: 1; 2.02: 1.1673; 2.04: 1.2; 2.05: 1; 2.06: < 1; 2.07: 1; 2.09: < 2 | 1
| 1 | 2 | 2.01: 1; 2.02: 1.1787; 2.04: 1; 2.05: < 1; 2.06: < 1; 2.07: 1; 2.08: 1; 2.09: < 2 | 19
| 1 | 3 | 2.01: 1; 2.02: 1.1725; 2.04: 1.3; 2.05: 1; 2.06: < 1; 2.07: 1; 2.09: < 2 | 16
| 1 | 4 | 2.01: 1; 2.02: 1.1817; 2.04: 1.3; 2.05: < 1; 2.06: < 1; 2.07: 1; 2.09: < 2 | 1.0349
| 1 | 5 | 2.01: 1; 2.02: 1.1834; 2.04: 1.1; 2.05: 1; 2.06: < 1; 2.07: 1; 2.09: < 2 | 168
| 1 | 6 | 2.01: 1; 2.02: 1.1738; 2.04: 1.3; 2.05: 1; 2.06: < 1; 2.07: 1; 2.09: < 2
| Mean |  | 2.01: 1; 2.02: 1.1763; 2.04: 1.2; 2.05: 1; 2.07: 1; 2.08: 1
| Median |  | 2.01: 1; 2.02: 1.1763; 2.04: 1.25; 2.05: 1; 2.07: 1; 2.08: 1
| Std Dev. |  | 2.01: 0; 2.02: 0.0061052; 2.04: 0.12649; 2.05: 0; 2.07: 0
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 0.519%; 2.04: 10.54%; 2.05: 0%; 2.07: 0%
| PDM3 |  | 2.01: -3.375%; 2.02: 13.66%; 2.04: 15.95%; 2.05: -3.375%; 2.07: -3.375%; 2.08: -3.375%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 1.43; 2.04: 1.62; 2.05: 1.35; 2.06: 4.05; 2.07: 0; 2.08: 0; 2.09: 0
| Analytical results for Hf in OREAS 163b (Indicative Value 0.6 ppm) | [22 empty] | Indicative
Hf | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.19; 2.08: 2.22 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: BF*OES/MS; 2.08: BF*MS | 2
2
| 1 | 1 | 2.01: < 0.5; 2.02: 0.6; 2.04: 0.63381; 2.05: < 10; 2.06: < 1; 2.07: 0.6; 2.08: 0.4 | 1
| 1 | 2 | 2.01: < 0.5; 2.02: 0.7; 2.04: 0.6163; 2.05: < 10; 2.06: < 1; 2.07: 0.6; 2.08: 0.6 | 23
| 1 | 3 | 2.01: < 0.5; 2.02: 0.6; 2.04: 0.65299; 2.05: < 10; 2.06: < 1; 2.07: 0.6; 2.08: 0.4 | 16
| 1 | 4 | 2.01: < 0.5; 2.02: 0.7; 2.04: 0.67575; 2.05: < 10; 2.06: < 1; 2.07: 0.7; 2.08: 0.5 | 0.60327
| 1 | 5 | 2.01: < 0.5; 2.02: 0.8; 2.04: 0.64996; 2.05: < 10; 2.06: < 1; 2.07: 0.6; 2.08: 0.4 | 29
| 1 | 6 | 2.01: < 0.5; 2.02: 0.8; 2.04: 0.64972; 2.05: < 10; 2.06: < 1; 2.07: 0.6; 2.08: 0.4
| Mean |  | 2.02: 0.7; 2.04: 0.64642; 2.07: 0.61667; 2.08: 0.45
| Median |  | 2.02: 0.7; 2.04: 0.64984; 2.07: 0.6; 2.08: 0.4
| Std Dev. |  | 2.02: 0.089443; 2.04: 0.019963; 2.07: 0.040825; 2.08: 0.083666
| Rel.Std.Dev. |  | 2.02: 12.78%; 2.04: 3.088%; 2.07: 6.62%; 2.08: 18.59%
| PDM3 |  | 2.02: 16.03%; 2.04: 7.153%; 2.07: 2.22%; 2.08: -25.41%
| Z-Score (Absolute) |  | 2.01: 2.12; 2.02: 0.48; 2.04: 0.17; 2.05: 25.33; 2.06: 0.67; 2.07: 0; 2.08: 0.96
| Analytical results for Hg in OREAS 163b (Indicative Value < 5 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: PF*OES | 2
2
| 1 | 1 | 2.01: < 5 | 1
| 1 | 2 | 2.01: < 5 | 1
| 1 | 3 | 2.01: < 5 | 16
| 1 | 4 | 2.01: < 5 | < 5
| 1 | 5 | 2.01: < 5 | 27
| 1 | 6 | 2.01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01
| Analytical results for Ho in OREAS 163b (Certified Value 0.23 ppm) | [22 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.06: 2.13; 2.07: 2.18; 2.08: 2.19; 2.09: 2.22; 2.10: 2.24 | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: BF*OES/MS; 2.09: BF*MS; 2.10: PF*OES/MS | 2
3
| 1 | 1 | 2.01: < 0.5; 2.02: 0.3; 2.04: 0.23458; 2.05: 0.2; 2.06: 0.23; 2.07: < 2; 2.08: 0.23; 2.09: 0.25; 2.10: 0.24 | 1
| 1 | 2 | 2.01: < 0.5; 2.02: 0.3; 2.04: 0.22703; 2.05: < 0.2; 2.06: 0.22; 2.07: < 2; 2.08: 0.24; 2.09: 0.23; 2.10: 0.25 | 3
| 1 | 3 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.23096; 2.05: < 0.2; 2.06: 0.23; 2.07: < 2; 2.08: 0.22; 2.09: 0.24; 2.10: 0.23 | 16
| 1 | 4 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.22253; 2.05: < 0.2; 2.06: 0.21; 2.07: < 2; 2.08: 0.23; 2.09: 0.24; 2.10: 0.24 | 0.23194
| 1 | 5 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.22829; 2.05: < 0.2; 2.06: 0.2; 2.07: < 2; 2.08: 0.22; 2.09: 0.26; 2.10: 0.23 | 169
| 1 | 6 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.22491; 2.05: 0.2; 2.06: 0.23; 2.07: < 2; 2.08: 0.23; 2.09: 0.25; 2.10: 0.24
| Mean |  | 2.02: 0.23333; 2.04: 0.22805; 2.05: 0.2; 2.06: 0.22; 2.08: 0.22833; 2.09: 0.245; 2.10: 0.23833
| Median |  | 2.02: 0.2; 2.04: 0.22766; 2.05: 0.2; 2.06: 0.225; 2.08: 0.23; 2.09: 0.245; 2.10: 0.24
| Std Dev. |  | 2.02: 0.05164; 2.04: 0.0043031; 2.05: 0; 2.06: 0.012649; 2.08: 0.0075277; 2.09: 0.010488; 2.10: 0.0075277
| Rel.Std.Dev. |  | 2.02: 22.13%; 2.04: 1.887%; 2.05: 0%; 2.06: 5.75%; 2.08: 3.297%; 2.09: 4.281%; 2.10: 3.158%
| PDM3 |  | 2.02: 0.5992%; 2.04: -1.678%; 2.05: -13.77%; 2.06: -5.149%; 2.08: -1.557%; 2.09: 5.629%; 2.10: 2.755%
| Z-Score (Absolute) |  | 2.01: 0.77; 2.02: NA; 2.04: 0.67; 2.05: NA; 2.06: 1.2; 2.07: 49.95; 2.08: 0.66; 2.09: 0.44; 2.10: 0
Results from laboratories 2.05 and 2.12 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 163b (Certified Value 1.13 ppm) | [22 empty] | Certified
In | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.12; 2.04: 2.13; 2.05: 2.14; 2.06: 2.16; 2.07: 2.23; 2.08: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES/MS; 2.04: PF*MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*OES/MS | 2
2
| 1 | 1 | 2.01: 1.1; 2.02: 1.2; 2.04: 1.1; 2.05: 1.0328; 2.06: 1; 2.07: 1; 2.08: 1.4 | 1
| 1 | 2 | 2.01: 1.2; 2.02: 1.2; 2.04: 1.1; 2.05: 1.0742; 2.06: 1; 2.07: 1.1; 2.08: 1.4 | 10
| 1 | 3 | 2.01: 1.1; 2.02: 1; 2.04: 1.2; 2.05: 1.0558; 2.06: 1.1; 2.07: 1.1; 2.08: 1.3 | 16
| 1 | 4 | 2.01: 1.2; 2.02: 1.3; 2.04: 1.1; 2.05: 1.0177; 2.06: 1.1; 2.07: 1; 2.08: 1.4 | 1.1293
| 1 | 5 | 2.01: 1; 2.02: 1.2; 2.04: 1.1; 2.05: 0.97049; 2.06: 1.1; 2.07: 1; 2.08: 1.3 | 170
| 1 | 6 | 2.01: 1.1; 2.02: 1.2; 2.04: 1.1; 2.05: 0.97925; 2.06: 1.1; 2.07: 1.1; 2.08: 1.3
| Mean |  | 2.01: 1.1167; 2.02: 1.1833; 2.04: 1.1167; 2.05: 1.0217; 2.06: 1.0667; 2.07: 1.05; 2.08: 1.35
| Median |  | 2.01: 1.1; 2.02: 1.2; 2.04: 1.1; 2.05: 1.0253; 2.06: 1.1; 2.07: 1.05; 2.08: 1.35
| Std Dev. |  | 2.01: 0.075277; 2.02: 0.098319; 2.04: 0.040825; 2.05: 0.041175; 2.06: 0.05164; 2.07: 0.054772; 2.08: 0.054772
| Rel.Std.Dev. |  | 2.01: 6.741%; 2.02: 8.309%; 2.04: 3.656%; 2.05: 4.03%; 2.06: 4.841%; 2.07: 5.216%; 2.08: 4.057%
| PDM3 |  | 2.01: -1.118%; 2.02: 4.786%; 2.04: -1.118%; 2.05: -9.527%; 2.06: -5.545%; 2.07: -7.021%; 2.08: 19.54%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.04: 0; 2.05: 0.96; 2.06: 0.51; 2.07: 0.67; 2.08: 2.36
| Analytical results for K in OREAS 163b (Certified Value 0.108 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | 3
3
| 1 | 1 | 0.09 | 0.27945 | 0.08 | 0.09 | 0.14336 | 0.058 | < 0.05 | 0.149 | 0.091 | 0.116 | 0.1 | 0.1 | 0.25881 | 0.15 | 0.2 | 0.2075 | 0.091 | 0.13 | 0.199 | 0.11 | 0.12 | 1
| 1 | 2 | 0.08 | 0.2725 | 0.08 | 0.09 | 0.14874 | 0.058 | < 0.05 | 0.141 | 0.1 | 0.091 | 0.1 | 0.1 | 0.23002 | 0.17 | 0.3 | 0.1919 | 0.083 | 0.14 | 0.207 | 0.11 | 0.09 | 0
| 1 | 3 | 0.08 | 0.27785 | 0.08 | 0.08 | 0.14662 | 0.075 | < 0.05 | 0.1 | 0.083 | 0.091 | < 0.1 | 0.1 | 0.26322 | 0.15 | 0.3 | 0.2274 | 0.091 | 0.13 | 0.189 | 0.12 | 0.11 | 16
| 1 | 4 | 0.09 | 0.27472 | 0.08 | 0.09 | 0.14353 | 0.075 | < 0.05 | 0.149 | 0.091 | 0.116 | < 0.1 | 0.1 | 0.25461 | 0.14 | 0.2 | 0.2217 | 0.091 | 0.14 | 0.199 | 0.11 | 0.1 | 0.10762
| 1 | 5 | 0.09 | 0.27212 | 0.08 | 0.09 | 0.14338 | 0.066 | < 0.05 | 0.149 | 0.066 | 0.091 | 0.1 | 0.1 | 0.24929 | 0.15 | 0.2 | 0.1868 | 0.091 | 0.14 | 0.196 | 0.1 | 0.11 | 171
| 1 | 6 | 0.09 | 0.27904 | 0.08 | 0.09 | 0.14317 | 0.075 | < 0.05 | 0.149 | 0.1 | 0.1 | 0.1 | 0.1 | 0.23187 | 0.15 | 0.2 | 0.1484 | 0.091 | 0.13 | 0.202 | 0.11 | 0.12
| Mean |  | 0.086667 | 0.27595 | 0.08 | 0.088333 | 0.1448 | 0.067833 |  | 0.1395 | 0.0885 | 0.10083 | 0.1 | 0.1 | 0.24797 | 0.15167 | 0.23333 | 0.19728 | 0.089667 | 0.135 | 0.19867 | 0.11 | 0.10833
| Median |  | 0.09 | 0.27629 | 0.08 | 0.09 | 0.14345 | 0.0705 |  | 0.149 | 0.091 | 0.0955 | 0.1 | 0.1 | 0.25195 | 0.15 | 0.2 | 0.1997 | 0.091 | 0.135 | 0.199 | 0.11 | 0.11
| Std Dev. |  | 0.005164 | 0.0032698 | 0 | 0.0040825 | 0.0023319 | 0.0083766 |  | 0.019614 | 0.012755 | 0.012254 | 0 | 1.5202e-17 | 0.013981 | 0.0098319 | 0.05164 | 0.028761 | 0.003266 | 0.0054772 | 0.0060222 | 0.0063246 | 0.01169
| Rel.Std.Dev. |  | 5.958% | 1.185% | 0% | 4.622% | 1.61% | 12.35% |  | 14.06% | 14.41% | 12.15% | 0% | 1.52e-14% | 5.638% | 6.483% | 22.13% | 14.58% | 3.642% | 4.057% | 3.031% | 5.75% | 10.79%
| PDM3 |  | -19.47% | 156.4% | -25.66% | -17.92% | 34.55% | -36.97% |  | 29.62% | -17.77% | -6.306% | -7.08% | -7.08% | 130.4% | 40.93% | 116.8% | 83.31% | -16.68% | 25.44% | 84.6% | 2.212% | 0.6629%
| Z-Score (Absolute) |  | 0.51 | 3.76 | 0.66 | 0.47 | 0.8 | 0.93 | 1.9 | 0.69 | 0.47 | 0.18 | NA | NA | 3.13 | 0.96 | NA | 1.99 | 0.43 | 0.58 | 2.02 | 0.02 | 0.02
Results from laboratories 2.12, 2.13 and 2.17 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for La in OREAS 163b (Certified Value 5.26 ppm) | [22 empty] | Certified
La | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab | 1
| No | No | 2.01: 2.02; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.20; 2.11: 2.21; 2.12: 2.22; 2.13: 2.24 | ppm
|  |  | 2.01: PF*OES; 2.02: PF*MS; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: PF*OES; 2.11: PF*OES; 2.12: BF*MS; 2.13: PF*OES/MS | 2
3
| 1 | 1 | 2.01: 3.78; 2.02: 5.5; 2.04: 4.9; 2.05: 15.3851; 2.06: 5.5; 2.07: 4.8; 2.08: 7; 2.09: 5.3; 2.10: 6; 2.11: < 10; 2.12: 5.7; 2.13: 5.31 | 1
| 1 | 2 | 2.01: 3.99; 2.02: 5.4; 2.04: 5; 2.05: 15.4994; 2.06: 5.4; 2.07: 4.7; 2.08: 8; 2.09: 5.4; 2.10: 6; 2.11: < 10; 2.12: 5.8; 2.13: 5.26 | 11
| 1 | 3 | 2.01: 3.95; 2.02: 5.5; 2.04: 5; 2.05: 15.4962; 2.06: 5.1; 2.07: 5; 2.08: 7; 2.09: 5.3; 2.10: 6; 2.11: < 10; 2.12: 5.4; 2.13: 5.43 | 16
| 1 | 4 | 2.01: 4.05; 2.02: 5.3; 2.04: 4.9; 2.05: 15.3105; 2.06: 5.1; 2.07: 4.6; 2.08: 8; 2.09: 5.1; 2.10: 6; 2.11: < 10; 2.12: 5.5; 2.13: 5.42 | 5.2607
| 1 | 5 | 2.01: 4.19; 2.02: 5.6; 2.04: 5; 2.05: 14.6324; 2.06: 5.2; 2.07: 4.8; 2.08: 8; 2.09: 5.5; 2.10: 6; 2.11: < 10; 2.12: 5.9; 2.13: 5.37 | 172
| 1 | 6 | 2.01: 4.21; 2.02: 5.4; 2.04: 5; 2.05: 15.6984; 2.06: 5.3; 2.07: 5; 2.08: 8; 2.09: 5.3; 2.10: 6; 2.11: < 10; 2.12: 5.6; 2.13: 5.36
| Mean |  | 2.01: 4.0283; 2.02: 5.45; 2.04: 4.9667; 2.05: 15.337; 2.06: 5.2667; 2.07: 4.8167; 2.08: 7.6667; 2.09: 5.3167; 2.10: 6; 2.12: 5.65; 2.13: 5.3583
| Median |  | 2.01: 4.02; 2.02: 5.45; 2.04: 5; 2.05: 15.4406; 2.06: 5.25; 2.07: 4.8; 2.08: 8; 2.09: 5.3; 2.10: 6; 2.12: 5.65; 2.13: 5.365
| Std Dev. |  | 2.01: 0.16055; 2.02: 0.10488; 2.04: 0.05164; 2.05: 0.36926; 2.06: 0.1633; 2.07: 0.16021; 2.08: 0.5164; 2.09: 0.13292; 2.10: 0; 2.12: 0.18708; 2.13: 0.064936
| Rel.Std.Dev. |  | 2.01: 3.986%; 2.02: 1.924%; 2.04: 1.04%; 2.05: 2.408%; 2.06: 3.101%; 2.07: 3.326%; 2.08: 6.736%; 2.09: 2.5%; 2.10: 0%; 2.12: 3.311%; 2.13: 1.212%
| PDM3 |  | 2.01: -23.43%; 2.02: 3.598%; 2.04: -5.589%; 2.05: 191.5%; 2.06: 0.1131%; 2.07: -8.441%; 2.08: 45.73%; 2.09: 1.064%; 2.10: 14.05%; 2.12: 7.4%; 2.13: 1.856%
| Z-Score (Absolute) |  | 2.01: 2.76; 2.02: 0.35; 2.04: 0.71; 2.05: 21.97; 2.06: 0.05; 2.07: 1.04; 2.08: NA; 2.09: 0.05; 2.10: NA; 2.11: 0.64; 2.12: 0.78; 2.13: 0.15
| Analytical results for Li in OREAS 163b (Certified Value 106 ppm) | [22 empty] | Certified
Li | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.18; 2.09: 2.20; 2.10: 2.21; 2.11: 2.23; 2.12: 2.24 | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*OES; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES; 2.10: PF*OES; 2.11: PF*MS; 2.12: PF*OES/MS | 0
0
| 1 | 1 | 2.01: 108; 2.02: 178.1695; 2.04: 113; 2.05: 106; 2.06: 102.0867; 2.07: 110; 2.08: 112; 2.09: 114; 2.10: 91; 2.11: 112; 2.12: 99 | 1
| 1 | 2 | 2.01: 105; 2.02: 178.8495; 2.04: 124; 2.05: 108; 2.06: 102.5305; 2.07: 106; 2.08: 92; 2.09: 114; 2.10: 92; 2.11: 109; 2.12: 108 | 0
| 1 | 3 | 2.01: 105; 2.02: 177.115; 2.04: 106; 2.05: 112; 2.06: 107.4682; 2.07: 106; 2.08: 105; 2.09: 114; 2.10: 92; 2.11: 108; 2.12: 103 | 16
| 1 | 4 | 2.01: 109; 2.02: 177.4805; 2.04: 107; 2.05: 109; 2.06: 104.8027; 2.07: 111; 2.08: 106; 2.09: 114; 2.10: 93; 2.11: 110; 2.12: 100 | 105.6452
| 1 | 5 | 2.01: 108; 2.02: 177.3105; 2.04: 116; 2.05: 108; 2.06: 105.8562; 2.07: 104; 2.08: 88; 2.09: 114; 2.10: 91; 2.11: 99; 2.12: 102 | 173
| 1 | 6 | 2.01: 109; 2.02: 177.011; 2.04: 109; 2.05: 105; 2.06: 100.1676; 2.07: 110; 2.08: 92; 2.09: 114; 2.10: 92; 2.11: 110; 2.12: 101
| Mean |  | 2.01: 107.3333; 2.02: 177.656; 2.04: 112.5; 2.05: 108; 2.06: 103.8187; 2.07: 107.8333; 2.08: 99.1667; 2.09: 114; 2.10: 91.8333; 2.11: 108; 2.12: 102.1667
| Median |  | 2.01: 108; 2.02: 177.3955; 2.04: 111; 2.05: 108; 2.06: 103.6666; 2.07: 108; 2.08: 98.5; 2.09: 114; 2.10: 92; 2.11: 109.5; 2.12: 101.5
| Std Dev. |  | 2.01: 1.8619; 2.02: 0.71372; 2.04: 6.775; 2.05: 2.4495; 2.06: 2.6992; 2.07: 2.8577; 2.08: 9.7245; 2.09: 0; 2.10: 0.75277; 2.11: 4.6043; 2.12: 3.1885
| Rel.Std.Dev. |  | 2.01: 1.735%; 2.02: 0.4017%; 2.04: 6.022%; 2.05: 2.268%; 2.06: 2.6%; 2.07: 2.65%; 2.08: 9.806%; 2.09: 0%; 2.10: 0.8197%; 2.11: 4.263%; 2.12: 3.121%
| PDM3 |  | 2.01: 1.598%; 2.02: 68.16%; 2.04: 6.489%; 2.05: 2.229%; 2.06: -1.729%; 2.07: 2.071%; 2.08: -6.132%; 2.09: 7.908%; 2.10: -13.07%; 2.11: 2.229%; 2.12: -3.293%
| Z-Score (Absolute) |  | 2.01: 0.07; 2.02: 10.09; 2.04: 0.67; 2.05: 0.02; 2.06: 0.58; 2.07: 0; 2.08: 1.25; 2.09: 0.89; 2.10: 2.31; 2.11: 0.02; 2.12: 0.82
| Analytical results for Lu in OREAS 163b (Indicative Value 0.058 ppm) | [22 empty] | Indicative
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.01: 2.04; 2.02: 2.06; 2.04: 2.13; 2.05: 2.18; 2.06: 2.19; 2.07: 2.22; 2.08: 2.24 | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES/MS; 2.04: PF*MS; 2.05: PF*MS; 2.06: BF*OES/MS; 2.07: BF*MS; 2.08: PF*OES/MS | 3
3
| 1 | 1 | 2.01: < 0.5; 2.02: 0.1019; 2.04: 0.06; 2.05: < 2; 2.06: 0.06; 2.07: 0.06; 2.08: 0.07 | 1
| 1 | 2 | 2.01: < 0.5; 2.02: 0.096623; 2.04: 0.06; 2.05: < 2; 2.06: 0.06; 2.07: 0.05; 2.08: 0.08 | 4
| 1 | 3 | 2.01: < 0.5; 2.02: 0.10344; 2.04: 0.06; 2.05: < 2; 2.06: 0.05; 2.07: 0.06; 2.08: 0.09 | 16
| 1 | 4 | 2.01: < 0.5; 2.02: 0.10146; 2.04: 0.05; 2.05: < 2; 2.06: 0.06; 2.07: 0.07; 2.08: 0.07 | 0.058222
| 1 | 5 | 2.01: < 0.5; 2.02: 0.096974; 2.04: < 0.05; 2.05: < 2; 2.06: 0.06; 2.07: 0.06; 2.08: 0.07 | 28
| 1 | 6 | 2.01: < 0.5; 2.02: 0.098448; 2.04: 0.06; 2.05: < 2; 2.06: 0.06; 2.07: 0.05; 2.08: 0.07
| Mean |  | 2.02: 0.099807; 2.04: 0.058; 2.06: 0.058333; 2.07: 0.058333; 2.08: 0.075
| Median |  | 2.02: 0.099953; 2.04: 0.06; 2.06: 0.06; 2.07: 0.06; 2.08: 0.07
| Std Dev. |  | 2.02: 0.0028399; 2.04: 0.0044721; 2.06: 0.0040825; 2.07: 0.0075277; 2.08: 0.0083666
| Rel.Std.Dev. |  | 2.02: 2.845%; 2.04: 7.711%; 2.06: 6.999%; 2.07: 12.9%; 2.08: 11.16%
| PDM3 |  | 2.02: 71.42%; 2.04: -0.3817%; 2.06: 0.1908%; 2.07: 0.1908%; 2.08: 28.82%
| Z-Score (Absolute) |  | 2.01: 5.24; 2.02: 0.74; 2.04: 0.67; 2.05: 27.72; 2.06: 0.5; 2.07: 0.5; 2.08: 0
| Analytical results for Mg in OREAS 163b (Certified Value 3.5 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 3.39 | 3.6301 | 3.3 | 3.36 | 3.676 | 3.462 | 3.293 | 3.437 | 3.57 | 3.456 | 3.51 | 3.43 | 3.5856 | 3.54 | 3.7 | 3.588 | 3.36 | 3.704 | 3.62 | 3.29 |  | 1
| 1 | 2 | 3.46 | 3.6035 | 3.48 | 3.43 | 3.7496 | 3.462 | 3.401 | 3.407 | 3.57 | 3.419 | 3.49 | 3.44 | 3.6106 | 3.57 | 3.66 | 3.624 | 3.35 | 3.72 | 3.58 | 3.31 |  | 0
| 1 | 3 | 3.47 | 3.6157 | 3.37 | 3.44 | 3.7097 | 3.443 | 3.377 | 3.45 | 3.528 | 3.486 | 3.52 | 3.49 | 3.5292 | 3.53 | 3.56 | 3.655 | 3.37 | 3.724 | 3.54 | 3.19 |  | 16
| 1 | 4 | 3.52 | 3.628 | 3.38 | 3.4 | 3.7273 | 3.462 | 3.431 | 3.492 | 3.588 | 3.498 | 3.54 | 3.43 | 3.2093 | 3.52 | 3.74 | 3.588 | 3.36 | 3.708 | 3.61 | 3.36 |  | 3.5042
| 1 | 5 | 3.42 | 3.6174 | 3.33 | 3.42 | 3.7173 | 3.456 | 3.498 | 3.462 | 3.558 | 3.51 | 3.53 | 3.46 | 3.5068 | 3.53 | 3.67 | 3.643 | 3.35 | 3.743 | 3.17 | 3.24 |  | 174
| 1 | 6 | 3.49 | 3.6167 | 3.36 | 3.42 | 3.71 | 3.431 | 3.462 | 3.462 | 3.588 | 3.474 | 3.54 | 3.43 | 3.1482 | 3.52 | 3.63 | 3.624 | 3.35 | 3.727 | 3.53 | 3.33
| Mean |  | 3.4583 | 3.6186 | 3.37 | 3.4117 | 3.715 | 3.4527 | 3.4103 | 3.4517 | 3.567 | 3.4738 | 3.5217 | 3.4467 | 3.4316 | 3.535 | 3.66 | 3.6203 | 3.3567 | 3.721 | 3.5083 | 3.2867
| Median |  | 3.465 | 3.6171 | 3.365 | 3.42 | 3.7136 | 3.459 | 3.416 | 3.456 | 3.57 | 3.48 | 3.525 | 3.435 | 3.518 | 3.53 | 3.665 | 3.624 | 3.355 | 3.722 | 3.56 | 3.3
| Std Dev. |  | 0.047081 | 0.0096164 | 0.061319 | 0.028577 | 0.024175 | 0.012925 | 0.07176 | 0.028472 | 0.02237 | 0.032756 | 0.019408 | 0.024221 | 0.20035 | 0.018708 | 0.061644 | 0.027689 | 0.008165 | 0.014057 | 0.16964 | 0.062183
| Rel.Std.Dev. |  | 1.361% | 0.2658% | 1.82% | 0.8376% | 0.6508% | 0.3744% | 2.104% | 0.8249% | 0.6271% | 0.9429% | 0.5511% | 0.7027% | 5.838% | 0.5292% | 1.684% | 0.7648% | 0.2432% | 0.3778% | 4.835% | 1.892%
| PDM3 |  | -1.309% | 3.263% | -3.83% | -2.641% | 6.015% | -1.471% | -2.679% | -1.5% | 1.792% | -0.8671% | 0.4979% | -1.642% | -2.072% | 0.8784% | 4.446% | 3.314% | -4.211% | 6.186% | 0.1174% | -6.208%
| Z-Score (Absolute) |  | 0.08 | 1.65 | 1.04 | 0.59 | 2.69 | 0.15 | 0.6 | 0.16 | 1.09 | 0.08 | 0.6 | 0.21 | 0.37 | 0.75 | 2.1 | 1.67 | 1.18 | 2.76 | 0.46 | 1.94
| Analytical results for Mn in OREAS 163b (Certified Value 0.048 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  | 3
3
| 1 | 1 | 0.049 | 0.049416 | 0.04499 | 0.048 | 0.046 | 0.046 | 0.046 | 0.046 | 0.046 | 0.0499 | 0.0479 | 0.049086 | 0.0491 | 0.05 | 0.046 | 0.0497 | 0.0427 | 0.0504 | 0.0454 |  |  | 1
| 1 | 2 | 0.049 | 0.049705 | 0.04571 | 0.049 | 0.046 | 0.046 | 0.046 | 0.046 | 0.046 | 0.0494 | 0.0482 | 0.049201 | 0.049 | 0.05 | 0.046 | 0.0496 | 0.0427 | 0.0494 | 0.0454 |  |  | 30
| 1 | 3 | 0.049 | 0.049802 | 0.04622 | 0.049 | 0.046 | 0.046 | 0.046 | 0.046 | 0.046 | 0.0492 | 0.0487 | 0.050518 | 0.0493 | 0.05 | 0.046 | 0.0497 | 0.0428 | 0.0487 | 0.0441 |  |  | 16
| 1 | 4 | 0.047 | 0.049742 | 0.04663 | 0.048 | 0.046 | 0.046 | 0.046 | 0.046 | 0.046 | 0.0485 | 0.0477 | 0.047684 | 0.0488 | 0.05 | 0.046 | 0.0496 | 0.0428 | 0.0503 | 0.0461 |  |  | 0.047588
| 1 | 5 | 0.048 | 0.049931 | 0.04528 | 0.048 | 0.046 | 0.046 | 0.046 | 0.046 | 0.054 | 0.0463 | 0.0483 | 0.050679 | 0.0494 | 0.05 | 0.046 | 0.0496 | 0.0428 | 0.0427 | 0.0445 |  |  | 175
| 1 | 6 | 0.049 | 0.04959 | 0.04662 | 0.048 | 0.046 | 0.046 | 0.046 | 0.046 | 0.046 | 0.049 | 0.0477 | 0.048125 | 0.0491 | 0.05 | 0.046 | 0.0497 | 0.0428 | 0.0493 | 0.0462
| Mean |  | 0.0485 | 0.049698 | 0.045908 | 0.048333 | 0.046 | 0.046 | 0.046 | 0.046 | 0.047333 | 0.048717 | 0.048083 | 0.049215 | 0.049117 | 0.05 | 0.046 | 0.04965 | 0.042767 | 0.048467 | 0.045283
| Median |  | 0.049 | 0.049724 | 0.045965 | 0.048 | 0.046 | 0.046 | 0.046 | 0.046 | 0.046 | 0.0491 | 0.04805 | 0.049144 | 0.0491 | 0.05 | 0.046 | 0.04965 | 0.0428 | 0.04935 | 0.0454
| Std Dev. |  | 0.00083666 | 0.00017797 | 0.00069309 | 0.0005164 | 7.6012e-18 | 7.6012e-18 | 7.6012e-18 | 7.6012e-18 | 0.003266 | 0.0012703 | 0.000392 | 0.0012158 | 0.0002137 | 7.6012e-18 | 7.6012e-18 | 5.4772e-05 | 5.164e-05 | 0.0028974 | 0.00084242
| Rel.Std.Dev. |  | 1.725% | 0.3581% | 1.51% | 1.068% | 1.652e-14% | 1.652e-14% | 1.652e-14% | 1.652e-14% | 6.9% | 2.608% | 0.8153% | 2.47% | 0.4351% | 1.52e-14% | 1.652e-14% | 0.1103% | 0.1207% | 5.978% | 1.86%
| PDM3 |  | 1.916% | 4.432% | -3.53% | 1.565% | -3.338% | -3.338% | -3.338% | -3.338% | -0.536% | 2.371% | 1.04% | 3.419% | 3.211% | 5.068% | -3.338% | 4.332% | -10.13% | 1.846% | -4.844%
| Z-Score (Absolute) |  | 0.17 | 0.67 | 0.91 | 0.1 | 0.67 | 0.67 | 0.67 | 0.67 | 0.14 | 0.26 | 0 | 0.47 | 0.43 | 0.8 | 0.67 | 0.65 | 2.22 | 0.16 | 1.17
| Analytical results for Mo in OREAS 163b (Certified Value 2.26 ppm) | [22 empty] | Certified
Mo | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.16; 2.10: 2.21; 2.11: 2.23; 2.12: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*OES; 2.11: PF*MS; 2.12: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 2; 2.02: 2.3; 2.04: 2; 2.05: 4; 2.07: < 2; 2.08: 1.7598; 2.09: 3; 2.10: < 10; 2.11: 3; 2.12: < 10 |  |  | 1
| 1 | 2 | 2.01: 2; 2.02: 2.5; 2.04: 2; 2.05: 3; 2.07: < 2; 2.08: 2.2089; 2.09: 3; 2.10: < 10; 2.11: 2; 2.12: < 10 |  |  | 31
| 1 | 3 | 2.01: 2; 2.02: 2.3; 2.04: 2; 2.05: 4; 2.07: 2; 2.08: 1.8552; 2.09: 2; 2.10: < 10; 2.11: 3; 2.12: < 10 |  |  | 16
| 1 | 4 | 2.01: 2; 2.02: 2.6; 2.04: 2; 2.05: 4; 2.07: < 2; 2.08: 1.8062; 2.09: 2; 2.10: < 10; 2.11: 3; 2.12: < 10 |  |  | 2.2642
| 1 | 5 | 2.01: 2; 2.02: 2.6; 2.04: 2; 2.05: 3; 2.07: < 2; 2.08: 1.8174; 2.09: 3; 2.10: < 10; 2.11: 3; 2.12: < 10 |  |  | 176
| 1 | 6 | 2.01: 2; 2.02: 2.2; 2.04: 2; 2.05: 3; 2.07: < 2; 2.08: 2.1469; 2.09: 3; 2.10: < 10; 2.11: 3; 2.12: < 10
| Mean |  | 2.01: 2; 2.02: 2.4167; 2.04: 2; 2.05: 3.5; 2.07: 2; 2.08: 1.9324; 2.09: 2.6667; 2.11: 2.8333
| Median |  | 2.01: 2; 2.02: 2.4; 2.04: 2; 2.05: 3.5; 2.07: 2; 2.08: 1.8363; 2.09: 3; 2.11: 3
| Std Dev. |  | 2.01: 0; 2.02: 0.17224; 2.04: 0; 2.05: 0.54772; 2.08: 0.19358; 2.09: 0.5164; 2.11: 0.40825
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 7.127%; 2.04: 0%; 2.05: 15.65%; 2.08: 10.02%; 2.09: 19.36%; 2.11: 14.41%
| PDM3 |  | 2.01: -11.67%; 2.02: 6.736%; 2.04: -11.67%; 2.05: 54.58%; 2.07: -11.67%; 2.08: -14.65%; 2.09: 17.78%; 2.11: 25.14%
| Z-Score (Absolute) |  | 2.01: 0.63; 2.02: 0.15; 2.04: 0.63; 2.05: 1.12; 2.07: 1.61; 2.08: 0.71; 2.09: 0.15; 2.10: 2.88; 2.11: 0.34; 2.12: 2.88
| Analytical results for Na in OREAS 163b (Indicative Value 0.022 wt.%) | [22 empty] | Indicative
Na | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.19; 2.05: 2.20 |  |  | wt.%
|  |  | 2.01: PF*OES; 2.02: BF*OES; 2.04: BF*OES/MS; 2.05: BF*OES |  |  | 3
3
| 1 | 1 | 2.01: 0.05; 2.02: 0.01; 2.04: 0.015; 2.05: < 0.01 |  |  | 1
| 1 | 2 | 2.01: 0.04; 2.02: 0.01; 2.04: 0.015; 2.05: < 0.01 |  |  | 23
| 1 | 3 | 2.01: 0.04; 2.02: 0.01; 2.04: 0.015; 2.05: < 0.01 |  |  | 16
| 1 | 4 | 2.01: 0.04; 2.02: 0.01; 2.04: 0.015; 2.05: < 0.01 |  |  | 0.022168
| 1 | 5 | 2.01: 0.04; 2.02: 0.01; 2.04: 0.015; 2.05: < 0.01 |  |  | 29
| 1 | 6 | 2.01: 0.04; 2.02: 0.01; 2.04: 0.015; 2.05: < 0.01
| Mean |  | 2.01: 0.041667; 2.02: 0.01; 2.04: 0.015
| Median |  | 2.01: 0.04; 2.02: 0.01; 2.04: 0.015
| Std Dev. |  | 2.01: 0.0040825; 2.02: 0; 2.04: 0
| Rel.Std.Dev. |  | 2.01: 9.798%; 2.02: 0%; 2.04: 0%
| PDM3 |  | 2.01: 87.96%; 2.02: -54.89%; 2.04: -32.33%
| Z-Score (Absolute) |  | 2.01: 4.01; 2.02: 0.33; 2.04: 0.33; 2.05: 1.02
| Analytical results for Nb in OREAS 163b (Certified Value 1.92 ppm) | [22 empty] | Certified
Nb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.19; 2.10: 2.20; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: PF*OES; 2.11: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 1.8; 2.02: 1.7; 2.04: 2.2171; 2.05: 2.8; 2.07: < 1; 2.08: < 2; 2.09: 1.5; 2.10: < 5; 2.11: < 5 |  |  | 1
| 1 | 2 | 2.01: 1.7; 2.02: 1.8; 2.04: 2.3287; 2.05: 2.7; 2.07: 2; 2.08: < 2; 2.09: 1.9; 2.10: < 5; 2.11: < 5 |  |  | 14
| 1 | 3 | 2.01: 1.8; 2.02: 1.8; 2.04: 2.318; 2.05: < 2.4; 2.07: 1; 2.08: < 2; 2.09: 1.9; 2.10: < 5; 2.11: < 5 |  |  | 16
| 1 | 4 | 2.01: 1.8; 2.02: 1.8; 2.04: 2.194; 2.05: 3.1; 2.07: < 1; 2.08: < 2; 2.09: 1.6; 2.10: < 5; 2.11: < 5 |  |  | 1.9153
| 1 | 5 | 2.01: 1.9; 2.02: 1.6; 2.04: 2.2842; 2.05: < 2.4; 2.07: 1; 2.08: < 2; 2.09: 1.7; 2.10: < 5; 2.11: < 5 |  |  | 177
| 1 | 6 | 2.01: 2.1; 2.02: 1.8; 2.04: 2.2177; 2.05: < 2.4; 2.07: 1; 2.08: < 2; 2.09: 1.8; 2.10: < 5; 2.11: < 5
| Mean |  | 2.01: 1.85; 2.02: 1.75; 2.04: 2.2599; 2.05: 2.8667; 2.07: 1.25; 2.09: 1.7333
| Median |  | 2.01: 1.8; 2.02: 1.8; 2.04: 2.2509; 2.05: 2.8; 2.07: 1; 2.09: 1.75
| Std Dev. |  | 2.01: 0.13784; 2.02: 0.083666; 2.04: 0.057702; 2.05: 0.20817; 2.07: 0.5; 2.09: 0.1633
| Rel.Std.Dev. |  | 2.01: 7.451%; 2.02: 4.781%; 2.04: 2.553%; 2.05: 7.262%; 2.07: 40%; 2.09: 9.421%
| PDM3 |  | 2.01: -3.41%; 2.02: -8.632%; 2.04: 17.99%; 2.05: 49.67%; 2.07: -34.74%; 2.09: -9.502%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.16; 2.04: 0.67; 2.05: 0.3; 2.07: 1.4; 2.08: 1.4; 2.09: 0.19; 2.10: 1.07; 2.11: 1.07
| Analytical results for Nd in OREAS 163b (Certified Value 5.5 ppm) | [22 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: 6.05; 2.02: 5.1; 2.04: 5.3893; 2.05: 6.9; 2.07: 5.1; 2.08: 6; 2.09: 5.4; 2.10: 5.5; 2.11: 5.4 |  |  | 1
| 1 | 2 | 2.01: 5.85; 2.02: 5.3; 2.04: 5.3464; 2.05: 6.6; 2.07: 5.1; 2.08: 6; 2.09: 5.4; 2.10: 6.1; 2.11: 5.7 |  |  | 5
| 1 | 3 | 2.01: 5.88; 2.02: 5.4; 2.04: 5.2688; 2.05: 6.1; 2.07: 5.5; 2.08: 6; 2.09: 5.4; 2.10: 5.3; 2.11: 5.4 |  |  | 16
| 1 | 4 | 2.01: 5.77; 2.02: 5.1; 2.04: 5.5402; 2.05: 6.6; 2.07: 4.9; 2.08: 6; 2.09: 5.2; 2.10: 5.6; 2.11: 5.6 |  |  | 5.4963
| 1 | 5 | 2.01: 5.79; 2.02: 5.4; 2.04: 5.5428; 2.05: 6.7; 2.07: 5.1; 2.08: 6; 2.09: 5.5; 2.10: 5.7; 2.11: 5.4 |  |  | 178
| 1 | 6 | 2.01: 5.77; 2.02: 5.8; 2.04: 5.5469; 2.05: 5.3; 2.07: 5.3; 2.08: 6; 2.09: 5.5; 2.10: 6.3; 2.11: 5.6
| Mean |  | 2.01: 5.8517; 2.02: 5.35; 2.04: 5.4391; 2.05: 6.3667; 2.07: 5.1667; 2.08: 6; 2.09: 5.4; 2.10: 5.75; 2.11: 5.5167
| Median |  | 2.01: 5.82; 2.02: 5.35; 2.04: 5.4648; 2.05: 6.6; 2.07: 5.1; 2.08: 6; 2.09: 5.4; 2.10: 5.65; 2.11: 5.5
| Std Dev. |  | 2.01: 0.10704; 2.02: 0.25884; 2.04: 0.12056; 2.05: 0.58538; 2.07: 0.20656; 2.08: 0; 2.09: 0.10954; 2.10: 0.37815; 2.11: 0.13292
| Rel.Std.Dev. |  | 2.01: 1.829%; 2.02: 4.838%; 2.04: 2.217%; 2.05: 9.194%; 2.07: 3.998%; 2.08: 0%; 2.09: 2.029%; 2.10: 6.577%; 2.11: 2.409%
| PDM3 |  | 2.01: 6.466%; 2.02: -2.662%; 2.04: -1.041%; 2.05: 15.84%; 2.07: -5.997%; 2.08: 9.164%; 2.09: -1.752%; 2.10: 4.616%; 2.11: 0.3707%
| Z-Score (Absolute) |  | 2.01: 1.26; 2.02: 0.43; 2.04: 0.13; 2.05: 3; 2.07: 1.05; 2.08: NA; 2.09: 0.26; 2.10: 0.92; 2.11: 0.13
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 163b (Certified Value 29.4 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 30 | 20.89 | 30 | 40 | 180 | 20 | 30 | 80 | 35 | 31.6886 | 28 | < 50 | 23 | 32 | 19 | 35 | 63 |  |  |  |  | 1
| 1 | 2 | 30 | 21.97 | 30 | 40 | 70 | 40 | 30 | 30 | 33 | 32.2701 | 27 | < 50 | 25 | 33 | 19 | 26 | 54 |  |  |  |  | 16
| 1 | 3 | 30 | 23.21 | 20 | 30 | 70 | 30 | 40 | 60 | 32 | 24.6403 | 25 | < 50 | 29 | 33 | 18 | 32 | < 50 |  |  |  |  | 16
| 1 | 4 | 30 | 22.44 | 20 | 30 | 60 | 30 | 30 | 90 | 30 | 27.0142 | 26 | < 50 | < 20 | 32 | 18 | 38 | 55 |  |  |  |  | 29.3903
| 1 | 5 | 30 | 22.28 | 20 | 30 | 80 | 20 | 40 | 60 | 32 | 33.3385 | 29 | < 50 | 21 | 33 | 19 | 25 | 55 |  |  |  |  | 179
| 1 | 6 | 30 | 26.33 | 30 | 40 | 70 | 40 | 140 | 60 | 33 | 25.001 | 21 | < 50 | 23 | 33 | 18 | 37 | 58
| Mean |  | 30 | 22.8533 | 25 | 35 | 88.3333 | 30 | 51.6667 | 63.3333 | 32.5 | 28.9921 | 26 |  | 24.2 | 32.6667 | 18.5 | 32.1667 | 57
| Median |  | 30 | 22.36 | 25 | 35 | 70 | 30 | 35 | 60 | 32.5 | 29.3514 | 26.5 |  | 23 | 33 | 18.5 | 33.5 | 55
| Std Dev. |  | 0 | 1.8627 | 5.4772 | 5.4772 | 45.3505 | 8.9443 | 43.5507 | 20.6559 | 1.6432 | 3.8907 | 2.8284 |  | 3.0332 | 0.5164 | 0.54772 | 5.5648 | 3.6742
| Rel.Std.Dev. |  | 0% | 8.151% | 21.91% | 15.65% | 51.34% | 29.81% | 84.29% | 32.61% | 5.056% | 13.42% | 10.88% |  | 12.53% | 1.581% | 2.961% | 17.3% | 6.446%
| PDM3 |  | 2.075% | -22.24% | -14.94% | 19.09% | 200.6% | 2.075% | 75.8% | 115.5% | 10.58% | -1.355% | -11.54% |  | -17.66% | 11.15% | -37.05% | 9.447% | 93.94%
| Z-Score (Absolute) |  | 0 | 0.96 | 0.67 | 0.67 | 7.87 | 0 | 2.92 | 4.5 | 0.34 | 0.14 | 0.54 | 0.67 | 1.1 | 0.36 | 1.55 | 0.29 | 2.92
| Analytical results for P in OREAS 163b (Certified Value 0.02 wt.%) | [22 empty] | Certified
P | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.07: 2.13; 2.08: 2.19; 2.09: 2.20; 2.10: 2.21 |  |  | wt.%
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*OES; 2.05: PF*OES/MS; 2.07: PF*OES; 2.08: BF*OES/MS; 2.09: PF*OES; 2.10: PF*OES |  |  | 3
3
| 1 | 1 | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.019002; 2.07: 0.02; 2.08: 0.017; 2.09: 0.0402; 2.10: 0.012 |  |  | 1
| 1 | 2 | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.019933; 2.07: 0.02; 2.08: 0.017; 2.09: 0.0404; 2.10: 0.022 |  |  | 0
| 1 | 3 | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.018781; 2.07: 0.02; 2.08: 0.017; 2.09: 0.0406; 2.10: 0.012 |  |  | 16
| 1 | 4 | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.018998; 2.07: 0.02; 2.08: 0.017; 2.09: 0.0406; 2.10: 0.023 |  |  | 0.019507
| 1 | 5 | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.018884; 2.07: 0.02; 2.08: 0.017; 2.09: 0.0403; 2.10: < 0.01 |  |  | 180
| 1 | 6 | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.018641; 2.07: 0.02; 2.08: 0.022; 2.09: 0.0405; 2.10: < 0.01
| Mean |  | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.01904; 2.07: 0.02; 2.08: 0.017833; 2.09: 0.040433; 2.10: 0.01725
| Median |  | 2.01: 0.02; 2.02: 0.02; 2.04: 0.02; 2.05: 0.018941; 2.07: 0.02; 2.08: 0.017; 2.09: 0.04045; 2.10: 0.017
| Std Dev. |  | 2.01: 0; 2.02: 0; 2.04: 0; 2.05: 0.00045836; 2.07: 0; 2.08: 0.0020412; 2.09: 0.0001633; 2.10: 0.0060759
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 0%; 2.04: 0%; 2.05: 2.407%; 2.07: 0%; 2.08: 11.45%; 2.09: 0.4039%; 2.10: 35.22%
| PDM3 |  | 2.01: 2.526%; 2.02: 2.526%; 2.04: 2.526%; 2.05: -2.397%; 2.07: 2.526%; 2.08: -8.581%; 2.09: 107.3%; 2.10: -11.57%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0; 2.04: 0; 2.05: 1.35; 2.07: 0; 2.08: 2.55; 2.09: 28.7; 2.10: 9.6
| Analytical results for Pb in OREAS 163b (Certified Value 56 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.21 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [6 empty] | 0
0
| 1 | 1 | 50 | 68.2 | < 100 | 100 | < 100 | < 100 | < 100 | 59.3 | 55 | 58.0997 | 62 | < 100 | 58 | 50 | < 100 | [6 empty] | 1
| 1 | 2 | 50 | 66.5 | < 100 | 100 | 100 | < 100 | < 100 | 58.6 | 54 | 55.8064 | 63 | < 100 | 58 | 49 | < 100 | [6 empty] | 17
| 1 | 3 | 50 | 69 | < 100 | 100 | < 100 | 100 | < 100 | 58.4 | 56 | 50.585 | 55 | < 100 | 58 | 44 | < 100 | [6 empty] | 16
| 1 | 4 | 50 | 68.3 | < 100 | 100 | < 100 | < 100 | 100 | 59.5 | 53 | 48.9248 | 59 | < 100 | 59 | 43 | < 100 | [6 empty] | 55.8834
| 1 | 5 | 50 | 68.7 | 100 | 100 | < 100 | < 100 | < 100 | 60.1 | 53 | 48.3709 | 60 | < 100 | 58 | 46 | < 100 | [6 empty] | 181
| 1 | 6 | 50 | 69.5 | < 100 | 100 | 100 | < 100 | 100 | 56.8 | 55 | 53.7171 | 54 | < 100 | 59 | 43 | < 100
| Mean |  | 50 | 68.3667 | 100 | 100 | 100 | 100 | 100 | 58.7833 | 54.3333 | 52.584 | 58.8333 |  | 58.3333 | 45.8333
| Median |  | 50 | 68.5 | 100 | 100 | 100 | 100 | 100 | 58.95 | 54.5 | 52.151 | 59.5 |  | 58 | 45
| Std Dev. |  | 2.01: 0; 2.02: 1.0309; 2.05: 0; 2.07: 0; 2.09: 0; 2.10: 1.1514; 2.11: 1.2111; 2.12: 3.93; 2.13: 3.656; 2.16: 0.5164; 2.17: 3.0605
| Rel.Std.Dev. |  | 2.01: 0%; 2.02: 1.508%; 2.05: 0%; 2.07: 0%; 2.09: 0%; 2.10: 1.959%; 2.11: 2.229%; 2.12: 7.474%; 2.13: 6.214%; 2.16: 0.8853%; 2.17: 6.677%
| PDM3 |  | -10.53% | 22.34% | 78.94% | 78.94% | 78.94% | 78.94% | 78.94% | 5.189% | -2.774% | -5.904% | 5.279% |  | 4.384% | -17.98%
| Z-Score (Absolute) |  | 0.67 | 0.81 | 0 | 3.37 | 0.67 | 0 | 0.67 | 0.04 | 0.32 | 0.47 | 0.04 | 0.67 | 0 | 1.01 | 0.67
| Analytical results for Pr in OREAS 163b (Certified Value 1.37 ppm) | [22 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: 1.4; 2.02: 1.3; 2.04: 1.8743; 2.05: 1.3; 2.07: 1.4; 2.08: < 2; 2.09: 1.31; 2.10: 1.38; 2.11: 1.41 |  |  | 1
| 1 | 2 | 2.01: 1.4; 2.02: 1.4; 2.04: 1.8348; 2.05: 1.3; 2.07: 1.34; 2.08: < 2; 2.09: 1.3; 2.10: 1.57; 2.11: 1.42 |  |  | 18
| 1 | 3 | 2.01: 1.5; 2.02: 1.4; 2.04: 1.8697; 2.05: 1.4; 2.07: 1.42; 2.08: < 2; 2.09: 1.32; 2.10: 1.37; 2.11: 1.36 |  |  | 16
| 1 | 4 | 2.01: 1.4; 2.02: 1.3; 2.04: 1.8697; 2.05: 1.4; 2.07: 1.3; 2.08: < 2; 2.09: 1.3; 2.10: 1.95; 2.11: 1.37 |  |  | 1.3689
| 1 | 5 | 2.01: 1.4; 2.02: 1.3; 2.04: 1.8152; 2.05: 1.4; 2.07: 1.31; 2.08: < 2; 2.09: 1.35; 2.10: 1.39; 2.11: 1.36 |  |  | 182
| 1 | 6 | 2.01: 1.4; 2.02: 1.5; 2.04: 1.7956; 2.05: 1.4; 2.07: 1.39; 2.08: < 2; 2.09: 1.35; 2.10: 1.78; 2.11: 1.37
| Mean |  | 2.01: 1.4167; 2.02: 1.3667; 2.04: 1.8432; 2.05: 1.3667; 2.07: 1.36; 2.09: 1.3217; 2.10: 1.5733; 2.11: 1.3817
| Median |  | 2.01: 1.4; 2.02: 1.35; 2.04: 1.8522; 2.05: 1.4; 2.07: 1.365; 2.09: 1.315; 2.10: 1.48; 2.11: 1.37
| Std Dev. |  | 2.01: 0.040825; 2.02: 0.08165; 2.04: 0.033132; 2.05: 0.05164; 2.07: 0.0502; 2.09: 0.023166; 2.10: 0.24369; 2.11: 0.026394
| Rel.Std.Dev. |  | 2.01: 2.882%; 2.02: 5.974%; 2.04: 1.798%; 2.05: 3.779%; 2.07: 3.691%; 2.09: 1.753%; 2.10: 15.49%; 2.11: 1.91%
| PDM3 |  | 2.01: 3.49%; 2.02: -0.1623%; 2.04: 34.65%; 2.05: -0.1623%; 2.07: -0.6494%; 2.09: -3.45%; 2.10: 14.94%; 2.11: 0.9334%
| Z-Score (Absolute) |  | 2.01: 0.75; 2.02: 0; 2.04: 7.14; 2.05: 0; 2.07: 0.1; 2.08: 5.49; 2.09: 0.67; 2.10: 3.1; 2.11: 0.22
| Analytical results for Rb in OREAS 163b (Certified Value 4.89 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.22 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 5 | 4.4 | 4.8735 | 5.9 | 4.5 | 5.4407 | 5.4 | 20 | 4.9 | 5 | 4.5 | < 10 | [9 empty] | 1
| 1 | 2 | 4.8 | 5.2 | 4.8842 | 4.7 | 4.3 | 6.4051 | 5.4 | 22 | 5 | 5.3 | 4.5 | < 10 | [9 empty] | 37
| 1 | 3 | 5 | 4.1 | 4.8055 | 5.5 | 4.9 | 6.5659 | 5.1 | 11 | 5.2 | 4.8 | 5 | < 10 | [9 empty] | 16
| 1 | 4 | 4.6 | 4.5 | 4.7996 | 6.7 | 4.8 | 7.0699 | 5 | 13 | 5 | 4.8 | 4.9 | < 10 | [9 empty] | 4.8892
| 1 | 5 | 4.8 | 4.6 | 4.8855 | 6.9 | 4.5 | 5.2204 | 5.3 | 38 | 4.9 | 4.9 | 3.3 | < 10 | [9 empty] | 183
| 1 | 6 | 4.7 | 4.8 | 4.9588 | 5.4 | 4.6 | 4.8482 | 4.9 | 28 | 4.9 | 4.7 | 4.4 | < 10
| Mean |  | 2.01: 4.8167; 2.02: 4.6; 2.04: 4.8678; 2.05: 5.85; 2.07: 4.6; 2.08: 5.925; 2.09: 5.1833; 2.10: 22; 2.11: 4.9833; 2.12: 4.9167; 2.13: 4.4333
| Median |  | 2.01: 4.8; 2.02: 4.55; 2.04: 4.8788; 2.05: 5.7; 2.07: 4.55; 2.08: 5.9229; 2.09: 5.2; 2.10: 21; 2.11: 4.95; 2.12: 4.85; 2.13: 4.5
| Std Dev. |  | 2.01: 0.16021; 2.02: 0.37417; 2.04: 0.059048; 2.05: 0.83367; 2.07: 0.21909; 2.08: 0.87664; 2.09: 0.2137; 2.10: 9.98; 2.11: 0.1169; 2.12: 0.2137; 2.13: 0.60553
| Rel.Std.Dev. |  | 2.01: 3.326%; 2.02: 8.134%; 2.04: 1.213%; 2.05: 14.25%; 2.07: 4.763%; 2.08: 14.8%; 2.09: 4.123%; 2.10: 45.36%; 2.11: 2.346%; 2.12: 4.346%; 2.13: 13.66%
| PDM3 |  | 2.01: -1.484%; 2.02: -5.915%; 2.04: -0.4372%; 2.05: 19.65%; 2.07: -5.915%; 2.08: 21.19%; 2.09: 6.016%; 2.10: 350%; 2.11: 1.925%; 2.12: 0.5617%; 2.13: -9.324%
| Z-Score (Absolute) |  | 0.25 | 0.8 | 0.12 | 2.36 | 0.8 | 2.55 | 0.67 | NA | 0.17 | 0 | 1.22 | 0.21
Results from laboratory 2.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Re in OREAS 163b (Indicative Value < 0.1 ppm) | [22 empty] | Indicative
Re | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.14; 2.04: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 1
| 1 | 2 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 38
| 1 | 3 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 16
| 1 | 4 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | < 0.1
| 1 | 5 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01 |  |  | 27
| 1 | 6 | 2.01: < 0.1; 2.02: < 0.1; 2.04: < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.04
| Analytical results for S in OREAS 163b (Certified Value 3.36 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.16 | 2.20 | 2.23 | 2.24 | [8 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 3.48 | 2.6 | 3.41 | 3.23 | 3.38 | 3.26 | 3.33 | 3.21 | 3.732 | 3.35 | 3.38 | 3.39 | 3.3 | [8 empty] | 1
| 1 | 2 | 3.53 | 2.78 | 3.39 | 3.32 | 3.32 | 3.3 | 3.24 | 3.18 | 3.6808 | 3.36 | 3.41 | 3.4 | 3.36 | [8 empty] | 0
| 1 | 3 | 3.55 | 2.77 | 3.41 | 3.29 | 3.35 | 3.24 | 3.36 | 3.21 | 3.6153 | 3.36 | 3.43 | 3.31 | 3.2 | [8 empty] | 16
| 1 | 4 | 3.53 | 2.8 | 3.4 | 3.32 | 3.34 | 3.29 | 3.4 | 3.23 | 3.2687 | 3.34 | 3.39 | 3.31 | 3.41 | [8 empty] | 3.3566
| 1 | 5 | 3.52 | 2.7 | 3.43 | 3.31 | 3.38 | 3.28 | 3.44 | 3.2 | 3.2442 | 3.36 | 3.55 | 3 | 3.25 | [8 empty] | 184
| 1 | 6 | 3.58 | 2.62 | 3.37 | 3.31 | 3.37 | 3.27 | 3.37 | 3.24 | 3.3749 | 3.34 | 3.32 | 3.27 | 3.36
| Mean |  | 3.5317 | 2.7117 | 3.4017 | 3.2967 | 3.3567 | 3.2733 | 3.3567 | 3.2117 | 3.486 | 3.3517 | 3.4133 | 3.28 | 3.3133
| Median |  | 3.53 | 2.735 | 3.405 | 3.31 | 3.36 | 3.275 | 3.365 | 3.21 | 3.4951 | 3.355 | 3.4 | 3.31 | 3.33
| Std Dev. |  | 0.033116 | 0.085888 | 0.020412 | 0.034448 | 0.024221 | 0.021602 | 0.068313 | 0.02137 | 0.21597 | 0.0098319 | 0.076594 | 0.14615 | 0.078401
| Rel.Std.Dev. |  | 0.9377% | 3.167% | 0.6001% | 1.045% | 0.7216% | 0.66% | 2.035% | 0.6654% | 6.195% | 0.2933% | 2.244% | 4.456% | 2.366%
| PDM3 |  | 5.215% | -19.21% | 1.342% | -1.786% | 0.00142% | -2.481% | 0.00142% | -4.318% | 3.853% | -0.1475% | 1.69% | -2.283% | -1.29%
| Z-Score (Absolute) |  | 1.97 | 7 | 0.55 | 0.6 | 0.05 | 0.86 | 0.05 | 1.53 | 1.47 | 0 | 0.67 | 0.78 | 0.42
| Analytical results for Sb in OREAS 163b (Certified Value 8.3 ppm) | [22 empty] | Certified
Sb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.12; 2.05: 2.13; 2.07: 2.14; 2.08: 2.16; 2.09: 2.21; 2.10: 2.23; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES; 2.10: PF*MS; 2.11: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 5; 2.02: 9.2; 2.04: 10; 2.05: 6.2; 2.07: 7.6173; 2.08: 10.8; 2.09: < 50; 2.10: 8.7; 2.11: 11 |  |  | 1
| 1 | 2 | 2.01: < 5; 2.02: 9.8; 2.04: 11; 2.05: 8.1; 2.07: 7.7444; 2.08: 9.4; 2.09: < 50; 2.10: 8.9; 2.11: 11 |  |  | 39
| 1 | 3 | 2.01: < 5; 2.02: 9.2; 2.04: 10; 2.05: 7.3; 2.07: 6.8257; 2.08: 9.8; 2.09: < 50; 2.10: 9.2; 2.11: 11 |  |  | 16
| 1 | 4 | 2.01: < 5; 2.02: 9.4; 2.04: 11; 2.05: 6.3; 2.07: 7.5471; 2.08: 9.3; 2.09: < 50; 2.10: 8.1; 2.11: 10 |  |  | 8.2988
| 1 | 5 | 2.01: < 5; 2.02: 9.2; 2.04: 10; 2.05: 7.1; 2.07: 7.3979; 2.08: 8.9; 2.09: < 50; 2.10: 7.7; 2.11: 11 |  |  | 185
| 1 | 6 | 2.01: < 5; 2.02: 9.3; 2.04: 10; 2.05: 6.8; 2.07: 6.6514; 2.08: 9.5; 2.09: < 50; 2.10: 8.4; 2.11: 11
| Mean |  | 2.02: 9.35; 2.04: 10.3333; 2.05: 6.9667; 2.07: 7.2973; 2.08: 9.6167; 2.10: 8.5; 2.11: 10.8333
| Median |  | 2.02: 9.25; 2.04: 10; 2.05: 6.95; 2.07: 7.4725; 2.08: 9.45; 2.10: 8.55; 2.11: 11
| Std Dev. |  | 2.02: 0.23452; 2.04: 0.5164; 2.05: 0.70333; 2.07: 0.45042; 2.08: 0.64936; 2.10: 0.54772; 2.11: 0.40825
| Rel.Std.Dev. |  | 2.02: 2.508%; 2.04: 4.997%; 2.05: 10.1%; 2.07: 6.172%; 2.08: 6.752%; 2.10: 6.444%; 2.11: 3.768%
| PDM3 |  | 2.02: 12.67%; 2.04: 24.52%; 2.05: -16.05%; 2.07: -12.07%; 2.08: 15.88%; 2.10: 2.425%; 2.11: 30.54%
| Z-Score (Absolute) |  | 2.01: 3.36; 2.02: 0.48; 2.04: NA; 2.05: 0.86; 2.07: 0.67; 2.08: 0.63; 2.09: 9.25; 2.10: 0; 2.11: NA
Results from laboratories 2.12 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 163b (Certified Value < 5 ppm) | [22 empty] | Certified
Sc | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.13; 2.05: 2.14; 2.07: 2.18; 2.08: 2.19; 2.09: 2.20; 2.10: 2.21 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*OES; 2.07: PF*MS; 2.08: BF*OES/MS; 2.09: PF*OES; 2.10: PF*OES |  |  | 2
2
| 1 | 1 | 2.01: < 20; 2.02: < 1; 2.04: < 5; 2.05: < 20; 2.07: < 2; 2.08: 1; 2.09: < 5; 2.10: < 5 |  |  | 1
| 1 | 2 | 2.01: < 20; 2.02: < 1; 2.04: < 5; 2.05: < 20; 2.07: < 2; 2.08: 1; 2.09: < 5; 2.10: < 5 |  |  | 19
| 1 | 3 | 2.01: < 20; 2.02: < 1; 2.04: < 5; 2.05: < 20; 2.07: < 2; 2.08: 1; 2.09: < 5; 2.10: < 5 |  |  | 16
| 1 | 4 | 2.01: < 20; 2.02: < 1; 2.04: < 5; 2.05: < 20; 2.07: < 2; 2.08: 1; 2.09: < 5; 2.10: < 5 |  |  | < 5
| 1 | 5 | 2.01: < 20; 2.02: < 1; 2.04: < 5; 2.05: < 20; 2.07: < 2; 2.08: 1; 2.09: < 5; 2.10: < 5 |  |  | 186
| 1 | 6 | 2.01: < 20; 2.02: < 1; 2.04: < 5; 2.05: < 20; 2.07: < 2; 2.08: 1; 2.09: < 5; 2.10: < 5
Mean | bar: 2.08
Median | bar: 2.08
| Std Dev. |  | 2.08: 0
Rel.Std.Dev. | bar: 2.08
PDM3
| Z-Score (Absolute) |  | 2.01: 3.37; 2.02: 0.9; 2.04: 0; 2.05: 3.37; 2.07: 0.67; 2.08: 0.67; 2.09: 0; 2.10: 0
| Analytical results for Se in OREAS 163b (Indicative Value < 10 ppm) | [22 empty] | Indicative
Se | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab |  |  | 1
| No | No | 2.01: 2.12; 2.02: 2.14; 2.04: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*OES/MS |  |  | 1
1
| 1 | 1 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 1
| 1 | 2 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 41
| 1 | 3 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 16
| 1 | 4 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | < 10
| 1 | 5 | 2.01: < 8; 2.02: < 20; 2.04: < 10 |  |  | 28
| 1 | 6 | 2.01: 12; 2.02: < 20; 2.04: < 10
| Mean |  | 2.01: 12
| Median |  | 2.01: 12
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 9.44; 2.04: 0.67
| Analytical results for Si in OREAS 163b (Certified Value 34.76 wt.%) | [22 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 31.635 | 33.71 | 35.71 | 32.91 | 36.7 | 34.59 | 34.78 | > 30 | 29.3 | 26.5791 | 34.7 | 29 | 34.938 | 34.97 | 38.549 | 36.7 | 34.3 |  |  |  |  | 1
| 1 | 2 | 31.8833 | 34.46 | 35.62 | 34.5 | 36.18 | 34.59 | 34.5 | > 30 | > 30 | 26.7269 | 35.4 | 31.2 | 34.817 | 34.98 | 38.876 | 36.4 | 34 |  |  |  |  | 0
| 1 | 3 | 31.4944 | 34.48 | 35.48 | 34.08 | 36.6 | 33.98 | 35.29 | > 30 | > 30 | 26.3154 | 35 | 28.7 | 34.747 | 34.95 | 37.491 | 36 | 33.1 |  |  |  |  | 16
| 1 | 4 | 31.618 | 33.92 | 35.81 | 34.78 | 37.02 | 34.69 | 35.29 | > 30 | > 30 | 27.4319 | 34.9 | 28.1 | 35.041 | 34.96 | 37.992 | 36.5 | 34.4 |  |  |  |  | 34.7556
| 1 | 5 | 31.2332 | 34.38 | 35.71 | 35.2 | 36.79 | 34.41 | 35.39 | > 30 | > 30 | 27.5463 | 34.9 | 31.5 | 34.826 | 34.97 | 37.616 | 32.2 | 33.5 |  |  |  |  | 187
| 1 | 6 | 31.5754 | 34.35 | 35.48 | 35.01 | 36.88 | 34.59 | 35.01 | > 30 | > 30 | 26.8092 | 34.8 | 27.6 | 34.817 | 34.96 | 38.979 | 35.9 | 34.3
| Mean |  | 31.5732 | 34.2167 | 35.635 | 34.4133 | 36.695 | 34.475 | 35.0433 |  | 29.3 | 26.9014 | 34.95 | 29.35 | 34.8643 | 34.965 | 38.2505 | 35.6167 | 33.9333
| Median |  | 31.5967 | 34.365 | 35.665 | 34.64 | 36.745 | 34.59 | 35.15 |  | 29.3 | 26.768 | 34.9 | 28.85 | 34.8215 | 34.965 | 38.2705 | 36.2 | 34.15
| Std Dev. |  | 0.21157 | 0.32179 | 0.13428 | 0.83522 | 0.29091 | 0.2589 | 0.34858 |  |  | 0.48653 | 0.2429 | 1.6257 | 0.10617 | 0.010488 | 0.64145 | 1.7011 | 0.52409
| Rel.Std.Dev. |  | 0.6701% | 0.9404% | 0.3768% | 2.427% | 0.7928% | 0.751% | 0.9947% |  |  | 1.809% | 0.695% | 5.539% | 0.3045% | 0.03% | 1.677% | 4.776% | 1.544%
| PDM3 |  | -9.157% | -1.551% | 2.53% | -0.9849% | 5.58% | -0.8075% | 0.8277% |  | -15.7% | -22.6% | 0.5592% | -15.55% | 0.3127% | 0.6024% | 10.06% | 2.477% | -2.366%
| Z-Score (Absolute) |  | 2.48 | 0.36 | 0.77 | 0.21 | 1.62 | 0.16 | 0.3 | NA | 4.3 | 6.22 | 0.22 | 4.26 | 0.16 | 0.24 | 2.87 | 0.76 | 0.59
| Analytical results for Sm in OREAS 163b (Certified Value 1.48 ppm) | [22 empty] | Certified
Sm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: 1.4; 2.02: 1.3; 2.04: 1.5141; 2.05: 1.6; 2.07: 1.6; 2.08: < 2; 2.09: 1.31; 2.10: 1.45; 2.11: 1.49 |  |  | 1
| 1 | 2 | 2.01: 1.4; 2.02: 1.4; 2.04: 1.5546; 2.05: 1; 2.07: 1.5; 2.08: < 2; 2.09: 1.44; 2.10: 1.33; 2.11: 1.53 |  |  | 5
| 1 | 3 | 2.01: 1.5; 2.02: 1.5; 2.04: 1.5677; 2.05: 1.7; 2.07: 1.4; 2.08: < 2; 2.09: 1.36; 2.10: 1.53; 2.11: 1.51 |  |  | 16
| 1 | 4 | 2.01: 1.5; 2.02: 1.7; 2.04: 1.5368; 2.05: 1.9; 2.07: 1.5; 2.08: < 2; 2.09: 1.37; 2.10: 1.53; 2.11: 1.51 |  |  | 1.4821
| 1 | 5 | 2.01: 1.4; 2.02: 1.3; 2.04: 1.5301; 2.05: 1.7; 2.07: 1.4; 2.08: < 2; 2.09: 1.36; 2.10: 1.53; 2.11: 1.42 |  |  | 188
| 1 | 6 | 2.01: 1.4; 2.02: 1.2; 2.04: 1.5286; 2.05: 1.5; 2.07: 1.6; 2.08: < 2; 2.09: 1.37; 2.10: 1.5; 2.11: 1.44
| Mean |  | 2.01: 1.4333; 2.02: 1.4; 2.04: 1.5387; 2.05: 1.5667; 2.07: 1.5; 2.09: 1.3683; 2.10: 1.4783; 2.11: 1.4833
| Median |  | 2.01: 1.4; 2.02: 1.35; 2.04: 1.5335; 2.05: 1.65; 2.07: 1.5; 2.09: 1.365; 2.10: 1.515; 2.11: 1.5
| Std Dev. |  | 2.01: 0.05164; 2.02: 0.17889; 2.04: 0.019366; 2.05: 0.30768; 2.07: 0.089443; 2.09: 0.041673; 2.10: 0.079099; 2.11: 0.043665
| Rel.Std.Dev. |  | 2.01: 3.603%; 2.02: 12.78%; 2.04: 1.259%; 2.05: 19.64%; 2.07: 5.963%; 2.09: 3.046%; 2.10: 5.351%; 2.11: 2.944%
| PDM3 |  | 2.01: -3.289%; 2.02: -5.538%; 2.04: 3.819%; 2.05: 5.707%; 2.07: 1.209%; 2.09: -7.675%; 2.10: -0.2531%; 2.11: 0.08424%
| Z-Score (Absolute) |  | 2.01: 0.5; 2.02: 0.88; 2.04: 0.67; 2.05: 0.99; 2.07: 0.24; 2.08: 5.35; 2.09: 1.23; 2.10: 0; 2.11: 0.06
| Analytical results for Sn in OREAS 163b (Certified Value 3.04 ppm) | [22 empty] | Certified
Sn | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.17; 2.10: 2.18; 2.11: 2.19; 2.12: 2.21; 2.13: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*OES; 2.13: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 5; 2.02: 3; 2.04: 3.4354; 2.05: 4; 2.07: 2; 2.08: < 100; 2.09: < 50; 2.10: 6; 2.11: 3; 2.12: < 50; 2.13: 7 |  |  | 1
| 1 | 2 | 2.01: < 5; 2.02: 3; 2.04: 3.1497; 2.05: 3.8; 2.07: 3; 2.08: < 100; 2.09: < 50; 2.10: 7; 2.11: 3; 2.12: < 50; 2.13: 7 |  |  | 26
| 1 | 3 | 2.01: < 5; 2.02: 3; 2.04: 3.078; 2.05: 3; 2.07: 3; 2.08: < 100; 2.09: < 50; 2.10: < 5; 2.11: 3; 2.12: < 50; 2.13: < 5 |  |  | 16
| 1 | 4 | 2.01: < 5; 2.02: 3; 2.04: 3.1197; 2.05: 3.2; 2.07: 2; 2.08: < 100; 2.09: 160; 2.10: 5; 2.11: 3; 2.12: < 50; 2.13: 6 |  |  | 3.0359
| 1 | 5 | 2.01: < 5; 2.02: 3; 2.04: 3.2825; 2.05: 3.6; 2.07: 3; 2.08: < 100; 2.09: < 50; 2.10: 6; 2.11: 3; 2.12: < 50; 2.13: 7 |  |  | 189
| 1 | 6 | 2.01: < 5; 2.02: 3; 2.04: 3.3131; 2.05: 2.1; 2.07: 3; 2.08: < 100; 2.09: < 50; 2.10: 6; 2.11: 3; 2.12: < 50; 2.13: 8
| Mean |  | 2.02: 3; 2.04: 3.2297; 2.05: 3.2833; 2.07: 2.6667; 2.09: 160; 2.10: 6; 2.11: 3; 2.13: 7
| Median |  | 2.02: 3; 2.04: 3.2161; 2.05: 3.4; 2.07: 3; 2.09: 160; 2.10: 6; 2.11: 3; 2.13: 7
| Std Dev. |  | 2.02: 0; 2.04: 0.13679; 2.05: 0.68823; 2.07: 0.5164; 2.10: 0.70711; 2.11: 0; 2.13: 0.70711
| Rel.Std.Dev. |  | 2.02: 0%; 2.04: 4.235%; 2.05: 20.96%; 2.07: 19.36%; 2.10: 11.79%; 2.11: 0%; 2.13: 10.1%
| PDM3 |  | 2.02: -1.184%; 2.04: 6.382%; 2.05: 8.149%; 2.07: -12.16%; 2.09: 5170%; 2.10: 97.63%; 2.11: -1.184%; 2.13: 130.6%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.24; 2.04: 0.05; 2.05: 0; 2.07: 0.53; 2.08: 40.21; 2.09: 38.06; 2.10: 1.84; 2.11: 0.24; 2.12: 18.69; 2.13: 2.55
| Analytical results for Sr in OREAS 163b (Certified Value 9.69 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.17 | 2.19 | 2.20 | 2.21 | 2.24 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | < 20 | 18.2 | 9.3 | 9.2 | 8.9936 | 23 | < 10 | < 20 | < 50 | 9.5 | 15 | 12 | 21 | [8 empty] | 1
| 1 | 2 | < 20 | 10.46 | 10.2 | 9.2 | 8.8047 | 22 | 11 | < 20 | < 50 | 9.6 | 16 | 11 | 21 | [8 empty] | 27
| 1 | 3 | < 20 | 11.3 | 9.5 | 9.1 | 8.5458 | 22 | 10 | < 20 | < 50 | 9.7 | 16 | 11 | < 20 | [8 empty] | 16
| 1 | 4 | < 20 | 11.61 | 9.4 | 9.3 | 8.5774 | 18 | < 10 | < 20 | < 50 | 9.6 | 16 | 12 | 21 | [8 empty] | 9.6895
| 1 | 5 | < 20 | 11.8 | 9.1 | 9.3 | 8.5806 | 21 | < 10 | < 20 | < 50 | 9.9 | 15 | 11 | 20 | [8 empty] | 190
| 1 | 6 | < 20 | 11.6 | 9.3 | 9.7 | 8.458 | 22 | < 10 | < 20 | < 50 | 9.7 | 16 | 11 | 21
| Mean |  | 2.02: 12.495; 2.04: 9.4667; 2.05: 9.3; 2.07: 8.66; 2.08: 21.3333; 2.09: 10.5; 2.12: 9.6667; 2.13: 15.6667; 2.14: 11.3333; 2.16: 20.8
| Median |  | 2.02: 11.605; 2.04: 9.35; 2.05: 9.25; 2.07: 8.579; 2.08: 22; 2.09: 10.5; 2.12: 9.65; 2.13: 16; 2.14: 11; 2.16: 21
| Std Dev. |  | 2.02: 2.8349; 2.04: 0.38297; 2.05: 0.20976; 2.07: 0.19958; 2.08: 1.7512; 2.09: 0.70711; 2.12: 0.13663; 2.13: 0.5164; 2.14: 0.5164; 2.16: 0.44721
| Rel.Std.Dev. |  | 2.02: 22.69%; 2.04: 4.045%; 2.05: 2.256%; 2.07: 2.305%; 2.08: 8.209%; 2.09: 6.734%; 2.12: 1.413%; 2.13: 3.296%; 2.14: 4.556%; 2.16: 2.15%
| PDM3 |  | 2.02: 28.95%; 2.04: -2.299%; 2.05: -4.019%; 2.07: -10.62%; 2.08: 120.2%; 2.09: 8.365%; 2.12: -0.2353%; 2.13: 61.69%; 2.14: 16.97%; 2.16: 114.7%
| Z-Score (Absolute) |  | 0.25 | 3.99 | 0.55 | 0.8 | 1.76 | NA | NA | 0.25 | 22.73 | 0.25 | NA | NA | NA
Results from laboratories 2.12, 2.13, 2.20, 2.21 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 163b (Certified Value 2.3 ppm) | [22 empty] | Certified
Ta | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.14; 2.09: 2.16; 2.10: 2.18; 2.11: 2.19; 2.12: 2.23; 2.13: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*MS; 2.13: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: 2.1; 2.02: 2; 2.04: 2.0077; 2.05: 5.8; 2.07: 1.7; 2.08: 4.6634; 2.09: 2.5; 2.10: 2.7; 2.11: 2; 2.12: 3.9; 2.13: < 3 |  |  | 1
| 1 | 2 | 2.01: 1.8; 2.02: 2.5; 2.04: 2.0221; 2.05: 6.8; 2.07: 4; 2.08: 4.799; 2.09: 2.3; 2.10: 2.8; 2.11: 3; 2.12: 2.4; 2.13: < 3 |  |  | 10
| 1 | 3 | 2.01: 1.9; 2.02: 2.6; 2.04: 2.032; 2.05: 5.9; 2.07: 1.7; 2.08: 5.6535; 2.09: 2.6; 2.10: 2.3; 2.11: 3.4; 2.12: 2.8; 2.13: 4 |  |  | 16
| 1 | 4 | 2.01: 2; 2.02: 2; 2.04: 2.0262; 2.05: 6.1; 2.07: 1.7; 2.08: 5.6689; 2.09: 2.6; 2.10: 1.5; 2.11: 2.1; 2.12: 2.2; 2.13: < 3 |  |  | 2.2975
| 1 | 5 | 2.01: 2.3; 2.02: 2.2; 2.04: 2.0941; 2.05: 6.3; 2.07: 1.9; 2.08: 4.9787; 2.09: 2.3; 2.10: 1.8; 2.11: 2.1; 2.12: 2.5; 2.13: 3 |  |  | 191
| 1 | 6 | 2.01: 2.4; 2.02: 2.1; 2.04: 2.0809; 2.05: 3.9; 2.07: 1.6; 2.08: 4.9773; 2.09: 2.4; 2.10: 1.8; 2.11: 2.5; 2.12: 2.5; 2.13: < 3
| Mean |  | 2.01: 2.0833; 2.02: 2.2333; 2.04: 2.0438; 2.05: 5.8; 2.07: 2.1; 2.08: 5.1235; 2.09: 2.45; 2.10: 2.15; 2.11: 2.5167; 2.12: 2.7167; 2.13: 3.5
| Median |  | 2.01: 2.05; 2.02: 2.15; 2.04: 2.0291; 2.05: 6; 2.07: 1.7; 2.08: 4.978; 2.09: 2.45; 2.10: 2.05; 2.11: 2.3; 2.12: 2.5; 2.13: 3.5
| Std Dev. |  | 2.01: 0.23166; 2.02: 0.2582; 2.04: 0.035024; 2.05: 0.99599; 2.07: 0.93595; 2.08: 0.43303; 2.09: 0.13784; 2.10: 0.53198; 2.11: 0.57067; 2.12: 0.61128; 2.13: 0.70711
| Rel.Std.Dev. |  | 2.01: 11.12%; 2.02: 11.56%; 2.04: 1.714%; 2.05: 17.17%; 2.07: 44.57%; 2.08: 8.452%; 2.09: 5.626%; 2.10: 24.74%; 2.11: 22.68%; 2.12: 22.5%; 2.13: 20.2%
| PDM3 |  | 2.01: -9.32%; 2.02: -2.791%; 2.04: -11.04%; 2.05: 152.5%; 2.07: -8.595%; 2.08: 123%; 2.09: 6.639%; 2.10: -6.419%; 2.11: 9.541%; 2.12: 18.25%; 2.13: 52.34%
| Z-Score (Absolute) |  | 2.01: 0.53; 2.02: 0; 2.04: 0.67; 2.05: 12.69; 2.07: 0.47; 2.08: 10.29; 2.09: 0.77; 2.10: 0.3; 2.11: 1.01; 2.12: 1.72; 2.13: 0.24
| Analytical results for Tb in OREAS 163b (Certified Value 0.24 ppm) | [22 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab |  |  | 1
| No | No | 2.01: 2.04; 2.02: 2.05; 2.04: 2.06; 2.05: 2.12; 2.07: 2.13; 2.08: 2.18; 2.09: 2.19; 2.10: 2.22; 2.11: 2.24 |  |  | ppm
|  |  | 2.01: PF*MS; 2.02: BF*MS; 2.04: PF*OES/MS; 2.05: PF*OES/MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: BF*OES/MS; 2.10: BF*MS; 2.11: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.25144; 2.05: 0.3; 2.07: 0.23; 2.08: < 2; 2.09: 0.23; 2.10: 0.24; 2.11: 0.33 |  |  | 1
| 1 | 2 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.24181; 2.05: 0.3; 2.07: 0.22; 2.08: < 2; 2.09: 0.24; 2.10: 0.26; 2.11: 0.34 |  |  | 11
| 1 | 3 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.23927; 2.05: 0.2; 2.07: 0.22; 2.08: < 2; 2.09: 0.23; 2.10: 0.24; 2.11: 0.35 |  |  | 16
| 1 | 4 | 2.01: < 0.5; 2.02: 0.3; 2.04: 0.25775; 2.05: 0.3; 2.07: 0.22; 2.08: < 2; 2.09: 0.24; 2.10: 0.25; 2.11: 0.33 |  |  | 0.24288
| 1 | 5 | 2.01: < 0.5; 2.02: 0.2; 2.04: 0.234; 2.05: 0.3; 2.07: 0.21; 2.08: < 2; 2.09: 0.24; 2.10: 0.25; 2.11: 0.32 |  |  | 192
| 1 | 6 | 2.01: < 0.5; 2.02: 0.3; 2.04: 0.24935; 2.05: 0.2; 2.07: 0.23; 2.08: < 2; 2.09: 0.24; 2.10: 0.28; 2.11: 0.34
| Mean |  | 2.02: 0.23333; 2.04: 0.2456; 2.05: 0.26667; 2.07: 0.22167; 2.09: 0.23667; 2.10: 0.25333; 2.11: 0.335
| Median |  | 2.02: 0.2; 2.04: 0.24558; 2.05: 0.3; 2.07: 0.22; 2.09: 0.24; 2.10: 0.25; 2.11: 0.335
| Std Dev. |  | 2.02: 0.05164; 2.04: 0.0087717; 2.05: 0.05164; 2.07: 0.0075277; 2.09: 0.005164; 2.10: 0.015055; 2.11: 0.010488
| Rel.Std.Dev. |  | 2.02: 22.13%; 2.04: 3.572%; 2.05: 19.36%; 2.07: 3.396%; 2.09: 2.182%; 2.10: 5.943%; 2.11: 3.131%
| PDM3 |  | 2.02: -3.93%; 2.04: 1.122%; 2.05: 9.794%; 2.07: -8.733%; 2.09: -2.557%; 2.10: 4.305%; 2.11: 37.93%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.04: 0.18; 2.05: 0.67; 2.07: 1.15; 2.08: 30.34; 2.09: 0.54; 2.10: 0.13; 2.11: 3.44
| Analytical results for Te in OREAS 163b (Indicative Value < 1 ppm) | [22 empty] | Indicative
Te | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab |  |  | 1
| No | No | 2.01: 2.12; 2.02: 2.13; 2.04: 2.14; 2.05: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*MS; 2.04: PF*MS; 2.05: PF*OES/MS |  |  | 2
2
| 1 | 1 | 2.01: < 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 1
| 1 | 2 | 2.01: < 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 12
| 1 | 3 | 2.01: < 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 16
| 1 | 4 | 2.01: < 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | < 1
| 1 | 5 | 2.01: < 6; 2.02: < 1; 2.04: < 1; 2.05: < 5 |  |  | 29
| 1 | 6 | 2.01: 6; 2.02: < 1; 2.04: < 1; 2.05: < 5
| Mean |  | 2.01: 6
| Median |  | 2.01: 6
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 2.01: 1.35; 2.02: 0.67; 2.04: 0.67; 2.05: 0.67
| Analytical results for Th in OREAS 163b (Certified Value 0.91 ppm) | [22 empty] | Certified
Th | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.05; 2.05: 2.06; 2.07: 2.12; 2.08: 2.13; 2.09: 2.14; 2.10: 2.16; 2.11: 2.19; 2.12: 2.23; 2.13: 2.24 |  |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.07: PF*OES/MS; 2.08: PF*MS; 2.09: PF*MS; 2.10: PF*MS; 2.11: BF*OES/MS; 2.12: PF*MS; 2.13: PF*OES/MS |  |  | 2
3
| 1 | 1 | 2.01: 1; 2.02: < 5; 2.04: 0.7; 2.05: 1.3021; 2.07: 0.8; 2.08: 0.9; 2.09: 0.87734; 2.10: 1; 2.11: 0.9; 2.12: 0.9; 2.13: 1 |  |  | 1
| 1 | 2 | 2.01: 1; 2.02: < 5; 2.04: 0.9; 2.05: 1.3706; 2.07: 1.2; 2.08: 0.9; 2.09: 0.97814; 2.10: 0.9; 2.11: 0.9; 2.12: 0.9; 2.13: 1 |  |  | 13
| 1 | 3 | 2.01: 1; 2.02: < 5; 2.04: 0.9; 2.05: 1.3158; 2.07: 0.9; 2.08: 1; 2.09: 0.90792; 2.10: 1; 2.11: 0.8; 2.12: 1; 2.13: 1 |  |  | 16
| 1 | 4 | 2.01: 1; 2.02: < 5; 2.04: 0.8; 2.05: 1.3224; 2.07: 1; 2.08: 0.9; 2.09: 0.8659; 2.10: 1; 2.11: 0.8; 2.12: 1; 2.13: 1 |  |  | 0.90608
| 1 | 5 | 2.01: 1; 2.02: < 5; 2.04: 0.8; 2.05: 1.3295; 2.07: 1.1; 2.08: 0.9; 2.09: 0.83936; 2.10: 0.9; 2.11: 0.8; 2.12: 0.9; 2.13: < 1 |  |  | 193
| 1 | 6 | 2.01: 1; 2.02: < 5; 2.04: 0.8; 2.05: 1.3642; 2.07: 1.1; 2.08: 0.9; 2.09: 1.0067; 2.10: 0.9; 2.11: 0.8; 2.12: 0.9; 2.13: 1
| Mean |  | 2.01: 1; 2.04: 0.81667; 2.05: 1.3341; 2.07: 1.0167; 2.08: 0.91667; 2.09: 0.91255; 2.10: 0.95; 2.11: 0.83333; 2.12: 0.93333; 2.13: 1
| Median |  | 2.01: 1; 2.04: 0.8; 2.05: 1.326; 2.07: 1.05; 2.08: 0.9; 2.09: 0.89263; 2.10: 0.95; 2.11: 0.8; 2.12: 0.9; 2.13: 1
| Std Dev. |  | 2.01: 0; 2.04: 0.075277; 2.05: 0.027392; 2.07: 0.1472; 2.08: 0.040825; 2.09: 0.066264; 2.10: 0.054772; 2.11: 0.05164; 2.12: 0.05164; 2.13: 0
| Rel.Std.Dev. |  | 2.01: 0%; 2.04: 9.218%; 2.05: 2.053%; 2.07: 14.48%; 2.08: 4.454%; 2.09: 7.261%; 2.10: 5.766%; 2.11: 6.197%; 2.12: 5.533%; 2.13: 0%
| PDM3 |  | 2.01: 10.37%; 2.04: -9.868%; 2.05: 47.24%; 2.07: 12.21%; 2.08: 1.169%; 2.09: 0.7146%; 2.10: 4.847%; 2.11: -8.029%; 2.12: 3.008%; 2.13: 10.37%
| Z-Score (Absolute) |  | 2.01: NA; 2.02: NA; 2.04: 1.25; 2.05: 4.73; 2.07: 1.06; 2.08: 0.1; 2.09: 0.14; 2.10: 0.29; 2.11: 1.06; 2.12: 0.1; 2.13: NA
Results from laboratory 2.01, 2.04 and 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 163b (Certified Value 0.03 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.03 | 0.028597 | 0.03 | 0.029 | 0.033883 | 0.036 | 0.03 | 0.012 | 0.036 | 0.036 | 0.03 | 0.03 | 0.028929 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026 | 0.03 | 0.03 |  | 1
| 1 | 2 | 0.03 | 0.028778 | 0.03 | 0.029 | 0.034709 | 0.03 | 0.03 | 0.042 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026423 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033 | 0.03 | 0.03 |  | 32
| 1 | 3 | 0.03 | 0.028647 | 0.03 | 0.029 | 0.034596 | 0.03 | 0.03 | 0.012 | 0.03 | 0.03 | 0.03 | 0.03 | 0.025585 | 0.03 | 0.03 | 0.03 | 0.03 | 0.025 | 0.03 | 0.029 |  | 16
| 1 | 4 | 0.03 | 0.028317 | 0.03 | 0.031 | 0.035662 | 0.03 | 0.042 | 0.03 | 0.036 | 0.03 | 0.03 | 0.03 | 0.029177 | 0.03 | 0.03 | 0.03 | 0.03 | 0.025 | 0.03 | 0.03 |  | 0.029636
| 1 | 5 | 0.03 | 0.028793 | 0.03 | 0.029 | 0.035992 | 0.03 | 0.042 | 0.012 | 0.03 | 0.126 | 0.03 | 0.03 | 0.02522 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026 | 0.03 | 0.029 |  | 194
| 1 | 6 | 0.03 | 0.028158 | 0.03 | 0.029 | 0.03428 | 0.03 | 0.03 | 0.042 | 0.036 | 0.132 | 0.03 | 0.03 | 0.024697 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026 | 0.03 | 0.03
| Mean |  | 0.03 | 0.028548 | 0.03 | 0.029333 | 0.034854 | 0.031 | 0.034 | 0.025 | 0.033 | 0.064 | 0.03 | 0.03 | 0.026672 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026833 | 0.03 | 0.029667
| Median |  | 0.03 | 0.028622 | 0.03 | 0.029 | 0.034652 | 0.03 | 0.03 | 0.021 | 0.033 | 0.033 | 0.03 | 0.03 | 0.026004 | 0.03 | 0.03 | 0.03 | 0.03 | 0.026 | 0.03 | 0.03
| Std Dev. |  | 0 | 0.00025711 | 0 | 0.0008165 | 0.00081324 | 0.0024495 | 0.0061968 | 0.0149 | 0.0032863 | 0.050438 | 0 | 0 | 0.00193 | 0 | 0 | 0 | 0 | 0.0030605 | 0 | 0.0005164
| Rel.Std.Dev. |  | 0% | 0.9006% | 0% | 2.784% | 2.333% | 7.902% | 18.23% | 59.6% | 9.959% | 78.81% | 0% | 0% | 7.236% | 0% | 0% | 0% | 0% | 11.41% | 0% | 1.741%
| PDM3 |  | 1.229% | -3.67% | 1.229% | -1.021% | 17.61% | 4.603% | 14.73% | -15.64% | 11.35% | 116% | 1.229% | 1.229% | -10% | 1.229% | 1.229% | 1.229% | 1.229% | -9.457% | 1.229% | 0.104%
| Z-Score (Absolute) |  | 0 | 1.96 | 0 | 0.9 | 6.55 | 1.31 | 5.36 | 6.77 | 4.01 | 45.78 | 0 | 0 | 4.49 | 0 | 0 | 0.03 | 0 | 4.27 | 0 | 0.45
| Analytical results for Tl in OREAS 163b (Certified Value 1.67 ppm) | [22 empty] | Certified
Tl | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.04; 2.04: 2.12; 2.05: 2.13; 2.06: 2.14; 2.07: 2.16; 2.08: 2.23; 2.09: 2.24 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: PF*MS; 2.08: PF*MS; 2.09: PF*OES/MS |  | 2
3
| 1 | 1 | 2.01: 1.7; 2.02: < 5; 2.04: 1.8; 2.05: 1.7; 2.06: 1.6074; 2.07: 1.8; 2.08: 1.6; 2.09: 1.75 |  | 1
| 1 | 2 | 2.01: 1.7; 2.02: < 5; 2.04: 1.5; 2.05: 1.7; 2.06: 1.5455; 2.07: 1.8; 2.08: 1.7; 2.09: 1.77 |  | 14
| 1 | 3 | 2.01: 1.6; 2.02: < 5; 2.04: 1.6; 2.05: 1.7; 2.06: 1.6353; 2.07: 1.7; 2.08: 1.8; 2.09: 1.72 |  | 16
| 1 | 4 | 2.01: 1.6; 2.02: < 5; 2.04: 1.8; 2.05: 1.6; 2.06: 1.5478; 2.07: 1.8; 2.08: 1.7; 2.09: 1.82 |  | 1.6747
| 1 | 5 | 2.01: 1.6; 2.02: < 5; 2.04: 1.8; 2.05: 1.6; 2.06: 1.6532; 2.07: 1.7; 2.08: 1.4; 2.09: 1.69 |  | 195
| 1 | 6 | 2.01: 1.6; 2.02: < 5; 2.04: 1.6; 2.05: 1.7; 2.06: 1.5587; 2.07: 1.8; 2.08: 1.6; 2.09: 1.74
| Mean |  | 2.01: 1.6333; 2.04: 1.6833; 2.05: 1.6667; 2.06: 1.5913; 2.07: 1.7667; 2.08: 1.6333; 2.09: 1.7483
| Median |  | 2.01: 1.6; 2.04: 1.7; 2.05: 1.7; 2.06: 1.583; 2.07: 1.8; 2.08: 1.65; 2.09: 1.745
| Std Dev. |  | 2.01: 0.05164; 2.04: 0.13292; 2.05: 0.05164; 2.06: 0.047079; 2.07: 0.05164; 2.08: 0.13663; 2.09: 0.04446
| Rel.Std.Dev. |  | 2.01: 3.162%; 2.04: 7.896%; 2.05: 3.098%; 2.06: 2.958%; 2.07: 2.923%; 2.08: 8.365%; 2.09: 2.543%
| PDM3 |  | 2.01: -2.471%; 2.04: 0.5148%; 2.05: -0.4804%; 2.06: -4.979%; 2.07: 5.491%; 2.08: -2.471%; 2.09: 4.396%
| Z-Score (Absolute) |  | 2.01: 0.49; 2.02: 9.67; 2.04: 0.1; 2.05: 0.1; 2.06: 0.98; 2.07: 1.07; 2.08: 0.49; 2.09: 0.86
| Analytical results for Tm in OREAS 163b (Certified Value 0.074 ppm) | [22 empty] | Certified
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.04; 2.02: 2.06; 2.04: 2.12; 2.05: 2.13; 2.06: 2.18; 2.07: 2.19; 2.08: 2.22; 2.09: 2.24 |  | ppm
|  |  | 2.01: PF*MS; 2.02: PF*OES/MS; 2.04: PF*OES/MS; 2.05: PF*MS; 2.06: PF*MS; 2.07: BF*OES/MS; 2.08: BF*MS; 2.09: PF*OES/MS |  | 3
3
| 1 | 1 | 2.01: < 1; 2.02: 0.094098; 2.04: < 0.1; 2.05: 0.07; 2.06: < 2; 2.07: 0.06; 2.08: 0.08; 2.09: 0.07 |  | 1
| 1 | 2 | 2.01: < 1; 2.02: 0.094654; 2.04: 0.1; 2.05: 0.07; 2.06: < 2; 2.07: 0.07; 2.08: 0.09; 2.09: 0.08 |  | 6
| 1 | 3 | 2.01: < 1; 2.02: 0.10374; 2.04: < 0.1; 2.05: 0.07; 2.06: < 2; 2.07: 0.07; 2.08: 0.08; 2.09: 0.09 |  | 16
| 1 | 4 | 2.01: < 1; 2.02: 0.10166; 2.04: < 0.1; 2.05: 0.07; 2.06: < 2; 2.07: 0.08; 2.08: 0.09; 2.09: 0.07 |  | 0.074333
| 1 | 5 | 2.01: < 1; 2.02: 0.10179; 2.04: 0.1; 2.05: 0.07; 2.06: < 2; 2.07: 0.08; 2.08: 0.09; 2.09: 0.08 |  | 196
| 1 | 6 | 2.01: < 1; 2.02: 0.096073; 2.04: < 0.1; 2.05: 0.07; 2.06: < 2; 2.07: 0.07; 2.08: 0.09; 2.09: 0.07
| Mean |  | 2.02: 0.098668; 2.04: 0.1; 2.05: 0.07; 2.07: 0.071667; 2.08: 0.086667; 2.09: 0.076667
| Median |  | 2.02: 0.098865; 2.04: 0.1; 2.05: 0.07; 2.07: 0.07; 2.08: 0.09; 2.09: 0.075
| Std Dev. |  | 2.02: 0.0041978; 2.04: 0; 2.05: 0; 2.07: 0.0075277; 2.08: 0.005164; 2.09: 0.008165
| Rel.Std.Dev. |  | 2.02: 4.254%; 2.04: 0%; 2.05: 0%; 2.07: 10.5%; 2.08: 5.958%; 2.09: 10.65%
| PDM3 |  | 2.02: 32.74%; 2.04: 34.53%; 2.05: -5.83%; 2.07: -3.587%; 2.08: 16.59%; 2.09: 3.139%
| Z-Score (Absolute) |  | 2.01: 21.16; 2.02: 0.86; 2.04: 0.76; 2.05: 0.59; 2.06: 46.44; 2.07: 0.51; 2.08: 0.25; 2.09: 0.25
| Analytical results for U in OREAS 163b (Certified Value 0.73 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.19 | 2.22 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 0.8 | 0.7 | 0.7 | 2.3938 | 0.8 | 0.72 | 0.69367 | 0.6 | 0.7 | 0.75 | 0.8 | 1 | [9 empty] | 1
| 1 | 2 | 0.8 | 0.7 | 0.7 | 2.3213 | 0.9 | 0.72 | 0.71502 | 0.6 | 0.7 | 0.75 | 0.8 | 1 | [9 empty] | 16
| 1 | 3 | 0.7 | 0.7 | 0.7 | 2.4018 | 0.9 | 0.72 | 0.7963 | 0.7 | 0.7 | 0.74 | 0.8 | 1 | [9 empty] | 16
| 1 | 4 | 0.7 | 0.7 | 0.7 | 2.3013 | 0.8 | 0.7 | 0.73609 | 0.6 | 0.7 | 0.78 | 0.7 | 1 | [9 empty] | 0.72855
| 1 | 5 | 0.7 | 0.7 | 0.8 | 2.5332 | 0.8 | 0.68 | 0.72136 | 0.6 | 0.7 | 0.85 | 0.7 | 1 | [9 empty] | 197
| 1 | 6 | 0.8 | 0.7 | 0.7 | 2.2959 | 0.9 | 0.68 | 0.78249 | 0.6 | 0.7 | 0.77 | 0.7 | 1
| Mean |  | 0.75 | 0.7 | 0.71667 | 2.3745 | 0.85 | 0.70333 | 0.74082 | 0.61667 | 0.7 | 0.77333 | 0.75 | 1
| Median |  | 0.75 | 0.7 | 0.7 | 2.3576 | 0.85 | 0.71 | 0.72873 | 0.6 | 0.7 | 0.76 | 0.75 | 1
| Std Dev. |  | 0.054772 | 1.2162e-16 | 0.040825 | 0.090182 | 0.054772 | 0.019664 | 0.040258 | 0.040825 | 1.2162e-16 | 0.040332 | 0.054772 | 0
| Rel.Std.Dev. |  | 7.303% | 1.737e-14% | 5.696% | 3.798% | 6.444% | 2.796% | 5.434% | 6.62% | 1.737e-14% | 5.215% | 7.303% | 0%
| PDM3 |  | 2.944% | -3.919% | -1.631% | 225.9% | 16.67% | -3.461% | 1.684% | -15.36% | -3.919% | 6.147% | 2.944% | 37.26%
| Z-Score (Absolute) |  | 0.17 | 0.73 | 0.43 | 29.39 | 1.96 | 0.67 | 0 | 2.23 | 0.73 | 0.58 | 0.17 | NA
Results from laboratory 2.24 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 163b (Certified Value 11.7 ppm) | [22 empty] | Certified
V | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.04: 2.04; 2.05: 2.12; 2.06: 2.13; 2.07: 2.14; 2.08: 2.17; 2.09: 2.19; 2.10: 2.20; 2.11: 2.21; 2.12: 2.24 |  | ppm
|  |  | 2.01: PF*OES/MS; 2.02: PF*OES; 2.04: PF*OES; 2.05: PF*OES/MS; 2.06: PF*OES; 2.07: PF*OES; 2.08: PF*OES; 2.09: BF*OES/MS; 2.10: BF*OES; 2.11: PF*OES; 2.12: PF*OES/MS |  | 1
1
| 1 | 1 | 2.01: < 20; 2.02: 12.5585; 2.04: 7.9; 2.05: 16; 2.06: 12; 2.07: < 20; 2.08: < 50; 2.09: 22; 2.10: 10; 2.11: < 10; 2.12: 13 |  | 1
| 1 | 2 | 2.01: < 20; 2.02: 14.6295; 2.04: 9.2; 2.05: 14; 2.06: 11; 2.07: < 20; 2.08: < 50; 2.09: 18; 2.10: 11; 2.11: < 10; 2.12: 12 |  | 36
| 1 | 3 | 2.01: < 20; 2.02: 11.8555; 2.04: 9; 2.05: 15; 2.06: 12; 2.07: < 20; 2.08: < 50; 2.09: 22; 2.10: 11; 2.11: < 10; 2.12: 13 |  | 16
| 1 | 4 | 2.01: < 20; 2.02: 11.01; 2.04: 8.4; 2.05: 16; 2.06: 11; 2.07: < 20; 2.08: < 50; 2.09: 26; 2.10: 10; 2.11: < 10; 2.12: 12 |  | 11.7208
| 1 | 5 | 2.01: < 20; 2.02: 12.0455; 2.04: 8.8; 2.05: 15; 2.06: 11; 2.07: < 20; 2.08: < 50; 2.09: 27; 2.10: 10; 2.11: < 10; 2.12: 11 |  | 198
| 1 | 6 | 2.01: < 20; 2.02: 10.0505; 2.04: 8.5; 2.05: 17; 2.06: 11; 2.07: < 20; 2.08: < 50; 2.09: 18; 2.10: 11; 2.11: < 10; 2.12: 13
| Mean |  | 2.02: 12.0249; 2.04: 8.6333; 2.05: 15.5; 2.06: 11.3333; 2.09: 22.1667; 2.10: 10.5; 2.12: 12.3333
| Median |  | 2.02: 11.9505; 2.04: 8.65; 2.05: 15.5; 2.06: 11; 2.09: 22; 2.10: 10.5; 2.12: 12.5
| Std Dev. |  | 2.02: 1.5509; 2.04: 0.46762; 2.05: 1.0488; 2.06: 0.5164; 2.09: 3.8166; 2.10: 0.54772; 2.12: 0.8165
| Rel.Std.Dev. |  | 2.02: 12.9%; 2.04: 5.416%; 2.05: 6.767%; 2.06: 4.556%; 2.09: 17.22%; 2.10: 5.216%; 2.12: 6.62%
| PDM3 |  | 2.02: 2.595%; 2.04: -26.34%; 2.05: 32.24%; 2.06: -3.306%; 2.09: 89.12%; 2.10: -10.42%; 2.12: 5.226%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.35; 2.04: 1.37; 2.05: 2.11; 2.06: 0; 2.07: 0.67; 2.08: 6.91; 2.09: 5.48; 2.10: 0.42; 2.11: 3.2; 2.12: 0.51
| Analytical results for W in OREAS 163b (Certified Value 2.55 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.19 | 2.21 | 2.23 | 2.24 | [9 empty] | ppm
|  |  | PF*OES | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 5 | 2 | < 5 | 2 | 2 | 3.1286 | 4 | < 100 | 2.2 | < 50 | 3 | < 5 | [9 empty] | 1
| 1 | 2 | < 5 | 2 | < 5 | 6.2 | 2 | 2.3723 | 3 | < 100 | 2.3 | < 50 | 3 | < 5 | [9 empty] | 17
| 1 | 3 | < 5 | 6 | < 5 | 2.1 | 2 | 3.2575 | 4 | < 100 | 2.2 | < 50 | 2 | 13 | [9 empty] | 16
| 1 | 4 | < 5 | 2 | < 5 | 3.2 | 2 | 2.4073 | 4 | < 100 | 2.4 | < 50 | 2 | < 5 | [9 empty] | 2.5489
| 1 | 5 | < 5 | 2 | < 5 | 2.3 | 2 | 3.0003 | 3 | < 100 | 2.3 | < 50 | 3 | < 5 | [9 empty] | 199
| 1 | 6 | < 5 | 2 | < 5 | 3.1 | 2 | 2.9482 | 3 | < 100 | 2.3 | < 50 | 3 | < 5
| Mean |  | 2.02: 2.6667; 2.05: 3.15; 2.06: 2; 2.07: 2.8524; 2.08: 3.5; 2.10: 2.2833; 2.12: 2.6667; 2.13: 13
| Median |  | 2.02: 2; 2.05: 2.7; 2.06: 2; 2.07: 2.9743; 2.08: 3.5; 2.10: 2.3; 2.12: 3; 2.13: 13
| Std Dev. |  | 2.02: 1.633; 2.05: 1.5783; 2.06: 0; 2.07: 0.3742; 2.08: 0.54772; 2.10: 0.075277; 2.12: 0.5164
| Rel.Std.Dev. |  | 2.02: 61.24%; 2.05: 50.1%; 2.06: 0%; 2.07: 13.12%; 2.08: 15.65%; 2.10: 3.297%; 2.12: 19.36%
| PDM3 |  | 2.02: 4.62%; 2.05: 23.58%; 2.06: -21.54%; 2.07: 11.91%; 2.08: 37.31%; 2.10: -10.42%; 2.12: 4.62%; 2.13: 410%
| Z-Score (Absolute) |  | 0.4 | 0.14 | 0.4 | 0.61 | 1.18 | 0.14 | 1.15 | 73.51 | 0.74 | 34.61 | 0.14 | 2.32
| Analytical results for Y in OREAS 163b (Certified Value 7 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.14 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*OES/MS | PF*OES | PF*OES | BF*MS | PF*MS | PF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 6.5 | 5.2 | 7.2 | 9.0593 | 7.1 | 7.3 | 7.1066 | 7 | 8 | 6.7 | 7 | < 5 | 6.8 | 7.3 | 7 | [6 empty] | 1
| 1 | 2 | 6.2 | 5 | 6.7 | 8.7228 | 8.2 | 7 | 7.3395 | 6.9 | 8 | 6.9 | 7 | < 5 | 6.7 | 7.2 | 8 | [6 empty] | 18
| 1 | 3 | 6.6 | 5.1 | 6.6 | 9.2292 | 6.8 | 7.3 | 7.0937 | 7 | 7 | 7.1 | 7 | < 5 | 6.8 | 7.2 | 7 | [6 empty] | 16
| 1 | 4 | 6.9 | 4.9 | 7.2 | 9.098 | 6.4 | 6.9 | 6.4983 | 6.9 | 8 | 7.2 | 7 | < 5 | 7 | 6.9 | 7 | [6 empty] | 7.0013
| 1 | 5 | 6.3 | 5 | 6.8 | 8.6808 | 7.5 | 6.8 | 6.9539 | 6.9 | 8 | 6.9 | 7 | < 5 | 6.9 | 6.6 | 7 | [6 empty] | 200
| 1 | 6 | 6.6 | 5 | 7 | 8.6766 | 7.8 | 7.1 | 6.4716 | 7 | 7 | 6.9 | 7 | < 5 | 7.3 | 7.3 | 7
| Mean |  | 6.5167 | 5.0333 | 6.9167 | 8.9111 | 7.3 | 7.0667 | 6.9106 | 6.95 | 7.6667 | 6.95 | 7 |  | 6.9167 | 7.0833 | 7.1667
| Median |  | 6.55 | 5 | 6.9 | 8.891 | 7.3 | 7.05 | 7.0238 | 6.95 | 8 | 6.9 | 7 |  | 6.85 | 7.2 | 7
| Std Dev. |  | 0.24833 | 0.10328 | 0.25626 | 0.24557 | 0.66332 | 0.20656 | 0.35226 | 0.054772 | 0.5164 | 0.17607 | 0 |  | 0.2137 | 0.27869 | 0.40825
| Rel.Std.Dev. |  | 3.811% | 2.052% | 3.705% | 2.756% | 9.087% | 2.923% | 5.097% | 0.7881% | 6.736% | 2.533% | 0% |  | 3.09% | 3.934% | 5.696%
| PDM3 |  | -6.922% | -28.11% | -1.209% | 27.28% | 4.266% | 0.9333% | -1.296% | -0.7331% | 9.503% | -0.7331% | -0.01893% |  | -1.209% | 1.171% | 2.362%
| Z-Score (Absolute) |  | 1.98 | 9.04 | 0.08 | 9.41 | 1.75 | 0.63 | 0.11 | 0.08 | NA | 0.08 | NA | 21.1 | 0.08 | 0.71 | NA
| Analytical results for Yb in OREAS 163b (Certified Value 0.42 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.12 | 2.13 | 2.18 | 2.19 | 2.22 | 2.24 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 0.5 | 0.38 | 0.4717 | 0.4 | 0.4 | < 2 | 0.41 | < 0.5 | 0.42 | [12 empty] | 1
| 1 | 2 | < 0.5 | 0.48 | 0.47487 | 0.8 | 0.4 | < 2 | 0.44 | < 0.5 | 0.44 | [12 empty] | 19
| 1 | 3 | < 0.5 | 0.33 | 0.48548 | 0.5 | 0.5 | < 2 | 0.41 | < 0.5 | 0.42 | [12 empty] | 16
| 1 | 4 | < 0.5 | 0.41 | 0.48791 | 0.4 | 0.4 | < 2 | 0.45 | < 0.5 | 0.45 | [12 empty] | 0.42482
| 1 | 5 | < 0.5 | 0.36 | 0.49647 | 0.4 | 0.4 | < 2 | 0.45 | < 0.5 | 0.44 | [12 empty] | 201
| 1 | 6 | < 0.5 | 0.35 | 0.47714 | 0.3 | 0.4 | < 2 | 0.41 | < 0.5 | 0.45
| Mean |  | 2.04: 0.385; 2.05: 0.48226; 2.06: 0.46667; 2.12: 0.41667; 2.14: 0.42833; 2.18: 0.43667
| Median |  | 2.04: 0.37; 2.05: 0.48131; 2.06: 0.4; 2.12: 0.4; 2.14: 0.425; 2.18: 0.44
| Std Dev. |  | 2.04: 0.053944; 2.05: 0.0093433; 2.06: 0.17512; 2.12: 0.040825; 2.14: 0.020412; 2.18: 0.013663
| Rel.Std.Dev. |  | 2.04: 14.01%; 2.05: 1.937%; 2.06: 37.53%; 2.12: 9.798%; 2.14: 4.766%; 2.18: 3.129%
| PDM3 |  | 2.04: -9.374%; 2.05: 13.52%; 2.06: 9.85%; 2.12: -1.92%; 2.14: 0.8266%; 2.18: 2.788%
| Z-Score (Absolute) |  | 2.78 | 0.67 | 0.84 | 0.6 | 0.18 | 8.9 | 0 | 2.78 | 0.13
| Analytical results for Zn in OREAS 163b (Certified Value 31.4 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.04 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | 2.17 | 2.20 | 2.21 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 1
1
| 1 | 1 | 30 | 57.9 | < 20 | 40 | 60 | 50 | < 20 | < 30 | 29 | < 20 | 27.59 | < 50 | 20 | 40 | 34 | 32 | [5 empty] | 1
| 1 | 2 | 30 | 55.6 | < 20 | 60 | 50 | 50 | < 20 | 30 | 29 | < 20 | 29.06 | < 50 | 20 | 39 | 33 | 29 | [5 empty] | 40
| 1 | 3 | 30 | 53.4 | < 20 | 50 | 60 | 30 | < 20 | 40 | 31 | < 20 | 28.77 | < 50 | 20 | 39 | 32 | 28 | [5 empty] | 16
| 1 | 4 | 30 | 51.6 | < 20 | 50 | 60 | 20 | < 20 | 30 | 30 | < 20 | 27.34 | < 50 | 21 | 40 | 34 | 30 | [5 empty] | 31.3796
| 1 | 5 | 30 | 53.1 | < 20 | 50 | 60 | 50 | < 20 | 30 | 32 | < 20 | 32.45 | < 50 | 21 | 39 | 29 | 31 | [5 empty] | 202
| 1 | 6 | 30 | 50.3 | < 20 | 50 | 60 | 30 | < 20 | < 30 | 27 | < 20 | 26.29 | < 50 | 21 | 39 | 36 | 33
| Mean |  | 30 | 53.65 |  | 50 | 58.3333 | 38.3333 |  | 32.5 | 29.6667 |  | 28.5833 |  | 20.5 | 39.3333 | 33 | 30.5
| Median |  | 30 | 53.25 |  | 50 | 60 | 40 |  | 30 | 29.5 |  | 28.18 |  | 20.5 | 39 | 33.5 | 30.5
| Std Dev. |  | 0 | 2.7443 |  | 6.3246 | 4.0825 | 13.2916 |  | 5 | 1.7512 |  | 2.145 |  | 0.54772 | 0.5164 | 2.3664 | 1.8708
| Rel.Std.Dev. |  | 0% | 5.115% |  | 12.65% | 6.999% | 34.67% |  | 15.38% | 5.903% |  | 7.504% |  | 2.672% | 1.313% | 7.171% | 6.134%
| PDM3 |  | -4.397% | 70.97% |  | 59.34% | 85.9% | 22.16% |  | 3.57% | -5.459% |  | -8.911% |  | -34.67% | 25.35% | 5.164% | -2.803%
| Z-Score (Absolute) |  | 0.01 | 1.8 | 1.5 | 1.53 | 2.16 | 0.64 | 1.5 | 0.24 | 0.01 | 1.5 | 0.09 | 0.37 | 0.71 | 0.72 | 0.24 | 0.05
| Analytical results for Zr in OREAS 163b (Certified Value 21.7 ppm) | [22 empty] | Certified
Zr | Round | Replicate | 2.01: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.12: Lab; 2.13: Lab; 2.14: Lab | [6 empty] | 1
| No | No | 2.01: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.13; 2.12: 2.18; 2.13: 2.19; 2.14: 2.20 | [6 empty] | ppm
|  |  | 2.01: PF*OES; 2.04: BF*MS; 2.05: PF*OES/MS; 2.06: PF*MS; 2.12: PF*MS; 2.13: BF*OES/MS; 2.14: BF*OES | [6 empty] | 1
2
| 1 | 1 | 2.01: 19.8; 2.04: 20; 2.05: 23.5605; 2.06: 20.7; 2.12: 22; 2.13: 23.4; 2.14: 23 | [6 empty] | 1
| 1 | 2 | 2.01: 20.2; 2.04: 20; 2.05: 24.3652; 2.06: 20.8; 2.12: 21; 2.13: 21.7; 2.14: 23 | [6 empty] | 41
| 1 | 3 | 2.01: 22; 2.04: 18; 2.05: 22.5288; 2.06: 22.9; 2.12: 21; 2.13: 22.5; 2.14: 23 | [6 empty] | 16
| 1 | 4 | 2.01: 20.2; 2.04: 20; 2.05: 24.4713; 2.06: 19.8; 2.12: 14; 2.13: 23.8; 2.14: 23 | [6 empty] | 21.6982
| 1 | 5 | 2.01: 17.3; 2.04: 22; 2.05: 23.4737; 2.06: 22.4; 2.12: 21; 2.13: 22.1; 2.14: 23 | [6 empty] | 203
| 1 | 6 | 2.01: 19.4; 2.04: 20; 2.05: 24.0236; 2.06: 20.7; 2.12: 22; 2.13: 22.8; 2.14: 23
| Mean |  | 2.01: 19.8167; 2.04: 20; 2.05: 23.7372; 2.06: 21.2167; 2.12: 20.1667; 2.13: 22.7167; 2.14: 23
| Median |  | 2.01: 20; 2.04: 20; 2.05: 23.7921; 2.06: 20.75; 2.12: 21; 2.13: 22.65; 2.14: 23
| Std Dev. |  | 2.01: 1.5211; 2.04: 1.2649; 2.05: 0.7178; 2.06: 1.1788; 2.12: 3.0605; 2.13: 0.78846; 2.14: 0
| Rel.Std.Dev. |  | 2.01: 7.676%; 2.04: 6.325%; 2.05: 3.024%; 2.06: 5.556%; 2.12: 15.18%; 2.13: 3.471%; 2.14: 0%
| PDM3 |  | 2.01: -8.671%; 2.04: -7.826%; 2.05: 9.397%; 2.06: -2.219%; 2.12: -7.058%; 2.13: 4.694%; 2.14: 6%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.04: 0.59; 2.05: 1.21; 2.06: 0; 2.12: 0.51; 2.13: 0.72; 2.14: 0.86
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C983: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 163b (Certified Value 1.64 wt.%) | [24 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 1.63 | 1.68 | 1.52 | 1.6 | 1.6454 | 1.6 | 1.73 | 1.64 | 1.65 | 1.64 | 1.63 | 1.66 | 1.5827 | 1.68 | 1.66 | 1.67 | 1.73 | 1.62 | 1.5 | 1.67 | 1.6 | 1.65 | 1.71 | 1
| 1 | 2 | 1.63 | 1.7 | 1.52 | 1.59 | 1.6344 | 1.61 | 1.74 | 1.63 | 1.63 | 1.65 | 1.6 | 1.648 | 1.5518 | 1.61 | 1.66 | 1.67 | 1.65 | 1.62 | 1.51 | 1.64 | 1.63 | 1.65 | 1.72 | 0
| 1 | 3 | 1.63 | 1.73 | 1.54 | 1.58 | 1.6451 | 1.61 | 1.75 | 1.63 | 1.63 | 1.66 | 1.6 | 1.66 | 1.591 | 1.81 | 1.67 | 1.68 | 1.65 | 1.62 | 1.52 | 1.62 | 1.63 | 1.64 | 1.72 | 16
| 1 | 4 | 1.62 | 1.69 | 1.54 | 1.59 | 1.6326 | 1.59 | 1.73 | 1.64 | 1.65 | 1.64 | 1.59 | 1.651 | 1.5631 | 1.69 | 1.66 | 1.68 | 1.68 | 1.61 | 1.5 | 1.65 | 1.63 | 1.64 | 1.73 | 1.6441
| 1 | 5 | 1.63 | 1.69 | 1.66 | 1.57 | 1.6439 | 1.62 | 1.74 | 1.64 | 1.65 | 1.65 | 1.61 | 1.648 | 1.5888 | 1.71 | 1.67 | 1.67 | 1.72 | 1.63 | 1.51 | 1.67 | 1.62 | 1.64 | 1.71 | 205
| 1 | 6 | 1.62 | 1.69 | 1.53 | 1.61 | 1.649 | 1.66 | 1.76 | 1.63 | 1.64 | 1.65 | 1.63 | 1.644 | 1.6027 | 1.74 | 1.67 | 1.67 | 1.68 | 1.62 | 1.51 | 1.68 | 1.7 | 1.64 | 1.73
| Mean |  | 1.6267 | 1.6967 | 1.5517 | 1.59 | 1.6417 | 1.615 | 1.7417 | 1.635 | 1.6417 | 1.6483 | 1.61 | 1.6518 | 1.58 | 1.7067 | 1.665 | 1.6733 | 1.685 | 1.62 | 1.5083 | 1.655 | 1.635 | 1.6433 | 1.72
| Median |  | 1.63 | 1.69 | 1.535 | 1.59 | 1.6445 | 1.61 | 1.74 | 1.635 | 1.645 | 1.65 | 1.605 | 1.6495 | 1.5857 | 1.7 | 1.665 | 1.67 | 1.68 | 1.62 | 1.51 | 1.66 | 1.63 | 1.64 | 1.72
| Std Dev. |  | 0.005164 | 0.017512 | 0.053821 | 0.014142 | 0.0066389 | 0.02429 | 0.01169 | 0.0054772 | 0.0098319 | 0.0075277 | 0.016733 | 0.0067057 | 0.018972 | 0.066533 | 0.0054772 | 0.005164 | 0.033912 | 0.0063246 | 0.0075277 | 0.022583 | 0.033912 | 0.005164 | 0.0089443
| Rel.Std.Dev. |  | 0.3175% | 1.032% | 3.469% | 0.8894% | 0.4044% | 1.504% | 0.6712% | 0.335% | 0.5989% | 0.4567% | 1.039% | 0.406% | 1.201% | 3.898% | 0.329% | 0.3086% | 2.013% | 0.3904% | 0.4991% | 1.365% | 2.074% | 0.3142% | 0.52%
| PDM3 |  | -1.058% | 3.2% | -5.62% | -3.288% | -0.1418% | -1.767% | 5.937% | -0.5508% | -0.1453% | 0.2602% | -2.071% | 0.4731% | -3.896% | 3.808% | 1.274% | 1.781% | 2.49% | -1.463% | -8.255% | 0.6657% | -0.5508% | -0.04395% | 4.619%
| Z-Score (Absolute) |  | 0.38 | 1.39 | 2.27 | 1.31 | 0 | 0.67 | 2.52 | 0.17 | 0 | 0.17 | 0.8 | 0.26 | 1.56 | 1.64 | 0.59 | 0.8 | 1.09 | 0.55 | 3.37 | 0.33 | 0.17 | 0.04 | 1.97
| Analytical results for S in OREAS 163b (Certified Value 3.37 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 3.62 | 2.59 | 3.27 | 3.35 | 3.4119 | 3.48 | 3.4 | 3.37 | 3.43 | 3.46 | 3.32 | 3.255 | 3.4539 | 3.48 | 3.41 | 3.4 | 3.3 | 3.43 | 3.25 | 3.13 | 3.36 | 3.43 | 3.28 | 1
| 1 | 2 | 3.6 | 2.63 | 3.26 | 3.43 | 3.4326 | 3.41 | 3.43 | 3.32 | 3.33 | 3.5 | 3.38 | 3.297 | 3.4302 | 3.67 | 3.38 | 3.36 | 3.41 | 3.34 | 3.28 | 3.12 | 3.39 | 3.39 | 3.31 | 0
| 1 | 3 | 3.62 | 2.69 | 3.29 | 3.32 | 3.4207 | 3.35 | 3.44 | 3.34 | 3.26 | 3.5 | 3.48 | 3.278 | 3.4442 | 3.59 | 3.36 | 3.42 | 3.7 | 3.45 | 3.3 | 3.13 | 3.36 | 3.43 | 3.28 | 16
| 1 | 4 | 3.62 | 2.64 | 3.25 | 3.36 | 3.4344 | 3.48 | 3.39 | 3.36 | 3.38 | 3.39 | 3.39 | 3.302 | 3.4703 | 3.52 | 3.4 | 3.42 | 3.38 | 3.36 | 3.3 | 3.16 | 3.36 | 3.45 | 3.33 | 3.3731
| 1 | 5 | 3.64 | 2.6 | 3.78 | 3.37 | 3.4245 | 3.46 | 3.41 | 3.27 | 3.4 | 3.43 | 3.39 | 3.286 | 3.4522 | 3.64 | 3.41 | 3.43 | 3.31 | 3.33 | 3.28 | 3.14 | 3.38 | 3.47 | 3.31 | 206
| 1 | 6 | 3.57 | 2.62 | 3.27 | 3.39 | 3.4094 | 3.27 | 3.38 | 3.33 | 3.4 | 3.52 | 3.34 | 3.298 | 3.471 | 3.62 | 3.43 | 3.43 | 3.78 | 3.37 | 3.26 | 3.11 | 3.37 | 3.44 | 3.3
| Mean |  | 3.6117 | 2.6283 | 3.3533 | 3.37 | 3.4222 | 3.4083 | 3.4083 | 3.3317 | 3.3667 | 3.4667 | 3.3833 | 3.286 | 3.4536 | 3.5867 | 3.3983 | 3.41 | 3.48 | 3.38 | 3.2783 | 3.1317 | 3.37 | 3.435 | 3.3017
| Median |  | 3.62 | 2.625 | 3.27 | 3.365 | 3.4226 | 3.435 | 3.405 | 3.335 | 3.39 | 3.48 | 3.385 | 3.2915 | 3.453 | 3.605 | 3.405 | 3.42 | 3.395 | 3.365 | 3.28 | 3.13 | 3.365 | 3.435 | 3.305
| Std Dev. |  | 0.024014 | 0.035449 | 0.20944 | 0.037417 | 0.010315 | 0.084242 | 0.023166 | 0.035449 | 0.06186 | 0.049666 | 0.055377 | 0.017584 | 0.015626 | 0.073121 | 0.024833 | 0.026833 | 0.20717 | 0.04899 | 0.020412 | 0.017224 | 0.012649 | 0.026646 | 0.019408
| Rel.Std.Dev. |  | 0.6649% | 1.349% | 6.246% | 1.11% | 0.3014% | 2.472% | 0.6797% | 1.064% | 1.837% | 1.433% | 1.637% | 0.5351% | 0.4525% | 2.039% | 0.7307% | 0.7869% | 5.953% | 1.449% | 0.6226% | 0.55% | 0.3753% | 0.7757% | 0.5878%
| PDM3 |  | 7.074% | -22.08% | -0.5849% | -0.09082% | 1.458% | 1.046% | 1.046% | -1.227% | -0.1896% | 2.775% | 0.3045% | -2.581% | 2.388% | 6.333% | 0.7492% | 1.095% | 3.17% | 0.2056% | -2.808% | -7.157% | -0.09082% | 1.836% | -2.117%
| Z-Score (Absolute) |  | 2.98 | 9.85 | 0.39 | 0.17 | 0.51 | 0.33 | 0.33 | 0.67 | 0.22 | 1.09 | 0 | 1.27 | 0.92 | 2.65 | 0.2 | 0.35 | 1.26 | 0.04 | 1.37 | 3.28 | 0.17 | 0.67 | 1.07
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 163b (Indicative Value 3.65 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3.8 | 8
| Mean |  | 3.65 | 16
| Median |  | 3.65 | 3.65
| Std Dev. |  | 0.21213 | 31
| Rel.Std.Dev. |  | 5.812%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 163b (Indicative Value 167 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 164 | 1
| 1 | 2 | 169 | 9
| Mean |  | 166.5 | 16
| Median |  | 166.5 | 166.5
| Std Dev. |  | 3.5355 | 32
| Rel.Std.Dev. |  | 2.123%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 163b (Indicative Value 36 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.5 | 1
| 1 | 2 | 36.5 | 27
| Mean |  | 36 | 16
| Median |  | 36 | 36
| Std Dev. |  | 0.70711 | 33
| Rel.Std.Dev. |  | 1.964%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 163b (Indicative Value 0.4 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 28
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 163b (Indicative Value 16 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16.2 | 1
| 1 | 2 | 15.8 | 29
| Mean |  | 16 | 16
| Median |  | 16 | 16
| Std Dev. |  | 0.28284 | 35
| Rel.Std.Dev. |  | 1.768%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 163b (Indicative Value 0.35 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.4 | 11
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.070711 | 36
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 163b (Indicative Value 9.58 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.6 | 1
| 1 | 2 | 9.56 | 31
| Mean |  | 9.58 | 16
| Median |  | 9.58 | 9.58
| Std Dev. |  | 0.028284 | 37
| Rel.Std.Dev. |  | 0.2952%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 163b (Indicative Value 227 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 229 | 1
| 1 | 2 | 225 | 32
| Mean |  | 227 | 16
| Median |  | 227 | 227
| Std Dev. |  | 2.8284 | 38
| Rel.Std.Dev. |  | 1.246%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 163b (Indicative Value 10 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 9 | 1
| 1 | 2 | 11 | 33
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 1.4142 | 39
| Rel.Std.Dev. |  | 14.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 163b (Indicative Value 0.9 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.92 | 1
| 1 | 2 | 0.87 | 13
| Mean |  | 0.895 | 16
| Median |  | 0.895 | 0.895
| Std Dev. |  | 0.035355 | 40
| Rel.Std.Dev. |  | 3.95%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 163b (Indicative Value 1.52 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.51 | 1
| 1 | 2 | 1.53 | 35
| Mean |  | 1.52 | 16
| Median |  | 1.52 | 1.52
| Std Dev. |  | 0.014142 | 41
| Rel.Std.Dev. |  | 0.9304%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 163b (Indicative Value 1.24 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.17 | 3
| Mean |  | 1.235 | 16
| Median |  | 1.235 | 1.235
| Std Dev. |  | 0.091924 | 42
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 163b (Indicative Value 0.56 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.54 | 1
| 1 | 2 | 0.57 | 4
| Mean |  | 0.555 | 16
| Median |  | 0.555 | 0.555
| Std Dev. |  | 0.021213 | 43
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 163b (Indicative Value 0.16 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.16 | 1
| 1 | 2 | 0.16 | 5
| Mean |  | 0.16 | 16
| Median |  | 0.16 | 0.16
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 163b (Indicative Value 1.35 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.3 | 17
| Mean |  | 1.35 | 16
| Median |  | 1.35 | 1.35
| Std Dev. |  | 0.070711 | 45
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 163b (Indicative Value 1.56 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.56 | 1
| 1 | 2 | 1.56 | 6
| Mean |  | 1.56 | 16
| Median |  | 1.56 | 1.56
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 163b (Indicative Value 0.78 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.75 | 1
| 1 | 2 | 0.8 | 19
| Mean |  | 0.775 | 16
| Median |  | 0.775 | 0.775
| Std Dev. |  | 0.035355 | 47
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 163b (Indicative Value 0.65 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.59 | 23
| Mean |  | 0.645 | 16
| Median |  | 0.645 | 0.645
| Std Dev. |  | 0.077782 | 31
| Rel.Std.Dev. |  | 12.06%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 163b (Indicative Value 0.23 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.23 | 1
| 1 | 2 | 0.22 | 3
| Mean |  | 0.225 | 16
| Median |  | 0.225 | 0.225
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 163b (Indicative Value 0.85 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.85 | 1
| 1 | 2 | 0.85 | 10
| Mean |  | 0.85 | 16
| Median |  | 0.85 | 0.85
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 163b (Indicative Value 4.76 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.73 | 1
| 1 | 2 | 4.79 | 11
| Mean |  | 4.76 | 16
| Median |  | 4.76 | 4.76
| Std Dev. |  | 0.042426 | 34
| Rel.Std.Dev. |  | 0.8913%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 163b (Indicative Value 0.055 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.05 | 4
| Mean |  | 0.055 | 16
| Median |  | 0.055 | 0.055
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 163b (Indicative Value 0.045 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0454 | 1
| 1 | 2 | 0.0449 | 30
| Mean |  | 0.04515 | 16
| Median |  | 0.04515 | 0.04515
| Std Dev. |  | 0.00035355 | 36
| Rel.Std.Dev. |  | 0.7831%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 163b (Indicative Value 1.5 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.6 | 31
| Mean |  | 1.5 | 16
| Median |  | 1.5 | 1.5
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 163b (Indicative Value 1.76 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.74 | 1
| 1 | 2 | 1.77 | 14
| Mean |  | 1.755 | 16
| Median |  | 1.755 | 1.755
| Std Dev. |  | 0.021213 | 38
| Rel.Std.Dev. |  | 1.209%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 163b (Indicative Value 5.19 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.22 | 1
| 1 | 2 | 5.16 | 5
| Mean |  | 5.19 | 16
| Median |  | 5.19 | 5.19
| Std Dev. |  | 0.042426 | 39
| Rel.Std.Dev. |  | 0.8175%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 163b (Indicative Value 26 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 26 | 16
| Mean |  | 26 | 16
| Median |  | 26 | 26
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 163b (Indicative Value 47.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 47 | 1
| 1 | 2 | 48 | 17
| Mean |  | 47.5 | 16
| Median |  | 47.5 | 47.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 1.489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 163b (Indicative Value 1.28 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.3 | 18
| Mean |  | 1.275 | 16
| Median |  | 1.275 | 1.275
| Std Dev. |  | 0.035355 | 42
| Rel.Std.Dev. |  | 2.773%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 163b (Indicative Value 4.33 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.25 | 37
| Mean |  | 4.325 | 16
| Median |  | 4.325 | 4.325
| Std Dev. |  | 0.10607 | 43
| Rel.Std.Dev. |  | 2.452%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 163b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 163b (Indicative Value 9.7 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.8 | 1
| 1 | 2 | 9.6 | 39
| Mean |  | 9.7 | 16
| Median |  | 9.7 | 9.7
| Std Dev. |  | 0.14142 | 45
| Rel.Std.Dev. |  | 1.458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 163b (Indicative Value 1.6 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.6 | 19
| Mean |  | 1.6 | 16
| Median |  | 1.6 | 1.6
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 163b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 163b (Indicative Value 1.44 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.48 | 1
| 1 | 2 | 1.4 | 5
| Mean |  | 1.44 | 16
| Median |  | 1.44 | 1.44
| Std Dev. |  | 0.056569 | 31
| Rel.Std.Dev. |  | 3.928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 163b (Indicative Value 3 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3.2 | 26
| Mean |  | 3 | 16
| Median |  | 3 | 3
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 163b (Indicative Value 8.75 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.7 | 1
| 1 | 2 | 8.8 | 27
| Mean |  | 8.75 | 16
| Median |  | 8.75 | 8.75
| Std Dev. |  | 0.070711 | 33
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 163b (Indicative Value 2.33 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.22 | 1
| 1 | 2 | 2.44 | 10
| Mean |  | 2.33 | 16
| Median |  | 2.33 | 2.33
| Std Dev. |  | 0.15556 | 34
| Rel.Std.Dev. |  | 6.677%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 163b (Indicative Value 0.23 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.22 | 1
| 1 | 2 | 0.23 | 11
| Mean |  | 0.225 | 16
| Median |  | 0.225 | 0.225
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 163b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 12
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 163b (Indicative Value 0.92 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.94 | 1
| 1 | 2 | 0.89 | 13
| Mean |  | 0.915 | 16
| Median |  | 0.915 | 0.915
| Std Dev. |  | 0.035355 | 37
| Rel.Std.Dev. |  | 3.864%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 163b (Indicative Value 0.029 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0295 | 1
| 1 | 2 | 0.0292 | 32
| Mean |  | 0.02935 | 16
| Median |  | 0.02935 | 0.02935
| Std Dev. |  | 0.00021213 | 38
| Rel.Std.Dev. |  | 0.7228%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 163b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 14
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 163b (Indicative Value 0.075 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.07 | 1
| 1 | 2 | 0.08 | 6
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 163b (Indicative Value 0.69 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.69 | 1
| 1 | 2 | 0.68 | 16
| Mean |  | 0.685 | 16
| Median |  | 0.685 | 0.685
| Std Dev. |  | 0.0070711 | 41
| Rel.Std.Dev. |  | 1.032%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 163b (Indicative Value 10.7 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.8 | 1
| 1 | 2 | 10.5 | 36
| Mean |  | 10.65 | 16
| Median |  | 10.65 | 10.65
| Std Dev. |  | 0.21213 | 42
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 163b (Indicative Value 2 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 2 | 17
| Mean |  | 2 | 16
| Median |  | 2 | 2
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 163b (Indicative Value 6.76 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.92 | 1
| 1 | 2 | 6.6 | 18
| Mean |  | 6.76 | 16
| Median |  | 6.76 | 6.76
| Std Dev. |  | 0.22627 | 44
| Rel.Std.Dev. |  | 3.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 163b (Indicative Value 0.44 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.43 | 1
| 1 | 2 | 0.45 | 19
| Mean |  | 0.44 | 16
| Median |  | 0.44 | 0.44
| Std Dev. |  | 0.014142 | 45
| Rel.Std.Dev. |  | 3.214%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 163b (Indicative Value 20 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 40
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 163b (Indicative Value 21.5 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.5 | 1
| 1 | 2 | 21.5 | 41
| Mean |  | 21.5 | 16
| Median |  | 21.5 | 21.5
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.