## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 993b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Umpire Labs (dry sample basis)
Classical Wet Chemistry
Cu, Copper (wt.%) | 23.4459 | 23.4192 | 23.4727 | 23.4177 | 23.4742
Pb Fire Assay (full corrections)
Au, Gold (ppm) | 33.3165 | 32.9066 | 33.7264 | 32.8741 | 33.7589
Acid Digestion (no HF)
Ag, Silver (ppm) | 40.6005 | 39.7303 | 41.4707 | 40.4163 | 40.7848
Geoanalytical Labs ('as received' sample basis)
Pb Fire Assay
Au, Gold (ppm) | 33.6477 | 32.3741 | 34.9212 | 33.3897 | 33.9057
Pd, Palladium (ppb) | 309.3386 | 297.2099 | 321.4672 | 301.6581 | 317.019
Pt, Platinum (ppb) | 44.2008 | 39.3035 | 49.0982 | 42.0794 | 46.3223
Oxidising Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 3.5288 | 3.4007 | 3.6569 | 3.4795 | 3.5782
CaO, Calcium oxide (wt.%) | 1.1573 | 1.111 | 1.2036 | 1.1364 | 1.1782
Co, Cobalt (ppm) | 193.0667 | 171.8353 | 214.298 | IND | IND
Cu, Copper (wt.%) | 23.6089 | 23.1047 | 24.1131 | 23.2742 | 23.9436
Fe, Iron (wt.%) | 24.6334 | 23.9909 | 25.2759 | 24.3464 | 24.9204
K2O, Potassium oxide (wt.%) | 0.84199 | 0.81227 | 0.87171 | 0.82727 | 0.85671
MgO, Magnesium oxide (wt.%) | 0.73395 | 0.68963 | 0.77827 | 0.72006 | 0.74783
Mo, Molybdenum (wt.%) | 0.2245 | 0.19927 | 0.24974 | 0.2111 | 0.2379
Ni, Nickel (ppm) | 233.4299 | 176.5787 | 290.2812 | 171.539 | 295.3208
P, Phosphorus (wt.%) | 0.028488 | 0.024805 | 0.032171 | IND | IND
S, Sulphur (wt.%) | 29.572 | 28.9688 | 30.1751 | 29.3136 | 29.8303
SiO2, Silicon dioxide (wt.%) | 11.5413 | 11.0764 | 12.0062 | 11.4121 | 11.6705
Sn, Tin (ppm) | < 100 | IND | IND | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.2018 | 0.18283 | 0.22078 | 0.19228 | 0.21133
W, Tungsten (ppm) | < 10 | IND | IND | IND | IND
Zn, Zinc (wt.%) | 0.33236 | 0.31077 | 0.35396 | 0.32687 | 0.33785
Thermogravimetry
LOI1000, Loss on ignition @1000 °C (wt.%) | 15.2068 | 14.9408 | 15.4727 | 15.0788 | 15.3347
Infrared Combustion
S, Sulphur (wt.%) | 29.6733 | 29.0294 | 30.3172 | 29.3745 | 29.972
Aqua Regia Digestion
Hg, Mercury (ppm) | 1.0665 | 0.99855 | 1.1344 | 1.0383 | 1.0946
Ion Selective Electrode
F, Fluorine (ppm) | 657.2879 | 630.2435 | 684.3323 | 631.1758 | 683.3999
4-Acid Digestion
Ag, Silver (ppm) | 40.357 | 39.2349 | 41.4792 | 39.7079 | 41.0062
As, Arsenic (ppm) | 102.5424 | 94.4192 | 110.6655 | 98.022 | 107.0628
Ba, Barium (ppm) | 95.5018 | 79.3623 | 111.6413 | 92.4111 | 98.5925
Be, Beryllium (ppm) | < 0.5 | IND | IND | IND | IND
Bi, Bismuth (ppm) | 22.8664 | 21.3398 | 24.393 | 22.3101 | 23.4227
Ca, Calcium (wt.%) | 0.82671 | 0.80241 | 0.85101 | 0.81072 | 0.8427
Cd, Cadmium (ppm) | 14.9714 | 14.1273 | 15.8155 | 14.3997 | 15.5431
Ce, Cerium (ppm) | 16.8095 | 15.33 | 18.2889 | 16.1957 | 17.4232
Co, Cobalt (ppm) | 153.9216 | 148.6799 | 159.1633 | 151.0778 | 156.7654
Cr, Chromium (ppm) | 57.2002 | 53.4838 | 60.9166 | 55.3989 | 59.0016
Cs, Caesium (ppm) | 0.64821 | 0.58035 | 0.71608 | 0.62279 | 0.67364
Cu, Copper (wt.%) | 23.5775 | 23.0951 | 24.0599 | 23.3658 | 23.7893
Dy, Dysprosium (ppm) | 1.0467 | 0.87611 | 1.2173 | IND | IND
Er, Erbium (ppm) | 0.59918 | 0.49236 | 0.706 | IND | IND
Fe, Iron (wt.%) | 24.5266 | 23.777 | 25.2761 | 24.1777 | 24.8754
Ga, Gallium (ppm) | 4.1404 | 3.877 | 4.4037 | 3.9706 | 4.3102
Gd, Gadolinium (ppm) | 1.2956 | 1.0675 | 1.5237 | IND | IND
Ge, Germanium (ppm) | 0.4495 | 0.1746 | 0.72439 | 0.3841 | 0.51489
Hf, Hafnium (ppm) | 0.70998 | 0.62718 | 0.79278 | 0.63995 | 0.78
Ho, Holmium (ppm) | 0.20192 | 0.18954 | 0.21429 | IND | IND
In, Indium (ppm) | 1.4539 | 1.3819 | 1.5259 | 1.398 | 1.5097
K, Potassium (wt.%) | 0.66846 | 0.64068 | 0.69624 | 0.65359 | 0.68334
La, Lanthanum (ppm) | 10.2212 | 9.2925 | 11.15 | 9.8284 | 10.6141
Li, Lithium (ppm) | 23.666 | 22.4143 | 24.9178 | 22.7407 | 24.5913
Mg, Magnesium (wt.%) | 0.4429 | 0.4218 | 0.46401 | 0.43194 | 0.45387
Mn, Manganese (wt.%) | 0.01386 | 0.013124 | 0.014596 | 0.013656 | 0.014064
Mo, Molybdenum (wt.%) | 0.21888 | 0.20896 | 0.22881 | 0.21544 | 0.22232
Na, Sodium (wt.%) | 0.42998 | 0.41634 | 0.44362 | 0.41774 | 0.44222
Nb, Niobium (ppm) | 1.8659 | 1.711 | 2.0209 | 1.7496 | 1.9822
Nd, Neodymium (ppm) | 7.3189 | 6.5786 | 8.0593 | 6.9697 | 7.6682
Ni, Nickel (ppm) | 255.658 | 246.1242 | 265.1918 | 250.1588 | 261.1572
P, Phosphorus (wt.%) | 0.020218 | 0.019009 | 0.021426 | 0.019207 | 0.021228
Pb, Lead (ppm) | 469.4268 | 448.1485 | 490.7052 | 461.9473 | 476.9064
Pr, Praseodymium (ppm) | 1.9361 | 1.6953 | 2.1768 | IND | IND
Rb, Rubidium (ppm) | 22.8846 | 21.4019 | 24.3674 | 22.0058 | 23.7635
Re, Rhenium (ppm) | 4.5386 | 4.0174 | 5.0598 | 4.3076 | 4.7696
S, Sulphur (wt.%) | 27.1863 | 25.6131 | 28.7596 | 26.7818 | 27.5909
Sb, Antimony (ppm) | 49.1187 | 45.4133 | 52.8242 | 47.4176 | 50.8198
Sc, Scandium (ppm) | 5.7534 | 5.2489 | 6.258 | 5.6478 | 5.8591
Se, Selenium (ppm) | 189.3073 | 172.1972 | 206.4173 | 182.2952 | 196.3194
Sn, Tin (ppm) | 1.9583 | 1.7515 | 2.165 | 1.778 | 2.1385
Sr, Strontium (ppm) | 159.4065 | 153.3899 | 165.4232 | 156.7364 | 162.0766
Ta, Tantalum (ppm) | 0.098687 | 0.091188 | 0.10618 | IND | IND
Tb, Terbium (ppm) | 0.19272 | 0.17441 | 0.21103 | IND | IND
Te, Tellurium (ppm) | 9.9405 | 8.8552 | 11.0258 | 9.2596 | 10.6213
Th, Thorium (ppm) | 2.0487 | 1.8866 | 2.2108 | 1.9988 | 2.0986
Ti, Titanium (wt.%) | 0.11804 | 0.11362 | 0.12247 | 0.11464 | 0.12145
Tl, Thallium (ppm) | 0.29124 | 0.25479 | 0.32769 | 0.27359 | 0.30889
U, Uranium (ppm) | 0.80139 | 0.73661 | 0.86618 | 0.74316 | 0.85962
V, Vanadium (ppm) | 62.787 | 59.5054 | 66.0685 | 60.8846 | 64.6893
W, Tungsten (ppm) | 2.4737 | 2.2724 | 2.675 | 2.313 | 2.6345
Y, Yttrium (ppm) | 5.6279 | 5.2309 | 6.0248 | 5.4018 | 5.8539
Yb, Ytterbium (ppm) | 0.58241 | 0.45738 | 0.70744 | IND | IND
Zn, Zinc (wt.%) | 0.33244 | 0.31749 | 0.34739 | 0.32697 | 0.33791
Zr, Zirconium (ppm) | 27.2724 | 25.4606 | 29.0841 | 26.4445 | 28.1003
3-Acid Digestion (no HF)
Ag, Silver (ppm) | 38.9118 | 36.4595 | 41.3641 | 38.2023 | 39.6213
As, Arsenic (ppm) | 106.9139 | 99.5966 | 114.2312 | 103.4202 | 110.4076
Co, Cobalt (ppm) | 151.0958 | 144.5591 | 157.6324 | 147.9187 | 154.2729
Ni, Nickel (ppm) | 252.9695 | 237.7716 | 268.1673 | 248.0862 | 257.8527
Pb, Lead (ppm) | 464.5927 | 439.2611 | 489.9244 | 457.4607 | 471.7248
Zn, Zinc (wt.%) | 0.34172 | 0.32559 | 0.35786 | 0.33575 | 0.34769
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 993b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Umpire Labs ('as received' sample basis)
Thermogravimetry
H2O- | wt.% | 0.44034
Geoanalytical Labs ('as received' sample basis)
Oxidising Fusion XRF
As | ppm | 108.3333 | Hg | ppm | < 100 | Se | ppm | 183.3333
Ba | ppm | 213.4683 | In | ppm | < 100 | Sr | ppm | 162.8724
Bi | ppm | 101.0667 | La2O3 | ppm | 119.4444 | Ta | ppm | 91.6667
Cd | ppm | 94 | Mn | wt.% | 0.013253 | Te | ppm | < 100
CeO2 | ppm | < 100 | Na | wt.% | 0.43067 | Tl | ppm | < 100
Cl | ppm | 226 | Nb | ppm | 212.5 | U | ppm | < 100
Cr | ppm | 123.4728 | Pb | ppm | 651.3754 | V | ppm | 81.5953
Cs | ppm | < 100 | Rb | ppm | < 50 | Y2O3 | ppm | 151.6667
Ga | ppm | < 100 | Re | ppm | < 100 | Zr | ppm | 133.3333
Ge | ppm | < 100 | Sb | ppm | 52.3333
HfO2 | ppm | < 100 | Sc | ppm | < 10
Infrared Combustion
C | wt.% | 0.15583
Ion Selective Electrode
Cl | ppm | 98
4-Acid Digestion
Al | wt.% | 1.8617 | Hg | ppm | 0.635 | Tm | ppm | < 0.1
B | ppm | 91.5 | Lu | ppm | 0.078703
Eu | ppm | 0.40467 | Sm | ppm | 1.5417
3-Acid Digestion (no HF)
Al | wt.% | 0.76755 | In | ppm | < 10 | Sc | ppm | 4.1667
Ba | ppm | 13.6356 | K | wt.% | 0.18947 | Se | ppm | 159.8333
Be | ppm | < 0.5 | La | ppm | 9.2448 | Sn | ppm | < 10
Bi | ppm | 13.9167 | Li | ppm | 20.145 | Sr | ppm | 94.5283
Ca | wt.% | 0.71538 | Mg | wt.% | 0.40492 | Te | ppm | 17.1667
Cd | ppm | 13.8333 | Mn | wt.% | 0.013113 | Ti | wt.% | 0.07977
Ce | ppm | 8.5 | Mo | wt.% | 0.22118 | Tl | ppm | < 5
Cr | ppm | 48.4425 | Na | wt.% | 0.09 | U | ppm | < 10
Cu | wt.% | 22.18 | Nb | ppm | < 10 | V | ppm | 47.5367
Fe | wt.% | 24.9908 | P | wt.% | 0.018242 | W | ppm | < 10
Ga | ppm | < 5 | Re | ppm | < 5 | Y | ppm | 4.8333
Ge | ppm | < 10 | S | wt.% | 30.5123 | Zr | ppm | 13.6667
Hg | ppm | < 1 | Sb | ppm | 12.1667
SI unit equivalents: ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 993b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Umpire Labs (dry sample basis)
Classical Wet Chemistry
Cu, wt.% | 23.4459 | 0.056417 | 23.3331 | 23.5588 | 23.2767 | 23.6152 | 0.2406% | 0.4813% | 0.7219% | 22.2736 | 24.6182
Pb Fire Assay (full corrections)
Au, ppm | 33.3165 | 0.56445 | 32.1876 | 34.4454 | 31.6231 | 35.0098 | 1.694% | 3.388% | 5.083% | 31.6507 | 34.9823
Acid Digestion (no HF)
Ag, ppm | 40.6005 | 1.5506 | 37.4993 | 43.7018 | 35.9486 | 45.2524 | 3.819% | 7.638% | 11.46% | 38.5705 | 42.6305
Geoanalytical Labs ('as received' sample basis)
Pb Fire Assay
Au, ppm | 33.6477 | 1.3006 | 31.0465 | 36.2489 | 29.7459 | 37.5495 | 3.865% | 7.731% | 11.6% | 31.9653 | 35.3301
Pd, ppb | 309.3386 | 13.8964 | 281.5457 | 337.1314 | 267.6493 | 351.0278 | 4.492% | 8.985% | 13.48% | 293.8716 | 324.8055
Pt, ppb | 44.2008 | 4.8553 | 34.4901 | 53.9115 | 29.6348 | 58.7669 | 10.98% | 21.97% | 32.95% | 41.9908 | 46.4109
Oxidising Fusion XRF
Al2O3, wt.% | 3.5288 | 0.11085 | 3.3071 | 3.7505 | 3.1963 | 3.8614 | 3.141% | 6.283% | 9.424% | 3.3524 | 3.7053
CaO, wt.% | 1.1573 | 0.053922 | 1.0495 | 1.2651 | 0.99554 | 1.3191 | 4.659% | 9.319% | 13.98% | 1.0994 | 1.2152
Co, ppm | 193.0667 | 14.6509 | 163.7649 | 222.3684 | 149.1141 | 237.0193 | 7.588% | 15.18% | 22.77% | 183.4133 | 202.72
Cu, wt.% | 23.6089 | 0.2974 | 23.0141 | 24.2037 | 22.7167 | 24.5011 | 1.26% | 2.519% | 3.779% | 22.4285 | 24.7893
Fe, wt.% | 24.6334 | 0.71146 | 23.2105 | 26.0563 | 22.499 | 26.7678 | 2.888% | 5.776% | 8.665% | 23.4017 | 25.8651
K2O, wt.% | 0.84199 | 0.021792 | 0.79841 | 0.88557 | 0.77662 | 0.90737 | 2.588% | 5.176% | 7.764% | 0.79989 | 0.88409
MgO, wt.% | 0.73395 | 0.031549 | 0.67085 | 0.79705 | 0.6393 | 0.82859 | 4.299% | 8.597% | 12.9% | 0.69725 | 0.77065
Mo, wt.% | 0.2245 | 0.016784 | 0.19094 | 0.25807 | 0.17415 | 0.27485 | 7.476% | 14.95% | 22.43% | 0.21328 | 0.23573
Ni, ppm | 233.4299 | 36.6752 | 160.0795 | 306.7804 | 123.4042 | 343.4556 | 15.71% | 31.42% | 47.13% | 221.7584 | 245.1014
P, wt.% | 0.028488 | 0.0033218 | 0.021844 | 0.035132 | 0.018522 | 0.038453 | 11.66% | 23.32% | 34.98% | 0.027064 | 0.029912
S, wt.% | 29.572 | 0.63323 | 28.3055 | 30.8384 | 27.6723 | 31.4716 | 2.141% | 4.283% | 6.424% | 28.0934 | 31.0505
SiO2, wt.% | 11.5413 | 0.56907 | 10.4031 | 12.6794 | 9.8341 | 13.2485 | 4.931% | 9.862% | 14.79% | 10.9642 | 12.1184
Sn, ppm | < 100 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
TiO2, wt.% | 0.2018 | 0.013134 | 0.17554 | 0.22807 | 0.1624 | 0.2412 | 6.508% | 13.02% | 19.52% | 0.19171 | 0.21189
W, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Zn, wt.% | 0.33236 | 0.016261 | 0.29984 | 0.36488 | 0.28358 | 0.38115 | 4.892% | 9.785% | 14.68% | 0.31575 | 0.34898
Thermogravimetry
LOI1000, wt.% | 15.2068 | 0.19397 | 14.8188 | 15.5947 | 14.6249 | 15.7887 | 1.276% | 2.551% | 3.827% | 14.4464 | 15.9671
Infrared Combustion
S, wt.% | 29.6733 | 0.72237 | 28.2285 | 31.118 | 27.5061 | 31.8404 | 2.434% | 4.869% | 7.303% | 28.1896 | 31.1569
Aqua Regia Digestion
Hg, ppm | 1.0665 | 0.069316 | 0.92783 | 1.2051 | 0.85852 | 1.2744 | 6.5% | 13% | 19.5% | 1.0131 | 1.1198
Ion Selective Electrode
F, ppm | 657.2879 | 27.9755 | 601.3368 | 713.239 | 573.3612 | 741.2145 | 4.256% | 8.512% | 12.77% | 624.4235 | 690.1523
4-Acid Digestion
Ag, ppm | 40.357 | 1.2573 | 37.8424 | 42.8717 | 36.5851 | 44.129 | 3.116% | 6.231% | 9.347% | 38.3392 | 42.3749
As, ppm | 102.5424 | 9.4946 | 83.5532 | 121.5316 | 74.0586 | 131.0261 | 9.259% | 18.52% | 27.78% | 97.4153 | 107.6695
Ba, ppm | 95.5018 | 20.3902 | 54.7215 | 136.2821 | 34.3314 | 156.6723 | 21.35% | 42.7% | 64.05% | 90.7267 | 100.2769
Be, ppm | < 0.5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Bi, ppm | 22.8664 | 1.9159 | 19.0345 | 26.6983 | 17.1186 | 28.6142 | 8.379% | 16.76% | 25.14% | 21.7231 | 24.0097
Ca, wt.% | 0.82671 | 0.024255 | 0.7782 | 0.87522 | 0.75395 | 0.89948 | 2.934% | 5.868% | 8.802% | 0.78538 | 0.86805
Cd, ppm | 14.9714 | 1.1073 | 12.7568 | 17.186 | 11.6495 | 18.2933 | 7.396% | 14.79% | 22.19% | 14.2228 | 15.72
Ce, ppm | 16.8095 | 1.0319 | 14.7457 | 18.8733 | 13.7137 | 19.9052 | 6.139% | 12.28% | 18.42% | 15.969 | 17.65
Co, ppm | 153.9216 | 6.8045 | 140.3127 | 167.5305 | 133.5082 | 174.335 | 4.421% | 8.841% | 13.26% | 146.2255 | 161.6177
Cr, ppm | 57.2002 | 4.8536 | 47.4931 | 66.9074 | 42.6396 | 71.7609 | 8.485% | 16.97% | 25.46% | 54.3402 | 60.0603
Cs, ppm | 0.64821 | 0.039496 | 0.56922 | 0.7272 | 0.52972 | 0.7667 | 6.093% | 12.19% | 18.28% | 0.6158 | 0.68062
Cu, wt.% | 23.5775 | 0.2558 | 23.0659 | 24.0891 | 22.8101 | 24.3449 | 1.085% | 2.17% | 3.255% | 22.3986 | 24.7564
Dy, ppm | 1.0467 | 0.092123 | 0.86245 | 1.2309 | 0.77032 | 1.3231 | 8.801% | 17.6% | 26.4% | 0.99436 | 1.099
Er, ppm | 0.59918 | 0.047369 | 0.50444 | 0.69392 | 0.45707 | 0.74129 | 7.906% | 15.81% | 23.72% | 0.56922 | 0.62914
Fe, wt.% | 24.5266 | 0.58112 | 23.3643 | 25.6888 | 22.7832 | 26.2699 | 2.369% | 4.739% | 7.108% | 23.3003 | 25.7529
Ga, ppm | 4.1404 | 0.27221 | 3.5959 | 4.6848 | 3.3237 | 4.957 | 6.575% | 13.15% | 19.72% | 3.9334 | 4.3474
Gd, ppm | 1.2956 | 0.13578 | 1.024 | 1.5671 | 0.88825 | 1.7029 | 10.48% | 20.96% | 31.44% | 1.2308 | 1.3604
Ge, ppm | 0.4495 | 0.17855 | 0.092397 | 0.80659 | 0 | 0.98514 | 39.72% | 79.44% | 119.2% | 0.42702 | 0.47197
Hf, ppm | 0.70998 | 0.066784 | 0.57641 | 0.84354 | 0.50963 | 0.91033 | 9.406% | 18.81% | 28.22% | 0.67448 | 0.74548
Ho, ppm | 0.20192 | 0.0067517 | 0.18842 | 0.21542 | 0.18166 | 0.22217 | 3.344% | 6.688% | 10.03% | 0.19182 | 0.21202
In, ppm | 1.4539 | 0.036939 | 1.38 | 1.5278 | 1.3431 | 1.5647 | 2.541% | 5.081% | 7.622% | 1.3812 | 1.5266
K, wt.% | 0.66846 | 0.0347 | 0.59906 | 0.73786 | 0.56436 | 0.77256 | 5.191% | 10.38% | 15.57% | 0.63504 | 0.70188
La, ppm | 10.2212 | 0.691 | 8.8392 | 11.6033 | 8.1482 | 12.2943 | 6.76% | 13.52% | 20.28% | 9.7102 | 10.7323
Li, ppm | 23.666 | 1.1934 | 21.2792 | 26.0529 | 20.0858 | 27.2463 | 5.043% | 10.09% | 15.13% | 22.4827 | 24.8493
Mg, wt.% | 0.4429 | 0.028407 | 0.38609 | 0.49972 | 0.35768 | 0.52812 | 6.414% | 12.83% | 19.24% | 0.42076 | 0.46505
Mn, wt.% | 0.01386 | 0.00069097 | 0.012478 | 0.015242 | 0.011787 | 0.015933 | 4.985% | 9.97% | 14.96% | 0.013167 | 0.014553
Mo, wt.% | 0.21888 | 0.014238 | 0.19041 | 0.24736 | 0.17617 | 0.2616 | 6.505% | 13.01% | 19.51% | 0.20794 | 0.22983
Na, wt.% | 0.42998 | 0.011528 | 0.40692 | 0.45304 | 0.3954 | 0.46456 | 2.681% | 5.362% | 8.043% | 0.40848 | 0.45148
Nb, ppm | 1.8659 | 0.15629 | 1.5533 | 2.1785 | 1.397 | 2.3348 | 8.376% | 16.75% | 25.13% | 1.7726 | 1.9592
Nd, ppm | 7.3189 | 0.4695 | 6.38 | 8.2579 | 5.9105 | 8.7274 | 6.415% | 12.83% | 19.24% | 6.953 | 7.6849
Ni, ppm | 255.658 | 10.1626 | 235.3329 | 275.9831 | 225.1703 | 286.1456 | 3.975% | 7.95% | 11.93% | 242.8751 | 268.4409
P, wt.% | 0.020218 | 0.001545 | 0.017128 | 0.023308 | 0.015583 | 0.024852 | 7.642% | 15.28% | 22.92% | 0.019207 | 0.021229
Pb, ppm | 469.4268 | 26.9491 | 415.5286 | 523.3251 | 388.5795 | 550.2742 | 5.741% | 11.48% | 17.22% | 445.9555 | 492.8982
Pr, ppm | 1.9361 | 0.18698 | 1.5621 | 2.31 | 1.3751 | 2.497 | 9.658% | 19.32% | 28.97% | 1.8393 | 2.0329
Rb, ppm | 22.8846 | 1.0493 | 20.786 | 24.9832 | 19.7367 | 26.0326 | 4.585% | 9.17% | 13.76% | 21.7404 | 24.0289
Re, ppm | 4.5386 | 0.61846 | 3.3017 | 5.7755 | 2.6832 | 6.394 | 13.63% | 27.25% | 40.88% | 4.3117 | 4.7655
S, wt.% | 27.1863 | 2.1856 | 22.8152 | 31.5575 | 20.6297 | 33.743 | 8.039% | 16.08% | 24.12% | 25.827 | 28.5457
Sb, ppm | 49.1187 | 4.7677 | 39.5833 | 58.6541 | 34.8157 | 63.4217 | 9.706% | 19.41% | 29.12% | 46.6628 | 51.5746
Sc, ppm | 5.7534 | 0.49749 | 4.7585 | 6.7484 | 4.261 | 7.2459 | 8.647% | 17.29% | 25.94% | 5.4658 | 6.0411
Se, ppm | 189.3073 | 17.8312 | 153.6448 | 224.9697 | 135.8135 | 242.801 | 9.419% | 18.84% | 28.26% | 179.8419 | 198.7726
Sn, ppm | 1.9583 | 0.14601 | 1.6662 | 2.2503 | 1.5202 | 2.3963 | 7.456% | 14.91% | 22.37% | 1.8603 | 2.0562
Sr, ppm | 159.4065 | 5.7031 | 148.0003 | 170.8127 | 142.2972 | 176.5158 | 3.578% | 7.155% | 10.73% | 151.4362 | 167.3769
Ta, ppm | 0.098687 | 0.011079 | 0.076528 | 0.12085 | 0.065448 | 0.13192 | 11.23% | 22.45% | 33.68% | 0.093752 | 0.10362
Tb, ppm | 0.19272 | 0.012664 | 0.1674 | 0.21805 | 0.15473 | 0.23072 | 6.571% | 13.14% | 19.71% | 0.18309 | 0.20236
Te, ppm | 9.9405 | 0.9814 | 7.9777 | 11.9033 | 6.9963 | 12.8846 | 9.873% | 19.75% | 29.62% | 9.4434 | 10.4375
Th, ppm | 2.0487 | 0.10033 | 1.848 | 2.2494 | 1.7477 | 2.3497 | 4.897% | 9.795% | 14.69% | 1.9463 | 2.1511
Ti, wt.% | 0.11804 | 0.0039311 | 0.11018 | 0.12591 | 0.10625 | 0.12984 | 3.33% | 6.66% | 9.991% | 0.11214 | 0.12395
Tl, ppm | 0.29124 | 0.02269 | 0.24586 | 0.33662 | 0.22317 | 0.35931 | 7.791% | 15.58% | 23.37% | 0.27668 | 0.3058
U, ppm | 0.80139 | 0.069125 | 0.66314 | 0.93964 | 0.59402 | 1.0088 | 8.626% | 17.25% | 25.88% | 0.76132 | 0.84146
V, ppm | 62.787 | 3.9551 | 54.8768 | 70.6972 | 50.9217 | 74.6523 | 6.299% | 12.6% | 18.9% | 59.6476 | 65.9263
W, ppm | 2.4737 | 0.17138 | 2.131 | 2.8165 | 1.9596 | 2.9879 | 6.928% | 13.86% | 20.78% | 2.35 | 2.5974
Y, ppm | 5.6279 | 0.29001 | 5.0478 | 6.2079 | 4.7578 | 6.4979 | 5.153% | 10.31% | 15.46% | 5.3465 | 5.9093
Yb, ppm | 0.58241 | 0.05853 | 0.46535 | 0.69947 | 0.40682 | 0.758 | 10.05% | 20.1% | 30.15% | 0.55329 | 0.61153
Zn, wt.% | 0.33244 | 0.016694 | 0.29905 | 0.36583 | 0.28236 | 0.38252 | 5.022% | 10.04% | 15.06% | 0.31582 | 0.34906
Zr, ppm | 27.2724 | 1.7235 | 23.8255 | 30.7193 | 22.102 | 32.4428 | 6.319% | 12.64% | 18.96% | 25.9088 | 28.636
3-Acid Digestion (no HF)
Ag, ppm | 38.9118 | 3.2938 | 32.3243 | 45.4993 | 29.0305 | 48.7931 | 8.465% | 16.93% | 25.39% | 36.9662 | 40.8574
As, ppm | 106.9139 | 8.0919 | 90.7302 | 123.0976 | 82.6383 | 131.1895 | 7.569% | 15.14% | 22.71% | 101.5682 | 112.2596
Co, ppm | 151.0958 | 6.31 | 138.4758 | 163.7158 | 132.1658 | 170.0258 | 4.176% | 8.352% | 12.53% | 143.541 | 158.6506
Ni, ppm | 252.9695 | 16.5323 | 219.9048 | 286.0341 | 203.3725 | 302.5664 | 6.535% | 13.07% | 19.61% | 240.321 | 265.6179
Pb, ppm | 464.5927 | 28.9564 | 406.68 | 522.5055 | 377.7237 | 551.4618 | 6.233% | 12.47% | 18.7% | 441.3631 | 487.8224
Zn, wt.% | 0.34172 | 0.016357 | 0.30901 | 0.37444 | 0.29265 | 0.3908 | 4.787% | 9.573% | 14.36% | 0.32464 | 0.35881
SI unit equivalents: ppb (parts per billion; 1 x 10-⁹) ≡ µg/kg; ppm (parts per million; 1 x 10-⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 993b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
3A*AAS | 3-acid (HNO3-HCIO4-HCI) digestion with atomic absorption spectroscopy
3A*MS | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma mass spectroscopy
3A*OES | 3-acid (HNO3-HCIO4-HCI) digestion with inductively coupled plasma optical emission spectroscopy
3A*OES/MS | 3-acid (HNO3-HCIO4-HCI) digestion with ICP-OES or ICP-MS finish
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AD*AAS | acid digestion (nonspecified) with atomic absorption spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*CVAAS | aqua regia digestion with cold vapour atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
at 105°C*Oven | moisture at 105°C with drying oven finish
at 105°C*TGA | moisture at 105°C by thermogravimetric analysis
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CIC | combustion ion chromatography
EG | electrogravimetry
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
ISE | ion specific electrode
LOI*Furnace | loss on ignition using a muffle furnace
LOI*TGA | loss on ignition by thermogravimetric analysis
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PPP*XRF | pressed powder pellet with X-ray fluorescence spectroscopy
SIT | short iodide titration
VOL*TITR | volumetric method with titration
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 993b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AH Knight, St Helens, Merseyside, UK
AH Knight, Tianjin, China
AHK Mongolia LLC, Ulaanbaatar, Mongolia
Alex Stewart International, Liverpool, UK
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Ulaanbaatar, Khan-Uul District, Mongolia
ALS, Vancouver, BC, Canada
ALS Inspection, Prescot, Merseyside, UK
American Assay Laboratories, Sparks, Nevada, USA
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bachelet, Angleur, Liege, Belgium
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas MET, Whyalla, SA, Australia
CERTIMIN, Lima, Peru
Erdenet Central Chemical Laboratory, Erdenet, Orkhon province, Mongolia
Independent, Perth, WA, Australia
Inspectorate (BV), Lima, Peru
Inspectorate (BV), Shanghai, Bao Shan District, China
Inspectorate (BV), Witham, Essex, UK
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek, Cupang, Muntinlupa, Philippines
Intertek, Perth, WA, Australia
Intertek LSI, Rotterdam, Zuid-Holland, Netherlands
Newcrest Laboratory Services, Orange, NSW, Australia
Ok Tedi Mine Lab, Mt Fubilan, Western Province, PNG
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 993b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
1 | 1 | ACTAn | 1 | 2 | 1 | 281115 | 1.05 | 33.5 | 33.6929
1 | 1 | ACTAn | 1 | 15 | 12 | 281116 | 1.09 | 33.5 | 33.6929
1 | 1 | ACTAn | 1 | 18 | 5 | 281117 | 1.05 | 33.7 | 33.7575
1 | 1 | ACTAn | 1 | 6 | 11 | 281118 | 1.07 | 34.2 | 33.9191
1 | 1 | ACTAn | 1 | 8 | 4 | 281119 | 1 | 34.6 | 34.0483
1 | 1 | ACTAn | 1 | 1 | 1 | 281120 | 1.02 | 33.4 | 33.6606
1 | 1 | ACTAn | 1 | 12 | 2 | 281121 | 1.03 | 32.9 | 33.499
1 | 1 | ACTAn | 1 | 7 | 4 | 281122 | 1.06 | 33.6 | 33.7252
1 | 1 | ACTAn | 1 | 13 | 7 | 281123 | 1.06 | 34 | 33.8545
1 | 1 | ACTAn | 1 | 5 | 11 | 281124 | 1.02 | 33.8 | 33.7898
1 | 1 | ACTAn | 1 | 20 | 10 | 281125 | 1.02 | 34.2 | 33.9191
1 | 1 | ACTAn | 1 | 9 | 9 | 281126 | 1.04 | 34 | 33.8545
1 | 1 | ACTAn | 1 | 17 | 5 | 281127 | 1.02 | 33.5 | 33.6929
1 | 1 | ACTAn | 1 | 4 | 6 | 281128 | 1.06 | 34 | 33.8545
1 | 1 | ACTAn | 1 | 10 | 9 | 281129 | 1.09 | 34 | 33.8545
1 | 1 | ACTAn | 1 | 11 | 2 | 281130 | 1.05 | 33.2 | 33.596
1 | 1 | ACTAn | 1 | 16 | 12 | 281131 | 1.04 | 34.2 | 33.9191
1 | 1 | ACTAn | 1 | 3 | 6 | 281132 | 1.02 | 34.4 | 33.9837
1 | 1 | ACTAn | 1 | 19 | 10 | 281133 | 1.05 | 33.2 | 33.596
1 | 1 | ACTAn | 1 | 14 | 7 | 281134 | 1.04 | 33.8 | 33.7898
Mean | [6 empty] | 1.044 | 33.785 | 33.785
Median | [7 empty] | 33.8 | 33.7898
SD | [7 empty] | 0.43803 | 0.14153
RSD | [7 empty] | 1.297% | 0.4189%
RSD Upscale Factor | [7 empty] | 0.32311
Parameters | :
Nominal Mass Value |  | 10 | gm

## Classical
| Analytical results for Cu in OREAS 993b (Certified Value 23.45 wt.%) | [19 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | 3.12 | 3.13 | 3.14 | 3.15 | 3.16 | 3.17 | 6.01 | wt.%
|  |  | SIT | SIT | SIT | SIT | SIT | VOL*TITR | SIT | SIT | SIT | SIT | SIT | SIT | SIT | SIT | SIT | SIT | EG | SIT | 2
3
| 1 | 1 | 23.454 | 23.4109 | 23.4755 | 23.478 | 23.4256 | 23.4273 | 23.4 | 23.5983 | 23.5495 | 23.34 | 23.5359 | 23.3844 | 23.47 | 23.327 | 23.49 | 23.4577 | 23.29 | 23.431 | 1
| 1 | 2 | 23.4238 | 23.421 | 23.4755 | 23.4591 | 23.3193 | 23.4273 | 23.33 | 23.5579 | 23.5696 | 23.23 | 23.6364 | 23.4445 | 23.47 | 23.4373 | 23.45 | 23.4592 | 23.274 | 23.4611 | 31
| 1 | 3 | 23.4238 | 23.431 |  | 23.4345 | 23.4152 | 23.4373 | 23.34 | 23.5579 | 23.5294 | 23.35 | 23.4053 | 23.5648 |  | 23.3571 |  | 23.4602 | 23.251 | 23.4811 | 16
| 1 | 4 | 3.01: 23.3937; 3.02: 23.4109; 3.04: 23.4157; 3.06: 23.4273; 6.01: 23.5111 | 23.4459
| 1 | 5 | 3.02: 23.431; 6.01: 23.4911 | 9
| 1 | 6 | 3.02: 23.4915; 6.01: 23.5011
| Mean |  | 23.4238 | 23.4327 | 23.4755 | 23.4468 | 23.3867 | 23.4298 | 23.3567 | 23.5714 | 23.5495 | 23.3067 | 23.5259 | 23.4646 | 23.47 | 23.3738 | 23.47 | 23.459 | 23.2717 | 23.4794
| Median |  | 23.4238 | 23.426 | 23.4755 | 23.4468 | 23.4152 | 23.4273 | 23.34 | 23.5579 | 23.5495 | 23.34 | 23.5359 | 23.4445 | 23.47 | 23.3571 | 23.47 | 23.4592 | 23.274 | 23.4861
| Std Dev. |  | 0.024604 | 0.030165 | 0 | 0.02733 | 0.058617 | 0.0050166 | 0.037859 | 0.02331 | 0.020111 | 0.066583 | 0.11585 | 0.091826 | 0 | 0.057051 | 0.028284 | 0.001239 | 0.019581 | 0.02932
| Rel.Std.Dev. |  | 0.105% | 0.1287% | 0% | 0.1166% | 0.2506% | 0.02141% | 0.1621% | 0.09889% | 0.0854% | 0.2857% | 0.4924% | 0.3913% | 0% | 0.2441% | 0.1205% | 0.005281% | 0.08414% | 0.1249%
| PDM3 |  | -0.09424% | -0.05636% | 0.1263% | 0.003753% | -0.2527% | -0.06876% | -0.3808% | 0.535% | 0.4418% | -0.594% | 0.3409% | 0.07951% | 0.1026% | -0.3077% | 0.1026% | 0.05591% | -0.7433% | 0.1428%
| Z-Score (Absolute) |  | 0.71 | 0.49 | 0.55 | 0.15 | 1.61 | 0.56 | 2.34 | 2.87 | 2.34 | 3.55 | 1.77 | 0.28 | 0.41 | 1.92 | 0.41 | 0.15 | 4.4 | 0.64
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fire Assay (Umpire)
| Analytical results for Au in OREAS 993b (Certified Value 33.32 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.00 | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | 3.12 | 3.13 | 3.15 | 3.16 | 3.17 | 6.01 | ppm
|  |  | INAA | FA*GRAV | FA*OES | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*AAS | FA*GRAV | FA*GRAV | FA*AAS | FA*AAS | FA*GRAV | FA*GRAV | FA*GRAV | 2
|  |  | 1.0g | 25g | 10g | 10g | 10g | 30g | 10g | 10g | 15g | 10g | 15g | 10g | 10g | 30g | 05g | 10g | 30g | 20g | 3
| 1 | 1 | 33.4 | 33.8501 | 32.6383 | 33.3367 | 32.4104 | 33.409 | 33.5006 | 33.5 | 34.2165 | 33.9869 | 32.9614 | 33.6012 | 33.7642 | 33.27 | 32.88 | 33.9997 | 32.1239 | 33.9795 | 1
| 1 | 2 | 33.5 | 33.147 | 32.8397 | 33.2663 | 33.4663 | 32.674 | 33.4103 | 33.9 | 34.0146 | 33.6853 | 32.2321 | 32.657 | 33.4636 | 33.19 | 33 | 34.1391 | 32.0337 | 33.8693 | 0
| 1 | 3 | 34.4 | 33.7496 | 32.9405 | 33.2763 | 32.6316 | 33.276 | 33.6109 | 33.7 | 34.2165 | 33.4842 | 32.7046 | 33.6012 | 33.4636 |  |  | 33.7474 | 32.0637 | 33.9093 | 16
| 1 | 4 | 1.00: 34; 3.01: 33.8501; 3.06: 33.4905; 3.12: 33.3634; 6.01: 33.9594 | 33.3165
| 1 | 5 | 1.00: 33.8; 3.01: 34.051; 3.12: 33.3634; 6.01: 33.9594 | 7
| 1 | 6 | 1.00: 34.2; 3.12: 33.3634; 6.01: 33.9394
|  | 7 | 1.00: 33.6
|  | 8 | 1.00: 34.6
|  | 9 | 1.00: 34
|  | 10 | 1.00: 34
|  | 11 | 1.00: 33.2
|  | 12 | 1.00: 32.9
|  | 13 | 1.00: 34
|  | 14 | 1.00: 33.8
|  | 15 | 1.00: 33.5
|  | 16 | 1.00: 34.2
|  | 17 | 1.00: 33.5
|  | 18 | 1.00: 33.7
|  | 19 | 1.00: 33.2
|  | 20 | 1.00: 34.2
| Mean |  | 33.785 | 33.7295 | 32.8062 | 33.2931 | 32.8361 | 33.1197 | 33.5031 | 33.7 | 34.1492 | 33.7188 | 32.6327 | 33.2865 | 33.4636 | 33.23 | 32.94 | 33.9621 | 32.0738 | 33.9361
| Median |  | 33.8 | 33.8501 | 32.8397 | 33.2763 | 32.6316 | 33.276 | 33.4955 | 33.7 | 34.2165 | 33.6853 | 32.7046 | 33.6012 | 33.4135 | 33.23 | 32.94 | 33.9997 | 32.0637 | 33.9494
| Std Dev. |  | 0.43803 | 0.34358 | 0.15388 | 0.038096 | 0.55684 | 0.39165 | 0.082481 | 0.2 | 0.11655 | 0.25305 | 0.36991 | 0.54516 | 0.15521 | 0.056569 | 0.084853 | 0.19855 | 0.045931 | 0.040403
| Rel.Std.Dev. |  | 1.297% | 1.019% | 0.469% | 0.1144% | 1.696% | 1.183% | 0.2462% | 0.5935% | 0.3413% | 0.7505% | 1.134% | 1.638% | 0.4638% | 0.1702% | 0.2576% | 0.5846% | 0.1432% | 0.1191%
| PDM3 |  | 1.406% | 1.24% | -1.532% | -0.07024% | -1.442% | -0.5908% | 0.56% | 1.151% | 2.499% | 1.208% | -2.052% | -0.09013% | 0.4415% | -0.2596% | -1.13% | 1.938% | -3.73% | 1.86%
| Z-Score (Absolute) |  | NA | 0.69 | 0.77 | 0 | 0.72 | 0.27 | 0.33 | 0.64 | 1.36 | 0.67 | 1.05 | 0.01 | 0.27 | 0.1 | 0.56 | 1.06 | 1.93 | 1.02
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AD (no HF)
| Analytical results for Ag in OREAS 993b (Certified Value 40.6 ppm) | [18 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | 3.12 | 3.13 | 3.14 | 3.15 | 3.17 | 6.01 | ppm
|  |  | 3A*AAS | 3A*AAS | 3A*OES | 3A*AAS | 3A*OES | 3A*AAS | 3A*AAS | AR*AAS | 3A*AAS | 4A*OES | 3A*OES | 3A*AAS | AR*OES/AAS | 3A*AAS | FA*GRAV | AD*AAS | AR*AAS | 1
|  |  | 3.08: 0.5g; 3.13: 0.4g; 3.15: 0.5g; 6.01: 01g | 2
| 1 | 1 | 40.1781 | 44.8272 | 40.2495 | 42.2351 | 37.1205 | 39.1291 | 38 | 42.3921 | 41.2267 | 42.3749 | 41.1552 | 42.08 | 41.1729 | 40.1324 | 39.42 | 36.945 | 41.0719 | 1
| 1 | 2 | 41.1826 | 44.4243 | 40.2495 | 42.2351 | 37.0305 | 39.1291 | 39 | 43.4015 | 42.2323 | 43.1797 | 39.8393 | 41.078 | 41.1729 | 40.1324 | 40.24 | 35.9727 | 41.0719 | 0
| 1 | 3 | 40.1781 | 44.4243 | 40.2495 | 43.2407 | 37.4708 | 38.1258 | 39 | 43.4015 | 41.2267 | 42.122 | 40.1105 | 40.0761 |  | 40.1324 |  | 36.4538 | 42.0736 | 16
| 1 | 4 | 3.01: 40.1781; 3.02: 44.2228; 3.04: 42.2351; 3.06: 39.1291; 6.01: 41.0719 | 40.6005
| 1 | 5 | 3.01: 40.1781; 6.01: 41.0719 | 11
| 1 | 6 | 3.01: 41.1826; 6.01: 41.0719
| Mean |  | 40.5129 | 44.4747 | 40.2495 | 42.4865 | 37.2073 | 38.8783 | 38.6667 | 43.065 | 41.5619 | 42.5589 | 40.3683 | 41.078 | 41.1729 | 40.1324 | 39.83 | 36.4572 | 41.2388
| Median |  | 40.1781 | 44.4243 | 40.2495 | 42.2351 | 37.1205 | 39.1291 | 39 | 43.4015 | 41.2267 | 42.3749 | 40.1105 | 41.078 | 41.1729 | 40.1324 | 39.83 | 36.4538 | 41.0719
| Std Dev. |  | 0.5187 | 0.25351 | 0 | 0.5028 | 0.23264 | 0.50166 | 0.57735 | 0.58274 | 0.58054 | 0.55234 | 0.69482 | 1.0019 | 0 | 0 | 0.57983 | 0.48613 | 0.40896
| Rel.Std.Dev. |  | 1.28% | 0.57% | 0% | 1.183% | 0.6252% | 1.29% | 1.493% | 1.353% | 1.397% | 1.298% | 1.721% | 2.439% | 0% | 0% | 1.456% | 1.333% | 0.9917%
| PDM3 |  | -0.2157% | 9.542% | -0.8644% | 4.645% | -8.358% | -4.242% | -4.763% | 6.07% | 2.368% | 4.824% | -0.5719% | 1.176% | 1.41% | -1.153% | -1.898% | -10.21% | 1.572%
| Z-Score (Absolute) |  | 0 | 2.55 | 0.17 | 1.27 | 2.12 | 1.05 | 1.19 | 1.64 | 0.67 | 1.32 | 0.09 | 0.36 | 0.42 | 0.24 | 0.44 | 2.61 | 0.47
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fire Assay
| Analytical results for Au in OREAS 993b (Certified Value 33.65 ppm) | [20 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.00 | 2.04 | 2.05 | 2.09 | 2.10 | 2.12 | 2.15 | 2.16 | 2.19 | [10 empty] | ppm
|  |  | INAA | FA*AAS | FA*GRAV | FA*AAS | FA*GRAV | FA*MS | FA*GRAV | FA*OES | FA*OES | [10 empty] | 2
|  |  | 1.0g | 30g | 30g | 40g | 30g | 50g | 30g | 30g | 30g | [10 empty] | 3
| 1 | 1 | 33.4 | 36.0954 | 34.1 | 34.29 | 33.1 | 32.285 | 33.4 | > 10 | 31.457 | [10 empty] | 1
| 1 | 2 | 33.5 | 36.3249 | 33.5 | 34.42 | 33.4 | 32.666 | 33.6 | > 10 | 31.279 | [10 empty] | 0
| 1 | 3 | 34.4 | 36.1612 | 33.3 | 34.43 | 33.2 | 32.857 | 33.5 | > 10 | 31.799 | [10 empty] | 16
| 1 | 4 | 34 | 35.9697 | 33.5 | 34.49 | 33.4 | 32.857 | 33.8 | > 10 | 32.13 | [10 empty] | 33.6477
| 1 | 5 | 33.8 | 36.2687 | 33.4 | 34.34 | 33.5 | 32.666 | 33.8 | > 10 | 32.481 | [10 empty] | 13
| 1 | 6 | 34.2 | 35.8636 | 33.8 | 34.34 | 33.4 | 35.428 | 33.7 | > 10 | 31.666
|  | 7 | 33.6
|  | 8 | 34.6
|  | 9 | 34
|  | 10 | 34
|  | 11 | 33.2
|  | 12 | 32.9
|  | 13 | 34
|  | 14 | 33.8
|  | 15 | 33.5
|  | 16 | 34.2
|  | 17 | 33.5
|  | 18 | 33.7
|  | 19 | 33.2
|  | 20 | 34.2
| Mean |  | 33.785 | 36.1139 | 33.6 | 34.385 | 33.3333 | 33.1265 | 33.6333 |  | 31.802
| Median |  | 33.8 | 36.1283 | 33.5 | 34.38 | 33.4 | 32.7615 | 33.65 |  | 31.7325
| Std Dev. |  | 0.43803 | 0.17581 | 0.29665 | 0.073959 | 0.15055 | 1.1467 | 0.1633 |  | 0.44265
| Rel.Std.Dev. |  | 1.297% | 0.4868% | 0.8829% | 0.2151% | 0.4517% | 3.462% | 0.4855% |  | 1.392%
| PDM3 |  | 0.4081% | 7.33% | -0.1417% | 2.191% | -0.9342% | -1.549% | -0.04264% |  | -5.485%
| Z-Score (Absolute) |  | NA | 3.58 | 0 | 1.12 | 0.38 | 0.67 | 0.05 | NA | 2.56
| Analytical results for Pd in OREAS 993b (Certified Value 309 ppb) | [20 empty] | Certified
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 5.01 | ppb
|  |  | FA*OES | FA*OES | FA*AAS | FA*MS | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*MS | 0
|  |  | 30g | 25g | 50g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 50g | 50g | 30g | 30g | 30g | 50g | 40g | 30g | 15g | 0
| 1 | 1 | 315 | 305 | 298 | 299 | 256 | 307 | 321 | 300 | 313 | 290 | 281 | 298.2993 | 296.2266 | 315 | 320 | 334 | 323 | 348 | 322 | 1
| 1 | 2 | 312 | 303 | 298 | 299 | 274 | 304 | 304 | 300 | 314 | 297 | 299 | 302.412 | 306.9155 | 315 | 316 | 315 | 324 | 335 | 320 | 0
| 1 | 3 | 316 | 309 | 298 | 295 | 313 | 304 | 325 | 300 | 319 | 287 | 298 | 299.7796 | 290.4051 | 311 | 319 | 323 | 321 | 330 | 331 | 16
| 1 | 4 | 310 | 313 | 300 | 297 | 259 | 307 | 314 | 300 | 312 | 294 | 298 | 307.0794 | 297.2436 | 311 | 316 | 320 | 329 | 347 | 310 | 309.3386
| 1 | 5 | 311 | 310 | 298 | 301 | 256 | 308 | 297 | 300 | 326 | 301 | 320 | 308.0852 | 289.5995 | 313 | 318 | 312 | 323 | 343 | 321 | 14
| 1 | 6 | 296 | 316 | 298 | 298 | 313 | 301 | 318 | 300 | 317 | 292 | 270 | 301.0032 | 308.7465 | 310 | 318 | 320 | 329 | 343
| Mean |  | 310 | 309.3333 | 298.3333 | 298.1667 | 278.5 | 305.1667 | 313.1667 | 300 | 316.8333 | 293.5 | 294.3333 | 302.7765 | 298.1895 | 312.5 | 317.8333 | 320.6667 | 324.8333 | 341 | 320.8
| Median |  | 311.5 | 309.5 | 298 | 298.5 | 266.5 | 305.5 | 316 | 300 | 315.5 | 293 | 298 | 301.7076 | 296.7351 | 312 | 318 | 320 | 323.5 | 343 | 321
| Std Dev. |  | 7.2388 | 4.8442 | 0.8165 | 2.0412 | 27.5445 | 2.6394 | 10.6849 | 0 | 5.1929 | 5.01 | 17.1891 | 3.975 | 8.0836 | 2.1679 | 1.6021 | 7.6333 | 3.3714 | 7.0711 | 7.4632
| Rel.Std.Dev. |  | 2.335% | 1.566% | 0.2737% | 0.6846% | 9.89% | 0.8649% | 3.412% | 0% | 1.639% | 1.707% | 5.84% | 1.313% | 2.711% | 0.6937% | 0.5041% | 2.38% | 1.038% | 2.074% | 2.326%
| PDM3 |  | 0.2138% | -0.001689% | -3.558% | -3.612% | -9.969% | -1.349% | 1.238% | -3.019% | 2.423% | -5.12% | -4.851% | -2.121% | -3.604% | 1.022% | 2.746% | 3.662% | 5.009% | 10.24% | 3.705%
| Z-Score (Absolute) |  | 0.04 | 0 | 0.67 | 0.68 | 1.89 | 0.26 | 0.23 | 0.57 | 0.46 | 0.97 | 0.92 | 0.4 | 0.68 | 0.19 | 0.52 | 0.69 | 0.95 | 1.94 | 0.7
| Analytical results for Pt in OREAS 993b (Certified Value 44.2 ppb) | [20 empty] | Certified
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 5.01 | ppb
|  |  | FA*OES | FA*OES | FA*AAS | FA*MS | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*MS | 1
|  |  | 30g | 25g | 50g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 50g | 50g | 30g | 30g | 30g | 50g | 40g | 30g | 15g | 1
| 1 | 1 | 46 | 41 | 50 | 42 | 34 | 44 | 49 | < 300 | 52 | 39 | 49 | 33.8729 | 38.6833 | 41 | 49 | 43 | 55 | 43 | 46 | 1
| 1 | 2 | 40 | 41 | 50 | 43 | 39 | 48 | 47 | < 300 | 45 | 39 | 49 | 39.0231 | 37.1719 | 48 | 47 | 44 | 60 | 45 | 46 | 0
| 1 | 3 | 43 | 42 | 50 | 43 | 46 | 39 | 51 | < 300 | 51 | 39 | 45 | 35.575 | 38.1833 | 42 | 48 | 45 | 50 | 40 | 47 | 16
| 1 | 4 | 49 | 42 | 50 | 37 | 37 | 47 | 50 | < 300 | 46 | 40 | 46 | 35.7269 | 38.3083 | 42 | 49 | 42 | 55 | 40 | 44 | 44.2008
| 1 | 5 | 46 | 40 | 49 | 44 | 35 | 43 | 44 | < 300 | 43 | 39 | 49 | 34.1227 | 37.8159 | 42 | 48 | 46 | 55 | 42 | 45 | 15
| 1 | 6 | 44 | 46 | 49 | 44 | 45 | 43 | 48 | < 300 | 49 | 39 | 53 | 38.7197 | 39.8874 | 43 | 48 | 42 | 50 | 44
| Mean |  | 44.6667 | 42 | 49.6667 | 42.1667 | 39.3333 | 44 | 48.1667 |  | 47.6667 | 39.1667 | 48.5 | 36.1734 | 38.3417 | 43 | 48.1667 | 43.6667 | 54.1667 | 42.3333 | 45.6
| Median |  | 45 | 41.5 | 50 | 43 | 38 | 43.5 | 48.5 |  | 47.5 | 39 | 49 | 35.6509 | 38.2458 | 42 | 48 | 43.5 | 55 | 42.5 | 46
| Std Dev. |  | 3.0768 | 2.0976 | 0.5164 | 2.6394 | 5.0859 | 3.2249 | 2.4833 |  | 3.559 | 0.40825 | 2.8107 | 2.2208 | 0.91401 | 2.5298 | 0.75277 | 1.633 | 3.7639 | 2.0656 | 1.1402
| Rel.Std.Dev. |  | 6.888% | 4.994% | 1.04% | 6.26% | 12.93% | 7.329% | 5.156% |  | 7.466% | 1.042% | 5.795% | 6.139% | 2.384% | 5.883% | 1.563% | 3.74% | 6.949% | 4.879% | 2.5%
| PDM3 |  | 1.054% | -4.979% | 12.37% | -4.602% | -11.01% | -0.4544% | 8.972% |  | 7.841% | -11.39% | 9.726% | -18.16% | -13.26% | -2.717% | 8.972% | -1.209% | 22.55% | -4.225% | 3.165%
| Z-Score (Absolute) |  | 0.14 | 0.3 | 0.96 | 0.28 | 0.74 | 0.03 | 0.72 | NA | 0.63 | 0.77 | 0.77 | 1.26 | 0.91 | 0.14 | 0.72 | 0.03 | 1.71 | 0.25 | 0.29
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Oxidising Fusion XRF
| Analytical results for Al2O3 in OREAS 993b (Certified Value 3.53 wt.%) | [13 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 3.4554 | 3.7 | 3.32 | 3.62 | 2.985 | 3.6 | 4.42 | 3.5359 | 3.5 | 3.741 | 3.3822 | 3.55 | 1
| 1 | 2 | 3.4365 | 3.63 | 3.39 | 3.64 | 3.004 | 3.5 | 4.58 | 3.5522 | 3.64 | 3.741 | 3.3633 | 3.51 | 0
| 1 | 3 | 3.4498 | 3.64 | 3.39 | 3.63 | 2.985 | 3.6 | 4.5 | 3.5516 | 3.53 | 3.685 | 3.3633 | 3.49 | 16
| 1 | 4 | 3.4442 | 3.65 | 3.44 | 3.7 | 3.023 | 3.6 | 4.49 | 3.5544 | 3.39 | 3.571 | 3.3444 | 3.39 | 3.5288
| 1 | 5 | 3.434 | 3.64 | 3.38 | 3.63 | 3.004 | 3.6 | 4.66 | 3.5729 | 3.51 | 3.685 | 3.42 | 3.49 | 17
| 1 | 6 | 3.4487 | 3.69 | 3.44 | 3.63 | 2.985 | 3.5 | 4.41 | 3.5479 | 3.47 | 3.533 | 3.4011
| Mean |  | 3.4448 | 3.6583 | 3.3933 | 3.6417 | 2.9977 | 3.5667 | 4.51 | 3.5525 | 3.5067 | 3.6593 | 3.379 | 3.486
| Median |  | 3.4465 | 3.645 | 3.39 | 3.63 | 2.9945 | 3.6 | 4.495 | 3.5519 | 3.505 | 3.685 | 3.3727 | 3.49
| Std Dev. |  | 0.0082274 | 0.029269 | 0.044572 | 0.029269 | 0.015513 | 0.05164 | 0.095917 | 0.011999 | 0.08165 | 0.087658 | 0.02782 | 0.058992
| Rel.Std.Dev. |  | 0.2388% | 0.8001% | 1.314% | 0.8037% | 0.5175% | 1.448% | 2.127% | 0.3378% | 2.328% | 2.395% | 0.8233% | 1.692%
| PDM3 |  | -2.382% | 3.67% | -3.84% | 3.198% | -15.05% | 1.072% | 27.8% | 0.6705% | -0.6281% | 3.698% | -4.245% | -1.214%
| Z-Score (Absolute) |  | 0.47 | 0.72 | 0.76 | 0.63 | 2.98 | 0.21 | 5.49 | 0.13 | 0.13 | 0.73 | 0.84 | 0.24
| Analytical results for As in OREAS 993b (Indicative Value 108 ppm) | [13 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.16 | 2.19 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [6 empty] | 0
0
| 1 | 1 | 100 | 200 | 100 | < 10 | 500 | < 100 | [6 empty] | 1
| 1 | 2 | 100 | 400 | 100 | < 10 | 600 | 200 | [6 empty] | 1
| 1 | 3 | 100 | 400 | 100 | < 10 | 500 | < 100 | [6 empty] | 16
| 1 | 4 | 100 | 400 | 200 | < 10 | 200 | 100 | [6 empty] | 108.3333
| 1 | 5 | 100 | 400 | < 100 | < 10 | 400 | 100 | [6 empty] | 7
| 1 | 6 | 100 | 400 | 200 | < 10 | 400 | 100
| Mean |  | 100 | 366.6667 | 140 |  | 433.3333 | 125
| Median |  | 100 | 400 | 100 |  | 450 | 100
| Std Dev. |  | 0 | 81.6497 | 54.7723 |  | 136.626 | 50
| Rel.Std.Dev. |  | 0% | 22.27% | 39.12% |  | 31.53% | 40%
| PDM3 |  | -7.692% | 238.5% | 29.23% |  | 300% | 15.38%
| Z-Score (Absolute) |  | 0.14 | 2.86 | 0.14 | 1.21 | 3.61 | 0.14
| Analytical results for Ba in OREAS 993b (Indicative Value 213 ppm) | [13 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF |  | 0
0
| 1 | 1 | 233 | 90 | < 90 | < 100 | < 10 | < 1000 | 90 | 273.6229 | 90 | < 100 | 457 |  | 1
| 1 | 2 | 224 | 180 | < 90 | < 100 | < 10 | < 1000 | 450 | 242.8348 | 180 | < 100 | 519 |  | 2
| 1 | 3 | 215 | 90 | < 90 | 100 | < 10 | < 1000 | 180 | 225.3695 | 90 | < 100 | 475 |  | 16
| 1 | 4 | 224 | 180 | < 90 | 100 | < 10 | < 1000 | 360 | 216.8607 | < 90 | < 100 | 403 |  | 213.4683
| 1 | 5 | 215 | 90 | < 90 | 100 | < 10 | < 1000 | 270 | 251.9033 | 180 | < 100 | 519 |  | 8
| 1 | 6 | 206 | 90 | < 90 | 100 | < 10 | < 1000 | < 90 | 269.9283 | 90 | < 100 | 519
| Mean |  | 219.5 | 120 |  | 100 |  |  | 270 | 246.7533 | 126 |  | 482
| Median |  | 219.5 | 90 |  | 100 |  |  | 270 | 247.369 | 90 |  | 497
| Std Dev. |  | 9.4393 | 46.4758 |  | 0 |  |  | 142.3025 | 23.0294 | 49.295 |  | 46.951
| Rel.Std.Dev. |  | 4.3% | 38.73% |  | 0% |  |  | 52.7% | 9.333% | 39.12% |  | 9.741%
| PDM3 |  | 2.826% | -43.79% |  | -53.15% |  |  | 26.48% | 15.59% | -40.97% |  | 125.8%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.5 | 0.24 | 0.77 | 2.57 | 0.75 | 0.86 | 0.05 | 0.47 | 2.44
| Analytical results for Bi in OREAS 993b (Indicative Value 101 ppm) | [13 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.16 | 2.19 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [6 empty] | 0
0
| 1 | 1 | 100 | < 100 | 100 | 66 | 200 | < 100 | [6 empty] | 1
| 1 | 2 | 100 | < 100 | 100 | 68 | 200 | < 100 | [6 empty] | 3
| 1 | 3 | 100 | < 100 | 100 | 56 | 200 | < 100 | [6 empty] | 16
| 1 | 4 | 100 | 100 | 100 | 65 | 100 | < 100 | [6 empty] | 101.0667
| 1 | 5 | 100 | < 100 | 100 | 67 | 200 | < 100 | [6 empty] | 9
| 1 | 6 | 100 | < 100 | 100 | 60 | 200 | < 100
| Mean |  | 100 | 100 | 100 | 63.6667 | 183.3333
| Median |  | 100 | 100 | 100 | 65.5 | 200
| Std Dev. |  | 0 |  | 0 | 4.6762 | 40.8248
| Rel.Std.Dev. |  | 0% |  | 0% | 7.345% | 22.27%
| PDM3 |  | -1.055% | -1.055% | -1.055% | -37.01% | 81.4%
| Z-Score (Absolute) |  | 0.59 | 0.76 | 0.59 | 0.59 | 3.29 | 1.03
| Analytical results for CaO in OREAS 993b (Certified Value 1.16 wt.%) | [13 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 1.1698 | 1.18 | 1.5 | 1.18 | 1.2313 | 1.2 | 1.06 | 1.066 | 1.16 | 1.399 | 1.21 | 1.13 | 1
| 1 | 2 | 1.1692 | 1.16 | 1.15 | 1.19 | 1.2453 | 1.2 | 0.99 | 1.0696 | 1.18 | 1.399 | 1.218 | 1.12 | 0
| 1 | 3 | 1.1646 | 1.17 | 1.11 | 1.19 | 1.2313 | 1.2 | 1.06 | 1.0774 | 1.18 | 1.343 | 1.204 | 1.11 | 16
| 1 | 4 | 1.1734 | 1.17 | 1.12 | 1.2 | 1.2313 | 1.2 | 1.1 | 1.0728 | 1.12 | 1.301 | 1.1941 | 1.16 | 1.1573
| 1 | 5 | 1.1672 | 1.17 | 1.15 | 1.18 | 1.2453 | 1.2 | 1.11 | 1.078 | 1.16 | 1.343 | 1.231 | 1.13 | 18
| 1 | 6 | 1.1691 | 1.18 | 1.15 | 1.18 | 1.1893 | 1.2 | 0.98 | 1.0702 | 1.14 | 1.315 | 1.234
| Mean |  | 1.1689 | 1.1717 | 1.1967 | 1.1867 | 1.229 | 1.2 | 1.05 | 1.0723 | 1.1567 | 1.35 | 1.2152 | 1.13
| Median |  | 1.1691 | 1.17 | 1.15 | 1.185 | 1.2313 | 1.2 | 1.06 | 1.0715 | 1.16 | 1.343 | 1.214 | 1.13
| Std Dev. |  | 0.0029178 | 0.0075277 | 0.14962 | 0.008165 | 0.020607 | 0 | 0.054406 | 0.0046906 | 0.023381 | 0.041294 | 0.015541 | 0.018708
| Rel.Std.Dev. |  | 0.2496% | 0.6425% | 12.5% | 0.6881% | 1.677% | 0% | 5.182% | 0.4374% | 2.021% | 3.059% | 1.279% | 1.656%
| PDM3 |  | 1% | 1.241% | 3.401% | 2.537% | 6.192% | 3.689% | -9.272% | -7.343% | -0.05516% | 16.65% | 5.001% | -2.359%
| Z-Score (Absolute) |  | 0.24 | 0.17 | 0.4 | 0.17 | 1.15 | 0.48 | 2.98 | 2.46 | 0.52 | 3.94 | 0.83 | 1.13
| Analytical results for Cd in OREAS 993b (Indicative Value 94 ppm) | [13 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.09 | 2.19 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | 99 | < 100 | [10 empty] | 1
| 1 | 2 | 92 | < 100 | [10 empty] | 4
| 1 | 3 | 92 | < 100 | [10 empty] | 16
| 1 | 4 | 98 | < 100 | [10 empty] | 94
| 1 | 5 | 88 | < 100 | [10 empty] | 10
| 1 | 6 | 95 | < 100
| Mean |  | 94
| Median |  | 93.5
| Std Dev. |  | 4.1473
| Rel.Std.Dev. |  | 4.412%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for CeO2 in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
CeO2 | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.05 | 2.06 | 2.16 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 100 | [9 empty] | 1
| 1 | 2 | < 100 | < 100 | 200 | [9 empty] | 5
| 1 | 3 | < 100 | < 100 | 200 | [9 empty] | 16
| 1 | 4 | < 100 | < 100 | < 100 | [9 empty] | < 100
| 1 | 5 | < 100 | < 100 | 200 | [9 empty] | 11
| 1 | 6 | < 100 | < 100 | 100
| Mean |  |  |  | 160
| Median |  |  |  | 200
| Std Dev. |  |  |  | 54.7723
| Rel.Std.Dev. |  |  |  | 34.23%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Cl in OREAS 993b (Indicative Value 226 ppm) | [13 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.19 | 2.20 | [10 empty] | ppm
|  |  | PPP*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | < 100 | 364 | [10 empty] | 1
| 1 | 2 | < 100 | 437 | [10 empty] | 6
| 1 | 3 | < 100 | 354 | [10 empty] | 16
| 1 | 4 | 100 | 406 | [10 empty] | 226
| 1 | 5 | < 100 | 354 | [10 empty] | 12
| 1 | 6 | < 100 | 447
| Mean |  | 100 | 393.6667
| Median |  | 100 | 385
| Std Dev. |  |  | 42.1932
| Rel.Std.Dev. |  |  | 10.72%
| PDM3 |  | -55.75% | 74.19%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Co in OREAS 993b (Certified Value 193 ppm) | [13 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.16 | 2.19 | 2.20 | [5 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [5 empty] | 0
0
| 1 | 1 | 100 | 200 | 200 | 200 | 200 | 100 | 162 | [5 empty] | 1
| 1 | 2 | 100 | 200 | 200 | 200 | 200 | 100 | 169 | [5 empty] | 0
| 1 | 3 | 100 | 200 | 200 | 200 | 200 | 100 | 162 | [5 empty] | 16
| 1 | 4 | 100 | 200 | 200 | 200 | 200 | 100 | 154 | [5 empty] | 193.0667
| 1 | 5 | 100 | 200 | 100 | 200 | 200 | 100 | 169 | [5 empty] | 19
| 1 | 6 | 100 | 200 | 200 | 200 | 200 | 200 | 176
| Mean |  | 100 | 200 | 183.3333 | 200 | 200 | 116.6667 | 165.3333
| Median |  | 100 | 200 | 200 | 200 | 200 | 100 | 165.5
| Std Dev. |  | 0 | 0 | 40.8248 | 0 | 0 | 40.8248 | 7.6333
| Rel.Std.Dev. |  | 0% | 0% | 22.27% | 0% | 0% | 34.99% | 4.617%
| PDM3 |  | -48.2% | 3.591% | -5.041% | 3.591% | 3.591% | -39.57% | -14.36%
| Z-Score (Absolute) |  | 3.37 | 0.67 | 0 | 0.67 | 0.67 | 2.7 | 0.73
| Analytical results for Cr in OREAS 993b (Indicative Value 123 ppm) | [13 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF |  |  | 0
0
| 1 | 1 | < 70 | 100 | 100 | 200 | < 10 | < 700 | < 70 | 140.3 | 100 | 100 |  |  | 1
| 1 | 2 | < 70 | 100 | 100 | 100 | < 10 | < 700 | < 70 | 126.6 | 100 | 100 |  |  | 7
| 1 | 3 | < 70 | 100 | 100 | 200 | < 10 | < 700 | < 70 | 154 | 100 | 200 |  |  | 16
| 1 | 4 | < 70 | 100 | 100 | 200 | < 10 | < 700 | < 70 | 154 | 100 | 200 |  |  | 123.4728
| 1 | 5 | < 70 | 100 | 100 | 100 | < 10 | < 700 | < 70 | 136.8 | 100 | 100 |  |  | 13
| 1 | 6 | < 70 | 100 | 100 | 200 | < 10 | < 700 | < 70 | 133.4 | 100 | 100
| Mean |  |  | 100 | 100 | 166.6667 |  |  |  | 140.85 | 100 | 133.3333
| Median |  |  | 100 | 100 | 200 |  |  |  | 138.55 | 100 | 100
| Std Dev. |  |  | 0 | 0 | 51.6398 |  |  |  | 11.1455 | 0 | 51.6398
| Rel.Std.Dev. |  |  | 0% | 0% | 30.98% |  |  |  | 7.913% | 0% | 38.73%
| PDM3 |  |  | -19.01% | -19.01% | 34.98% |  |  |  | 14.07% | -19.01% | 7.986%
| Z-Score (Absolute) |  | 0.83 | 0 | 0 | 0.84 | 1.2 | 3.06 | 0.83 | 0.52 | 0 | 0.42
| Analytical results for Cs in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Cs | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 8
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 993b (Certified Value 23.61 wt.%) | [13 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 |  |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF | OxBF*XRF |  |  | 2
3
| 1 | 1 | 23.7 | 25.1 | 23.8 | 23.67 | 23.4 | 25.0763 | > 20 | 22.3 | 23.3 | 23.67 |  |  | 1
| 1 | 2 | 23.6 | 23.7 | 23.9 | 23.38 | 23.8 | 25.3008 | > 20 | 21.98 | 23.7 | 23.41 |  |  | 31
| 1 | 3 | 23.8 | 23.4 | 23.6 | 23.25 | 23.8 | 25.1302 | > 20 | 22.05 | 23 | 23.56 |  |  | 16
| 1 | 4 | 23.7 | 24.4 | 23.7 | 23.36 | 23.8 | 25.1442 | > 20 | 22.36 | 23 | 23.56 |  |  | 23.6089
| 1 | 5 | 23.8 | 24.4 | 23.2 | 23.42 | 23.6 | 25.3962 | > 20 | 22.45 | 23.6 | 23.52 |  |  | 20
| 1 | 6 | 23.8 | 23.2 | 23.7 | 23.31 | 23.8 | 25.1897 | > 20 | 22.6 | 23.9
| Mean |  | 23.7333 | 24.0333 | 23.65 | 23.3983 | 23.7 | 25.2062 |  | 22.29 | 23.4167 | 23.544
| Median |  | 23.75 | 24.05 | 23.7 | 23.37 | 23.8 | 25.167 |  | 22.33 | 23.45 | 23.56
| Std Dev. |  | 0.08165 | 0.72296 | 0.2429 | 0.14552 | 0.16733 | 0.11985 |  | 0.23681 | 0.37639 | 0.093434
| Rel.Std.Dev. |  | 0.344% | 3.008% | 1.027% | 0.6219% | 0.706% | 0.4755% |  | 1.062% | 1.607% | 0.3969%
| PDM3 |  | 0.527% | 1.798% | 0.1741% | -0.8919% | 0.3859% | 6.766% |  | -5.586% | -0.8143% | -0.2749%
| Z-Score (Absolute) |  | 0.24 | 1.11 | 0 | 0.73 | 0.14 | 4.5 | NA | 3.93 | 0.67 | 0.31
| Analytical results for Fe in OREAS 993b (Certified Value 24.63 wt.%) | [13 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 23.4315 | 24.9 | 23.6 | 25.2 | 24.04 | 24.97 | 28.502 | 24.4927 | 24.4 | 23.67 | 25.76 | 25.07 | 1
| 1 | 2 | 23.4182 | 25.2 | 24.5 | 25.5 | 24.31 | 24.62 | 28.663 | 24.609 | 24.9 | 23.95 | 26.034 | 24.74 | 32
| 1 | 3 | 23.4525 | 25 | 24.1 | 25.5 | 24.33 | 24.62 | 28.901 | 24.5963 | 24.7 | 24.01 | 25.782 | 24.85 | 16
| 1 | 4 | 23.4252 | 24.9 | 24.5 | 25.5 | 24.45 | 24.9 | 29.02 | 24.6202 | 23.4 | 23.86 | 25.364 | 24.87 | 24.6334
| 1 | 5 | 23.4357 | 25.1 | 24.4 | 25.2 | 24.69 | 25.11 | 30.328 | 24.7291 | 24.3 | 23.84 | 26.292 | 24.82 | 21
| 1 | 6 | 23.4567 | 25 | 24.5 | 25.5 | 24.67 | 24.69 | 27.796 | 24.5608 | 24.1 | 24.08 | 26.413
| Mean |  | 23.4366 | 25.0167 | 24.2667 | 25.4 | 24.415 | 24.8183 | 28.8683 | 24.6013 | 24.3 | 23.9017 | 25.9408 | 24.87
| Median |  | 23.4336 | 25 | 24.45 | 25.5 | 24.39 | 24.795 | 28.782 | 24.6027 | 24.35 | 23.905 | 25.908 | 24.85
| Std Dev. |  | 0.01518 | 0.1169 | 0.36148 | 0.15492 | 0.24525 | 0.20488 | 0.83438 | 0.077717 | 0.52536 | 0.14497 | 0.38616 | 0.12227
| Rel.Std.Dev. |  | 0.06477% | 0.4673% | 1.49% | 0.6099% | 1.005% | 0.8255% | 2.89% | 0.3159% | 2.162% | 0.6065% | 1.489% | 0.4916%
| PDM3 |  | -4.858% | 1.556% | -1.489% | 3.112% | -0.8866% | 0.7507% | 17.19% | -0.1302% | -1.353% | -2.971% | 5.308% | 0.9604%
| Z-Score (Absolute) |  | 2.01 | 0.48 | 0.7 | 1.09 | 0.47 | 0.17 | 6.57 | 0.17 | 0.65 | 1.28 | 1.95 | 0.25
| Analytical results for Ga in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Ga | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 9
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 15
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Ge | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 10
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 16
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for HfO2 in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.14 | 2.16 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 100 | < 100 | 2025 | < 100 | [8 empty] | 1
| 1 | 2 | < 100 | < 100 | 2040 | < 100 | [8 empty] | 11
| 1 | 3 | < 100 | < 100 | 2070 | < 100 | [8 empty] | 16
| 1 | 4 | < 100 | < 100 | 2010 | < 100 | [8 empty] | < 100
| 1 | 5 | < 100 | < 100 | 2085 | < 100 | [8 empty] | 17
| 1 | 6 | < 100 | < 100 | 2015 | < 100
| Mean |  |  |  | 2040.8333
| Median |  |  |  | 2032.5
| Std Dev. |  |  |  | 30.5641
| Rel.Std.Dev. |  |  |  | 1.498%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Hg in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Hg | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 35
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 7
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
In | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 2
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 8
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 993b (Certified Value 0.842 wt.%) | [13 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.808 | 0.87 | 0.81 | 0.86 | 0.7348 | 0.9 | 0.95 | 0.83732 | 0.83 | 1.048 | 0.86298 |  | 1
| 1 | 2 | 0.808 | 0.86 | 0.83 | 0.87 | 0.7264 | 0.9 | 0.94 | 0.84145 | 0.85 | 1.024 | 0.85996 |  | 0
| 1 | 3 | 0.8154 | 0.87 | 0.82 | 0.86 | 0.7336 | 0.9 | 0.96 | 0.83881 | 0.84 | 1.012 | 0.85996 |  | 16
| 1 | 4 | 0.8108 | 0.85 | 0.83 | 0.87 | 0.7481 | 0.9 | 0.95 | 0.83944 | 0.8 | 0.976 | 0.84695 |  | 0.84199
| 1 | 5 | 0.816 | 0.85 | 0.83 | 0.86 | 0.7348 | 0.9 | 0.99 | 0.846 | 0.82 | 1.024 | 0.87394 |  | 22
| 1 | 6 | 0.8152 | 0.87 | 0.84 | 0.86 | 0.736 | 0.9 | 0.97 | 0.83542 | 0.82 | 0.988 | 0.87803
| Mean |  | 0.81223 | 0.86167 | 0.82667 | 0.86333 | 0.73562 | 0.9 | 0.96 | 0.83974 | 0.82667 | 1.012 | 0.86364
| Median |  | 0.813 | 0.865 | 0.83 | 0.86 | 0.7348 | 0.9 | 0.955 | 0.83913 | 0.825 | 1.018 | 0.86147
| Std Dev. |  | 0.003766 | 0.0098319 | 0.010328 | 0.005164 | 0.007019 | 1.2162e-16 | 0.017889 | 0.0036753 | 0.017512 | 0.026291 | 0.011129
| Rel.Std.Dev. |  | 0.4637% | 1.141% | 1.249% | 0.5981% | 0.9542% | 1.351e-14% | 1.863% | 0.4377% | 2.118% | 2.598% | 1.289%
| PDM3 |  | -3.534% | 2.337% | -1.82% | 2.535% | -12.63% | 6.889% | 14.02% | -0.2676% | -1.82% | 20.19% | 2.571%
| Z-Score (Absolute) |  | 1.08 | 0.31 | 0.67 | 0.35 | 3.23 | NA | 3.07 | 0.31 | 0.67 | 4.52 | 0.36
Results from laboratory 2.10 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for La2O3 in OREAS 993b (Indicative Value 119 ppm) | [13 empty] | Indicative
La2O3 | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.05 | 2.06 | 2.16 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 100 | < 100 | 200 | [9 empty] | 1
| 1 | 2 | < 100 | 100 | 300 | [9 empty] | 3
| 1 | 3 | < 100 | 100 | 200 | [9 empty] | 16
| 1 | 4 | < 100 | 100 | 100 | [9 empty] | 119.4444
| 1 | 5 | < 100 | 200 | 200 | [9 empty] | 9
| 1 | 6 | 100 | 100 | 100
| Mean |  | 100 | 120 | 183.3333
| Median |  | 100 | 100 | 200
| Std Dev. |  | 0 | 44.7214 | 75.2773
| Rel.Std.Dev. |  | 0% | 37.27% | 41.06%
| PDM3 |  | -16.28% | 0.4651% | 53.49%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.21
| Analytical results for MgO in OREAS 993b (Certified Value 0.734 wt.%) | [13 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF | OxBF*XRF | 3
3
| 1 | 1 | 0.5857 | 0.76 | 0.74 | 0.74 | 0.73 | 0.8 | 0.71 | 0.77581 | 0.88 | 0.912 | 0.68401 | 0.8 | 1
| 1 | 2 | 0.5873 | 0.75 | 0.75 | 0.74 | 0.763 | 0.8 | 0.7 | 0.78228 | 0.88 | 0.912 | 0.68699 | 0.8 | 0
| 1 | 3 | 0.5828 | 0.74 | 0.73 | 0.74 | 0.73 | 0.8 | 0.72 | 0.77746 | 0.89 | 0.895 | 0.6719 | 0.8 | 16
| 1 | 4 | 0.5891 | 0.75 | 0.73 | 0.74 | 0.73 | 0.8 | 0.73 | 0.77428 | 0.85 | 0.862 | 0.67704 | 0.8 | 0.73395
| 1 | 5 | 0.5843 | 0.72 | 0.76 | 0.73 | 0.746 | 0.8 | 0.77 | 0.78332 | 0.87 | 0.879 | 0.68699 | 0.8 | 23
| 1 | 6 | 0.5825 | 0.72 | 0.77 | 0.73 | 0.713 | 0.8 | 0.71 | 0.77594 | 0.87 | 0.862 | 0.65897
| Mean |  | 0.58528 | 0.74 | 0.74667 | 0.73667 | 0.73533 | 0.8 | 0.72333 | 0.77818 | 0.87333 | 0.887 | 0.67765 | 0.8
| Median |  | 0.585 | 0.745 | 0.745 | 0.74 | 0.73 | 0.8 | 0.715 | 0.7767 | 0.875 | 0.887 | 0.68053 | 0.8
| Std Dev. |  | 0.0025957 | 0.016733 | 0.01633 | 0.005164 | 0.017108 | 1.2162e-16 | 0.025033 | 0.0037316 | 0.013663 | 0.022926 | 0.01093 | 0
| Rel.Std.Dev. |  | 0.4435% | 2.261% | 2.187% | 0.701% | 2.326% | 1.52e-14% | 3.461% | 0.4795% | 1.564% | 2.585% | 1.613% | 0%
| PDM3 |  | -20.26% | 0.8246% | 1.733% | 0.3704% | 0.1888% | 9% | -1.446% | 6.027% | 18.99% | 20.85% | -7.671% | 9%
| Z-Score (Absolute) |  | 3.76 | 0.04 | 0.2 | 0.04 | 0.08 | NA | 0.37 | 0.98 | 3.32 | 3.66 | 1.49 | NA
Results from laboratories 2.10 and 5.01 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Mn in OREAS 993b (Indicative Value 0.013 wt.%) | [13 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.01015 | 0.02 | < 0.01 | 0.01 | < 0.001 | < 0.08 | < 0.008 | 0.0055103 | 0.02 | < 0.01 | 0.0122 |  | 1
| 1 | 2 | 0.01076 | 0.02 | < 0.01 | 0.01 | < 0.001 | < 0.08 | < 0.008 | 0.0059014 | 0.02 | < 0.01 | 0.013 |  | 4
| 1 | 3 | 0.00937 | 0.02 | < 0.01 | 0.01 | < 0.001 | < 0.08 | < 0.008 | 0.0059014 | 0.02 | < 0.01 | 0.0122 |  | 16
| 1 | 4 | 0.00937 | 0.02 | < 0.01 | 0.02 | < 0.001 | < 0.08 | < 0.008 | 0.0044299 | 0.02 | < 0.01 | 0.0101 |  | 0.013253
| 1 | 5 | 0.00937 | 0.01 | < 0.01 | 0.01 | < 0.001 | < 0.08 | < 0.008 | 0.0047435 | 0.02 | < 0.01 | 0.013 |  | 10
| 1 | 6 | 0.01015 | 0.02 | < 0.01 | 0.02 | < 0.001 | < 0.08 | < 0.008 | 0.0051501 | 0.02 | < 0.01 | 0.0158
| Mean |  | 0.0098617 | 0.018333 |  | 0.013333 |  |  |  | 0.0052728 | 0.02 |  | 0.012717
| Median |  | 0.00976 | 0.02 |  | 0.01 |  |  |  | 0.0053302 | 0.02 |  | 0.0126
| Std Dev. |  | 0.00058284 | 0.0040825 |  | 0.005164 |  |  |  | 0.00060867 | 0 |  | 0.0018465
| Rel.Std.Dev. |  | 5.91% | 22.27% |  | 38.73% |  |  |  | 11.54% | 0% |  | 14.52%
| PDM3 |  | -25.59% | 38.33% |  | 0.6068% |  |  |  | -60.21% | 50.91% |  | -4.046%
| Z-Score (Absolute) |  | 0 | 1.18 | 0.67 | 0.48 | 1.3 | 4 | 0.83 | 0.64 | 1.41 | 0.67 | 0.4
| Analytical results for Mo in OREAS 993b (Certified Value 0.225 wt.%) | [13 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.16 | 2.19 | 5.01 | [5 empty] | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.241 | 0.193 | 0.228 | 0.1999 | 0.244 | 0.31 | 0.24 | [5 empty] | 1
| 1 | 2 | 0.233 | 0.214 | 0.234 | 0.2027 | 0.252 | 0.33 | 0.22 | [5 empty] | 29
| 1 | 3 | 0.238 | 0.204 | 0.245 | 0.2007 | 0.248 | 0.35 | 0.22 | [5 empty] | 16
| 1 | 4 | 0.242 | 0.213 | 0.228 | 0.2028 | 0.221 | 0.32 | 0.23 | [5 empty] | 0.2245
| 1 | 5 | 0.24 | 0.209 | 0.232 | 0.2048 | 0.238 | 0.29 | 0.22 | [5 empty] | 24
| 1 | 6 | 0.251 | 0.21 | 0.224 | 0.2052 | 0.228 | 0.3
| Mean |  | 0.24083 | 0.20717 | 0.23183 | 0.20268 | 0.2385 | 0.31667 | 0.226
| Median |  | 0.2405 | 0.2095 | 0.23 | 0.20275 | 0.241 | 0.315 | 0.22
| Std Dev. |  | 0.0059133 | 0.0077825 | 0.0073326 | 0.0021217 | 0.011996 | 0.021602 | 0.0089443
| Rel.Std.Dev. |  | 2.455% | 3.757% | 3.163% | 1.047% | 5.03% | 6.822% | 3.958%
| PDM3 |  | 7.274% | -7.722% | 3.265% | -9.719% | 6.235% | 41.05% | 0.6669%
| Z-Score (Absolute) |  | 0.67 | 1.85 | 0 | 2.18 | 0.5 | 6.36 | 0.44
| Analytical results for Na in OREAS 993b (Indicative Value 0.431 wt.%) | [13 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.04 | 2.12 | 2.14 | 2.19 | [8 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | PPP*XRF | [8 empty] | 3
3
| 1 | 1 | 0.43724 | 0.512 | 0.42763 | 0.52 | [8 empty] | 1
| 1 | 2 | 0.43769 | 0.282 | 0.42797 | 0.52 | [8 empty] | 5
| 1 | 3 | 0.43769 | 0.415 | 0.42997 | 0.49 | [8 empty] | 16
| 1 | 4 | 0.44525 | 0.312 | 0.42363 | 0.48 | [8 empty] | 0.43067
| 1 | 5 | 0.4385 | 0.26 | 0.43197 | 0.52 | [8 empty] | 11
| 1 | 6 | 0.43264 | 0.942 | 0.41929 | 0.48
| Mean |  | 0.43817 | 0.45383 | 0.42674 | 0.50167
| Median |  | 0.43769 | 0.3635 | 0.4278 | 0.505
| Std Dev. |  | 0.0040532 | 0.25707 | 0.00459 | 0.020412
| Rel.Std.Dev. |  | 0.925% | 56.64% | 1.076% | 4.069%
| PDM3 |  | 1.742% | 5.38% | -0.9106% | 16.49%
| Z-Score (Absolute) |  | 0.39 | 0.39 | 0.96 | 2.78
| Analytical results for Nb in OREAS 993b (Indicative Value 213 ppm) | [13 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.16 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 50 | 260 | < 100 | 130 | [8 empty] | 1
| 1 | 2 | < 50 | 270 | < 100 | 140 | [8 empty] | 6
| 1 | 3 | < 50 | 290 | < 100 | 140 | [8 empty] | 16
| 1 | 4 | < 50 | 300 | < 100 | 150 | [8 empty] | 212.5
| 1 | 5 | < 50 | 290 | < 100 | 150 | [8 empty] | 12
| 1 | 6 | < 50 | 280 | < 100 | 150
| Mean |  |  | 281.6667 |  | 143.3333
| Median |  |  | 285 |  | 145
| Std Dev. |  |  | 14.7196 |  | 8.165
| Rel.Std.Dev. |  |  | 5.226% |  | 5.696%
| PDM3 |  |  | 32.55% |  | -32.55%
| Z-Score (Absolute) |  | 0.82 | 2.11 | 0.53 | 0.53
| Analytical results for Ni in OREAS 993b (Certified Value 233 ppm) | [13 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.14 | 2.16 | 2.19 | [5 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | [5 empty] | 0
0
| 1 | 1 | 220 | 240 | 300 | < 10 | 247.5 | 270 | 200 | [5 empty] | 1
| 1 | 2 | 180 | 280 | 200 | < 10 | 216.1 | 270 | 200 | [5 empty] | 0
| 1 | 3 | 190 | 250 | 200 | < 10 | 267.2 | 290 | 200 | [5 empty] | 16
| 1 | 4 | 220 | 280 | 200 | < 10 | 286.8 | 230 | 200 | [5 empty] | 233.4299
| 1 | 5 | 170 | 270 | 200 | < 10 | 239.7 | 260 | 200 | [5 empty] | 25
| 1 | 6 | 220 | 250 | 300 | < 10 | 216.1 | 240 | 200
| Mean |  | 200 | 261.6667 | 233.3333 |  | 245.5667 | 260 | 200
| Median |  | 205 | 260 | 200 |  | 243.6 | 265 | 200
| Std Dev. |  | 22.8035 | 17.224 | 51.6398 |  | 28.0817 | 21.9089 | 0
| Rel.Std.Dev. |  | 11.4% | 6.582% | 22.13% |  | 11.44% | 8.427% | 0%
| PDM3 |  | -14.32% | 12.1% | -0.04138% |  | 5.199% | 11.38% | -14.32%
| Z-Score (Absolute) |  | 0.79 | 0.67 | 0 | 5.43 | 0.29 | 0.63 | 0.79
| Analytical results for P in OREAS 993b (Certified Value 0.028 wt.%) | [13 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF |  |  | 3
3
| 1 | 1 | 0.02749 | 0.031 | 0.031 | 0.03 | 0.02 | 0.04 | < 0.004 | 0.03142 | 0.022 | 0.02 |  |  | 1
| 1 | 2 | 0.02754 | 0.031 | 0.031 | 0.03 | 0.02 | < 0.04 | < 0.004 | 0.03055 | 0.022 | 0.02 |  |  | 32
| 1 | 3 | 0.02745 | 0.031 | 0.031 | 0.03 | 0.02 | 0.04 | < 0.004 | 0.03164 | 0.022 | 0.02 |  |  | 16
| 1 | 4 | 0.02749 | 0.026 | 0.031 | 0.03 | 0.02 | 0.04 | < 0.004 | 0.03098 | 0.022 | 0.02 |  |  | 0.028488
| 1 | 5 | 0.02745 | 0.031 | 0.031 | 0.03 | 0.02 | 0.04 | < 0.004 | 0.03164 | 0.022 | 0.02 |  |  | 26
| 1 | 6 | 0.0274 | 0.031 | 0.031 | 0.03 | 0.02 | 0.04 | < 0.004 | 0.03142 | 0.022 | 0.02
| Mean |  | 0.02747 | 0.030167 | 0.031 | 0.03 | 0.02 | 0.04 |  | 0.031275 | 0.022 | 0.02
| Median |  | 0.02747 | 0.031 | 0.031 | 0.03 | 0.02 | 0.04 |  | 0.03142 | 0.022 | 0.02
| Std Dev. |  | 4.7749e-05 | 0.0020412 | 0 | 0 | 0 | 0 |  | 0.00042922 | 3.8006e-18 | 0
| Rel.Std.Dev. |  | 0.1738% | 6.767% | 0% | 0% | 0% | 0% |  | 1.372% | 1.728e-14% | 0%
| PDM3 |  | -3.573% | 5.893% | 8.818% | 5.308% | -29.79% | 40.41% |  | 9.783% | -22.77% | -29.79%
| Z-Score (Absolute) |  | 0.17 | 0.17 | 0.27 | 0.19 | 1.23 | 1.62 | 3.77 | 0.38 | 0.97 | 1.23
| Analytical results for Pb in OREAS 993b (Indicative Value 651 ppm) | [13 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.14 | 2.16 | 2.19 | 2.20 |  |  |  | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF |  |  |  | 0
0
| 1 | 1 | 480 | < 50 | 500 | 800 | < 1000 | 886.6 | 590 | 600 | 743 |  |  |  | 1
| 1 | 2 | 440 | < 50 | 500 | 800 | < 1000 | 933 | 660 | < 100 | 808 |  |  |  | 7
| 1 | 3 | 460 | < 50 | 400 | 800 | < 1000 | 891.2 | 660 | 600 | 724 |  |  |  | 16
| 1 | 4 | 460 | < 50 | 500 | 800 | < 1000 | 886.6 | 450 | < 100 | 668 |  |  |  | 651.3754
| 1 | 5 | 470 | < 50 | 400 | 900 | < 1000 | 891.2 | 570 | 600 | 808 |  |  |  | 13
| 1 | 6 | 470 | < 50 | 400 | 900 | < 1000 | 905.1 | 540 | 500 | 854
| Mean |  | 463.3333 |  | 450 | 833.3333 |  | 898.95 | 578.3333 | 575 | 767.5
| Median |  | 465 |  | 450 | 800 |  | 891.2 | 580 | 600 | 775.5
| Std Dev. |  | 13.6626 |  | 54.7723 | 51.6398 |  | 18.0133 | 79.3515 | 50 | 68.0522
| Rel.Std.Dev. |  | 2.949% |  | 12.17% | 6.197% |  | 2.004% | 13.72% | 8.696% | 8.867%
| PDM3 |  | -28.87% |  | -30.92% | 27.93% |  | 38.01% | -11.21% | -11.73% | 17.83%
| Z-Score (Absolute) |  | 0.25 | 3.2 | 0.34 | 2.25 | 0 | 2.69 | 0.53 | 0.67 | 1.8
| Analytical results for Rb in OREAS 993b (Indicative Value < 50 ppm) | [13 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.05 | 2.06 | 2.16 | 2.19 | [8 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [8 empty] | 0
0
| 1 | 1 | < 50 | < 50 | 120 | < 100 | [8 empty] | 1
| 1 | 2 | < 50 | < 50 | 140 | < 100 | [8 empty] | 8
| 1 | 3 | < 50 | < 50 | 130 | < 100 | [8 empty] | 16
| 1 | 4 | < 50 | < 50 | 70 | < 100 | [8 empty] | < 50
| 1 | 5 | < 50 | < 50 | 100 | < 100 | [8 empty] | 14
| 1 | 6 | < 50 | < 50 | 100 | < 100
| Mean |  |  |  | 110
| Median |  |  |  | 110
| Std Dev. |  |  |  | 25.2982
| Rel.Std.Dev. |  |  |  | 23%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67 | 3.91 | 0.67
| Analytical results for Re in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Re | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 9
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 15
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 993b (Certified Value 29.57 wt.%) | [13 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 |  |  | wt.%
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF |  |  | 2
3
| 1 | 1 | > 20 | > 20 | 29.4 | 30.235 | 27.85 | 28.989 | > 20 | 30.18 | 29.98 | 29.75 |  |  | 1
| 1 | 2 | > 20 | > 20 | 29.5 | 30.167 | 28.13 | 29.1446 | > 20 | 30.12 | 30.03 | 29.39 |  |  | 34
| 1 | 3 | > 20 | > 20 | 29.6 | 30.096 | 28.26 | 29.0769 | > 20 | 30.24 | 29.97 | 29.61 |  |  | 16
| 1 | 4 | > 20 | > 20 | 29.6 | 30.092 | 28.51 | 29.2294 | > 20 | 30.17 | 29.53 | 29.74 |  |  | 29.572
| 1 | 5 | > 20 | > 20 | 29.6 | 30.283 | 29.63 | 29.2258 | > 20 | 30.21 | 30.54 | 29.66 |  |  | 27
| 1 | 6 | > 20 | > 20 | 29.5 | 30.18 | 26.99 | 29.1974 | > 20 | 30.32 | 28.98
| Mean |  |  |  | 29.5333 | 30.1755 | 28.2283 | 29.1438 |  | 30.2067 | 29.8383 | 29.63
| Median |  |  |  | 29.55 | 30.1735 | 28.195 | 29.171 |  | 30.195 | 29.975 | 29.66
| Std Dev. |  |  |  | 0.08165 | 0.075474 | 0.86386 | 0.09527 |  | 0.068605 | 0.52883 | 0.14612
| Rel.Std.Dev. |  |  |  | 0.2765% | 0.2501% | 3.06% | 0.3269% |  | 0.2271% | 1.772% | 0.4931%
| PDM3 |  |  |  | -0.1306% | 2.041% | -4.544% | -1.448% |  | 2.146% | 0.9008% | 0.1963%
| Z-Score (Absolute) |  | NA | NA | 0.13 | 0.76 | 1.94 | 0.67 | NA | 0.8 | 0.29 | 0
| Analytical results for Sb in OREAS 993b (Indicative Value 52 ppm) | [13 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.16 | 2.19 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [6 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | 53 | < 50 | < 100 | [6 empty] | 1
| 1 | 2 | < 50 | 80 | < 100 | 55 | 70 | < 100 | [6 empty] | 10
| 1 | 3 | < 50 | 100 | 100 | 53 | < 50 | < 100 | [6 empty] | 16
| 1 | 4 | < 50 | 90 | < 100 | 56 | < 50 | < 100 | [6 empty] | 52.3333
| 1 | 5 | < 50 | 70 | < 100 | 58 | < 50 | < 100 | [6 empty] | 16
| 1 | 6 | < 50 | 60 | < 100 | 56 | < 50 | < 100
| Mean |  |  | 80 | 100 | 55.1667 | 70
| Median |  |  | 80 | 100 | 55.5 | 70
| Std Dev. |  |  | 15.8114 |  | 1.9408
| Rel.Std.Dev. |  |  | 19.76% |  | 3.518%
| PDM3 |  |  | 52.87% | 91.08% | 5.414% | 33.76%
| Z-Score (Absolute) |  | 1.55 | 1.03 | 0.32 | 0.15 | 1.13 | 0.15
| Analytical results for Sc in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.09 | 2.19 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | [10 empty] | 1
1
| 1 | 1 | < 10 | < 100 | [10 empty] | 1
| 1 | 2 | < 10 | < 100 | [10 empty] | 11
| 1 | 3 | < 10 | < 100 | [10 empty] | 16
| 1 | 4 | < 10 | < 100 | [10 empty] | < 10
| 1 | 5 | < 10 | < 100 | [10 empty] | 17
| 1 | 6 | < 10 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Se in OREAS 993b (Indicative Value 183 ppm) | [13 empty] | Indicative
Se | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | PPP*XRF | [11 empty] | 0
0
| 1 | 1 | 200 | [11 empty] | 1
| 1 | 2 | 200 | [11 empty] | 1
| 1 | 3 | 200 | [11 empty] | 16
| 1 | 4 | 200 | [11 empty] | 183.3333
| 1 | 5 | 200 | [11 empty] | 7
| 1 | 6 | 100
| Mean |  | 183.3333
| Median |  | 200
| Std Dev. |  | 40.8248
| Rel.Std.Dev. |  | 22.27%
| PDM3 |  | 1.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 993b (Certified Value 11.54 wt.%) | [13 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 | 5.01 | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 8.44 | 11.55 | 10.7 | 11.4 | 12.407 | 11.4 | 12.18 | 11.3018 | 11.3 | 10.439 | 12.044 | 11.61 | 1
| 1 | 2 | 8.33 | 11.5 | 11 | 11.45 | 12.493 | 11.3 | 12.17 | 11.3191 | 11.5 | 10.525 | 11.937 | 11.56 | 0
| 1 | 3 | 8.49 | 11.6 | 10.8 | 11.5 | 12.578 | 11.3 | 12.31 | 11.3182 | 11.4 | 10.461 | 11.958 | 11.52 | 16
| 1 | 4 | 8.372 | 11.55 | 10.95 | 11.5 | 12.578 | 11.4 | 12.16 | 11.2997 | 11 | 10.503 | 11.83 | 12.02 | 11.5413
| 1 | 5 | 8.41 | 11.6 | 11.3 | 11.45 | 12.792 | 11.5 | 12.85 | 11.3712 | 11.35 | 10.503 | 12.108 | 11.61 | 28
| 1 | 6 | 8.3 | 11.65 | 11.15 | 11.5 | 12.728 | 11.5 | 12.12 | 11.2776 | 11.15 | 10.461 | 12.129
| Mean |  | 8.3903 | 11.575 | 10.9833 | 11.4667 | 12.596 | 11.4 | 12.2983 | 11.3146 | 11.2833 | 10.482 | 12.001 | 11.664
| Median |  | 8.391 | 11.575 | 10.975 | 11.475 | 12.578 | 11.4 | 12.175 | 11.31 | 11.325 | 10.482 | 12.001 | 11.61
| Std Dev. |  | 0.070602 | 0.05244 | 0.22061 | 0.040825 | 0.1434 | 0.089443 | 0.27781 | 0.031586 | 0.18074 | 0.033051 | 0.1139 | 0.20256
| Rel.Std.Dev. |  | 0.8415% | 0.453% | 2.009% | 0.356% | 1.138% | 0.7846% | 2.259% | 0.2792% | 1.602% | 0.3153% | 0.9491% | 1.737%
| PDM3 |  | -27.3% | 0.292% | -4.834% | -0.6466% | 9.139% | -1.224% | 6.559% | -1.964% | -2.235% | -9.178% | 3.983% | 1.063%
| Z-Score (Absolute) |  | 6.03 | 0.28 | 0.89 | 0.07 | 2.3 | 0.07 | 1.71 | 0.24 | 0.3 | 1.88 | 1.12 | 0.46
| Analytical results for Sn in OREAS 993b (Certified Value < 100 ppm) | [13 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.14 | 2.16 | 2.19 | [5 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | [5 empty] | 0
0
| 1 | 1 | < 50 | < 50 | < 100 | < 100 | < 160 | 50 | < 100 | [5 empty] | 1
| 1 | 2 | < 50 | 50 | < 100 | < 100 | < 160 | 90 | < 100 | [5 empty] | 25
| 1 | 3 | < 50 | 120 | < 100 | < 100 | < 160 | < 50 | < 100 | [5 empty] | 16
| 1 | 4 | < 50 | 70 | < 100 | < 100 | < 160 | 80 | < 100 | [5 empty] | < 100
| 1 | 5 | < 50 | 70 | < 100 | < 100 | < 160 | < 50 | < 100 | [5 empty] | 29
| 1 | 6 | < 50 | 60 | < 100 | < 100 | < 160 | < 50 | < 100
| Mean |  |  | 74 |  |  |  | 73.3333
| Median |  |  | 70 |  |  |  | 80
| Std Dev. |  |  | 27.0185 |  |  |  | 20.8167
| Rel.Std.Dev. |  |  | 36.51% |  |  |  | 28.39%
PDM3
| Z-Score (Absolute) |  | 20.23 | 12.81 | 0 | 0 | 23.28 | 0.67 | 0
| Analytical results for Sr in OREAS 993b (Indicative Value 163 ppm) | [13 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.05 | 2.06 | 2.09 | 2.12 | 2.14 | 2.16 | 2.19 | [5 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | [5 empty] | 0
0
| 1 | 1 | 200 | 100 | 100 | 250 | 80.3 | 200 | 200 | [5 empty] | 1
| 1 | 2 | 200 | 200 | 100 | 250 | 135.3 | 300 | 100 | [5 empty] | 2
| 1 | 3 | 200 | 100 | 100 | 250 | 105.7 | 300 | 100 | [5 empty] | 16
| 1 | 4 | 200 | 200 | 100 | 250 | 63.4 | 200 | 100 | [5 empty] | 162.8724
| 1 | 5 | 200 | 200 | 100 | 170 | 97.2 | 200 | 100 | [5 empty] | 8
| 1 | 6 | 200 | 200 | 100 | 170 | 105.7 | 200 | 100
| Mean |  | 200 | 166.6667 | 100 | 223.3333 | 97.9333 | 233.3333 | 116.6667
| Median |  | 200 | 200 | 100 | 250 | 101.45 | 200 | 100
| Std Dev. |  | 0 | 51.6398 | 0 | 41.3118 | 24.5802 | 51.6398 | 40.8248
| Rel.Std.Dev. |  | 0% | 30.98% | 0% | 18.5% | 25.1% | 22.13% | 34.99%
| PDM3 |  | 22.8% | 2.33% | -38.6% | 37.12% | -39.87% | 43.26% | -28.37%
| Z-Score (Absolute) |  | 0.38 | 0 | 0.76 | 0.67 | 0.79 | 0.76 | 0.57
| Analytical results for Ta in OREAS 993b (Indicative Value 92 ppm) | [13 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.07 | 2.19 | [10 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | [10 empty] | 0
0
| 1 | 1 | 100 | < 100 | [10 empty] | 1
| 1 | 2 | < 100 | < 100 | [10 empty] | 3
| 1 | 3 | 100 | < 100 | [10 empty] | 16
| 1 | 4 | 100 | < 100 | [10 empty] | 91.6667
| 1 | 5 | 100 | < 100 | [10 empty] | 9
| 1 | 6 | 100 | < 100
| Mean |  | 100
| Median |  | 100
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 9.091%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Te | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 4
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 10
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 993b (Certified Value 0.202 wt.%) | [13 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.19 | 2.20 |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | PPP*XRF | OxBF*XRF |  | 3
3
| 1 | 1 | 0.193 | 0.21 | 0.19 | 0.19 | 0.1668 | 0.3 | 0.19 | 0.1615 | 0.22 | 0.2 | 0.21802 |  | 1
| 1 | 2 | 0.1901 | 0.21 | 0.2 | 0.2 | 0.1668 | 0.3 | 0.18 | 0.159 | 0.23 | 0.2 | 0.22002 |  | 0
| 1 | 3 | 0.1941 | 0.19 | 0.19 | 0.2 | 0.1501 | 0.3 | 0.19 | 0.1785 | 0.22 | 0.2 | 0.21802 |  | 16
| 1 | 4 | 0.1951 | 0.2 | 0.2 | 0.19 | 0.1668 | 0.3 | 0.19 | 0.174 | 0.21 | 0.167 | 0.21101 |  | 0.2018
| 1 | 5 | 0.1923 | 0.19 | 0.2 | 0.21 | 0.1668 | 0.3 | 0.21 | 0.166 | 0.22 | 0.2 | 0.22503 |  | 30
| 1 | 6 | 0.1908 | 0.21 | 0.2 | 0.19 | 0.1668 | 0.3 | 0.18 | 0.165 | 0.22 | 0.183 | 0.22803
| Mean |  | 0.19257 | 0.20167 | 0.19667 | 0.19667 | 0.16402 | 0.3 | 0.19 | 0.16733 | 0.22 | 0.19167 | 0.22002
| Median |  | 0.19265 | 0.205 | 0.2 | 0.195 | 0.1668 | 0.3 | 0.19 | 0.1655 | 0.22 | 0.2 | 0.21902
| Std Dev. |  | 0.0019096 | 0.0098319 | 0.005164 | 0.008165 | 0.0068177 | 0 | 0.010954 | 0.0074811 | 0.0063246 | 0.013866 | 0.0059739
| Rel.Std.Dev. |  | 0.9917% | 4.875% | 2.626% | 4.152% | 4.157% | 0% | 5.766% | 4.471% | 2.875% | 7.234% | 2.715%
| PDM3 |  | -4.577% | -0.06761% | -2.545% | -2.545% | -18.72% | 48.66% | -5.849% | -17.08% | 9.017% | -5.023% | 9.028%
| Z-Score (Absolute) |  | 0.24 | 0.81 | 0.24 | 0.24 | 3.54 | NA | 0.53 | 3.15 | 2.93 | 0.32 | 2.94
Results from laboratory 2.10 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Tl in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
Tl | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 5
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 11
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 993b (Indicative Value < 100 ppm) | [13 empty] | Indicative
U | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.19 | [11 empty] | ppm
|  |  | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | [11 empty] | 1
| 1 | 2 | < 100 | [11 empty] | 6
| 1 | 3 | < 100 | [11 empty] | 16
| 1 | 4 | < 100 | [11 empty] | < 100
| 1 | 5 | < 100 | [11 empty] | 12
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 993b (Indicative Value 82 ppm) | [13 empty] | Indicative
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.12 | 2.14 | 2.16 | 2.19 |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF |  |  |  | 0
0
| 1 | 1 | 62.7 | 100 | < 100 | 100 | < 35 | < 60 | 44.8 | 100 | < 100 |  |  |  | 1
| 1 | 2 | 62.2 | 100 | 100 | 100 | < 35 | < 60 | 61.6 | 100 | < 100 |  |  |  | 7
| 1 | 3 | 62.7 | 100 | < 100 | 100 | < 35 | < 60 | 56 | 100 | < 100 |  |  |  | 16
| 1 | 4 | 59.4 | 100 | 100 | 100 | < 35 | < 60 | 58.8 | 100 | < 100 |  |  |  | 81.5953
| 1 | 5 | 61.1 | < 100 | 100 | 100 | < 35 | < 60 | 42 | 100 | < 100 |  |  |  | 13
| 1 | 6 | 62.7 | 100 | 100 | 100 | < 35 | < 60 | 50.4 | 100 | < 100
| Mean |  | 61.8 | 100 | 100 | 100 |  |  | 52.2667 | 100
| Median |  | 62.45 | 100 | 100 | 100 |  |  | 53.2 | 100
| Std Dev. |  | 1.3297 | 0 | 0 | 0 |  |  | 7.8533 | 0
| Rel.Std.Dev. |  | 2.152% | 0% | 0% | 0% |  |  | 15.03% | 0%
| PDM3 |  | -24.26% | 22.56% | 22.56% | 22.56% |  |  | -35.94% | 22.56%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.49 | 0.86 | 1 | 0.76 | 0.22 | 0.86 | 0.27
| Analytical results for W in OREAS 993b (Certified Value < 10 ppm) | [13 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.05 | 2.06 | 2.07 | 2.09 | 2.16 | 2.19 | [6 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | [6 empty] | 1
1
| 1 | 1 | < 10 | 10 | < 80 | < 10 | 40 | < 100 | [6 empty] | 1
| 1 | 2 | < 10 | < 10 | < 80 | < 10 | 70 | < 100 | [6 empty] | 32
| 1 | 3 | < 10 | < 10 | < 80 | < 10 | 50 | < 100 | [6 empty] | 16
| 1 | 4 | < 10 | < 10 | < 80 | < 10 | < 10 | < 100 | [6 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 80 | < 10 | < 10 | < 100 | [6 empty] | 31
| 1 | 6 | < 10 | < 10 | < 80 | < 10 | 20 | < 100
| Mean |  |  | 10 |  |  | 45
| Median |  |  | 10 |  |  | 45
| Std Dev. | [5 empty] | 20.8167
| Rel.Std.Dev. | [5 empty] | 46.26%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.63 | 1.03 | 0.67 | 0.63 | 1.53
| Analytical results for Y2O3 in OREAS 993b (Indicative Value 152 ppm) | [13 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.05 | 2.06 | 2.16 | [9 empty] | ppm
|  |  | OxBF*XRF | OxBF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | < 50 | < 50 | 160 | [9 empty] | 1
| 1 | 2 | < 50 | < 50 | 180 | [9 empty] | 8
| 1 | 3 | < 50 | < 50 | 180 | [9 empty] | 16
| 1 | 4 | < 50 | < 50 | 100 | [9 empty] | 151.6667
| 1 | 5 | < 50 | < 50 | 150 | [9 empty] | 14
| 1 | 6 | < 50 | < 50 | 140
| Mean |  |  |  | 151.6667
| Median |  |  |  | 155
| Std Dev. |  |  |  | 29.9444
| Rel.Std.Dev. |  |  |  | 19.74%
| PDM3 |  |  |  | -2.22e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Zn in OREAS 993b (Certified Value 0.332 wt.%) | [13 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.14 | 2.16 | 2.19 |  |  |  | wt.%
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF |  |  |  | 3
3
| 1 | 1 | 0.33676 | 0.337 | 0.317 | 0.35 | 0.32 | 0.4 | 0.34554 | 0.333 | 0.3 |  |  |  | 1
| 1 | 2 | 0.33588 | 0.339 | 0.336 | 0.36 | 0.32 | 0.4 | 0.34831 | 0.346 | 0.33 |  |  |  | 0
| 1 | 3 | 0.33507 | 0.335 | 0.328 | 0.36 | 0.32 | 0.4 | 0.35026 | 0.341 | 0.29 |  |  |  | 16
| 1 | 4 | 0.33515 | 0.337 | 0.334 | 0.35 | 0.32 | 0.4 | 0.34914 | 0.315 | 0.31 |  |  |  | 0.33236
| 1 | 5 | 0.33636 | 0.337 | 0.331 | 0.35 | 0.31 | 0.4 | 0.34797 | 0.329 | 0.32 |  |  |  | 32
| 1 | 6 | 0.33499 | 0.337 | 0.334 | 0.35 | 0.31 | 0.4 | 0.35002 | 0.322 | 0.29
| Mean |  | 0.3357 | 0.337 | 0.33 | 0.35333 | 0.31667 | 0.4 | 0.34854 | 0.331 | 0.30667
| Median |  | 0.33552 | 0.337 | 0.3325 | 0.35 | 0.32 | 0.4 | 0.34873 | 0.331 | 0.305
| Std Dev. |  | 0.00074767 | 0.0012649 | 0.006957 | 0.005164 | 0.005164 | 6.0809e-17 | 0.0017258 | 0.011576 | 0.01633
| Rel.Std.Dev. |  | 0.2227% | 0.3753% | 2.108% | 1.462% | 1.631% | 1.52e-14% | 0.4951% | 3.497% | 5.325%
| PDM3 |  | 1.004% | 1.395% | -0.7112% | 6.309% | -4.723% | 20.35% | 4.867% | -0.4103% | -7.732%
| Z-Score (Absolute) |  | 0.17 | 0.26 | 0.24 | 1.43 | 1.19 | NA | 1.09 | 0.17 | 1.91
Results from laboratory 2.10 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Zr in OREAS 993b (Indicative Value 133 ppm) | [13 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.14 | 2.16 | 2.19 |  |  |  |  | ppm
|  |  | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 70 | 100 | 400 | 100 | 200 | < 370 | 200 | < 100 |  |  |  |  | 1
| 1 | 2 | < 70 | 100 | 400 | 100 | 200 | < 370 | 200 | < 100 |  |  |  |  | 9
| 1 | 3 | < 70 | 100 | 400 | 100 | 200 | < 370 | 200 | 100 |  |  |  |  | 16
| 1 | 4 | < 70 | 100 | 400 | 100 | 200 | < 370 | 200 | < 100 |  |  |  |  | 133.3333
| 1 | 5 | < 70 | 100 | 400 | 100 | 200 | < 370 | 200 | < 100 |  |  |  |  | 15
| 1 | 6 | < 70 | 100 | 500 | 100 | 100 | < 370 | 200 | 100
| Mean |  |  | 100 | 416.6667 | 100 | 183.3333 |  | 200 | 100
| Median |  |  | 100 | 400 | 100 | 200 |  | 200 | 100
| Std Dev. |  |  | 0 | 40.8248 | 0 | 40.8248 |  | 0 | 0
| Rel.Std.Dev. |  |  | 0% | 9.798% | 0% | 22.27% |  | 0% | 0%
| PDM3 |  |  | -25% | 212.5% | -25% | 37.5% |  | 50% | -25%
| Z-Score (Absolute) |  | 1.39 | 0.55 | 3.65 | 0.55 | 0.55 | 0.58 | 0.77 | 0.99
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C378: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C487: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C505: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C523: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C541: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C577: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C595: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C613: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C631: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C649: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C667: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C703: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C721: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C740: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C758: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C776: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C794: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C812: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C830: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C849: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 993b (Indicative Value 0.44 wt.%) | [19 empty] | Indicative
H2O- | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 3.01 | 3.02 | 3.03 | 3.04 | 3.05 | 3.06 | 3.07 | 3.08 | 3.09 | 3.10 | 3.11 | 3.12 | 3.13 | 3.14 | 3.15 | 3.16 | 3.17 | 6.01 | wt.%
|  |  | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*TGA | at 105°C*Oven | at 105°C*Oven | at 105°C*TGA | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | 3
3
| 1 | 1 | 0.44 | 0.73 | 0.62 | 0.57 | 0.2271 | 0.33 | 0.2 | 0.85 | 0.55 | 0.69 | 0.45 | 0.19 | 0.42 | 0.33 | 0.3003 | 0.41023 | 0.23 | 0.15 | 1
| 1 | 2 | 3.01: 0.44; 3.03: 0.62; 3.04: 0.54; 3.07: 0.2; 3.08: 1; 3.13: 0.42; 3.16: 0.69386; 6.01: 0.15 | 13
| 1 | 3 | 3.01: 0.45; 3.04: 0.56; 3.07: 0.22; 6.01: 0.2 | 16
| 1 | 4 | 6.01: 0.2 | 0.44034
| Mean |  | 0.44333 | 0.73 | 0.62 | 0.55667 | 0.2271 | 0.33 | 0.20667 | 0.925 | 0.55 | 0.69 | 0.45 | 0.19 | 0.42 | 0.33 | 0.3003 | 0.55204 | 0.23 | 0.175 | 19
| Median |  | 0.44 | 0.73 | 0.62 | 0.56 | 0.2271 | 0.33 | 0.2 | 0.925 | 0.55 | 0.69 | 0.45 | 0.19 | 0.42 | 0.33 | 0.3003 | 0.55204 | 0.23 | 0.175
| Std Dev. |  | 3.01: 0.0057735; 3.03: 0; 3.04: 0.015275; 3.07: 0.011547; 3.08: 0.10607; 3.13: 0; 3.16: 0.20056; 6.01: 0.028868
| Rel.Std.Dev. |  | 3.01: 1.302%; 3.03: 0%; 3.04: 2.744%; 3.07: 5.587%; 3.08: 11.47%; 3.13: 0%; 3.16: 36.33%; 6.01: 16.5%
| PDM3 |  | 0.68% | 65.78% | 40.8% | 26.42% | -48.43% | -25.06% | -53.07% | 110.1% | 24.9% | 56.7% | 2.194% | -56.85% | -4.619% | -25.06% | -31.8% | 25.37% | -47.77% | -60.26%
| Z-Score (Absolute) |  | 0.05 | 1.26 | 0.79 | 0.53 | 0.86 | 0.43 | 0.95 | 2.08 | 0.5 | 1.09 | 0.08 | 1.02 | 0.05 | 0.43 | 0.55 | 0.51 | 0.85 | 1.08
| Analytical results for LOI1000 in OREAS 993b (Certified Value 15.21 wt.%) | [19 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.18 | 2.19 | [6 empty] | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*Furnace | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | LOI*Furnace | LOI*TGA | LOI*TGA | [6 empty] | 2
3
| 1 | 1 | 15.3 | 15.418 | 15.37 | 15.11 | 14.99 | 15.14 | 15.04 | 15.77 | 15.0849 | 15.79 | 14.3 | 15.18 | [6 empty] | 1
| 1 | 2 | 15.29 | 15.399 | 15.52 | 15.36 | 14.85 | 15.11 | 14.99 | 15.6 | 14.9566 | 15.31 | 14.4 | 15.22 | [6 empty] | 0
| 1 | 3 | 15.15 | 15.314 | 15.21 | 15.37 | 14.84 | 15.13 | 15.17 | 16.12 | 15.0997 | 15.99 | 14.4 | 15.19 | [6 empty] | 16
| 1 | 4 | 15.14 | 15.279 | 15.41 | 15.38 | 15.39 | 15.11 | 15.04 | 15.72 | 15.0465 | 16.01 | 14.5 | 15.24 | [6 empty] | 15.2068
| 1 | 5 | 15.17 | 15.352 | 14.91 | 15.12 | 15.57 | 15.13 | 15.11 | 15.93 | 15.0373 | 15.29 | 14.8 | 15.19 | [6 empty] | 34
| 1 | 6 | 15.3 | 15.323 | 15.43 | 15.07 | 14.59 | 15.14 | 15.2 | 15.49 | 15.0607 | 15.38 | 14.4 | 15.21
| Mean |  | 15.225 | 15.3475 | 15.3083 | 15.235 | 15.0383 | 15.1267 | 15.0917 | 15.7717 | 15.0476 | 15.6283 | 14.4667 | 15.205
| Median |  | 15.23 | 15.3375 | 15.39 | 15.24 | 14.92 | 15.13 | 15.075 | 15.745 | 15.0536 | 15.585 | 14.4 | 15.2
| Std Dev. |  | 0.079183 | 0.053024 | 0.22004 | 0.14896 | 0.37 | 0.013663 | 0.082321 | 0.2271 | 0.050314 | 0.34061 | 0.17512 | 0.022583
| Rel.Std.Dev. |  | 0.5201% | 0.3455% | 1.437% | 0.9778% | 2.46% | 0.09032% | 0.5455% | 1.44% | 0.3344% | 2.179% | 1.211% | 0.1485%
| PDM3 |  | 0.1199% | 0.9255% | 0.6679% | 0.1857% | -1.108% | -0.5267% | -0.7569% | 3.715% | -1.047% | 2.772% | -4.867% | -0.01158%
| Z-Score (Absolute) |  | 0.05 | 0.7 | 0.49 | 0.11 | 0.93 | 0.47 | 0.65 | 2.93 | 0.88 | 2.18 | 3.94 | 0.05
Note on C15: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C33: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 993b (Indicative Value 0.156 wt.%) | [17 empty] | Indicative
C | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 2.12 | 2.19 | [14 empty] | wt.%
|  |  | IRC | IRC | [14 empty] | 3
3
| 1 | 1 | 0.17 | 0.14 | [14 empty] | 1
| 1 | 2 | 0.18 | 0.14 | [14 empty] | 15
| 1 | 3 | 0.15 | 0.15 | [14 empty] | 16
| 1 | 4 | 0.15 | 0.15 | [14 empty] | 0.15583
| 1 | 5 | 0.18 | 0.14 | [14 empty] | 21
| 1 | 6 | 0.18 | 0.14
| Mean |  | 0.16833 | 0.14333
| Median |  | 0.175 | 0.14
| Std Dev. |  | 0.01472 | 0.005164
| Rel.Std.Dev. |  | 8.744% | 3.603%
| PDM3 |  | 8.021% | -8.021%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 993b (Certified Value 29.67 wt.%) | [17 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 | 2.19 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 30.3 | 29.37 | 30 | 31.506 | 29 | 28.8 | 29.6 | 28.3 | 30.57 | 29.31 | 28.27 | 30.26 | 29.25 | 29.43 | 29.3 | 31.1 | 1
| 1 | 2 | 30.3 | 29.9 | 30.3 | 31.391 | 29.1 | 29.2 | 29.5 | 28.9 | 30.68 | 29.82 | 28.48 | 30.5 | 29.3907 | 29.78 | 29.5 | 30.8 | 34
| 1 | 3 | 30.5 | 29.41 | 30.6 | 31.666 | 29 | 29.4 | 29.4 | 29.4 | 30.96 | 29.82 | 28.47 | 30.51 | 29.4207 | 29.82 | 29.5 | 30.9 | 16
| 1 | 4 | 32 | 29.32 | 30.4 | 31.415 | 29 | 29.3 | 29.3 | 28.8 | 30.68 | 29.54 | 28.28 | 30.86 | 29.3 | 29.73 | 29.2 | 31 | 29.6733
| 1 | 5 | 30.2 | 29.78 | 30.4 | 31.544 | 29.2 | 29.4 | 29.9 | 29.3 | 30.74 | 29.33 | 28.49 | 30.5 | 29.3971 | 29.44 | 29.3 | 30.1 | 36
| 1 | 6 | 29.3 | 29.38 | 30.5 | 31.602 | 29.2 | 29.6 | 29.1 | 29.2 | 30.79 | 28.83 | 28.38 | 30.55 | 29.415 | 29.72 | 29.1 | 31.1
| Mean |  | 30.4333 | 29.5267 | 30.3667 | 31.5207 | 29.0833 | 29.2833 | 29.4667 | 28.9833 | 30.7367 | 29.4417 | 28.395 | 30.53 | 29.3622 | 29.6533 | 29.3167 | 30.8333
| Median |  | 30.3 | 29.395 | 30.4 | 31.525 | 29.05 | 29.35 | 29.45 | 29.05 | 30.71 | 29.435 | 28.425 | 30.505 | 29.3939 | 29.725 | 29.3 | 30.95
| Std Dev. |  | 0.8756 | 0.24736 | 0.20656 | 0.10628 | 0.098319 | 0.27142 | 0.27325 | 0.40702 | 0.13186 | 0.37413 | 0.10095 | 0.19204 | 0.07028 | 0.17294 | 0.16021 | 0.37771
| Rel.Std.Dev. |  | 2.877% | 0.8377% | 0.6802% | 0.3372% | 0.3381% | 0.9269% | 0.9273% | 1.404% | 0.429% | 1.271% | 0.3555% | 0.629% | 0.2394% | 0.5832% | 0.5465% | 1.225%
| PDM3 |  | 2.561% | -0.494% | 2.337% | 6.226% | -1.988% | -1.314% | -0.6962% | -2.325% | 3.584% | -0.7805% | -4.308% | 2.887% | -1.048% | -0.06716% | -1.202% | 3.909%
| Z-Score (Absolute) |  | 1.36 | 0.04 | 1.27 | 2.95 | 0.6 | 0.31 | 0.04 | 0.75 | 1.8 | 0.08 | 1.6 | 1.5 | 0.2 | 0.23 | 0.26 | 1.95
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Hg in OREAS 993b (Certified Value 1.07 ppm) | [17 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*CVAAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*CVAAS | AR*MS | 2
|  |  | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.1g | 0.5g | 10g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 40g | 3
| 1 | 1 | 1.16 | 0.969 | 1.05 | 1.085 | 0.983 | 1.135 | 1.07 | 1.067 | 0.89 | 0.93 | 1.0839 | 1.1989 | 1.055 | 1.09 | 1.047 | 1 | 1
| 1 | 2 | 1.04 | 0.975 | 1.02 | 1.09 | 1.08 | 1.095 | 1.01 | 1.073 | 0.97 | 1.006 | 1.0864 | 1.203 | 1.05 | 1.2 | 1.053 | 1 | 35
| 1 | 3 | 1.09 | 1.014 | 1.04 | 1.11 | 1.085 | 1.09 | 1.03 | 1.076 | 0.98 | 1.137 | 1.1063 | 1.198 | 1.06 | 1.21 | 1.047 | 1 | 16
| 1 | 4 | 1.18 | 0.956 | 1.09 | 1.11 | 1.06 | 1.105 | 0.97 | 1.076 | 0.93 | 1.157 | 1.1001 | 1.1923 | 1.07 | 1.13 | 1.043 | 1 | 1.0665
| 1 | 5 | 1.14 | 1.031 | 1.05 | 1.085 | 1.08 | 1.12 | 0.97 | 1.067 | 0.99 | 1.078 | 1.0683 | 1.2026 | 1.095 | 1.085 | 1.053 | 1 | 38
| 1 | 6 | 1.19 | 0.975 | 1.07 | 1.13 | 1.075 | 1.1 | 1 | 1.072 | 1.01 | 0.95 | 1.0578 | 1.2009 | 1.06 | 1.12 | 1.047 | 1
| Mean |  | 1.1333 | 0.98667 | 1.0533 | 1.1017 | 1.0605 | 1.1075 | 1.0083 | 1.0718 | 0.96167 | 1.043 | 1.0838 | 1.1993 | 1.065 | 1.1392 | 1.0483 | 1
| Median |  | 1.15 | 0.975 | 1.05 | 1.1 | 1.0775 | 1.1025 | 1.005 | 1.0725 | 0.975 | 1.042 | 1.0852 | 1.1999 | 1.06 | 1.125 | 1.047 | 1
| Std Dev. |  | 0.05785 | 0.029111 | 0.024221 | 0.018074 | 0.038929 | 0.016956 | 0.038166 | 0.0040702 | 0.044008 | 0.095754 | 0.01842 | 0.0039526 | 0.016125 | 0.05389 | 0.0039328 | 0
| Rel.Std.Dev. |  | 5.104% | 2.95% | 2.299% | 1.641% | 3.671% | 1.531% | 3.785% | 0.3797% | 4.576% | 9.181% | 1.7% | 0.3296% | 1.514% | 4.731% | 0.3751% | 0%
| PDM3 |  | 6.27% | -7.482% | -1.231% | 3.301% | -0.5593% | 3.848% | -5.451% | 0.5034% | -9.827% | -2.2% | 1.627% | 12.45% | -0.1373% | 6.817% | -1.7% | -6.232%
| Z-Score (Absolute) |  | 1.14 | 1.23 | 0.15 | 0.63 | 0.04 | 0.72 | 0.88 | 0.15 | 1.63 | 0.32 | 0.34 | 2.2 | 0.04 | 1.23 | 0.23 | 1.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## ISE
| Analytical results for Cl in OREAS 993b (Indicative Value 98 ppm) | [14 empty] | Indicative
Cl | Round | Replicate | Lab | [12 empty] | 1
| No | No | 5.01 | [12 empty] | ppm
|  |  | CIC | [12 empty] | 0
0
| 1 | 1 | 100 | [12 empty] | 1
| 1 | 2 | 100 | [12 empty] | 6
| 1 | 3 | 120 | [12 empty] | 16
| 1 | 4 | 80 | [12 empty] | 98
| 1 | 5 | 90 | [12 empty] | 23
| Mean |  | 98
| Median |  | 100
| Std Dev. |  | 14.8324
| Rel.Std.Dev. |  | 15.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for F in OREAS 993b (Certified Value 657 ppm) | [14 empty] | Certified
F | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.16 | 2.17 | 2.18 | 5.01 | ppm
|  |  | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | ISE | 0
0
| 1 | 1 | 644 | 669 | 870 | 680 | 650 | 690 | 649 | 660 | 590 | 670 | 689 | 630 | 620 | 1
| 1 | 2 | 659 | 592 | 841 | 690 | 640 | 630 | 645 | 660 | 584 | 680 | 665 | 680 | 640 | 0
| 1 | 3 | 655 | 703 | 899 | 690 | 620 | 610 | 639 | 649 | 604 | 730 | 685 | 670 | 620 | 16
| 1 | 4 | 648 | 600 | 904 | 660 | 640 | 710 | 645 | 681 | 569 | 700 | 678 | 660 | 620 | 657.2879
| 1 | 5 | 656 | 605 | 849 | 690 | 640 | 660 | 661 | 653 | 542 | 670 | 670 | 680 | 620 | 40
| 1 | 6 | 635 | 658 | 899 | 680 | 630 | 690 | 657 | 668 | 569 | 630 | 689 | 670
| Mean |  | 649.5 | 637.8333 | 877 | 681.6667 | 636.6667 | 665 | 649.3333 | 661.8333 | 576.3333 | 680 | 679.3333 | 665 | 624
| Median |  | 651.5 | 631.5 | 884.5 | 685 | 640 | 675 | 647 | 660 | 576.5 | 675 | 681.5 | 670 | 620
| Std Dev. |  | 9.0056 | 45.2434 | 27.6622 | 11.6905 | 10.328 | 38.8587 | 8.2381 | 11.4441 | 21.4352 | 33.4664 | 10.1325 | 18.7083 | 8.9443
| Rel.Std.Dev. |  | 1.387% | 7.093% | 3.154% | 1.715% | 1.622% | 5.843% | 1.269% | 1.729% | 3.719% | 4.922% | 1.492% | 2.813% | 1.433%
| PDM3 |  | -1.185% | -2.96% | 33.43% | 3.709% | -3.137% | 1.173% | -1.21% | 0.6915% | -12.32% | 3.455% | 3.354% | 1.173% | -5.064%
| Z-Score (Absolute) |  | 0.46 | 0.89 | 7.99 | 0.74 | 0.93 | 0.12 | 0.46 | 0 | 3.17 | 0.67 | 0.65 | 0.12 | 1.4
Note on C16: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C34: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 993b (Certified Value 40.4 ppm) | [18 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 40.637 | 39.9 | 41.1 | 41.89 | 41.5 | 38.8 | 41 | 35.9 | 39 | 39.657 | 41.8 | 41.5 | 41.124 | 38.1298 | 41 | 40.8 | 40.8 | 1
| 1 | 2 | 41.412 | 39.7 | 41 | 42.22 | 41.3 | 38.4 | 41 | 36.4 | 38.9 | 39.176 | 42.3 | 41.3 | 40.8429 | 38.8805 | 40 | 37.6 | 40.4 | 0
| 1 | 3 | 39.78 | 40.4 | 41.3 | 42.13 | 43.3 | 39.4 | 41 | 35.8 | 38.9 | 38.856 | 41.1 | 41.3 | 41.0717 | 38.1041 | 41 | 38.6 | 40.2 | 16
| 1 | 4 | 41.045 | 40.7 | 41.6 | 41.93 | 40.2 | 38.7 | 41 | 36.3 | 38.6 | 40.206 | 42.1 | 41.6 | 41.1028 | 38.3778 | 41 | 38.4 | 39.6 | 40.357
| 1 | 5 | 39.739 | 40.6 | 42 | 42.08 | 41 | 38.1 | 40 | 37 | 38.6 | 39.214 | 42.4 | 40.9 | 40.8216 | 38.1071 | 41 | 39.1 | 39.6 | 42
| 1 | 6 | 39.494 | 40.3 | 41.6 | 41.76 | 41.6 | 39 | 42 | 36.1 | 38.6 | 40.395 | 41 | 40.3 | 40.7686 | 38.5044 | 40 | 39.9
| Mean |  | 40.3512 | 40.2667 | 41.4333 | 42.0017 | 41.4833 | 38.7333 | 41 | 36.25 | 38.7667 | 39.584 | 41.7833 | 41.15 | 40.9553 | 38.3506 | 40.6667 | 39.0667 | 40.12
| Median |  | 40.2085 | 40.35 | 41.45 | 42.005 | 41.4 | 38.75 | 41 | 36.2 | 38.75 | 39.4355 | 41.95 | 41.3 | 40.9573 | 38.2538 | 41 | 38.85 | 40.2
| Std Dev. |  | 0.79047 | 0.39328 | 0.37238 | 0.17105 | 1.0226 | 0.45461 | 0.63246 | 0.43243 | 0.18619 | 0.61372 | 0.6047 | 0.48062 | 0.16073 | 0.30787 | 0.5164 | 1.1413 | 0.52154
| Rel.Std.Dev. |  | 1.959% | 0.9767% | 0.8987% | 0.4072% | 2.465% | 1.174% | 1.543% | 1.193% | 0.4803% | 1.55% | 1.447% | 1.168% | 0.3925% | 0.8028% | 1.27% | 2.922% | 1.3%
| PDM3 |  | -0.01456% | -0.2239% | 2.667% | 4.075% | 2.791% | -4.023% | 1.593% | -10.18% | -3.941% | -1.916% | 3.534% | 1.965% | 1.482% | -4.972% | 0.7672% | -3.197% | -0.5874%
| Z-Score (Absolute) |  | 0 | 0.05 | 0.67 | 1.03 | 0.71 | 1.01 | 0.4 | 2.56 | 0.99 | 0.48 | 0.89 | 0.5 | 0.38 | 1.25 | 0.2 | 0.8 | 0.14
| Analytical results for Al in OREAS 993b (Indicative Value 1.86 wt.%) | [18 empty] | Indicative
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 2
2
| 1 | 1 | 1.91 | 1.86 | 1.81 | 1.9875 | 1.86 | 1.84 | 1.59 | 1.66 | 1.61 | 1.84 | 2.19 | 1.969 | 1.882 | 1.8569 | 1.32 | 1.86 | 1.88 | 1
| 1 | 2 | 1.96 | 1.87 | 1.81 | 1.9775 | 1.9 | 1.83 | 1.57 | 1.54 | 1.64 | 1.83 | 2.19 | 1.9642 | 1.8786 | 1.8343 | 1.26 | 1.78 | 1.81 | 19
| 1 | 3 | 1.87 | 1.86 | 1.8 | 1.9846 | 1.9 | 1.88 | 1.8 | 1.5 | 1.63 | 1.8 | 2.2 | 1.9758 | 1.8636 | 1.8565 | 1.35 | 1.84 | 1.83 | 16
| 1 | 4 | 1.95 | 1.88 | 1.82 | 1.9816 | 1.84 | 1.84 | 1.61 | 1.69 | 1.63 | 1.86 | 2.2 | 1.9702 | 1.8821 | 1.8671 | 1.4 | 1.78 | 1.83 | 1.8617
| 1 | 5 | 1.86 | 1.88 | 1.8 | 1.9762 | 1.86 | 1.82 | 1.67 | 1.54 | 1.63 | 1.82 | 2.19 | 1.9279 | 1.8499 | 1.86 | 1.4 | 1.88 | 1.82 | 25
| 1 | 6 | 1.85 | 1.88 | 1.8 | 1.984 | 1.89 | 1.85 | 1.73 | 1.69 | 1.61 | 1.91 | 2.18 | 1.9137 | 1.8521 | 1.8741 | 1.26 | 1.81
| Mean |  | 1.9 | 1.8717 | 1.8067 | 1.9819 | 1.875 | 1.8433 | 1.6617 | 1.6033 | 1.625 | 1.8433 | 2.1917 | 1.9535 | 1.8681 | 1.8582 | 1.3317 | 1.825 | 1.834
| Median |  | 1.89 | 1.875 | 1.805 | 1.9828 | 1.875 | 1.84 | 1.64 | 1.6 | 1.63 | 1.835 | 2.19 | 1.9666 | 1.8711 | 1.8585 | 1.335 | 1.825 | 1.83
| Std Dev. |  | 0.047329 | 0.0098319 | 0.008165 | 0.0043588 | 0.0251 | 0.020656 | 0.089536 | 0.085946 | 0.012247 | 0.038297 | 0.0075277 | 0.025963 | 0.014877 | 0.013504 | 0.063377 | 0.041833 | 0.027019
| Rel.Std.Dev. |  | 2.491% | 0.5253% | 0.4519% | 0.2199% | 1.339% | 1.121% | 5.388% | 5.36% | 0.7537% | 2.078% | 0.3435% | 1.329% | 0.7964% | 0.7267% | 4.759% | 2.292% | 1.473%
| PDM3 |  | 2.057% | 0.5354% | -2.956% | 6.457% | 0.7145% | -0.9865% | -10.74% | -13.88% | -12.71% | -0.9865% | 17.72% | 4.929% | 0.3414% | -0.1898% | -28.47% | -1.971% | -1.488%
| Z-Score (Absolute) |  | 1.04 | 0.52 | 0.67 | 2.55 | 0.58 | 0 | 3.34 | 4.41 | 4.02 | 0 | 6.41 | 2.03 | 0.45 | 0.27 | 9.41 | 0.34 | 0.17
| Analytical results for As in OREAS 993b (Certified Value 103 ppm) | [18 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 106.4 | 114 | 109 | 88.977 | 105 | 105 | 110 | 82.7 | 107 | 24.3 | 34.1 | 152 | 115.7972 | 95.9628 | 120 | 111 | 90 | 1
| 1 | 2 | 108.4 | 115 | 109 | 83.921 | 104 | 107 | 110 | 92.3 | 107 | 22.4 | 36 | 130.2 | 113.5182 | 94.5849 | 110 | 103.5 | 90 | 1
| 1 | 3 | 104.8 | 115 | 106 | 90.277 | 114 | 105 | 120 | 95.8 | 106 | 24.1 | 31.4 | 170.2 | 115.5075 | 95.2682 | 120 | 104.5 | 92 | 16
| 1 | 4 | 107.8 | 117 | 112 | 87.97 | 102.5 | 105 | 110 | 87.1 | 107 | 22.2 | 32.2 | 120.2 | 113.5371 | 98.7317 | 120 | 105.5 | 86 | 102.5424
| 1 | 5 | 103.4 | 115 | 106 | 88.847 | 104 | 103 | 120 | 89.3 | 106 | 23.2 | 31.5 | 152.9 | 114.1091 | 97.5061 | 120 | 106 | 88 | 43
| 1 | 6 | 104.2 | 116 | 106 | 91.166 | 105 | 104 | 120 | 82.7 | 107 | 23.8 | 36.5 | 115.9 | 114.5276 | 97.2434 | 110 | 111.5
| Mean |  | 105.8333 | 115.3333 | 108 | 88.5263 | 105.75 | 104.8333 | 115 | 88.3167 | 106.6667 | 23.3333 | 33.6167 | 140.2333 | 114.4994 | 96.5495 | 116.6667 | 107 | 89.2
| Median |  | 105.6 | 115 | 107.5 | 88.912 | 104.5 | 105 | 115 | 88.2 | 107 | 23.5 | 33.15 | 141.1 | 114.3184 | 96.6031 | 120 | 105.75 | 90
| Std Dev. |  | 2.0216 | 1.0328 | 2.4495 | 2.5242 | 4.1443 | 1.3292 | 5.4772 | 5.2423 | 0.5164 | 0.88468 | 2.264 | 21.4057 | 0.97388 | 1.5495 | 5.164 | 3.4059 | 2.2804
| Rel.Std.Dev. |  | 1.91% | 0.8955% | 2.268% | 2.851% | 3.919% | 1.268% | 4.763% | 5.936% | 0.4841% | 3.792% | 6.735% | 15.26% | 0.8506% | 1.605% | 4.426% | 3.183% | 2.556%
| PDM3 |  | 3.209% | 12.47% | 5.322% | -13.67% | 3.128% | 2.234% | 12.15% | -13.87% | 4.022% | -77.25% | -67.22% | 36.76% | 11.66% | -5.844% | 13.77% | 4.347% | -13.01%
| Z-Score (Absolute) |  | 0.01 | 0.7 | 0.16 | 1.26 | 0 | 0.07 | NA | 1.28 | 0.07 | 6.04 | 5.29 | 2.53 | 0.64 | 0.67 | NA | 0.09 | 1.21
Results from laboratories 2.07 and 2.15 were removed due to their 10 ppm reading resolution.
| Analytical results for B in OREAS 993b (Indicative Value 92 ppm) | [18 empty] | Indicative
B | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.09 | ppm
|  |  | 2.01: 4A*OES | 0
0
| 1 | 1 | 2.01: 90 | 1
| 1 | 2 | 2.01: 93 | 20
| 1 | 3 | 2.01: 93 | 16
| 1 | 4 | 2.01: 94 | 91.5
| 1 | 5 | 2.01: 90 | 26
| 1 | 6 | 2.01: 89
| Mean |  | 2.01: 91.5
| Median |  | 2.01: 91.5
| Std Dev. |  | 2.01: 2.0736
| Rel.Std.Dev. |  | 2.01: 2.266%
PDM3 | bar: 2.01
Z-Score (Absolute) | bar: 2.01
| Analytical results for Ba in OREAS 993b (Certified Value 96 ppm) | [18 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | 5.01 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 23 | 105 | 106 | 83.8376 | 90 | 30 | 8 | 93 | 68 | 126 | 112.4901 | 78.9755 | 110 | 93 |  |  |  | 1
| 1 | 2 | 15 | 106 | 105 | 84.2444 | 100 | 20 | 17 | 93 | 44 | 129 | 112.4239 | 80.8099 | 100 | 92 |  |  |  | 2
| 1 | 3 | 17 | 107 | 106 | 83.6192 | 100 | 40 | 48 | 94 | 48 | 129 | 110.1673 | 74.2905 | 100 | 95.5 |  |  |  | 16
| 1 | 4 | 22 | 106 | 104 | 83.5452 | 110 | 20 | 7 | 95 | 85 | 130 | 109.2487 | 78.9138 | 90 | 97 |  |  |  | 95.5018
| 1 | 5 | 19 | 107 | 106 | 83.4288 | 100 | 10 | 42 | 92 | 56 | 128 | 110.1258 | 75.3479 | 100 | 97.5 |  |  |  | 44
| 1 | 6 | 14 | 107 | 106 | 83.1404 | 90 | 20 | 6 | 93 | 48 | 129 | 110.8566 | 80.632 | 100
| Mean |  | 18.3333 | 106.3333 | 105.5 | 83.6359 | 98.3333 | 23.3333 | 21.3333 | 93.3333 | 58.1667 | 128.5 | 110.8854 | 78.1616 | 100 | 95
| Median |  | 18 | 106.5 | 106 | 83.5822 | 100 | 20 | 12.5 | 93 | 52 | 129 | 110.512 | 78.9446 | 100 | 95.5
| Std Dev. |  | 3.6697 | 0.8165 | 0.83666 | 0.37626 | 7.5277 | 10.328 | 18.8432 | 1.0328 | 15.6769 | 1.3784 | 1.3202 | 2.7293 | 6.3246 | 2.4238
| Rel.Std.Dev. |  | 20.02% | 0.7679% | 0.793% | 0.4499% | 7.655% | 44.26% | 88.33% | 1.107% | 26.95% | 1.073% | 1.191% | 3.492% | 6.325% | 2.551%
| PDM3 |  | -80.8% | 11.34% | 10.47% | -12.42% | 2.965% | -75.57% | -77.66% | -2.271% | -39.09% | 34.55% | 16.11% | -18.16% | 4.71% | -0.5254%
| Z-Score (Absolute) |  | 2.35 | 0.6 | 0.57 | 0.16 | NA | 2.18 | 2.25 | 0.16 | 1.02 | 1.34 | 0.75 | 0.35 | NA | 0.22
Results from laboratories 2.05 and 2.16 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 993b (Certified Value < 0.5 ppm) | [18 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.03 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 1 | < 0.5 | 0.19 | < 0.5 | 0.2 | < 5 | < 1 | < 1 | < 0.5 | < 0.5 | 0.21 | [6 empty] | 1
| 1 | 2 | < 1 | < 0.5 | 0.19 | < 0.5 | 0.3 | < 5 | < 1 | < 1 | < 0.5 | < 0.5 | 0.21 | [6 empty] | 25
| 1 | 3 | < 1 | < 0.5 | 0.19 | < 0.5 | 0.2 | < 5 | < 1 | < 1 | < 0.5 | < 0.5 | 0.21 | [6 empty] | 16
| 1 | 4 | < 1 | < 0.5 | 0.19 | < 0.5 | 0.2 | < 5 | < 1 | < 1 | < 0.5 | < 0.5 | 0.2 | [6 empty] | < 0.5
| 1 | 5 | < 1 | < 0.5 | 0.18 | < 0.5 | 0.2 | < 5 | < 1 | < 1 | < 0.5 | < 0.5 | 0.21 | [6 empty] | 45
| 1 | 6 | < 1 | < 0.5 | 0.19 | < 0.5 | 0.3 | < 5 | < 1 | < 1 | < 0.5 | < 0.5 | 0.2
| Mean |  | 2.03: 0.18833; 2.05: 0.23333; 2.11: 0.20667
| Median |  | 2.03: 0.19; 2.05: 0.2; 2.11: 0.21
| Std Dev. |  | 2.03: 0.0040825; 2.05: 0.05164; 2.11: 0.005164
| Rel.Std.Dev. |  | 2.03: 2.168%; 2.05: 22.13%; 2.11: 2.499%
PDM3
| Z-Score (Absolute) |  | 3.89 | 0 | 0.96 | 0 | 0.26 | 35.01 | 3.89 | 3.89 | 0 | 0 | 0.67
| Analytical results for Bi in OREAS 993b (Certified Value 22.9 ppm) | [18 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 22.89 | 23.6 | 23.5 | 22.643 | 24.8 | 31 | 40 | 21.7 | 11 | 25.33 | 23.9 | 24.72 | 21.2329 | 21.2221 | 60 | 23.7 |  | 1
| 1 | 2 | 20.35 | 24.7 | 23.6 | 22.386 | 23.9 | < 2 | < 10 | 22.5 | 10 | 24.61 | 23.8 | 22.77 | 21.4854 | 21.0025 | 10 | 22.1 |  | 3
| 1 | 3 | 20.12 | 23.7 | 23.5 | 22.385 | 25.2 | 22 | 10 | 22.8 | 10 | 24.36 | 23.6 | 24.14 | 21.2558 | 20.9985 | 20 | 23.6 |  | 16
| 1 | 4 | 20.24 | 24.1 | 23.6 | 22.588 | 23.9 | 39 | < 10 | 21.9 | 10 | 26.27 | 23.5 | 23.74 | 21.4955 | 21.1007 | 50 | 23.3 |  | 22.8664
| 1 | 5 | 19.64 | 24 | 23.3 | 22.418 | 24.3 | 9 | < 10 | 22.1 | 10 | 24.91 | 23.4 | 22.93 | 21.2225 | 21.1174 | 40 | 23.6 |  | 46
| 1 | 6 | 21.7 | 23.8 | 23.7 | 22.331 | 23.9 | 15 | 10 | 22 | 9 | 25.01 | 21.5 | 21.46 | 21.0381 | 21.4667 | 50 | 22.9
| Mean |  | 20.8233 | 23.9833 | 23.5333 | 22.4585 | 24.3333 | 23.2 | 20 | 22.1667 | 10 | 25.0817 | 23.2833 | 23.2933 | 21.2884 | 21.1513 | 38.3333 | 23.2
| Median |  | 20.295 | 23.9 | 23.55 | 22.402 | 24.1 | 22 | 10 | 22.05 | 10 | 24.96 | 23.55 | 23.335 | 21.2444 | 21.1091 | 45 | 23.45
| Std Dev. |  | 1.2247 | 0.39707 | 0.13663 | 0.12599 | 0.55377 | 12.0499 | 17.3205 | 0.40825 | 0.63246 | 0.67107 | 0.89312 | 1.1592 | 0.17483 | 0.1753 | 19.4079 | 0.61319
| Rel.Std.Dev. |  | 5.881% | 1.656% | 0.5806% | 0.561% | 2.276% | 51.94% | 86.6% | 1.842% | 6.325% | 2.676% | 3.836% | 4.977% | 0.8212% | 0.8288% | 50.63% | 2.643%
| PDM3 |  | -8.935% | 4.885% | 2.917% | -1.784% | 6.415% | 1.459% | -12.54% | -3.06% | -56.27% | 9.688% | 1.823% | 1.867% | -6.901% | -7.5% | 67.64% | 1.459%
| Z-Score (Absolute) |  | 1.02 | 0.59 | 0.36 | 0.19 | 0.76 | 1.69 | NA | 0.34 | 6.51 | 1.14 | 0.23 | 0.24 | 0.78 | 0.85 | NA | 0.19
Results from laboratories 2.07 and 2.15 were removed due to their 10 ppm reading resolution.
| Analytical results for Ca in OREAS 993b (Certified Value 0.827 wt.%) | [18 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.86 | 0.85 | 0.814 | 0.83606 | 0.83 | 0.81 | 0.86 | 0.58 | 0.79 | 0.79 | 0.89 | 0.8258 | 0.83186 | 0.85909 | 0.82 | 0.82 | 0.84 | 1
| 1 | 2 | 0.88 | 0.86 | 0.807 | 0.83565 | 0.84 | 0.8 | 0.86 | 0.69 | 0.81 | 0.78 | 0.88 | 0.8154 | 0.82826 | 0.83671 | 0.82 | 0.8 | 0.82 | 27
| 1 | 3 | 0.85 | 0.85 | 0.806 | 0.83235 | 0.85 | 0.82 | 0.86 | 0.69 | 0.8 | 0.77 | 0.89 | 0.8283 | 0.81897 | 0.8605 | 0.7 | 0.81 | 0.83 | 16
| 1 | 4 | 0.88 | 0.86 | 0.814 | 0.83866 | 0.81 | 0.8 | 0.86 | 0.72 | 0.81 | 0.8 | 0.91 | 0.8261 | 0.83481 | 0.86423 | 0.82 | 0.79 | 0.83 | 0.82671
| 1 | 5 | 0.84 | 0.86 | 0.802 | 0.83423 | 0.83 | 0.8 | 0.86 | 0.65 | 0.8 | 0.78 | 0.88 | 0.808 | 0.81198 | 0.84371 | 0.82 | 0.83 | 0.83 | 47
| 1 | 6 | 0.84 | 0.86 | 0.808 | 0.83556 | 0.84 | 0.81 | 0.86 | 0.66 | 0.79 | 0.82 | 0.89 | 0.8024 | 0.81663 | 0.85976 | 0.81 | 0.8
| Mean |  | 0.85833 | 0.85667 | 0.8085 | 0.83542 | 0.83333 | 0.80667 | 0.86 | 0.665 | 0.8 | 0.79 | 0.89 | 0.81767 | 0.82375 | 0.854 | 0.79833 | 0.80833 | 0.83
| Median |  | 0.855 | 0.86 | 0.8075 | 0.8356 | 0.835 | 0.805 | 0.86 | 0.675 | 0.8 | 0.785 | 0.89 | 0.8206 | 0.82362 | 0.85942 | 0.82 | 0.805 | 0.83
| Std Dev. |  | 0.018348 | 0.005164 | 0.0047223 | 0.0020894 | 0.013663 | 0.008165 | 0 | 0.048477 | 0.0089443 | 0.017889 | 0.010954 | 0.010789 | 0.0091719 | 0.01105 | 0.048339 | 0.01472 | 0.0070711
| Rel.Std.Dev. |  | 2.138% | 0.6028% | 0.5841% | 0.2501% | 1.64% | 1.012% | 0% | 7.29% | 1.118% | 2.264% | 1.231% | 1.319% | 1.113% | 1.294% | 6.055% | 1.821% | 0.8519%
| PDM3 |  | 3.825% | 3.623% | -2.203% | 1.053% | 0.801% | -2.425% | 4.027% | -19.56% | -3.231% | -4.441% | 7.655% | -1.094% | -0.3581% | 3.301% | -3.433% | -2.223% | 0.3978%
| Z-Score (Absolute) |  | 0.98 | 0.93 | 0.43 | 0.33 | 0.27 | 0.49 | 1.03 | 4.51 | 0.67 | 0.96 | 1.88 | 0.17 | 0 | 0.86 | 0.72 | 0.44 | 0.18
| Analytical results for Cd in OREAS 993b (Certified Value 15 ppm) | [18 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 16.03 | 16.5 | 15.9 | 14.6516 | 15.8 | 14.2 | 20 | 13 | 14 | 15.1 | 10.1 | 15.8 | 16.0133 | 14.8055 | 20 | 16.4 |  | 1
| 1 | 2 | 15.41 | 16.1 | 15.9 | 14.0108 | 15.1 | 13.8 | 20 | 12.7 | 14.1 | 15.15 | 10.3 | 14.04 | 15.405 | 14.4225 | 20 | 15.95 |  | 4
| 1 | 3 | 15.19 | 17.3 | 15.8 | 14.134 | 15.75 | 13.8 | 20 | 12.4 | 14 | 14.92 | 10 | 14.96 | 16.0373 | 14.388 | 20 | 15.8 |  | 16
| 1 | 4 | 15.44 | 17.4 | 16.1 | 14.1132 | 15.35 | 13.6 | 20 | 13.7 | 13.9 | 15.19 | 10.3 | 15.07 | 15.7059 | 14.6269 | 20 | 15.75 |  | 14.9714
| 1 | 5 | 14.01 | 18 | 15.8 | 13.9212 | 15.25 | 14.4 | 20 | 12.8 | 14 | 14.6 | 10.2 | 14.36 | 15.6895 | 14.5474 | 20 | 15.65 |  | 48
| 1 | 6 | 16.48 | 15.1 | 16.1 | 14.2964 | 15.75 | 14.1 | 20 | 13.3 | 13.6 | 15.61 | 9.8 | 13.43 | 15.5325 | 14.5773 | 20 | 16.15
| Mean |  | 15.4267 | 16.7333 | 15.9333 | 14.1879 | 15.5 | 13.9833 | 20 | 12.9833 | 13.9333 | 15.095 | 10.1167 | 14.61 | 15.7306 | 14.5613 | 20 | 15.95
| Median |  | 15.425 | 16.9 | 15.9 | 14.1236 | 15.55 | 13.95 | 20 | 12.9 | 14 | 15.125 | 10.15 | 14.66 | 15.6977 | 14.5624 | 20 | 15.875
| Std Dev. |  | 0.84106 | 1.0482 | 0.13663 | 0.25991 | 0.30332 | 0.29944 | 0 | 0.46224 | 0.17512 | 0.33255 | 0.19408 | 0.84024 | 0.2536 | 0.15078 | 0 | 0.28107
| Rel.Std.Dev. |  | 5.452% | 6.264% | 0.8575% | 1.832% | 1.957% | 2.141% | 0% | 3.56% | 1.257% | 2.203% | 1.918% | 5.751% | 1.612% | 1.036% | 0% | 1.762%
| PDM3 |  | 3.041% | 11.77% | 6.425% | -5.233% | 3.531% | -6.6% | 33.59% | -13.28% | -6.934% | 0.8257% | -32.43% | -2.414% | 5.071% | -2.739% | 33.59% | 6.537%
| Z-Score (Absolute) |  | 0.44 | 1.45 | 0.83 | 0.51 | 0.5 | 0.67 | NA | 1.44 | 0.71 | 0.19 | 3.66 | 0.19 | 0.68 | 0.22 | NA | 0.85
| Analytical results for Ce in OREAS 993b (Certified Value 16.8 ppm) | [18 empty] | Certified
Ce | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.05; 2.05: 2.08; 2.06: 2.10; 2.07: 2.11; 2.08: 2.14; 2.09: 2.16 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.06: 4A*OES/MS; 2.07: 4A*OES/MS; 2.08: 4A*MS; 2.09: 4A*OES/MS |  | 1
2
| 1 | 1 | 2.01: 16.49; 2.02: 18.5; 2.03: 16.127; 2.04: 18.2; 2.05: 15.5; 2.06: 18.53; 2.07: 31; 2.08: 17.0337; 2.09: 16.65 |  | 1
| 1 | 2 | 2.01: 15.78; 2.02: 18.7; 2.03: 16.047; 2.04: 17.45; 2.05: 16.3; 2.06: 17.72; 2.07: 22; 2.08: 16.5647; 2.09: 15.65 |  | 29
| 1 | 3 | 2.01: 15.38; 2.02: 18.7; 2.03: 16.314; 2.04: 19.05; 2.05: 16.6; 2.06: 17.45; 2.07: 34; 2.08: 16.9447; 2.09: 16.05 |  | 16
| 1 | 4 | 2.01: 15.94; 2.02: 18.9; 2.03: 16.656; 2.04: 17.1; 2.05: 16.1; 2.06: 19.47; 2.07: 30; 2.08: 16.8951; 2.09: 16.2 |  | 16.8095
| 1 | 5 | 2.01: 15.16; 2.02: 18.5; 2.03: 16.326; 2.04: 17.5; 2.05: 16.8; 2.06: 18.33; 2.07: 33; 2.08: 16.4675; 2.09: 15.85 |  | 49
| 1 | 6 | 2.01: 16.54; 2.02: 18.6; 2.03: 15.991; 2.04: 18.3; 2.05: 16; 2.06: 18.86; 2.07: 28; 2.08: 17.0317; 2.09: 16.65
| Mean |  | 2.01: 15.8817; 2.02: 18.65; 2.03: 16.2435; 2.04: 17.9333; 2.05: 16.2167; 2.06: 18.3933; 2.07: 29.6667; 2.08: 16.8229; 2.09: 16.175
| Median |  | 2.01: 15.86; 2.02: 18.65; 2.03: 16.2205; 2.04: 17.85; 2.05: 16.2; 2.06: 18.43; 2.07: 30.5; 2.08: 16.9199; 2.09: 16.125
| Std Dev. |  | 2.01: 0.56386; 2.02: 0.15166; 2.03: 0.24408; 2.04: 0.71671; 2.05: 0.46224; 2.06: 0.74067; 2.07: 4.3205; 2.08: 0.24537; 2.09: 0.41201
| Rel.Std.Dev. |  | 2.01: 3.55%; 2.02: 0.8132%; 2.03: 1.503%; 2.04: 3.996%; 2.05: 2.85%; 2.06: 4.027%; 2.07: 14.56%; 2.08: 1.459%; 2.09: 2.547%
| PDM3 |  | 2.01: -5.52%; 2.02: 10.95%; 2.03: -3.367%; 2.04: 6.686%; 2.05: -3.527%; 2.06: 9.422%; 2.07: 76.49%; 2.08: 0.07984%; 2.09: -3.775%
| Z-Score (Absolute) |  | 2.01: 0.87; 2.02: 2.83; 2.03: 0.39; 2.04: 1.87; 2.05: 0.42; 2.06: 2.48; 2.07: NA; 2.08: 0.39; 2.09: 0.48
Results from laboratory 2.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 993b (Certified Value 154 ppm) | [18 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 156.4 | 157 | 163 | 159.18 | 150.5 | 154 | 162 | 142 | 148 | 159.1 | 155.4 | 131 | 160.5845 | 151.9454 | 157 | 151.5 | 146.5 | 1
| 1 | 2 | 144.9 | 157 | 160 | 157.71 | 151.5 | 150 | 163 | 139 | 148 | 156.5 | 157.7 | 120.1 | 160.1229 | 151.5714 | 157 | 146 | 141 | 0
| 1 | 3 | 147.2 | 158 | 162 | 159.07 | 157 | 156 | 166 | 138 | 147 | 154.7 | 157.2 | 128.2 | 160.6646 | 151.6922 | 157 | 151 | 143.5 | 16
| 1 | 4 | 149.2 | 160 | 165 | 159.05 | 149.5 | 152 | 163 | 145 | 145 | 158.5 | 158.2 | 124.5 | 158.3767 | 151.3262 | 159 | 150 | 143 | 153.9216
| 1 | 5 | 144.8 | 157 | 165 | 157.78 | 150.5 | 151 | 166 | 141 | 145 | 158.9 | 155.2 | 119.7 | 159.019 | 150.3397 | 157 | 149.5 | 142.5 | 50
| 1 | 6 | 157.8 | 160 | 162 | 158.11 | 150.5 | 156 | 163 | 145 | 145 | 163.8 | 155.8 | 114.8 | 161.7101 | 151.6232 | 159 | 146
| Mean |  | 150.05 | 158.1667 | 162.8333 | 158.4833 | 151.5833 | 153.1667 | 163.8333 | 141.6667 | 146.3333 | 158.5833 | 156.5833 | 123.05 | 160.0796 | 151.4163 | 157.6667 | 149 | 143.3
| Median |  | 148.2 | 157.5 | 162.5 | 158.58 | 150.5 | 153 | 163 | 141.5 | 146 | 158.7 | 156.5 | 122.3 | 160.3537 | 151.5973 | 157 | 149.75 | 143
| Std Dev. |  | 5.7152 | 1.472 | 1.9408 | 0.69032 | 2.7279 | 2.5626 | 1.7224 | 2.9439 | 1.5055 | 3.0662 | 1.2781 | 5.9969 | 1.2069 | 0.56391 | 1.0328 | 2.429 | 2.0187
| Rel.Std.Dev. |  | 3.809% | 0.9306% | 1.192% | 0.4356% | 1.8% | 1.673% | 1.051% | 2.078% | 1.029% | 1.934% | 0.8163% | 4.874% | 0.7539% | 0.3724% | 0.6551% | 1.63% | 1.409%
| PDM3 |  | -2.515% | 2.758% | 5.79% | 2.964% | -1.519% | -0.4905% | 6.439% | -7.962% | -4.93% | 3.029% | 1.729% | -20.06% | 4.001% | -1.628% | 2.433% | -3.197% | -6.901%
| Z-Score (Absolute) |  | 0.4 | 0.63 | 1.23 | 0.67 | 0.2 | 0 | 1.35 | 1.46 | 0.87 | 0.69 | 0.43 | 3.82 | 0.88 | 0.22 | 0.57 | 0.53 | 1.25
| Analytical results for Cr in OREAS 993b (Certified Value 57 ppm) | [18 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 0
1
| 1 | 1 | 50 | 64 | 58 | 51.2385 | 58 | 56 | 60 | 70 | 50 | 57 | 57 | 62.6544 | 52.212 | 60 | 58 |  |  | 1
| 1 | 2 | 54 | 63 | 58 | 51.111 | 59 | 57 | 60 | 70 | 51 | 56 | 57 | 62.0217 | 51.7083 | 60 | 57 |  |  | 7
| 1 | 3 | 53 | 64 | 57 | 51.927 | 60 | 56 | 60 | 66 | 51 | 54 | 57 | 61.0652 | 52.4248 | 60 | 59 |  |  | 16
| 1 | 4 | 54 | 64 | 58 | 52.811 | 57 | 55 | 60 | 58 | 50 | 59 | 58 | 61.9731 | 53.3772 | 60 | 56 |  |  | 57.2002
| 1 | 5 | 52 | 65 | 58 | 51.7825 | 58 | 56 | 60 | 64 | 50 | 57 | 57 | 61.0745 | 55.3375 | 60 | 60 |  |  | 51
| 1 | 6 | 50 | 65 | 57 | 50.261 | 59 | 57 | 60 | 68 | 51 | 59 | 58 | 63.0604 | 53.5792 | 60 | 55
| Mean |  | 52.1667 | 64.1667 | 57.6667 | 51.5218 | 58.5 | 56.1667 | 60 | 66 | 50.5 | 57 | 57.3333 | 61.9749 | 53.1065 | 60 | 57.5
| Median |  | 52.5 | 64 | 58 | 51.5105 | 58.5 | 56 | 60 | 67 | 50.5 | 57 | 57 | 61.9974 | 52.901 | 60 | 57.5
| Std Dev. |  | 1.8348 | 0.75277 | 0.5164 | 0.86404 | 1.0488 | 0.75277 | 0 | 4.5607 | 0.54772 | 1.8974 | 0.5164 | 0.81001 | 1.3034 | 0 | 1.8708
| Rel.Std.Dev. |  | 3.517% | 1.173% | 0.8955% | 1.677% | 1.793% | 1.34% | 0% | 6.91% | 1.085% | 3.329% | 0.9007% | 1.307% | 2.454% | 0% | 3.254%
| PDM3 |  | -8.8% | 12.18% | 0.8154% | -9.927% | 2.272% | -1.807% | 4.895% | 15.38% | -11.71% | -0.3501% | 0.2327% | 8.347% | -7.157% | 4.895% | 0.524%
| Z-Score (Absolute) |  | 0.82 | 1.09 | 0.05 | 0.93 | 0.19 | 0.19 | NA | 1.38 | 1.09 | 0.05 | 0 | 0.74 | 0.67 | NA | 0.03
Results from laboratories 2.07 and 2.15 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 993b (Certified Value 0.65 ppm) | [18 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 0.7 | 0.6 | 0.7 | 0.585 | 0.63 | 0.67 | 0.7 | 0.8 | 0.63137 | 0.63384 | 0.69 | [6 empty] | 1
| 1 | 2 | 0.6 | 0.7 | 0.7 | 0.6412 | 0.62 | 0.62 | 0.7 | 0.9 | 0.64232 | 0.62443 | 0.69 | [6 empty] | 8
| 1 | 3 | 0.6 | 0.7 | 0.7 | 0.6134 | 0.66 | 0.69 | 0.7 | 0.9 | 0.62956 | 0.62185 | 0.66 | [6 empty] | 16
| 1 | 4 | 0.6 | 0.7 | 0.7 | 0.6524 | 0.59 | 0.65 | 0.7 | 0.9 | 0.63932 | 0.61412 | 0.67 | [6 empty] | 0.64821
| 1 | 5 | 0.6 | 0.6 | 0.7 | 0.5876 | 0.62 | 0.65 | 0.7 | 1 | 0.63842 | 0.61649 | 0.64 | [6 empty] | 52
| 1 | 6 | 0.7 | 0.6 | 0.6 | 0.6048 | 0.62 | 0.64 | 0.7 | 0.8 | 0.63609 | 0.61049 | 0.66
| Mean |  | 0.63333 | 0.65 | 0.68333 | 0.61407 | 0.62333 | 0.65333 | 0.7 | 0.88333 | 0.63618 | 0.6202 | 0.66833
| Median |  | 0.6 | 0.65 | 0.7 | 0.6091 | 0.62 | 0.65 | 0.7 | 0.9 | 0.63725 | 0.61917 | 0.665
| Std Dev. |  | 0.05164 | 0.054772 | 0.040825 | 0.027703 | 0.022509 | 0.024221 | 1.2162e-16 | 0.075277 | 0.0048902 | 0.0083891 | 0.019408
| Rel.Std.Dev. |  | 8.154% | 8.427% | 5.974% | 4.511% | 3.611% | 3.707% | 1.737e-14% | 8.522% | 0.7687% | 1.353% | 2.904%
| PDM3 |  | -2.295% | 0.2759% | 5.418% | -5.268% | -3.838% | 0.7901% | 7.989% | 36.27% | -1.856% | -4.321% | 3.104%
| Z-Score (Absolute) |  | 0.42 | 0 | 0.84 | 0.91 | 0.67 | 0.08 | 1.26 | 5.9 | 0.35 | 0.75 | 0.46
| Analytical results for Cu in OREAS 993b (Certified Value 23.58 wt.%) | [18 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*AAS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS |  | 2
3
| 1 | 1 | 23.327 | 23.0802 | 23.7 | 21.2591 | 22.7 | 24 | 23.6 | 23.9 | 23.65 | 23.488 | > 1 | 23.3377 | 23.8777 | 23.7996 | 23.5 | 23.6 |  | 1
| 1 | 2 | 23.241 | 23.3015 | 23.9 | 21.2482 | 23.9 | 23.8 | 23.5 | 23.3 | 23.5 | 24.058 | > 1 | 24.1009 | 23.823 | 23.434 | 23.5 | 23.6 |  | 31
| 1 | 3 | 23.405 | 23.0764 | 23.4 | 21.2858 | 23.4 | 23.6 | 23.7 | 23.9 | 23.63 | 23.759 | > 1 | 23.177 | 23.9396 | 23.7485 | 23.5 | 23.5 |  | 16
| 1 | 4 | 23.701 | 23.7937 | 23.7 | 21.2981 | 23.4 | 23.4 | 23.5 | 23.6 | 23.31 | 23.68 | > 1 | 23.9323 | 23.6871 | 23.6212 | 23.5 | 23.7 |  | 23.5775
| 1 | 5 | 23.843 | 23.0623 | 23.2 | 21.2837 | 23 | 23.8 | 23.6 | 22.9 | 23.47 | 23.504 | > 1 | 23.7181 | 23.7547 | 23.8037 | 23.5 | 23.6 |  | 53
| 1 | 6 | 23.708 | 23.5787 | 23.5 | 21.2782 | 23.6 | 23.8 | 23.6 | 23.1 | 23.49 | 23.873 | > 1 | 23.1637 | 23.7491 | 23.9532 | 23.5 | 23.6
| Mean |  | 23.5375 | 23.3155 | 23.5667 | 21.2755 | 23.3333 | 23.7333 | 23.5833 | 23.45 | 23.5083 | 23.727 |  | 23.5716 | 23.8052 | 23.7267 | 23.5 | 23.6
| Median |  | 23.553 | 23.1909 | 23.6 | 21.281 | 23.4 | 23.8 | 23.6 | 23.45 | 23.495 | 23.7195 |  | 23.5279 | 23.7889 | 23.7741 | 23.5 | 23.6
| Std Dev. |  | 0.2445 | 0.30815 | 0.25033 | 0.018483 | 0.4274 | 0.20656 | 0.075277 | 0.41833 | 0.12336 | 0.21953 |  | 0.40213 | 0.093049 | 0.17878 | 0 | 0.063246
| Rel.Std.Dev. |  | 1.039% | 1.322% | 1.062% | 0.08687% | 1.832% | 0.8703% | 0.3192% | 1.784% | 0.5247% | 0.9252% |  | 1.706% | 0.3909% | 0.7535% | 0% | 0.268%
| PDM3 |  | -0.1697% | -1.111% | -0.046% | -9.764% | -1.036% | 0.6609% | 0.02469% | -0.5408% | -0.2934% | 0.634% |  | -0.025% | 0.9657% | 0.6328% | -0.3288% | 0.09538%
| Z-Score (Absolute) |  | 0.17 | 1.45 | 0 | 13.24 | 1.35 | 0.96 | 0.1 | 0.67 | 0.34 | 0.93 | NA | 0.03 | 1.38 | 0.92 | 0.39 | 0.19
| Analytical results for Dy in OREAS 993b (Certified Value 1.05 ppm) | [18 empty] | Certified
Dy | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.06: 2.14 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.06: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 1.1; 2.02: 1.3; 2.03: 0.9871; 2.04: 1.2; 2.05: 1.1; 2.06: 0.99435 |  | 1
| 1 | 2 | 2.01: 0.9; 2.02: 1.1; 2.03: 0.9653; 2.04: 1.1; 2.05: 1.1; 2.06: 0.96922 |  | 0
| 1 | 3 | 2.01: 0.9; 2.02: 1.1; 2.03: 1.0704; 2.04: 1.1; 2.05: 1.1; 2.06: 0.96812 |  | 16
| 1 | 4 | 2.01: 0.9; 2.02: 1.4; 2.03: 1.0704; 2.04: 1.1; 2.05: 1.2; 2.06: 0.9643 |  | 1.0467
| 1 | 5 | 2.01: 0.9; 2.02: 1.1; 2.03: 1.0158; 2.04: 1; 2.05: 1.1; 2.06: 0.96908 |  | 54
| 1 | 6 | 2.01: 1; 2.02: 1.1; 2.03: 0.9772; 2.04: 1; 2.05: 1.1; 2.06: 0.98965
| Mean |  | 2.01: 0.95; 2.02: 1.1833; 2.03: 1.0144; 2.04: 1.0833; 2.05: 1.1167; 2.06: 0.97578
| Median |  | 2.01: 0.9; 2.02: 1.1; 2.03: 1.0014; 2.04: 1.1; 2.05: 1.1; 2.06: 0.96915
| Std Dev. |  | 2.01: 0.083666; 2.02: 0.13292; 2.03: 0.046506; 2.04: 0.075277; 2.05: 0.040825; 2.06: 0.012772
| Rel.Std.Dev. |  | 2.01: 8.807%; 2.02: 11.23%; 2.03: 4.585%; 2.04: 6.949%; 2.05: 3.656%; 2.06: 1.309%
| PDM3 |  | 2.01: -9.238%; 2.02: 13.05%; 2.03: -3.088%; 2.04: 3.501%; 2.05: 6.685%; 2.06: -6.775%
| Z-Score (Absolute) |  | 2.01: 0.95; 2.02: 1.29; 2.03: 0.33; 2.04: 0.33; 2.05: 0.65; 2.06: 0.7
| Analytical results for Er in OREAS 993b (Certified Value 0.6 ppm) | [18 empty] | Certified
Er | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.06: 2.14 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.06: 4A*MS |  | 2
3
| 1 | 1 | 2.01: 0.6; 2.02: 0.6; 2.03: 0.5964; 2.04: 0.5; 2.05: 0.6; 2.06: 0.56833 |  | 1
| 1 | 2 | 2.01: 0.6; 2.02: 0.7; 2.03: 0.663; 2.04: 0.7; 2.05: 0.6; 2.06: 0.55707 |  | 0
| 1 | 3 | 2.01: 0.5; 2.02: 0.6; 2.03: 0.6303; 2.04: 0.6; 2.05: 0.6; 2.06: 0.55977 |  | 16
| 1 | 4 | 2.01: 0.5; 2.02: 0.6; 2.03: 0.6703; 2.04: 0.7; 2.05: 0.6; 2.06: 0.55821 |  | 0.59918
| 1 | 5 | 2.01: 0.6; 2.02: 0.6; 2.03: 0.6315; 2.04: 0.6; 2.05: 0.6; 2.06: 0.57076 |  | 55
| 1 | 6 | 2.01: 0.6; 2.02: 0.6; 2.03: 0.5867; 2.04: 0.6; 2.05: 0.6; 2.06: 0.57804
| Mean |  | 2.01: 0.56667; 2.02: 0.61667; 2.03: 0.6297; 2.04: 0.61667; 2.05: 0.6; 2.06: 0.56536
| Median |  | 2.01: 0.6; 2.02: 0.6; 2.03: 0.6309; 2.04: 0.6; 2.05: 0.6; 2.06: 0.56405
| Std Dev. |  | 2.01: 0.05164; 2.02: 0.040825; 2.03: 0.033819; 2.04: 0.075277; 2.05: 0; 2.06: 0.0083647
| Rel.Std.Dev. |  | 2.01: 9.113%; 2.02: 6.62%; 2.03: 5.371%; 2.04: 12.21%; 2.05: 0%; 2.06: 1.48%
| PDM3 |  | 2.01: -5.426%; 2.02: 2.919%; 2.03: 5.094%; 2.04: 2.919%; 2.05: 0.1373%; 2.06: -5.643%
| Z-Score (Absolute) |  | 2.01: 1.89; 2.02: 0.38; 2.03: 0.97; 2.04: 0.38; 2.05: 0.38; 2.06: 1.95
| Analytical results for Eu in OREAS 993b (Indicative Value 0.4 ppm) | [18 empty] | Indicative
Eu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.08; 2.04: 2.10; 2.05: 2.14 |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS; 2.05: 4A*MS |  | 2
2
| 1 | 1 | 2.01: 0.4; 2.02: 0.4; 2.03: 0.37; 2.04: 0.4; 2.05: 0.39807 |  | 1
| 1 | 2 | 2.01: 0.4; 2.02: 0.4; 2.03: 0.44; 2.04: 0.4; 2.05: 0.39843 |  | 21
| 1 | 3 | 2.01: 0.3; 2.02: 0.5; 2.03: 0.44; 2.04: 0.4; 2.05: 0.39734 |  | 16
| 1 | 4 | 2.01: 0.4; 2.02: 0.5; 2.03: 0.42; 2.04: 0.4; 2.05: 0.40518 |  | 0.40467
| 1 | 5 | 2.01: 0.4; 2.02: 0.4; 2.03: 0.43; 2.04: 0.4; 2.05: 0.40462 |  | 27
| 1 | 6 | 2.01: 0.4; 2.02: 0.6; 2.03: 0.4; 2.04: 0.4; 2.05: 0.40848
| Mean |  | 2.01: 0.38333; 2.02: 0.46667; 2.03: 0.41667; 2.04: 0.4; 2.05: 0.40202
| Median |  | 2.01: 0.4; 2.02: 0.45; 2.03: 0.425; 2.04: 0.4; 2.05: 0.40152
| Std Dev. |  | 2.01: 0.040825; 2.02: 0.08165; 2.03: 0.027325; 2.04: 6.0809e-17; 2.05: 0.0046644
| Rel.Std.Dev. |  | 2.01: 10.65%; 2.02: 17.5%; 2.03: 6.558%; 2.04: 1.52e-14%; 2.05: 1.16%
| PDM3 |  | 2.01: -5.273%; 2.02: 15.32%; 2.03: 2.964%; 2.04: -1.154%; 2.05: -0.6553%
| Z-Score (Absolute) |  | 2.01: 0.86; 2.02: 2.98; 2.03: 0.67; 2.04: 0.09; 2.05: 0
| Analytical results for Fe in OREAS 993b (Certified Value 24.53 wt.%) | [18 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 25.51 | 24.22 | 27.6 | 23.4319 | 24.1 | 24.5 | 24.7 | 22.7 | 24.3 | 25.05 | > 10 | 24.3839 | 26.2726 | 25.1329 | 24.3 | 24.8 |  | 1
| 1 | 2 | 25.74 | 24.6 | 27.1 | 23.5184 | 24.4 | 24.4 | 24.6 | 22.6 | 24.41 | 24.96 | > 10 | 24.2651 | 26.2411 | 25.2252 | 24 | 23.8 |  | 32
| 1 | 3 | 24.87 | 24.2 | 27.2 | 23.5327 | 24.8 | 25.1 | 24.8 | 23.4 | 24.39 | 24.34 | > 10 | 24.4076 | 26.0744 | 25.1822 | 24.5 | 24.5 |  | 16
| 1 | 4 | 25.82 | 24.51 | 26.8 | 23.4258 | 23.7 | 24.6 | 24.6 | 21.8 | 24.5 | 25.58 | > 10 | 24.3218 | 26.492 | 25.4377 | 24.4 | 24 |  | 24.5266
| 1 | 5 | 24.82 | 24.58 | 26.9 | 23.4358 | 24.2 | 24.2 | 24.6 | 20.7 | 24.48 | 24.73 | > 10 | 23.9847 | 25.6798 | 25.2095 | 24.6 | 25.2 |  | 56
| 1 | 6 | 24.77 | 24.31 | 27.5 | 23.4572 | 24.5 | 24.7 | 24.6 | 22 | 24.43 | 25.93 | > 10 | 23.5272 | 25.8572 | 25.5844 | 24 | 24.2
| Mean |  | 25.255 | 24.4033 | 27.1833 | 23.467 | 24.2833 | 24.5833 | 24.65 | 22.2 | 24.4183 | 25.0983 |  | 24.1484 | 26.1028 | 25.2953 | 24.3 | 24.4167
| Median |  | 25.19 | 24.41 | 27.15 | 23.4465 | 24.3 | 24.55 | 24.6 | 22.3 | 24.42 | 25.005 |  | 24.2934 | 26.1577 | 25.2173 | 24.35 | 24.35
| Std Dev. |  | 0.48829 | 0.18162 | 0.31885 | 0.046803 | 0.37639 | 0.30605 | 0.083666 | 0.92736 | 0.071391 | 0.57548 |  | 0.3403 | 0.29648 | 0.17623 | 0.25298 | 0.52313
| Rel.Std.Dev. |  | 1.933% | 0.7443% | 1.173% | 0.1994% | 1.55% | 1.245% | 0.3394% | 4.177% | 0.2924% | 2.293% |  | 1.409% | 1.136% | 0.6967% | 1.041% | 2.143%
| PDM3 |  | 2.97% | -0.5025% | 10.83% | -4.32% | -0.9918% | 0.2314% | 0.5032% | -9.486% | -0.4414% | 2.331% |  | -1.542% | 6.427% | 3.134% | -0.9238% | -0.4481%
| Z-Score (Absolute) |  | 2.09 | 0.04 | 6.91 | 2.38 | 0.34 | 0.41 | 0.58 | 5.54 | 0 | 1.7 | NA | 0.67 | 4.21 | 2.19 | 0.3 | 0
| Analytical results for Ga in OREAS 993b (Certified Value 4.14 ppm) | [18 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 4.56 | 4.7 | 4.5 | 15.367 | 3.74 | < 10 | 4.3 | 22 | 4.26 | 5 | < 5 | 4.2984 | 4.0557 | 4.08 |  |  |  | 1
| 1 | 2 | 3.87 | 3.9 | 4.2 | 15.969 | 3.66 | < 10 | 4.8 | 23 | 4.1 | 5 | < 5 | 4.1279 | 4.0286 | 3.96 |  |  |  | 9
| 1 | 3 | 3.95 | 3.9 | 4.4 | 16.172 | 3.96 | < 10 | 4.6 | 23 | 4.1 | 5 | < 5 | 4.3347 | 4.0282 | 3.82 |  |  |  | 16
| 1 | 4 | 4.04 | 4.5 | 4.4 | 16.34 | 3.59 | < 10 | 4.2 | 22 | 4.31 | 6 | < 5 | 4.2494 | 4.0347 | 3.95 |  |  |  | 4.1404
| 1 | 5 | 3.85 | 4.4 | 4.3 | 16.501 | 3.68 | 10 | 4.3 | 22 | 4.24 | 5 | < 5 | 4.1278 | 4.0135 | 3.84 |  |  |  | 57
| 1 | 6 | 4.5 | 3.9 | 4.4 | 16.396 | 3.71 | < 10 | 4.2 | 22 | 4.37 | 5 | < 5 | 4.2067 | 4.0544 | 3.98
| Mean |  | 4.1283 | 4.2167 | 4.3667 | 16.1242 | 3.7233 | 10 | 4.4 | 22.3333 | 4.23 | 5.1667 |  | 4.2242 | 4.0359 | 3.9383
| Median |  | 3.995 | 4.15 | 4.4 | 16.256 | 3.695 | 10 | 4.3 | 22 | 4.25 | 5 |  | 4.2281 | 4.0316 | 3.955
| Std Dev. |  | 0.31884 | 0.36009 | 0.10328 | 0.41525 | 0.1266 |  | 0.24495 | 0.5164 | 0.11027 | 0.40825 |  | 0.086289 | 0.016433 | 0.096003
| Rel.Std.Dev. |  | 7.723% | 8.54% | 2.365% | 2.575% | 3.4% |  | 5.567% | 2.312% | 2.607% | 7.902% |  | 2.043% | 0.4072% | 2.438%
| PDM3 |  | -0.2908% | 1.843% | 5.466% | 289.4% | -10.07% | 141.5% | 6.271% | 439.4% | 2.165% | 24.79% |  | 2.024% | -2.524% | -4.88%
| Z-Score (Absolute) |  | 0.34 | 0.01 | 0.54 | 44.09 | 1.84 | NA | 0.67 | 67.08 | 0.04 | NA | 6.37 | 0.01 | 0.68 | 1.04
Results from laboratory 2.11 were removed due to their 1 ppm reading resolution. Results from laboratory 2.06 were removed due to their 10 ppm reading resolution.
| Analytical results for Gd in OREAS 993b (Certified Value 1.3 ppm) | [18 empty] | Certified
Gd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.06: 2.14 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.06: 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.01: 1.3; 2.02: 1.4; 2.03: 1.4341; 2.04: 1.1; 2.05: 1.4; 2.06: 1.1191 |  |  |  | 1
| 1 | 2 | 2.01: 1.2; 2.02: 1.2; 2.03: 1.3635; 2.04: 1.3; 2.05: 1.3; 2.06: 1.1128 |  |  |  | 0
| 1 | 3 | 2.01: 1.1; 2.02: 1.3; 2.03: 1.4454; 2.04: 1.2; 2.05: 1.3; 2.06: 1.1263 |  |  |  | 16
| 1 | 4 | 2.01: 1.2; 2.02: 1.2; 2.03: 1.5556; 2.04: 1.4; 2.05: 1.5; 2.06: 1.1174 |  |  |  | 1.2956
| 1 | 5 | 2.01: 1.2; 2.02: 1.4; 2.03: 1.4822; 2.04: 1.3; 2.05: 1.4; 2.06: 1.1545 |  |  |  | 58
| 1 | 6 | 2.01: 1.3; 2.02: 1.3; 2.03: 1.5924; 2.04: 1.3; 2.05: 1.4; 2.06: 1.1377
| Mean |  | 2.01: 1.2167; 2.02: 1.3; 2.03: 1.4789; 2.04: 1.2667; 2.05: 1.3833; 2.06: 1.128
| Median |  | 2.01: 1.2; 2.02: 1.3; 2.03: 1.4638; 2.04: 1.3; 2.05: 1.4; 2.06: 1.1227
| Std Dev. |  | 2.01: 0.075277; 2.02: 0.089443; 2.03: 0.083941; 2.04: 0.10328; 2.05: 0.075277; 2.06: 0.015639
| Rel.Std.Dev. |  | 2.01: 6.187%; 2.02: 6.88%; 2.03: 5.676%; 2.04: 8.154%; 2.05: 5.442%; 2.06: 1.386%
| PDM3 |  | 2.01: -6.091%; 2.02: 0.3407%; 2.03: 14.15%; 2.04: -2.232%; 2.05: 6.773%; 2.06: -12.94%
| Z-Score (Absolute) |  | 2.01: 0.54; 2.02: 0.13; 2.03: 1.58; 2.04: 0.13; 2.05: 0.81; 2.06: 1.26
| Analytical results for Ge in OREAS 993b (Certified Value 0.45 ppm) | [18 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.08 | 2.11 | 2.13 | 2.14 | 2.16 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 0.29 | 1.5 | < 0.1 | 0.45 | 1 | 4.3 | 0.33353 | 0.71016 | 0.5 | [8 empty] | 1
| 1 | 2 | 0.22 | 1.4 | < 0.1 | 0.42 | 1 | 4.5 | 0.24766 | 0.72054 | 0.47 | [8 empty] | 10
| 1 | 3 | 0.24 | 1.4 | < 0.1 | 0.68 | 1.1 | 4.3 | 0.25993 | 0.7091 | 0.48 | [8 empty] | 16
| 1 | 4 | 0.22 | 1.4 | < 0.1 | 0.43 | 0.9 | 4.4 | 0.3952 | 0.73241 | 0.48 | [8 empty] | 0.4495
| 1 | 5 | 0.21 | 1.3 | < 0.1 | 0.46 | 1.1 | 4.3 | 0.36914 | 0.74982 | 0.44 | [8 empty] | 59
| 1 | 6 | 0.2 | 1.4 | < 0.1 | 0.44 | 0.9 | 4.4 | 0.35681 | 0.72059 | 0.55
| Mean |  | 2.01: 0.23; 2.02: 1.4; 2.04: 0.48; 2.05: 1; 2.06: 4.3667; 2.08: 0.32705; 2.09: 0.72377; 2.10: 0.48667
| Median |  | 2.01: 0.22; 2.02: 1.4; 2.04: 0.445; 2.05: 1; 2.06: 4.35; 2.08: 0.34517; 2.09: 0.72056; 2.10: 0.48
| Std Dev. |  | 2.01: 0.032249; 2.02: 0.063246; 2.04: 0.098995; 2.05: 0.089443; 2.06: 0.08165; 2.08: 0.060252; 2.09: 0.015324; 2.10: 0.036697
| Rel.Std.Dev. |  | 2.01: 14.02%; 2.02: 4.518%; 2.04: 20.62%; 2.05: 8.944%; 2.06: 1.87%; 2.08: 18.42%; 2.09: 2.117%; 2.10: 7.54%
| PDM3 |  | 2.01: -48.83%; 2.02: 211.5%; 2.04: 6.786%; 2.05: 122.5%; 2.06: 871.5%; 2.08: -27.24%; 2.09: 61.02%; 2.10: 8.269%
| Z-Score (Absolute) |  | 0.91 | NA | 1.86 | 0.4 | NA | NA | 0.4 | 1.68 | 0.44
Results from laboratories 2.03, 2.08 and 2.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 993b (Certified Value 0.71 ppm) | [18 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 0.72 | 0.8 | 0.8 | 0.6048 | 0.7 | 0.7 | 0.72 | 0.8 | 0.66659 | 0.66338 | 0.8 | [6 empty] | 1
| 1 | 2 | 0.68 | 0.7 | 0.8 | 0.5464 | 0.7 | 0.7 | 0.7 | 0.8 | 0.67813 | 0.62287 | 0.7 | [6 empty] | 0
| 1 | 3 | 0.64 | 0.7 | 0.8 | 0.6111 | 0.7 | 0.7 | 0.7 | 0.8 | 0.67241 | 0.63429 | 0.8 | [6 empty] | 16
| 1 | 4 | 0.62 | 0.7 | 0.8 | 0.6761 | 0.7 | 0.7 | 0.81 | 0.8 | 0.70766 | 0.64958 | 0.8 | [6 empty] | 0.70998
| 1 | 5 | 0.68 | 0.8 | 0.8 | 0.5266 | 0.7 | 0.7 | 0.72 | 0.8 | 0.70729 | 0.65811 | 0.8 | [6 empty] | 60
| 1 | 6 | 1.02 | 0.7 | 0.7 | 0.6138 | 0.7 | 0.7 | 0.71 | 0.8 | 0.68965 | 0.66167 | 0.7
| Mean |  | 0.72667 | 0.73333 | 0.78333 | 0.59647 | 0.7 | 0.7 | 0.72667 | 0.8 | 0.68696 | 0.64832 | 0.76667
| Median |  | 0.68 | 0.7 | 0.8 | 0.60795 | 0.7 | 0.7 | 0.715 | 0.8 | 0.68389 | 0.65385 | 0.8
| Std Dev. |  | 0.14787 | 0.05164 | 0.040825 | 0.053504 | 1.2162e-16 | 1.2162e-16 | 0.041793 | 1.2162e-16 | 0.017626 | 0.016413 | 0.05164
| Rel.Std.Dev. |  | 20.35% | 7.042% | 5.212% | 8.97% | 1.737e-14% | 1.737e-14% | 5.751% | 1.52e-14% | 2.566% | 2.532% | 6.736%
| PDM3 |  | 2.351% | 3.29% | 10.33% | -15.99% | -1.405% | -1.405% | 2.351% | 12.68% | -3.242% | -8.685% | 7.985%
| Z-Score (Absolute) |  | 0 | 0.11 | 0.96 | 2.21 | 0.45 | 0.45 | 0 | 1.25 | 0.67 | 1.33 | 0.68
| Analytical results for Hg in OREAS 993b (Indicative Value 0.64 ppm) | [18 empty] | Indicative
Hg | Round | Replicate | 2.01: Lab; 2.02: Lab |  |  |  | 1
| No | No | 2.01: 2.12; 2.02: 2.14 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.01: 0.58; 2.02: < 1 |  |  |  | 1
| 1 | 2 | 2.01: 0.62; 2.02: < 1 |  |  |  | 35
| 1 | 3 | 2.01: 0.68; 2.02: < 1 |  |  |  | 16
| 1 | 4 | 2.01: 0.69; 2.02: < 1 |  |  |  | 0.635
| 1 | 5 | 2.01: 0.65; 2.02: < 1 |  |  |  | 25
| 1 | 6 | 2.01: 0.59; 2.02: < 1
| Mean |  | 2.01: 0.635
| Median |  | 2.01: 0.635
| Std Dev. |  | 2.01: 0.045935
| Rel.Std.Dev. |  | 2.01: 7.234%
| PDM3 |  | 2.01: -2.22e-14%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.67
| Analytical results for Ho in OREAS 993b (Certified Value 0.2 ppm) | [18 empty] | Certified
Ho | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.06: 2.14 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.06: 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.01: 0.2; 2.02: 0.3; 2.03: 0.2022; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19596 |  |  |  | 1
| 1 | 2 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.22173; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19637 |  |  |  | 0
| 1 | 3 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.21418; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19364 |  |  |  | 16
| 1 | 4 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.22667; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19422 |  |  |  | 0.20192
| 1 | 5 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.22132; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19568 |  |  |  | 61
| 1 | 6 | 2.01: 0.2; 2.02: 0.3; 2.03: 0.20935; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19919
| Mean |  | 2.01: 0.2; 2.02: 0.23333; 2.03: 0.21591; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19584
| Median |  | 2.01: 0.2; 2.02: 0.2; 2.03: 0.21774; 2.04: 0.2; 2.05: 0.2; 2.06: 0.19582
| Std Dev. |  | 2.01: 3.0405e-17; 2.02: 0.05164; 2.03: 0.0090869; 2.04: 3.0405e-17; 2.05: 3.0405e-17; 2.06: 0.0019522
| Rel.Std.Dev. |  | 2.01: 1.52e-14%; 2.02: 22.13%; 2.03: 4.209%; 2.04: 1.52e-14%; 2.05: 1.52e-14%; 2.06: 0.9968%
| PDM3 |  | 2.01: -0.9506%; 2.02: 15.56%; 2.03: 6.928%; 2.04: -0.9506%; 2.05: -0.9506%; 2.06: -3.01%
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 10.81; 2.03: 5.16; 2.04: 0; 2.05: 0; 2.06: 1.35
| Analytical results for In in OREAS 993b (Certified Value 1.45 ppm) | [18 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.46 | 1.43 | 1.46 | 1.0596 | 1.495 | 1.4 | < 2 | 1.47 | 1.44 | 1.4763 | 1.2822 | 1.55 | [5 empty] | 1
| 1 | 2 | 1.29 | 1.48 | 1.45 | 1.1424 | 1.435 | 1.4 | < 2 | 1.46 | 1.47 | 1.4477 | 1.2655 | 1.425 | [5 empty] | 2
| 1 | 3 | 1.33 | 1.44 | 1.44 | 1.243 | 1.535 | 1.4 | < 2 | 1.43 | 1.46 | 1.4765 | 1.2531 | 1.465 | [5 empty] | 16
| 1 | 4 | 1.32 | 1.44 | 1.42 | 1.144 | 1.395 | 1.3 | < 2 | 1.5 | 1.47 | 1.4618 | 1.3293 | 1.43 | [5 empty] | 1.4539
| 1 | 5 | 1.31 | 1.55 | 1.48 | 1.148 | 1.42 | 1.4 | < 2 | 1.48 | 1.49 | 1.4556 | 1.2798 | 1.465 | [5 empty] | 62
| 1 | 6 | 1.36 | 1.4 | 1.46 | 1.0905 | 1.475 | 1.4 | < 2 | 1.52 | 1.45 | 1.491 | 1.3165 | 1.51
| Mean |  | 1.345 | 1.4567 | 1.4517 | 1.1379 | 1.4592 | 1.3833 |  | 1.4767 | 1.4633 | 1.4681 | 1.2877 | 1.4742
| Median |  | 1.325 | 1.44 | 1.455 | 1.1432 | 1.455 | 1.4 |  | 1.475 | 1.465 | 1.469 | 1.281 | 1.465
| Std Dev. |  | 0.06091 | 0.052409 | 0.020412 | 0.062573 | 0.052002 | 0.040825 |  | 0.031411 | 0.017512 | 0.015956 | 0.029472 | 0.048106
| Rel.Std.Dev. |  | 4.529% | 3.598% | 1.406% | 5.499% | 3.564% | 2.951% |  | 2.127% | 1.197% | 1.087% | 2.289% | 3.263%
| PDM3 |  | -7.49% | 0.1907% | -0.1532% | -21.73% | 0.3627% | -4.853% |  | 1.566% | 0.6493% | 0.9804% | -11.43% | 1.394%
| Z-Score (Absolute) |  | 3.46 | 0.08 | 0.08 | 10.04 | 0.16 | 2.25 | 14.41 | 0.71 | 0.29 | 0.44 | 5.28 | 0.63
| Analytical results for K in OREAS 993b (Certified Value 0.668 wt.%) | [18 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.66 | 0.62 | 0.661 | 0.67084 | 0.7 | 0.67 | 0.65 | 0.63 | 0.62 | 0.7 | 0.69 | 0.65319 | 0.72007 | 0.67 | 0.7 | 0.74 |  | 1
| 1 | 2 | 0.67 | 0.62 | 0.663 | 0.67077 | 0.72 | 0.68 | 0.63 | 0.59 | 0.63 | 0.69 | 0.7 | 0.65116 | 0.70145 | 0.67 | 0.67 | 0.71 |  | 24
| 1 | 3 | 0.62 | 0.62 | 0.665 | 0.67694 | 0.7 | 0.69 | 0.67 | 0.62 | 0.61 | 0.68 | 0.7 | 0.64699 | 0.7119 | 0.49 | 0.7 | 0.73 |  | 16
| 1 | 4 | 0.66 | 0.64 | 0.669 | 0.67317 | 0.7 | 0.68 | 0.65 | 0.59 | 0.63 | 0.69 | 0.7 | 0.65528 | 0.71413 | 0.67 | 0.68 | 0.72 |  | 0.66846
| 1 | 5 | 0.64 | 0.63 | 0.657 | 0.67747 | 0.69 | 0.67 | 0.64 | 0.55 | 0.62 | 0.69 | 0.7 | 0.63736 | 0.71063 | 0.66 | 0.71 | 0.72 |  | 63
| 1 | 6 | 0.64 | 0.62 | 0.662 | 0.6768 | 0.71 | 0.69 | 0.6 | 0.62 | 0.61 | 0.72 | 0.69 | 0.64124 | 0.71586 | 0.66 | 0.68
| Mean |  | 0.64833 | 0.625 | 0.66283 | 0.67433 | 0.70333 | 0.68 | 0.64 | 0.6 | 0.62 | 0.695 | 0.69667 | 0.64754 | 0.71234 | 0.63667 | 0.69 | 0.724
| Median |  | 0.65 | 0.62 | 0.6625 | 0.67499 | 0.7 | 0.68 | 0.645 | 0.605 | 0.62 | 0.69 | 0.7 | 0.64908 | 0.71302 | 0.665 | 0.69 | 0.72
| Std Dev. |  | 0.018348 | 0.0083666 | 0.0040208 | 0.0031309 | 0.010328 | 0.0089443 | 0.023664 | 0.029665 | 0.0089443 | 0.013784 | 0.005164 | 0.00705 | 0.0062779 | 0.072019 | 0.015492 | 0.011402
| Rel.Std.Dev. |  | 2.83% | 1.339% | 0.6066% | 0.4643% | 1.468% | 1.315% | 3.698% | 4.944% | 1.443% | 1.983% | 0.7412% | 1.089% | 0.8813% | 11.31% | 2.245% | 1.575%
| PDM3 |  | -3.011% | -6.502% | -0.8419% | 0.8784% | 5.217% | 1.726% | -4.258% | -10.24% | -7.25% | 3.97% | 4.219% | -3.13% | 6.564% | -4.756% | 3.222% | 8.308%
| Z-Score (Absolute) |  | 0.48 | 1.04 | 0.14 | 0.14 | 0.83 | 0.27 | 0.68 | 1.63 | 1.16 | 0.63 | 0.67 | 0.5 | 1.04 | 0.76 | 0.51 | 1.32
| Analytical results for La in OREAS 993b (Certified Value 10.2 ppm) | [18 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.04 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 10 | 10.9 | 9.75 | 11 | 10 | 8.8 | 8 | 11.3 | 16 | 10.3163 | 10.8 | [6 empty] | 1
| 1 | 2 | 8.9 | 11.1 | 9.96 | 10.6 | 10 | 9.7 | 9 | 10.3 | 16 | 9.9706 | 10.4 | [6 empty] | 25
| 1 | 3 | 8.9 | 10.8 | 9.8 | 11.2 | 10 | 10 | 9 | 10.5 | 17 | 9.9278 | 10.2 | [6 empty] | 16
| 1 | 4 | 9.2 | 10.8 | 9.82 | 10.2 | 10 | 9.5 | 8 | 11.6 | 16 | 10.3263 | 10.3 | [6 empty] | 10.2212
| 1 | 5 | 8.8 | 10.9 | 9.8 | 10.6 | 10 | 9.9 | 8 | 10.8 | 16 | 10.251 | 10.1 | [6 empty] | 64
| 1 | 6 | 9.7 | 11 | 9.69 | 11 | 10 | 9.2 | 8 | 11.1 | 16 | 10.2079 | 10.7
| Mean |  | 9.25 | 10.9167 | 9.8033 | 10.7667 | 10 | 9.5167 | 8.3333 | 10.9333 | 16.1667 | 10.1666 | 10.4167
| Median |  | 9.05 | 10.9 | 9.8 | 10.8 | 10 | 9.6 | 8 | 10.95 | 16 | 10.2294 | 10.35
| Std Dev. |  | 0.49295 | 0.1169 | 0.090037 | 0.36697 | 0 | 0.4535 | 0.5164 | 0.49261 | 0.40825 | 0.17443 | 0.27869
| Rel.Std.Dev. |  | 5.329% | 1.071% | 0.9184% | 3.408% | 0% | 4.765% | 6.197% | 4.506% | 2.525% | 1.716% | 2.675%
| PDM3 |  | -9.502% | 6.804% | -4.089% | 5.336% | -2.165% | -6.893% | -18.47% | 6.967% | 58.17% | -0.5342% | 1.912%
| Z-Score (Absolute) |  | 1.26 | 0.76 | 0.59 | 0.58 | NA | 0.94 | NA | 0.78 | NA | 0.15 | 0.15
Results from laboratories 2.09 and 2.11 were removed due to their 1 ppm reading resolution. Results from laboratory 2.06 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 993b (Certified Value 23.7 ppm) | [18 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 24.8 | 24 | 24 | 24.1033 | 25.4 | 30 | < 100 | 22 | 22 | 20 | 23.8 | 23.7854 | 22.2382 | < 100 | 25 |  |  | 1
| 1 | 2 | 21.1 | 24 | 24 | 23.5962 | 25.7 | 30 | < 100 | 23 | 24 | 19.4 | 23.9 | 22.2522 | 21.7185 | < 100 | 24.2 |  |  | 26
| 1 | 3 | 20.7 | 24 | 24.8 | 25.0891 | 25.9 | 30 | < 100 | 22 | 23 | 18.9 | 23.6 | 23.6947 | 21.9817 | < 100 | 25 |  |  | 16
| 1 | 4 | 22.5 | 24 | 24.4 | 25.2198 | 25 | 30 | < 100 | 23 | 24 | 19.6 | 24.2 | 23.1271 | 22.265 | < 100 | 24.4 |  |  | 23.666
| 1 | 5 | 21.8 | 24 | 24.7 | 23.9446 | 25.2 | 30 | < 100 | 23 | 23 | 19.3 | 23.1 | 23.4165 | 22.2279 | < 100 | 25.5 |  |  | 65
| 1 | 6 | 23.6 | 24 | 24.3 | 25.4222 | 25.7 | 30 | < 100 | 23 | 23 | 19.2 | 23.4 | 22.5646 | 22.3106 | < 100 | 24.3
| Mean |  | 22.4167 | 24 | 24.3667 | 24.5625 | 25.4833 | 30 |  | 22.6667 | 23.1667 | 19.4 | 23.6667 | 23.1401 | 22.1237 |  | 24.7333
| Median |  | 22.15 | 24 | 24.35 | 24.5962 | 25.55 | 30 |  | 23 | 23 | 19.35 | 23.7 | 23.2718 | 22.233 |  | 24.7
| Std Dev. |  | 1.5587 | 0 | 0.33862 | 0.77134 | 0.34303 | 0 |  | 0.5164 | 0.75277 | 0.37417 | 0.38816 | 0.61971 | 0.22943 |  | 0.51251
| Rel.Std.Dev. |  | 6.953% | 0% | 1.39% | 3.14% | 1.346% | 0% |  | 2.278% | 3.249% | 1.929% | 1.64% | 2.678% | 1.037% |  | 2.072%
| PDM3 |  | -5.279% | 1.411% | 2.961% | 3.788% | 7.679% | 26.76% |  | -4.223% | -2.11% | -18.03% | 0.002712% | -2.222% | -6.517% |  | 4.51%
| Z-Score (Absolute) |  | 0.92 | 0.11 | 0.35 | 0.48 | 1.08 | NA | 17.08 | 0.76 | 0.44 | 2.89 | 0.11 | 0.45 | 1.12 | 17.08 | 0.59
Results from laboratory 2.06 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 993b (Indicative Value 0.079 ppm) | [18 empty] | Indicative
Lu | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.08; 2.04: 2.10; 2.05: 2.14 |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS; 2.05: 4A*MS |  |  | 3
3
| 1 | 1 | 2.01: < 0.1; 2.02: 0.07; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.079435 |  |  | 1
| 1 | 2 | 2.01: < 0.1; 2.02: 0.07; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.083132 |  |  | 20
| 1 | 3 | 2.01: < 0.1; 2.02: 0.1; 2.03: 0.1; 2.04: < 0.1; 2.05: 0.07755 |  |  | 16
| 1 | 4 | 2.01: < 0.1; 2.02: 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.082043 |  |  | 0.078703
| 1 | 5 | 2.01: < 0.1; 2.02: 0.09; 2.03: 0.1; 2.04: < 0.1; 2.05: 0.081675 |  |  | 26
| 1 | 6 | 2.01: < 0.1; 2.02: 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.082818
| Mean |  | 2.02: 0.088333; 2.03: 0.1; 2.05: 0.081109
| Median |  | 2.02: 0.095; 2.03: 0.1; 2.05: 0.081859
| Std Dev. |  | 2.02: 0.01472; 2.03: 0; 2.05: 0.0021758
| Rel.Std.Dev. |  | 2.02: 16.66%; 2.03: 0%; 2.05: 2.683%
| PDM3 |  | 2.02: 12.24%; 2.03: 27.06%; 2.05: 3.057%
| Z-Score (Absolute) |  | 2.01: 0.67; 2.02: 0.88; 2.03: 0; 2.04: 0.67; 2.05: 0.58
| Analytical results for Mg in OREAS 993b (Certified Value 0.443 wt.%) | [18 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.42 | 0.47 | 0.465 | 0.4799 | 0.43 | 0.42 | 0.41 | 0.45 | 0.44 | 0.46 | 0.53 | 0.5055 | 0.45137 | 0.40517 | 0.41 | 0.44 | 0.44 | 1
| 1 | 2 | 0.43 | 0.48 | 0.461 | 0.47484 | 0.44 | 0.41 | 0.4 | 0.28 | 0.46 | 0.45 | 0.53 | 0.5016 | 0.4614 | 0.4016 | 0.39 | 0.42 | 0.43 | 27
| 1 | 3 | 0.42 | 0.47 | 0.46 | 0.47214 | 0.44 | 0.42 | 0.43 | 0.28 | 0.45 | 0.44 | 0.54 | 0.5064 | 0.45711 | 0.40132 | 0.4 | 0.43 | 0.44 | 16
| 1 | 4 | 0.43 | 0.48 | 0.467 | 0.47927 | 0.42 | 0.41 | 0.41 | 0.43 | 0.46 | 0.46 | 0.54 | 0.5081 | 0.46288 | 0.40908 | 0.42 | 0.42 | 0.43 | 0.4429
| 1 | 5 | 0.41 | 0.48 | 0.469 | 0.47303 | 0.43 | 0.41 | 0.41 | 0.3 | 0.45 | 0.45 | 0.53 | 0.4993 | 0.45094 | 0.41799 | 0.42 | 0.44 | 0.43 | 66
| 1 | 6 | 0.41 | 0.48 | 0.463 | 0.47798 | 0.44 | 0.42 | 0.42 | 0.45 | 0.45 | 0.46 | 0.53 | 0.4923 | 0.45515 | 0.42534 | 0.4 | 0.43
| Mean |  | 0.42 | 0.47667 | 0.46417 | 0.47619 | 0.43333 | 0.415 | 0.41333 | 0.365 | 0.45167 | 0.45333 | 0.53333 | 0.5022 | 0.45648 | 0.41008 | 0.40667 | 0.43 | 0.434
| Median |  | 0.42 | 0.48 | 0.464 | 0.47641 | 0.435 | 0.415 | 0.41 | 0.365 | 0.45 | 0.455 | 0.53 | 0.50355 | 0.45613 | 0.40712 | 0.405 | 0.43 | 0.43
| Std Dev. |  | 0.0089443 | 0.005164 | 0.0034881 | 0.0033067 | 0.008165 | 0.0054772 | 0.010328 | 0.086429 | 0.0075277 | 0.008165 | 0.005164 | 0.005833 | 0.0049827 | 0.0096887 | 0.012111 | 0.0089443 | 0.0054772
| Rel.Std.Dev. |  | 2.13% | 1.083% | 0.7515% | 0.6944% | 1.884% | 1.32% | 2.499% | 23.68% | 1.667% | 1.801% | 0.9682% | 1.161% | 1.092% | 2.363% | 2.978% | 2.08% | 1.262%
| PDM3 |  | -5.171% | 7.623% | 4.801% | 7.516% | -2.161% | -6.3% | -6.676% | -17.59% | 1.979% | 2.355% | 20.42% | 13.39% | 3.064% | -7.41% | -8.182% | -2.913% | -2.01%
| Z-Score (Absolute) |  | 0.42 | 1.28 | 0.91 | 1.27 | 0.02 | 0.57 | 0.62 | 2.07 | 0.53 | 0.58 | 2.98 | 2.05 | 0.67 | 0.72 | 0.82 | 0.12 | 0
| Analytical results for Mn in OREAS 993b (Certified Value 0.014 wt.%) | [18 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.0139 | 0.0144 | 0.0139 | 0.014089 | 0.0137 | 0.0136 | 0.01 | 0.0141 | 0.0402 | 0.0151 | 0.0168 | 0.0138 | 0.014536 | 0.012566 | 0.01 | 0.0134 | 0.0137 | 1
| 1 | 2 | 0.0141 | 0.0144 | 0.0139 | 0.013939 | 0.0139 | 0.0135 | 0.02 | 0.0148 | 0.0404 | 0.0148 | 0.0163 | 0.0136 | 0.013961 | 0.012302 | 0.01 | 0.0131 | 0.01325 | 4
| 1 | 3 | 0.0136 | 0.0144 | 0.0139 | 0.01408 | 0.0137 | 0.0139 | 0.01 | 0.015 | 0.0401 | 0.0147 | 0.0164 | 0.0137 | 0.014133 | 0.012213 | 0.01 | 0.0132 | 0.01355 | 16
| 1 | 4 | 0.0141 | 0.0145 | 0.014 | 0.014046 | 0.0135 | 0.0135 | 0.01 | 0.0141 | 0.0401 | 0.0152 | 0.0163 | 0.013 | 0.015135 | 0.01284 | 0.01 | 0.0129 | 0.01335 | 0.01386
| 1 | 5 | 0.0135 | 0.0145 | 0.014 | 0.014033 | 0.0136 | 0.0138 | 0.01 | 0.0134 | 0.0403 | 0.0146 | 0.0162 | 0.0129 | 0.015193 | 0.012861 | 0.01 | 0.0138 | 0.0134 | 67
| 1 | 6 | 0.0136 | 0.0145 | 0.0139 | 0.01408 | 0.0138 | 0.0138 | 0.01 | 0.0144 | 0.0401 | 0.0157 | 0.0161 | 0.0137 | 0.01464 | 0.01236 | 0.01 | 0.013
| Mean |  | 0.0138 | 0.01445 | 0.013933 | 0.014044 | 0.0137 | 0.013683 | 0.011667 | 0.0143 | 0.0402 | 0.015017 | 0.01635 | 0.01345 | 0.014599 | 0.012524 | 0.01 | 0.013233 | 0.01345
| Median |  | 0.01375 | 0.01445 | 0.0139 | 0.014063 | 0.0137 | 0.0137 | 0.01 | 0.01425 | 0.04015 | 0.01495 | 0.0163 | 0.01365 | 0.014588 | 0.012463 | 0.01 | 0.01315 | 0.0134
| Std Dev. |  | 0.00026833 | 5.4772e-05 | 5.164e-05 | 5.6297e-05 | 0.00014142 | 0.00017224 | 0.0040825 | 0.00057271 | 0.00012649 | 0.00040702 | 0.0002429 | 0.0003937 | 0.000504 | 0.00027885 | 0 | 0.0003266 | 0.00017678
| Rel.Std.Dev. |  | 1.944% | 0.379% | 0.3706% | 0.4008% | 1.032% | 1.259% | 34.99% | 4.005% | 0.3147% | 2.71% | 1.486% | 2.927% | 3.452% | 2.227% | 0% | 2.468% | 1.314%
| PDM3 |  | -0.4353% | 4.254% | 0.5267% | 1.329% | -1.157% | -1.277% | -15.83% | 3.172% | 190% | 8.343% | 17.96% | -2.96% | 5.332% | -9.643% | -27.85% | -4.524% | -2.96%
| Z-Score (Absolute) |  | 0 | 0.77 | 0.16 | 0.29 | 0.12 | 0.14 | 2.54 | 0.59 | 31.41 | 1.45 | 3.03 | 0.42 | 0.95 | 1.52 | 4.52 | 0.67 | 0.42
| Analytical results for Mo in OREAS 993b (Certified Value 0.219 wt.%) | [18 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.22861 | 0.2335 | 0.225 | 0.23043 | 0.199 | 0.21 | 0.232 | 0.207 | 0.1967 | 0.22738 | > 0.1 | 0.1953 | 0.23265 | 0.18246 | 0.226 | 0.21 | 0.232 | 1
| 1 | 2 | 0.22951 | 0.233 | 0.229 | 0.23029 | 0.1965 | 0.209 | 0.232 | 0.217 | 0.1999 | 0.22516 | > 0.1 | 0.20562 | 0.23032 | 0.1828 | 0.223 | 0.191 | 0.2226 | 29
| 1 | 3 | 0.22355 | 0.2342 | 0.228 | 0.23265 | 0.214 | 0.207 | 0.233 | 0.212 | 0.1951 | 0.22166 | > 0.1 | 0.20655 | 0.23429 | 0.18069 | 0.229 | 0.198 | 0.2252 | 16
| 1 | 4 | 0.23198 | 0.236 | 0.228 | 0.23218 | 0.1885 | 0.207 | 0.231 | 0.21 | 0.1966 | 0.23244 | > 0.1 | 0.20176 | 0.23055 | 0.18835 | 0.226 | 0.1925 | 0.224 | 0.21888
| 1 | 5 | 0.222 | 0.2346 | 0.227 | 0.23238 | 0.1985 | 0.209 | 0.231 | 0.215 | 0.1946 | 0.22383 | > 0.1 | 0.20804 | 0.23395 | 0.18437 | 0.229 | 0.195 | 0.233 | 68
| 1 | 6 | 0.22098 | 0.2345 | 0.227 | 0.23099 | 0.203 | 0.209 | 0.232 | 0.21 | 0.1965 | 0.23204 | > 0.1 | 0.19426 | 0.23572 | 0.18991 | 0.226 | 0.211
| Mean |  | 0.22611 | 0.2343 | 0.22733 | 0.23149 | 0.19992 | 0.2085 | 0.23183 | 0.21183 | 0.19657 | 0.22709 |  | 0.20192 | 0.23291 | 0.18476 | 0.2265 | 0.19958 | 0.22736
| Median |  | 0.22608 | 0.23435 | 0.2275 | 0.23159 | 0.19875 | 0.209 | 0.232 | 0.211 | 0.19655 | 0.22627 |  | 0.20369 | 0.2333 | 0.18358 | 0.226 | 0.1965 | 0.2252
| Std Dev. |  | 0.0045181 | 0.0010354 | 0.0013663 | 0.0010419 | 0.0083989 | 0.0012247 | 0.00075277 | 0.003656 | 0.0018522 | 0.0044059 |  | 0.0059179 | 0.0021558 | 0.0036119 | 0.0022583 | 0.0087887 | 0.0047946
| Rel.Std.Dev. |  | 1.998% | 0.4419% | 0.601% | 0.4501% | 4.201% | 0.5874% | 0.3247% | 1.726% | 0.9423% | 1.94% |  | 2.931% | 0.9256% | 1.955% | 0.997% | 4.404% | 2.109%
| PDM3 |  | 3.3% | 7.044% | 3.861% | 5.758% | -8.665% | -4.743% | 5.917% | -3.221% | -10.2% | 3.748% |  | -7.749% | 6.41% | -15.59% | 3.48% | -8.817% | 3.873%
| Z-Score (Absolute) |  | 0.02 | 0.74 | 0.1 | 0.48 | 2.44 | 1.64 | 0.51 | 1.34 | 2.75 | 0.07 | NA | 2.25 | 0.61 | 3.84 | 0.02 | 2.47 | 0.1
| Analytical results for Na in OREAS 993b (Certified Value 0.43 wt.%) | [18 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.434 | 0.42 | 0.434 | 0.43731 | 0.43 | 0.43 | 0.43 | 0.35 | 0.4 | 0.392 | 0.44 | 0.4401 | 0.46582 | 0.41502 | 0.44 | 0.43 | 0.45 | 1
| 1 | 2 | 0.444 | 0.42 | 0.432 | 0.43771 | 0.44 | 0.42 | 0.43 | 0.38 | 0.41 | 0.405 | 0.44 | 0.4358 | 0.46945 | 0.40655 | 0.43 | 0.42 | 0.44 | 5
| 1 | 3 | 0.42 | 0.42 | 0.43 | 0.43773 | 0.43 | 0.44 | 0.43 | 0.33 | 0.4 | 0.393 | 0.44 | 0.4465 | 0.46125 | 0.40481 | 0.43 | 0.43 | 0.45 | 16
| 1 | 4 | 0.438 | 0.43 | 0.438 | 0.44527 | 0.42 | 0.43 | 0.43 | 0.37 | 0.41 | 0.399 | 0.45 | 0.4413 | 0.47075 | 0.41284 | 0.43 | 0.42 | 0.44 | 0.42998
| 1 | 5 | 0.422 | 0.43 | 0.439 | 0.4385 | 0.43 | 0.42 | 0.43 | 0.34 | 0.41 | 0.392 | 0.44 | 0.4335 | 0.44952 | 0.42584 | 0.44 | 0.44 | 0.44 | 69
| 1 | 6 | 0.424 | 0.43 | 0.43 | 0.43265 | 0.43 | 0.43 | 0.43 | 0.34 | 0.4 | 0.404 | 0.44 | 0.4279 | 0.45298 | 0.41995 | 0.44 | 0.42
| Mean |  | 0.43033 | 0.425 | 0.43383 | 0.4382 | 0.43 | 0.42833 | 0.43 | 0.35167 | 0.405 | 0.3975 | 0.44167 | 0.43752 | 0.46163 | 0.41417 | 0.435 | 0.42667 | 0.444
| Median |  | 0.429 | 0.425 | 0.433 | 0.43772 | 0.43 | 0.43 | 0.43 | 0.345 | 0.405 | 0.396 | 0.44 | 0.43795 | 0.46353 | 0.41393 | 0.435 | 0.425 | 0.44
| Std Dev. |  | 0.0097502 | 0.0054772 | 0.00392 | 0.0040524 | 0.0063246 | 0.0075277 | 0 | 0.019408 | 0.0054772 | 0.0060249 | 0.0040825 | 0.0065307 | 0.008756 | 0.0079735 | 0.0054772 | 0.008165 | 0.0054772
| Rel.Std.Dev. |  | 2.266% | 1.289% | 0.9036% | 0.9248% | 1.471% | 1.757% | 0% | 5.519% | 1.352% | 1.516% | 0.9243% | 1.493% | 1.897% | 1.925% | 1.259% | 1.914% | 1.234%
| PDM3 |  | 0.0823% | -1.158% | 0.8963% | 1.911% | 0.004773% | -0.3828% | 0.004773% | -18.21% | -5.809% | -7.554% | 2.718% | 1.753% | 7.36% | -3.677% | 1.168% | -0.7705% | 3.261%
| Z-Score (Absolute) |  | 0.03 | 0.45 | 0.34 | 0.74 | 0 | 0.15 | 0 | 7.03 | 2.24 | 2.92 | 1.05 | 0.67 | 2.84 | 1.42 | 0.45 | 0.3 | 1.26
| Analytical results for Nb in OREAS 993b (Certified Value 1.87 ppm) | [18 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 2.08 | 1.8 | 2 | 2.0619 | 1.8 | 2.5 | < 5 | 1.7 | 2 | 2.2449 | 1.682 | 1.9 | [5 empty] | 1
| 1 | 2 | 1.88 | 1.8 | 2 | 2.0935 | 1.7 | 2.5 | < 5 | 1.69 | 1.8 | 1.8216 | 1.779 | 1.7 | [5 empty] | 6
| 1 | 3 | 1.78 | 1.8 | 2.1 | 2.1033 | 1.9 | 2.5 | < 5 | 1.65 | 1.8 | 1.8547 | 1.7029 | 1.8 | [5 empty] | 16
| 1 | 4 | 1.72 | 1.8 | 2 | 2.2502 | 1.7 | 2.7 | < 5 | 1.69 | 2 | 2.2956 | 1.8437 | 1.8 | [5 empty] | 1.8659
| 1 | 5 | 1.86 | 1.7 | 2 | 2.0619 | 1.7 | 2.3 | < 5 | 1.7 | 1.9 | 1.8083 | 1.7111 | 1.8 | [5 empty] | 70
| 1 | 6 | 2.19 | 1.8 | 2 | 2.0434 | 1.8 | 3.1 | < 5 | 1.76 | 1.9 | 1.7994 | 1.6748 | 1.8
| Mean |  | 1.9183 | 1.7833 | 2.0167 | 2.1024 | 1.7667 | 2.6 |  | 1.6983 | 1.9 | 1.9707 | 1.7323 | 1.8
| Median |  | 1.87 | 1.8 | 2 | 2.0777 | 1.75 | 2.5 |  | 1.695 | 1.9 | 1.8382 | 1.707 | 1.8
| Std Dev. |  | 0.18071 | 0.040825 | 0.040825 | 0.075745 | 0.08165 | 0.27568 |  | 0.035449 | 0.089443 | 0.23329 | 0.065955 | 0.063246
| Rel.Std.Dev. |  | 9.42% | 2.289% | 2.024% | 3.603% | 4.622% | 10.6% |  | 2.087% | 4.708% | 11.84% | 3.807% | 3.514%
| PDM3 |  | 2.809% | -4.426% | 8.079% | 12.67% | -5.319% | 39.34% |  | -8.981% | 1.827% | 5.618% | -7.163% | -3.532%
| Z-Score (Absolute) |  | 0.05 | 0.63 | 0.54 | 0.97 | 0.72 | 3.47 | 2.97 | 1.06 | 0.05 | 0.31 | 0.89 | 0.55
| Analytical results for Nd in OREAS 993b (Certified Value 7.32 ppm) | [18 empty] | Certified
Nd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.08: Lab; 2.09: Lab | [5 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.08: 2.12; 2.09: 2.14 | [5 empty] | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.08: 4A*MS; 2.09: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.01: 7.4; 2.02: 7.5; 2.03: 7.2288; 2.04: 6.5; 2.05: 7.9; 2.08: 9; 2.09: 6.9291 | [5 empty] | 1
| 1 | 2 | 2.01: 6.5; 2.02: 7.8; 2.03: 7.5429; 2.04: 7.1; 2.05: 7.6; 2.08: 8; 2.09: 6.9343 | [5 empty] | 0
| 1 | 3 | 2.01: 6.6; 2.02: 7.8; 2.03: 7.583; 2.04: 7.4; 2.05: 7.7; 2.08: 9; 2.09: 6.8957 | [5 empty] | 16
| 1 | 4 | 2.01: 6.6; 2.02: 8; 2.03: 7.4313; 2.04: 7.5; 2.05: 8.2; 2.08: 9; 2.09: 6.7842 | [5 empty] | 7.3189
| 1 | 5 | 2.01: 6.7; 2.02: 7.7; 2.03: 7.6153; 2.04: 7.1; 2.05: 7.8; 2.08: 9; 2.09: 6.8406 | [5 empty] | 71
| 1 | 6 | 2.01: 7.1; 2.02: 7.8; 2.03: 7.3345; 2.04: 7.1; 2.05: 8; 2.08: 8; 2.09: 6.9623
| Mean |  | 2.01: 6.8167; 2.02: 7.7667; 2.03: 7.456; 2.04: 7.1167; 2.05: 7.8667; 2.08: 8.6667; 2.09: 6.891
| Median |  | 2.01: 6.65; 2.02: 7.8; 2.03: 7.4871; 2.04: 7.1; 2.05: 7.85; 2.08: 9; 2.09: 6.9124
| Std Dev. |  | 2.01: 0.35449; 2.02: 0.1633; 2.03: 0.15235; 2.04: 0.34881; 2.05: 0.21602; 2.08: 0.5164; 2.09: 0.066902
| Rel.Std.Dev. |  | 2.01: 5.2%; 2.02: 2.103%; 2.03: 2.043%; 2.04: 4.901%; 2.05: 2.746%; 2.08: 5.958%; 2.09: 0.9709%
| PDM3 |  | 2.01: -6.863%; 2.02: 6.117%; 2.03: 1.872%; 2.04: -2.764%; 2.05: 7.484%; 2.08: 18.41%; 2.09: -5.846%
| Z-Score (Absolute) |  | 2.01: 0.73; 2.02: 0.75; 2.03: 0.26; 2.04: 0.26; 2.05: 0.9; 2.08: NA; 2.09: 0.62
| Analytical results for Ni in OREAS 993b (Certified Value 256 ppm) | [18 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 250.5 | 259 | 259 | 257.07 | 258 | 257 | 270 | 221 | 239 | 264 | 261.9 | 247.5 | 266.5506 | 229.5751 | 260 | 252 | 253 | 1
| 1 | 2 | 256.1 | 260 | 261 | 258.91 | 267 | 255 | 280 | 228 | 244 | 265.8 | 265.4 | 246.4 | 260.7707 | 232.1876 | 260 | 247 | 243 | 0
| 1 | 3 | 248.3 | 261 | 257 | 257.85 | 262 | 262 | 280 | 226 | 242 | 256.4 | 265.2 | 245.8 | 263.6982 | 220.1835 | 270 | 256 | 249 | 16
| 1 | 4 | 254.1 | 266 | 260 | 252.98 | 253 | 256 | 270 | 229 | 241 | 267.1 | 266.9 | 250.5 | 259.7594 | 235.1432 | 260 | 250 | 250 | 255.658
| 1 | 5 | 243.7 | 265 | 259 | 257.69 | 258 | 253 | 270 | 232 | 241 | 260.9 | 265.7 | 244.3 | 267.2494 | 232.745 | 260 | 260 | 248 | 72
| 1 | 6 | 244.7 | 260 | 261 | 258.25 | 263 | 258 | 260 | 227 | 243 | 270.3 | 262.7 | 242.2 | 265.241 | 233.8059 | 250 | 248
| Mean |  | 249.5667 | 261.8333 | 259.5 | 257.125 | 260.1667 | 256.8333 | 271.6667 | 227.1667 | 241.6667 | 264.0833 | 264.6333 | 246.1167 | 263.8782 | 230.6067 | 260 | 252.1667 | 248.6
| Median |  | 249.4 | 260.5 | 259.5 | 257.77 | 260 | 256.5 | 270 | 227.5 | 241.5 | 264.9 | 265.3 | 246.1 | 264.4696 | 232.4663 | 260 | 251 | 249
| Std Dev. |  | 4.9762 | 2.9269 | 1.5166 | 2.1204 | 4.8751 | 3.0605 | 7.5277 | 3.656 | 1.7512 | 4.8963 | 1.918 | 2.8266 | 3.0665 | 5.4331 | 6.3246 | 4.9967 | 3.6469
| Rel.Std.Dev. |  | 1.994% | 1.118% | 0.5844% | 0.8247% | 1.874% | 1.192% | 2.771% | 1.609% | 0.7246% | 1.854% | 0.7248% | 1.148% | 1.162% | 2.356% | 2.433% | 1.981% | 1.467%
| PDM3 |  | -2.383% | 2.415% | 1.503% | 0.5738% | 1.764% | 0.4597% | 6.262% | -11.14% | -5.473% | 3.296% | 3.511% | -3.732% | 3.215% | -9.799% | 1.698% | -1.366% | -2.761%
| Z-Score (Absolute) |  | 0.73 | 0.46 | 0.23 | 0 | 0.29 | 0.03 | 1.41 | 2.9 | 1.5 | 0.67 | 0.73 | 1.07 | 0.65 | 2.57 | 0.28 | 0.48 | 0.83
| Analytical results for P in OREAS 993b (Certified Value 0.02 wt.%) | [18 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.018 | 0.02 | 0.021 | 0.02 | 0.021 | 0.022 | 0.07 | 0.1208 | 0.018 | 0.022 | 0.0227 | 0.019342 | 0.018162 | 0.021 | 0.023 | 0.0195 |  | 1
| 1 | 2 | 0.019 | 0.02 | 0.021 | 0.02 | 0.021 | 0.02 | 0.05 | 0.1206 | 0.018 | 0.022 | 0.0237 | 0.018735 | 0.018517 | 0.02 | 0.022 | 0.0185 |  | 32
| 1 | 3 | 0.018 | 0.02 | 0.02 | 0.021 | 0.022 | 0.021 | 0.05 | 0.12 | 0.018 | 0.022 | 0.0258 | 0.018668 | 0.018206 | 0.02 | 0.022 | 0.0195 |  | 16
| 1 | 4 | 0.019 | 0.02 | 0.02 | 0.02 | 0.021 | 0.022 | 0.07 | 0.1201 | 0.019 | 0.022 | 0.0238 | 0.019765 | 0.018256 | 0.02 | 0.022 | 0.019 |  | 0.020218
| 1 | 5 | 0.018 | 0.02 | 0.021 | 0.02 | 0.021 | 0.019 | 0.05 | 0.1194 | 0.019 | 0.022 | 0.0225 | 0.019198 | 0.018432 | 0.02 | 0.023 | 0.019 |  | 73
| 1 | 6 | 0.018 | 0.02 | 0.021 | 0.02 | 0.022 | 0.017 | 0.072 | 0.1193 | 0.019 | 0.021 | 0.0219 | 0.01943 | 0.01845 | 0.02 | 0.022
| Mean |  | 0.018333 | 0.02 | 0.020667 | 0.020167 | 0.021333 | 0.020167 | 0.060333 | 0.12003 | 0.0185 | 0.021833 | 0.0234 | 0.01919 | 0.018337 | 0.020167 | 0.022333 | 0.0191
| Median |  | 0.018 | 0.02 | 0.021 | 0.02 | 0.021 | 0.0205 | 0.06 | 0.12005 | 0.0185 | 0.022 | 0.0232 | 0.01927 | 0.018344 | 0.02 | 0.022 | 0.019
| Std Dev. |  | 0.0005164 | 0 | 0.0005164 | 0.00040825 | 0.0005164 | 0.0019408 | 0.011343 | 0.00060882 | 0.00054772 | 0.00040825 | 0.0013828 | 0.00042205 | 0.0001471 | 0.00040825 | 0.0005164 | 0.00041833
| Rel.Std.Dev. |  | 2.817% | 0% | 2.499% | 2.024% | 2.421% | 9.624% | 18.8% | 0.5072% | 2.961% | 1.87% | 5.909% | 2.199% | 0.8022% | 2.024% | 2.312% | 2.19%
| PDM3 |  | -9.32% | -1.076% | 2.221% | -0.252% | 5.519% | -0.252% | 198.4% | 493.7% | -8.496% | 7.992% | 15.74% | -5.084% | -9.302% | -0.252% | 10.46% | -5.528%
| Z-Score (Absolute) |  | 0.87 | 0.08 | 0.24 | 0 | 0.56 | 0 | 19.12 | 47.53 | 0.79 | 0.79 | 1.54 | 0.46 | 0.87 | 0 | 1.03 | 0.51
| Analytical results for Pb in OREAS 993b (Certified Value 469 ppm) | [18 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 459.35 | 480 | 506 | 469.57 | 462 | 452 | 500 | 417 | 498 | 498.28 | 480 | 468 | 484.2174 | 430.8458 | 480 | 472 | 438 | 1
| 1 | 2 | 469.22 | 479 | 516 | 469.87 | 470 | 451 | 530 | 395 | 497 | 495.91 | 485 | 455 | 484.4569 | 435.6461 | 480 | 460 | 426 | 7
| 1 | 3 | 454.27 | 480 | 516 | 471.87 | 474 | 460 | 500 | 396 | 499 | 482.51 | 485 | 465 | 486.3612 | 429.7678 | 490 | 469 | 426 | 16
| 1 | 4 | 470.19 | 479 | 514 | 465.45 | 461 | 448 | 500 | 434 | 494 | 518.09 | 489 | 461 | 485.6379 | 443.2683 | 480 | 457 | 428 | 469.4268
| 1 | 5 | 451.42 | 481 | 503 | 464 | 461 | 449 | 500 | 405 | 493 | 478.26 | 482 | 451 | 487.0172 | 434.8692 | 490 | 482 | 426 | 74
| 1 | 6 | 449.82 | 482 | 500 | 466.36 | 464 | 461 | 500 | 427 | 494 | 491.45 | 461 | 443 | 485.8428 | 441.9162 | 480 | 460
| Mean |  | 459.045 | 480.1667 | 509.1667 | 467.8533 | 465.3333 | 453.5 | 505 | 412.3333 | 495.8333 | 494.0833 | 480.3333 | 457.1667 | 485.5889 | 436.0522 | 483.3333 | 466.6667 | 428.8
| Median |  | 456.81 | 480 | 510 | 467.965 | 463 | 451.5 | 500 | 411 | 495.5 | 493.68 | 483.5 | 458 | 485.7404 | 435.2576 | 480 | 464.5 | 426
| Std Dev. |  | 8.8755 | 1.169 | 7.0545 | 3.0332 | 5.4283 | 5.6125 | 12.2474 | 16.2931 | 2.4833 | 14.0607 | 9.9532 | 9.347 | 1.0831 | 5.561 | 5.164 | 9.5009 | 5.2154
| Rel.Std.Dev. |  | 1.933% | 0.2435% | 1.386% | 0.6483% | 1.167% | 1.238% | 2.425% | 3.951% | 0.5008% | 2.846% | 2.072% | 2.045% | 0.2231% | 1.275% | 1.068% | 2.036% | 1.216%
| PDM3 |  | -2.212% | 2.288% | 8.466% | -0.3352% | -0.872% | -3.393% | 7.578% | -12.16% | 5.625% | 5.252% | 2.323% | -2.612% | 3.443% | -7.11% | 2.962% | -0.588% | -8.655%
| Z-Score (Absolute) |  | 0.38 | 0.54 | 1.8 | 0 | 0.11 | 0.63 | 1.62 | 2.42 | 1.22 | 1.14 | 0.54 | 0.47 | 0.77 | 1.39 | 0.67 | 0.05 | 1.7
| Analytical results for Pr in OREAS 993b (Certified Value 1.94 ppm) | [18 empty] | Certified
Pr | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.06: 2.14 | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.01: 2; 2.02: 2.17; 2.03: 2.142; 2.04: 1.8; 2.05: 1.9; 2.06: 1.7204 | 1
| 1 | 2 | 2.01: 1.6; 2.02: 2.04; 2.03: 2.1634; 2.04: 1.8; 2.05: 1.9; 2.06: 1.6851 | 0
| 1 | 3 | 2.01: 1.7; 2.02: 2.14; 2.03: 2.1407; 2.04: 1.9; 2.05: 1.9; 2.06: 1.7329 | 16
| 1 | 4 | 2.01: 1.7; 2.02: 2.16; 2.03: 2.2946; 2.04: 1.8; 2.05: 2.1; 2.06: 1.7531 | 1.9361
| 1 | 5 | 2.01: 1.7; 2.02: 2.09; 2.03: 2.2031; 2.04: 1.9; 2.05: 1.9; 2.06: 1.7779 | 75
| 1 | 6 | 2.01: 1.9; 2.02: 2.17; 2.03: 2.1321; 2.04: 1.9; 2.05: 2; 2.06: 1.7833
| Mean |  | 2.01: 1.7667; 2.02: 2.1283; 2.03: 2.1793; 2.04: 1.85; 2.05: 1.95; 2.06: 1.7421
| Median |  | 2.01: 1.7; 2.02: 2.15; 2.03: 2.1527; 2.04: 1.85; 2.05: 1.9; 2.06: 1.743
| Std Dev. |  | 2.01: 0.15055; 2.02: 0.052694; 2.03: 0.061986; 2.04: 0.054772; 2.05: 0.083666; 2.06: 0.037174
| Rel.Std.Dev. |  | 2.01: 8.522%; 2.02: 2.476%; 2.03: 2.844%; 2.04: 2.961%; 2.05: 4.291%; 2.06: 2.134%
| PDM3 |  | 2.01: -8.75%; 2.02: 9.93%; 2.03: 12.56%; 2.04: -4.446%; 2.05: 0.7194%; 2.06: -10.02%
| Z-Score (Absolute) |  | 2.01: 0.62; 2.02: 1.06; 2.03: 1.29; 2.04: 0.23; 2.05: 0.23; 2.06: 0.73
| Analytical results for Rb in OREAS 993b (Certified Value 22.9 ppm) | [18 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 22.3 | 24.3 | 22 | 23.2247 | 23 | 22.4 | 24.4 | 34.6 | 22.8144 | 21.318 | 23.7 | [6 empty] | 1
| 1 | 2 | 21.9 | 24.4 | 21.4 | 23.6134 | 22.5 | 24.2 | 24.1 | 35.4 | 22.4163 | 20.9067 | 22 | [6 empty] | 8
| 1 | 3 | 20.9 | 23.9 | 21.6 | 23.1637 | 24.5 | 25.1 | 23.3 | 36.5 | 22.9284 | 21.0288 | 22.9 | [6 empty] | 16
| 1 | 4 | 21.6 | 24.3 | 22.3 | 24.2585 | 22.1 | 22.7 | 23.9 | 36.5 | 22.8903 | 21.9049 | 22.3 | [6 empty] | 22.8846
| 1 | 5 | 21.7 | 23.9 | 21.5 | 23.3109 | 22.9 | 22.9 | 24.5 | 36.8 | 22.7479 | 21.9257 | 23 | [6 empty] | 76
| 1 | 6 | 23.9 | 23.6 | 21.6 | 23.2235 | 22.9 | 22.9 | 24.6 | 35.4 | 22.4635 | 22.1383 | 22.9
| Mean |  | 22.05 | 24.0667 | 21.7333 | 23.4658 | 22.9833 | 23.3667 | 24.1333 | 35.8667 | 22.7101 | 21.5371 | 22.8
| Median |  | 21.8 | 24.1 | 21.6 | 23.2678 | 22.9 | 22.9 | 24.25 | 35.95 | 22.7812 | 21.6114 | 22.9
| Std Dev. |  | 1.0154 | 0.31411 | 0.34448 | 0.42006 | 0.81588 | 1.0501 | 0.48442 | 0.86178 | 0.2189 | 0.5199 | 0.5933
| Rel.Std.Dev. |  | 4.605% | 1.305% | 1.585% | 1.79% | 3.55% | 4.494% | 2.007% | 2.403% | 0.9639% | 2.414% | 2.602%
| PDM3 |  | -3.647% | 5.165% | -5.031% | 2.539% | 0.4313% | 2.106% | 5.457% | 56.73% | -0.7624% | -5.889% | -0.3698%
| Z-Score (Absolute) |  | 0.67 | 0.78 | 0.9 | 0.35 | 0 | 0.28 | 0.83 | 9.31 | 0.2 | 1.04 | 0.13
| Analytical results for Re in OREAS 993b (Certified Value 4.54 ppm) | [18 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 4.991 | 4.99 | 4.6 | 4.01 | 5.29 | 5.009 | > 1 | 3.596 | 5.035 | 3.5591 | 4.43 | [6 empty] | 1
| 1 | 2 | 4.823 | 4.97 | 4.61 | 3.88 | 5.11 | 4.906 | > 1 | 3.706 | 5.0156 | 3.5517 | 3.62 | [6 empty] | 9
| 1 | 3 | 4.73 | 5.08 | 4.75 | 4.84 | 5.04 | 4.915 | > 1 | 3.844 | 5.0428 | 3.5832 | 4.21 | [6 empty] | 16
| 1 | 4 | 4.816 | 5.29 | 4.89 | 3.38 | 5.22 | 5.23 | > 1 | 3.867 | 5.0301 | 3.8118 | 3.99 | [6 empty] | 4.5386
| 1 | 5 | 4.538 | 5.11 | 4.77 | 3.84 | 5.1 | 4.837 | > 1 | 3.716 | 5.0415 | 3.7737 | 3.72 | [6 empty] | 77
| 1 | 6 | 5.212 | 5.19 | 4.73 | 4.02 | 5.31 | 5.152 | > 1 | 3.405 | 5.0175 | 3.7216 | 4.85
| Mean |  | 2.01: 4.8517; 2.02: 5.105; 2.03: 4.725; 2.04: 3.995; 2.05: 5.1783; 2.06: 5.0082; 2.08: 3.689; 2.09: 5.0304; 2.10: 3.6669; 2.11: 4.1367
| Median |  | 2.01: 4.8195; 2.02: 5.095; 2.03: 4.74; 2.04: 3.945; 2.05: 5.165; 2.06: 4.962; 2.08: 3.711; 2.09: 5.0326; 2.10: 3.6524; 2.11: 4.1
| Std Dev. |  | 2.01: 0.23002; 2.02: 0.12128; 2.03: 0.1084; 2.04: 0.47547; 2.05: 0.11089; 2.06: 0.15383; 2.08: 0.1709; 2.09: 0.01167; 2.10: 0.116; 2.11: 0.46094
| Rel.Std.Dev. |  | 2.01: 4.741%; 2.02: 2.376%; 2.03: 2.294%; 2.04: 11.9%; 2.05: 2.141%; 2.06: 3.072%; 2.08: 4.633%; 2.09: 0.232%; 2.10: 3.163%; 2.11: 11.14%
| PDM3 |  | 2.01: 6.898%; 2.02: 12.48%; 2.03: 4.107%; 2.04: -11.98%; 2.05: 14.1%; 2.06: 10.35%; 2.08: -18.72%; 2.09: 10.84%; 2.10: -19.21%; 2.11: -8.856%
| Z-Score (Absolute) |  | 0.12 | 0.6 | 0.12 | 1.51 | 0.74 | 0.42 | NA | 2.1 | 0.46 | 2.14 | 1.24
| Analytical results for S in OREAS 993b (Certified Value 27.19 wt.%) | [18 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS |  |  | 2
3
| 1 | 1 | 25.09 | 27.9 | 31.3843 | 24.2 | 26.2 | 25.9 | 30.1 | 29.94 | > 10 | > 10 | 27.6125 | 28.5954 | 23.8357 | 25.9 | 26.1 |  |  | 1
| 1 | 2 | 25.84 | 27.7 | 31.3433 | 26.7 | 26 | 26.4 | 29 | 29.71 | > 10 | > 10 | 28.2669 | 29.5903 | 23.2825 | 25.1 | 25.9 |  |  | 34
| 1 | 3 | 25.56 | 27.7 | 31.34 | 25.8 | 26.6 | 25.7 | 29.7 | 29.98 | > 10 | > 10 | 28.0571 | 28.6015 | 23.9278 | 25.7 | 26.2 |  |  | 16
| 1 | 4 | 23.65 | 28 | 31.3823 | 26.3 | 25.1 | 25.7 | 29.7 | 29.69 | > 10 | > 10 | 28.0169 | 29.3016 | 23.979 | 25.6 | 25.9 |  |  | 27.1863
| 1 | 5 | 25.75 | 28 | 31.3951 | 24.4 | 26.6 | 26.6 | 28.8 | 29.92 | > 10 | > 10 | 27.2543 | 28.8478 | 23.7141 | 26.2 | 25.6 |  |  | 78
| 1 | 6 | 25.53 | 27.7 | 31.3286 | 25.1 | 25.8 | 25.7 | 29.5 | 29.87 | > 10 | > 10 | 27.1098 | 29.1679 | 23.8701 | 25.8 | 26.2
| Mean |  | 25.2367 | 27.8333 | 31.3623 | 25.4167 | 26.05 | 26 | 29.4667 | 29.8517 |  |  | 27.7196 | 29.0174 | 23.7682 | 25.7167 | 25.9833
| Median |  | 25.545 | 27.8 | 31.3628 | 25.45 | 26.1 | 25.8 | 29.6 | 29.895 |  |  | 27.8147 | 29.0078 | 23.8529 | 25.75 | 26
| Std Dev. |  | 0.81941 | 0.15055 | 0.028121 | 1.0187 | 0.5648 | 0.4 | 0.48442 | 0.12287 |  |  | 0.4694 | 0.40287 | 0.25443 | 0.3656 | 0.23166
| Rel.Std.Dev. |  | 3.247% | 0.5409% | 0.08966% | 4.008% | 2.168% | 1.538% | 1.644% | 0.4116% |  |  | 1.693% | 1.388% | 1.07% | 1.422% | 0.8916%
| PDM3 |  | -7.172% | 2.38% | 15.36% | -6.509% | -4.18% | -4.364% | 8.388% | 9.804% |  |  | 1.961% | 6.735% | -12.57% | -5.406% | -4.425%
| Z-Score (Absolute) |  | 0.33 | 0.72 | 2.15 | 0.26 | 0 | 0.02 | 1.38 | 1.54 | NA | NA | 0.67 | 1.2 | 0.92 | 0.13 | 0.03
| Analytical results for Sb in OREAS 993b (Certified Value 49.1 ppm) | [18 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 57.45 | 54 | 54.8 | 46.7258 | 57.1 | 46 | 50 | 7.7 | 41 | 52.22 | 45.2 | 45.49 | 52.661 | 43.941 | 60 | 51.3 |  | 1
| 1 | 2 | 49.7 | 53.5 | 54.5 | 47.3462 | 55.3 | 47 | 50 | 10.4 | 42 | 51.29 | 46.2 | 41.42 | 52.6581 | 44.1064 | 60 | 50.5 |  | 10
| 1 | 3 | 49.34 | 54.6 | 55.3 | 46.9003 | 55.9 | 50 | 40 | 23.2 | 42 | 50.06 | 46.5 | 44.33 | 52.1354 | 41.7909 | 50 | 49.6 |  | 16
| 1 | 4 | 50.09 | 54.3 | 53.6 | 46.7168 | 54.3 | 44 | 50 | 6.4 | 41 | 52.22 | 45.6 | 42.97 | 52.8175 | 45.5397 | 50 | 50.5 |  | 49.1187
| 1 | 5 | 49.1 | 54.1 | 54.8 | 47.2035 | 55.7 | 47 | 40 | 38.5 | 43 | 50.54 | 46 | 41.8 | 52.0603 | 44.961 | 50 | 48.4 |  | 79
| 1 | 6 | 53.92 | 53.9 | 54.6 | 46.8518 | 56.9 | 48 | 40 | 7.9 | 41 | 52.66 | 43.4 | 39.73 | 52.6337 | 45.3795 | 60 | 52.1
| Mean |  | 51.6 | 54.0667 | 54.6 | 46.9574 | 55.8667 | 47 | 45 | 15.6833 | 41.6667 | 51.4983 | 45.4833 | 42.6233 | 52.4943 | 44.2864 | 55 | 50.4
| Median |  | 49.895 | 54.05 | 54.7 | 46.876 | 55.8 | 47 | 45 | 9.15 | 41.5 | 51.755 | 45.8 | 42.385 | 52.6459 | 44.5337 | 55 | 50.5
| Std Dev. |  | 3.372 | 0.37238 | 0.56214 | 0.25987 | 1.0386 | 2 | 5.4772 | 12.7711 | 0.8165 | 1.0413 | 1.1179 | 2.0856 | 0.31486 | 1.3856 | 5.4772 | 1.2931
| Rel.Std.Dev. |  | 6.535% | 0.6887% | 1.03% | 0.5534% | 1.859% | 4.255% | 12.17% | 81.43% | 1.96% | 2.022% | 2.458% | 4.893% | 0.5998% | 3.129% | 9.959% | 2.566%
| PDM3 |  | 5.052% | 10.07% | 11.16% | -4.4% | 13.74% | -4.313% | -8.385% | -68.07% | -15.17% | 4.845% | -7.401% | -13.22% | 6.872% | -9.838% | 11.97% | 2.609%
| Z-Score (Absolute) |  | 0.48 | 0.88 | 0.97 | 0.29 | 1.18 | 0.28 | NA | 5.42 | 1.16 | 0.46 | 0.53 | 1 | 0.62 | 0.73 | NA | 0.28
Results from laboratories 2.07 and 2.15 were removed due to their 10 ppm reading resolution.
| Analytical results for Sc in OREAS 993b (Certified Value 5.75 ppm) | [18 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 6.5 | 6 | 6.03 | 5.4 | 5 | 5 | 6 | 6 | 5.9 | 11 | 6.2554 | 6.0659 | 5 | 6.1 |  |  |  | 1
| 1 | 2 | 6.7 | 6 | 6.02 | 5.2 | 5 | 5 | 6 | 6 | 5.9 | 11.1 | 6.1233 | 5.9667 | 5 | 5.7 |  |  |  | 11
| 1 | 3 | 6.4 | 6 | 5.99 | 5.9 | 5 | 6 | 5 | 6 | 5.7 | 11.1 | 6.3075 | 5.8498 | 5 | 5.5 |  |  |  | 16
| 1 | 4 | 6.6 | 6 | 6.02 | 5.3 | 5 | 5 | 6 | 6 | 6 | 11.1 | 6.3897 | 6.2283 | 5 | 5.6 |  |  |  | 5.7534
| 1 | 5 | 6.3 | 6 | 5.98 | 5.2 | 5 | 5 | 6 | 6 | 5.9 | 11 | 6.2744 | 6.2048 | 5 | 5.4 |  |  |  | 80
| 1 | 6 | 6.3 | 6 | 6.02 | 5.3 | 5 | 5 | 6 | 6 | 6.2 | 11 | 6.0641 | 6.2786 | 5 | 5.7
| Mean |  | 6.4667 | 6 | 6.01 | 5.3833 | 5 | 5.1667 | 5.8333 | 6 | 5.9333 | 11.05 | 6.2357 | 6.099 | 5 | 5.6667
| Median |  | 6.45 | 6 | 6.02 | 5.3 | 5 | 5 | 6 | 6 | 5.9 | 11.05 | 6.2649 | 6.1353 | 5 | 5.65
| Std Dev. |  | 0.1633 | 0 | 0.02 | 0.26394 | 0 | 0.40825 | 0.40825 | 0 | 0.1633 | 0.054772 | 0.12067 | 0.16784 | 0 | 0.24221
| Rel.Std.Dev. |  | 2.525% | 0% | 0.3328% | 4.903% | 0% | 7.902% | 6.999% | 0% | 2.752% | 0.4957% | 1.935% | 2.752% | 0% | 4.274%
| PDM3 |  | 12.4% | 4.285% | 4.459% | -6.433% | -13.1% | -10.2% | 1.389% | 4.285% | 3.127% | 92.06% | 8.383% | 6.006% | -13.1% | -1.508%
| Z-Score (Absolute) |  | 1.18 | 0.08 | 0.1 | 1.38 | 2.29 | 1.9 | 0.32 | 0.08 | 0.08 | 12.05 | 0.64 | 0.31 | 2.29 | 0.71
| Analytical results for Se in OREAS 993b (Certified Value 189 ppm) | [18 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.08 | 2.09 | 2.10 | 2.12 | 2.13 | 2.14 | 2.16 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | [6 empty] | 0
0
| 1 | 1 | 179 | 194 | 211 | 186 | 207 | 152 | 196.6 | 209.84 | 190.319 | 176.3998 | 192 | [6 empty] | 1
| 1 | 2 | 185.1 | 190 | 210 | 178 | 189 | 156 | 193.6 | 226.35 | 191.4472 | 177.1307 | 172 | [6 empty] | 1
| 1 | 3 | 178.6 | 193 | 205 | 208 | 195 | 157 | 191.2 | 222.17 | 190.0333 | 176.279 | 179 | [6 empty] | 16
| 1 | 4 | 182.9 | 197 | 200 | 175 | 207 | 156 | 201.4 | 224.69 | 188.9751 | 174.4409 | 172 | [6 empty] | 189.3073
| 1 | 5 | 175.1 | 201 | 202 | 180 | 181 | 153 | 198.2 | 222.16 | 192.6954 | 175.1448 | 166 | [6 empty] | 81
| 1 | 6 | 175.2 | 190 | 204 | 188 | 205 | 155 | 204.1 | 220.66 | 191.771 | 175.7728 | 201
| Mean |  | 179.3167 | 194.1667 | 205.3333 | 185.8333 | 197.3333 | 154.8333 | 197.5167 | 220.9783 | 190.8735 | 175.8614 | 180.3333
| Median |  | 178.8 | 193.5 | 204.5 | 183 | 200 | 155.5 | 197.4 | 222.165 | 190.8831 | 176.0259 | 175.5
| Std Dev. |  | 4.0405 | 4.2622 | 4.3665 | 11.9066 | 10.8382 | 1.9408 | 4.7952 | 5.8243 | 1.3469 | 0.96068 | 13.4858
| Rel.Std.Dev. |  | 2.253% | 2.195% | 2.127% | 6.407% | 5.492% | 1.253% | 2.428% | 2.636% | 0.7057% | 0.5463% | 7.478%
| PDM3 |  | -5.277% | 2.567% | 8.466% | -1.835% | 4.24% | -18.21% | 4.337% | 16.73% | 0.8273% | -7.103% | -4.74%
| Z-Score (Absolute) |  | 0.74 | 0.21 | 0.93 | 0.32 | 0.41 | 2.31 | 0.42 | 1.93 | 0 | 0.96 | 0.67
| Analytical results for Sm in OREAS 993b (Indicative Value 1.54 ppm) | [18 empty] | Indicative
Sm | Round | Replicate | 2.01: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab |  |  |  | 1
| No | No | 2.01: 2.01; 2.03: 2.02; 2.04: 2.04; 2.05: 2.08; 2.06: 2.10; 2.07: 2.14 |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.03: 4A*MS; 2.04: 4A*MS; 2.05: 4A*OES/MS; 2.06: 4A*OES/MS; 2.07: 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.01: 1.3; 2.03: 1.3; 2.04: 1.5152; 2.05: 1.6; 2.06: 1.6; 2.07: 1.3038 |  |  |  | 1
| 1 | 2 | 2.01: 1.2; 2.03: 1.3; 2.04: 1.5562; 2.05: 1.4; 2.06: 1.5; 2.07: 1.302 |  |  |  | 21
| 1 | 3 | 2.01: 1.3; 2.03: 1.6; 2.04: 1.5511; 2.05: 1.6; 2.06: 1.5; 2.07: 1.3059 |  |  |  | 16
| 1 | 4 | 2.01: 1.2; 2.03: 1.7; 2.04: 1.5279; 2.05: 1.7; 2.06: 1.6; 2.07: 1.3022 |  |  |  | 1.5417
| 1 | 5 | 2.01: 1.2; 2.03: 1.6; 2.04: 1.5408; 2.05: 1.4; 2.06: 1.6; 2.07: 1.3138 |  |  |  | 27
| 1 | 6 | 2.01: 1.3; 2.03: 1.5; 2.04: 1.51; 2.05: 1.7; 2.06: 1.6; 2.07: 1.3179
| Mean |  | 2.01: 1.25; 2.03: 1.5; 2.04: 1.5335; 2.05: 1.5667; 2.06: 1.5667; 2.07: 1.3076
| Median |  | 2.01: 1.25; 2.03: 1.55; 2.04: 1.5344; 2.05: 1.6; 2.06: 1.6; 2.07: 1.3048
| Std Dev. |  | 2.01: 0.054772; 2.03: 0.16733; 2.04: 0.018967; 2.05: 0.13663; 2.06: 0.05164; 2.07: 0.0066873
| Rel.Std.Dev. |  | 2.01: 4.382%; 2.03: 11.16%; 2.04: 1.237%; 2.05: 8.721%; 2.06: 3.296%; 2.07: 0.5114%
| PDM3 |  | 2.01: -18.92%; 2.03: -2.706%; 2.04: -0.5313%; 2.05: 1.618%; 2.06: 1.618%; 2.07: -15.19%
| Z-Score (Absolute) |  | 2.01: 3.6; 2.03: 0.23; 2.04: 0.23; 2.05: 0.67; 2.06: 0.67; 2.07: 2.83
| Analytical results for Sn in OREAS 993b (Certified Value 1.96 ppm) | [18 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | 5.01 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | [5 empty] | 2
3
| 1 | 1 | 2.1 | 2 | 2 | 1.7 | 2 | < 10 | 1.8 | 2.1 | 1.9841 | < 2.5 | 2.3 | 6 | [5 empty] | 1
| 1 | 2 | 1.8 | 2 | 2 | 1.7 | 2 | < 10 | 1.8 | 2.1 | 1.8841 | < 2.5 | 2.1 | 6 | [5 empty] | 25
| 1 | 3 | 1.8 | 2 | 2 | 1.8 | 2 | < 10 | 1.7 | 2.2 | 1.8924 | < 2.5 | 2.8 | 6 | [5 empty] | 16
| 1 | 4 | 1.9 | 2 | 2 | 1.7 | 2 | < 10 | 1.8 | 2 | 1.9638 | < 2.5 | 2 | 8 | [5 empty] | 1.9583
| 1 | 5 | 1.8 | 2 | 2.1 | 1.8 | 2 | < 10 | 1.8 | 2.3 | 1.9087 | < 2.5 | 1.9 | 6 | [5 empty] | 82
| 1 | 6 | 2 | 2 | 2 | 1.8 | 2 | < 10 | 1.8 | 2 | 1.993 | < 2.5 | 2.3
| Mean |  | 1.9 | 2 | 2.0167 | 1.75 | 2 |  | 1.7833 | 2.1167 | 1.9377 |  | 2.2333 | 6.4
| Median |  | 1.85 | 2 | 2 | 1.75 | 2 |  | 1.8 | 2.1 | 1.9362 |  | 2.2 | 6
| Std Dev. |  | 0.12649 | 0 | 0.040825 | 0.054772 | 0 |  | 0.040825 | 0.1169 | 0.048312 |  | 0.32042 | 0.89443
| Rel.Std.Dev. |  | 6.657% | 0% | 2.024% | 3.13% | 0% |  | 2.289% | 5.523% | 2.493% |  | 14.35% | 13.98%
| PDM3 |  | -2.975% | 2.132% | 2.983% | -10.63% | 2.132% |  | -8.933% | 8.089% | -1.051% |  | 14.05% | 226.8%
| Z-Score (Absolute) |  | 0.4 | 0 | 0.07 | 1.01 | 0 | 12.14 | 0.88 | 0.47 | 0.25 | 3.03 | 0.94 | 17.8
| Analytical results for Sr in OREAS 993b (Certified Value 159 ppm) | [18 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 164 | 166.7 | 154 | 166.1878 | 158.5 | 159 | 200 | 147 | 158 | 157 | 170 | 157.423 | 153.3869 | 200 | 161.5 |  |  | 1
| 1 | 2 | 167 | 166.4 | 156 | 165.2093 | 160.5 | 156 | 200 | 144 | 159 | 155 | 168 | 156.2215 | 150.0503 | 200 | 154 |  |  | 2
| 1 | 3 | 158 | 166.9 | 154 | 166.4556 | 166.5 | 161 | 200 | 155 | 159 | 149 | 170 | 158.0518 | 150.11 | 200 | 161 |  |  | 16
| 1 | 4 | 167 | 166.6 | 152 | 165.8891 | 155 | 157 | 200 | 154 | 160 | 159 | 170 | 158.2285 | 155.3858 | 200 | 157.5 |  |  | 159.4065
| 1 | 5 | 160 | 165 | 157 | 165.6316 | 158 | 157 | 200 | 140 | 157 | 157 | 168 | 156.4132 | 155.5368 | 100 | 163.5 |  |  | 83
| 1 | 6 | 159 | 166.3 | 156 | 166.6925 | 161 | 160 | 200 | 149 | 158 | 159 | 168 | 155.736 | 154.9991 | 200 | 159
| Mean |  | 162.5 | 166.3167 | 154.8333 | 166.011 | 159.9167 | 158.3333 | 200 | 148.1667 | 158.5 | 156 | 169 | 157.0124 | 153.2448 | 183.3333 | 159.4167
| Median |  | 162 | 166.5 | 155 | 166.0385 | 159.5 | 158 | 200 | 148 | 158.5 | 157 | 169 | 156.9181 | 154.193 | 200 | 160
| Std Dev. |  | 4.0373 | 0.67946 | 1.8348 | 0.54679 | 3.8654 | 1.9664 | 0 | 5.7764 | 1.0488 | 3.7417 | 1.0954 | 1.0336 | 2.5678 | 40.8248 | 3.3677
| Rel.Std.Dev. |  | 2.485% | 0.4085% | 1.185% | 0.3294% | 2.417% | 1.242% | 0% | 3.899% | 0.6617% | 2.398% | 0.6482% | 0.6583% | 1.676% | 22.27% | 2.113%
| PDM3 |  | 1.941% | 4.335% | -2.869% | 4.143% | 0.32% | -0.6732% | 25.47% | -7.051% | -0.5687% | -2.137% | 6.018% | -1.502% | -3.865% | 15.01% | 0.006362%
| Z-Score (Absolute) |  | 0.74 | 1.44 | 0.67 | 1.38 | 0.26 | 0.03 | NA | 1.9 | 0 | 0.46 | 1.93 | 0.27 | 0.97 | NA | 0.17
Results from laboratories 2.07 and 2.15 were removed due to their 100 ppm reading resolution.
| Analytical results for Ta in OREAS 993b (Certified Value 0.099 ppm) | [18 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.1 | 0.2 | 0.1 | 0.12764 | 0.09 | < 0.1 | 9 | < 0.1 | 0.2 | 0.085165 | 0.096114 | 0.09 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.1 | 0.1 | 0.11744 | 0.09 | 0.1 | 9 | < 0.1 | 0.2 | 0.086434 | 0.10063 | 0.09 | [5 empty] | 3
| 1 | 3 | 0.1 | 0.1 | 0.1 | 0.1092 | 0.09 | 0.1 | 9 | < 0.1 | 0.2 | 0.084799 | 0.098906 | 0.09 | [5 empty] | 16
| 1 | 4 | 0.1 |  | 0.1 | 0.12024 | 0.09 | < 0.1 | 9 | < 0.1 | 0.2 | 0.091686 | 0.10867 | 0.09 | [5 empty] | 0.098687
| 1 | 5 | 0.2 | 0.1 | 0.1 | 0.1376 | 0.09 | 0.1 | 9 | 0.1 | 0.2 | 0.084969 | 0.097117 | 0.09 | [5 empty] | 84
| 1 | 6 | 0.2 | 0.1 | 0.1 | 0.11148 | 0.09 | < 0.1 | 9 | < 0.1 | 0.2 | 0.092869 | 0.098121 | 0.09
| Mean |  | 0.15 | 0.12 | 0.1 | 0.1206 | 0.09 | 0.1 | 9 | 0.1 | 0.2 | 0.087654 | 0.099925 | 0.09
| Median |  | 0.15 | 0.1 | 0.1 | 0.11884 | 0.09 | 0.1 | 9 | 0.1 | 0.2 | 0.085799 | 0.098513 | 0.09
| Std Dev. |  | 0.054772 | 0.044721 | 1.5202e-17 | 0.0106 | 1.5202e-17 | 1.6997e-17 | 0 |  | 3.0405e-17 | 0.003647 | 0.0045524 | 1.5202e-17
| Rel.Std.Dev. |  | 36.51% | 37.27% | 1.52e-14% | 8.789% | 1.689e-14% | 1.7e-14% | 0% |  | 1.52e-14% | 4.161% | 4.556% | 1.689e-14%
| PDM3 |  | 52% | 21.6% | 1.331% | 22.21% | -8.802% | 1.331% | 9020% | 1.331% | 102.7% | -11.18% | 1.255% | -8.802%
| Z-Score (Absolute) |  | 1.68 | 0.67 | 0 | 0.69 | 0.33 | 0.84 | NA | 1.4 | 3.36 | 0.41 | 0 | 0.33
Results from laboratory 2.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 993b (Certified Value 0.19 ppm) | [18 empty] | Certified
Tb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab | [5 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.08; 2.04: 2.10; 2.05: 2.14 | [5 empty] | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS; 2.05: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.01: 0.2; 2.02: 0.19; 2.03: 0.2; 2.04: 0.2; 2.05: 0.16854 | [5 empty] | 1
| 1 | 2 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.16562 | [5 empty] | 0
| 1 | 3 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.17216 | [5 empty] | 16
| 1 | 4 | 2.01: 0.2; 2.02: 0.19; 2.03: 0.2; 2.04: 0.2; 2.05: 0.16903 | [5 empty] | 0.19272
| 1 | 5 | 2.01: 0.2; 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.1703 | [5 empty] | 85
| 1 | 6 | 2.01: 0.2; 2.02: 0.19; 2.03: 0.2; 2.04: 0.2; 2.05: 0.16606
| Mean |  | 2.01: 0.2; 2.02: 0.195; 2.03: 0.2; 2.04: 0.2; 2.05: 0.16862
| Median |  | 2.01: 0.2; 2.02: 0.195; 2.03: 0.2; 2.04: 0.2; 2.05: 0.16879
| Std Dev. |  | 2.01: 3.0405e-17; 2.02: 0.0054772; 2.03: 3.0405e-17; 2.04: 3.0405e-17; 2.05: 0.0024938
| Rel.Std.Dev. |  | 2.01: 1.52e-14%; 2.02: 2.809%; 2.03: 1.52e-14%; 2.04: 1.52e-14%; 2.05: 1.479%
| PDM3 |  | 2.01: 3.775%; 2.02: 1.181%; 2.03: 3.775%; 2.04: 3.775%; 2.05: -12.51%
Z-Score (Absolute) | bar: 2.01 to 2.05
| Analytical results for Te in OREAS 993b (Certified Value 9.94 ppm) | [18 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 9.31 | 9.8 | 9.9 | 9.0731 | 11.1 | 9.7 | 38 | 10.81 | 9.2 | 11.22 | 10.6138 | 8.7391 | 10.7 |  |  |  |  | 1
| 1 | 2 | 9.56 | 9.5 | 10 | 9.1082 | 10.5 | 10.9 | 41 | 10.61 | 8.9 | 10.07 | 9.2994 | 8.2941 | 10.75 |  |  |  |  | 4
| 1 | 3 | 9.51 | 10.2 | 10.2 | 9.495 | 12.05 | 11.8 | 40 | 10.55 | 8.7 | 10.22 | 9.3442 | 8.0587 | 10.55 |  |  |  |  | 16
| 1 | 4 | 9.6 | 10.8 | 9.9 | 9.3115 | 10.3 | 8.2 | 41 | 10.91 | 7.8 | 9.38 | 10.7967 | 8.6111 | 10.5 |  |  |  |  | 9.9405
| 1 | 5 | 9.4 | 10.8 | 10.5 | 9.257 | 11.05 | 12.3 | 39 | 10.66 | 8.1 | 9.99 | 9.5441 | 8.615 | 10.1 |  |  |  |  | 86
| 1 | 6 | 10.53 | 11.1 | 10.1 | 9.5455 | 11.2 | 8.9 | 40 | 10.95 | 9.1 | 10.07 | 9.5397 | 8.517 | 11.4
| Mean |  | 9.6517 | 10.3667 | 10.1 | 9.2984 | 11.0333 | 10.3 | 39.8333 | 10.7483 | 8.6333 | 10.1583 | 9.8563 | 8.4725 | 10.6667
| Median |  | 9.535 | 10.5 | 10.05 | 9.2843 | 11.075 | 10.3 | 40 | 10.735 | 8.8 | 10.07 | 9.5419 | 8.5641 | 10.625
| Std Dev. |  | 0.44332 | 0.63456 | 0.22804 | 0.19418 | 0.61455 | 1.6334 | 1.169 | 0.16546 | 0.56451 | 0.59677 | 0.66759 | 0.25124 | 0.42622
| Rel.Std.Dev. |  | 4.593% | 6.121% | 2.258% | 2.088% | 5.57% | 15.86% | 2.935% | 1.539% | 6.539% | 5.875% | 6.773% | 2.965% | 3.996%
| PDM3 |  | -2.905% | 4.288% | 1.605% | -6.459% | 10.99% | 3.617% | 300.7% | 8.127% | -13.15% | 2.192% | -0.8464% | -14.77% | 7.306%
| Z-Score (Absolute) |  | 0.63 | 0.32 | 0.04 | 1.1 | 1.2 | 0.23 | NA | 0.82 | 1.99 | 0.04 | 0.36 | 2.2 | 0.71
Results from laboratory 2.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 993b (Certified Value 2.05 ppm) | [18 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 2.1 | 2.1 | 2.13 | 2.4314 | 1.99 | < 20 | 2 | 2.1 | 3 | 2.1546 | 1.9202 | 2.04 | [5 empty] | 1
| 1 | 2 | 1.9 | 2.1 | 2.1 | 2.3968 | 1.94 | < 20 | 2.1 | 2 | 2.9 | 2.1439 | 1.9179 | 1.94 | [5 empty] | 0
| 1 | 3 | 1.9 | 2.1 | 2.12 | 2.3945 | 2.23 | < 20 | 2.2 | 2 | 2.9 | 2.136 | 1.9041 | 2.12 | [5 empty] | 16
| 1 | 4 | 1.9 | 2.2 | 2.1 | 2.5061 | 1.94 | < 20 | 2 | 2.2 | 3 | 2.1497 | 1.9731 | 2.02 | [5 empty] | 2.0487
| 1 | 5 | 1.9 | 2.3 | 2.12 | 2.4532 | 1.98 | < 20 | 2.1 | 2 | 3 | 2.1287 | 1.9361 | 2.04 | [5 empty] | 87
| 1 | 6 | 2.5 | 2.1 | 2.11 | 2.3616 | 1.94 | < 20 | 2 | 2.1 | 2.8 | 2.136 | 1.9093 | 2.02
| Mean |  | 2.0333 | 2.15 | 2.1133 | 2.4239 | 2.0033 |  | 2.0667 | 2.0667 | 2.9333 | 2.1415 | 1.9268 | 2.03
| Median |  | 1.9 | 2.1 | 2.115 | 2.4141 | 1.96 |  | 2.05 | 2.05 | 2.95 | 2.14 | 1.919 | 2.03
| Std Dev. |  | 0.24221 | 0.083666 | 0.012111 | 0.051287 | 0.11325 |  | 0.08165 | 0.08165 | 0.08165 | 0.0096821 | 0.025197 | 0.057619
| Rel.Std.Dev. |  | 11.91% | 3.891% | 0.5731% | 2.116% | 5.653% |  | 3.951% | 3.951% | 2.784% | 0.4521% | 1.308% | 2.838%
| PDM3 |  | -0.7497% | 4.945% | 3.155% | 18.32% | -2.214% |  | 0.8773% | 0.8773% | 43.18% | 4.529% | -5.952% | -0.9124%
| Z-Score (Absolute) |  | 0.64 | 0.67 | 0.26 | 3.75 | 0.97 | 88.9 | 0.26 | 0.26 | 9.48 | 0.58 | 1.83 | 0.67
| Analytical results for Ti in OREAS 993b (Certified Value 0.118 wt.%) | [18 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | 5.01 |  |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.121 | 0.119 | 0.11574 | 0.121 | 0.12 | 0.119 | 0.1 | 0.115 | 0.124 | 0.1218 | 0.11155 | 0.117 | 0.12 |  |  |  |  | 1
| 1 | 2 | 0.122 | 0.119 | 0.11401 | 0.121 | 0.12 | 0.117 | 0.1 | 0.115 | 0.123 | 0.12118 | 0.11131 | 0.111 | 0.12 |  |  |  |  | 28
| 1 | 3 | 0.117 | 0.122 | 0.11642 | 0.12 | 0.12 | 0.118 | 0.1 | 0.111 | 0.123 | 0.11981 | 0.11043 | 0.114 | 0.12 |  |  |  |  | 16
| 1 | 4 | 0.123 | 0.121 | 0.11696 | 0.117 | 0.12 | 0.122 | 0.1 | 0.116 | 0.127 | 0.12151 | 0.11319 | 0.11 | 0.11 |  |  |  |  | 0.11804
| 1 | 5 | 0.117 | 0.122 | 0.11533 | 0.119 | 0.12 | 0.121 | 0.1 | 0.114 | 0.12 | 0.11959 | 0.11406 | 0.117 | 0.11 |  |  |  |  | 88
| 1 | 6 | 0.118 | 0.121 | 0.11443 | 0.121 | 0.12 | 0.121 | 0.1 | 0.117 | 0.124 | 0.12102 | 0.11584 | 0.112
| Mean |  | 0.11967 | 0.12067 | 0.11548 | 0.11983 | 0.12 | 0.11967 | 0.1 | 0.11467 | 0.1235 | 0.12082 | 0.11273 | 0.1135 | 0.116
| Median |  | 0.1195 | 0.121 | 0.11553 | 0.1205 | 0.12 | 0.12 | 0.1 | 0.115 | 0.1235 | 0.1211 | 0.11237 | 0.113 | 0.12
| Std Dev. |  | 0.0026583 | 0.0013663 | 0.0011329 | 0.0016021 | 0 | 0.0019664 | 1.5202e-17 | 0.0020656 | 0.0022583 | 0.0009091 | 0.0020171 | 0.0030166 | 0.0054772
| Rel.Std.Dev. |  | 2.221% | 1.132% | 0.981% | 1.337% | 0% | 1.643% | 1.52e-14% | 1.801% | 1.829% | 0.7524% | 1.789% | 2.658% | 4.722%
| PDM3 |  | 1.374% | 2.221% | -2.171% | 1.515% | 1.657% | 1.374% | -15.29% | -2.861% | 4.622% | 2.351% | -4.5% | -3.85% | -1.732%
| Z-Score (Absolute) |  | 0 | 0.18 | 0.77 | 0.03 | 0.06 | 0 | 3.62 | 0.92 | 0.7 | 0.21 | 1.28 | 1.13 | 0.67
| Analytical results for Tl in OREAS 993b (Certified Value 0.29 ppm) | [18 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.16 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.3 | 0.3 | 0.31 | 0.2594 | 0.28 | 10 | 0.31 | 0.31 | 0.6 | 0.347 | 0.29938 | 0.25357 | 0.3 |  |  |  |  | 1
| 1 | 2 | 0.25 | 0.3 | 0.3 | 0.2895 | 0.28 | 10 | 0.3 | 0.3 | 0.5 | 0.322 | 0.29604 | 0.25651 | 0.29 |  |  |  |  | 5
| 1 | 3 | 0.26 | 0.3 | 0.31 | 0.2475 | 0.32 | 10 | 0.29 | 0.3 | 0.5 | 0.414 | 0.29804 | 0.25356 | 0.3 |  |  |  |  | 16
| 1 | 4 | 0.28 | 0.3 | 0.29 | 0.2754 | 0.3 | 10 | 0.34 | 0.31 | 0.5 | 0.314 | 0.30399 | 0.26592 | 0.27 |  |  |  |  | 0.29124
| 1 | 5 | 0.26 | 0.4 | 0.31 | 0.2736 | 0.28 | < 10 | 0.28 | 0.3 | 0.6 | 0.319 | 0.30427 | 0.271 | 0.28 |  |  |  |  | 89
| 1 | 6 | 0.32 | 0.4 | 0.3 | 0.2478 | 0.28 | < 10 | 0.32 | 0.3 | 0.5 | 0.275 | 0.30185 | 0.27467 | 0.28
| Mean |  | 0.27833 | 0.33333 | 0.30333 | 0.26553 | 0.29 | 10 | 0.30667 | 0.30333 | 0.53333 | 0.33183 | 0.3006 | 0.26254 | 0.28667
| Median |  | 0.27 | 0.3 | 0.305 | 0.2665 | 0.28 | 10 | 0.305 | 0.3 | 0.5 | 0.3205 | 0.30062 | 0.26121 | 0.285
| Std Dev. |  | 0.027142 | 0.05164 | 0.008165 | 0.016817 | 0.016733 | 0 | 0.021602 | 0.005164 | 0.05164 | 0.046456 | 0.0033274 | 0.009248 | 0.012111
| Rel.Std.Dev. |  | 9.751% | 15.49% | 2.692% | 6.333% | 5.77% | 0% | 7.044% | 1.702% | 9.682% | 14% | 1.107% | 3.523% | 4.225%
| PDM3 |  | -4.432% | 14.45% | 4.152% | -8.827% | -0.4258% | 3334% | 5.297% | 4.152% | 83.12% | 13.94% | 3.212% | -9.855% | -1.57%
| Z-Score (Absolute) |  | 1.14 | NA | 0.54 | 2.01 | 0.36 | NA | 0.77 | 0.54 | NA | 2.46 | 0.36 | 2.21 | 0.58
Results from laboratories 2.02 and 2.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 2.06 were removed due to their 10 ppm reading resolution.
| Analytical results for Tm in OREAS 993b (Indicative Value < 0.1 ppm) | [18 empty] | Indicative
Tm | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab |  |  |  |  | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.08; 2.04: 2.10; 2.05: 2.14 |  |  |  |  | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*OES/MS; 2.04: 4A*OES/MS; 2.05: 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2.01: < 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.078329 |  |  |  |  | 1
| 1 | 2 | 2.01: < 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.07792 |  |  |  |  | 19
| 1 | 3 | 2.01: < 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.078059 |  |  |  |  | 16
| 1 | 4 | 2.01: < 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.07773 |  |  |  |  | < 0.1
| 1 | 5 | 2.01: < 0.1; 2.02: 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.080067 |  |  |  |  | 25
| 1 | 6 | 2.01: < 0.1; 2.03: < 0.1; 2.04: < 0.1; 2.05: 0.080125
| Mean |  | 2.02: 0.1; 2.05: 0.078705
| Median |  | 2.02: 0.1; 2.05: 0.078194
| Std Dev. |  | 2.05: 0.0010951
| Rel.Std.Dev. |  | 2.05: 1.391%
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.05
| Analytical results for U in OREAS 993b (Certified Value 0.8 ppm) | [18 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.9 | 0.9 | 0.76 | 0.72 | 0.8 | < 10 | < 50 | 0.8 | 0.8 | 0.9 | 1.17 | 0.80626 | 0.54589 | < 50 | 0.7 |  |  | 1
| 1 | 2 | 0.7 | 0.9 | 0.79 | 0.7236 | 0.8 | < 10 | < 50 | 0.8 | 0.8 | 0.9 | 0.88 | 0.79421 | 0.53782 | < 50 | 0.7 |  |  | 6
| 1 | 3 | 0.7 | 0.9 | 0.8 | 0.70093 | 0.8 | < 10 | < 50 | 0.8 | 0.8 | 0.9 | 1.11 | 0.78082 | 0.54773 | < 50 | 0.8 |  |  | 16
| 1 | 4 | 0.7 | 0.9 | 0.78 | 0.73591 | 0.8 | < 10 | < 50 | 0.8 | 0.8 | 0.9 | 0.9 | 0.81653 | 0.55531 | < 50 | 0.7 |  |  | 0.80139
| 1 | 5 | 0.7 | 0.8 | 0.78 | 0.69112 | 0.8 | < 10 | < 50 | 0.8 | 0.8 | 0.9 | 0.92 | 0.79603 | 0.53095 | < 50 | 0.7 |  |  | 90
| 1 | 6 | 0.8 | 0.9 | 0.78 | 0.66975 | 0.8 | < 10 | < 50 | 0.8 | 0.8 | 0.9 | 0.83 | 0.77169 | 0.5757 | < 50 | 0.7
| Mean |  | 0.75 | 0.88333 | 0.78167 | 0.70688 | 0.8 |  |  | 0.8 | 0.8 | 0.9 | 0.96833 | 0.79426 | 0.5489 |  | 0.71667
| Median |  | 0.7 | 0.9 | 0.78 | 0.71046 | 0.8 |  |  | 0.8 | 0.8 | 0.9 | 0.91 | 0.79512 | 0.54681 |  | 0.7
| Std Dev. |  | 0.083666 | 0.040825 | 0.013292 | 0.024319 | 1.2162e-16 |  |  | 1.2162e-16 | 1.2162e-16 | 0 | 0.13761 | 0.01633 | 0.015589 |  | 0.040825
| Rel.Std.Dev. |  | 11.16% | 4.622% | 1.7% | 3.44% | 1.52e-14% |  |  | 1.52e-14% | 1.52e-14% | 0% | 14.21% | 2.056% | 2.84% |  | 5.696%
| PDM3 |  | -6.413% | 10.22% | -2.461% | -11.79% | -0.1737% |  |  | -0.1737% | -0.1737% | 12.3% | 20.83% | -0.8902% | -31.51% |  | -10.57%
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.15 | 0.75 | 0 | 33.99 | 195.82 | 0 | 0 | 0.81 | 1.36 | 0.05 | 2.03 | 195.82 | 0.67
| Analytical results for V in OREAS 993b (Certified Value 63 ppm) | [18 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 58 | 64 | 62.9152 | 60 | 65 | 100 | 62 | 62 | 60 | 72 | 77 | 69.4177 | 58.4629 | 100 | 61 | 66 |  | 1
| 1 | 2 | 60 | 64 | 62.464 | 62 | 65 | 100 | 64 | 63 | 60 | 72 | 76 | 70.5504 | 55.7803 | 100 | 59 | 66 |  | 7
| 1 | 3 | 57 | 63 | 63.2608 | 63 | 65 | 100 | 67 | 63 | 58 | 71 | 77 | 64.5185 | 56.9539 | 100 | 60 | 64 |  | 16
| 1 | 4 | 59 | 64 | 59.8432 | 60 | 63 | 100 | 58 | 63 | 61 | 72 | 77 | 66.0472 | 60.2623 | 100 | 58 | 64 |  | 62.787
| 1 | 5 | 56 | 65 | 61.2448 | 61 | 63 | 100 | 61 | 63 | 60 | 70 | 76 | 67.4511 | 60.1526 | 100 | 60 | 64 |  | 91
| 1 | 6 | 56 | 63 | 62.9344 | 62 | 65 | 100 | 61 | 63 | 61 | 72 | 75 | 68.2771 | 58.0469 | 100 | 60
| Mean |  | 57.6667 | 63.8333 | 62.1104 | 61.3333 | 64.3333 | 100 | 62.1667 | 62.8333 | 60 | 71.5 | 76.3333 | 67.7103 | 58.2765 | 100 | 59.6667 | 64.8
| Median |  | 57.5 | 64 | 62.6896 | 61.5 | 65 | 100 | 61.5 | 63 | 60 | 72 | 76.5 | 67.8641 | 58.2549 | 100 | 60 | 64
| Std Dev. |  | 1.633 | 0.75277 | 1.3164 | 1.2111 | 1.0328 | 0 | 3.0605 | 0.40825 | 1.0954 | 0.83666 | 0.8165 | 2.2056 | 1.7624 | 0 | 1.0328 | 1.0954
| Rel.Std.Dev. |  | 2.832% | 1.179% | 2.119% | 1.975% | 1.605% | 0% | 4.923% | 0.6497% | 1.826% | 1.17% | 1.07% | 3.257% | 3.024% | 0% | 1.731% | 1.691%
| PDM3 |  | -8.155% | 1.667% | -1.078% | -2.315% | 2.463% | 59.27% | -0.9879% | 0.07385% | -4.439% | 13.88% | 21.58% | 7.841% | -7.184% | 59.27% | -4.97% | 3.206%
| Z-Score (Absolute) |  | 1.23 | 0.24 | 0.17 | 0.36 | 0.36 | 8.85 | 0.16 | 0 | 0.67 | 2.06 | 3.21 | 1.16 | 1.08 | NA | 0.75 | 0.47
Results from laboratory 2.15 were removed due to their 100 ppm reading resolution.
| Analytical results for W in OREAS 993b (Certified Value 2.47 ppm) | [18 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.6 | 3 | 2.7 | 2.4936 | 2.5 | < 10 | < 80 | 2.9 | 22 | 2.6 | 2.3 | 2.6776 | 2.2714 | < 80 | 2.5 | 37 |  | 1
| 1 | 2 | 2.4 | 2 | 2.7 | 2.4254 | 2.4 | < 10 | < 80 | 3 | 23 | 2.5 | 2.4 | 2.5855 | 2.1211 | < 80 | 2.4 | 35 |  | 32
| 1 | 3 | 2.2 | 3 | 2.9 | 2.5211 | 2.7 | < 10 | < 80 | 3.1 | 22 | 2.5 | 2.3 | 2.6145 | 2.128 | < 80 | 2.5 | 36 |  | 16
| 1 | 4 | 2.4 | 2 | 2.7 | 2.648 | 2.5 | < 10 | < 80 | 2.9 | 25 | 2.6 | 2.3 | 2.6769 | 2.22 | < 80 | 2.4 | 36 |  | 2.4737
| 1 | 5 | 2.4 | 2 | 2.7 | 2.545 | 2.4 | < 10 | < 80 | 3.1 | 25 | 2.5 | 2.3 | 2.5477 | 2.2045 | < 80 | 2.4 | 35 |  | 92
| 1 | 6 | 3.2 | 3 | 2.8 | 2.5163 | 2.5 | < 10 | < 80 | 3.1 | 26 | 2.5 | 2.2 | 2.5369 | 2.3482 | < 80 | 2.4
| Mean |  | 2.5333 | 2.5 | 2.75 | 2.5249 | 2.5 |  |  | 3.0167 | 23.8333 | 2.5333 | 2.3 | 2.6065 | 2.2155 |  | 2.4333 | 35.8
| Median |  | 2.4 | 2.5 | 2.7 | 2.5187 | 2.5 |  |  | 3.05 | 24 | 2.5 | 2.3 | 2.6 | 2.2123 |  | 2.4 | 36
| Std Dev. |  | 0.35024 | 0.54772 | 0.083666 | 0.072837 | 0.10954 |  |  | 0.098319 | 1.7224 | 0.05164 | 0.063246 | 0.061363 | 0.086532 |  | 0.05164 | 0.83666
| Rel.Std.Dev. |  | 13.83% | 21.91% | 3.042% | 2.885% | 4.382% |  |  | 3.259% | 7.227% | 2.038% | 2.75% | 2.354% | 3.906% |  | 2.122% | 2.337%
| PDM3 |  | 2.409% | 1.062% | 11.17% | 2.068% | 1.062% |  |  | 21.95% | 863.5% | 2.409% | -7.023% | 5.367% | -10.44% |  | -1.633% | 1347%
| Z-Score (Absolute) |  | 0 | NA | 0.67 | 0.03 | 0.1 | 7.68 | 116.6 | 1.5 | NA | 0 | 0.73 | 0.23 | 0.99 | 116.6 | 0.31 | NA
Results from laboratories 2.02, 2.09 and 5.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 993b (Certified Value 5.63 ppm) | [18 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 5.6 | 5.7 | 5.8 | 5.2401 | 5.3 | 5.3 | 5 | 5.8 | 8 | 5.9323 | 5.451 | 5.9 | [5 empty] | 1
| 1 | 2 | 5 | 5.8 | 5.8 | 5.2534 | 5.2 | 5.8 | 5 | 5.7 | 8.4 | 5.9802 | 5.4612 | 5.6 | [5 empty] | 33
| 1 | 3 | 5.2 | 5.8 | 5.9 | 5.3655 | 5.8 | 6 | 5 | 5.6 | 8.2 | 5.8673 | 5.4428 | 5.6 | [5 empty] | 16
| 1 | 4 | 5.3 | 6 | 5.8 | 5.2506 | 5.1 | 5.7 | 5 | 5.9 | 8.3 | 6.2039 | 5.6075 | 5.6 | [5 empty] | 5.6279
| 1 | 5 | 5.2 | 6 | 5.9 | 5.175 | 5.2 | 5.6 | 5 | 5.7 | 8.3 | 6.0118 | 5.6459 | 5.6 | [5 empty] | 93
| 1 | 6 | 5.7 | 5.7 | 5.8 | 5.3326 | 5.3 | 5.5 | 5 | 6 | 8 | 6.1588 | 5.6919 | 5.8
| Mean |  | 5.3333 | 5.8333 | 5.8333 | 5.2695 | 5.3167 | 5.65 | 5 | 5.7833 | 8.2 | 6.0257 | 5.5501 | 5.6833
| Median |  | 5.25 | 5.8 | 5.8 | 5.252 | 5.25 | 5.65 | 5 | 5.75 | 8.25 | 5.996 | 5.5343 | 5.6
| Std Dev. |  | 0.26583 | 0.13663 | 0.05164 | 0.068696 | 0.24833 | 0.2429 | 0 | 0.1472 | 0.16733 | 0.13081 | 0.11118 | 0.13292
| Rel.Std.Dev. |  | 4.984% | 2.342% | 0.8853% | 1.304% | 4.671% | 4.299% | 0% | 2.545% | 2.041% | 2.171% | 2.003% | 2.339%
| PDM3 |  | -5.233% | 3.651% | 3.651% | -6.368% | -5.53% | 0.3934% | -11.16% | 2.763% | 45.7% | 7.07% | -1.382% | 0.9857%
| Z-Score (Absolute) |  | 1.57 | 0.67 | 0.67 | 1.86 | 1.65 | 0.15 | NA | 0.45 | 11.31 | 1.54 | 0.6 | 0
Results from laboratory 2.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 993b (Certified Value 0.58 ppm) | [18 empty] | Certified
Yb | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.08: Lab | [5 empty] | 1
| No | No | 2.01: 2.01; 2.02: 2.02; 2.03: 2.04; 2.04: 2.08; 2.05: 2.10; 2.08: 2.14 | [5 empty] | ppm
|  |  | 2.01: 4A*MS; 2.02: 4A*MS; 2.03: 4A*MS; 2.04: 4A*OES/MS; 2.05: 4A*OES/MS; 2.08: 4A*MS | [5 empty] | 2
2
| 1 | 1 | 2.01: 0.6; 2.02: 0.6; 2.03: 0.59759; 2.04: 0.6; 2.05: 0.6; 2.08: 0.5177 | [5 empty] | 1
| 1 | 2 | 2.01: 0.5; 2.02: 0.7; 2.03: 0.59591; 2.04: 0.6; 2.05: 0.6; 2.08: 0.5484 | [5 empty] | 0
| 1 | 3 | 2.01: 0.5; 2.02: 0.7; 2.03: 0.54141; 2.04: 0.6; 2.05: 0.6; 2.08: 0.541 | [5 empty] | 16
| 1 | 4 | 2.01: 0.5; 2.02: 0.6; 2.03: 0.57669; 2.04: 0.7; 2.05: 0.6; 2.08: 0.56046 | [5 empty] | 0.58241
| 1 | 5 | 2.01: 0.5; 2.02: 0.5; 2.03: 0.64736; 2.04: 0.6; 2.05: 0.6; 2.08: 0.52736 | [5 empty] | 94
| 1 | 6 | 2.01: 0.6; 2.02: 0.5; 2.03: 0.58152; 2.04: 0.7; 2.05: 0.6; 2.08: 0.53145
| Mean |  | 2.01: 0.53333; 2.02: 0.6; 2.03: 0.59008; 2.04: 0.63333; 2.05: 0.6; 2.08: 0.53773
| Median |  | 2.01: 0.5; 2.02: 0.6; 2.03: 0.58872; 2.04: 0.6; 2.05: 0.6; 2.08: 0.53623
| Std Dev. |  | 2.01: 0.05164; 2.02: 0.089443; 2.03: 0.034618; 2.04: 0.05164; 2.05: 0; 2.08: 0.015421
| Rel.Std.Dev. |  | 2.01: 9.682%; 2.02: 14.91%; 2.03: 5.867%; 2.04: 8.154%; 2.05: 0%; 2.08: 2.868%
| PDM3 |  | 2.01: -8.427%; 2.02: 3.02%; 2.03: 1.317%; 2.04: 8.743%; 2.05: 3.02%; 2.08: -7.672%
| Z-Score (Absolute) |  | 2.01: 1.92; 2.02: 0.15; 2.03: 0.15; 2.04: 1.19; 2.05: 0.15; 2.08: 1.79
| Analytical results for Zn in OREAS 993b (Certified Value 0.332 wt.%) | [18 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 5.01 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.34639 | 0.356 | 0.33678 | 0.324 | 0.324 | 0.349 | 0.306 | 0.336 | 0.33385 | 0.3137 | 0.3456 | 0.35043 | 0.30923 | 0.35 | 0.334 | 0.31805 |  | 1
| 1 | 2 | 0.35189 | 0.357 | 0.33592 | 0.323 | 0.32 | 0.351 | 0.298 | 0.3342 | 0.33174 | 0.3211 | 0.3262 | 0.34876 | 0.30126 | 0.349 | 0.324 | 0.3066 |  | 0
| 1 | 3 | 0.33912 | 0.359 | 0.33509 | 0.34 | 0.328 | 0.353 | 0.296 | 0.3393 | 0.32433 | 0.321 | 0.3393 | 0.35627 | 0.29985 | 0.351 | 0.329 | 0.31405 |  | 16
| 1 | 4 | 0.35362 | 0.364 | 0.3352 | 0.327 | 0.323 | 0.346 | 0.32 | 0.3312 | 0.34051 | 0.3244 | 0.341 | 0.34597 | 0.30979 | 0.351 | 0.32 | 0.3104 |  | 0.33244
| 1 | 5 | 0.33681 | 0.362 | 0.33638 | 0.322 | 0.322 | 0.349 | 0.302 | 0.3346 | 0.33047 | 0.3212 | 0.3255 | 0.35742 | 0.30778 | 0.35 | 0.34 | 0.3102 |  | 95
| 1 | 6 | 0.33738 | 0.362 | 0.33502 | 0.325 | 0.327 | 0.351 | 0.314 | 0.3346 | 0.3454 | 0.3068 | 0.3187 | 0.35629 | 0.32685 | 0.349 | 0.326
| Mean |  | 0.3442 | 0.36 | 0.33573 | 0.32683 | 0.324 | 0.34983 | 0.306 | 0.33498 | 0.33438 | 0.31803 | 0.33272 | 0.35252 | 0.30913 | 0.35 | 0.32883 | 0.31186
| Median |  | 0.34275 | 0.3605 | 0.33556 | 0.3245 | 0.3235 | 0.35 | 0.304 | 0.3346 | 0.3328 | 0.32105 | 0.33275 | 0.35335 | 0.30851 | 0.35 | 0.3275 | 0.3104
| Std Dev. |  | 0.007478 | 0.0031623 | 0.00074239 | 0.0066758 | 0.0030332 | 0.0024014 | 0.0093808 | 0.0026415 | 0.0075174 | 0.0065381 | 0.010667 | 0.0047698 | 0.0096384 | 0.00089443 | 0.0072226 | 0.0043493
| Rel.Std.Dev. |  | 2.173% | 0.8784% | 0.2211% | 2.043% | 0.9362% | 0.6864% | 3.066% | 0.7886% | 2.248% | 2.056% | 3.206% | 1.353% | 3.118% | 0.2556% | 2.196% | 1.395%
| PDM3 |  | 3.538% | 8.29% | 0.9898% | -1.687% | -2.539% | 5.232% | -7.954% | 0.7647% | 0.5842% | -4.334% | 0.08286% | 6.041% | -7.013% | 5.282% | -1.085% | -6.191%
| Z-Score (Absolute) |  | 0.55 | 1.36 | 0.11 | 0.35 | 0.49 | 0.84 | 1.42 | 0.07 | 0.04 | 0.8 | 0.04 | 0.98 | 1.26 | 0.85 | 0.24 | 1.12
| Analytical results for Zr in OREAS 993b (Certified Value 27.3 ppm) | [18 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.08 | 2.09 | 2.10 | 2.11 | 2.13 | 2.14 | 2.16 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 26.7 | 30 | 27.9 | 29.62 | 28.4 | 28 | 29 | 26.5 | 27.8 | 26.0444 | 26.2621 | 25.6 | [5 empty] | 1
| 1 | 2 | 26.3 | 38 | 27.8 | 29.67 | 27.3 | 29 | 30 | 25.2 | 28.5 | 24.6361 | 26.1854 | 24.1 | [5 empty] | 9
| 1 | 3 | 25 | 27 | 27 | 29.65 | 28.1 | 29 | 30 | 25.2 | 27.8 | 24.9842 | 25.8634 | 24.4 | [5 empty] | 16
| 1 | 4 | 26.5 | 27 | 27.7 | 29.9 | 26.8 | 28 | 30 | 28.3 | 28.6 | 26.2915 | 26.7524 | 24.2 | [5 empty] | 27.2724
| 1 | 5 | 24.3 | 27 | 27.7 | 29.74 | 26.9 | 27 | 30 | 26.3 | 27.8 | 24.8712 | 26.7015 | 24 | [5 empty] | 96
| 1 | 6 | 25.6 | 28 | 27.2 | 29.96 | 27.7 | 28 | 30 | 27.8 | 27.8 | 25.0813 | 26.7986 | 25
| Mean |  | 25.7333 | 29.5 | 27.55 | 29.7567 | 27.5333 | 28.1667 | 29.8333 | 26.55 | 28.05 | 25.3181 | 26.4272 | 24.55
| Median |  | 25.95 | 27.5 | 27.7 | 29.705 | 27.5 | 28 | 30 | 26.4 | 27.8 | 25.0327 | 26.4818 | 24.3
| Std Dev. |  | 0.94375 | 4.3243 | 0.36194 | 0.14123 | 0.64704 | 0.75277 | 0.40825 | 1.2911 | 0.38859 | 0.67932 | 0.38014 | 0.6253
| Rel.Std.Dev. |  | 3.667% | 14.66% | 1.314% | 0.4746% | 2.35% | 2.673% | 1.368% | 4.863% | 1.385% | 2.683% | 1.438% | 2.547%
| PDM3 |  | -5.643% | 8.168% | 1.018% | 9.109% | 0.9568% | 3.279% | 9.39% | -2.649% | 2.851% | -7.166% | -3.099% | -9.982%
| Z-Score (Absolute) |  | 0.83 | 0.9 | 0 | 1.02 | 0 | 0.29 | 1.06 | 0.46 | 0.23 | 1.03 | 0.51 | 1.38
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C620: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C911: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1077: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1132: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 3-Acid
| Analytical results for Ag in OREAS 993b (Certified Value 38.9 ppm) | [13 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 | 2.19 | 5.01 | ppm
|  |  | 3A*OES | 3A*OES/MS | 3A*OES | 3A*OES | 3A*OES | AR*AAS | 3A*MS | 3A*OES | 3A*OES | 3A*OES | AR*AAS | AR*OES | 1
[8 empty] | 0.4g |  |  |  |  | 01g | 0.2g | 2
| 1 | 1 | 38.8 | 40.6 | 45.11 | 36 | 38.6 | 36.3 | 37.4 | 38.4391 | 41.7 | 31 | 42.55 | 41.4 | 1
| 1 | 2 | 37.9 | 40.9 | 43.29 | 37 | 38.3 | 37.6 | 37.9 | 38.4195 | 41.8 | 32 | 43.02 | 40.8 | 0
| 1 | 3 | 37.2 | 42.3 | 43.34 | 36 | 38.7 | 37.4 | 36.6 | 37.554 | 41.5 | 32 | 43.69 | 39.6 | 16
| 1 | 4 | 39.3 | 40.9 | 43.39 | 37 | 38.5 | 36.8 | 37.7 | 37.4441 | 42 | 30 | 42.03 | 40.6 | 38.9118
| 1 | 5 | 39.3 | 40.5 | 43.7 | 37 | 38.5 | 36.4 | 37.1 | 37.4075 | 41.3 | 32 | 42.84 | 40.2 | 98
| 1 | 6 | 37.9 | 40.8 | 44.13 | 37 | 38.7 | 36.6 | 36.4 | 37.8163 | 42.1 | 32 | 43.06
| Mean |  | 38.4 | 41 | 43.8267 | 36.6667 | 38.55 | 36.85 | 37.1833 | 37.8467 | 41.7333 | 31.5 | 42.865 | 40.52
| Median |  | 38.35 | 40.85 | 43.545 | 37 | 38.55 | 36.7 | 37.25 | 37.6852 | 41.75 | 32 | 42.93 | 40.6
| Std Dev. |  | 0.86255 | 0.65727 | 0.70293 | 0.5164 | 0.15166 | 0.53572 | 0.59805 | 0.47339 | 0.30111 | 0.83666 | 0.55493 | 0.67231
| Rel.Std.Dev. |  | 2.246% | 1.603% | 1.604% | 1.408% | 0.3934% | 1.454% | 1.608% | 1.251% | 0.7215% | 2.656% | 1.295% | 1.659%
| PDM3 |  | -1.315% | 5.366% | 12.63% | -5.77% | -0.9298% | -5.299% | -4.442% | -2.737% | 7.251% | -19.05% | 10.16% | 4.133%
| Z-Score (Absolute) |  | 0.03 | 0.88 | 1.87 | 0.63 | 0.03 | 0.57 | 0.45 | 0.22 | 1.14 | 2.44 | 1.54 | 0.72
| Analytical results for Al in OREAS 993b (Indicative Value 0.768 wt.%) | [13 empty] | Indicative
Al | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | wt.%
|  |  | 3A*OES | 3A*OES | [10 empty] | 3
3
| 1 | 1 | 0.81 | 0.72698 | [10 empty] | 1
| 1 | 2 | 0.8 | 0.73777 | [10 empty] | 19
| 1 | 3 | 0.77 | 0.72944 | [10 empty] | 16
| 1 | 4 | 0.83 | 0.73897 | [10 empty] | 0.76755
| 1 | 5 | 0.81 | 0.72823 | [10 empty] | 29
| 1 | 6 | 0.79 | 0.73926
| Mean |  | 0.80167 | 0.73344
| Median |  | 0.805 | 0.7336
| Std Dev. |  | 0.020412 | 0.0057979
| Rel.Std.Dev. |  | 2.546% | 0.7905%
| PDM3 |  | 4.444% | -4.444%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for As in OREAS 993b (Certified Value 107 ppm) | [13 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 | 5.01 |  | ppm
|  |  | 3A*OES | 3A*OES/MS | 3A*OES | 3A*OES | 3A*OES | 3A*OES | 3A*MS | 3A*OES | 3A*OES | 3A*OES | AR*OES |  | 0
[13 empty] | 0.2g |  | 0
| 1 | 1 | 104 | 96 | 114.46 | < 100 | 104 | 106 | 90.5 | 111.6654 | 114 | 108 | 110 |  | 1
| 1 | 2 | 103 | 94 | 114.7 | 100 | 104 | 118 | 97.6 | 111.3338 | 113 | 111 | 108 |  | 1
| 1 | 3 | 100 | 93 | 113.18 | < 100 | 105 | 109 | 93.3 | 114.0242 | 114 | 108 | 102 |  | 16
| 1 | 4 | 107 | 92 | 112.94 | 100 | 104 | 126 | 99.4 | 113.5125 | 113 | 108 | 110 |  | 106.9139
| 1 | 5 | 107 | 98 | 111.29 | 100 | 104 | 103 | 100.8 | 110.0022 | 113 | 111 | 104 |  | 99
| 1 | 6 | 102 | 95 | 112.78 | < 100 | 105 | 129 | 90.6 | 114.9463 | 115 | 111
| Mean |  | 103.8333 | 94.6667 | 113.225 | 100 | 104.3333 | 115.1667 | 95.3667 | 112.5807 | 113.6667 | 109.5 | 106.8
| Median |  | 103.5 | 94.5 | 113.06 | 100 | 104 | 113.5 | 95.45 | 112.589 | 113.5 | 109.5 | 108
| Std Dev. |  | 2.7869 | 2.1602 | 1.2431 | 0 | 0.5164 | 10.8336 | 4.5045 | 1.8756 | 0.8165 | 1.6432 | 3.6332
| Rel.Std.Dev. |  | 2.684% | 2.282% | 1.098% | 0% | 0.4949% | 9.407% | 4.723% | 1.666% | 0.7183% | 1.501% | 3.402%
| PDM3 |  | -2.881% | -11.46% | 5.903% | -6.467% | -2.414% | 7.719% | -10.8% | 5.3% | 6.316% | 2.419% | -0.1065%
| Z-Score (Absolute) |  | 0.61 | 1.91 | 0.72 | NA | 0.54 | 1 | 1.81 | 0.63 | 0.78 | 0.19 | 0.19
Results from laboratory 2.08 were removed due to their 100 ppm reading resolution.
| Analytical results for Ba in OREAS 993b (Indicative Value 13.6 ppm) | [13 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 5.01 | [9 empty] | ppm
|  |  | 3A*OES | 3A*OES | AR*OES | [9 empty] | 1
[5 empty] | 0.2g | [9 empty] | 1
| 1 | 1 | 11 | 13.36 | 16 | [9 empty] | 1
| 1 | 2 | 11 | 15.22 | 18 | [9 empty] | 2
| 1 | 3 | 11 | 16.62 | 16 | [9 empty] | 16
| 1 | 4 | 11 | 12.32 | 15.5 | [9 empty] | 13.6356
| 1 | 5 | 8 | 12.88 | 15.5 | [9 empty] | 30
| 1 | 6 | 10 | 13.04
| Mean |  | 10.3333 | 13.9067 | 16.2
| Median |  | 11 | 13.2 | 16
| Std Dev. |  | 1.2111 | 1.6558 | 1.0368
| Rel.Std.Dev. |  | 11.72% | 11.91% | 6.4%
| PDM3 |  | -24.22% | 1.988% | 18.81%
| Z-Score (Absolute) |  | 1.05 | 0 | 0.67
| Analytical results for Be in OREAS 993b (Indicative Value < 0.5 ppm) | [13 empty] | Indicative
Be | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | < 0.5 | [11 empty] | 1
| 1 | 2 | < 0.5 | [11 empty] | 25
| 1 | 3 | < 0.5 | [11 empty] | 16
| 1 | 4 | < 0.5 | [11 empty] | < 0.5
| 1 | 5 | < 0.5 | [11 empty] | 31
| 1 | 6 | < 0.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 993b (Indicative Value 13.9 ppm) | [13 empty] | Indicative
Bi | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 5.01 | [10 empty] | ppm
|  |  | 3A*OES | AR*OES | [10 empty] | 1
|  |  |  | 0.2g | [10 empty] | 1
| 1 | 1 | 18 | 22 | [10 empty] | 1
| 1 | 2 | 15 | 12 | [10 empty] | 3
| 1 | 3 | 17 | 12 | [10 empty] | 16
| 1 | 4 | 18 | 12 | [10 empty] | 13.9167
| 1 | 5 | 18 | 8 | [10 empty] | 32
| 1 | 6 | 15
| Mean |  | 16.8333 | 13.2
| Median |  | 17.5 | 12
| Std Dev. |  | 1.472 | 5.2154
| Rel.Std.Dev. |  | 8.744% | 39.51%
| PDM3 |  | 20.96% | -5.15%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ca in OREAS 993b (Indicative Value 0.715 wt.%) | [13 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 5.01 | [9 empty] | wt.%
|  |  | 3A*OES | 3A*OES | AR*OES | [9 empty] | 3
[5 empty] | 0.2g | [9 empty] | 3
| 1 | 1 | 0.82 | 0.71678 | 0.66 | [9 empty] | 1
| 1 | 2 | 0.75 | 0.71132 | 0.66 | [9 empty] | 27
| 1 | 3 | 0.72 | 0.72401 | 0.64 | [9 empty] | 16
| 1 | 4 | 0.82 | 0.71667 | 0.64 | [9 empty] | 0.71538
| 1 | 5 | 0.78 | 0.72247 | 0.63 | [9 empty] | 33
| 1 | 6 | 0.79 | 0.72959
| Mean |  | 0.78 | 0.72014 | 0.646
| Median |  | 0.785 | 0.71962 | 0.64
| Std Dev. |  | 0.039497 | 0.0064936 | 0.013416
| Rel.Std.Dev. |  | 5.064% | 0.9017% | 2.077%
| PDM3 |  | 9.033% | 0.6654% | -9.698%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.84
| Analytical results for Cd in OREAS 993b (Indicative Value 13.8 ppm) | [13 empty] | Indicative
Cd | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | 14 | [11 empty] | 1
| 1 | 2 | 14 | [11 empty] | 4
| 1 | 3 | 13 | [11 empty] | 16
| 1 | 4 | 14 | [11 empty] | 13.8333
| 1 | 5 | 14 | [11 empty] | 34
| 1 | 6 | 14
| Mean |  | 13.8333
| Median |  | 14
| Std Dev. |  | 0.40825
| Rel.Std.Dev. |  | 2.951%
| PDM3 |  | 2.442e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 993b (Indicative Value 8.5 ppm) | [13 empty] | Indicative
Ce | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | 11 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 29
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | 10 | [11 empty] | 8.5
| 1 | 5 | 10 | [11 empty] | 35
| 1 | 6 | 10
| Mean |  | 10.25
| Median |  | 10
| Std Dev. |  | 0.5
| Rel.Std.Dev. |  | 4.878%
| PDM3 |  | 20.59%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 993b (Certified Value 151 ppm) | [13 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 | 5.01 |  | ppm
|  |  | 3A*OES | 3A*OES/MS | 3A*OES | 3A*OES | 3A*OES | 3A*OES | 3A*MS | 3A*OES | 3A*OES | 3A*OES | AR*OES |  | 0
[13 empty] | 0.2g |  | 0
| 1 | 1 | 149 | 158 | 163.23 | 150 | 147 | 100 | 143.8 | 161.4319 | 152 | 148 | 150 |  | 1
| 1 | 2 | 147 | 159 | 161.64 | 150 | 146 | 200 | 145.3 | 154.4911 | 153 | 150 | 151.5 |  | 0
| 1 | 3 | 144 | 157 | 161.38 | 150 | 146 | 200 | 141.7 | 149.2738 | 154 | 148 | 145 |  | 16
| 1 | 4 | 154 | 155 | 161.37 | 150 | 143 | 200 | 144.8 | 155.7051 | 151 | 148 | 144.5 |  | 151.0958
| 1 | 5 | 153 | 155 | 162.52 | 160 | 144 | 100 | 142.7 | 152.9149 | 151 | 150 | 147 |  | 100
| 1 | 6 | 147 | 158 | 162.99 | 150 | 144 | 200 | 139.7 | 158.6261 | 151 | 148
| Mean |  | 149 | 157 | 162.1883 | 151.6667 | 145 | 166.6667 | 143 | 155.4071 | 152 | 148.6667 | 147.6
| Median |  | 148 | 157.5 | 162.08 | 150 | 145 | 200 | 143.25 | 155.0981 | 151.5 | 148 | 147
| Std Dev. |  | 3.8471 | 1.6733 | 0.83204 | 4.0825 | 1.5492 | 51.6398 | 2.09 | 4.2759 | 1.2649 | 1.0328 | 3.07
| Rel.Std.Dev. |  | 2.582% | 1.066% | 0.513% | 2.692% | 1.068% | 30.98% | 1.462% | 2.751% | 0.8322% | 0.6947% | 2.08%
| PDM3 |  | -1.387% | 3.908% | 7.341% | 0.3778% | -4.034% | 10.31% | -5.358% | 2.853% | 0.5984% | -1.608% | -2.314%
| Z-Score (Absolute) |  | 0 | 1.35 | 2.22 | NA | 0.67 | NA | 1.01 | 1.08 | 0.51 | 0.06 | 0.24
Results from laboratory 2.08 were removed due to their 10 ppm reading resolution. Results from laboratory 2.11 were removed due to their 100 ppm reading resolution.
| Analytical results for Cr in OREAS 993b (Indicative Value 48.4 ppm) | [13 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | ppm
|  |  | 3A*OES | 3A*OES | [10 empty] | 1
1
| 1 | 1 | 44 | 53.03 | [10 empty] | 1
| 1 | 2 | 43 | 54.08 | [10 empty] | 7
| 1 | 3 | 41 | 55.62 | [10 empty] | 16
| 1 | 4 | 44 | 53.46 | [10 empty] | 48.4425
| 1 | 5 | 44 | 51.38 | [10 empty] | 36
| 1 | 6 | 44 | 53.74
| Mean |  | 43.3333 | 53.5517
| Median |  | 44 | 53.6
| Std Dev. |  | 1.2111 | 1.3849
| Rel.Std.Dev. |  | 2.795% | 2.586%
| PDM3 |  | -10.55% | 10.55%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cu in OREAS 993b (Indicative Value 22.18 wt.%) | [13 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.19 | [9 empty] | wt.%
|  |  | 3A*OES | 3A*OES | AR*AAS | [9 empty] | 2
[5 empty] | 01g | [9 empty] | 2
| 1 | 1 | > 1 | 22.8889 | 21.44 | [9 empty] | 1
| 1 | 2 | > 1 | 22.8033 | 21.81 | [9 empty] | 31
| 1 | 3 | > 1 | 22.9276 | 21.78 | [9 empty] | 16
| 1 | 4 | > 1 | 22.5994 | 21.28 | [9 empty] | 22.18
| 1 | 5 | > 1 | 22.7968 | 21.52 | [9 empty] | 37
| 1 | 6 | > 1 | 22.6434 | 21.67
| Mean |  |  | 22.7766 | 21.5833
| Median |  |  | 22.8 | 21.595
| Std Dev. |  |  | 0.1309 | 0.20695
| Rel.Std.Dev. |  |  | 0.5747% | 0.9588%
| PDM3 |  |  | 2.69% | -2.69%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Fe in OREAS 993b (Indicative Value 24.99 wt.%) | [13 empty] | Indicative
Fe | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 2.19 | [9 empty] | wt.%
|  |  | 3A*OES | 3A*OES | AR*AAS | [9 empty] | 2
[5 empty] | 01g | [9 empty] | 2
| 1 | 1 | 23.54 | 26.8237 | 24.52 | [9 empty] | 1
| 1 | 2 | 23.33 | 26.793 | 25.12 | [9 empty] | 32
| 1 | 3 | 22.6 | 26.7532 | 25 | [9 empty] | 16
| 1 | 4 | 24.06 | 26.5796 | 24.35 | [9 empty] | 24.9908
| 1 | 5 | 24.22 | 26.919 | 24.62 | [9 empty] | 38
| 1 | 6 | 23.18 | 26.6356 | 24.79
| Mean |  | 23.4883 | 26.7507 | 24.7333
| Median |  | 23.435 | 26.7731 | 24.705
| Std Dev. |  | 0.59566 | 0.12488 | 0.29296
| Rel.Std.Dev. |  | 2.536% | 0.4668% | 1.184%
| PDM3 |  | -6.012% | 7.042% | -1.03%
| Z-Score (Absolute) |  | 0.67 | 1.09 | 0
| Analytical results for Ga in OREAS 993b (Indicative Value < 5 ppm) | [13 empty] | Indicative
Ga | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | < 5 | [11 empty] | 1
| 1 | 2 | < 5 | [11 empty] | 9
| 1 | 3 | < 5 | [11 empty] | 16
| 1 | 4 | < 5 | [11 empty] | < 5
| 1 | 5 | < 5 | [11 empty] | 39
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
Ge | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 10
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | < 10 | [11 empty] | < 10
| 1 | 5 | < 10 | [11 empty] | 40
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Hg in OREAS 993b (Indicative Value < 1 ppm) | [13 empty] | Indicative
Hg | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | < 1 | [11 empty] | 1
| 1 | 2 | < 1 | [11 empty] | 35
| 1 | 3 | < 1 | [11 empty] | 16
| 1 | 4 | < 1 | [11 empty] | < 1
| 1 | 5 | < 1 | [11 empty] | 41
| 1 | 6 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
In | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 2
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | < 10 | [11 empty] | < 10
| 1 | 5 | < 10 | [11 empty] | 29
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 993b (Indicative Value 0.189 wt.%) | [13 empty] | Indicative
K | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | wt.%
|  |  | 3A*OES | 3A*OES | [10 empty] | 3
3
| 1 | 1 | 0.19 | 0.19634 | [10 empty] | 1
| 1 | 2 | 0.19 | 0.19519 | [10 empty] | 24
| 1 | 3 | 0.18 | 0.19342 | [10 empty] | 16
| 1 | 4 | 0.2 | 0.18588 | [10 empty] | 0.18947
| 1 | 5 | 0.18 | 0.18993 | [10 empty] | 30
| 1 | 6 | 0.18 | 0.19289
| Mean |  | 0.18667 | 0.19228
| Median |  | 0.185 | 0.19315
| Std Dev. |  | 0.008165 | 0.0038238
| Rel.Std.Dev. |  | 4.374% | 1.989%
| PDM3 |  | -1.48% | 1.48%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for La in OREAS 993b (Indicative Value 9.24 ppm) | [13 empty] | Indicative
La | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | ppm
|  |  | 3A*OES | 3A*OES | [10 empty] | 2
2
| 1 | 1 | 10 | 8.922 | [10 empty] | 1
| 1 | 2 | 9 | 8.603 | [10 empty] | 25
| 1 | 3 | 9 | 8.757 | [10 empty] | 16
| 1 | 4 | 10 | 9.021 | [10 empty] | 9.2448
| 1 | 5 | 10 | 8.636 | [10 empty] | 31
| 1 | 6 | 10 | 8.999
| Mean |  | 9.6667 | 8.823
| Median |  | 10 | 8.8395
| Std Dev. |  | 0.5164 | 0.18314
| Rel.Std.Dev. |  | 5.342% | 2.076%
| PDM3 |  | 4.563% | -4.563%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Li in OREAS 993b (Indicative Value 20.1 ppm) | [13 empty] | Indicative
Li | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | ppm
|  |  | 3A*OES | 3A*OES | [10 empty] | 1
1
| 1 | 1 | 34 | 7.43 | [10 empty] | 1
| 1 | 2 | 33 | 7.25 | [10 empty] | 26
| 1 | 3 | 32 | 6.97 | [10 empty] | 16
| 1 | 4 | 34 | 6.77 | [10 empty] | 20.145
| 1 | 5 | 34 | 7.27 | [10 empty] | 32
| 1 | 6 | 32 | 7.05
| Mean |  | 33.1667 | 7.1233
| Median |  | 33.5 | 7.15
| Std Dev. |  | 0.98319 | 0.23889
| Rel.Std.Dev. |  | 2.964% | 3.354%
| PDM3 |  | 64.64% | -64.64%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mg in OREAS 993b (Indicative Value 0.405 wt.%) | [13 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | wt.%
|  |  | 3A*OES | 3A*OES | [10 empty] | 3
3
| 1 | 1 | 0.38 | 0.43626 | [10 empty] | 1
| 1 | 2 | 0.37 | 0.43264 | [10 empty] | 27
| 1 | 3 | 0.36 | 0.43438 | [10 empty] | 16
| 1 | 4 | 0.39 | 0.43116 | [10 empty] | 0.40492
| 1 | 5 | 0.39 | 0.43257 | [10 empty] | 33
| 1 | 6 | 0.37 | 0.43199
| Mean |  | 0.37667 | 0.43317
| Median |  | 0.375 | 0.43261
| Std Dev. |  | 0.012111 | 0.0018485
| Rel.Std.Dev. |  | 3.215% | 0.4268%
| PDM3 |  | -6.977% | 6.977%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mn in OREAS 993b (Indicative Value 0.013 wt.%) | [13 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | wt.%
|  |  | 3A*OES | 3A*OES | [10 empty] | 3
3
| 1 | 1 | 0.0122 | 0.014218 | [10 empty] | 1
| 1 | 2 | 0.0119 | 0.014178 | [10 empty] | 4
| 1 | 3 | 0.0116 | 0.014207 | [10 empty] | 16
| 1 | 4 | 0.0124 | 0.014107 | [10 empty] | 0.013113
| 1 | 5 | 0.0124 | 0.014036 | [10 empty] | 34
| 1 | 6 | 0.0119 | 0.014213
| Mean |  | 0.012067 | 0.01416
| Median |  | 0.01205 | 0.014193
| Std Dev. |  | 0.00032042 | 7.3336e-05
| Rel.Std.Dev. |  | 2.655% | 0.5179%
| PDM3 |  | -7.981% | 7.981%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mo in OREAS 993b (Indicative Value 0.221 wt.%) | [13 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 5.01 | [9 empty] | wt.%
|  |  | 3A*OES | 3A*OES | AR*OES | [9 empty] | 3
[5 empty] | 0.2g | [9 empty] | 3
| 1 | 1 | 0.217 | 0.23862 | 0.2132 | [9 empty] | 1
| 1 | 2 | 0.2163 | 0.23675 | 0.2142 | [9 empty] | 29
| 1 | 3 | 0.2066 | 0.23417 | 0.209 | [9 empty] | 16
| 1 | 4 | 0.2217 | 0.23637 | 0.2074 | [9 empty] | 0.22118
| 1 | 5 | 0.2229 | 0.23512 | 0.2102 | [9 empty] | 35
| 1 | 6 | 0.2139 | 0.237
| Mean |  | 0.2164 | 0.23634 | 0.2108
| Median |  | 0.21665 | 0.23656 | 0.2102
| Std Dev. |  | 0.0058822 | 0.0015492 | 0.0028496
| Rel.Std.Dev. |  | 2.718% | 0.6555% | 1.352%
| PDM3 |  | -2.161% | 6.854% | -4.693%
| Z-Score (Absolute) |  | 0 | 2.4 | 0.67
| Analytical results for Na in OREAS 993b (Indicative Value 0.09 wt.%) | [13 empty] | Indicative
Na | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | wt.%
|  |  | 3A*OES | [11 empty] | 3
3
| 1 | 1 | 0.09 | [11 empty] | 1
| 1 | 2 | 0.09 | [11 empty] | 5
| 1 | 3 | 0.09 | [11 empty] | 16
| 1 | 4 | 0.09 | [11 empty] | 0.09
| 1 | 5 | 0.09 | [11 empty] | 36
| 1 | 6 | 0.09
| Mean |  | 0.09
| Median |  | 0.09
| Std Dev. |  | 1.5202e-17
| Rel.Std.Dev. |  | 1.689e-14%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
Nb | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 6
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | < 10 | [11 empty] | < 10
| 1 | 5 | < 10 | [11 empty] | 37
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 993b (Certified Value 253 ppm) | [13 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 | 5.01 |  | ppm
|  |  | 3A*OES | 3A*OES/MS | 3A*OES | 3A*OES | 3A*OES | 3A*OES | 3A*MS | 3A*OES | 3A*OES | 3A*OES | AR*OES |  | 0
[13 empty] | 0.2g |  | 0
| 1 | 1 | 217 | 259 | 261.94 | 270 | 236 | 300 | 199.9 | 250.546 | 258 | 272 | 256 |  | 1
| 1 | 2 | 214 | 261 | 262.36 | 280 | 237 | 300 | 201.9 | 249.049 | 254 | 272 | 260 |  | 0
| 1 | 3 | 211 | 257 | 259.74 | 270 | 237 | 300 | 197.9 | 245.1421 | 257 | 268 | 248 |  | 16
| 1 | 4 | 224 | 261 | 260.04 | 280 | 239 | 300 | 202.2 | 257.0473 | 253 | 268 | 248 |  | 252.9695
| 1 | 5 | 224 | 259 | 259.55 | 270 | 240 | 300 | 198.6 | 248.9396 | 253 | 270 | 250 |  | 101
| 1 | 6 | 214 | 259 | 261.48 | 270 | 241 | 300 | 194.8 | 257.117 | 254 | 264
| Mean |  | 217.3333 | 259.3333 | 260.8517 | 273.3333 | 238.3333 | 300 | 199.2167 | 251.3068 | 254.8333 | 269 | 252.4
| Median |  | 215.5 | 259 | 260.76 | 270 | 238 | 300 | 199.25 | 249.7975 | 254 | 269 | 250
| Std Dev. |  | 5.5015 | 1.5055 | 1.2201 | 5.164 | 1.9664 | 0 | 2.7636 | 4.8166 | 2.137 | 3.0332 | 5.3666
| Rel.Std.Dev. |  | 2.531% | 0.5805% | 0.4677% | 1.889% | 0.8251% | 0% | 1.387% | 1.917% | 0.8386% | 1.128% | 2.126%
| PDM3 |  | -14.09% | 2.516% | 3.116% | 8.05% | -5.786% | 18.59% | -21.25% | -0.6572% | 0.7368% | 6.337% | -0.2251%
| Z-Score (Absolute) |  | 2.17 | 0.34 | 0.43 | 1.18 | 0.92 | NA | 3.26 | 0.14 | 0.07 | 0.92 | 0.07
Results from laboratory 2.11 were removed due to their 100 ppm reading resolution.
| Analytical results for P in OREAS 993b (Indicative Value 0.018 wt.%) | [13 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab | [9 empty] | 1
| No | No | 2.01 | 2.04 | 5.01 | [9 empty] | wt.%
|  |  | 3A*OES | 3A*OES | AR*OES | [9 empty] | 3
[5 empty] | 0.2g | [9 empty] | 3
| 1 | 1 | 0.0176 | 1.1325 | 0.02 | [9 empty] | 1
| 1 | 2 | 0.0175 | 1.1309 | 0.019 | [9 empty] | 32
| 1 | 3 | 0.0169 | 1.1205 | 0.018 | [9 empty] | 16
| 1 | 4 | 0.0178 | 1.1185 | 0.019 | [9 empty] | 0.018242
| 1 | 5 | 0.0183 | 1.1168 | 0.0185 | [9 empty] | 38
| 1 | 6 | 0.0174 | 1.1167
| Mean |  | 0.017583 | 1.1226 | 0.0189
| Median |  | 0.01755 | 1.1195 | 0.019
| Std Dev. |  | 0.00046224 | 0.0071632 | 0.00074162
| Rel.Std.Dev. |  | 2.629% | 0.6381% | 3.924%
| PDM3 |  | -3.609% | 6054% | 3.609%
| Z-Score (Absolute) |  | 0.67 | 565.27 | 0
| Analytical results for Pb in OREAS 993b (Certified Value 465 ppm) | [13 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 | 5.01 |  | ppm
|  |  | 3A*OES | 3A*OES/MS | 3A*OES | 3A*OES | 3A*OES | 3A*OES | 3A*MS | 3A*OES | 3A*OES | 3A*OES | AR*OES |  | 0
[13 empty] | 0.2g |  | 0
| 1 | 1 | 423 | 502 | 492.12 | 440 | 483 | 500 | 449 | 465.7686 | 485 | 495 | 426 |  | 1
| 1 | 2 | 415 | 511 | 486.02 | 460 | 483 | 500 | 450 | 450.8675 | 480 | 495 | 420 |  | 7
| 1 | 3 | 406 | 490 | 486.46 | 440 | 487 | 500 | 446 | 452.8813 | 486 | 500 | 408 |  | 16
| 1 | 4 | 431 | 501 | 482.39 | 460 | 486 | 500 | 449 | 458.4148 | 480 | 495 | 416 |  | 464.5927
| 1 | 5 | 434 | 497 | 485.76 | 460 | 484 | 500 | 444 | 458.4922 | 482 | 495 | 416 |  | 102
| 1 | 6 | 417 | 495 | 488.58 | 460 | 487 | 500 | 434 | 465.6106 | 482 | 500
| Mean |  | 421 | 499.3333 | 486.8883 | 453.3333 | 485 | 500 | 445.3333 | 458.6725 | 482.5 | 496.6667 | 417.2
| Median |  | 420 | 499 | 486.24 | 460 | 485 | 500 | 447.5 | 458.4535 | 482 | 495 | 416
| Std Dev. |  | 10.4881 | 7.174 | 3.2458 | 10.328 | 1.8974 | 0 | 5.9889 | 6.2136 | 2.51 | 2.582 | 6.5727
| Rel.Std.Dev. |  | 2.491% | 1.437% | 0.6667% | 2.278% | 0.3912% | 0% | 1.345% | 1.355% | 0.5202% | 0.5199% | 1.575%
| PDM3 |  | -9.383% | 7.478% | 4.799% | -2.424% | 4.393% | 7.621% | -4.145% | -1.274% | 3.854% | 6.904% | -10.2%
| Z-Score (Absolute) |  | 1.57 | 0.91 | 0.52 | 0.55 | 0.46 | NA | 0.8 | 0.38 | 0.38 | 0.83 | 1.69
Results from laboratory 2.11 were removed due to their 100 ppm reading resolution.
| Analytical results for Re in OREAS 993b (Indicative Value < 5 ppm) | [13 empty] | Indicative
Re | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | < 5 | [11 empty] | 1
| 1 | 2 | < 5 | [11 empty] | 9
| 1 | 3 | < 5 | [11 empty] | 16
| 1 | 4 | < 5 | [11 empty] | < 5
| 1 | 5 | < 5 | [11 empty] | 39
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 993b (Indicative Value 30.51 wt.%) | [13 empty] | Indicative
S | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | wt.%
|  |  | 3A*OES | 3A*OES | [10 empty] | 2
2
| 1 | 1 | > 15 | 30.3425 | [10 empty] | 1
| 1 | 2 | > 15 | 30.5294 | [10 empty] | 34
| 1 | 3 | > 15 | 30.6412 | [10 empty] | 16
| 1 | 4 | > 15 | 30.3662 | [10 empty] | 30.5123
| 1 | 5 | > 15 | 30.7555 | [10 empty] | 40
| 1 | 6 | > 15 | 30.439
| Mean |  |  | 30.5123
| Median |  |  | 30.4842
| Std Dev. |  |  | 0.16222
| Rel.Std.Dev. |  |  | 0.5317%
| PDM3 |  |  | 1.11e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Sb in OREAS 993b (Indicative Value 12.2 ppm) | [13 empty] | Indicative
Sb | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | 13 | [11 empty] | 1
| 1 | 2 | 14 | [11 empty] | 10
| 1 | 3 | 11 | [11 empty] | 16
| 1 | 4 | 11 | [11 empty] | 12.1667
| 1 | 5 | 12 | [11 empty] | 41
| 1 | 6 | 12
| Mean |  | 12.1667
| Median |  | 12
| Std Dev. |  | 1.169
| Rel.Std.Dev. |  | 9.609%
| PDM3 |  | -2.776e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 993b (Indicative Value 4.17 ppm) | [13 empty] | Indicative
Sc | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | 4 | [11 empty] | 1
| 1 | 2 | 4 | [11 empty] | 11
| 1 | 3 | 4 | [11 empty] | 16
| 1 | 4 | 5 | [11 empty] | 4.1667
| 1 | 5 | 4 | [11 empty] | 29
| 1 | 6 | 4
| Mean |  | 4.1667
| Median |  | 4
| Std Dev. |  | 0.40825
| Rel.Std.Dev. |  | 9.798%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 993b (Indicative Value 160 ppm) | [13 empty] | Indicative
Se | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 0
0
| 1 | 1 | 162 | [11 empty] | 1
| 1 | 2 | 160 | [11 empty] | 1
| 1 | 3 | 152 | [11 empty] | 16
| 1 | 4 | 163 | [11 empty] | 159.8333
| 1 | 5 | 164 | [11 empty] | 30
| 1 | 6 | 158
| Mean |  | 159.8333
| Median |  | 161
| Std Dev. |  | 4.4008
| Rel.Std.Dev. |  | 2.753%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
Sn | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 25
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | < 10 | [11 empty] | < 10
| 1 | 5 | < 10 | [11 empty] | 31
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 993b (Indicative Value 95 ppm) | [13 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | ppm
|  |  | 3A*OES | 3A*OES | [10 empty] | 0
0
| 1 | 1 | 107 | 84.33 | [10 empty] | 1
| 1 | 2 | 105 | 83.55 | [10 empty] | 2
| 1 | 3 | 102 | 84.09 | [10 empty] | 16
| 1 | 4 | 110 | 83.54 | [10 empty] | 94.5283
| 1 | 5 | 104 | 83.69 | [10 empty] | 32
| 1 | 6 | 103 | 84.14
| Mean |  | 105.1667 | 83.89
| Median |  | 104.5 | 83.89
| Std Dev. |  | 2.9269 | 0.33888
| Rel.Std.Dev. |  | 2.783% | 0.404%
| PDM3 |  | 11.25% | -11.25%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Te in OREAS 993b (Indicative Value 17.2 ppm) | [13 empty] | Indicative
Te | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | 17 | [11 empty] | 1
| 1 | 2 | 17 | [11 empty] | 4
| 1 | 3 | 17 | [11 empty] | 16
| 1 | 4 | 17 | [11 empty] | 17.1667
| 1 | 5 | 18 | [11 empty] | 33
| 1 | 6 | 17
| Mean |  | 17.1667
| Median |  | 17
| Std Dev. |  | 0.40825
| Rel.Std.Dev. |  | 2.378%
| PDM3 |  | -1.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 993b (Indicative Value 0.08 wt.%) | [13 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | wt.%
|  |  | 3A*OES | 3A*OES | [10 empty] | 3
3
| 1 | 1 | 0.08 | 0.079044 | [10 empty] | 1
| 1 | 2 | 0.08 | 0.079048 | [10 empty] | 28
| 1 | 3 | 0.07 | 0.080411 | [10 empty] | 16
| 1 | 4 | 0.08 | 0.078868 | [10 empty] | 0.07977
| 1 | 5 | 0.08 | 0.079489 | [10 empty] | 34
| 1 | 6 | 0.08 | 0.080383
| Mean |  | 0.078333 | 0.079541
| Median |  | 0.08 | 0.079269
| Std Dev. |  | 0.0040825 | 0.00069449
| Rel.Std.Dev. |  | 5.212% | 0.8731%
| PDM3 |  | -1.801% | -0.288%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Tl in OREAS 993b (Indicative Value < 5 ppm) | [13 empty] | Indicative
Tl | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | < 5 | [11 empty] | 1
| 1 | 2 | < 5 | [11 empty] | 5
| 1 | 3 | < 5 | [11 empty] | 16
| 1 | 4 | < 5 | [11 empty] | < 5
| 1 | 5 | < 5 | [11 empty] | 35
| 1 | 6 | < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
U | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 6
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | < 10 | [11 empty] | < 10
| 1 | 5 | < 10 | [11 empty] | 36
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 993b (Indicative Value 47.5 ppm) | [13 empty] | Indicative
V | Round | Replicate | Lab | Lab | [10 empty] | 1
| No | No | 2.01 | 2.04 | [10 empty] | ppm
|  |  | 3A*OES | 3A*OES | [10 empty] | 1
1
| 1 | 1 | 48 | 45.71 | [10 empty] | 1
| 1 | 2 | 47 | 48.31 | [10 empty] | 7
| 1 | 3 | 45 | 48.7 | [10 empty] | 16
| 1 | 4 | 49 | 48.53 | [10 empty] | 47.5367
| 1 | 5 | 48 | 46.83 | [10 empty] | 37
| 1 | 6 | 46 | 49.36
| Mean |  | 47.1667 | 47.9067
| Median |  | 47.5 | 48.42
| Std Dev. |  | 1.472 | 1.3622
| Rel.Std.Dev. |  | 3.121% | 2.843%
| PDM3 |  | -0.7783% | 0.7783%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for W in OREAS 993b (Indicative Value < 10 ppm) | [13 empty] | Indicative
W | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | < 10 | [11 empty] | 1
| 1 | 2 | < 10 | [11 empty] | 32
| 1 | 3 | < 10 | [11 empty] | 16
| 1 | 4 | < 10 | [11 empty] | < 10
| 1 | 5 | < 10 | [11 empty] | 38
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 993b (Indicative Value 4.83 ppm) | [13 empty] | Indicative
Y | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 2
2
| 1 | 1 | 5 | [11 empty] | 1
| 1 | 2 | 5 | [11 empty] | 33
| 1 | 3 | 4 | [11 empty] | 16
| 1 | 4 | 5 | [11 empty] | 4.8333
| 1 | 5 | 5 | [11 empty] | 39
| 1 | 6 | 5
| Mean |  | 4.8333
| Median |  | 5
| Std Dev. |  | 0.40825
| Rel.Std.Dev. |  | 8.447%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 993b (Certified Value 0.342 wt.%) | [13 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.17 | 2.18 |  |  | wt.%
|  |  | 3A*OES | 3A*OES/MS | 3A*OES | 3A*OES | 3A*OES | 3A*OES | 3A*MS | 3A*OES | 3A*OES | 3A*OES |  |  | 3
3
| 1 | 1 | 0.297 | 0.362 | 0.34 | 0.316 | 0.3348 | 0.35 | 0.3157 | 0.35072 | 0.3499 | 0.353 |  |  | 1
| 1 | 2 | 0.2962 | 0.366 | 0.33733 | 0.324 | 0.3341 | 0.36 | 0.3179 | 0.34278 | 0.3526 | 0.352 |  |  | 0
| 1 | 3 | 0.2873 | 0.36 | 0.33633 | 0.315 | 0.3364 | 0.36 | 0.3098 | 0.34084 | 0.3534 | 0.348 |  |  | 16
| 1 | 4 | 0.3049 | 0.366 | 0.33382 | 0.328 | 0.3302 | 0.35 | 0.3201 | 0.35229 | 0.3495 | 0.346 |  |  | 0.34172
| 1 | 5 | 0.3074 | 0.364 | 0.33371 | 0.326 | 0.3331 | 0.35 | 0.3123 | 0.34868 | 0.3503 | 0.357 |  |  | 103
| 1 | 6 | 0.2945 | 0.365 | 0.33601 | 0.323 | 0.332 | 0.37 | 0.3055 | 0.35488 | 0.3521 | 0.345
| Mean |  | 0.29788 | 0.36383 | 0.3362 | 0.322 | 0.33343 | 0.35667 | 0.31355 | 0.34836 | 0.3513 | 0.35017
| Median |  | 0.2966 | 0.3645 | 0.33617 | 0.3235 | 0.3336 | 0.355 | 0.314 | 0.3497 | 0.3512 | 0.35
| Std Dev. |  | 0.0073057 | 0.0024014 | 0.0023509 | 0.0053292 | 0.0021787 | 0.008165 | 0.0054199 | 0.0055015 | 0.0016087 | 0.0046224
| Rel.Std.Dev. |  | 2.453% | 0.66% | 0.6993% | 1.655% | 0.6534% | 2.289% | 1.729% | 1.579% | 0.4579% | 1.32%
| PDM3 |  | -12.83% | 6.47% | -1.616% | -5.772% | -2.426% | 4.373% | -8.245% | 1.943% | 2.802% | 2.471%
| Z-Score (Absolute) |  | 2.56 | 1.24 | 0.35 | 1.17 | 0.51 | 0.83 | 1.66 | 0.35 | 0.52 | 0.45
| Analytical results for Zr in OREAS 993b (Indicative Value 13.7 ppm) | [13 empty] | Indicative
Zr | Round | Replicate | Lab | [11 empty] | 1
| No | No | 2.01 | [11 empty] | ppm
|  |  | 3A*OES | [11 empty] | 1
1
| 1 | 1 | 14 | [11 empty] | 1
| 1 | 2 | 14 | [11 empty] | 9
| 1 | 3 | 13 | [11 empty] | 16
| 1 | 4 | 14 | [11 empty] | 13.6667
| 1 | 5 | 14 | [11 empty] | 40
| 1 | 6 | 13
| Mean |  | 13.6667
| Median |  | 14
| Std Dev. |  | 0.5164
| Rel.Std.Dev. |  | 3.779%
| PDM3 |  | -2.442e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C415: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C451: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C469: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C633: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C651: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C669: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C777: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C795: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.